696 |  |  | Japan | Higashi-Ohmi General Medical C | 341015 | 4429500% |  | 01CS:KK:End exam後次の患者様登録画面に切り替わったが何も入力できない。Logout/Loginしたが青い画面のまま先に進まない。>>Alt+Ctrl+DelよりLogoutしていただき再起動にて復旧。 | Problem Reported by Customer: 01CS: KK: End exam After the exam, I switched to the next patient registration screen, but I can't enter anything. I logged out / logged in, but the blue screen remains and I can't proceed. >> Log out from Alt + Ctrl + Del and restart to recover. T2 Activities: Problem Description: Target device: CT Failure location: 01Cosole When did it first occur? (Within ~): 1 Hour How often do problems occur ?: First Occurrence At what point in the end-user workflow is the problem occurring ?: Setting up Protocol What happened recently ?: N.A. Does the customer create a specific system event report for the problem ?: No. The customer was unable to create a log file or report. Troubleshooting Action: Result after response: Improved Failure location: 01Cosole What steps were taken during the diagnosis ?: There is no particular error regarding the matter that could not be entered on the new patient registration screen. There was no particular error when logging in after that, but it seems that the app could not operate normally. Recovered by restarting the Host. I was allowed to wait and see. Repair Action: [Phenomenon confirmation] Confirmed | 4 |  | End exam After the exam, I switched to the next patient registration screen, but I can't enter anything. I logged out / logged in, but the blue screen remains and I can't proceed. | [12 empty] | 117343905 | Fixed | 788953 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00631367 | [5 empty] | Engineering | [6 empty] | New Yellow Screen | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
697 |  |  | Japan | Maki Hospital | 341017 | 4429300% |  | 05CS:KKI: The X-ray irradiation sound is loud. >> I asked FSE Yamamoto to respond. | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 05DS/KKI / The X-ray irradiation sound is loud. Safety Questions:Case: 0117330273 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: KAZUMASA KIMURA Last updated by: KAZUMASA KIMURA 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: subject equipment: CT Failure location:: 05 Peripheral When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user's work flow did the problem occur?: During mechanical movements (During geometry/mechanical movements) What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem? No.The customer has not created a log file or report. Troubleshooting Action: Result after troubleshooting: Improved Failure location:: 05 Peripheral What steps did you perform to diagnose the problem? Call description: The X-ray irradiation sound is loud. Response I reported the issue to FSE Yoshiyuki Yamamoto. Repair Action: [Confirmation of problem] Confirmed. Work details: Call description: The X-ray irradiation sound is loud. Response I reported the issue to FSE Yoshiyuki Yamamoto. 【Confirmation content】 Good. OneEMS internal remarks: --Troubleshooting Action-- Result of response: Problem solved. Failure location:: 05 Peripheral What steps did you perform to diagnose the problem? Call description: The X-ray irradiation sound is loud. Response I reported the issue to FSE Yoshiyuki Yamamoto. --Repair Action-- --Subject-- 05CS:KKI:Mr. Hoshi: The X-ray irradiation sound is loud. >> I reported the issue to FSE Yoshiyuki Yamamoto. --Problem Description-- Equipment affected: CT Failure location:: 05 Peripheral When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user's work flow did the problem occur?: During mechanical movements (During geometry/mechanical movements) What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem? No.The customer has not created a log file or report. --4/8Elements-- Q1. Was the reported problem confirmed? Confirmed. Q2. Were any other problems confirmed? None. Q3. What did you do to resolve the problem?: Call description: The X-ray irradiation sound is loud. Response I reported the issue to FSE Yoshiyuki Yamamoto. Q4: Is the equipment operating normally following the response? Good. Q5. How did the examination finish? None. Q6. Were there any mistaken findings (misdiagnosis, mistaken treatment) due to the problem? None. Q7. Was there any unintended unnecessary exposure? None. Q8. Were there any artifacts? None. ----------------------------------------------------------------------------- | 4 |  | Sound Adjustment | [12 empty] | 117330273 | In Process | 780484 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Service | [6 empty] | Sound Adjustment | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
698 |  |  | Japan | Maki Hospital | 341017 | 4429300% |  | 07CS:KKI: The telephone set power is not on.>> I told him that the telephone line would become available from April 13. | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Application Support Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer: 07CS/KKI/ The telephone set power is not on.>> I told him that the telephone line would become available from April 13. Safety Questions:Case: 0117330257 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: KAZUMASA KIMURA Last updated by: KAZUMASA KIMURA 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: subject equipment: CT Location of fault: 07 Other When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence When, in the end user's work flow, is the problem occurring?: During Examination ・Measurement preparation (Setting up Protocol) What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem? No.The customer has not created a log file or report. Troubleshooting Action: Result after troubleshooting: Improved Location of fault: 07 Other What steps did you perform to diagnose the problem? Call description: The telephone set power is not on. Response I told him that the telephone line would become available from April 13. Repair Action: [Confirmation of problem] Confirmed. Work details: Call description: The telephone set power is not on. Response I told him that the telephone line would become available from April 13. 【Confirmation content】 Good. OneEMS internal remarks: --Troubleshooting Action-- Result of response: Problem solved. Location of fault: 07 Other What steps did you perform to diagnose the problem? Call description: The telephone set power is not on. Response I told him that the telephone line would become available from April 13. --Repair Action-- --Subject-- 07CS:KKI:Mr. Hoshi: The telephone set power is not on.>> I told him that the telephone line would become available from April 13. --Problem Description-- Equipment affected: CT Location of fault: 07 Other When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence When, in the end user's work flow, is the problem occurring?: During Examination ・Measurement preparation (Setting up Protocol) What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem? No.The customer has not created a log file or report. --4/8Elements-- Q1. Was the reported problem confirmed? Confirmed. Q2. Were any other problems confirmed? None. Q3. What did you do to resolve the problem?: Call description: The telephone set power is not on. Response I told him that the telephone line would become available from April 13. Q4: Is the equipment operating normally following the response? Good. Q5. How did the examination finish? None. Q6. Were there any mistaken findings (misdiagnosis, mistaken treatment) due to the problem? None. Q7. Was there any unintended unnecessary exposure? None. Q8. Were there any artifacts? None. SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 5 |  | The telephone set power is not on.>> I told him that the telephone line would become available from April 13. | [12 empty] | 117330257 | Fixed | 780468 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Telephone | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
699 |  |  | Portugal | Centro Hospitalar de Vila Nova | 40030 | 4429300% |  | Aquando da emissão de rx diz: XRay failed, try again" | Work Order #: WO-03289261 08-Apr-2021 Jaime Granja The system fully meets specification and returned to use Regulatory Questions : N/A Problem Description by Engineer : Customer Function/Role: Radiology Technician Device Use: Diagnostics Expected and Unexpected Results: X ray error just before an acquisition. Then it was repeated and went uneventful User Impact: None Patient Impact: None Current Software Version: 4.1.10.25202 Diagnostic performed by Engineer : Coincident with the reported error, an mA error was detected in the system error file. As a preventive measure, a filament calibration was carried out following the recommendation of the service manuals for these cases. Monitoring of some exams that ran without incident. | 4 |  | As a preventive measure, a filament calibration was carried out following the recommendation of the service manuals for these cases. Monitoring of some exams that ran without incident. | [12 empty] | 117333504 | Incomplete | 782625 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | NotEngineering-Calibration | [5 empty] | Others | [6 empty] | Xray failed | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
700 |  |  | Puerto Rico | HOSPITAL DR SUSONI INC | 40013 | 4429300% |  | The table when go up, make noise | Work Order #: WO-03282937 Caldelas Aldana, Carlos Manuel The system partly meets specification and is returned to use with restrictions Internal Comments : Error # and/or description of error : noise in patient suport Completed or Recommended Repair Action "Internal" : n/a Diagnostic performed by Engineer : abnormal noise detected in vertical movement in patient support. no errors appear. Problem Description by Engineer : Problem description by engineer : abnormal noise in patient support When was the first occurrence? Within... : 07-Apr-2021 11:00:00 How often is issue occurring? : First occurrence Malfunction area : Table Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: abnormal noise in couch Patient Impact/Outcome: n/a User Impact: noise in patient support Current Software Version: Not provided Follow Up Required : What are the next Troubleshooting steps : check lubrication on patient support on vertical movment related mechanism. Discard covers problems. Error # and/or description of error : noise in patient suport Completed or Recommended Repair Action "Internal" : n/a ----------------------------------------------------------------------------- Work Order #: WO-03283341 07-Apr-2021 Ramos Rivera, Pablo The system fully meets specification and returned to use Problem Description by Engineer : Problem description by engineer : abnormal noise in patient support When was the first occurrence? Within... : 07-Apr-2021 11:00:00 How often is issue occurring? : First occurrence Malfunction area : Table Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: abnormal noise in couch when moving vertically Patient Impact/Outcome: n/a User Impact: noise in patient support Current Software Version: Not provided Resolution : Found that the vertical movement guides needed lubrication. Lubricated the guides and now there is no noise. The exam room temperature is high and humidity is too low. ----------------------------------------------------------------------------- | 3 |  | Found that the vertical movement guides needed lubrication. Lubricated the guides and now there is no noise. The exam room temperature is high and humidity is too low. | [12 empty] | 117334112 | Closed | 778984 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | E-125059 | [5 empty] | Engineering | [6 empty] | Couch Noise | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | Lubricant | Couch-HW-Lubricant | Couch-HW-Lubricant | E-125059 | Lubricant | Couch-HW | Couch-HW-Lubricant | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
701 |  |  | Spain | S.C.S. HOSP. GRAL. DE FUERTEVE | 40020 | 4429200% |  | Images are not transferred to the PACS. | Work Order #: WO-03273671 Carlos Saumell Lasala The system fully meets specification and returned to use Resolution: After the transfer restart, shipments to Pacs are resumed, and the client confirms to me after two days that these shipments continue to work well. Diagnostic performed by Engineer: The transfer spooler is restarted remotely, without asking the client to restart the host computer. Problem Description by Engineer: Image transfers from Tac to Pacs have been blocked. Regulatory Questions: N / A ----------------------------------------------------------------------------- | 3 |  | The transfer spooler is restarted remotely, without asking the client to restart the host computer. | [12 empty] | 117325824 | Closed | 777378 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS Transfer | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
702 |  |  | Spain | MAZ ZARAGOZA M.A.T.E.P.S.S.11 | 40042 | 4429200% |  | It makes a strange sound | Work Order #: WO-03274440 07-Apr-2021 Diego Soria The system fully meets specification and returned to use Regulatory Questions : N/A Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : In high speed rotation study it makes sharp noise in the gantry I open the gantry and turn the rotor, the noise comes from the rear area on the motor belt. Resolution: I have checked the condition of the belt for the transmission of rotation in the gantry. I only hear the noise when turning at high speeds. It is not a loud sound. I reattach the strap and run tests. I do not hear the noise again, I inform that I will come back to check and inform me if the reason for the notice is reproduced. The system is working properly. Follow Up Required: Check if the reason for the ad is reproduced. | 3 |  | I reattach the strap and run tests. I do not hear the noise again, I inform that I will come back to check and inform me if the reason for the notice is reproduced. The system is working properly. | [12 empty] | 117326472 | Closed | 776799 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | ECR-077012 | [5 empty] | Engineering | [6 empty] | Belt | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Belt | Gantry-HW-Belt | Gantry-HW-Belt | ECR-077012 | Belt | Gantry-HW | Gantry-HW-Belt | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
703 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 341006 | 4429500% |  | The screen of the device is frozen | The screen of the device is frozen | 3 |  | The screen of the device is frozen Ask field to collect bugreport | [12 empty] | 117345727 | Closed | 789690 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00631367 | [5 empty] | Engineering | [6 empty] | New Yellow Screen | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
704 |  |  | Turkey | HITACHI HEALTHCARE TURKEY | 337000 | 4429500% |  | Customer wants back up | Customer wants back up | 3 |  | A full back up, including new protocols, was purchased and delivered to the customer. The system was tested with the phantom in the clinic and delivered ready to use. | [12 empty] | 117346244 | In Progress | 786249 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Backup | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
705 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 341006 | 4429400% |  | Freezes during shooting | Work Order #: WO-03286997 Soner Saglam The system fully meets specification and returned to use Problem Description by Engineer : SW'de donma. Resolution : There is no RSN connection. Cannot log from the system. With the field visit, the connection problem should be resolved and the issue should be examined by taking bugrep. Follow Up Required: No RSN connection. Cannot log from the system. With the field visit, the connection problem should be resolved and the issue should be examined by taking bugrep. ----------------------------------------------------------------------------- | 3 |  | Freezing in SW. | [12 empty] | 117338353 | Closed | 781324 | TRUE | Investigation | Closed - Done |  | CTUDT00631367 | [5 empty] | Engineering | [6 empty] | New Yellow Screen | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
706 |  |  | United States | Unity Health Systems | 40037 | 4429700% |  | UPS beeping intermittent | No further issues after resetting UPS. OK to close. | 4 |  | No further issues after resetting UPS. OK to close. | [12 empty] | 117351331 | In Progress | 791159 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | UPS | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
707 |  |  | United States | Leesburg Regional Medical Cent | 40022 | 4429500% |  | 0903CDY70 ring artifact | Expected and Unexpected results: ring artifact Patient Impact/Outcome: none User Impact: ring artifact Current Software Version: 4.1.10 Diagnostic performed by Engineer : Angelo said they are getting rings on images. I downloaded vis Secure FTP a few images he gave me and confirmed Module 7 or 36 is showing in the images. I was unable to download any bad detector tables or air tables. System locks up with this message "Unexpected close was received from the remote system for subsystem SFTP. Please contact the server's administrator for further information. " Customer ran air cals and a test scan and stated he did not see the ring but I am not comfortable not being able to see what is going on so dispatching the case. Internal Comments : Error # and/or description of error : ring artifact Completed or Recommended Repair Action "Internal" : Contact the customer and confirm whether the air calibration resolved the issue. If not use the image quality T/S manual and determine best fix for the module. Might move it our or replace it. Confirm PN listed below prior to ordering. 4598005911281 TDMS V2 Module Class B FRU Follow Up Required : Error # and/or description of error : ring artifact Completed or Recommended Repair Action "Internal" : Contact the customer and confirm whether the air calibration resolved the issue. If not use the image quality T/S manual and determine best fix for the module. Might move it our or replace it. Confirm PN listed below prior to ordering. 4598005911281 TDMS V2 Module Class B FRU ----------------------------------------------------------------------------- Work Order #: WO-03297843 10-Apr-2021 Neuman Terry The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Problem description by engineer : ring artifact When was the first occurrence? Within... : 09-Apr-2021 How often is issue occurring? : First occurrence Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: ring artifact Patient Impact/Outcome: none User Impact: ring artifact Current Software Version: 4.1.10 Follow Up Required : replace detector module Diagnostic performed by Engineer : using FEE test determined module 36 to be bad, moved to module 41 and simptons followed ----------------------------------------------------------------------------- | 3 |  | using FEE test determined module 36 to be bad, moved to module 41 and simptons followed | [12 empty] | 117348642 | In Progress | 787933 | TRUE | Closed - Done | Closed - Done |  | CTUDT00632110 | [5 empty] | Engineering | [6 empty] | Artifact | [5 empty] | Refresh |  |  | System issue | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | DMS Module | DMS-HW | DMS-HW-DMS Module | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
708 |  |  | United States |  | 40053 | 4429300% |  | XR-29 needed | Work Order #: WO-03284690 Boyd Craig The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: XR-29 needed User Impact: None Patient Impact: None Current Software Version: Resolution : Request XR-29. Certificate emailed to customer. Diagnostic performed by Engineer : New installation. Customer requesting XR-29 for system. ----------------------------------------------------------------------------- Work Order #: WO-03288357 Hirchak, Michele The system fully meets specification and returned to use Diagnostic performed by Engineer : Emailed certificate to customer and FSE. MBH Problem description by Clinical Support Specialist : Customer Function/Role: TECH Device Use: In Use Expected and Unexpected Results: XR-29 certificate request User Impact: NONE Current Software Version: 4.1 No rescan, no loss of data, no reinjection. ----------------------------------------------------------------------------- | 3 |  | New installation. Customer requesting XR-29 for system. | [12 empty] | 117335435 | Closed | 779727 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Customer | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
709 |  |  | United States | Clay County Medical Center | 40003 | 4429200% |  | PRS and Network/HIS issues. | Work Order #: WO-03277360 06-Apr-2021 Stauffer Jeffrey The system fully meets specification and returned to use Diagnostic performed by Engineer : Looked at network settings in LAN Config and saw missing information. Problem Description by Engineer : Customer Function/Role: Radiology Tech Device Use: General ct exams Expected and Unexpected Results: Unexpected cannot access PRS or Worklist. User Impact: None Patient Impact: None Current Software Version: 4.1.10 Resolution : Changed LAN Config settings. ----------------------------------------------------------------------------- | 3 |  | Worklist Configuration | [12 empty] | 117328833 | Closed | 775302 | TRUE | Investigation | Closed - Done | [7 empty] | Service | [6 empty] | Worklist Configuration | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
710 |  |  | United States | Select Specialty Hospitals | 310051 | 4429700% |  | 3-System Restricted | Work Order #: WO-03301930 Morrow, Dan The system does not meet specification and is out of use Diagnostic performed by Engineer : Downloaded Bug Report. Horizontal encoders not synced. Dispatching FSE. Problem Description by Engineer : Problem description by engineer : Estops keep opening reference case 0117342543 When was the first occurrence? Within... : 11-Apr-2021 How often is issue occurring? : Multiple times a day Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Unable to use Patient Impact/Outcome: none User Impact: Unable to use Current Software Version: 3.5.5.25007 Internal Comments : Error # and/or description of error : S_MC_DEVICE_HOR_ENCODERS_NOT_SYNCED S_MANMOT_HORIZONTAL_START_ERROR Completed or Recommended Repair Action "Internal" : Downloaded Bug Report. Horizontal encoders not synced. Dispatching FSE. Follow Up Required : What are the next Troubleshooting steps : ERROR_STRING : S_MC_DEVICE_HOR_ENCODERS_NOT_SYNCED Explanation : Horizontal abs and relative encoders do not match. Action : 1. Reset couch and retry, check encoders/baseboard/cable. 2. Check Fuse F3 0.5A in CCB this is 15Vdc protection for Absolute Encoder Experience: 1. Found table MOTOR COUPLING between Absolute encoder and lead screw broken. MOTOR COUPLING replaced. 2. Absolute encoder losing its calibration. Replaced Absolute encoder. 3. Coupling loose between encoders, had to tighten set screws and calibrate horizontal- 4.. Found J1 – RS232/422 connector when touched would trip out the ESTOP and generate this error. Inspection found issue to be due to the ACS pcb where the JI socket is mounted on. Replaced the ACS successfully. 5. E-stop opened at end of max in or out from table movement. Checked connector for calibrating table horizontal movement, found in one case this was still connected. 6. Replaced both encoders to solve this Error # and/or description of error : S_MC_DEVICE_HOR_ENCODERS_NOT_SYNCED S_MANMOT_HORIZONTAL_START_ERROR Completed or Recommended Repair Action "Internal" : Downloaded Bug Report. Horizontal encoders not synced. Dispatching FSE. ----------------------------------------------------------------------------- | 3 |  | Replaced absolute and positive encoders, tested system, ready for use. 1.00000,453567399251,Assy, Absolute Encoder 1.00000,453567399231,Assy, Position Encoder | [12 empty] | 117351913 | In Progress | 791610 | #N/A | #N/A | Closed - Done |  | NotEngineering-PartReplacement | [5 empty] | Engineering | [6 empty] | Couch E-Stop | [5 empty] | Non-Refresh |  |  | Upgrade issue | system | Couch-HW | Encoder | Couch-HW-Encoder | Couch-HW-Encoder | Couch-HW-Encoder | Encoder | Couch-HW | Couch-HW-Encoder | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | y
711 |  |  | United States | Select Specialty Hospitals | 310051 | 4429400% |  | E-stop tripping | Work Order #: WO-03291941 Boyd Craig The system does not meet specification and is out of use Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: E-stop tripping User Impact: None Patient Impact: None Current Software Version: Diagnostic performed by Engineer : Power cycle system. Cycle E-stops. No PRS access. E-stop still tripping after CT box key is turned. Follow Up Required : Hold till morning. FSE on site to evaluate E-stop circuit. Check error logs. No remote PRS access. ----------------------------------------------------------------------------- Work Order #: WO-03292104 09-Apr-2021 Siker Shawn The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: E-stop tripping User Impact: None Patient Impact: None Current Software Version: Resolution : Recalibrated horizontal encoders on table and tested system. Performed air cals and iq check, system is operational. ----------------------------------------------------------------------------- | 2 |  | Recalibrated horizontal encoders on table and tested system. Performed air cals and iq check, system is operational. | [12 empty] | 117342543 | Closed | 783902 | TRUE | Closed - Done | Closed - Done |  | NotEngineering_CouchEstop_Redocalibration | [5 empty] | Engineering | [6 empty] | Couch E-Stop | [5 empty] | Refresh |  |  | Upgrade issue | System | NotEngineering | ShouldbeCustomer | NotEngineering-ShouldbeCustomer | NotEngineering-ShouldbeCustomer | NotEngineering-ShouldbeCustomer | NotEngineering-ShouldbeCustomer | NotEngineering | NotEngineering-ShouldbeCustomer | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
712 |  |  | United States | The Christ Hospital-Ft Wright | 300205 | 4429100% |  | need password for the exam card editor/ appliction support | need password for the exam card editor/ appliction support Attached is a quote for service(s)/part(s) requested on case #0117320598 for $1,155.00. | 5 |  | Password | [12 empty] | 117320598 | New | 773992 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Application | [6 empty] | Password | [5 empty] | Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
713 |  |  | United States | Lakeland Regional Medical Cent | 320474 | 4429200% |  | Q/O System is not sending Dose report to MPPS Cerner | Work Order #: WO-03274839 06-Apr-2021 Watson Adam The system fully meets specification and returned to use Resolution : After a patient is scanned, the system is generating the dose structured report and automatically sending to Dose Wise and MPPS Cerner. The MPPS Cerner port is failing to receive. Checked PWC CT against a known working CT’s IP configuration set up and they are the same. Worked with hospital IT to determine issue and found the system is attempting to send but the problem is on Cerner side. Hospital IT has opened a ticket on the Cerner side to address the issue. Problem Description by Engineer : Customer Function/Role: | 3 |  | MPPS Cerner | [12 empty] | 117326874 | Closed | 773899 | TRUE | Investigation | Investigation | [7 empty] | Others | [6 empty] | MPPS Cerner | [5 empty] | Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
714 |  |  | Argentina | Hospital Dr. Bernardo Houssay | 40027 | 4429800% |  | THERE IS NO HORIZONTAL MOVEMENT DURING STUDY | problem Description by Engineer: Horizontal movement failure in study We contact engineer Christian, the equipment does not have a remote connection. Support is requested to extract the error log manually, but by metrics they do not have authorization. Follow Up Required: Follow Up Required What are the next Troubleshooting steps : Error log is checked and the failure is replicated. Check electronics in the patient table, ACS (AMC), perform flash to CCB, horizontal calibration. Mechanical check. Resolution: Firmware was re-installed in the bed controller. Diagnostic performed by Engineer: The equipment's software status logs are checked. | 3 |  | Firmware was re-installed in the bed controller. Diagnostic performed by Engineer: The equipment's software status logs are checked. | [12 empty] | 117355700 | Closed | 792858 | TRUE | Closed - Done | Closed - Done |  | NotEngineering_Hwconnection | [5 empty] | Engineering | [6 empty] | Horizontal Movement | [8 empty] | System issue | System | NotEngineering | ShouldbeOthers | NotEngineering-ShouldbeOthers | NotEngineering-ShouldbeOthers | NotEngineering-ShouldbeOthers | NotEngineering-ShouldbeOthers | NotEngineering | NotEngineering-ShouldbeOthers | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
715 |  |  | Chile | HOSPITAL BARROS LUCO TRUDEAU | 341005 | 4430100% |  | CL1688-TEST TIME OFFSET | Expected and Unexpected Results: Time is synchronized between host and CIRS. Patient Impact/Outcome: It does not impact. User Impact: It does not impact. Current Software Version: 4.1.10 Diagnostic Performed by Engineer: Remote connection is performed, time is configured in CIRS, working correctly. |  |  | Remote connection is performed, time is configured in CIRS, working correctly. | [12 empty] | 117383785 | Closed | 807917 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | NotEngineering - TimeSync | [5 empty] | Service | [6 empty] | Time Sync | [8 empty] | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
716 |  |  | China | JX Yugan Chudong hospital | 337500 | 4430500% |  | Ingenutiy Core CT Flatspring fragment |  |  |  | Design improvement and Manufuture check the assembly process | [14 empty] | 815389 | TRUE | Closed - Duplicate | Closed - Duplicate |  | CTUDT00632833 | [5 empty] | Engineering | [6 empty] | Flat Spring | [8 empty] | System issue | system | Gantry-HW | FlatSpring | Gantry-HW-FlatSpring | Gantry-HW-FlatSpring | CTUDT00632833 | FlatSpring | Gantry-HW | Gantry-HW-FlatSpring | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
717 |  |  | China | Lindian County Hospital | 345038 | 4430100% |  | WeChat:02N/Core128/Tube high temperature | Customer Symptom Code: Monitor Alarm Failure Reason: Corrective Maintenance Engineer Failure Code: Incompatible Use Repair Activity Code: Instructed Customer Problem Reported by Customer : Tube high temperature Troubleshooting Action: instruct customer modify the absolute Min mAs to 50 MAX 260 Repair Action: instruct to customer Test/Inspection: Functional Check PASSED. Device returned to full functionality. |  |  | instruct customer modify the absolute Min mAs to 50 MAX 260 确认是否是由于球管真的过热，能否录视频显示详细参数 | [12 empty] | 117376681 | Fixed | 809605 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD_Tube hot_bugreport | [5 empty] | Engineering | [6 empty] | Tube Overheat | [8 empty] | System issue | System | Gantry-SW | TubHot | Gantry-SW-TubHot | Gantry-SW-TubHot | TBD_Tube hot_bugreport | TubHot | Gantry-SW | Gantry-SW-TubHot | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
718 |  |  | China | NMG Alashan Youqi People Hosp | 345031 | 4430100% |  | consult image quality | Problem Description: The customer thinks the images are a bit blurry, especially after the head images are enlarged it appears blurry. Troubleshooting Action: Guided the customer to not try to use filter functions to reconstruct images. Modify different scan conditions and see if they have changed Repair Action: The customer agrees to try different reconstruction conditions to reconstruct images Test/Inspection: N/A Test Equipment: N/A SAP Transfer:Case Transferred to SAP |  |  | Guided the customer to not try to use filter functions to reconstruct images. Modify different scan conditions and see if they have changed | [12 empty] | 117376317 | Fixed | 809525 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | NotEngineering - NotcorrectParameter | [5 empty] | Application | [6 empty] | Image Quality | [8 empty] | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
719 |  |  | China | SD Chest Hosp. | 40025 | 4430100% |  | Can not start up,blue screen | Problem Reported by Customer : Can not start up,blue screen T2 Activities : Problem Description: 客户问题概述： 计算机软件相关故障。 客户问题详细描述或故障现象： 设备蓝屏，启动不到登录界面 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 设备蓝屏，启动不到登录界面 远程诊断与测试结果： 设备主机启动有问题，所以PRS不通，与工程师沟通后需到场检查 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 检查软件，如有必要进行重装系统 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 15.04.2021 13:11:53 UTC SUP JCOUSER (86038405)服务内容：软件卡死后无法进入，用IST权限进入后配置。 服务结果：维修后，设备恢复正常使用。 External Remarks : 16.04.2021 10:49:48 UTC SUP JCOUSER (86038451)服务内容：软件卡死后无法进入，用IST权限进入后配置。 服务结果：维修后，设备恢复正常使用。 |  |  | 服务内容：软件卡死后无法进入，用IST权限进入后配置。 服务结果：维修后，设备恢复正常使用。 | [12 empty] | 117374910 | In Process | 809156 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00631367 | [5 empty] | Engineering | [6 empty] | Blue screen | [8 empty] | System issue | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
720 |  |  | China | Jinggangshan Hospital, the First Affiliated Hospital of Nanchang University, Jiangxi | 345002 | 4430100% |  | Data loss | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer : Data loss Safety Questions : Case: 0117372980 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: Data loss Troubleshooting Action: Remotely instruct the customer to full rebuild data. Repair Action: Remotely instruct the customer to full rebuild data. Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- |  |  | Remotely instruct the customer to full rebuild data. | [12 empty] | 117372980 | Fixed | 808309 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Recon | [8 empty] | System issue | System | TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
721 |  |  | China |  | 345036 | 4430000% |  | Can not log in | BU Q&R Feedback: No SAP Customer Symptom Code: MARF Customer Symptom Code: Monitor Alarm Failure Reason: Corrective Maintenance Engineer Failure Code: Incompatible Use Repair Activity Code: Instructed Customer Problem Reported by Customer : Can not log in Safety Questions : Case: 0117368253 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Bai Jingwen Last updated by: Bai Jingwen 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: consulting the olevel_service password Troubleshooting Action: instruct the new system olevel service password to customer Repair Action: instruct to customer Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: NA SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | Lu, Xueying |  | instruct the new system olevel service password to customer | [12 empty] | 117368253 | Fixed | 804981 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [14 empty] | Password | [8 empty] | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
722 |  |  | China | SD Jinan No.4 Peo. Hosp. | 40038 | 4430000% |  | clean the database | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Configuration Error Repair Activity Code: Database Recreation / Recovery Problem Reported by Customer: 02N/Ingenuity CT 728326 Next Step Action: Q6 = DICOM image database have a lot of image cannot delete, need clean the database T2 Activities:Problem Description: Some files in the image database cannot be deleted Troubleshooting Action: Use PRS to clean older DICOM data Repair Action: Clean host database | Lu, Xueying |  | Use PRS to clean older DICOM data collect bugreport | [12 empty] | 117368007 | Fixed | 804834 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Disk full | [8 empty] | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
723 |  |  | China | HEN Puyang People's Hosp | 40031 | 4430000% |  | The screen turn yellow | T2 Activities : Problem Description: 客户问题概述： 每周一次黄屏故障。 客户问题详细描述或故障现象： 每周一次黄屏故障。 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 升级后，大概每周一次黄屏故障。 远程诊断与测试结果： 到场收集bugreport，提交反馈 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 到场收集bugreport，提交反馈 Test Equipment: N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 16.04.2021 13:38:26 UTC SUP JCOUSER (86038950)服务内容：扫描软件，界面黄屏，屏幕冻结，鼠标点击无响应。 服务结果：重启主机，重置软件，设备恢复正常，继续观察。 其它重要事项：无 Internal Remarks : 16.04.2021 13:38:28 UTC Yongping Luo (86038950) $Q7:Confirmed report problem，Patient examination procedure interface appears yellow screen, and the interface is frozen . $Q8:Perform visiual check on site ,Patient examination procedure interface appears yellow screen, and the interface is frozen, Restart the host, reset softwarethe APP return to normaly. $Q9:Restart the host, reset software ,Device return to full function. $Q10: test and inspection data refer to PA. Device returned to full functionality. $Q11:Calibrated Tools refer to PA. ----------------------------------------------------------------------------- |  |  | 服务内容：扫描软件，界面黄屏，屏幕冻结，鼠标点击无响应。 服务结果：重启主机，重置软件，设备恢复正常，继续观察。 | [12 empty] | 117367559 | Fixed | 804519 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00631367 | [5 empty] | Engineering | [6 empty] | New Yellow Screen | [8 empty] | System issue | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
724 |  |  | China | SD Chest Hosp. | 40025 | 4430000% |  | yellow screen and consulting change password | T2 Activities : Problem Description: 客户问题概述： 其他类型故障。 客户问题详细描述或故障现象： 设备黄屏后重启主机登录提示修改密码，需要工程师到场解决 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 设备黄屏后重启主机登录提示修改密码，需要工程师到场解决 远程诊断与测试结果： 黄屏故障，重启已恢复会正常，现场查看采集故障信息，另密码还有6天过期， PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 修改密码，查看黄屏原因 SAP Transfer : Case Transferred to SAP Next Step Action : HUANG ZHAOKUN External Remarks : 15.04.2021 13:03:25 UTC Zhaokun Huang (86038405)服务内容：远程指导客户修改即将过期的登录密码。 服务结果：设置后，机器恢复正常使用。 Internal Remarks : 16.04.2021 16:30:33 UTC Zhaokun Huang (86038405) “RDI Accuracy: NO" $Q7: (Confirmed reported problem, CT login password soon to expire.) $Q8: ( CT login password need change between 180 days.) $Q9: ( Remote support customer change CT login password.) $Q10: ( Test and inspection data refer to: No PA Required. Device returned to full functionality) $Q11: Calibrated Tools name ( N/A) ^Current Tube installed date: N/A ^Current Tube usage: N/A ----------------------------------------------------------------------------- |  |  | 服务内容：远程指导客户修改即将过期的登录密码。 服务结果：设置后，机器恢复正常使用 | [12 empty] | 117366742 | In Process | 803874 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00631367 | [5 empty] | Engineering | [6 empty] | New Yellow Screen | [8 empty] | System issue | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
725 |  |  | China | JS No.2 Chinese Medicine Hosp. | 345027 | 4430000% |  | The detector is selected to have sound | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : The detector is selected to have sound T2 Activities : Problem Description: 客户问题概述： 其他类型故障。 客户问题详细描述或故障现象： 机架旋转起来声音异常，需工程师到场检查 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 机架旋转有噪音 远程诊断与测试结果： 机架不旋转声音正常，旋转起来声音异常 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 检查旋转相关部件，是否和盖子有摩擦，检查皮带以及张紧轮等 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 18.04.2021 14:45:38 UTC Liangdong Qi (86036024)服务内容：调整机架前盖. 服务结果： 维修后，设备恢复正常工作. Internal Remarks : 18.04.2021 14:45:41 UTC Liangdong Qi (86036024) "RDI Accuracy: Yes“ $Q7: (Confirmed reported problem, there was abnormal sound sometimes when gantry rotating.) $Q8: (No error report, chencked the front cover of gantry and found it could have friction with rotor part.) $Q9: (Adjusted the front cover of gantry.) $Q10: (Test and Inspection data refer to PA & Device return to fully function) $Q11: Calibrated Tools (NA). ----------------------------------------------------------------------------- |  |  | 服务内容：调整机架前盖. 服务结果： 维修后，设备恢复正常工作. | [12 empty] | 117366077 | Closed | 803468 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | NotEngineering | [5 empty] | Others | [6 empty] | Gantry Noise | [8 empty] | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
726 |  |  | China | GD Dongguan Dalingshan Hosp. | 345013 | 4429900% |  | Change the setting | T2 Activities:Problem Description: Summary of customer's problems: Other types of error. Detailed description of Customer's problems or malfunctions: Unable to enter Edit protocol. Not able to edit the scan protocol It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: Unable to enter Edit protocol. Not able to edit the scan protocol Remote troubleshoot and test result: Cannot enter Edit protocol. Unable to edit the scan protocol. Password error, cannot be solved remotely PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Checked the configuration and change the protocol according to the customer's requirement SAP Transfer:Dispatched for On site Support ----------------------------------------------------------------------------- |  |  | Checked the configuration and change the protocol according to the customer's requirement | [12 empty] | 117360309 | In Process | 799382 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Application | [6 empty] | User permission | [8 empty] | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
727 |  |  | China |  | 345009 | 4429900% |  | need second training | BU Q&R Feedback: No Reason: Customer Information Next Step Action: Jun Yang 13817647801 notified the trainer to follow up Problem Reported by Customer: need second training Safety Questions:Case: 0117359722 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Tang Nan Last updated by: Tang Nan 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action: 013523259896 feedback, second training is going to be carried out on Friday, but need Shenfei Yun engineer to be also on site on Friday Shenfeiyun is a new type of CT workstation. This is a product that Shenzhou Company purchased from Philips and resold to the hospital. for service request, have Shenzhou Company follow up on its own. If you encounter customer request service online for ShenfeiYun workstation, currently resort all to contact: 15901552977 Ren Xiao zhe I am very busy online and will process again in the afternoon Next Step Action: 15901552977 Ren Xiaoze said he would coordinate this matter, sent SMS Next Step Action: Customer called in to cancel training, informed 15901552977 Ren Xiao zhe and YANG JUN 13817647801 ----------------------------------------------------------------------------- |  |  | Customer need second training | [12 empty] | 117359722 | Closed | 799004 | FALSE | NAC? | Closed - Non-Complaint | [14 empty] | Application Training | [8 empty] | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
728 |  |  | China | JS Nanjing M.U.A Yifu Hosp. | 345028 | 4429800% |  | the system can not scan | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 05R/Core128/ FSE Notes : This PO is a Reference Number only Safety Questions : Case: 0117354733 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Jianqiang Last updated by: Wang Jianqiang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 扫描失败。 客户问题详细描述或故障现象： 保养时发现，扫描失败，有打火现象 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： "S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT 5:0 APC ARC//S_MASTER_SEQUENCER_NO_XRAY//S_MASTER_SEQUENCER_EXIT_STATE_ERR " 远程诊断与测试结果： 5:0 APM Arc PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： "Recommendation: Review the above *** HV Error Info: *** and follow the X-Ray System Troubleshooting Guide under AEL help (?) / HW Troubleshooting. 1. Put silicone grease on high voltage connector 2. Perform short tube condition 3. Run no load test in order to find fault part (Arcs are indication of a very high probability of a faulty tube, KPM and APM are very rarely the cause of arcs). 4. Check the Power Brush Blocks / slip ring for poor connections 5. Performed Filament Calibrations followed by Long Tube Conditioning." 建议备件12NC及名称： PHILIPS MRC 880 989000086352 //Spellman 80 KW RoHS APM Assy 453567560012 //459800089031 //453566495091 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 15.04.2021 13:58:05 UTC Liuping Yao (86036129)服务内容： 设备不能扫描，球管打火。 灯丝校正和长训球管。 服务结果： 维修后，设备工作正常。 Internal Remarks : 15.04.2021 13:58:08 UTC Liuping Yao (86036129) "RDI Accuracy:Yes" $Q7: (Confirm reported problem. System cannot scan.) $Q8: (Reviewed the error log and found system cannot scan for tube arcing.) $Q9: (Check the HV connectors and nothing abnormal was found. Perform filament calibration and long tube conditioning.) $Q10: (Test and inspection data refer to PA. Device returned to full functionality.) $Q11: Calibrated tools refer to PA. ----------------------------------------------------------------------------- | WANG LIN, LV |  | 服务内容： 设备不能扫描，球管打火。 灯丝校正和长训球管。 服务结果： 维修后，设备工作正常。 | [12 empty] | 117354733 | Closed | 795709 | TRUE | Closed - Done | Closed - Done |  | CTUDT00630085 | [5 empty] | Engineering | [6 empty] | Tube Arc | [8 empty] | System issue | System | IC-HW | Tub Arc | IC-HW-Tub Arc | IC-HW-Tub Arc | CTUDT00630085 | Tub Arc | IC-HW | IC-HW-Tub Arc | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
729 |  |  | China | Zhenjiang M.C.A. Hosp. | 345023 | 4429800% |  | PRS Connection | BU Q&R Feedback: No Reason: Customer Relations SAP Transfer:Dispatched for On site Support Problem Reported by Customer:PRS Connection Next Step Action: FSE Ma Shen shen call in and disp him SAP Transfer:Case Transferred to SAP Safety Questions:Case: 0117353237 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Yu Zhao Last updated by: Yu Zhao 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A External Remarks: 12.04.2021 10:42:21 UTC SUP JCOUSER (86036526) Service details: PRS shows M2M is not connected, re register PRS Service outcome: after the repair, the equipment is back to normal. Internal Remarks: 12.04.2021 10:42:24 UTC Shenshen Ma (86036526) RDI Accuracy: Yes $Q7: (Confirmed reported problem, M2M can't find in PRS) $Q8: (Login the system find PSA can't find) $Q9: (Re-register the PRS) $Q10: (Test and Inspection refer to PA. device return to full function) $Q11:Calibrated Tools(N/A). ----------------------------------------------------------------------------- | Zhai, Lisa |  | PRS Connection | [12 empty] | 117353237 | Closed | 794717 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Service | [6 empty] | PRS Connection | [8 empty] | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
730 |  |  | Japan |  | 341017 | 4430200% |  | CT:APP:AK:ﾎｼ様:質問 | Problem Reported by Customer: Q: I have scanned two people in the same order. I want to divide it, so please tell me how. A: We went with you from change Patient Detail. FSE Notes: This PO is a Reference Number only OneEMS internal remarks: Q1: Did you confirm the reported problem? :Confirmed Q2: Did you confirm any other problems? ;Nothing Q3: What did you do to solve the problem? 1: If resolved by phone: Explanation Q4: Is the device in good condition after the support? : Good Safety Questions: Case: 0117393778 | Xu, Xiaoting |  | Application question | [12 empty] | 117393778 | Fixed | 812529 | FALSE | NAC? | Closed - Non-Complaint | [14 empty] | Q&A | [8 empty] | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
731 |  |  | Japan | Saiseikai Gosu Hospital | 343000 | 4430200% |  | 04CM:KKI:頭部ヘリカル撮影にてリングアーチファクト発生 | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 04DM:KKI:頭部ヘリカル撮影にてリングアーチファクト発生 Safety Questions : Case: 0117393444 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: KAZUMASA KIMURA Last updated by: KAZUMASA KIMURA 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | WANG LIN, LV |  | Ring artifact occurs in head helical photography | [12 empty] | 117393444 | In Process | 812498 | FALSE | Investigation | Closed - Done |  | TBD-NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Ring Artifact | [8 empty] | System issue | System | TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
732 |  |  | Japan |  | 341017 | 4430200% |  | CT:APP:AK:ﾄｶﾞｼ様:質問 | Problem Reported by Customer: Q: I have scanned two people in the same order. I want to divide it, so please tell me how. A: We went with you from change Patient Detail. OneEMS internal remarks: Q1: Did you confirm the reported problem? :Confirmed Q2: Did you confirm any other problems? ;Nothing Q3: What did you do to solve the problem? 1: If resolved by phone: Explanation Q4: Is the device in good condition after the support? : Good | Xu, Xiaoting |  | Application question | [12 empty] | 117393646 | Fixed | 812511 | FALSE | Duplicate with 812529 | Closed - Non-Complaint | [14 empty] | Q&A | [8 empty] | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
733 |  |  | Japan | Higashi-Ohmi General Medical C | 341015 | 4430000% |  | 01CS: TM: When I tried to shoot, Gantry cannot … | Engineer's Comment: '============================== $ version = "" v4.1.10.25210 "" 1:56 incoming call Call content When I tried to shoot, the message Gantry cannot Initialize.Change scan parameter was displayed at the initialization stage and the shooting did not run. I asked the patient to do the test with the phantom, but it also did not run with an error. Also, pressing the Zero Set button does not change the value. When the sleeper is operated, the value of the sleeper position changes. I was a well-built patient, can that cause it? '============================== Troubleshooting Action: What steps are taken during the system diagnosis: Post-response results: Improved Failure location: 01Cosole What steps were taken during the diagnosis ?: '================================== Correspondence Check the log. Judged that it is unlikely to depend on the patient. We had you restart including CT-Host, CIRS, TEAL. After restarting, ZeroSet also responds and shooting is possible. | Zhu, Chun |  | We had you restart including CT-Host, CIRS, TEAL. After restarting, ZeroSet also responds and shooting is possible. | [12 empty] | 117371702 | In Process | 807299 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-NotEnoughInformation | [5 empty] | Others | [6 empty] | Gantry cannot Initialize | [8 empty] | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
734 |  |  | Japan | Maki Hospital | 341017 | 4429900% |  | JY: Adjusting scan irradiation noise | Problem confirmation: I confirmed it. Work details: There is no way to adjust the beeping (during scans) noise on the application. I adjusted the part of the gantry that makes noise. Other problems/Measures taken: There is one. With regard to the error that occurs when images are sent to the PACS, this can be reduced by changing the application template from PACS to Remote System at external facilities. So I changed the template. Equipment status The equipment can be used. Work status Work finished. | Xu, Xiaoting |  | Since I was able to adjust the Ready sound at the start of scanning on the software, I adjusted it on the software. | [12 empty] | 117360732 | New | 799679 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Volume adjustment | [8 empty] | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
735 |  |  | Japan | Saiseikai Gosu Hospital | 343000 | 4429800% |  | 04DM:TM:Vertical lines appeared in one Surview scan.Since then, there has been no abnormality in subsequent [Truncated] | T2 Activities:Problem Description: Problem description by engineer : Equipment affected: CT Problem location: 04 Gantry Exclude HV When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user's work flow did the problem occur?: (During Acquisition/Scan) What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem?: NA Engineer's Comment: '============================== Call description: Vertical lines appeared in one Surview scan.Since then, there has been no abnormality in subsequent scans. '============================== External Remarks : 13.04.2021 09:45:15 UTC YUKI KOBAYASHI (81100095)[4 Elements] 1. Was the reported problem confirmed?: Confirmed. 2. Were any other problems confirmed?: 3. What did you do to resolve the problem?: I am canceling this SWO as it duplicates the content of SWO53390892. We responded to this issue on April 12 under that SWO. 4. Is the equipment operating normally following the response? Good. 13-Apr-2021 81100095 FSE Handover Tool1.3.4.32612.DB13 20210115 ----------------------------------------------------------------------------- | LIU, Ankor |  | Vertical lines appeared in one Surview scan.Since then, there has been no abnormality in subsequent scans. Recommended 12Nc: PR1 459800911271: TDMS V2 M.Det Class A FRU QTY: 1 SLOC: Osaka (U615) | [12 empty] | 117352301 | Fixed | 794133 | TRUE | Investigation | Closed - Done |  | CTUDT00632110 | [5 empty] | Engineering | [6 empty] | Vertical line | [8 empty] | System issue | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | DMS Module | DMS-HW | DMS-HW-DMS Module | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
736 |  |  | Spain | MAZ ZARAGOZA M.A.T.E.P.S.S.11 | 40042 | 4429800% | 15 | Unit is down | They inform me that at the end of a study, and taking the patient off the table, the gantry is turned off. It takes them a couple of starts to get the system working again. In the error log I see that when removing the patient there is an e-stop due to movement of the table not allowed. Then when starting up again, the main contactor does not close properly and you have to try again. The rest of the morning they work without problems, during the diagnosis, I carry out test studies, the system works correctly. | 2 |  | In the error log I see that when removing a patient there is an e-stop due to not allowed couch movement. Then when starting the equipment again, the main contactor does not close correctly and they have to try again. Check if have bugreport can share engineering investigation, and perform couch calibration if necessary | [12 empty] | 117357337 | Closed | 795837 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Power Issue | [5 empty] | Engineering | [6 empty] | Couch E-Stop | [8 empty] | System issue | System | Gantry-HW | PowerOff | Gantry-HW-PowerOff | Gantry-HW-PowerOff | TBD-PowerOff | PowerOff | Gantry-HW | Gantry-HW-PowerOff | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
737 |  |  | Sweden | Avesta lasarett | 40040 | 4430200% |  | Problems with phantom positioning at IQ Check | Users have problems with monthly IQ check on our two latest CTs, they center the phantom according to the laser but after scan the system asks the user to move the phantom about 6.9mm to the right, if the users do as described they get a new message to move the phantom about 22mm to left, do as the instruction states so at the next scan they get to move the phantom about 12mm to the right and then they get stuck in this loop by moving the phantom back and forth, a user had tried to push through this IQ check but gave up after 15 minutes moving back and forth. They showed me that if you redo the IQ check and at the instructions move 6.9mm to the right, only move about 1.5mm to the right, so it goes on. Thank you if your technicians can take a look at this when you come out and install FCO on DT Ludvika SE5333 and DT2 Avesta SE5301. Follow Up Required : When doing monthly check, user have problems centering phantom. Agreed to have a FSE look at this when doing FCO upgrade on WO-03073973, scheduled 28APR21. |  |  | The guidance of centering phantom is not accurate | [12 empty] | 117397366 | In Progress | 813853 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00630107 | [5 empty] | Engineering | [6 empty] | Phantom positioning | [8 empty] | System issue | System | DMS-SW | Phantom Detect | DMS-SW-Phantom Detect | DMS-SW-Phantom Detect | CTUDT00630107 | Phantom Detect | DMS-SW | DMS-SW-Phantom Detect | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
738 |  |  | Sweden | Ludvika Lasarett | 40049 | 4430200% |  | Problems with phantom positioning at IQ Check | Work Order #: WO-03334617 Knudsen Martin The system partly meets specification and is returned to use with restrictions Diagnostic performed by Engineer : When doing monthly check, user have problems centering phantom. Agreed to have a FSE look at this when doing FCO upgrade on WO-033282269, not yet scheduled. Problem Description by Engineer : When doing monthly check, user have problems centering phantom. Follow Up Required : When doing monthly check, user have problems centering phantom. Agreed to have a FSE look at this when doing FCO upgrade on WO-033282269, not yet scheduled. No parts. -MK ----------------------------------------------------------------------------- |  |  | The guidance of centering phantom is not accurate | [12 empty] | 117397371 | In Progress | 813857 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00630107 | [5 empty] | Engineering | [6 empty] | Phantom positioning | [8 empty] | System issue | System | DMS-SW | Phantom Detect | DMS-SW-Phantom Detect | DMS-SW-Phantom Detect | CTUDT00630107 | Phantom Detect | DMS-SW | DMS-SW-Phantom Detect | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
739 |  |  | Turkey | T.C Saglik Bakanligi Bagcilar -> Bağcılar State Hospital | 40010 | 4430500% |  | Spill Guard order | Work Order #: WO-03341490 Soner Saglam The system fully meets specification and returned to use Problem Description by Engineer: ordering spill guards for opaque. Resolution: Spill guard ordered. Follow Up Required: Spill guard has been ordered. ----------------------------------------------------------------------------- |  |  | Parts Consumed Description: 1.00000,453566136722,APRON PACKING-SPILL GUARD | [12 empty] | 117403995 | In Progress | 815531 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Others | [6 empty] | SPILL GUARD | [8 empty] | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
740 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 341006 | 4430500% |  | Spill Guard siparişi | Work Order #: WO-03341457 Soner Saglam The system fully meets specification and returned to use Problem Description by Engineer : Opak kaçmasını engellemek için spill guard siparişi. Resolution : Spill guard sipariş edildi. Follow Up Required : Spill guard sipariş edildi. |  |  | Parts Consumed Description: 1.00000,453566136722,APRON PACKING-SPILL GUARD | [12 empty] | 117403969 | In Progress | 815517 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Others | [6 empty] | SPILL GUARD | [8 empty] | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
741 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 341006 | 4430500% |  | There are artifacts in the images | "Work Order #: WO-03341105 Soner Saglam The system partly meets specification and is returned to use with restrictions Resolution: Aplane analyzed, no crack. Images were examined. There are 2 distant non-homogeneous rind structures. There is a high probability that opaque material got into the detector. Control with site visit is required. Problem Description by Engineer: Images contain ring artifacts. Follow Up Required: Aplane analyzed, no cracks. Images were examined. There are 2 distant non-homogeneous rind structures. There is a high probability that opaque material got into the detector. Control with site visit is required. |  |  | There is a high probability that opaque material got into the detector. | [12 empty] | 117403475 | In Progress | 815406 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | ECR-075975 | [5 empty] | Engineering | [6 empty] | Ring Artifact | [8 empty] | System issue | System | Gantry-HW | Lexan Ring | Gantry-HW-Lexan Ring | Gantry-HW-Lexan Ring | ECR-075975 | Lexan Ring | Gantry-HW | Gantry-HW-Lexan Ring | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
742 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 341006 | 4430200% |  | Device freezes. | Work Order #: WO-03332907 Turgay Baylan Diagnostic performed by Engineer : N/A Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: ----------------------------------------------------------------------------- |  |  | The bug report was taken and shared with Remote Service Engineer. The system has been delivered in working condition. | [12 empty] | 117395678 | In Progress | 813185 | FALSE | Investigation | Closed - Done |  | CTUDT00631367 | [5 empty] | Engineering | [6 empty] | New Yellow Screen | [8 empty] | System issue | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
743 |  |  | Turkey | T.C Saglik Bakanligi Bagcilar -> Bağcılar State Hospital | 40010 | 4430200% |  | artefact in thin section | "Work Order #: WO-03334803 Soner Saglam The system fully meets specification and returned to use Problem Description by Engineer: Artifact in thin sections. Resolution: UHR filter may have got opaque, site visit required. In the examination, it was observed that contrast material got into the UHR filter of the device. Contrast material was cleaned and air cal was applied. The system has been tested and delivered in working condition. |  |  | opaque material got into the detector. | [12 empty] | 117397528 | In Progress | 810380 | TRUE | Closed - Done | Closed - Done |  | ECR-075975 | [5 empty] | Engineering | [6 empty] | Ring Artifact | [8 empty] | System issue | System | Gantry-HW | Lexan Ring | Gantry-HW-Lexan Ring | Gantry-HW-Lexan Ring | ECR-075975 | Lexan Ring | Gantry-HW | Gantry-HW-Lexan Ring | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
744 |  |  | Turkey | T.C Saglik Bakanligi Bagcilar -> Bağcılar State Hospital | 40010 | 4430100% |  | There are artifacts in images. | "Work Order #: WO-03325607 Soner Saglam The system partly meets specification and is returned to use with restrictions Problem Description by Engineer: Images have ring artefact. Resolution: There are 2 ring artefacts in the images that are not symmetrical with the screen. Site visit required. The Compensator analysis program failed to analyze. Follow Up Required: There are 2 ring artefacts that are not symmetrical with the screen. Site visit required. The Compensator analysis program failed to analyze. -------------------------------------------------- --------------------------- Work Order #: WO-03325888 15-Apr-2021 Ugur, Demirbas The system fully meets specification and returned to use Problem Description by Engineer: Images have ring artefact. Resolution: The contrast material was detected on the detector and cleaned. Air cal was done and the system was delivered in working condition. |  |  | opaque material got into the detector. | [12 empty] | 117373824 | Incomplete | 804443 | TRUE | Closed - Done | Closed - Done |  | ECR-075975 | [5 empty] | Engineering | [6 empty] | Lexan Ring/Contrast | [8 empty] | System issue | System | Gantry-HW | Lexan Ring | Gantry-HW-Lexan Ring | Gantry-HW-Lexan Ring | ECR-075975 | Lexan Ring | Gantry-HW | Gantry-HW-Lexan Ring | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
745 |  |  | United States | McLaren Macomb | 40021 | 4430100% |  | Tube Heat Issue | Work Order #: WO-03328908 Mujica, David The system fully meets specification and returned to use Problem description by Clinical Support Specialist : Customer Function/Role: Technologist Device Use: In use Expected and Unexpected Results: Tube Heat Issues User Impact: none Patient Impact: none Current Software Version: 4.1 Diagnostic performed by Engineer : Michael was aided to decrease tube heat of their CT PE/Abd/Pel Exam Cards. Cycle times in the bolus tracker were decreased by 33% and a lower DRI were used to decrease tube heat. Michael was thankful for the aide and has successfully implemented these changes to the Exam Cards. Michael has no further questions or need for clinical support at this time. No rescan, no loss of data, no reinjection. |  |  | Cycle times in the bolus tracker were decreased by 33% and a lower DRI were used to decrease tube heat. | [12 empty] | 117381762 | In Progress | 806332 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD_Tube hot_bugreport | [5 empty] | Engineering | [6 empty] | Tube Overheat | [8 empty] | System issue | System | Gantry-SW | TubHot | Gantry-SW-TubHot | Gantry-SW-TubHot | TBD_Tube hot_bugreport | TubHot | Gantry-SW | Gantry-SW-TubHot | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
746 |  |  | United States | Clay County Medical Center | 40003 | 4430000% |  | Fails QA scans | Work Order #: WO-03323757 14-Apr-2021 Stauffer Jeffrey The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Radiology Technician Device Use: General CT Exams Expected and Unexpected Results: Expected that the system would pass Daily QA scans. Unexpected - hard to get to pass Daily QA scan. System has you move phantoms multiple times. User Impact: Delayed scans. Patient Impact: None Current Software Version: 4.17.10 Diagnostic performed by Engineer : Looked and phantoms and found the head phantom was tilted down. Found head phantom needed moved to the left and the body phantom needed moving to the right. Resolution : Tilted head phantom was tilted back. Moved head phantom to the left and the body phantom to the right. System then ran and passed the daily QA scan. Tightened adjustment pins on phantoms for repeatable scans. ----------------------------------------------------------------------------- | WANG LIN, LV |  | Tilted head phantom was tilted back. Moved head phantom to the left and the body phantom to the right. System then ran and passed the daily QA scan. Tightened adjustment pins on phantoms for repeatable scans. | [12 empty] | 117372009 | Closed | 803376 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00630107 | [5 empty] | Engineering | [6 empty] | Phantom positioning | [8 empty] | System issue | System | DMS-SW | Phantom Detect | DMS-SW-Phantom Detect | DMS-SW-Phantom Detect | CTUDT00630107 | Phantom Detect | DMS-SW | DMS-SW-Phantom Detect | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
747 |  |  | United States |  | 40015 | 4429900% |  | customer has question about acr | Work Order #: WO-03314706 Hirchak, Michele The system fully meets specification and returned to use Problem description by Clinical Support Specialist : Customer Function/Role: TECH Device Use: In Use Expected and Unexpected Results: ACR Current Software Version: 4.1.7 No rescan, no loss of data, no reinjection Diagnostic performed by Engineer : Talked about where to find the mA station. Discussed using the EC manager as a calculator. Also, emailed ACR guide. Keeping open until Friday. MBH |  |  | customer has question about acr | [12 empty] | 117363982 | Closed | 797612 | FALSE | NAC? | Closed - Non-Complaint | [14 empty] | Q&A | [8 empty] | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
748 |  |  | United States | Holston Medical Group | 40053 | 4428500% |  | What did not work as they expected? Physics phantom does not securely attach to the head phantom | Physics phantom does not securely attach to the head phantom | XU, Miley |  | Design improvement and add instruction to guide phantom attachment | [14 empty] | 797302 | TRUE | Investigation | Closed - Duplicate |  | NotEngineering-Clinical CTUDT00632803 | [5 empty] | Engineering | [6 empty] | Physics Phantom | [8 empty] | System issue | system | Couch-HW | Phantom | Couch-HW-Phantom | Couch-HW-Phantom | CTUDT00632803 | Phantom | Couch-HW | Couch-HW-Phantom | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
749 |  |  | United States | Select Specialty Hospitals | 310051 | 4430200% | 15 | Table cuts off the Gantry | Work Order #: WO-03337215 16-Apr-2021 Siker Shawn The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : Gullwing switch needs replacing Follow Up Required : Replace gullwing switch Thomas Marx will return monday to replace switch | 2 |  | 1.00000,453566147251,SWITCH FRNT INTERLCK 1.00000,459800712461,FRU, AMC upgrade KITS Collect bugreport for eng investigation | [12 empty] | 117399414 |  | 812254 | TRUE | Closed - Done | Closed - Done |  | TBD_CouchEstop | [5 empty] | Engineering | [6 empty] | Couch E-Stop | [5 empty] | Refresh |  |  | Upgrade issue | System | Couch-SW | TBD_CouchEstop | Couch-SW-TBD_CouchEstop | Couch-SW-TBD_CouchEstop | TBD_CouchEstop | TBD_CouchEstop | Couch-SW | Couch-SW-TBD_CouchEstop | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
750 |  |  | United States | NORTHWEST RADIOLOGY/MRI | 300057 | 4430200% |  | 4.1.10 Rings on images | Work Order #: WO-03337555 Hirchak, Michele The system fully meets specification and returned to use Problem description by Clinical Support Specialist : Customer Function/Role: Tech Device Use: In Use Expected and Unexpected Results: Rings on high resolution and standard resolution. User Impact: None Patient Impact: None Current Software Version: 4.1.10 No rescan, no loss of data, no reinjection. Diagnostic performed by Engineer : LOTS did not work. Customer has rings on high res hip/knee/ankle protocol. Had her run a recon with higher levels of iDose and without iDose. Had her run another air cal. As well. Conf RSE on the line. He will page the FSE for Monday. MBH ----------------------------------------------------------------------------- Work Order #: WO-03337594 Weaver David The system fully meets specification and returned to use Problem Description by Engineer : DESCRIPTION OF PROBLEM (EXPECTED/UNEXPECTED RESULTS) Rings on images CALLER FUNCTION/ROLE Technician/Sonographer/Therapist DEVICE INTENDED USE Diagnostic WHEN WAS THE FIRST OCCURRENCE? WITHIN . . . <1 Hour HOW OFTEN IS ISSUE OCCURRING? First occurrence WHERE IN THE END-USER WORKFLOW IS THE ISSUE OCCURRING? Pre-acquisition/scan WHAT RECENT EVENTS HAVE OCCURRED? During normal use NEGATIVE OPERATOR IMPACT Can not scan patients NEGATIVE PATIENT IMPACT/PROCEDURE OUTCOME delay in scanning HAS CUSTOMER CREATED A SPECIFIC SYSTEM EVENT REPORT OF ISSUE? No DEVICE SOFTWARE VERSION 4.1.10 Diagnostic performed by Engineer : Rings on images on all scans. Resolution : Rings on images on all scans. ----------------------------------------------------------------------------- | 5 |  | Called to f/u with Bridgette. She said all rings are gone. OK to close | [12 empty] | 117399597 |  | 812141 | TRUE | Closed - Done | Closed - Done |  | NotEngieering-Clinical | [5 empty] | Application | [6 empty] | Ring Artifact | [5 empty] | Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
751 |  |  | United States |  | 310167 | 4430000% |  | Accidentally hit E-Stop button | Work Order #: WO-03323734 Mujica, David The system fully meets specification and returned to use Problem description by Clinical Support Specialist : Customer Function/Role: Technologist Device Use: In use Expected and Unexpected Results: E-stop triggered User Impact: none Patient Impact: none Current Software Version: 4.1 Diagnostic performed by Engineer : Kinn accidentally hit the e-stop button on the gantry and was aided to disengage the e-stop button and shut down the CT system to restart the gantry. The shut down was successful and by turning the key on the console, the gantry successfully was initiated. The CT scanner is working properly. Kinn has no further need for clinical support. No rescan, no loss of data, no reinjection. ----------------------------------------------------------------------------- | 4 |  | Accidentally hit E-Stop button | [12 empty] | 117371981 |  | 802950 | #N/A | #N/A | Closed - No Further Investigation | [14 empty] | E-Stop | [5 empty] | Non-Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | y
752 |  |  | United States |  | 310034 | 4429900% |  | nemo / dennis | Work Order #: WO-03314367 Parts Consumed : 2.00000,453567560301,Nut, Hex, Nylon 6/6 2.00000,453567560291,Set Screw-Nylon ----------------------------------------------------------------------------- | E, Sravanthi |  | nemo / dennis | [12 empty] | 117363723 |  | 801854 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [14 empty] | nemo | [5 empty] | Non-Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | y
753 |  |  | Netherlands |  | 52101 | 4430000% |  | System is slow after restart due to blue screen. | Diagnostic performed by Engineer: system slow after restart bluescreen, application now restarts but crashes after exam card selection Resolution: Acornis restored from the C drive together with a service backup. System tested and released for use. | 2 |  | restored from the C drive together with a service backup. System tested and released for use. Andy, please Communicate with Local team to inform the Exteranl tool is not allowed | [12 empty] | 117369335 |  | 801611 | #N/A | #N/A | Evaluation - Pending Translation | [14 empty] | Blue screen | [5 empty] | Non-Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | y
754 |  |  | United States |  | 310034 | 4429800% |  | 0308CDY70 intermittent couch issues | Work Order #: WO-03307520 Mcginnis, Mike The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Problem description by engineer : intermittent couch issues How often is issue occurring? : Intermittent Malfunction area : Table Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected results: intermittent couch issues Patient Impact/Outcome: none User Impact: intermittent couch issues Current Software Version: 3.6.8.25004 Follow Up Required : Error # and/or description of error : 2021/04/12 10:57:17.656 S_MC_DEVICE_COUCH_VERTICAL_HW_FAULT ERROR_WARNING 2021/04/12 10:57:17.750 S_MC_DEVICE_COUCH_VERTICAL_HW_FAULT ERROR_NORMAL 2021/04/12 10:57:18.828 S_COUCHMGR_UNEXPECTED_EVENT_ERROR ERROR_NORMAL 2021/04/12 10:57:34.531 S_TIMER_WATCHDOG_CAUSED_LAST_RESET ERROR_NORMAL 2021/04/12 10:57:36.640 S_MP_MPSHELL_SERIAL_LIFE_SYNC_ERROR ERROR_NORMAL Completed or Recommended Repair Action "Internal" : Check the green bar safety switch and all cables associated with it. 453567535032 KIT-ENCODER BELT; 453567247021 Limit Switch, Safety Bar Brkt Limit switch located at the end of the green vertical support; 453566491851 CBL, Vertical Limits/Brake; 453567484781 Inverter Panel Assembly Is the complete unit; 453566502431 Vertical Endcoder Assy; 459800033291 Limit Switch Assembly 459800033291 Internal Comments : Error # and/or description of error : 2021/04/12 10:57:17.656 S_MC_DEVICE_COUCH_VERTICAL_HW_FAULT ERROR_WARNING 2021/04/12 10:57:17.750 S_MC_DEVICE_COUCH_VERTICAL_HW_FAULT ERROR_NORMAL 2021/04/12 10:57:18.828 S_COUCHMGR_UNEXPECTED_EVENT_ERROR ERROR_NORMAL 2021/04/12 10:57:34.531 S_TIMER_WATCHDOG_CAUSED_LAST_RESET ERROR_NORMAL 2021/04/12 10:57:36.640 S_MP_MPSHELL_SERIAL_LIFE_SYNC_ERROR ERROR_NORMAL Completed or Recommended Repair Action "Internal" : Check the green bar safety switch and all cables associated with it. 453567535032 KIT-ENCODER BELT; 453567247021 Limit Switch, Safety Bar Brkt Limit switch located at the end of the green vertical support; 453566491851 CBL, Vertical Limits/Brake; 453567484781 Inverter Panel Assembly Is the complete unit; 453566502431 Vertical Endcoder Assy; 459800033291 Limit Switch Assembly 459800033291 Diagnostic performed by Engineer : Dennis, biomed stated they have been having issues with the couch motion. He has replaced vertical string pot, inverter and couch board. ----------------------------------------------------------------------------- Work Order #: WO-03307843 Hanlin Christopher +I756:I967 |  |  | 2021-04-12 | [12 empty] | 117357843 |  | 793409 | #N/A | #N/A | Evaluation | [14 empty] | Non-Refresh | [5 empty] | Non-Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | y
755 |  |  | Argentina | Hospital Dr. Bernardo Houssay | 40027 | 4430500% |  | A DENSE, RING-SHAPED IMAGE IS DISPLAYED IN THE ACQUISITION. | Work Order #: WO-03344466 Alec, Preciado The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Problem description by engineer : Artifact after air calibration. When was the first occurrence? Within... : April 19, 2021 10:00:00 How often is issue occurring? : Multiple times a day Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected Results: Clean image, no artifact Patient Impact/Outcome: Delayed medical care and diagnosis. User Impact: Delayed patient care. Current Software Version: Not provided. Diagnostic Performed by Engineer: We contact user, field service engineer to provide feedback who will go to the site. Internal Comments : Error # and/or description of error : Artifact in the image. Completed or Recommended Repair Action "Internal" : Ring-type artifact is displayed in the image. We suggest them air calibration, bad detector, tables FS0- FS1, Aplane, compensator and external objects check-up. It does not have PRS. Follow Up Required : What are the next Troubleshooting steps : Ring-type artifact is displayed in the image. We suggest them air calibration, bad detector, tables FS0- FS1, Aplane, compensator and external objects check-up. Error # and/or description of error : Artifact in the image. Completed or Recommended Repair Action "Internal" : Ring-type artifact is displayed in the image. We suggest them air calibration, bad detector, tables FS0- FS1, Aplane, compensator and external objects check-up. It does not have PRS. ----------------------------------------------------------------------------- Work Order #: WO-03346062 Meza, Edgardo Arturo Problem Description by Engineer : Problem description by engineer : Artifact after air calibration. When was the first occurrence? Within... : April 19, 2021 10:00:00 How often is issue occurring? : Multiple times a day Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected Results: Clean image, no artifact Patient Impact/Outcome: Delayed medical care and diagnosis. User Impact: Delayed patient care. Current Software Version: Not provided. Diagnostic performed by Engineer: N/A | 3 |  | Detector module No. 2 damaged is identified. | [12 empty] | 117406681 | In Progress | 817035 |  | Investigation | Closed - Done |  | CTUDT00632110 | [5 empty] | Engineering | [6 empty] | DMS Module | [5 empty] | Refresh |  |  | System issue | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | DMS Module | DMS-HW | DMS-HW-DMS Module | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
756 |  |  | China | SD Tai'an T.C.M. Hosp. No.2 Ho | 345037 | 4431100% |  | 02n/Ingenuity Core128/Elite China/artifact | Problem Reported by Customer : 02n/Ingenuity Core128/Elite China Next Step Action : FSE 陈涛18560023293 call in and disp him Next Step Action : 已回访 SAP Transfer : Case Transferred to SAP T2 Activities : Problem Description: 客户问题概述： 图像伪影 客户问题详细描述或故障现象： 图像伪影 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： NA 远程诊断与测试结果： 报修图像伪影，之前出现过类似伪影，是由于钨片的Flat Spring断裂掉到准直器造成，这次需要订三个一起换掉。 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 检查Flat Spring，检查原始数据分析是否其他原因造成伪影，做图像质量校准 建议备件12NC及名称： " 453567268731 Flat Spring" SAP Transfer : Dispatched for Onsite Support Next Step Action : PPD 453567268731*3 正在核实替代号，请先派工，金晓-86-13370562316 | 5 |  | Replace flat spring (*3) | [12 empty] | 117438939 | In Process | 839450 |  | Investigation | Closed - Done |  | CTUDT00632833 | [5 empty] | Engineering | [6 empty] | Flat Spring | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | FlatSpring | Gantry-HW-FlatSpring | Gantry-HW-FlatSpring | CTUDT00632833 | FlatSpring | Gantry-HW | Gantry-HW-FlatSpring | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
757 |  |  | China | JS No.2 Chinese Medicine Hosp. | 345027 | 4431100% |  | 02n/Ingenuity Core128/Elite China | BU Q&R Feedback: No SAP Customer Symptom Code: BOOT Customer Symptom Code: Boot up/Power-up Problem Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer : 02n/Ingenuity Core128/Elite China Safety Questions : Case: 0117438385 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Cao Last updated by: Liu Cao 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: Call for logger checking Troubleshooting Action: S_tilt_C_no_move error, not more repeat. Check P1/P2 is same at 497. Could be motor or PCB issue. Repair Action: No more error repeat, FSE will follow this issue Test Equipment: N/A ----------------------------------------------------------------------------- | 3 |  | No more error repeat, FSE will follow this issue, FSE check the belt tension force is too high | [12 empty] | 117438385 | Fixed | 838961 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | ECR-077012 | [5 empty] | Engineering | [6 empty] | Tension Force | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Belt | Gantry-HW-Belt | Gantry-HW-Belt | ECR-077012 | Belt | Gantry-HW | Gantry-HW-Belt | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
758 |  |  | China | JL Songyuan Central Hosptial | 345008 | 4431000% |  | 02N/Core128/Yellow screen | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : Yellow screen SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Safety Questions : Case: 0117437773 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Pan Chuchu Last updated by: Pan Chuchu 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action : FSE 华天-18946514633申请开此维修单给他后续找TSE申请新的Escalation Internal Remarks : 25.04.2021 01:28:45 UTC Tian Hua (86038446) RDI Accuracy ：No $Q7: (Confirmed reported problem&Computer Yellow screen ) $Q8: (Check the computer ，Software problem and Update Escalation ） $Q9: (Update Escalation） $Q10: (Test and inspection data refer to: PA. Device returned to full functionality) $Q11 Calibrated Tools name (FLUKE287C) S/N (27420003) Next Calibration date (03/17/2022) External Remarks : 25.04.2021 01:28:43 UTC Tian Hua (86038446)服务内容： 软件偶发黄屏，上Escalation。 服务结果： 重启后，设备恢复正常工作。 其它重要事项： 无。 ----------------------------------------------------------------------------- | 5 |  | New Yellow Screen issue on HF3, waiting workaround publish | [12 empty] | 117437773 | Closed | 838623 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00631367 | [5 empty] | Engineering | [6 empty] | New Yellow Screen | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
759 |  |  | China | SD Laiyang People Hospital | 345005 | 4431000% |  | 02N/Ingenuity Core128/Elite China/software crash | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ Safety Questions : Case: 0117437575 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Guoping Last updated by: Liu Guoping 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action : Ma jun SAP Transfer : Dispatched for Onsite Support SAP Transfer : Case Transferred to SAP T2 Activities : Problem Description: 客户问题概述： 计算机软件相关故障。 客户问题详细描述或故障现象： 计算机反应速度慢，黄屏 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 计算机反应速度慢，黄屏 远程诊断与测试结果： 怀疑软件问题 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 1检查计算机剩余空间是否合适 2 检查床是否有异物 3 重装软件 建议备件12NC及名称： ----------------------------------------------------------------------------- | 3 |  | New Yellow Screen issue on HF3, waiting workaround publish | [12 empty] | 117437575 | In Process | 838420 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00631367 | [5 empty] | Engineering | [6 empty] | New Yellow Screen | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
760 |  |  | China | JL Songyuan Central Hosptial | 345008 | 4431000% |  | escalation/EQ:83704366 | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer:escalation Safety Questions:Case: 0117437412 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Zhang Jing Last updated by: Zhang Jing 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 3 |  | New Yellow Screen | [12 empty] | 117437412 | New | 838302 |  |  | Closed - Non-Complaint | [14 empty] | New Yellow Screen | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
761 |  |  | China | Guangzhou No8 Peoples Hospital | 345032 | 4430800% |  | 02N/Ingenuity Core128/ring artifact | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity Core128/ Next Step Action : 02036549013 查老师来电 SAP Transfer : Dispatched for Onsite Support T2 Activities : Problem Description: 客户问题概述： 图像伪影。 客户问题详细描述或故障现象： 图像半环形伪影 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 图像半环形伪影 远程诊断与测试结果： 有坏通道 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 根据伪影手册做相关的伪影故障排除 建议备件12NC及名称： 459800911281 TDMS V2 M.Det Class B FRU Safety Questions : Case: 0117423105 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Jianqiang Last updated by: Wang Jianqiang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP Next Step Action : PPD TDMS V2 M.Det Class B FRU 459800911281 1 已安排24H到医院，FSE朱光录-86-13902259150 Next Step Action : zhu guanglu External Remarks : 23.04.2021 04:04:54 UTC Guanglu Zhu (86034061)服务内容: 高密度环形伪影，12模块探测器有坏通道，更换探测器模块，完成相应校正，恢复正常 服务结果: 服务完成后,设备恢复正常工作 Internal Remarks : 25.04.2021 05:19:51 UTC Guanglu Zhu (86034061) $ RDI Accuracy: yes $Q7: ( Confirmed reported problem+ring artifact in images ) $Q8: (analysis rawdata confirm detector 177 has bad ) $Q9: (replace 12 module and finish recommended callibration ) $Q10: (Test and Inspection data refer to PA & Device status return to fully function) $Q11: Calibrated Tools name (na ) S/N (na ) Next Calibration date (na ) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date: na ^Current Tube exposure:NA ----------------------------------------------------------------------------- | 3 |  | Ring Artifact due to module 12, issue fixed by replacing module | [12 empty] | 117423105 | Closed | 831697 |  | Closed - Done | Closed - Done |  | CTUDT00632110 | [5 empty] | Engineering | [6 empty] | DMS Module | [5 empty] | Refresh |  |  | System issue | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | DMS Module | DMS-HW | DMS-HW-DMS Module | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
762 |  |  | China | SD Chest Hosp. | 40025 | 4430700% |  | 02N/Ingenuity CT 728326/MODIFT TITLE | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer : 02N/Ingenuity CT 728326/ Safety Questions : Case: 0117418579 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Cao Last updated by: Liu Cao 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: Need modify image title Troubleshooting Action: Need modify image title Repair Action: Instruct customer Test/Inspection: N/A Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | Instruct customer to modify image title | [12 empty] | 117418579 | Fixed | 828482 |  | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Application | [6 empty] | Image title | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
763 |  |  | China | Guiyang Fuyou Hosp | 40033 | 4430600% |  | 02N/Ingenuity CT 728326/check system | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity CT 728326/ SAP Transfer : Case Transferred to SAP T2 Activities : Problem Description: 客户问题概述： 新装机，需要工程师检查场地，教客户开关机，日常保养 设备本身技术问题： 是。 Troubleshooting Action: 远程诊断与测试结果： 新装机，需要工程师检查场地，教客户开关机，日常保养 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 新装机，需要工程师检查场地，教客户开关机，日常保养 客户需求备注： 联系报修人，确认时间 SAP Transfer : Dispatched for Onsite Support Safety Questions : Case: 0117411517 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Weng Jinming Last updated by: Weng Jinming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action : zhou zhongxiang External Remarks : 22.04.2021 16:13:07 UTC SUP JCOUSER (86038997)服务内容：确认报修故障，新装机，需要工程师检查场地，教客户开关机，日常保养 。现场指导客户。 服务结果：服务后设备恢复正常。 ----------------------------------------------------------------------------- | 5 |  | New installed system, intruct customer to power on/off and routine maintainance of system. | [12 empty] | 117411517 | In Process | 824020 |  |  | Closed - Non-Complaint | [14 empty] | NAC? | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
764 |  |  | China | Wuhan Ren 'ai Hospital, Hubei Province | 345004 | 4430600% |  | praise/Ingenuity Core128/HU FANGTAO | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer: praise Next Step Action: Cheng Zhenhua follow Safety Questions:Case: 0117410771 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Tang Nan Last updated by: Tang Nan 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 3 |  | No issue report from customer | [12 empty] | 117410771 | Closed | 823559 |  |  | Closed - Non-Complaint | [14 empty] | NAC? | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
765 |  |  | China | Wuhan Ren 'ai Hospital, Hubei Province | 345004 | 4430600% |  | 02N/Ingenuity Core128/modify CT ip | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity Core128 FSE Notes : This PO is a Reference Number only T2 Activities : Problem Description: 客户问题概述： 需要修改host IP 客户问题详细描述或故障现象： 需要修改host IP 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 需要修改host IP 远程诊断与测试结果： 远程指导客户进了LAN CONFIG，网关改不了， PRS TR状态： 不通，请工程师到场检查。 Repair Action: 建议现场维修步骤： 根据客户需要修改IP地址，重新配置PACS 打印机等等相关连接设备 Safety Questions : Case: 0117409738 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Jianqiang Last updated by: Wang Jianqiang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 21.04.2021 15:25:00 UTC SUP JCOUSER (86036493)[服务内容]：医院IP变动，重新添加节点后恢复。 [服务结果]：维修后，设备恢复正常。 Internal Remarks : 21.04.2021 15:25:03 UTC FangTao Hu (86036493) RDI Accuracy：Yes $Q7: (Confirm reported problem &components connection fault ) $Q8: (hospital IT change the IP of PACS and prints leads to connection broken) $Q9: (Add ports and adjust IP adress ,system return to normal ) $Q10: (No PA required ,Device return to full functionality) $Q11: Calibrated Tools name (N/A) S/N (N/A) Next Calibration date (N/A) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date: N/A ^Current Tube usage: scan second_N/A ----------------------------------------------------------------------------- | 3 |  | Change host IP per customer IT assignment | [12 empty] | 117409738 | Closed | 823037 |  | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | IP Config | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
766 |  |  | China | The Second People's Hospital | 345039 | 4430600% |  | wechat:02n/Ingenuity Core128/Elite China/Run slow | BU Q&R Feedback: No SAP Customer Symptom Code: HANG Customer Symptom Code: Hang/Crash Reason: Corrective Maintenance Engineer Failure Code: Wear-out/Breakage/Mechanical Failure Repair Activity Code: Reboot/Restart Problem Reported by Customer:run slow Safety Questions:Case: 0117409659 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Lili Last updated by: Lu Lili 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: host software run slow Troubleshooting Action: reboot computer Repair Action: reboot computer Test/Inspection: customer confirm can scan patient and image ok.functional Check PASSED. The device returned to full functionality. Test Equipment: N/A SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | host software run slow, instruct customer to reboot system, The device returned to full functionality. Intermitent issue, no bugreport saved, similar with YS | [12 empty] | 117409659 | Closed | 822981 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00631367 | [5 empty] | Engineering | [6 empty] | New Yellow Screen | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
767 |  |  | China | HEN Puyang People's Hosp | 40031 | 4430300% |  | 02n/Ingenuity CT 728326/tilt error | Next Step Action : 客户013903937617再次来电，转接韩工 T2 Activities : Problem Description: 客户问题概述： 角度问题 客户问题详细描述或故障现象： 角度问题 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： S TILT C POT1 PO2 KO, S TILT C POT CALIB KO 远程诊断与测试结果： 打角度后掉电，无法恢复，重启无效。 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 检查角度电位器，检查皮带张进度，检查角度电位器到GMP路相关 建议备件12NC及名称： " 459800441121 DUAL POT ASSEMBLY" SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : PPD DUAL POT ASSEMBLY 459800441121 1 已安排24H发货，FSE罗永平-86-15226131133 Next Step Action : Guo Haibin onsite，已发派工短信。 External Remarks : 17.04.2021 10:09:56 UTC Haibin Guo (86038308)"SUAT. Ready for Clinical Use local time" : "2021-04-17 18:30" External Remarks : 17.04.2021 10:22:18 UTC SUP JCOUSER (86038308)故障描述：设备无法开机；垂直编码器不匹配；重新做倾斜校正，调整编码器 服务内容：重新做倾斜校正，调整编码器 服务结果：服务 完成后，设备恢复正常使用 其他事项：无 Internal Remarks : 20.04.2021 05:28:50 UTC Haibin Guo (86038308) RDI Accuracy: YES Q7:Can not confirmed report problem，infact problem was system cannot power up $Q8:perform visiual check on site and found S TILT C POT1 PO2 KO, S TILT C POT CALIB KO and conform the tilt pot need be calibrated or replaced | 2 |  | After adjust the dual pot and belt as well as tilt calibration, system works well, check and update PM accordingly | [12 empty] | 117400375 | Closed | 818618 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | ECR-077012 | [5 empty] | Engineering | [6 empty] | Dual Pot belt | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Belt | Gantry-HW-Belt | Gantry-HW-Belt | ECR-077012 | Belt | Gantry-HW | Gantry-HW-Belt | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
768 |  |  | China | XIFENG CHINESE MEDICINE HOSPIT | 345014 | 4430500% |  | 02n/Ingenuity Core128/The Gantry cann't powerup | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 02n/Ingenuity Core128/Elite China SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support T2 Activities:Problem Description: Summary of customer's problems: There is an E-stop closing prompt box. After twisting the key, there is no relay suction sound. It's a technology problem with the equipment itself: Yes. Troubleshooting Action: remote diagnosis and testing result: Suspect power supply, doctor's operation issue PRS TR status: it's not through, ask engineer for maintenance on site. Repair Action: Suggest onsite service procedures: Suggest engineer to check power supply on site, and how the doctor can operate it.。。 Safety Questions:Case: 0117402506 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Weng Jinming Last updated by: Weng Jinming N/A External Remarks : 20.04.2021 04:02:24 UTC Guanglin Tang (86039448)"SUAT. Ready for Clinical Use local time" : "2021-04-19 11:45" External Remarks: 20.04.2021 05:14:11 UTC SUP JCOUSER (86039448) error description: prompted ESTOP close. After turned the key there is no relay suction sound. Error analysis: After turning key, HOST and GHOST failed to connect.Guided the customer to reboot the software and gantry, and the connection succeeded. Failure result: boot up success. Internal Remarks : 20.04.2021 05:14:14 UTC Guanglin Tang (86039448) "RDI Accuracy:YSE" $Q7: (Confirmed reported problem,The gantry will not boot.) $Q8: (Troubleshooting,HOST and GHOST did not connect successfully, The connection was successful after the reboot.) $Q10: (Test and Inspection data:Refer to PA.Device status: Device return to functionality.) $Q11: Calibrated Tools name: na. For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date:NA ^Current Tube usage: NA ----------------------------------------------------------------------------- | 2 |  | Check if have CAN trace of issue, CAN trace received, send to Jerry review | [12 empty] | 117402506 | Closed | 818116 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-NoEnoughInfo | [5 empty] | Engineering | [6 empty] | System Reboot | [5 empty] | Refresh |  |  | System issue | System | Gantry-SW | TBD-NoEnoughInfo | Gantry-SW-TBD-NoEnoughInfo | Gantry-SW-TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo | Gantry-SW | Gantry-SW-TBD-NoEnoughInfo | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
769 |  |  | China | SX Datong Yanggao People Hosp. | 345025 | 4430400% |  | 02N/Ingenuity Core128/Elite China/fault | BU Q&R Feedback: No SAP Customer Symptom Code: HANG Customer Symptom Code: Hang/Crash Reason: Corrective Maintenance Engineer Failure Code: Configuration Error Repair Activity Code: Reboot/Restart Problem Reported by Customer: fault Safety Questions:Case: 0117402095 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Mengen Xie Last updated by: Mengen Xie 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: The printer face cannot display the added image Troubleshooting Action: reboot host Repair Action: reboot host, Look again Test/Inspection: customer confirm can scan patient and image ok.functional Check PASSED. The device returned to full functionality. Test Equipment: N/A SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | The printer face cannot display the added image, reboot host, Look again, customer confirm can scan patient and image ok.functional Check PASSED. The device returned to full functionality. Intermitent issue, no bugreport saved | [12 empty] | 117402095 | Fixed | 817735 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-NotEnoughInformation | [5 empty] | Others | [6 empty] | Print | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
770 |  |  | France | SCP GROUPE D IMAGERIE MEDICALE | 40019 | 4431200% |  | Ordered part: | Work Order #: WO-03382675 Aurelien ZAJAC The system partly meets specification and is returned to use with restrictions Regulatory Questions : N/A Problem Description by Engineer : Customer Function/Role: user Device Use: yes Expected and Unexpected Results: n/a User Impact: no Patient Impact: no Current Software Version: Diagnostic performed by Engineer: part ordered. Resolution: Spare part requested for replacement. Follow Up Required: Order: CT Box: 459801705022 Qty: 1 Part to be delivered to the scanner console to the attention of JL Codron. ----------------------------------------------------------------------------- Work Order #: WO-03382693 ----------------------------------------------------------------------------- | 4 |  | CT Box request for replacement. Collect picture or video to indicate button issue? | [12 empty] | 117442322 | In Progress | 839666 |  |  | Closed - Non-Complaint | [14 empty] | CT Box | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
771 |  |  | France | SCP GROUPE D IMAGERIE MEDICALE | 40019 | 4430800% |  | Button is out of order. | Work Order #: WO-03364005 Aurelien ZAJAC The system fully meets specification and returned to use Regulatory Questions : N/A Problem Description by Engineer : Customer Function/Role: user Device Use: yes Expected and Unexpected Results: n/a User Impact: no Patient Impact: no Current Software Version: 4.1.7 Diagnostic performed by Engineer: The customer reports that the bed lowering button is malfunctioning. The CT box must be replaced. Resolution: Service on site required. Follow Up Required: Order: CT Box: 459800484611 Qty: 1. Part to be delivered to the scanner console to the attention of JL Codron. ----------------------------------------------------------------------------- | 3 |  | CT Box request for replacement. Collect picture or video to indicate button issue? | [12 empty] | 117425041 | In Progress | 828627 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632185 | [5 empty] | Engineering | [6 empty] | CT Box | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | CT Box | Gantry-HW-CT Box | Gantry-HW-CT Box | CTUDT00632185 | CT Box | Gantry-HW | Gantry-HW-CT Box | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
772 |  |  | Japan | Maki Hospital | 341017 | 4430900% |  | アーチファクト | In the image of the phantom, the central part is black. In the shooting data of the fruit, you can see a small ring-shaped artifact in the center. Regarding fruit artifacts, the person in charge of the app has reported to the customer that it is conspicuous by reducing the FOV. '================================== --Repair Action-- --Subject-- 04CS: TM: Hoshi-sama: An artifact has occurred. The central part is blackened by the phantom. Fine ring artifacts in the center of the fruit photography data >> Inspection continues. Scheduled for inspection at a later date | 4 |  | Regarding fruit artifacts, the person in charge of the app has reported to the customer that it is conspicuous by reducing the FOV. | [12 empty] | 117431484 | Fixed | 835976 |  | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Mis-use | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
773 |  |  | Japan | Higashi-Ohmi General Medical C | 341015 | 4430600% |  | Application inquiry >> Koizumi will respond | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer: Application inquiry >> Koizumi will respond ----------------------------------------------------------------------------- | 5 |  | Application inquiry >> Koizumi will respond | [12 empty] | 117410792 | In Process | 823572 |  | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Application | [6 empty] | Q&A | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
774 |  |  | Japan | Tomakomai City Hospital | 341014 | 4430500% |  | JY: ECG TESTER failure, replacement | BU Q&R Feedback: No Reason: Corrective Maintenance Safety Questions:Case: 0117402641 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: YAN Joanna Last updated by: YAN Joanna 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Problem Reported by Customer: JY: ECG tester failure, replacement SAP Transfer: Case Transferred to SAP OneEMS internal remarks: Hashimoto DB SAP Transfer: Dispatched for Onsite Support External Remarks:22.04.2021 05:52:37 UTC HODAKA HASHIMOTO (81100553)[Details of request] ECG tester problem. Problem confirmation: I confirmed it. Work details: I delivered an ECG simulator. Other problems/Measures taken: None. Equipment status The equipment can be used. Work status Work finished. 22-Apr-2021 81100553 FSE Handover Tool1.3.4.32612.DB13 20200715 Internal Remarks:22.04.2021 05:52:40 UTC HODAKA HASHIMOTO (81100553) Replacement parts: Replacement parts: As below 12Nc:453566095882 Name:ECG TESTER QTY:1 [Report on mistakes in parts order] No mistakes. Situation monitoring information: Not monitoring the situation TS Request: No parts to return. Additional CT information: Q1. How did the examination finish? Q1. No (not applicable) Q2. Were there any mistaken findings (misdiagnosis, mistaken treatment) due to the problem? None. Q3. Were there any artifacts? None. Q4. Was there any unintended unnecessary exposure? None. Repair sticker: I affixed a Repaired sticker. [Other information to share] 1. Was treatment/quarantine carried out for a patient who has or is suspected of having new coronavirus? Yes 2. Is it necessary to go into treatment rooms/beds where there are infected patients?(Including workplace and routes of movement) No Did you agree on how to complete the report with the person in charge of devices? YES Contacted the customer regarding the status the next day. 22-Apr-2021 81100553 FSE Handover Tool1.3.4.32612.DB13 20200715 ----------------------------------------------------------------------------- | 5 |  | ECG tester failure, replacement, order 12Nc:453566095882 Name:ECG TESTER QTY:1 Check why need order this part? | [12 empty] | 117402641 | New | 818199 |  | Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | ECG | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
775 |  |  | Japan | Nikko Memorial Hospital | 341011 | 4430500% |  | JY: ECG TESTER failure, replacement | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: JY: ECG tester failure, replacement Safety Questions:Case: 0117402620 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: YAN Joanna Last updated by: YAN Joanna 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A OneEMS internal remarks: Hashimoto DB SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Support External Remarks:22.04.2021 05:51:26 UTC HODAKA HASHIMOTO (81100553)[Details of request] ECG tester problem. Problem confirmation: I confirmed it. Work details: I delivered an ECG simulator. Other problems/Measures taken: None. Equipment status The equipment can be used. Work status Work finished. 22-Apr-2021 81100553 FSE Handover Tool1.3.4.32612.DB13 20200715 Internal Remarks:22.04.2021 05:51:29 UTC HODAKA HASHIMOTO (81100553) Replacement parts: Replacement parts: As below 12Nc:453566095882 Name:ECG TESTER QTY:1 [Report on mistakes in parts order] No mistakes. Situation monitoring information: Not monitoring the situation TS Request: No parts to return. Additional CT information: Q1. How did the examination finish? Q1. No (not applicable) Q2. Were there any mistaken findings (misdiagnosis, mistaken treatment) due to the problem? None. Q3. Were there any artifacts? None. Q4. Was there any unintended unnecessary exposure? None. Repair sticker: I affixed a Repaired sticker. [Other information to share] 1. Was treatment/quarantine carried out for a patient who has or is suspected of having new coronavirus? Yes 2. Is it necessary to go into treatment rooms/beds where there are infected patients?(Including workplace and routes of movement) No Did you agree on how to complete the report with the person in charge of devices? YES Contacted the customer regarding the status the next day. 22-Apr-2021 81100553 FSE Handover Tool1.3.4.32612.DB13 20200715 ----------------------------------------------------------------------------- | 5 |  | ECG tester failure, replacement, order 12Nc:453566095882 Name:ECG TESTER QTY:1 | [12 empty] | 117402620 | New | 818184 |  | Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | ECG | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
776 |  |  | Sweden | Ludvika Lasarett | 40049 | 4431200% |  | Saturday at 4.00 PM they had issues with initialize gantry. Doesn´t work this morning. Bug rep made. | Work Order #: WO-03381271 Knudsen Martin The system does not meet specification and is out of use Problem Description by Engineer : Gets gantry error when using system Diagnostic performed by Engineer : Logfiles show Heater error and Ref Converter error. Logged on to system and did Diagnostic's, indicating RCOM or DMC error. Will order RCOM and DMC for replacement. Follow Up Required : Heater issue and Ref conv issue. Alle related to RCOM/DMC. VD shows these errors. Begin with replacing RCOM. Please schedule FSE and order parts. -MK | 2 |  | Gets gantry error when using system, schedule FSE and order parts, collect bugreport for analysis | [12 empty] | 117440693 | In Progress | 839559 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-Bugrep | [5 empty] | Engineering | [6 empty] | RCOM/DMC | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | TBD-Bugrep | Gantry-HW-TBD-Bugrep | Gantry-HW-TBD-Bugrep | TBD-Bugrep | TBD_Bugreport | Gantry-HW | Gantry-HW-TBD-Bugrep | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
777 |  |  | United States | Central Vermont Medical Center | 40024 | 4431200% |  | Gating is not working | Work Order #: WO-03383304 Miller, John M The system partly meets specification and is returned to use with restrictions Diagnostic performed by Engineer : Collect information from customer. Problem Description by Engineer : Problem description by engineer : Customer states the Gating is not connecting. When was the first occurrence? Within... : 26-Apr-2021 How often is issue occurring? : Continuously Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Gating not connecting Patient Impact/Outcome: Can't have Gating Scan. User Impact: Can't use gating Current Software Version: Not provided Internal Comments : Error # and/or description of error : Gating not connecting Completed or Recommended Repair Action "Internal" : Customer states they have tried power cycling and checking all connections. Suggest checking the cabling from the EKG monitor to the gantry. Verify the connections at the Usb hub and test the Ibox. 459801413791 New EXTenders USB, released and available!!! Material Number 459801731761 - USB EXTENDER CABLE FROM IBOX TO CONSOLE Material Number 459800166652 - I-Box-2 Base board assy Material Number 459801787001 - Interconnect cable from GHOST to BELLOW Follow Up Required : Error # and/or description of error : Gating not connecting Completed or Recommended Repair Action "Internal" : Customer states they have tried power cycling and checking all connections. Suggest checking the cabling from the EKG monitor to the gantry. Verify the connections at the Usb hub and test the Ibox. 459801413791 New EXTenders USB, released and available!!! Material Number 459801731761 - USB EXTENDER CABLE FROM IBOX TO CONSOLE Material Number 459800166652 - I-Box-2 Base board assy ----------------------------------------------------------------------------- Work Order #: WO-03383421 Daigle Jim Problem Description by Engineer : Problem description by engineer : Customer states the Gating is not connecting. When was the first occurrence? Within... : 26-Apr-2021 How often is issue occurring? : Continuously Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Gating not connecting Patient Impact/Outcome: Can't have Gating Scan. User Impact: Can't use gating Current Software Version: Not provided Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 3 |  | Replaced iBox box cable and or30 m cable, Material Number 459801731761 - USB EXTENDER CABLE FROM IBOX TO CONSOLE Material Number 459800166652 - I-Box-2 Base board check which part is finally fix the issue? | [12 empty] | 117442880 | In Progress | 839886 |  | Investigation | Closed - Done |  | NotEngineering-CabelReplace | [5 empty] | Others | [6 empty] | Ibox and USB Cable | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
778 |  |  | United States | Lackey Memorial Hospital | 40015 | 4430900% |  | The FSE is on site and the table was making a noise. Was not receiving an error message or anything. | Work Order #: WO-03376194 25-Apr-2021 Denison Chuck The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: X-Ray Supervisor Device Use: Philips C.T. scanner Expected and Unexpected Results: Adjustments to the table top assembly. User Impact: None. system up. Patient Impact: None System up. Current Software Version:4.1.10 Diagnostic performed by Engineer : Took Table top apart. Made repairs to the Cover support and made adjustments to correct the table performance. | 5 |  | Took Table top apart. Made repairs to the Cover support and made adjustments to correct the table performance. | [12 empty] | 117435989 | In Progress | 837830 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | NotEngineering-Assembly | [5 empty] | Others | [6 empty] | Couch Noise | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
779 |  |  | United States | McLaren Macomb | 40021 | 4430900% |  | 1002CDN81 constancy phantom was dropped and broken | Work Order #: WO-03374555 Sheffield, Scott The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: biomed Device Use: not is use Expected and Unexpected Results: phantom was dropped and broken User Impact: can't do qa Patient Impact: none Current Software Version: Diagnostic performed by Engineer : Customer dropped phantom and broke it. Customer gave po # 4500017862 for the phantom. Follow Up Required : Customer dropped phantom and broke it. Customer gave po # 4500017862 for the phantom. 453567135965 - EXTENDED PHANTOM PERFORMANCE KIT without the low contrast pins, used in BigBore v4.2 or above, Ingenuity/ Ingenuity Core/Ingenuity Core128,Ingenuity Flex; for others, please order 453567135962; must order 1pcs 459801778031 for phantom replacement; order 453566200321 to replenish wat ----------------------------------------------------------------------------- Work Order #: WO-03375192 Mona Tony Problem Description by Engineer : Customer Function/Role: biomed Device Use: not is use Expected and Unexpected Results: phantom was dropped and broken User Impact: can't do qa Patient Impact: none Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 5 |  | Is there any injure on operator? Is Body section dropped? Any picture or details could share? | [12 empty] | 117434793 | In Progress | 834467 |  | Investigation | Closed - Done |  | CTUDT00632803 | [5 empty] | Engineering | [6 empty] | Phantom drop | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | Phantom | Couch-HW-Phantom | Couch-HW-Phantom | CTUDT00632803 | Phantom | Couch-HW | Couch-HW-Phantom | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | NotEngineering-CTUDT00632803 | y | y
780 |  |  | United States | The Regional Medical Center | 40005 | 4430700% |  | Customer feels that the post-threshold delay is too long on PE exams | Enhancement request Customer feels that the post-threshold delay is too long on PE exams. They are bolus tracking at the level of the carina (mid-chest). The couch has to move to the top of the chest to begin the scan. The post-threshold delay is usually around 7 seconds. This does include the breathing instructions. They feel that this time is too long when trying to scan while there is a good contrast bolus in the pulmonary arteries. They commented that on other vendor's scanners that they have worked on, the table moves much faster into scan position after performing the bolus tracking scan. They would like to see the couch move to the beginning of the scan faster. Even if the breath-hold instructions are removed, the post-threshold delay is still approximately 6 seconds. The expectation is that the post-threshold delay should be no more than 2-3 seconds. It was in clinical use when they encountered the concern. Procedure was continued. It happens every time a PE exam is performed. They feel they get poor contrast enhancement on some PE exams due to this long delay. No actions taken. There is no fix for the issue, the system works as designed. | CFM |  | Customer feels that the post-threshold delay is too long on PE exams, share to Silva and Lilian to review | [14 empty] | 833710 |  | Investigation | Closed - Done | [7 empty] | Application | [6 empty] | RFC_post-threshold delay | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
781 |  |  | United States | The Regional Medical Center | 40005 | 4430700% |  | Customer would like the ability to change the result direction during a basic interventional procedure. | Enhancement request Customer would like the ability to change the result direction during a basic interventional procedure. When the patient is on the head head first during an interventional procedure when they use the arrow keys or the wheel to scroll through the images, as they press the up arrow key or roll the mouse wheel upwards, it scroll towards the feet. Arrow down or rolling the wheel down goes towards the head. This is confusing to the operator and physician during the procedure. They would like the ability to reverse the result direction so when using the mouse or wheel in an upward fashion makes the images scroll towards the head, and scrolling down will have the images scroll towards the feet. The Radiologist, Dr. Forte said any other scanner he has worked on has allowed for the ability to reverse the results. It was in clinical use when they encountered the concern. Procedure was continued. It happens every time the patient is positioned head first. Customer states they feel the fact that the images scroll in the opposite direction, they could mistake which way the needle needs to be manipulated (superior vs inferior). No actions taken. No fix for the issue, the system works as designed. | CFM |  | Customer would like the ability to change the result direction during a basic interventional procedure. share to Silva and Lilian to review | [14 empty] | 833690 |  | Investigation | Closed - Done | [7 empty] | Application | [6 empty] | RFC_interventional | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
782 |  |  | United States | Seton Medical Center Harker He | 40006 | 4430800% |  | Scan Stopped During a Calcium Score | Work Order #: WO-03369736 Paine Matthew The system partly meets specification and is returned to use with restrictions Follow Up Required : Gantry Acq System Can't Comply without Resends. Error S_MASTER_SEQUENCER_ACQ_NOT_OK. CIRS Error- SafetyCheckThread ACQ board reported error. Recommend Running ACQ Diagnostics. Check Health of RAID in CIRS. Run DataPath Diagnostics. Customer states issue only happens during Cardiac Studies. Diagnostic performed by Engineer : Customer states that they recently purchased the Cardiac Package. Customer states that during a Cardiac Calcium Score, they receive a Gantry ACQ Can't Comply Message. Customer states a Bug Report was made for Incident and Authorized Remote Access through PRS. Logs indicate Issue with ACQ Cancelling the Scan. Problem Description by Engineer : Customer Function/Role: Technologist Device Use: Diagnostic Expected and Unexpected Results: Gantry ACQ Can't Comply During Calcium Score User Impact: Unknown Patient Impact: Unknown Current Software Version:4.1.10 Internal Comments : Contacted Primary FSE Tom Beato. Recommend checking the Raid of CIRS, Run ACQ Diagnostics, Datapath Diagnostics. Try to Duplicate Issue if possible by performing Cardiac Scan's. No Resends seen. ----------------------------------------------------------------------------- Work Order #: WO-03370392 23-Apr-2021 Beato Tom The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Technologist Device Use: Diagnostic Expected and Unexpected Results: Gantry ACQ Can't Comply During Calcium Score User Impact: Unknown Patient Impact: Unknown Current Software Version:4.1.10 Diagnostic performed by Engineer : Checked the view logs and bug report. Verified one ACQ not ok error. Performed DMS diagnostics and data path diagnostics tests. All results passed. Will order loopback test kit to perform Acquisitor diagnostics. Follow Up Required : Perform ACQ diagnostics | 4 |  | Checked the view logs and bug report. Verified one ACQ not ok error. Performed DMS diagnostics and data path diagnostics tests. All results passed. Will order loopback test kit to perform Acquisitor diagnostics. Ask to share bugreport for analysis. To team review May cause rescan on surview!! | [12 empty] | 117430254 |  | 830714 |  | Closed - Done | Closed - Done |  | TBD_Bugreport | [5 empty] | Engineering | [6 empty] | Scan stop | [5 empty] | Refresh |  |  | System issue | System | TBD-Bugreport | TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | TBD_Bugreport | TBD-Bugreport | TBD-Bugreport-TBD_Bugreport | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
783 |  |  | United States | Unity Health Systems | 40037 | 4430700% |  | v 4.1.10.2521 Tube heat on everything | Work Order #: WO-03358780 Hirchak, Michele The system partly meets specification and is returned to use with restrictions Problem description by Clinical Support Specialist : Customer Function/Role: Tech Device Use: In Use Expected and Unexpected Results: Tube heat on all scans. User Impact: None Patient Impact: None Current Software Version: 4.1.10 No rescan, no loss of data, no reinjection. Diagnostic performed by Engineer : Customer stated that on most scans they get tube heat messages. It tells them they have to wait up to 30 minutes to be able to scan. Techs have decreased technique and still get the message. I went over protocols with her, they all seem fine. Also, she said when they try to use the select Current Patient option, the screen goes blank and they have to log out and log back in in order to get it to work properly. I have sent this to an RSE to evaluate. This is not a protocol issue. MBH ----------------------------------------------------------------------------- Work Order #: WO-03359133 Mark Finnigan The system does not meet specification and is out of use Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: Issues with tube heat User Impact: Unknown. Patient Impact: Current Software Version: Diagnostic performed by Engineer : The case was started by Michelle Hirchak in clinical applications. The customer is seeing excessive tube heat warnings and also complained about a issue with the software . When the customer selects the "Use Current" for the patient selection, the instruction does not execute and the application goes back to the original step. Brought NSS Rich Hladun into the case. He downloaded bugreports and looked at some parameters. He would like the case dispatched and have a FSE onsite to work with. Dispatching the case. Follow Up Required : The case was started by Michelle Hirchak in clinical applications. The customer is seeing excessive tube heat warnings and also complained about a issue with the software . When the customer selects the "Use Current" for the patient selection, the instruction does not execute and the application goes back to the original step. Brought NSS Rich Hladun into the case. He downloaded bugreports and looked at some parameters. He would like the case dispatched and have a FSE onsite to work with. Dispatching the case. ----------------------------------------------------------------------------- Work Order #: WO-03360517 Taylor, Durrance Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: Issues with tube heat User Impact: Unknown. Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- Work Order #: WO-03367700 22-Apr-2021 Lockwood Thomas The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : System resistricted when awaiting tube heat to countdown. CT Unit is investigating ----------------------------------------------------------------------------- | 3 |  | Tube heat and software stuck issue, yellow scrren? needs bugreport for analysis. | [12 empty] | 117420104 |  | 826206 |  | Evaluation | Closed - No Further Investigation |  | TBD_Tube hot_bugreport | [5 empty] | Engineering | [6 empty] | Tube Overheat | [5 empty] | Refresh |  |  | System issue | System | Gantry-SW | TubHot | Gantry-SW-TubHot | Gantry-SW-TubHot | TBD_Tube hot_bugreport | TubHot | Gantry-SW | Gantry-SW-TubHot | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
784 |  |  | United States | Advent Health Wesley Chapel | 300106 | 4430800% |  | 0706CDY73 Attempting to print cd on scanner but saying its locked due to scanning in progress but not scanning... | Work Order #: WO-03367666 Sheffield, Scott The system partly meets specification and is returned to use with restrictions Internal Comments : Error # and/or description of error : CD is locked because scan is in progress Completed or Recommended Repair Action "Internal" : Check CDR configuration. Diagnostic performed by Engineer : I had customer reboot host computer. Still got same problem after reboot. Applications sent me an alternate method to burn a cd and that had the same result. I called FSE he is going to check the cd configuration. Problem Description by Engineer : Problem description by engineer : When trying to burn Cd getting message saying: CD is locked because scan is in progress Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No further information is available as per the customer Expected and Unexpected results: When trying to burn Cd getting message saying; CD is locked because scan is in progress Patient Impact/Outcome: No further information is available as per the customer User Impact: delay in burning a cd Current Software Version: 4.1.7.25030 Follow Up Required : Error # and/or description of error : CD is locked because scan is in progress Completed or Recommended Repair Action "Internal" : Check CDR configuration. ----------------------------------------------------------------------------- Work Order #: WO-03368698 22-Apr-2021 Sher Mark The system fully meets specification and returned to use Problem Description by Engineer : Problem description by engineer : When trying to burn Cd getting message saying: CD is locked because scan is in progress Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No further information is available as per the customer Expected and Unexpected results: When trying to burn Cd getting message saying; CD is locked because scan is in progress Patient Impact/Outcome: No further information is available as per the customer User Impact: delay in burning a cd Current Software Version: 4.1.7.25030 Diagnostic performed by Engineer : changed cd burner drive letter to G from e and tested system.. issue resolved ----------------------------------------------------------------------------- | 3 |  | changed cd burner drive letter to G from e and tested system.. issue resolved | [12 empty] | 117428491 |  | 830419 |  | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | CD burner | [5 empty] | Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
785 |  |  | United States | Advent Health Wesley Chapel | 300106 | 4430700% |  | Q/O: images numbers not the image number expected | Work Order #: WO-03359751 22-Apr-2021 Sher Mark The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : king with applications customer corrected converted protocol form 4.1.7 upgrade to windows 10 .. slice thickness on step and shoot were wrong, not converted | 3 |  | Step and Shot, need align with Caroline to know more detail 10mm -> 15mm bring biopy issue and pedal scan no image. | [12 empty] | 117420814 |  | 830143 |  | Investigation | Investigation |  | TBD-CTDUT | [5 empty] | Engineering | [6 empty] | Step and Shot | [5 empty] | Refresh |  |  | Upgrade issue | System | NotEngineering | ShouldbeClinical | NotEngineering-ShouldbeClinical | NotEngineering-ShouldbeClinical | NotEngineering-ShouldbeClinical | NotEngineering-ShouldbeClinical | NotEngineering | NotEngineering-ShouldbeClinical | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
786 |  |  | United States |  | 300106 | 4430500% |  | Q/O Need bore mylar replaced and phantom | Work Order #: WO-03342940 20-Apr-2021 Sher Mark The system fully meets specification and returned to use Parts Consumed : 1.00000,459800048742,RING-LEXAN GN5 1.00000,455012004391,CALIBRATION PHANTOM 1.00000,453567122551,SURVIEW GAUGE 1.00000,453567122541,LASER ALIGNMENT GAUGE Diagnostic performed by Engineer : received parts for Pm, checked on system Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: ----------------------------------------------------------------------------- | 3 |  | 1.00000,459800048742,RING-LEXAN GN5 1.00000,455012004391,CALIBRATION PHANTOM 1.00000,453567122551,SURVIEW GAUGE 1.00000,453567122541,LASER ALIGNMENT GAUGE | [12 empty] | 117405409 |  | 819987 |  |  | Closed - Non-Complaint | [14 empty] | Mylar/Phantom | [5 empty] | Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
787 |  |  | United States | Upson Regional Medical Center | 333009 | 4430100% |  | System jumping to another result during the planning step on exam card utilizing bolus tracking. This results in breakin |  | CFM |  | Mouse Jump | [14 empty] | 816865 |  | Investigation | Investigation |  | CTUDT00632986 | [5 empty] | Engineering | [6 empty] | Mouse Jump | [5 empty] | Refresh |  |  | Upgrade issue | System | Console-SW | ResultSelectionJump | Console-SW-ResultSelectionJump | Console-SW-ResultSelectionJump | CTUDT00632986 | ResultSelectionJump | Console-SW | Console-SW-ResultSelectionJump | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
788 |  |  | United States | Upson Regional Medical Center | 310115 | 4430000% |  | System jumping to another result during the planning step on exam card utilizing bolus tracking. |  | CFM |  | Mouse Jump | [14 empty] | 816831 |  | Investigation | Investigation |  | CTUDT00632986 | [5 empty] | Engineering | [6 empty] | Mouse Jump | [5 empty] | Refresh |  |  | Upgrade issue | System | Console-SW | ResultSelectionJump | Console-SW-ResultSelectionJump | Console-SW-ResultSelectionJump | CTUDT00632986 | ResultSelectionJump | Console-SW | Console-SW-ResultSelectionJump | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
789 |  |  | United States | Upson Regional Medical Center | 333009 | 4430100% |  | A system lag between mouse clicks during the scan plan setup with an exam card that has more than about 10 results. This |  | CFM |  | Performace | [14 empty] | 816806 |  | Investigation | Investigation |  | TBD-Performance | [5 empty] | Engineering | [6 empty] | Performace | [5 empty] | Refresh |  |  | Upgrade issue | System | Console-SW | TBD-Performance | Console-SW-TBD-Performance | Console-SW-TBD-Performance | TBD-Performance | TBD-Performance | Console-SW | Console-SW-TBD-Performance | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
790 |  |  | United States | Upson Regional Medical Center | 310115 | 4430100% |  | A system lag between mouse clicks during the scan plan setup with an exam card that has more than about 10 results. This |  | CFM |  | Performace | [14 empty] | 816778 |  | Investigation | Investigation |  | CTUDT00632986 | [5 empty] | Engineering | [6 empty] | Performace | [5 empty] | Refresh |  |  | Upgrade issue | System | Console-SW | SerielsJump | Console-SW-SerielsJump | Console-SW-SerielsJump | Duplicate | ResultSelectionJump | Console-SW | Console-SW-SerielsJump | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
791 |  |  | United States |  | 310034 | 4430800% |  | NEMO | Work Order #: WO-03369989 ----------------------------------------------------------------------------- | 3 |  | Non-Refresh | [12 empty] | 117430454 |  | 835227 |  | #N/A | Evaluation | [14 empty] | Non-Refresh | [5 empty] | Non-Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | y
792 |  |  | United States |  | 310034 | 4430700% |  | nemo | Work Order #: WO-03361780 ----------------------------------------------------------------------------- | 3 |  | Non-Refresh | [12 empty] | 117422240 |  | 831122 |  | #N/A | Evaluation | [14 empty] | Non-Refresh | [5 empty] | Non-Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | y
793 |  |  | United States |  | 310034 | 4430700% |  | nemo | Work Order #: WO-03361745 Parts Consumed : 1.00000,459800484511,CONTROL PANEL ASSEMBLY-FRONT RH ----------------------------------------------------------------------------- | 3 |  | Non-Refresh | [12 empty] | 117422216 |  | 831104 |  |  | Closed - Non-Complaint | [14 empty] | Non-Refresh | [5 empty] | Non-Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | y
794 |  |  | United States |  | 310167 | 4430700% |  | Failing to send images | Work Order #: WO-03361513 Morrow, Dan The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Problem description by engineer : Failing to send Images. Customer states that IT checked their end and issue with scanner. When was the first occurrence? Within... : 21-Apr-2021 How often is issue occurring? : First occurrence Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Unable to send images Patient Impact/Outcome: none User Impact: Unable to send images Current Software Version: 4.1.7.25028 Diagnostic performed by Engineer : No PRS connection. Customer checked lan cable and is OK. Dispatching FSE. Follow Up Required : What are the next Troubleshooting steps : Check LAN connection to Host. Verify port has connection. Check NIC card. Error # and/or description of error : No PRS. Completed or Recommended Repair Action "Internal" : No PRS connection. Customer checked lan cable and is OK. Dispatching FSE. Internal Comments : Error # and/or description of error : No PRS. Completed or Recommended Repair Action "Internal" : No PRS connection. Customer checked lan cable and is OK. Dispatching FSE. ----------------------------------------------------------------------------- | 3 |  | Non-Refresh | [12 empty] | 117422081 |  | 826822 |  | #N/A | Investigation | [14 empty] | Non-Refresh | [5 empty] | Non-Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | y
795 |  |  | China | ZJ YUEQING PEOPLE HOSPITAL | 40054 | 4432100% |  | Application/Ingenuity CT | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer : Application FSE Notes : This PO is a Reference Number only Next Step Action : 邮件ASP鸿乐嘉跟进， 短信协调员 曾彩花(MATC/US 东区/南区) 17875512591 放射科 林拉子 13362770096 申请首次培训，装机单53515511派工ASP鸿乐嘉 | 5 |  | NAC_APP training | [12 empty] | 117494304 |  | 872990 |  | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Application | [6 empty] | NAC_APP training | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
796 |  |  | China | XJ Urumqi No1 People Hosp XJ Urumqi Children Hosp | 40058 | 4431500% |  | 02N/Ingenuity CT 728326/stand by | BU Q&R Feedback: No Reason: Customer Relations Problem Reported by Customer: stand by SAP Transfer:Dispatched for On site Support External Remarks: 01.05.2021 04:34:19 UTC Jianjun Sang (86038298) Service details: equipment acceptance. Service outcome: Now the device returned to service fully. Internal Remarks:01.05.2021 04:34:21 UTC Jianjun Sang (86038298) $Q7: (Confirm with reported problem: equipment acceptance.) $Q8: (equipment acceptance, check the equipment.) $Q9: (equipment acceptance.) $Q10: (Test and inspection data refer to: PA. The device returned to full functionality. $Q11: Calibrated Tools name (N/A) S/N (N/A) Next Calibration date (N/A) | 5 |  | NAC_FSE support equipment acceptance test | [12 empty] | 117464487 |  | 855913 |  | #N/A | Closed - non-complaint | [7 empty] | Service | [6 empty] | NAC_FSE support equipment acceptance test | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
797 |  |  | United States | Unity Health Systems | 40037 | 4430700% |  | Excessive Tube Heat Warnings and issue with software |  | CFM |  | Excessive Tube Heat Warnings and issue with software | [14 empty] | 843526 |  | Investigation | Closed - Done |  | TBD_Tube hot_bugreport | [5 empty] | Engineering | [6 empty] | Tube Overheat | [5 empty] | Refresh |  |  | System issue | System | Gantry-SW | TubHot | Gantry-SW-TubHot | Gantry-SW-TubHot | TBD_Tube hot_bugreport | TubHot | Gantry-SW | Gantry-SW-TubHot | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
798 |  |  | Turkey | HITACHI HEALTHCARE TURKEY | 337000 | 4431300% |  | Compensator analysis. | Work Order #: WO-03391860 Soner Saglam The system fully meets specification and returned to use Problem Description by Engineer : Compensator crack check. Resolution : No cracks were found. | 3 |  | Compensator analysis | [12 empty] | 117450887 |  | 848977 |  | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Service | [6 empty] | Compensator analysis | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
799 |  |  | Chile | CMD Clinica Integrales Rancagu | 40008 | 4431300% |  | CL1675-THE GANTRY IS RESTARTING | Work Order #: WO-03394637 Valenzuela Velázquez, Eduardo Ariel The system fully meets specification and returned to use. Problem Description by Engineer : Problem description by engineer : Gantry shuts off Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected Results: No gantry errors. Patient Impact/Outcome: Cancelled examinations. User Impact: Canceled examinations. Current Software Version: 4.1.10 Follow Up Required: Error Number and/or Description: S_MC_DEVICE_HOR_ENCODERS_NOT_SYNCED Completed or Recommended Repair Action "Internal" : Absolute encoder coupler replacement. Internal Comments : Error # and/or description of error : S_MC_DEVICE_HOR_ENCODERS_NOT_SYNCED Completed or Recommended Repair Action "Internal" : Absolute encoder coupler replacement. Diagnostic Performed by Engineer: An error is displayed in the position sensors of the horizontal couch motion. A broken coupling part is found ion the couch. With client's help, the part is repaired. Equipment is left working but the part should be replaced. | 4 |  | Manuf WI update to control the process, need wait ECO # for WI. | [12 empty] | 117453038 |  | 850654 |  | Closed - Done | Closed - Done |  | CTUDT00632179 | [5 empty] | Engineering | [6 empty] | Coupling | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | Motor Coupling | Couch-HW-Motor Coupling | Couch-HW-Motor Coupling | CTUDT00632179 | Motor Coupling | Couch-HW | Couch-HW-Motor Coupling | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
800 |  |  | Poland | Agencja Rezerw Materiałowych | 341010 | 4432000% |  | Physicist training | Work Order #: WO-03433693 Robert Kamola Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A | 4 |  | APP training | [12 empty] | 117488965 |  | 868913 |  | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Application | [6 empty] | APP training | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
801 |  |  | Turkey | T.C Saglik Bakanligi Bagcilar -> Bağcılar State Hospital | 40010 | 4431900% |  | CTbox key malfunction. | Work Order #: WO-03427934 Soner Saglam The system fully meets specification and returned to use Problem Description by Engineer : Unable to use the CTbox down key. Resolution : Ctbox replacement. Follow Up Required : Ctbox replacement. | 3 |  | William do investigation | [12 empty] | 117484484 |  | 862117 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632185 | [5 empty] | Engineering | [6 empty] | CT Box | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | CT Box | Gantry-HW-CT Box | Gantry-HW-CT Box | CTUDT00632185 | CT Box | Gantry-HW | Gantry-HW-CT Box | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
802 |  |  | Turkey | T.C Saglik Bakanligi Bagcilar -> Bağcılar State Hospital | 40010 | 4431600% |  | Loss of communication errors. | Work Order #: WO-03418930 Soner Saglam The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : e-stop is activated. logger.mdb was analyzed. S_COUCHMGR_HOR_MOTION_STOP_ERROR, S_MANMOT_HORIZONTAL_STOP_ERROR errors were found. Horizantal axis check and calibration are recommended. NOTE: Full gantry recover was advised to the customer. The customer reported that they restarted the gantry without any success. In the examination, it was found that the device gave an error about horizontal movement and opened an estop. The coupling was checked and horizontal encoder calibration was done. It was checked whether there was any congestion in the horizontal motion. 10 trial patients were enrolled. No problems were observed. The system has been delivered in operation. | 3 |  | Check and collect bugreport for anlysis. To ChangXian, if known issue? | [12 empty] | 117477048 |  | 857371 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632187 | [5 empty] | Engineering | [6 empty] | Couch SW issue | [5 empty] | Refresh |  |  | System issue | System | Couch-SW | Estop | Couch-SW-Estop | Couch-SW-Estop | CTUDT00632187 | Estop | Couch-SW | Couch-SW-Estop | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
803 |  |  | Turkey | T.C Saglik Bakanligi Bagcilar -> Bağcılar State Hospital | 40010 | 4431300% |  | Compensator analysis. | Work Order #: WO-03391863 Soner Saglam The system fully meets specification and returned to use Problem Description by Engineer : Compensator crack check. Resolution : No cracks were found. ----------------------------------------------------------------------------- | 3 |  | Compensator analysis | [12 empty] | 117450896 |  | 848980 |  | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Service | [6 empty] | Compensator analysis | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
804 |  |  | United States | St Barnabas Hospital | 40028 | 4431700% |  | 0711CFXXX Transfer queue failing | Work Order #: WO-03424316 Mcginnis, Mike The system fully meets specification and returned to use Problem Description by Engineer : Problem description by engineer : Transfer queue failing When was the first occurrence? Within... : 01-May-2021 How often is issue occurring? : First occurrence Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Transfer queue failing Patient Impact/Outcome: none User Impact: had to place service call Current Software Version: 4.1.10 Diagnostic performed by Engineer : Iris said they had 5 cases stuck in the transfer Q. I had her delete the transfer of all images and we then went back in and transferred all of the one at a time. I discovered speaking to her that on one of those patients they had an accession number incorrect and had changed it. Everything passed when sending the corrected patient. Internal Comments : Error # and/or description of error : Transfer queue failing Completed or Recommended Repair Action "Internal" : Removed failed Q and sent all exams to the pacs with no issues. Resolution : Removed failed Q and sent all exams to the pacs with no issues. ----------------------------------------------------------------------------- | 3 |  | Transfer queue failing | [12 empty] | 117481437 |  | 860994 |  | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Transfer queue failing | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
805 |  |  | United States | St Barnabas Hospital | 40028 | 4431400% |  | Table horizontal not working. | Work Order #: WO-03404474 Mark Finnigan The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: Customer experienced an issue with table horizontal motion. User Impact: Unknown Patient Impact: Current Software Version: Diagnostic performed by Engineer : Philips CT applications support Susan Shick was doing onsite applications support when she encountered an issue where table horizontal motion was not working. She called for support. The emote connection is established but the refresh passwords were not allowing login. When I called the site back Susan had gotten the table working. I recommended that she take the system down to the breaker when she had a chance. I also called the FSE Mohamed Abbas and let him know about the PRS connectivity issue so that he can get that resolved. Internal Comments : No additional calls from Susan Sheck from the site. Left her a VM to call back if she needs additional support | 2 |  | Check and collect bugreport for anlysis. | [12 empty] | 117462864 |  | 850492 |  | Closed - Done | Closed - Done |  | TBD_Bugreport | [5 empty] | Engineering | [6 empty] | Horizontal | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | TBD-bugrep | Couch-HW-TBD-bugrep | Couch-HW-TBD-bugrep | TBD_CouchEstop | TBD_Bugreport | Couch-HW | Couch-HW-TBD-bugrep | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
806 |  |  | China |  | 40050 | 4431600% |  | 02G/Ingenuity CT 728326/add printer dicom node | BU Q&R Feedback: No SAP Customer Symptom Code: CENV Customer Symptom Code: Customer Environmental Concern Reason: Corrective Maintenance Engineer Failure Code: Configuration Error Repair Activity Code: Configuration Change Problem Reported by Customer: 02G/Ingenuity CT 728326 Safety Questions:Case: 0117475950 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Gaoge Last updated by: Wang Gaoge 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action: Q6 = add DICOM printer SAP Transfer:Case Transferred to SAP T2 Activities:Problem Description: need to add printer node Troubleshooting Action: Instruct customer to enter admin account and add printer node Repair Action: Instruct customer to enter admin account and add printer node Test/Inspection: N/A Test Equipment: N/A ----------------------------------------------------------------------------- | Zhai, Lisa |  | NAC | [12 empty] | 117475950 |  | 859811 |  |  | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
807 |  |  | United States |  | 40021 | 4431500% |  | LS - 065137 | Work Order #: WO-03412685 30-Apr-2021 Mona Tony The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : Replaced phantom Parts Consumed : 1.00000,459801415171,HARMONIZED PHANTOM PHYSICS SECTION ----------------------------------------------------------------------------- | E, Sravanthi |  | NAC | [12 empty] | 117470844 |  | 857952 |  |  | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
808 |  |  | United States | Southwest Mississippi Regional | 40055 | 4431400% |  | 1203CDY72 System stops scan during surview | Work Order #: WO-03403330 Weaver David The system partly meets specification and is returned to use with restrictions Follow Up Required : S_GEN_FAULT_CATHODE_TUBE_ARC, S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT Clean high voltage connectors. If this happens with low (or very) high mA and the arcing is in the beginning of scan (Look in szInfo for a short exposure time like GXray EXP=18 (=18 milliseconds)) it could be related to the filament calibration. You might want to run filament calibration followed by long tube reconditioning in such a case. If it happens not in the beginning of scan a long tube reconditioning followed by an hour rest (to degas the tube) might be beneficial. Also look for soft arcs if these are on both anode and cathode side. If it happens all the time run a no load test both with tube connected and tube disconnected. Start with low kV and gradually increase. If there is arcing with tube connected and not when tube is disconnected then the problem is most likely in the tube. Check for burn marks on the HT connectors. Also you may want to listen to arcing sounds coming from one of the power modules (only with no load test!!!). Problem Description by Engineer : DESCRIPTION OF PROBLEM (EXPECTED/UNEXPECTED RESULTS) System stops scan during surview CALLER FUNCTION/ROLE Technician/Sonographer/Therapist DEVICE INTENDED USE Diagnostic WHEN WAS THE FIRST OCCURRENCE? WITHIN . . . <1 Week HOW OFTEN IS ISSUE OCCURRING? Intermittent WHERE IN THE END-USER WORKFLOW IS THE ISSUE OCCURRING? During acquisition/scan WHAT RECENT EVENTS HAVE OCCURRED? During normal use NEGATIVE OPERATOR IMPACT Scan must be restarted NEGATIVE PATIENT IMPACT/PROCEDURE OUTCOME Delay in scanninf HAS CUSTOMER CREATED A SPECIFIC SYSTEM EVENT REPORT OF ISSUE? Yes 4.2 Diagnostic performed by Engineer : System stops scan during surview ----------------------------------------------------------------------------- Work Order #: WO-03403375 30-Apr-2021 Donald Vic The system fully meets specification and returned to use Resolution : Following the recommendations of RSE I ran the long tube conditioning twice and no arcs were present during the long tube conditioning. Diagnostic performed by Engineer : Following the recommendations of RSE I ran the long tube conditioning twice and no arcs were present during the long tube conditioning. Problem Description by Engineer : DESCRIPTION OF PROBLEM (EXPECTED/UNEXPECTED RESULTS) System stops scan during surview CALLER FUNCTION/ROLE Technician/Sonographer/Therapist DEVICE INTENDED USE Diagnostic WHEN WAS THE FIRST OCCURRENCE? WITHIN . . . <1 Week HOW OFTEN IS ISSUE OCCURRING? Intermittent WHERE IN THE END-USER WORKFLOW IS THE ISSUE OCCURRING? During acquisition/scan WHAT RECENT EVENTS HAVE OCCURRED? During normal use NEGATIVE OPERATOR IMPACT Scan must be restarted NEGATIVE PATIENT IMPACT/PROCEDURE OUTCOME Delay in scanninf HAS CUSTOMER CREATED A SPECIFIC SYSTEM EVENT REPORT OF ISSUE? Yes 4.2 ----------------------------------------------------------------------------- | 4 |  | Peter check if have Hazad ? | [12 empty] | 117459746 |  | 849718 |  | Closed - Done | Closed - Done |  | CTUDT00630085 | [5 empty] | Engineering | [6 empty] | Tube Arc | [5 empty] | Refresh |  |  | System issue | System | IC-HW | Tub Arc | IC-HW-Tub Arc | IC-HW-Tub Arc | CTUDT00630085 | Tub Arc | IC-HW | IC-HW-Tub Arc | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
809 |  |  | United States | St Barnabas Hospital | 40028 | 4431300% |  | Machine beeping | Work Order #: WO-03391382 28-Apr-2021 Joseph Casale The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : beep sound from ac unit not system rear gantry controls cables in wrong possition corrected | 4 |  | beep sound from ac unit not system rear gantry controls cables in wrong possition corrected | [12 empty] | 117450493 |  | 848652 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | NotEngineering | [5 empty] | Operation | [6 empty] | Gantry Noise | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
810 |  |  | China |  | 40023 | 4431400% |  | 02N/Ingenuity CT 728326/change password | BU Q&R Feedback: No SAP Customer Symptom Code: HANG Customer Symptom Code: Hang/Crash Reason: Corrective Maintenance Engineer Failure Code: Configuration Error Repair Activity Code: Configuration Change Problem Reported by Customer: 02N/Ingenuity CT 728326 T2 Activities:Problem Description: Password expired Troubleshooting Action: Help customers reset and modify passwords Repair Action: Help customers reset and modify passwords Test/Inspection: customer confirm can scan patient and image ok.functional Check PASSED. The device returned to full functionality. Test Equipment: N/A Safety Questions:Case: 0117455662 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Bo Last updated by: Wang Bo 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | Zhai, Lisa |  | NAC | [12 empty] | 117455662 |  | 851693 |  |  | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
811 |  |  | United States | Odessa Regional Medical Center | 40029 | 4431400% |  | 1102CDN83 Power outage last night. UPS failure. | Work Order #: WO-03400353 Morrow, Dan The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Problem description by engineer : Power outage and UPS did not remain on for console processing. Console turned off with loss of power. When was the first occurrence? Within... : 28-Apr-2021 How often is issue occurring? : First occurrence Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: Console remain on Patient Impact/Outcome: none User Impact: Unable to process images Current Software Version: Not provided Follow Up Required : What are the next Troubleshooting steps : Test UPS. Possible control or battery failure. 459800306382 UPS (Emerson), 230V, 50/60Hz, 3 KVA 459800306392 EXTERNAL BATT PACK, 3KVA UPS, 72VDC Error # and/or description of error : No PRS Completed or Recommended Repair Action "Internal" : Emerson Staco installed. Dispatching FSE to test UPS. Diagnostic performed by Engineer : Emerson Staco installed. Dispatching FSE to test UPS. Internal Comments : Error # and/or description of error : No PRS Completed or Recommended Repair Action "Internal" : Emerson Staco installed. Dispatching FSE to test UPS. ----------------------------------------------------------------------------- | 3 |  | 459800306382 UPS (Emerson), 230V, 50/60Hz, 3 KVA 459800306392 EXTERNAL BATT PACK, 3KVA UPS, 72VDC | [12 empty] | 117459016 |  | 849131 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-NoEnoughInfo | [5 empty] | Engineering | [6 empty] | UPS | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | TBD-NoEnoughInfo | Gantry-HW-TBD-NoEnoughInfo | Gantry-HW-TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo | Gantry-HW | Gantry-HW-TBD-NoEnoughInfo | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
812 |  |  | China | Guangzhou No8 Peoples Hospital | 345032 | 4431200% |  | 02N/Ingenuity Core128/Elite China | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 02N/Ingenuity Core128/Elite China SAP Transfer:Dispatched for On site Support Safety Questions:Case: 0117439674 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Guoping Last updated by: Liu Guoping 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Case Transferred to SAP External Remarks: 29.04.2021 05:27:23 UTC Zhenhua Liu (86034125) Failure description: machine noise Service details: check machine movement, wipe slip ring Service result: After service, the machine is back to normal. Other important issues: None Internal Remarks : 29.04.2021 05:27:29 UTC Zhenhua Liu (86034125) "RDI Accuracy:Yes" $Q1: () $Q2: () $Q3: () $Q4: () $Q7: (Confirmation & Description of problem) Confirmed reported problem, gantry make some noise when rotor run in highest speed, but could not repeat the abnormality. $Q8: (Troubleshooting & Results) check the rotor and case, main belt, but could not find anything strange, but find out the slip ring is quite dirty. $Q9: (Repair steps, clean the slip ring.) $Q10: (Test and inspection data refer to PA . Device returned to full functionality ) $Q11: Calibrated Tools name ( refer to PA ) S/N ( ) Next Calibration date ( refer to PA ) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date: N/A ^Current Tube usage: N/A ----------------------------------------------------------------------------- | 3 |  | Check if the brush is installed correctly | [12 empty] | 117439674 |  | 841678 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | NotEngineering-OPS | [5 empty] | Operation | [6 empty] | Brush/Slipring | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
813 |  |  | United States | McLaren Macomb | 40021 | 4431900% |  | 0899CDY71 IBox Fails Diags | Work Order #: WO-03428412 Carpenter, Kevin The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Customer Function/Role: Biomed Jim Device Use: Diag Expected and Unexpected Results: iBox failed diags, even after gantry power cycle User Impact: None Patient Impact: Can't do gated Pt studies Current Software Version: IDK Diagnostic performed by Engineer : Confirmed that trained biomed had power cycled gantry but that iBox still failed diags. CT RSE colab'ed with pFSE Tony Mona ~ agree to ship iBox Part to his home FunLoc Follow Up Required : Biomed reports iBox failed diags, even after gantry power cycle. pFSE Tony Mona req's iBox PCB part be shipped to his home / funloc ----------------------------------------------------------------------------- | 3 |  | 459800166651,I-BOX-2 BASE BOARD ASSY | [12 empty] | 117484842 |  | 862222 |  | Investigation | Closed - Done |  | TBD-NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Ibox | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | TBD-NoEnoughInfo | Gantry-HW-TBD-NoEnoughInfo | Gantry-HW-TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo | Gantry-HW | Gantry-HW-TBD-NoEnoughInfo | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
814 |  |  | United States | Lackey Memorial Hospital | 40015 | 4431900% |  | Not sending to PACS | Work Order #: WO-03429838 04-May-2021 Denison Chuck The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: X-Ray Supervisor Device Use: Philips C.T. scanner. Expected and Unexpected Results: None was a follow up. User Impact: None C.T. Up. Patient Impact: None C.T. Up. Current Software Version:4.1.10 Diagnostic performed by Engineer : Done follow up on table adjustments. and checked the performance of the FCO that was performed on C.T. scanner. ----------------------------------------------------------------------------- | 5 |  | FCO740? | [12 empty] | 117485900 |  | 866397 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | HF3 | [5 empty] | Engineering | [6 empty] | Couch adjustment | [5 empty] | Refresh |  |  | System issue | System | Console-SW | HF3 | Console-SW-HF3 | Console-SW-HF3 | HF3 | HF3 | Console-SW | Console-SW-HF3 | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
815 |  |  | United States |  | 40015 | 4431400% |  | needs help with a protocol | Work Order #: WO-03400997 Hirchak, Michele The system fully meets specification and returned to use Diagnostic performed by Engineer : Michea had questions on CTA Runoff. Threshold, pre and post injection delay, flow rate, saline flush, feet to knee scan and IV placement were discussed and explored on their Ingenuity CT system. Michea will discuss what other parameters could be changed with the radiologist to produce diagnostic imaging when performing CTA Runoffs on larger patients for approval. Michea had no further questions on this issue. No rescan, no loss of data, no reinjection. Problem description by Clinical Support Specialist : Customer Function/Role: Technologist Device Use: In use Expected and Unexpected Results: Runoff Questions User Impact: none Patient Impact: none Current Software Version: 4.1 ----------------------------------------------------------------------------- | Dantha, Rajashekar |  | NAC | [12 empty] | 117459462 |  | 849489 |  |  | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
816 |  |  | Japan | Kinan Hospital | 340000 | 4431200% |  | 02DS:TM:(via FSE): The patient couch does not become free. >> FSE response. | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02DS:TM:(via FSE): The patient couch does not become free. >> FSE response. Safety Questions:Case: 0117441020 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: TAKEYUKI MIYAKAWA Last updated by: TAKEYUKI MIYAKAWA 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Support SAP Transfer:OBJ: Case_Activity__c - unable to obtain exclusive access to this record T2 Activities:Problem Description: Problem description by engineer : Equipment affected: CT Failure location: 02Couch When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user's work flow did the problem occur?: (During Acquisition/Scan) What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem?: NA Engineer's Comment: '============================== Call description: (via FSE): The patient couch does not become free. '============================== Troubleshooting Action: What steps are taken during the system diagnosis: Result of response: Problem not solved. Failure location: 02Couch What steps did you perform to diagnose the problem? '================================== Response FSE will respond. '================================== Error # and/or description of error: NA Repair Action: Completed or Recommended Repair Action "External": The service engineer for the site visited and worked on the equipment. OneEMS internal remarks : Problem description by engineer : Equipment affected: CT Failure location: 02Couch When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user's work flow did the problem occur?: (During Acquisition/Scan) What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem?: NA Engineer's Comment: '============================== Call description: (via FSE): The patient couch does not become free. '============================== What steps are taken during the system diagnosis: Result of response: Problem not solved. Failure location: 02Couch What steps did you perform to diagnose the problem? '================================== Response FSE will respond. '================================== Error # and/or description of error: NA Completed or Recommended Repair Action "External" : The service engineer for the site visited and worked on the equipment. Completed or Recommended Repair Action "Internal": --Troubleshooting Action-- Result of response: Problem not solved. Failure location: 02Couch What steps did you perform to diagnose the problem? '================================== Response FSE will respond. '================================== --Repair Action-- Recommended 12Nc: PR1 453566494823:CLUTCH - HORIZONTAL DRIVE QTY:1 SLOC: Osaka (U615) --Subject-- 02DS:TM:via FSE: The patient couch does not become free. >> FSE response. --Problem Description-- Equipment affected: CT Failure location: 02Couch When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user's work flow did the problem occur?: (During Acquisition/Scan) What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem?: NA Engineer's Comment: '============================== Call description: (via FSE): The patient couch does not become free. '============================== SAP Message: Components Added Succesfully to SWO!- SAP Message: LOGI Task added Successfully!- External Remarks:29.04.2021 11:39:09 UTC TETSUO OHYAMA (81100250) [Request details] The bed does not become free. Problem confirmation: I confirmed it. Work details: There was a problem with the clutch in the horizontal direction of the bed. I replaced the clutch part. There is no problem with the operation of the equipment. Other problems/Measures taken: None. Equipment status The equipment can be used. Work status Work finished. 26-Apr-2021 81100250 FSE Handover Tool1.3.4.32612.DB13 20201001 Internal Remarks : 29.04.2021 11:39:11 UTC TETSUO OHYAMA (81100250) Replacement parts: Replacement parts: As below 12Nc:453566494823 Description:CLUTCH - HORIZONTAL DRIVE QTY:1 [Report on mistakes in parts order] No mistakes. Situation monitoring information: Not monitoring the situation TS Request: No parts to return. Additional CT information: Q1. How did the examination finish? Q1. No (not applicable) Q2. Were there any mistaken findings (misdiagnosis, mistaken treatment) due to the problem? None. Q3. Were there any artifacts? None. Q4. Was there any unintended unnecessary exposure? None. Repair sticker: I affixed a Repaired sticker. [Other information to share] 1. Was treatment/quarantine carried out for a patient who has or is suspected of having new coronavirus? Yes 2. Is it necessary to go into treatment rooms/beds where there are infected patients?(Including workplace and routes of movement) No Did you agree on how to complete the report with the person in charge of devices? YES On-site the next day 26-Apr-2021 81100250 FSE Handover Tool1.3.4.32612.DB13 20201001 ----------------------------------------------------------------------------- | 3 |  | 453566494823:CLUTCH - HORIZONTAL DRIVE | [12 empty] | 117441020 |  | 842585 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD_Bugreport | [5 empty] | Engineering | [6 empty] | CLUTCH | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | TBD-bugrep | Couch-HW-TBD-bugrep | Couch-HW-TBD-bugrep | TBD_CouchEstop | TBD_Bugreport | Couch-HW | Couch-HW-TBD-bugrep | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
817 |  |  | China | JS Yangzhou No.1 People Hosp. Xiqu Hosp. | 345022 | 4431700% |  | 02N/Ingenuity Core128/Elite China/cannot scan | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: Cannot scan Safety Questions:Case: 0117481105 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingli Last updated by: Lu Mingli 服务内容：删除无用文件，腾空D盘空间. 服务结果： 维修后，设备恢复正常工作. | 3 |  | PACS transfer failed and bring D disk full, after disk clean and implement FCO740, system works well | [12 empty] | 117481105 |  | 861292 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | HF3 | [5 empty] | Engineering | [6 empty] | Disk full | [5 empty] | Refresh |  |  | System issue | System | Console-SW | HF3 | Console-SW-HF3 | Console-SW-HF3 | HF3 | HF3 | Console-SW | Console-SW-HF3 | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
818 |  |  | China | JS No.2 Chinese Medicine Hosp. | 345027 | 4431700% |  | 02N/Ingenuity Core128/Elite China | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 02N/Ingenuity Core128/Elite China SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Safety Questions:Case: 0117480558 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Zhang Jing Last updated by: Zhang Jing 服务内容：重启主机和医院网络系统，调整系统参数. 服务结果： 维修后，设备恢复正常工作 | 3 |  | the worklist lost and system freeze occasionally No error report, the network system of hospital could have issue. The system parameter of host server need to be changed. Rebooted the network system of hospital and changed the system parameter of host server. | [12 empty] | 117480558 |  | 861090 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | NotEngineering-Customer | [5 empty] | Customer | [6 empty] | Worklist | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
819 |  |  | China | JS No.2 Chinese Medicine Hosp. | 345027 | 4431300% |  | Escalation/EQ#88392971 | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: EQ#88392971 Next Step Action: FSE Qi Liangdong -86-18662702811 call to order Safety Questions:Case: 0117447512 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Cheng Zhenhua Last updated by: Cheng Zhenhua On this case the host was not rebooted for more than 45 hours (almost 2 days) please instruct the user to reboot the host on daily bases Suggest to perform workaround. | 4 |  | Workaround performed | [12 empty] | 117447512 |  | 846706 |  | Evaluation | Closed - No Further Investigation |  | CTUDT00631367 | [5 empty] | Engineering | [6 empty] | New Yellow Screen | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
820 |  |  | China | Tangshan | 40044 | 4431300% |  | 02N/Ingenuity CT 728326 | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 02N/Ingenuity CT 728326 Safety Questions:Case: 0117449055 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Guoping Last updated by: Liu Guoping 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 3 |  | Unknown | [12 empty] | 117449055 |  | 847759 |  | Evaluation | Closed - non-complaint | [7 empty] | Others | [6 empty] | Unknown | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
821 |  |  | Japan |  | 341015 | 4431300% |  | CT:APP:NO:Mr. Hisaki: Question. | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Application Support Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer : Q: He wants to know how to create function analysis data A: I had him create it from "Edit Phase". Safety Questions:Case: 0117449279 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Ootani Natsuki Last updated by: Ootani Natsuki 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | Zhu, Chun |  | NAC | [12 empty] | 117449279 |  | 847927 |  |  | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
822 |  |  | Japan |  | 341017 | 4431300% |  | CT:APP:NO:Mr. Noda: Question | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Application Support Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer: Q: There was a mistake in the patient name so he wants to know how to correct it. A: I had him do it with me from Change Patient Detail. Safety Questions:Case: 0117449259 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Ootani Natsuki Last updated by: Ootani Natsuki 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | Zhu, Chun |  | NAC | [12 empty] | 117449259 |  | 847904 |  |  | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
823 |  |  | China | HLJ Harbin Senhai Hospital | 345007 | 4431200% |  | ESTOP Open | Summary of customer's problems: The customer didn't transfer. The engineer said it was intermittently power off Detailed description of Customer's problems or malfunctions: Intermittently power off problem It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: Intermittently power off, no PRS Remote troubleshoot and test result: Go on site to confirm the log. If necessary order spare parts to be replaced $Q7: (Confirmed reported problem: gantry failure.) $Q8: (check log, couch horizontal failure ) $Q9: (do horizontal calibration) $Q10( Test and inspection data refer to : PA . Device returned to full functionality.) $Q11( Calibrated tools name (N/A) External Remarks: 27.04.2021 02:19:40 UTC Jun Jin (86039442) service details: bed horizontal adjustment Service Result: The device returned to normal. External Remarks: 27.04.2021 02:35:21 UTC Yadong Liu (86039473) service details: bed horizontal adjustment Service Result: The device returned to normal. | 3 |  | bed horizontal adjustment | [12 empty] | 117440075 |  | 841942 |  | Investigation | Closed - Done |  | TBD_Bugreport | [5 empty] | Engineering | [6 empty] | Horizontal calibration | [5 empty] | Refresh |  |  | System issue | System | Couch-SW | TBD-bugrep | Couch-SW-TBD-bugrep | Couch-SW-TBD-bugrep | TBD_CouchEstop | TBD_Bugreport | Couch-SW | Couch-SW-TBD-bugrep | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
824 |  |  | China | Huaying Mountain Guangneng General Hospital | 345000 | 4431600% |  | Wechat:02N/Ingenuity Core128/Elite China/Abnormal belt noise | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : Abnormal belt noise Safety Questions:Case: 0117474672 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Cheng Zhenhua Last updated by: Cheng Zhenhua 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Case Transferred to SAP T2 Activities:Problem Description: Summary of customer's problems: The device makes a strange sound. Detailed description of Customer's problems or malfunctions: There is noise at the back of the gantry, like the belt is bad. It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: There is noise at the back of the gantry, like the belt is bad. Remote troubleshoot and test result: Contacted the field service engineer, the belt is suspected. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Analyze the source of the noise. If necessary replace the belt. We recommend spare parts 12NC and names: BELT, 5-RIB POLY-VEE, 198" 453566492401 SAP Transfer:Dispatched for On site Support Next Step Action: PPD 453566492401 *1, have arranged for 24H delivery, Field Service Engineer Li Quanfei - 86-13699457740 External Remarks:05.05.2021 08:59:28 UTC Quanfei Li (86001471) Service details: Check the equipment and confirm rotation drive belt is making noises. After replacing the belt and adjusting the position, the machine is back to normal. Service outcome: Device returns to normal working condition after service Internal Remarks: 05.05.2021 08:59:34 UTC Quanfei Li (86001471) $Q7: Confirmation the report problem. Abnormal belt noise. $Q:8:check the device.found the noise from the belt. Need to replace it. $Q9:Replace the "BELT, 5-RIB POLY-VEE, 198" (12nc453566492401)"; the device returned to full service. $Q10: (Test and inspection data refer to: PA. Device returned to full functionality.) $Q11: Calibrated Tools:refer to N/A ----------------------------------------------------------------------------- | 3 |  | BELT, 5-RIB POLY-VEE, 198" (12nc453566492401) | [12 empty] | 117474672 |  | 866174 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | ECR-077012 | [5 empty] | Engineering | [6 empty] | Belt | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Belt | Gantry-HW-Belt | Gantry-HW-Belt | ECR-077012 | Belt | Gantry-HW | Gantry-HW-Belt | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
825 |  |  | China | Huaying Mountain Guangneng General Hospital | 345000 | 4431400% |  | 02N/Ingenuity Core128/Elite China/Ask for equipment parameters | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ Safety Questions:Case: 0117454467 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Bo Last updated by: Wang Bo 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action: mail ASP Kangmei chengda follow Next Step Action: Deng Shaoyun of our company has replied, please be aware ----------------------------------------------------------------------------- | 3 |  | NAC_Ask for equipment parameters | [12 empty] | 117454467 |  | 851173 |  | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Customer | [6 empty] | NAC_Ask for equipment parameters | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
826 |  |  | Poland | Agencja Rezerw Materiałowych | 341010 | 4431200% |  | Starting the device. | Work Order #: WO-03383399 28-Apr-2021 Mr. Robert Kamola The system fully meets specification and returned to use Problem Description by Engineer: Starting the CT scanner pre-installed in the container. Tools & Equipment : MULTIMETER,2021-11-04 00:00:00 ELECTRICAL SAFETY TESTER,2021-09-14 00:00:00 Resolution: Starting the CT scanner pre-installed in the container. Calibrations and system tests were performed. Image quality tests with a positive result. . Assistance during acceptance tests. . Check of the CT application. Backup upload - system tests and calibrations were performed again. . The device was handed over for operation. ----------------------------------------------------------------------------- Work Order #: WO-03383386 27-Apr-2021 Mr. Mariusz Pachla The system fully meets specification and returned to use Diagnostic performed by Engineer: Starting the device after the container has been transported to the target location. Continuation of service by Robert Kamola. Tools & Equipment : MULTIMETER,2021-06-10 00:00:00 Problem Description by Engineer: Starting the device. ----------------------------------------------------------------------------- | 5 |  | NAC_Calibrations and system tests | [12 empty] | 117442996 |  | 843833 |  | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Others | [6 empty] | NAC_Calibrations and system tests | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
827 |  |  | Germany | Gesundheit Nord gGmbh | 40002 | 4431300% |  | Error message occurred | Work Order #: WO-03390628 Carsten Becker The system fully meets specification and returned to use Problem Description by Engineer: Mr. Glogner-Umbeer (user) He does not have any info regarding the reported error Follow-up required: Setup of SSH access with the correct credentials Setup of the PSA Call-back in the remote in order to test access Download of the indicated bug reports for offline error analysis Diagnostic performed by Engineer: System can indeed be reached by remote but dial-in via SSH (FTP or Telnet) not possible since the known credentials do not apply Error message: "Password authentication failed. Please verify that the username and password are correct." In addition, remote access via M2M is not set up on the system. ----------------------------------------------------------------------------- | 3 |  | COPSSH setting might need to check | [12 empty] | 117449763 |  | 844076 |  | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Service | [6 empty] | PRS Config | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
828 |  |  | China | GD Dongguan Dalingshan Hosp. | 345013 | 4431900% |  | WeChat:02N/Ingenuity Core128/Elite China/gantry is disconnected automatically | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : gantry is disconnected automatically External Remarks: 04.05.2021 09:47:52 UTC Shaoxiang Huang (86034120) Service details: [When the support is tilting, it loses power. Replace the tilt encoder and do the corresponding calibrations.】 Service outcome: The machine is in normal use. Precautions: [None]服务内容： 【支架打角度的时候出现掉电，更换角度倾斜编码器并且做对应的校正。】 | 3 |  | Replace the tilt encoder and do the corresponding calibrations | [12 empty] | 117483552 |  | 864446 |  | Investigation | Closed - Done |  | NotEngineering-partreplacement | [5 empty] | Others | [6 empty] | Tilt | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
829 |  |  | United States | Clay County Medical Center | 40003 | 4432100% |  | Hard to run IQ Check | Work Order #: WO-03440191 05-May-2021 Stauffer Jeffrey The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: General CT exams Expected and Unexpected Results: Expected - the use of the IQ Check would be easy and pass. Unexpected - the IQ Check intermittently fails for no apparent reason and it tells to move the phantoms in inconsistent manners. User Impact: Delays in getting IQ Check to pass. Patient Impact: None Current Software Version:4.1.10 Diagnostic performed by Engineer : Attempted to perform IQ Check as customer does. IQ Check would pass intermittently. Sometimes system software would tell to move the phantom one direction then back the other direction. Follow Up Required : Contact NSS to discuss issue and file Customer Feedback form. | 4 |  | Sometimes system software would tell to move the phantom one direction then back the other direction. | [12 empty] | 117494009 |  | 869677 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00630107 | [5 empty] | Engineering | [6 empty] | Phantom positioning | [5 empty] | Refresh |  |  | System issue | System | DMS-SW | Phantom Detect | DMS-SW-Phantom Detect | DMS-SW-Phantom Detect | CTUDT00630107 | Phantom Detect | DMS-SW | DMS-SW-Phantom Detect | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
830 |  |  | Peru | Hogar Clinica San Juan de Dios | 40034 | 4431200% |  | CT installation in the city of Huancayo- For PO | CT installation in the city of Huancayo- For PO | 5 |  | NAC_CT Installation? | [12 empty] | 117446425 |  | 846218 |  | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Service | [6 empty] | NAC_CT Installation? | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
831 |  |  | China |  | 40059 | 4431200% |  | WeChat:02G/Ingenuity CT 728326/Move the distribution box location | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer: Move the distribution box location Safety Questions:Case: 0117440143 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Cheng Zhenhua Last updated by: Cheng Zhenhua 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action: have requested to add operation to IA#53232184 and dispatch FSE Lin Zhitao Reason: Customer need to cooperate with migration of power supply cabinet Next Step Action: OP added 0030. ----------------------------------------------------------------------------- | Lu, Xueying |  | NAC | [12 empty] | 117440143 |  | 842007 |  |  | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
832 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 341006 | 4431300% |  | Compensator analysis. | Work Order #: WO-03391871 Soner Saglam The system fully meets specification and returned to use Problem Description by Engineer : Compensator crack check. Resolution : No cracks were found. ----------------------------------------------------------------------------- | 3 |  | Compensator analysis | [12 empty] | 117450924 |  | 849003 |  | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Service | [6 empty] | Compensator analysis | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
833 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 341006 | 4431300% |  | Table movement is blocked. | Work Order #: WO-03390062 Soner Saglam The system fully meets specification and returned to use Problem Description by Engineer : E-stop was activated. Resolution : Likely e-stop activation during cleaning. The system is functional. ----------------------------------------------------------------------------- | 3 |  | Likely e-stop activation during cleaning. The system is functional. | [12 empty] | 117449191 |  | 843956 |  | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | E-Stop button | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
834 |  |  | United States |  | 300033 | 4432100% |  | 0799CFXXX BU reports daily CTHost power down restart is overdue and RADAR shows it is generally not done, which may... | Work Order #: WO-03442627 Carpenter, Kevin The system fully meets specification and returned to use Diagnostic performed by Engineer : BU reports daily CTHost power down restart is overdue and RADAR shows it is generally not done, which may cause CT system to not preform optimally. Problem Description by Engineer : Customer Function/Role: BU / RADAR Device Use: Diag Expected and Unexpected Results: Remote ProActive BU Alert User Impact: None Patient Impact: None Current Software Version: IDK Resolution : BU reports daily CTHost power down restart is overdue and RADAR shows it is generally not done, which may cause CT system to not preform optimally. CT RSE will recheck system in a few days ~ if CTHost has still not been restarted then this case will be dispatched to the Field so the local FSE can speak to the customer. ----------------------------------------------------------------------------- | 5 |  | Non-Refresh | [12 empty] | 117496283 |  | 869716 |  | #N/A | Investigation | [14 empty] | Non-Refresh | [5 empty] | Non-Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | y
835 |  |  | United Kingdom |  | 336286 | 4432000% |  | Tube warm up failed - error insure system is not in a working state | Work Order #: WO-03433992 Jeffrey Thompson The system fully meets specification and returned to use Regulatory Questions : N/A Problem Description by Engineer : Problem description by engineer : Spoke to Melanie who confirmed the issue, they tried to do a short tube conditioning and the system showed an error ensure system is in a working state. They then did a QA and this was successful and then could complete the short tube conditioning. Bug report has been completed and advised will check logs. ***** No patient re-scans or re-injections ***** When was the first occurrence? Within... : 04-May-2021 07:05:00 How often is issue occurring? : First occurrence Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Radiographer Device Use: No patient involved Expected and Unexpected results: Attempted short tube conditioning and unable to carry this out Patient Impact/Outcome: No patient re-scans or re-injections User Impact: Unable to complete short tube conditioning Current Software Version: 4.1.7.25026 Diagnostic performed by Engineer : Logs checked and can see the horizontal appears to have been left on a positive limit S_HMC_AXIS_AT_POSITIVE_LIMIT (Couch is at horizontal outer limit. (soft limit)) Which may have caused the issue. Always good to move the pallet a few millimetres forward, although on second check of the logs the encoder may have shown it had been moved. This was followed by a S_HMC_AXIS_DEST_COLLISION error which could highlight that also the vertical height of the table may also not have been correct for the short tube conditioning to complete. Hence when the QA carried out it would have moved the table and moved it into specification for the short tube conditioning. Internal Comments : Error # and/or description of error : S_HMC_AXIS_AT_POSITIVE_LIMIT (Couch is at horizontal outer limit. (soft limit)) Which may have caused the issue. Always good to move the pallet a few millimetres forward, although on second check of the logs the encoder may have shown it had been moved. This was followed by a S_HMC_AXIS_DEST_COLLISION error which could highlight that also the horizontal position and possibly the vertical height of the table may also not have been correct for the short tube conditioning to complete. Hence when the QA carried out it would have moved the table and moved it into specification for the short tube conditioning. Unable to check due to cantrace issues. Completed or Recommended Repair Action "Internal" : No further action required. System is in full clinical use with no further issues reported. Resolution : No further problems reported - case closed. ----------------------------------------------------------------------------- | 4 |  | Non-Refresh | [12 empty] | 117489232 |  | 869069 |  | #N/A | Opened | [14 empty] | Non-Refresh | [5 empty] | Non-Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | y
836 |  |  | Spain | Policlínica Nuestra Señora del Rosario | 336190 | 4432000% |  | The CT is blocked. Does not recognize the password | Work Order #: WO-03433461 Emilio Olmo Soria The system fully meets specification and returned to use. Diagnostic performed by Engineer: The login of the equipment is blocked; after waiting it unlocks by itself, possibly the passwords were entered several times wrongly and it got blocked for a while. Now it lets you enter and the equipment has started correctly. Problem Description by Engineer: The CT is blocked. Does not recognize the password Regulatory Questions : N/A Resolution: The equipment has been unlocked and they are already working. | 2 |  | Password | [12 empty] | 117488731 |  | 864494 |  | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Application | [6 empty] | Password | [5 empty] | Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
837 |  |  | United States | Select Specialty Hospital | 32261 | 4432000% |  | 0699CDN84 Gantry will not expose | Work Order #: WO-03432556 Mark Finnigan The system does not meet specification and is out of use Problem Description by Engineer : Customer Function/Role: Tech Device Use: Unknown Expected and Unexpected Results: Customer reports that they can not scan. User Impact: Unknown Patient Impact: Current Software Version: Diagnostic performed by Engineer : The case came in on the Web Portal so there was no discussion with the customer. Called the provided phone number and it went to voicemail. Tried to log into the site via PRS and the connection will not establish. All ports are blocked so no remote diagnosis is possible. Dispatching the case for first thing in the morning 5/4. Will leave the customer a VM stating this. Follow Up Required : The case came in on the Web Portal so there was no discussion with the customer. Called the provided phone number and it went to voicemail. Tried to log into the site via PRS and the connection will not establish. All ports are blocked so no remote diagnosis is possible. Dispatching the case for first thing in the morning 5/4. Will leave the customer a VM stating this. ----------------------------------------------------------------------------- Found communication error between gantry and CIRS in error logs. Reseated fiber optics on stator transmitter. Tried multiple times to restart, and allow system to shut off completely for over 30 minutes then restarted. Attempted to run quality checks, and system will not allow the gantry to expose. System went down for a period of 1 hour, and then I turned it back on and ran my checks. The checks worked, and then I attempted to scan a patient. The system would not allow me to scan patient. I then tried to run checks again, but system would not allow me to scan. I checked both computers, and all other accessories that operate the machine. All systems are fine, and the tube heat utility says that the tube heat is fine. Every time I try to expose for checks now it will only give me the generalized error that short tube conditioning Failed. | 3 |  | Gantry CIRS communication | [12 empty] | 117487898 |  | 863920 |  | Closed - Done | Closed - Done |  | TBD_Bugreport | [5 empty] | Engineering | [6 empty] | Gantry CIRS communication | [5 empty] | Refresh |  |  | Upgrade issue | System | TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
838 |  |  | United States | Advent Health Wesley Chapel | 300099 | 4431900% |  | 1302CFXXX Issue transferring to PACS | Work Order #: WO-03429011 Powell, Jamie The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : Logged in via remote console and took a look at customer's queue manager which was empty.. Tried transferring to PACS which was successful. Stayed on the line while customer scanned another patient and she was able to transfer to PACS with no issues. Resolution : Logged in via remote console and took a look at customer's queue manager which was empty.. Tried transferring to PACS which was successful. Stayed on the line while customer scanned another patient and she was able to transfer to PACS with no issues. ----------------------------------------------------------------------------- | 3 |  | PACS Transfer | [12 empty] | 117485312 |  | 862476 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS Transfer | [5 empty] | Refresh |  |  | Upgrade issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
839 |  |  | United States |  | 68116 | 4431900% |  | Remote Alert: "CT-PRO-GANTRY-COUCH_Vertical_Motion | Work Order #: WO-03428856 Powell, Jamie The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: M2M Device Use: Diagnostic Expected and Unexpected Results: N/A User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : Remote Alert: "CT-PRO-GANTRY-COUCH_Vertical_Motion Follow Up Required : "CT-PRO-GANTRY-COUCH_Vertical_Motion Dear colleague, We ask your attention for the following Proactive Alert. To prevent repeating cases, keep the case open until the system issue has been solved. PLANNING ADVICE Please contact customer in order to make an appointment. In order to prevent a customer call, please follow-up . PROBLEM DESCRIPTION The Logfiles show problems with the vertical couch movement. This model alerts when too many of following errors are showing: S_VMC_AXIS_UNALLOWED_MOVE S_VMC_AXIS_AT_COLLISION S_MC_DEVICE_VERT_POT_OPEN S_MC_DEVICE_VERT_POT_OPEN_OR_SHORTED Caused by a mismatch between expected and actual vertical position of the couch, or a vertical movement happened without being commanded. The attached graph is showing which were found and which one triggered the alert. Although not alerting on below errors, these are also shown for your reference: S_MC_AXIS_VERT_INVERTER_FAULT S_MC_DEVICE_INVERTER_FAILURE 0x20000DD S_MC_DEVICE_ESTOP_IS_OPEN S_MC_DEVICE_COUCH_VERTICAL_HW_FAULT S_VMC_AXIS_HW_DETECTED_OBSTRUCTION S_VMC_AXIS_STOP_REASON_UNKNOWN S_VMC_AXIS_WRONG_DEST_REACHED CUSTOMER IMPACT Workflow-, scan interrupts or estop opens. OCCURRENCES OVERVIEW See the attached PDF graph (Attachments) in OneEMS / ServiceMax for this case. RECOMMENDED ACTION Verify the vertical movement over the whole range. Verify limitswitches, cabling and clean string pot by carefully extending it over its complete range. Parts involved are the vertical string pot, vertical drive inverter and CCB. When the graph also shows: S_MC_AXIS_VERT_INVERTER_FAULT S_MC_DEVICE_INVERTER_FAILURE S_MC_DEVICE_COUCH_VERTICAL_HW_FAULT, Table moves in small steps.. Additionally verify: - O-Ring attached to the vertical Encoder. - the switch of the safety green bar. - the cable to the switch of the green safety green bar. - and adjust the vertical separation limit switch - CK1 ESTIMATED TECHNICAL REPAIR TIME Depending on findings. ADDITIONAL INFORMATION For more guidance, see Section 2.2 Vertical Motion Troubleshooting, of the Brilliance/Ingenuity, Patient Support Theory and Troubleshooting Guide, 459800414331. If you have any questions or remarks, please let us know. Kind Regards, Remote.Monitoring.Team.CT@philips.com IS CS – Remote Services" ----------------------------------------------------------------------------- | 5 |  | Non-Refresh | [12 empty] | 117485192 |  | 862341 |  | #N/A | Evaluation | [14 empty] | Non-Refresh | [5 empty] | Non-Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | y
840 |  |  | United States |  | 68114 | 4431800% |  | 0704CDN84 Host not booting up. | Work Order #: WO-03425284 Miller, John M The system does not meet specification and is out of use Problem Description by Engineer : Host won't boot. Customer Function/Role:Technician/Sonographer/Therapist Device Use: Diagnostic Expected and Unexpected Results:Host won't boot. User Impact: None Patient Impact: None Current Software Version: 4.1.7 Follow Up Required : Customer states the Host won't boot all the way up. Hangs in a Blue Screen. Suggest running the Dell or HP Diagnostic (Dell press F12 during bootup, HP press Esc then F2) to determine the health of the PC. Also check the health of the local drive. Also suggest re-seating memory modules. Suggest replacing hard drive and reload software on new drive. Diagnostic performed by Engineer : Host won't boot. Internal Comments : Customer states the Host won't boot all the way up. Hangs in a Blue Screen. Suggest running the Dell or HP Diagnostic (Dell press F12 during bootup, HP press Esc then F2) to determine the health of the PC. Also check the health of the local drive. Also suggest re-seating memory modules. Suggest replacing hard drive and reload software on new drive. ----------------------------------------------------------------------------- Work Order #: WO-03425289 Scheyer William Problem Description by Engineer : Host won't boot. Customer Function/Role:Technician/Sonographer/Therapist Device Use: Diagnostic Expected and Unexpected Results:Host won't boot. User Impact: None Patient Impact: None Current Software Version: 4.1.7 Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 2 |  | Non-Refresh | [12 empty] | 117482235 |  | 861389 |  | #N/A | Evaluation | [14 empty] | Non-Refresh | [5 empty] | Non-Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | y
841 |  |  | United States | Lakeland Regional Medical Cent | 320474 | 4431800% |  | 1101CFXXX Can't get gantry on | Work Order #: WO-03424664 Mcginnis, Mike The system fully meets specification and returned to use Internal Comments : Error # and/or description of error : Can't get gantry on Completed or Recommended Repair Action "Internal" : Jeff said the tech called him in. She told him the screen froze so they restarted the system and now they can't get the gantry on. I questioned him and the host was on but the fans in the gantry were not. I had him check for a start/stop switch in the control room. He found it and hit start and the gantry can't on. Closed ESTOP ok so we ran a test scan and he said everything looks good at this point. Resolution : Jeff said the tech called him in. She told him the screen froze so they restarted the system and now they can't get the gantry on. I questioned him and the host was on but the fans in the gantry were not. I had him check for a start/stop switch in the control room. He found it and hit start and the gantry can't on. Closed ESTOP ok so we ran a test scan and he said everything looks good at this point. Diagnostic performed by Engineer : Jeff said the tech called him in. She told him the screen froze so they restarted the system and now they can't get the gantry on. I questioned him and the host was on but the fans in the gantry were not. I had him check for a start/stop switch in the control room. He found it and hit start and the gantry can't on. Closed ESTOP ok so we ran a test scan and he said everything looks good at this point. Problem Description by Engineer : Problem description by engineer : Can't get gantry on When was the first occurrence? Within... : 01-May-2021 How often is issue occurring? : First occurrence Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Can't get gantry on Patient Impact/Outcome: none User Impact: system was down Current Software Version: 4.1.7.25026 ----------------------------------------------------------------------------- | 2 |  | E-Stop button | [12 empty] | 117481691 |  | 861092 |  | Closed - Done | Closed - Done | [7 empty] | Customer | [6 empty] | E-Stop button | [5 empty] | Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
842 |  |  | United States |  | 300167 | 4431600% |  | 0599CFXXX E-stop not closing | Work Order #: WO-03422666 Boyd Craig The system does not meet specification and is out of use Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: E-stop not closing User Impact: None Patient Impact: None Current Software Version: Diagnostic performed by Engineer : E-stop not closing when turning key on CT Box. Power cycle system. Excise E-stops. Boot system. OK. Resolution : Power cycle system. Excise E-stops. Boot system. OK. | 2 |  | Non-Refresh | [12 empty] | 117480067 |  | 859309 |  | #N/A | Investigation | [14 empty] | Non-Refresh | [5 empty] | Non-Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | y
843 |  |  | United States | Upson Regional Medical Center | 310115 | 4431600% |  | 0612CDY72 E-Stop Opened during contrast injection | Work Order #: WO-03420543 Weaver David The system does not meet specification and is out of use Problem Description by Engineer : DESCRIPTION OF PROBLEM (EXPECTED/UNEXPECTED RESULTS) E-Stop Opened during contrast injection. CALLER FUNCTION/ROLE Technician/Sonographer/Therapist DEVICE INTENDED USE Diagnostic WHEN WAS THE FIRST OCCURRENCE? WITHIN . . . Same day HOW OFTEN IS ISSUE OCCURRING? First occurrence WHERE IN THE END-USER WORKFLOW IS THE ISSUE OCCURRING? During acquisition/scan WHAT RECENT EVENTS HAVE OCCURRED? During normal use NEGATIVE OPERATOR IMPACT Delay in scanning NEGATIVE PATIENT IMPACT/PROCEDURE OUTCOME Can not be scanned HAS CUSTOMER CREATED A SPECIFIC SYSTEM EVENT REPORT OF ISSUE? No 4.1.7 Diagnostic performed by Engineer : E-Stop Opened during contrast injection. Follow Up Required : Saw S_MP_CTRL_NO_MORE_AT_SPEED in logger.mdb. Speed is off more than 2%. Parameter 1 is wanted speed in RPM/100 (IT IS AT 8000). Parameter 2 is actual speed in RPM/100. (IT IS AT 0) If P2=0 then rotation lost its speed and position. This may be caused by Ctrl link errors. Check the counters in the EAL Report. The Rotor did not arrive to the desired speed. Recommendation: 1. Perform Spindle block commission. 2. Check the belt tension. 3. Perform Rotor balancing test. 4. Replace the Spindle block. 5. Replace GMP. ----------------------------------------------------------------------------- Seeing a lot of GPU configuration errors in the logs, and the FSE stated that there were GPU issues a few months back, but were resolved with reseating the board. As there is currently no stock available, requested the FSE reseat the board again to see if the issue goes away temporarily. Work Order #: WO-03420710 04-May-2021 Quach Duy The system fully meets specification and returned to use Problem Description by Engineer : DESCRIPTION OF PROBLEM (EXPECTED/UNEXPECTED RESULTS) E-Stop Opened during contrast injection. CALLER FUNCTION/ROLE Technician/Sonographer/Therapist DEVICE INTENDED USE Diagnostic WHEN WAS THE FIRST OCCURRENCE? WITHIN . . . Same day HOW OFTEN IS ISSUE OCCURRING? First occurrence WHERE IN THE END-USER WORKFLOW IS THE ISSUE OCCURRING? During acquisition/scan WHAT RECENT EVENTS HAVE OCCURRED? During normal use NEGATIVE OPERATOR IMPACT Delay in scanning NEGATIVE PATIENT IMPACT/PROCEDURE OUTCOME Can not be scanned HAS CUSTOMER CREATED A SPECIFIC SYSTEM EVENT REPORT OF ISSUE? No 4.1.7 Diagnostic performed by Engineer : intermittent estop open. view logger. checked gantry belt. checked tension on belt. logger state framing err. Tools & Equipment : MULTIMETER,2021-07-31 00:00:00 Resolution : clean power/signal brushed and slip rings ----------------------------------------------------------------------------- | 3 |  | E-Stop_contrast injection | [12 empty] | 117478456 |  | 858161 |  | Closed - Done | Closed - Done |  | ECR-075975 | [5 empty] | Engineering | [6 empty] | E-Stop_contrast injection | [5 empty] | Refresh |  |  | Upgrade issue | System | Gantry-HW | Lexan Ring | Gantry-HW-Lexan Ring | Gantry-HW-Lexan Ring | ECR-075975 | Lexan Ring | Gantry-HW | Gantry-HW-Lexan Ring | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
844 |  |  | United States |  | 310034 | 4431500% |  | NEMO | Work Order #: WO-03412540 Parts Consumed : 4.00000,453567027551,CORNER, PEDESTAL COVERS ----------------------------------------------------------------------------- | 3 |  | Non-Refresh | [12 empty] | 117470756 |  | 857895 |  |  | Closed - Non-Complaint | [14 empty] | Non-Refresh | [5 empty] | Non-Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | y
845 |  |  | United States |  | 310165 | 4431400% |  | Part Order Per NSS | Work Order #: WO-03404261 Foltz John Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 5 |  | Non-Refresh | [12 empty] | 117462685 |  | 854751 |  | #N/A | Evaluation | [14 empty] | Non-Refresh | [5 empty] | Non-Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | y
846 |  |  | United States |  | 310165 | 4431400% |  | 0308CDN84 Remote Alert: "CT-PRO-GANTRY-S_MC_DEVICE_COUCH_VERTICAL_HW_FAULT | Work Order #: WO-03400813 Holeman, David The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: M2M/RADAR Device Use: No Patient Involved Expected and Unexpected Results: Remote Alert: "CT-PRO-GANTRY-S_MC_DEVICE_COUCH_VERTICAL_HW_FAULT User Impact: None Patient Impact: None Current Software Version: 4.1.7 Diagnostic performed by Engineer : Remote Alert: "CT-PRO-GANTRY-S_MC_DEVICE_COUCH_VERTICAL_HW_FAULT Follow Up Required : PROBLEM DESCRIPTION This model alerts when too many of the following error shows in the logs: S_MC_DEVICE_COUCH_VERTICAL_HW_FAULT Couch does not move up or down smoothly but operator needs to press the button several times to lower or raise the couch. This usually happens when (un)loading the patient and is often the combination of a mechanical obstruction and a (heavy) load. If the couch only moves in small steps of 0.5 or 1mm then the O-ring to the vertical encoder is broken. The attached graph is showing the details of the occurrences and the alert reason. CUSTOMER IMPACT When vertical position is eventually reached, customer can continue ok with scanning OCCURRENCES OVERVIEW See the attached PDF graph (Attachments) in OneEMS / ServiceMax for this case. RECOMMENDED ACTION First check the couch vertical motion to operate smoothly over its complete range. Preferably this should be done with a load simulating a patient in the fully out position. This allows the couch to go completely down. In case the problem occurs all the time: - Check / replace the O-Ring attached to the vertical encoder. - Check / replace the switch of the safety green bar. - For Extended couch only: Check/ replace the two vertical safety switches in series with the green safety bar switch. - Check / replace the cable to the switch of the green safety green bar. - Check CK1 at the Inverter PCB. - Replace the Inverter PCB ----------------------------------------------------------------------------- | 5 |  | Non-Refresh | [12 empty] | 117458748 |  | 849370 |  | #N/A | Evaluation | [14 empty] | Non-Refresh | [5 empty] | Non-Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | y
847 |  |  | United States | Advent Health Wesley Chapel | 300099 | 4431300% |  | 4.1.10 CTA head and neck images freze and boulus tracking isssues |  | 5 |  | CTA head and neck images freze and boulus tracking isssues | [12 empty] | 117451229 |  | 849298 |  | Closed - Done | Closed - Done |  | TBD-bugrep | [5 empty] | Engineering | [6 empty] | CTA head and neck images freze and boulus tracking isssues | [5 empty] | Refresh |  |  | Upgrade issue | System | Console-SW | TBD-bugrep | Console-SW-TBD-bugrep | Console-SW-TBD-bugrep | TBD-bugrep | TBD_Bugreport | Console-SW | Console-SW-TBD-bugrep | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
848 |  |  | Netherlands | Rijnstate Ziekenhuis Arnhem | 52101 | 4431400% |  | Table makes sudden movement in certain position | Work Order #: WO-03400251 NIELS MELICK VAN The system fully meets specification and returned to use Follow Up Required: check table movement Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer: Check table movement ----------------------------------------------------------------------------- Work Order #: WO-03400291 30-Apr-2021 EDDY ELFRING The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : Checked firmware couch controller and calibrated the horizontal table movement. Made a service backup. ----------------------------------------------------------------------------- | 5 |  | Checked firmware couch controller and calibrated the horizontal table movement. Made a service backup. | [12 empty] | 117458894 |  | 848976 |  | Investigation | Investigation |  | TBD-bugrep | [5 empty] | Engineering | [6 empty] | Table sudden movement in certain position | [5 empty] | Refresh |  |  | Upgrade issue | System | Couch-SW | TBD-bugrep | Couch-SW-TBD-bugrep | Couch-SW-TBD-bugrep | TBD_CouchEstop | TBD_Bugreport | Couch-SW | Couch-SW-TBD-bugrep | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
849 |  |  | United States | Advent Health Wesley Chapel | 300099 | 4431400% |  | 0711CDN82 both host monitors go off during scans | Work Order #: WO-03399245 Sheffield, Scott The system partly meets specification and is returned to use with restrictions Internal Comments : Error # and/or description of error : host system event logs Type : Error Date : 4/28/2021 Time : 7:43:04 AM Event : 1001 Source : BugCheck Category : None User : N/A Computer : 128-CT-ED Description: The computer has rebooted from a bugcheck. The bugcheck was: 0x00000119 (0x0000000000000002, 0xffffffffc000000d, 0xffff8281e8e5c2a0, 0xffffb40820b60820). A dump was saved in: C:\Windows\MEMORY.DMP. Report Id: 05acd35c-5605-46ff-8885-89d49869366a. Type : Error Date : 4/28/2021 Time : 8:10:32 AM Event : 10001 Source : DCOM Category : None User : \S-1-5-21-3853370449-718592429-2989568474-1007 Computer : 128-CT-ED Description: Unable to start a DCOM Server: microsoft.windows.immersivecontrolpanel as Unavailable/Unavailable. The error: "5" Happened while starting this command: "C:\Windows\ImmersiveControlPanel\SystemSettings.exe" -ServerName:microsoft.windows.immersivecontrolpanel Type : Error Date : 4/28/2021 Time : 8:10:24 AM Event : 10001 Source : DCOM Category : None User : \S-1-5-21-3853370449-718592429-2989568474-1007 Computer : 128-CT-ED Description: Unable to start a DCOM Server: CortanaUI as Unavailable/Unavailable. The error: "5" Happened while starting this command: "C:\windows\SystemApps\Microsoft.Windows.Cortana_cw5n1h2txyewy\SearchUI.exe" -ServerName:CortanaUI.AppXa50dqqa5gqv4a428c9y1jjw7m3btvepj.mca Type : Error Date : 4/28/2021 Time : 7:43:03 AM Event : 6008 Source : EventLog Category : None User : N/A Computer : 128-CT-ED Description: The previous system shutdown at 7:26:55 AM on ?4/?28/?2021 was unexpected. Type : Error Date : 4/28/2021 Time : 7:43:34 AM Event : 7034 Source : Service Control Manager Category : None User : N/A Computer : 128-CT-ED Description: The GantryServiceHost service terminated unexpectedly. It has done this 3 time(s). host application event logs Type : Error Date : 4/28/2021 Time : 8:10:24 AM Event : 5973 Source : Microsoft-Windows-Immersive-Shell Category : Activation of app %1 failed with error: %2 See the Microsoft-Windows-TWinUI/Oper User : \S-1-5-21-3853370449-718592429-2989568474-1007 Computer : 128-CT-ED Description: Activation of app Microsoft.Windows.Cortana_cw5n1h2txyewy!CortanaUI failed with error: -2147024891 See the Microsoft-Windows-TWinUI/Operational log for additional information. Type : Error Date : 4/28/2021 Time : 8:10:08 AM Event : 1000 Source : Application Error Category : Application Crashing Events User : N/A Computer : 128-CT-ED Description: Faulting application name: Philips.PmsCT.Service.AutomaticDownloadApplication.exe, version: 1.0.0.0, time stamp: 0x5deb7d46 Faulting module name: KERNELBASE.dll, version: 6.2.14393.3085, time stamp: 0x5d1d7bde Exception code: 0xe0434f4d Fault offset: 0x000dc1e2 Faulting process id: 0x(null) Faulting application start time: 0x(null) Faulting application path: (null) Faulting module path: (null) Report Id: (null) Faulting package full name: (null) Faulting package-relative application ID: (null) Type : Error Date : 4/28/2021 Time : 7:43:42 AM Event : 5973 Source : Microsoft-Windows-Immersive-Shell Category : Activation of app %1 failed with error: %2 See the Microsoft-Windows-TWinUI/Oper User : \S-1-5-21-3853370449-718592429-2989568474-1007 Computer : 128-CT-ED Description: Activation of app Microsoft.Windows.Cortana_cw5n1h2txyewy!CortanaUI failed with error: -2147024891 See the Microsoft-Windows-TWinUI/Operational log for additional information. Completed or Recommended Repair Action "Internal" : It appears that the host software is crashing. Consult with NSS about needing to reload host computer software. Diagnostic performed by Engineer : I had Russ create a bug report. The problem happened at 7:30 am today but has been happening since Sunday he said. I downloaded logs and found alot of errors in host system and application event logs. It appears that the host software is crashing. Below are some errors from the host event logs. Problem Description by Engineer : Problem description by engineer : both host monitors go off during scans. This only happens during scans. Both host monitors display go black and the monitor power buttons turn amber color. He said that the host computer has lights on when this happens he has to turn off host computer and turn it back on. The monitors work after rebooting and he is able to finish the scan. It happened today at 7:30 am. I had him create a bug report. FSE has been working on this issue but no previous work order for this issue. Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: host computer monitors go black during scans. Patient Impact/Outcome: delay User Impact: delay Current Software Version: 4.1.7.25030 Follow Up Required : Error # and/or description of error : host system event logs Type : Error Date : 4/28/2021 Time : 7:43:04 AM Event : 1001 Source : BugCheck Category : None User : N/A Computer : 128-CT-ED Description: The computer has rebooted from a bugcheck. The bugcheck was: 0x00000119 (0x0000000000000002, 0xffffffffc000000d, 0xffff8281e8e5c2a0, 0xffffb40820b60820). A dump was saved in: C:\Windows\MEMORY.DMP. Report Id: 05acd35c-5605-46ff-8885-89d49869366a. Type : Error Date : 4/28/2021 Time : 8:10:32 AM Event : 10001 Source : DCOM Category : None User : \S-1-5-21-3853370449-718592429-2989568474-1007 Computer : 128-CT-ED Description: Unable to start a DCOM Server: microsoft.windows.immersivecontrolpanel as Unavailable/Unavailable. The error: "5" Happened while starting this command: "C:\Windows\ImmersiveControlPanel\SystemSettings.exe" -ServerName:microsoft.windows.immersivecontrolpanel Type : Error Date : 4/28/2021 Time : 8:10:24 AM Event : 10001 Source : DCOM Category : None User : \S-1-5-21-3853370449-718592429-2989568474-1007 Computer : 128-CT-ED Description: Unable to start a DCOM Server: CortanaUI as Unavailable/Unavailable. The error: "5" Happened while starting this command: "C:\windows\SystemApps\Microsoft.Windows.Cortana_cw5n1h2txyewy\SearchUI.exe" -ServerName:CortanaUI.AppXa50dqqa5gqv4a428c9y1jjw7m3btvepj.mca Type : Error Date : 4/28/2021 Time : 7:43:03 AM Event : 6008 Source : EventLog Category : None User : N/A Computer : 128-CT-ED Description: The previous system shutdown at 7:26:55 AM on ?4/?28/?2021 was unexpected. Type : Error Date : 4/28/2021 Time : 7:43:34 AM Event : 7034 Source : Service Control Manager Category : None User : N/A Computer : 128-CT-ED Description: The GantryServiceHost service terminated unexpectedly. It has done this 3 time(s). host application event logs Type : Error Date : 4/28/2021 Time : 8:10:24 AM Event : 5973 Source : Microsoft-Windows-Immersive-Shell Category : Activation of app %1 failed with error: %2 See the Microsoft-Windows-TWinUI/Oper User : \S-1-5-21-3853370449-718592429-2989568474-1007 Computer : 128-CT-ED Description: Activation of app Microsoft.Windows.Cortana_cw5n1h2txyewy!CortanaUI failed with error: -2147024891 See the Microsoft-Windows-TWinUI/Operational log for additional information. Type : Error Date : 4/28/2021 Time : 8:10:08 AM Event : 1000 Source : Application Error Category : Application Crashing Events User : N/A Computer : 128-CT-ED Description: Faulting application name: Philips.PmsCT.Service.AutomaticDownloadApplication.exe, version: 1.0.0.0, time stamp: 0x5deb7d46 Faulting module name: KERNELBASE.dll, version: 6.2.14393.3085, time stamp: 0x5d1d7bde Exception code: 0xe0434f4d Fault offset: 0x000dc1e2 Faulting process id: 0x(null) Faulting application start time: 0x(null) Faulting application path: (null) Faulting module path: (null) Report Id: (null) Faulting package full name: (null) Faulting package-relative application ID: (null) Type : Error Date : 4/28/2021 Time : 7:43:42 AM Event : 5973 Source : Microsoft-Windows-Immersive-Shell Category : Activation of app %1 failed with error: %2 See the Microsoft-Windows-TWinUI/Oper User : \S-1-5-21-3853370449-718592429-2989568474-1007 Computer : 128-CT-ED Description: Activation of app Microsoft.Windows.Cortana_cw5n1h2txyewy!CortanaUI failed with error: -2147024891 See the Microsoft-Windows-TWinUI/Operational log for additional information. Completed or Recommended Repair Action "Internal" : It appears that the host software is crashing. Consult with NSS about needing to reload host computer software. ----------------------------------------------------------------------------- Work Order #: WO-03399885 30-Apr-2021 Garcia Marrero, Jorge The system fully meets specification and returned to use Problem Description by Engineer : Problem description by engineer : both host monitors go off during scans. This only happens during scans. Both host monitors display go black and the monitor power buttons turn amber color. He said that the host computer has lights on when this happens he has to turn off host computer and turn it back on. The monitors work after rebooting and he is able to finish the scan. It happened today at 7:30 am. I had him create a bug report. FSE has been working on this issue but no previous work order for this issue. Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: host computer monitors go black during scans. Patient Impact/Outcome: delay User Impact: delay Current Software Version: 4.1.7.25030 Diagnostic performed by Engineer : Re-loading of operating system and applications in host computer. Replacement of ordered parts. System is fully operational. ----------------------------------------------------------------------------- | 3 |  | 1.00000,459801633871,16GB DDR4-2666 (2x8GB) RegRAM 1.00000,459801127432,PCIe CAN CARD PCIe-402-1 1.00000,459801536801,NVIDIA Quadro P400 2GB (3)mDP GFX 1.00000,459800962351,LCC CRC RACK SDB BD ASSY Need check bugreport | [12 empty] | 117458073 |  | 848690 |  | Closed - Done | Closed - Done |  | TBD-bugrep | [5 empty] | Engineering | [6 empty] | Black screen | [5 empty] | Refresh |  |  | Upgrade issue | System | Console-HW | BlackScreen | Console-HW-BlackScreen | Console-HW-BlackScreen | BlackScreen | BlackScreen | Console-HW | Console-HW-BlackScreen | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
850 |  |  | United States |  | 68116 | 4431300% |  | 0999CDY72 Ring artifacts | Work Order #: WO-03394635 Boyd Craig The system partly meets specification and is returned to use with restrictions Follow Up Required : FSE on site to evaluate artifact. Recommendations: Clean gantry window. Check for anything in the field of view. Perform air calibrations. Perform Test scans. Analyze images. Run bad detector test. Diagnostic performed by Engineer : PRS accessed. No errors showing. Customer is third party and not on site. Requesting case be paged out. Issue started a few days ago. Run Compensator Analyzer. No indication of crack in the compensator was found. Problem Description by Engineer : Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected Results: Ring artifacts User Impact: None Patient Impact: None Current Software Version: ----------------------------------------------------------------------------- Work Order #: WO-03394698 28-Apr-2021 Blackford Cable The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected Results: Ring artifacts User Impact: None Patient Impact: None Current Software Version: Diagnostic performed by Engineer : ring artifacts in iac protical ----------------------------------------------------------------------------- | 3 |  | Non-Refresh | [12 empty] | 117453042 |  | 846067 |  | #N/A | Investigation | [14 empty] | Non-Refresh | [5 empty] | Non-Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | y
851 |  |  | United States | Wellstar Kennestone Hospital | 68117 | 4431300% |  | 0101CFXXX Failed during bolus scan | Work Order #: WO-03394484 Mcginnis, Mike The system fully meets specification and returned to use Problem Description by Engineer : Problem description by engineer : Failed during bolus scan How often is issue occurring? : First occurrence Malfunction area : Others Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Failed during bolus scan Patient Impact/Outcome: patient was rescanned, customer stated no one was harmed User Impact: had to restart the scan Current Software Version: 4.1.7.25037 Diagnostic performed by Engineer : Spoke to Melissa, Philips Clinical Applications. She was on site doing follow up training. They had an issue earlier today where a bolus scan stopped. They restarted the host and cirs and scanned the patient ok. Said the system has been working ok since then. The wanted me to review the bug report and let the NSS Sam Pierce know of the incident as he will be on site tomorrow morning to do a Win10 upgrade on the system. I downloaded the bug report and log files via our scripts. Neither of them contained the cirs event logs. See error information found below. Notified NSS Sam Pierce of the findings Internal Comments : Error # and/or description of error : 2021/04/27 14:58:47.730__001 S_MASTER_SEQUENCER_ACQ_NOT_OK ERROR_NORMAL Completed or Recommended Repair Action "Internal" : Their was not enough information to determine the cause of the error. No CIRS logs in bug report or script logs. Notified NSS Sam Pierce of the findings ----------------------------------------------------------------------------- | 3 |  | Need bugreport for analysis | [12 empty] | 117452801 |  | 845982 |  | Closed - Done | Closed - Done |  | TBD-Bugrep | [5 empty] | Engineering | [6 empty] | bolus scan | [5 empty] | Refresh |  |  | Upgrade issue | System | Gantry-SW | TBD-Bugrep | Gantry-SW-TBD-Bugrep | Gantry-SW-TBD-Bugrep | TBD-Bugrep | TBD-Bugrep | Gantry-SW | Gantry-SW-TBD-Bugrep | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
852 |  |  | United States |  | 300106 | 4431200% |  | LS - 061663 AND LS - 061661 | Work Order #: WO-03386674 Parts Consumed : 1.00000,459800041131,Cushion-Extended Patient Support 1.00000,459800126661,Cushion-47.24 Inch Mattress Pad ----------------------------------------------------------------------------- | 5 |  | Non-Refresh | [12 empty] | 117445782 |  | 845664 |  |  | Closed - Non-Complaint | [14 empty] | Non-Refresh | [5 empty] | Non-Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | y
853 |  |  | United States | Lakeland Regional Medical Cent | 320474 | 4431300% |  | Q/O: Cannot connect to system remotely | Work Order #: WO-03392258 27-Apr-2021 Watson Adam The system fully meets specification and returned to use Resolution : Attempted to remotely log into system but found a connection could not be established. Went on site, reset remoteservice password, and resolved the remote connection issue. Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: ----------------------------------------------------------------------------- | 5 |  | PRS Config | [12 empty] | 117451257 |  | 845148 |  | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Service | [6 empty] | PRS Config | [5 empty] | Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
854 |  |  | Argentina | Hospital Dr. Bernardo Houssay | 40027 | 4432400% |  | THE INJECTOR PUMP GETS ERROR 311 AND IS LOCKED | It is validated with Ing Bianco, data is obtained from the failing injector, a visit from the injector supplier is required for review. MEDRAD STELLANT SN: 60764296 207916 | 4 |  | It is validated with Ing Bianco, data is obtained from the failing injector, a visit from the injector supplier is required for review. | [12 empty] | 117518232 |  | 885667 |  | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Injector |  |  | Completed | [5 empty] | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
855 |  |  | China | JS No.2 Chinese Medicine Hosp. | 345027 | 4431700% |  | 02N/Ingenuity Core128/Elite China system freeze |  | CFM |  | Split from PR 885945 | [14 empty] | 885945 |  | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Split |  |  | Completed | [5 empty] | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
856 |  |  | China | SD Jinan No.4 Peo. Hosp. | 40038 | 4432500% |  | 02N/Ingenuity CT 728326/yellow screen | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity CT 728326/ T2 Activities : Problem Description: 客户问题概述： 计算机软件相关故障。 客户问题详细描述或故障现象： 工程师来电开单。 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 已和工程师确认故障。 远程诊断与测试结果： 重启无效。 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议做workaruond升级。 建议备件12NC及名称： N/A Safety Questions : Case: 0117518647 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Bo Last updated by: Wang Bo 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 10.05.2021 01:10:40 UTC Tao Chen (86038373)故障现象: CT host主机黄屏 服务内容: 运行workaround方案 服务结果: 维修后,设备恢复正常. Internal Remarks : 10.05.2021 01:17:36 UTC Tao Chen (86038373) “RDI Accuracy: YES" $Q7: (Confirmed reported problem, CT host yellow screen) $Q8: ( Software bug of CT 5000 host) $Q9: ( Do workaround for CT host) $Q10: ( Test and inspection data refer to: PA. Device returned to full functionality) $Q11: Calibrated Tools name ( N/A) ^Current Tube installed date: N/A ^Current Tube usage: N/A ----------------------------------------------------------------------------- | 3 |  | Follow up case 0116861304 to implement workaround proactively | [12 empty] | 117518647 |  | 885874 |  | Closed - Duplicate | Closed - Duplicate | [14 empty] | Yellow Screen Workaround | [8 empty] | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
857 |  |  | China | SD Chest Hosp. | 40025 | 4432400% |  | 02N/Ingenuity CT 728326/ yellow screen | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity CT 728326 Next Step Action : Huang Zhaokun call in SAP Transfer : Case Transferred to SAP External Remarks : 08.05.2021 11:12:10 UTC SUP JCOUSER (86038405)服务内容：执行CT5000黄屏workround。 服务结果：维修后，设备正常工作。 Internal Remarks : 10.05.2021 11:45:43 UTC Zhaokun Huang (86038405) "RDI Accuracy: NO." $Q7: ( Confirmed reported problem & When End exam , Yellow Screen.) $Q8: ( Execute Yellow Screen Workaround from BIU & Win10 BUG.) $Q9: ( Execute Yellow Screen Workaround from BIU.) $Q10: ( Test and inspection data refer to: PA . Device returned to full functionality.) $Q11: Calibrated Tools name (N/A) ^Current Tube installed date: N/A ^Current Tube usage: N/A Safety Questions : Case: 0117517816 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Jianqiang Last updated by: Wang Jianqiang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 其他类型故障。 客户问题详细描述或故障现象： host主机黄屏 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： host主机黄屏 远程诊断与测试结果： host主机黄屏 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 根据技术支持建议修改文件 SAP Transfer : Dispatched for Onsite Support ----------------------------------------------------------------------------- | 3 |  | Follow up case 0117366742 to implement workaround proactively | [12 empty] | 117517816 |  | 885388 |  | Closed - Duplicate | Closed - Duplicate | [14 empty] | Yellow Screen Workaround | [8 empty] | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
858 |  |  | China | NMG Xinganmeng Hos. | 345030 | 4432400% |  | 02N/Ingenuity Core128/Elite China/consultation | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: Instructed Customer Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ FSE Notes : This PO is a Reference Number only Safety Questions : Case: 0117517783 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wennian Han Last updated by: Wennian Han 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: consultation Troubleshooting Action: customer want to know the max KV and MA value Repair Action: answer customer Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: NA SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | NAC? | [12 empty] | 117517783 |  | 885377 |  | Closed - Non-Complaint | Closed - Non-Complaint | [14 empty] | Consultation | [8 empty] | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
859 |  |  | China | HEN Puyang People's Hosp | 40031 | 4432400% |  | 02N/Ingenuity CT 728326/estop open | BU Q&R Feedback: No Reason: Corrective Maintenance SAP Transfer : Case Transferred to SAP Problem Reported by Customer : 02N/Ingenuity CT 728326/ T2 Activities : Problem Description: 客户问题概述： ESTOP断开。 客户问题详细描述或故障现象： 机架无法启动。 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： lifesync P1=8 ,wait_for_acs_init 远程诊断与测试结果： 重启无效。 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议检查AMC启动状态，检查AMC到CCC通讯I/O连线，必要时更换AMC排除故障。 建议备件12NC及名称： PPD 459800712461 External Remarks : 10.05.2021 12:54:28 UTC Yongping Luo (86038950)服务内容： 水平移动床，设备急停，弹出钥匙框。 服务结果：重新校准水平位置，设备恢复正常。 其它重要事项：无 Internal Remarks : 10.05.2021 12:54:30 UTC Yongping Luo (86038950) $Q7:Confirmed report problem，ESTOP Open. $Q8:Perform visiual check on site ,when Move the table the Estop open, recalibration the couch horizontal position,system return full function. $Q9:recalibration the couch horizontal position ,Device return to full function. $Q10: test and inspection data refer to PA. Device returned to full functionality. $Q11:Calibrated Tools refer to PA. Safety Questions : Case: 0117517194 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Bo Last updated by: Wang Bo 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Dispatched for Onsite Support Next Step Action : PPD 459800712461*1 24H FSE:罗永平-86-15226131133 已安排出货 ----------------------------------------------------------------------------- | 3 |  | 459800712461 FRU, AMC upgrade KITS: check if replace the AMC on field? Parts was good returned @Zhou, Robin@ZHANG, Chang Xian please support investigation and check if this is known issue? CAN trace and logger saved: https://teams.microsoft.com/_#/files/%E5%B8%B8%E8%A7%84?threadId=19%3A76cc42a0c8fa4bd7877f7de12141f84c%40thread.tacv2&ctx=channel&context=Couch%2520E-Stop_8_May_40031&rootfolder=%252Fsites%252FBigBoreRTRefreshFOKCoreTeam%252FShared%2520Documents%252FGeneral%252FIngenuity%2520Bugreps%252FChina%2520KM%252FCouch%2520E-Stop_8_May_40031 | [12 empty] | 117517194 |  | 885116 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD_CouchEstop | [5 empty] | Engineering | [6 empty] | Horizontal calibration |  |  | Completed | [5 empty] | System issue | System | Couch-SW | TBD_CouchEstop | Couch-SW-TBD_CouchEstop | Couch-SW-TBD_CouchEstop | TBD_CouchEstop | TBD_CouchEstop | Couch-SW | Couch-SW-TBD_CouchEstop | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
860 |  |  | China | SD Laiyang People Hospital | 345005 | 4432400% |  | 02N/Ingenuity Core128/Elite China | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity Core128/Elite China SAP Transfer : Dispatched for Onsite Support SAP Transfer : Case Transferred to SAP T2 Activities : Problem Description: 客户问题概述： 图像伪影。 客户问题详细描述或故障现象： 图像每隔几层会有弧形伪影 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 图像弧形伪影 远程诊断与测试结果： 怀疑有坏通道 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 检查屏蔽坏通道，如有必要进行更换模块 建议备件12NC及名称： 459800911281 V2 探测器模块 Safety Questions : Case: 0117517173 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Jianqiang Last updated by: Wang Jianqiang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 3 |  | 459800911281 TDMS V2 M.Det Class B FRU 1 24H已出货 Check if replace DMS on field? 图像有伪影，检查发现第4块module有坏通道，屏蔽无效，更换module，第40块有坏通道屏蔽恢复。 Further step: engineering inputs for service manual update to deliver the IQ troubleshooting knowledge on Refresh version based on DMS team investigation on the failure parts return for analysis. | [12 empty] | 117517173 |  | 885101 |  | Closed - Done | Closed - Done |  | CTUDT00632110 | [5 empty] | Engineering | [6 empty] | DMS Module |  |  | Completed | [5 empty] | System issue | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | DMS Module | DMS-HW | DMS-HW-DMS Module | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
861 |  |  | China | SHX XIANYANG CENTRAL HOSP. | 40050 | 4432300% |  | 02G/Ingenuity CT 728326/transfer failed | BU Q&R Feedback: No SAP Customer Symptom Code: HANG Customer Symptom Code: Hang/Crash Reason: Corrective Maintenance Engineer Failure Code: Configuration Error Repair Activity Code: Reboot/Restart T2 Activities: Problem Description: transfer failed Troubleshooting Action: reboot gantry and software Repair Action: reboot gantry and software Test/Inspection: customer confirm can scan patient and image ok.functional Check PASSED. The device returned to full functionality. Test Equipment: N/A Problem Reported by Customer: 02G/Ingenuity CT 728326 Safety Questions:Case: 0117511285 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Bo Last updated by: Wang Bo 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | 2021-05-07 | [12 empty] | 117511285 |  | 883130 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | Image Transfer |  |  | Completed | [5 empty] | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
862 |  |  | China | JX Province Children Hosp. | 40041 | 4432300% |  | application/02N/Ingenuity CT 728326 | BU Q&R Feedback: No Reason: Customer Information Safety Questions:Case: 0117511124 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Qian Wang Last updated by: Qian Wang 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Problem Reported by Customer: 02N/Ingenuity CT 728326 Next Step Action: Chen Xianfeng 13971190179 ----------------------------------------------------------------------------- | 5 |  | NAC | [12 empty] | 117511124 |  | 883040 |  | Closed - Non-Complaint | Closed - Non-Complaint | [14 empty] | Application | [8 empty] | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
863 |  |  | China | HEN Puyang People's Hosp | 40031 | 4432300% |  | 02n/Ingenuity CT 728326/couch fault | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02n/Ingenuity CT 728326/ SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : PPD FRU, CBL, HANDLE TAPE SWITCH 459800623252 2 已安排24H发货，FSE罗永平-86-15226131133 External Remarks : 09.05.2021 13:50:50 UTC Yongping Luo (86038950)服务内容： 检查床水平按键不移动，只能手动移动，查看报错S HMC TAPE SWITCH PRESSED 。 服务结果：更换TAPE Switch后，设备恢复正常。 其它重要事项：无 Internal Remarks : 09.05.2021 13:50:53 UTC Yongping Luo (86038950) $Q7:Confirmed report problem，the couch cannot movement by press the button. $Q8:Perform visiual check on site ,find the Error code S HMC TAPE SWITCH PRESSED , replace the Tape switch,Device return to full function . $Q9:replace the Tape switch ,Device return to full function. $Q10: test and inspection data refer to PA. Device returned to full functionality. $Q11:Calibrated Tools refer to PA. Safety Questions : Case: 0117510466 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wennian Han Last updated by: Wennian Han 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 床故障 客户问题详细描述或故障现象： 床故障 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： S HMC TAPE SWITCH PRESSED 远程诊断与测试结果： 床水平按键不动，水平灯是亮的，有压带开关错 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 检查压带开关相关 建议备件12NC及名称： " 459800623252 FRU, CBL, HANDLE TAPE SWITCH" ----------------------------------------------------------------------------- | 3 |  | 459800623252 FRU, CBL, HANDLE TAPE SWITCH | [12 empty] | 117510466 |  | 882684 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | Tape Switch |  |  | Completed | [5 empty] | System issue | System | Couch-HW | TabSwitch | Couch-HW-TabSwitch | Couch-HW-TabSwitch | CTUDT00632127 | TabSwitch | Couch-HW | Couch-HW-TabSwitch | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
864 |  |  | China | SHX Ankang T.C.M.Hosp. | 345041 | 4432200% |  | 02N/Ingenuity Core128/Elite China/Image quality issue | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ Safety Questions:Case: 0117506321 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingsong Last updated by: Lu Mingsong 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer's problems: Image quality problem Detailed description of Customer's problems or malfunctions: Image quality problem It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: N/A Remote troubleshoot and test result: N/A PRS TR status: it's not through, ask engineer for maintenance on site. Repair Action: Suggest onsite service procedures: Carry out air calibration and module calibration. We recommend spare parts 12NC and names: N/A Memo of Customer request: The customer asked FSE to be on the field. SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support ----------------------------------------------------------------------------- | 3 |  | 服务内容： 图像质量差，重新做空气矫正，H值校正，phantom校正，调整扫描参数。 服务结果：设备恢复正常。 | [12 empty] | 117506321 |  | 879228 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | ShouldbeCustomer | [5 empty] | Engineering | [6 empty] | IQ |  |  | Completed | [5 empty] | System issue | System | NotEngineering | ShouldbeCustomer | NotEngineering-ShouldbeCustomer | NotEngineering-ShouldbeCustomer | ShouldbeCustomer | NotEngineering-ShouldbeCustomer | NotEngineering | NotEngineering-ShouldbeCustomer | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
865 |  |  | China | JX Yugan Chudong hospital | 337500 | 4432200% |  | wechat:02n/Ingenuity Core128/Elite China/Bed drop button discontinuous failure | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: Bed drop button discontinuous failure Safety Questions:Case: 0117503991 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Lili Last updated by: Lu Lili 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer's problems: CTBOX key of downward failure Detailed description of Customer's problems or malfunctions: CTBOX key of downward failure It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: NA Remote troubleshoot and test result: The customer said ctbox downward key was pressed, but did not respond. The downward key on gantry panel can work normally. The machine is working normally right now.Consider CTBOX itself.12NC on the CTBOX label sent by the customer cannot be found in espf, nor can be found in the system by parts group. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Check CT BOX, check if operation is correct etc.CFM is recommended for related problems. We recommend spare parts 12NC and names: " 459801705022 CT BOX" SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Next Step Action: PPD 459801705022 1 out of stock and has been ordered. FSE xie sheng -86-13970873443 ----------------------------------------------------------------------------- | 5 |  | 459801705022 CT BOX Wait failure part return to Suzhou for analysis | [12 empty] | 117503991 |  | 877637 |  | Investigation | Closed - Done |  | CTUDT00632185 | [5 empty] | Engineering | [6 empty] | CT Box |  |  | Completed | [5 empty] | System issue | System | Gantry-HW | CT Box | Gantry-HW-CT Box | Gantry-HW-CT Box | CTUDT00632185 | CT Box | Gantry-HW | Gantry-HW-CT Box | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
866 |  |  | Japan | Maki Hospital | 341017 | 4432600% |  | This is an enhancement request for Quick viewer or CT viewer on console. |  | CFM | [16 empty] | 885927 |  | Evaluation | Closed - non-complaint | [7 empty] | Customer | [6 empty] | RFC |  |  | Completed | [5 empty] | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
867 |  |  | Japan | Maki Hospital | 341017 | 4432600% |  | This is an enhancement request for CT viewer on console. |  | CFM | [16 empty] | 885918 |  | Evaluation | Closed - non-complaint | [7 empty] | Customer | [6 empty] | RFC |  |  | Completed | [5 empty] | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
868 |  |  | Japan | Maki Hospital | 341017 | 4432400% |  | 01CS: KK: When I create a MIP in a series that shot the head Axial, Slice is divided and recognized. | BU Q&R Feedback: No Reason: Application Support Problem Reported by Customer : 01CS:KK:When a MIP is created in a series of photographs of the head Axial, Slice is divided and recognized. ----------------------------------------------------------------------------- | 4 |  | Without sufficient information | [12 empty] | 117517635 |  | 885285 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632610 | [5 empty] | Engineering | [6 empty] | MIP |  |  | Completed | [5 empty] | System issue | System | Console-SW | MIP | Console-SW-MIP | Console-SW-MIP | CTUDT00632610 | EC Upgrade |  |  | No | y | y | y | y | y | No | y | 0 | y | y | y | y | y | y | 0 | y | y
869 |  |  | Poland | Agencja Rezerw Materiałowych | 341010 | 4432300% |  | DICOM configuration. | Work Order #: WO-03460167 07-May-2021 Mr. Robert Kamola The system fully meets specification and returned to use Problem Description by Engineer: DICOM configuration. Resolution: DICOM configuration of the new CT scanner for RIS/PACS and ISP descriptive stations. Testing network connections. Device is operational. ----------------------------------------------------------------------------- | 4 |  | NAC | [12 empty] | 117513096 |  | 882301 |  | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Service | [6 empty] | System Configuration |  |  | Completed | [5 empty] | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
870 |  |  | United States | The Regional Medical Center | 40005 | 4432600% |  | Need to power on the CT system | Work Order #: WO-03467911 N, Manjunatha The system fully meets specification and returned to use Diagnostic performed by Engineer : Tara came is to find the gantry is switched off. She has tried to power cycle at the mains breaker. Walked though on how to create bugreport. Logs shows issue on the gantry tilt. Problem Description by Engineer : Problem description by engineer : Cannot close Estop / Gantry is off When was the first occurrence? Within... : 10-May-2021 How often is issue occurring? : Continuously Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: Not provided Patient Impact/Outcome: None User Impact: Cannot start up gantry Current Software Version: 4.1.7 Resolution : The call was transferred into wrong queue as we support ISP product hence informed customer that will create case and will inform CT technician to call back for providing further support on this regards, Follow Up Required : What are the next Troubleshooting steps : Read Pot1 and Pot2 in service mode. If one is more than 1000, check the connector of that Pot. Otherwise, recalibrate Pot1 and Pot2 at 0 degrees with a value of 512 +/- 10. Error # and/or description of error : 2021/05/10 01:00:33.721 S_TILT_C_POT_CALIB_KO ERROR_FATAL EStop Initializing Logger.mdb Checked The Calibration between POT1 and POT2 is too big, or the POT connector is OFF Completed or Recommended Repair Action "Internal" : Read Pot1 and Pot2 in service mode. If one is more than 1000, check the connector of that Pot. Otherwise, recalibrate Pot1 and Pot2 at 0 degrees with a value of 512 +/- 10. Internal Comments : Error # and/or description of error : 2021/05/10 01:00:33.721 S_TILT_C_POT_CALIB_KO ERROR_FATAL EStop Initializing Logger.mdb Checked The Calibration between POT1 and POT2 is too big, or the POT connector is OFF Completed or Recommended Repair Action "Internal" : Read Pot1 and Pot2 in service mode. If one is more than 1000, check the connector of that Pot. Otherwise, recalibrate Pot1 and Pot2 at 0 degrees with a value of 512 +/- 10. ----------------------------------------------------------------------------- Work Order #: WO-03468683 10-May-2021 Holt Douglas Problem Description by Engineer : Problem description by engineer : Cannot close Estop / Gantry is off When was the first occurrence? Within... : 10-May-2021 How often is issue occurring? : Continuously Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: Not provided Patient Impact/Outcome: None User Impact: Cannot start up gantry Current Software Version: 4.1.7 Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- Work Order #: WO-03468666 Ali, Adam The system does not meet specification and is out of use Problem Description by Engineer : Problem description by engineer : Cannot close Estop / Gantry is off When was the first occurrence? Within... : 10-May-2021 How often is issue occurring? : Continuously Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: Not provided Patient Impact/Outcome: None User Impact: Cannot start up gantry Current Software Version: 4.1.7 Follow Up Required : What are the next Troubleshooting steps : Read Pot1 and Pot2 in service mode. If one is more than 1000, check the connector of that Pot. Otherwise, recalibrate Pot1 and Pot2 at 0 degrees with a value of 512 +/- 10. Error # and/or description of error : 2021/05/10 01:00:33.721 S_TILT_C_POT_CALIB_KO ERROR_FATAL EStop Initializing Logger.mdb Checked The Calibration between POT1 and POT2 is too big, or the POT connector is OFF Completed or Recommended Repair Action "Internal" : Read Pot1 and Pot2 in service mode. If one is more than 1000, check the connector of that Pot. Otherwise, recalibrate Pot1 and Pot2 at 0 degrees with a value of 512 +/- 10. Internal Comments : Error # and/or description of error : 2021/05/10 01:00:33.721 S_TILT_C_POT_CALIB_KO ERROR_FATAL EStop Initializing Logger.mdb Checked The Calibration between POT1 and POT2 is too big, or the POT connector is OFF Completed or Recommended Repair Action "Internal" : Read Pot1 and Pot2 in service mode. If one is more than 1000, check the connector of that Pot. Otherwise, recalibrate Pot1 and Pot2 at 0 degrees with a value of 512 +/- 10. Diagnostic performed by Engineer : Tara came is to find the gantry is switched off. She has tried to power cycle at the mains breaker. Walked though on how to create bugreport. Logs shows issue on the gantry tilt. ----------------------------------------------------------------------------- | 3 |  | Reseated connections in the e stop circuit | [12 empty] | 117520499 |  | 885976 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | ECR-077012 | [5 empty] | Engineering | [6 empty] | Dual pot |  |  | Completed | [5 empty] | System issue | System | Gantry-HW | Belt | Gantry-HW-Belt | Gantry-HW-Belt | ECR-077012 | Belt | Gantry-HW | Gantry-HW-Belt | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
871 |  |  | United States | St Barnabas Hospital | 40028 | 4432500% |  | 0711CFXXX Software question | Work Order #: WO-03466761 Boyd Craig The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: Software question User Impact: None Patient Impact: None Current Software Version: Resolution : Instructed tech on manual reconstructions. Images good. Diagnostic performed by Engineer : Tech needed to make some manual reconstructions. ----------------------------------------------------------------------------- | 3 |  | NAC | [12 empty] | 117519067 |  | 885610 |  | Closed - Non-Complaint | Closed - Non-Complaint | [14 empty] | Guidance | [8 empty] | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
872 |  |  | United States | St Barnabas Hospital | 40028 | 4432300% |  | recon thin slices | Work Order #: WO-03462463 Curry, Dyan The system fully meets specification and returned to use Problem description by Clinical Support Specialist : Customer Function/ Role: technologist Device Use: in use name Expected and Unexpected Results: recon thins User Impact: none Patient Impact: none Current Software version: 4.1.7 Diagnostic performed by Engineer : helped customer recon thin images in IRS and save to local. ----------------------------------------------------------------------------- | 4 |  | helped customer recon thin images in IRS and save to local. NAC | [12 empty] | 117514938 |  | 880838 |  | Closed - Non-Complaint | Closed - Non-Complaint | [14 empty] | Guidance | [8 empty] | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
873 |  |  | United States | Unity Health Systems | 40037 | 4432200% |  | 0711CFXXX system intermittently freezing when using perfusion exam card | Work Order #: WO-03456966 Swain Wesley The system fully meets specification and returned to use Diagnostic performed by Engineer : Download and analyze bugreport. Consulted Dyan Curry from ClinApps Problem Description by Engineer : Customer Function/Role: CT User Device Use: Rapid Perfusion Exam Card Expected and Unexpected Results: System freezing up intermittently User Impact: Delay Patient Impact: Delay Current Software Version: 4.1.10 Resolution : Dyan Curry from ClinApps informed that this is a known issue with rapid protocol. Dyan provided workaround as follows. Scan Surview, Plan 2nd perfusion study, Plan 1st perfusion study. Scan normally. Angela will integrate this into their workflow. | 3 |  | Dyan provided workaround as follows. Scan Surview, Plan 2nd perfusion study, Plan 1st perfusion study. Scan normally. Angela will integrate this into their workflow. Is it a known issue with defect tracking? | [12 empty] | 117509271 |  | 876664 |  | Investigation | Closed - Done |  | TBD_Bugreport | [5 empty] | Engineering | [6 empty] | Freezing |  |  | Completed | [5 empty] | System issue | System | Console-SW | TBD-bugrep | Console-SW-TBD-bugrep | Console-SW-TBD-bugrep | TBD_Bugreport | TBD_Bugreport |  |  | No | y | y | y | y | y | No | y | 0 | y | y | y | y | y | y | 0 | y | y
874 |  |  | United States | Leesburg Regional Medical Cent | 40022 | 4432200% |  | 0308CDY72 Noise during table vertical motion | Work Order #: WO-03454978 Swain Wesley The system partly meets specification and is returned to use with restrictions Follow Up Required : inspect covers for rubbing. inspect / lubricate vertical guide rails. inspect lubricate ball screw and thrust bearings. Unable to provide part numbers for replacements as PRS will not allow telnet session and was unable to determine couch type. Diagnostic performed by Engineer : rubbing noise when operating table vertical drive Problem Description by Engineer : Customer Function/Role: CT User Device Use: table vertical motion Expected and Unexpected Results: noisy vertical drive User Impact: none Patient Impact: none Current Software Version: n/a ----------------------------------------------------------------------------- | 3 |  | cleaned of old grease from verticle scissors tubes and relubricated tested and no more noise, followed up next day still good | [12 empty] | 117507740 |  | 875081 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00631545/CTUDT00630971 | [5 empty] | Engineering | [6 empty] | Couch Noise |  |  | Completed | [5 empty] | System issue | System | Couch-HW | Table Noise | Couch-HW-Table Noise | Couch-HW-Table Noise | CTUDT00631545/CTUDT00630971 | Table Noise | Couch-HW | Couch-HW-Table Noise | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
875 |  |  | United States | Clay County Medical Center | 40003 | 4432000% |  | The customer has had multiple issues running the IQ check. It will tell them to move the phantom multiple times | Phantom detection issue | CFM |  | Bugreport received https://www.clouddrive.philips.com.cn/file/pkokusrr69689w09ew0y63k5dq7h7acj# Password:123 | [14 empty] | 873986 |  | Investigation | Closed - Done |  | CTUDT00631545/CTUDT00630971 | [5 empty] | Engineering | [6 empty] | Phantom detection |  |  | Completed | [5 empty] | System issue | System | DMS-SW | Phantom Detect | DMS-SW-Phantom Detect | DMS-SW-Phantom Detect | CTUDT00631545/CTUDT00630971 | Phantom Detect | DMS-SW | DMS-SW-Phantom Detect | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
876 |  |  | United States | Bibb Medical Center | 310175 | 4432300% |  | Q/O: Image Artifact | Work Order #: WO-03462944 10-May-2021 Foster Bobby The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : checked compensator, cleaned lexan ring and detector cover | 5 |  | Resolution Description: checked compensator, cleaned lexan ring and detector cover | [12 empty] | 117515516 |  | 884589 |  | Closed - Done | Closed - Done |  | ECR-075975 | [5 empty] | Engineering | [6 empty] | Artifact |  |  | Complete |  |  | Refresh |  |  | Upgrade issue | System | Gantry-HW | Lexan Ring | Gantry-HW-Lexan Ring | Gantry-HW-Lexan Ring | ECR-075975 | Lexan Ring | Gantry-HW | Gantry-HW-Lexan Ring | No | y | y | y | y | y | No | y | 0 | y | y | y | y | y | y | 0 | y | y
877 |  |  | United States | Advent Health Wesley Chapel | 300099 | 4432400% | 18 | 0705CDN84 customer locked out of system, multiple issues | Work Order #: WO-03465132 08-May-2021 Stoutt Chad The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Problem description by engineer : Customer stated they have had nothing but problems since the upgrade. She said there are " a Slew" of bug reports on the system. The keyboard has been intermittent and now they are completely locked out of the system. When was the first occurrence? Within... : 07-May-2021 How often is issue occurring? : Multiple times a day Malfunction area : Others Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: customer locked out of system, multiple issues Patient Impact/Outcome: moved to another scanner User Impact: system down Current Software Version: 4.1.7.25030 Diagnostic performed by Engineer : Onsite 5/8/2021 to diagnose problem. System was in use upon arrival. Customer indicated that a few buttons on keyboard were not working. Keyboard was replaced with part ordered by PFSE. Customer also indicated they are experiencing intermittent blank/black screens (both) after a surview or scan. Customer then has to force reboot on Host PC and perform offline recons. Suspected Host PC video card to be the cause of the issue. Follow Up Required : Replace host PC video card ----------------------------------------------------------------------------- Work Order #: WO-03465120 Mcginnis, Mike The system does not meet specification and is out of use Problem Description by Engineer : Problem description by engineer : Customer stated they have had nothing but problems since the upgrade. She said there are " a Slew" of bug reports on the system. The keyboard has been intermittent and now they are completely locked out of the system. When was the first occurrence? Within... : 07-May-2021 How often is issue occurring? : Multiple times a day Malfunction area : Others Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: customer locked out of system, multiple issues Patient Impact/Outcome: moved to another scanner User Impact: system down Current Software Version: 4.1.7.25030 Diagnostic performed by Engineer : Spoke to Diana. She stated they have had nothing but problems since the upgrade. She said there are " a Slew" of bug reports on the system. The keyboard has been intermittent and now they are completely locked out of the system. I am getting errors when I attempt to remote to the system to review logs. Internal Comments : Error # and/or description of error : customer locked out of system, multiple issues Completed or Recommended Repair Action "Internal" : Reset the password, according to the customer you may have to install a different keyboard. She said one was supposed to be there today but they haven't seen it so one is on the way. Once you are back into the system review the log files and troubleshoot per your findings. Could not connect to the system for review of the logs remotely Follow Up Required : Error # and/or description of error : customer locked out of system, multiple issues Completed or Recommended Repair Action "Internal" : Reset the password, according to the customer you may have to install a different keyboard. She said one was supposed to be there today but they haven't seen it so one is on the way. Once you are back into the system review the log files and troubleshoot per your findings. Could not connect to the system for review of the logs remotely ----------------------------------------------------------------------------- Work Order #: WO-03466034 10-May-2021 Sher Mark The system fully meets specification and returned to use Problem Description by Engineer : Problem description by engineer : Customer stated they have had nothing but problems since the upgrade. She said there are " a Slew" of bug reports on the system. The keyboard has been intermittent and now they are completely locked out of the system. When was the first occurrence? Within... : 07-May-2021 How often is issue occurring? : Multiple times a day Malfunction area : Others Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: customer locked out of system, multiple issues Patient Impact/Outcome: moved to another scanner User Impact: system down Current Software Version: 4.1.7.25030 Follow Up Required : outstanding NFO part to be returned.. I have in possession Diagnostic performed by Engineer : bypassed computer cabinet back interface and ran directly to host for the monitor video to correct video errors ----------------------------------------------------------------------------- Work Order #: WO-03471813 08-May-2021 Watson Adam Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 2 |  | Part#:459801536801 Description: NVIDIA Quadro P400 2GB (3)mDP GFX Quantity: 1 | [12 empty] | 117517087 |  | 882762 |  | Closed - Non-Complaint | Closed - Non-Complaint | [14 empty] | Black screen | [5 empty] | Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
878 |  |  | United States | Advent Health Wesley Chapel | 300099 | 4432200% |  | Surview Scan Plan -Exam Aborts | Work Order #: WO-03456576 Carpenter, Kevin The system partly meets specification and is returned to use with restrictions Follow Up Required : Philips Onsite Clin Apps ( Mellissa ) relays on going issues with CTHost : Since Wni10 upgrade : During Plan Scan on Surview - Exam Aborts 12:08pm Monitors go blank ( sometimes just one ) 12:58pm Keyboard - control key / num lock not working - then letter "B" She shook keyboard and letter ' B ' began working again. Cust says service has already worked on this and CTHost may have already been replaced once. FSE while bad keyboard likely would not cause monitors to go blank ~ a bad keyboard has been know to cause bizarre problems. Please consult NSS, but in mean time please do replace keyboard. CT RSE is ordering one to site. Diagnostic performed by Engineer : Reviewed logs - ordered keyboard to site ( per onsite Philips Clin Apps ) ~ Advised FSE to consult with NSS as this is a repeat problem. Diagnostic performed by Engineer : Reviewed CTHost EVT logs from very large bugrep ( took a long time to d/l ) ~ Forwarded same to FSE assigned to onsite, pFSE and regional NSS. Problem Description by Engineer : Customer Function/Role: Philips Onsite Clin Apps ( Mellissa ) Device Use: Diag Expected and Unexpected Results: CTHost Lockups and odd behavior User Impact: None Patient Impact: None Current Software Version: IDK ----------------------------------------------------------------------------- Work Order #: WO-03456629 10-May-2021 Sher Mark The system fully meets specification and returned to use Diagnostic performed by Engineer : worked with Adam on differend wo to correct vlicker issues Problem Description by Engineer : Customer Function/Role: Philips Onsite Clin Apps ( Mellissa ) Device Use: Diag Expected and Unexpected Results: CTHost Lockups and odd behavior User Impact: None Patient Impact: None Current Software Version: IDK ----------------------------------------------------------------------------- | 3 |  | Parts Consumed Description: 1.00000,459800918431,Cherry USB keyboard English Silver | [12 empty] | 117508913 |  | 876293 |  | Investigation | Investigation |  | NotEngineering-partreplace | [5 empty] | Engineering | [6 empty] | Keyboard |  |  | Complete |  |  | Refresh |  |  | Upgrade issue | System | Console-HW | Keyboard | Console-HW-Keyboard | Console-HW-Keyboard | NotEngineering-partreplace | Keyboard | Console-HW | Console-HW-Keyboard | No | y | y | y | y | y | No | y | 0 | y | y | y | y | y | y | 0 | y | y
879 |  |  | Argentina | Hospital Dr. Bernardo Houssay | 40027 | 4432600% |  | A CIRCLE IS DISPLAYED IN THE IMAGE THAT PREVENTS THE STUDIES FROM BEING CLEARLY SEEN. | Work Order #: WO-03476495 Alec, Preciado The system partly meets specification and is returned to use with restrictions Follow Up Required : What are the next Troubleshooting steps : Check the artifact, error log, checked raw data in the system, identify damaged detector and add it to the bad detector list, run corresponding calibrations to exclude it. Error # and/or description of error : ring artifact in images Completed or Recommended Repair Action "Internal" : It is confirmed with field service engineer Edgardo that the system showed a damaged detector last time in a module different from the one that was replaced, a visit on site is required to run tests on modules. Diagnostic performed by Engineer: We contact the user and the field service engineer, we identify that there is a detector in a module that generates an artifact, a visit to the site is required to check it. Problem Description by Engineer : Problem description by engineer : A CIRCLE IS DISPLAYED IN THE IMAGE THAT PREVENTS THE STUDIES FROM BEING CLEARLY SEEN. When was the first occurrence? Within... : 11-May-2021 09:00:00 How often is issue occurring? : Multiple times a day Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected results: System without artifacts Patient Impact/Outcome: Delay in attention User Impact: Delay in work load. Current Software Version: 4.1.7 Internal Comments : Error # and/or description of error : ring artifact in images Completed or Recommended Repair Action "Internal" : It is confirmed with field service engineer Edgardo that the system showed a damaged detector last time in a module different from the one that was replaced, a visit on site is required to run tests on modules. ----------------------------------------------------------------------------- Work Order #: WO-03478687 11-May-2021 Meza, Edgardo Arturo The system fully meets specification and returned to use. Problem Description by Engineer: A RING IS DISPLAYED IN THE IMAGE PREVENTING THE STUDIES FROM BEING CLEARLY SEEN. Resolution: Air calibration is performed and it is tested with a patient. Artifact is removed. Diagnostic Performed by Engineer: Calibration table is corrupt. ----------------------------------------------------------------------------- | 4 |  | Calibration table is corrupt | [12 empty] | 117523859 |  | 891772 |  | Closed - Done | Closed - Done |  | NotEngineering | [5 empty] | Engineering | [6 empty] | Artifact |  |  | Completed |  |  | Refresh |  |  | System issue | System | NotEngineering | ShouldbeCustomer | NotEngineering-ShouldbeCustomer | NotEngineering-ShouldbeCustomer | NotEngineering | NotEngineering-ShouldbeCustomer | NotEngineering | NotEngineering-ShouldbeCustomer | No | y | y | y | y | y | No | y | 0 | y | y | y | y | y | y | 0 | y | y
880 |  |  | Chile | HOSPITAL BARROS LUCO TRUDEAU | 341005 | 4432600% |  | CL1688 - STUDIO SETTING PROBLEMS, TUBE TURNING PROBLEMS, DOES NOT START STUDIES AND HAS NOISE IN THE GANTRY | Work Order #: WO-03479792 Alec, Preciado The system fully meets specification and returned to use Diagnostic performed by Engineer : Se contacta a usuario turno de la tarde, indica que ha trabajado sin problema por 4 días, revisión de logs, se contacta a encargada del área, se identifica ruido en banda de rotor y rotor. Indican que la temperatura del área es de 20°C, aún así perciben problemas de desempeño de tubo. Problem Description by Engineer : Problem description by engineer : CL1688 - PROBLEMAS DE SETEO DE ESTUDIOS, PROBLEMAS CON GIRO DE TUBO, NO INICIA LOS ESTUDIOS Y TIENE RUIDO EN EL GANTRY When was the first occurrence? Within... : 11-May-2021 15:00:00 How often is issue occurring? : First occurrence Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Not provided Expected and Unexpected results: No further information is available as per the customer Patient Impact/Outcome: no impacta User Impact: no impacta Current Software Version: 4.1.10 Internal Comments : Error # and/or description of error : WARNINGS Optical slip ring signal error. S_DMC_SLIP_RING S_COUCHMG_MOTION_START_ERROR Completed or Recommended Repair Action "Internal" : Se identifican Warnings, no se identifican errores normales o fatales que indicaran que el sistema presenta problemas. Follow Up Required : What are the next Troubleshooting steps : Se requiere limpieza del punto de encuentro entre rotor y banda de motor, presenta ruido por suciedad o un desajuste, buscar por debris en la parte inferior de la banda. Revisar el Gap del tensor, de acuerdo a manual de reemplazo, así como revisión por desgaste o rotura en banda. Se requiere corroborar las condiciones de temperatura de la sala y realizar pruebas de desempeño al tubo, así como revisión de protocolos con usuario para identificar si el tubo presenta mensaje de calentamiento de tubo. Realizar calibración horizontal, se identifican algunos warnings relacionados con los limites y el avance de la carbon top. Error # and/or description of error : WARNINGS Optical slip ring signal error. S_DMC_SLIP_RING S_COUCHMG_MOTION_START_ERROR Completed or Recommended Repair Action "Internal" : Se identifican Warnings, no se identifican errores normales o fatales que indicaran que el sistema presenta problemas. ----------------------------------------------------------------------------- Work Order #: WO-03484019 Osorio, Juan Problem Description by Engineer : Problem description by engineer : CL1688 - PROBLEMAS DE SETEO DE ESTUDIOS, PROBLEMAS CON GIRO DE TUBO, NO INICIA LOS ESTUDIOS Y TIENE RUIDO EN EL GANTRY When was the first occurrence? Within... : 11-May-2021 15:00:00 How often is issue occurring? : First occurrence Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Not provided Expected and Unexpected results: No further information is available as per the customer Patient Impact/Outcome: no impacta User Impact: no impacta Current Software Version: 4.1.10 Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 3 |  | Cleaning of the meeting point between rotor and motor belt is required, it presents noise due to dirt or a mismatch, look for debris in the lower part of the belt. Check the Gap of the tensioner, according to the replacement manual, as well as check for wear or breakage in the belt. It is required to corroborate the temperature conditions of the room and perform performance tests on the tube, as well as a review of protocols with the user to identify if the tube presents a tube heating message. Perform horizontal calibration, some warnings related to the limits and the advance of the carbon top are identified. | [12 empty] | 117523046 |  | 891053 |  | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Multi issue |  |  | Completed |  |  | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
881 |  |  | China | ZJ YUEQING PEOPLE HOSPITAL | 40054 | 4433100% |  | 02N/Ingenuity CT 728326 | BU Q&R Feedback: No SAP Customer Symptom Code: CENV Customer Symptom Code: Customer Environmental Concern Reason: Corrective Maintenance Engineer Failure Code: Environmental Cause Repair Activity Code: Reboot/Restart Problem Reported by Customer : 02N/Ingenuity CT 728326 Safety Questions : Case: 0117548137 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Weng Jinming Last updated by: Weng Jinming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Troubleshooting Action: instruct customer to restart the system Repair Action: instruct customer to restart the system Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | instruct customer to restart the system | [12 empty] | 117548137 |  | 907990 |  | Closed - Non-Complaint | Closed - Non-Complaint | [14 empty] | E-Stop | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
882 |  |  | China | Lindian County Hospital | 345038 | 4433100% |  | 02N/Ingenuity Core128/Elite China/ECG signal Loss | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ Safety Questions : Case: 0117548523 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Weng Jinming Last updated by: Weng Jinming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action : PPD 459800166651*1 459801731761*1备件已系统下单，安排24H到场，FSE：倪刚-86-18686766464（替代号已确认） External Remarks : 17.05.2021 00:29:48 UTC Gang Ni (86038229)服务内容：连接ibox电缆 服务结果：设备恢复 T2 Activities : Problem Description: 客户问题概述： CT机操作台没有心电信号 客户问题详细描述或故障现象： 心电监护仪正常 设备本身技术问题： 是。 Troubleshooting Action: 远程诊断与测试结果： 怀疑心电电缆，I-box，驱动软件。。。问题 PRS TR状态： 全通。 Repair Action: 远程诊断与测试结果： 怀疑心电电缆，I-box，驱动软件。。。问题 PRS TR状态： 全通。 T2 Activities : Problem Description: 客户问题概述： CT机操作台没有心电信号 客户问题详细描述或故障现象： 心电监护仪正常 设备本身技术问题： 是。 Troubleshooting Action: 远程诊断与测试结果： 怀疑心电电缆，I-box，驱动软件。。。问题 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议工程师到场检查心电电缆，驱动软件，I-box，教客户使用飞利浦微信报修 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support ----------------------------------------------------------------------------- | 3 |  | Reconnect ibox cable | [12 empty] | 117548523 |  | 907958 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Ibox |  |  | Completed |  |  | Refresh |  |  | System issue | System | Gantry-HW | TBD-NoEnoughInfo | Gantry-HW-TBD-NoEnoughInfo | Gantry-HW-TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo | Gantry-HW | Gantry-HW-TBD-NoEnoughInfo | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
883 |  |  | China | XJ Yili County Friendship Hospital | 40048 | 4433100% |  | 02N/Ingenuity CT 728326/The couch move stuck | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity CT 728326 Safety Questions : Case: 0117548506 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Weng Jinming Last updated by: Weng Jinming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support T2 Activities : Problem Description: 客户问题概述： 新装机，床水平运动卡顿 设备本身技术问题： 是。 Troubleshooting Action: 远程诊断与测试结果： 怀疑外壳，丝杠。。问题 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议工程师到场检查外壳，丝杠，教客户使用飞利浦微信报修 ----------------------------------------------------------------------------- | 3 |  | 床水平运动卡顿, 建议工程师到场检查外壳，丝杠 | [12 empty] | 117548506 |  | 907940 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | AssemblyAdjust | [5 empty] | Engineering | [6 empty] | Vertical Transducer |  |  | Completed |  |  | Refresh |  |  | System issue | System | Couch-HW | Table move stuck | Couch-HW-Table move stuck | Couch-HW-Table move stuck | AssemblyAdjust | Table move stuck | Couch-HW | Couch-HW-Table move stuck | No | y | y | y | y | y | No | y | 0 | y | y | y | y | y | y | 0 | y | y
884 |  |  | China | Tongling Boai Hospital Co LTD | 345034 | 4433000% |  | 02N/Ingenuity Core128/TILE LOSE POWER | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Replace Part Problem Reported by Customer : 02N/Ingenuity Core128/ Safety Questions : Case: 0117543360 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Cao Last updated by: Liu Cao 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP T2 Activities : Problem Description: 客户问题概述： 机架打角度掉电，不能归零。 客户问题详细描述或故障现象： 机架打角度掉电，不能归零。 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： S_TILT_C_NO_MOVE S_ROTOR_MCO 远程诊断与测试结果： 检查pot示数，需要时进行更换； pot正常时检查倾斜电机液压杆两侧是否平衡，右侧电源板指示，接线等 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 1.检查pot示数，需要时进行更换； 2.pot正常时检查倾斜电机液压杆两侧是否平衡，右侧电源板指示，接线等 PPD 24H: 459800441121 DUAL POT ASSEMBLY 1 Test Equipment: N/A SAP Transfer : Dispatched for Onsite Support Next Step Action : PPD备件DUAL POT ASSEMBLY 459800441121 1已安排24H到场地，FSE：周凯-86-15056051758 ----------------------------------------------------------------------------- | 3 |  | 服务内容：打倾角后机架不能上电，mco旋转电机过温，处理温控接头。 服务结果：服务后设备恢复正常。 | [12 empty] | 117543360 |  | 906008 |  | Closed - Done | Closed - Done |  | NotEngineering | [5 empty] | Engineering | [6 empty] | Tilt |  |  | Completed |  |  | Refresh |  |  | System issue | System | NotEngineering | ShouldbeService | NotEngineering-ShouldbeService | NotEngineering-ShouldbeService | NotEngineering | NotEngineering-ShouldbeService | NotEngineering | NotEngineering-ShouldbeService | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
885 |  |  | China | SD Tai'an T.C.M. Hosp. No.2 Ho | 345037 | 4432900% |  | 02N/Ingenuity Core128/Elite China/Replace Flat Spring | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 05R/Ingenuity Core128/Elite China Safety Questions : Case: 0117538657 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingsong Last updated by: Lu Mingsong 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | replace flat spring | [12 empty] | 117538657 |  | 902531 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632833 | [5 empty] | Engineering | [6 empty] | Flat Spring |  |  | Completed |  |  | Refresh |  |  | System issue | System | Gantry-HW | FlatSpring | Gantry-HW-FlatSpring | Gantry-HW-FlatSpring | CTUDT00632833 | FlatSpring | Gantry-HW | Gantry-HW-FlatSpring | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
886 |  |  | China | HLJ Harbin Senhai Hospital | 345007 | 4432900% | 19 | 02N/Ingenuity Core128/Elite China/warm up error | T2 Activities : Problem Description: 客户问题概述： 预热失败。 客户问题详细描述或故障现象： 预热失败。 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 不能预热 Repair Action: 建议现场维修步骤： 根据后台报错做相关troubleshooting Safety Questions : Case: 0117538024 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 14.05.2021 13:56:57 UTC Haijun Du (86038210)"SUAT. Ready for Clinical Use local time" : "2021-05-14 18:00" 服务内容：更换阴极高压。 服务结果：维修后，设备恢复正常工作。 Internal Remarks : 14.05.2021 13:57:26 UTC Haijun Du (86038210) $Q7: (Confirmed reported problem &can't scan) $Q8: (check the CT log，S_RHOST_HVG_HSS_GENERATOR_NON_ARC_FAULT 1:1 Grid error) $Q9: (ok after replaced CPM) | 2 |  | $Q7: (Confirmed reported problem &can't scan) $Q8: (check the CT log，S_RHOST_HVG_HSS_GENERATOR_NON_ARC_FAULT 1:1 Grid error) $Q9: (ok after replaced CPM) 服务内容：更换阴极高压。 服务结果：维修后，设备恢复正常工作。 | [12 empty] | 117538024 |  | 902103 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-APM/CPM | [5 empty] | Engineering | [6 empty] | CPM |  |  | Completed |  |  | Refresh |  |  | System issue | System | IC-HW | APM/CPM | IC-HW-APM/CPM | IC-HW-APM/CPM | TBD-APM/CPM | APM/CPM | IC-HW | IC-HW-APM/CPM | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
887 |  |  | China | Wuhan Ren 'ai Hospital, Hubei Province | 345004 | 4432900% |  | 02N/Ingenuity Core128/Elite China/system can not log in | BU Q&R Feedback: No SAP Customer Symptom Code: BOOT Customer Symptom Code: Boot up/Power-up Problem Reason: Corrective Maintenance Engineer Failure Code: Incompatible Use Repair Activity Code: Instructed Customer Safety Questions : Case: 0117537938 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Gaoge Last updated by: Wang Gaoge 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Problem Reported by Customer : 02N/Ingenuity Core128/Elite China T2 Activities : Problem Description: 主机启动后登录CT账户自动注销，不能登录 Troubleshooting Action: 指导客户开机后等待三五分钟，让后台进程完全启动后再次登录，此时登录正常。 Repair Action: 指导客户等待几分钟让后台进程完成启动后再登录 Test/Inspection: N/A Test Equipment: N/A Next Step Action : Q6 = system can not log in and the system log out when input the user name and password SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | 指导客户开机后等待三五分钟，让后台进程完全启动后再次登录，此时登录正常。 | [12 empty] | 117537938 |  | 902014 |  | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Login |  |  | Completed |  |  | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
888 |  |  | China | ZJ YUEQING PEOPLE HOSPITAL | 40054 | 4432800% |  | 02N/Ingenuity CT 728326change cover | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity CT 728326 Safety Questions : Case: 0117533931 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Tao Fengming Last updated by: Tao Fengming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 其他类型故障。 客户问题详细描述或故障现象： 更换外壳 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 更换外壳 远程诊断与测试结果： 需要更换外壳 Repair Action: 建议现场维修步骤： 更换外壳 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : huang yi Internal Remarks : 13.05.2021 03:05:19 UTC Yi Huang (86036249) RDI Accuracy:Yes $Q7: (Confirmed reported problem&system can't boot up) $Q8: (check errorlog found lifesync 8, no keyprompt, found CCC LED change from 22 to 00) $Q9: (replace CCC) $Q10: (Test and inspection data refer to: PA.Device returned to full functionality.) $Q11: Calibrated tools name(N/A) S/N (N/A) Next calibration date (N/A) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date:N/A ^Current Tube usage: second:N/A External Remarks : 13.05.2021 03:05:17 UTC SUP JCOUSER (86036249)服务内容：机架不能上电，lifesync8，床控制板状态异常，更换控制板。 服务结果：维修后，设备恢复正常工作。 其它重要事项：无 ----------------------------------------------------------------------------- | 3 |  | 服务内容：机架不能上电，lifesync8，床控制板状态异常，更换控制板。 服务结果：维修后，设备恢复正常工作。 | [12 empty] | 117533931 |  | 899292 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD_CouchEstop | [5 empty] | Engineering | [6 empty] | CCB |  |  | Completed |  |  | Refresh |  |  | System issue | System | Couch-SW | TBD-bugrep | Couch-SW-TBD-bugrep | Couch-SW-TBD-bugrep | TBD_CouchEstop | TBD_Bugreport | Couch-SW | Couch-SW-TBD-bugrep | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
889 |  |  | China | SHX XIANYANG CENTRAL HOSP. | 40050 | 4432800% |  | 02G/Ingenuity CT 728326/transfer image failed | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer : 02G/Ingenuity CT 728326 T2 Activities : Problem Description: 客户问题概述： 图像传输故障 客户问题详细描述或故障现象： 图像传输故障 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： NA 远程诊断与测试结果： 工程师已经指导客户检查网络正常，节点信息这次，服务模式提示密码不对无法reset spooler PRS TR状态： 不通，请工程师到场检查。 Repair Action: 建议现场维修步骤： 需要到场检查图像传输相关，用工程师权限设置olevel账户密码 Safety Questions : Case: 0117533853 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wennian Han Last updated by: Wennian Han 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action : 设备暂未入保，在IA# 53341371上加operation，派工fse 刘军，以邮件IA Mgr批复 Next Step Action : Op-20已加 ----------------------------------------------------------------------------- | 3 |  | 需要到场检查图像传输相关，用工程师权限设置olevel账户密码, reset spooler | [12 empty] | 117533853 |  | 899231 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | Image Transfer |  |  | Completed |  |  | Refresh |  |  | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
890 |  |  | China | ZJ YUEQING PEOPLE HOSPITAL | 40054 | 4432800% |  | 02N/Ingenuity CT 728326/replace shell | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: Instructed Customer Problem Reported by Customer : replace shell Safety Questions : Case: 0117532354 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Tao Fengming Last updated by: Tao Fengming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: need check option Troubleshooting Action: need customer check option Repair Action: guide customer check option Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | NAC | [12 empty] | 117532354 |  | 898301 |  |  | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
891 |  |  | China | SD Laiyang People Hospital | 345005 | 4432800% |  | 02N/Ingenuity Core128/add node | BU Q&R Feedback: No Reason: Corrective Maintenance SAP Transfer : Dispatched for Onsite Support Problem Reported by Customer : 02N/Ingenuity Core128/ Safety Questions : Case: 0117532244 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingsong Last updated by: Lu Mingsong 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 添加节点 客户问题详细描述或故障现象： 添加节点 设备本身技术问题： 否。 Troubleshooting Action: 关键故障代码或现象： N/A 远程诊断与测试结果： olevel_service账号不能登录，无法使用lanconfig PRS TR状态： 不通，请工程师到场检查。 Repair Action: 建议现场维修步骤： 现场添加节点. 建议备件12NC及名称： N/A 客户需求备注： 无 SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 5 |  | olevel_service账号不能登录，无法使用lanconfig PRS TR状态： 不通，请工程师到场检查。 | [12 empty] | 117532244 |  | 898208 |  | Closed - Non-Complaint | Closed - Non-Complaint | [14 empty] | LAN Config | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
892 |  |  | China | ZJ YUEQING PEOPLE HOSPITAL | 40054 | 4432700% | 19 | 02G/Ingenuity CT 728326/can't power on | T2 Activities:Problem Description: Summary of customer's problems: Unable to turn the machine on. Detailed description of Customer's problems or malfunctions: Unable to turn the machine on. It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: NA Remote troubleshoot and test result: The trainer was training on site to demonstrate the normal switch on and off. When turned on, he found that the machine could not be powered on. host has no key box. The customer had contacted the installation engineer. Under instruction, he had reboot several times but could not recover PRS TR status: Only M2M is not connected, ask FSE check onsite. Repair Action: Suggest onsite service procedures: Check customer operation to see if it's correct, check CAN communication related, such as main control board and power supply Memo of Customer request: The customer said they should have prepared for machine launch ceremony tomorrow. The situation is very urgent now and need an engineer to go to the site as soon as possible Next Step Action: Device is yet under warranty. Contacted MM. manager Yuan inform the site has yet been allocated to a FSE. Ask installation engineer to follow up first. from System query found IA#53515511. Dispatched ASP honglejia, mailed honglejia to follow up first ----------------------------------------------------------------------------- | 2 |  | ok after replace CCB | [12 empty] | 117529382 |  | 895802 |  | Evaluation | Closed - No Further Investigation |  | TBD_CouchEstop | [5 empty] | Engineering | [6 empty] | CCB |  |  | Completed |  |  | Refresh |  |  | System issue | System | Couch-SW | TBD-bugrep | Couch-SW-TBD-bugrep | Couch-SW-TBD-bugrep | TBD_CouchEstop | TBD_Bugreport | Couch-SW | Couch-SW-TBD-bugrep | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
893 |  |  | China | SD Laiyang People Hospital | 345005 | 4432700% |  | 02N/Ingenuity Core128/add the workstation | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer: 01/Access CT Safety Questions:Case: 0117527641 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Weng Jinming Last updated by: Weng Jinming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Need add workstation Troubleshooting Action: FSE_admin cannot login with password Philips! 1/Philips!/Philips1 Repair Action: Customer will try to confirm password and call back Test/Inspection: N/A Test Equipment: N/A SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | NAC | [12 empty] | 117527641 |  | 894459 |  |  | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
894 |  |  | China | SD Anqiu T.C.M Hosp. | 345017 | 4432700% |  | 02N/Core128/Rebuild the image failure | BU Q&R Feedback: No Reason: Corrective Maintenance Next Step Action : FSE Yang xingbo 013583235930 call in T2 Activities:Problem Description: Summary of customer's problems: Occasionally there are no images after scanning. Detailed description of Customer's problems or malfunctions: After scanning, there is raw data, but no reconstructed images. It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: N/A Remote troubleshoot and test result: N/A PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: The field service engineer asked to open escalation on a service order. We recommend spare parts 12NC and names: N/A Memo of Customer request: None Problem Reported by Customer: 02N/Core128/ Safety Questions:Case: 0117526548 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingsong Last updated by: Lu Mingsong 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support ----------------------------------------------------------------------------- | 5 |  | 扫描结束后有原始数据，没有重建出图像 工程师要求开单上ESCALATION | [12 empty] | 117526548 |  | 893672 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-NoEnoughInfo-TBD-NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Recon image |  |  | Completed |  |  | Refresh |  |  | System issue | System | TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
895 |  |  | China | NMG Zhengxiang Baiqi Hosp. | 345033 | 4432700% |  | 02N/Ingenuity Core128/Elite China/gantry can not power on | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure SAP Transfer : Case Transferred to SAP Problem Reported by Customer : 02N/Ingenuity Core128/Elite China Safety Questions : Case: 0117525846 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingsong Last updated by: Lu Mingsong 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: gantry can not power on Troubleshooting Action: Remotely instruct the customer to restart the system. Repair Action: Remotely instruct the customer to restart the system. Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: N/A ----------------------------------------------------------------------------- | 3 |  | Remotely instruct the customer to restart the system. Test/Inspection: Functional Check PASSED. Device returned to full functionality. | [12 empty] | 117525846 |  | 893210 |  | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Reboot |  |  | Completed |  |  | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
896 |  |  | China | Guiyang Fuyou Hosp | 40033 | 4432600% |  | 02N/Ingenuity CT 728326/Consult factory date | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer : Consult factory date Next Step Action : mail to IB 确认 Safety Questions : Case: 0117519675 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Qi Xiaoyu Last updated by: Qi Xiaoyu 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 3 |  | Consult factory date | [12 empty] | 117519675 |  | 888636 |  | Closed - Non-Complaint | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
897 |  |  | China | JL Songyuan Central Hosptial | 345008 | 4432600% |  | Application/Ingenuity Core128 | BU Q&R Feedback: No Reason: Customer Information Field Service Engineer Notes:This PO is a Reference Number only Problem Reported by Customer : Application Safety Questions:Case: 0117519599 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Pan Chuchu Last updated by: Pan Chuchu 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action: Installation Order IA#52827232 dispatched ASP Huicong, Jilin Province Songyuan City Central Hospital's Ingenuity Core128, the CT Room, Teacher Xia informed that training has not been done yet, now they need to apply for clinical training for the first time. Please ask Huicong to assist the customer with applying. CT Room, Teacher Xia 18904381395 ----------------------------------------------------------------------------- | 5 |  | apply for clinical training for the first time | [12 empty] | 117519599 |  | 888541 |  | Closed - Non-Complaint | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
898 |  |  | Japan | Maki Hospital | 341017 | 4433100% |  | 07CS:TM:Quality assureranceのメニューに入るとWarningメッセージが表示される>>Tiltが0.5度かかっていた為でした | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Application Support Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer : Quality assureranceのメニューに入るとWarningメッセージが表示される>>Tiltが0.5度かかっていた為でした T2 Activities : Problem Description: Problem description by engineer : 対象装置:CT 故障個所:07その他 最初に発生したのはいつか?(~以内):1 Hour 問題発生の頻度は?:First Occurrence エンドユーザｰのワークフローのどの時点において問題が発生しているのか?:検査・測定準備中(Setting up Protocol) 最近どのようなイベントが発生したか?:N.A. 顧客は問題について具体的なシステムイベントレポートを作成しているか?:N.A. Engineer's Comment: '============================== コール内容 Quality assureranceのメニューに入るとWarningメッセージが表示される '============================== Troubleshooting Action: What steps are taken during the system diagnosis : 対応後の結果:改善された 故障個所:07その他 診断中にどのようなステップを実施したか?: '================================== 対応 メッセージを読んでいただいたところTiltを0に戻してくださいというものでした。 Tiltが0.5度かかっていた為でしたので０にしていただきました '================================== Error # and/or description of error : NA Repair Action: Completed or Recommended Repair Action "External" : 【現象確認】 確認済み 【作業内容詳細】 質問内容を回答 【確認内容】 良好 OneEMS internal remarks : Problem description by engineer : 対象装置:CT 故障個所:07その他 最初に発生したのはいつか?(~以内):1 Hour 問題発生の頻度は?:First Occurrence エンドユーザｰのワークフローのどの時点において問題が発生しているのか?:検査・測定準備中(Setting up Protocol) 最近どのようなイベントが発生したか?:N.A. 顧客は問題について具体的なシステムイベントレポートを作成しているか?:N.A. Engineer's Comment: '============================== コール内容 Quality assureranceのメニューに入るとWarningメッセージが表示される '============================== What steps are taken during the system diagnosis : 対応後の結果:改善された 故障個所:07その他 診断中にどのようなステップを実施したか?: '================================== 対応 メッセージを読んでいただいたところTiltを0に戻してくださいというものでした。 Tiltが0.5度かかっていた為でしたので０にしていただきました '================================== Error # and/or description of error : NA Completed or Recommended Repair Action "External" : 【現象確認】 確認済み 【作業内容詳細】 質問内容を回答 【確認内容】 良好 Completed or Recommended Repair Action "Internal" : --Troubleshooting Action-- 対応後の結果:改善された 故障個所:07その他 診断中にどのようなステップを実施したか?: '================================== 対応 メッセージを読んでいただいたところTiltを0に戻してくださいというものでした。 Tiltが0.5度かかっていた為でしたので０にしていただきました '================================== --Repair Action-- --Subject-- 07CS:TM:ﾎｼ様:Quality assureranceのメニューに入るとWarningメッセージが表示される>>Tiltが0.5度かかっていた為でした --Problem Description-- 対象装置:CT 故障個所:07その他 最初に発生したのはいつか?(~以内):1 Hour 問題発生の頻度は?:First Occurrence エンドユーザｰのワークフローのどの時点において問題が発生しているのか?:検査・測定準備中(Setting up Protocol) 最近どのようなイベントが発生したか?:N.A. 顧客は問題について具体的なシステムイベントレポートを作成しているか?:N.A. Engineer's Comment: '============================== コール内容 Quality assureranceのメニューに入るとWarningメッセージが表示される '============================== --4/8Elements-- Q1 報告された問題は確認できましたか? 確認済み Q2 他の問題は確認できましたか? 無し Q3 問題を解決するために何をおこないましたか? 質問内容を回答 Q4 対応後の装置は順調ですか? 良好 Q5 その検査はどの様に終了しましたか? 無し Q6 問題により間違った所見（誤診、誤治療）の有無: 無し Q7 意図しない不要な曝射の有無: 無し Q8 アーチファクトの有無: 無し Safety Questions : Case: 0117548087 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: TAKEYUKI MIYAKAWA Last updated by: TAKEYUKI MIYAKAWA 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 5 |  | It was because Tilt took 0.5 degrees | [12 empty] | 117548087 |  | 907991 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | NotEngineering | [5 empty] | Engineering | [6 empty] | Tilt |  |  | Completed |  |  | Refresh |  |  | System issue | System | NotEngineering | ShouldbeCustomer | NotEngineering-ShouldbeCustomer | NotEngineering-ShouldbeCustomer | NotEngineering | NotEngineering-ShouldbeCustomer | NotEngineering | NotEngineering-ShouldbeCustomer | No | y | y | y | y | y | No | y | 0 | y | y | y | y | y | y | 0 | y | y
899 |  |  | Japan | Maki Hospital | 341017 | 4433000% |  | 07CS:TM:操作室からの声が伝わりづらい。良い方法は無いか>>担当FSEに内容を報告。またアプリ担当者より現状で考えられる対応方法を回答 | BU Q&R Feedback: No Reason: Application Support Problem Reported by Customer : 操作室からの声が伝わりづらい。マイク部分にフィルムを貼っているかの確認。良い方法が無いかの確認 Safety Questions : Case: 0117544211 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: TAKEYUKI MIYAKAWA Last updated by: TAKEYUKI MIYAKAWA 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: Problem description by engineer : 対象装置:CT 故障個所:07その他 最初に発生したのはいつか?(~以内):1 Hour 問題発生の頻度は?:Unknow エンドユーザｰのワークフローのどの時点において問題が発生しているのか?:その他(Other) 最近どのようなイベントが発生したか?:N.A. 顧客は問題について具体的なシステムイベントレポートを作成しているか?:N.A. Engineer's Comment: '============================== コール内容 操作室からの声が伝わりづらい。良い方法は無いか '============================== Troubleshooting Action: What steps are taken during the system diagnosis : 対応後の結果:改善された 故障個所:07その他 診断中にどのようなステップを実施したか?: '================================== 対応 担当FSEに内容を報告。 またアプリ担当者より現状で考えられる対応方法を回答 （マイクから少し距離をおいて話していただく、音量を調整いただく等） '================================== Error # and/or description of error : NA Repair Action: Completed or Recommended Repair Action "External" : 【現象確認】 確認済み 【作業内容詳細】 質問内容を回答 【確認内容】 良好 OneEMS internal remarks : Problem description by engineer : 対象装置:CT 故障個所:07その他 最初に発生したのはいつか?(~以内):1 Hour 問題発生の頻度は?:Unknow エンドユーザｰのワークフローのどの時点において問題が発生しているのか?:その他(Other) 最近どのようなイベントが発生したか?:N.A. 顧客は問題について具体的なシステムイベントレポートを作成しているか?:N.A. Engineer's Comment: '============================== コール内容 操作室からの声が伝わりづらい。良い方法は無いか '============================== What steps are taken during the system diagnosis : 対応後の結果:改善された 故障個所:07その他 診断中にどのようなステップを実施したか?: '================================== 対応 担当FSEに内容を報告。 またアプリ担当者より現状で考えられる対応方法を回答 （マイクから少し距離をおいて話していただく、音量を調整いただく等） '================================== Error # and/or description of error : NA Completed or Recommended Repair Action "External" : 【現象確認】 確認済み 【作業内容詳細】 質問内容を回答 【確認内容】 良好 Completed or Recommended Repair Action "Internal" : --Troubleshooting Action-- 対応後の結果:改善された 故障個所:07その他 診断中にどのようなステップを実施したか?: '================================== 対応 担当FSEに内容を報告。 またアプリ担当者より現状で考えられる対応方法を回答 （マイクから少し距離をおいて話していただく、音量を調整いただく等） '================================== --Repair Action-- --Subject-- 07CS:TM:ﾎｼ様:操作室からの声が伝わりづらい。良い方法は無いか>>担当FSEに内容を報告。またアプリ担当者より現状で考えられる対応方法を回答 --Problem Description-- 対象装置:CT 故障個所:07その他 最初に発生したのはいつか?(~以内):1 Hour 問題発生の頻度は?:Unknow エンドユーザｰのワークフローのどの時点において問題が発生しているのか?:その他(Other) 最近どのようなイベントが発生したか?:N.A. 顧客は問題について具体的なシステムイベントレポートを作成しているか?:N.A. Engineer's Comment: '============================== コール内容 操作室からの声が伝わりづらい。良い方法は無いか '============================== --4/8Elements-- Q1 報告された問題は確認できましたか? 確認済み Q2 他の問題は確認できましたか? 無し Q3 問題を解決するために何をおこないましたか? 質問内容を回答 Q4 対応後の装置は順調ですか? 良好 Q5 その検査はどの様に終了しましたか? 無し Q6 問題により間違った所見（誤診、誤治療）の有無: 無し Q7 意図しない不要な曝射の有無: 無し Q8 アーチファクトの有無: 無し ----------------------------------------------------------------------------- | 5 |  | Check if the film is stuck on the microphone part. Check if there is a good way | [12 empty] | 117544211 |  | 906473 |  | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Engineering | [6 empty] | FilmStuck |  |  | Completed |  |  | Refresh |  |  | System issue | System | Console-SW | FilmStuck | Console-SW-FilmStuck | Console-SW-FilmStuck | 0 | FilmStuck | Console-SW | Console-SW-FilmStuck | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
900 |  |  | Japan | Maki Hospital | 341017 | 4433000% |  | CT:APP:AK:ﾎｼ様:質問 | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Application Support Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer : Q:①頭部アキシャル撮影で、360°と420°があるかと思いますが他のユーザーさんはどちらが多いですか、また見た目で違うのが分かりますでしょうか？ ②16×0.625で撮影していますが40×0.625で撮影されている施設はありますか？ A:①撮影の撮影時間を考えると360°が多く選択されています。また肉眼で画像の大きな違いははっきりわからない程度と思います。 ②16列と40列で使い分けているご施設は有りました。通常は16列ですがOPE後は早く撮影し脳全体を確認すためだけに40列で撮影すると聞いたことがあります。 Safety Questions : Case: 0117544070 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: ATSUSHI KOIZUMI Last updated by: ATSUSHI KOIZUMI 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A OneEMS internal remarks : Q1:報告された問題は確認できましたか？：確認済 Q2：他の問題は確認できましたか？；無 Q3：問題を解決するために何を行いましたか？ １：電話で解決した場合：説明 Q4：対応後の装置は順調ですか？：良好 FSE Notes : This PO is a Reference Number only SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 5 |  | resolved by phone: Explanation | [12 empty] | 117544070 |  | 906407 |  | Closed - Non-Complaint | Closed - Non-Complaint | [14 empty] | QA | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
901 |  |  | Japan | Maki Hospital | 341017 | 4433000% |  | 04DM:KK:サジタル画像でZ方向に線状のアーチファクト。アキシャル画像でも感度ムラのようなアーチファクトがでた。 | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 04DM:KK:サジタル画像でZ方向に線状のアーチファクト。アキシャル画像でも感度ムラのようなアーチファクトがでた。 ----------------------------------------------------------------------------- | 4 |  | tbd | [12 empty] | 117544052 |  | 906396 |  | Investigation | Closed - Done |  | TBD-NoEnoughInfo-TBD-NoEnoughInfo | [5 empty] | Engineering | [6 empty] | linear artifact |  |  | Completed |  |  | Refresh |  |  | System issue | System | TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
902 |  |  | Japan | Higashi-Ohmi General Medical C | 341015 | 4432900% |  | HY: Cannot install the pediatric attachment. | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : HY: Cannot install the pediatric attachment. Safety Questions:Case: 0117539014 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Hui Yi Last updated by: Wang Hui Yi 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A OneEMS internal remarks: Yamamoto DB SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Support ----------------------------------------------------------------------------- | 5 |  | Cannot install the pediatric attachment. | [12 empty] | 117539014 |  | 902823 |  | Evaluation | Closed - No Further Investigation |  | NotEngineering | [5 empty] | Engineering | [6 empty] | Pediatric |  |  | Completed |  |  | Refresh |  |  | System issue | System | NotEngineering | ShouldbeClinical | NotEngineering-ShouldbeClinical | NotEngineering-ShouldbeClinical | NotEngineering | NotEngineering-ShouldbeClinical | NotEngineering | NotEngineering-ShouldbeClinical | No | y | y | y | y | y | No | y | 0 | y | y | y | y | y | y | 0 | y | y
903 |  |  | Japan | Maki Hospital | 341017 | 4432900% |  | 04CS:RS: A shower artifact occurred on head conventional scan images. >> The artifact has a point of converg [Truncated] | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 04CS:RS: A shower artifact occurred on head conventional scan images. >> Safety Questions:Case: 0117538815 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: RUI SATAKE Last updated by: RUI SATAKE 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A OneEMS internal remarks: --Troubleshooting Action-- Result of response: Problem solved. Problem location: 04 Gantry Exclude HV What steps did you perform to diagnose the problem? I asked him to perform re-reconstruction with an FOV of 500 mm, and when I checked the images, there was an artifact with a point of convergence at the left edge of the tabletop. The artifact is the kind that can occur with arcing, but there was no arcing recorded in the log. Also, given that when he performed another conventional scan, a shower artifact occurred in the same place, and he said that he had not repositioned the patient to perform a helical scan, the patient setting is the same in the FOV, but he returned the tilt to zero, so the X-ray irradiation line would have changed. Given the above, it is possible that the tabletop or an accessory such as an additional cushion caused the problem. I told Mr. Matsumoto the above findings and told him that we recommend the use of the head extension for head scans, and he agreed to monitor the situation. --Repair Action-- --Subject-- 04CS:RS:Mr. Matsumoto: A shower artifact occurred on head conventional scan images. >> The artifact has a point of convergence near the tabletop. --Problem Description-- Equipment affected: CT Problem location: 04 Gantry Exclude HV When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user's work flow did the problem occur?: During reconstruction/post processing What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem?: NA --4/8Elements-- Q1. Was the reported problem confirmed? Confirmed. Q2. Were any other problems confirmed? None. Q3. What did you do to resolve the problem?: I asked him to perform re-reconstruction with an FOV of 500 mm, and when I checked the images, there was an artifact with a point of convergence at the left edge of the tabletop. The artifact is the kind that can occur with arcing, but there was no arcing recorded in the log. Also, given that when he performed another conventional scan, a shower artifact occurred in the same place, and he said that he had not repositioned the patient to perform a helical scan, the patient setting is the same in the FOV, but he returned the tilt to zero, so the X-ray irradiation line would have changed. Given the above, it is possible that the tabletop or an accessory such as an additional cushion caused the problem. I told Mr. Matsumoto the above findings and told him that we recommend the use of the head extension for head scans, and he agreed to monitor the situation. Q4: Is the equipment operating normally following the response? Good. Q5. How did the examination finish? None. Q6. Were there any mistaken findings (misdiagnosis, mistaken treatment) due to the problem? None. Q7. Was there any unintended unnecessary exposure? None. Q8. Were there any artifacts? None. T2 Activities:Problem Description: subject equipment: CT Problem location: 04 Gantry Exclude HV When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user's work flow did the problem occur?: During reconstruction/post processing What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem?: NA Troubleshooting Action: Result after troubleshooting: Improved Problem location: 04 Gantry Exclude HV What steps did you perform to diagnose the problem? I asked him to perform re-reconstruction with an FOV of 500 mm, and when I checked the images, there was an artifact with a point of convergence at the left edge of the tabletop. The artifact is the kind that can occur with arcing, but there was no arcing recorded in the log. Also, given that when he performed another conventional scan, a shower artifact occurred in the same place, and he said that he had not repositioned the patient to perform a helical scan, the patient setting is the same in the FOV, but he returned the tilt to zero, so the X-ray irradiation line would have changed. Given the above, it is possible that the tabletop or an accessory such as an additional cushion caused the problem. I told Mr. Matsumoto the above findings and told him that we recommend the use of the head extension for head scans, and he agreed to monitor the situation. Repair Action: [Confirmation of problem] Confirmed. Work details: I asked him to perform re-reconstruction with an FOV of 500 mm, and when I checked the images, there was an artifact with a point of convergence at the left edge of the tabletop. The artifact is the kind that can occur with arcing, but there was no arcing recorded in the log. Also, given that when he performed another conventional scan, a shower artifact occurred in the same place, and he said that he had not repositioned the patient to perform a helical scan, the patient setting is the same in the FOV, but he returned the tilt to zero, so the X-ray irradiation line would have changed. Given the above, it is possible that the tabletop or an accessory such as an additional cushion caused the problem. I told Mr. Matsumoto the above findings and told him that we recommend the use of the head extension for head scans, and he agreed to monitor the situation. 【Confirmation content】 Good. ----------------------------------------------------------------------------- | 4 |  | Given the above, it is possible that the tabletop or an accessory such as an additional cushion caused the problem. I told Mr. Matsumoto the above findings and told him that we recommend the use of the head extension for head scans, and he agreed to monitor the situation. | [12 empty] | 117538815 |  | 902670 |  | Closed - Done | Closed - Done |  | NotEngineering | [5 empty] | Engineering | [6 empty] | Artifact |  |  | Completed |  |  | Refresh |  |  | System issue | System | NotEngineering | ShouldbeCustomer | NotEngineering-ShouldbeCustomer | NotEngineering-ShouldbeCustomer | NotEngineering | NotEngineering-ShouldbeCustomer | NotEngineering | NotEngineering-ShouldbeCustomer | No | y | y | y | y | y | No | y | 0 | y | y | y | y | y | y | 0 | y | y
904 |  |  | Japan | Saiseikai Gosu Hospital | 343000 | 4432700% |  | HY:ガントリーベルト交換 | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : HY:ガントリーベルト交換 Safety Questions : Case: 0117526745 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Hui Yi Last updated by: Wang Hui Yi 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A OneEMS internal remarks : 河重DB SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support ----------------------------------------------------------------------------- | 5 |  | TBD | [12 empty] | 117526745 |  | 893816 |  | Evaluation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Not clear |  |  | Completed |  |  | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
905 |  |  | Japan | Maki Hospital | 341017 | 4432600% |  | CT:APP:NO: Mr. Hoshi: Inquiry. | BU Q&R Feedback: No Reason: Application Support Problem Reported by Customer : Q: Thin slice data from a head axial scan was split into two series. A: It is possible to merge the split series data using the CT Viewer and create cor/sag images. ----------------------------------------------------------------------------- | 5 |  | Kevin investigation | [12 empty] | 117520965 |  | 889502 |  | Evaluation | Closed - non-complaint | [7 empty] | TBD | [6 empty] | conventional |  |  | Completed |  |  | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
906 |  |  | Japan | Maki Hospital | 341017 | 4432700% |  | In conventional head exam, when thin slice data read into CT Viewer, the data is displayed separately. |  | WANG, Peter |  | Kevin investigation | [14 empty] | 889106 |  | Investigation | Closed - Done | [7 empty] | TBD | [6 empty] | conventional |  |  | Completed |  |  | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
907 |  |  | Poland | Agencja Rezerw Materiałowych | 341010 | 4433000% |  | Artefakty kolorowe | Work Order #: WO-03494007 Robert Kamola Diagnostic performed by Engineer : N/A Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: ----------------------------------------------------------------------------- Work Order #: WO-03495119 Robert Kamola Diagnostic performed by Engineer : N/A Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: ----------------------------------------------------------------------------- | 4 |  | DMS replace | [12 empty] | 117544532 |  | 906631 |  | Closed - Done | Closed - Done |  | CTUDT00632110 | [5 empty] | Engineering | [6 empty] | Artifact |  |  | Completed |  |  | Refresh |  |  | System issue | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | DMS Module | DMS-HW | DMS-HW-DMS Module | No | y | y | y | y | y | No | y | 0 | y | y | y | y | y | y | 0 | y | y
908 |  |  | Poland | Agencja Rezerw Materiałowych | 341010 | 4433000% |  | Podpięcie strzykawki | Work Order #: WO-03494014 Robert Kamola Diagnostic performed by Engineer : N/A Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: ----------------------------------------------------------------------------- | 4 |  | Attaching the syringe | [12 empty] | 117544535 |  | 906632 |  | Closed - Non-Complaint | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
909 |  |  | Poland | Uniwersytecki Szpital Kliniczny | 341001 | 4432700% |  | Configuration of the DVD robot. | Work Order #: WO-03475733 14-May-2021 Marcin Knyps The system fully meets specification and returned to use Problem Description by Engineer: Change of device configuration, adding a recorder. Resolution: Adding a recorder node. 172.16.18.67 AROBOT1 6001. Recording test - correct result. A change of configuration ended successfully. ----------------------------------------------------------------------------- | 4 |  | Change of device configuration, adding a recorder. | [12 empty] | 117527848 |  | 894609 |  | Closed - Non-Complaint | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
910 |  |  | Portugal | Centro Hospitalar de Vila Nova | 40030 | 4432600% |  | Checking the couch cover on the left side of the equipment |  | 3 |  | Checking the table cover on the left side of the equipment | [12 empty] | 117522287 |  | 890531 |  | Evaluation | Evaluation |  | TBD-CouchCover | [5 empty] | Engineering | [6 empty] | Table cover |  |  | Completed |  |  | Refresh |  |  | System issue | System | Couch-HW | Noise | Couch-HW-Noise | Couch-HW-Noise | TBD-CouchCover | Noise | Couch-HW | Couch-HW-Noise | No | y | y | y | y | y | No | y | 0 | y | y | y | y | y | y | 0 | y | y
911 |  |  | Sweden | Ludvika Lasarett | 40049 | 4432800% |  | SE5333 - Long time to reconstruct an investigation. |  | 3 |  | Customer will call again if the issue is to reoccur. | [12 empty] | 117534933 |  | 899920 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-NoEnoughInfo-TBD-NoEnoughInfo | [5 empty] | Engineering | [6 empty] | long recon time |  |  | Completed |  |  | Refresh |  |  | System issue | System | TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
912 |  |  | United States | St Barnabas Hospital | 40028 | 4433100% | 19 | 0711CFXXX cannot transfer images to the PACS | Problem Description by Engineer : Problem description by engineer : cannot transfer images to the PACS When was the first occurrence? Within... : 15-May-2021 Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: No further information is available as per the customer Patient Impact/Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer Current Software Version: Not provided Resolution : Advised to reset the spooler and the PACS Diagnostic performed by Engineer : advised the customer to reset the spooler in Olevel Internal Comments : Error # and/or description of error : cannot transfer images to the PACS Completed or Recommended Repair Action "Internal" : advised the customer to reset the spooler in Olevel, follow up a few hours later and there was no answer. | 2 |  | reset the spooler, Suggest to install HF3 (FCO72800740). | [12 empty] | 117548905 |  | 907654 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS Transfer |  |  | Completed |  |  | Refresh |  |  | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
913 |  |  | United States | St Barnabas Hospital | 40028 | 4433000% |  | 3D cell greyed out when saving as batch | Work Order #: WO-03496169 Hirchak, Michele The system fully meets specification and returned to use Problem description by Clinical Support Specialist : Customer Function/Role: Tech Device Use: In Use Expected and Unexpected Results: Save a 3D image. Patient Impact: None Current Software Version: 4. No rescan, no loss of data, no reinjection. Diagnostic performed by Engineer : Customer already figured it out by the time the call got transferred to Apps. She is all set. MBH ----------------------------------------------------------------------------- | 3 |  | Customer already figured it out by the time the call got transferred to Apps. She is all set. MBH | [12 empty] | 117546527 |  | 904454 |  | Investigation | Closed - Done |  | NotEngineering | [5 empty] | Engineering | [6 empty] | 3D |  |  | Completed |  |  | Refresh |  |  | System issue | System | NotEngineering | ShouldbeCustomer | NotEngineering-ShouldbeCustomer | NotEngineering-ShouldbeCustomer | NotEngineering | NotEngineering-ShouldbeCustomer | NotEngineering | NotEngineering-ShouldbeCustomer | No | y | y | y | y | y | No | y | 0 | y | y | y | y | y | y | 0 | y | y
914 |  |  | United States | Unity Health Systems | 40037 | 4432800% |  | Monthly QA for JUne | Work Order #: WO-03483802 Lockwood Thomas Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 5 |  | NAC | [12 empty] | 117535034 |  | 899978 |  | Closed - Non-Complaint | Closed - Non-Complaint | [14 empty] | QA | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
915 |  |  | United States | Unity Health Systems | 40037 | 4432800% |  | Monthly QA for May and June | Work Order #: WO-03483746 Lockwood Thomas Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 5 |  | NAC | [12 empty] | 117534992 |  | 899957 |  | Closed - Non-Complaint | Closed - Non-Complaint | [14 empty] | QA | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
916 |  |  | United States | The Regional Medical Center | 40005 | 4432800% | 19 | 0614CDY73 Gantry will not come on | Work Order #: WO-03485838 Maier, Donnie The system does not meet specification and is out of use Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: Diag Expected and Unexpected Results: Cant close estop User Impact: No Patient Impact: No Current Software Version: 4.1.10 Diagnostic performed by Engineer : Checked ULV to see following errors. 2021/05/12 14:36:32.132 S_TILT_C_POT_CALIB_KO ESTOP 1.344502E+08 0 0 0 0.00146990740740741 0.0352777777777778 2.11666666666667 2021/05/12 14:36:31.945 S_TILT_C_POT_CALIB_KO WARN 511 90 0 3997697 0.00145833333333333 0.035 2.1 PFSE Doug Holt was onsite and I conferenced in my colleague on NSS side (Don Parr). After having PFSE check connections to pot and belt tension still having issues with calibrating. This is a reoccurring intermittent issue from a couple days ago so new pot assembly was already onsite and PFSE will try installing and calibrating. Will call back if needs further assistance. Follow Up Required : Install new pot assembly and try to recalibrate. Pot assembly is still onsite from previous closed case from a couple days ago. ----------------------------------------------------------------------------- Work Order #: WO-03486157 12-May-2021 Holt Douglas The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: Diag Expected and Unexpected Results: Cant close estop User Impact: No Patient Impact: No Current Software Version: 4.1.10 Tools & Equipment : MULTIMETER,2022-08-30 00:00:00 Follow Up Required : Verify system operation Resolution : Replaced dual pot assembly and calibrated the tilt ----------------------------------------------------------------------------- | 2 |  | Replaced dual pot assembly and calibrated the tilt | [12 empty] | 117536646 |  | 897217 |  | Closed - Done | Closed - Done |  | ECR-077012 | [5 empty] | Engineering | [6 empty] | Dual pot |  |  | Completed |  |  | Refresh |  |  | System issue | System | Gantry-HW | Belt | Gantry-HW-Belt | Gantry-HW-Belt | ECR-077012 | Belt | Gantry-HW | Gantry-HW-Belt | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
917 |  |  | United States | McLaren Macomb | 40021 | 4432800% |  | 4.1.1 CTA 's are intermittently not combining results | Work Order #: WO-03484293 Giles Brian The system fully meets specification and returned to use Problem description by Clinical Support Specialist : Customer Function/Role: tech Device Use: In Use Expected and Unexpected Results: CTA 's are intermittently not combining results User Impact: None Patient Impact: None Current Software Version: 4.1.1 Diagnostic performed by Engineer : I followed up with Michael and left a voicemail. that I will be closing the case Diagnostic performed by Engineer : I brought Michele Hrchack (RAS) on board. We had customer look at parameters page. X/Y was off by51mm. Tried recon and it didn't work. Then we went into CT Viewer and had them merge the disconnected series. That did work and customer was happy. | 5 |  | I brought Michele Hrchack (RAS) on board. We had customer look at parameters page. X/Y was off by51mm. Tried recon and it didn't work. Then we went into CT Viewer and had them merge the disconnected series. That did work and customer was happy. | [12 empty] | 117535382 |  | 896220 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | ShouldbeClinical | [5 empty] | Engineering | [6 empty] | CTA |  |  | Completed |  |  | Refresh |  |  | System issue | System | Console-SW | CTA | Console-SW-CTA | Console-SW-CTA | ShouldbeClinical | CTA | Console-SW | Console-SW-CTA | No | y | y | y | y | y | No | y | 0 | y | y | y | y | y | y | 0 | y | y
918 |  |  | United States | McLaren Macomb | 40021 | 4432700% |  | Getting tube heat message on scans with bolus tracking | Work Order #: WO-03477043 Hirchak, Michele The system fully meets specification and returned to use Diagnostic performed by Engineer : Customer getting tube heat message on scans with bolus tracking. Had her change from 30 trackers to 20. Also, turned off AST and changed the RT to 0.75. Talked about changing the time on the fly if they get the message. Looked at 100kV PE study / Abd Pelvis and the 120kV PE with Abd/Pelvis. Keeping case open. MBH Problem description by Clinical Support Specialist : Called to f/u with Allison. She was busy. Resolution : Spoke to Allison. Tube heat messages have stopped. She is all set. OK to close. MBH Problem description by Clinical Support Specialist : Customer Function/Role: Tech Device Use: In Use Expected and Unexpected Results: Tube heat on scans with bolus tracking. Patient Impact: None Current Software Version: 4.1.10 No rescan, no loss of data, no reinjection | 5 |  | Customer getting tube heat message on scans with bolus tracking. Had her change from 30 trackers to 20. Also, turned off AST and changed the RT to 0.75. | [12 empty] | 117528731 |  | 890964 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | Tube Heat | [5 empty] | Engineering | [6 empty] | Tube Overheat |  |  | Completed |  |  | Refresh |  |  | System issue | System | Gantry-SW | Tube Heat | Gantry-SW-Tube Heat | Gantry-SW-Tube Heat | Tube Heat | Tube Heat | Gantry-SW | Gantry-SW-Tube Heat | No | y | y | y | y | y | No | y | 0 | y | y | y | y | y | y | 0 | y | y
919 |  |  | United States | Advent Health Wesley Chapel | 300099 | 4433000% |  | 0711CDX-Console goes black when going to next series of scan | Console goes black when going to next series of scan - Have to start all over when this happens | [15 empty] | 117547277 |  | 905109 |  | Investigation | Opened |  | BlackScreen | [5 empty] | Engineering | [6 empty] | Black screen |  |  | Complete |  |  | Refresh |  |  | Upgrade issue | System | Console-HW | BlackScreen | Console-HW-BlackScreen | Console-HW-BlackScreen | BlackScreen | BlackScreen | Console-HW | Console-HW-BlackScreen | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
920 |  |  | United States | Advent Health Wesley Chapel | 300099 | 4432400% |  | 0705CDN84 customer locked out of system, multiple issues | customer locked out of system, multiple issues |  |  | Single case. Replace Video card & Keyboard | [14 empty] | 900856 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | NotEngineering-partreplace | [5 empty] | Engineering | [6 empty] | Video card & Keyboard |  |  | Complete |  |  | Refresh |  |  | Upgrade issue | System | Console-HW | Video card & Keyboard | Console-HW-Video card & Keyboard | Console-HW-Video card & Keyboard | NotEngineering-partreplace | Video card & Keyboard | Console-HW | Console-HW-Video card & Keyboard | No | y | y | y | y | y | No | y | 0 | y | y | y | y | y | y | 0 | y | y
921 |  |  | United States | Advent Health Wesley Chapel | 300099 | 4433000% |  | 0711CDX DISPLAYS SYSTEM ERROR WHILE PLOTTING FOR SCAN AFTER SURVIEW | DISPLAYS SYSTEM ERROR WHILE PLOTTING FOR SCAN AFTER SURVIEW | [15 empty] | 117546481 |  | 904438 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | BlackScreen | [5 empty] | Engineering | [6 empty] | Black screen |  |  | Complete |  |  | Refresh |  |  | Upgrade issue | System | Console-HW | BlackScreen | Console-HW-BlackScreen | Console-HW-BlackScreen | BlackScreen | BlackScreen | Console-HW | Console-HW-BlackScreen | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
922 |  |  | United States | Advent Health Wesley Chapel | 300099 | 4432200% |  | Lost system functionality post iPatient S upgrade (v4.1.7 to v4.1.10) - unable to utilize Same Series tool | •Feedback Type: Technical Issue •Other Details: NA •Submitter Email: melissa.lyter@philips.com •CFM Form Received Date (Outlook email received date): 05/12/2021 Lost system functionality post iPatient S upgrade- Ability to utilize Same Series tool On the v4.1.7 software, when the customer wanted to save multiple individual ROI images in the calcium scoring application or save multiple individual images in the brain perfusion application, they would use the Same Series option. This option is no longer available when using the same workflow wtih v.4.1.10. The customer's workflow was to save multiple individual screenshots for calcium scoring and brain perfusion studies. They would use the save image or save display option to do this. When labeling the new image series, as long as they kept the description name the same, the Same series option would appear already checked to keep the images in the same series. Now in v4.1.10, the Same series function no longer appears even though they keep the description name the same. This means when saving this way the images are save separately. This seems to apply to all post-processing applications. Yes the system was in clinical use at the time. The problem occurred during diagnosis. The procedure was continued. The issue occurs every time they try to save multiple images. No impact to the patient. The customer has to use the workaround of saving the individual images to the Film application. From the Filming application, they can choose to save all the images together in the same series. |  |  | Check the behavior on 4.1.10 -- Kevin | [14 empty] | 894938 |  | Investigation | Investigation | [7 empty] | Customer | [6 empty] | Refresh diff |  |  | Complete |  |  | Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
923 |  |  | United States | Eisenhower Medical Center | 68116 | 4432800% |  | Table is screaching in in and out | Table is screaching in in and out | [15 empty] | 117532201 |  | 896762 |  | #N/A | Evaluation | [14 empty] | Table noise | [5 empty] | Non-Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | #REF!
924 |  |  | United States | Advent Health Wesley Chapel | 300106 | 4430500% |  | After a 4.1.10 upgrade Interventional/Biopsy's are giving incorrect coordinates | •Feedback Type: Technical Issue •Other Details: NA •Submitter Email: David.Mujica@philips.com •CFM Form Received Date (Outlook email received date): 04/21/2021 After a 4.1.10 upgrade over this past weekend 4/7-4/18, Interventional/Biopsy's are giving incorrect coordinates as described by CT Technologist Russ Golden. Once it was realized that the Interventional/Biopsy was not working as intended all study's were stopped and placed on hold and it was reported by the customer to CT Remote Clinical Support and to the Remote Service Engineering group. The device was being used to perform a Biopsy on a patient and was in clinical use. Biopsy's are frequently performed at this site. There was no negative impact to the user or patient as reported by the customer. The resolve was to create a case and dispatch the Field Service Engineer to fix the problem on site. The device being used was an Ingenuity CT system 4.1.10. Install Product . Reason for delay: Missing event date |  |  | CTUDT00632916 IS CREATED TO COVER CCT wrong position issue | [14 empty] | 905009 |  | Investigation | Opened |  | CTUDT00632916 | [5 empty] | Engineering | [6 empty] | Biopsy |  |  | Complete |  |  | Refresh |  |  | Upgrade issue | System | Console-SW | Biopsy | Console-SW-Biopsy | Console-SW-Biopsy | CTUDT00632916 | Biopsy | Console-SW | Console-SW-Biopsy | No | y | y | y | y | y | No | y | 0 | y | y | y | y | y | y | 0 | y | y
925 |  |  | United States | Advent Health Wesley Chapel | 300106 | 4432200% |  | Lost system functionality post iPatient S upgrade (v4.1.7 to v4.1.10) - unable to utilize Same Series tool | •Feedback Type: Technical Issue •Other Details: NA •Submitter Email: melissa.lyter@philips.com •CFM Form Received Date (Outlook email received date): 05/12/2021 Lost system functionality post iPatient S upgrade- Ability to utilize Same Series tool On the v4.1.7 software, when the customer wanted to save multiple individual ROI images in the calcium scoring application or save multiple individual images in the brain perfusion application, they would use the Same Series option. This option is no longer available when using the same workflow wtih v.4.1.10. The customer's workflow was to save multiple individual screenshots for calcium scoring and brain perfusion studies. They would use the save image or save display option to do this. When labeling the new image series, as long as they kept the description name the same, the Same series option would appear already checked to keep the images in the same series. Now in v4.1.10, the Same series function no longer appears even though they keep the description name the same. This means when saving this way the images are save separately. This seems to apply to all post-processing applications. Yes the system was in clinical use at the time. The problem occurred during diagnosis. The procedure was continued. The issue occurs every time they try to save multiple images. No impact to the patient. The customer has to use the workaround of saving the individual images to the Film application. From the Filming application, they can choose to save all the images together in the same series. |  |  | Check the behavior on 4.1.10 -- Kevin | [14 empty] | 894943 |  | Investigation | Investigation | [7 empty] | Customer | [6 empty] | Refresh diff |  |  | Complete |  |  | Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
926 |  |  | Australia | Maia Financial Pty Limited | 341000 | 4433400% | 20 | Scanner cant initialize. | Work Order #: WO-03510806 Ian Fishwick The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: CT Radiographer Device Use: Initialise CT Scan Expected and Unexpected Results: Site indicated that they could not initialise a CT Spine Scan. User Impact: No Patient Impact: None Reported Current Software Version:4.1.10 Diagnostic performed by Engineer : _______________________________________________________________ PROBLEM DESCRIPTION Site indicated that they could not initialise a CT Spine Scan. ________________________________________________________________ DATE THE ISSUE OCCURRED: 19 May 2021 ________________________________________________________________ TROUBLESHOOTING: System SW Version is: 4.1.10 I spoke to CT Radiographer Elizabeth and she has rebooted the system including the CT Host and Reconstructor plus Power cycle of the Gantry, after this action the system completed the Short Tube Conditioning and also a test scan without error. ________________________________________________________________ LOG REVIEW: A Bugreport was saved the previous evening when the original issue occurred, this will be downloaded and reviewed to understand the original issue. Review of Bugreport indicates that the system failed to initialise a Scan at 20:12:14.463 due to the following error. 2021/05/18 20:12:15.021 S_DMC_UHR_COMB_TMOT ERROR_WARNING Description: UHR movement Time Out The previous series and this series are not programed as UHR so UHR should not be in motion. There are also at same time reported from the CIRS an Optical Datapath error. and information in this error indicates that the Optical Channel link was not known reporting 0x0 instead of 0x4321 Error #12 - CheckGantryConnection: ACQ board or Gantry-to-CIRS datapath failure. Gantry Power Cycle has cleared both errors. Review of Gantry on time at time of the event indicates that last Gantry Power Cycle was 53 days ago, Philips Recommends gantry Power cycle once in a 7 day period. I have requested that Philips Engineer attending next Scheduled system PM on the 6th June 2021 to run the sites CT team through the complete User maintenance procedures including CT Host reboot daily, CIRS reboot once per week and Gantry Power Cycle once per week. ________________________________________________________________ CUSTOMER UPDATE: Site have advised that after the Gantry Power cycle that they have been able to complete all new CT Scans for the Clinical Scanning day 19th May 2021, the site will continue to monitor and report any further issues. ________________________________________________________________ RECOMMENDATION: Monitoring is recommended at this stage, further investigation will be required should this issue recur. If you have any further questions don’t hesitate to contact me. Ian Fishwick Philips ANZ Remote Service Engineer. Email, ian.fishwick@philips.com. You may receive a short survey re our service today, and your feedback is gratefully used to drive continuous improvement in our service delivery. ________________________________________________________________ CURRENT SYSTEM STATUS: System is in clinical use. ________________________________________________________________ ----------------------------------------------------------------------------- | 2 |  | RECOMMENDATION: Monitoring is recommended at this stage, further investigation will be required should this issue recur. If you have any further questions don’t hesitate to contact me. | [12 empty] | 117559515 |  | 915891 |  | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Power cycle |  |  | Completed |  |  | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
927 |  |  | Chile | Scotiabank Chile | 40008 | 4433500% |  | CL1675-ON SCANNER SCREEN SHOWS A FLASHING ICON. WHEN TAKING AN EXAM, THE SCREEN SHOWS A CT ICON LIKE FLASHING AND NOT CORRECT, IT MAY BE ERRATIC OPERATION | Problem description by engineer: Low imaging disk space. Information to support the complaint handling process Customer Function / Role: BioMed Device Use: Diagnostic Expected and Unexpected results: No disk space errors. Patient Impact / Outcome: Delay in exams. User Impact: Delay in exams. Current Software Version: 4.1.10 It presents 1% of available space in Local disk of images. There is a group of event log files that are taking up a lot of space in the d: \ Pms \ Log \ Data folder. The files are erased and the available disk capacity is good. The client is requested to restart the computer daily. Error # and / or description of error: Space available in local dioco 1%. | 4 |  | Delete event log files that are taking up a lot of space in the d: \ Pms \ Log \ Data folder. The files are erased and the available disk capacity is good. | [12 empty] | 117568913 |  | 924783 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00633269 | [5 empty] | Engineering | [6 empty] | Low disk |  |  | Completed |  |  | Refresh |  |  | System issue | System | Console-SW | DiskFull | Console-SW-DiskFull | Console-SW-DiskFull | CTUDT00632261 | DiskFull | Console-SW | Console-SW-DiskFull | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
928 |  |  | China | 山东省泰安市中医二院 | 345037 | 4433800% |  | 02N/Ingenuity Core128/Elite China/add printer | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ Safety Questions : Case: 0117583119 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Qian Wang Last updated by: Qian Wang 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 22.05.2021 13:50:55 UTC Tao Chen (86038373)故障现象: 添加打印机 服务内容: 设置打印机 服务结果: 维修后,设备恢复正常. Next Step Action : chen tao Internal Remarks : 22.05.2021 13:50:58 UTC Tao Chen (86038373) “RDI Accuracy: YES" $Q7: (Confirmed reported problem, add printer) $Q8: ( Get IP, port and AE from hospital IT) $Q9: (Add new DICOM printer) $Q10: ( Test and inspection data refer to: PA. Device returned to full functionality) $Q11: Calibrated Tools name ( N/A) ^Current Tube installed date: N/A ^Current Tube usage: N/A ----------------------------------------------------------------------------- | 5 |  | add printer | [12 empty] | 117583119 |  | 933175 |  | Closed - Non-Complaint | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
929 |  |  | China | 山东安丘市中医医院 | 345017 | 4433700% |  | 02NIngenuity Core128/Elite China/image transfer error | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : Ingenuity Core128/Elite China SAP Transfer : Case Transferred to SAP Next Step Action : FSE Yang xingbo 013583235930来电 Safety Questions : Case: 0117578656 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Jianqiang Last updated by: Wang Jianqiang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 图像传输慢 客户问题详细描述或故障现象： 图像传输慢 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 图像传输慢 远程诊断与测试结果： 现场诊断 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 检查网络，检查软件 SAP Transfer : Dispatched for Onsite Support ----------------------------------------------------------------------------- | 5 |  | Intermittent problem. After the exam is completed, the patient image cannot be automatically transmitted to the PACS system and needs to be sent manually. Escalation 0117578777 raised | [12 empty] | 117578656 |  | 931413 |  | Evaluation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS Transfer |  |  | Completed |  |  | Refresh |  |  | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
930 |  |  | China | 陕西省咸阳中心医院 (咸阳市第二人民医院） | 40050 | 4433700% |  | Application/Ingenuity CT 728326 | BU Q&R Feedback: No Reason: Application Support Problem Reported by Customer : Application/Ingenuity CT 728326 Next Step Action : 客户要求换个培训老师再次培训，已反馈售前培训经理跟进 Safety Questions : Case: 0117577477 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Bai Jingwen Last updated by: Bai Jingwen 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 5 |  | Application training | [12 empty] | 117577477 |  | 930780 |  | Evaluation | Closed - non-complaint | [7 empty] | TBD | [6 empty] | NAC |  |  | Completed |  |  | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
931 |  |  | China | 济南市第四人民医院 | 40038 | 4433700% |  | 02N/Ingenuity CT 728326/filming lost title | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity CT 728326/ T2 Activities : Problem Description: 客户问题概述： 其他类型故障。 客户问题详细描述或故障现象： 主机打印图像时会丢扫描时间title。 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： log无报错。 远程诊断与测试结果： 重启无效。 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议互换扫描时间title，更改字体大小间距排除问题，检查重建柜时区和时间设置。 建议备件12NC及名称： N/A Safety Questions : Case: 0117576901 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Bo Last updated by: Wang Bo 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 21.05.2021 14:40:11 UTC Tao Chen (86038373)故障现象: Film图像上间歇性无患者信息，Film图像间距大 服务内容: 设置layout为全页打印，修改打印排版 服务结果: 维修后,设备恢复正常. Internal Remarks : 22.05.2021 14:20:31 UTC Tao Chen (86038373) “RDI Accuracy: YES" $Q7: (Confirmed reported problem, no patient film title on CT images intermittently) $Q8: ( Check patient film process, found film layout page was single when do film printing) $Q9: ( Adjust film page layout to full page) $Q10: ( Test and inspection data refer to: PA. Device returned to full functionality) $Q11: Calibrated Tools name ( N/A) ^Current Tube installed date: N/A ^Current Tube usage: N/A ----------------------------------------------------------------------------- | 5 |  | Check patient film process, found film layout page was single when do film printing) $Q9: ( Adjust film page layout to full page) | [12 empty] | 117576901 |  | 930474 |  | Closed - Done | Closed - Done | [7 empty] | Application | [6 empty] | Image title |  |  | Completed |  |  | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
932 |  |  | China | 莱阳市人民医院 | 345005 | 4433700% |  | 02N/Ingenuity Core128/Rebuild the image failure | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer : 02N/Ingenuity Core128/ Safety Questions : Case: 0117576807 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingsong Last updated by: Lu Mingsong 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: Rebuild the image failure Troubleshooting Action: Remotely instruct the customer to restart the system. Repair Action: Remotely instruct the customer to restart the system. Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- |  |  | Remotely instruct the customer to restart the system. | [12 empty] | 117576807 |  | 930400 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-NoEnoughInfo-TBD-NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Recon |  |  | Completed |  |  | Refresh |  |  | System issue | System | TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
933 |  |  | China | 四川华蓥市华蓥山广能集团总医院 | 345000 | 4433600% |  | 02N/Ingenuity Core128/Elite China/Password expired | BU Q&R Feedback: No SAP Customer Symptom Code: HANG Customer Symptom Code: Hang/Crash Reason: Corrective Maintenance Engineer Failure Code: Configuration Error Repair Activity Code: Configuration Change Problem Reported by Customer : 02N/Ingenuity Core128/Elite China Safety Questions : Case: 0117572025 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Bo Last updated by: Wang Bo 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: host software password expired. Troubleshooting Action: Assist customer to reset password Repair Action: Assist customer to reset password Test/Inspection: customer confirm can scan patient and image ok.functional Check PASSED. Device returned to full functionality. Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- |  |  | Assist customer to reset password | [12 empty] | 117572025 |  | 926787 |  | Closed - Non-Complaint | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
934 |  |  | China | 河南省濮阳市人民医院 | 40031 | 4433600% |  | 02N/Ingenuity CT 728326/Estop open | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity CT 728326/ T2 Activities : Problem Description: 客户问题概述： 机架倾斜运动故障。 客户问题详细描述或故障现象： 扫描打角度时掉电。 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： S_TILT_C_NO_MOVE 远程诊断与测试结果： 重启恢复。diffPOT12 =-1 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议检查pot皮带张紧度，滑轮六角螺丝固定，limit switch。必要时更换pot并做tilt calibration。 建议备件12NC及名称： PPD 459800441121 SAP Transfer : Case Transferred to SAP Internal Remarks : 23.05.2021 07:34:01 UTC Yongping Luo (86038950) $Q7:Confirmed report problem，the gantry maybe lose power and estop open when tilt movement. $Q8:Perform visiual check on site ,replace the gantry tilt potentiometer, device return to fully function. $Q9:replace the tilt potentiometer,Device return to full function. $Q10: test and inspection data refer to PA. Device returned to full functionality. $Q11:Calibrated Tools refer to PA. External Remarks : 23.05.2021 07:33:58 UTC SUP JCOUSER (86038950)服务内容：机架倾斜时，偶发掉电，无法扫描，更换tilt 电位器。 服务结果：更换tilt电位器，设备恢复正常。 其它重要事项：无 Safety Questions : Case: 0117570465 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Bo Last updated by: Wang Bo 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Dispatched for Onsite Support Next Step Action : PPD 459800441121*1 已安排24H到医院 罗永平-86-15226131133 ----------------------------------------------------------------------------- | 5 |  | Delete event log files that are taking up a lot of space in the d: \ Pms \ Log \ Data folder. The files are erased and the available disk capacity is good. | [12 empty] | 117570465 |  | 925749 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | ECR-077012 | [5 empty] | Engineering | [6 empty] | Dual pot |  |  | Completed |  |  | Refresh |  |  | System issue | System | Gantry-HW | Belt | Gantry-HW-Belt | Gantry-HW-Belt | ECR-077012 | Belt | Gantry-HW | Gantry-HW-Belt | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
935 |  |  | China | 山东省泰安市中医二院 | 345037 | 4433600% |  | 02N/Ingenuity Core128/Elite China/workaround for yellow screen | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Calibration/Adjustment Problem Reported by Customer : 02N/Ingenuity Core128/Elite China Safety Questions : Case: 0117569866 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Gaoge Last updated by: Wang Gaoge 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action : Q6 = workaround for yellow screen SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support T2 Activities : Problem Description: 客户问题概述： 计算机软件相关故障。 客户问题详细描述或故障现象： 计算机主机黄屏，间歇性故障 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 主机黄屏，按钮无反应 远程诊断与测试结果： 软件defect PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 1，执行软件升级 Test/Inspection: N/A Test Equipment: N/A ----------------------------------------------------------------------------- |  |  | Yellow screen workaround installation | [12 empty] | 117569866 |  | 925522 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632520 | [5 empty] | Engineering | [6 empty] | New Yellow Screen |  |  | Completed |  |  | Refresh |  |  | System issue | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #REF!
936 |  |  | China | 吉林省松原市中心医院 | 345008 | 4433500% |  | 02N/Ingenuity Core128/Elite China/transfer error | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity Core128/Elite China SAP Transfer : Case Transferred to SAP T2 Activities : Problem Description: 客户问题概述： 图像传输失败。 客户问题详细描述或故障现象： 图像老师传输一半之后就不传了 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 传输故障 远程诊断与测试结果： 传输问题，客户要求工程师到场 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 检查网络速度,检查软件 Safety Questions : Case: 0117564872 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Jianqiang Last updated by: Wang Jianqiang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Dispatched for Onsite Support External Remarks : 19.05.2021 14:17:42 UTC Xiaodong Chi (86038306)服务内容： 图像无法传输到pacs，重新启动主机。 服务结果： 维修后，设备恢复正常工作。 其它重要事项： 无。 Internal Remarks : 19.05.2021 14:17:45 UTC Xiaodong Chi (86038306) "RDI Accuracy:No" $Q7: (Confirmed reported problem&the image can not transfer to pacs ) $Q8: (review the system logger，the device need be restart） $Q9: (restart the device) $Q10: (Test and inspection data refer to: PA. Device returned to full functionality) $Q11 Calibrated Tools name (FLUKE287) S/N (16490094) Next Calibration date (10/19/2021) ----------------------------------------------------------------------------- | Gao, Mia |  | review the system logger，the device need be restart） restart the device) | [12 empty] | 117564872 |  | 921576 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS Transfer |  |  | Completed |  |  | Refresh |  |  | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
937 |  |  | China | 江苏省第二中医院 | 345027 | 4433500% |  | Wechat:02N/Ingenuity Core128/Elite China/Rack power down | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : Rack power down Safety Questions : Case: 0117563937 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Cheng Zhenhua Last updated by: Cheng Zhenhua 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 机架倾斜运动故障。 客户问题详细描述或故障现象： 打角度掉电 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： " S_TILT_C_NO_MOVE ERROR_FATAL//S_MPS_MAIN_DRIVE_NOT_CLOSING " 远程诊断与测试结果： "Description: A move request is issued but the Tilt is not moving." PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： "Recommendation: Check the Tilt motors, power supply, and see if anything is in the way of the Tilt. Check for any other S_TILT_ errors." 建议备件12NC及名称： DUAL POT ASSEMBLY 459800441121 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : PPD 459800441121 1已发18H至医院FSE齐良东-86-18662702811 Next Step Action : QI LIANGDONG Internal Remarks : 21.05.2021 08:20:14 UTC Liangdong Qi (86036024) "RDI Accuracy: Yes“ $Q7: (Confirmed reported problem, the MDC was opened when tilt gantry occasionally.) $Q8: (Analysis error report, found "A move request is issued but the Tilt is not moving". Checked the potentiometer belt of tilt, it was too tighten.) $Q9: (Adjusted the belt of tilt potentiometer and did tilt calibration.) $Q10: (Test and Inspection data refer to PA & Device return to fully function) $Q11: Calibrated Tools (NA). External Remarks : 21.05.2021 08:20:08 UTC Liangdong Qi (86036024)服务内容：调整倾斜电位器皮带并校正. 服务结果： 维修后，设备恢复正常工作. ----------------------------------------------------------------------------- |  |  | Adjusted the belt of tilt potentiometer and did tilt calibration. | [12 empty] | 117563937 |  | 921008 |  | Closed - Done | Closed - Done |  | ECR-077012 | [5 empty] | Engineering | [6 empty] | Dual pot |  |  | Completed |  |  | Refresh |  |  | System issue | System | Gantry-HW | Belt | Gantry-HW-Belt | Gantry-HW-Belt | ECR-077012 | Belt | Gantry-HW | Gantry-HW-Belt | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
938 |  |  | China | Jizhong Energy jingxing Mining Group Co., Ltd. General Hospital | 345003 | 4433400% |  | PRS connect/02N/Ingenuity Core128/Elite China | BU Q&R Feedback: No Reason: Customer Relations Problem Reported by Customer : PRS connect /02N/Ingenuity Core128/Elite China SAP Transfer:Case Transferred to SAP Safety Questions:Case: 0117559119 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Zhang Jing Last updated by: Zhang Jing 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Dispatched for On site Support Internal Remarks: 19.05.2021 07:15:54 UTC Litu Yang (86038906) RDI accuracy Yes $Q7: (Confirmed the reported problem. install prs) $Q8: (There has no error，just need to reinstall PRS system) $Q9: (Install PRS system) $Q10: (Test and inspection data refer to: PA. Device returned to full functionality.) $Q11: Calibrated Tools name (NA) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date: 09/15/2020 ^Current Tube usage: scan second_33217 External Remarks: 19.05.2021 07:15:51 UTC Litu Yang (86038906) Service Details: [No error, just install the PRS device.] Service results: The device returned to normal working conditions.] ----------------------------------------------------------------------------- | 5 |  | PRS config | [12 empty] | 117559119 |  | 917546 |  |  | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
939 |  |  | China | Fujian Provincial Hospital | 345011 | 4433400% |  | 02N/Ingenuity Core128/Elite China/add pacs | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 02N/Ingenuity Core128/Elite China Next Step Action: site1 Li Feng 018650369579 called in Safety Questions:Case: 0117558426 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Bo Last updated by: Wang Bo 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer's problems: Other types of error. Detailed description of Customer's problems or malfunctions: Add AI workstation. It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: The log has no errors. Remote troubleshoot and test result: Restart but didn't work. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Suggest to cooperate with customer to add AI workstation. We recommend spare parts 12NC and names: N/A SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support External Remarks: 19.05.2021 02:13:42 UTC SUP JCOUSER (86036224) service details: mobile CT need to add a network AI analysis node.Work on the field to add this network AI node and modify the scan protocol.CT images can be transmitted to this AI analysis node normally. service result: device returns to normal working condition Internal Remarks : 19.05.2021 02:13:44 UTC Feng Li (86036224) RDI Accuracy: Yes $Q7: (Confirmed reported problem. The CT need add an AI node.) $Q8: (On site to add an AI node and change scan protocol.) $Q9: (The image can transfer to AI node.) $Q10: (Test and inspection data refer to: PA.+Device returned to full functionality.) $Q11: Calibrated Tools name (FLUKE 287C Multimeter) S/N (16490128) NextCalibration date (10/20/2021) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date: 09/13/2020 ^Current Tube Usage: Scan Second_10375 ----------------------------------------------------------------------------- | 5 |  | add pacs | [12 empty] | 117558426 |  | 917094 |  |  | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
940 |  |  | China | Shandong Province Chest Hospital | 40025 | 4433400% |  | 02N/Ingenuity CT 728326// ASP engineer ask to nemo filter | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity CT 728326 Safety Questions : Case: 0117557592 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Jianqiang Last updated by: Wang Jianqiang 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | Zhai, Lisa |  | NEMO | [12 empty] | 117557592 |  | 916502 |  |  | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
941 |  |  | China | Zhejiang Province Yueqing No. 1 People’s Hospital | 40054 | 4433400% |  | WeChat:02N/Ingenuity CT 728326/Frame without electricity | BU Q&R Feedback: No SAP Customer Symptom Code: HANG Customer Symptom Code: Hang/Crash Reason: Corrective Maintenance Engineer Failure Code: Wear-out/Breakage/Mechanical Failure Repair Activity Code: Reboot/Restart Problem Reported by Customer : Frame without electricity T2 Activities : Problem Description: gantry can not boot up Troubleshooting Action: gantry can not boot up Repair Action: Assist the customer to restart the machine Test/Inspection: customer confirm can scan patient and image ok.functional Check PASSED. Device returned to full functionality. Test Equipment: N/A Safety Questions : Case: 0117557549 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Bo Last updated by: Wang Bo 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | Zhai, Lisa |  | Assist the customer to restart the machine | [12 empty] | 117557549 |  | 916466 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD_CouchEstop | [5 empty] | Engineering | [6 empty] | E-Stop can't close |  |  | Completed |  |  | Refresh |  |  | System issue | System | Couch-SW | TBD-bugrep | Couch-SW-TBD-bugrep | Couch-SW-TBD-bugrep | TBD_CouchEstop | TBD_Bugreport | Couch-SW | Couch-SW-TBD-bugrep | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
942 |  |  | China | 南昌大学第一附属医院(井冈山医院) | 345002 | 4433300% |  | WeChat：02N/Ingenuity Core128/Elite China/can't run | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Calibration/Adjustment Safety Questions : Case: 0117553262 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Cao Last updated by: Liu Cao 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Problem Reported by Customer : WeChat：02N/Ingenuity Core128/Elite China/can't run T2 Activities : Problem Description: 客户问题概述： 账户锁死，已到场 处理 客户问题详细描述或故障现象： 账户锁死，已到场 处理 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 账户锁死，已到场 处理 远程诊断与测试结果： 记录维修 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 反馈账户锁问题已到场处理，开单记录 Test/Inspection: N/A Test Equipment: N/A SAP Transfer : Dispatched for Onsite Support Next Step Action : yan xuehui SAP Transfer : Case Transferred to SAP External Remarks : 17.05.2021 17:25:21 UTC SUP JCOUSER (86038347)服务内容：账户过期锁定，重置账户；软件卡顿，传输队列堵住，清理队列后恢复。 服务结果：维修后，设备恢复正常工作。 Internal Remarks : 20.05.2021 00:29:23 UTC Xuehui Yan (86038347) RDI Accuracy:Yes $Q7: Confirmed reported problem can't run, additionally it was the account locked due to expired & software running slowly. $Q8: Check the CT and IT user login, indicate that expired; Check the queue, found the queue was stuck. $Q9: Reset Host account and reset spooler. $Q10: Test and inspection data refer to: PA. Device returned to full functionality. $Q11: Calibrated Tools name (N/A) S/N (N/A) Next Calibration date (N/A). ^Current Tube installed date: MM/DD/ YYYY. ^Current Tube usage: scan second_N/A. ----------------------------------------------------------------------------- | Zhai, Lisa |  | 服务内容：账户过期锁定，重置账户；软件卡顿，传输队列堵住，清理队列后恢复。 服务结果：维修后，设备恢复正常工作。 | [12 empty] | 117553262 |  | 913282 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | User account lock + PACS Transfer |  |  | Completed |  |  | Refresh |  |  | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
943 |  |  | China | 内蒙古兴安盟人民医院 | 345030 | 4433300% |  | PRS Connection/02n/Ingenuity Core128/Elite China | BU Q&R Feedback: No Reason: Customer Relations Next Step Action : FSE：苏瑞杰-13304719335 call in and disp him Problem Reported by Customer : PRS Connection Safety Questions : Case: 0117550653 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Lili Last updated by: Lu Lili 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 17.05.2021 03:15:41 UTC Ruijie Su (86039043)服务内容： 1.检查华为4G路由的上网情况。 2.重新配置PRS参数并测试。 服务结果： 设备开机正常，曝光正常，PRS正常。 Internal Remarks : 17.05.2021 03:15:44 UTC Ruijie Su (86039043) $Q1: (Y) $Q2: (N) $Q3: (NA) $Q4: (NA) $Q5: (NA) $Q6: (PRS Connection) $Q7: (Confirmed reported problem, no error ，only PRS Connection) $Q8: (check log,no log，PRS Connection) $Q9: (check the router，reconfigure the PRS) $Q10: (Test and inspection data refer to: PA. Device returned to full functionality.) $Calibrated Tools :NA ----------------------------------------------------------------------------- | 5 |  | PRS config | [12 empty] | 117550653 |  | 911488 |  |  | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
944 |  |  | China | DaKuai Town Health, Fengquan District, Xinxiang City, Henan Province | 345009 | 4433300% |  | PRS connection/02N/Ingenuity Core128/Elite China | BU Q&R Feedback: No Reason: Customer Relations Problem Reported by Customer:PRS Connection Safety Questions:Case: 0117550359 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Qi Xiaoyu Last updated by: Qi Xiaoyu 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Dispatched for On site Support External Remarks: 18.05.2021 04:05:48 UTC SUP JCOUSER (86038308) Fault Description: There is no malfunction, link PRS device Service Details: There is no malfunction, link PRS device Service result: After the service, the equipment starts as normal other issues: none. Internal Remarks: 19.05.2021 02:24:21 UTC Haibin Guo (86038308) RDI Accuracy: N/A Q7:conform report problem and just PRS connection $Q8:no CM action and just PRS connection $Q9:Replace the 4G router. $Q10: No test and inspection data. Device returned to full functionality. $Q11:Calibrated Tools N/A SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | 5 |  | PRS config | [12 empty] | 117550359 |  | 911320 |  |  | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
945 |  |  | Germany | Gesundheit Nord gGmbH Klinikverbund | 40002 | 4433300% |  | Error message: emergency stop switch | Work Order #: WO-03501207 Eduard Berthold The system fully meets specification and returned to use Problem Description by Engineer: Callback to Raik Glogner-Umbeer: Error message emergency stop switch cannot be closed. Resolution: The following activities were carried out for error analysis / troubleshooting: Error log diagnosis:> no indications in the log file It is very likely that a tape switch was pressed. Customer briefed! | 4 |  | It is very likely that a tape switch was pressed. Customer briefed! | [12 empty] | 117551170 |  | 908944 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD_CouchEstop | [5 empty] | Engineering | [6 empty] | E-Stop can't close |  |  | Completed |  |  | Refresh |  |  | System issue | System | Couch-SW | TBD-bugrep | Couch-SW-TBD-bugrep | Couch-SW-TBD-bugrep | TBD_CouchEstop | TBD_Bugreport | Couch-SW | Couch-SW-TBD-bugrep | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
946 |  |  | Japan | Maki Hospital | 341017 | 4433600% |  | JS: Investigation of an artifact that appeared on head images. | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: JS: Investigation of an artifact that appeared on head images. Safety Questions:Case: 0117570693 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: SHA Joy Last updated by: SHA Joy 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A OneEMS internal remarks: Yamamoto DB SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Support ----------------------------------------------------------------------------- | 5 |  | TBD | [12 empty] | 117570693 |  | 925897 |  | Investigation | Closed - Done | [7 empty] | TBD | [6 empty] | TBD |  |  | Completed |  |  | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
947 |  |  | Japan | Independent administrative institution National Hospital Organization | 341015 | 4433400% |  | 01CS:KKI:After EndStudy, the previous patient's data remains on the patient registration screen >> I asked t [Truncated] | [Phenomenon confirmation] Confirmed [Details of work] Call content After EndStudy, registration error occurs because the previous patient data remains on the patient registration screen. Correspondence Request that the RIS side handle it or recommend an examination after confirming it when registering a patient. [Confirmation details] Good | 5 |  | Request that the RIS side handle it or recommend an examination after confirming it when registering a patient. | [12 empty] | 117558498 |  | 917153 |  | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | registration |  |  | Completed |  |  | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
948 |  |  | Poland | Government Agency. | 341010 | 4433700% |  | Connection of the syringe. | Work Order #: WO-03530711 21-May-2021 Mr. Robert Kamola Diagnostic performed by Engineer : N/A Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: ----------------------------------------------------------------------------- | 4 |  | Connection of the syringe. | [12 empty] | 117578658 |  | 931416 |  | Evaluation | Closed - non-complaint | [7 empty] | TBD | [6 empty] | NAC |  |  | Completed |  |  | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
949 |  |  | Poland | Government Agency. | 341010 | 4433600% |  | Parts Supply Only | Work Order #: WO-03526538 Parts Consumed : 1.00000,459800042241,Cushion-Head/ Armrest ----------------------------------------------------------------------------- | 5 |  | Connection of the syringe. | [12 empty] | 117574335 |  | 928259 |  | Closed - Non-Complaint | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
950 |  |  | Spain | S.C.S. HOSP. GRAL. DE FUERTEVENTURA | 40020 | 4433600% |  | The images are not transferred. | They confirm to me in later days that the transfers to the Pacs, from the Tac, are working normally. The transfer queue is unlocked remotely to release the shipments. They did not pass the studies towards the Pacs. |  |  | The transfer queue is unlocked remotely to release the shipments. | [12 empty] | 117572758 |  | 927512 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS Transfer |  |  | Completed |  |  | Refresh |  |  | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
951 |  |  | Turkey | ADIYAMANOZELSAGINSTURTASSANVETICA.S | 341006 | 4433700% |  | sarı ekran problemi | Work Order #: WO-03530557 Soner Saglam The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Device freezes intermitent SW and gives yellow screen. Follow Up Required: The process from Helpdesk will be applied. Resolution: The process from Helpdesk will be applied. ----------------------------------------------------------------------------- |  |  | Recommended work around procedure for freezing of the device has been carried out. The device has been delivered in working condition. | [12 empty] | 117578514 |  | 926651 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632520 | [5 empty] | Engineering | [6 empty] | New Yellow Screen |  |  | Completed |  |  | Refresh |  |  | System issue | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
952 |  |  | Turkey | ADIYAMANOZELSAGINSTURTASSANVETICA.S | 341006 | 4433600% |  | HCOR kalibrasyonu | Work Order #: WO-03524594 Emre Gül The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : The user was given support on the phone and the malfunction was examined. In order to increase the quality of the image, it is necessary to clean the bore blood and opaque material and make the HCOR calibration. An appointment will be made with the customer for this job. The device is operational. Follow Up Required: Bore / Aplane will be cleaned from blood and opaque material and HCOR calibration will be done. ----------------------------------------------------------------------------- |  |  | Bore / Aplane will be cleaned from blood and opaque material and HCOR calibration will be done. | [12 empty] | 117572904 |  | 924254 |  | Evaluation - Pending Translation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | opaque material |  |  | Completed |  |  | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
953 |  |  | Turkey | ADIYAMAN OZEL SAGINSTURTAS SAN VE TIC A.S | 341006 | 4428500% |  | No connection between the CT scanner and monitor; Replaced the HDMI cable |  | CFM |  | No connection between the CT scanner and monitor; Replaced the HDMI cable | [14 empty] | 921401 |  | Closed - Done | Closed - Done | [7 empty] | Customer | [6 empty] | DVI Adapter |  |  | Completed |  |  | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
954 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 337000 | 4433300% |  | Table movement is difficult and noisy. |  |  |  | The places that needed to be oiled were oiled and the movements were made again, and it was observed that the sound was gone. The device was delivered actively by shooting trials with the Phantom. | [12 empty] | 117553681 |  | 909482 |  | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Couch Noise |  |  | Completed |  |  | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
955 |  |  | United States | Philips Medical Capital | 40028 | 4433500% |  | Questions during Biopsy |  | 5 |  | Q&A | [12 empty] | 117568585 |  | 919980 |  | Closed - Non-Complaint | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
956 |  |  | United States | Lackey Memorial Hospital | 40015 | 4424500% |  | Found system grounding issue during service; Upsized the ground and installed a separate ground rod to resolve. |  | CFM |  | FSE along with the hospital electrician upsized the ground and installed a separate rod to correct the grounding issue.The device works as specified after the correction. | [14 empty] | 908786 |  | Closed - Done | Closed - Done | [7 empty] | Customer | [6 empty] | Grounding |  |  | Completed |  |  | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
957 |  |  | United States | Advent Health Wesley Chapel | 300106 | 4433600% |  | all monitors (3) blanked out during scanning |  | CFM |  | Received bugreport and timestamp of issue happened. System works well after replace Host PC. | [12 empty] | CFM |  | 922906 |  | Closed - Duplicate | Closed - Duplicate | [14 empty] | Black screen | [5 empty] | Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
958 |  |  | Netherlands | Stichting Rijnstate Ziekenhuis | 52101 | 4433600% | 20 | Host computer is frozen. | Work Order #: WO-03522547 JOHNNY BASTIAENS The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : The Host computer froze after logging in repeatedly with wrong information. Resolution : Had the user log in with the admin account of the hospital. | 1 |  | Host PC is blocked after repeated login with wrong credentials Let the user log in with the hospital's admin account | [12 empty] | 117570919 |  | 921414 |  | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Password |  |  | Complete |  |  | Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
959 |  |  | United States | AdventHealth Zephyrhills | 300106 | 4433400% |  | Screens go blank and then back on | Work Order #: WO-03512405 Boyd Craig The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: Screens go blank and then back on User Impact: None Patient Impact: None Current Software Version: Diagnostic performed by Engineer : Screens go blank and then back on. FSE is aware of issue and requested customer call in for case number. PRS accessed. ULV: No normal errors seen at time of issue 10:00 AM. Follow Up Required : FSE on site to evaluate Host video connections. Recommendations: Check all video connections. Replace video cables or video card if needed.Display Port to DVI-I cable adapter 455300015471, nVidea Quadro FX580 Display Card 455300004321. ----------------------------------------------------------------------------- Work Order #: WO-03512457 20-May-2021 Sher Mark The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: Screens go blank and then back on User Impact: None Patient Impact: None Current Software Version: Diagnostic performed by Engineer : due to heavy workload no access to the machine was provided BIU you requested different parts be installed Follow Up Required : replace cable Parts Consumed : 2.00000,455300015471,Display Port to DVI-I cable adapter 3.00000,452213232132,CABLE BH-X0 - OCD ----------------------------------------------------------------------------- Work Order #: WO-03526065 22-May-2021 Sher Mark Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: Screens go blank and then back on User Impact: None Patient Impact: None Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 3 |  | Parts Consumed : 2.00000,455300015471,Display Port to DVI-I cable adapter 3.00000,452213232132,CABLE BH-X0 - OCD | [12 empty] | 117561516 |  | 914886 |  | Closed - Done | Closed - Done |  | CTUDT00632173 | [5 empty] | Engineering | [6 empty] | DVI Adapter |  |  | Complete |  |  | Refresh |  |  | Upgrade issue | System | Console-HW | Adaptor | Console-HW-Adaptor | Console-HW-Adaptor | CTUDT00632173 | Adaptor | Console-HW | Console-HW-Adaptor | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
960 |  |  | United States | Christ Hospital | 300205 | 4434300% |  | Errors during scan due to xray tube arcing; Advised Biomed to check the candlestick and run long tube conditioning. | Work Order #: WO-03563437 Morrow, Dan The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Problem description by engineer : Errors during scan When was the first occurrence? Within... : 27-May-2021 How often is issue occurring? : First occurrence Malfunction area : Tube&Gen Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected results: Error Patient Impact/Outcome: none User Impact: Retry scan Current Software Version: 4.1.7.25026 Diagnostic performed by Engineer : Reviewed logs. XRT arcing event. Advised Biomed to check the candlestick wells and run LTC. If issue persists replace the XRT. Biomed to perform work on system. Internal Comments : Error # and/or description of error : S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT S_GEN_FAULT_ANODE_TUBE_ARC S_GEN_FAULT_CATHODE_TUBE_ARC Completed or Recommended Repair Action "Internal" : Reviewed logs. XRT arcing event. Advised Biomed to check the candlestick wells and run LTC. If issue persists replace the XRT. Biomed to perform work on system. ----------------------------------------------------------------------------- | 3 |  | Advised Biomed to check the candlestick wells and run LTC. If issue persists replace the XRT. Biomed to perform work on system. | [12 empty] | 117610856 |  | 948774 |  | Closed - Done | Closed - Done |  | CTUDT00630085 | [5 empty] | Engineering | [6 empty] | Tube arc |  |  | Complete |  |  | Refresh |  |  | Upgrade issue | System | IC-HW | Tub Arc | IC-HW-Tub Arc | IC-HW-Tub Arc | CTUDT00630085 | Tub Arc | IC-HW | IC-HW-Tub Arc | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
961 |  |  | United States | Select Specialty Hospitals | 310051 | 4434200% |  | how to set contrast time | Work Order #: WO-03556318 Curry, Dyan The system fully meets specification and returned to use Diagnostic performed by Engineer : customer does not have password for ECM . Customer says she will wait to speak to April Panos, Clin Ed for password . OK to close w/o Problem description by Clinical Support Specialist : Customer Function/ Role: technologist Device Use: in use name Expected and Unexpected Results: how to add contrast to exam card. User Impact: none Patient Impact: none Current Software version: 4.1.7 | 4 |  | 2021-05-26 | [12 empty] | 117603726 |  | 948676 |  | Closed - Non-Complaint | Closed - Non-Complaint | [14 empty] | Password | [5 empty] | Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
962 |  |  | United States | Select Specialty Hospitals | 310051 | 4434100% |  | CT Brain Repeat for motion | Work Order #: WO-03548032 Mujica, David The system fully meets specification and returned to use Problem description by Clinical Support Specialist : Customer Function/Role: Technologist Device Use: In use Expected and Unexpected Results: Repeat scan due to motion User Impact: none Patient Impact: none Current Software Version: 4.1 Diagnostic performed by Engineer : Voncile indicated that she needed to repeat a portion of a CT head scan due to patient motion and wanted to know the quickest way to do so. Repeat was explained, z location, and drag explored. A simulation was done to increase understanding and application. Voncile successfully repeated a mock scan using the repeat function. Voncile had no further questions or need for CT clinical support. | 4 |  | Repeat scan due to motion | [12 empty] | 117595808 |  | 938746 |  | Closed - Done | Closed - Done | [7 empty] | Others | [6 empty] | Patient Motion |  |  | Complete |  |  | Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
963 |  |  | United States | Select Specialty Hospitals | 310051 | 4434000% |  | how to get table location | Work Order #: WO-03541643 Curry, Dyan The system fully meets specification and returned to use Problem description by Clinical Support Specialist : Customer Function/ Role: technologist Device Use: in use name Expected and Unexpected Results: how to plan scan location User Impact: none Patient Impact: none Current Software version: 4.4 Diagnostic performed by Engineer : advised customer to scroll though previous images for scan location needed. She can plan new scan and insert needed numbers in start/stop location. | 4 |  | 2021-05-24 | [12 empty] | 117589186 |  | 934985 |  |  | Closed - Non-Complaint | [14 empty] | Q&A | [5 empty] | Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
964 |  |  | Chile | BANCO DE CHILE | 341005 | 4434200% | 21 | CL1688 THE SCAN IS ABORTED BEFORE FINISHING DUE TO OVEREXPOSURE | CL1688 THE EXAM IS ABORT BY ITSELF BEFORE ENDING BY OVEREXPOSURE Message indicates tube is very hot: change technique. Resolution Description: Error log is reviewed. A message appears on the screen indicating that the tube is very hot. Falla occurs only with large patients. There is an indicator on the screen that shows that the technique used is not recommended for the patient: asterisks appear in planned slices. The protocol is reviewed with an application specialist and changes are requested in the kV, scan time and filter parameters for this type of patient. Error message on screen, does not allow exam. Information to support the complaint handling process Customer Function / Role: Technician / Sonographer Device Use: Diagnostic Expected and Unexpected results: No errors stopping the exam. Patient Impact / Outcome: Delay in exams. User Impact: Delay in exams. | 2 |  | The protocol is reviewed with an application specialist and changes are requested in the kV, scan time and filter parameters for this type of patient. Collect log and scan parameters for eng investigation | [12 empty] | 117603303 |  | 948395 |  | Opened | Closed - Done |  | Tube Heat | [5 empty] | Engineering | [6 empty] | Tube Overheat |  |  | Completed |  |  | Refresh |  |  | System issue | System | Gantry-SW | Tube heat | Gantry-SW-Tube heat | Gantry-SW-Tube heat | Tube Heat | Tube Heat | Gantry-SW | Gantry-SW-Tube heat | No | y | y | y | y | y | No | y | 0 | y | y | y | y | y | y | 0 | y | y
965 |  |  | China | Inner Mongolia Alxa Right Banner People's Hospital | 345031 | 4434400% |  | 02N/Ingenuity Core128/Elite China/UPS fault | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer:UPS fault Safety Questions:Case: 0117617330 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Zixuan Last updated by: Wang Zixuan 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 3 |  | fse suruijie 告知设备已恢复，不用派工 | [12 empty] | 117617330 |  | 955544 |  | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | UPS |  |  | Completed |  |  | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
966 |  |  | China | 苏州广慈肿瘤医院有限公司 | 345019 | 4434400% |  | 02N/Ingenuity Core128/Elite China/password failed | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity Core128/Elite China Safety Questions : Case: 0117614148 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wennian Han Last updated by: Wennian Han 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 密码问题 客户问题详细描述或故障现象： 密码问题 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： NA 远程诊断与测试结果： 重启主机后登录提示密码不对，客户告知没有输错密码。工程师指导客户用IT账户也不能登录。 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 用IST权限重置账户；检查是否软件问题。 客户需求备注： 客户着急需要尽快到场 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 28.05.2021 07:12:11 UTC Haotian Xu (86090280) 服务内容：CT账号被锁定，无法登陆，pacs回传配合修改参数 服务结果：设备运行正常。 ----------------------------------------------------------------------------- | 3 |  | Checked the CT and IT account and password Change IT and CT password according to service manual align with MPM and TSE to highlight the Password policy | [12 empty] | 117614148 |  | 954449 |  | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Application | [6 empty] | Password |  |  | Completed |  |  | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
967 |  |  | China | 苏州广慈肿瘤医院有限公司 | 345019 | 4434400% |  | 02N/Ingenuity Core128/scan interrupt | BU Q&R Feedback: No SAP Customer Symptom Code: BOOT Customer Symptom Code: Boot up/Power-up Problem Reason: Corrective Maintenance Engineer Failure Code: Electrical/Electronic Failure Repair Activity Code: Reboot/Restart Problem Reported by Customer : 02N/Ingenuity Core128 Safety Questions : Case: 0117613883 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Weng Jinming Last updated by: Weng Jinming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Troubleshooting Action: instruct customer to restart the system Repair Action: instruct customer to restart the system Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: N/A SAP Transfer : Case Transferred to SAP | 3 |  | instruct customer to restart the system align with MPM and TSE to highlight the restart system regularly | [12 empty] | 117613883 |  | 954305 |  | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Scan |  |  | Completed |  |  | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
968 |  |  | China | 青海省海北州第二医院 | 345029 | 4434400% |  | Wechat：02N/Ingenuity Core128/Elite China/Black screen | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : Black screen Safety Questions : Case: 0117613609 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Dai Suyang Last updated by: Dai Suyang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 其他类型故障。 客户问题详细描述或故障现象： 右侧显示器黑屏 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 右侧显示器黑屏 远程诊断与测试结果： 重启无效，一直不亮，显示器本身故障 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 更换显示器 建议备件12NC及名称： 459801140821 19" LCD display monitor T2 Activities : Problem Description: 客户问题概述： 其他类型故障。 客户问题详细描述或故障现象： 右侧显示器黑屏 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 右侧显示器黑屏 远程诊断与测试结果： 重启时提示没有视频输入，应该是显卡故障 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 更换显卡 建议备件12NC及名称： 459801536801 NVIDIA Quadro P400 2GB (3)mDP GFX Next Step Action : PPD 459801536801 1数量已核实,SWO产生后发货FSE许永华-86-13327661334 SAP Transfer : Dispatched for Onsite Support SAP Transfer : Case Transferred to SAP External Remarks : 30.05.2021 06:03:47 UTC Yonghua Xu (86001334)放射科CT室 联系人：郭荣寿 联系电话：15309708576 故障现象：右侧显示器黑屏 。 解决方案/维修措施：右侧显示器黑屏,显卡控制面板里面识别不到右侧显示器，和左侧显示器替换数据线及电源线发现显卡接右侧显示器信号输出口数据输出异常 。判断显卡故障。更换显卡。 完成相应校准，完成备份。 服务后设备正常工作。 Internal Remarks : 30.05.2021 06:03:50 UTC Yonghua Xu (86001334) “RDI Accuracy: Yes” $Q7:Confirmation reported problem,The right display is black. $Q8:The right display is black.Replace the data line and power line with the left display and find that the graphics card is connected to the right display signal output data abnormal.Determine graphics card failure. $Q9:Replace the graphics card. $Q10:Test and Inspection data refer to PA , Do a few patient, Observation of a certain period of time. The system is ok. Device returned to full functionality. $Q11:Calibrated tools name（NA）S/N（NA) Next calibration date（NA） ----------------------------------------------------------------------------- | 3 |  | 右侧显示器黑屏,显卡控制面板里面识别不到右侧显示器，和左侧显示器替换数据线及电源线发现显卡接右侧显示器信号输出口数据输出异常 。判断显卡故障。更换显卡。 | [12 empty] | 117613609 |  | 954140 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | NotEngineering-partreplace | [5 empty] | Engineering | [6 empty] | DisplayCard |  |  | Completed |  |  | Refresh | 459801536801 | NVIDIA Quadro P400 2GB (3)mDP GFX | System issue | System | Console-HW | DisplayCard | Console-HW-DisplayCard | Console-HW-DisplayCard | NotEngineering-partreplace | DisplayCard | Console-HW | Console-HW-DisplayCard | No | y | y | y | y | y | No | y | 0 | y | y | y | y | y | y | 0 | y | y
969 |  |  | Japan | Maki Hospital | 341017 | 4434300% |  | 02CS:KK: Patient restraint strap order. | BU Q&R Feedback: No Reason: Corrective Maintenance Safety Questions:Case: 0117605937 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: KENICHI KOJIMA Last updated by: KENICHI KOJIMA 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Problem Reported by Customer : 02CS:KK: Patient restraint strap order. T2 Activities:Problem Description: subject equipment: CT Failure location: 02Couch When did the problem first occur (within what period of time)?: Unknown How often does the problem occur?: Unknown At what point in the end user's work flow did the problem occur?: Other What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem? No.The customer has not created a log file or report. Troubleshooting Action: Result after troubleshooting: Improved Failure location: 02Couch What steps did you perform to diagnose the problem? Patient restraint strap order. OneEMS internal remarks: --Troubleshooting Action-- Result of response: Problem solved. Failure location: 02Couch What steps did you perform to diagnose the problem? Patient restraint strap order. --Repair Action-- Recommended 12Nc: --None-- 453567009951:L-6"PATIENT RESTRAINT STRAP/ 42" QTY:2 SLOC:--None-- --Subject-- 02CS:KK:Person in charge on site: Patient restraint strap order. --Problem Description-- Equipment affected: CT Failure location: 02Couch When did the problem first occur (within what period of time)?: Unknown How often does the problem occur?: Unknown At what point in the end user's work flow did the problem occur?: Other What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem? No.The customer has not created a log file or report. --4/8Elements-- Q1. Was the reported problem confirmed? Cannot confirm Q2. Were any other problems confirmed? None. Q3. What did you do to resolve the problem?: Patient restraint strap order. Q4: Is the equipment operating normally following the response? Good. Q5. How did the examination finish? None. Q6. Were there any mistaken findings (misdiagnosis, mistaken treatment) due to the problem? None. Q7. Was there any unintended unnecessary exposure? None. Q8. Were there any artifacts? None. SAP Transfer: Case Transferred to SAP SAP Message: Components Added Succesfully to SWO!- SAP Message: LOGI Task added Successfully!- ----------------------------------------------------------------------------- | Yang, Cancan |  | 2021-05-27 | [12 empty] | 117605937 |  | 950122 |  |  | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh | 453567009951 | L-6" PATIENT RESTRAINT STRAP/ 42" QTY:2 | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
970 |  |  | United States | Unity Health Systems | 40037 | 4434200% |  | 0711CFXXX March 21st issue it initialized for a while until a computer reboot and it is happening again | Work Order #: WO-03555078 Atz, Michael The system fully meets specification and returned to use Problem Description by Engineer : Problem description by engineer : March 21st issue it initialized for a while until a computer reboot and it is happening again When was the first occurrence? Within... : 21-May-2021 How often is issue occurring? : Intermittent Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: No further information is available as per the customer Patient Impact/Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer Current Software Version: Not provided Diagnostic performed by Engineer : I advised the tech to do a daily Host reboot and a weekly CIRS/Gantry reboot. Tech power-cycled the Host/IRS computers and gantry at wall breaker. DAYS_SinceGantryPwrCycle 56.549849537037 I spoke to the FSE tom Lockwood and he has concurred that a weekly Gantry reboot is the issue Resolution : I advised the tech to do a daily Host reboot and a weekly CIRS/Gantry reboot. Tech power-cycled the Host/IRS computers and gantry at wall breaker. DAYS_SinceGantryPwrCycle 56.549849537037 I spoke to the FSE tom Lockwood and he has concurred that a weekly Gantry reboot is the issue Internal Comments : Error # and/or description of error : DAYS_SinceGantryPwrCycle 56.549849537037 Completed or Recommended Repair Action "Internal" : I advised the tech to do a daily Host reboot and a weekly CIRS/Gantry reboot. Tech power-cycled the Host/IRS computers and gantry at wall breaker. DAYS_SinceGantryPwrCycle 56.549849537037 I spoke to the FSE tom Lockwood and he has concurred that a weekly Gantry reboot is the issue ----------------------------------------------------------------------------- | 3 |  | that a weekly Gantry reboot is the issue 459801656021_b the common step by step sheet for customer, make it visible and train customer to follow it | [12 empty] | 117602764 |  | 943369 |  | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Reboot |  |  | Completed |  |  | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
971 |  |  | China | Jiangsu Province No. 2 Chinese Medicine Hospital | 345027 | 4434100% |  | Wechat:02N/Ingenuity Core128/Elite China/Gantry stop scanning | BU Q&R Feedback: No SAP Customer Symptom Code: MARF Customer Symptom Code: Monitor Alarm Failure Reason: Corrective Maintenance Engineer Failure Code: Electrical/Electronic Failure Repair Activity Code: Reboot/Restart Problem Reported by Customer : Gantry stop scanning Safety Questions:Case: 0117592397 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Caiqin Zhang Last updated by: Caiqin Zhang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer's problems: Scanning failed. Detailed description of Customer's problems or malfunctions: Scan error, scan failed It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: NORMAL ErrNum: S_DMC_COLL_ACOLLZ_TMOT P1:0/P2:0 Remote troubleshoot and test result: The collimator aperture movement timed out. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: 1. After restarting the gantry, the problem disappeared. Test/Inspection: N/A Test Equipment: N/A Next Step Action: Q6 = scan failed and report error SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | After restarting the gantry, the problem disappeared. | [12 empty] | 117592397 |  | 941082 |  | Investigation | Investigation |  | CollimatorTimerOout | [5 empty] | Engineering | [6 empty] | Collimator Timeout |  |  | Completed |  |  | Refresh |  |  | System issue | System | Gantry-SW | CollimatorTimerOout | Gantry-SW-CollimatorTimerOout | Gantry-SW-CollimatorTimerOout | CollimatorTimerOout | CollimatorTimerOout | Gantry-SW | Gantry-SW-CollimatorTimerOout | No | y | y | y | y | y | No | y | 0 | y | y | y | y | y | y | 0 | y | y
972 |  |  | China | Hebei Province Tangshan City Workers Hospital | 40044 | 4434100% |  | 02N/Ingenuity CT/Voice interference | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 02N/Ingenuity CT T2 Activities:Problem Description: Summary of customer's problems: Other types of error. Detailed description of Customer's problems or malfunctions: The customer talked from CT box, there was a noise from gantry. CT box buttons failed. It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: The log has no errors. Remote troubleshoot and test result: Restart but didn't work. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Suggest to replace CT box and adjust sound volume appropriately. We recommend spare parts 12NC and names: PPD 459801705022 SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Next Step Action: PPD 459801705022 1 This spare part is out of stock and needs to be ordered FSE Hu Yue-86-13703597273 Safety Questions:Case: 0117590842 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Bo Last updated by: Wang Bo 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 3 |  | 客户从CT BOX说话，机架发出时有杂音，CT box 按键失灵。 建议更换CT box并适当调节音量大小。 建议备件12NC及名称： PPD 459801705022 CT Box 459800367081 - Noise Reduction Gantry Audio Board Assy | [12 empty] | 117590842 |  | 940255 |  | Evaluation - Pending Translation | Closed - No Further Investigation |  | CTUDT00632185 | [5 empty] | Engineering | [6 empty] | CT Box |  |  | Completed |  |  | Refresh | 459801705022 | CT Box | System issue | System | Gantry-HW | CT Box | Gantry-HW-CT Box | Gantry-HW-CT Box | CTUDT00632185 | CT Box | Gantry-HW | Gantry-HW-CT Box | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
973 |  |  | China | 扎兰屯市人民医院 | 40045 | 4434200% |  | PRS connection/02N/Ingenuity CT 728326 | BU Q&R Feedback: No Reason: Customer Relations Problem Reported by Customer : PRS connection Safety Questions : Case: 0117598973 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingli Last updated by: Lu Mingli 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action : FSE：苏瑞杰-13304719335 call in and disp him FSE Notes : This PO is a Reference Number only SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 27.05.2021 06:29:08 UTC Ruijie Su (86039043)服务内容： 1.检查华为4G路由的上网情况。 2.重新配置PRS参数并测试。 服务结果： 设备开机正常，曝光正常，PRS正常。 Internal Remarks : 27.05.2021 06:29:11 UTC Ruijie Su (86039043) $Q1: (Y) $Q2: (N) $Q3: (NA) $Q4: (NA) $Q5: (NA) $Q6: (PRS Connection) $Q7: (Confirmed reported problem, no error ，only PRS Connection) $Q8: (check log,no log，PRS Connection) $Q9: (check the router，reconfigure the PRS) $Q10: (Test and inspection data refer to: PA. Device returned to full functionality.) $Calibrated Tools :NA ----------------------------------------------------------------------------- | Lu, Xueying |  | 2021-05-26 | [12 empty] | 117598973 |  | 945379 |  |  | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
974 |  |  | Japan | Maki Hospital | 341017 | 4434300% |  | Via Nakagawa in Applications.Regarding a band artifact on head images. | ① How should I think about the position setting when shooting plastic surgery (limbs)? (2) When shooting the elbow, there are two shooting methods, raising the arm and lowering the arm. What are the shooting conditions? ③ A black band appears inside the bone in the data of the southern condition when the sinuses are photographed. App person in charge answer (1) Decide which one to position in front and how to position it separately from the trunk. For example, in the case of elbows, the inside is the front and the upper arm is the head. (2) If the part other than the elbow enters the imaging range, the dose will be changed according to the DRI setting. I think that the current setting (DRI25) is okay, but please raise it a little and see the situation at about 26. ③ I checked the image. There is a dark band because the BHC function is not used. Select a function such as UA or UB. ④ Others: It seems that a dark band appears in the Mask and Delay images of the head CTA. It seems that the original data is also used for diagnosis, so I changed it to UB. I explained Dynaminc_Viewer and Brain_Viewer. | 5 |  | NAC? | [12 empty] | 117605901 |  | 950101 |  | Evaluation | Closed - Done | [7 empty] | Application | [6 empty] | Q&A |  |  | Completed |  |  | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
975 |  |  | United States | Philips Medical Capital | 40055 | 4434400% |  | 0998CDY73 FSE Rich is aware ring artifact over the weekend there is contrast in the machine | Work Order #: WO-03568810 Atz, Michael The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Problem description by engineer : FSE Rich is aware ring artifact over the weekend there is contrast in the machine, he wants to have it scheduled for Wednesday June 2nd When was the first occurrence? Within... : 28-May-2021 How often is issue occurring? : Multiple times per week Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: No further information is available as per the customer Patient Impact/Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer Current Software Version: Not provided Diagnostic performed by Engineer : The customer states that the ring goes away once the images are processed, FSE Rich is aware ring artifact over the weekend there is contrast in the machine, he wants to have it scheduled for Wednesday June 2nd Internal Comments : Error # and/or description of error : none Completed or Recommended Repair Action "Internal" : The customer states that the ring goes away once the images are processed, FSE Rich is aware ring artifact over the weekend there is contrast in the machine, he wants to have it scheduled for Wednesday June 2nd Follow Up Required : Error # and/or description of error : none Completed or Recommended Repair Action "Internal" : The customer states that the ring goes away once the images are processed, FSE Rich is aware ring artifact over the weekend there is contrast in the machine, he wants to have it scheduled for Wednesday June 2nd ----------------------------------------------------------------------------- Work Order #: WO-03568890 Turley Rich Problem Description by Engineer : Problem description by engineer : FSE Rich is aware ring artifact over the weekend there is contrast in the machine, he wants to have it scheduled for Wednesday June 2nd When was the first occurrence? Within... : 28-May-2021 How often is issue occurring? : Multiple times per week Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: No further information is available as per the customer Patient Impact/Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer Current Software Version: Not provided Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 3 |  | there is contrast in the machine | [12 empty] | 117617633 |  | 952541 |  | Investigation | Investigation |  | ECR-075975 | [5 empty] | Engineering | [6 empty] | artifact |  |  | Completed |  |  | Refresh |  |  | System issue | System | Gantry-HW | Lexan Ring | Gantry-HW-Lexan Ring | Gantry-HW-Lexan Ring | ECR-075975 | Lexan Ring | Gantry-HW | Gantry-HW-Lexan Ring | No | y | y | y | y | y | No | y | 0 | y | y | y | y | y | y | 0 | y | y
976 |  |  | Japan | Independent administrative institution National Hospital Organization | 341015 | 4434100% |  | 07CS:TM: ECG monitor lead wire procurement. >> | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer:07CS:TM: ECG monitor lead wire procurement. >> SAP Transfer: Case Transferred to SAP Safety Questions:Case: 0117593113 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: TAKEYUKI MIYAKAWA Last updated by: TAKEYUKI MIYAKAWA 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Message: Components Added Succesfully to SWO!- SAP Message: LOGI Task added Successfully!- ----------------------------------------------------------------------------- | Zhu, Chun |  | 2021-05-25 | [12 empty] | 117593113 |  | 941487 |  |  | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
977 |  |  | United States | Philips Medical Capital | 40055 | 4434200% |  | 0711CFXXX One study not going to PACS | Work Order #: WO-03556434 Boyd Craig The system fully meets specification and returned to use Diagnostic performed by Engineer : One study not going to PACS. PACS issues on 05-21-21. Patient information manually entered at that time. All other studies going to PACS. Unable to connect to target message when trying to send study. Out of resources. Bring applications, David Mujica, on call. Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: One study not going to PACS User Impact: None Patient Impact: None Current Software Version: Resolution : Found patient header information differed and made changes. Study went to PACS after changes make. ----------------------------------------------------------------------------- Work Order #: WO-03556488 Mujica, David The system fully meets specification and returned to use Problem description by Clinical Support Specialist : Customer Function/Role: Technologist Device Use: In use Expected and Unexpected Results: CT study not going to PACS User Impact: none Patient Impact: none Current Software Version: 4.1 Diagnostic performed by Engineer : Brent indicated that a CT study that was done on May 22nd is not going to PACS from the CT scanner. The error says "resource allocation error, source not mapped" Reconstruction form RAW data performed to possibly send the new reconstructed series to PACS, this was unsuccessful and the error did manifest again. The accession number was adjusted in patient details and changed to the correct accession number. After this was performed the study did successfully sent to PACS. Brent had no further questions or need for further Clinical support. No rescan, no loss of data, no reinjection as stated by the customer. ----------------------------------------------------------------------------- | 3 |  | The accession number was adjusted in patient details and changed to the correct accession number. After this was performed the study did successfully sent to PACS. | [12 empty] | 117603805 |  | 944143 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS Transfer |  |  | Completed |  |  | Refresh |  |  | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
978 |  |  | China | Medical center (contagious disease, He nan, Xu chang) | 40047 | 4434100% |  | 02N/Ingenuity CT 728326/stand by | BU Q&R Feedback: No Reason: Customer Relations Problem Reported by Customer: stand by Safety Questions:Case: 0117591335 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Zhang Jing Last updated by: Zhang Jing 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Case Transferred to SAP External Remarks: 26.05.2021 08:45:15 UTC SUP JCOUSER (86038955) Service Details: The hospital conducted an environmental evaluation and the field service engineer is on standby on site. Service result: After the service, the equipment starts as normal Others: None as of now. Internal Remarks:26.05.2021 08:45:18 UTC Yi Zhang (86038955) RDI Accuracy:Yes $Q7: (Confirm the problem, standby . ) $Q8: (Cooperate with hospital environmental impact assessment by standby. ) $Q9: (Cooperate with hospital environmental impact assessment by standby, then the issue was solved! ) $Q10: (Test and Inspection data refer to PA. Device returned to full functionality.) $Q11: (Calibrated Tools :N/A) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date: 00/ 00/ 0000 ^Current Tube usage: SAP Transfer:Dispatched for On site Support ----------------------------------------------------------------------------- | Gao, Mia |  | 2021-05-25 | [12 empty] | 117591335 |  | 940507 |  |  | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
979 |  |  | China | Shandong Anqiu City Traditional Chinese Medicine hospital | 345017 | 4434300% |  | 02NIngenuity Core128/Elite China/Image reconstruction failure | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 02N Ingenuity Core128/Elite China/ Safety Questions:Case: 0117605512 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingsong Last updated by: Lu Mingsong 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer's problems: Image reconstruction failed. Detailed description of Customer's problems or malfunctions: Intermittently unable to reconstruct images. Shut down and reboot, then restore It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: N/A Remote troubleshoot and test result: N/A PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Suspect it's software bug, suggest to collect bugrep and open ESCALATION We recommend spare parts 12NC and names: N/A Memo of Customer request: None SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support ----------------------------------------------------------------------------- | 5 |  | Image reconstruction failure: Unable to rebuild images intermittently after scan complete, normal after shutting down and restarting. Customer temperature: between Yellow and Red? Bugreport: https://www.clouddrive.philips.com.cn/file/iwcqt0xve89fuw15bwxxuzmfejk2f1m7# Password:123 | [12 empty] | 117605512 |  | 949841 |  | Evaluation | Closed - Done |  | CIRS-SW-CannotRecon | [5 empty] | Engineering | [6 empty] | Recon |  |  | Completed |  |  | Refresh |  |  | System issue | System | CIRS-SW | CannotRecon | CIRS-SW-CannotRecon | CIRS-SW-CannotRecon | CIRS-SW-CannotRecon | CannotRecon | CIRS-SW | CIRS-SW-CannotRecon | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
980 |  |  | China | 山东省泰安市中医二院 | 345037 | 4434000% |  | 02N/Ingenuity Core128/Elite China/change parameter | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: Instructed Customer Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ Safety Questions : Case: 0117585968 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wennian Han Last updated by: Wennian Han 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: change parameter Troubleshooting Action: need change printer parameter Repair Action: guide customer to change printer parameter Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: NA SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | Gao, Mia |  | 2021-05-24 | [12 empty] | 117585968 |  | 936400 |  |  | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
981 |  |  | Japan | Maki Hospital | 341017 | 4434000% |  | CT:APP:RK: Question: Mr. Noda. | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Application Support Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer : Q: He wanted to know how to create an MPR image from skull bones series in the same way as for brain parenchyma. A: Please use CT Viewer to load the series you want to create an MPR for. Then set the angle and determine the range as you normally would.Then select Save Preset from the Batch preset and save it, and apply this to the next series, and the angle will be the same.However, please keep in mind that the window from the series will be carried over. OneEMS internal remarks: Q1: Did you confirm the reported problem?: Confirmed Q2. Was there any other problem?None Q3. What did you do to resolve the problem?: 1: If resolved by telephone: explanation Q4: Is the equipment operating normally following the response?: Good Safety Questions:Case: 0117585700 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: RYUSUKE KUJIRAI Last updated by: RYUSUKE KUJIRAI 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | Xu, Xiaoting |  | 2021-05-24 | [12 empty] | 117585700 |  | 936213 |  |  | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
982 |  |  | United States | Central Vermont Medical Center | 40024 | 4434000% |  | 0308CDY73 Vertical movement noise | Work Order #: WO-03540631 Miller, John M The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Problem description by engineer : Customer states the couch vertical movement is making a lot of noise. When was the first occurrence? Within... : 24-May-2021 How often is issue occurring? : Continuously Malfunction area : Table Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Noise in couch movement Patient Impact/Outcome: None User Impact: Noise in couch movement Current Software Version: Not provided Diagnostic performed by Engineer : Collect information from customer. Follow Up Required : Error # and/or description of error : Vertical noise. Completed or Recommended Repair Action "Internal" : Suggest checking for debris in the Vertical movement, as well as covers binding. Suggest checking for proper lubrication of the the vertical drive screw and debris on the Thompson rods. Inspect the vertical spindle for wear and all parts in the vertical drive system. Refer to the Couch PM manual for proper lubrication and Ball screw inspection. Suggest checking the safety bar switch and its wiring. Check / replace limit switches. On Bariatric Couch, check for leaking piston. Clean any oil residue from the piston and replace them. Suggest checking, repairing, cleaning and/or replacing Vertical Encoder O-Ring. Internal Comments : Error # and/or description of error : Vertical noise. Completed or Recommended Repair Action "Internal" : Suggest checking for debris in the Vertical movement, as well as covers binding. Suggest checking for proper lubrication of the the vertical drive screw and debris on the Thompson rods. Inspect the vertical spindle for wear and all parts in the vertical drive system. Refer to the Couch PM manual for proper lubrication and Ball screw inspection. Suggest checking the safety bar switch and its wiring. Check / replace limit switches. On Bariatric Couch, check for leaking piston. Clean any oil residue from the piston and replace them. Suggest checking, repairing, cleaning and/or replacing Vertical Encoder O-Ring. ----------------------------------------------------------------------------- Work Order #: WO-03540665 24-May-2021 Morgan Russ Van The system fully meets specification and returned to use Problem Description by Engineer : Problem description by engineer : Customer states the couch vertical movement is making a lot of noise. When was the first occurrence? Within... : 24-May-2021 How often is issue occurring? : Continuously Malfunction area : Table Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Noise in couch movement Patient Impact/Outcome: None User Impact: Noise in couch movement Current Software Version: Not provided Diagnostic performed by Engineer : noisy couch , cleaned and lubricated ----------------------------------------------------------------------------- | 3 |  | noisy couch , cleaned and lubricated | [12 empty] | 117588255 |  | 934368 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-Service | [5 empty] | Engineering | [6 empty] | Couch noise |  |  | Completed |  |  | Refresh |  |  | System issue | System | Couch-HW | Vertical spring | Couch-HW-Vertical spring | Couch-HW-Vertical spring | TBD-Service | Vertical spring | Couch-HW | Couch-HW-Vertical spring | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
983 |  |  | Chile | Scotiabank Chile | 40008 | 4435000% |  | PROBLEM WITH VERTICAL MOVEMENT, GUIDE PISTON OUT OF ADJUSTMENT | Work Order #: WO-03599146 Alec, Preciado The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Problem description by engineer : PROBLEM WITH VERTICAL MOVEMENT, GUIDE PISTON OUT OF ADJUSTMENT When was the first occurrence? Within... : 03-Jun-2021 12:00:00 How often is issue occurring? : First occurrence Malfunction area : Table Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected results: Restore the patient table's proper operation. Patient Impact/Outcome: It does not impact. User Impact: It does not impact. Current Software Version: 4.1.10 Diagnostic performed by Engineer: Support is provided with the user, it is identified that the system shows noise in the vertical movement, derived from possible loss of mechanical adjustment in the piston. Internal Comments : Error # and/or description of error : NA Completed or Recommended Repair Action "Internal" : We checked the system with the user and it shares the sound's audio, which is characteristic of a vertical movement piston not being adjusted. It is in the patient table's base. Follow Up Required : What are the next Troubleshooting steps : It is necessary to check the patient table pistons. By the sound, it indicates that it is not adjusted. If so, perform activities to adjust again and afterwards grease the patient table according to the maintenance manual. Error # and/or description of error : NA Completed or Recommended Repair Action "Internal" : We checked the system with the user and it shares the sound's audio, which is characteristic of a vertical movement piston not being adjusted. It is in the patient table's base. ----------------------------------------------------------------------------- Work Order #: WO-03599244 Osorio, Juan Problem Description by Engineer : Problem description by engineer : PROBLEM WITH VERTICAL MOVEMENT, GUIDE PISTON OUT OF ADJUSTMENT When was the first occurrence? Within... : 03-Jun-2021 12:00:00 How often is issue occurring? : First occurrence Malfunction area : Table Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected results: Restore the patient table's proper operation. Patient Impact/Outcome: It does not impact. User Impact: It does not impact. Current Software Version: 4.1.10 Diagnostic performed by Engineer: N/A ----------------------------------------------------------------------------- | 3 |  | Planning Notes changed from "DL1 CLIENT TO BE CONFIRMED ¦ FSE J OSORIO ¦ CL1675 ¦ TUESDAY 8 JUN TENT" to "DL1 ASSIGNED ¦ FSE J OSORIO ¦ CL1675 ¦ FRIDAY 11 JUN | [12 empty] | 117648281 |  | 971398 |  | Opened | Closed - No Further Investigation | [7 empty] | Engineering | [6 empty] | Couch Noise |  |  | Completed |  |  | Refresh |  |  | System issue | System | Couch-HW | Table Cover | Couch-HW-Table Cover | Couch-HW-Table Cover | 0 | Table Cover | Couch-HW | Couch-HW-Table Cover | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
984 |  |  | Chile | Scotiabank Chile | 40008 | 4435000% |  | CL1675-CIRCULAR ARTIFACT IN THE PICTURE AND NOISE IN THE BED | Work Order #: WO-03597437 Alec, Preciado The system fully meets specification and returned to use Problem Description by Engineer : Problem description by engineer : ARTEFACTO CIRCULAR EN LA IMAGEN When was the first occurrence? Within... : 02-Jun-2021 11:00:00 How often is issue occurring? : First occurrence Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected results: imagenes sin problema de artefactos Patient Impact/Outcome: retraso en atencion User Impact: retraso en carga de trabajo Current Software Version: 4.1.7 Diagnostic performed by Engineer : Se realiza revision de tablas de aire en el sistema, se identifica señal no homogenea, se realiza nuevamente una calibración de aire y se revisa imagen, se corrige artefacto. Internal Comments : Error # and/or description of error : NA Completed or Recommended Repair Action "Internal" : calibracion de aire realizada, se corrige artefacto en sistema. Follow Up Required : PROBLEMA CON MOVIMIENTO VERTICAL, PISTON GUIA FUERA DE AJUSTE Resolution : Se realiza revision de tablas de aire en el sistema, se identifica señal no homogenea, se realiza nuevamente una calibración de aire y se revisa imagen, se corrige artefacto. ----------------------------------------------------------------------------- | 3 |  | A review of air tables is performed in the system, a non-homogeneous signal is identified, an air calibration is performed again and an image is reviewed, an artifact is corrected. | [12 empty] | 117646963 |  | 971330 |  | Opened - Pending Translation | Closed - Done |  | NotEngineering | [5 empty] | Engineering | [6 empty] | Artifact |  |  | Completed |  |  | Refresh |  |  | System issue | System | NotEngineering | ShouldbeCustomer | NotEngineering-ShouldbeCustomer | NotEngineering-ShouldbeCustomer | NotEngineering | NotEngineering-ShouldbeCustomer | NotEngineering | NotEngineering-ShouldbeCustomer | No | y | y | y | y | y | No | y | 0 | y | y | y | y | y | y | 0 | y | y
985 |  |  | Chile | Scotiabank Chile | 40008 | 4431300% |  | Burning smell during servicing, verified components in rotor-stator and power sector worked properly | Burning smell during servicing, verified components in rotor-stator and power sector worked properly | CFM |  | After confirmed, it's not from CT, but from air condition | [14 empty] | 962707 |  | Closed - Done | Closed - Done | [7 empty] | Others | [6 empty] | Burning smell |  |  | Completed |  |  | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
986 |  |  | Puerto Rico | METRO MAYAGUEZ INC | 40009 | 4435100% |  | Makes noise when tube is spinning | Work Order #: WO-03605095 Caldelas Aldana, Carlos Manuel The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Problem description by engineer : noise when rotating When was the first occurrence? Within... : 04-Jun-2021 11:00:00 How often is issue occurring? : Continuously Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: noise when rotating Patient Impact/Outcome: noise in system User Impact: noise in gantry Current Software Version: Not provided Diagnostic performed by Engineer : validated system is having abnormal noise like something hitting inside the gantry when rotating during scans. no errors appear for the moment. Internal Comments : Error # and/or description of error : noise when rotation Completed or Recommended Repair Action "Internal" : n/a Follow Up Required : What are the next Troubleshooting steps : open and inspect gantry to validate what is loose or causing the noise when rotating Error # and/or description of error : noise when rotation Completed or Recommended Repair Action "Internal" : n/a ----------------------------------------------------------------------------- | 3 |  | During the remote support session, it was identified that a field service engineer visit is required to continue with technical support. The planning team will be in touch with you shortly to coordinate the technical visit. | [12 empty] | 117654113 |  | 975210 |  | Opened | Closed - No Further Investigation |  | TBD-GantryNoise | [5 empty] | Engineering | [6 empty] | Gantry Noise |  |  | Completed |  |  | Refresh |  |  | System issue | System | Gantry-HW | Noise | Gantry-HW-Noise | Gantry-HW-Noise | TBD-GantryNoise | Noise | Gantry-HW | Gantry-HW-Noise | No | y | y | y | y | y | No | y | 0 | y | y | y | y | y | y | 0 | y | y
987 |  |  | United States | Lackey Memorial Hospital | 40015 | 4434800% |  | customer says there is a noise inside the gantry when using xray tube | Work Order #: WO-03582737 01-Jun-2021 Denison Chuck The system fully meets specification and returned to use Diagnostic performed by Engineer : Checked the Gantry internal. Made adjustments to front cover. installed plate into the bottom of the Gantry. Problem Description by Engineer : Customer Function/Role: Head of x-ray Dept. Device Use: Philips C.T. Expected and Unexpected Results: Scanner is working. User Impact: None system is up. Patient Impact: System is up. Current Software Version:4.1.10 ----------------------------------------------------------------------------- | 4 |  | Checked the Gantry internal. Made adjustments to front cover. installed plate into the bottom of the Gantry. | [12 empty] | 117632842 |  | 961582 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | NotEngineering | [5 empty] | Engineering | [6 empty] | Gantry Noise |  |  | Completed |  |  | Refresh |  |  | System issue | System | NotEngineering | ShouldbeOPS | NotEngineering-ShouldbeOPS | NotEngineering-ShouldbeOPS | NotEngineering | NotEngineering-ShouldbeOPS | NotEngineering | NotEngineering-ShouldbeOPS | No | y | y | y | y | y | No | y | 0 | y | y | y | y | y | y | 0 | y | y
988 |  |  | China | 河南省濮阳市人民医院 | 40031 | 4434900% |  | Wechat:02N/Ingenuity CT 728326/Data transmission failure | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: Calibration/Adjustment Problem Reported by Customer : Data transmission failure Safety Questions : Case: 0117638092 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Bai Jingwen Last updated by: Bai Jingwen 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: Cannot transfer image to ISP Troubleshooting Action: Check ISP free space is 0%, need clean Repair Action: Customer will clean data and try again Test/Inspection: N/A Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | Check ISP free space is 0%, need clean Customer will clean data and try again | [12 empty] | 117638092 |  | 968863 |  | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | ISP |  |  | Completed |  |  | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
989 |  |  | United States | Unity Health Systems | 40037 | 4434800% |  | Gantry can't comply due to table hardware issue; Performed hardware reset with a power cycle to resolve. | Work Order #: WO-03581616 Ali, Adam The system fully meets specification and returned to use Problem Description by Engineer : Problem description by engineer : Gantry can't comply When was the first occurrence? Within... : 01-Jun-2021 How often is issue occurring? : Continuously Malfunction area : Table Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No further information is available as per the customer Expected and Unexpected results: Not provided Patient Impact/Outcome: No further information is available as per the customer User Impact: Unable to scan Current Software Version: 4.1 Resolution : System went down with a 'Gantry cannot comply error'. Bruce made a bugreport & also power cycled the system. The error has cleared up and tests scanned ran with no issues. I dialed into system, checked logs & also downloaded bugreport. Logs shows a table hardware issue caused an obstruction error which was cleared once hardware reset with a power cycle. Internal Comments : Error # and/or description of error : 2021/06/01 05:36:02.838__02 S_HMC_HOR_MOTOR_RELAY_OPEN NORMAL 0 0 3 0 Hor Motor Relay open 5.71928240740741 137.262777777778 8235.76666666667 2021/06/01 05:34:59.635__01 S_HMC_CTRL_A_POSITION_ERROR_LIMIT_EXCEEDED NORMAL 25 0 3 0 Check for couch obstruction 5.71854166666667 137.245 8234.7 2021/06/01 05:28:47.473__02 S_HMC_AXIS_WRONG_DEST_REACHED NORMAL 50223 50270 3 0 EndStatus: 0 5.71424768518519 137.141944444444 8228.51666666667 ----------------------------------------------------------------------------- | 4 |  | System went down with a 'Gantry cannot comply error'. Bruce made a bugreport & also power cycled the system. The error has cleared up and tests scanned ran with no issues. Bruce made a bugreport, collect bugreport for further analysis | [12 empty] | 117631872 |  | 959979 |  | Closed - Done | Closed - Done |  | TBD_Bugreport | [5 empty] | Engineering | [6 empty] | Gantry comply |  |  | Completed |  |  | Refresh |  |  | System issue | System | Console-SW | Bugrep | Console-SW-Bugrep | Console-SW-Bugrep | TBD_Bugreport | TBD_Bugreport | Console-SW | Console-SW-Bugrep | No | y | y | y | y | y | No | y | 0 | y | y | y | y | y | y | 0 | y | y
990 |  |  | Sweden | Falu Lasarett | 40049 | 4435100% |  | Application crash after reconstruction. Email attached from cust. | Work Order #: WO-03602652 Knudsen Martin The system fully meets specification and returned to use Problem Description by Engineer : Friday: User experienced yellow screen and needed to reboot system. Have worked since whole week. Diagnostic performed by Engineer : No bugreport was created. Philips asks to close case since there is no bugreport. If error comes back, they will give a new call. ----------------------------------------------------------------------------- | 3 |  | Yellow Screen on HF3, sync with L3 to see if need workaround | [12 empty] | 117652074 |  | 973711 |  | Opened - Pending Translation | Closed - Done |  | CTUDT00632520 | [5 empty] | Engineering | [6 empty] | New Yellow Screen |  |  | Completed |  |  | Refresh |  |  | System issue | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
991 |  |  | China | 浙江省乐清第一人民医院 | 40054 | 4434900% |  | Wechat:02N/Ingenuity CT 728326/run slowly | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : run slowly Safety Questions : Case: 0117637652 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Tao Fengming Last updated by: Tao Fengming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 系统卡顿 设备本身技术问题： 是。 Troubleshooting Action: 远程诊断与测试结果： 怀疑软件，显卡。。。问题 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议工程师到场检查主机散热，显卡，重新安装软件 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 03.06.2021 09:34:45 UTC SUP JCOUSER (86036249)服务内容：软件死机黄屏，清洁垃圾数据，清理操作台电脑，修改主机到重建柜网段为10网段。 服务结果：维修后，设备恢复正常工作。 其它重要事项：无 Internal Remarks : 03.06.2021 09:34:51 UTC Yi Huang (86036249) RDI Accuracy:Yes $Q7: (Confirmed reported problem&system yellow screen and hung up) $Q8: (check the hosp network IP is 192.168, same to host to irs) $Q9: (clean the console and change network to 10.168 address) $Q10: (Test and inspection data refer to: PA.Device returned to full functionality.) $Q11: Calibrated tools name(N/A) S/N (N/A) Next calibration date (N/A) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date:N/A ^Current Tube usage: second:N/A ----------------------------------------------------------------------------- | 3 |  | $Q7: (Confirmed reported problem&system yellow screen and hung up) $Q8: (check the hosp network IP is 192.168, same to host to irs) $Q9: (clean the console and change network to 10.168 address) | [12 empty] | 117637652 |  | 968645 |  | Opened - Pending Translation | Closed - No Further Investigation |  | CTUDT00632520 | [5 empty] | Engineering | [6 empty] | New Yellow Screen |  |  | Completed |  |  | Refresh |  |  | System issue | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
992 |  |  | China | 浙江省乐清第一人民医院 | 40054 | 4434800% |  | 02N/Ingenuity CT 728326/report error | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: Instructed Customer Safety Questions : Case: 0117631325 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Qi Xiaoyu Last updated by: Qi Xiaoyu 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Problem Reported by Customer : 02N/Ingenuity CT 728326 T2 Activities : Problem Description: report error Troubleshooting Action: report error Repair Action: Customer click off the error, and then scan normally. Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: NA SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | Customer click off the error, and then scan normally. | [12 empty] | 117631325 |  | 963986 |  | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Unknown |  |  | Completed |  |  | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
993 |  |  | China | 浙江省乐清第一人民医院 | 40054 | 4434800% |  | Wechat:02N/Ingenuity CT 728326/CIRS boot issue | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer : Correct the error Safety Questions : Case: 0117628950 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Tao Fengming Last updated by: Tao Fengming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: Cannot scan or STC Troubleshooting Action: CIRS disconnection, reboot but still failed. Check logger has no error and CIRS ping sucess. Wait and try again, issue solved Repair Action: Wait for CIRS boot and try again, issue solved Test/Inspection: N/A Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | CIRS disconnection, reboot but still failed. Check logger has no error and CIRS ping sucess. Wait and try again, issue solved Repair Action Wait for CIRS boot and try again, issue solved ask FSE to check if Service node setting changed or not, and see if have implement in OPS | [12 empty] | 117628950 |  | 962567 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | NotEngineering | [5 empty] | Engineering | [6 empty] | Cannot scan |  |  | Completed |  |  | Refresh |  |  | System issue | System | NotEngineering | ShouldbeService | NotEngineering-ShouldbeService | NotEngineering-ShouldbeService | NotEngineering | NotEngineering-ShouldbeService | NotEngineering | NotEngineering-ShouldbeService | No | y | y | y | y | y | No | y | 0 | y | y | y | y | y | y | 0 | y | y
994 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 337000 | 4434600% |  | Device does not start shooting | Work Order #: WO-03573165 Can Özgenlik The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : Cihazın verdiği hatalar incelendi. Hataların sonucu olarak slipring ile ilgili hatalar görüldü. Öncelike slipring temizlenmeli ve yağlanmalıdır. ----------------------------------------------------------------------------- Work Order #: WO-03573167 03-Jun-2021 Tunay Uysal The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Cihaz çekime başlamıyor Resolution : Yapılan incelemede slipring data iletim hatası görüldü. Slipring bakımı yapıldı. Sistem klinikte bulunan fantom ile test edilerek kullanıma hazır bir şekilde teslim edildi. ----------------------------------------------------------------------------- | 3 |  | In the examination, slipring data transmission error was observed. Slipring maintenance done. The system was tested with the phantom in the clinic and delivered ready to use. Errors given by the device have been examined. Errors related to slipring were seen as a result of the errors. First of all, the slipring should be cleaned and lubricated. | [12 empty] | 117621514 |  | 957525 |  | Evaluation - Pending Translation | Closed - No Further Investigation |  | NotEngineering-SlipringClean | [5 empty] | Engineering | [6 empty] | Slipring |  |  | Completed |  |  | Refresh |  |  | System issue | System | Gantry-HW | SlipRingClean | Gantry-HW-SlipRingClean | Gantry-HW-SlipRingClean | NotEngineering-SlipringClean | SlipRingClean | Gantry-HW | Gantry-HW-SlipRingClean | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
995 |  |  | Turkey | ADIYAMANOZELSAGINSTURTASSANVETICA.S | 341006 | 4435000% |  | There is a data issue | Work Order #: WO-03596130 Ugur,Demirbas The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : There are fluctuations in the ECG signal, BPM is OK. Diagnostic performed by Engineer : possible cable problem between gantry and ECG Follow Up Required: If necessary, the cable between the ECG gantry needs to be changed. otherwise, ekg probes, ekg communication line ibox should be checked. It is useful to go with the ECG simulator. | 3 |  | If necessary, the cable between the ECG gantry needs to be changed. otherwise, ekg probes, ekg communication line ibox should be checked. It is useful to go with the ECG simulator. | [12 empty] | 117645950 |  | 969776 |  | Opened - Pending Translation | Evaluation - Pending Translation | [7 empty] | Others | [6 empty] | ECG |  |  | Completed |  |  | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
996 |  |  | Turkey | ADIYAMAN OZEL SAGINSTURTAS SAN VE TIC A.S | 341006 | 4435000% |  | Remote control is out of use. | Work Order #: WO-03596142 Ugur, Demirbas The system partly meets specification and is returned to use with restrictions Diagnostic performed by Engineer : Keypad replacement is required. Problem Description by Engineer: Console down button is out of use. Follow Up Required: Keypad replacement. ----------------------------------------------------------------------------- | 3 |  | down button not working Under failuare part anaysis | [12 empty] | 117645969 |  | 969613 |  | Opened | Evaluation |  | CTUDT00632185 | [5 empty] | Engineering | [6 empty] | CT Box |  |  | Completed |  |  | Refresh |  |  | System issue | System | Gantry-HW | CT Box | Gantry-HW-CT Box | Gantry-HW-CT Box | CTUDT00632185 | CT Box | Gantry-HW | Gantry-HW-CT Box | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
997 |  |  | China | Zhejiang Province Yueqing No. 1 People’s Hospital | 40054 | 4434800% |  | 02N/Ingenuity CT 728326/equipment information request | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Environmental Cause Repair Activity Code: Instructed Customer Problem Reported by Customer: 02N/Ingenuity CT 728326 T2 Activities:Problem Description: customer ask equipment information. Troubleshooting Action: customer ask equipment information. Repair Action: check the information, and reply to customer. Test/Inspection: system is ok. Test Equipment: N/A Safety Questions:Case: 0117629563 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | 2021-06-01 | [12 empty] | 117629563 |  | 962976 |  |  | Closed - Non-Complaint | [14 empty] | 2021-06-01 | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
998 |  |  | Japan | 社会医療法人 母恋 日鋼記念病院 | 341011 | 4434700% |  | Since after installation, detector failure 5times . Installed date:Dec-28-2020 |  | CFM |  | Duplicate Check customer temp, and whether need BIU support | [14 empty] | 968521 |  | Investigation | Investigation |  | CTUDT00632110 | [5 empty] | Engineering | [6 empty] | Artifact |  |  | Completed |  |  | Refresh |  |  | System issue | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | 0 |  |  | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | y | y | y | y | y | 0 | y | #REF!
999 |  |  | Japan | 社会医療法人 母恋 日鋼記念病院 | 341011 | 4434700% |  | 04DS:KKI: Many DMS module failures | Detector failure frequency is high [Phenomenon confirmation] I was able to confirm. [Details of work] Since the frequency of detector failures from the time of installation is high, countermeasures have been taken. We have replaced the motherboard of the data measurement system (DMS). After the replacement, we made various adjustments and confirmed that there was no problem. We checked the operation and conducted a test scan to confirm that there were no problems. [Other failures / treatments] None [Device status] Equipment can be used [Work status] done [Other] Parts used Part number: 459800263491 / Part name: DMS MOTHER BOARD RIGHT / Serial number: ー / Batch number: ー / Quantity: 1 / List price: ー | 5 |  | We have replaced the motherboard of the data measurement system (DMS). After the replacement, we made various adjustments and confirmed that there was no problem. 459800263491 / Part name: DMS MOTHER BOARD RIGHT | [12 empty] | 117625053 |  | 959140 |  | Evaluation | Closed - No Further Investigation |  | DMSMotherboard | [5 empty] | Engineering | [6 empty] | DMS MB |  |  | Completed |  |  | Refresh | 459800263491 | DMS MOTHER BOARD RIGHT | System issue | System | DMS-HW | DMS Motherboard | DMS-HW-DMS Motherboard | DMS-HW-DMS Motherboard | should be DEFOA | DMS Motherboard | DMS-HW | DMS-HW-DMS Motherboard | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
1000 |  |  | United States | Philips Medical Capital | 40005 | 4434800% |  | 4.1.7 Brain Perfusion training | Work Order #: WO-03585212 Giles Brian The system fully meets specification and returned to use Problem description by Clinical Support Specialist : Customer Function/Role: Tech Device Use: In Use Expected and Unexpected Results: Wants Brain Perfusion Training User Impact: None Patient Impact: None Current Software Version: 4.1.7 Diagnostic performed by Engineer : Sent customer to Fisher, Kelly Clinical Services Project Manager Cust Svc Clin Ed - IS Southeast Zone Diagnostic performed by Engineer : Contacted Dianne and she was communicating with Kelly Fisher from scheduling. Case closed. ----------------------------------------------------------------------------- | 5 |  | 2021-06-01 | [12 empty] | 117634665 |  | 961938 |  |  | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1001 |  |  | Japan | 社会医療法人 母恋 日鋼記念病院 | 341011 | 4434600% |  | 01CS: KKI: Air cal. Does not run >> No hardware error> It seems that it did not run due to the problem of bed position. Wait and see | Original Problem Reported by Customer 01CS: KKI: Air cal. Does not run >> No hardware error> It seems that it did not run due to the problem of bed position. Wait and see Problem Description Problem description by engineer: Target device: CT Failure location: 01Cosole When did it first occur? (Within ~): 1 Hour How often do problems occur ?: First Occurrence At what point in the end-user workflow is the problem occurring ?: Before inspection / measurement starts (Pre-Acquistion / Scan) What happened recently ?: N.A. Does the customer create a specific system event report for the problem ?: No. The customer was unable to create a log file or report. Troubleshooting Action What steps are taken during the system diagnosis: Post-response results: Improved Failure location: 01Cosole What steps were taken during the diagnosis ?: Call content Air cal. Doesn't run Correspondence As a result of log confirmation, there is no hardware error It seems that he did not run due to the problem of bed position. We have decided that it is not a problem and will wait and see. Error # and / or description of error: N / A Repair Action Completed or Recommended Repair Action "External": [Phenomenon confirmation] Confirmed [Details of work] Call content Air cal. Doesn't run Correspondence As a result of log confirmation, there is no hardware error It seems that he did not run due to the problem of bed position. We have decided that it is not a problem and will wait and see. [Confirmation details] Good | 4 |  | there is no hardware error It seems that he did not run due to the problem of bed position. We have decided that it is not a problem and will wait and see. | [12 empty] | 117622122 |  | 958191 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | NotEngineering | [5 empty] | Engineering | [6 empty] | Air cal. |  |  | Completed |  |  | Refresh |  |  | System issue | System | NotEngineering | ShouldbeOthers | NotEngineering-ShouldbeOthers | NotEngineering-ShouldbeOthers | NotEngineering | NotEngineering-ShouldbeOthers | NotEngineering | NotEngineering-ShouldbeOthers | No | y | y | y | y | y | No | y | 0 | y | y | y | y | y | y | 0 | y | y
1002 |  |  | Japan | 社会医療法人 母恋 日鋼記念病院 | 341011 | 4434600% |  | 04DS:KKI: Ring artifact appeared >> Appeared near M18 or M25 > Did not appear in subsequent test scan | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 01DS: KKI: Artifact appeared Safety Questions:Case: 0117622065 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: KAZUMASA KIMURA Last updated by: KAZUMASA KIMURA 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Support SAP Message: Components Added Succesfully to SWO!- SAP Message: LOGI Task added Successfully!- T2 Activities:Problem Description: Problem description by engineer : Equipment affected: CT Problem location: 04 Gantry Exclude HV When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user's work flow did the problem occur?: Post-Acquisition/Scan What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem? No.The customer has not created a log file or report. Troubleshooting Action: What steps are taken during the system diagnosis: Result of response: Problem solved. Problem location: 04 Gantry Exclude HV What steps did you perform to diagnose the problem? Call description: Ring artifact occurrence Artifacts appear even when re-scanning immediately. [Response] Artifacts did not appear in subsequent test scans. [Image confirmation results] The problem occurred near M18 or M25. Artifacts are coming from detectors close to where they appeared in March. Error # and/or description of error: N/A Repair Action: Completed or Recommended Repair Action "External": The service engineer for the site visited and worked on the equipment. OneEMS internal remarks : Problem description by engineer : Equipment affected: CT Problem location: 04 Gantry Exclude HV When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user's work flow did the problem occur?: Post-Acquisition/Scan What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem? No.The customer has not created a log file or report. What steps are taken during the system diagnosis: Result of response: Problem solved. Problem location: 04 Gantry Exclude HV What steps did you perform to diagnose the problem? Call description: Ring artifact occurrence Artifacts appear even when re-scanning immediately. [Response] Artifacts did not appear in subsequent test scans. [Image confirmation results] The problem occurred near M18 or M25. Artifacts are coming from detectors close to where they appeared in March. Error # and/or description of error: N/A Completed or Recommended Repair Action "External" : The service engineer for the site visited and worked on the equipment. Completed or Recommended Repair Action "Internal": --Troubleshooting Action-- Result of response: Problem solved. Problem location: 04 Gantry Exclude HV What steps did you perform to diagnose the problem? Call description: Ring artifact occurrence Artifacts appear even when re-scanning immediately. [Response] Artifacts did not appear in subsequent test scans. [Image confirmation results] The problem occurred near M18 or M25. Artifacts are coming from detectors close to where they appeared in March. --Repair Action-- Recommended 12Nc: PR1 459800911271:TDMS V2 M.Det Class A FRU QTY:2 SLOC: Sapporo (U612) --Subject-- 04DS:KKI: Mr./Ms. Ando: Ring artifact appeared >> Appeared near M18 or M25 > Did not appear in subsequent test scan --Problem Description-- Equipment affected: CT Problem location: 04 Gantry Exclude HV When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user's work flow did the problem occur?: Post-Acquisition/Scan What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem? No.The customer has not created a log file or report. ----------------------------------------------------------------------------- | 4 |  | The problem occurred near M18 or M25. Artifacts are coming from detectors close to where they appeared in March. --Repair Action-- Recommended 12Nc: PR1 459800911271:TDMS V2 M.Det Class A FRU QTY:2 SLOC: Sapporo (U612) finally system works after replace DMS MB | [12 empty] | 117622065 |  | 958147 |  | Investigation | Closed - Done |  | CTUDT00632110 | [5 empty] | Engineering | [6 empty] | Artifact |  |  | Completed |  |  | Refresh |  |  | System issue | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | 0 |  |  | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | y | y | y | y | y | 0 | y | #REF!
1003 |  |  | Japan | 苫小牧市立病院 | 341014 | 4435200% |  | 03CS:KK:Air Calibration中にE-Stop... | Original Problem Reported by Customer 03CS: KK: E-Stop Open occurred during Air Calibration. >> Restored by closing again. >> S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT, S_GEN_FAULT_CATHODE_TUBE_ARC, szInfo: SoftArcReg 0x211 * Problem Description Problem description by engineer: Target device: CT Failure location: 03 High Voltage Energy When did it first occur? (Within ~): 1 Hour How often do problems occur ?: First Occurrence At what point in the end-user workflow is the problem occurring ?: Setting up Protocol What happened recently ?: N.A. Does the customer create a specific system event report for the problem ?: No. The customer was unable to create a log file or report. Troubleshooting Action What steps are taken during the system diagnosis: Post-response results: Improved Failure location: 03 High Voltage Energy What steps were taken during the diagnosis ?: Restored by E-Stop re-Close. It seems that an arc was generated, S_TIMER_WATCHDOG_CAUSED_LAST_RESET was applied, and E-Stop Open was performed. I was allowed to wait and see. | 4 |  | Restored by E-Stop re-Close. It seems that an arc was generated, S_TIMER_WATCHDOG_CAUSED_LAST_RESET was applied, and E-Stop Open was performed. I was allowed to wait and see. | [12 empty] | 117655296 |  | 978314 |  | Evaluation - Pending Translation | Closed - No Further Investigation |  | CTUDT00630085 | [5 empty] | Engineering | [6 empty] | Tube Arc |  |  | Completed |  |  | Refresh |  |  | System issue | System | IC-HW | Tub Arc | IC-HW-Tub Arc | IC-HW-Tub Arc | CTUDT00630085 | Tub Arc | IC-HW | IC-HW-Tub Arc | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
1004 |  |  | Japan | Maki Hospital | 341017 | 4435100% |  | The scan stopped. | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : A scan stopped. ----------------------------------------------------------------------------- | 4 |  | No info | [12 empty] | 117650914 |  | 976801 |  | Evaluation | Investigation | [7 empty] | TBD | [6 empty] | TBD |  |  | Completed |  |  | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1005 |  |  | China | Jiangsu Province Yangzhou City No. 1 People's Hospital, West District Hospital | 345022 | 4434700% |  | 02N/Ingenuity Core128/Elite China/dicom transfer problem | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity Core128/Elite China SAP Transfer : Case Transferred to SAP Safety Questions : Case: 0117622669 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 其他类型故障。 客户问题详细描述或故障现象： 图像传输失败 设备本身技术问题： 否。 Troubleshooting Action: 关键故障代码或现象： CT 传PACS有图传不过去 远程诊断与测试结果： 客户要求工程师到场处理 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 查看传输失败的图像，检查pacs设置 SAP Transfer : Dispatched for Onsite Support T2 Activities : Problem Description: 客户问题概述： 其他类型故障。 客户问题详细描述或故障现象： estop 断开 设备本身技术问题： 否。 Troubleshooting Action: 关键故障代码或现象： S_TILT_C_TO_BAD_SERIAL 远程诊断与测试结果： 通讯故障，R2D引起 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 更换R2D 测试 建议备件12NC及名称： 459800467343 R2D Assembly Next Step Action : 013951434175再次来电，联系 qi liangdong回电 External Remarks : 06.06.2021 09:05:07 UTC Liangdong Qi (86036024)服务内容：删除垃圾文件，修改配置参数. 服务结果： 维修后，设备恢复正常工作. Internal Remarks : 06.06.2021 09:05:10 UTC Liangdong Qi (86036024) "RDI Accuracy: Yes“ $Q7: (Confirmed reported problem, the Dose structured Report tansfer failed from CT host to pacs.) $Q8: (No error report, checked the pacs and found it can't accepted the DSR.) $Q9: (Agjusted the image transfer parameter, disable rhe Dose Structure Report option.) $Q10: (Test and Inspection data refer to PA & Device return to fully function) $Q11: Calibrated Tools (NA). ----------------------------------------------------------------------------- | 3 |  | $Q7: (Confirmed reported problem, the Dose structured Report tansfer failed from CT host to pacs.) $Q8: (No error report, checked the pacs and found it can't accepted the DSR.) $Q9: (Agjusted the image transfer parameter, disable rhe Dose Structure Report option.) | [12 empty] | 117622669 |  | 958409 |  | Opened - Pending Translation | Closed - No Further Investigation |  | CTUDT00633122 | [5 empty] | Engineering | [6 empty] | Audit Repository |  |  | Completed |  |  | Refresh |  |  | System issue | System | Console-SW | AuditRepository | Console-SW-AuditRepository | Console-SW-AuditRepository | CTUDT00633122 | AuditRepository | Console-SW | Console-SW-AuditRepository | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
1006 |  |  | China | Nanjing Medical University Affiliated Yifu Hospital | 345028 | 4434900% |  | 02N/Core128/software crash | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 02N/Core128/ SAP Transfer:Dispatched for On site Support SAP Transfer:Case Transferred to SAP T2 Activities:Problem Description: Summary of customer's problems: Often crash Detailed description of Customer's problems or malfunctions: Often crash It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: NA Remote troubleshoot and test result: The customer said the machine often crashes, with yellow screen PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: patch according to Yellow screen workaround. Collect bugreport if necessary Safety Questions:Case: 0117636136 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created by: Wennian Han Last updated by: Wennian Han 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 3 |  | Yellow screen workaround | [12 empty] | 117636136 |  | 967619 |  | Opened | Closed - No Further Investigation |  | CTUDT00632520 | [5 empty] | Engineering | [6 empty] | New Yellow Screen |  |  | Completed |  |  | Refresh |  |  | System issue | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #REF!
1007 |  |  | China | 林甸县医院 | 345038 | 4435000% |  | Wechat:02N/Ingenuity Core128/Elite China/Consultation questions | BU Q&R Feedback: No SAP Customer Symptom Code: CENV Customer Symptom Code: Customer Environmental Concern Reason: Corrective Maintenance Engineer Failure Code: Environmental Cause Repair Activity Code: Instructed Customer Problem Reported by Customer : Consultation questions Safety Questions : Case: 0117643840 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Zhou Xiaohong Last updated by: Zhou Xiaohong 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Troubleshooting Action: answer customer question Repair Action: answer customer question Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | answer customer question | [12 empty] | 117643840 |  | 972650 |  | Opened - Pending Translation | Closed - non-complaint | [7 empty] | TBD | [6 empty] | NAC |  |  | Completed |  |  | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1008 |  |  | United States | Select Medical Corporation | 310050 | 4435000% |  | add MWL to CT | Work Order #: WO-03596780 Franklin Rex Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 5 |  | add Worklist to CT | [12 empty] | 117646423 |  | 974268 |  | Opened | Closed - non-complaint | [7 empty] | TBD | [6 empty] | NAC |  |  | Complete |  |  | Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1009 |  |  | United States | AdventHealth Zephyrhills | 300099 | 4435000% |  | 1302CFXXX Unable to send to pacs & worklist not updating | Work Order #: WO-03594294 Ali, Adam The system fully meets specification and returned to use Problem Description by Engineer : Problem description by engineer : Unable to send to pacs & worklist not updating When was the first occurrence? Within... : 03-Jun-2021 How often is issue occurring? : Continuously Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No further information is available as per the customer Expected and Unexpected results: Not provided Patient Impact/Outcome: No further information is available as per the customer User Impact: Cannot send images to PACS or pull new patient info from worklist Current Software Version: 4.1.7.25030 Resolution : Restarted Host computer. Worklist is not updating, also images are transferring to PACS. Internal Comments : Error # and/or description of error : Unable to send to pacs & worklist not updating Diagnostic performed by Engineer : Unable to send to pacs & worklist not updating. Restarted Host & CIRS computers. Worklist is updating, also images are transferring to PACS. ----------------------------------------------------------------------------- | 3 |  | Restarted Host & CIRS computers.6/3/2021 collect bugreport if have | [12 empty] | 117643789 |  | 968474 |  | Evaluation | Evaluation |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS Transfer |  |  | Complete |  |  | Refresh |  |  | Upgrade issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
1010 |  |  | United States | Select Specialty Hospitals | 310051 | 4434700% | 22 | Device shutdown due to power went out and shut down the CIRS causing a disconnect; Verified system's normal operation. | Work Order #: WO-03578022 01-Jun-2021 Siker Shawn The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : Power went out Friday night and shut down the cirs causing a disconnect, I verified system was working with test scans and handed back to the customer. ----------------------------------------------------------------------------- | 2 |  | Power went out Friday night and shut down the cirs causing a disconnect at Monday, I verified system was working with test scans and handed back to the customer. | [12 empty] | 117627250 |  | 960508 |  | Closed - Done | Closed - Done | [7 empty] | Customer | [6 empty] | CIRS disconnect |  |  | Complete |  |  | Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1011 |  |  | Argentina | INSTITUTO NACIONAL DE | 40027 | 4435500% |  | Detector replacement | Work Order #: WO-03636155 Meza, Edgardo Arturo The system partly meets specification and is returned to use with restrictions Problem Description by Engineer: Intermittent rings in the image. Follow Up Required: Replace detector. Diagnostic performed by Engineer: No signal in the detector module. ----------------------------------------------------------------------------- | 5 |  | No signal in the detector module. Replace DMS module | [12 empty] | 117670171 |  | 989472 | TRUE | Investigation | Investigation |  | CTUDT00632110 | [5 empty] | Engineering | [6 empty] | DMS Module | [5 empty] | Refresh |  |  | System issue | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | DMS Module | DMS-HW | DMS-HW-DMS Module | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | y
1012 |  |  | Argentina | INSTITUTO NACIONAL DE | 40027 | 4435500% |  | TRAINING ON EQUIPMENT IS REQUESTED | Work Order #: WO-03621433 Alec, Preciado The system fully meets specification and returned to use. Problem Description by Engineer : Problem description by engineer : They request a second visit by Applications department to install a new system. Information to support the complaint handling process Customer Function/Role: BioMed Device Use: No patient involved. Expected and Unexpected results: No further information is available as per the customer Patient Impact/Outcome: It does not impact. User Impact: It does not impact. Current Software Version: 4.1.7 Diagnostic Performed by Engineer: They request a second visit by Applications department for a new system installed on March. The client says that they already contacted the application expert Fernando Dominguez. Internal Comments : Error # and/or description of error : NA Completed or Recommended Repair Action "Internal" : They request a second visit by Applications department for a new system installed on March. The client says that they already contacted the application expert Fernando Dominguez. Follow Up Required: They request a second visit by Applications department for a new system installed on March. The client says that they already contacted the application expert Fernando Dominguez. Resolution: They request a second visit by Applications department for a new system installed on March. The client says that they already contacted the application expert Fernando Dominguez. ----------------------------------------------------------------------------- | 5 |  | They request a second visit by Applications department for a new system installed on March. | [12 empty] | 117669723 |  | 985226 |  |  | Closed - Non-Complaint | [7 empty] | NAC | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1013 |  |  | China | 山东安丘市中医医院 | 345017 | 4436000% |  | 02N/Ingenuity Core128/Elite China/couch fault | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : couch fault Next Step Action : 派工 yang xingbo 13583235930 Safety Questions : Case: 0117692556 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Guomei Li Last updated by: Guomei Li 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support ----------------------------------------------------------------------------- | Closed - Done |  | 服务内容：调整tape switch 位置。 服务结果：维修后，设备恢复正常工作。 | [12 empty] | 117692556 |  | 1001027 | TRUE | Closed - Done | Closed - Done |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | Tape switch | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | TabSwitch | Couch-HW-TabSwitch | Couch-HW-TabSwitch | CTUDT00632127 | TabSwitch | Couch-HW | Couch-HW-TabSwitch | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | y
1014 |  |  | China | 浙江省乐清第一人民医院 | 40054 | 4435900% |  | 02N/Ingenuity CT 728326/Connect the door machine linkage | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : Connect the door machine linkage Safety Questions : Case: 0117692011 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Zixuan Last updated by: Wang Zixuan 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 12.06.2021 12:40:31 UTC Yi Huang (86036249)服务内容：连接门控开关线。 服务结果：维修后，设备恢复正常工作。 其它重要事项：无 Internal Remarks : 12.06.2021 12:40:34 UTC Yi Huang (86036249) RDI Accuracy:Yes $Q7: (Confirmed reported problem&connect door switch cable) $Q8: (connect door switch cable) $Q9: (connect door switch cable) $Q10: (Test and inspection data refer to: PA.Device returned to full functionality.) $Q11: Calibrated tools name(N/A) S/N (N/A) Next calibration date (N/A) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date:N/A ^Current Tube usage: second:N/A ----------------------------------------------------------------------------- | Closed - No Further Investigation |  | 服务内容：连接门控开关线。 服务结果：维修后，设备恢复正常工作。 | [12 empty] | 117692011 |  | 1000613 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Service | [6 empty] | System Installation | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1015 |  |  | China | 扎兰屯市人民医院 | 40045 | 4435800% |  | 02N/Ingenuity CT 728326/cann't log in | BU Q&R Feedback: No SAP Customer Symptom Code: CENV Customer Symptom Code: Customer Environmental Concern Reason: Corrective Maintenance Engineer Failure Code: Environmental Cause Repair Activity Code: Configuration Change Problem Reported by Customer : 02N/Ingenuity CT 728326 Safety Questions : Case: 0117685942 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Weng Jinming Last updated by: Weng Jinming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Troubleshooting Action: instruct customer to change password Repair Action: instruct customer to change password Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | Closed - No Further Investigation |  | instruct customer to change password | [12 empty] | 117685942 |  | 999974 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Password | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1016 |  |  | China | 内蒙古兴安盟人民医院 | 345030 | 4435600% |  | 02N/Ingenuity Core128/Elite China/scan room leakage of rain | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Calibration/Adjustment Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ Safety Questions : Case: 0117672104 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Gaoge Last updated by: Wang Gaoge 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 其他类型故障。 客户问题详细描述或故障现象： 房间漏雨到设备上，需要工程师到场检查 设备本身技术问题： 否。 Troubleshooting Action: 关键故障代码或现象： 设备上漏水了 远程诊断与测试结果： 到场根据实际情况清理设备 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 1，清理水迹，检查设备 Test/Inspection: N/A Test Equipment: N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Internal Remarks : 11.06.2021 01:08:19 UTC Ruijie Su (86039043) $Q7: Confirmed reported problem,Air conditioning leaking。 . $Q8: on sit service ,Air conditioning leaking cpm arc. . need clean apm plug . $Q9: clean apm plug . .system is ok! $Q10:Test and inspection data refer to:PA. Device returned to full functionality. $Q11: Calibrated Toolsname ( Multimeter Fluke 287C) S/N (31060028) Next Calibrationdate (06/1do bed calibration and air calibration31/10/2020) 2( Torque Wrench Kit ) S/N E201602048 Next Calibration date22/03 /2022 For Mainland China CV/ XR, CTengineers only.Please fill in: ^ Current Tubeinstalled date 10/01/2017 ^ Current Tube usage: scan second 6589 External Remarks : 11.06.2021 01:08:16 UTC Ruijie Su (86039043)服务结果： [检查发现打火清洁CPM高压插头后恢复 ，设备恢复正常工作。] 其它重要事项：无 ----------------------------------------------------------------------------- | 5 |  | 服务结果： [检查发现打火清洁CPM高压插头后恢复 ，设备恢复正常工作。] on sit service ,Air conditioning leaking cpm arc. . need clean apm plug . $Q9: clean apm plug . .system is ok! $Q10:Test and inspection data refer to:PA. Device returned to full functionality. | [12 empty] | 117672104 |  | 990986 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Rain leakage | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1017 |  |  | China | Shandong Province Chest Hospital | 40025 | 4435500% |  | WeChat:02N/Ingenuity CT 728326/Couch fault | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : Couch fault Safety Questions : Case: 0117665311 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Bo Last updated by: Wang Bo 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 扫描床水平运动故障。 客户问题详细描述或故障现象： 机架掉电。 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： VELOCITY_ENCODER_FAILURE,HOR_MOTION_STOP_ERROR,COUCH IS JOGGING. 远程诊断与测试结果： 重启无效。 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议做床水平校准，检查床水平读数，velocity读数偏差过大，cantrace里床表现卡顿检查床侧板是否变形。检查水平电机皮带松紧。短接tapeswitch。 建议备件12NC及名称： PPD 453566489521，459800712461 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : PPD备件FRU, AMC upgrade KITS 459800712461*1&MCG /PITTMAN MOTOR ASSEMBLY 453566489521*1已安排24H到场地，FSE：黄兆坤-86-18654536122 External Remarks : 08.06.2021 11:59:35 UTC SUP JCOUSER (86038405)服务内容：扫描床水平校正，垂直校正，调节扫描床运动阻力。 服务结果：维修后，设备恢复正常使用。 Internal Remarks : 12.06.2021 08:46:26 UTC Zhaokun Huang (86038405) "RDI Accuracy: YES." $Q7: ( Confirmed reported problem & Couch fault ) $Q8: ( Check error log , report "S_COUCHMGR_HOR_MOTION_STOP_ERROR". Manual push couch , find resistance serious .) $Q9: ( Adjust Couch movement resistance.) $Q10: ( Test and inspection data refer to: PA . Device returned to full functionality.) $Q11: Calibrated Tools name (FLUKE287 ) S/N (23350170) Next Calibration date (04/07/2022) ^Current Tube installed date: N/A ^Current Tube usage: N/A ----------------------------------------------------------------------------- | Closed - No Further Investigation |  | Check error log , report "S_COUCHMGR_HOR_MOTION_STOP_ERROR". Manual push couch , find resistance serious 服务内容：扫描床水平校正，垂直校正，调节扫描床运动阻力。 服务结果：维修后，设备恢复正常使用。 | [12 empty] | 117665311 |  | 986914 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | AssemblyAdjust | [5 empty] | Engineering | [6 empty] | Couch Vertical Movement | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | AssemblyAdjust | Couch-HW-AssemblyAdjust | Couch-HW-AssemblyAdjust | AssemblyAdjust |  | Couch-HW | Couch-HW-AssemblyAdjust | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | y
1018 |  |  | China | Guangxi Province Teng County People’s Hospital | 40036 | 4435500% |  | 02N/Ingenuity CT 728326/application | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer:application Next Step Action : yang jun &guo fan follow Safety Questions:Case: 0117664378 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingli Last updated by: Lu Mingli 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | Closed - Non-Complaint |  | NAC | [12 empty] | 117664378 |  | 986558 |  |  | Closed - Non-Complaint | [7 empty] | NAC | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1019 |  |  | United States | Unity Health Systems | 40037 | 4435500% |  | Gantry Cannot Initialize | Work Order #: WO-03623534 08-Jun-2021 Taylor, Durrance The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: tech Device Use: patient scan Expected and Unexpected Results: patient scanned/gantry cannot initialize User Impact: delayed Patient Impact: delayed Current Software Version: 4.1.10 Resolution : Found the couch horizontal movement had some resistance when table was inside gantry. Adjusted the horizontal drive screw assy and verified normal operation. | 2 |  | Found the couch horizontal movement had some resistance when table was inside gantry. Adjusted the horizontal drive screw assy and verified normal operation. | [12 empty] | 117671365 |  | 986152 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | AssemblyAdjust | [5 empty] | Engineering | [6 empty] | Couch horizontal drive | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | AssemblyAdjust | Couch-HW-AssemblyAdjust | Couch-HW-AssemblyAdjust | AssemblyAdjust |  | Couch-HW | Couch-HW-AssemblyAdjust | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | y
1020 |  |  | Peru | BANCO BBVA PERU | 40034 | 4435600% |  | 0117675674-THE TUBE HEATING IS NOT DONE | Work Order #: WO-03628027 Caldelas Aldana, Carlos Manuel The system does not meet specification and is out of use Problem Description by Engineer : Problem description by engineer : It does not allow them to discharge. When was the first occurrence? Within... : 09-Jun-2021 07:00:00 How often is issue occurring? : First occurrence Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected Results: Discharge is aborted. Patient Impact/Outcome: N/A User Impact: It does not allow scanning. Current Software Version: 4.1.10 Diagnostic Performed by Engineer: It is checked that even after a total shut down, the equipment does not allow scans of any type, not even the X-ray tube heating. We try to access remotely but without success, so we cannot get further details about the diagnosis. Internal Comments : Error # and/or description of error : It does not scan Completed or Recommended Repair Action "Internal" : N/A Follow Up Required : What are the next Troubleshooting steps : Check logs to determine if the cause of the failure is control link, high voltage, etc. Repair and reconfigure PRS connection correctly for this software version and it is working, run functional tests. Error # and/or description of error : It does not scan Completed or Recommended Repair Action "Internal" : N/A ----------------------------------------------------------------------------- Work Order #: WO-03629139 Antonella Haro Problem Description by Engineer : Problem description by engineer : It does not allow them to discharge. When was the first occurrence? Within... : 09-Jun-2021 07:00:00 How often is issue occurring? : First occurrence Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected Results: Discharge is aborted. Patient Impact/Outcome: N/A User Impact: It does not allow scanning. Current Software Version: 4.1.10 Diagnostic performed by Engineer: N/A ----------------------------------------------------------------------------- | 2 |  | Error log is reviewed. Bows are found. High voltage chunks are checked, the one that goes to the generator is very dirty. They are cleaned and lubricated. NO LOAD TEST is performed and the equipment does not generate arcs, generators are discarded. X-ray tube problems, parameters of the well grounded will be reviewed with maintenance personnel from the clinic. Change X-ray tube. After excessive cleaning, lubrication and downtime. The equipment remains operational. Air calibration and scans are performed on patients normally. However, on June 11 artifacts are reported. It is required to measure the wells to earth and diagnose the detectors. Image artifacts. Carry out bad detectors test, check collimator status. | [12 empty] | 117675674 |  | 989268 | TRUE | Closed - Done | Closed - Done |  | ShouldbeService | [5 empty] | Engineering | [6 empty] | Tube | [5 empty] | Refresh |  |  | System issue | System | IC-HW | Tube | IC-HW-Tube | IC-HW-Tube | ShouldbeService |  | IC-HW | IC-HW-Tube | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | y
1021 |  |  | China | Guangxi Province Teng County People’s Hospital | 40036 | 4435500% |  | Application/Ingenuity CT 728326/Application Support | BU Q&R Feedback: No Reason: Application Support Problem Reported by Customer : Application Support Next Step Action: Emailed CT Pre-Sales Training Department Manager Yang Jun, Guo Fan to follow up. Safety Questions:Case: 0117664375 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Zhou Xiaohong Last updated by: Zhou Xiaohong 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action: The customer requests a 2nd on-site application training. ----------------------------------------------------------------------------- | Closed - Non-Complaint |  | NAC | [12 empty] | 117664375 |  | 986555 |  |  | Closed - Non-Complaint | [7 empty] | NAC | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1022 |  |  | China | 江苏大学附属医院 | 345023 | 4435400% |  | 02N/Ingenuity Core128/Elite China/adjust couch | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ Safety Questions : Case: 0117660478 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wennian Han Last updated by: Wennian Han 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 调床 客户问题详细描述或故障现象： 调床 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： NA 远程诊断与测试结果： 报修床围裙刮擦，客户需要来人调节 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 调节床围裙；围裙是活动部件，需和客户沟通好相关。 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 08.06.2021 07:56:00 UTC SUP JCOUSER (86036528)服务内容：扫描床降到最低的时候有异响，调整扫描床底部弹簧位置。 服务结果：服务后，设备恢复正常 Internal Remarks : 08.06.2021 08:04:32 UTC Yang Liu (86036528) RDI Accuracy: Yes $Q7: (Confirm reported problem,the couch has abnormal noise at the lowest level.) $Q8: ( found there is friction between the spring and the spring post） $Q9: (adjust the spring position according to service manual .) $Q10: (Test and inspection data refer to PA. Device returned to full functionality.) $Q11: Calibrated tools refer to PA. ----------------------------------------------------------------------------- | Closed - No Further Investigation |  | 服务内容：扫描床降到最低的时候有异响，调整扫描床底部弹簧位置。 服务结果：服务后，设备恢复正常 | [12 empty] | 117660478 |  | 983565 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00631058 | [5 empty] | Engineering | [6 empty] | Couch Spring | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | Vertical spring | Couch-HW-Vertical spring | Couch-HW-Vertical spring | CTUDT00631058 |  | Couch-HW | Couch-HW-Vertical spring | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | Couch Spring ECR | y | y
1023 |  |  | China | 湖北省武汉仁爱医院 | 345004 | 4435400% |  | 02N/Ingenuity Core128/Elite China/gantry can not close estop | BU Q&R Feedback: No SAP Customer Symptom Code: BOOT Customer Symptom Code: Boot up/Power-up Problem Reason: Corrective Maintenance Engineer Failure Code: Electrical/Electronic Failure Repair Activity Code: Instructed Customer Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ Safety Questions : Case: 0117657807 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Gaoge Last updated by: Wang Gaoge 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 有钥匙框，但是拧钥匙闭合estop没反应，机架不能正常上电 Troubleshooting Action: 指导客户检查紧急开关，找打机架背面的紧急开关被按下，恢复紧急开关后设备恢复正常 Repair Action: 指导客户检查恢复紧急开关 Test/Inspection: N/A Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | Closed - No Further Investigation |  | 指导客户检查紧急开关，找打机架背面的紧急开关被按下，恢复紧急开关后设备恢复正常 | [12 empty] | 117657807 |  | 981894 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | E-Stop button | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1024 |  |  | China | 河北省唐山市工人医院 | 40044 | 4435400% |  | 02N/Ingenuity CT/image artifacts | BU Q&R Feedback: No SAP Customer Symptom Code: ACQP Customer Symptom Code: Imaging/Patient Recording/Acquisition Problem Reason: Corrective Maintenance Engineer Failure Code: Connection/Cable/Interconnect Repair Activity Code: Calibration/Adjustment Problem Reported by Customer : 02N/Ingenuity CT/ Safety Questions : Case: 0117657597 Safety Question Status: In Process PSA: false PSE: false PSF Status: NA Created By: Wang Gaoge Last updated by: Wang Gaoge 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Safety Questions : Case: 0117657597 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Gaoge Last updated by: Wang Gaoge 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 图像有一个半环圆形伪影在图像中心位置 Troubleshooting Action: 检查图像发现有探测器坏道引起伪影，指导客户做空气校正后伪影消失 Repair Action: 指导客户做空气校正 Test/Inspection: N/A Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | Investigation |  | 图像有一个半环圆形伪影在图像中心位置 Troubleshooting Action: 检查图像发现有探测器坏道引起伪影，指导客户做空气校正后伪影消失 | [12 empty] | 117657597 |  | 981699 | TRUE | Investigation | Investigation |  | NotEngineering | [5 empty] | Engineering | [6 empty] | Artifact | [5 empty] | Refresh |  |  | System issue | System | NotEngineering | ShouldbeCustomer | NotEngineering-ShouldbeCustomer | NotEngineering-ShouldbeCustomer | NotEngineering |  | NotEngineering | NotEngineering-ShouldbeCustomer | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | y
1025 |  |  | Japan | 苫小牧市立病院 | 341014 | 4435800% |  | 03DM:KKI:撮影開始時に、エラー発生＞＞SoftArcReg 0x211* /S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT＞継続して検査可 | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 03DM:KKI:撮影開始時に、エラー発生＞＞SoftArcReg 0x211* /S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT＞継続して検査可 Safety Questions : Case: 0117686469 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: KAZUMASA KIMURA Last updated by: KAZUMASA KIMURA 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Message : LOGI Task added Successfully!- SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support SAP Message : Components Added Succesfully to SWO!- ----------------------------------------------------------------------------- | 4 |  | There are replacement parts 12Nc: 989000086352 Name: PHILIPS MRC 880 QTY: 1 We replaced the high-voltage cable on June 14, but the arc occurred again, so we replaced the X-ray tube, which is thought to be the cause. After replacement, make various adjustments. We have confirmed that there is no problem with the overall performance. 12Nc: 989000005011 Name: HV-ANODECABLE PHILIPS MRC880 CT 1,84m QTY: 1 12Nc: 989000003302 Name: HV CABLE PHILIPS MRC CT-O4S 0,95m QTY: 1 | [12 empty] | 117686469 |  | 1001235 | TRUE | Investigation | Investigation |  | NotEngineering | [5 empty] | Engineering | [6 empty] | Tube | [5 empty] | Refresh |  |  | System issue | System | NotEngineering | ShouldbeOthers | NotEngineering-ShouldbeOthers | NotEngineering-ShouldbeOthers | NotEngineering |  | NotEngineering | NotEngineering-ShouldbeOthers | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | y
1026 |  |  | China | 内蒙古阿拉善右旗人民医院 | 345031 | 4435700% |  | 02N/Ingenuity Core128/Elite China/ can not bootup | BU Q&R Feedback: No Reason: Corrective Maintenance Safety Questions : Case: 0117679178 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Zhang Caiqin Last updated by: Zhang Caiqin 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Problem Reported by Customer : 02N/Ingenuity Core128/Elite China T2 Activities : Problem Description: 客户问题概述： 扫描失败。 客户问题详细描述或故障现象： 无法扫描 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： S_MASTER_SEQUENCER_GDC_LOW_AFTER_TIMEOUT 远程诊断与测试结果： "Rail voltage did not go up to at least 500VDC after the soft start contactor needed to be pulled in. The rotor will start to rotate but the Main drive contactor will drop. The E stop GUI will not disappear from the screen. " PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： "Recommendation: Caution: Due to the 480V AC you may have to Kill Main Power wall, in some cases the blue button did not remove the power. 1. Try to close the E stop when the gantry cover switch is midway (cover open) so there is no rotor movement to see if this is caused by rotor movement, it it is, check / clean the slip rings and the brush block. 2. Check the 480V AC fuses to the slipring F351,352,353. 3 .Check the 480V AC fuses on the APM and the KPM F1.2,3. 4. Check the relays on the ACCU. 5. Check CB302. 6.. Check the soft start contactor (SS)and the activation signal (120V) at TB352 / 2: a. If the signal is not there, skip to step 8. b. Check the Aux. Contact side mounted N.C. that is mechanically attached to the right side of the MKD. The signal should be on the bottom of that contactor when the MKD is not pulled in. If the signal is at the top but not at the bottom, replace the Aux contactor. c.. Check the 2 varistors at the bottom of the SS contactor. d. Check the voltage at the bottom of the SS contactor (red wire - voltage, white wire - neutral) if the voltage is there, but the contactor does not pull in, replace the contactor. d. If the SS contactor pulls in check the Soft Start Resistors. 7. If the soft start is OK, with the cover switch midway, check for 480V AC on both the APM and the KPM, if the voltage is not there check the brush block and slip rings, if the voltage is there: a. check the cable from the RHOST to the DMC. (the DMC is reporting to the GHOST on DC level) b. Check / replace the RHOST. c. Check / replace the System controll card. d. Check / replace the DMC. 8. If the soft start signal is missing: a. Check / replace the flat cable from the ACCU to the GHOST. b. Check / replace the ACCU. c. Check / replace the GHSOT Timing (use the Stargate 164 Stationary Power drawings): 1. E stop close + ""START"" relay on will pull in ""PUP"" relay. 2. ""PUP"" relay will pull in ""MK"" which will act as a self holding replay to ""PUP"" check for system LED on ACCU to be on. 3 """"PUP"" relay will send the ""5GNORM"" to the GHOST. 4. The GHOST will send the ""S/SON"" which will pull in the ""S/SSSR"" relay. 5. ""S/SSSR"" relay will send the ""120 S/SON"" to TB 352 / 2 which will pull in the ""S/S (soft start) relay. 6. When the DC voltage is over 500V, the GHOST will send the ""MKDON"" signal. 7. The ""MKDON"" signal will pull in the MKDSSR"" relay. 8. The ""MKDSSR"" relay will send the ""120MKDON"" to TB 352 /1 which will pull in the ""MKD"" (the main 480V AC contactor). 9. ""MKD"" relay will open the ""S/S"" relay by removing the 120 S/SON. (via Aux. Contact side mounted N.C. that is mechanically attached to the right side of the MKD) Experience: 1. The following troubleshooting steps were performed: a. I arrived on site and found 2 out of 3 APM 480 vac input fuses blown. b. I replaced the 30a fuses. I applied power to the gantry and closed the estop. A few seconds into the APM/CPM power charging routine, I heard what I thought might be an arc and the estop opened. The same APM fuses were blown. c. Next, I removed the HSS blue connector and replaced the 30a fuses. The estop closed and the system powered up for about 15 seconds before I heard multiple arcs coming from the APM. No fuses blew this time. d. I replaced the APM. No further problems. " 建议备件12NC及名称： Fuse, 30A/600V 459800066152 1 F1, F2, F3//FUSE,2.50A,500V.406"DIA X1.50" 459800726611//Fuse, 50A, 600VAC, Dual Element TD 459800043431 F351, F352, F353 000090 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 12.06.2021 09:01:23 UTC Sheng Cui (86038219)服务内容：机架无法上电，Softstart 无法吸合，复位交流接触器，更换滑环高压保险和滑环保险。 服务结果：维修后，设备恢复正常工作 Internal Remarks : 12.06.2021 09:01:48 UTC Sheng Cui (86038219) RDI Accuracy（Yes） $Q7: (Confirmation & Description of problem:gantry boot up failed )$Q8: (On site check logeer message GDC low ,need to replace slip ring 50A fuse and HV generator 20A fuse.) $Q9: (Replace slip ring 50A fuse and HV generator 20A fuse. ) $Q10: (Test and inspection data refer to: PA. Device returned to full functionality.) $Q11: Calibrated Tools name (FLUKE287 ) S/N (12470209 ) Next Calibration date (07/12/2021）: For Mainland China CV / XR, CT engineers only. Please fill in: Current Tube installed date Tube installed date: ##09/ 4/ 2020 ^Current Tube usage: scan second_##401202S ----------------------------------------------------------------------------- | 2 |  | 机架无法上电，Softstart 无法吸合，复位交流接触器，更换滑环高压保险和滑环保险。 | [12 empty] | 117679178 |  | 995706 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Operation | [6 empty] | Power Brush assembly WI | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1027 |  |  | Japan | 社会福祉法人恩賜財団済生会 | 343000 | 4435700% |  | 07CS:TM:FSE経由。消耗品劣化 | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 07CS:TM:FSE経由。消耗品劣化 SAP Transfer : Case Transferred to SAP SAP Message : Components Added Succesfully to SWO!- SAP Message : LOGI Task added Successfully!- Safety Questions : Case: 0117680995 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: TAKEYUKI MIYAKAWA Last updated by: TAKEYUKI MIYAKAWA 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: Problem description by engineer : 対象装置:CT 故障個所:07その他 最初に発生したのはいつか?(~以内):Unknow 問題発生の頻度は?:Unknow エンドユーザｰのワークフローのどの時点において問題が発生しているのか?:その他(Other) 最近どのようなイベントが発生したか?:N.A. 顧客は問題について具体的なシステムイベントレポートを作成しているか?:N.A. Engineer's Comment: '============================== コール内容 FSE経由。消耗品劣化 '============================== Troubleshooting Action: What steps are taken during the system diagnosis : 対応後の結果:改善された 故障個所:07その他 診断中にどのようなステップを実施したか?: '================================== 対応 Parts手配 '================================== Error # and/or description of error : NA OneEMS internal remarks : Problem description by engineer : 対象装置:CT 故障個所:07その他 最初に発生したのはいつか?(~以内):Unknow 問題発生の頻度は?:Unknow エンドユーザｰのワークフローのどの時点において問題が発生しているのか?:その他(Other) 最近どのようなイベントが発生したか?:N.A. 顧客は問題について具体的なシステムイベントレポートを作成しているか?:N.A. Engineer's Comment: '============================== コール内容 FSE経由。消耗品劣化 '============================== What steps are taken during the system diagnosis : 対応後の結果:改善された 故障個所:07その他 診断中にどのようなステップを実施したか?: '================================== 対応 Parts手配 '================================== Error # and/or description of error : NA Completed or Recommended Repair Action "Internal" : --Troubleshooting Action-- 対応後の結果:改善された 故障個所:07その他 診断中にどのようなステップを実施したか?: '================================== 対応 Parts手配 '================================== --Repair Action-- Recommended 12Nc: Out of Stock 453567104591:SINGLE STRAP QTY:5 SLOC:BackOrder --Subject-- 07CS:TM:Yasuhiro Kawashige:FSE経由。消耗品劣化 --Problem Description-- 対象装置:CT 故障個所:07その他 最初に発生したのはいつか?(~以内):Unknow 問題発生の頻度は?:Unknow エンドユーザｰのワークフローのどの時点において問題が発生しているのか?:その他(Other) 最近どのようなイベントが発生したか?:N.A. 顧客は問題について具体的なシステムイベントレポートを作成しているか?:N.A. Engineer's Comment: '============================== コール内容 FSE経由。消耗品劣化 '============================== --4/8Elements-- Q1 報告された問題は確認できましたか? 確認済み Q2 他の問題は確認できましたか? 無し Q3 問題を解決するために何をおこないましたか? PARTS手配 Q4 対応後の装置は順調ですか? 良好 Q5 その検査はどの様に終了しましたか? 無し Q6 問題により間違った所見（誤診、誤治療）の有無: 無し Q7 意図しない不要な曝射の有無: 無し Q8 アーチファクトの有無: 無し ----------------------------------------------------------------------------- | 5 |  | Deterioration of consumables 453567104591: SINGLE STRAP QTY: 5 SLOC: BackOrder --Subject-- | [12 empty] | 117680995 |  | 996578 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | NotEngineering | [5 empty] | Engineering | [6 empty] | Strap | [5 empty] | Refresh |  |  | System issue | System | NotEngineering | ShouldbeOthers | NotEngineering-ShouldbeOthers | NotEngineering-ShouldbeOthers | NotEngineering |  | ShouldbeOthers | NotEngineering-ShouldbeOthers | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | y
1028 |  |  | Japan | Maki Hospital | 341017 | 4435700% |  | Applications request. >> Koizumi will respond. | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Application Support Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer : アプリ依頼>>小泉さん対応 Safety Questions : Case: 0117680088 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: TAKEYUKI MIYAKAWA Last updated by: TAKEYUKI MIYAKAWA 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: Problem description by engineer : 対象装置:CT 故障個所:07その他 最初に発生したのはいつか?(~以内):Unknow 問題発生の頻度は?:Unknow エンドユーザｰのワークフローのどの時点において問題が発生しているのか?:その他(Other) 最近どのようなイベントが発生したか?:N.A. 顧客は問題について具体的なシステムイベントレポートを作成しているか?:N.A. Engineer's Comment: '============================== コール内容 患者名を間違えたので、修正方法を教えてください。 '============================== Troubleshooting Action: What steps are taken during the system diagnosis : 対応後の結果:改善された 故障個所:07その他 診断中にどのようなステップを実施したか?: '================================== アプリ担当者回答 change Patient Detailから一緒に行っていただきました。 '================================== Error # and/or description of error : NA Repair Action: Completed or Recommended Repair Action "External" : 【現象確認】 確認済み 【作業内容詳細】 質問内容を回答 【確認内容】 良好 OneEMS internal remarks : Problem description by engineer : 対象装置:CT 故障個所:07その他 最初に発生したのはいつか?(~以内):Unknow 問題発生の頻度は?:Unknow エンドユーザｰのワークフローのどの時点において問題が発生しているのか?:その他(Other) 最近どのようなイベントが発生したか?:N.A. 顧客は問題について具体的なシステムイベントレポートを作成しているか?:N.A. Engineer's Comment: '============================== コール内容 患者名を間違えたので、修正方法を教えてください。 '============================== What steps are taken during the system diagnosis : 対応後の結果:改善された 故障個所:07その他 診断中にどのようなステップを実施したか?: '================================== アプリ担当者回答 change Patient Detailから一緒に行っていただきました。 '================================== Error # and/or description of error : NA Completed or Recommended Repair Action "External" : 【現象確認】 確認済み 【作業内容詳細】 質問内容を回答 【確認内容】 良好 Completed or Recommended Repair Action "Internal" : --Troubleshooting Action-- 対応後の結果:改善された 故障個所:07その他 診断中にどのようなステップを実施したか?: '================================== アプリ担当者回答 change Patient Detailから一緒に行っていただきました。 '================================== --Repair Action-- --Subject-- 07CS:TM:ﾄｶﾞﾜ様:アプリ問い合わせ。患者名を間違えたので、修正方法を教えてください。>>アプリ担当者対応>>change Patient Detailから一緒に行っていただきました。 --Problem Description-- 対象装置:CT 故障個所:07その他 最初に発生したのはいつか?(~以内):Unknow 問題発生の頻度は?:Unknow エンドユーザｰのワークフローのどの時点において問題が発生しているのか?:その他(Other) 最近どのようなイベントが発生したか?:N.A. 顧客は問題について具体的なシステムイベントレポートを作成しているか?:N.A. Engineer's Comment: '============================== コール内容 患者名を間違えたので、修正方法を教えてください。 '============================== --4/8Elements-- Q1 報告された問題は確認できましたか? 確認済み Q2 他の問題は確認できましたか? 無し Q3 問題を解決するために何をおこないましたか? 質問内容を回答 Q4 対応後の装置は順調ですか? 良好 Q5 その検査はどの様に終了しましたか? 無し Q6 問題により間違った所見（誤診、誤治療）の有無: 無し Q7 意図しない不要な曝射の有無: 無し Q8 アーチファクトの有無: 無し SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | Closed - Non-Complaint |  | NAC | [12 empty] | 117680088 |  | 995964 |  |  | Closed - Non-Complaint | [7 empty] | NAC | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1029 |  |  | Japan | Independent administrative institution National Hospital Organization | 341013 | 4435600% |  | 03DS:KKI: An error occurred during a contrast-enhanced tracker scan. >> S_GEN_MGR_MA_OUT_OF_RANGE/S_GEN_FAUL [Truncated] | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 03CS:KKI: An error occurred during a contrast-enhanced tracker scan. >> S_GEN_MGR_MA_OUT_OF_RANGE/S_GEN_FAULT_ANODE_TUBE_ARC > He was able to continue the examination. Safety Questions:Case: 0117673753 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: KAZUMASA KIMURA Last updated by: KAZUMASA KIMURA 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Support External Remarks : 10.06.2021 06:21:07 UTC USHIO TAKADA (81100646) Request: Arcing occurred during a scan. Problem confirmation: I confirmed it. I confirmed it in the log. Work details: Arcing occurred, so today I adjusted the filament output. When I finished, I performed a test scan and confirmed that arcing did not occur and there was no problem with the scanner operation. Other problems/Measures taken: None. Equipment status The equipment can be used. Work status Work finished. 09-Jun-2021 81100646 FSE Handover Tool1.3.4.32612.DB13 20210115 Internal Remarks:11.06.2021 07:03:30 UTC USHIO TAKADA (81100646) Replacement parts: No replacement parts. [Report on mistakes in parts order] No mistakes. Situation monitoring information: Not monitoring the situation TS Request: No parts to return. Additional CT information: Q1. How did the examination finish? Q1. No (not applicable) Q2. Were there any mistaken findings (misdiagnosis, mistaken treatment) due to the problem? None. Q3. Were there any artifacts? None. Q4. Was there any unintended unnecessary exposure? None. Repair sticker: I affixed a Repaired sticker. [Other information to share] 1. Was treatment/quarantine carried out for a patient who has or is suspected of having new coronavirus? No 2. Is it necessary to go into treatment rooms/beds where there are infected patients?(Including workplace and routes of movement) No Did you agree on how to complete the report with the person in charge of devices? Yes. I completed normal reporting that day. 11-Jun-2021 81100646 FSE Handover Tool1.3.4.32612.DB13 20210115 ----------------------------------------------------------------------------- | 4 |  | An arc occurred during the scan. [Phenomenon confirmation] I was able to confirm. I confirmed it in the log. [Details of work] Since an arc was generated, we have adjusted the filament output today. After the work, a test scan is carried out to confirm that no arc is generated and that there is no problem in the operation of the equipment. | [12 empty] | 117673753 |  | 991971 | TRUE | Investigation | Investigation |  | CTUDT00630085 | [5 empty] | Engineering | [6 empty] | Tube Arc | [5 empty] | Refresh |  |  | System issue | System | IC-HW | Tub Arc | IC-HW-Tub Arc | IC-HW-Tub Arc | CTUDT00630085 | Tub Arc | IC-HW | IC-HW-Tub Arc | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | y
1030 |  |  | China | Hebei Province Geriatric Hospital | 345015 | 4435500% |  | PRS connection/02N/Ingenuity Core128/Elite China/ | BU Q&R Feedback: No Reason: Customer Relations Problem Reported by Customer : PRS connection Safety Questions : Case: 0117665791 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingli Last updated by: Lu Mingli 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action : FSE 孙建平 018031820728 call in and disp him SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 08.06.2021 06:27:32 UTC Jianping Sun (86039023)服务内容： [连接PRS。] 服务结果： [设备正常工作。] Internal Remarks : 10.06.2021 14:31:35 UTC Jianping Sun (86039023) RDI Accuracy：yes. $Q7:(Confirmed the reported problem. PRS install） $Q8:(PRS install) $Q9:(PRS install ) $Q10:(Test and Inspection data refer to PA.Device returened to full futionnallity.) $Q11:(Calibrated tools name（FLUKE287）S/N（20060013）Next calibration date:(05/10/2022)) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date:N/A ^Current Tube usage:N/A ----------------------------------------------------------------------------- | Zhai, Lisa |  | 2021-06-09 14:23 | [12 empty] | 117665791 |  | 987202 |  |  | Closed - Non-Complaint | [14 empty] | 2021-06-12 00:46 | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1031 |  |  | Japan | 独立行政法人国立病院機構 | 341015 | 4435500% |  | 07CS:RS:心臓のアプリに関して担当者と連絡希望>>APPサポート時間外の為、明日ご連絡いただけるとのこと | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Customer Information Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer : 07CS:RS:心臓のアプリに関して担当者と連絡希望>>APPサポート時間外の為、明日ご連絡いただけるとのこと Safety Questions : Case: 0117668167 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: RUI SATAKE Last updated by: RUI SATAKE 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | Closed - Non-Complaint |  | NAC | [12 empty] | 117668167 |  | 987965 |  |  | Closed - Non-Complaint | [7 empty] | NAC | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1032 |  |  | Japan | Social Medical Corporation BOKOI, Nikko Memorial Hospital | 341011 | 4435400% |  | JY: The bolus tracking was not triggered even when the threshold was exceeded. | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : JY: The bolus tracking was not triggered even when the threshold was exceeded. Safety Questions:Case: 0117658722 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: YAN Joanna Last updated by: YAN Joanna 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A OneEMS internal remarks: Sugahara DB SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Support ----------------------------------------------------------------------------- | 4 |  | The case was reviewed by Console team. Here is the investigation details. Observations: The tracker scan completed successfully. Not observed a big delay between end of tracker to the clinical scan start. It is possible that there could be a extra tracker shot can happen The skip command to gantry also went in the same time. All the series 15, 16 and 17 are successfully completed without any error. Conclusion: Not seen any delay or issue in tracker and subsequent clinical scans System is working as expected. | [12 empty] | 117658722 |  | 982397 | TRUE | Evaluation | Closed - No Further Investigation |  | NotDefect | [5 empty] | Engineering | [6 empty] | Bolus Tracking | [5 empty] | Refresh |  |  | System issue | System | Console-SW | NotDefect | Console-SW-NotDefect | Console-SW-NotDefect | NotDefect |  | Console-SW | Console-SW-NotDefect | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | y
1033 |  |  | Sweden | Falu Lasarett | 40049 | 4435500% |  | Application does not work, background turns yellow, but works when restart is done | Work Order #: WO-03618921 Knudsen Martin The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Intermittent user get's yellow screen and application dissapears. Follow Up Required : Intermittent yellow screen and application disappear. No errors found in logfiles. SW reload is first step. Please book FSE 7 hours for this. No parts. -MK Diagnostic performed by Engineer : Intermittent user get's yellow screen and application dissapears. Reboot solves issue. Checked logfiles, and was not able to see any errors related. Advised to do a SW reload of the Host computer. ----------------------------------------------------------------------------- | Evaluation |  | Yellow screen issue, SW version is not clear, suggest field to implement new SP1 SR FCO72800758 | [12 empty] | 117667610 |  | 987581 | TRUE | Evaluation | Evaluation |  | CTUDT00632520 | [5 empty] | Engineering | [6 empty] | New Yellow Screen | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | y
1034 |  |  | China | 内蒙古阿拉善右旗人民医院 | 345031 | 4435700% |  | 02N/Ingenuity Core128/Elite China/The gantry cann't power up | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity Core128/Elite China T2 Activities : Problem Description: 客户问题概述： 机架不能上电 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： S_MASTER_SEQUENCER_GDC_LOW_AFTER_TIMEOUT 远程诊断与测试结果： 联系工程师苏瑞杰，怀疑电源刷短路，保险丝问题 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议工程师到场检查清理滑环，检查电源是，保险丝 建议备件12NC及名称： 保险丝：459800043431，459800448931 Safety Questions : Case: 0117679129 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Weng Jinming Last updated by: Weng Jinming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : PPD 459800448931*2 459800043431*3备件已系统下单安排24H到场，FSE：苏瑞杰-86-13304719335 External Remarks : 11.06.2021 07:02:43 UTC Sheng Cui (86038219)"SUAT. Ready for Clinical Use local time" : "2021-06-11 16:00" External Remarks : 12.06.2021 01:25:13 UTC Sheng Cui (86038219)服务内容：机架无法上电，Softstart 无法吸合，复位交流接触器，更换滑环高压保险和滑环保险。 服务结果：维修后，设备恢复正常工作 Internal Remarks : 12.06.2021 01:25:16 UTC Sheng Cui (86038219) RDI Accuracy（Yes） $Q7: (Confirmation & Description of problem:gantry boot up failed )$Q8: (On site check logeer message GDC low ,need to replace slip ring 50A fuse and HV generator 20A fuse.) $Q9: (Replace slip ring 50A fuse and HV generator 20A fuse. ) $Q10: (Test and inspection data refer to: PA. Device returned to full functionality.) $Q11: Calibrated Tools name (FLUKE287 ) S/N (12470209 ) Next Calibration date (07/12/2021）: For Mainland China CV / XR, CT engineers only. Please fill in: Current Tube installed date Tube installed date: ##09/ 4/ 2020 ^Current Tube usage: scan second_##401202S ----------------------------------------------------------------------------- | 2 |  | 机架无法上电，Softstart 无法吸合，复位交流接触器，更换滑环高压保险和滑环保险。 459800066152 FUSE, 30A/600V AJT30 3 24H 459800726611 FUSE, 2.5A/500V FNQ-2-1/2 2 24H 459800043431 Fuse, 50A, 600VAC, Dual Element TD" 3 24H 459800448931 20 AMP FUSE 1 24H已出货 | [12 empty] | 117679129 |  | 995664 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Operation | [6 empty] | Power Brush assembly WI | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1035 |  |  | Turkey | ADIYAMANOZELSAGINSTURTASSANVETICA.S | 341006 | 4435800% |  | artifact occurred in cardiac extraction | Work Order #: WO-03641560 Turgay Baylan The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : Case escalation was performed in order to examine the images in the cardiac extraction. ----------------------------------------------------------------------------- | Investigation |  | There is any image problem for 130 HR, but customer complains images for 82 HR. | [12 empty] | 117688189 |  | 998252 | TRUE | Investigation | Investigation - Pending Translation |  | TBD_Bugreport | [5 empty] | Engineering | [6 empty] | Cardiac | [5 empty] | Refresh |  |  | System issue | System | TBD-Bugreport | TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | TBD_Bugreport |  | TBD-Bugreport | TBD-Bugreport-TBD_Bugreport | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | y
1036 |  |  | Turkey | Republic of Turkey Ministry of Health, Bagcilar | 40010 | 4435600% |  | The belt is noisy. | Work Order #: WO-03627268 09-Jun-2021 Ugur Altunoz The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : Gantry main drive belt has a crack on it and it should be replaced. Follow Up Required : Gantry main drive belt replacement. ----------------------------------------------------------------------------- | Evaluation |  | There is a crack on the gantry main drive belt and there is a sound. Belt replacement is required. Gantry main drive belt will be replaced. | [12 empty] | 117675108 |  | 989138 | TRUE | Evaluation | Evaluation |  | ECR-077012 | [5 empty] | Engineering | [6 empty] | Belt | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Belt | Gantry-HW-Belt | Gantry-HW-Belt | ECR-077012 | Belt | Gantry-HW | Gantry-HW-Belt | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | y
1037 |  |  | Turkey | Republic of Turkey Ministry of Health, Bagcilar | 40010 | 4435500% |  | It interrupts the scanning. | "Work Order #: WO-03620575 Salih Mendi The system does not meet specification and is out of use Problem Description by Engineer : Customer Function/Role: technician Device Use: out of use Expected and Unexpected Results: S_CTRL_LINK_GHOST_CRC_FRAMING_ERROR User Impact: No user impact Patient Impact: no patient impact Current Software Version:NA Diagnostic performed by Engineer : Upon the user's complaint that ""The device interrupts the recording"", the system was examined by connecting to the system via remote access. Upon inspection, it was observed that the system gave the error S_CTRL_LINK_GHOST_CRC_FRAMING_ERROR. A site visit is required for detailed examination of Slipring, Ghost, DMC and laser systems. Follow Up Required: A site visit is required for detailed examination of Slipring, Ghost, DMC and laser systems. ---------------------------------------------------------------- --------------------------- Work Order #: WO-03621501 09-Jun-2021 Ugur Altunoz The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: technician Device Use: out of use Expected and Unexpected Results: S_CTRL_LINK_GHOST_CRC_FRAMING_ERROR User Impact: No user impact Patient Impact: no patient impact Current Software Version:NA Resolution : Due to CRC frame errors in the device, slipring was cleaned and the images were checked. Trial shots were made on the device and it was delivered in an active condition. | Closed - Done |  | Due to CRC frame errors in the device, slipring was cleaned and the images were checked. Trial shots were made on the device and it was delivered in an active condition. | [12 empty] | 117669046 |  | 984579 | TRUE | Closed - Done | Closed - Done |  | NotEngineering-SlipringClean | [5 empty] | Engineering | [6 empty] | Slipring | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | SlipRingClean | Gantry-HW-SlipRingClean | Gantry-HW-SlipRingClean | NotEngineering-SlipringClean |  | Gantry-HW | Gantry-HW-SlipRingClean | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | y
1038 |  |  | United States | Philips Medical Capital | 40055 | 4436000% |  | 0609CDN84 knocking noise when gantry tilted | Work Order #: WO-03647989 Mcginnis, Mike The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Problem description by engineer : knocking noise when gantry tilted When was the first occurrence? Within... : 13-Jun-2021 How often is issue occurring? : First occurrence Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: knocking noise when gantry tilted Patient Impact/Outcome: none User Impact: knocking noise when gantry tilted Current Software Version: Not provided Follow Up Required : Error # and/or description of error : knocking noise when gantry tilted Completed or Recommended Repair Action "Internal" : Check the covers for loose or missing bolts. If none found open the cover and isolate the noise and repair as needed. Internal Comments : Error # and/or description of error : knocking noise when gantry tilted Completed or Recommended Repair Action "Internal" : Check the covers for loose or missing bolts. If none found open the cover and isolate the noise and repair as needed. Diagnostic performed by Engineer : Melissa said when she tilts the gantry posterior she hears a knocking noise from the gantry. She said she feels comfortable using the system but would like it looked at. ----------------------------------------------------------------------------- | Closed - No Further Investigation |  | Tilted the gantry to reproduce the noise experienced by the customer. I heard the covers rubbing against each other and cusing a growling noise. Adjusted the cover to resolve the noises. | [12 empty] | 117693213 |  | 1000880 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-GantryNoise | [5 empty] | Engineering | [6 empty] | Gantry cover misalignment | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Noise | Gantry-HW-Noise | Gantry-HW-Noise | TBD-GantryNoise |  | Gantry-HW | Gantry-HW-Noise | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | y
1039 |  |  | United States | Department of Veterans Affairs | 40018 | 4435700% |  | 5-Scheduled Activity | Work Order #: WO-03639047 Powell, Jamie The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : Perform safety clip check/inspection/repair for FSN/FCO 72800752.. Follow Up Required : Perform safety clip check/inspection/repair for FSN/FCO 72800752.. ----------------------------------------------------------------------------- | 5 |  | FCO72800752 | [12 empty] | 117685359 |  | 995412 | TRUE | Evaluation | Closed - non-complaint | [7 empty] | Others | [6 empty] | FCO72800752 | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1040 |  |  | China | Beijing City Hospital of Traditional Chinese and Western Medicine | 40032 | 4435400% |  | 02N/Ingenuity CT 728326//ask engineer onsite information to customer about installation | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 02N/Ingenuity CT 728326 Next Step Action: 013811972406 Yu called T2 Activities:Problem Description: Summary of customer's problems: Other types of error. Detailed description of Customer's problems or malfunctions: Yu called in, need labor dispatch to the field to give the director an introduction of the machine's leftovers after installation, and the simple switch on/off. It's a technology problem with the equipment itself: No. Troubleshooting Action: key fault code or symptom: Sales called in, need labor dispatch to the field to give the director an introduction of the machine's leftovers after installation, and the simple switch on/off. Remote troubleshoot and test result: The customer needs an engineer to go to the field. Contact Director Zhang PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: introduction on site. Safety Questions:Case: 0117660029 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Jianqiang Last updated by: Wang Jianqiang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support External Remarks: 09.06.2021 07:02:13 UTC Guochao Fan (86039456) service details: work with the hospital to organize and clean the scan room, completed service result: device returned to normal working condition after repairing Internal Remarks: 09.06.2021 07:02:18 UTC Guochao Fan (86039456) "RDI Accuracy:Yes" $Q7: Confirm reported problems:Equipment safety hazard detection $Q8: Trouble shooting steps:for Equipment safety hazard detection. Equipment normal $Q9: NA $Q10: Test and Inspection data refer to PA. Device status:Device returned to full functionally $Q11: N/A ----------------------------------------------------------------------------- | Zhai, Lisa |  | 2021-06-08 17:27 | [12 empty] | 117660029 |  | 983225 |  |  | Closed - Non-Complaint | [14 empty] | 2021-06-11 23:46 | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1041 |  |  | China | 陕西省咸阳中心医院 (咸阳市第二人民医院） | 40050 | 4435400% |  | 02N/Ingenuity CT 728326/asking parameters | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: Instructed Customer Problem Reported by Customer : 02N/Ingenuity CT 728326 FSE Notes : This PO is a Reference Number only Safety Questions : Case: 0117658800 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: customer asking CT parameters Troubleshooting Action: customer asking CT parameters,max KV,max mA, answer it. Repair Action: customer asking CT parameters,max KV,max mA, answer it. Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | Zhai, Lisa |  | 2021-06-08 14:13 | [12 empty] | 117658800 |  | 982471 |  |  | Closed - Non-Complaint | [14 empty] | 2021-06-10 14:45 | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1042 |  |  | Japan | Tomakomai City Hospital | 341014 | 4435400% |  | He wanted to contact Sugahara about the arcing that occurred the other day. >> Sugahara called him back. | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer: He wanted to contact Sugahara about the arcing that occurred the other day. >> Sugahara called him back. Safety Questions:Case: 0117658774 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: TAKEYUKI MIYAKAWA Last updated by: TAKEYUKI MIYAKAWA 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | Zhu, Chun |  | 2021-06-08 14:08 | [12 empty] | 117658774 |  | 982452 |  |  | Closed - Non-Complaint | [14 empty] | 2021-06-14 09:45 | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1043 |  |  | United States | Willis Knighton Hospital | 40004 | 4435600% |  | Infant Phantom has Bad Seal | Work Order #: WO-03631273 11-Jun-2021 Williams Spence The system fully meets specification and returned to use Parts Consumed : 1.00000,453567047262,INFANT PHANTOM Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : Phantom leaked. Replaced defective infant phantom ----------------------------------------------------------------------------- | 5 |  | 1.00000,453567047262,INFANT PHANTOM Phantom leaked. Replaced defective infant phantom | [12 empty] | 117678304 |  | 995113 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | NotEngineering-PartReplace | [5 empty] | Engineering | [6 empty] | Infant phantom | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | Phantom | Couch-HW-Phantom | Couch-HW-Phantom | NotEngineering-PartReplace |  | Couch-HW | Couch-HW-Phantom | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | y
1044 |  |  | United States | Odessa Regional Medical Center | 40029 | 4435500% |  | The contrast issue, performed air calibration and rebooted the system | Work Order #: WO-03623673 Atz, Michael The system fully meets specification and returned to use Internal Comments : Error # and/or description of error : none Completed or Recommended Repair Action "Internal" : Advised to run an air cal now that the contrast is cleaned off of the gantry and to reboot the system to the wall contrast on the gantry she will call back if further service is needed. Diagnostic performed by Engineer : Advised to run an air cal now that the contrast is cleaned off of the gantry and to reboot the system to the wall contrast on the gantry she will call back if further service is needed. Problem Description by Engineer : Problem description by engineer : contrast on the gantry now a ring artifact air cal fails When was the first occurrence? Within... : 08-Jun-2021 How often is issue occurring? : Unknown Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: No further information is available as per the customer Patient Impact/Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer Current Software Version: Not provided Resolution : system is operational ----------------------------------------------------------------------------- | Closed - Done |  | Advised to run an air cal now that the contrast is cleaned off of the gantry and to reboot the system to the wall contrast on the gantry she will call back if further service is needed. | [12 empty] | 117671425 |  | 985996 | TRUE | Closed - Done | Closed - Done | [7 empty] | Customer | [6 empty] | Contrast | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1045 |  |  | United States | Philips Medical Capital | 40005 | 4435400% |  | New EMR Procedure Codes | Work Order #: WO-03612662 Mujica, David The system fully meets specification and returned to use Problem description by Clinical Support Specialist : Customer Function/Role: Technologist/Manager Device Use: In use Expected and Unexpected Results: map accession numbers to new orders on new system User Impact: none Patient Impact: none Current Software Version: 4.1 Diagnostic performed by Engineer : Kelly indicated that she needed help associating newly populated orders to existing protocols/exam cards on their CT system. Kelly was not in front of the CT system and could not be instructed on where to make these changes. A follow up call is required when Kelly is in front of the CT system, so she can be shown where this change to the existing exam cards/protocols can be made. The follow up is scheduled for Thursday 6/10. No further clinical instruction is needed at this time. ----------------------------------------------------------------------------- | 5 |  | NAC | [12 empty] | 117661514 |  | 984359 |  |  | Closed - Non-Complaint | [7 empty] | NAC | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1046 |  |  | Chile | Scotiabank Chile | 40008 | 4436100% |  | CL1675-THE COUCH REFERENCE LINE IS NOT SHOWN IN THE SCANS | Work Order #: WO-03653507 Valenzuela Velasquez, Eduardo Ariel The system fully meets specification and returned to use Problem Description by Engineer : Problem description by engineer : No aparece linea de referencia en imágenes de surview. Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected results: Linea de referencia visible en surview. Patient Impact/Outcome: Dificultad para centrar exámenes. User Impact: Dificultad para centrar exámenes. Current Software Version: 4.1.10 Resolution : The options to enable the reference are reviewed with the support of applications. It is activated with the right mouse button on the surview image. Tests are performed and the requested reference is achieved. Internal Comments: Error # and / or description of error: There is no referral line on the surview.--------------- |  |  | The options to enable the reference are reviewed with the support of applications. It is activated with the right mouse button on the surview image. | [12 empty] | 117698098 |  | 1004078 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Application | [6 empty] | Reference line | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1047 |  |  | China | 陕西省咸阳中心医院 (咸阳市第二人民医院） | 40050 | 4436400% |  | 02n/Ingenuity CT 728326/Add the PACS | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02n/Ingenuity CT 728326 Safety Questions : Case: 0117716668 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Weng Jinming Last updated by: Weng Jinming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 医院需要添加PACS 客户问题详细描述或故障现象： 用户名，密码一直不断 设备本身技术问题： 是。 Troubleshooting Action: 远程诊断与测试结果： 联系不上工程师，网络不通，不能远程授权 PRS TR状态： 不通，请工程师到场检查。 Repair Action: 建议现场维修步骤： 需要工程师到场设置网络 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support ----------------------------------------------------------------------------- | 5 |  | confirmed report problem,customer want at ingenuity ct add PACS, customer cannot transfer image to pacs $Q8:remort contact the pacs fse ,on the ct login IT account ,and can add pacs dicom parameter , test pacs is normal ,device normal work | [12 empty] | 117716668 |  | 1020727 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | PACS config | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1048 |  |  | China | 河南省濮阳市人民医院 | 40031 | 4436400% |  | 02N/Ingenuity CT 728326/ Login interface locked | BU Q&R Feedback: No SAP Customer Symptom Code: CENV Customer Symptom Code: Customer Environmental Concern Reason: Corrective Maintenance Engineer Failure Code: Environmental Cause Repair Activity Code: Instructed Customer Safety Questions : Case: 0117715562 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Guomei Li Last updated by: Guomei Li 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Problem Reported by Customer : Login interface locked T2 Activities : Problem Description: CT账户密码输错三次锁定了 Troubleshooting Action: 指导客户使用IT账户重置密码或等待半小时后重新输入正确密码 Repair Action: 指导客户重置密码或等待一会儿后输入正确密码 Test/Inspection: N/A Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- |  |  | 指导客户使用IT账户重置密码或等待半小时后重新输入正确密码 | [12 empty] | 117715562 |  | 1019849 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Password | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1049 |  |  | China | 河南省濮阳市人民医院 | 40031 | 4436400% |  | Wechat:02N/Ingenuity CT 728326/The code is locked | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Software Defect or Virus Repair Activity Code: Instructed Customer Problem Reported by Customer : The code is locked Safety Questions : Case: 0117715480 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Bai Jingwen Last updated by: Bai Jingwen 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: CT password locked Troubleshooting Action: waiting for password unlock. Repair Action: waiting for password unlock. system is ok. Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- |  |  | waiting for password unlock. system is ok. | [12 empty] | 117715480 |  | 1019781 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Password | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1050 |  |  | China | 浙江省乐清第一人民医院 | 40054 | 4436200% |  | Wechat:02N/Ingenuity CT 728326/gantry has no power | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : gantry has no power Safety Questions : Case: 0117699800 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Tao Fengming Last updated by: Tao Fengming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 无法加电 客户问题详细描述或故障现象： 无法加电 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： lifesync p1=8, 远程诊断与测试结果： 无法加电，不出钥匙框，CCC无响应，重启无效 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 检查CCC本身及供电，检查接线，检查CCC上power fail是否亮，若亮参考newsletter排查 建议备件12NC及名称： " 453566503742 ASSY, 4 OUTPUT DC SUPPLY； 459800477572 Couch control board with CONFIG-B CPM" SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : PPD 459800477572 1&453566503742 1已安排12H到医院 黄逸-86-15958725282 External Remarks : 16.06.2021 13:57:59 UTC SUP JCOUSER (86036249)服务内容：机架不能上电，lifesync8，测量四输出正常，CCC板在00和22切换，更换AMC。 服务结果：维修后，设备恢复正常工作。 其它重要事项：无 Internal Remarks : 16.06.2021 13:58:02 UTC Yi Huang (86036249) RDI Accuracy:Yes $Q7: (Confirmed reported problem&estop can't close) $Q8: (check errorlog found lifesync 22 and ccc config end bkst parameter is 2) $Q9: (replace the AMC) $Q10: (Test and inspection data refer to: PA.Device returned to full functionality.) $Q11: Calibrated tools name(N/A) S/N (N/A) Next calibration date (N/A) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date:N/A ^Current Tube usage: second:N/A External Remarks : 17.06.2021 01:44:56 UTC Yi Huang (86036249)"SUAT. Ready for Clinical Use local time" : "2021-06-15 21:00" ----------------------------------------------------------------------------- | 2 |  | System works well after replace the AMC, CTUDT00632187 | [12 empty] | 117699800 |  | 1008995 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632187 | [5 empty] | Engineering | [6 empty] | Horizontal Movement | [5 empty] | Refresh |  |  | System issue | System | Couch-SW | Estop | Couch-SW-Estop | Couch-SW-Estop | CTUDT00632187 |  | Couch-SW | Couch-SW-Estop | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | y
1051 |  |  | China | 莱阳市人民医院 | 345005 | 4436500% |  | CT systsm stuck, cannot scan. |  | WANG, Peter |  | 2021-06-18 14:58 | [14 empty] | 1019538 | TRUE | Investigation | Closed - Done | [7 empty] | Engineering | [6 empty] | systsm stuck | [5 empty] | Refresh |  |  | System issue | System | TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | 0 |  |  |  | No | 0 | 0 | 0 | 0 | 0 | No | y | y | y | y | y | y | y | y | y | y | #REF!
1052 |  |  | China | 苏州广慈肿瘤医院有限公司 | 345019 | 4436300% |  | Wechat:02N/Ingenuity Core128/Elite China/update PACS | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : update PACS Safety Questions : Case: 0117709352 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Tao Fengming Last updated by: Tao Fengming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 医院PACS系统更新升级，需要设备工程师协助 设备本身技术问题： 是。 Troubleshooting Action: 远程诊断与测试结果： 医院PACS系统更新升级，需要设备工程师协助 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 医院PACS系统更新升级，需要设备工程师协助 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 17.06.2021 07:23:16 UTC Haotian Xu (86090280) 服务内容：主机传输pacs异常，配合pacs工程师修改参数，测试。 服务结果：设备正常运行。 Internal Remarks : 17.06.2021 10:25:45 UTC Haotian Xu (86090280) RDI Accuracy: Yes $Q7: (Confirm reported problem,host delivered image to PACS slowly and sometimes failed.) $Q8: ( check the network and pacs parameter） $Q9: (change the pacs parameter accourding to service manual .) $Q10: (Test and inspection data refer to PA. Device returned to full functionality.) $Q11: Calibrated tools refer to PA. ----------------------------------------------------------------------------- | 5 |  | 服务内容：主机传输pacs异常，配合pacs工程师修改参数，测试。 服务结果：设备正常运行。 | [12 empty] | 117709352 |  | 1015512 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | PACS config | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1053 |  |  | China | 新疆伊犁州友谊医院 | 40048 | 4436300% |  | 02N/Ingenuity CT 728326/ISP workflow error | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer : 02N/Ingenuity CT 728326 Safety Questions : Case: 0117707422 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Cao Last updated by: Liu Cao 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: Cannot login ISP LX with workflow error Troubleshooting Action: System is newly installed don't has network connection Repair Action: Connect network and get IP,then issue solved Test/Inspection: N/A Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | Lu, Xueying |  | System is newly installed don't has network connection Repair Action: Connect network and get IP,then issue solved | [12 empty] | 117707422 |  | 1014076 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Network setting | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1054 |  |  | China | 河南许昌市中心医院(传染病) | 40047 | 4436300% |  | 02N/Ingenuity CT 728326/can not close estop | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: Instructed Customer Problem Reported by Customer : 02N/Ingenuity CT 728326 Safety Questions : Case: 0117707257 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Zhang Caiqin Last updated by: Zhang Caiqin 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: can not close estop Troubleshooting Action: guide customer check emergency switch, it is pressed Repair Action: release this switch, then close estop loop normally Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: NA SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | Gao, Mia |  | Troubleshooting Action: guide customer check emergency switch, it is pressed Repair Action: release this switch, then close estop loop normally | [12 empty] | 117707257 |  | 1013977 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | E-Stop button | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1055 |  |  | China | Hebei Province Tangshan City Workers Hospital | 40044 | 4436200% |  | WeChat:02N/Ingenuity CT 728326/Couch fault | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : Couch fault Safety Questions : Case: 0117699667 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Jianqiang Last updated by: Wang Jianqiang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 床偶发报错位置错误 客户问题详细描述或故障现象： 床偶发报错位置错误 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 床偶发报错位置错误 远程诊断与测试结果： 重启可恢复，但是偶发影响使用 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 检查拉线定位器，检查CCC 的控制，如有必要进行更换 建议备件12NC及名称： BRILLIANCE COUCH CNTL PCB ASSY 459800477572// Vertical Transducer 459800623331 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : PPD备件Vertical Transducer 459800623331*1&Couch control board with CONFIG-B CPM 459800477572*1已安排48H到场地，请派工，FSE：徐国梁-86-13869524913 External Remarks : 17.06.2021 11:32:18 UTC Guoliang Xu (86038283) 服务内容：垂直位置故障，更换电位器。 服务结果：维修后观察使用中。 Internal Remarks : 17.06.2021 11:32:21 UTC Guoliang Xu (86038283) $Q1: (Y) $Q2: (N) $Q3: () $Q4: () $Q5: () $Q6: （Couch fault ）RDI Accuracy ：Yes $Q7:Confirmed reported problem，couch position error . $Q8: Couc vertical position fault，.$Q9: Replaced pot transducer. $Q10: Test and inspection data refer to: PA. Device returned to full functionality. $Q11: Calibrated Tools name ( Multimeter Fluke 287C) S/N (15930163) NextCalibration date (04/28/2022) For Mainland China CV / XR, CTengineers only. Please fill in: ^ Current Tube installed date03/13/2021 ^Current Tube usage: scan second_20041 ----------------------------------------------------------------------------- | Lu, Xueying |  | Confirmed reported problem，couch position error . $Q8: Couc vertical position fault，.$Q9: Replaced pot transducer (Vertical Transducer 459800623331*1). $Q10: Test and inspection data refer to: PA. Device returned to full functionality. | [12 empty] | 117699667 |  | 1008884 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | NotEngineering-partreplace | [5 empty] | Engineering | [6 empty] | Vertical Transducer | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | VerticalTransducer | Couch-HW-VerticalTransducer | Couch-HW-VerticalTransducer | NotEngineering-partreplace |  | Couch-HW | Couch-HW-VerticalTransducer | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | y
1056 |  |  | China | 宁安市第二人民医院 | 345039 | 4436300% |  | 02N/Ingenuity Core128/Elite China/forgot the admin password | BU Q&R Feedback: No SAP Customer Symptom Code: CENV Customer Symptom Code: Customer Environmental Concern Reason: Corrective Maintenance Engineer Failure Code: Environmental Cause Repair Activity Code: Database Recreation / Recovery Next Step Action : ASP yiyuanji 015604188333 call in Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ Safety Questions : Case: 0117708798 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Gaoge Last updated by: Wang Gaoge 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: olevel_service账户密码忘记了 Troubleshooting Action: 通过远程重置olevel_service账户密码 重置后的密码是 ！1 Repair Action: 重置账户密码 Test/Inspection: N/A Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | Gao, Mia |  | 2021-06-17 16:33 | [12 empty] | 117708798 |  | 1015072 |  |  | Closed - Non-Complaint | [14 empty] | 2021-06-20 11:30 | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1057 |  |  | Japan | 紀南病院 | 340000 | 4436500% |  | SY:ガントリから異音がする | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : SY:ガントリから異音がする Safety Questions : Case: 0117721677 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Smile YE Last updated by: Smile YE 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A OneEMS internal remarks : 大山DB SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support ----------------------------------------------------------------------------- | 5 |  | There are replacement parts 12Nc: 453566492401 Name: POLY-VEE BELT QTY: 1 | [12 empty] | 117721677 |  | 1024609 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | ECR-077012 | [5 empty] | Engineering | [6 empty] | Belt | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Belt | Gantry-HW-Belt | Gantry-HW-Belt | ECR-077012 | Belt | Gantry-HW | Gantry-HW-Belt | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | y
1058 |  |  | Japan | 社会福祉法人恩賜財団済生会 | 343000 | 4436400% |  | 03CS:RS:Short TC時に「Maximum number of arc...」とポップアップし中断>>再実施にて完了>> S_GEN_FAULT_ANODE_TUBE_ARC | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 03CS:RS:Short TC時に「Maximum number of arc...」とポップアップし中断>>再実施にて完了>> ----------------------------------------------------------------------------- | 4 |  | Short TC pops up "Maximum number of arc ..." and interrupts Plan various calibrations at the next inspection. | [12 empty] | 117720230 |  | 1023889 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00630085 | [5 empty] | Engineering | [6 empty] | Tube Arc | [5 empty] | Refresh |  |  | System issue | System | IC-HW | Tub Arc | IC-HW-Tub Arc | IC-HW-Tub Arc | CTUDT00630085 | Tub Arc | IC-HW | IC-HW-Tub Arc | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | y
1059 |  |  | Japan | Independent administrative institution National Hospital Organization | 341013 | 4436200% |  | 01CS:KKI:Cannot login to the CT-Host. >> Restarting the CT-Host resolved the problem. | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 01CS:KKI:Cannot login to the CT-Host. >> Restarting the CT-Host resolved the problem. Safety Questions:Case: 0117706588 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: KAZUMASA KIMURA Last updated by: KAZUMASA KIMURA 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 4 |  | Restored by CT host restart | [12 empty] | 117706588 |  | 1013625 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Login | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1060 |  |  | Japan | Tomakomai City Hospital | 341014 | 4436200% |  | 03CM:TM: The equipment stopped due to an error during a scan>> SoftArcReg/S_GEN_FAULT_ANODE_TUBE_ARC/S_GEN_F [Truncated] | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: The chest scan stopped due to an error >> Safety Questions:Case: 0117699766 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: TAKEYUKI MIYAKAWA Last updated by: TAKEYUKI MIYAKAWA 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Problem description by engineer : Equipment affected: CT Failure location: 03High Voltage Energy When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user's work flow did the problem occur?: (During Acquisition/Scan) What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem?: NA Engineer's Comment: '============================== $version = "v4.1.10.25210" 9:58 Call description: An "X-ray system cannot comply" error message was displayed during a chest scan and the equipment stopped. The necessary range of data was captured, so he didn't rescan. He asked me to check the log. '============================== Troubleshooting Action: What steps are taken during the system diagnosis: Result of response: Problem solved. Failure location: 03High Voltage Energy What steps did you perform to diagnose the problem? '================================== Results of checking the log: The following error occurred. 2021/06/15_09:48:27.317__03 <*HV*> szInfo: SoftArcReg 0x211* P1:0/P2: 2021/06/15_09:48:27.368 <*HV*> Type: NORMAL ErrNum: S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT szInfo: P1:0/P2: 2021/06/15_09:48:27.376 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_ANALOG_FAULT szInfo:2:5 Analog Fault P1:2/P2:32 2021/06/15_09:48:27.378 [CANCEL] CStateShot:Series 285 cancelled, due to GANTRY_CANCEL_SOURCE_NORMAL_ERROR, in state CStateShot 2021/06/15_09:48:27.415 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_ANODE_TUBE_ARC szInfo: 5:0 APM Arc P1:5/P2: 2021/06/15_09:48:27.418 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_CATHODE_TUBE_ARC szInfo: 5:5 CPM Arc P1:5/P2: 2021/06/15_09:48:29.472 <*ERROR*> Type: NORMAL ErrNum: S_MASTER_SEQUENCER_EXIT_STATE_ERR P1:4096/P2:8369 '================================== Error # and/or description of error: NA Repair Action: Completed or Recommended Repair Action "External": [Confirmation of problem] Confirmed. Work details: I checked the log and reported to the customer. 【Confirmation content】 Good. OneEMS internal remarks : Problem description by engineer : Equipment affected: CT Failure location: 03High Voltage Energy When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user's work flow did the problem occur?: (During Acquisition/Scan) What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem?: NA Engineer's Comment: '============================== $version = "v4.1.10.25210" 9:58 Call description: An "X-ray system cannot comply" error message was displayed during a chest scan and the equipment stopped. The necessary range of data was captured, so he didn't rescan. He asked me to check the log. '============================== What steps are taken during the system diagnosis: Result of response: Problem solved. Failure location: 03High Voltage Energy What steps did you perform to diagnose the problem? '================================== Results of checking the log: The following error occurred. 2021/06/15_09:48:27.317__03 <*HV*> szInfo: SoftArcReg 0x211* P1:0/P2: 2021/06/15_09:48:27.368 <*HV*> Type: NORMAL ErrNum: S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT szInfo: P1:0/P2: 2021/06/15_09:48:27.376 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_ANALOG_FAULT szInfo:2:5 Analog Fault P1:2/P2:32 2021/06/15_09:48:27.378 [CANCEL] CStateShot:Series 285 cancelled, due to GANTRY_CANCEL_SOURCE_NORMAL_ERROR, in state CStateShot 2021/06/15_09:48:27.415 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_ANODE_TUBE_ARC szInfo: 5:0 APM Arc P1:5/P2: 2021/06/15_09:48:27.418 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_CATHODE_TUBE_ARC szInfo: 5:5 CPM Arc P1:5/P2: 2021/06/15_09:48:29.472 <*ERROR*> Type: NORMAL ErrNum: S_MASTER_SEQUENCER_EXIT_STATE_ERR P1:4096/P2:8369 '================================== Error # and/or description of error: NA Completed or Recommended Repair Action "External" : [Confirmation of problem] Confirmed. Work details: I checked the log and reported to the customer. 【Confirmation content】 Good. Completed or Recommended Repair Action "Internal": --Troubleshooting Action-- Result of response: Problem solved. Failure location: 03High Voltage Energy What steps did you perform to diagnose the problem? '================================== Results of checking the log: The following error occurred. 2021/06/15_09:48:27.317__03 <*HV*> szInfo: SoftArcReg 0x211* P1:0/P2: 2021/06/15_09:48:27.368 <*HV*> Type: NORMAL ErrNum: S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT szInfo: P1:0/P2: 2021/06/15_09:48:27.376 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_ANALOG_FAULT szInfo:2:5 Analog Fault P1:2/P2:32 2021/06/15_09:48:27.378 [CANCEL] CStateShot:Series 285 cancelled, due to GANTRY_CANCEL_SOURCE_NORMAL_ERROR, in state CStateShot 2021/06/15_09:48:27.415 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_ANODE_TUBE_ARC szInfo: 5:0 APM Arc P1:5/P2: 2021/06/15_09:48:27.418 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_CATHODE_TUBE_ARC szInfo: 5:5 CPM Arc P1:5/P2: 2021/06/15_09:48:29.472 <*ERROR*> Type: NORMAL ErrNum: S_MASTER_SEQUENCER_EXIT_STATE_ERR P1:4096/P2:8369 '================================== --Repair Action-- --Subject-- 03CM:TM:Mr. Yuki: The equipment stopped due to an error during a scan>> SoftArcReg/S_GEN_FAULT_ANODE_TUBE_ARC/S_GEN_FAULT_CATHODE_TUBE_ARC --Problem Description-- Equipment affected: CT Failure location: 03High Voltage Energy When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user's work flow did the problem occur?: (During Acquisition/Scan) What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem?: NA Engineer's Comment: '============================== $version = "v4.1.10.25210" 9:58 Call description: An "X-ray system cannot comply" error message was displayed during a chest scan and the equipment stopped. The necessary range of data was captured, so he didn't rescan. He asked me to check the log. '============================== --4/8Elements-- Q1. Was the reported problem confirmed? Confirmed. Q2. Were any other problems confirmed? None. Q3. What did you do to resolve the problem?: I checked the log and reported to the customer. Q4: Is the equipment operating normally following the response? Good. Q5. How did the examination finish? None. Q6. Were there any mistaken findings (misdiagnosis, mistaken treatment) due to the problem? None. Q7. Was there any unintended unnecessary exposure? None. Q8. Were there any artifacts? None. ----------------------------------------------------------------------------- | 4 |  | Tube replacement | [12 empty] | 117699766 |  | 1008972 | TRUE | Investigation | Investigation |  | NotEngineering | [5 empty] | Engineering | [6 empty] | Tube | [5 empty] | Refresh |  |  | System issue | System | NotEngineering | ShouldbeOthers | NotEngineering-ShouldbeOthers | NotEngineering-ShouldbeOthers | NotEngineering |  | NotEngineering | NotEngineering-ShouldbeOthers | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | y
1061 |  |  | Japan | Tomakomai City Hospital | 341014 | 4436200% |  | 04DS:TM:A dark band ring artifact in head area images>>Improved with AirCal>>There is a single black ring ar [Truncated] | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: A ring artifact occurred SAP Transfer: Dispatched for Onsite Support SAP Message: Components Added Succesfully to SWO!- SAP Message: LOGI Task added Successfully!- SAP Transfer: Case Transferred to SAP External Remarks:15.06.2021 09:36:29 UTC HIDEO KONISHI (81100219) Request: Ring artifact occurrence Problem confirmation: I confirmed it. Work details: I checked where ring artifacts could be found, such as the detector and bowtie filter, but there were no problems. Also, as the ring artifact had disappeared after the second air calibration, an abnormal calibration could have been created during the first air calibration. The second calibration could have corrected to the correct table. There are no problems now, so he can keep using the system as is. Other problems/Measures taken: None. Equipment status The equipment can be used. Work status Work finished. 14-Jun-2021 81100219 FSE Handover Tool1.3.4.32612.DB13 20210415 SUAT: System Up Again Time: June 15, 2021, 13:30" T2 Activities:Problem Description: Problem description by engineer : Equipment affected: CT Problem location: 04 Gantry Exclude HV When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user's work flow did the problem occur?: (During Acquisition/Scan) What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem?: NA Engineer's Comment: '============================== $version = "v4.1.10.25210" 9:04 Call description: Ring artifacts were found in head images. They can be clearly seen in the head images but there are no problems with the chest scan. Air calibration was performed in the morning '============================== Troubleshooting Action: What steps are taken during the system diagnosis: Result of response: Problem solved. Problem location: 04 Gantry Exclude HV What steps did you perform to diagnose the problem? '================================== Result I had him do air calibration again and he said that improved the situation. (However, after that, the scan stopped due to arcing.） Image confirmation In head images, there is a single black ring artifact near Module 14 or Module 29. It is not a complete circle, it is an arc artifact, and it's not that data is completely missing either. There were no abnormalities during the morning AirCal, but when I looked at the AirCal images after the problem occurred, there were clear rings on 2 of the 4 images. '================================== Error # and/or description of error: NA Repair Action: Completed or Recommended Repair Action "External": The service engineer for the site visited and worked on the equipment. OneEMS internal remarks : Problem description by engineer : Equipment affected: CT Problem location: 04 Gantry Exclude HV When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user's work flow did the problem occur?: (During Acquisition/Scan) What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem?: NA Engineer's Comment: '============================== $version = "v4.1.10.25210" 9:04 Call description: Ring artifacts were found in head images. They can be clearly seen in the head images but there are no problems with the chest scan. Air calibration was performed in the morning '============================== What steps are taken during the system diagnosis: Result of response: Problem solved. Problem location: 04 Gantry Exclude HV What steps did you perform to diagnose the problem? '================================== Result I had him do air calibration again and he said that improved the situation. (However, after that, the scan stopped due to arcing.） Image confirmation In head images, there is a single black ring artifact near Module 14 or Module 29. It is not a complete circle, it is an arc artifact, and it's not that data is completely missing either. There were no abnormalities during the morning AirCal, but when I looked at the AirCal images after the problem occurred, there were clear rings on 2 of the 4 images. '================================== Error # and/or description of error: NA Completed or Recommended Repair Action "External" : The service engineer for the site visited and worked on the equipment. Completed or Recommended Repair Action "Internal": --Troubleshooting Action-- Result of response: Problem solved. Problem location: 04 Gantry Exclude HV What steps did you perform to diagnose the problem? '================================== Result I had him do air calibration again and he said that improved the situation. (However, after that, the scan stopped due to arcing.） Image confirmation In head images, there is a single black ring artifact near Module 14 or Module 29. It is not a complete circle, it is an arc artifact, and it's not that data is completely missing either. There were no abnormalities during the morning AirCal, but when I looked at the AirCal images after the problem occurred, there were clear rings on 2 of the 4 images. '================================== --Repair Action-- Recommended 12Nc: PR1 459800795292:KIT - BR ROHS COLLIMATOR FIELD SERVICE QTY:1 SLOC: Sapporo (U612) PR1 459800911271:TDMS V2 M.Det Class A FRU QTY:1 SLOC: Sapporo (U612) --Subject-- 04DS:TM:A dark band ring artifact in head area images>>Improved with AirCal>>There is a single black ring artifact near Module 14 or Module 29 --Problem Description-- Equipment affected: CT Problem location: 04 Gantry Exclude HV When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user's work flow did the problem occur?: (During Acquisition/Scan) What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem?: NA Engineer's Comment: '============================== $version = "v4.1.10.25210" 9:04 Call description: Ring artifacts were found in head images. They can be clearly seen in the head images but there are no problems with the chest scan. Air calibration was performed in the morning '============================== Safety Questions:Case: 0117699673 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: TAKEYUKI MIYAKAWA Last updated by: TAKEYUKI MIYAKAWA 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Internal Remarks : 15.06.2021 09:36:33 UTC HIDEO KONISHI (81100219) Replacement parts: No replacement parts. [Report on mistakes in parts order] No mistakes. Situation monitoring information: Not monitoring the situation TS Request: No parts to return. Additional CT information: Q1. How did the examination finish? As it stopped before X-ray emission, no extra irradiation Q2. Were there any mistaken findings (misdiagnosis, mistaken treatment) due to the problem? None. Q3. Were there any artifacts? None. Q4. Was there any unintended unnecessary exposure? None. Repair sticker: I affixed a Repaired sticker. [Other information to share] 1. Was treatment/quarantine carried out for a patient who has or is suspected of having new coronavirus? Yes 2. Is it necessary to go into treatment rooms/beds where there are infected patients?(Including workplace and routes of movement) No Did you agree on how to complete the report with the person in charge of devices? Yes. I completed normal reporting that day. 14-Jun-2021 81100219 FSE Handover Tool1.3.4.32612.DB13 20210415 ----------------------------------------------------------------------------- | 2 |  | Air calibration was performed | [12 empty] | 117699673 |  | 1008890 | TRUE | Investigation | Investigation |  | NotEngineering | [5 empty] | Engineering | [6 empty] | Artifact | [5 empty] | Refresh |  |  | System issue | System | NotEngineering | ShouldbeCustomer | NotEngineering-ShouldbeCustomer | NotEngineering-ShouldbeCustomer | NotEngineering |  | NotEngineering | NotEngineering-ShouldbeCustomer | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | y
1062 |  |  | Japan | Social Medical Corporation BOKOI, Nikko Memorial Hospital | 341011 | 4435800% |  | JY: SR FCO implementation | BU Q&R Feedback: No Reason: Lifecycle Solution Problem Reported by Customer: JY: SR FCO implementation Safety Questions:Case: 0117686447 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: YAN Joanna Last updated by: YAN Joanna 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A OneEMS internal remarks: Hashimoto DB FCO72800737 SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Support External Remarks:15.06.2021 14:25:56 UTC HODAKA HASHIMOTO (81100553)[Details of request] FCO72800737 implemented Problem confirmation: I confirmed it. Work details: I performed FCO72800737. KEYBOARD AND RECON SW UPDATE FOR INGENUITY, BIG BORE AND BR64 Also, I implemented it on February 11th. When I replaced the X-ray tube, I reported that the gantry rpm was about 360,000 rotations. Other problems/Measures taken: None. Equipment status The equipment can be used. Work status Work finished. 15-Jun-2021 81100553 FSE Handover Tool1.3.4.32612.DB13 20210415 Internal Remarks:15.06.2021 14:25:59 UTC HODAKA HASHIMOTO (81100553) Replacement parts: No replacement parts. [Report on mistakes in parts order] No mistakes. Situation monitoring information: Not monitoring the situation TS Request: No parts to return. Additional CT information: Q1. How did the examination finish? Q1. No (not applicable) Q2. Were there any mistaken findings (misdiagnosis, mistaken treatment) due to the problem? None. Q3. Were there any artifacts? None. Q4. Was there any unintended unnecessary exposure? None. Repair sticker: I affixed a Repaired sticker. [Other information to share] 1. Was treatment/quarantine carried out for a patient who has or is suspected of having new coronavirus? No 2. Is it necessary to go into treatment rooms/beds where there are infected patients?(Including workplace and routes of movement) No Did you agree on how to complete the report with the person in charge of devices? Yes. I completed normal reporting that day. 15-Jun-2021 81100553 FSE Handover Tool1.3.4.32612.DB13 20210415 ----------------------------------------------------------------------------- | 5 |  | FCO72800737 | [12 empty] | 117686447 |  | 1006868 |  |  | Closed - Non-Complaint | [7 empty] | Others | [6 empty] | FCO72800737 | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1063 |  |  | United States | Philips Medical Capital | 40005 | 4436500% |  | 0402CFXXX no power to gentry | Work Order #: WO-03678282 Ali, Adam The system fully meets specification and returned to use Problem Description by Engineer : Problem description by engineer : No power to system When was the first occurrence? Within... : 18-Jun-2021 How often is issue occurring? : Continuously Malfunction area : Others Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: Not provided Patient Impact/Outcome: No further information is available as per the customer User Impact: System is out of use Current Software Version: Not provided Resolution : Angela walked in to find the power to the gantry & computers was out. She has tried resetting the main breakers. The UPS was beeping earlier but has since stopped. Found the green/blue light on the TEAL is off. Angela will get in touch with the facilities so they can check incoming power from the hospital. I called back later to find they had a flicker in power due to the generator that caused hospital breaker to trip. System is up and running now. Internal Comments : Error # and/or description of error : No power to system ----------------------------------------------------------------------------- | 2 |  | Found the green/blue light on the TEAL is off. | [12 empty] | 117721290 |  | 1020309 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Power | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1064 |  |  | Poland | Rzadowa Agencja | 341008 | 4436400% |  | Security of the tomograph during transport | Work Order #: WO-03672267 Paweł Kierzek Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- Work Order #: WO-03672272 Mateusz Ramadan Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 4 |  | Security of the tomograph during transport Launch of a mobile tomograph in a new location - Katowice Głowackiego 10. Preparation from the service side of the tomograph for transport - container relocation. Work completed successfully. | [12 empty] | 117715858 |  | 1020053 | TRUE | Evaluation | Evaluation | [7 empty] | Others | [6 empty] | Cabin CT transport | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1065 |  |  | Turkey | Republic of Turkey Ministry of Health, Bagcilar | 40010 | 4436500% |  | spill guard order | A new one was ordered because the previous spill guard was torn immediately. |  |  | A new one was ordered because the previous spill guard was torn immediately. | [12 empty] | 117722467 |  | 1024981 | TRUE | Evaluation | Closed - non-complaint | [7 empty] | Others | [6 empty] | spill guard | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1066 |  |  | United States | Philips Medical Capital | 40055 | 4436600% |  | cant log into system its locked the user out | Work Order #: WO-03710811 Boyd Craig The system does not meet specification and is out of use Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: log into system its locked the user out User Impact: None Patient Impact: None Current Software Version: Diagnostic performed by Engineer : User has tried to log in with wrong passwords and is locked out. Customer will call back with PO for service if needed. ----------------------------------------------------------------------------- |  |  | Customer was able to get the emergency IT log on password and log into system. Customer said to close the case. | [12 empty] | 117753160 |  | 1026299 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Password | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1067 |  |  | United States | Philips Medical Capital | 40028 | 4436300% |  | How to make sag and coronals | Work Order #: WO-03668359 Hirchak, Michele The system fully meets specification and returned to use Problem description by Clinical Support Specialist : Customer Function/Role: Tech Device Use: In Use Expected and Unexpected Results: Need sagittals and coronals. User Impact: None Patient Impact: None Current Software Version: 4.1.7 Diagnostic performed by Engineer : Customer said she did not get her sag and coronal results at the time of scan. Walked her through recon and they were still in the list, but were not complete. Had her hit start recon, and they reconned without issue. She is all set. MBH ----------------------------------------------------------------------------- | 5 |  | Q&A | [12 empty] | 117712156 |  | 1017705 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Application | [6 empty] | Q&A | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1068 |  |  | United States | Philips Medical Capital | 40055 | 4436300% |  | 0703CDN84 Screen goes blank after a exam | Work Order #: WO-03663569 Ali, Adam The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Problem description by engineer : Left screen goes blank after an exam When was the first occurrence? Within... : 16-Jun-2021 How often is issue occurring? : Multiple times per week Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Not provided Patient Impact/Outcome: None User Impact: Cannot restart and have to go a hard shutdown the host pc Current Software Version: 4.1 Diagnostic performed by Engineer : Left screen goes blank after an examination, Leah cannot restart the software due to no visibility to the left screen. This has happened a couple of times now since last week. Unable to download hostlogs due an error 'Keyboard-interactive authentication with the SSH2 server failed.' Internal Comments : Error # and/or description of error : Left screen goes blank after an examination Completed or Recommended Repair Action "Internal" : Check host eventlogs for further investigation Inspect/ reseat graphics card, cables. Recommended parts: NVIDIA Quadro P400 2GB (3)mDP GFX 459801536801 CT Gen5.3 Host, Z4 459801493151 Follow Up Required : What are the next Troubleshooting steps : Check host eventlogs for further investigation Inspect/ reseat graphics card, cables. Recommended parts: NVIDIA Quadro P400 2GB (3)mDP GFX 459801536801 CT Gen5.3 Host, Z4 459801493151 Error # and/or description of error : Left screen goes blank after an examination Completed or Recommended Repair Action "Internal" : Check host eventlogs for further investigation Inspect/ reseat graphics card, cables. Recommended parts: NVIDIA Quadro P400 2GB (3)mDP GFX 459801536801 CT Gen5.3 Host, Z4 459801493151 ----------------------------------------------------------------------------- Work Order #: WO-03664336 Donald Vic Problem Description by Engineer : Problem description by engineer : Left screen goes blank after an exam When was the first occurrence? Within... : 16-Jun-2021 How often is issue occurring? : Multiple times per week Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Not provided Patient Impact/Outcome: None User Impact: Cannot restart and have to go a hard shutdown the host pc Current Software Version: 4.1 Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- |  |  | Left screen goes blank after an examination, Leah cannot restart the software due to no visibility to the left screen. This has happened a couple of times now since last week. Unable to download hostlogs due an error 'Keyboard-interactive authentication with the SSH2 server failed.' | [12 empty] | 117707968 |  | 1010412 | TRUE | Evaluation | Investigation |  | TBD-Adapter | [5 empty] | Engineering | [6 empty] | Blank screen | [5 empty] | Refresh |  |  | System issue | System | TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-Adapter |  | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | y
1069 |  |  | United States | Philips Medical Capital | 40028 | 4436100% |  | 9599CDX- Work List filtering is not setup properly - brand new scanner | Work Order #: WO-03654361 Dangos John The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: Diagnostics Expected and Unexpected Results: Customer having His/Ris lissues where they are unable to use filter option to customize their viewing. User Impact: No Patient Impact: No Current Software Version:4.1 Diagnostic performed by Engineer : Customer having His/Ris lissues where they are unable to use filter option to customize their viewing. Follow Up Required : FSE to communicate with Hospital PACS IT team regarding proper setup of his/ris and match with Scanner HSI/RIS settings thru the LANCONFIG app. ----------------------------------------------------------------------------- |  |  | Customer having His/Ris lissues where they are unable to use filter option to customize their viewing. Work with BU and Carestream to resolve the query. Verify WL Query. | [12 empty] | 117698805 |  | 1003539 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | HIS/RIS | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1070 |  |  | United States | Willis Knighton Hospital | 40004 | 4436100% |  | 0308CDN81 Table making noise when moved vertically | Work Order #: WO-03652082 Mcginnis, Mike The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Problem description by engineer : Table making noise when moved vertically When was the first occurrence? Within... : 14-Jun-2021 How often is issue occurring? : First occurrence Malfunction area : Table Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Table making noise when moved vertically Patient Impact/Outcome: none User Impact: hears noise from couch Current Software Version: Not provided Diagnostic performed by Engineer : Spoke to Jennifer. She said the couch is making a terrible noise when they raise or lower it. She said you can hear it from the other scanner. Internal Comments : Error # and/or description of error : Table making noise when moved vertically Completed or Recommended Repair Action "Internal" : Insure the couch is lubricated per the PM manual, check for any cover binding. Check the vertical brake, motor, ball screw and bearings. If defective components found replace parts as necessary Follow Up Required : Error # and/or description of error : Table making noise when moved vertically Completed or Recommended Repair Action "Internal" : Insure the couch is lubricated per the PM manual, check for any cover binding. Check the vertical brake, motor, ball screw and bearings. If defective components found replace parts as necessary ----------------------------------------------------------------------------- Work Order #: WO-03652118 15-Jun-2021 Williams Spence The system fully meets specification and returned to use Problem Description by Engineer : Problem description by engineer : Table making noise when moved vertically When was the first occurrence? Within... : 14-Jun-2021 How often is issue occurring? : First occurrence Malfunction area : Table Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Table making noise when moved vertically Patient Impact/Outcome: none User Impact: hears noise from couch Current Software Version: Not provided Resolution : Lubricated lower couch slides ----------------------------------------------------------------------------- | Dantha, Rajashekar |  | Lubricated lower couch slides | [12 empty] | 117697033 |  | 1002464 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Lubricated | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1071 |  |  | China | Laiyang City People's Hospital | 345005 | 4436900% |  | Wechat:02N/Ingenuity Core128/Transfer Image | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Configuration Error Repair Activity Code: Instructed Customer Problem Reported by Customer : Transfer Image Safety Questions:Case: 0117763367 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Zhou Tingting Last updated by: Zhou Tingting 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: setting images auto save device. Troubleshooting Action: setting images auto save device. Repair Action: instruct customer to set the auto save device, system is ok. Test/Inspection: Functional Check PASSED. The device returned to full functionality. Test Equipment: N/A SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | Troubleshooting Action: setting images auto save device. Repair Action: instruct customer to set the auto save device, system is ok. | [12 empty] | 117763367 |  | 1036037 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Application | [6 empty] | Q&A | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1072 |  |  | Japan | Maki Hospital | 341017 | 4437200% |  | SY: CT application support package. | BU Q&R Feedback: No Reason: Application Support Problem Reported by Customer: SY: CT application support package. Safety Questions:Case: 0117781946 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: YE Smile Last updated by: YE Smile 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A OneEMS internal remarks : Otani DB SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Support ----------------------------------------------------------------------------- | 5 |  | Implementation of study sessions [Details of work] --We held a study session on parameters unique to Philips. | [12 empty] | 117781946 |  | 1048773 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Application | [6 empty] | application | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1073 |  |  | Japan | Maki Hospital | 341017 | 4436900% |  | SY: Replace patient restraint strap. | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : SY: Replace patient restraint strap. Safety Questions:Case: 0117762965 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Smile YE Last updated by: Smile YE 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A OneEMS internal remarks: Kawazoe DB SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Support ----------------------------------------------------------------------------- | 5 |  | Fixed band replacement | [12 empty] | 117762965 |  | 1035757 | TRUE | Evaluation | Closed - non-complaint | [7 empty] | Service | [6 empty] | patient restraint strap | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1074 |  |  | Japan | Independent administrative institution National Hospital Organization | 341015 | 4437000% |  | 01CS:KKI:The CT-Host froze. >> Restarting the CT-Host restored everything. > There was no hardware error. | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 01CS:KKI:The CT-Host froze. >> Restarting the CT-Host restored everything. > There was no hardware error. Safety Questions:Case: 0117768352 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: KAZUMASA KIMURA Last updated by: KAZUMASA KIMURA 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 4 |  | We have not confirmed any hardware errors. Probably the performance has dropped because the reboot hasn't been done for about a week. We are telling you that it can be alleviated by restarting the CT host every day. No. The customer was unable to create a log file or report. | [12 empty] | 117768352 |  | 1039733 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Reboot | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1075 |  |  | United States | Leesburg Regional Medical Center | 40022 | 4437300% |  | 0711CFXXX Locked out of system due to entering the incorrect password | Work Order #: WO-03747860 Miller, John M The system does not meet specification and is out of use Diagnostic performed by Engineer : Collect information from customer. Problem Description by Engineer : Problem description by engineer : Customer states they are locked out of the system due to password incorrect input. When was the first occurrence? Within... : 26-Jun-2021 How often is issue occurring? : First occurrence Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: Can't Scan Patient Impact/Outcome: Not provided User Impact: Can't Scan Current Software Version: 4.1.10 Internal Comments : Error # and/or description of error : Wrong password entered Completed or Recommended Repair Action "Internal" : Local service will need to reset the password. Customer did not want case sent to local service over the weekend. Suggest customer contact IT department, as they should have a password to reset the CT login. Resolution : Error # and/or description of error : Wrong password entered Completed or Recommended Repair Action "Internal" : Local service will need to reset the password. Customer did not want case sent to local service over the weekend. Suggest customer contact IT department, as they should have a password to reset the CT login. ----------------------------------------------------------------------------- | 2 |  | Local service will need to reset the password. Customer did not want case sent to local service over the weekend. Suggest customer contact IT department, as they should have a password to reset the CT login. Not a defect, customer can reset password by themselves according IFU. | [12 empty] | 117788182 |  | 1049644 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Customer | [6 empty] | Password | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1076 |  |  | China | 山东省胸科医院 | 40025 | 4437400% |  | 02N/Ingenuity CT/cannot login | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Calibration/Adjustment Problem Reported by Customer : 02N/Ingenuity CT/ Safety Questions : Case: 0117788721 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Cao Last updated by: Liu Cao 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 无法登录 客户问题详细描述或故障现象： 卡死在Philips桌面 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： FSE 反馈设备为CT5000，异常关机后有时会出现，需要密码狗登录恢复 重装无效 远程诊断与测试结果： 到场恢复系统，需要时收集信息提交escalation PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 到场恢复系统，需要时收集信息提交escalation Test Equipment: N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : Site1 FSE 黄兆坤 18654536122 一直在通话中 External Remarks : 27.06.2021 03:44:26 UTC Zhaokun Huang (86038405)"SUAT. Ready for Clinical Use local time" : "2021-06-27 12:30" External Remarks : 27.06.2021 03:50:39 UTC SUP JCOUSER (86038405)服务内容：用IST权限登录软件后，重新启动Host。 服务结果：维修后，设备恢复正常。 ----------------------------------------------------------------------------- | 2 |  | 服务内容：用IST权限登录软件后，重新启动Host。 服务结果：维修后，设备恢复正常。 收集信息提交escalation Collect issue detail and bugreport, issue happen before and after HF3 https://www.clouddrive.philips.com.cn/file/8uzf62w7mue2ua8ihcovq1cmq23butql# 密码:123 | [12 empty] | 117788721 |  | 1051124 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632553 | [5 empty] | Engineering | [6 empty] | Blue screen | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Blue-Screen | Console-SW-Blue-Screen | Console-SW-Blue-Screen | CTUDT00632553 | Blue-Screen | Console-SW | Console-SW-Blue-Screen | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | y
1077 |  |  | Japan | Independent administrative institution National Hospital Organization | 341015 | 4436800% |  | CT:APP:NO:Mr. Hisaki: Question. | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Application Support Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer: Q: (1) When he was editing the ECG, he tried to reconstruct it by shifting the phase for only one heartbeat, but the scan did not run. (2) What is the CT Viewer VIP? A: (1) It is likely that reconstruction was not possible due to data loss. (2) It stands for Volume Intensity Projection, and the maximum value is allocated to the voxel that is closest to the eye point. Safety Questions:Case: 0117756315 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Ootani Natsuki Last updated by: Ootani Natsuki 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 5 |  | Q: (1) When he was editing the ECG, he tried to reconstruct it by shifting the phase for only one heartbeat, but the scan did not run. (2) What is the CT Viewer VIP? A: (1) It is likely that reconstruction was not possible due to data loss. (2) It stands for Volume Intensity Projection, and the maximum value is allocated to the voxel that is closest to the eye point. | [12 empty] | 117756315 |  | 1031070 | TRUE | Investigation | Investigation | [7 empty] | Application | [6 empty] | Q&A | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1078 |  |  | Japan | 苫小牧市立病院 | 341014 | 4437100% |  | 01DS: KK: Treatment plan examination shows an artifact from the shoulder to the thyroid gland. >> It seems that even if you check the examination of the past treatment plan, it is out. >> Surveyed on FSE onsite. | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 01DS:KK:治療計画の検査で肩から甲状腺に向かってアーチファクトが発生する。>>過去の治療計画の検査を確認していても出ているそうです。>>FSEオンサイトにて調査。 ----------------------------------------------------------------------------- | 4 |  | It was said that an artifact would appear when shooting the treatment plan. Even if you go back in time, it is said that similar artifacts have occurred in the shooting of the treatment plan. It was a field survey by FSE on-site. No. The customer was unable to create a log file or report. | [12 empty] | 117776693 |  | 1045269 | TRUE | Investigation | Investigation |  | TBD_Bugreport | [5 empty] | Engineering | [6 empty] | Artifact | [5 empty] | Refresh |  |  | System issue | System | TBD-Bugreport | TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | TBD_Bugreport |  | TBD-Bugreport | TBD-Bugreport-TBD_Bugreport | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | y
1079 |  |  | Japan | Tomakomai City Hospital | 341014 | 4436900% |  | 03CM: RS: An error occurred during the first scan of the morning >> It was possible to continue with the sca [Truncated] | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 03CM: RS: An error occurred during the first scan of the morning >> It was possible to continue with the scan >> S_MASTER_SEQUENCER_NO_XRAY/S_GEN_FAULT_ANODE_TUBE_ARC Safety Questions:Case: 0117761490 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: RUI SATAKE Last updated by: RUI SATAKE 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 4 |  | You can confirm the occurrence of S_GEN_FAULT_ANODE_TUBE_ARC. It seems that the tube was replaced last Sunday, so I think there is a high possibility of initial discharge. The inspection will continue as it is, and we ask for your understanding by follow-up. | [12 empty] | 117761490 |  | 1034816 | TRUE | Investigation - Pending Translation | Closed - Done |  | CTUDT00630085 | [5 empty] | Engineering | [6 empty] | Tube Arc | [5 empty] | Refresh |  |  | System issue | System | IC-HW | Tub Arc | IC-HW-Tub Arc | IC-HW-Tub Arc | CTUDT00630085 | Tub Arc | IC-HW | IC-HW-Tub Arc | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | y
1080 |  |  | Japan | Independent administrative institution National Hospital Organization | 341013 | 4436900% |  | 01CS:RS: The scan screen was closed during a scan. >> He recreated the plan and continued with the scan. >> [Truncated] | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 01CS:RS: The scan screen was closed during a scan. >> He recreated the plan and continued with the scan. >> Safety Questions:Case: 0117762419 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: RUI SATAKE Last updated by: RUI SATAKE 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 4 |  | The scan reads the surface again and continues the inspection. No noticeable error was found in the logs. Judged as temporary application performance degradation. follow-up. | [12 empty] | 117762419 |  | 1035326 | TRUE | Investigation - Pending Translation | Investigation |  | TBD_Bugreport | [5 empty] | Engineering | [6 empty] | scan screen | [5 empty] | Refresh |  |  | System issue | System | TBD-Bugreport | TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | TBD_Bugreport |  | TBD-Bugreport | TBD-Bugreport-TBD_Bugreport | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | y
1081 |  |  | Japan | Social Medical Corporation BOKOI, Nikko Memorial Hospital | 341011 | 4436900% |  | Via FSE.Contrast agents disappeared from the list.Contrast agents were listed in the Referring Physician sec [Truncated] | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: Via FSE.Contrast agents disappeared from the list.Contrast agents were listed in the Referring Physician section, but only OMNI300 was listed this morning. Safety Questions:Case: 0117762375 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: TAKEYUKI MIYAKAWA Last updated by: TAKEYUKI MIYAKAWA 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 5 |  | As a result, it seems likely that it disappeared due to the operation | [12 empty] | 117762375 |  | 1035292 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | NotEngineering | [5 empty] | Engineering | [6 empty] | contrast agent | [5 empty] | Refresh |  |  | System issue | System | NotEngineering | ShouldbeService | NotEngineering-ShouldbeService | NotEngineering-ShouldbeService | NotEngineering |  | NotEngineering | NotEngineering-ShouldbeService | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | y
1082 |  |  | China | Jiangsu Xuzhou Medical School Affiliated No. 3 Hospital | 40039 | 4436800% |  | 02G/Ingenuity CT/blue screen | BU Q&R Feedback: No SAP Customer Symptom Code: BOOT Customer Symptom Code: Boot up/Power-up Problem Reason: Corrective Maintenance Engineer Failure Code: Electrical/Electronic Failure Repair Activity Code: Reboot/Restart Next Step Action: NTD: This device is not yet under warranty. Please do not add Onsite directly. If we need to dispatch labor, please inform C1 Problem Reported by Customer: 02G/Ingenuity CT/ Next Step Action: Machine install order is 52828478. If need to dispatch, must inform IB to add OP to installation order for FSE Safety Questions:Case: 0117756731 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Weng Jinming Last updated by: Weng Jinming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Troubleshooting Action: instruct customer to restart the system Repair Action: instruct customer to restart the system Test/Inspection: Functional Check PASSED. The device returned to full functionality. Test Equipment: N/A ----------------------------------------------------------------------------- | 3 |  | instruct customer to restart the system | [12 empty] | 117756731 |  | 1031334 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD_Bugreport | [5 empty] | Engineering | [6 empty] | Reboot | [5 empty] | Refresh |  |  | System issue | System | NotEngineering | TBD | NotEngineering-TBD | NotEngineering-TBD | TBD_Bugreport |  | NotEngineering | NotEngineering-TBD | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | y
1083 |  |  | China | 济南市第四人民医院 | 40038 | 4437100% |  | 02N/Ingenuity CT 728326/image transfer performance slow | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Calibration/Adjustment Problem Reported by Customer : 02N/Ingenuity CT 728326 Safety Questions : Case: 0117775438 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Gaoge Last updated by: Wang Gaoge 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 图像传输失败。 客户问题详细描述或故障现象： 主机传图慢 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 主机传图慢 远程诊断与测试结果： 重置传输缓存，检查网络环境，如果还是慢，考虑软件defective PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 1，教会客户重置传输缓存 2 检查医院网络是否千兆节点是否正常 Test/Inspection: N/A Test Equipment: N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 25.06.2021 00:24:14 UTC Tao Chen (86038373)故障现象: 软件运行慢，传图慢 服务内容: 重置spooler，清理垃圾文件 服务结果: 维修后,设备恢复正常. Internal Remarks : 25.06.2021 01:05:43 UTC Tao Chen (86038373) “RDI Accuracy: YES" $Q7: (Confirmed reported problem, software work slow) $Q8: ( Reboot CT host, found it was still slow) $Q9: ( Reset spooler and clean useless file for CT host) $Q10: ( Test and inspection data refer to: PA. Device returned to full functionality) $Q11: Calibrated Tools name ( N/A) ^Current Tube installed date: N/A ^Current Tube usage: N/A ----------------------------------------------------------------------------- | 3 |  | $Q7: (Confirmed reported problem, software work slow) $Q8: ( Reboot CT host, found it was still slow) $Q9: ( Reset spooler and clean useless file for CT host) | [12 empty] | 117775438 |  | 1044463 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-SW-Slow | [5 empty] | Engineering | [6 empty] | SW slow | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Performance | Console-SW-Performance | Console-SW-Performance | TBD-SW-Slow | Performance | Console-SW | Console-SW-Performance | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | y
1084 |  |  | United States | Odessa Regional Medical Center | 40029 | 4437200% |  | Temperture Control Warning | Work Order #: WO-03746425 McCain, Jimmie Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 5 |  | Parts Consumed Description: 1.00000,459800550141,BR64T RoHS HEATER REPLACEMENT KIT 1.00000,459800550141,BR64T RoHS HEATER REPLACEMENT KIT 1.00000,459800980091,BR64 Air Fan Replacement Kit During PM | [12 empty] | 117786463 |  | 1050377 | TRUE | Evaluation | Evaluation | [7 empty] | Others | [6 empty] | Heater | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1085 |  |  | United States | Philips Medical Capital | 40028 | 4437100% |  | 9512CFFB1 4.1 Export Dose Logs | Work Order #: WO-03737120 Giles Brian The system fully meets specification and returned to use Problem description by Clinical Support Specialist : Customer Function/Role: Therapist Device Use: In Use Expected and Unexpected Results: Dose Log archive User Impact: None Patient Impact: None Current Software Version: 4.1 Diagnostic performed by Engineer : Customer states that the CT is not reading the USB. I suggested that she try to insert it again and look for the blinking light with the folder Browser stated the USB is in. She will check the PC she downloads it on. If she doesn't call back she ask for me to close it out ----------------------------------------------------------------------------- | 5 |  | Customer states that the CT is not reading the USB. I suggested that she try to insert it again and look for the blinking light with the folder Browser stated the USB is in. She will check the PC she downloads it on. If she doesn't call back she ask for me to close it out | [12 empty] | 117778050 |  | 1041813 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | USB | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1086 |  |  | United States | Lackey Memorial Hospital | 40015 | 4436900% |  | down button on control box is not working | Work Order #: WO-03723108 22-Jun-2021 Denison Chuck The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Customer Function/Role: C.T. tech Device Use: Philips C.T. Expected and Unexpected Results: System is up. User Impact: None Patient Impact: None Current Software Version:4.1.10 Diagnostic performed by Engineer : Need to replace the C.T. control box. Down button is not working. Follow Up Required : Need to order parts and install next day. C.T. Box at console. ----------------------------------------------------------------------------- Work Order #: WO-03725579 22-Jun-2021 Denison Chuck Problem Description by Engineer : Customer Function/Role: C.T. tech Device Use: Philips C.T. Expected and Unexpected Results: System is up. User Impact: None Patient Impact: None Current Software Version:4.1.10 Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 5 |  | Need to replace the C.T. control box. Down button is not working. | [12 empty] | 117765203 |  | 1034306 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632185 | [5 empty] | Engineering | [6 empty] | CT Box | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | CT Box | Gantry-HW-CT Box | Gantry-HW-CT Box | CTUDT00632185 | CT Box | Gantry-HW | Gantry-HW-CT Box | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | y
1087 |  |  | Japan | 独立行政法人国立病院機構 | 341015 | 4436800% |  | 07CS:KK:中国語のAuto Voiceはありますか？>>Mandarin chineseがあると回答。 | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Application Support Engineer Failure Code: Unknown Cause Repair Activity Code: Instructed Customer Problem Reported by Customer : 07CS:KK:中国語のAuto Voiceはありますか？>>Mandarin chineseがあると回答。 T2 Activities : Problem Description: 対象装置:CT 故障個所:07その他 最初に発生したのはいつか?(~以内):Unknow 問題発生の頻度は?:Unknow エンドユーザｰのワークフローのどの時点において問題が発生しているのか?:その他(Other) 最近どのようなイベントが発生したか?:N.A. 顧客は問題について具体的なシステムイベントレポートを作成しているか?:N.A. Troubleshooting Action: 対応後の結果:その他 その他:お問合せ 故障個所:07その他 診断中にどのようなステップを実施したか?: Mandarin Chineseがあります。と回答。 Repair Action: 【現象確認】 問合せの為、N.A. 【作業内容詳細】 Mandarin Chineseがあります。と回答。 【確認内容】 良好 OneEMS internal remarks : --Troubleshooting Action-- 対応後の結果:その他 その他:お問合せ 故障個所:07その他 診断中にどのようなステップを実施したか?: Mandarin Chineseがあります。と回答。 --Repair Action-- --Subject-- 07CS:KK:ﾋｻｷ様:中国語のAuto Voiceはありますか？>>Mandarin chineseがあると回答。 --Problem Description-- 対象装置:CT 故障個所:07その他 最初に発生したのはいつか?(~以内):Unknow 問題発生の頻度は?:Unknow エンドユーザｰのワークフローのどの時点において問題が発生しているのか?:その他(Other) 最近どのようなイベントが発生したか?:N.A. 顧客は問題について具体的なシステムイベントレポートを作成しているか?:N.A. --4/8Elements-- Q1 報告された問題は確認できましたか? 問合せの為、N.A. Q2 他の問題は確認できましたか? 無し Q3 問題を解決するために何をおこないましたか? Mandarin Chineseがあります。と回答。 Q4 対応後の装置は順調ですか? 良好 Q5 その検査はどの様に終了しましたか? 無し Q6 問題により間違った所見（誤診、誤治療）の有無: 無し Q7 意図しない不要な曝射の有無: 無し Q8 アーチファクトの有無: 無し Safety Questions : Case: 0117756833 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: KENICHI KOJIMA Last updated by: KENICHI KOJIMA 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | Zhu, Chun |  | 2021-06-22 18:42 | [12 empty] | 117756833 |  | 1031418 |  |  | Closed - Non-Complaint | [14 empty] | 2021-06-23 13:30 | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1088 |  |  | United States | Mount Grant General Hospital | 40012 | 4437000% |  | Pitch factor questions | Work Order #: WO-03732938 Hirchak, Michele The system fully meets specification and returned to use Problem description by Clinical Support Specialist : Customer Function/Role: Physicist Device Use: In Use Expected and Unexpected Results: CTDI measurement for ACR different than what the protocol shows. User Impact: None Patient Impact: None Current Software Version: 4.1.7 Diagnostic performed by Engineer : Customer sent some screen shots with the helical and the axial converted protocol. Customer is missing the mA station, and the scan angle for the axial is 420 and should be 360. Also the collimation is different between the 2. I emailed him the screen shots and told him where to fix that. I also sent that to James Hubenschmidt in CLE so he can reach out to the customer and further assist. Called customer and told him I would close this w/o due to 7 day time limit. Will follow up next week. MBH ----------------------------------------------------------------------------- | 5 |  | Pitch factor questions | [12 empty] | 117773844 |  | 1043465 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Application | [6 empty] | Q&A | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1089 |  |  | United States | Philips Medical Capital | 40005 | 4436800% |  | System estop open due to setscrew loose on tilt pot2; Made necessary adjustments, tightened setscrew, ran calibration. | Work Order #: WO-03715902 Mcginnis, Mike The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Problem description by engineer : ESTOP opened When was the first occurrence? Within... : 21-Jun-2021 How often is issue occurring? : First occurrence Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: ESTOP opened Patient Impact/Outcome: none User Impact: had to reset system Current Software Version: Not provided Follow Up Required : Ship TO: The Regional Medical Center 3000 Saint Matthews Rd ORANGEBURG SC 29118-1442 ATT: CT Dept / Philips Error # and/or description of error : ESTOP opened 2021/06/21 09:37:32.705 S_TILT_C_POT1_POT2_KO Completed or Recommended Repair Action "Internal" : Check all connections, replace dual pot assy and calibrate tilt. DUAL POT ASSEMBLY 459800441121 Internal Comments : Error # and/or description of error : ESTOP opened 2021/06/21 09:37:32.705 S_TILT_C_POT1_POT2_KO Completed or Recommended Repair Action "Internal" : Check all connections, replace dual pot assy and calibrate tilt. DUAL POT ASSEMBLY 459800441121 Diagnostic performed by Engineer : Diane said they just completed a scan, Nurse was in the room removing connections to the patient. When she hit the end study ESTOP opened. She power cycled the system, brought it up and ran a test scan with no issues. I downloaded log files and found the tilt error below. She said the Nurse was at the end furthest away from the gantry at the time of the issue. ----------------------------------------------------------------------------- Work Order #: WO-03716017 21-Jun-2021 Beasley Donald The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Problem description by engineer : ESTOP opened When was the first occurrence? Within... : 21-Jun-2021 How often is issue occurring? : First occurrence Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: ESTOP opened Patient Impact/Outcome: none User Impact: had to reset system Current Software Version: Not provided Diagnostic performed by Engineer : checked error logs Follow Up Required : install and calibrate the tilt pot assy Parts Consumed : 1.00000,459800441121,DUAL POT ASSEMBLY ----------------------------------------------------------------------------- Work Order #: WO-03722050 23-Jun-2021 Beasley Donald The system fully meets specification and returned to use Problem Description by Engineer : Problem description by engineer : ESTOP opened When was the first occurrence? Within... : 21-Jun-2021 How often is issue occurring? : First occurrence Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: ESTOP opened Patient Impact/Outcome: none User Impact: had to reset system Current Software Version: Not provided Resolution : found that the setscrew was loose on the tilt pot #2 ----------------------------------------------------------------------------- Work Order #: WO-03725845 23-Jun-2021 Stanczyk, Mark The system fully meets specification and returned to use Problem Description by Engineer : E-stop opens during tilt Resolution : Assisted DB in repair. Removed gantry covers and saw a slight in 90 angle. Jog'd tilt FORWARD ever so slightly and started cal's. Pot1 was close but could use a nudge. Pot2 however we found the set screw to be loose. made necessary adjustments,tightened setscrew, finished cal. easy peasy ----------------------------------------------------------------------------- | 3 |  | Removed gantry covers and saw a slight in 90 angle. Jog'd tilt FORWARD ever so slightly and started cal's. Pot1 was close but could use a nudge. Pot2 however we found the set screw to be loose. made necessary adjustments,tightened setscrew, finished cal. easy peasy | [12 empty] | 117758400 |  | 1028247 | TRUE | Closed - Done | Closed - Done | [7 empty] | Others | [6 empty] | Dual pot | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1090 |  |  | Germany | Gesundheit Nord gGmbH Klinikverbund | 40002 | 4436900% |  | Gantry briefly turned off and on again twice during operation | Work Order #: WO-03722535 Eduard Berthold The system partly meets specification and is returned to use with restrictions Problem Description by Engineer: Call-back to the customer, Raik Glogner-Umbeer Gantry turned off 2 times Diagnostic performed by engineer: The following actions were performed for error analysis / error correction: Error log diagnosis: S_VMC_AXIS_UNALLOWED_MOVE 12:35, 11:38 Experience: Replacing the Vertical Transducer (String Pot) resolved the error 459800623331 Vertical transducer Follow-up required: CT still in operation Check and clean brake Replacement as first thing: 459800623331 Vertical transducer 2 hours labor The customer wants info tomorrow morning, but I expect a system failure tomorrow! ----------------------------------------------------------------------------- | 3 |  | Replacing the Vertical Transducer (String Pot) resolved the error 459800623331 Vertical transducer | [12 empty] | 117764743 |  | 1032671 | TRUE | Investigation | Investigation |  | NotEngineering-partreplace | [5 empty] | Engineering | [6 empty] | Vertical transducer | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | VerticalTransducer | Couch-HW-VerticalTransducer | Couch-HW-VerticalTransducer | NotEngineering-partreplace |  | Couch-HW | Couch-HW-VerticalTransducer | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | y
1091 |  |  | China | 林甸县医院 | 345038 | 4437800% |  | 02N/Ingenuity Core128/Elite China/need push IST | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: Calibration/Adjustment Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ Safety Questions : Case: 0117818159 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wennian Han Last updated by: Wennian Han 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: need push IST Troubleshooting Action: site engineer and customer forget IT account and olevel account password, so PM engineer can not enter in olevel account. want to push ISP Repair Action: Push IST failed, FSE feedback he will reset the password when next time onsite service Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: NA SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 5 |  | Troubleshooting Action: site engineer and customer forget IT account and olevel account password, so PM engineer can not enter in olevel account. want to push IST Repair Action: Push IST failed, FSE feedback he will reset the password when next time onsite service | [12 empty] | 117818159 |  | 1069652 | TRUE | Opened - Pending Translation | Closed - non-complaint | [7 empty] | Customer | [6 empty] | Password | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1092 |  |  | China | 山东省泰安市中医二院 | 345037 | 4437800% |  | 02N/Ingenuity Core128/Elite China/timeout | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ Safety Questions : Case: 0117817735 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wennian Han Last updated by: Wennian Han 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 扫描失败。 客户问题详细描述或故障现象： 扫描超时 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： S MASTER SEQUENCER PREP DC READY TIMEOUT 远程诊断与测试结果： 扫描超时，重启后恢复，客户需要工程师到场检查设备 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 检查供电，检查DC相关 SAP Transfer : Case Transferred to SAP External Remarks : 01.07.2021 15:09:28 UTC Tao Chen (86038373)故障现象: CT扫描间歇性报错 服务内容: 重刷Rhost应用程序，做灯丝, 空气校正 服务结果: 维修后,设备恢复正常. Internal Remarks : 01.07.2021 15:16:00 UTC Tao Chen (86038373) “RDI Accuracy: YES" $Q7: (Confirmed reported problem, CT scanning failed intermittently) $Q8: ( Check gantry log, found "S_MASTER_SEQUENCER_PREP_DC_SERIES_READY_TIMEOUT" error) $Q9: ( Do boot flash burner and application flash burner for Rhost, restore NVROM) $Q10: ( Test and inspection data refer to: PA. Device returned to full functionality) $Q11: Calibrated Tools name ( N/A) ^Current Tube installed date: N/A ^Current Tube usage: N/A SAP Transfer : Dispatched for Onsite Support ----------------------------------------------------------------------------- | 3 |  | Do boot flash burner and application flash burner for Rhost, restore NVROM, filament and air calibration | [12 empty] | 117817735 |  | 1069371 | TRUE | Opened - Pending Translation | Closed - Done |  | Timeout | [5 empty] | Engineering | [6 empty] | Timeout | [5 empty] | Refresh |  |  | System issue | System | TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | Timeout |  | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | y
1093 |  |  | China | 正镶白旗医院 | 345033 | 4438000% |  | 02N/Core128/reset password | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Core128/ Next Step Action : FSE：苏瑞杰-13304719335 call in FSE Notes : This PO is a Reference Number only Safety Questions : Case: 0117830412 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: reset password Troubleshooting Action: reset password Repair Action: reset password SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support ----------------------------------------------------------------------------- | 5 |  | reset password | [12 empty] | 117830412 |  | 1075696 | TRUE | Opened - Pending Translation | Closed - non-complaint | [7 empty] | Customer | [6 empty] | Password | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1094 |  |  | China | 息烽县中医医院 | 345014 | 4437900% |  | Wechat:02N/Ingenuity Core128/Ct box down button stuck | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : Wechat:02N/Ingenuity Core128/Ct box down button stuck T2 Activities : Problem Description: 报修为工作站 EQ87453892已修改 SAP Transfer : Case Transferred to SAP Safety Questions : Case: 0117826032 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wennian Han Last updated by: Wennian Han 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 控制面板按键失灵。 客户问题详细描述或故障现象： ctbox按键问题 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： NA 远程诊断与测试结果： CTBOX下降键不灵，其他键正常。机架面板下降键控制正常。 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 检查CTBOX；检查接线、供电等。 建议备件12NC及名称： " 459801705022 CT BOX ASSEMBLY_BEST" SAP Transfer : Dispatched for Onsite Support Next Step Action : PPD CT BOX ASSEMBLY_BEST 459801705022 1 备件无货，已安排订货，FSE：唐广林-86-18084189777 Next Step Action : tang guanglin ----------------------------------------------------------------------------- | 5 |  | 459801705022 CT BOX ASSEMBLY_BEST Ct box down button stuck | [12 empty] | 117826032 |  | 1074175 | TRUE | Opened - Pending Translation | Evaluation - Pending Translation |  | CTUDT00632185 | [5 empty] | Engineering | [6 empty] | CT Box | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | CT Box | Gantry-HW-CT Box | Gantry-HW-CT Box | CTUDT00632185 | CT Box | Gantry-HW | Gantry-HW-CT Box | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | y
1095 |  |  | Japan | 医療法人清翠会 牧病院 | 341017 | 4438000% |  | 01CS: KK: Image transfer to PACS is FAIL, but PACS is able to receive images. >> It seems that it was not sent to a device called Radamés. | BU Q&R Feedback: No SAP Customer Symptom Code: IONW Customer Symptom Code: Interoperability/Network Issue Reason: Application Support Engineer Failure Code: Unknown Cause Repair Activity Code: Instructed Customer Problem Reported by Customer : 01CS:KK:PACSへの画像転送がFAILになるがPACSでは画像受信できている。>>送れていなかったのはRadamesというデバイスに対してだったようです。 T2 Activities : Problem Description: Problem description by engineer : 対象装置:CT 故障個所:01Cosole 最初に発生したのはいつか?(~以内):1 Hour 問題発生の頻度は?:First Occurrence エンドユーザｰのワークフローのどの時点において問題が発生しているのか?:ネットワーク関連(Network) 最近どのようなイベントが発生したか?:N.A. 顧客は問題について具体的なシステムイベントレポートを作成しているか?:いいえ。お客様にログファイル、レポートは作成して頂けませんでした。 Troubleshooting Action: What steps are taken during the system diagnosis : 対応後の結果:改善された 故障個所:01Cosole 診断中にどのようなステップを実施したか?: PACSには画像は送られていました。 Radmesというデバイスへの転送がFailしていたようです。 お客様の勘違いだったそうです。様子見となりました。 Error # and/or description of error : NA Repair Action: Completed or Recommended Repair Action "External" : 【現象確認】 問合せの為、N.A. 【作業内容詳細】 PACSには画像は送られていました。 Radmesというデバイスへの転送がFailしていたようです。 お客様の勘違いだったそうです。様子見となりました。 【確認内容】 良好 OneEMS internal remarks : Problem description by engineer : 対象装置:CT 故障個所:01Cosole 最初に発生したのはいつか?(~以内):1 Hour 問題発生の頻度は?:First Occurrence エンドユーザｰのワークフローのどの時点において問題が発生しているのか?:ネットワーク関連(Network) 最近どのようなイベントが発生したか?:N.A. 顧客は問題について具体的なシステムイベントレポートを作成しているか?:いいえ。お客様にログファイル、レポートは作成して頂けませんでした。 What steps are taken during the system diagnosis : 対応後の結果:改善された 故障個所:01Cosole 診断中にどのようなステップを実施したか?: PACSには画像は送られていました。 Radmesというデバイスへの転送がFailしていたようです。 お客様の勘違いだったそうです。様子見となりました。 Error # and/or description of error : NA Completed or Recommended Repair Action "External" : 【現象確認】 問合せの為、N.A. 【作業内容詳細】 PACSには画像は送られていました。 Radmesというデバイスへの転送がFailしていたようです。 お客様の勘違いだったそうです。様子見となりました。 【確認内容】 良好 Completed or Recommended Repair Action "Internal" : --Troubleshooting Action-- 対応後の結果:改善された 故障個所:01Cosole 診断中にどのようなステップを実施したか?: PACSには画像は送られていました。 Radmesというデバイスへの転送がFailしていたようです。 お客様の勘違いだったそうです。様子見となりました。 --Repair Action-- --Subject-- 01CS:KK:ﾏﾂﾓﾄ様:PACSへの画像転送がFAILになるがPACSでは画像受信できている。>>送れていなかったのはRadamesというデバイスに対してだったようです。 --Problem Description-- 対象装置:CT 故障個所:01Cosole 最初に発生したのはいつか?(~以内):1 Hour 問題発生の頻度は?:First Occurrence エンドユーザｰのワークフローのどの時点において問題が発生しているのか?:ネットワーク関連(Network) 最近どのようなイベントが発生したか?:N.A. 顧客は問題について具体的なシステムイベントレポートを作成しているか?:いいえ。お客様にログファイル、レポートは作成して頂けませんでした。 --4/8Elements-- Q1 報告された問題は確認できましたか? 問合せの為、N.A. Q2 他の問題は確認できましたか? 無し Q3 問題を解決するために何をおこないましたか? PACSには画像は送られていました。 Radmesというデバイスへの転送がFailしていたようです。 お客様の勘違いだったそうです。様子見となりました。 Q4 対応後の装置は順調ですか? 良好 Q5 その検査はどの様に終了しましたか? 無し Q6 問題により間違った所見（誤診、誤治療）の有無: 無し Q7 意図しない不要な曝射の有無: 無し Q8 アーチファクトの有無: 無し Safety Questions : Case: 0117830498 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: KENICHI KOJIMA Last updated by: KENICHI KOJIMA 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 4 |  | Images were sent to PACS. It seems that the transfer to a device called Radmes was failing. It seems that it was a customer's misunderstanding. It was a wait-and-see. The customer was unable to create a log file or report. | [12 empty] | 117830498 |  | 1075720 | TRUE | Evaluation - Pending Translation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | misunderstanding | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1096 |  |  | Japan | Social Medical Corporation BOKOI, Nikko Memorial Hospital | 341011 | 4437700% |  | 03DS:KK: Scanning stopped during a contrast scan.>> He was able to perform the examination with a rescan.>>S [Truncated] | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 03DS:KK: Scanning stopped during a contrast scan.>> He was able to perform the examination with a rescan.>>S_STARTER_FAULT_NOT_UP_TO_SPEED SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Support T2 Activities:Problem Description: Problem description by engineer : Equipment affected: CT Failure location: 03High Voltage Energy When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user's work flow did the problem occur?: (During Acquisition/Scan) What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem? No.The customer has not created a log file or report. Troubleshooting Action: What steps are taken during the system diagnosis: Result of response: Problem solved. Failure location: 03High Voltage Energy What steps did you perform to diagnose the problem? Stopped during a contrast examination. S_STARTER_FAULT_NOT_UP_TO_SPEED and S_ROTOR_STOP_ANODE_NO_MORE_SPINNING had occurred. The examination was possible with a rescan. Sugawara will respond today. Error # and/or description of error: NA Repair Action: Completed or Recommended Repair Action "External": The service engineer for the site visited and worked on the equipment. OneEMS internal remarks : Problem description by engineer : Equipment affected: CT Failure location: 03High Voltage Energy When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user's work flow did the problem occur?: (During Acquisition/Scan) What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem? No.The customer has not created a log file or report. What steps are taken during the system diagnosis: Result of response: Problem solved. Failure location: 03High Voltage Energy What steps did you perform to diagnose the problem? Stopped during a contrast examination. S_STARTER_FAULT_NOT_UP_TO_SPEED and S_ROTOR_STOP_ANODE_NO_MORE_SPINNING had occurred. The examination was possible with a rescan. Sugawara will respond today. Error # and/or description of error: NA Completed or Recommended Repair Action "External" : The service engineer for the site visited and worked on the equipment. Completed or Recommended Repair Action "Internal": --Troubleshooting Action-- Result of response: Problem solved. Failure location: 03High Voltage Energy What steps did you perform to diagnose the problem? Stopped during a contrast examination. S_STARTER_FAULT_NOT_UP_TO_SPEED and S_ROTOR_STOP_ANODE_NO_MORE_SPINNING had occurred. The examination was possible with a rescan. Sugawara will respond today. --Repair Action-- Recommended 12Nc: --None--453567560042:80KW High Speed Starter Assembly Qty: 1SLOC:--None-- --Subject-- 03DS:KK:Mr. Tsukakoshi: Scanning stopped during a contrast scan.>> He was able to perform the examination with a rescan.>>S_STARTER_FAULT_NOT_UP_TO_SPEED --Problem Description-- Equipment affected: CT Failure location: 03High Voltage Energy When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user's work flow did the problem occur?: (During Acquisition/Scan) What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem? No.The customer has not created a log file or report. SAP Message: LOGI Task added Successfully!- Safety Questions:Case: 0117806825 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: KENICHI KOJIMA Last updated by: KENICHI KOJIMA 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Message: Components Added Succesfully to SWO!- Internal Remarks: 30.06.2021 08:30:14 UTC MOTOTSUGU SUGAWARA (81100170) Replacement parts: Replacement parts: As below 12Nc:453567560042 Name: 80KW High Speed Starter Assembly QTY:1 [Report on mistakes in parts order] No mistakes. Situation monitoring information: Not monitoring the situation TS Request: No parts to return. Additional CT information: Q1. How did the examination finish? As it stopped before X-ray emission, no extra irradiation Q2. Were there any mistaken findings (misdiagnosis, mistaken treatment) due to the problem? None. Q3. Were there any artifacts? None. Q4. Was there any unintended unnecessary exposure? None. Repair sticker: I affixed a Repaired sticker. [Other information to share] 1. Was treatment/quarantine carried out for a patient who has or is suspected of having new coronavirus? No 2. Is it necessary to go into treatment rooms/beds where there are infected patients?(Including workplace and routes of movement) No Did you agree on how to complete the report with the person in charge of devices? Yes. I completed normal reporting that day. 30-Jun-2021 81100170 FSE Handover Tool1.3.4.32612.DB13 20210115 External Remarks : 30.06.2021 08:30:11 UTC MOTOTSUGU SUGAWARA (81100170) [Replacement parts] Replacement parts: As below 12Nc:453567560042 Name: 80KW High Speed Starter Assembly QTY:1 [Report on mistakes in parts order] No mistakes. Situation monitoring information: Not monitoring the situation TS Request: No parts to return. Additional CT information: Q1. How did the examination finish? As it stopped before X-ray emission, no extra irradiation Q2. Were there any mistaken findings (misdiagnosis, mistaken treatment) due to the problem? None. Q3. Were there any artifacts? None. Q4. Was there any unintended unnecessary exposure? None. Repair sticker: I affixed a Repaired sticker. [Other information to share] 1. Is treatment or quarantining of confirmed or suspected Covid-19 patients taking place? No 2. Is it necessary to go into treatment rooms/beds where there are infected patients?(Including work area as well as movement corridors.) No Did you agree on how to complete the report with the person in charge of devices? Yes. I completed normal reporting that day. 30-Jun-2021 81100170 FSE Handover Tool1.3.4.32612.DB13 20210115 ----------------------------------------------------------------------------- | 4 |  | Stopped during a contrast examination. S_STARTER_FAULT_NOT_UP_TO_SPEED and S_ROTOR_STOP_ANODE_NO_MORE_SPINNING had occurred. Replacement parts: As below 12Nc:453567560042 Name: 80KW High Speed Starter Assembly QTY:1 | [12 empty] | 117806825 |  | 1064632 | TRUE | Evaluation | Closed - Done |  | NotEngineering | [5 empty] | Engineering | [6 empty] | High Speed Starter Assembly | [5 empty] | Refresh |  |  | System issue | System | NotEngineering | ShouldbeOthers | NotEngineering-ShouldbeOthers | NotEngineering-ShouldbeOthers | NotEngineering |  | NotEngineering | NotEngineering-ShouldbeOthers | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | y
1097 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 337000 | 4437900% |  | WS designation | Work Order #: WO-03785420 Soner Saglam The system fully meets specification and returned to use Problem Description by Engineer : User requested to define WS to CT. Resolution : It is understood that the specified WS is already defined.----------------------------------------------------------------------------- | 3 |  | Problem Description by Engineer : User requested to define WS to CT. Resolution : It is understood that the specified WS is already | [12 empty] | 117826878 |  | 1070764 | TRUE | Opened - Pending Translation | Closed - non-complaint | [7 empty] | Customer | [6 empty] | misunderstanding | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1098 |  |  | Japan | 医療法人清翠会 牧病院 | 341017 | 4437800% |  | CT:APP:RK：質問：ｲﾇｲさま | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Application Support Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer : Q:brain2DとheadでMPRを作成すると印象が結構変わりますが、プロトコルでAutoで作成したMPRは変化が小さくなっているように感じるのですが、slice厚はどのようになっていますか＞ A:"brain2Dは最小slice厚1.25ｍｍ、headでは最小slice厚0.8mmです。しかしプロトコルでautoで作成されるMPRのslice厚が不明であったため、thin slice settingより確認いただきました。 確認できたslice厚はbrain2D0.625mm,head0.8mmであったため、これがMPRの見た目を変えている可能性があることをお伝えしました。" OneEMS internal remarks : Q1:報告された問題は確認できましたか？：確認済 Q2：他の問題は確認できましたか？；無 Q3：問題を解決するために何を行いましたか？ １：電話で解決した場合：説明 Q4：対応後の装置は順調ですか？：良好 Safety Questions : Case: 0117820388 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: RYUSUKE KUJIRAI Last updated by: RYUSUKE KUJIRAI 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 44379.8236 | [14 empty] | 117820388 |  | 1071024 |  |  | Closed - Non-Complaint | [20 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1099 |  |  | Japan | 医療法人清翠会 牧病院 | 341017 | 4437800% |  | CT:APP:RK:質問:ﾎｼさま | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Application Support Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer : Q:頭のscantypeでbrain2Dとheadがありますが、2つは違いますか？ A:はい。それぞれは設定できる項目が異なります。選択できるフィルターや最小slice厚など OneEMS internal remarks : Q1:報告された問題は確認できましたか？：確認済 Q2：他の問題は確認できましたか？；無 Q3：問題を解決するために何を行いましたか？ １：電話で解決した場合：説明 Q4：対応後の装置は順調ですか？：良好 Safety Questions : Case: 0117820343 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: RYUSUKE KUJIRAI Last updated by: RYUSUKE KUJIRAI 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 44379.8201 | [14 empty] | 117820343 |  | 1070998 |  |  | Closed - Non-Complaint | [20 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1100 |  |  | China | Dongguan Donghua Hospital Co., Ltd. | 40059 | 4437700% |  | WeChat:02N/Ingenuity CT 728326/Add node | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Calibration/Adjustment Problem Reported by Customer:Add node Safety Questions:Case: 0117808304 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Cao Last updated by: Liu Cao 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer's problems: Plan to add devices tomorrow and need assistance in adding nodes Detailed description of Customer's problems or malfunctions: Plan to add devices tomorrow and need assistance in adding nodes It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: NA Remote troubleshoot and test result: Contact the customer to confirm the time to assist in adding nodes PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Contact the customer to confirm the time to assist in adding nodes Test/Inspection: N/A Test Equipment: N/A SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support ----------------------------------------------------------------------------- | 5 |  | 计划明后天添加设备 需协助添加节点 | [12 empty] | 117808304 |  | 1064976 | FALSE | Opened | Closed - non-complaint | [7 empty] | NAC | [6 empty] | PACS Config | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1101 |  |  | United States | Philips Medical Capital | 40055 | 4437900% |  | 9507CFFD7 NYN Changing Filter | Work Order #: WO-03788656 Mujica, David The system fully meets specification and returned to use Problem description by Clinical Support Specialist : Customer Function/Role: Technologist Device Use: In use Expected and Unexpected Results: Changing to Bone filter User Impact: none Patient Impact: none Current Software Version: 4.1 Diagnostic performed by Engineer : Chelsey was assisted in changing a filter on their CT Soft Tissue neck Exam Card via Exam Card Manager, show all parameters, Bone Recon. Exploration of filters in show all parameters was explored and changes were made successfully and saved. Chelsey had no further questions or need for clinical support. No rescan, no loss of data, no reinjection per customer. ----------------------------------------------------------------------------- | 5 |  | assisted in changing a filter on their CT Soft Tissue neck Exam Card via Exam Card Manager | [12 empty] | 117829629 |  | 1072793 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | NAC | [6 empty] | Application support | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1102 |  |  | China | 浙江省乐清第一人民医院 | 40054 | 4437700% |  | Wechat:02N/Ingenuity CT 728326/scan interrput | BU Q&R Feedback: No SAP Customer Symptom Code: BOOT Customer Symptom Code: Boot up/Power-up Problem Reason: Corrective Maintenance Engineer Failure Code: Electrical/Electronic Failure Repair Activity Code: Replace Part Problem Reported by Customer : scan interrput Safety Questions : Case: 0117811180 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Weng Jinming Last updated by: Weng Jinming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Troubleshooting Action: repeat case0117811185 Repair Action: repeat case0117811185 Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | repeat case0117811185 | [12 empty] | 117811180 |  | 1065963 | FALSE | Evaluation | Closed - Duplicate | [7 empty] | NAC | [6 empty] | Duplicate | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1103 |  |  | China | Zhejiang Province Yueqing No. 1 People’s Hospital | 40054 | 4437700% |  | 02N/Ingenuity CT 728326/scan interrupt | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 02N/Ingenuity CT 728326 Safety Questions:Case: 0117811185 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer's problems: Other types of error. Detailed description of Customer's problems or malfunctions: Scanning was interrupted. It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: INvalid dfs settings Remote troubleshoot and test result: R2D malfunction. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Replace R2D. 0 R2D Assembly 459800467343 SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Internal Remarks: 01.07.2021 04:05:56 UTC Yi Huang (86036249) RDI Accuracy:Yes $Q7: (Confirmed reported problem&system scan stop ) $Q8: (check errorlog found acq not ok&invalid dfs settings, R2D LED5 is red) $Q9: (replace R2d) $Q10: (Test and inspection data refer to: PA. Device returned to full functionality.) $Q11: Calibrated tools name(N/A) S/N (N/A) Next calibration date (N/A). For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date:N/A ^Current Tube usage: second:N/A External Remarks: 01.07.2021 04:05:45 UTC SUP JCOUSER (86036249) service details: enhanced scan stopped. Checked the log and found DFS setting abnormal. Replace R2D. service result: device returns to normal working condition after service Other significant matters: none Next step Action: PPD parts R2D Assembly 459800467341 1 Have arranged for 12H to the site. (alternative number has been confirmed) FSE: Huang Yi-86-15958725282 ----------------------------------------------------------------------------- | 3 |  | R2D Assembly 459800467341 $Q7: (Confirmed reported problem&system scan stop) $Q8: (check errorlog found acq not ok&invalid dfs settings, R2D LED5 is red) $Q9: (replace R2d) $Q10: (Test and inspection data refer to: PA.Device returned to full | [12 empty] | 117811185 |  | 1065964 | TRUE | Opened | Closed - Done |  | CTUDT00630864 | [5 empty] | Engineering | [6 empty] | R2D | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | R2D | Gantry-HW-R2D | Gantry-HW-R2D | CTUDT00630864 | R2D | Gantry-HW | Gantry-HW-R2D | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | y
1104 |  |  | China | Guangxi Province Teng County People’s Hospital | 40036 | 4437800% |  | application/Ingenuity CT 728326 | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer:application Safety Questions:Case: 0117817621 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Zixuan Last updated by: Wang Zixuan 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 3 |  | 杨俊13817647801 & 郭凡 13501285722均未接通，已邮件 | [12 empty] | 117817621 |  | 1069326 | TRUE | Opened | Closed - non-complaint | [7 empty] | Application | [6 empty] | Application | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1105 |  |  | China | 北京市中西医结合医院 | 40032 | 4437800% |  | 02N/Ingenuity CT 728326/ECG Loss | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/IntelliSpace IX/LX Workstations/ T2 Activities : Problem Description: 客户问题概述： 新装机，没有安装心电系统 设备本身技术问题： 是。 Troubleshooting Action: 远程诊断与测试结果： 需要工程师到场，安装心电系统 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 需要工程师到场，安装心电系统 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Safety Questions : Case: 0117817822 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Weng Jinming Last updated by: Weng Jinming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A External Remarks : 01.07.2021 13:55:50 UTC SUP JCOUSER (86038892)服务内容：机器ECG无法使用，更换EKG线缆后正常。 服务结果：服务后设备恢复正常。 Internal Remarks : 01.07.2021 13:55:53 UTC Xiang Zhang (86038892) "RDI Accuracy: Yes" $Q7: Confirm reported problems:The system has no ECG signal. $Q8: Trouble shooting steps:The ECG cable is in bad condition,the ECG cable needs to be replaced $Q9: Function returns to normal after the ECG cable is replaced。 $Q10: Test and Inspection data refer to PA& Device status:Device returned to full functionally $Q11: Calibrated Tools refer to PA ----------------------------------------------------------------------------- | 5 |  | The system has no ECG signal. The ECG cable is in bad condition,the ECG cable needs to be replaced | [12 empty] | 117817822 |  | 1069439 | TRUE | Opened - Pending Translation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | ECG | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1106 |  |  | China | Beijing City Hospital of Traditional Chinese and Western Medicine | 40032 | 4437600% |  | 02N/Ingenuity CT 728326/CT box fault | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : Ingenuity CT 728326 Safety Questions : Case: 0117798433 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingsong Last updated by: Lu Mingsong 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： CT BOX按键故障 客户问题详细描述或故障现象： CT BOX按键故障 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： N/A 远程诊断与测试结果： N/A PRS TR状态： 无PRS。 Repair Action: 建议现场维修步骤： 更换CT BOX。 建议备件12NC及名称： CT BOX ASSEMBLY_BEST：459801705022 客户需求备注： 无 SAP Transfer : Dispatched for Onsite Support External Remarks : 02.07.2021 12:52:58 UTC Guochao Fan (86039456)服务内容：更换CTbox 服务结果：服务后设备恢复正常。 Internal Remarks : 02.07.2021 12:53:07 UTC Guochao Fan (86039456) "RDI Accuracy: Yes" $Q7: Confirm reported problems: ctbox Out of bed key malfunction $Q8: Trouble shooting steps:for Out of bed key malfunction .The ctbox is in bad condition,the ctbox needs to be replaced $Q9: Function returns to normal after the ctbox is replaced。 $Q10: Test and Inspection data refer to PA& Device status:Device returned to full functionally $Q11: N/A SAP Transfer : Case Transferred to SAP Next Step Action : PPD 459801705022 *1 ,已安排订货，FSE范国超-86-13520729230 ----------------------------------------------------------------------------- | 3 |  | $Q7: Confirm reported problems: ctbox Out of bed key malfunction $Q8: Trouble shooting steps:for Out of bed key malfunction .The ctbox is in bad condition,the ctbox needs to be replaced $Q9: Function returns to normal after the ctbox is replaced CT BOX ASSEMBLY_BEST：459801705022 | [12 empty] | 117798433 |  | 1059417 | TRUE | Opened - Pending Translation | Closed - No Further Investigation |  | CTUDT00632185 | [5 empty] | Engineering | [6 empty] | CT Box | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | CT Box | Gantry-HW-CT Box | Gantry-HW-CT Box | CTUDT00632185 | CT Box | Gantry-HW | Gantry-HW-CT Box | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | y
1107 |  |  | China | Beijing City Hospital of Traditional Chinese and Western Medicine | 40032 | 4437600% |  | 05R/Ingenuity CT 728326/password failure | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Calibration/Adjustment SAP Transfer:Case Transferred to SAP Problem Reported by Customer: Password failure Safety Questions:Case: 0117797920 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Tao Fengming Last updated by: Tao Fengming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer's problems: Other types of error. Detailed description of Customer's problems or malfunctions: User password has too many error, account is locked It's a technology problem with the equipment itself: No. Troubleshooting Action: key fault code or symptom: Password is locked, fail to log in Remote troubleshoot and test result: No PRS, admin password has also been forgot, need to reset password on site PRS TR status: No PRS. Repair Action: Suggest onsite service procedures: 1. Reset the password. Test/Inspection: N/A Test Equipment: N/A SAP Transfer:Dispatched for On site Support External Remarks: 29.06.2021 05:38:01 UTC HongHong Han (86039436) service details: change login password Service Result: The device returned to working normally. Internal Remarks: 29.06.2021 05:38:11 UTC HongHong Han (86039436) "RDI Accuracy:Yes" $Q7: Confirmed reported problem, need to change password $Q8: need to change password $Q9: change ct and it password $Q10: Test and Inspection data refer to PA & Device status returned to fully functionality. $Q11:N/A Next Step Action: dispatch engineer han honghong ----------------------------------------------------------------------------- | 3 |  | $Q7: Confirmed reported problem, need to change password $Q8: need to change password $Q9: change ct and it password | [12 empty] | 117797920 |  | 1059165 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Password | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1108 |  |  | China | Henan Province Puyang City People’s Hospital | 40031 | 4437500% |  | Wechat:02N/Ingenuity CT 728326/tube over heat | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer:tube overheating Safety Questions:Case: 0117791222 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Jianqiang Last updated by: Wang Jianqiang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Case Transferred to SAP T2 Activities:Problem Description: Summary of customer's problems: The old tube overheated recently, it used to be OK Detailed description of Customer's problems or malfunctions: The old tube overheated recently, it used to be OK It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: Tube overheat Remote troubleshoot and test result: Remote instruct to modify scan MAS. Customer request engineer's on site. Contact engineer who said no service had been done yet PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Check mAs, check the minimum/maximum values of the setting, check filter mesh SAP Transfer:Dispatched for On site Support ----------------------------------------------------------------------------- | 5 |  | 最近老报球管过热，以前是好的 收集bugreport，calibration info | [12 empty] | 117791222 |  | 1054371 | TRUE | Opened | Closed - Done |  | Tube Heat | [5 empty] | Engineering | [6 empty] | Tube Overheat | [5 empty] | Refresh |  |  | System issue | System | Gantry-SW | Tube over heat | Gantry-SW-Tube over heat | Gantry-SW-Tube over heat | Tube Heat | Tube Heat | Gantry-SW | Gantry-SW-Tube over heat | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | y
1109 |  |  | United States | Central Vermont Medical Center | 40024 | 4437500% |  | 0711CFXXX Storage issue on local | Work Order #: WO-03754268 Morrow, Dan The system fully meets specification and returned to use Problem Description by Engineer : Problem description by engineer : Local database percentage dropping. Only 2 days of studies. How often is issue occurring? : First occurrence Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: Database issue Patient Impact/Outcome: none User Impact: Database issue Current Software Version: 4.1 Diagnostic performed by Engineer : Logged into system and checked database. Orphaned studies found. Advised customer to perform a full database rebuild. Once complete the orphaned studies should populate in the local and be ready for deletion. I will check back with the customer within 24 hours. Internal Comments : Error # and/or description of error : No Error Completed or Recommended Repair Action "Internal" : Logged into system and checked database. Orphaned studies found. Advised customer to perform a full database rebuild. Once complete the orphaned studies should populate in the local and be ready for deletion. I will check back with the customer within 24 hours. Resolution : Database rebuild fixed issue. OK to close case. ----------------------------------------------------------------------------- | 3 |  | Logged into system and checked database. Orphaned studies found. Advised customer to perform a full database rebuild. Once complete the orphaned studies should populate in the local and be ready for deletion. I will check back with the customer within 24 hours. Resolution : Database rebuild fixed issue. OK to close case. | [12 empty] | 117794973 |  | 1056146 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | orphaned studies | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1110 |  |  | Japan | 独立行政法人国立病院機構 | 341015 | 4437600% |  | 01CS:KKI:問い合わせ:心臓解析の時間分解能に関して | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Application Support Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer : 01CS:KKI:問い合わせ:心臓解析の時間分解能に関して Safety Questions : Case: 0117799961 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: KAZUMASA KIMURA Last updated by: KAZUMASA KIMURA 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 対象装置:CT 故障個所:01Cosole 最初に発生したのはいつか?(~以内):1 Hour 問題発生の頻度は?:First Occurrence エンドユーザｰのワークフローのどの時点において問題が発生しているのか?:その他(Other) 最近どのようなイベントが発生したか?:N.A. 顧客は問題について具体的なシステムイベントレポートを作成しているか?:いいえ。お客様にログファイル、レポートは作成して頂けませんでした。 Troubleshooting Action: 対応後の結果:改善された 故障個所:01Cosole 診断中にどのようなステップを実施したか?: 【コール内容】 ①時間分解能が悪くなる心拍は？グラフがあったら送ってください。 ②他のサイトの心臓CT造影条件を教えてください。 【対応】 ①71bpm前後が悪くなります。大谷さんからお送りします。 ②調べられる範囲で調べ報告します。 Repair Action: 【現象確認】 確認済み 【作業内容詳細】 【コール内容】 ①時間分解能が悪くなる心拍は？グラフがあったら送ってください。 ②他のサイトの心臓CT造影条件を教えてください。 【対応】 ①71bpm前後が悪くなります。大谷さんからお送りします。 ②調べられる範囲で調べ報告します。 【確認内容】 良好 OneEMS internal remarks : --Troubleshooting Action-- 対応後の結果:改善された 故障個所:01Cosole 診断中にどのようなステップを実施したか?: 【コール内容】 ①時間分解能が悪くなる心拍は？グラフがあったら送ってください。 ②他のサイトの心臓CT造影条件を教えてください。 【対応】 ①71bpm前後が悪くなります。大谷さんからお送りします。 ②調べられる範囲で調べ報告します。 --Repair Action-- --Subject-- 01CS:KKI:ﾋｻｷ様:問い合わせ:心臓解析の時間分解能に関して --Problem Description-- 対象装置:CT 故障個所:01Cosole 最初に発生したのはいつか?(~以内):1 Hour 問題発生の頻度は?:First Occurrence エンドユーザｰのワークフローのどの時点において問題が発生しているのか?:その他(Other) 最近どのようなイベントが発生したか?:N.A. 顧客は問題について具体的なシステムイベントレポートを作成しているか?:いいえ。お客様にログファイル、レポートは作成して頂けませんでした。 --4/8Elements-- Q1 報告された問題は確認できましたか? 確認済み Q2 他の問題は確認できましたか? 無し Q3 問題を解決するために何をおこないましたか? 【コール内容】 ①時間分解能が悪くなる心拍は？グラフがあったら送ってください。 ②他のサイトの心臓CT造影条件を教えてください。 【対応】 ①71bpm前後が悪くなります。大谷さんからお送りします。 ②調べられる範囲で調べ報告します。 Q4 対応後の装置は順調ですか? 良好 Q5 その検査はどの様に終了しましたか? 無し Q6 問題により間違った所見（誤診、誤治療）の有無: 無し Q7 意図しない不要な曝射の有無: 無し Q8 アーチファクトの有無: 無し SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 44377.6813 | [14 empty] | 117799961 |  | 1060260 |  |  | Closed - Non-Complaint | [20 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1111 |  |  | China | Shandong Province Tai'an City T.C.M. No. 2 Hospital | 345037 | 4437600% |  | 02N/Ingenuity Core128/Elite China/Add printer | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer:add printer Safety Questions:Case: 0117799692 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Zixuan Last updated by: Wang Zixuan 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Add printer Troubleshooting Action: Add printer Repair Action: Add printer Test/Inspection: N/A Test Equipment: N/A SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | 44377.6507 | [14 empty] | 117799692 |  | 1060079 |  |  | Closed - Non-Complaint | [20 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1112 |  |  | United States | Lackey Memorial Hospital | 40015 | 4437900% |  | 0711CDN89 Prompted them to change the password and now she can't get into the exam card manager. Telling her she's not authorized. Customer requested clinical apps. | Work Order #: WO-03788299 Curry, Dyan The system fully meets specification and returned to use Problem description by Clinical Support Specialist : Customer Function/ Role: technologist Device Use: in use name Expected and Unexpected Results: ECM password expired, customer can not access User Impact: none Patient Impact: none Current Software version: 4.1.10 Diagnostic performed by Engineer : transferred call to Daniel M, RSE to page FSE with dongle. ----------------------------------------------------------------------------- Work Order #: WO-03788362 Morrow, Dan The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Customer Function/Role: Tech Device Use: No Patient involved Expected and Unexpected Results: Customer needs EC password reset User Impact: none Patient Impact: none Current Software Version: 4.1.10 Diagnostic performed by Engineer : Customer prompted to reset EC password. Now locked out of EC manager. Unable to reset remotely. Dispatching FSE. Follow Up Required : Log into service and reset customer EC manager password. ----------------------------------------------------------------------------- | 5 |  | Log into service and reset customer EC manager password. | [12 empty] | 117829310 |  | 1072635 | TRUE | Opened | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Password | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1113 |  |  | Turkey | Republic of Turkey Ministry of Health, Bagcilar | 40010 | 4437500% |  | Sound is coming from the table. | Work Order #: WO-03753000 Salih Mendi The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Customer Function/Role: technician Device Use: out of use Expected and Unexpected Results: Noise during patient table vertical and vertical movement. User Impact: No user impact. Patient Impact: No patient impact. Current Software Version:NA Follow Up Required : Noise during vertical and horizontal movements was found. Therefore, the site should be visited to lubricate the table. System is partially operational. Diagnostic performed by Engineer : The user was remotely supported for the noise complaint. In diagnosis, noise during vertical and horizontal movements was found. Therefore, the site should be visited to lubricate the table. System is partially operational. ----------------------------------------------------------------------------- Work Order #: WO-03753317 29-Jun-2021 Ugur Altunoz The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: technician Device Use: out of use Expected and Unexpected Results: Noise during patient table vertical and vertical movement. User Impact: No user impact. Patient Impact: No patient impact. Current Software Version:NA Resolution : The noise problem during table vertical movement was solved after lubricating the side rails. The device is operational. ----------------------------------------------------------------------------- | 3 |  | The noise problem during table vertical movement was solved after lubricating the side rails. | [12 empty] | 117793879 |  | 1052062 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | E-125059 | [5 empty] | Engineering | [6 empty] | Couch Noise | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | Lubricant | Couch-HW-Lubricant | Couch-HW-Lubricant | E-125059 | Lubricant | Couch-HW | Couch-HW-Lubricant | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | y
1114 |  |  | Chile | Scotiabank Chile | 40008 | 4435000% |  | Couch noise when reaching the top of table, lubricated vertical guide piston |  |  |  | Couch noise when reaching the top of table, lubricated vertical guide piston | [14 empty] | 1076312 | TRUE | Investigation | Closed - Done |  | E-125059 | [5 empty] | Engineering | [6 empty] | Couch Noise | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | Lubricant | Couch-HW-Lubricant | Couch-HW-Lubricant | E-125059 | Lubricant | Couch-HW | Couch-HW-Lubricant | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | y
1115 |  |  | Japan | 独立行政法人国立病院機構 | 341015 | 4437500% |  | CT:APP:NO:ﾋｻｷ様:質問 | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Application Support Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer : Q:メタボ検査用プロトコルを作りたいのですが、どのプロトコルから作ればよいですか？ A:Referenceセットから体幹部コンべのプロトコルを選択いただき、編集して保存していただきました。 Safety Questions : Case: 0117792699 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Ootani Natsuki Last updated by: Ootani Natsuki 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 44376.7125 | [14 empty] | 117792699 |  | 1055327 |  |  | Closed - Non-Complaint | [20 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1116 |  |  | Germany | Gesundheit Nord gGmbH Klinikverbund | 40002 | 4437700% |  | After a biopsy, the system does not display a dose | Work Order #: WO-03771553 Schwidrogitz Michaela The system fully meets specification and returned to use Problem description by Clinical Support Specialist: The customer reports that after a puncture, no dose info page was created by the system. She created a bug report and phoned Engineering. Diagnostic performed by Engineer: Asked the customer to perform a restart. The dose info page was not displayed afterwards. It occurred for the first time and so far once. Customer will still contact Ms. Schirmer in the FU appointment next week. The customer will call again if it should occur reoccur. ----------------------------------------------------------------------------- | 3 |  | The customer reports that after a puncture, no dose info page was created by the system. Collect bugreport for engineering investigation. | [12 empty] | 117810739 |  | 1061794 | TRUE | Opened | Closed - No Further Investigation |  | TBD_Bugreport | [5 empty] | Engineering | [6 empty] | Biopsy dose | [5 empty] | Refresh |  |  | System issue | System | TBD-Bugreport | TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | TBD_Bugreport |  | TBD-Bugreport | TBD-Bugreport-TBD_Bugreport | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | y
1117 |  |  | United States | Leesburg Regional Medical Center | 40022 | 4437500% |  | 0711CDY73 FSE is aware over the weekend scanner locked out duplicate see case 0117788182 | Work Order #: WO-03753586 Atz, Michael The system partly meets specification and is returned to use with restrictions Internal Comments : Error # and/or description of error : none Completed or Recommended Repair Action "Internal" : FSE is aware over the weekend scanner locked out due to password duplicate see case 0117788182 -- Customer used admin PW as a temporary Problem Description by Engineer : Problem description by engineer : FSE is aware over the weekend scanner locked out due to password duplicate see case 0117788182 When was the first occurrence? Within... : 26-Jun-2021 How often is issue occurring? : Continuously Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: No further information is available as per the customer Patient Impact/Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer Current Software Version: Not provided Diagnostic performed by Engineer : FSE is aware over the weekend scanner locked out due to password duplicate see case 0117788182 Follow Up Required : Error # and/or description of error : none Completed or Recommended Repair Action "Internal" : FSE is aware over the weekend scanner locked out due to password duplicate see case 0117788182 -- Customer used admin PW as a temporary ----------------------------------------------------------------------------- Work Order #: WO-03753629 30-Jun-2021 Le David The system fully meets specification and returned to use Problem Description by Engineer : Problem description by engineer : FSE is aware over the weekend scanner locked out due to password duplicate see case 0117788182 When was the first occurrence? Within... : 26-Jun-2021 How often is issue occurring? : Continuously Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: No further information is available as per the customer Patient Impact/Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer Current Software Version: Not provided Resolution : Instructed customer to re log into system. System only times out for 5 mins after incorrect attempts have been reached. Notified customer to setup Emergency Account as alternate login. ----------------------------------------------------------------------------- | 44375.8868 | [14 empty] | 117794397 |  | 1052169 |  |  | Closed - Duplicate | [20 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1118 |  |  | United States | AdventHealth Zephyrhills | 300099 | 4438000% |  | Taking a long time to send images to PACS | Work Order #: WO-03790386 Boyd Craig The system partly meets specification and is returned to use with restrictions Diagnostic performed by Engineer : On going issue. Study taking up to an hour to send over to PACS. PRS accessed. No normal errors seen. Customer requesting service today. PO WC0010001651531. Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: Taking a long time to send images to PACS User Impact: None Patient Impact: None Current Software Version: Follow Up Required : Customer requesting service today. PO WC0010001651531. FSE on site to evaluate DICOM settings and check images going to PACS. Work with IT department at facility. Please request port 80 be opened for remote access. ----------------------------------------------------------------------------- Work Order #: WO-03790399 03-Jul-2021 Garcia Marrero, Jorge The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: Taking a long time to send images to PACS User Impact: None Patient Impact: None Current Software Version: Diagnostic performed by Engineer : Proactively ordering Host computer for networks studies transfer slow. Follow Up Required : Installation of host computer. ----------------------------------------------------------------------------- Work Order #: WO-03790759 03-Jul-2021 Garcia Marrero, Jorge Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: Taking a long time to send images to PACS User Impact: None Patient Impact: None Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 3 | [14 empty] | 117831044 |  | 1075189 | #N/A | #N/A | Opened | [7 empty] | Customer | [6 empty] | Network Switch | [5 empty] | Non-Refresh |  |  | Upgrade issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | #REF!
1119 |  |  | United States | AdventHealth Zephyrhills | 300099 | 4437900% |  | Gantry won't initialize | Work Order #: WO-03788701 Powell, Jamie The system does not meet specification and is out of use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : Gantry won't initialize.. Customer cycled power multiple times but still having issues. Follow Up Required : Gantry won't initialize.. Customer cycled power multiple times but still having issues. Internal Comments : Spoke to FSE Mark Sher who was already onsite.. States he blue button the gantry and the system is up and scanning. Mark will look at bugrep later. ----------------------------------------------------------------------------- Work Order #: WO-03788737 02-Jul-2021 Sher Mark The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : system had several issues Blue button the gantry to clear initiation error, customr allowed contrast to get in gantry csusing image artifacts.. rings,, troubleshot pacs transfer issue ----------------------------------------------------------------------------- Work Order #: WO-03789543 03-Jul-2021 Garcia Marrero, Jorge The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : Relieve Mark Sher to continue network troubleshooting. The network cable from the hospital network port to the CT host computer was disconnected and used my laptop as a PACS. Transfer of studies from the CT to my mock PACS went well with normal transfer rate. After 8:00 PM the hospital network normalize and was able to transfer studies to any DICOM node inside the hospital network without issues and release the scanner to clinical usage.. There's an issue in the hospital network that unable the scanner to send studies to other DICOM nodes. ----------------------------------------------------------------------------- | 2 | [14 empty] | 117829654 |  | 1072828 | #N/A | #N/A | Opened |  | CustomerNetIssue | [5 empty] | Engineering | [6 empty] | Multi issue | [5 empty] | Refresh |  |  | Upgrade issue | System | NotEngineering | ShouldbeCustomer | NotEngineering-ShouldbeCustomer | NotEngineering-ShouldbeCustomer | CustomerNetIssue |  |  |  | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | y | y | y | y | y | 0 | y | #REF!
1120 |  |  | United States | AdventHealth Zephyrhills | 300106 | 4437700% |  | 0308CDN8A Couch making noise when moving vertically | Work Order #: WO-03773705 Powell, Jamie The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : Couch making noise when moving vertically Follow Up Required : Couch making noise when moving vertically. Grease vertical drive screw and sliding rails ----------------------------------------------------------------------------- Work Order #: WO-03773715 01-Jul-2021 Sher Mark The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : Greased guide rails and tested system ----------------------------------------------------------------------------- | 3 | [14 empty] | 117815027 |  | 1063029 | #N/A | #N/A | Evaluation | [7 empty] | Others | [6 empty] | lubrecate | [5 empty] | Non-Refresh |  |  | Upgrade issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | #REF!
1121 |  |  | United States | Select Specialty Hospitals | 300037 | 4437600% |  | Parts Quote | [16 empty] | 117802415 | 1061640 | 1061640 | #N/A | #N/A | Closed - Non-Complaint | [23 empty] | Upgrade issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1122 |  |  | United States | Select Specialty Hospitals | 300037 | 4437100% |  | Cracked compensator message popped out ; Replaced the A-plane collimator and performed calibration. | Work Order #: WO-03738673 Atz, Michael The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Customer Function/Role: tech Device Use: diag Expected and Unexpected Results: N/A User Impact: N/A Patient Impact: N/A Current Software Version: N/A Diagnostic performed by Engineer : CRACKEDWEDGE BR ROHS COLLIMATOR FIELD SERVICE Follow Up Required : On-site recommendations according to the system's PM manual: Re-confirm the crack, running the Compensator Analyzer Tool. (See the Compensator Analyzer Tool Guide) If confirmed by the tool, replace the compensator/wedge cassette. If not confirmed by the tool, perform compensator visual inspection. (See PM manual, Appendix A: Visual Collimator Compensator Inspection). If visually confirmed, replace the compensator/wedge cassette. After replacement, perform IQ tests. ----------------------------------------------------------------------------- Work Order #: WO-03738802 29-Jun-2021 Beard Michael The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: tech Device Use: diag Expected and Unexpected Results: N/A User Impact: N/A Patient Impact: N/A Current Software Version: N/A Resolution : Arrived onsite. replaced A plane collimator following instructions in DMR 4598 000 72481 Rev X. Performed calibrations following instructions in DMR 4598 015 87311 Rev E. Tested and verified system is performing correctly. Returned back over to customer for clinical use. ----------------------------------------------------------------------------- | 3 | [14 empty] | 117779333 |  | 1042877 | #N/A | #N/A | Closed - Done | [7 empty] | Engineering | [6 empty] | A-Plane | [5 empty] | Non-Refresh |  |  | Upgrade issue | system | Gantry-HW | aplane | Gantry-HW-aplane | NoEnoughInfo | 0 |  | Gantry-HW | Gantry-HW-aplane | No | 0 | 0 | 0 | 0 | 0 | No | y | y | y | 0 | 0 | y | y | y |  |  | y
1123 |  |  | Spain | POLICLINICA NTRA SRA DEL ROSARIO SL | 336190 | 4436400% |  | Gantry control panel key is broken | Work Order #: WO-03672241 20-Jun-2021 Miguel Ribot Martínez The system fully meets specification and returned to use. Parts Consumed : 1.00000,453567003451,TENSIONER ASSY 1.00000,459800484501,CONTROL PANEL ASSEMBLY-FRONT LH 2.00000,453567055351,AIR FILTER 1.00000,453566492401,POLY-VEE BELT Diagnostic Performed by Engineer: During the last preventive maintenance, we saw that the key to enter the couch in the gantry on the left control panel is worn out. The air filters have to be replaced with new ones. A new belt has also been ordered for the gantry with its tensor. All these materials are sent to the Polyclinic and will be replaced at the next scheduled visit. Regulatory Questions : N/A Problem Description by Engineer: Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: ----------------------------------------------------------------------------- | [15 empty] | 117716200 | 1020312 | 1020312 | #N/A | #N/A | Closed - No Further Investigation | [20 empty] | Non-Refresh |  |  | Upgrade issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | #REF!
1124 |  |  | United States | Philips Medical Capital | 300057 | 4436300% |  | 0699CDY72 System Stuck On Initializing Cannot Scan | Work Order #: WO-03665769 Holeman, David The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Problem description by engineer : System Stuck On Initializing Cannot Scan When was the first occurrence? Within... : 15-Mar-2021 08:00:00 How often is issue occurring? : Intermittent Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: System Stuck On Initializing Cannot Scan Patient Impact/Outcome: None User Impact: System Stuck On Initializing Cannot Scan Current Software Version: 4.1.7.25026 Follow Up Required : Error # and/or description of error : S_HMC_CTRL_A_MOTOR_SHORT_CIRCUIT_OR_OVERTEMP NORMAL (P1 = 26) S_DMC_SHOT_COMM_COLLAPSE NORMAL (P1 = 4098) Error #12 - CheckGantryConnection: ACQ board or Gantry-to-CIRS datapath failure. Unexpected gantry Confirm Resend: 1 (1 = error) Reading resends: 147 CRC error count Chan A: 0 Chan B: 147 Chan C: 0 Chan D: 147 Loss of Signal Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 Line Fault Indicator Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 On-board overflow retry count: 41 Error #12 - : ACQ board was inactive for at least 328 ms Completed or Recommended Repair Action "Internal" : For S_HMC_CTRL_A_MOTOR_SHORT_CIRCUIT_OR_OVERTEMP NORMAL (P1 = 26) 1. Check for an obstruction or mechanical load problem. 2. On a BedRock couch make sure the power to the AMC is between 14VAC and 62VAC. 3. Check / replace the horizontal motor Note: On a BedRock table, It maybe normal to see error For S_DMC_SHOT_COMM_COLLAPSE NORMAL (P1 = 4098) Inspect data transfer between gantry and IRS/Perform Visual Diagnostics. Run ACQ board diagnostics. Connect the optical cable directly from the ACQ PCB to the RCOM. Try to make scans without rotation to see if scanning is possible without above error. Clean optical slipring. Inspect Rotor transmitter and Stator receivers, cabling/connections and incoming power to those modules (i.e. 12vdc from FRC). Repair/replace as necessary. Note: These couch & datapath errors seem to be intermittent, but occurring around the same time, thus one of these errors may be triggering the other error. Diagnostic performed by Engineer : CT Tech reports "System stuck on initializing Cannot scan" and RECON Box error. Tech states she had this error on 6/15 in AM also. Tech is currently power-cycling the Host/IRS/IMR and gantry. Inspected logger which indicates: S_HMC_CTRL_A_MOTOR_SHORT_CIRCUIT_OR_OVERTEMP NORMAL (P1 = 26) and S_DMC_SHOT_COMM_COLLAPSE NORMAL (P1 = 4098) before the power-cycle. Tech is able to scan successfully (no errors noted) post-power cycle. Tech has clinical scans to perform and requested RSE to call back at 9am. RSE called Tech back and Tech reports 2 clinical scans OK, but is concerned about the intermittent issues which she reports are occurring more frequently. RSE bug report analysis from 6/14/21 indicates the same couch and datapath issues around 1412. Internal Comments : Error # and/or description of error : S_HMC_CTRL_A_MOTOR_SHORT_CIRCUIT_OR_OVERTEMP NORMAL (P1 = 26) S_DMC_SHOT_COMM_COLLAPSE NORMAL (P1 = 4098) Error #12 - CheckGantryConnection: ACQ board or Gantry-to-CIRS datapath failure. Unexpected gantry Confirm Resend: 1 (1 = error) Reading resends: 147 CRC error count Chan A: 0 Chan B: 147 Chan C: 0 Chan D: 147 Loss of Signal Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 Line Fault Indicator Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 On-board overflow retry count: 41 Error #12 - : ACQ board was inactive for at least 328 ms Completed or Recommended Repair Action "Internal" : For S_HMC_CTRL_A_MOTOR_SHORT_CIRCUIT_OR_OVERTEMP NORMAL (P1 = 26) 1. Check for an obstruction or mechanical load problem. 2. On a BedRock couch make sure the power to the AMC is between 14VAC and 62VAC. 3. Check / replace the horizontal motor Note: On a BedRock table, It maybe normal to see error For S_DMC_SHOT_COMM_COLLAPSE NORMAL (P1 = 4098) Inspect data transfer between gantry and IRS/Perform Visual Diagnostics. Run ACQ board diagnostics. Connect the optical cable directly from the ACQ PCB to the RCOM. Try to make scans without rotation to see if scanning is possible without above error. Clean optical slipring. Inspect Rotor transmitter and Stator receivers, cabling/connections and incoming power to those modules (i.e. 12vdc from FRC). Repair/replace as necessary. Note: These couch & datapath errors seem to be intermittent, but occurring around the same time, thus one of these errors may be triggering the other error. ----------------------------------------------------------------------------- Work Order #: WO-03666488 Williams John Problem Description by Engineer : Problem description by engineer : System Stuck On Initializing Cannot Scan When was the first occurrence? Within... : 15-Mar-2021 08:00:00 How often is issue occurring? : Intermittent Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: System Stuck On Initializing Cannot Scan Patient Impact/Outcome: None User Impact: System Stuck On Initializing Cannot Scan Current Software Version: 4.1.7.25026 Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 4 | [14 empty] | 117710049 |  | 1011733 | #N/A | #N/A | Evaluation | [7 empty] | Others | [6 empty] | Optic cable | [5 empty] | Non-Refresh |  |  | Upgrade issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | #REF!
1125 |  |  | United States | Upson Regional Medical Center | 310115 | 4436300% |  | System error during scanning due to faulty signal brush block and tapeswitch; Replaced signal brush block and tapeswitch | Work Order #: WO-03665959 Miller, John M The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Problem description by engineer : Customer states they had CT System errors last night during scanning. When was the first occurrence? Within... : 15-Jun-2021 How often is issue occurring? : Intermittent Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Treatment Expected and Unexpected results: System errors Patient Impact/Outcome: Delay in Scan User Impact: Delay in Scan Current Software Version: 4.1.10.2508 Follow Up Required : Error # and/or description of error : TmStamp 19548 ERR_TYPE NORMAL ErrNum S_STARTER_FAULT_NOT_UP_TO_SPEED vxwErrNo 3997698 File d_sn/stargate2/gantry/source/rhost/rhostSupport/source/RhostHVPSHssFaultISR.cpp Line 285 Param1 64 Param2 0 Info EalTime 6/16/2021 4:22:26 AM TmStamp 116484 ERR_TYPE NORMAL ErrNum S_CTRL_LINK_GHOST_GOT_CRC_FRAMING_ERROR vxwErrNo 0 File y_bld_main_Bld_sn/stargate2/gantry/source/gaf/Control_Link/src/CntlLinkTask.cpp Line 370 Param1 0 Param2 0 Info EalTime 6/15/2021 5:31:40 PM TmStamp 21769 ERR_TYPE NORMAL ErrNum S_HMC_TAPE_SWITCH_PRESSED vxwErrNo 0 File /gantry/source/galaxy/CouchMgrVertMotion/Src/StoppedToInitVertMoveFailTrans.cpp Line 406 Param1 0 Param2 0 Info EalTime 6/14/2021 3:11:31 PM TmStamp 27631 ERR_TYPE NORMAL ErrNum S_HMC_VELOCITY_ENCODER_FAILURE vxwErrNo 0 File S:\sy_bld_main_Bld_sn/stargate2/gantry/source/galaxy/hmc/src/HMCSMachine.cpp Line 1558 Param1 0 Param2 0 Info DeltaAbs:68544 Rel:68778 Vel:94943 6877 EalTime 6/14/2021 4:49:17 PM Type : Error Date : 6/15/2021 Time : 9:27:11 AM Event : 0 Source : PmsCT Category : None User : N/A Computer : HOST-4CE9110D8L Description: Message :XML parsing: line 135, character 57, illegal xml characterStack Trace : at System.Data.SqlClient.SqlConnection.OnError(SqlException exception, Boolean breakConnection, Action`1 wrapCloseInAction) at System.Data.SqlClient.TdsParser.ThrowExceptionAndWarning(TdsParserStateObject stateObj, Boolean callerHasConnectionLock, Boolean asyncClose) at System.Data.SqlClient.TdsParser.TryRun(RunBehavior runBehavior, SqlCommand cmdHandler, SqlDataReader dataStream, BulkCopySimpleResultSet bulkCopyHandler, TdsParserStateObject stateObj, Boolean& dataReady) at System.Data.SqlClient.TdsParser.Run(RunBehavior runBehavior, SqlCommand cmdHandler, SqlDataReader dataStream, BulkCopySimpleResultSet bulkCopyHandler, TdsParserStateObject stateObj) at System.Data.SqlClient.SqlBulkCopy.RunParser(BulkCopySimpleResultSet bulkCopyHandler) at System.Data.SqlClient.SqlBulkCopy.CopyBatchesAsyncContinuedOnSuccess(BulkCopySimpleResultSet internalResults, String updateBulkCommandText, CancellationToken cts, TaskCompletionSource`1 source) at System.Data.SqlClient.SqlBulkCopy.CopyBatchesAsyncContinued(BulkCopySimpleResultSet internalResults, String updateBulkCommandText, CancellationToken cts, TaskCompletionSource`1 source) at System.Data.SqlClient.SqlBulkCopy.CopyBatchesAsync(BulkCopySimpleResultSet internalResults, String updateBulkCommandText, CancellationToken cts, TaskCompletionSource`1 source) at System.Data.SqlClient.SqlBulkCopy.WriteToServerInternalRestContinuedAsync(BulkCopySimpleResultSet internalResults, CancellationToken cts, TaskCompletionSource`1 source) at System.Data.SqlClient.SqlBulkCopy.WriteToServerInternalRestAsync(CancellationToken cts, TaskCompletionSource`1 source) at System.Data.SqlClient.SqlBulkCopy.WriteToServerInternalAsync(CancellationToken ctoken) at System.Data.SqlClient.SqlBulkCopy.WriteRowSourceToServerAsync(Int32 columnCount, CancellationToken ctoken) at System.Data.SqlClient.SqlBulkCopy.WriteToServer(DataTable table, DataRowState rowState) at Philips.PmsCT.Logger.LogServer.DBInterface.AddLogBuffer(DataTable dataTable, Int32 targetId) Completed or Recommended Repair Action "Internal" : Suggest checking tape switches for proper operation and replace as needed. Test tape switches and check on Couch Control Board leds to determine which tape switch is sticking. Suggest cleaning the signal brush block, slip ring and/or replace brushes. Signal Brush, over time, may loose strength and not hold the slip ring properly. If cleaning helps and then the issue comes back something is contaminating the slip ring. The drive belt on one occasion and misaligned brush block on another, causing a accumulation on the slip rings. Also if there is a Black build-up on the rings, this can come from Covid-19 cleaning process. Make sure the Customer power off the Gantry prior to clean and leave the power off until dry. Recommendation: If this error is seen together with "S_CTRL_LINK_GHOST_GOT_CRC_FRAMING_ERROR" or S_CTRL_LINK_DMC_GOT_CRC_FRAMING_" it is caused by missing AP0, AP90 signals on the GHost.due to the Control Link problems. Mostly cleaning of signal slipring and brushes solves this problem.. On a BB Check / Replace the BI-directional optical control links parts. Check DMC for 33 on 7 segment display. 88 means the DMC is rebooting. Check Ghost for 1 (Normal state), anything else ghost is rebooting. Replace Ghost. Replace DMC Velocity encoder, check for proper couch movement, obstructions and proper lubrication. System errors, possible software reload. Internal Comments : Error # and/or description of error : TmStamp 19548 ERR_TYPE NORMAL ErrNum S_STARTER_FAULT_NOT_UP_TO_SPEED vxwErrNo 3997698 File d_sn/stargate2/gantry/source/rhost/rhostSupport/source/RhostHVPSHssFaultISR.cpp Line 285 Param1 64 Param2 0 Info EalTime 6/16/2021 4:22:26 AM TmStamp 116484 ERR_TYPE NORMAL ErrNum S_CTRL_LINK_GHOST_GOT_CRC_FRAMING_ERROR vxwErrNo 0 File y_bld_main_Bld_sn/stargate2/gantry/source/gaf/Control_Link/src/CntlLinkTask.cpp Line 370 Param1 0 Param2 0 Info EalTime 6/15/2021 5:31:40 PM TmStamp 21769 ERR_TYPE NORMAL ErrNum S_HMC_TAPE_SWITCH_PRESSED vxwErrNo 0 File /gantry/source/galaxy/CouchMgrVertMotion/Src/StoppedToInitVertMoveFailTrans.cpp Line 406 Param1 0 Param2 0 Info EalTime 6/14/2021 3:11:31 PM TmStamp 27631 ERR_TYPE NORMAL ErrNum S_HMC_VELOCITY_ENCODER_FAILURE vxwErrNo 0 File S:\sy_bld_main_Bld_sn/stargate2/gantry/source/galaxy/hmc/src/HMCSMachine.cpp Line 1558 Param1 0 Param2 0 Info DeltaAbs:68544 Rel:68778 Vel:94943 6877 EalTime 6/14/2021 4:49:17 PM Type : Error Date : 6/15/2021 Time : 9:27:11 AM Event : 0 Source : PmsCT Category : None User : N/A Computer : HOST-4CE9110D8L Description: Message :XML parsing: line 135, character 57, illegal xml characterStack Trace : at System.Data.SqlClient.SqlConnection.OnError(SqlException exception, Boolean breakConnection, Action`1 wrapCloseInAction) at System.Data.SqlClient.TdsParser.ThrowExceptionAndWarning(TdsParserStateObject stateObj, Boolean callerHasConnectionLock, Boolean asyncClose) at System.Data.SqlClient.TdsParser.TryRun(RunBehavior runBehavior, SqlCommand cmdHandler, SqlDataReader dataStream, BulkCopySimpleResultSet bulkCopyHandler, TdsParserStateObject stateObj, Boolean& dataReady) at System.Data.SqlClient.TdsParser.Run(RunBehavior runBehavior, SqlCommand cmdHandler, SqlDataReader dataStream, BulkCopySimpleResultSet bulkCopyHandler, TdsParserStateObject stateObj) at System.Data.SqlClient.SqlBulkCopy.RunParser(BulkCopySimpleResultSet bulkCopyHandler) at System.Data.SqlClient.SqlBulkCopy.CopyBatchesAsyncContinuedOnSuccess(BulkCopySimpleResultSet internalResults, String updateBulkCommandText, CancellationToken cts, TaskCompletionSource`1 source) at System.Data.SqlClient.SqlBulkCopy.CopyBatchesAsyncContinued(BulkCopySimpleResultSet internalResults, String updateBulkCommandText, CancellationToken cts, TaskCompletionSource`1 source) at System.Data.SqlClient.SqlBulkCopy.CopyBatchesAsync(BulkCopySimpleResultSet internalResults, String updateBulkCommandText, CancellationToken cts, TaskCompletionSource`1 source) at System.Data.SqlClient.SqlBulkCopy.WriteToServerInternalRestContinuedAsync(BulkCopySimpleResultSet internalResults, CancellationToken cts, TaskCompletionSource`1 source) at System.Data.SqlClient.SqlBulkCopy.WriteToServerInternalRestAsync(CancellationToken cts, TaskCompletionSource`1 source) at System.Data.SqlClient.SqlBulkCopy.WriteToServerInternalAsync(CancellationToken ctoken) at System.Data.SqlClient.SqlBulkCopy.WriteRowSourceToServerAsync(Int32 columnCount, CancellationToken ctoken) at System.Data.SqlClient.SqlBulkCopy.WriteToServer(DataTable table, DataRowState rowState) at Philips.PmsCT.Logger.LogServer.DBInterface.AddLogBuffer(DataTable dataTable, Int32 targetId) Completed or Recommended Repair Action "Internal" : Suggest checking tape switches for proper operation and replace as needed. Test tape switches and check on Couch Control Board leds to determine which tape switch is sticking. Suggest cleaning the signal brush block, slip ring and/or replace brushes. Signal Brush, over time, may loose strength and not hold the slip ring properly. If cleaning helps and then the issue comes back something is contaminating the slip ring. The drive belt on one occasion and misaligned brush block on another, causing a accumulation on the slip rings. Also if there is a Black build-up on the rings, this can come from Covid-19 cleaning process. Make sure the Customer power off the Gantry prior to clean and leave the power off until dry. Recommendation: If this error is seen together with "S_CTRL_LINK_GHOST_GOT_CRC_FRAMING_ERROR" or S_CTRL_LINK_DMC_GOT_CRC_FRAMING_" it is caused by missing AP0, AP90 signals on the GHost.due to the Control Link problems. Mostly cleaning of signal slipring and brushes solves this problem.. On a BB Check / Replace the BI-directional optical control links parts. Check DMC for 33 on 7 segment display. 88 means the DMC is rebooting. Check Ghost for 1 (Normal state), anything else ghost is rebooting. Replace Ghost. Replace DMC Velocity encoder, check for proper couch movement, obstructions and proper lubrication. System errors, possible software reload. Diagnostic performed by Engineer : Collect information from customer and view log files. ----------------------------------------------------------------------------- Work Order #: WO-03666180 18-Jun-2021 Lee John The system fully meets specification and returned to use Problem Description by Engineer : Problem description by engineer : Customer states they had CT System errors last night during scanning. When was the first occurrence? Within... : 15-Jun-2021 How often is issue occurring? : Intermittent Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Treatment Expected and Unexpected results: System errors Patient Impact/Outcome: Delay in Scan User Impact: Delay in Scan Current Software Version: 4.1.10.2508 Resolution : Installed brush block and tape switch Parts Consumed : 1.00000,459800623252,FRU, CBL, HANDLE TAPE SWITCH 1.00000,459800325761,Brush, Signal Block Assembly ----------------------------------------------------------------------------- | 3 | [14 empty] | 117710321 |  | 1011535 | #N/A | #N/A | Closed - Done | [7 empty] | Others | [6 empty] | Signal brush | [5 empty] | Non-Refresh |  |  | Upgrade issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | #REF!
1126 |  |  | United States | AdventHealth Zephyrhills | 300106 | 4436200% |  | 0499CFXXX Short TC Failed | Work Order #: WO-03657944 Holeman, David The system fully meets specification and returned to use Diagnostic performed by Engineer : CT Tech reports short TC and Quick IQ Check scan failures. Tech states he has restarted Host and zeroed couch = same issue. Inspected logger which indicates multiple: S_MASTER_SEQUENCER_SUITE_DOOR_OPEN errors. Problem Description by Engineer : Problem description by engineer : Short TC - Scan Failure When was the first occurrence? Within... : 15-Jun-2021 07:00:00 How often is issue occurring? : Continuously Malfunction area : Others Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: Not provided Patient Impact/Outcome: None User Impact: Short TC - Scan Failure Current Software Version: 4.1.7.25030 Internal Comments : Error # and/or description of error : S_MASTER_SEQUENCER_SUITE_DOOR_OPEN Resolution : Advised Tech to inspect/exercise all gantry room doors and re-attempt scans. After inspected gantry room doors, Tech found one door ajar and secured the door. Tech is now able to scan successfully. ----------------------------------------------------------------------------- | 3 | [14 empty] | 117702890 |  | 1005779 | #N/A | #N/A | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Mis-use | [5 empty] | Non-Refresh |  |  | Upgrade issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | #REF!
1127 |  |  | United States | AdventHealth Zephyrhills | 300099 | 4435800% |  | Recurring Intermittent Issue Transferring To PACS | Work Order #: WO-03641902 Holeman, David The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Problem description by engineer : Recurring Intermittent Issue Transferring To PACS When was the first occurrence? Within... : 03-May-2021 08:00:00 How often is issue occurring? : Intermittent Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: Recurring Intermittent Issue Transferring To PACS Patient Impact/Outcome: None User Impact: Delayed Transfers To PACS Current Software Version: 4.1.7.25030 Diagnostic performed by Engineer : CT Tech reports slow, then stopped transfer of images to PACS. Tech states around 0633 on 6/11/21 at around 175 images transferred, the transfer just stopped for about 15 minutes, then resumed sending and finished transfer. Tech states this occurs about once a day since Windows 10 upgrade (May 2021). Tech states a work-around to send faster is to access the ED local directory from the Main CT and send to PACS from the Main CT (an transfer is OK). Inspected last Host restart = 6/11/21 at 0509. Inspected LAN info (for PACS): name= PACS InetAddress=10.189.15.12 PortNumber=104 AETitle=WCHWFM1 Guided Tech to login as olevel and perform spooler reset. Communication with PFSE as an FYI. To monitor site and call back customer early next week to determine if issue persists. Internal Comments : Error # and/or description of error : Recurring Intermittent Issue Transferring To PACS Diagnostic performed by Engineer : Called customer back but Russ, CT Tech who placed call is not available. Spoke with another CT Tech who is unaware of any PACS transfer issues. Left message for Russ. RSE to callback on 6/15/21. Diagnostic performed by Engineer : RSE spoke with Russ, CT Tech who reports no further PACS transfer issues since spooler reset. RSE to follow-up with Tech on 6/16/21. Resolution : Called and spoke with Russ, CT Tech. No further issues with study transfers to PACS post-spooler reset. Closing call. ----------------------------------------------------------------------------- | 4 | [14 empty] | 117688366 |  | 997571 | #N/A | #N/A | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS Transfer | [5 empty] | Refresh |  |  | Upgrade issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | y
1128 |  |  | United States | Christ Hospital | 300205 | 4435700% |  | 0711CFXXX Q : How to Block EKG in LanConfig. from going to PACS | Work Order #: WO-03637667 Carpenter, Kevin The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Philips CT Clin Apps Onsite [ Rhonda ] Device Use: Diag Expected and Unexpected Results: None User Impact: None Patient Impact: None Current Software Version: IDK Diagnostic performed by Engineer : Assisted CT Clin Apps Onsite [ Rhonda ] reconfig'ing LanConfig such that cardiac studies did not send EKC waveforms to PACS. Resolution : Assisted CT Clin Apps Onsite [ Rhonda ] reconfig'ing LanConfig such that cardiac studies did not send EKC waveforms to PACS. ----------------------------------------------------------------------------- | [15 empty] | 117684427 | 997378 | 997378 | #N/A | #N/A | Closed - Non-Complaint | [23 empty] | Upgrade issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1129 |  |  | United States | Westgate Hospital | 40001 | 4437900% |  | Password Locked Out | Work Order #: WO-03787839 Carpenter, Kevin Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 2 |  | Password Locked Out | [12 empty] | 117828844 |  | 1075243 | TRUE | Opened | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Password | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1130 |  |  | Chile | Scotiabank Chile | 40008 | 4438600% |  | CL1675 - EQUIPMENT NEEDS TO BE CONFIGURED FOR SEVERAL TASKS IN RIS PACS. ON AUGUST 21ST AT 2:00 P.M. LOCAL TIME | Work Order #: WO-03826461 Valenzuela Cariaga, Luciano The system fully meets specification and returned to use. Follow Up Required : What are the next Troubleshooting steps : The customer requests a quote for work during non-working hours (Saturday August 21) for migration of its complete RIS PACS system and local image network. Error # and/or description of error : The customer requests a quote for work during non-working hours (Saturday August 21) for migration of its complete RIS PACS system and local image network. Completed or Recommended Repair Action "Internal" : The customer requests a quote for work during non-working hours (Saturday August 21) for migration of its complete RIS PACS system and local image network. Internal Comments : Error # and/or description of error : The customer requests a quote for work during non-working hours (Saturday August 21) for migration of its complete RIS PACS system and local image network. Completed or Recommended Repair Action "Internal" : The customer requests a quote for work during non-working hours (Saturday August 21) for migration of its complete RIS PACS system and local image network. Diagnostic performed by Engineer: The customer requests a quote for work during non-working hours (Saturday August 21) for migration of its complete RIS PACS system and local image network. Problem Description by Engineer : Problem description by engineer : The customer requests a quote for work during non-working hours (Saturday August 21) for migration of its complete RIS PACS system and local image network. Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected results: Equipment with new RIS PACS configured Patient Impact/Outcome: Not affected. User Impact: Migration of RIS PACS Current Software Version: 4.1.10 ----------------------------------------------------------------------------- | 3 |  | complete RIS PACS system and local image network. | [12 empty] | 117864190 |  | 1096609 | TRUE | Investigation | Investigation | [7 empty] | Service | [6 empty] | RIS PACS Config | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1131 |  |  | China | JX Province Children Hosp. | 40041 | 4438900% |  | The sheave has moved,which made the sheave squeeze against the motor | The sheave has moved,which made the sheave squeeze against the motor |  |  | ECR-076490 Optimize torque for pulley bushing | [14 empty] | 1098246 | TRUE | Investigation | Investigation |  | ECR-076490 | [5 empty] | Engineering | [6 empty] | Sheave | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Sheave | Gantry-HW-Sheave | Gantry-HW-Sheave | ECR-076490 |  | Gantry-HW | Gantry-HW-Sheave | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | #N/A
1132 |  |  | China | 息烽县中医医院 | 345014 | 4438500% |  | 02N/Ingenuity Core128/Elite China/Rebuild list | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer : 02N/Ingenuity Core128/Elite China Safety Questions : Case: 0117851128 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingsong Last updated by: Lu Mingsong 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: Rebuild list Troubleshooting Action: Remotely instruct the customer to Rebuild list. Repair Action: Remotely instruct the customer to Rebuild list. Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | Troubleshooting Action: Remotely instruct the customer to Rebuild list. Repair Action: Remotely instruct the customer to Rebuild list. Test/Inspection: Functional Check PASSED. Device returned to full functionality. | [12 empty] | 117851128 |  | 1092621 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Application | [6 empty] | Q&A | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1133 |  |  | China | 江西省儿童医院 | 40041 | 4438400% | 27 | 02N/Ingenuity CT/Can not scan | T2 Activities : Problem Description: 客户问题概述： 不能扫描 客户问题详细描述或故障现象： 不能扫描，重建柜故障 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： S_ROTOR_SEEK_REF_TIMEOUT_ERROR S_MP_CTRL_DRIVER_SK_FAULT_ISQT_PN_CURRENT_OUT 找不到零位，电机过温 S_DMC_SHOT_COMM_COLLAPSE P1=4098 OVERRUN&CIRS OPTIC CHANNEL 远程诊断与测试结果： 1.检查机架旋转情况，检查刹车 2.确认重建柜情况，重启重建柜 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 1.检查机架旋转情况，检查刹车 2.确认重建柜情况，重启重建柜 Test Equipment: N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 07.07.2021 15:48:34 UTC Jiawei Li (86091272)"SUAT. Ready for Clinical Use local time" : "2021-07-07 22:00“ 服务内容：机架不能旋转，检查发现 服务结果： Internal Remarks : 07.07.2021 16:19:44 UTC Jiawei Li (86091272) $Q1:(Y) $Q2:(N) $Q3:(NA) $Q4:(NA) $Q5:(NA) $Q6:（Unable to scan，CRIS Error ) "RDI Accuracy:Yes" $Q7:(Unable to duplicate reported problem ，gantry cannot rotate) $Q8:( The squeeze against the motor, causing the motor to fail to operate ) $Q9:( Reposition the sheave , solved the problem.) $Q10:(Test and inspection data refer to PA. Device returned to full functionality.) $Q11:(Calibrated Tools: No tools used.) External Remarks : 07.07.2021 16:19:41 UTC SUP JCOUSER (86091272)服务内容：机架不能旋转，检查发现电机滑轮位置移动后与电机摩擦，导致电机无法正常运转。重新调整滑轮位置。 服务结果：维修后，设备恢复正常使用。 ----------------------------------------------------------------------------- | 2 |  | William/Alan investigate the issue is same as issue two years ago from Zhengjiang site, does torque added for motor related parts? Check if system produced before ECR cut in to OPS? Mengya TBIN to OPER. May 13, 2021 at 9:31 AM Partial Gantry SN5047是2020年5月27日完成quality 检验的， ECR-076490 Optimize torque for pulley bushing | [12 empty] | 117846204 |  | 1092188 | TRUE | Closed - Done | Closed - Done |  | ECR-076490 | [5 empty] | Engineering | [6 empty] | Sheave | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Sheave | Gantry-HW-Sheave | Gantry-HW-Sheave | ECR-076490 |  | Gantry-HW | Gantry-HW-Sheave | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | #N/A
1134 |  |  | China | Jiangsu Xuzhou Medical School Affiliated No. 3 Hospital | 40039 | 4438300% |  | 02G/Ingenuity CT 728326/yellow screen | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer : 02G/Ingenuity CT 728326 T2 Activities:Problem Description: Summary of customer's problems: Other types of error. Detailed description of Customer's problems or malfunctions: Yellow screen. It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: Yellow screen. Remote troubleshoot and test result: Software bug. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Update the patch Safety Questions:Case: 0117838123 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action: confirmed with FSE Zhang Chuanxi, no need to add operation to follow up ----------------------------------------------------------------------------- | 3 |  | Yellow Screen | [12 empty] | 117838123 |  | 1083128 | TRUE | Investigation | Investigation |  | CTUDT00632520 | [5 empty] | Engineering | [6 empty] | Yellow Screen | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1135 |  |  | China | Jiangbei location, Southeast University Affiliated Zhongda Hospital | 345021 | 4438300% |  | Wechat:02N/Ingenuity Core128/Elite China/image transfer failed | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: Instructed Customer Problem Reported by Customer : Wechat:02N/Ingenuity Core128/Elite China/image transfer failed Safety Questions:Case: 0117837355 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created by: Wennian Han Last updated by: Wennian Han 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: image transfer failed Troubleshooting Action: image transfer failed Repair Action: guide customer to check and found HOST net cable failed, customer replaced a new net cable, then can transfer Test/Inspection: Functional Check PASSED. The device returned to full functionality. Test Equipment: NA SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | Troubleshooting Action: image transfer failed Repair Action: guide customer to check and found HOST net cable failed, customer replaced a new net cable, then can transfer | [12 empty] | 117837355 |  | 1082495 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Net Cable | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1136 |  |  | China | Lindian County Hospital | 345038 | 4438200% |  | 02N/Core128/rotation noise | BU Q&R Feedback: No Reason: Corrective Maintenance T2 Activities:Problem Description: Summary of customer's problems: NOISE IN ROTATION Detailed description of Customer's problems or malfunctions: NOISE IN ROTATION It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: High-speed rotation noise. Remote troubleshoot and test result: Opened the cover to check the source of the noise. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Opened the cover to check the source of the noise. Test Equipment: N/A SAP Transfer:Case Transferred to SAP Next Step Action: NTD: Service order opened through Ourmail.cn. Details: Contacted the field service engineer. No need to transfer. 13199082229 Doctor Teacher Li Lindian County Hospital 89753207 There is noise inside when rotating at high speed. Safety problem: Y N Safety Questions:Case: 0117835046 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Cao Last updated by: Liu Cao 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Problem Reported by Customer: 02N/Core128/ T2 Activities:Problem Description: The call didn't go through. SAP Transfer:Dispatched for On site Support External Remarks: 07.07.2021 14:45:07 UTC SUP JCOUSER (86038210) service details: gantry noise. Replace the pulley and adjust the rotation part cable position. service result: device returns to normal working condition after service Internal Remarks : 07.07.2021 14:57:15 UTC Haijun Du (86038210) RDI Accuracy: no $Q7:(Confirmed reported problem. gantry noise.) $Q8: (the CT gantry have noise when do CT rotation by 0.5s, TENSIONER ASSY failed) $Q9: ( ok after replacement the TENSIONER ASSY and adjust rotor cable position.) $Q10: (Test and inspection data refer to: PA. Device returned to full functionality) $Q11 Calibrated Tools name： refer to PA ----------------------------------------------------------------------------- | 3 |  | $Q7: (Confirmed reported problem，gantry noise) $Q8: (the CT gantry have noise when do CT rotation by 0.5s，TENSIONER ASSY failed) $Q9: ( ok after replacement the TENSIONER ASSY and adjust rotor cable position.) | [12 empty] | 117835046 |  | 1079668 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | ECR-077012 | [5 empty] | Engineering | [6 empty] | Tensioner | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Belt | Gantry-HW-Belt | Gantry-HW-Belt | ECR-077012 | Belt | Gantry-HW | Gantry-HW-Belt | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1137 |  |  | China | 青海省海北州第二医院 | 345029 | 4438200% |  | Wechat:02N/Ingenuity Core128/Elite China/check device | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : Wechat:02N/Ingenuity Core128/Elite China Safety Questions : Case: 0117833759 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wennian Han Last updated by: Wennian Han 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 其他类型故障。 客户问题详细描述或故障现象： 检查设备 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： NA 远程诊断与测试结果： 客户告知：CT机房漏水，主机架进水，现已烘干，请工程师前来检测开机 PRS TR状态： 不通，请工程师到场检查。 Repair Action: 建议现场维修步骤： 检查设备，确认有无其他问题 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 06.07.2021 09:41:50 UTC Yonghua Xu (86001334)科室：放射科CT室 联系人：郭荣寿 联系电话：15309708576 故障现象：CT机房漏水，主机架进水，现已烘干，需要工程师前来检测开机 。 解决方案/维修措施：CT机房漏水，主机架进水，现已烘干，需要工程师前来检测开机。检测设备机架、床 、Teal柜还有UPS干燥正常，通电测试设备预热扫描正常。 完成相应校准，完成备份。 服务后设备正常工作。 Internal Remarks : 06.07.2021 09:41:53 UTC Yonghua Xu (86001334) “RDI Accuracy: No” $Q7:Confirmation reported problem,The CT scanning room is leaking, and the gantry is flooded. It has been dried, and an engineer is required to check and start the machine. $Q8:The CT scanning room is leaking, and the gantry is flooded. It has been dried, and an engineer is required to check and start the machine. Check that the equipment gantry, couch, TEAL and UPS are dry properly. $Q9:Check that the equipment gantry, couch, TEAL and UPS are dry properly.Power on test equipment tube warming and scan normal. $Q10:Test and Inspection data refer to PA , Do a few patient, Observation of a certain period of time. The system is ok. Device returned to full functionality. $Calibrated Tools name (NA) S/N (NA) Next Calibration date (NA) ----------------------------------------------------------------------------- | 3 |  | CT机房漏水，主机架进水，现已烘干，需要工程师前来检测开机。检测设备机架、床 、Teal柜还有UPS干燥正常，通电测试设备预热扫描正常。 | [12 empty] | 117833759 |  | 1078731 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1138 |  |  | China | Guangxi Province Teng County People’s Hospital | 40036 | 4438200% |  | application/Ingenuity CT 728326 | BU Q&R Feedback: No Reason: Application Support Problem Reported by Customer: Secondary training of high pressure injector Safety Questions:Case: 0117832536 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Zhang Jing Last updated by: Zhang Jing 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 5 |  | Application training | [12 empty] | 117832536 |  | 1077907 | TRUE | Evaluation | Evaluation | [7 empty] | Application | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1139 |  |  | Japan | 独立行政法人国立病院機構 | 341013 | 4438600% |  | 01DM: RS: Scan did not proceed in a frozen state after Prep with CTA >> Manual start responded so inspection was possible >> No noticeable hard error | Problem Reported by Customer: 01CS: RS: Scan did not proceed in a frozen state after Prep with CTA >> Manual start responded so inspection was possible >> OneEMS internal remarks: --Troubleshooting Action-- Post-response results: Improved Failure location: 01Cosole What steps were taken during the diagnosis ?: In the subsequent inspection, there was no operation delay and the operation was good. No noticeable error was found in the logs. Judged as application operation delay. FSE compatible. --Repair Action-- --Subject-- 01DM: RS: Mr. Mamamoto: Scan did not proceed in a frozen state after Prep with CTA >> Manual start responded so inspection was possible >> No noticeable hard error --Problem Description-- Target device: CT Failure location: 01Cosole When did it first occur? (Within ~): 1 Hour How often do you have problems ?: Multiple times this Month At what point in the end-user workflow is the problem occurring ?: During inspection / measurement (During Acqustion / Scan) What happened recently ?: N.A. Does the customer create a specific system event report for the problem ?: N.A. SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Support Safety Questions: Case: 0117856986 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: RUI SATAKE Last updated by: RUI SATAKE 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm / alert capability, did it alarm / alert as it should have at the time the issue was discovered? N / A 4. Was this an out of box failure? N / A T2 Activities: Problem Description: Target device: CT Failure location: 01Cosole When did it first occur? (Within ~): 1 Hour How often do you have problems ?: Multiple times this Month At what point in the end-user workflow is the problem occurring ?: During inspection / measurement (During Acqustion / Scan) What happened recently ?: N.A. Does the customer create a specific system event report for the problem ?: N.A. Troubleshooting Action: Result after response: Improved Failure location: 01Cosole What steps were taken during the diagnosis ?: In the subsequent inspection, there was no operation delay and the operation was good. No noticeable error was found in the logs. Judged as application operation delay. FSE compatible. Repair Action: Service representative visit work. -------------------------------------------------- --------------------------- [Replacement parts] No replacement parts [Misorder Parts Report] No mistakes [Waiting information] No wait-and-see [TS Requests] No return parts [CT additional information] Q1 How did the inspection end ?: Not applicable because of Q1 No Q2 Presence or absence of incorrect findings (misdiagnosis, mistreatment) due to problems: None Q3 Presence or absence of artifacts: None Q4 Presence or absence of unintended unnecessary exposure: None [Repair certificate] Repair certificate is pasted [Other information sharing matters] 1. Treatment for patients infected with or suspected of being infected with the new coronavirus / Is it quarantined? Okay 2. Do I need to enter the treatment room or bed where the infected patient is? (Workplace and transfer (Including dynamic path) No Did you agree on how to complete the report with the person in charge of the equipment? YES Normal report completed on the day 09-Jul-2021 81100269 FSE Handover Tool1.3.4.32612.DB13 20210115 16.07.2021 01:54:08 UTC YUKO YAMAMOTO (81100269) [Request details] The left screen display of the console freezes during Tracker scanning (button operation is possible) [Phenomenon confirmation] I couldn't confirm. [Details of work] Since the manual start button was functioning normally, It is probable that a situation similar to a PC freeze in the screen display area had occurred. As a response, we carried out PC-related work such as deleting unnecessary files on the HOST PC. [Other failures / treatments] None [Device status] Equipment can be used [Work status] done | 4 |  | Since the manual start button was functioning normally, It is probable that a situation similar to a PC freeze in the screen display area had occurred. As a response, we carried out PC-related work such as deleting unnecessary files on the HOST PC. | [12 empty] | 117856986 |  | 1096904 | TRUE | Closed - Done | Closed - Done | [7 empty] | Others | [6 empty] | Disk clean | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1140 |  |  | Poland | Rzadowa Agencja | 341008 | 4438400% |  | commissioning and implementation into operation | Resolution Description: Switching on the power supply in the modular computed tomography laboratory. System calibration, Constancy and QuickIQ image quality test in accordance with the manufacturer's specifications. Due to the lack of a configured computer network, the DICOM settings for external systems have not been entered. The apparatus is technically ready for use. | 4 |  | System acceptance test | [12 empty] | 117846595 |  | 1089244 | TRUE | Evaluation | Evaluation | [7 empty] | Service | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1141 |  |  | Turkey | T.C Saglik Bakanligi Bagcilar | 40010 | 4438500% |  | Too much noise coming from the device | "Work Order #: WO-03811957 Soner Saglam The system partially meets specification and is returned to use with restrictions Resolution : Sound was estimated as DMS fan failure, part order was entered. Problem Description by Engineer : It has been reported that the gantry is very loud. Follow Up Required : Sound was estimated as DMS fan failure, parts order entered. | 3 |  | Sound was estimated as DMS fan failure, parts order entered. | [12 empty] | 117851395 |  | 1088133 | TRUE | Evaluation | Evaluation |  | CTUDT00632110 | [5 empty] | Engineering | [6 empty] | DMS fan | [5 empty] | Refresh |  |  | System issue | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | DMS Module | DMS-HW | DMS-HW-DMS Module | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1142 |  |  | China | Jilin Province Zhaonan City TCM Hospital | 345006 | 4438300% |  | 02n/Ingenuity Core128/Elite China/need parts | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 02n/Ingenuity Core128/Elite China Safety Questions:Case: 0117839325 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Lili Last updated by: Lu Lili 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action: FSE Hua Tian 018946514633 call in nemo SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | Gao, Mia |  | 2021-07-06 | [12 empty] | 117839325 |  | 1083964 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | 2021-07-07 17:31 | [6 empty] | 4438300% | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1143 |  |  | Turkey | Republic of Turkey Ministry of Health, Bagcilar | 40010 | 4438300% |  | An error message was displayed, device is up. | Work Order #: WO-03797735 Soner Saglam The system partly meets specification and is returned to use with restrictions Resolution : Log was analyzed. S_MPLTRANSPORTREC_DEFRAG_TOUCAN_ERROR error was found. Slipring and signal brush cleaning is recommended. Problem Description by Engineer : A gantry error was displayed. Follow Up Required : Log was analyzed. S_MPLTRANSPORTREC_DEFRAG_TOUCAN_ERROR error was found. Slipring and signal brush cleaning is recommended. ----------------------------------------------------------------------------- | 3 |  | Log has been examined. S_MPLTRANSPORTREC_DEFRAG_TOUCAN_ERROR is seen. Slipring and signal brush cleaning and control is recommended. | [12 empty] | 117838741 |  | 1081861 | TRUE | Evaluation | Evaluation |  | NotEngineering-SlipringClean | [5 empty] | Engineering | [6 empty] | Slip ring clean | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | SlipRingClean | Gantry-HW-SlipRingClean | Gantry-HW-SlipRingClean | NotEngineering-SlipringClean | SlipRingClean | Gantry-HW | Gantry-HW-SlipRingClean | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1144 |  |  | United Kingdom | STOCKPORT NHS FOUNDATION TRUST | 40016 | 4438600% |  | Interventional monitor and table upgrade | Work Order #: WO-03822493 09-Jul-2021 Martin Greenbank The system fully meets specification and returned to use Regulatory Questions : N/A Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : INST - Installation started but not yet completed Installed new extended couch and joystick option. Need to plan another day to complete the ceiling mounted monitor. System tested ok Follow Up Required : replan 2 men for 6h to complete ceiling mounted monitor when parts available ----------------------------------------------------------------------------- Work Order #: WO-03822496 09-Jul-2021 Warren Thorpe The system fully meets specification and returned to use Regulatory Questions : N/A Problem Description by Engineer : CT table needed upgraded from a standard couch to an extended couch. Biopsy monitor and extras needed to be installed. Resolution : New extended table installed successfully with additional biopsy features. Ceiling mounted monitor arm still needs to be fitted but all cables ran and ready to attach. Follow Up Required : Install ceiling mounted monitor arm. Please plan Martin G and myself to install. ----------------------------------------------------------------------------- Product/Part - MONITOR SUPPORT Part number - 453566207542 | 5 |  | New extended table installed successfully with additional biopsy features. Ceiling mounted monitor arm still needs to be fitted but all cables ran and ready to attach. Follow Up Required : Install ceiling mounted monitor arm. Please plan Martin G and myself to install. | [12 empty] | 117861311 |  | 1095469 | TRUE | Evaluation | Evaluation | [7 empty] | Service | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1145 |  |  | United States | Unity Health Systems | 40037 | 4438800% |  | 0609CDN84 Clicking noise right side gantry when rotor moving | Work Order #: WO-03828644 Maier, Donnie The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: Diagnostic Expected and Unexpected Results: Clicking noise from right side of gantry User Impact: No Patient Impact: No Current Software Version: 4.1.10.25210 Diagnostic performed by Engineer : Talked cu thru a complete power cycle of the system as it had been over 10 days for the gantry. After power cycle had cu run a mock scan and he can still hear the noise. No normal errors in ULV. Cu is just going to use the other CT Scanner and wants HTM 7/12 AM Dispatch. I provided case number to the CT Tech. Follow Up Required : Identify the source of the noise and correct. Tie wrap, etc. ----------------------------------------------------------------------------- | 3 |  | Clicking noise from right side of gantry FSE to be follow up Follow Up Required : Identify the source of the noise and correct. Tie wrap, etc. | [12 empty] | 117865876 |  | 1098098 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-GantryNoise | [5 empty] | Engineering | [6 empty] | Gantry Noise | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Noise | Gantry-HW-Noise | Gantry-HW-Noise | TBD-GantryNoise | Noise | Gantry-HW | Gantry-HW-Noise | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1146 |  |  | United States | Odessa Regional Medical Center | 40029 | 4438300% |  | Acquiring Dose Info | Work Order #: WO-03799935 Mujica, David The system fully meets specification and returned to use Problem description by Clinical Support Specialist : Customer Function/Role: Technologist Device Use: In use Expected and Unexpected Results: Get dose info for Peds User Impact: none Patient Impact: none Current Software Version: 4.1 Diagnostic performed by Engineer : Maya wanted to report all Pediatric dose information from the last year off of their CT system. It was explained to Maya that unless that patient information is on the directory then that dose information can be acquired from the CT system. Maya indicated that they delete CT exams frequently and this would not be a method to find the dose information for all Pediatrics from the last year. It was suggested to look in their PACS system as it would have the dose information stored. Maya took a look into PACS and successfully found the dose information. Maya had no further questions or need for clinical support. No rescan, no loss of data, no reinjection. ----------------------------------------------------------------------------- | E, Sravanthi |  | 2021-07-06 | [12 empty] | 117840721 |  | 1080377 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | 2021-07-06 21:57 | [6 empty] | 4438300% | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1147 |  |  | United States | Odessa Regional Medical Center | 40029 | 4438600% |  | Curved MPR | Work Order #: WO-03825534 Curry, Dyan The system fully meets specification and returned to use Problem description by Clinical Support Specialist : Customer Function/ Role: technologist Device Use: in use name Expected and Unexpected Results: curved MPRs of cspine User Impact: none Patient Impact: none Current Software version: 4.1.7 Diagnostic performed by Engineer : helped customer created MPR curved images in planar , CT Viewer. Saved batches of each plane. ----------------------------------------------------------------------------- | 5 |  | helped customer created MPR curved images in planar , CT Viewer. | [12 empty] | 117863525 |  | 1095973 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Application | [6 empty] | Q&A | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1148 |  |  | United States | McLaren Macomb | 40021 | 4438400% |  | customer requesting support | Work Order #: WO-03810117 Hirchak, Michele The system fully meets specification and returned to use Problem description by Clinical Support Specialist : Customer Function/Role: TECH Device Use: In Use Expected and Unexpected Results: XR-29 certificate. User Impact: None Patient Impact: None Current Software Version: 4.1.7 Diagnostic performed by Engineer : Customer requesting XR-29 Certificate. Emailed document and uploaded to TEAMS. MBH Resolution : Called Allison to make sure she rec'd the certificate. OK to close. MBH ----------------------------------------------------------------------------- | 5 |  | XR-29 certificate | [12 empty] | 117849487 |  | 1087150 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Customer | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1149 |  |  | United States | Philips Medical Capital | 40028 | 4438400% |  | 4.1 Accession number issues | Work Order #: WO-03809653 Mujica, David The system fully meets specification and returned to use Problem description by Clinical Support Specialist : Customer Function/Role: Technologist Device Use: In use Expected and Unexpected Results: Accession numbers not associating User Impact: none Patient Impact: none Current Software Version: 4.1 Diagnostic performed by Engineer : Iris indicated that the accession numbers were not associating to the CT exams performed. The exam cards were explored that were used and after investigation it was noticed that the user was picking all the exams and then launching each protocol which would not associate the proper accession numbers to each exam ordered. Iris was told to only launch the CT studies ordered per scan. Iris will comply and do this going forward. Iris had no further questions or need for clinical support at this time. No rescan, no loss of data, no reinjection per Iris. ----------------------------------------------------------------------------- | 5 |  | user was picking all the exams and then launching each protocol which would not associate the proper accession numbers to each exam ordered. Iris was told to only launch the CT studies ordered per scan. Iris will comply and do this going forward. | [12 empty] | 117849118 |  | 1086689 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Application | [6 empty] | Mis-use | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1150 |  |  | Chile | Scotiabank Chile | 40008 | 4438900% |  | CL1675-BURNING SMELL | Integration:Confirmation of order 53875179 saved ----------------------------------------------------------------------------- Work Order #: WO-03834204 González, Carlos The system meets the specification for the performed service and is returned to use. Problem Description by Engineer : Problem description by engineer : Client says that there was a burning smell in the Tomography room. When was the first occurrence? Within... : 12-Jul-2021 12:00:00 How often is issue occurring? : First occurrence Malfunction area : Others Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected results: No further information is available as per the customer Patient Impact/Outcome: No further information is available as per the customer. User Impact: No further information is available as per the customer Current Software Version: 4 Diagnostic Performed by Engineer: User reported a burning smell in Tomography room. Field service engineer Juan Osorio from Chile went previously to check the equipment and no damage was found inside. A preventive maintenance was even done on 08/07 and no damage was found either. Possibly the smell comes from the air conditioning system. Equipment operational. Internal Comments : Error # and/or description of error : NA Repair Action: NA Resolution: Equipment is OK. ----------------------------------------------------------------------------- | 4 |  | User reported a burning smell in Tomography room. Field service engineer Juan Osorio from Chile went previously to check the equipment and no damage was found inside. | [12 empty] | 117871717 |  | 1100496 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | False alarm | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1151 |  |  | China | 北京市中西医结合医院 | 40032 | 4439500% |  | 02N/Ingenuity CT 728326/connect PACS | BU Q&R Feedback: No SAP Customer Symptom Code: CENV Customer Symptom Code: Customer Environmental Concern Reason: Corrective Maintenance Engineer Failure Code: Environmental Cause Repair Activity Code: Configuration Change Problem Reported by Customer : connect PACS Safety Questions : Case: 0117907046 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Weng Jinming Last updated by: Weng Jinming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Troubleshooting Action: next Thursday ，connect PACS Repair Action: next Thursday ，connect PACS Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: N/A ----------------------------------------------------------------------------- | 5 |  | PACS Config | [12 empty] | 117907046 |  | 1125199 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Customer | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1152 |  |  | China | 浙江省乐清第一人民医院 | 40054 | 4439500% |  | 02N/Ingenuity CT 728326/Cirs fault | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity CT 728326 Safety Questions : Case: 0117906849 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Zhou Xiaohong Last updated by: Zhou Xiaohong 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 主机检查不能连接重建计算机，重建柜需要连接USB键盘，点回车 设备本身技术问题： 是。 Troubleshooting Action: 远程诊断与测试结果： 怀疑重建计算机软件问题 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议工程师到场检查，重新安装重建计算机软件 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : 已派工黄逸 15958725282 ----------------------------------------------------------------------------- | 3 |  | $Q7: (Confirmed reported problem&IRS disconnected) $Q8: (check errorlog found teermal sensor not programmed, all fans works normally) $Q9: (update BIOS version) | [12 empty] | 117906849 |  | 1125038 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | ShouldbeService | [5 empty] | Engineering | [6 empty] | CIRS | [5 empty] | Refresh |  |  | System issue | System | NotEngineering | ShouldbeService | NotEngineering-ShouldbeService | NotEngineering-ShouldbeService | ShouldbeService | ShouldbeService | NotEngineering | NotEngineering-ShouldbeService | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1153 |  |  | China | 江苏大学附属医院 | 345023 | 4439300% |  | Wechat:02N/Ingenuity Core128/Elite China/can not scan | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer : can not scan Safety Questions : Case: 0117895563 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingsong Last updated by: Lu Mingsong 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: can not scan Troubleshooting Action: Remotely instruct the customer to restart the system. Repair Action: Remotely instruct the customer to restart the system. Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | Troubleshooting Action: Remotely instruct the customer to restart the system. Repair Action: Remotely instruct the customer to restart the system. Test/Inspection: Functional Check PASSED. Device returned to full functionality. | [12 empty] | 117895563 |  | 1122766 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Reboot | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1154 |  |  | China | 浙江省乐清第一人民医院 | 40054 | 4439300% |  | 02N/Ingenuity CT 728326/account is lock | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: Instructed Customer Problem Reported by Customer : 02N/Ingenuity CT 728326 Safety Questions : Case: 0117893670 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wennian Han Last updated by: Wennian Han 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: account is lock Troubleshooting Action: enter wrong password, account is lock Repair Action: wait for some time, then enter wright password, then can enter in account Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: NA SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | Repair Action: wait for some time, then enter wright password, then can enter in account Test/Inspection: Functional Check PASSED. Device returned to full functionality. | [12 empty] | 117893670 |  | 1121897 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Password | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1155 |  |  | China | 福建省医科大学附属协和医院 | 345010 | 4439200% |  | 02N/Core128//CT account locked | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Calibration/Adjustment Problem Reported by Customer : 02N/Core128/ Next Step Action : NTD：协助接线同事开单 内容：ct账号被锁定 需要工程师账号解锁18650369579工程师来电要转RSE，中断了 T2 Activities : Problem Description: CT 账户被锁定 Troubleshooting Action: 需要重置密码 Repair Action: 到场重置密码，教育客户保存好密码和使用IT账户重置密码 Test/Inspection: N/A Test Equipment: N/A SAP Transfer : Case Transferred to SAP Safety Questions : Case: 0117888014 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Gaoge Last updated by: Wang Gaoge 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Dispatched for Onsite Support Internal Remarks : 15.07.2021 11:02:10 UTC Feng Li (86036224) RDI Accuracy: Yes $Q7: (Confirmed reported problem.The Ingenuity Core 128 CT user account is locked.) $Q8: (On site use Philips_service account to un-lock the user account and modify the password.) $Q9: (On site use Philips_service account to un-lock the user account and modify the password.) $Q10: (Test and inspection data refer to: PA.+Device returned to full functionality.) $Q11: Calibrated Tools name (FLUKE 287C Multimeter) S/N (16490128) NextCalibration date (10/20/2021) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date: NA ^Current Tube usage: NA External Remarks : 15.07.2021 11:02:07 UTC SUP JCOUSER (86036224)服务内容：CT医生用户账户被锁了。到场使用工程师账户重置医生账户，并修改密码。 服务结果：维修后，设备恢复正常工作。 ----------------------------------------------------------------------------- | 3 |  | Repair Action: 到场重置密码，教育客户保存好密码和使用IT账户重置密码 | [12 empty] | 117888014 |  | 1118020 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | customer | [6 empty] | Password | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1156 |  |  | China | 内蒙古自治区第四医院 | 40023 | 4439100% |  | 02N/Ingenuity CT 728326/ECG loss | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : connect fault Safety Questions : Case: 0117881084 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Weng Jinming Last updated by: Weng Jinming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 没有心电信号 客户问题详细描述或故障现象： 拔插USB，没有连接成功 设备本身技术问题： 是。 Troubleshooting Action: 远程诊断与测试结果： 怀疑USB接口问题 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议工程师到场检查USB接口，教客户使用飞利浦微信报修 SAP Transfer : Dispatched for Onsite Support SAP Transfer : Case Transferred to SAP Next Step Action : PPD IBOX2 USB EXTENDER Kit 459801413791 1 无货，替代号咨询RSE不能用，已安排订货，FSE：苏瑞杰-86-13304719335 Next Step Action : PPD 459801731761 替代号已确认可以出货，已安排24H到医院 External Remarks : 16.07.2021 14:19:07 UTC Ruijie Su (86039043)服务结果： [更换I BOX 板和Ibox线，设备恢复正常工作。] 其它重要事项：无 Internal Remarks : 16.07.2021 14:19:10 UTC Ruijie Su (86039043) $Q7: Confirmed reported problem,connect fault $Q8: On site checked ECG connect fault . Need replace Ibox and cable $Q9: replace I-BOX and cable system is ok! $Q10:Test and inspection data refer to: PA. Device returned to not full functionality. $ $Q11: Calibrated Toolsname ( Multimeter Fluke 287C) S/N (31060028) Next Calibrationdate (29/10/2021) 2( Torque Wrench Kit ) S/N E201602090 Next Calibration date10/09 /2021 For Mainland China CV / XR, CT engineers only. Please fill in: ^ Current Tube usage: scan second 53865 External Remarks : 16.07.2021 14:26:37 UTC Zhen Liang (86039057)服务结果： [更换I BOX 板和Ibox线，设备恢复正常工作。] 其它重要事项：无 Internal Remarks : 16.07.2021 14:26:52 UTC Zhen Liang (86039057)$Q7: Confirmed reported problem,connect fault $Q8: On site checked ECG connect fault . Need replace Ibox and cable $Q9: replace I-BOX and cable system is ok! $Q10:Test and inspection data refer to: PA. Device returned to not full functionality. $ $Q11: Calibrated Toolsname ( Multimeter Fluke 287C) S/N (31060028) Next Calibrationdate (29/10/2021) 2( Torque Wrench Kit ) S/N E201602090 Next Calibration date10/09 /2021 For Mainland China CV / XR, CT engineers only. Please fill in: ^ Current Tube usage: scan second 53865 Next Step Action : FSE：苏瑞杰-13304719335 备件DOA申请reopen，40-OP已派梁工 Next Step Action : 梁震 15849133876来电要求转单苏瑞杰 13304719335 Next Step Action : 已转派苏瑞杰 13304719335 External Remarks : 18.07.2021 22:39:41 UTC Ruijie Su (86039043))服务结果： [兼差I BOX 板和Ibox线，设备恢复正常工作。] 其它重要事项：无 Internal Remarks : 18.07.2021 22:39:45 UTC Ruijie Su (86039043)$Q7: Confirmed reported problem,connect fault $Q8: On site checked ECG connect fault . Need adjust Ibox and cable $Q9: adjust I-BOX and cable system is ok! $Q10:Test and inspection data refer to: PA. Device returned to not full functionality. $ $Q11: Calibrated Toolsname ( Multimeter Fluke 287C) S/N (31060028) Next Calibrationdate (29/10/2021) 2( Torque Wrench Kit ) S/N E201602090 Next Calibration date10/09 /2021 For Mainland China CV / XR, CT engineers only. Please fill in: ^ Current Tube usage: scan second 53865 ----------------------------------------------------------------------------- | 5 |  | 更换I BOX 板和Ibox线后IBOX仍然有断开报错 FSE说没有其他设备连接到IBOX， 也没有呼吸门控。 | [12 empty] | 117881084 |  | 1113582 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Ibox | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | TBD-NoEnoughInfo | Gantry-HW-TBD-NoEnoughInfo | Gantry-HW-TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo | Gantry-HW | Gantry-HW-TBD-NoEnoughInfo | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1157 |  |  | China | 苏州广慈肿瘤医院有限公司 | 345019 | 4439100% |  | 02N/Core128/cann't manual scan | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Core128/ Next Step Action : NTD：协助接线同事开单 内容：扫描的时候检测程序列无法手动触发，不能保障成功率 Safety Questions : Case: 0117880809 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Weng Jinming Last updated by: Weng Jinming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 偶发增强扫描，不能手动触发 客户问题详细描述或故障现象： 不能提供具体时间 设备本身技术问题： 是。 Troubleshooting Action: 远程诊断与测试结果： 很难判断问题，联系工程师徐浩天，到场检查 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议工程师到场检查log，判断问题，教客户使用飞利浦微信报修 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : Xu Haotian External Remarks : 16.07.2021 10:28:44 UTC Haotian Xu (86090280) 服务内容：做增强手动触发偶发无法使用，检查log报错和软件系统，配合培训老师 与客户沟通 服务结果：服务后，设备恢复正常 ----------------------------------------------------------------------------- | 5 |  | $Q7 (Confirmed reported problem , unable to manually scan ) $Q8: (Checked the log and the protocols,and found the problem is a software bug) $Q9: (connect to the trainning teacher and costomers) $Q10: (Test and inspection scan refer to PA . device returned to full functionality) | [12 empty] | 117880809 |  | 1113397 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | CTA | [5 empty] | Refresh |  |  | System issue | System | Console-SW | CTA | Console-SW-CTA | Console-SW-CTA |  | CTA | Console-SW | Console-SW-CTA | No | 0 | 0 | 0 | 0 | 0 | No | y | y | 0 | 0 | 0 | n | y | y | y | y | #N/A
1158 |  |  | China | Shaanxi Province Xianyang Central Hospital (Xianyang City No. 2 People's Hospital) | 40050 | 4439000% |  | 02N/Ingenuity CT/gantry noise | BU Q&R Feedback: No Reason: Corrective Maintenance Next Step Action: NTD: assist the operator to open the service order Shaanxi Province Xianyang City Central Hospital53636103 Teacher Ma 13809141954 gantry has abnormal noise, yes Problem Reported by Customer: 05R/Brilliance CT 64/ Safety Questions:Case: 0117874106 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Bo Last updated by: Wang Bo 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer's problems: device has abnormal smell. Detailed description of Customer's problems or malfunctions: The gantry is making an abnormal sound. It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: The log has no errors. Remote troubleshoot and test result: Restart but didn't work. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: It is recommended to check the noise source, utility transformer, rotor brake part, rotating belt and pressure wheel. We recommend spare parts 12NC and names: N/A SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Next Step Action:GUO YUJIE External Remarks: 13.07.2021 15:20:53 UTC SUP JCOUSER (86038990) Service details: Gantry noise. Check to see if it's new machine belt and motor friction noise. Add talcum powder to restore. service result: device returns to normal working condition after service ----------------------------------------------------------------------------- | 3 |  | 服务内容：机架异响，检查是新机器皮带和电机摩擦声，加滑石粉恢复。 服务结果：服务后，设备正常工作。 | [12 empty] | 117874106 |  | 1108663 | TRUE | Evaluation | Evaluation |  | TBD-GantryNoise | [5 empty] | Engineering | [6 empty] | Gantry Noise | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Noise | Gantry-HW-Noise | Gantry-HW-Noise | TBD-GantryNoise | Noise | Gantry-HW | Gantry-HW-Noise | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1159 |  |  | China | Jinan City No. 4 People's Hospital | 40038 | 4438600% |  | 02n/Ingenuity CT 728326/Check the internal cable machine | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: Check the internal cable machine Next Step Action: FSE chentao called to request for repair Safety Questions:Case: 0117858853 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Lili Last updated by: Lu Lili 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer's problems: check device. Detailed description of Customer's problems or malfunctions: check device. It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: N/A Remote troubleshoot and test result: N/A PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: The equipment warranty is about to expire. The customer asked FSE to go to the hospital to check the equipment We recommend spare parts 12NC and names: N/A Memo of Customer request: None SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Next Step Action : Chen Tao External Remarks:09.07.2021 12:52:41 UTC Tao Chen (86038373) Error description: Check the internal cable of the rack Service content: Check and secure the internal cables of the rack Service outcome: Device returns to normal working condition after servicing. Internal Remarks : 09.07.2021 13:18:22 UTC Tao Chen (86038373) “RDI Accuracy: YES" $Q7: (Confirmed reported problem, check cables in the gantry) $Q8: ( Check cables of gantry rotor and stator according to customer requirement) $Q9: ( Fasten cables for the gantry) $Q10: (Test and inspection data refer to: PA. Device returned to full functionality) $Q11: Calibrated Tools name ( N/A) ^Current Tube installed date: N/A ^Current Tube usage: N/A ----------------------------------------------------------------------------- | 5 |  | $Q7: (Confirmed reported problem, check cables in the gantry) $Q8: ( Check cables of gantry rotor and stator according to customer requirement) $Q9: ( Fasten cables for the gantry) | [12 empty] | 117858853 |  | 1105707 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Others | [6 empty] | Gantry Cable | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1160 |  |  | China | Dongguan Donghua Hospital Co., Ltd. | 40059 | 4438900% |  | WeChat:02N/Ingenuity CT 728326/adjust CIRS time | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer:adjust CIRS time Safety Questions:Case: 0117867602 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Guoping Last updated by: Liu Guoping 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support External Remarks: 12.07.2021 14:27:31 UTC Zhenhua Liu (86034125) Error description: CIRS time is not in synch with host Service details: synchronizing the time of the CIRS with host Service result: After service, the machine is back to normal. Other important issues: None Internal Remarks : 12.07.2021 14:27:33 UTC Zhenhua Liu (86034125) "RDI Accuracy:Yes" $Q1: () $Q2: () $Q3: () $Q4: () $Q7: (Confirmation & Description of problem) Confirmed reported problem, the time of CIRS is not sync with host $Q8: (Troubleshooting & Results) time of CIRS is not sync with host $Q9: (Repair steps) sync the time of cirs and host $Q10: (Test and inspection data refer to PA . Device returned to full functionality ) $Q11: Calibrated Tools name ( refer to PA ) S/N ( ) Next Calibration date ( refer to PA ) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date: N/A ^Current Tube usage: N/A ----------------------------------------------------------------------------- | 5 |  | $Q7: (Confirmation & Description of problem) Confirmed reported problem, the time of CIRS is not sync with host $Q8: (Troubleshooting & Results) time of CIRS is not sync with host $Q9: (Repair steps) sync the time of cirs and host $Q10: (Test and inspection data refer to PA . Device returned to full functionality ) | [12 empty] | 117867602 |  | 1102201 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Service | [6 empty] | Time Sync | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1161 |  |  | China | Xifeng County Traditional Chinese Medicine Hospital | 345014 | 4438900% |  | Wechat:02N/Ingenuity Core128/Elite China/ACQ error | BU Q&R Feedback: No Reason: Corrective Maintenance T2 Activities:Problem Description: Summary of customer's problems: Scanning failed. Detailed description of Customer's problems or malfunctions: Report ACQ error It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: DatapathCheck(127): Excessive acquisition Holdoff or stalled buffered stream m_seriesUid = 0x4D timeSinceLastReadingSec = 1.31259 bufStreamWorkquePtr = 0 acquisitorHoldoffMs = 0 m_maxHoldoffTimeSec = 1.28 SafetyCheck(873): 0x4d SafetyCheckThread: HaltAcquisition WorkItem index = 33 SafetyCheck.cpp(469): Error #12 - Unknown error 0x4d SafetyCheckThread: Datapath error. Data Path Diagnostics End Series Series UID 0x4d Real-time stream Number of channels: 4 Physical channel order: 0x4321 Unexpected gantry Confirm Resend: 0 (1 = error) Reading resends: 0 CRC error count Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 Loss of Signal Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 Line Fault Indicator Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 Resend bit-mapped angle Chan A: 0x0000000000000000000000000000000000000000000000000000000000000000 Chan B: 0x0000000000000000000000000000000000000000000000000000000000000000 Chan C: 0x0000000000000000000000000000000000000000000000000000000000000000 Chan D: 0x0000000000000000000000000000000000000000000000000000000000000000 DMS recoverable error count: 0 Duplicate reading count: 0 On-board overflow retry count: 0 On-board overflow max retry period: 0 ms On-board overflow total retry time: 0 ms Remote troubleshoot and test result: Recon removed ACQ_OK signal to stop the scan. There are too many reasons it can either be the Reconstructor or Gantry causing this error. We are still improving the AEL in which we can tackle this error. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: "1. Check for CIRS errors and follow their recommendations. 2. Run Visual Diagnostics/Data Path test. 3. Check during a scan that Enable Acq and Acq Ok Led on the GHOST goes on. If not or any doubt because of intermittent error, run the Acq test with the electrical loopback to verify the Acquisitor PCB (verification of signals comm Ok, enable acq and acq ok). Check / Reseat the internal CRC cable connected to the Acquisitor J4 connector were the cable with D15 connector is mounted. Also check the cable from the CRC rack to the Ghost. 3. Check the Raid at the server 1 for health indication and/or errors on one of the disks. 4. Connect the optical fiber cable directly between the RCOM and ACQ when ACQ_OK is now visible at the RCOM (no S_MASTER_SEQUENCER_ACQ_NOT_OK error in log) you know that there is an issue at the optical slip ring. 5. Check the CIRS.log if you find ""bad dynamic rotor position"" and system has no resends, datapath test runs without errors. The problem was solved after DMC-16 replacement, probably bad control link data caused by the DMC-16 pcb. 6. Check the CIRS. log when you find ""bad table motion"" indicating couch pulses do not match rotor pulses causing the acq error. Check Couch and /or Angulation system. 7. If encountered without a CIRS. error reported and no surrounding Normal errors reported in logger, this indicates a problem with the ACQ_OK signal coming from the ACQ card to the Ghost. Look at the ACQ_OK led on the Ghost and see if it is lit. If it is lit and there are no CTRL link errors then it is a problem with the Ghost. If CTRL link error the check slipring signal brushes and DMC, If it is out then the Ghost never received the ACQ_OK signal from the ACQ card. Check the connections from the ACQ Card to the Ghost, verify cable integrity and replace if bad, replace ACQ card, replace Ghost. Experience: 1. Check the Raid at the server 1 for errors on one of the disks. 2. Check the CIRS.log for ""ERROR - ReadingType changed during shot"" and / or ""ERROR - Number of slices changed during shot"". During a shot the dms is sending temperature data which interrupts the shot. . In the Raw data it shows the X-Ray signal jump to 0 and back to 1 - causing the H.W to send Temperature reading again. So it’s a problem with the “X-Ray On” signal and the parts involved are: Dmc, Rcom, Rhost to DMC Cable 4535-670-25791, and Dmc to Rcom Cable. Replacing DMC and RCOM did not solve, but replacing the RHOST solved the problem. 3. On a TDMS system, got high resends and at times the order of the channels will changed by itself from the normal 4321 (in the CIRS log ""Unexpected"" messages) to another order. Visual Diagnostic reported that an index pulse was received by the RCOM but was not by the Tx box. Found that the wire from pin 12 (AUX_0_INDEX_N) that connects to DMC16 J23 in the Aux. Shut Down Adapter cable 453-567-525791 was broken. This cable adapts to the long cable that connects the DMC to the Tx box. 4. On FRC TX module power connection was problematic. Socket seems seated but it was not working. There were 15V while not connected but it was dropping to 3V while connected. Connection socket seated and sealed fixed the problem. 5. ACQ _NOT _OK with parm 1=0 happened in high speed and High MA scan, no other Xray errors but, in CIRS Errorlog a few lines before the ACQ error, you can see :AcqWorkItem.cpp216ERROR - Max number of consecutive Readings with tube arcs (HV drops) exceeded, BufferSel = 1, ReadingOffset = 0x139bc00. Checked in the Can log and found a lot of soft arcs- Tube was replaced and didn't fix this problem. The RHOST fixed it. 6. Detector module 10 was swapped with module 2 location to see if ring would move to outside. Tried to perform air calibrations on system and received this error. Error went away when module 10 was replaced with new one and performing calibration procedure for module replacement. Also moved module 2 back to original location. Use FEE Test and view the plots to look for any dead modules that may hang up the datapath." SAP Transfer:Dispatched for On site Support Problem Reported by Customer: Wechat:02N/Ingenuity Core128/Elite China/ACQ error Safety Questions:Case: 0117866756 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Jianqiang Last updated by: Wang Jianqiang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | To be followed with field and get bugreport | [12 empty] | 117866756 |  | 1101522 | TRUE | Closed - Done | Closed - Done |  | CTUDT00630408 | [5 empty] | Engineering | [6 empty] | ACQ Error | [5 empty] | Refresh |  |  | System issue | System | CIRS-SW | ACQ_NOT_OK | CIRS-SW-ACQ_NOT_OK | CIRS-SW-ACQ_NOT_OK | ACQ_NOT_OK | ACQ_NOT_OK | CIRS-SW | CIRS-SW-ACQ_NOT_OK | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1162 |  |  | Japan | Social Welfare Organization Saiseikai Imperial Gift Foundation, Inc. | 343000 | 4439100% |  | 01CS:TM:CT-Host power was off although nothing was being done. >> There was normal operation once power was [Truncated] | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer: 01CS:TM:CT-Host power was off. >> Safety Questions:Case: 0117883620 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: TAKEYUKI MIYAKAWA Last updated by: TAKEYUKI MIYAKAWA 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Problem description by engineer : Equipment affected: CT Problem location: 01 Console When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user's work flow did the problem occur?: Other What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem?: NA Engineer's Comment: '============================== $version = "v4.1.10.25210" 22:21 Call description: Normally the CT-Host power is left on, but it turned off at some time. The CIRS power was still on. '============================== Troubleshooting Action: What steps are taken during the system diagnosis: Result of response: Problem solved. Problem location: 01 Console What steps did you perform to diagnose the problem? '================================== Results of checking the log: I checked the event log and found that the shutdown was running normally at 16:20 today.There was no error recorded. Also, just before that, the key switch was off.(I cannot identify the problem from the log, but it seems he has also logged out) CIRS power was still on. I think, from the situation, that the user shut down the CT-Host. '================================== Error # and/or description of error: NA Repair Action: [Confirmation of problem] Confirmed. Work details: I checked the log and reported to the customer. 【Confirmation content】 Good. OneEMS internal remarks : Problem description by engineer : Equipment affected: CT Problem location: 01 Console When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user's work flow did the problem occur?: Other What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem?: NA Engineer's Comment: '============================== $version = "v4.1.10.25210" 22:21 Call description: Normally the CT-Host power is left on, but it turned off at some time. The CIRS power was still on. '============================== What steps are taken during the system diagnosis: Result of response: Problem solved. Problem location: 01 Console What steps did you perform to diagnose the problem? '================================== Results of checking the log: I checked the event log and found that the shutdown was running normally at 16:20 today.There was no error recorded. Also, just before that, the key switch was off.(I cannot identify the problem from the log, but it seems he has also logged out) CIRS power was still on. I think, from the situation, that the user shut down the CT-Host. '================================== Error # and/or description of error: NA Repair Action "External" : [Confirmation of problem] Confirmed. Work details: I checked the log and reported to the customer. 【Confirmation content】 Good. Repair Action "Internal" : --Troubleshooting Action-- Result of response: Problem solved. Problem location: 01 Console What steps did you perform to diagnose the problem? '================================== Results of checking the log: I checked the event log and found that the shutdown was running normally at 16:20 today.There was no error recorded. Also, just before that, the key switch was off.(I cannot identify the problem from the log, but it seems he has also logged out) CIRS power was still on. I think, from the situation, that the user shut down the CT-Host. '================================== --Repair Action-- --Subject-- 01CS:TM:Mr. Yamaguchi:CT-Host power was off although nothing was being done. >> There was normal operation once power was turned on.It looks like a shutdown was done normally in the evening. --Problem Description-- Equipment affected: CT Problem location: 01 Console When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user's work flow did the problem occur?: Other What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem?: NA Engineer's Comment: '============================== $version = "v4.1.10.25210" 22:21 Call description: Normally the CT-Host power is left on, but it turned off at some time. The CIRS power was still on. '============================== --4/8Elements-- Q1. Was the reported problem confirmed? Confirmed. Q2. Were any other problems confirmed? None. Q3. What did you do to resolve the problem?: I checked the log and reported to the customer. Q4: Is the equipment operating normally following the response? Good. Q5. How did the examination finish? None. Q6. Were there any mistaken findings (misdiagnosis, mistaken treatment) due to the problem? None. Q7. Was there any unintended unnecessary exposure? None. Q8. Were there any artifacts? None. SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 4 |  | from the situation, that the user shut down the CT-Host. | [12 empty] | 117883620 |  | 1115498 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Host off | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1163 |  |  | Japan | 医療法人清翠会 牧病院 | 341017 | 4438900% |  | 01DS: KK: Cannot transfer images to PACS. >> It does not return by restarting the Host. It became possible to transfer after a while. | 01DS: KK: Cannot transfer images to PACS. >> It does not return by restarting the Host. It became possible to transfer after a while. " BU Q & R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 01DS: KK: Unable to transfer images to PACS. >> It does not return by restarting the Host. It became possible to transfer after a while. SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Support T2 Activities: Problem Description: Problem description by engineer: Target device: CT Failure location: 01Cosole When did it first occur? (Within ~): 1 Hour How often do problems occur ?: First Occurrence At what point in the end-user workflow is the problem occurring ?: Network What happened recently ?: N.A. Does the customer create a specific system event report for the problem ?: No. The customer was unable to create a log file or report. Troubleshooting Action: What steps are taken during the system diagnosis: Post-response results: Improved Failure location: 01Cosole What steps were taken during the diagnosis ?: Images are no longer transferred to PACS. I restarted the Host, but it didn't work. And incoming call. When I checked the log, the PACS queue was accumulated in Waiting. Ping went to PACS. When I contacted the customer with the result of log confirmation, it started to be transferred. Consult with FSE. It was investigated after raising Bugrep. Error # and / or description of error: NA Repair Action: Service representative visit work. OneEMS internal remarks: Problem description by engineer: Target device: CT Failure location: 01Cosole When did it first occur? (Within ~): 1 Hour How often do problems occur ?: First Occurrence At what point in the end-user workflow is the problem occurring ?: Network What happened recently ?: N.A. Does the customer create a specific system event report for the problem ?: No. The customer was unable to create a log file or report. What steps are taken during the system diagnosis: Post-response results: Improved Failure location: 01Cosole What steps were taken during the diagnosis ?: Images are no longer transferred to PACS. I restarted the Host, but it didn't work. And incoming call. When I checked the log, the PACS queue was accumulated in Waiting. Ping went to PACS. When I contacted the customer with the result of log confirmation, it started to be transferred. Consult with FSE. It was investigated after raising Bugrep. Error # and / or description of error: NA Repair Action "" External "": Visit work for service personnel. Repair Action "" Internal "": --Troubleshooting Action-- Post-response results: Improved Failure location: 01Cosole What steps were taken during the diagnosis ?: Images are no longer transferred to PACS. I restarted the Host, but it didn't work. And incoming call. When I checked the log, the PACS queue was accumulated in Waiting. Ping went to PACS. When I contacted the customer with the result of log confirmation, it started to be transferred. Consult with FSE. It was investigated after raising Bugrep. --Repair Action-- --Subject-- 01DS: KK: Noda: I can't transfer images to PACS. >> It does not return by restarting the Host. It became possible to transfer after a while. --Problem Description-- Target device: CT Failure location: 01Cosole When did it first occur? (Within ~): 1 Hour How often do problems occur ?: First Occurrence At what point in the end-user workflow is the problem occurring ?: Network What happened recently ?: N.A. Does the customer create a specific system event report for the problem ?: No. The customer was unable to create a log file or report. External Remarks: 14.07.2021 00:49:45 UTC YOSHIYUKI YAMAMOTO (81100236) [Request] PACS transfer is stopped halfway. [Phenomenon confirmation] I was able to confirm. [Details of work] I was able to confirm the phenomenon from the log. About the phenomenon that the image transfer that occurred this time stops in the middle, I got the log file at the time of occurrence. We will send it to the Dutch headquarters for analysis. It is a function that the software has that the transfer resumed and completed without doing anything This is due to the fact that the "function that times out 15 minutes after the transfer stagnation occurs and restarts the transfer" has worked. [Other failures / treatments] Yes When QRing to PACS, all series were searched and frozen, so We have changed the settings so that you can sort by any search. [Device status] Equipment can be used [Work status] done 12-Jul-2021 81100236 FSE Handover Tool1.3.4.32612.DB13 20210415 Internal Remarks: 14.07.2021 00:49:49 UTC YOSHIYUKI YAMAMOTO (81100236) [Replacement parts] No replacement parts [Misorder Parts Report] No mistakes [Waiting information] No wait-and-see [TS Requests] No return parts [CT additional information] Q1 How did the inspection end ?: No extra irradiation due to stop before X-ray irradiation Q2 Presence or absence of incorrect findings (misdiagnosis, mistreatment) due to problems: None Q3 Presence or absence of artifacts: None Q4 Presence or absence of unintended unnecessary exposure: None [Repair certificate] Repair certificate is pasted [Other information sharing matters] 1. Treatment for patients infected with or suspected of being infected with the new coronavirus / Is it quarantined? Okay 2. Do I need to enter the treatment room or bed where the infected patient is? (Workplace and transfer (Including dynamic path) No Did you agree on how to complete the report with the person in charge of the equipment? YES Normal report completed on the day 12-Jul-2021 81100236 FSE Handover Tool1.3.4.32612.DB13 20210415 | 4 |  | About the phenomenon that the image transfer that occurred this time stops in the middle, I got the log file at the time of occurrence. We will send it to the Dutch headquarters for analysis. It is a function that the software has that the transfer resumed and completed without doing anything This is due to the fact that the "function that times out 15 minutes after the transfer stagnation occurs and restarts the transfer" has worked. | [12 empty] | 117866778 |  | 1101553 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS Transfer | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1164 |  |  | Puerto Rico | DORADO HEALTH INC DBA | 40007 | 4439200% |  | Noise on Gantry | Work Order #: WO-03867300 Alec, Preciado The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Internal Comments : Error # and/or description of error : belt wear down and tensiones out of spec. Repair Action : Parts requested by the FSE, belt and tension assembly Diagnostic performed by Engineer : Parts requested by the FSE, belt and tension assembly Problem Description by Engineer : Problem description by engineer : The rotor belt and belt tensioner was defective, please order both parts When was the first occurrence? Within... : 16-Jul-2021 10:00:00 How often is issue occurring? : First occurrence Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected results: banda de rotor reemplazada y ajustada Patient Impact/Outcome: no impacta User Impact: no impacta Current Software Version: 4.1.7 Resolution : Belt and tensioner assembly replaced Follow Up Required : Error # and/or description of error : belt wear down and tensiones out of spec. Repair Action : Parts requested by the FSE, belt and tension assembly What are the next Troubleshooting steps : Replace belt and tensioner and adjust as indicated in the repair replacement manual. Technician Name : Alec, Preciado ----------------------------------------------------------------------------- | 3 |  | The rotor belt and belt tensioner was defective, please order both parts | [12 empty] | 117892755 |  | 1120398 | TRUE | Evaluation | Evaluation |  | ECR-077012 | [5 empty] | Engineering | [6 empty] | Belt | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Belt | Gantry-HW-Belt | Gantry-HW-Belt | ECR-077012 | Belt | Gantry-HW | Gantry-HW-Belt | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1165 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 337000 | 4439300% |  | Ring artifact problem | Resolution: It was seen that there was a ring artifact in the images taken on the device. When the inside of the gantry was checked, opaque material was seen on the detector sheath, under the mylar and on the collimator. Air calibration was performed by cleaning the opaque material. A test shot was made with the Phantom, no problem was seen, the device was delivered in an active manner. | 3 |  | opaque material was seen on the detector sheath | [12 empty] | 117895561 |  | 1119110 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Contrast intrusion | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1166 |  |  | United States | Philips Medical Capital | 40055 | 4439200% | 28 | 0603CDY70 Customer is getting and error when trying to run QC. | Problem Description by Engineer : Problem description by engineer : Spoke with Brent.The customer is unable to scan. They are seeing a CIRS error. Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: No further information is available as per the customer Patient Impact/Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer Current Software Version: 4.1.10 Diagnostic performed by Engineer : Used logger.mdb, UNLV, and AEL to locate the error messages. Follow Up Required : Error # and/or description of error : S_DMC_SHOT_COMM_COLLAPSE. This indicates that no frames are being sent betweenthe DMS and the CIRS. This is a datapath issue. Parts required : Part : 1, 12NC : 453567057167, Description : FRC Assembly, Quantity : 1, Priority : 1, Chance Of Success : 6.1 Part : 2, 12NC : 453567538334, Description : Module - Rotor Electronics TX, Quantity : 1, Priority : 2, Chance Of Success : ? Part : 3, 12NC : 453567538322, Description : Module, Optical RX Stator Moog Velocity, Quantity : 1, Priority : 3, Chance Of Success : ? Repair Action : The S_DMC_SHOT_COMM_COLLAPSE indicates a datapath issue. The CIRS event log shows no line faults or data dropouts so it appears that the frames are never being sent from the DMS/RCOM. Look for power from the FRC to the transmitter to be present. Use the diagnostic LEDs on the TX to assist. Run datapath diagnostics to look for the source of the errors. Bypass the optical slip ring with a quad fibre from the RCOM and do a non rotation scan. What are the next Troubleshooting steps : Look for power from the FRC to the transmitter to be present. Use the diagnostic LEDs on the TX to assist. Run datapath diagnostics to look for the source of the errors. Bypass the optical slip ring with a quad fibre from the RCOM and do a non rotation scan. Technician Name : Mark Finnigan Internal Comments : Error # and/or description of error : S_DMC_SHOT_COMM_COLLAPSE. This indicates that no frames are being sent betweenthe DMS and the CIRS. This is a datapath issue. Repair Action : The S_DMC_SHOT_COMM_COLLAPSE indicates a datapath issue. The CIRS event log shows no line faults or data dropouts so it appears that the frames are never being sent from the DMS/RCOM. Look for power from the FRC to the transmitter to be present. Use the diagnostic LEDs on the TX to assist. Run datapath diagnostics to look for the source of the errors. Bypass the optical slip ring with a quad fibre from the RCOM and do a non rotation scan. ----------------------------------------------------------------------------- | 2 |  | 40055 site is a false alarm. A wall box power recycle fixed he problem. No parts replaced. No more errors on site since day1 of the problem. system is monitoring. | [12 empty] | 117890421 |  | 1115563 | TRUE | Evaluation | Evaluation | [7 empty] | Others | [6 empty] | Power cycle | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1167 |  |  | United States | Philips Medical Capital | 40053 | 4438900% |  | Intermittent Issues with EKG Connection | Internal Remarks : FADB: (FSE Aware of Double Booking) Title: Internal Remarks Initials: Twila Actions Taken: Per message sent in Teams from FSE/McDowell Preston - Moved WO from 7/14/21 to 7/21/21 (anytime). I kept the same time. Outcome: Moved WO per FSE's Request IP: 88993323 Next Steps: None ----------------------------------------------------------------------------- Work Order #: WO-03832660 McDowell Preston Diagnostic performed by Engineer : N/A Problem Description by Engineer : Customer Function/Role: Technologist Device Use: Diagnostic Expected and Unexpected Results: Customer States Intermittent EKG Waveform Not Present on Host PC Monitor. Customer has to Restart Host PC to get EKG Waveform User Impact: Unknown Patient Impact: Unknown Current Software Version:4.1.10 ----------------------------------------------------------------------------- Work Order #: WO-03840684 Weaver David The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Description of Problem Intermittent Issues with EKG Connection CALLER FUNCTION/ROLE Technician/Sonographer/Therapist DEVICE INTENDED USE Diagnostic WHEN WAS THE FIRST OCCURRENCE? WITHIN . . . <1 Week HOW OFTEN IS ISSUE OCCURRING? Intermittent WHERE IN THE END-USER WORKFLOW IS THE ISSUE OCCURRING? Pre-acquisition/scan WHAT RECENT EVENTS HAVE OCCURRED? During normal use NEGATIVE OPERATOR IMPACT EKG not workin sometimes NEGATIVE PATIENT IMPACT/PROCEDURE OUTCOME Delay in scanning HAS CUSTOMER CREATED A SPECIFIC SYSTEM EVENT REPORT OF ISSUE? No 4.1.10 Diagnostic performed by Engineer : Intermittent Issues with EKG Connection Resolution : Intermittent Issues with EKG Connection. ----------------------------------------------------------------------------- Work Order #: WO-03832183 Paine Matthew The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Customer Function/Role: Technologist Device Use: Diagnostic Expected and Unexpected Results: Customer States Intermittent EKG Waveform Not Present on Host PC Monitor. Customer has to Restart Host PC to get EKG Waveform User Impact: Unknown Patient Impact: Unknown Current Software Version:4.1.10 Diagnostic performed by Engineer : Customer states that they have to restart the Host PC everytime the go to use the EKG Monitor. After Restarting the Host PC, the EKG Waveform is now present on the Host PC. Follow Up Required : Issue: Customer States Intermittent Issue with EKG Waveform Not Present on Host PC Monitor, but is Present on the EKG Monitor. Customer States Restart of Host PC Resolves Issue. Recommendation: Check All Cabling to CRC Rack. Use Upgrade Installation Manual for Proper USB Cable Locations. Check to Ensure that Cabling from SBD to Host PC is Correct. Using the wrong USB can cause Intermittent Issue's if not Installed Correctly. If Cabling is correct, check USB Extender Version. If Upgarde Kit is needed: 459801731761-USB Extender Cable from IBOX To Console Internal Comments : Called and left Primary FSE Preston McDowell a VM. FSE needs to Inspect Cabling for EKG from Ibox to CRC Rack and from CRC Rack to Host PC to ensure proper cabling for USB2 Devices. ----------------------------------------------------------------------------- | 4 |  | Product/Part - USB EXTENDER CABLE FROM IBOX TO CONSOLE Part number - 459801731761 | [12 empty] | 117870146 |  | 1099343 | TRUE | Closed - Done | Closed - Done | [7 empty] | Others | [6 empty] | USB Cable | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | n | y | y | y | y | #N/A
1168 |  |  | Viet Nam | Công ty TNHH Thương mại và | 40046 | 4438900% |  | Vert Encoder got defective | Product/Part - VERT ENCODER ASSY Part number - 453566502431 | 4 |  | Product/Part - VERT ENCODER ASSY Part number - 453566502431 | [12 empty] | 117866928 |  | 1101707 | TRUE | Closed - Done | Closed - Done |  | TBD-Encoder | [5 empty] | Engineering | [6 empty] | Vertical Encoder | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | VerticalEncoder | Couch-HW-VerticalEncoder | Couch-HW-VerticalEncoder | TBD-Encoder |  | Couch-HW | Couch-HW-VerticalEncoder | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | #N/A
1169 |  |  | Chile | Scotiabank Chile | 40008 | 4439700% |  | CL1675 - ESPEJO - EQUIPMENT NEEDS TO BE CONFIGURED FOR SEVERAL RIS PACS TASKS. ON AUGUST 21ST AT 2:00 P.M. LOCAL TIME |  | 3 |  | RIS/PACS config | [12 empty] | 117917617 |  | 1135736 | TRUE | Opened | Opened | [7 empty] | Customer | [6 empty] | RIS/PACS config | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1170 |  |  | China | 四川华蓥市华蓥山广能集团总医院 | 345000 | 4440100% | 29 | 02N/Ingenuity Core128/Elite China/scan fault | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity Core128/Elite China Safety Questions : Case: 0117957832 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Bo Last updated by: Wang Bo 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 扫描失败。 客户问题详细描述或故障现象： 定位像扫描失败。 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： ACQ NOT OK，CRC ABCD=700+ RESEND=500+ fiber optical disconnected 远程诊断与测试结果： 重启无效。 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议现场做ACQ自诊断，测量光强，检查接受发射盒供电和工作状态，清洁光路，bypass Rcom-CIRS端尝试排除故障。 建议备件12NC及名称： PPD 453567538334，453567538322 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : PPD 453567538322*1+453567538334 *1 已安排备件24H到医院 李全飞-86-13699457740 External Remarks : 25.07.2021 07:59:39 UTC Haijun Chen (86001310)"SUAT. Ready for Clinical Use local time" : "2021-07-24 23:00" External Remarks : 25.07.2021 08:02:43 UTC SUP JCOUSER (86001310)服务内容:机器不能扫描，发射盒有问题，更换发射盒。 服务结果: 维修后，设备恢复正常工作。 Internal Remarks : 25.07.2021 08:02:46 UTC Haijun Chen (86001310) "RDI Accuracy:Yes" $Q7:Confirmed reported problem.Can't scan. $Q8: Check the logs of CIRS Error: ACQ NOT OK and find the TX have problem. $Q9: Replace TX(12NC:453567538334) $Q10: Test and inspection data refer to: PA. Device returned to full functionality. $Q11: NA. current tube install dated:(N/A) current tube usage(:N/A) | 2 |  | $Q7:Confirmed reported problem.Can't scan. $Q8: Check the logs of CIRS Error: ACQ NOT OK and find the TX have problem. $Q9: Replace TX(12NC:453567538334) $Q10: Test and inspection data refer to: PA. Device returned to full functionality. | [12 empty] | 117957832 |  | 1148076 | TRUE | Opened | Opened |  | TBD_OBS | [5 empty] | Engineering | [6 empty] | TX | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Other | Gantry-HW-Other | Gantry-HW-Other | TBD_OBS | TBD-NoEnoughInfo | Gantry-HW | Gantry-HW-Other | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1171 |  |  | China | 北京市中西医结合医院 | 40032 | 4439900% |  | Wechat:02N/Ingenuity CT 728326/connect PACS | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : connect PACS Safety Questions : Case: 0117927110 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wennian Han Last updated by: Wennian Han 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 添加PACS 设备本身技术问题： 是。 Troubleshooting Action: 远程诊断与测试结果： 添加PACS PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 添加PACS，教客户使用飞利浦微信报修 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 22.07.2021 11:20:53 UTC Guochao Fan (86039456)服务内容：添加pacs节点。 服务结果：服务后设备恢复正常。 Internal Remarks : 22.07.2021 11:20:56 UTC Guochao Fan (86039456) "RDI Accuracy: Yes" $Q7: Confirm reported problems:Adding pacs $Q8: N/A $Q9:N/A $Q10: Test and Inspection data refer to PA& Device status:Device returned to full functionally $Q11: N/A ----------------------------------------------------------------------------- | 5 |  | RIS/PACS config | [12 empty] | 117927110 |  | 1142280 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Customer | [6 empty] | PACS Config | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1172 |  |  | China | 东莞市中西医结合医院 | 345013 | 4439800% |  | WeChat:02N/Ingenuity Core128/ECG problem | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : error External Remarks : 23.07.2021 13:12:44 UTC Zongxian Li (86001504)服务内容：更换导联线。 服务结果：工作正常。 Safety Questions : Case: 0117920811 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wennian Han Last updated by: Wennian Han 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP Internal Remarks : 23.07.2021 13:12:47 UTC Zongxian Li (86001504) RDI Accuracy:Yes $Q7: (Confirmed the reported problem, ECG signal abnoarmal) $Q8: (Checked the leadset in MP5 broken) $Q9: (Replaced leadset) $Q10: (Test and inspection data refer to PA. Device returned to full functionality.) $Q11: $Calibrated Tools name (N/A) S/N (N/A) Next Calibration date (N/A) SAP Transfer : Dispatched for Onsite Support T2 Activities : Problem Description: 客户问题概述： ECG相关故障。 客户问题详细描述或故障现象： 心电问题 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： NA 远程诊断与测试结果： 客户告知总是出現锯齿样心电图，有时是心电监护仪的心电图与我操作电脑的心电图不一致，检查过电极片，擦拭酒精，排除了心率问题。需要到场检查心电 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 检查电极片，接地，导联线等线缆、IBOX程序，IBOX板、检查IBOX到主机通路，以及转换盒状态，检查心电监护仪。 建议备件12NC及名称： " 3 Leadset, Grabber, IEC, ICU 989803145101" Next Step Action : PPD 989803145101*1 已安排24h到医院 FSE李宗贤-86-13751567086 ----------------------------------------------------------------------------- | 3 |  | 3 Leadset, Grabber, IEC, ICU 989803145101 | [12 empty] | 117920811 |  | 1138158 | TRUE | Investigation | Investigation | [7 empty] | Others | [6 empty] | ECG | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1173 |  |  | China | Shandong Province Chest Hospital | 40025 | 4439700% |  | 02N/Ingenuity CT 728326/Unable to start normally | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : Unable to start normally Safety Questions : Case: 0117916251 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Zhou Xiaohong Last updated by: Zhou Xiaohong 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 软件登录故障 客户问题详细描述或故障现象： 软件登录故障 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： N/A 远程诊断与测试结果： N/A PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 软件bug，已开ESCALATION，需要FSE现场用管理员账户登录，恢复CT账号。 建议备件12NC及名称： N/A 客户需求备注： 无 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Internal Remarks : 25.07.2021 00:00:59 UTC Zhaokun Huang (86038405) "RDI Accuracy: No." $Q7: ( Confirmed reported problem & Software can't login in) $Q8: ( Login in Software use smart card and reload software. & It is app bug from TSE and fixed wait for SP2.) $Q9: ( Login in Software use smart card and reload software . ) $Q10: ( Test and inspection data refer to: PA . Device returned to full functionality.) $Q11: Calibrated Tools name (N/A) ^Current Tube installed date: N/A ^Current Tube usage: N/A External Remarks : 21.07.2021 10:00:16 UTC Zhaokun Huang (86038405)服务内容：按照手册，重新安装软件。 服务结果：维修后，设备恢复正常。 ----------------------------------------------------------------------------- | 3 |  | blue screen $Q7: ( Confirmed reported problem & Software can't login in) $Q8: ( Login in Software use smart card and reload software. & It is app bug from TSE and fixed wait for SP2.) $Q9: ( Login in Software use smart card and reload software . ) | [12 empty] | 117916251 |  | 1134828 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632553 | [5 empty] | Engineering | [6 empty] | Blue screen | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Blue-Screen | Console-SW-Blue-Screen | Console-SW-Blue-Screen | CTUDT00632553 | Blue-Screen | Console-SW | Console-SW-Blue-Screen | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1174 |  |  | China | Nanchang University First Affiliated hospital (Jinggangshan Hospital) | 345002 | 4439700% |  | 02N/Ingenuity Core128/Elite China/ | BU Q&R Feedback: No Reason: Corrective Maintenance Safety Questions:Case: 0117915495 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Gao Luqi Last updated by: Gao Luqi 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ ----------------------------------------------------------------------------- | 3 |  | Not clear | [12 empty] | 117915495 |  | 1134208 | TRUE | Evaluation | Evaluation | [7 empty] | TBD | [6 empty] | 4439700% | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | n | y | y | y | y | #N/A
1175 |  |  | China | The Second Hospital of Haibei Prefecture in Qinghai Province | 345029 | 4439800% |  | 02n/Ingenuity Core128/Elite China/Add printer | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Configuration Error Repair Activity Code: Instructed Customer Problem Reported by Customer: Add printer Safety Questions:Case: 0117921776 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities: Problem Description: add printer Troubleshooting Action: add printer Repair Action: instruct customer add printer.system is ok. Test/Inspection: Functional Check PASSED. The device returned to full functionality. Test Equipment: N/A ----------------------------------------------------------------------------- | Lu, Xueying |  | 4439800% | [12 empty] | 117921776 |  | 1138854 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | 2021-07-22 16:40 | [6 empty] | 4439800% | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1176 |  |  | China | 陕西省安康中医医院 | 345041 | 4439600% |  | WeChat:02N/Ingenuity Core128/Elite China/Host crash frequently | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : WeChat:02N/Ingenuity Core128/Elite China/Host crash frequently Safety Questions : Case: 0117908321 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Cao Last updated by: Liu Cao 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 其他类型故障。 客户问题详细描述或故障现象： 软件死机 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 软件死机 远程诊断与测试结果： 软件问题 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 重装系统 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 21.07.2021 15:11:12 UTC Yujie Guo (86038990)服务内容：扫描软件间歇性卡顿，重装系统恢复。 服务结果：服务后，设备正常工作。 ----------------------------------------------------------------------------- | 3 |  | 服务内容：扫描软件间歇性卡顿，重装系统恢复。 服务结果：服务后，设备正常工作。 Check if can get bugreport? | [12 empty] | 117908321 |  | 1128823 | TRUE | Closed - Done | Closed - Done |  | TBD-SW-Slow | [5 empty] | Engineering | [6 empty] | SW slow | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Performance | Console-SW-Performance | Console-SW-Performance | TBD-SW-Slow | Performance | Console-SW | Console-SW-Performance | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1177 |  |  | China | Jiangsu Xuzhou Medical School Affiliated No. 3 Hospital | 40039 | 4439600% |  | 02G/Ingenuity CT 728326/The bed makes a strange noise | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer: The bed makes a strange noise Safety Questions:Case: 0117908206 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Xianyi Last updated by: Wang Xianyi 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer's problems: The device makes a strange sound. Detailed description of Customer's problems or malfunctions: The bed has a strange sound It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: The bed has abnormal sound Remote troubleshoot and test result: The bed has abnormal sound and needs to be checked on the field PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Check the noises in the bed. Look for sound source and troubleshoot Next Step Action: contacted FSE Zhang Chuanxi to confirm, add Operation Next Step Action: mailed IB, add Operation to FSE Zhang Chuanxi Next Step Action : OP added 0030 ----------------------------------------------------------------------------- | 3 |  | Repair Action: Suggest onsite service procedures: Check the noises in the bed. Look for sound | [12 empty] | 117908206 |  | 1128707 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | 4439600% | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | n | y | y | y | y | #N/A
1178 |  |  | China | Shanxi YangGao County People's Hospital | 345025 | 4439800% |  | 02N/Ingenuity Core128/Elite China/Print problem | BU Q&R Feedback: No SAP Customer Symptom Code: CENV Customer Symptom Code: Customer Environmental Concern Reason: Corrective Maintenance Engineer Failure Code: Environmental Cause Repair Activity Code: Instructed Customer Problem Reported by Customer: 02N/Ingenuity Core128/Elite China Safety Questions:Case: 0117920393 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Weng Jinming Last updated by: Weng Jinming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Troubleshooting Action: answer customer question Repair Action: answer customer question Test/Inspection: Functional Check PASSED. The device returned to full functionality. Test Equipment: N/A SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | Lu, Xueying |  | 4439800% | [12 empty] | 117920393 |  | 1137889 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | 2021-07-22 10:20 | [6 empty] | 4439800% | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1179 |  |  | Germany | Gesundheit Nord gGmbH Klinikverbund | 40002 | 4439900% |  | Verzögerte Bildübertragung ins PACS | Customer comment: 0034W000028wmrPQAQ Customer comment: Rs with Mr. Halama, T for 2h ## Appointment with Mr. Glogner-Umbeer for Jul 28, 1:30 p.m. ## -------------------------------------------------- --------------------------- Work Order #: WO-03921170 Halama, Thomas Diagnostic performed by Engineer: N / A Problem Description by Engineer: Radiologists got the impression that the images were from the end of the scan take longer than usual to reach the Pacs. -------------------------------------------------- --------------------------- Work Order #: WO-03912460 Zurstegge, Uwe The system meets the specification for the performed service and is returned to use Diagnostic performed by Engineer: phone call with Raik Glogner-Umbeer Medizintechnik. House IT Mr. Gloistein has commissioned the Pacs company Ashvin analyze the Dicomtransfer CT-Pacs. Follow Up Required: Check Cir's Recon Times, Delete lan spooler. Contacting the Pacs company Ashvin, whether log files regarding Transfer Ingenuity Pacs have already been evaluated. Telephone number Haus IT H. Gloistein 0421/49781181 Problem Description by Engineer: Radiologists got the impression that the images were from the end of the scan take longer than usual to reach the Pacs. | 4 |  | Delete lan spooler. Contacting the Pacs company Ashvin, whether log files regarding Transfer Ingenuity Pacs have already been evaluated. | [12 empty] | 117949498 |  | 1142449 | TRUE | Evaluation - Pending Translation | Evaluation - Pending Translation | [7 empty] | TBD | [6 empty] | 4439900% | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1180 |  |  | Japan | Social Medical Corporation BOKOI, Nikko Memorial Hospital | 341011 | 4439800% |  | 02DS:KK: "Tape switch pressed" appeared and scanning is not possible.>> Releasing the tape switch solved the [Truncated] | BU Q&R Feedback: No Reason: Corrective Maintenance T2 Activities:Problem Description: Problem description by engineer : Equipment affected: CT Failure location: 02Couch When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user's work flow did the problem occur?: (Pre-Acquisition/Scan) What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem? No.The customer has not created a log file or report. Troubleshooting Action: What steps are taken during the system diagnosis: Result of response: Problem solved. Failure location: 02Couch What steps did you perform to diagnose the problem? It seems the tape switch was not released and a scan did not run.The customer said that after he operated the tape switch a number of times to release it, operation returned to normal.An FSE will respond. Error # and/or description of error: NA Repair Action: A service engineer visited the site. OneEMS internal remarks : Problem description by engineer : Equipment affected: CT Failure location: 02Couch When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user's work flow did the problem occur?: (Pre-Acquisition/Scan) What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem? No.The customer has not created a log file or report. What steps are taken during the system diagnosis: Result of response: Problem solved. Failure location: 02Couch What steps did you perform to diagnose the problem? It seems the tape switch was not released and a scan did not run.The customer said that after he operated the tape switch a number of times to release it, operation returned to normal.An FSE will respond. Error # and/or description of error: NA Repair Action "External" : The service engineer for the site visited and worked on the equipment. Repair Action "Internal" : --Troubleshooting Action-- Result of response: Problem solved. Failure location: 02Couch What steps did you perform to diagnose the problem? It seems the tape switch was not released and a scan did not run.The customer said that after he operated the tape switch a number of times to release it, operation returned to normal.An FSE will respond. --Repair Action-- Recommended 12Nc: --None-- 459800623252:FRU, CBL, HANDLE TAPE SWITCH QTY:2 SLOC:--None-- --Subject-- 02DS:KK:FSE Hashimoto: "Tape switch pressed" appeared and scanning is not possible.>> Releasing the tape switch solved the problem.>> An FSE will respond. --Problem Description-- Equipment affected: CT Failure location: 02Couch When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user's work flow did the problem occur?: (Pre-Acquisition/Scan) What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem? No.The customer has not created a log file or report. Problem Reported by Customer : 02DS:KK: "Tape switch pressed" appeared and scanning is not possible.>> Releasing the tape switch solved the problem.>> An FSE will respond. SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Support Safety Questions:Case: 0117920858 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: KENICHI KOJIMA Last updated by: KENICHI KOJIMA 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Message: Components Added Succesfully to SWO!- SAP Message: LOGI Task added Successfully!- ----------------------------------------------------------------------------- | 4 |  | It seems the tape switch was not released and a scan did not run.The customer said that after he operated the tape switch a number of times to release it, operation returned to normal.An FSE will respond. | [12 empty] | 117920858 |  | 1138182 | TRUE | Evaluation | Evaluation | [7 empty] | TBD | [6 empty] | Tape switch | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1181 |  |  | Japan | Social Medical Corporation BOKOI, Nikko Memorial Hospital | 341011 | 4439800% |  | 01CS:KK: The Exam Summary was not created for a particular examination, and it was also no automatically transferred. | BU Q&R Feedback: No SAP Customer Symptom Code: HANG Customer Symptom Code: Hang/Crash Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: Reboot/Restart Problem Reported by Customer: 01CS:KK: The Exam Summary was not created for a particular examination, and it was also no automatically transferred. T2 Activities:Problem Description: Problem description by engineer : Equipment affected: CT Problem location: 01 Console When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user's work flow did the problem occur?: Post-Acquisition/Scan What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem? No.The customer has not created a log file or report. Troubleshooting Action: What steps are taken during the system diagnosis: Result of response: Problem solved. Problem location: 01 Console What steps did you perform to diagnose the problem? Could not find any particular error.Data transfer was restored after restarting the host.I had the FSE contact the customer and have them take a wait-and-see approach. Error # and/or description of error: NA Repair Action: [Confirmation of problem] Confirmed. Work details: Could not find any particular error.Data transfer was restored after restarting the host.I had the FSE contact the customer and have them take a wait-and-see approach. 【Confirmation content】 Good. OneEMS internal remarks : Problem description by engineer : Equipment affected: CT Problem location: 01 Console When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user's work flow did the problem occur?: Post-Acquisition/Scan What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem? No.The customer has not created a log file or report. What steps are taken during the system diagnosis: Result of response: Problem solved. Problem location: 01 Console What steps did you perform to diagnose the problem? Could not find any particular error.Data transfer was restored after restarting the host.I had the FSE contact the customer and have them take a wait-and-see approach. Error # and/or description of error: NA Repair Action "External" : [Confirmation of problem] Confirmed. Work details: Could not find any particular error.Data transfer was restored after restarting the host.I had the FSE contact the customer and have them take a wait-and-see approach. 【Confirmation content】 Good. Repair Action "Internal" : --Troubleshooting Action-- Result of response: Problem solved. Problem location: 01 Console What steps did you perform to diagnose the problem? Could not find any particular error.Data transfer was restored after restarting the host.I had the FSE contact the customer and have them take a wait-and-see approach. --Repair Action-- --Subject-- 01CS:KK: Mr. Nishiya: The Exam Summary was not created for a particular examination, and it was also no automatically transferred. --Problem Description-- Equipment affected: CT Problem location: 01 Console When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user's work flow did the problem occur?: Post-Acquisition/Scan What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem? No.The customer has not created a log file or report. --4/8Elements-- Q1. Was the reported problem confirmed? Confirmed. Q2. Were any other problems confirmed? None. Q3. What did you do to resolve the problem?: Could not find any particular error.Data transfer was restored after restarting the host.I had the FSE contact the customer and have them take a wait-and-see approach. Q4: Is the equipment operating normally following the response? Good. Q5. How did the examination finish? None. Q6. Were there any mistaken findings (misdiagnosis, mistaken treatment) due to the problem? None. Q7. Was there any unintended unnecessary exposure? None. Q8. Were there any artifacts? None. Safety Questions:Case: 0117920389 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: KENICHI KOJIMA Last updated by: KENICHI KOJIMA 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 4 |  | Could not find any particular error.Data transfer was restored after restarting the host.I had the FSE contact the customer and have them take a wait-and-see approach. Monitoring and collect bugreport/problem details for investigation. | [12 empty] | 117920389 |  | 1137888 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | TBD | [6 empty] | 4439800% | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1182 |  |  | Japan | Maki Hospital | 341017 | 4439600% |  | 01DM:TM: Data transfer from the CT-Host has stopped.The data is remaining in QM>> Restarting the CT-Host sol [Truncated] | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 01DS:TM: Image transfer stopped>> SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Support Safety Questions:Case: 0117907956 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: TAKEYUKI MIYAKAWA Last updated by: TAKEYUKI MIYAKAWA 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Problem description by engineer : Equipment affected: CT Problem location: 01 Console When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user's work flow did the problem occur?: Network What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem?: NA Engineer's Comment: '============================== $version = "v4.1.10.25210" 10:46 Call description: Data transfer is stopped. Restarting the CT-Host did not change the situation. He called us when he was restarting the system after he shut down the CT host. '============================== Troubleshooting Action: What steps are taken during the system diagnosis: Result of response: Problem solved. Problem location: 01 Console What steps did you perform to diagnose the problem? '================================== Result Transfer became possible after restarting. However, the same problem occurs frequently, which interferes with work. '================================== Error # and/or description of error: NA Repair Action: A service engineer visited the site. OneEMS internal remarks : Problem description by engineer : Equipment affected: CT Problem location: 01 Console When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user's work flow did the problem occur?: Network What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem?: NA Engineer's Comment: '============================== $version = "v4.1.10.25210" 10:46 Call description: Data transfer is stopped. Restarting the CT-Host did not change the situation. He called us when he was restarting the system after he shut down the CT host. '============================== What steps are taken during the system diagnosis: Result of response: Problem solved. Problem location: 01 Console What steps did you perform to diagnose the problem? '================================== Result Transfer became possible after restarting. However, the same problem occurs frequently, which interferes with work. '================================== Error # and/or description of error: NA Repair Action "External" : The service engineer for the site visited and worked on the equipment. Repair Action "Internal" : --Troubleshooting Action-- Result of response: Problem solved. Problem location: 01 Console What steps did you perform to diagnose the problem? '================================== Result Transfer became possible after restarting. However, the same problem occurs frequently, which interferes with work. '================================== --Repair Action-- --Subject-- 01DM:TM:Mr. Hoshi: Data transfer from the CT-Host has stopped.The data is remaining in QM>> Restarting the CT-Host solved the problem. --Problem Description-- Equipment affected: CT Problem location: 01 Console When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user's work flow did the problem occur?: Network What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem?: NA Engineer's Comment: '============================== $version = "v4.1.10.25210" 10:46 Call description: Data transfer is stopped. Restarting the CT-Host did not change the situation. He called us when he was restarting the system after he shut down the CT host. '============================== ----------------------------------------------------------------------------- | 4 |  | Data transfer from the CT-Host has stopped.The data is remaining in QM>> Restarting the CT-Host solved the problem. v4.1.10.25210 Monitoring and collect bugreport/problem details for investigation. | [12 empty] | 117907956 |  | 1128573 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | TBD | [6 empty] | 4439600% | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1183 |  |  | Poland | Government Agency. | 341009 | 4439800% |  | Starting and implementing it for use. | Integration:VBAKKOM has been processed successfully-VBAPKOM has been processed successfully-KONVKOM has been processed successfully-KONVKOM has been processed successfully-KONVKOM has been processed successfully-KONVKOM has been processed successfully-REQ2RES DMR 0018133501 has been saved-MPG is present on the product code : P7 Integration: Confirmation of order 53913979 saved. ----------------------------------------------------------------------------- Work Order #: WO-03889096 21-Jul-2021 Marcin Knyps The system meets the specification for the performed service and is returned to use Problem Description by Engineer: Connection of a modular CT laboratory. Resolution: Powering on in the module CT laboratory. System calibration, Constancy and QuickIQ image quality test in accordance with the manufacturer's specifications. Due to the lack of a configured computer network, the DICOM settings for external systems have not been entered. The device is technically ready for use. Tools & Equipment : MULTIMETER,2022-04-14 00:00:00 ----------------------------------------------------------------------------- | 4 |  | Powering on in the module CT laboratory. System calibration, Constancy and QuickIQ image quality test in accordance with the manufacturer's specifications. Due to the lack of a configured computer network, the DICOM settings for external systems have not been entered. The device is technically ready for use. | [12 empty] | 117922996 |  | 1135626 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | 2021-07-21 21:50 | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1184 |  |  | Portugal | Centro Hospitalar de Vila Nova | 40030 | 4439900% |  | Password access to Exam Card Manager protocols has expired |  | 3 |  | Password access to Exam Card Manager protocols has expired | [12 empty] | 117950281 |  | 1143936 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Password | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | n | y | y | y | y | #N/A
1185 |  |  | Puerto Rico | DORADO HEALTH INC DBA | 40007 | 4440000% |  | No esta realizando las reconstrucciones sagitales completas | Work Order #: WO-03923749 Valenzuela Velasquez, Eduardo Ariel Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 3 |  | Not clear | [12 empty] | 117956108 |  | 1147519 | TRUE | Evaluation | Evaluation | [7 empty] | TBD | [6 empty] | 4440000% | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1186 |  |  | Puerto Rico | HOSPITAL DR SUSONI INC | 40013 | 4439600% |  | the gantry at 0 has angulation, it must be given 2.5 degrees | Work Order #: WO-03878989 23-Jul-2021 Rios, William The system meets the specification for the performed service and is returned to use Resolution : Verify system, found that Gantry was tilt to the front about 3.5 grade, calibrate the Tilt and the problem was fixed, make functional test. Unit OK Problem Description by Engineer : Problem description by engineer : tilt of gantry abnormal When was the first occurrence? Within... : 19-Jul-2021 11:00:00 How often is issue occurring? : Multiple times a day Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: tilt of gantry visually incorrect Patient Impact/Outcome: n/a User Impact: they see tilt is not correct Current Software Version: unknown ----------------------------------------------------------------------------- Work Order #: WO-03875700 Caldelas Aldana, Carlos Manuel The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Problem description by engineer : tilt of gantry abnormal When was the first occurrence? Within... : 19-Jul-2021 11:00:00 How often is issue occurring? : Multiple times a day Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: tilt of gantry visually incorrect Patient Impact/Outcome: n/a User Impact: they see tilt is not correct Current Software Version: unknown Diagnostic performed by Engineer : they visually check that gantry is not correctly in tilt, they chck and when its supposed to be at 0 degrees it has in screen more than 2. and when system is at 0, they gantry looks tilted. No errors io system. Internal Comments : Error # and/or description of error : tilt in gantry Repair Action : n/a Resolution : FSE visit Follow Up Required : Error # and/or description of error : tilt in gantry Repair Action : n/a What are the next Troubleshooting steps : check gantry covers and check gantry tilt Technician Name : Caldelas Aldana, Carlos Manuel ----------------------------------------------------------------------------- | 3 |  | Resolution Description: Verify system, found that Gantry was tilt to the front about 3.5 grade, calibrate the Tilt and the problem was fixed, make functional test. Unit OK | [12 empty] | 117911037 |  | 1128075 | TRUE | Evaluation | Evaluation | [7 empty] | Others | [6 empty] | Tilt Calibration | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | n | y | y | y | y | #N/A
1187 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 337000 | 4439700% |  | There are problems in images; Cleaned the opaque substances and performed air calibration. | Integration: Confirmation of order 53907744 has been saved. ----------------------------------------------------------------------------- Work Order #: WO-03881271 20-Tem-2021 Can Ozgenlik The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : Opaque substance leaked inside the device. The opaque substances on the detector sleeve were cleaned and air calibration was completed. Test scan with Phantom was completed successfully. The device was delivered in working condition. ----------------------------------------------------------------------------- | 3 |  | Contrast intrusion • Improved Lexan ring 459801485251 is on stock, need check if Turkey site having contrast intrusion issue because of lexan ring gap and recommend to replace accordingly -- Andy | [12 empty] | 117916099 |  | 1130919 | TRUE | Closed - Done | Closed - Done |  | ECR-075975 | [5 empty] | Engineering | [6 empty] | Lexan Ring | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Lexan Ring | Gantry-HW-Lexan Ring | Gantry-HW-Lexan Ring | ECR-075975 | Lexan Ring | Gantry-HW | Gantry-HW-Lexan Ring | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | n | y | y | 0 | y | #N/A
1188 |  |  | United States | Doctors Hospital LLC | 40014 | 4439900% |  | System Removal | Integration : Confirmation of order 53919742 saved ----------------------------------------------------------------------------- Work Order #: WO-03917663 22-Jul-2021 Lynn Matthew The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: n/a Device Use: n/a Expected and Unexpected Results: n/a User Impact: n/a Patient Impact: n/a Current Software Version: n/a Internal Comments : THis work order/case is for a life solutions quote, #LS - 070005 Customer has provided a cashiers check in the amount of $24197.16, per the life solutions quote. Diagnostic performed by Engineer : Customer requested entire system removal. A life solutions quote was generated and accepted by our customer Resolution : Assisted Philips installation/deinstallation personnel with disassembly and removal of CT system. ----------------------------------------------------------------------------- | 5 |  | Assisted Philips installation/deinstallation personnel with disassembly and removal of CT system. | [12 empty] | 117950736 |  | 1145559 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | 2021-07-24 04:25 | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1189 |  |  | United States | Odessa Regional Medical Center | 40029 | 4439800% |  | MPR formats | Integration : Confirmation of order 53916107 saved ----------------------------------------------------------------------------- Work Order #: WO-03891810 Curry, Dyan The system fully meets specification and returned to use Problem description by Clinical Support Specialist : Customer Function/ Role: technologist Device Use: in use name Expected and Unexpected Results; help creating cor and sag of crooked elbow User Impact: none Patient Impact: none Current Software version: 4.1.7 Diagnostic performed by Engineer : helped customer use planar to straighten and create sag and cor mprs of twisted elbow. ----------------------------------------------------------------------------- | 5 |  | helped customer use planar to straighten and create sag and cor mprs of twisted elbow. | [12 empty] | 117925122 |  | 1137532 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | 2021-07-22 06:35 | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1190 |  |  | United States | Philips Medical Capital | 40028 | 4439600% |  | ACR questions | Integration : Confirmation of order 53904475 saved ----------------------------------------------------------------------------- Work Order #: WO-03878577 Giles Brian The system meets the specification for the performed service and is returned to use Problem description by Clinical Support Specialist : Customer Function/Role: TECH Device Use: In Use Expected and Unexpected Results: iDose question User Impact: None Patient Impact: None Current Software Version: 4.1.7 Diagnostic performed by Engineer : Customer wanted to know if iDose saved dose to the patient. It does and should be used. ----------------------------------------------------------------------------- Work Order #: WO-03877549 Mujica, David The system meets the specification for the performed service and is returned to use Problem description by Clinical Support Specialist : Customer Function/Role: Technologist Device Use: In use Expected and Unexpected Results: ACR questions User Impact: none Patient Impact: none Current Software Version: 4.1 Diagnostic performed by Engineer : Iris needed help gathering information for the ACR on a Head CT on their Ingenuity CT system. Number of channel, slice, increment, rotation and modulation explored. All of Iris's questions were answered and the information was gathered for ACR on their Head CT exam card. Iris had no further questions or need for clinical support. No rescan, no loss of data, no reinjection per Iris. ----------------------------------------------------------------------------- | 5 |  | 4439600% | [12 empty] | 117912371 |  | 1127499 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | 2021-07-20 01:15 | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1191 |  |  | Chile | BANCO DE CHILE | 341005 | 4440700% |  | CL1688-THE GANTRY WENT OFF TWICE THIS WEEK | Expected and Unexpected results: the gantry is turned off Patient Impact / Outcome: study aborat User Impact: the gantry is turned off when in use Current Software Version: Not provided Diagnostic performed by Engineer: they comment that gantry opens estop when being used in the study, now the equipment is operational at the time of the remote support call. The equipment is accessed remotely to inspect logs from the last 3 days without finding any record or reason for failure. I went to review together with the user team to see the reason for failure to open the estop. Check electrical supply is not having variations. SANDOVAL IS COMMUNICATED THAT NOBODY CAME ON FRIDAY TO THE SITE, FOLLOW-UP MAIL IS SENT | 3 |  | Next step: review together with the user team to see the reason for failure to open the estop. Check electrical supply is not having variations. | [12 empty] | 118006346 | In Progress | 1169625 | TRUE | Opened - Pending Translation | Opened - Pending Translation | [7 empty] | TBD | [6 empty] | Gantry off | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1192 |  |  | China | 东莞市中西医结合医院 | 345013 | 4440900% |  | 02N/Ingenuity Core128/Elite China/nemo cable | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity Core128/Elite China Safety Questions : Case: 0118018928 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: nemo part Troubleshooting Action: FSE call in nemo DP to DVI cable. Repair Action: nemo DP to DVI cable. Test/Inspection: N/A Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | FSE call in nemo DP to DVI cable | [12 empty] | 118018928 | In Process | 1175422 | TRUE | Opened - Pending Translation | Opened - Pending Translation |  | CTUDT00632173 | [5 empty] | Engineering | [6 empty] | DVI Adaptor | [5 empty] | Refresh |  |  | System issue | System | Console-HW | Adaptor | Console-HW-Adaptor | Console-HW-Adaptor | CTUDT00632173 | Adaptor | Console-HW | Console-HW-Adaptor | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1193 |  |  | China | 河南省新乡市凤泉区大块镇卫生院 | 345009 | 4440800% | 30 | 02n/Ingenuity Core128/Elite China/There is a problem with the program | BU Q&R Feedback: No SAP Customer Symptom Code: BOOT Customer Symptom Code: Boot up/Power-up Problem Reason: Corrective Maintenance Engineer Failure Code: Software Defect or Virus Repair Activity Code: Instructed Customer Problem Reported by Customer : There is a problem with the program Next Step Action : GUO GAIBIN call for,need RSE to deal Priority Safety Questions : Case: 0118015051 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Lili Last updated by: Lu Lili T2 Activities : Problem Description: CT application can not login Troubleshooting Action: CT application can not login, indicate need to create bugrep, instruct customer use olevel mode to login, it's ok, then use CT mode login , this time can login application successfully. system is ok now. Repair Action: instruct customer use olevel mode to login, it's ok, then use CT mode login , this time can login application successfully. system is ok now. Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: N/A SAP Transfer : Case Transferred to SAP | 2 |  | CT application can not login, indicate need to create bugrep, instruct customer use olevel mode to login, it's ok, then use CT mode login , this time can login application successfully. system is ok now. Any screenshot and bugreport for the issue? Check Blue screen defect and if related to this issue? | [12 empty] | 118015051 | Fixed | 1172857 | TRUE | Opened - Pending Translation | Opened - Pending Translation |  | CTUDT00632553 | [5 empty] | Engineering | [6 empty] | Login failed | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Blue-Screen | Console-SW-Blue-Screen | Console-SW-Blue-Screen | CTUDT00632553 | Blue-Screen | Console-SW | Console-SW-Blue-Screen | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1194 |  |  | China | 山东省胸科医院 | 40025 | 4440800% |  | 02N/Ingenuity CT 728326/estop open | BU Q&R Feedback: No SAP Customer Symptom Code: BOOT Customer Symptom Code: Boot up/Power-up Problem Reason: Corrective Maintenance Engineer Failure Code: Electrical/Electronic Failure Repair Activity Code: Reboot/Restart Problem Reported by Customer : 02N/Ingenuity CT 728326 Safety Questions : Case: 0118009760 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang T2 Activities : Problem Description: estop open Troubleshooting Action: 1,check error logger, velocity encoder error cause the estop open. 2,instruct customer power off gantry power, then power on ,system is ok. Repair Action: reboot gantry power ,system is ok. Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | 1,check error logger, velocity encoder error cause the estop open. 2,instruct customer power off gantry power, then power on ,system is ok. Repair Action: reboot gantry power ,system is ok. | [12 empty] | 118009760 | Closed | 1172131 | TRUE | Opened - Pending Translation | Opened - Pending Translation |  | TBD-Encoder | [5 empty] | Engineering | [6 empty] | velocity encoder | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | VelocityEncoder | Couch-HW-VelocityEncoder | Couch-HW-VelocityEncoder |  |  | Couch-HW | Couch-HW-VelocityEncoder | No | 0 | 0 | 0 | 0 | 0 | No | y | y | y | 0 | 0 | y | y | y | 0 | y | #N/A
1195 |  |  | China | 河南省新乡市凤泉区大块镇卫生院 | 345009 | 4440800% | 30 | 02N/Ingenuity Core128/Elite China/Lifting fault | T2 Activities : Problem Description: 客户问题概述： 其他类型故障。 客户问题详细描述或故障现象： 床降到最低位后升不起来了 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 床降到最低位后升不起来了 远程诊断与测试结果： 怀疑压到限位开关了 PRS TR状态： 不通，请工程师到场检查。 Repair Action: 建议现场维修步骤： 维修模式，把床升起来，重做垂直校正。 External Remarks : 31.07.2021 10:52:11 UTC Haibin Guo (86038308)"SUAT. Ready for Clinical Use local time" : "2021-07-31 16:00" Internal Remarks : 01.08.2021 05:46:24 UTC Haibin Guo (86038308) RDI Accuracy: YES Q7:Confirmed report problem，Lifting fault $Q8:perform visiual check on site and conform the couch can not lift up. when switch to service mode, then the couch can move up normally. then we switch to normal mode, the couch can lift up normaly too $Q9:move up the couch in service mode and do vertical calibration $Q10: test and inspection data refer to PA. Device returned to full functionality. $Q11:Calibrated Tools N/A External Remarks : 01.08.2021 05:46:18 UTC SUP JCOUSER (86038308)故障描述：床在最低位置升不起来 服务内容：使用服务模式升床并重新完成垂直校正 服务结果：服务完成后，设备恢复正常使用 其他事项：无 SAP Transfer : Dispatched for Onsite Support | 2 |  | perform visiual check on site and conform the couch can not lift up. when switch to service mode, then the couch can move up normally. then we switch to normal mode, the couch can lift up normaly too move up the couch in service mode and do vertical calibration Any bugreport and detail for this issue? | [12 empty] | 118008848 | Closed | 1172026 | TRUE | Opened - Pending Translation | Opened - Pending Translation |  | TBD-bugrep | [5 empty] | Engineering | [6 empty] | Couch fault | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | TBD-bugrep | Couch-HW-TBD-bugrep | Couch-HW-TBD-bugrep |  |  | Couch-HW | Couch-HW-TBD-bugrep | No | 0 | 0 | 0 | 0 | 0 | No | y | y | y | 0 | 0 | y | y | y | 0 | y | #N/A
1196 |  |  | China | 东莞市中西医结合医院 | 345013 | 4440700% |  | wechat:02n/Ingenuity Core128/Elite China/system error | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : system error Safety Questions : Case: 0118000001 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Lili Last updated by: Lu Lili T2 Activities : Problem Description: 客户问题概述： ECG相关故障。 客户问题详细描述或故障现象： 心电无波形信号。 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： log无报错。 远程诊断与测试结果： 重启无效。 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议现场检查MP5到机架接线，用MP5模拟波形测试MP5是否信号能到IBOX板，如果能到检查firmware IBOX重新尝试连接。 建议备件12NC及名称： N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : li zongxian ----------------------------------------------------------------------------- | 3 |  | 心电无波形信号。建议现场检查MP5到机架接线，用MP5模拟波形测试MP5是否信号能到IBOX板，如果能到检查firmware IBOX重新尝试连接。 Check if USB cable have 5vdc power cable, or Bizlink USB cable. | [12 empty] | 118000001 | In Process | 1171215 | TRUE | Opened - Pending Translation | Opened - Pending Translation | [7 empty] | Others | [6 empty] | ECG | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | n | y | y | y | y | #N/A
1197 |  |  | China | 苏州广慈肿瘤医院有限公司 | 345019 | 4440700% |  | 02N/Ingenuity Core128/Elite China | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity Core128/Elite China Next Step Action : FSE xu haotian 018761995321来电报修设备 Safety Questions : Case: 0117998855 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 其他类型故障。 客户问题详细描述或故障现象： 扫描中断 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： S_MASTER_SEQUENCER_TIMER_SETTIME_ERR 远程诊断与测试结果： DMC，RHOST，Ghost通讯故障 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 检查滑环，信号刷，检查高压之间的连线 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : Xu Haotian External Remarks : 01.08.2021 09:22:02 UTC Haotian Xu (86090280) 服务内容：做增强检测时报错，报错代码是sequence_timer_setting_error，检查 hss上的接口，收集bugrep 服务结果：服务后，设备恢复正常 ----------------------------------------------------------------------------- | 3 |  | 远程诊断与测试结果： DMC，RHOST，Ghost通讯故障 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 检查滑环，信号刷，检查高压之间的连线 做增强检测时报错，报错代码是sequence_timer_setting_error，检查 hss上的接口，收集bugrep | [12 empty] | 117998855 | In Process | 1170882 | TRUE | Opened - Pending Translation | Opened - Pending Translation | [7 empty] | TBD | [6 empty] | Data path | [5 empty] | Refresh |  |  | System issue |  |  |  | - | - |  |  | 0 | - | No | y | y | y | 0 | 0 | No | y | y | y | y | 0 | n | y | y | y | y | #N/A
1198 |  |  | China | 息烽县中医医院 | 345014 | 4440400% |  | Wechat:02N/Ingenuity Core128/Elite China/ACQ error | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Replace Part Problem Reported by Customer : ACQ error Safety Questions : Case: 0117970209 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Gaoge Last updated by: Wang Gaoge 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 扫描失败。 客户问题详细描述或故障现象： 扫描时报错gantry acquisition system cannot comply 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： S_MASTER_SEQUENCER_ACQ_NOT_OK 远程诊断与测试结果： CIRS日志中显示光路重发数量为0 ，报错时间有ERROR - Invalid DFS setting ，考虑R2D频率导致采集异常 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 1，检查R2D的频率和幅值 ，偶发报错，需要仔细诊断和排查 建议备件12NC及名称： "R2D Assembly 459800467343" Test/Inspection: N/A Test Equipment: N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : PPD 459800467341*1 已确认替代号可出24h到医院 FSE:唐广林-86-18084189777 External Remarks : 29.07.2021 10:03:50 UTC SUP JCOUSER (86029494)服务内容：确认报修故障，扫描中断。现场检查，发现R2D故障，更换R2D后设备恢复正常。 服务结果：维修后设备恢复正常。 ----------------------------------------------------------------------------- | 3 |  | 服务内容：确认报修故障，扫描中断。现场检查，发现R2D故障，更换R2D后设备恢复正常。 William: check if new 459800467344 can improve this issue? | [12 empty] | 117970209 | In Process | 1157518 | TRUE | Opened - Pending Translation | Opened - Pending Translation |  | CTUDT00630864 | [5 empty] | Engineering | [6 empty] | R2D | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | R2D | Gantry-HW-R2D | Gantry-HW-R2D | CTUDT00630864 | R2D | Gantry-HW | Gantry-HW-R2D | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1199 |  |  | China | 哈尔滨市森海医院 | 345007 | 4440300% |  | 02N/Ingenuity Core128/Elite China/printer issue | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ Safety Questions : Case: 0117961156 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Cao Last updated by: Liu Cao 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: Cannot print Troubleshooting Action: Print connection failed, temp network issue Repair Action: Temporary network issue , take monitoring Test/Inspection: N/A Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | Problem Description: Cannot print Troubleshooting Action: Print connection failed, temp network issue Repair Action: Temporary network issue , take monitoring | [12 empty] | 117961156 | Fixed | 1151904 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Network Connection | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | n | y | y | y | y | #N/A
1200 |  |  | China | Zhalantun City People's Hospital | 40045 | 4440300% |  | 02N/Ingenuity CT 728326/scan interrupt | BU Q&R Feedback: No SAP Customer Symptom Code: ACQP Customer Symptom Code: Imaging/Patient Recording/Acquisition Problem Reason: Corrective Maintenance Engineer Failure Code: Electrical/Electronic Failure Repair Activity Code: Reboot/Restart Problem Reported by Customer: 02N/Ingenuity CT 728326 Safety Questions:Case: 0117960989 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: scan interrupt Troubleshooting Action: check error logger, ACQ not ok, check eventlogger ,no logger found, this error only occurred once, so need monitor the system's status. Repair Action: instruct customer monitor the system status. now system is ok. Test/Inspection: Functional Check PASSED. The device returned to full functionality. Test Equipment: N/A SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | Troubleshooting Action: check error logger, ACQ not ok, check eventlogger ,no logger found, this error only occurred once, so need monitor the system's status. Repair Action: instruct customer monitor the system status. now system is ok. | [12 empty] | 117960989 | Fixed | 1151827 | TRUE | Investigation | Investigation |  | CTUDT00630408 | [5 empty] | Engineering | [6 empty] | ACQ | [5 empty] | Refresh |  |  | System issue | System | CIRS-SW | ACQ_NOT_OK | CIRS-SW-ACQ_NOT_OK | CIRS-SW-ACQ_NOT_OK | ACQ_NOT_OK | ACQ_NOT_OK | CIRS-SW | CIRS-SW-ACQ_NOT_OK | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | n | y | y | 0 | y | #N/A
1201 |  |  | Japan | 社会福祉法人恩賜財団済生会 | 343000 | 4440800% |  | 01CS: RS: Axial of the head is noisy when the coronal cross section is created by ISP | Canceled | 5 |  | Canceled | [12 empty] | 118016702 | New | 1173973 | TRUE | Opened | Opened | [7 empty] | Others | [6 empty] | Image Quality | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1202 |  |  | Japan | 医療法人清翠会 牧病院 | 341017 | 4440800% |  | SY: CT application support package | Canceled | 5 |  | Canceled | [12 empty] | 118016430 | Canceled | 1173747 | TRUE | Opened | Opened | [7 empty] | Application | [6 empty] | Application | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1203 |  |  | Japan | 社会医療法人 母恋 日鋼記念病院 | 341011 | 4440800% |  | 02DS: KK: Tape Switch pressed and cannot scan. >> Recovered by releasing Tape Switch. ＞＞ FSE compatible | 02DS: KK: Tape Switch pressed and cannot scan. >> Recovered by releasing Tape Switch. ＞＞ FSE compatible | 4 |  | Tape Switch | [12 empty] | 118016314 | New | 1173636 | TRUE | Opened | Opened | [7 empty] | Customer | [6 empty] | Tape Switch | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1204 |  |  | Japan | 社会福祉法人恩賜財団済生会 | 343000 | 4440800% |  | HY: Gantry belt replacement | HY: Gantry belt replacement | 5 |  | Belt | [12 empty] | 118014821 | New | 1172681 | TRUE | Opened | Opened |  | ECR-077012 | [5 empty] | Engineering | [6 empty] | Belt | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Belt | Gantry-HW-Belt | Gantry-HW-Belt | ECR-077012 | Belt | Gantry-HW | Gantry-HW-Belt | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1205 |  |  | Japan | 社会福祉法人恩賜財団済生会 | 343000 | 4440400% |  | 01CS: RS: The image of the coronal cross section of the Axial of the head created by ISP became noisy >> It seems that the image of the soft tissue condition and the bone condition were created together >> The bone condition was recalculated and resolved | 01CS: RS: Moriwaki-sama: The image of the Axial of the head created with the coronal cross section by ISP became noisy >> It seems that the image of the soft tissue condition and the bone condition were created together >> The bone condition is recalculated separately Eliminate --Problem Description-- Target device: CT Failure location: 01Cosole When did it first occur? (Within ~): 1 Hour How often do problems occur ?: First Occurrence At what point in the end-user workflow is the problem occurring ?: Other What happened recently ?: N.A. Does the customer create a specific system event report for the problem ?: N.A. --4 / 8Elements-- Q1 Did you confirm the reported problem? Verified Q2 Did you confirm other problems? None Q3 What did you do to solve the problem? It seems that it was created by combining the image of the soft tissue condition and the bone condition >> Bone condition is divided and recalculated and eliminated Is the device after Q4 compatible? Good Q5 How did the inspection end? None Q6 Presence or absence of incorrect findings (misdiagnosis, mistreatment) due to problems: None Q7 Presence or absence of unintended unnecessary exposure: None Q8 Presence or absence of artifacts: None 与此原文有关的更多信息要查看其他翻译信息，您必须输入相应原文 | 4 |  | It seems that it was created by combining the image of the soft tissue condition and the bone condition >> Bone condition is divided and recalculated and eliminated | [12 empty] | 117977319 | Fixed | 1161155 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | ShouldbeApplication | [5 empty] | Engineering | [6 empty] | Image Quality | [5 empty] | Refresh |  |  | System issue | System | NotEngineering | ShouldbeApplication | NotEngineering-ShouldbeApplication | NotEngineering-ShouldbeApplication | ShouldbeApplication |  | NotEngineering | NotEngineering-ShouldbeApplication | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | #N/A
1206 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 337000 | 4440500% |  | Mylar order. | "Work Order #: WO-03944507 Can Ozgenlik The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer: Mylar order. Resolution : The Mylar cover which connects the gantry rear door and front door will be replaced with a newer model. This will prevent liquid leak into the inside, such as blood, opaque substance, etc. Device is operational. Follow Up Required: Mylar replacement. | 3 |  | The Mylar cover which connects the gantry rear door and front door will be replaced with a newer model. This will prevent liquid leak into the inside, such as blood, opaque substance, etc. | [12 empty] | 117979182 | In Progress | 1157769 | TRUE | Evaluation | Evaluation |  | ECR-075975 | [5 empty] | Engineering | [6 empty] | Lexan ring | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Lexan Ring | Gantry-HW-Lexan Ring | Gantry-HW-Lexan Ring | ECR-075975 | Lexan Ring | Gantry-HW | Gantry-HW-Lexan Ring | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1207 |  |  | United States | Westgate Hospital | 40001 | 4440700% |  | Add A New Pacs Destination | Internal Remarks : Planner Initials: JJ Action Taken: FSE-Peter Leon called in to create a Case/WO and dispatch to him I created Case#0118006668/WO-03967361 and dispatched as requested Outcome: Dispatched Next Steps: N/A Last Updated: 1:54PM EDT 7/30/21 ----------------------------------------------------------------------------- Work Order #: WO-03967361 Leon Peter Diagnostic performed by Engineer : N/A Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: ----------------------------------------------------------------------------- | 5 |  | Add A New Pacs Destination | [12 empty] | 118006668 | In Progress | 1171663 | TRUE | NAC | NAC | [7 empty] | Customer | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1208 |  |  | Japan | 医療法人清翠会 牧病院 | 341017 | 4440400% |  | CT:APP:RK:質問:ﾎｼさま | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Application Support Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer : Q:"①たぶん、体動によるものであると思いますが画質がいつもと若干異なるので確認してください。 ②画像を転送をすると、転送先が前回の設定を覚えてしまうのですが覚えないようにできますか？" A:"①画像を拝見しました。やはり体動が原因で画像が変わっていました。 ②申し訳ありません。前回の転送先を覚えてしまう仕様になっておりまして変更できません。" OneEMS internal remarks : Q1:報告された問題は確認できましたか？：確認済 Q2：他の問題は確認できましたか？；無 Q3：問題を解決するために何を行いましたか？ １：電話で解決した場合：説明 Q4：対応後の装置は順調ですか？：良好 Safety Questions : Case: 0117972537 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: RYUSUKE KUJIRAI Last updated by: RYUSUKE KUJIRAI 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | Zhu, Chun |  | 4440400% | [12 empty] | 117972537 | Fixed | 1158116 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | 2021-07-28 18:39 | [6 empty] | 4440400% | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1209 |  |  | United States | Unity Health Systems | 40037 | 4440700% |  | Password incorrect. | Integration : Confirmation of order 53957845 saved ----------------------------------------------------------------------------- Work Order #: WO-03966342 30-Jul-2021 Lockwood Thomas The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : Reset IT password. Created updated password for master user to edit Exam cards ----------------------------------------------------------------------------- | 5 |  | Password | [12 empty] | 118004053 | Closed | 1169693 | TRUE | Opened | Opened | [7 empty] | Customer | [6 empty] | Password | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1210 |  |  | United States | Philips Medical Capital | 40028 | 4440700% |  | Split Study | Work Order #: WO-03967829 Mujica, David The system meets the specification for the performed service and is returned to use Problem description by Clinical Support Specialist : Customer Function/Role: Technologist Device Use: In use Expected and Unexpected Results: Split Accession numbers User Impact: none Patient Impact: none Current Software Version: 4.1 Diagnostic performed by Engineer : Iris indicated that her multi-study CT were not splitting to their appropriate accession numbers. Her Exam Card was accessed and Iris was shown where to make the association to change her Procedure in the Exam Card/Protocol to associate each accession number to the procedure description and save the Exam Card/Protocol so all accession number associated to each study accordingly. The EC was successfully changed and saved. Iris had no further questions or need for clinical support. No rescan, no loss of data, no reinjection per Iris. ----------------------------------------------------------------------------- | 4 |  | change her Procedure in the Exam Card/Protocol to associate each accession number to the procedure description and save the Exam Card/Protocol so all accession number associated to each study accordingly. The EC was successfully changed and saved. | [12 empty] | 118007049 | In Progress | 1169328 | TRUE | Opened | Opened | [7 empty] | Application | [6 empty] | Exam cards | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1211 |  |  | United States | Philips Medical Capital | 40053 | 4440500% |  | 0802CDY72 no ecg signal showing on host monitor | Internal Remarks : FADB: (FSE Aware of Double Booking) Title: Internal Remarks Initials: Twila Actions Taken: Called PFSE/McDowell Preston - He accepted WO for 7/29/21 @ 5hr Labor - No Travel Time I Dispatched WO for 1-6pm EST. Outcome: Dispatched per PFSE's Request IP: 88993323 Next Steps: None Internal Remarks : Assisted FSE with Flashing replacement iBox 2. Internal Remarks : FADB: (FSE Aware of Double Booking) Title: Internal Remarks Initials: Twila Actions Taken: Dispatched per WO FU Note - Recommended Service Engineer: McDowell Preston Follow Up Date/Time Requested: 8/2/2021 3:00 PM EST. Outcome: Dispatched per WO FU Note IP: 88993323 Next Steps: None ----------------------------------------------------------------------------- Work Order #: WO-03959714 29-Jul-2021 McDowell Preston Problem Description by Engineer : Problem description by engineer : no ecg signal showing on host monitor Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No further information is available as per the customer Expected and Unexpected results: no ecg signal showing on host monitor Patient Impact/Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer Current Software Version: Not provided Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- Work Order #: WO-03950575 29-Jul-2021 McDowell Preston The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Problem description by engineer : no ecg signal showing on host monitor Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No further information is available as per the customer Expected and Unexpected results: no ecg signal showing on host monitor Patient Impact/Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer Current Software Version: Not provided Follow Up Required : replaced patient monitor interface cable. Programmed the Ibox board and verified proper readings at the control desk . At this time, it is working, but will monitor the issue and await other part to arrive from overseas before case closure. Diagnostic performed by Engineer : replaced patient monitor interface cable. Programmed the Ibox board and verified proper readings at the control desk . At this time, it is working, but will monitor the issue and await other part to arrive from overseas before case closure. ----------------------------------------------------------------------------- Work Order #: WO-03959547 29-Jul-2021 McDowell Preston The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : Troubleshot with NSS as well. I box internal tests pass just fine, but it still shows ECG leads disconnected. Will order new cables for Thursday. Customer can utilize scanner for other patients without the patient monitor. ----------------------------------------------------------------------------- Work Order #: WO-03950493 Sheffield, Scott The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Problem description by engineer : no ecg signal showing on host monitor Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No further information is available as per the customer Expected and Unexpected results: no ecg signal showing on host monitor Patient Impact/Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer Current Software Version: Not provided Diagnostic performed by Engineer : Ongoing problem with the ecg waveform not showing on host monitor. Follow Up Required : Error # and/or description of error : none Repair Action : I spoke with primary FSE, he is already talking with NSS about this issue and FSE will be ordering parts himself. Technician Name : Sheffield, Scott Internal Comments : Error # and/or description of error : none Repair Action : I spoke with primary FSE, he is already talking with NSS about this issue and FSE will be ordering parts himself. ----------------------------------------------------------------------------- | 3 |  | replaced patient monitor interface cable. Programmed the Ibox board and verified proper readings at the control desk . At this time, it is working, but will monitor the issue and await other part to arrive from overseas before case closure. | [12 empty] | 117985264 | In Progress | 1160487 | TRUE | Evaluation | Evaluation | [7 empty] | Others | [6 empty] | ECG | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1212 |  |  | Japan | Maki Hospital | 341017 | 4440300% |  | SY: CT application support package. | BU Q&R Feedback: No Reason: Application Support Problem Reported by Customer: SY: CT application support package. Safety Questions:Case: 0117961832 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: YE Smile Last updated by: YE Smile 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A OneEMS internal remarks : Otani DB SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Support Internal Remarks : 29.07.2021 04:27:13 UTC NATSUKI OOTANI (81100565) 1. Are patients with COVID-19 infection or patients who are suspected of having COVID-19 infection being treated or quarantined ? Yes 2. Is it necessary to enter the treatment room or the bed of infected patients?(Including workplace and transfer routes) No External Remarks: 29.07.2021 04:27:10 UTC NATSUKI OOTANI (81100565) [Request details] A study group meeting was held about cardiac CT basics. Work details: I explained about cardiac CT using slides. ----------------------------------------------------------------------------- | Xu, Xiaoting |  | 4440300% | [12 empty] | 117961832 | New | 1152384 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | 2021-07-27 14:53 | [6 empty] | 4440300% | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1213 |  |  | United States | Unity Health Systems | 40037 | 4440500% |  | Monthly QA | Integration : Confirmation of order 53945608 saved ----------------------------------------------------------------------------- Work Order #: WO-03947431 28-Jul-2021 Taylor, Durrance The system meets the specification for the performed service and is returned to use Resolution : Monthly QA completed Problem Description by Engineer : Customer Function/Role: tech Device Use: scheduled activity Expected and Unexpected Results: Monthly QA User Impact: none Patient Impact: none Current Software Version: 4.1.7 | 5 |  | Monthly QA | [12 empty] | 117982135 | Closed | 1160277 | TRUE | NAC | NAC | [7 empty] | Customer | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1214 |  |  | United States | Ascension Seton Medical Center | 40006 | 4440400% | 30 | STC Failing Error during scan | Work Order #: WO-03940410 27-Jul-2021 Beato Tom The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: No Patient Involved Expected and Unexpected Results: STC Failed Error during scan User Impact: No Patient Impact: No Current Software Version: 4.1.10 Resolution : Attempted to perform short tube conditioning on system. Scan aborted. Performed complete shutdown of PCs and gantry. Restarted system. Attempted to scan but CIRS not initialized. Waited for CIRS to initialize. Able to scan after 10 minute wait. Performed short tube conditioning successfully. Analyzed bug report. Error occurred because gantry was not restarted in 80 plus days and scan was attempted without CIRS initialized. System is fully operational. ----------------------------------------------------------------------------- Work Order #: WO-03940262 Maier, Donnie The system does not meet the specification for the performed service and is not returned to use Diagnostic performed by Engineer : Cu stated they had already power cycled everything a few time. ULV shows error below before and after I talked Bonnie thru a complete power cycle of the system (Host/CIRS/Gantry) and still unable to perform STC. Unable to pull logs to troubleshoot further, dispatching FSE to analyze CIRS event logs, etc. Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: No Patient Involved Expected and Unexpected Results: STC Failed Error during scan User Impact: No Patient Impact: No Current Software Version: 4.1.10 Follow Up Required : Analyze logs locally as I am unable to pull them remotely. I already talked cu thru a complete power cycle (Host/CIRS/Gantry) but still getting the error during STC. Gantry hadn't been power cycled in over 80 days according to ULV. ----------------------------------------------------------------------------- | 2 |  | Error occurred because gantry was not restarted in 80 plus days and scan was attempted without CIRS initialized. System is fully operational. Instruct customer to restart whole system weekly | [12 empty] | 117975026 | Closed | 1154684 | TRUE | Evaluation | Evaluation | [7 empty] | Customer | [6 empty] | Reboot | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1215 |  |  | United States | Philips Medical Capital | 40053 | 4440300% |  | EKG disconnected intermittently, replaced I-Box | Integration : Confirmation of order 53932202 saved ----------------------------------------------------------------------------- Work Order #: WO-03934417 27-Jul-2021 McDowell Preston The system meets the specification for the performed service and is returned to use Resolution : THe patient monitor readings have been intermittent. Per NSS, replaced the Ibox board today and verified proper operation...though it was tested prior to replacement and found working as well. Will monitor system for issues, and inform customer. Problem Description by Engineer : Description of Event Still have intermittent EKG connectivity problem Ref Case 0117870146 CALLER FUNCTION/ROLE Technician/Sonographer/Therapist DEVICE INTENDED USE Diagnostic <1 Hour Intermittent WHERE IN THE END-USER WORKFLOW IS THE ISSUE OCCURRING? Pre-acquisition/scan WHAT RECENT EVENTS HAVE OCCURRED? During normal use NEGATIVE OPERATOR IMPACT N/A NEGATIVE PATIENT IMPACT/PROCEDURE OUTCOME N/A HAS CUSTOMER CREATED A SPECIFIC SYSTEM EVENT REPORT OF ISSUE? No System Release Version 3.6.8 Parts Consumed : 1.00000,459800166651,I-BOX-2 BASE BOARD ASSY ----------------------------------------------------------------------------- Work Order #: WO-03934258 Weaver David The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Description of Event Still have intermittent EKG connectivity problem Ref Case 0117870146 CALLER FUNCTION/ROLE Technician/Sonographer/Therapist DEVICE INTENDED USE Diagnostic <1 Hour Intermittent WHERE IN THE END-USER WORKFLOW IS THE ISSUE OCCURRING? Pre-acquisition/scan WHAT RECENT EVENTS HAVE OCCURRED? During normal use NEGATIVE OPERATOR IMPACT N/A NEGATIVE PATIENT IMPACT/PROCEDURE OUTCOME N/A HAS CUSTOMER CREATED A SPECIFIC SYSTEM EVENT REPORT OF ISSUE? No System Release Version 3.6.8 Diagnostic performed by Engineer : Still have intermittent EKG connectivity problem Ref Case 0117870146. Follow Up Required : Verify iBox 5v power at board is >5v, suggest 5.15v. IF PART REQUIRED - 12 NC RECOMMENDED: 459800484881 - PNEUMO CHEST BELLOWS 459800484391 - BELLOWS PULMO CABLE This is the Pulmo Digital assembly, attaches to gantry and the Chest bellows attach to this: 459800484361 - PULMO TOOL KIT (SG164) Pulmo Bellows Assy Replacement Kit 455011503261 ----------------------------------------------------------------------------- | 3 |  | THe patient monitor readings have been intermittent. Per NSS, replaced the Ibox board today and verified proper operation...though it was tested prior to replacement and found working as well. Will monitor system for issues, and inform customer. | [12 empty] | 117966766 | Closed | 1150710 | TRUE | Closed - Done | Closed - Done | [7 empty] | Others | [6 empty] | ECG | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1216 |  |  | United States | Philips Medical Capital | 300057 | 4440500% |  | Error when trying to do tube warm up | Integration : Confirmation of order 53944768 saved ----------------------------------------------------------------------------- Work Order #: WO-03946309 28-Jul-2021 Hartowicz, Ethan The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: CT will not scan, CIRS error Device Use: CIRS used for reconstruction Expected and Unexpected Results: system would not perform short tube User Impact: system can not scan Patient Impact: NA Current Software Version: 4.1.10.25208 Resolution : Checked errors and logs and found inconsistent 4 channel connection. 0x4321 -> 0x0. I found loose fiber connections on the MOOG. Secured these connections and system returned to normal working condition. ----------------------------------------------------------------------------- Work Order #: WO-03946187 Powell, Jamie The system does not meet the specification for the performed service and is not returned to use Diagnostic performed by Engineer : Recurring system not ready error when scanning. Physical channel order: 0x0 Issue might be with transmitter or power supply.. Check power supply. Check for errors on transmitter.. Check optical.. Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Follow Up Required : Recurring system not ready error when scanning. Physical channel order: 0x0 Issue might be with transmitter or power supply.. Check power supply. Check for errors on transmitter.. Check optical.. MODULE, ROTOR ELECTRONICS TX 453567538334 Internal Comments : Spoke to FSE John Williams who states another FSE Ethan is already on his way to the site to troubleshoot the issue. ----------------------------------------------------------------------------- |  |  | Fiber, check if have fiber cable loose picture. | [12 empty] | 117980984 | Closed | 1158740 |  |  | Evaluation | [20 empty] | Non-Refresh |  |  | Upgrade issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | #N/A
1217 |  |  | United States | Select Specialty Hospitals | 310051 | 4440300% |  | QO: need positioning Pads ordered | Internal Remarks : FADB: (FSE Aware of Double Booking) Title: Internal Remarks Initials: Twila Actions Taken: Per message sent in Teams from FSE/Swett, Mike - Create Case and Dispatch WO for 7/27/21 @ 10-11am CST Travel Time 1hr each way Outcome: Created Case and Dispatched WO per FSE's Request IP: 56915900 Next Steps: None Integration : Confirmation of order 53932767 saved ----------------------------------------------------------------------------- Work Order #: WO-03935087 27-Jul-2021 Swett Mike The system meets the specification for the performed service and is returned to use Diagnostic performed by Engineer : replaced head holder pads Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Parts Consumed : 1.00000,459800033851,Cushion- Small Wedge 1.00000,459800033821,Cushion - Flat 1.00000,459800033841,Cushion- Medium Wedge 1.00000,459800033831,Cushion- Large Wedge ----------------------------------------------------------------------------- |  |  | 2021-07-26 | [12 empty] | 117967451 | Closed | 1155705 |  |  | Closed - Non-Complaint | [20 empty] | Non-Refresh |  |  | Upgrade issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | #N/A
1218 |  |  | United States | Select Specialty Hospitals | 310051 | 4440300% |  | LS Parts Quote |  |  |  | 2021-07-26 | [12 empty] | 117965767 | New | 1154651 |  |  | Evaluation | [20 empty] | Non-Refresh |  |  | Upgrade issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | #N/A
1219 |  |  | United States | Christ Hospital | 300205 | 4439700% |  | 0711CDN85 windows error message | Work Order #: WO-03885044 27-Jul-2021 Hanlin Christopher The system meets the specification for the performed service and is returned to use Tools & Equipment : MULTIMETER,2022-06-30 00:00:00 Resolution : 4.1.10 software reload needed. Reloaded 4.1.10 software. Discovered missing FCO patch disk when restoring software and had to return to finalize software configuration. Problem Description by Engineer : Problem description by engineer : windows error message When was the first occurrence? Within... : 20-Jul-2021 How often is issue occurring? : Continuously Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: No further information is available as per the customer Patient Impact/Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer Current Software Version: Not provided ----------------------------------------------------------------------------- Work Order #: WO-03885022 Atz, Michael The system does not meet the specification for the performed service and is not returned to use Follow Up Required : Error # and/or description of error : windows error message Repair Action : windows error message customer has rebooted several times and the issue remains Technician Name : Atz, Michael Internal Comments : Error # and/or description of error : windows error message Repair Action : windows error message customer has rebooted several times and the issue remains Problem Description by Engineer : Problem description by engineer : windows error message When was the first occurrence? Within... : 20-Jul-2021 How often is issue occurring? : Continuously Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: No further information is available as per the customer Patient Impact/Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer Current Software Version: Not provided Diagnostic performed by Engineer : windows error message ----------------------------------------------------------------------------- |  |  | Check with filed to collect log and screen to support investigation | [12 empty] | 117918979 | Incomplete | 1132321 |  |  | Evaluation |  | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | SW error | [5 empty] | Refresh |  |  | Upgrade issue | System | TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo |  |  | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | No | 0 | 0 | 0 | 0 | 0 | No | y | y | y | 0 | 0 | y | y | y | 0 | y | #N/A
1220 |  |  | United States | Select Specialty Hospitals | 300037 | 4439300% |  | 0699CD72 getting compensator cracked warnings | Internal Remarks : Assisted FSE with false Cracked Compensator errors during Air Calibrations. Collimator was recently replaced. Suggested FSE submit a CFR and collect a bug report. Integration : Confirmation of order 53897942 saved ----------------------------------------------------------------------------- Work Order #: WO-03869287 24-Jul-2021 Beard Michael The system meets the specification for the performed service and is returned to use Resolution : Inspected system. Performed a compensator analysis. Result show no cracks in the collimator. Will follow up with NSS to find a solution for why the system is giving a crack in the collimator error message during air cal. Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: compensator messages User Impact: System functional Patient Impact: None Current Software Version: 4.8 ----------------------------------------------------------------------------- Work Order #: WO-03869107 Karlen James The system meets the specification for the performed service and is returned to use Follow Up Required : check with TSS on messages and compensator test for Windows 10. These may be false messages Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: compensator messages User Impact: System functional Patient Impact: None Current Software Version: 4.8 Diagnostic performed by Engineer : cracked compensator messages on W10 system. Can't remotely test compensator. Collimator was recently replaced. Suggest escalating to TSS ----------------------------------------------------------------------------- |  |  | Check with L3 to sync the latest status of the system. Silva check the required test for collimator replacement Beam and DMS CHECK is used to replace conpenstator crach test | [12 empty] | 117903846 | Closed | 1121245 |  |  | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Compensator | [5 empty] | Refresh |  |  | Upgrade issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | n | y | y | y | y | #N/A
1221 |  |  | United States | Upson Regional Medical Center | 333009 | 4439300% |  | 0711CDY73 "Blue Screen" Login Issue Post-Host Restart | Integration : Confirmation of order 53897121 saved ----------------------------------------------------------------------------- Work Order #: WO-03868162 19-Jul-2021 Lee John The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Problem description by engineer : 0711CDY73 "Blue Screen" Login Issue Post-Host Restart When was the first occurrence? Within... : 02-Jul-2021 08:00:00 How often is issue occurring? : Multiple times per month Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: 0711CDY73 "Blue Screen" Login Issue Post-Host Restart - Delayed Login Patient Impact/Outcome: None User Impact: Delayed Login Current Software Version: 4.1.10.25208 Resolution : Loaded software patches ----------------------------------------------------------------------------- Work Order #: WO-03868139 Holeman, David The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Problem description by engineer : 0711CDY73 "Blue Screen" Login Issue Post-Host Restart When was the first occurrence? Within... : 02-Jul-2021 08:00:00 How often is issue occurring? : Multiple times per month Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: 0711CDY73 "Blue Screen" Login Issue Post-Host Restart - Delayed Login Patient Impact/Outcome: None User Impact: Delayed Login Current Software Version: 4.1.10.25208 Diagnostic performed by Engineer : CT Tech reports for last couple of weeks when she performs a Host restart in the AM, when the HP z440 Host boots back up, she only sees "blue screens" and does not see login screen/box. Tech states she hast to proceed with multiple restarts of the Host and finally the login screen is present. Tech states there are 2 Ingenuity CT's at site and both are exhibiting this issue. Internal Comments : Error # and/or description of error : 0711CDY73 "Blue Screen" Login Issue Post-Host Restart Repair Action : Complete the COPSSH User Activation on the CT Console located in the Ingenuity Family and Brilliance 64 v4.1.10 Software Installation Instructions. Inspect/configure the Host Windows Firewall settings to ensure full PRS/M2M Connectivity. Inspect functionality of restarting z440 Host and logging in. Recommend to proactively install FCO72800740 - FCO TITLE: SOFTWARE UPDATE FOR ING BR64 TO 4.1.10.XX210 Case Number0117119675 WO-03074011 Follow Up Required : Error # and/or description of error : 0711CDY73 "Blue Screen" Login Issue Post-Host Restart Repair Action : Complete the COPSSH User Activation on the CT Console located in the Ingenuity Family and Brilliance 64 v4.1.10 Software Installation Instructions. Inspect/configure the Host Windows Firewall settings to ensure full PRS/M2M Connectivity. Inspect functionality of restarting z440 Host and logging in. Recommend to proactively install FCO72800740 - FCO TITLE: SOFTWARE UPDATE FOR ING BR64 TO 4.1.10.XX210 Case Number0117119675 WO-03074011 Technician Name : Holeman, David ----------------------------------------------------------------------------- |  |  | Link to "Blue screen" defect CTUDT00632553 | [12 empty] | 117903119 | Closed | 1120460 |  |  | Closed - No Further Investigation |  | CTUDT00632553 | [5 empty] | Engineering | [6 empty] | Blue Screen | [5 empty] | Refresh |  |  | Upgrade issue | System | Console-SW | Blue-Screen | Console-SW-Blue-Screen | Console-SW-Blue-Screen | CTUDT00632553 | Blue-Screen | Console-SW | Console-SW-Blue-Screen | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1222 |  |  | United States | Upson Regional Medical Center | 310115 | 4439300% |  | 0711CDY73 "Blue Screen" Login Issue Post-Host Restart | Internal Remarks : Caller: Susan Jaudon IP & Issue: 84628582/ When logging in - getting a blue screen Event Type: Incident Service Request: Remote Technical Did Customer Decline Remote Support? No Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Created Remote W/O Routed to RTAC Next Step Action: N/A Verbal PO Number: Verbal PO Approver: Verbal PO CAP: Integration : Confirmation of order 53896790 saved ----------------------------------------------------------------------------- Work Order #: WO-03868080 19-Jul-2021 Lee John The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Problem description by engineer : "Blue Screen" Login Issue Post-Host Restart When was the first occurrence? Within... : 02-Jul-2021 08:00:00 How often is issue occurring? : Multiple times per month Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: "Blue Screen" Login Issue Post-Host Restart Patient Impact/Outcome: None User Impact: Delayed Login Current Software Version: 4.1.10.25208 Resolution : Loaded backups ----------------------------------------------------------------------------- Work Order #: WO-03867221 Holeman, David The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Problem description by engineer : "Blue Screen" Login Issue Post-Host Restart When was the first occurrence? Within... : 02-Jul-2021 08:00:00 How often is issue occurring? : Multiple times per month Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: "Blue Screen" Login Issue Post-Host Restart Patient Impact/Outcome: None User Impact: Delayed Login Current Software Version: 4.1.10.25208 Diagnostic performed by Engineer : CT Tech reports for last couple of weeks when she performs a Host restart in the AM, when the HP z440 Host boots back up, she only sees "blue screens" and does not see login screen/box. Tech states she hast to proceed with multiple restarts of the Host and finally the login screen is present. Tech states there are 2 Ingenuity CT's at site and both are exhibiting this issue. Follow Up Required : Error # and/or description of error : "Blue Screen" Login Issue Post-Host Restart Repair Action : Complete the COPSSH User Activation on the CT Console located in the Ingenuity Family and Brilliance 64 v4.1.10 Software Installation Instructions. Inspect/configure the Host Windows Firewall settings to ensure full PRS/M2M Connectivity. Inspect functionality of restarting z440 Host and loggin in. Recommend to proactively install FCO72800740 - FCO TITLE: SOFTWARE UPDATE FOR ING BR64 TO 4.1.10.XX210 Case Number 0117119674 WO-03074010 Technician Name : Holeman, David Internal Comments : Error # and/or description of error : "Blue Screen" Login Issue Post-Host Restart Repair Action : Complete the COPSSH User Activation on the CT Console located in the Ingenuity Family and Brilliance 64 v4.1.10 Software Installation Instructions. Inspect/configure the Host Windows Firewall settings to ensure full PRS/M2M Connectivity. Inspect functionality of restarting z440 Host and loggin in. Recommend to proactively install FCO72800740 - FCO TITLE: SOFTWARE UPDATE FOR ING BR64 TO 4.1.10.XX210 Case Number 0117119674 WO-03074010 ----------------------------------------------------------------------------- |  |  | Link to "Blue screen" defect CTUDT00632553 | [12 empty] | 117902663 | Closed | 1120413 |  |  | Evaluation |  | CTUDT00632553 | [5 empty] | Engineering | [6 empty] | Blue Screen | [5 empty] | Refresh |  |  | Upgrade issue | System | Console-SW | Blue-Screen | Console-SW-Blue-Screen | Console-SW-Blue-Screen | CTUDT00632553 | Blue-Screen | Console-SW | Console-SW-Blue-Screen | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1223 |  |  | United States | Christ Hospital | 300205 | 4439000% |  | nemo | Integration : Confirmation of order 53882134 saved Integration : Confirmation of order 53882134 saved ----------------------------------------------------------------------------- Work Order #: WO-03843612 Parts Consumed : 1.00000,459801731761,USB EXTENDER CABLE FROM IBOX TO CONSOLE 1.00000,455300027201,CBL,USB A/M TO B/M 0.5M ----------------------------------------------------------------------------- | [15 empty] | 117879514 | Closed | 1112781 |  |  | Closed - Non-Complaint | [20 empty] | Non-Refresh |  |  | Upgrade issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | #N/A
1224 |  |  | United States | Christ Hospital | 300205 | 4439000% |  | Gate to CT signal missing | Integration : Confirmation of order 53881648 saved ----------------------------------------------------------------------------- Work Order #: WO-03849843 Paine Matthew The system meets the specification for the performed service and is returned to use Diagnostic performed by Engineer : Customer states that EKG Waveform is Present on Monitor, but EKG Waveform is Not Present on Host PC Monitor. Problem Description by Engineer : Customer Function/Role: Clinical Engineer Device Use: Diagnostic Expected and Unexpected Results: Gate to CT signal missing User Impact: Unknown Patient Impact: Unknown Current Software Version:4.1.10 Resolution : Had Customer observe the Ibox 7-Segment Display's while Heart Simulator was connected to EKG Monitor to ensure Ibox Functionality. The Displays changed as the Waveform did ruling out issue with Ibox. Had Customer check for connection of USB Hub Extender on Host PC Side. Customer found that USB Hub Extender had come Disconnected from Signal Distribution Board. Customer Connected USB Extender and now EKG Waveform Present on Host PC Monitor. ----------------------------------------------------------------------------- Work Order #: WO-03843020 Weaver David The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, no follow-up requested Diagnostic performed by Engineer : Gate to CT signal missing Problem Description by Engineer : Description of Problem Gate to CT signal missing CALLER FUNCTION/ROLE Biomed DEVICE INTENDED USE Diagnostic WHEN WAS THE FIRST OCCURRENCE? WITHIN . . . Same day HOW OFTEN IS ISSUE OCCURRING? Continuously WHERE IN THE END-USER WORKFLOW IS THE ISSUE OCCURRING? Pre-acquisition/scan WHAT RECENT EVENTS HAVE OCCURRED? During normal use NEGATIVE OPERATOR IMPACT No signal from gate to CT Host NEGATIVE PATIENT IMPACT/PROCEDURE OUTCOME Delay in scanning HAS CUSTOMER CREATED A SPECIFIC SYSTEM EVENT REPORT OF ISSUE? No 4.1.10 Resolution : Gate to CT signal missing ----------------------------------------------------------------------------- | [15 empty] | 117879056 | Closed | 1107668 |  |  | Closed - No Further Investigation | [20 empty] | Non-Refresh |  |  | Upgrade issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | #N/A
1225 |  |  | United States | AdventHealth Zephyrhills | 300106 | 4438900% |  | 1302CDY73 problems sending images to pacs | Internal Remarks : I spoke to primary FSE Mark, he can't go today and he said to send this call to secondary FSE Garcia Marrero, Jorge. Integration : Confirmation of order 53874315 saved ----------------------------------------------------------------------------- Work Order #: WO-03833277 22-Jul-2021 Garcia Marrero, Jorge The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Problem description by engineer : problems sending images to pacs Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No further information is available as per the customer Expected and Unexpected results: problems sending images to pacs Patient Impact/Outcome: none User Impact: delay Current Software Version: 4.1.7.25030 Diagnostic performed by Engineer : System is having intermittent network delays when sending studies. Host computer was verify remotely with Philips Remote Service. Host computer is fully operational. ----------------------------------------------------------------------------- Work Order #: WO-03833141 Sheffield, Scott The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Problem description by engineer : problems sending images to pacs Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No further information is available as per the customer Expected and Unexpected results: problems sending images to pacs Patient Impact/Outcome: none User Impact: delay Current Software Version: 4.1.7.25030 Diagnostic performed by Engineer : Russ said they have been having this problem in the past and again today. Russ said that it has been freezing up sending images to pacs, now it is starting to send but is sending images very slow to pacs. He is not having problems sending anywhere else and no other systems are having problems sending to pacs. Internal Comments : Error # and/or description of error : none Repair Action : Reading previous case notes the problem was with customer network. When I spoke with FSE Mark he said it is problem with customer network. Follow Up Required : Error # and/or description of error : none Repair Action : Reading previous case notes the problem was with customer network. When I spoke with FSE Mark he said it is problem with customer network. Technician Name : Sheffield, Scott ----------------------------------------------------------------------------- | [15 empty] | 117870905 | Closed | 1099754 |  |  | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS Transfer | [5 empty] | Refresh |  |  | Upgrade issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1226 |  |  | United States | AdventHealth Zephyrhills | 300106 | 4438200% |  | 1302CDN84 DICOM transfer issues sending to PACS | Work Order #: WO-03795988 Mark Finnigan The system partly meets specification and is returned to use with restrictions Follow Up Required : Spoke with Jesse. She states that the DICOM transfers are taking an excessive amount of time to complete. She added that her FSE Jorge worked on this issue this weekend but there is no case showing such activity. Tried to have Jesses log in as olevel but the system was not accepting the username and password. Dispatching the case to the field as there is nothing else that can be done to resolve the issue remotely. According to the customer this system is running 4.1.1. Diagnostic performed by Engineer : Spoke with Jesse. She states that the DICOM transfers are taking an excessive amount of time to complete. She added that her FSE Jorge worked on this issue this weekend but there is no case showing such activity. Tried to have Jesses log in as olevel but the system was not accepting the username and password. Dispatching the case to the field as there is nothing else that can be done to resolve the issue remotely. According to the customer this system is running 4.1.1. Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: The DICOM transfers to PACS are taking an excessive amount of time. User Impact: Patient Impact: Current Software Version: ----------------------------------------------------------------------------- Work Order #: WO-03796009 06-Jul-2021 Sher Mark The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: The DICOM transfers to PACS are taking an excessive amount of time. User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : customer confirms problem is in their network and has observe lost packets customers network is being handled by their IT dapartment ----------------------------------------------------------------------------- | [15 empty] | 117836591 | Closed | 1077479 |  |  | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS Transfer | [5 empty] | Refresh |  |  | Upgrade issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1227 |  |  | Australia | Maia Financial Pty Limited | 341000 | 4441100% |  | Scan cannot initialise | Work Order #: WO-04024450 05-Aug-2021 David BROWN The system meets the specification for the performed service and is returned to use Diagnostic performed by Engineer : Check over system. Carry out replacement of R2D. Carry out full checks and setups. Carry out functional testing. Return to service. Problem Description by Engineer : ________________________________________________________________ CUSTOMER REPORTED PROBLEM OR EVENT DESCRIPTION: 3 Aug 2021 Customer called to report that system would not initialize in preparation for a surview scan, despite shutting down to the wall power. ________________________________________________________________ Parts Consumed : 1.00000,459800467341,R2D Assembly ----------------------------------------------------------------------------- Work Order #: WO-04015383 03-Aug-2021 David BROWN The system does not meet the specification for the performed service and is not returned to use Problem Description by Engineer : ________________________________________________________________ CUSTOMER REPORTED PROBLEM OR EVENT DESCRIPTION: 3 Aug 2021 Customer called to report that system would not initialize in preparation for a surview scan, despite shutting down to the wall power. ________________________________________________________________ Diagnostic performed by Engineer : check over system. all systems seem to be running fine. reset CIRS and full system shut down. system failed during scan. check with Steve Teixera for further instructions. Follow Up Required : N/A Follow Up Required : check angulation of Rotor controller and check r2d signal ----------------------------------------------------------------------------- Work Order #: WO-04021651 04-Aug-2021 David BROWN The system does not meet the specification for the performed service and is not returned to use Diagnostic performed by Engineer : Check over system. Parts have been ordered by RSE. Check system fault with diagnostics by RSE. Order R2D for delivery. Delivery stalled as Border is closed and slowing same day delivery. Return Thursday. Problem Description by Engineer : ________________________________________________________________ CUSTOMER REPORTED PROBLEM OR EVENT DESCRIPTION: 3 Aug 2021 Customer called to report that system would not initialize in preparation for a surview scan, despite shutting down to the wall power. ________________________________________________________________ ----------------------------------------------------------------------------- Work Order #: WO-04015214 Steve, Teixeira The system does not meet the specification for the performed service and is not returned to use Problem Description by Engineer : ________________________________________________________________ CUSTOMER REPORTED PROBLEM OR EVENT DESCRIPTION: 3 Aug 2021 Customer called to report that system would not initialize in preparation for a surview scan, despite shutting down to the wall power. ________________________________________________________________ Diagnostic performed by Engineer : ________________________________________________________________ CUSTOMER REPORTED PROBLEM OR EVENT DESCRIPTION: 3 Aug 2021 Customer called to report that system would not initialize in preparation for a surview scan, despite shutting down to the wall power. ________________________________________________________________ TROUBLESHOOTING: Review of the logs shows Rotor positional WARNING errors (S_MP_POSLOOP_MAX_MIN_ACCEL_ERROR,S_MPS_MOTION_TIMEOUT) on the first attempt at the surview. NORMAL error "S_DMC_SHOT_COMM_COLLAPSE" followed on the 2nd attempt, at 09:21:48.180 3 Aug 2021. Review of the CIRS logs shows numerous resends starting at 09:21:47 and persists throughout the logs. After a shutdown the logs continue to make reference to Rotor positional and Data path errors. ________________________________________________________________ TECHNICAL SUPPORT: 3 Aug 2021 / 4 Aug 2021 Assisted the Engineer on site with troubleshooting. 5 Aug 2021 Assisted the Engineer with handover checks. ________________________________________________________________ ONSITE WORK ORDER: Please refer to onsite work order WO-04015383 for further details of the actions taken to address this issue. ________________________________________________________________ Follow Up Required : TROUBLESHOOTING: Review of the logs shows Rotor positional WARNING errors (S_MP_POSLOOP_MAX_MIN_ACCEL_ERROR,S_MPS_MOTION_TIMEOUT) on the first attempt at the surview. NORMAL error "S_DMC_SHOT_COMM_COLLAPSE" followed on the 2nd attempt, at 09:21:48.180 3 Aug 2021. Review of the CIRS logs shows numerous resends starting at 09:21:47 and persists throughout the logs. After a shutdown the logs continue to make reference to Rotor positional and Data path errors. NEXT STEP ACTION: Engineer required on site to isolate the fault after consultation with the Case RSE. ----------------------------------------------------------------------------- | 2 |  | Carry out replacement of R2D. Carry out full checks and setups. Carry out functional testing. 459800467341,R2D Assembly 可能和供应商水洗工艺相关ECR-077785 | [12 empty] | 118046426 | Closed | 1177833 |  |  | Opened |  | CTUDT00630864 | [5 empty] | Engineering | [6 empty] | R2D | [9 empty] | System | Gantry-HW | R2D | Gantry-HW-R2D | Gantry-HW-R2D | CTUDT00630864 | R2D | Gantry-HW | Gantry-HW-R2D | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1228 |  |  | Chile | BANCO DE CHILE | 341005 | 4440900% |  | CL1688 - SYSTEM IS BLOCKED AND WITH IMAGE ERRORS. | Integration:Confirmation of order 53964748 saved ----------------------------------------------------------------------------- Work Order #: WO-03986980 Valenzuela Velázquez, Eduardo Ariel The system meets the specification for the performed service and is returned to use. Problem Description by Engineer : Problem description by engineer : When performing an acquisition, the equipment is not displaying images on the screen. Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected Results: No errors in reconstructed images. Patient Impact/Outcome: Equipment is down. User Impact: Equipment stopped. Current Software Version: 4.1.10 Diagnostic Performed by Engineer: There is a problem with reconstructed images, they are black and without information. Internal Comments : Error # and/or description of error : It does not display acquisition images. Repair Action: Reconstruction tests of previous examinations are performed and it works properly. Air calibration is performed and the equipment is tested, resulting without errors. The equipment is working. Resolution: Reconstruction tests of previous examinations are performed and it works properly. Air calibration is performed and the equipment is tested, resulting without errors. The equipment is working. ----------------------------------------------------------------------------- | 2 |  | Reconstruction tests of previous examinations are performed and it works properly. Air calibration is performed and the equipment is tested, resulting without errors. | [12 empty] | 118019851 | Closed | 1177302 |  |  | Opened |  | TBD-NoEnoughInfo-TBD-NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Recon image | [9 empty] | System | TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1229 |  |  | China | 东莞市中西医结合医院 | 345013 | 4441500% |  | Wechat：02N/Ingenuity Core128/Elite China/space is full | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ T2 Activities : Problem Description: 客户问题概述： 计算机软件相关故障。 客户问题详细描述或故障现象： 空间不足 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： NA 远程诊断与测试结果： 主机空间不足，客户已经删除多天病人剩余空间还是很少，告知工程师已经远程指导过，还是不行，需要到场处理。联系工程师怀疑是有异常文件占用空间 PRS TR状态： 不通，请工程师到场检查。 Repair Action: 建议现场维修步骤： 检查是否异常文件占用空间；检查是否需要做full rebuild;检查软件,必要时收集bug report,寻求技术支持；维护PRS，若路由器问题需要备件建议开PRS单订备件。 SAP Transfer : Case Transferred to SAP External Remarks : 08.08.2021 04:42:30 UTC SUP JCOUSER (86001504)服务内容：删除D:\Pms\log\data下的300G日志。 服务结果：工作正常。 Internal Remarks : 08.08.2021 04:42:38 UTC Zongxian Li (86001504) RDI Accuracy:Yes $Q7: (Confirmed the reported problem, D disk was full) $Q8: (Not many patient datas in D disk, the D:\Pms\log\data was too large) $Q9: (Deleted the unnecessariy log files in it) $Q10: (Test and inspection data refer to N/A, no PA required. Device returned to full functionality.) $Q11: $Calibrated Tools name (N/A) S/N (N/A) Next Calibration date (N/A) Safety Questions : Case: 0118076705 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Gao Luqi Last updated by: Gao Luqi 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Dispatched for Onsite Support Next Step Action : Li Zongxian ----------------------------------------------------------------------------- | 3 |  | log is too big, CTUDT00633269 | [12 empty] | 118076705 | In Process | 1196136 |  |  | Opened - Pending Translation |  | CTUDT00632261 | [5 empty] | Engineering | [6 empty] | Disk full | [9 empty] | System | Console-SW | DiskFull | Console-SW-DiskFull | Console-SW-DiskFull | CTUDT00632261 | DiskFull | Console-SW | Console-SW-DiskFull | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1230 |  |  | China | 河北省唐山市工人医院 | 40044 | 4441400% |  | 02N/Ingenuity CT 728326/cannot exposure | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Calibration/Adjustment Safety Questions : Case: 0118071432 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Gaoge Last updated by: Wang Gaoge 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 扫描失败。 客户问题详细描述或故障现象： 扫描报错，不能扫描 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 无PRS，看不到具体报错 远程诊断与测试结果： 需要到场检查具体报错 PRS TR状态： 无PRS。 Repair Action: 建议现场维修步骤： 根据报错和故障现象进行维修 Test/Inspection: N/A Test Equipment: N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Problem Reported by Customer : cannot exposure ----------------------------------------------------------------------------- | 3 |  | 服务内容：扫描间温度过热，不让扫描，空调坏了，建议修空调 服务结果：设备恢复正常 | [12 empty] | 118071432 | In Process | 1195213 |  |  | Opened - Pending Translation | [7 empty] | Customer | [6 empty] | Customer | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | n | y | y | y | y | #N/A
1231 |  |  | China | SD Chest Hosp EAST | 40025 | 4441500% |  | three issues for 728323 Ingenuity CT 5000 | Problem Describe: Issue 1： Patient get off the couch when scan end，operator use unload or fast out button to drop the couch. At this moment，Estop opened and log report "S_COUCHMGR_HOR_MOTION_STOP_ERROR" . From the Cantrace log no "COUCH_STOP_HOR_ANS and COUCH_STOP_HOR_JOG_ANS" This problem intermediate occur. Issue 2： Host PC reboot,Customer login software in "CT" account intermittent failed. When this issue occur,the PC screen display blue. Need use IST account login can solve it. Issue 3： When print film in CT software , films quantity intermittent lost. In conclusion, Issue 1 and 2 is bug from L3 reply and need time to fix. But the customer temperature is high. | SHU, Tim |  | What's detail? | [14 empty] | 1194004 |  |  | Investigation | [7 empty] | TBD | [6 empty] | 2021-08-03 | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1232 |  |  | China | 莱阳市人民医院 | 345005 | 4441300% |  | 02N/Ingenuity Core128/Elite China/power off | BU Q&R Feedback: No Reason: Corrective Maintenance External Remarks : 06.08.2021 06:00:42 UTC Xingbo Yang (86038250)服务内容：间歇性掉电，更换FRC电源。 服务结果：维修后，设备恢复正常工作。 Internal Remarks : 06.08.2021 06:00:45 UTC Xingbo Yang (86038250) "RDI Accuracy:yes" $Q7: (Confirmed reported problem & gantry e-stop open) $Q8: ( review error logger 'lifesync p1=2 '& indicate FRC assy defect ) $Q9: ( replace FRC assy 453567057167 ) $Q10: (Test and inspection data refer to: PA & Device returned to full functionality) $Q11: Calibrated tools name(NA) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date: 9/10 / 2020 ^Current Tube usage: scan second_246932 Problem Reported by Customer : 02N/Ingenuity Core128/Elite China SAP Transfer : Case Transferred to SAP T2 Activities : Problem Description: 客户问题概述： ESTOP断开。 客户问题详细描述或故障现象： 偶发掉电 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： lifesync p1=2 远程诊断与测试结果： 机器静置时偶发掉电，丢rhost PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 检查供电，RHOST，检查接线 建议备件12NC及名称： " 453567057167 FRC ASSEMBLY" SAP Transfer : Dispatched for Onsite Support Next Step Action : PPD备件453567057167安排24H到医院，杨兴波-86-13583235930 Safety Questions : Case: 0118062481 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wennian Han Last updated by: Wennian Han 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action : YANG XINGBO ----------------------------------------------------------------------------- | 3 |  | $Q7: (Confirmed reported problem & gantry e-stop open) $Q8: ( review error logger 'lifesync p1=2 '& indicate FRC assy defect ) $Q9: ( replace FRC assy 453567057167 ) | [12 empty] | 118062481 | Closed | 1191118 |  |  | Opened - Pending Translation | [7 empty] | TBD | [6 empty] | FRC | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1233 |  |  | China | Shandong Province Chest Hospital | 40025 | 4441100% |  | 02N/Ingenuity CT/HOST Print failed | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Calibration/Adjustment Internal Remarks : 08.08.2021 02:34:25 UTC Zhaokun Huang (86038405) "RDI Accuracy: YES." $Q7: ( Confirmed reported problem & When Print loss films.) $Q8: ( Ping Printers from Host , passed. Reconfiguration Printers in Lanconfig.) $Q9: ( Reconnected cables between router and printers . Reconfiguration Printers in Lanconfig ) $Q10: ( Test and inspection data refer to: PA . Device returned to fully functionality.) $Q11: Calibrated Tools name (N/A) ^Current Tube installed date: N/A ^Current Tube usage: N/A Problem Reported by Customer : 02N/Ingenuity CT/ T2 Activities : Problem Description: 客户问题概述： 其他类型故障。 客户问题详细描述或故障现象： 打印两张胶片只有一张正常打印 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 打印胶片没有完全打印 远程诊断与测试结果： 考虑软件设置问题 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 1，检查打印机的设置，重新删除添加打印机节点 Test/Inspection: N/A Test Equipment: N/A Safety Questions : Case: 0118048554 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Gaoge Last updated by: Wang Gaoge 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : HUANG ZHAOKUN External Remarks : 03.08.2021 12:38:12 UTC SUP JCOUSER (86038405)服务内容：采集bugreport，重新配置HOST打印机。 服务结果：维修后，设备恢复正常使用。 ----------------------------------------------------------------------------- | 3 |  | 采集bugreport，Software team is under investigation https://teams.microsoft.com/_#/files/%E5%B8%B8%E8%A7%84?threadId=19%3A76cc42a0c8fa4bd7877f7de12141f84c%40thread.tacv2&ctx=channel&context=Escalation%2520%25230118138741%2520new%2520to%2520Queue%2520-%2520customer%2520temperature%2520Yellow!!&rootfolder=%252Fsites%252FBigBoreRTRefreshFOKCoreTeam%252FShared%2520Documents%252FGeneral%252FIngenuity%2520Bugreps%252FChina%2520KM%252FEscalation%2520%25230118138741%2520new%2520to%2520Queue%2520-%2520customer%2520temperature%2520Yellow!! | [12 empty] | 118048554 | Closed | 1183210 |  |  | Opened - Pending Translation |  | CTUDT00633684 | [5 empty] | Engineering | [6 empty] | Printer | [9 empty] | System | Console-SW | Imagemissing-Printer | Console-SW-Imagemissing-Printer | Console-SW-Imagemissing-Printer |  |  | Console-SW | Console-SW-Imagemissing-Printer | No | 0 | 0 | 0 | 0 | 0 | No | y | y | y | 0 | 0 | y | y | y | 0 | y | #N/A
1234 |  |  | China | ZhengXiang White Banner Hospital | 345033 | 4441100% |  | Wechat:02n/Ingenuity Core128/Elite China/Cannot log on | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer:cannot log on Safety Questions:Case: 0118046688 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Tao Fengming Last updated by: Tao Fengming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer's problems: CT login password is locked It's a technology problem with the equipment itself: Yes. Troubleshooting Action: remote diagnosis and testing result: Need an engineer to go to the field to reset the password PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Need an engineer to go to the field to reset the password SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support External Remarks : 04.08.2021 08:36:45 UTC Ruijie Su (86039043)"SUAT. Ready for Clinical Use local time" : "2021-08-04 17:30" External Remarks: 04.08.2021 08:42:05 UTC Ruijie Su (86039043) service outcome: [reset password and it is back to normal. The equipment is back to normal.] Other significant matters: none Internal Remarks: 04.08.2021 08:42:15 UTC Ruijie Su (86039043) $Q7: (Confirmed reported problem: cannot log on ) $Q8: on sit service, Cannot log on, need to reset password. . $Q9: Reset Password . .system is ok! $Q10:Test and inspection data refer to:PA. Device returned to full functionality. $Q11: Calibrated Tools name N/A. Current Tube usage: scan second: 16693 ----------------------------------------------------------------------------- | 2 |  | CT login password is locked Local service communication | [12 empty] | 118046688 | Closed | 1182226 |  |  | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Password | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1235 |  |  | China | 黑龙江省桦川县人民医院 | 345018 | 4441000% |  | 02N/Core128estop open | BU Q&R Feedback: No SAP Customer Symptom Code: HANG Customer Symptom Code: Hang/Crash Reason: Corrective Maintenance Engineer Failure Code: Wear-out/Breakage/Mechanical Failure Repair Activity Code: Reboot/Restart Problem Reported by Customer : 02N/Core128/ Safety Questions : Case: 0118021641 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Bo Last updated by: Wang Bo 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 机架掉电。 Troubleshooting Action: 机架掉电，PRS不通，通过与客户沟通怀疑供电不稳造成，重启机架后正常，FSE反馈再观察。 Repair Action: 机架掉电，PRS不通，通过与客户沟通怀疑供电不稳造成，重启机架后正常，FSE反馈再观察。 Test/Inspection: customer confirm can scan patient and image ok.functional Check PASSED. Device returned to full functionality. Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | 机架掉电，PRS不通，通过与客户沟通怀疑供电不稳造成，重启机架后正常，FSE反馈再观察。 | [12 empty] | 118021641 | Fixed | 1178417 |  |  | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Customer | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | n | y | y | y | y | #N/A
1236 |  |  | China | 江西省余干楚东医院(新院） | 337500 | 4441000% |  | WeChat:02N/Ingenuity Core128/Elite China/Serious artifact | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Calibration/Adjustment Problem Reported by Customer : Serious artifact Safety Questions : Case: 0118021350 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Cheng Zhenhua Last updated by: Cheng Zhenhua 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 图像伪影。 客户问题详细描述或故障现象： 近期颈胸交界处伪影干扰过大 设备本身技术问题： 否。 Troubleshooting Action: 关键故障代码或现象： 图像有伪影 远程诊断与测试结果： 颈胸间伪影属于特殊部位伪影，只能减弱，不能消除，需要到场检查图像质量，排除设备干扰 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 1，做图像质量校正 2 联合技师尝试优化扫描条件和方式降低伪影影响 Test/Inspection: N/A Test Equipment: N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Internal Remarks : 02.08.2021 11:21:58 UTC Jiawei Li (86091272) "RDI Accuracy:Yes" $Q7:(Confirmed reported problem，Image has artifacts) $Q8:(artifacts of special-site ,cannot to be removed) $Q9:( Instruct the physician to pay attention to patient positioning and scan condition selection .) $Q10:(Test and inspection data refer to PA. Device returned to full functionality.) $Q11:(Calibrated Tools: No tools used.) External Remarks : 02.08.2021 11:21:55 UTC Jiawei Li (86091272)服务内容： 颈肩交接的骨头处有伪影干扰，由于人体生理构造原因，此处属于特殊部位伪影，无法消除。指导医院注意患者摆位及扫描条件选择。 服务结果：设备正常使用。 ----------------------------------------------------------------------------- | 3 |  | $Q8:(artifacts of special-site ,cannot to be removed) $Q9:( Instruct the physician to pay attention to patient positioning and scan condition selection .) $Q10:(Test and inspection data refer to PA. Device returned to full functionality.) 颈肩交接的骨头处有伪影干扰，由于人体生理构造原因，此处属于特殊部位伪影，无法消除。指导医院注意患者摆位及扫描条件选择。 | [12 empty] | 118021350 | Closed | 1178263 |  |  | Investigation | [7 empty] | Application | [6 empty] | Application | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | n | y | y | y | y | #N/A
1237 |  |  | China | Ning'an City No. 2 People's Hospital | 345039 | 4441000% |  | Wechat:02N/Ingenuity Core128/Elite China/X-ray fault | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Replace Part SAP Transfer:Case Transferred to SAP Problem Reported by Customer:X-ray fault Safety Questions:Case: 0118020289 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Bai Jingwen Last updated by: Bai Jingwen 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer's problems: Scanning failed. Detailed description of Customer's problems or malfunctions: Scan error, preheat failed It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: " <*HV*> szInfo: 1:1 Grid Comm Fault 2021/08/01 20:32:48.070 <*HV*> szInfo: 8:2 No Fil Current" Remote troubleshoot and test result: Consider the failure of the filament driv and grid control, very likely cathode failure. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: 1. check system status, measure tube light. 2. perform HV diagnosis We recommend spare parts 12NC and names: "Spellman 80 KW RoHS CPM Assy 453567560002" Test/Inspection: N/A Test Equipment: N/A SAP Transfer:Dispatched for On site Support Next Step Action: PPD parts 453567560022 have arranged for 24H to the hospital, Du Haijun -86-18604602671 Next Step Action: Liu Yadong External Remarks : 03.08.2021 14:59:20 UTC Yadong Liu (86039473)"SUAT. Ready for Clinical Use local time" : "2021-08-03 22:00" External Remarks: 04.08.2021 00:36:31 UTC Yadong Liu (86039473) service details: exposure and scan failed. Replace cathode and it is back to normal service result: device returns to normal working condition after service Internal Remarks: 04.08.2021 00:36:34 UTC Yadong Liu (86039473) RDI Accuracy: Yes $Q7:(Confirmed reported problem & X-ray fault,grid comm fault) $Q8: (check error logger, 1:1 Grid Comm Fault and 8:2 No Fil Current, check filament of tube is ok, CPM fault) $Q9: (Replace CPM, cperform relative test, system is ok) $Q10: (Test and inspection data refer to: PA. Device returned to full functionality.) $Q11:Calibrated tools refer to:PA ----------------------------------------------------------------------------- | 2 |  | $Q7:(Confirmed reported problem & X-ray fault,grid comm fault) $Q8: (check error logger, 1:1 Grid Comm Fault and 8:2 No Fil Current, check filament of tube is ok, CPM fault) $Q9: (Replace CPM, cperform relative test, system is ok) | [12 empty] | 118020289 | Closed | 1177704 |  |  | Closed - No Further Investigation |  | TBD-APM/CPM | [5 empty] | Engineering | [6 empty] | CPM | [9 empty] | System | IC-HW | APM/CPM | IC-HW-APM/CPM | IC-HW-APM/CPM | TBD-APM/CPM | APM/CPM | IC-HW | IC-HW-APM/CPM | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1238 |  |  | China | 福建省医科大学附属协和医院 | 345010 | 4440900% |  | 02n/Ingenuity Core128/Elite China/Unable to expose | BU Q&R Feedback: No Reason: Corrective Maintenance SAP Transfer : Dispatched for Onsite Support Problem Reported by Customer : Unable to expose Next Step Action : FSE 李枫 018650369579 call in and disp him Safety Questions : Case: 0118019003 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Lili Last updated by: Lu Lili 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP External Remarks : 02.08.2021 05:58:52 UTC SUP JCOUSER (86036224)服务内容：CT不能曝光。到场检查发现重建柜那边接受到的光路顺序为0x0。打开机架前后盖，检测TX和RX，无异常。重启机架后 ，重建柜得到正确的光路顺序为0x1234。 服务结果：维修后，设备恢复正常工作。 Internal Remarks : 02.08.2021 05:58:55 UTC Feng Li (86036224) RDI Accuracy: Yes $Q7: (Confirmed reported problem.The ASP PM Engineer on site do PM find the CT cannot scan.) $Q8: (Philips FSE on site check and check the Gantry TX and RX.) $Q9: (The channel is wrong. 0X0. Restart the gantry to get the right channel, 0X1234.) $Q10: (Test and inspection data refer to: PA.+Device returned to full functionality.) $Q11: Calibrated Tools name (FLUKE 287C Multimeter) S/N (16490128) NextCalibration date (10/20/2021) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date: NA ^Current Tube usage: NA ----------------------------------------------------------------------------- | 3 |  | $Q8: (Philips FSE on site check and check the Gantry TX and RX.) $Q9: (The channel is wrong. 0X0. Restart the gantry to get the right channel, 0X1234.) did customer restart system regulaly | [12 empty] | 118019003 | Closed | 1175573 |  |  | Closed - No Further Investigation | [7 empty] | TBD | [6 empty] | Reboot | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1239 |  |  | Germany | Gesundheit Nord gGmbH Klinikverbund | 40002 | 4441300% |  | Soft tissue / sagittal images are not completely recalculated | Internal Remarks : involviert Herr Halama Next Step Action : Anruf wird unter der angegebenen Tel.Nr. nicht beantwortet automatische Ansage: "Der Teilnehmer ist abwesend" Next Step Action : Anruf wird unter der angegebenen Tel.Nr. nicht beantwortet automatische Ansage: "Der Teilnehmer ist abwesend" ----------------------------------------------------------------------------- Work Order #: WO-04037036 Carsten Becker Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- MPR not reconstructed successfully Customer Temperature as Yellow | 3 |  | MPR not reconstructed successfully Customer Temperature as Yellow Reminder sent to the help desk that the escalation is still unprocessed The call will be made on the specified phone number. not answered automatic announcement: "The participant is absent" @Silva, need link to MPR defect | [12 empty] | 118066890 | In Progress | 1192687 |  |  | Opened - Pending Translation |  | CTUDT00632116 | [5 empty] | Engineering | [6 empty] | MPR | [9 empty] | System | CIRS-SW | MPR Recon | CIRS-SW-MPR Recon | CIRS-SW-MPR Recon | CTUDT00632116 | MPR Recon | CIRS-SW | CIRS-SW-MPR Recon | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | n | y | y | 0 | y | #N/A
1240 |  |  | Germany | Gesundheit Nord gGmbH Klinikverbund | 40002 | 4441100% |  | It takes a long time until images arrive in PACS when sending | Internal Remarks: Mr. Halama clarifies with Mr. Kraschewski and contact us when he has received the computer. -------------------------------------------------- --------------------------- Work Order #: WO-04018729 Carsten Becker The system meets the specification for the performed service and is returned to use Diagnostic performed by Engineer: LogFile download and analysis Problem Description by Engineer: via FSE Patient data are not automatically sent to the PACS Follow Up Required: after consultation with R. Kraschewski (TSS) Installation of a sniff PC in the customer network is required to clarify the cause. The FSE used (e.g. T. Halama) should contact him so that he can send him the computer and coordinate the further procedure. Internal Comments: No errors listed in the transfer log. All images were successfully transferred The following messages were stored in the default log for the period in question 2021.08.02 10:59:14 TCPLayer: recvCompleted-EndReceive returned with size = 0 2021.08.02 11:00:53 2021.08.02 11:05:21 2021.08.02 11:10:44 2021/08/02 11:14:54 AM DicomSession: DICOM Error. Status = 0110 RemoteDevice: MPPS failed with status 0110 Generic MPPSClient: Exception while sending Create Message to RIS. Failure status returned 0110 DicomSession: servicesClose called but association already disposed ExamManager MPPS Set message sending failed SSDE Computation failed as the WED profile is empty or incomplete Request RR from T. Halama on box Internal Comments: Mr. Glogner-Umbeer has information that there may be a network problem, as a DICOM error message is issued when communicating with the PACS as well as the RIS server (which both have the same IP address and are therefore physically one computer) DICOM Error Code 0110 = Processing failure | 3 |  | 2021-08-03 | [12 empty] | 118049952 | In Progress | 1184114 |  |  | Opened - Pending Translation | [7 empty] | TBD | [6 empty] | 2021-08-03 | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1241 |  |  | Japan | Maki Hospital | 341017 | 4441300% |  | 04DS:KK: A scan stopped yesterday during a contrast-enhanced examination.>>He was able to scan when he retri [Truncated] | Work Order #: WO-04035332 Kojima, Kenichi The system meets the specification for the performed service and is returned to use Problem Description by Engineer : 昨日の造影検査にてスキャン停止。>>再スキャンでスキャン可。>>S_MPLTRANSPORTREC_DEFRAG_TOUCAN_ERROR Diagnostic performed by Engineer : 再スキャンにて検査可。Ctrl linkのカウントにも大きな異常は無し。 造影検査の停止の為FSE対応希望。三宅さんに対応依頼。 Loggerより ============================================================ 2021/07/30 11:13:53.202 [START STUDY] GSE::CState::StateBkst::STUDY_STARTED was sent to GPC 2021/07/30 11:17:21.399__01 <WARN> Type: WARN ErrNum: S_HMC_AXIS_AT_POSITIVE_LIMIT P1:215600/P2:215895 2021/07/30 11:18:40.194__06 [---] ------ Series 13 Load Structure ----- 2021/07/30 11:18:58.382__02 [--o] ----------------- Series 13 Ready ----------------- 2021/07/30 11:19:03.353__01 [-oo] - - - - - - Shot 0 Starts - - - - - - 2021/07/30 11:19:13.387 [ooo] ----------------- Series 13 Complete ----------------- 2021/07/30 11:20:04.216 [---] ------ Series 14 Load Structure ----- 2021/07/30 11:20:22.479__02 [--o] ----------------- Series 14 Ready ----------------- 2021/07/30 11:20:28.991__01 [-oo] - - - - - - Shot 0 Starts - - - - - - 2021/07/30 11:20:45.917 [ooo] ----------------- Series 14 Complete ----------------- 2021/07/30 11:33:56.619__07 [---] ------ Series 15 Load Structure ----- 2021/07/30 11:33:56.693 [CANCEL] Series 0 cancelled, due to GANTRY_CANCEL_SOURCE_NORMAL_ERROR, in state CStateSeriesLoading 2021/07/30 11:33:56.693__01 [CANCEL] CStateSeriesLoading::GantrySeriesCancelAns had GANTRY_CANCEL_SOURCE_NORMAL_ERROR 2021/07/30 11:33:56.703 <*ERROR*> Type: NORMAL ErrNum: S_MPLTRANSPORTREC_DEFRAG_TOUCAN_ERROR P1:3.289968E+07/P2:24576 2021/07/30 11:36:40.546__10 [---] ------ Series 18 Load Structure ----- 2021/07/30 11:36:40.613 [---] ------ Series 19 Load Structure ----- 2021/07/30 11:36:40.707__01 [---] ------ Series 20 Load Structure ----- 2021/07/30 11:36:56.790__02 [--o] ----------------- Series 18 Ready ----------------- 2021/07/30 11:37:46.664__01 [-oo] - - - - - - Shot 0 Starts - - - - - - 2021/07/30 11:37:55.522 [ooo] ----------------- Series 18 Complete ----------------- Internal Comments : [Subject] 04DS:KK:ﾎｼ様:昨日の造影検査にてスキャン停止。>>再スキャンでスキャン可。>>S_MPLTRANSPORTREC_DEFRAG_TOUCAN_ERROR [ディスパッチ] 1 RC -----------------FSE登録情報----------------- -訪問FSE1氏名 三宅 正浩 -FSE1訪問日時 2021/07/3118:00 -FSE1作業時間 2 ----------------------------------------------------------------------------- Work Order #: WO-04035453 06-Aug-2021 Miyake, Masahiro The system meets the specification for the performed service and is returned to use Regulatory Questions : N/A Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : [作業内容詳細] エラーログ等にて故障原因を調査しました結果、ガントリー内部基板内ソフトウェアにて一部通信エラーが初性しており、エラー検知による撮影停止と思われます。 ハードウェアチェックを行いましたが、異常は見られませんでしたので、基板内部ソフトウェア再インストールを実施しております。 作業実施後テストスキャン等実施正常動作を確認しております。 ご迷惑をお掛けしまして誠に申し訳ありませんでした。 Internal Comments : [CX関連] 機器担当者と報告完了方法を合意できましたか？ YES 当日通常報告完了 [作業報告書（紙）提出] 作業報告書は紙で提出しましたか？ YES Q1その検査はどの様に終了しましたか?: ・中断した個所から再撮 Q2 問題により間違った所見（誤診、誤治療）の有無: ・無し Q3 アーチファクトの有無: ・無し Q4 意図しない不要な曝射の有無: ・無し ----------------------------------------------------------------------------- | 4 |  | 2021-08-05 | [12 empty] | 118064473 | Fixed | 1188036 |  |  | Investigation - Pending Translation | [7 empty] | TBD | [6 empty] | 2021-08-05 | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1242 |  |  | Japan | Social Welfare Organization Saiseikai Imperial Gift Foundation, Inc. | 343000 | 4441300% |  | 04DS:KKI:Artifact occurred on Surview images only.There is no problem with the main scan images. | Work Order #: WO-04033110 05-Aug-2021 Kawashige, Yasuhiro Regulatory Questions : N/A Diagnostic performed by Engineer: N/A Problem Description by Engineer: Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: ----------------------------------------------------------------------------- Work Order #: WO-04032952 Kimura, Kazumasa The system meets the specification for the performed service and is returned to use. Problem Description by Engineer: Artifact occurred on Surview images only.There is no problem with the main scan images. Diagnostic performed by Engineer : Artifact occurred on Surview images only.There is no problem with the main scan images. Internal Comments: [Subject] 04DS:KKI:Mr. Moriwaki:Artifact occurred on Surview images only.There is no problem with the main scan images. [Internal Comments] The problem occurs intermittently. Shipped part: 459800911271:TDMS V2 M.Det Class A FRU QTY:2 ZREP/Traceable Okayama (U616) Dispatched: 1 RC FSE Registration Information: Name of FSE 1 who visited the site: Yasuhiro Kawashige FSE 1 visit date and time: 2021/08/0519:00 FSE 1 hours worked: 2 Dispatch instructions: Dispatch instruction #1: Shipping conditions: TDD (date and time specified) Delivery date: Specified date: 2021/08/05 19:00 Shipping destination: Directly to site. To be collected by: Yasuhiro Kawashige Phone number: 070-2196-4606 Other: call mobile phone upon arrival Person making request: Kazumasa Kimura ----------------------------------------------------------------------------- | 4 |  | 2021-08-05 | [12 empty] | 118061985 | In Progress | 1186767 |  |  | Investigation | [7 empty] | TBD | [6 empty] | 2021-08-05 | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1243 |  |  | Japan | Independent administrative institution National Hospital Organization | 341015 | 4441200% |  | 04DS:KKI: Error occurred during a contrast-enhanced scan >> szInfo: SoftArcReg 0x211*/S_RHOST_HVG_HSS_GENERA [Truncated] | Work Order #: WO-04024958 ----------------------------------------------------------------------------- Work Order #: WO-04024933 04-Aug-2021 Yamamoto, Yuko The system meets the specification for the performed service and is returned to use. Problem Description by Engineer: Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Regulatory Questions : N/A Diagnostic performed by Engineer: N/A ----------------------------------------------------------------------------- Work Order #: WO-04025225 Kimura, Kazumasa The system meets the specification for the performed service and is returned to use. Diagnostic performed by Engineer: Call description: Error occurred during a contrast-enhanced scan [Response] He was able to continue with the examination. Problem Description by Engineer: Error occurred during a contrast-enhanced scan >> szInfo: SoftArcReg 0x211*/S_RHOST_HVG_HSS_GENERATOR_ARC_FAULTDevice Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Internal Comments: [Subject] 04DS:KKI: Mr./Ms. Hisaki: Error occurred during a contrast-enhanced scan >> szInfo: SoftArcReg 0x211*/S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT [Internal Comments] [Results of checking the log] An arc has occurred. szInfo: SoftArcReg 0x211* S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT S_GEN_FAULT_ANODE_TUBE_ARC S_GEN_FAULT_CATHODE_TUBE_ARC LongTC and Filament cal. are scheduled to be performed today. Dispatched: 1 RC FSE Registration Information: Name of FSE 1 who visited the site: L1 Engineer Shigeya Yoshimura FSE 1 visit date and time: 2021/08/0417:30 FSE 1 hours worked: 2 - Name of dispatched supervisor Yuko Yamamoto ¦ L2 (Level 2) Dispatched supervisor [Dispatched supervision] Dispatched supervisor case L1 FSE_ Shigeya Yoshimura ¦ L2 L2 FSE_ Yutaka Yamamoto ¦ L2 L1 FSE_ This system has work performed under the supervision of the L2FSE, and it is necessary for the L2FSE (engaged in another operation) to approve the parts order, manage technical escalation, and receive final approval before starting use on the patient. L2 FSE The role of dispatched supervisors is to approve all parts orders, manage all technical escalations, provide guidance and instructions for service work to the L1FSE who has been assigned, and review and approve the work of the FSE. The L2 FSE makes final approval of the of the matters recorded in the case, and Close (CNF) means that the repair of the SWO and the recorded content (including all recorded comments) were officially approved. ----------------------------------------------------------------------------- | 4 |  | 2021-08-04 | [12 empty] | 118055041 | In Progress | 1183073 |  |  | Investigation | [7 empty] | TBD | [6 empty] | 2021-08-04 | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1244 |  |  | Japan | Kinan Hospital | 340000 | 4441200% |  | 01DS: KKI: "Recon Box initialize error" occurred during a scan. Scanning is not possible >>S_DMC_SHOT_COMM_C [Truncated] | Work Order #: WO-04023668 Kimura, Kazumasa The system does not meet the specification for the performed service and is not returned to use Problem Description by Engineer: "Recon Box initialize error" occurred during a scan. Scanning is not possible >> S_DMC_SHOT_COMM_COLLAPSE > Restarting the system did not improve the situation Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer: "Recon Box initialize error" occurred during a scan. Scanning is not possible >> S_DMC_SHOT_COMM_COLLAPSE > Restarting the system did not improve the situation Internal Comments: [Subject] 01DS: KKI: Mr. Kadosaka: "Recon Box initialize error" occurred during a scan. Scanning is not possible >>S_DMC_SHOT_COMM_COLLAPSE > Restarting the system did not improve the situation [Internal Comments] [Response] Even after a system restart, Surview was possible but the main scan was interrupted by an error. The situation has not improved, so I think the system is down. [Results of checking the log] The following error occurred. S_DMC_SHOT_COMM_COLLAPSE The physical channel order was 0 x 0, and a communication problem occurred between the CIRS and the gantry. Shipped part: 453567538332: Module - Rotor Electronics TX QTY: 1 ZREP/Traceable Tokyo (U611) 453567538322: Module, Optical RX Stator Moog Velocity QTY: 1 ZREP/Traceable Tokyo (U611) 453567057167: FRC ASSEMBLY QTY: 1 ZREP/Traceable Osaka (U615) Dispatched: 1 RC FSE Registration Information: Name of FSE 1 who visited the site: Shigeya Yoshimura FSE 1 visit date and time: 2021/08/0413:00 FSE 1 hours worked: 4.0 Dispatch instructions: Dispatch instruction #1: Shipping conditions: ASAP (Arrange shipment) Delivery date: ASAP Shipping destination: Directly to site. To be collected by: Shigeya Yoshimura Phone number: 090-9020-2081 Other: call mobile phone upon arrival Person making request: Kazumasa Kimura ----------------------------------------------------------------------------- | 2 |  | Shipped part: 453567538332: Module - Rotor Electronics TX QTY: 1 ZREP/Traceable Tokyo (U611) 453567538322: Module, Optical RX Stator Moog Velocity QTY: 1 ZREP/Traceable Tokyo (U611) 453567057167: FRC ASSEMBLY QTY: 1 ZREP/Traceable Osaka (U615) Part/Product - ACQ4eG2 Acquisitor FRU Part Number - 459800245334 Actual Quantity - 1.00000 Part/Product - BR64 LC RCOM ASSY Part Number - 459801073611 Actual Quantity - 1.00000 This case was originally reported as system down since the error occurrence was very frequent, but still intermittent actually. After the receiver and ACQ card were replaced, it has not happened for two weeks. | [12 empty] | 118053754 | In Progress | 1182507 |  |  | Investigation |  | ACQCardReplacement | [5 empty] | Engineering | [6 empty] | Recon Box initialize error | [9 empty] | System | Console-HW | ACQCard | Console-HW-ACQCard | Console-HW-ACQCard |  |  | Console-HW | Console-HW-ACQCard | No | 0 | 0 | 0 | 0 | 0 | No | y | y | y | 0 | 0 | y | y | y | 0 | y | #N/A
1245 |  |  | China | ZhengXiang White Banner Hospital | 345033 | 4441100% |  | 02N/Ingenuity Core128/Elite China/Login password error | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ Safety Questions:Case: 0118047344 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingsong Last updated by: Lu Mingsong 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Login password error Troubleshooting Action: Case Number : 0118046688 follow up Repair Action: Case Number : 0118046688 follow up Test/Inspection: Functional Check PASSED. The device returned to full functionality. Test Equipment: N/A SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | Gao, Mia |  | 2021-08-03 | [12 empty] | 118047344 | Closed | 1182515 |  |  | Closed - Duplicate | [7 empty] | 2021-08-04 11:39 | [6 empty] | 2021-08-03 | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1246 |  |  | Peru | BANCO BBVA PERU | 40034 | 4441200% |  | Bad detector module 24 - 459800594931 | Work Order #: WO-04038394 [REDACTED] The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled. Problem Description by Engineer : Customer Function/Role: Technician Device Use: In use Expected and Unexpected Results: Artifacts in image User Impact: Bad quality images Patient Impact: Bad quality images Current Software Version: 4.1.10 Diagnostic Performed by Engineer: Artifacts are displayed in the image due to two damaged detectors in module 24. Due to quality policies, the detector must be replaced. Follow Up Required: Change module 24. Bring long Allen wrench, according to the manual. ----------------------------------------------------------------------------- | 4 |  | 2021-08-04 | [12 empty] | 118060169 | In Progress | 1189887 |  |  | Opened | [7 empty] | TBD | [6 empty] | 2021-08-04 | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1247 |  |  | Puerto Rico | METRO MAYAGUEZ INC | 40009 | 4441300% |  | Head support is broken | Work Order #: WO-04038494 Valenzuela Velasquez, Eduardo Ariel The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Problem description by engineer : Head support is broken. Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Head support good no cracks. Patient Impact/Outcome: No impact. User Impact: No impact. Current Software Version: 4.1.10 Internal Comments : Error # and/or description of error : Head support is broken. Repair Action : Head support replacement. Diagnostic performed by Engineer : The head support is broken and must be replaced. Follow Up Required : Error # and/or description of error : Head support is broken. Repair Action : Head support replacement. Technician Name : Valenzuela Velasquez, Eduardo Ariel ----------------------------------------------------------------------------- | 3 |  | 2021-08-05 | [12 empty] | 118068093 | In Progress | 1190730 |  |  | Evaluation | [7 empty] | TBD | [6 empty] | 2021-08-05 | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1248 |  |  | Slovakia | S&T Slovakia s.r.o. | 40062 | 4441200% |  | CT Ingenuity, SN 40062 |  | 3 |  | 2021-08-04 | [12 empty] | 118055255 | New | 1187383 |  |  | Opened | [7 empty] | TBD | [6 empty] | 2021-08-04 | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1249 |  |  | Turkey | T.C Saglik Bakanligi Bagcilar | 40010 | 4441200% |  | Gantry dönerken ses yapıyor | Integration : Confirmation of order 53977716 saved ----------------------------------------------------------------------------- Work Order #: WO-04027444 05-Aug-2021 Ergun Aydin The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: technician Device Use: in use Expected and Unexpected Results: noise while gantry rotation User Impact: no user impact Patient Impact: no patient impact Current Software Version:NA Diagnostic performed by Engineer : rotor dönüşünü sonlandırırken gürültü geliyor Resolution : Rotor kayış değişiminden sonra gantry den gürültü geldiği söylendi. Kayış kontrol edildi. Problem görülmedi. Bahsi geçen gürültü eski-yeni kayış farkından dolayıdır. Söylenen gürültü sistem normal çalışması sırasında oluşan, mekanik olarak rotoru durdurma sırasındaki fren sesidir. Herhangi bir düzeltmeye gerek yoktur. Sistem çalışır ve hasta alır durumda teslim edildi. ----------------------------------------------------------------------------- Work Order #: WO-04026830 Salih Mendi The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Customer Function/Role: technician Device Use: in use Expected and Unexpected Results: noise while gantry rotation User Impact: no user impact Patient Impact: no patient impact Current Software Version:NA Diagnostic performed by Engineer : Kullanıcının, gantry dönüşü sırasında oluşan gürültü şikayeti üzerine, kullanıcıya uzaktan destek verildi. Yapılan incelemede, gürültü kaynağı tespit edilemediğinden, saha ziyareti gerçekleştirilerek ayrıntılı inceleme yapılması gerekmektedir. Sistem faal durumdadır. Follow Up Required : Yapılan incelemede, gürültü kaynağı tespit edilemediğinden, saha ziyareti gerçekleştirilerek ayrıntılı inceleme yapılması gerekmektedir. Sistem faal durumdadır. ----------------------------------------------------------------------------- | 3 |  | 2021-08-04 | [12 empty] | 118057076 | Closed | 1183909 |  |  | Opened - Pending Translation | [7 empty] | TBD | [6 empty] | 2021-08-04 | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1250 |  |  | Turkey | Republic of Turkey Ministry of Health, Bagcilar | 40010 | 4441000% |  | the table is having difficulty moving | Integration: Confirmation of order 53967425 has been saved. ----------------------------------------------------------------------------- Work Order #: WO-04014224 02-Aug-2021 Serkan Yavaşça The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : In diagnosis, we found that the table had been impacted on the front left and the holder inside the top cover was broken. The broken part was affixed. Table was lubricated. System was delivered in working condition. ----------------------------------------------------------------------------- | 3 |  | 2021-08-02 | [12 empty] | 118045756 | Closed | 1177417 |  |  | Investigation | [7 empty] | Customer | [6 empty] | 2021-08-02 | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | n | y | y | y | y | #N/A
1251 |  |  | United States | McLaren Macomb | 40021 | 4441400% |  | 0399CFXXX Customer Requesting Support - Parts ID. | Internal Remarks : Interaction Type: New Case Caller: Allison Adamczyk IP & Issue: 85380634 Customer Requesting Support - Parts ID. Event Type: Part ID Service Request: Remote Technical Did Customer Decline Remote Support? No Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Remote Technical Next Step Action: N/A Verbal PO Number: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Integration : Confirmation of order 53991395 saved ----------------------------------------------------------------------------- Work Order #: WO-04046517 Boyd Craig The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Problem description by engineer : Customer needs couch restraint and cushion part number. When was the first occurrence? Within... : 06-Aug-2021 How often is issue occurring? : Unknown Malfunction area : Accessories Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: None Patient Impact/Outcome: Not provided User Impact: None Current Software Version: Not provided Resolution : Researched part numbers for customer. Customer already has the table cushion pad for the Brilliance. 453567009951 6" Patient Restraint strap CT Yes 453567009961 14" Patient Restraint strap CT Yes Customer is okay with closing the call. Diagnostic performed by Engineer : Collect information from customer. Internal Comments : Error # and/or description of error : Part numbers Repair Action : Researched part numbers for customer. Customer already has the table cushion pad for the Brilliance. 453567009951 6" Patient Restraint strap CT Yes 453567009961 14" Patient Restraint strap CT Yes Customer is okay with closing the call. ----------------------------------------------------------------------------- | 5 |  | 2021-08-06 | [12 empty] | 118074964 | Closed | 1193725 |  |  | Opened | [7 empty] | TBD | [6 empty] | 2021-08-06 | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1252 |  |  | United States | Holston Medical Group | 40053 | 4441200% |  | When using angle greater than 10 for axial Brain CT scanner does not create complete result |  |  |  | @Silva, link to MPR defect | [14 empty] | 1192574 |  |  | Investigation |  | CTUDT00632116 | [5 empty] | Engineering | [6 empty] | MPR | [9 empty] | System | CIRS-SW | MPR Recon | CIRS-SW-MPR Recon | CIRS-SW-MPR Recon | CTUDT00632116 | MPR Recon | CIRS-SW | CIRS-SW-MPR Recon | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1253 |  |  | United States | Willis Knighton Hospital | 40004 | 4441200% |  | Upper Pallet Float Motion Stiff |  | 4 |  | 2021-08-04 | [12 empty] | 118060392 | Closed | 1189984 |  |  | Evaluation | [7 empty] | TBD | [6 empty] | 2021-08-04 | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1254 |  |  | United States | Philips Medical Capital | 40028 | 4441200% |  | ACR phantom scanning | Work Order #: WO-04029582 Hirchak, Michele The system meets the specification for the performed service and is returned to use Diagnostic performed by Engineer : Customer working with Physicist to set up axial protocols to match the helical clinical scans. Customer unsure on how to build an axial protocol. Physicist sent me his ACR worksheet so I can help Iris later to build protocols. I will call her later.> MBH Problem description by Clinical Support Specialist : Customer Function/Role: Tech Device Use: In Use Expected and Unexpected Results:ACR set up axial protocol for Phantom scanning. User Impact: None Patient Impact: None Current Software Version: 4.1.10 Problem description by Clinical Support Specialist : Called Iris multiple times over the last 2 days. Emailed physicist. I was able to review his ACR document and found multiple errors. I sent him an additional email with corrections, and told him we can't be sure on everything since I have never been able to check with the tech regarding the actual protocol parameters. I will keep trying Iris in the morning. MBH ----------------------------------------------------------------------------- | 5 |  | 2021-08-04 | [12 empty] | 118059474 | In Progress | 1185766 |  |  | Evaluation | [7 empty] | TBD | [6 empty] | 2021-08-04 | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1255 |  |  | United States | Philips Medical Capital | 40053 | 4441100% |  | MPR help | Integration : Confirmation of order 53974846 saved ----------------------------------------------------------------------------- Work Order #: WO-04022710 Curry, Dyan The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Diagnostic performed by Engineer : CFR # 1192574 Problem description by Clinical Support Specialist : Customer Function/ Role: technologist Device Use: in use name Expected and Unexpected Results: Axial head not creating complete set of mprs User Impact: none Patient Impact: none Current Software version: 4.1.10 Diagnostic performed by Engineer : customer states cor and sag reformats a partial of needed images for axial heads. Advised customer tilting gantry can conflict with creating MPRS. Customer is able to create thin data set , load thins in Ct Viewer and create full reformat series. Customer also says EC was loaded on scanner from another scanner using EC sharing. Built new EC using reference set. Customer will use this new EC for next exams. Follow up tomorrow afternoon. Diagnostic performed by Engineer : customer says new EC card worked with no issue . Advised customer CFR will be created for science department to take a closer look for updates. CFR created, waiting for reference number ----------------------------------------------------------------------------- | 4 |  | @Silva, link to MPR defect | [12 empty] | 118052925 | Closed | 1181874 |  |  | Duplicate |  | CTUDT00632116 | [5 empty] | Engineering | [6 empty] | MPR | [9 empty] | System | CIRS-SW | MPR Recon | CIRS-SW-MPR Recon | CIRS-SW-MPR Recon | CTUDT00632116 | MPR Recon | CIRS-SW | CIRS-SW-MPR Recon | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1256 |  |  | United States | Unity Health Systems | 40037 | 4441000% |  | Monthly QA | Work Order #: WO-04013475 06-Aug-2021 Lockwood Thomas The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : Performed NY State QA testing ----------------------------------------------------------------------------- | 5 |  | 2021-08-02 | [12 empty] | 118045329 | In Progress | 1181364 |  |  | Evaluation | [7 empty] | TBD | [6 empty] | 2021-08-02 | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1257 |  |  | United States | Philips Medical Capital | 300057 | 4441400% |  | Gantry Error. Delayed imaging after injections. | Internal Remarks : Interaction Type: New Case Caller: Michelle Kelso IP & Issue: 54851181 ; Gantry Error. Delayed imaging after injections. Event Type: Incident Service Request: Onsite Did Customer Decline Remote Support? Yes Preferred Service Time: Week of 08/09/2021 Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Create case and onsite work order Next Step Action: N/A Verbal PO Number: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- Work Order #: WO-04048250 Williams John Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 4 |  | 1.00000,459800170271,Upgrade Kit, Commander-SK Brilliance 64 | [12 empty] | 118076167 |  | 1195981 |  |  | Opened | [20 empty] | Non-Refresh | [12 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | #N/A
1258 |  |  | United States | Christ Hospital | 300205 | 4441300% |  | Ecg not functioning | Integration : Check the settlement rule-Entry 1 in table OPERATION (method CHANGE) was not found-Error during processing of BAPI methods-MPG is present on the product code : P9-MPG is present on the product code : P9 Integration : VBAKKOM has been processed successfully-VBAPKOM has been processed successfully-VBAPKOM has been processed successfully-VBAPKOM has been processed successfully-VBAPKOM has been processed successfully-KONVKOM has been processed successfully-KONVKOM has been processed successfully-KONVKOM has been processed successfully-KONVKOM has been processed successfully-REQ2RES DMR 18144623 has been saved-MPG is present on the product code : P9-MPG is present on the product code : P9 ----------------------------------------------------------------------------- Work Order #: WO-04040554 Hanlin Christopher Diagnostic performed by Engineer : N/A Problem Description by Engineer : Problem description by engineer : ECG disconnecting When was the first occurrence? Within... : 01-Aug-2021 How often is issue occurring? : Intermittent Malfunction area : Console Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected results: can't do cardiac exams Patient Impact/Outcome: cardiac exams moved to another scanner User Impact: Can't do cardiac exams Current Software Version: Not provided ----------------------------------------------------------------------------- Work Order #: WO-04040376 Mcginnis, Mike The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, no follow-up requested Internal Comments : Error # and/or description of error : ECG disconnecting. Biomed has replaced USB extenders, cable from MP5 to gantry, used the new USB cable from i-box to host and replaced i-box. He said the host sometimes sounds like a piano playing with the usb disconnects. He also stated sometimes the host won't boot up after a power cycle. Repair Action : Check 5v power supply. If good replace host and insure bios setting for usb is correct. Run new usb cable from gantry to console. Cable already on site Diagnostic performed by Engineer : Spoke to Biomed. He said they have been having ecg disconnects since the last upgrade. Problem Description by Engineer : Problem description by engineer : ECG disconnecting When was the first occurrence? Within... : 01-Aug-2021 How often is issue occurring? : Intermittent Malfunction area : Console Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected results: can't do cardiac exams Patient Impact/Outcome: cardiac exams moved to another scanner User Impact: Can't do cardiac exams Current Software Version: Not provided Follow Up Required : Error # and/or description of error : ECG disconnecting. Biomed has replaced USB extenders, cable from MP5 to gantry, used the new USB cable from i-box to host and replaced i-box. He said the host sometimes sounds like a piano playing with the usb disconnects. He also stated sometimes the host won't boot up after a power cycle. Repair Action : Check 5v power supply. If good replace host and insure bios setting for usb is correct. Run new usb cable from gantry to console. Cable already on site What are the next Troubleshooting steps : Check 5v power supply. If good replace host and insure bios setting for usb is correct. Cable already on site Technician Name : Mcginnis, Mike ----------------------------------------------------------------------------- Replaced USB connection with straight through cable recommended by RSE. Connected directly to Host and iBox. Installed, then de-installed iBox and restored system to original iBox. Programmed iBox and tested operation. Will monitor. | 3 |  | Replaced USB connection with straight through cable recommended by RSE. Connected directly to Host and iBox. Installed, then de-installed iBox and restored system to original iBox. Programmed iBox and tested operation. Will monitor. | [12 empty] | 118069547 |  | 1190332 |  |  | Investigation | [20 empty] | Non-Refresh | [12 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | #N/A
1259 |  |  | United States | Christ Hospital | 300205 | 4441100% |  | 459801190772 X1 459801190762 X1 459800166652 X1 | Integration : Confirmation of order 53974620 saved Integration : Confirmation of order 53974620 saved ----------------------------------------------------------------------------- Work Order #: WO-04022342 Parts Consumed : 1.00000,459801190772,ICRON USB RG2302 REMOTE EXT MODULE (REX) 1.00000,459801190762,ICRON USB RG2302 LOCAL EXT MODULE (LEX) 1.00000,459800166651,I-BOX-2 BASE BOARD ASSY ----------------------------------------------------------------------------- | 3 |  | Replaced USB connection with straight through cable recommended by RSE. Connected directly to Host and iBox. Installed, then de-installed iBox and restored system to original iBox. Programmed iBox and tested operation. Will monitor. | [12 empty] | 118052759 |  | 1186147 |  |  | Closed - Non-Complaint | [20 empty] | Non-Refresh | [12 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | #N/A
1260 |  |  | United States | Christ Hospital | 300205 | 4441100% |  | ECG gating not working | Work Order #: WO-04022239 Karlen James The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: Biomed Device Use: Diagnostic Expected and Unexpected Results: ECG intermittent User Impact: ECG gating is intermittent Patient Impact: gated scans not performed Current Software Version: 4.10 Diagnostic performed by Engineer : assisted Dennis with part numbers for the IBOX2 and USB converter boxes Resolution : assisted Dennis with part numbers for the IBOX2 and USB converter boxes ----------------------------------------------------------------------------- | 3 |  | Replaced USB connection with straight through cable recommended by RSE. Connected directly to Host and iBox. Installed, then de-installed iBox and restored system to original iBox. Programmed iBox and tested operation. Will monitor. | [12 empty] | 118052451 |  | 1181636 |  |  | Closed - Done | [20 empty] | Non-Refresh | [12 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | #N/A
1261 |  |  | United States | Philips Medical Capital | 300057 | 4441000% |  | error when doing short tc | Work Order #: WO-04007702 Williams John Diagnostic performed by Engineer : N/A Problem Description by Engineer : Problem description by engineer : error when doing short tc. Message failed to get DMS.. check gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: error when doing short tc Patient Impact/Outcome: no patient involved User Impact: delay Current Software Version: Not provided ----------------------------------------------------------------------------- Work Order #: WO-04007575 Sheffield, Scott The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Problem description by engineer : error when doing short tc. Message failed to get DMS.. check gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: error when doing short tc Patient Impact/Outcome: no patient involved User Impact: delay Current Software Version: Not provided Diagnostic performed by Engineer : Bridget said that the problem started about 7:45 am and the last time was 8:04 am she rebooted computers and still had problem then she power cycled system. After power cycle I had her do a chest scan of table top and system scanned with no errors. I had her try the short tc again and it did the short tc this time. System is working at this time but she wants FSE to check system. I downloaded log and noted error below. Internal Comments : Error # and/or description of error : 2021/08/02 07:57:32.920__004 S_CONT_AXIS_OVERWRITE_TARGET ERROR_NORMAL 2021/08/02 08:01:27.812__001 S_HMC_CTRL_A_MOTOR_SHORT_CIRCUIT_OR_OVERTEMP ERROR_NORMAL Repair Action : Recommendation: 1. Check for an obstruction or mechanical load problem. 2. On a BedRock couch make sure the power to the AMC is between 14VAC and 62VAC. 3. Check / replace the horizontal motor Follow Up Required : Error # and/or description of error : 2021/08/02 07:57:32.920__004 S_CONT_AXIS_OVERWRITE_TARGET ERROR_NORMAL 2021/08/02 08:01:27.812__001 S_HMC_CTRL_A_MOTOR_SHORT_CIRCUIT_OR_OVERTEMP ERROR_NORMAL Repair Action : Recommendation: 1. Check for an obstruction or mechanical load problem. 2. On a BedRock couch make sure the power to the AMC is between 14VAC and 62VAC. 3. Check / replace the horizontal motor Technician Name : Sheffield, Scott ----------------------------------------------------------------------------- Couch motor overtempt errors. Re-flash couch controller (hot fix). | 3 |  | Couch motor overtempt, Hot fix 1 | [12 empty] | 118040308 |  | 1175874 |  |  | Closed - Done |  | CTUDT00630138/630160 | [5 empty] | Engineering | [6 empty] | HF1 | [5 empty] | Refresh |  |  |  | System | Couch-SW | Horizontal Move Block | Couch-SW-Horizontal Move Block | Couch-SW-Horizontal Move Block | CTUDT00630138/630160 | Horizontal Move Block | Couch-SW | Couch-SW-Horizontal Move Block | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1262 |  |  | Argentina | INSTITUTO NACIONAL DE | 40027 | 4442200% |  | PRESENT ARTIFACTS IN THE PICTURES | Work Order #: WO-04123291 Meza, Edgardo Arturo The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : white line in survey and thick ring in helix image Diagnostic performed by Engineer : Modulo detector con problemas Follow Up Required : Detector module with problems ----------------------------------------------------------------------------- Product/Part - TDMS V2 M.Det Class B FRU Part number - 459800911281 Requested Qty - 1.0000 Product/Part - TDMS V2 M.Det Class A FRU Part number - 459800911271 Requested Qty - 1.0000 | 3 |  | System year and detail information provided and Silva escalate to DMS team | [12 empty] | 118136979 |  | 1219731 |  |  | Opened - Pending Translation |  | CTUDT00632110 | [5 empty] | Engineering | [6 empty] | DMS Module | [9 empty] | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | DMS Module | DMS-HW | DMS-HW-DMS Module | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1263 |  |  | China | 无锡市江原医院 | 345020 | 4442300% |  | 02N/Ingenuity Core128/Elite China/Rack failed to start | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : Rack failed to start Safety Questions : Case: 0118137518 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Jianqiang Last updated by: Wang Jianqiang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 扫描失败。 客户问题详细描述或故障现象： 放好病人点go 后，报错 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： " S_HMC_CTRL_A_POSITION_ERROR_LIMIT_EXCEEDED " 远程诊断与测试结果： ACS Code 25 - External problem is preventing the controller to track the position correctly.设备新设备查看AMC以及接线，伴随ACQ 报错，但是eventlog下载不下来，也连不上S1 ，现场检查 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 检查AMC 电源线等，A broken ACS cable caused this error. In a few instances the motion controller power cable L22455 part number 453566491791 was intermittent. The VP light on the ACS would flicker as the table was moved vertically. Also check 40 VAC to ACS SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support ----------------------------------------------------------------------------- | 3 |  | 453566491791 Cable, | [12 empty] | 118137518 |  | 1219903 |  |  | Opened - Pending Translation |  | TBD-AMC | [5 empty] | Engineering | [6 empty] | AMC Cable | [9 empty] | System | Couch-HW | AMC | Couch-HW-AMC | Couch-HW-AMC | TBD-AMC | AMC | Couch-HW | Couch-HW-AMC | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1264 |  |  | China | 山东省胸科医院 | 40025 | 4442100% |  | 02N/Ingenuity CT 728326/Print fault | T2 Activities : Problem Description: 客户问题概述： 打印故障 客户问题详细描述或故障现象： 间隙性打印图像失败 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： N/A 远程诊断与测试结果： N/A PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 收集bug report，上ESCALATION 建议备件12NC及名称： N/A 客户需求备注： 无 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support ----------------------------------------------------------------------------- | 5 |  | bugreport and share to USP team is under investigation | [12 empty] | 118131496 |  | 1217861 |  |  | Opened - Pending Translation |  | CTUDT00633684 | [5 empty] | Engineering | [6 empty] | Print error | [9 empty] | System | Console-SW | Imagemissing-Printer | Console-SW-Imagemissing-Printer | Console-SW-Imagemissing-Printer |  |  | Console-SW | Console-SW-Imagemissing-Printer | No | 0 | 0 | 0 | 0 | 0 | No | y | y | y | 0 | 0 | y | y | y | 0 | y | #N/A
1265 |  |  | China | 吉林省洮南市中医院 | 345006 | 4442000% |  | 02N/Ingenuity Core128/Elite China/artifacts | T2 Activities : Problem Description: 客户问题概述： 其他类型故障。 客户问题详细描述或故障现象： 伪影 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 环形伪影 远程诊断与测试结果： 探测器通道故障 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 分析原始数据，找到伪影位置，屏蔽通道，测试伪影是否消失 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 13.08.2021 08:58:06 UTC Tian Hua (86038446)服务内容： 伪影，有通道响应低，校准。 服务结果： 维修后，设备恢复正常工作。 其它重要事项： 无。 Internal Remarks : 13.08.2021 08:58:09 UTC Tian Hua (86038446) RDI Accuracy ：YES $Q7: (Confirmed reported problem&Artifact) $Q8: (Review logger and check the image，The image artifact） $Q9: (Calibration the image and test） $Q10: (Test and inspection data refer to: PA. Device returned to full functionality) $Q11 Calibrated Tools name (FLUKE287C) S/N (27420003) Next Calibration date (03/17/2022) ----------------------------------------------------------------------------- | 3 |  | Keep obseravation | [12 empty] | 118125302 |  | 1213859 |  |  | Pending Review No Further Investigation |  | ShouldbeCustomer | [5 empty] | Engineering | [6 empty] | Ring Artifact | [9 empty] | System | NotEngineering | ShouldbeCustomer | NotEngineering-ShouldbeCustomer | NotEngineering-ShouldbeCustomer | ShouldbeCustomer |  | NotEngineering | NotEngineering-ShouldbeCustomer | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | #N/A
1266 |  |  | China | Jinan City No. 4 People's Hospital | 40038 | 4441400% |  | PRS connection/02N/Ingenuity CT 728326 | Error description: PRS is not through. Service content: Replace mobile network card Service outcome: Device returns to normal working condition after servicing. Internal Remarks : 11.08.2021 03:04:07 UTC Tao Chen (86038373) “RDI Accuracy: YES" $Q7: (Confirmed reported problem, PRS disconnected) $Q8: ( Check 4G signal, found it was low) $Q9: ( Replaced 4G SIM card with new one) $Q10: (Test and inspection data refer to: PA. Device returned to full functionality) $Q11: Calibrated Tools name ( N/A) ^Current Tube installed date: N/A ^Current Tube usage: N/A ----------------------------------------------------------------------------- | 5 |  | ############# | [12 empty] | 118073654 |  | 1204418 |  |  | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | 4G SIM | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1267 |  |  | China | 四川省宜宾市第二中医院 | 345035 | 4441700% |  | 02n/Ingenuity Core128/Elite China/Unable | Problem Reported by Customer : Unable to print T2 Activities : Problem Description: Unable to print Troubleshooting Action: Unable to print Repair Action: FSE had guide customer to adjust the network, and then can print Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: NA SAP Transfer : Case Transferred to SAP | 3 |  | ############# | [12 empty] | 118080367 |  | 1199554 |  |  | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Network | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1268 |  |  | Germany | Gesundheit Nord gGmbH Klinikverbund Bremen Klinikum Bremen-Ost | 40002 | 4441400% |  | Duplicate of PR#1192687 - Automatic MPR was not finished successfully | Duplicate of PR#1192687 - Automatic MPR was not finished successfully | ZHANG, Yue Xing |  | 2021-08-09 19:40 | [14 empty] | 1197624 |  |  | Duplicate |  | CTUDT00632116 | [5 empty] | Engineering | [6 empty] | MPR | [9 empty] | System | CIRS-SW | MPR Recon | CIRS-SW-MPR Recon | CIRS-SW-MPR Recon | CTUDT00632116 | MPR Recon | CIRS-SW | CIRS-SW-MPR Recon | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1269 |  |  | Poland | Government Agency. | 341009 | 4442000% |  | Repair after PM. | On 08/11/2021, during the review, I found the need to reinstall the software. 08/13/2021 - software reinstallation. After that it works fine. The fault has been removed | 4 |  | software reinstallation. After that it works fine. The fault has been removed | [12 empty] | 118125603 |  | 1214044 |  |  | Opened |  | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Reload SW | [9 empty] | System | TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo |  |  | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | No | 0 | 0 | 0 | 0 | 0 | No | y | y | y | 0 | 0 | y | y | y | 0 | y | #N/A
1270 |  |  | Slovakia | S&T Slovakia s.r.o. | 40062 | 4442000% |  | intermittent interrupting scan | Work Order #: WO-04110787 Parts Consumed : null,459800467341,R2D Assembly Resolution : after replacement of R2D system working Diagnostic performed by Engineer : visual diagnostic, log check, discussion with helpdesk Problem Description by Engineer : intermittent interrupting scan | 5 |  | ############# | [12 empty] | 118127842 |  | 1211292 |  |  | Opened |  | CTUDT00630864 | [5 empty] | Engineering | [6 empty] | R2D | [9 empty] | System | Gantry-HW | R2D | Gantry-HW-R2D | Gantry-HW-R2D | CTUDT00630864 | R2D | Gantry-HW | Gantry-HW-R2D | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1271 |  |  | Spain | S.C.S. HOSP. GRAL. DE FUERTEVENTURA | 40020 | 4441700% |  | Problems with the image transfer. Francisco Gordillo 677980652 - 928862109 | Integration:Confirmation of order 54002137 saved ----------------------------------------------------------------------------- Work Order #: WO-04058402 [REDACTED] The system meets the specification for the performed service and is returned to use. Diagnostic performed by Engineer: I talk to [REDACTED] and [REDACTED]. Studies are not reaching the station. It is connected to network. I ask them to reconstruct the data base. It seems that it has not worked. Follow Up Required: Check the station and reload the software as required. Regulatory Questions : N/A Problem Description by Engineer: The studies are missing in the station. ----------------------------------------------------------------------------- | 3 |  | ############# | [12 empty] | 118082763 |  | 1201346 |  |  | Duplicate | [7 empty] | TBD | [6 empty] | In progress | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1272 |  |  | Spain | S.C.S. HOSP. GRAL. DE FUERTEVENTURA | 40020 | 4441700% |  | The images are not sent to the reconstruction stations. | Reported Problem: Images cannot be reconstructed on workstations. When did the breakdown occur?: Sunday When will the team be available for intervention? What is the best day or time ?: Immediately. Was the equipment in clinical use when the failure occurred ?: Yes Equipment Location: Floor 1 | 2 |  | ############# | [12 empty] | 118082580 |  | 1201218 |  |  | Evaluation | [7 empty] | TBD | [6 empty] | In progress | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1273 |  |  | United States | Philips Medical Capital | 40028 | 4442100% |  | 0711CFXXX gantry cannot comply | Integration : Confirmation of order 54022694 saved Integration : According to factory calendar 20210815 (Posting date) is not a workday (check entry) ----------------------------------------------------------------------------- Work Order #: WO-04121033 Atz, Michael The system meets the specification for the performed service and is returned to use Internal Comments : Error # and/or description of error : none Repair Action : gantry cannot comply earlier today looked at the logs no issue found they scanned all day without an issue Diagnostic performed by Engineer : gantry cannot comply earlier today looked at the logs Problem Description by Engineer : Problem description by engineer : gantry cannot comply earlier today When was the first occurrence? Within... : 13-Aug-2021 Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: No further information is available as per the customer Patient Impact/Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer Current Software Version: Not provided Resolution : gantry cannot comply earlier today looked at the logs no issue found they scanned all day without an issue ----------------------------------------------------------------------------- | 3 |  | Seems system power up no having right status | [12 empty] | 118135611 |  | 1217234 |  |  | Opened | [7 empty] | Others | [6 empty] | False alarm | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1274 |  |  | United States | Leesburg Regional Medical Center | 40022 | 4442100% |  | 0711CFXXX node needs to be changed | Problem Description by Engineer : Problem description by engineer : node needs to be changed When was the first occurrence? Within... : 13-Aug-2021 How often is issue occurring? : Continuously Malfunction area : Tube&Gen Information to support the complaint handling process Customer Function/Role: BioMed Device Use: No patient involved Expected and Unexpected results: No further information is available as per the customer Patient Impact/Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer Current Software Version: Not provided Diagnostic performed by Engineer : node needs to be changed the new IP 172.25.30.25 dose wise port 1044 Tera recon 172.25.2.155 nighthawk 172.25.30.16 irapid 172.25.2.23 Internal Comments : Error # and/or description of error : none Repair Action : node needs to be changed the new IP 172.25.30.25 dose wise port 1044 Tera recon 172.25.2.155 nighthawk 172.25.30.16 irapid 172.25.2.23 Resolution : node needs to be changed the new IP 172.25.30.25 dose wise port 1044 Tera recon 172.25.2.155 nighthawk 172.25.30.16 irapid 172.25.2.23 ----------------------------------------------------------------------------- | 3 |  | 2021-08-14 04:50 | [12 empty] | 118135412 |  | 1217035 |  |  | Opened | [7 empty] | TBD | [6 empty] | NAC-PACS SET UP | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | n | y | y | y | NAC | #N/A
1275 |  |  | United States | Philips Medical Capital | 40028 | 4442000% |  | 0705CDY71Letter P on Keyboard Not Working | Work Order #: WO-04111366 Paine Matthew The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Customer Function/Role: Technologist Device Use: Diagnostic Expected and Unexpected Results: P Key on Keyboard Does Not Work User Impact: N/A Patient Impact: N/A Current Software Version:N/A Diagnostic performed by Engineer : Clinical Applications Support Remote Michelle Hirchak was working on another case with Customer when issue was determined. Researched Part Number to Order for FSE for FVF. Follow Up Required : Issue-P Key is Not Working On the Keyboard Recommendation to FSE-Replace Defective Keyboard Cherry USB keyboard English Silver 459800918431 Internal Comments : Contacted Primary FSE Mohamed Abbas. Got Contact Information and Verified Site Address with him to Ship Keyboard Directly to Site to have and FSE Install. ----------------------------------------------------------------------------- | 3 |  | 2021-08-13 00:05 | [12 empty] | 118128240 |  | 1211692 |  |  | Evaluation | [7 empty] | Others | [6 empty] | Keyboard | [35 empty] | #N/A | bar: No to NAC
1276 |  |  | United States | Unity Health Systems | 40037 | 4441800% |  | 4.1.1 CAP questions. | Work Order #: WO-04071844 Mujica, David The system meets the specification for the performed service and is returned to use Problem description by Clinical Support Specialist : Customer Function/Role: Technologist Device Use: In use Expected and Unexpected Results: Accession Numbers split User Impact: none Patient Impact: none Current Software Version: 4.1 Diagnostic performed by Engineer : Andrea indicated that on her CAP CT Exam Card a part of the abdomen (coronals) are getting assigned to the Chest Accession number. The Exam Card manager was accessed and her CAP CT EC was assessed. All description associations were labeled correctly. It was suggested that test orders be submitted so this Exam Card/Protocol could be tested. Andrea got in touch with the PACS administrator to put test orders in and test this Protocol. After the test there was no way to determine if they were associating correctly until real orders are put in. A rebuild from factory EC's was suggested, also a deletion of the coronals and a copy paste of the recon's that associated correctly was also suggested. Andrea will make the suggested changes and monitor this issue going forward. Andrea had no further questions or need for clinical support. No rescan, no loss of data, no reinjection per Andrea. ----------------------------------------------------------------------------- | 5 |  | 2021-08-11 01:00 | [12 empty] | 118092323 |  | 1202759 |  |  | Investigation | [7 empty] | Application | [6 empty] | NAC | [35 empty] | #N/A | bar: No to NAC
1277 |  |  | United States | Philips Medical Capital | 40005 | 4441700% |  | protocol assistance | Integration : Confirmation of order 53999971 saved ----------------------------------------------------------------------------- Work Order #: WO-04062129 Curry, Dyan The system meets the specification for the performed service and is returned to use Problem description by Clinical Support Specialist : Customer Function/ Role: technologist Device Use: in use name Expected and Unexpected Results: ECM password expired User Impact: none Patient Impact: none Current Software version: 4.1.10 Diagnostic performed by Engineer : customer logged out of scanner . When she logged back , scanner gave her option to reset password. she is good to go ----------------------------------------------------------------------------- | 4 |  | 2021-08-10 04:35 | [12 empty] | 118085559 |  | 1199051 |  |  | Closed - Non-Complaint | [7 empty] | 2021-08-09 | [6 empty] | 2021-08-09 | [35 empty] | #N/A | bar: No to NAC
1278 |  |  | Spain | POLICLINICA NTRA SRA DEL ROSARIO SL | 336190 | 4442000% |  | Colchoneta y reposacabezas deteriorados | Integration : Confirmation of order 54014217 saved ----------------------------------------------------------------------------- Work Order #: WO-04110109 Parts Consumed : 1.00000,459800126661,Cushion-47.24 Inch Mattress Pad 1.00000,453566502591,PATIENT CUSHIONS EXP ----------------------------------------------------------------------------- | 5 | [14 empty] | 118127322 |  | 1215227 |  |  | Opened - Pending Translation | [7 empty] | 2021-08-13 21:05 | [6 empty] | 2021-08-12 | [5 empty] | Non-Refresh | [29 empty] | #N/A | bar: No to y
1279 |  |  | Spain | POLICLINICA NTRA SRA DEL ROSARIO SL | 336190 | 4442100% |  | The left gantry buttons are damaged and need to be replaced. | Work Order #: WO-04116786 [REDACTED] The system meets the specification for the performed service and is returned to use. Regulatory Questions : N/A Problem Description by Engineer: The left gantry buttons are damaged. Follow Up Required: Replace the keypad. Diagnostic performed by Engineer: I talked to Beatriz; the left keypad of the gantry is worn out and it is not working well, the membranes are very worn out. ----------------------------------------------------------------------------- | 3 | [14 empty] | 118132597 |  | 1214934 |  |  | Opened | [7 empty] | 2021-08-13 19:40 | [6 empty] | 2021-08-13 | [5 empty] | Non-Refresh | [29 empty] | #N/A | bar: No to y
1280 |  |  | United States | Christ Hospital | 300205 | 4442000% |  | intermittent gantry cannot comply | Integration : Confirmation of order 54015456 saved Integration : According to factory calendar 20210815 (Posting date) is not a workday (check entry) ----------------------------------------------------------------------------- Work Order #: WO-04110671 Swain Wesley The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: Biomed Device Use: Morning qa Expected and Unexpected Results: gantry cannot comply User Impact: none Patient Impact: none Current Software Version: n/a Diagnostic performed by Engineer : biomed power cycled gantry using blue button. Afterwards the system is operating normally but would like the logs looked at. Unable to login and download logs via prs. Jeff to email bugrep to me for analysis Internal Comments : talked with Jeff 3:15pm as i still have not received bugrep. he stated he is busy and will send it tomorrow morning. 8/13 Jeff did not send bugrep in and i was unable to reach him. Closing work order as system is operational. Resolution : System operating normally after power cycle using blue button by biomed Jeff. | 4 | [14 empty] | 118127739 |  | 1212520 |  |  | Evaluation | [7 empty] | 2021-08-13 03:35 | [6 empty] | 2021-08-12 | [5 empty] | Non-Refresh | [29 empty] | #N/A | bar: No to y
1281 |  |  | United States | Advent Health Wesley Chapel | 300106 | 4441900% |  | Scans taking up to 10 minuets to get start scan button |  | SHU, Tim | [16 empty] | 1212298 |  |  | Investigation | [7 empty] | TBD | [6 empty] | Duplicate | [5 empty] | Refresh | [6 empty] | - | - |  |  | 0 | - | No | y | y | y | 0 | 0 | No | y | y | y | y | 0 | n | y | y | y | y | #N/A
1282 |  |  | United States | AdventHealth Zephyrhills | 300106 | 4442000% |  | 0711CDN84 Long wait before scan | Integration : Confirmation of order 54014453 saved Integration : According to factory calendar 20210814 (Posting date) is not a workday (check entry) ----------------------------------------------------------------------------- Work Order #: WO-04109858 12-Aug-2021 Sher Mark The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Problem description by engineer : Long wait before scan would initialize When was the first occurrence? Within... : 11-Aug-2021 How often is issue occurring? : Multiple times per week Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Scan does not start immediately Patient Impact/Outcome: No further information is available as per the customer User Impact: have to wait for a long time to start a scan Current Software Version: 4.1.7.25030 Diagnostic performed by Engineer : reviewed logs in ATP, created CFM and uploaded bug reports ----------------------------------------------------------------------------- Work Order #: WO-04109412 Ali, Adam The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Problem description by engineer : Long wait before scan would initialize When was the first occurrence? Within... : 11-Aug-2021 How often is issue occurring? : Multiple times per week Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Scan does not start immediately Patient Impact/Outcome: No further information is available as per the customer User Impact: have to wait for a long time to start a scan Current Software Version: 4.1.7.25030 Diagnostic performed by Engineer : When customer click the 'go' button, they have to wait about 3 to 10 mins before scan would start. Other times they get 'System is still initializing. Please wait?', another time While running air calibration the system froze and did not complete air calibrations. Downloaded 3 different bugreports since yesterday. Logs do not show a hardware issue, unable to download host logs for further analysis. Internal Comments : Error # and/or description of error : "Problem Description" When clicking go, the scanner freezes for 3 to 10 mins Time of Event" "8/12/2021 6:56:00 AM 2021/08/12 06:54:32.000__001 CirsEvt_ImgNodeMgr_699 CIRS_Warning Pre Series CIRS Event Log Unchecked 2021/08/12 06:57:35.000__100 CirsEvt_ImgNodeMgr_307 CIRS_Warning Pre Series CIRS Event Log Unchecked "Problem Description" Initialization issue "Time of Event" "8/11/2021 11:39:00 AM" 2021/08/11 11:37:47.812__001 HOST_USR_POPUP USR_POPUP Pre Series Usplog.log Unchecked Showing message to user : System is still initializing. Please wait? "Problem Description" While running air calibration the system froze and did not complete air calibrations "Time of Event" "8/11/2021 7:02:00 AM" 2021/08/11 07:00:01.000__001 CirsEvt_ImgNodeMgr_699 CIRS_Warning Pre Series CIRS Event Log Unchecked Host logs fail to download d:\>c:\pms\system\unzip -j D:\PMS\BugRep\packages\*_cdm_test.zip -d D:\RSE\Host_Evts cd d:\ c:\pms\system\unzip -j D:\PMS\BugRep\packages\*_cdm_test.zip -d D:\RSE\Host_Evts unzip: cannot find any matches for wildcard specification "D:/PMS/BugRep/packages/*_cdm_test.zip". No zipfiles found. Repair Action : Check host event logs Run test scans and try to reproduce issue Check disks, host hardware, fans, filters , cooling for host tower Clean database Software reinstall if it software shows errors or running slow. Follow Up Required : Error # and/or description of error : "Problem Description" When clicking go, the scanner freezes for 3 to 10 mins Time of Event" "8/12/2021 6:56:00 AM 2021/08/12 06:54:32.000__001 CirsEvt_ImgNodeMgr_699 CIRS_Warning Pre Series CIRS Event Log Unchecked 2021/08/12 06:57:35.000__100 CirsEvt_ImgNodeMgr_307 CIRS_Warning Pre Series CIRS Event Log Unchecked "Problem Description" Initialization issue "Time of Event" "8/11/2021 11:39:00 AM" 2021/08/11 11:37:47.812__001 HOST_USR_POPUP USR_POPUP Pre Series Usplog.log Unchecked Showing message to user : System is still initializing. Please wait? "Problem Description" While running air calibration the system froze and did not complete air calibrations "Time of Event" "8/11/2021 7:02:00 AM" 2021/08/11 07:00:01.000__001 CirsEvt_ImgNodeMgr_699 CIRS_Warning Pre Series CIRS Event Log Unchecked Host logs fail to download d:\>c:\pms\system\unzip -j D:\PMS\BugRep\packages\*_cdm_test.zip -d D:\RSE\Host_Evts cd d:\ c:\pms\system\unzip -j D:\PMS\BugRep\packages\*_cdm_test.zip -d D:\RSE\Host_Evts unzip: cannot find any matches for wildcard specification "D:/PMS/BugRep/packages/*_cdm_test.zip". No zipfiles found. Repair Action : Check host event logs Run test scans and try to reproduce issue Check disks, host hardware, fans, filters , cooling for host tower Clean database Software reinstall if it software shows errors or running slow. What are the next Troubleshooting steps : Check host event logs Run test scans and try to reproduce issue Check disks, host hardware, fans, filters , cooling for host tower Clean database Software reinstall if it software shows errors or running slow. Technician Name : Ali, Adam ----------------------------------------------------------------------------- | 3 |  | Check the software version, RAM size, and if reinstall software, collect bugreport for analysis | [12 empty] | 118126749 |  | 1210743 |  |  | Evaluation |  | TBD_Bugreport | [5 empty] | Engineering | [6 empty] | Long intialize | [5 empty] | Refresh |  |  |  | System | TBD-Bugreport | TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | TBD_Bugreport |  | TBD-Bugreport | TBD-Bugreport-TBD_Bugreport | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | #N/A
1283 |  |  | United States | Select Specialty Hospitals | 310051 | 4441800% |  | how to burn CD | The system meets the specification for the performed service and is returned to use Diagnostic performed by Engineer : Went over steps to copy images to crd, open cdr folder and press record. Also went over direct record option. Advised customer to clear CDR folder after each cd burning Problem description by Clinical Support Specialist : Customer Function/ Role: technologist Device Use: in use name Expected and Unexpected Results: how to burn cd User Impact: none Patient Impact: none Current Software version: 4.1.10 | 4 | [14 empty] | 118089837 |  | 1202109 |  |  | Evaluation | [7 empty] | Application | [6 empty] | NAC | [5 empty] | Non-Refresh | [29 empty] | #N/A | bar: No to y
1284 |  |  | United States | Wellstar Kennestone Hospital | 68117 | 4441700% |  | the gating not working on EKG | IP & Issue: 78989232 the gating not working on EKG Internal Remarks : Parts ID of IBOX to Console extender cable. Next Step Action : IBox looks that has a problem , approved to order . Next Step Action : FSE is trying to JTAG the new IBox , the file doesn't show on the dvd. He will take the file from the Host computer . Internal Remarks : Emailed the 14.7 Xilinx zip file to Bill Internal Remarks : Approved TPR for 459800861311 - Xilinx 14.7 service Programming Kit FSE needs the 14.7 DVD, but it is not available for download via InCenter. Advised the FSE to try and get it from Xilinix directly, but it's a 15 GB download. ----------------------------------------------------------------------------- Work Order #: WO-04072829 11-Aug-2021 Davis Paul Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- Work Order #: WO-04062228 10-Aug-2021 Scheyer William Problem Description by Engineer : EKG function not working Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- Work Order #: WO-04072748 10-Aug-2021 Bourne Barrett The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, no follow-up requested Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : Assisted with running a new IBOX cable. ----------------------------------------------------------------------------- Work Order #: WO-04058453 09-Aug-2021 Scheyer William The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : EKG function not working Diagnostic performed by Engineer : Found EKG function intermittent. Ordered replacement cable. Follow Up Required : replace IBOX2 USB cable ----------------------------------------------------------------------------- Product/Part - I-BOX-2 BASE BOARD ASSY Part number - 459800166651 Actual Qty - 1.0000 Product/Part - ECG Monitor Trigger 7M Ca Part number - 455012301901 Actual Qty - 1.0000 Part/Product - USB EXTENDER CABLE FROM IBOX TO CONSOLE Part Number - 459801731761 Actual Quantity - 1.00000 | 3 | [14 empty] | 118082792 |  | 1199078 |  |  | Investigation | [7 empty] | 2021-08-10 04:50 | [6 empty] | 2021-08-09 | [5 empty] | Non-Refresh | [29 empty] | #N/A | bar: No to y
1285 |  |  | China | Shandong Province Chest Hospital | 40025 | 4442500% |  | Wechat：02N/Ingenuity CT/Machine fault | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : Machine fault Next Step Action : NTD：微信报修 T2 Activities : Problem Description: 客户问题概述： 其他类型故障。 客户问题详细描述或故障现象： 调整网络 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： NA 远程诊断与测试结果： 客户告知明天全院换PACS，调整网络，不需要远程，需要工程师明天到场配合 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 配合医院调整网络，连PACS， External Remarks : 18.08.2021 08:20:52 UTC SUP JCOUSER (86038405)服务内容：配合医院更换PACS，更改IP等工作。 服务结果：配置后，设备正常使用。 Internal Remarks : 23.08.2021 02:42:15 UTC Zhaokun Huang (86038405) "RDI Accuracy: YES." $Q7: ( Confirmed reported problem & Need Config PACS/Worklist.) $Q8: ( Customer request engineer onsite configuration PACS/Worklist.) $Q9: ( Config PACS/Worklist and backup. ) $Q10: ( Test and inspection data refer to: NO PA Required . Device returned to fully functionality.) $Q11: Calibrated Tools name (N/A) ^Current Tube installed date: N/A ^Current Tube usage: N/A ----------------------------------------------------------------------------- | 5 |  | 客户告知明天全院换PACS，调整网络，不需要远程，需要工程师明天到场配合 | [12 empty] | 118150523 |  | 1229915 |  |  | NAC | [7 empty] | Customer | [6 empty] | PACS Setup | [35 empty] | #N/A | bar: No to NAC
1286 |  |  | Poland | Government Agency. | 341008 | 4441400% |  | Starting the tomography computer and application training. |  | 4 |  | App training | [12 empty] | 118073782 |  | 1221599 |  |  | NAC | [7 empty] | Application | [6 empty] | Application | [35 empty] | #N/A | bar: No to NAC
1287 |  |  | Japan | Independent administrative institution National Hospital Organization | 341013 | 4442700% |  | 01DS:KKI: Cannot log in to the Exam Manager. |  | 4 |  | Reset the IT Password and reset the Exam card editing password. | [12 empty] | 118162854 |  | 1237710 |  |  | Duplicate | [7 empty] | Customer | [6 empty] | Duplicate | [35 empty] | #N/A | bar: No to NAC
1288 |  |  | Japan | Independent administrative institution National Hospital Organization | 341013 | 4442500% |  | 01DS: TM: CIRS won't turn on.No response even when the power switch is pressed. | Work Order #: WO-04142507 18-Aug-2021 Takada, Ushio Regulatory Questions : N/A Problem Description by Engineer: Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer: N/A ----------------------------------------------------------------------------- Work Order #: WO-04142290 Yamamoto, Yuko Regulatory Questions : N/A Problem Description by Engineer: Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer: N/A ----------------------------------------------------------------------------- Work Order #: WO-04142275 Miyakawa, Takeyuki The system does not meet the specification for the performed service and is not returned to use Problem Description by Engineer: CIRS won't turn on.No response even when the power switch is pressed. Internal Comments: [Subject] 01DS: TM: Mr. Yamamoto: CIRS won't turn on.No response even when the power switch is pressed. [Internal Comments] Response I determined that the system went down. Shipped part: 459801493161: CT Gen5.3 IDose CIRS, Z8 QTY: 1 ZREP/Traceable BackOrder 459801647771: HP Power Supply Z8 G4 1125W QTY: 1 ZSID BackOrder 459801633881: 64GB DDR4-2666 (8x8GB) RegRAM QTY: 1 ZSID BackOrder Dispatched: 1 RC FSE Registration Information: Name of FSE 1 who visited the site: Yutaka Yamamoto FSE 1 visit date and time: 2021/08/189:30 FSE 1 hours worked: 2 - Name of FSE 2 who visited the site Shio Takada FSE2 visit date and time 2021/08/189:30 FSE2 hours worked 2 Dispatch instructions: Dispatch instruction #1: Shipping conditions: NBD (mixed loading, next day or thereafter) Delivery date: 19 August 2021. No time specified. Shipping destination: Kyoto Office To be collected by: Yutaka Yamamoto Phone number: 090-1730-2299 Other: If the item is not in stock in Japan , please place an overseas order Person making request: Diagnostic performed by Engineer: A call came in from Mr. Yamamoto at 8:04 Call description: CIRS won't turn on.Even when the power switch is pressed the system does not respond. CT-Host was on Response I determined that the system went down. ----------------------------------------------------------------------------- | 2 |  | 459801493161: CT Gen5.3 IDose CIRS, Z8 QTY: 1 ZREP/Traceable BackOrder 459801647771: HP Power Supply Z8 G4 1125W QTY: 1 ZSID BackOrder 459801633881: 64GB DDR4-2666 (8x8GB) RegRAM QTY: 1 ZSID BackOrder Product/Part - CT Gen5.3 Hawk CIRS Master Z8 Part number - 459801592003 This was a system down case due to CIRS hardware problem. The CIRS could not be turned on, nothing happened by pressing power button. FSE found the problem was caused by the mother board, but the Z8 computer (459801493161 CT Gen5.3 IDose CIRS, Z8) was out of stock in local warehouse. So while waiting for the delivery from SPS, they installed the one for CT7500 (459801592003 CT Gen5.3 Hawk CIRS Master Z8) instead. So the system is currently available, but they need to replace the computer again. Therefore the case is still open. | [12 empty] | 118153000 |  | 1230114 |  |  | Evaluation | [7 empty] | TBD | [6 empty] | CIRS | [9 empty] | TBD | [8 empty] | No | 0 | y | y | y | y | No | y | y | y | y | y | n | y | y | y | y | #N/A
1289 |  |  | China | Dongguan traditional chinese and western medicine Integrated Hospital Hospital | 345013 | 4442500% |  | Duplicate to PR#1223313 default file-host*. XML files in D:\pms\log\data up to 300G+N26:T26 made D disk full in one or two days. |  | ZHANG, Yue Xing |  | 2021-08-17 | [14 empty] | 1224189 |  |  | Duplicate |  | CTUDT00633269 | [5 empty] | Engineering | [6 empty] | 2021-08-16 | [9 empty] | System | Console-SW | DiskFull | Console-SW-DiskFull | Console-SW-DiskFull | CTUDT00633269 |  | Console-SW | Console-SW-DiskFull | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | #N/A
1290 |  |  | China | 济南市第四人民医院 | 40038 | 4443000% |  | 02N/Ingenuity CT 728326/CT BOX Button failure | T2 Activities : Problem Description: 客户问题概述： CT BOX按键故障 客户问题详细描述或故障现象： CT BOX按键故障 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： N/A 远程诊断与测试结果： N/A PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 更换CT BOX 建议备件12NC及名称： "CT BOX ASSEMBLY_BEST : 459801705022" 客户需求备注： 无 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : PPD 459801705022 *1 已安排备件24H手工单出货 陈涛-86-18560023293 | 3 |  | CT BOX ASSEMBLY_BEST : 459801705022 $Q8: ( Do couch down movement with CT box, found it was failed intermittently) $Q9: ( Replaced CT box with new one, part number was 459801705022) | [12 empty] | 118174884 |  | 1243873 |  |  | Opened - Pending Translation |  | CTUDT00632185 | [5 empty] | Engineering | [6 empty] | CT BOX | [9 empty] | System | Gantry-HW | CT Box | Gantry-HW-CT Box | Gantry-HW-CT Box | CTUDT00632185 | CT Box | Gantry-HW | Gantry-HW-CT Box | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1291 |  |  | China | Hebei Province Tangshan City Workers Hospital | 40044 | 4442400% |  | 02N/Ingenuity CT 728326/artifact | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : artifact Safety Questions : Case: 0118138744 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingsong Last updated by: Lu Mingsong 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 图像放射状伪影 客户问题详细描述或故障现象： 图像放射状伪影 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： N/A 远程诊断与测试结果： N/A PRS TR状态： 无PRS。 Repair Action: 建议现场维修步骤： 执行Bad Detector Test,确定坏模块位置，更换探测器模块 建议备件12NC及名称： 12NC: 459800594911, Description: TDMS Module Class B FRU 客户需求备注： 无 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : PPD备件TDMS Module Class B FRU 459800594911 1需要TSE加批。批复安排24H到场地，请先派工，FSE：徐国梁-86-13869524913 T2 Activities : Problem Description: 客户问题概述： 图像放射状伪影 客户问题详细描述或故障现象： 图像放射状伪影 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： N/A 远程诊断与测试结果： N/A PRS TR状态： 无PRS。 Repair Action: 建议现场维修步骤： 做原始数据分析，执行Bad Detector Test。 建议备件12NC及名称： N/A 客户需求备注： 要求工程师到现场检查 External Remarks : 19.08.2021 12:55:13 UTC JunLei Wang (86039063) 服务内容：设备伪影，更换探测器 服务结果：设备恢复正常 Internal Remarks : 19.08.2021 12:55:16 UTC JunLei Wang (86039063) "RDI Accuracy: Yes" $Q7: (Confirmation & Description of problem) Confirmed reported problem, have artifact $Q8: (Troubleshooting & Results) check the detector ,need to replace $Q9: (Repair steps) repalace the detector $Q10: (Test and Inspection data & Device status) Test and inspection data refer to: PA. Device returned to full functionality $Q11: Calibrated Tools name to PA ----------------------------------------------------------------------------- | 3 |  | 12NC: 459800594911, Description: TDMS Module Class B FRU | [12 empty] | 118138744 |  | 1222990 |  |  | Investigation |  | CTUDT00632110 | [5 empty] | Engineering | [6 empty] | DMS Module | [9 empty] | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | DMS Module | DMS-HW | DMS-HW-DMS Module | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1292 |  |  | China | Jilin Province Songyuan City Central Hospital | 345008 | 4442500% |  | 02N/Ingenuity Core128/Elite China/ scan aborted | T2 Activities:Problem Description: Summary of customer's problems: Scanning failed. Detailed description of Customer's problems or malfunctions: Can't scan after maintenance. The maintenance engineer called asking to look at the log It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: Scan start failure and reboot to restore. Remote troubleshoot and test result: "no frames were generated by DMS during scan. Watch for more occurrences of this error. It is possible to get infrequent occurrences of this error while the scanner is operating correctly." PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: "1. Check the CIRS log for ACQ status and information. This maybe a data path problem. 2. Connect the optical cable directly from the ACQ PCB to the RCOM. Try to make scans without rotation to see if scanning is possible without above error. 3. Clean optical and slipring. 4. Check data path. Run VD 5. On a BB, check Optical Transceiver Rotor and Stator. T2 Activities : Problem Description: scan aborted Troubleshooting Action: no frames were generated by DMS during scan. Repair Action: reboot gantry, the device return full function Test/Inspection: Functional Check PASSED. The device returned to full functionality. Test Equipment: NA | 3 |  | reboot gantry, the device return full function | [12 empty] | 118149741 |  | 1229404 |  |  | Closed - No Further Investigation |  | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Scan abort | [9 empty] | System | DMS-HW | NoEnoughInfo | DMS-HW-NoEnoughInfo | DMS-HW-NoEnoughInfo |  |  | DMS-HW | DMS-HW-NoEnoughInfo | No | 0 | 0 | 0 | 0 | 0 | No | y | y | y | 0 | 0 | y | y | y | 0 | y | #N/A
1293 |  |  | China | Dongguan City Integrated Traditional Chinese and Western medicine Hospital | 345013 | 4442400% |  | 02N/Ingenuity Core128/Elite China/escalation | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 02N/Ingenuity Core128/Elite China/escalation Safety Questions:Case: 0118139355 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Bo Last updated by: Wang Bo 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer's problems: Other types of error. Detailed description of Customer's problems or malfunctions: Storage space for disc D can be full in one day. It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: The log has no errors. Remote troubleshoot and test result: Restart but didn't work.Have communicated with the engineer. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Suggest escalation, ask for TSE help. We recommend spare parts 12NC and names: N/A SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support ----------------------------------------------------------------------------- system is generating more number of logs because Verbose logging is enabled in the system. Because of this, every 2-3 min a single defaultFile is getting generated. This is expected as per the design and the usage of the system. Please change verbosity logging back to event to not to generate more number of logs. This will be enabled to get more logs for debugging purpose. Please follow following steps to revert this change. 1. Login as service user. 2. Launch c:\Pms\System\LogConfigurator.exe 3. Change the value Verbose to Event in the Default Level: Combo box of Source setting. 4. Save and Exit. | 5 |  | Please change verbosity logging back to event to not to generate more number of logs. This will be enabled to get more logs for debugging purpose. CTUDT00633269 fix in SP2 | [12 empty] | 118139355 |  | 1223313 |  |  | Opened |  | CTUDT00633269 | [5 empty] | Engineering | [6 empty] | Verbose logging | [9 empty] | System | Console-SW | DiskFull | Console-SW-DiskFull | Console-SW-DiskFull | CTUDT00632261 | DiskFull | Console-SW | Console-SW-DiskFull | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1294 |  |  | China | Dongguan City Integrated Traditional Chinese and Western medicine Hospital | 345013 | 4442400% |  | Wechat:02N/Ingenuity Core128/Elite China/ECG fault | BU Q&R Feedback: No Reason: Corrective Maintenance Safety Questions:Case: 0118139175 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Bai Jingwen Last updated by: Bai Jingwen Problem Reported by Customer:ECG fault T2 Activities:Problem Description: Summary of customer's problems: ECG related failure. Detailed description of Customer's problems or malfunctions: ECG waveform interference. It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: The log has no errors. Remote troubleshoot and test result: Restart but didn't work.Have communicated with the engineer. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: It is recommended to replace the parts to eliminate the interference of ECG waveforms. We recommend spare parts 12NC and names: PPD 455012301901, 453567036891, 459800485201, 455012303732 SAP Transfer:Dispatched for On site Support SAP Transfer:Case Transferred to SAP Next Step Action: PPD 459800485201*1+453567036891*1+455012301901*1 24H FSE: Li Zongxian -86-13751567086 arranged shipment ----------------------------------------------------------------------------- | 3 |  | $Q7: (Confirmed the reported problem, ECG wave abnormal) $Q8:( The wave was normal in ECG monitor, considered the cable to ibox and the PS of ibox) $Q9: (Replaced the 5V PS and ecg cable) 更换电源和信号线 | [12 empty] | 118139175 |  | 1223188 |  |  | Opened | [7 empty] | Others | [6 empty] | ECG | [12 empty] | - | - |  |  | 0 | - | No | y | y | y | 0 | 0 | No | y | y | y | y | 0 | n | y | y | y | y | #N/A
1295 |  |  | China | 黑龙江省桦川县人民医院 | 345018 | 4442900% |  | 02N/Ingenuity Core128/Elite China/Add node | Problem Reported by Customer : 02N/Ingenuity Core128/Elite China T2 Activities : Problem Description: 客户问题概述： 添加相机节点 客户问题详细描述或故障现象： 添加相机节点 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 远程指导客户用olevel_service账号登录失败 远程诊断与测试结果： N/A PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 现场恢复olevel_service账号密码，添加节点 建议备件12NC及名称： N/A 客户需求备注： 客户很着急，接到工单后及时联系客户，CT.Dr.Zhang：15246304333 Safety Questions : Case: 0118172888 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingsong Last updated by: Lu Mingsong | 3 |  | 现场恢复olevel_service账号密码，添加节点 | [12 empty] | 118172888 |  | 1242806 |  |  | Opened - Pending Translation | [7 empty] | Customer | [6 empty] | PACS Setup | [35 empty] | #N/A | bar: No to NAC
1296 |  |  | China | 黑龙江省桦川县人民医院 | 345018 | 4442700% |  | 02N/Ingenuity Core128/Elite China/Spare parts need to be replaced | BU Q&R Feedback: No SAP Customer Symptom Code: CENV Customer Symptom Code: Customer Environmental Concern Reason: Corrective Maintenance Engineer Failure Code: Environmental Cause Repair Activity Code: Configuration Change Problem Reported by Customer : Spare parts need to be replaced. Safety Questions : Case: 0118161518 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Lili Last updated by: Lu Lili 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Troubleshooting Action: wrong case ,should be workstation,They didn't have Philips workstation Repair Action: wrong case ,should be workstation,They didn't have Philips workstation Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | wrong case ,should be workstation,They didn't have Philips workstation | [12 empty] | 118161518 |  | 1236679 |  |  | Opened - Pending Translation | [7 empty] | Service | [6 empty] | NAC | [35 empty] | #N/A | bar: No to NAC
1297 |  |  | China | 苏州广慈肿瘤医院有限公司 | 345019 | 4442600% |  | WeChat：02N/Ingenuity Core128/Elite China/gantry noise | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : gantry noise SAP Transfer : Case Transferred to SAP Next Step Action : PPD 453566492401*1 453567003451*1 12H FSE:徐昊天-86-18761995321 已安排发货 External Remarks : 20.08.2021 09:40:00 UTC Haotian Xu (86090280) 服务内容： 旋转异响，更换皮带 服务结果： 服务后，设备恢复正常使用。 External Remarks : 20.08.2021 11:08:11 UTC Yang Liu (86036528)服务内容： 旋转异响，更换皮带 服务结果： 服务后，设备恢复正常使用。 Safety Questions : Case: 0118153519 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Zixuan Last updated by: Wang Zixuan 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 设备异响。 客户问题详细描述或故障现象： 机架旋转时异响。 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： log无报错。 远程诊断与测试结果： 重启无效。 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议检查异响来源，重点检查机架压轮和rotor皮带。 建议备件12NC及名称： PPD 453567003451,453566492401 SAP Transfer : Dispatched for Onsite Support Next Step Action : XU HAOTIAN Internal Remarks : 20.08.2021 11:37:43 UTC Yang Liu (86036528) RDI Accuracy: Yes $Q7 (Confirmed reported problem , The problem is the gantry has abnormal sound) $Q8: ( check the gantry ，found the belt has abnormal sound, ) $Q9: (replace the belt according to service manual) $Q10: (Test and inspection scan refer to PA . device returned to full functionality $Q11: Calibrated Tools name (refer to PA ) Internal Remarks : 21.08.2021 06:32:15 UTC Haotian Xu (86090280)RDI Accuracy: Yes $Q7 (Confirmed reported problem , The problem is the gantry has abnormal sound) $Q8: ( check the gantry ，found the belt has abnormal sound, ) $Q9: (replace the belt according to service manual) $Q10: (Test and inspection scan refer to PA . device returned to full functionality $Q11: Calibrated Tools name (refer to PA ) ----------------------------------------------------------------------------- | 3 |  | $Q8: ( check the gantry ，found the belt has abnormal sound, ) $Q9: (replace the belt according to service manual) Check with Yiming to collect customer feedback of the issue | [12 empty] | 118153519 |  | 1232193 |  |  | Opened - Pending Translation |  | ECR-077012 | [5 empty] | Engineering | [6 empty] | Belt | [9 empty] | System | Gantry-HW | Belt | Gantry-HW-Belt | Gantry-HW-Belt | ECR-077012 | Belt | Gantry-HW | Gantry-HW-Belt | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1298 |  |  | China | 正镶白旗医院 | 345033 | 4442700% |  | 02N/Ingenuity Core128/Elite China/consulting scan parameter | BU Q&R Feedback: No SAP Customer Symptom Code: MARF Customer Symptom Code: Monitor Alarm Failure Reason: Corrective Maintenance Engineer Failure Code: Incompatible Use Repair Activity Code: Instructed Customer Safety Questions : Case: 0118163251 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Jianqiang Last updated by: Wang Jianqiang 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Problem Reported by Customer : 02N/Ingenuity Core128/Elite China T2 Activities : Problem Description: consulting scan slice Troubleshooting Action: information to customer how to change slice Repair Action: instruct to customer Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: NA SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | instruct customer how to change slice | [12 empty] | 118163251 |  | 1238100 |  |  | Opened - Pending Translation | [7 empty] | Application | [6 empty] | NAC | [35 empty] | #N/A | bar: No to NAC
1299 |  |  | China | 陕西省安康中医医院 | 345041 | 4442600% |  | WeChat:02N/Ingenuity Core128/Elite China/noisy | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : noisy Safety Questions : Case: 0118154307 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Weng Jinming Last updated by: Weng Jinming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 飞利浦64排CT的机 是8月12号 做过一次保养 但是目前做病人 升降床的时候 有噪声 类似于有东西卡在里面的感觉 设备本身技术问题： 是。 Troubleshooting Action: 远程诊断与测试结果： 怀疑滑轨，丝杠问题 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议工程师到场检查床丝杠，滑轨，外壳。。。 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 19.08.2021 15:27:11 UTC Jun Liu (86039454)服务原因：床水平运动阻力大 服务内容：修正被撞部分水平 服务结果：设备工作正常 Internal Remarks : 19.08.2021 16:18:21 UTC Jun Liu (86039454) RDI Accuracy :Yes $Q7: (Confirmed reported problem. & The couch has unexpected noise when in and out) $Q8: (1.Check the couchhorizontal lead screw and found the . & We need mechanical repair to so that the patient table horizontal move normally .) $Q9: (1. Mechanical repair and calibraiton the CT. 2. Test scan for CT was OK. ) $Q10: (Test and inspection data refer to PA. & Device returned to full functionality.) $Q11：Calibrated Tools name (Multimeter Fluke 287) S/N (22940012) Next Calibration date 06/17/2022 Internal Remarks : 19.08.2021 16:21:49 UTC Jun Liu (86039454)RDI Accuracy :Yes $Q7: (Confirmed reported problem. & The couch has unexpected noise when in and out) $Q8: (1.Check the couch horizontal lead screw and found the left LH subframe cover support was knocked out of shape. & We need mechanical repair to so that the patient table horizontal move normally .) $Q9: (1. Mechanical repair and calibraiton the CT. 2. Test scan for CT was OK. ) $Q10: (Test and inspection data refer to PA. & Device returned to full functionality.) $Q11：Calibrated Tools name (Multimeter Fluke 287) S/N (22940012) Next Calibration date 06/17/2022 ----------------------------------------------------------------------------- | 3 |  | $Q8: (1.Check the couch horizontal lead screw and found the left LH subframe cover support was knocked out of shape. & We need mechanical repair to so that the patient table horizontal move normally .) | [12 empty] | 118154307 |  | 1232608 |  |  | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Couch movement | [35 empty] | #N/A | bar: No to NAC
1300 |  |  | China | 宣城市骨科医院 | 345042 | 4442700% |  | 02N/Ingenuity Core128/PACS issue | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Safety Questions : Case: 0118162391 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Cao Last updated by: Liu Cao 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Problem Reported by Customer : 02N/Ingenuity Core128/PACS issue T2 Activities : Problem Description: PACS issue, transfer stuck Troubleshooting Action: PACS issue, transfer stuck Repair Action: Adjust setting with pacs engineer Test/Inspection: N/A Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | Troubleshooting Action: PACS issue, transfer stuck Repair Action: Adjust setting with pacs engineer | [12 empty] | 118162391 |  | 1237263 |  |  | Opened - Pending Translation | [7 empty] | Others | [6 empty] | PACS Vendor | [35 empty] | #N/A | bar: No to NAC
1301 |  |  | China | 鄄城马爱云妇儿医院 | 345044 | 4442700% |  | 02N/Ingenuity Core128/Elite China/Change the information | BU Q&R Feedback: No SAP Customer Symptom Code: HANG Customer Symptom Code: Hang/Crash Reason: Corrective Maintenance Engineer Failure Code: Configuration Error Repair Activity Code: Configuration Change Problem Reported by Customer : Change the information Safety Questions : Case: 0118161710 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Qian Wang Last updated by: Qian Wang T2 Activities : Problem Description: 客户需要修改医院抬头 Troubleshooting Action: 远程告诉客户如何修改医院名称在软件上。 Repair Action: 远程告诉客户如何修改医院名称在软件上。 Test/Inspection: customer confirm can scan patient and image ok.functional Check PASSED. Device returned to full functionality. Test Equipment: N/A | 3 |  | 远程告诉客户如何修改医院名称在软件上。 | [12 empty] | 118161710 |  | 1236792 |  |  | Opened - Pending Translation | [7 empty] | Application | [6 empty] | NAC | [35 empty] | #N/A | bar: No to NAC
1302 |  |  | Japan | 独立行政法人国立病院機構 | 341013 | 4442800% |  | 01DS: KKI: Cannot log in to Examcard Manager | Work Order #: WO-04157913 20-Aug-2021 Yamamoto, Yuko The system meets the specification for the performed service and is returned to use Resolution: Reset the IT Password and reset the Exam card editing password. | 5 |  | Reset the IT Password and reset the Exam card editing password. | [12 empty] | 118167220 |  | 1240804 |  |  | Evaluation - Pending Translation | [7 empty] | Customer | [6 empty] | Password | [35 empty] | #N/A | bar: No to NAC
1303 |  |  | Japan | Independent administrative institution National Hospital Organization | 341013 | 4442700% |  | 03CS:TM:Via FSE.It stopped due to an error during SortTC. >> Can be re-implemented. >> SoftArcRe/"S_GEN_FAUL [Truncated] | Integration : Confirmation of order 54043997 saved ----------------------------------------------------------------------------- Work Order #: WO-04152996 Miyakawa, Takeyuki The system meets the specification for the performed service and is returned to use. Problem Description by Engineer : via FSE.It stopped due to an error during SortTC. >> Can be re-implemented. >> SoftArcRe/"S_GEN_FAULT_ANODE_TUBE_ARC/S_GEN_FAULT_CATHODE_TUBE_ARC" Diagnostic performed by Engineer : via FSE.It stopped due to an error during SortTC. >> Can be re-implemented. >> SoftArcRe/"S_GEN_FAULT_ANODE_TUBE_ARC/S_GEN_FAULT_CATHODE_TUBE_ARC" Internal Comments: [Subject] 03CS:TM:Via FSE:Via FSE.It stopped due to an error during SortTC. >> Can be re-implemented. >> SoftArcRe/"S_GEN_FAULT_ANODE_TUBE_ARC/S_GEN_FAULT_CATHODE_TUBE_ARC" [Internal Comments] Results of checking the log: 2021/08/19_07:53:32.059 <*HV*> szInfo: SoftArcReg 0x211* P1:0/P2: 2021/08/19_07:53:32.106 <*HV*> Type: NORMAL ErrNum: S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT szInfo: P1:0/P2: 2021/08/19_07:53:32.124 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_ANALOG_FAULT szInfo:2:5 Analog Fault P1:2/P2:32 2021/08/19_07:53:32.126 [CANCEL] CStateShot:Series 3 cancelled, due to GANTRY_CANCEL_SOURCE_NORMAL_ERROR, in state CStateShot 2021/08/19_07:53:32.134__03 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_ANODE_TUBE_ARC szInfo: 5:0 APM Arc P1:5/P2: 2021/08/19_07:53:32.139 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_CATHODE_TUBE_ARC szInfo: 5:5 CPM Arc P1:5/P2: 2021/08/19_07:53:34.164__01 <*ERROR*> Type: NORMAL ErrNum: S_MASTER_SEQUENCER_EXIT_STATE_ERR P1:4096/P2:8369 [4/8 Elements] Q5. How did the examination end? None. Q6. Did any mistaken findings (mistaken diagnosis or treatment) result from the problem? No. Q7. Was there any unintended unnecessary exposure? No Q8 Artifacts? None ----------------------------------------------------------------------------- | 4 |  | It stopped due to an error during SortTC. >> Can be re-implemented. | [12 empty] | 118161118 |  | 1234044 |  |  | Closed - No Further Investigation |  | CTUDT00630085 | [5 empty] | Engineering | [6 empty] | Tube Arc | [9 empty] | System | IC-HW | Tub Arc | IC-HW-Tub Arc | IC-HW-Tub Arc | CTUDT00630085 | Tub Arc | IC-HW | IC-HW-Tub Arc | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1304 |  |  | Japan | 独立行政法人国立病院機構 | 341015 | 4443100% |  | 01CS: RS: The operation became slow, such as updating the tracker image on the Bolus Track around 11:30 on the 18th. | "Work Order #: WO-04172697 Satake, Rui The system meets the specification for the performed service and is returned to use Problem Description by Engineer: The operation became slow, such as updating the tracker image on the Bolus Track around 11:30 on the 18th. Resolution: There was no default log left, but there were no noticeable errors in the Logger or Host events. In the back grant, the tracking was triggered and the scan was performed. Consider the possible cause of poor host performance. Please note that there is no reproducibility after the 18th, so please follow up. Internal Comments: [Subject] 01CS: RS: Mr. Hisaki: The operation became slow, such as updating the tracker image on the Bolus Track around 11:30 on the 18th. [Internal Comments] Error code: None [4/8 Elements] Q5 How did the test end ?: None Q6 Presence or absence of incorrect findings (misdiagnosis, mistreatment) due to problems: None Q7 Presence or absence of unintended unnecessary exposure: None Q8 Presence / absence of artifacts: None | 4 |  | Consider the possible cause of poor host performance. Please note that there is no reproducibility after the 18th, so please follow up. Keep monitoring | [12 empty] | 118180196 |  | 1244172 |  |  | Opened - Pending Translation |  | TBD-SW-Slow | [5 empty] | Engineering | [6 empty] | System slow | [9 empty] | System | Console-SW | Performance | Console-SW-Performance | Console-SW-Performance | TBD-SW-Slow | Performance | Console-SW | Console-SW-Performance | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1305 |  |  | Japan | Maki Hospital | 341017 | 4442700% |  | 01DS:KKI: When two orders are raised for the same patient on the MWM, both are selected. | Work Order #: WO-04152337 20-Aug-2021 Yamamoto, Yoshiyuki The system meets the specification for the performed service and is returned to use. Regulatory Questions : N/A Problem Description by Engineer: Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : Work details: When data was imported from the worklist, two items of data were automatically selected and after one of those was selected, if it was double-clicked without moving the cursor, with the first click the data was reselected, and with the second click it was retrieved. This is why two items of data were retrieved. As the data is automatically selected at the point when it appears on the worklist, please select the data that is not necessary and then retrieve the necessary data. Internal Comments: [CX related] Did you agree on how to complete the report with the person in charge of devices? Yes. I completed normal reporting that day. [Work report (paper) submission] Did you submit a work report on paper? <Yes> Q1. How did the examination finish? ・As it stopped before X-ray emission, no extra irradiation Q2. Were there any mistaken findings (misdiagnosis, mistaken treatment) due to the problem? None. Q3. Were there any artifacts? None. Q4. Was there any unintended unnecessary exposure? None. ----------------------------------------------------------------------------- | 4 |  | When data was imported from the worklist, two items of data were automatically selected and after one of those was selected, if it was double-clicked without moving the cursor, with the first click the data was reselected, and with the second click it was retrieved. This is why two items of data were retrieved. Try to reproduce the issue | [12 empty] | 118162946 |  | 1237062 |  |  | Evaluation |  | TBD-WorklistUsage | [5 empty] | Engineering | [6 empty] | Cursor | [9 empty] | System | Console-SW | WorklistUsage | Console-SW-WorklistUsage | Console-SW-WorklistUsage |  |  | Console-SW | Console-SW-WorklistUsage | No | 0 | 0 | 0 | 0 | 0 | No | y | y | y | 0 | 0 | y | y | y | 0 | y | #N/A
1306 |  |  | Japan | Maki Hospital | 341017 | 4442400% |  | AF: An artifact appears on images. | Work Order #: WO-04125507 Ohyama, Tetsuo Regulatory Questions : N/A Problem Description by Engineer: Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer: N/A ----------------------------------------------------------------------------- | 4 |  | Calibration and image data is available in One EMS, to be check with engeineering | [12 empty] | 118138756 |  | 1223001 |  |  | Investigation |  | Image Artifact | [5 empty] | Engineering | [6 empty] | Image Artifact | [9 empty] | System | TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | Image Artifact |  | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | #N/A
1307 |  |  | Peru | BANCO BBVA PERU | 40034 | 4442600% |  | 40034-PROBLEMS TO ACCESS LAN AND REMOTE CONFIGURATION - REINSTALL SOFTWARE | 40034-PROBLEMS TO ACCESS LAN AND REMOTE CONFIGURATION - REINSTALL SOFTWARE | 5 |  | PROBLEMS TO ENTER THE LAN AND REMOTE SETUP - REINSTALL SOFTWARE need check with local service if it is really software issue | [12 empty] | 118160599 |  | 1236139 |  |  | Opened |  | CTUDT00631055 | [5 empty] | Engineering | [6 empty] | LAN Config | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Lan Config | Console-SW-Lan Config | Console-SW-Lan Config | CTUDT00631055 | Lan Config | Console-SW | Console-SW-Lan Config | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | 0 | 0 | y | #N/A
1308 |  |  | Sweden | Falu Lasarett | 40049 | 4442700% |  | software and application hangup, bug rep done 09:00 today. | Work Order #: WO-04152219 BERG NIKLAS Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 3 |  | @Silva, need engineering investigation on the software hangup issue https://teams.microsoft.com/_#/files/%E5%B8%B8%E8%A7%84?threadId=19%3A76cc42a0c8fa4bd7877f7de12141f84c%40thread.tacv2&ctx=channel&context=118162766_software%2520and%2520application%2520hangup%2520bug%2520rep%2520done%25209AM%2520today&rootfolder=%252Fsites%252FBigBoreRTRefreshFOKCoreTeam%252FShared%2520Documents%252FGeneral%252FIngenuity%2520Bugreps%252F118162766_software%2520and%2520application%2520hangup%2520bug%2520rep%2520done%25209AM%2520today | [12 empty] | 118162766 |  | 1237645 |  |  | Opened - Pending Translation |  | CTUDT00632520 | [5 empty] | Engineering | [6 empty] | Software hang | [9 empty] | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1309 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 337000 | 4442300% |  | E-stop is not turning off. | Work Order #: WO-04124886 17-Aug-2021 Can Özgenlik The system meets the specification for the performed service and is returned to use Problem Description by Engineer : estop does not close Resolution: All estop buttons on the device were checked, no problem was found. The front and rear sides of the gantry were examined and all cards and leds were found to be normal. In the errors, it was seen that there was a ct box problem in detail. When the CTbox was opened, a socket was found to be loose. After the socket was inserted, it was seen that the key was turned without any problem and the estop circuit was closed. The test shot was taken with Pahntom and it was delivered in an active way. | 3 |  | In the errors, it was seen that there was a ct box problem in detail. When the CTbox was opened, a socket was found to be loose. After the socket was inserted, it was seen that the key was turned without any problem and the estop circuit was closed. | [12 empty] | 118138256 |  | 1222547 |  |  | Evaluation - Pending Translation | [7 empty] | Others | [6 empty] | E-Stop | [12 empty] | - | - |  |  | 0 | - | No | y | y | y | 0 | 0 | No | y | y | y | y | 0 | n | y | y | y | y | #N/A
1310 |  |  | United States | McLaren Macomb | 40021 | 4442700% |  | Tech support | Work Order #: WO-04155045 Powell, Jamie The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Diagnostic performed by Engineer : Issue with getting EKG wave on the CT.. It works fine on the monitors. Customer has ran ibox diag (which passes), checked ekg monitor and ohmed out cables, reseated usb extenders but still having issues. Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Follow Up Required : Issue with getting EKG wave on the CT.. It works fine on the monitors. Customer has ran ibox diag (which passes), checked ekg monitor and ohmed out cables, reseated usb extenders but still having issues. | 3 |  | Check I-Box firmware & software. Program and update I-Box | [12 empty] | 118165167 |  | 1235433 |  |  | Investigation |  | TBD-NoEnoughInfo | [5 empty] | Engineering | [6 empty] | IBOX | [9 empty] | System | Gantry-HW | TBD-NoEnoughInfo | Gantry-HW-TBD-NoEnoughInfo | Gantry-HW-TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo | Gantry-HW | Gantry-HW-TBD-NoEnoughInfo | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1311 |  |  | United States | Unity Health Systems | 40037 | 4442700% |  | 0699CFXXX gantry cannot comply shut down the system | Problem Description by Engineer : Problem description by engineer : gantry cannot comply shut down the system Work Order #: WO-04154954 19-Aug-2021 Shultz Bob The system does not meet the specification for the performed service and is not returned to use Diagnostic performed by Engineer : Acq system can’t comply Problem Description by Engineer : Problem description by engineer : gantry cannot comply shut down the system When was the first occurrence? Within... : 19-Aug-2021 How often is issue occurring? : Unknown Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: No further information is available as per the customer Patient Impact/Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer Current Software Version: Not provided Follow Up Required : Replace xmitter or receiver ----------------------------------------------------------------------------- Work Order #: WO-04154724 Atz, Michael The system meets the specification for the performed service and is returned to use Resolution : customer rans test scans and the system is now operational Internal Comments : Error # and/or description of error : S_DMC_SHOT_COMM_COLLAPSE NORMAL Explanation no frames were generated by DMS during scan. Watch for more occurrences of this error. It is possible to get infrequent occurrences of this error while the scanner is operating correctly. Repair Action : gantry cannot comply shut down the system to the wall ran a STC Diagnostic performed by Engineer : gantry cannot comply shut down the system to the wall Problem Description by Engineer : Problem description by engineer : gantry cannot comply shut down the system When was the first occurrence? Within... : 19-Aug-2021 How often is issue occurring? : Unknown Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: No further information is available as per the customer Patient Impact/Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer Current Software Version: Not provided Follow Up Required : CCB they would like for a FSE to come onsite ----------------------------------------------------------------------------- | 3 |  | Part/Product - MODULE, ROTOR ELECTRONICS TX Part Number - 453567538334 Part/Product - Module, Optical RX Stator Moog Velocity Part Number - 453567538322 | [12 empty] | 118164944 |  | 1235004 |  |  | Investigation |  | TBD_OBS | [5 empty] | Engineering | [6 empty] | TX & RX | [9 empty] | System | Gantry-HW | Other | Gantry-HW-Other | Gantry-HW-Other | TBD_OBS | TBD-NoEnoughInfo | Gantry-HW | Gantry-HW-Other | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1312 |  |  | United States | Philips Medical Capital | 40055 | 4442500% |  | 0609CDN83 Noise coming from gantry | Work Order #: WO-04139719 20-Aug-2021 Turley Rich The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: Noise coming from gantry User Impact: None Patient Impact: None Current Software Version: Resolution : Checked system found warning label detached and rubbing rotor. Had issues restarting system had to shutdown to wall wait 5 minutes allow cirs and host to come fully up before starting gantry ----------------------------------------------------------------------------- Work Order #: WO-04139671 Boyd Craig The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: Noise coming from gantry User Impact: None Patient Impact: None Current Software Version: Follow Up Required : FSE needed on site to evaluate noise coming from gantry. Diagnostic performed by Engineer : Clicking noise coming from gantry when rotating. Started this morning. FSE, Rich Turley, is aware of issue. ----------------------------------------------------------------------------- | 3 |  | Checked system found warning label detached and rubbing rotor. Had issues restarting system had to shutdown to wall wait 5 minutes allow cirs and host to come fully up before starting gantry | [12 empty] | 118151226 |  | 1226084 |  |  | Evaluation | [7 empty] | Operation | [6 empty] | Gantry Noise | [35 empty] | #N/A | bar: No to NAC
1313 |  |  | United States | Lakeland Regional Medical Center | 300143 | 4442300% |  | Delays in scanner in regards to timing. | Internal Remarks : Interaction Type: New Case Caller: Mark Germany IP & Issue: 55151787. Delays in scanner in regards to timing. Event Type: Incident Service Request: Onsite Did Customer Decline Remote Support? No Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Created case and onsite workorder. Next Step Action: N/A Verbal PO Number: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- |  |  | Per customer issue resolved with hard reboot- cancel call | [12 empty] | 118138318 |  | 1222770 |  |  | Investigation | [20 empty] | Non-Refresh | [29 empty] | #N/A | bar: No to y
1314 |  |  | China | 河南许昌市中心医院(传染病) | 40047 | 4443500% |  | 02N/Ingenuity CT 728326/couch can not move up and down | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Replace Part Problem Reported by Customer : 02N/Ingenuity CT 728326 SAP Transfer : Dispatched for Onsite Support Safety Questions : Case: 0118234247 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Gaoge Last updated by: Wang Gaoge 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 扫描床垂直运动故障。 客户问题详细描述或故障现象： 床不能正常上升 设备本身技术问题： 是 Troubleshooting Action: 关键故障代码或现象： 按生降床按钮床没有反应，水平移动正常 远程诊断与测试结果： invert 故障 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 1，建议客户先重启一下机架尝试 2 更换vertical controller 建议备件12NC及名称： vertical control （459800623432） Test/Inspection: N/A Test Equipment: N/A SAP Transfer : Case Transferred to SAP Next Step Action : PPD 459800623432 *1, 已安排24H发货，FSE黄勇-86-18039573296 Internal Remarks : 29.08.2021 11:58:23 UTC Yong Huang (86038368) RDI Accuracy: Yes $Q7: (Confirm reported problem,the couch cannot up and down ) $Q8: (Performed visual inspection ,found CCC display 02， need to reset the CCC ) $Q9: (Reset the CCC ) $Q10: (Test and inspection data refer to: PA. Device returned to full functionality.) $Q11: Calibrated Tools (N/A); External Remarks : 29.08.2021 11:58:20 UTC SUP JCOUSER (86038368)服务内容：[床不能升降，检查发现床控制板故障，重刷床控制板] 服务结果：[维修后，设备恢复正常工作] 其他重要事项：[无] ----------------------------------------------------------------------------- | 3 |  | 服务内容：[床不能升降，检查发现床控制板故障，重刷床控制板] 服务结果：[维修后，设备恢复正常工作] check with field to get log for analysis | [12 empty] | 118234247 |  | 1264588 |  |  | Opened - Pending Translation |  | NotEngineering | [5 empty] | Engineering | [6 empty] | Vertical movement | [9 empty] | System | Couch-SW | Fwupdate | Couch-SW-Fwupdate | Couch-SW-Fwupdate | NotEngineering |  | Couch-SW | Couch-SW-Fwupdate | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | n | y | y | 0 | y | #N/A
1315 |  |  | China | 山东省泰安市中医二院 | 345037 | 4443500% |  | 02N/Ingenuity Core128/Elite China/scan error | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : scan error Next Step Action : FSE 陈涛 018560023293 call in and disp him SAP Transfer : Dispatched for Onsite Support External Remarks : 28.08.2021 12:20:51 UTC Tao Chen (86038373)故障现象: CT扫描间歇性报错 服务内容: 检查并重启CIRS和CT host计算机 服务结果: 维修后,设备恢复正常. Next Step Action : chen tao Safety Questions : Case: 0118232585 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Qian Wang Last updated by: Qian Wang 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 扫描失败。 客户问题详细描述或故障现象： 机架扫描中断。 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 已和TSE确认。Error data path 远程诊断与测试结果： 重启无效。 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议根据TSE建议维修故障。 建议备件12NC及名称： N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 5 |  | 故障现象: CT扫描间歇性报错 服务内容: 检查并重启CIRS和CT host计算机 服务结果: | [12 empty] | 118232585 |  | 1264013 |  |  | Opened - Pending Translation |  | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Reboot | [9 empty] | System | TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo |  |  |  |  | No | 0 | 0 | 0 | 0 | 0 | No | y | y | y | y | y | n | y | y | 0 | y | #N/A
1316 |  |  | China | 陕西省咸阳中心医院 (咸阳市第二人民医院） | 40050 | 4443500% |  | wechat:02N/Ingenuity CT 728326/Abnormal sound is heard in the rack | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : Abnormal sound is heard in the rack Safety Questions : Case: 0118232297 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Zixuan Last updated by: Wang Zixuan 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 机架旋转有噪声，一个月前，报修同样问题 设备本身技术问题： 是。 Troubleshooting Action: 远程诊断与测试结果： 联系工程师郭宇杰，需要调整皮带 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议工程师到场检查皮带，教客户使用飞利浦微信报修 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 30.08.2021 01:19:09 UTC yifei zhang (86039086)服务原因：机架异响 服务内容：调整张紧轮位置 服务结果：设备工作正常 Internal Remarks : 30.08.2021 01:19:12 UTC yifei zhang (86039086) $Q7: (Confirmed reported problem. & The bearing have hardware problem ) $Q8: (1.Check Belt& We need to remove the tensioning wheel) $Q9: (1. remove the tensioning wheel) . 2. Test scan for CT was OK. ) $Q10: (Test and inspection data refer to PA. & Device returned to full functionality.) $Q11：Calibrated Tools name (N.A.） ----------------------------------------------------------------------------- | 3 |  | $Q7: (Confirmed reported problem. & The bearing have hardware problem ) $Q8: (1.Check Belt& We need to remove the tensioning wheel) $Q9: (1. remove the tensioning wheel) . 2. Test scan for CT was OK. ) | [12 empty] | 118232297 |  | 1263862 |  |  | Opened - Pending Translation | [7 empty] | Engineering | [6 empty] | Belt ajustment | [9 empty] | System | Gantry-HW | Belt | Gantry-HW-Belt | Gantry-HW-Belt | ECR-077012 | Belt | Gantry-HW | Gantry-HW-Belt | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | n | y | y | y | y | #N/A
1317 |  |  | China | Dongguan City Integrated Traditional Chinese and Western medicine Hospital | 345013 | 4443300% |  | WeChat：02N/Ingenuity Core128/Elite China/Failure to connect | T2 Activities:Problem Description: Summary of customer's problems: There is no ECG signal. Detailed description of Customer's problems or malfunctions: There is no ECG signal. It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: N/A Remote troubleshoot and test result: An I-Box malfunction is suspected. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Check the I-Box input power supply and the I-Box board. Replace if necessary. We recommend spare parts 12NC and names: I-Box-2 Base board assy：459800166651 Memo of Customer request: None SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Next Step Action: PPD part: 459800166651 arranged, 24H to Dongguan City Hospital of Integrated Chinese and Western Medicine, Li Zongxian - 86-13751567086; please dispatch, thanks. External Remarks: 27.08.2021 04:46:10 UTC Zongxian Li (86001504) service details: replace ECG board. Service's result: device operating normally Internal Remarks:27.08.2021 04:46:13 UTC Zongxian Li (86001504) RDI Accuracy:Yes $Q7: (Confirmed the reported problem, ibox disconnected) $Q8: (reboot the ibox, it didn't respond to host) $Q9: (Replaced the ibox) $Q10: (Test and inspection data refer to PA. Device returned to full functionality._ $Q11: $Calibrated Tools name (N/A) S/N (N/A) Next Calibration date (N/A | 3 |  | I-Box-2 Base board assy：459800166651 check how FSE diagnost the IBOX issue? Have 5v dc cable? | [12 empty] | 118206661 |  | 1256414 |  |  | Opened |  | TBD-NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Ibox | [9 empty] | System | Gantry-HW | TBD-NoEnoughInfo | Gantry-HW-TBD-NoEnoughInfo | Gantry-HW-TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo | Gantry-HW | Gantry-HW-TBD-NoEnoughInfo | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | n | y | y | 0 | y | #N/A
1318 |  |  | China | 扎兰屯市人民医院 | 40045 | 4443300% |  | 02N/Ingenuity CT 728326/report errors | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: Instructed Customer Problem Reported by Customer : report errors Safety Questions : Case: 0118204929 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wennian Han Last updated by: Wennian Han 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: report errors Troubleshooting Action: report that tube is too hot, room temperature is high, also use it with continuous high-dose scan Repair Action: advise customer to reduce room temperature, optimize scanning method, keep monitoring for a long time, CT is normal Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: NA SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | report that tube is too hot, room temperature is high, also use it with continuous high-dose scan Repair Action: advise customer to reduce room temperature, optimize scanning method, keep monitoring for a long time, CT is normal | [12 empty] | 118204929 |  | 1255447 |  |  | Closed - No Further Investigation |  | Not Engineering | [5 empty] | Engineering | [6 empty] | Room temp | [9 empty] | System | NotEngineering | ShouldnotEngineering | NotEngineering-ShouldnotEngineering | NotEngineering-ShouldnotEngineering | Not Engineering |  | NotEngineering | NotEngineering-ShouldnotEngineering | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | n | y | y | 0 | y | #N/A
1319 |  |  | China | Shaanxi Province Ankang T.C.M. Hospital | 345041 | 4443200% |  | 02N/Ingenuity Core128/Elite China/couch can't move | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity Core128/Elite China SAP Transfer : Dispatched for Onsite Support SAP Transfer : Case Transferred to SAP Next Step Action : PPD 床拐角 CORNER, PEDESTAL COVERS, TOP 453567027541 1 已发货 张翼飞-86-17621702391 24H External Remarks : 30.08.2021 02:08:05 UTC yifei zhang (86039086)服务原因：偶发病床不能移动 服务内容：查看日志，后续观察使用 服务结果：设备工作正常 Internal Remarks : 30.08.2021 02:08:08 UTC yifei zhang (86039086) $Q7: (Confirmed reported problem. & couch prblem ) $Q8: (1.Check log& We need to observation use) $Q9: (1.Check log，observation use) . 2. Test scan for CT was OK. ) $Q10: (Test and inspection data refer to PA. & Device returned to full functionality.) $Q11：Calibrated Tools name (N.A.） Safety Questions : Case: 0118194524 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Dai Suyang Last updated by: Dai Suyang 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： ESTOP断开。 客户问题详细描述或故障现象： 床不能动，水平垂直都不行，重启后已恢复正常 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： PRS远程用户名密码登陆错误，尝试了remoteservice Philips1,philips_service service_only 都无法登陆 远程诊断与测试结果： 设备重启后已恢复正常，但是客户要求工程师到场查看原因。 PRS TR状态： bu通。 Repair Action: 建议现场维修步骤： 按照PRS手册重新配置PRS,重置密码remoteservice Philips1 。查看log信息，排除设备隐患，与客户做好良好沟通。 ----------------------------------------------------------------------------- | 3 |  | ESTOP断开。 客户问题详细描述或故障现象： 床不能动，水平垂直都不行，重启后已恢复正常,后续观察使用 PPD 床拐角 CORNER, PEDESTAL COVERS, TOP 453567027541 1 已发货 张翼飞-86-17621702391 24H check field to get bugreport/log | [12 empty] | 118194524 |  | 1250776 |  |  | Opened - Pending Translation | [7 empty] | Others | [6 empty] | Couch movement | bar: No; No
1320 |  |  | China | Fujian Province Medical University Affiliated No. 1 Hospital | 345012 | 4443100% |  | 02N/Ingenuity Core128/Elite China/add AI | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : /add AI Safety Questions:Case: 0118178045 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Tao Fengming Last updated by: Tao Fengming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action: When requesting service on behalf of the customer, please make sure to provide: Device Number: 87219903 Contact Person: Radiology Department, Dr. Chen Ruiquan 18959128263 Hospital Name: The First Affiliated Hospital of Fujian Medical University Fault Details: Work together to add an AI node. Safety issue: None, there is no safety issue. Li Feng T2 Activities:Problem Description: Summary of customer's problems: Other types of error. Detailed description of Customer's problems or malfunctions: Add an AI node. It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: Add an AI node. Remote troubleshoot and test result: The customer's olevel account can't be accessed, version 4.1.10 PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Reset the olevel_service account and record for later service use. Do the related node configuration according to the manual. SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Internal Remarks : 25.08.2021 06:42:19 UTC Feng Li (86036224) RDI Accuracy: Yes $Q7: (Confirmed reported problem.The Ingenuity Core128/Elite China need update time and add AI node.) $Q8: (On site to update Host & CIRS time. And Lan-config to add AI node.) $Q9: (On site to update Host & CIRS time. And Lan-config to add AI node.) $Q10: (Test and inspection data refer to: PA.+Device returned to full functionality.) $Q11: Calibrated Tools name (FLUKE 287C Multimeter) S/N (16490128) NextCalibration date (10/20/2021) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date: 09/16/2020 ^Current Tube usage: 36536 External Remarks: 25.08.2021 06:42:16 UTC SUP JCOUSER (86036224) Service Details: The on-board CT needed to have the time synchronized and add an AI node.On site, synchronized the time of the host computer and the reconstruction cabinet computer.Added the AI node transmission, testing is normal. service result: device returns to normal working condition after service ----------------------------------------------------------------------------- | 3 |  | The on-board CT needed to have the time synchronized and add an AI node.On site, synchronized the time of the host computer and the reconstruction cabinet computer.Added the AI node transmission, testing is normal. | [12 empty] | 118178045 |  | 1247274 |  |  | Closed - Non-Complaint | [14 empty] | Time sync | bar: No; No
1321 |  |  | China | Jiangxi Province Yugan Chu East Hospital (New Branch) | 337500 | 4443100% |  | Wechat:02N/Ingenuity Core128/Elite China/Couch fault | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : Couch fault Safety Questions : Case: 0118177211 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Bai Jingwen Last updated by: Bai Jingwen 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Dispatched for Onsite Support T2 Activities : Problem Description: 客户问题概述： 床升降噪声 客户问题详细描述或故障现象： 床升降噪声 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： N/A 远程诊断与测试结果： N/A PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 1.确认噪音来源；2.润滑两侧滑轨、检查丝杆 建议备件12NC及名称： N/A 客户需求备注： 无 SAP Transfer : Case Transferred to SAP External Remarks : 24.08.2021 15:52:33 UTC Jiawei Li (86091272)服务内容： 床升降异响，检查床丝杆及两侧滑轨，按照维修手册，完成床两侧滑轨和丝杆润滑。异响消失。 服务结果：维修后，设备恢复正常使用。 Internal Remarks : 24.08.2021 15:52:35 UTC Jiawei Li (86091272) $Q1:(N) $Q2:(N) $Q3:(NA) $Q4:(NA) $Q5:(NA) $Q6:(Big noise during couch vertical move.) "RDI Accuracy:Yes" $Q7:(Confirmed reported problem.Big noise during couch vertical move.) $Q8:(Checked and found couch slide noise.) $Q9:(Lubricated couch slides solved the problem.) $Q10:(Test and inspection data refer to PA. Device returned to full functionality.) $Q11:(Calibrated Tools: No tools used.) ----------------------------------------------------------------------------- | 3 |  | 服务内容： 床升降异响，检查床丝杆及两侧滑轨，按照维修手册，完成床两侧滑轨和丝杆润滑。异响消失。 459801813061 - FRU, LUBRICANT, GADUS S3 V460D 2 is available in SPS, old stock will use up | [12 empty] | 118177211 |  | 1247067 |  |  | Opened |  | E-125059 | [5 empty] | Engineering | [6 empty] | Lubrication | [9 empty] | System | Couch-HW | Lubricant | Couch-HW-Lubricant | Couch-HW-Lubricant | E-125059 | Lubricant | Couch-HW | Couch-HW-Lubricant | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | n | y | y | 0 | y | #N/A
1322 |  |  | Italy | ASST SPEDALI CIVILI | 40026 | 4443800% |  | Gantry non emette raggi. | Internal Remarks : Allego richiesta cliente per completamento pratica ----------------------------------------------------------------------------- Work Order #: WO-04239661 CENTEMERO, LORENZO The system does not meet the specification for the performed service and is not returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Customer reports scanning error. After running stop, the closing and stop mask no longer appears. From the initially error logs of GENERATOR ALE ERROR and after restarting RDTH HC TIMEOUT ERROR with RHOST_STARTUP_UNKNOWN__GENERATOR_ID. Suspected problem with RHOST or system control board, I order both. System stopped Follow Up Required: Customer reports scanning error. After running stop, the closing and stop mask no longer appears. From the initially error logs of GENERATOR ALE ERROR and after restarting RDTH HC TIMEOUT ERROR with RHOST_STARTUP_UNKNOWN__GENERATOR_ID. Suspected problem with RHOST or system control board, I order both. System stopped Diagnostic performed by Engineer: Scan problem confirmed, intervention will follow System cntrl board replaced and successful operation tests performed Product/Part - 80 KW System Control Board Part number - 453567560022 | 3 |  | System cntrl board replaced and successful operation tests performed Product/Part - 80 KW System Control Board Part number - 453567560022 | [12 empty] | 118243669 |  | 1267033 |  |  | Opened - Pending Translation |  | NotEngineering | [5 empty] | Engineering | [6 empty] | SCB | [9 empty] | System | Couch-HW | SCB | Couch-HW-SCB | Couch-HW-SCB | NotEngineering |  | Couch-HW | Couch-HW-SCB | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | n | y | y | 0 | y | #N/A
1323 |  |  | Japan | 社会福祉法人恩賜財団済生会 | 343000 | 4443800% |  | 07CS:TM:use previous surviewの表示がされるようになった>>担当FSEへ報告。どのような操作による挙動か調査を行います | Work Order #: WO-04240829 Miyakawa, Takeyuki The system meets the specification for the performed service and is returned to use Diagnostic performed by Engineer : 担当FSEへ報告。どのような操作による挙動か調査を行います Problem Description by Engineer : use previous surviewの表示がされるようになった>>担当FSEへ報告。どのような操作による挙動か調査を行います Internal Comments : [Subject] 07CS:TM:モリワキ様:use previous surviewの表示がされるようになった>>担当FSEへ報告。どのような操作による挙動か調査を行います [Internal Comments] どのような操作による挙動か調査を行います [4/8 Elements] Q5 その検査はどの様に終了しましたか?:無しQ6 問題により間違った所見（誤診、誤治療）の有無:無し Q7 意図しない不要な曝射の有無:無し Q8 アーチファクトの有無:無し ----------------------------------------------------------------------------- | 5 |  | Keep monitoring | [12 empty] | 118241370 |  | 1267344 |  |  | Opened - Pending Translation | [7 empty] | TBD | [6 empty] | Surview | bar: No; No
1324 |  |  | Japan | Kinan Hospital | 340000 | 4443400% |  | 01DS:RS: He set a high threshold and tried to start a bolus tracking scan manually but it did not work. | Work Order #: WO-04219242 Satake, Rui The system meets the specification for the performed service and is returned to use. Resolution: I did not find any hardware error in the log. It is possible that the application was not working properly. Monitoring the situation. Problem Description by Engineer: Set the threshold value high and tried to start BolusTrack manually, but there was no response. Internal Comments: [Subject] 01CS:RS:Mr./Ms. Kadosaka: Set the threshold value high and tried to start BolusTrack manually, but there was no response. [Internal Comments] Error code: None [4/8 Elements] Q5. How did the examination end? None. Q6. Did any mistaken findings (mistaken diagnosis or treatment) result from the problem? No. Q7. Was there any unintended unnecessary exposure? No Q8 Artifacts? None | 3 |  | Check with field to get bugreport for analysis | [12 empty] | 118226641 |  | 1260704 |  |  | Evaluation |  | Shouldbecustomer | [5 empty] | Engineering | [6 empty] | Bolus tracking | [9 empty] | System | Console-SW | ShouldbeCustomer | Console-SW-ShouldbeCustomer | Console-SW-ShouldbeCustomer | Shouldbecustomer |  | Console-SW | Console-SW-ShouldbeCustomer | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | n | y | y | 0 | y | #N/A
1325 |  |  | Japan | Maki Hospital | 341017 | 4443500% |  | •What were they expecting the device to do? Get patient data with MWM connection | Check CFM detail with Peter | SHU, Tim |  | Kevin share the introduction to Peter/Tim and communicate with submitter and check if have detail step in IFU | [14 empty] | 1260215 |  |  | Investigation | [7 empty] | Application | [6 empty] | Worklist | bar: No; No
1326 |  |  | Japan | Independent administrative institution National Hospital Organization | 341015 | 4443100% |  | 01CS:RS: Cannot transfer raw NMRS data to an external device. >> Restarting the host solved the problem. | cancel CASE due to a mistake in the site. | 4 |  | Mistake | [12 empty] | 118181278 |  | 1248107 |  |  | Closed - Non-Complaint | [14 empty] | NAC | bar: No; No
1327 |  |  | Japan | Tomakomai City Hospital | 341014 | 4443200% |  | 02DS:KK: The error "Patient table cannot comply" appears frequently.>> It is possible to close the error and [Truncated] | Work Order #: WO-04179994 ----------------------------------------------------------------------------- Work Order #: WO-04179865 Kojima, Kenichi The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : The error "Patient table cannot comply" appears frequently.>> It is possible to close the error and continue the examination.>>S_HMC_ABSOLUTE_POSITION_ENCODER_FAILURE Internal Comments: [Subject] 02DS:KK:Mr. Masuda: The error "Patient table cannot comply" appears frequently.>> It is possible to close the error and continue the examination.>>S_HMC_ABSOLUTE_POSITION_ENCODER_FAILURE Shipped part: 453567399251:Assy, Absolute Encoder QTY:1 ZSID --None-- 453567399231:Assy, Position Encoder QTY:1 ZSID --None-- 453567099401:COUPLING/SPIDER QTY:1 ZSID --None-- Dispatched: 1 RC FSE Registration Information: Name of FSE 1 who visited the site: Hodaka Hashimoto FSE 1 visit date and time: 2021/08/2417:00 FSE 1 hours worked: 2 Diagnostic performed by Engineer : "S_HMC_ABSOLUTE_POSITION_ENCODER_FAILURE" had occurred. On-site reposnse by FSE. ----------------------------------------------------------------------------- | 4 |  | An error occurred during horizontal operation of the bed. I lost sight of the location information because the coupling connecting the units viewing the horizontal position information was cracked and slipped. We replaced the coupling and calibrated the location information. After the replacement, we checked the operation and confirmed that there was no problem. We performed a test scan and confirmed that there were no problems. [Parts used] Part number: <453567099401> / Part name: <COUPLING / SPIDER> / Serial number: <ー> / Batch number: <-> / Quantity: <1> | [12 empty] | 118195806 |  | 1246979 |  |  | Evaluation |  | NotEngineering | [5 empty] | Engineering | [6 empty] | Absolute encoder | [9 empty] | System | Couch-HW | Encoder | Couch-HW-Encoder | Couch-HW-Encoder | NotEngineering |  | Couch-HW | Couch-HW-Encoder | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | n | y | y | 0 | y | #N/A
1328 |  |  | Puerto Rico | METRO MAYAGUEZ INC | 40009 | 4443800% |  | When trying to expose, user gets an error “acquisition cannot comply” and then “gantry can't comply” | Work Order #: WO-04243624 Caldelas Aldana, Carlos Manuel The system does not meet the specification for the performed service and is not returned to use Problem Description by Engineer : Problem description by engineer : Aborting aquisitions When was the first occurrence? Within... : 30-Aug-2021 12:00:00 How often is issue occurring? : First occurrence Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: aborting study Patient Impact/Outcome: cant perform estudy User Impact: cant perform aquisition Current Software Version: 4.1.10 Diagnostic performed by Engineer : validated that system is having error of aquisition cant comply. user mentions probably liquid or something went inside rotor. Analisis with remote conection of logs shows errors of refence detector errors. Internal Comments : Error # and/or description of error : S_DMC_REFERENCE_CONV_DATA S_DMC_DMS_VOLTAGE_OUT_OF_SPEC Repair Action : N/A Resolution : FSE Visit Follow Up Required : Error # and/or description of error : S_DMC_REFERENCE_CONV_DATA S_DMC_DMS_VOLTAGE_OUT_OF_SPEC Repair Action : N/A What are the next Troubleshooting steps : clean inside DMS and collimator to discard short circuit. Check if milar or any cover needs to be replaced (new follow up case if needed) Physically look at the RCOM LED’s reading, and/or measuring the PS’s output. Measure the (Vicors) PS output and Check the respective DMS Power Supply cabling. Replace as needed. Run (DMS) VD and check/follow-up accordingly with the results/outcome. If all the above does not address the issue, check / replace a motherboard Replace reference converter if needed replace reference detector if needed Check / Replace the RCOM. Technician Name : Caldelas Aldana, Carlos Manuel ----------------------------------------------------------------------------- Product/Part - REF DETECTOR ASSY Part number - 455012004371 Actual Qty - 1.0000 Product/Part - BR. REFERENCE CONVERTER ASSY ROHS Part number - 459800258961 Actual Qty - 1.0000 Visual inspection revealed liquid contras from the patient felt inside the gantry. As the gantry rotated, the contrast felt inside the gantry components damaging the Reference detector and the Reference detector converter. These parts have to be replaced. | 2 |  | The Reference Detector and the Reference Detector Converter were ordered. The parts were delivered to FedEx. The parts were installed. The CT was tested and left fully operational. | [12 empty] | 118246886 |  | 1268612 |  |  | Opened |  | NotEngineering | [5 empty] | Engineering | [6 empty] | REF DETECTOR | [9 empty] | System | Gantry-HW | Reference-Detector | Gantry-HW-Reference-Detector | Gantry-HW-Reference-Detector | NotEngineering |  | Gantry-HW | Gantry-HW-Reference-Detector | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | n | y | y | 0 | y | #N/A
1329 |  |  | Puerto Rico | DORADO HEALTH INC DBA | 40007 | 4443200% |  | Noise from the couch during vertical movement | Work Order #: WO-04192386 Rios, William Problem Description by Engineer : Problem description by engineer : Noise during vertical movement. Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: No noise during vertical movement of the couch. Patient Impact/Outcome: Noise during movement. User Impact: Noise during movement. Current Software Version: 4.1.7 Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- Work Order #: WO-04191193 Valenzuela Velasquez, Eduardo Ariel The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Problem description by engineer : Noise during vertical movement. Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: No noise during vertical movement of the couch. Patient Impact/Outcome: Noise during movement. User Impact: Noise during movement. Current Software Version: 4.1.7 Follow Up Required : Error # and/or description of error : Noise from the patient support. Repair Action : - Check/lubricate couch vertical movement. Technician Name : Valenzuela Velasquez, Eduardo Ariel Internal Comments : Error # and/or description of error : Noise from the patient support. Repair Action : - Check/lubricate couch vertical movement. Diagnostic performed by Engineer : Customer reports internal noise during vertical movement of the patient support. The case will be sent to planning. ----------------------------------------------------------------------------- | 4 |  | Lubricate all patient support, perform functional test to the system. Unit OK 459801813061 - FRU, LUBRICANT, GADUS S3 V460D 2 is available in SPS, old stock will use up | [12 empty] | 118198669 |  | 1249298 |  |  | Evaluation |  | E-125059 | [5 empty] | Engineering | [6 empty] | Lubrication | [9 empty] | System | Couch-HW | Lubricant | Couch-HW-Lubricant | Couch-HW-Lubricant | E-125059 | Lubricant | Couch-HW | Couch-HW-Lubricant | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | n | y | y | 0 | y | #N/A
1330 |  |  | United States | Lackey Memorial Hospital | 40015 | 4443100% |  | 0602CDN84 the gentry randomly turns off, lowering and raising the table/ stopped scanning during scan | Internal Remarks : Interaction Type: Customer Call Back Caller: Lane Sulton for Michea Mclemore IP & Issue: 85117259 ccb- requesting for service asap, pls call CT SCAN DEPT-5707703056 for ETA Event Type: Incident Service Request: Remote Technical Did Customer Decline Remote Support? No Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: case Next Step Action: N/A Verbal PO Number: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : FADB: (FSE Aware of Double Booking) Title: Internal Remarks Initials: Twila Actions Taken: Called Recommended Service Engineer: Denison Chuck - He accepted WO for today (8/23/21 @ 10am CST). Travel Time 1hr to Site Outcome: Dispatched per FSE's Request IP: 85117259 Next Steps: None Next Step Action : Looks like Estop opened. In the bugrep it didn't show why. The customer is scanning . FSe will monitor for errors. Internal Remarks : FADB Planner Initials: LV Actions Taken: Received call FSE Richard Turley, WO scheduled with FSE for 8/23at 2p CST, customer is aware Outcome: Dispatched Next Steps: N/A Last Updated: 5:58PM EST Internal Remarks : FADB: (FSE Aware of Double Booking) Title: Internal Remarks Initials: Twila Actions Taken: Dispatched per WO FU Note - Recommended Service Engineer: Denison Chuck Follow Up Date/Time Requested: 9/1/2021 2:00 AM CST Outcome: Dispatched per WO FU Note IP: 85117259 Next Steps: None ----------------------------------------------------------------------------- Work Order #: WO-04224112 25-Aug-2021 Denison Chuck Diagnostic performed by Engineer : N/A Problem Description by Engineer : Problem description by engineer : Customer states over the weekend the gantry shut off intermittently. When was the first occurrence? Within... : 22-Aug-2021 How often is issue occurring? : Intermittent Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Gantry turning off. Patient Impact/Outcome: Delay in having Scan User Impact: Delay in Scan Current Software Version: Not provided ----------------------------------------------------------------------------- Work Order #: WO-04197673 25-Aug-2021 Denison Chuck The system meets the specification for the performed service and is returned to use Diagnostic performed by Engineer : Wrong Part > not the correct Vertical Motor. Will ship back. Problem Description by Engineer : Problem description by engineer : Customer states over the weekend the gantry shut off intermittently. When was the first occurrence? Within... : 22-Aug-2021 How often is issue occurring? : Intermittent Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Gantry turning off. Patient Impact/Outcome: Delay in having Scan User Impact: Delay in Scan Current Software Version: Not provided Follow Up Required : Will try to order the correct Vertical motor for the patient support. ----------------------------------------------------------------------------- Work Order #: WO-04216194 26-Aug-2021 Donald Vic The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, no follow-up requested Problem Description by Engineer : Customer Function/Role: Tech Device Use: normal use Expected and Unexpected Results: the table is making a noise when raising the table User Impact: no impact Patient Impact: no impact Current Software Version: 4.1.7 Diagnostic performed by Engineer : The system is working and just making noise. We are sourcing a new motor for the vertical drive. ----------------------------------------------------------------------------- Work Order #: WO-04177875 25-Aug-2021 Turley Rich The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : assisted with diagnostics ----------------------------------------------------------------------------- Work Order #: WO-04177965 24-Aug-2021 Denison Chuck The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Problem description by engineer : Customer states over the weekend the gantry shut off intermittently. When was the first occurrence? Within... : 22-Aug-2021 How often is issue occurring? : Intermittent Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Gantry turning off. Patient Impact/Outcome: Delay in having Scan User Impact: Delay in Scan Current Software Version: Not provided Diagnostic performed by Engineer : Waited to get the C.T. scanner. Could not get work done to the Patient support because of covid patient. Follow Up Required : Need to change out the Patient support vertical motor assembly. ----------------------------------------------------------------------------- Work Order #: WO-04174624 23-Aug-2021 Denison Chuck The system meets the specification for the performed service and is returned to use Diagnostic performed by Engineer : Worked with Ronnie Dober NSS On logs. Gantry Error. We could not come up on what caused the error. Also worked on the Patient couch support. Needs motor replaced. Problem Description by Engineer : Problem description by engineer : Customer states over the weekend the gantry shut off intermittently. When was the first occurrence? Within... : 22-Aug-2021 How often is issue occurring? : Intermittent Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Gantry turning off. Patient Impact/Outcome: Delay in having Scan User Impact: Delay in Scan Current Software Version: Not provided Follow Up Required : Need to come back to site to replace the Vertical Motor and worm drive. ----------------------------------------------------------------------------- Work Order #: WO-04174128 Miller, John M The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Problem description by engineer : Customer states over the weekend the gantry shut off intermittently. When was the first occurrence? Within... : 22-Aug-2021 How often is issue occurring? : Intermittent Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Gantry turning off. Patient Impact/Outcome: Delay in having Scan User Impact: Delay in Scan Current Software Version: Not provided Follow Up Required : Error # and/or description of error : TmStamp 106389 ERR_TYPE NORMAL ErrNum S_MASTER_SEQUENCER_SHOT_OUT_OF_SEQUENCE_OR_ERR vxwErrNo 0 File S:\sy_bld_main_Bld_sn/stargate2/gantry/source/ghost/Sequencer/src/ShotTrans.cpp Line 774 Param1 128 Param2 0 Info EalTime 8During Shot state Param1 has a response bit map: Bit 0 (1) 0x01 XOFF_GUARD - if 1 the RHOST did not respond on a request from GHOST to Switch off X-ray. Bit 1 (2) 0x02 GRANT_GUARD - if 1 the RHOST did not respond on a request from GHOST to go to Grant State Bit 2 (4) 0x04 ARMED_GUARD - if 1 the RHOST did not respond on a request from GHOST to go to Armed State. Bit 3 (8) 0x08 FILAMENT_GUARD - if 1 the RHOST did not respond on a request from GHOST to preheat the filament or Filament timer was set twice. Bit 4 (16) 0x10 EXECUTING_GUARD. Bit 5 (32) 0x20 PRE_SHOT_DELAY_GUARD - If 1 Could mean that the RHOST did not responded for GEN_GRANT or GEN ARM. Bit 6 (64) 0x40 XON_GUARD - If 1 RHOST did not detect Xray ON Bit 7(128) 0x80 XARM_GUARD - If 1 the RHOST ARMED state timed out. Bit 8 (256) 0x100 DC_TAKESHOT_GUARD - If 1 the DMC did not start the shot. Bit 9 (512) 0x200 COUCH_GO_HOR_MOVE_GUARD - If 1 the couch did not move (correctly) during surview or Helix scan. 8/21/2021 11:27:55 PM TmStamp 204122 ERR_TYPE FATAL ErrNum S_MPL_VX_SEND_ERROR vxwErrNo 3997697 File S:\sy_bld_main_Bld_sn/stargate2/gantry/source/gaf/mpl/src/MPL.cpp Line 598 Param1 -1 Param2 429062 Info 0 EalTime 8/23/2021 2:38:40 AM ERR_TYPE NORMAL ErrNum S_MP_MPSHELL_SERIAL_LIFE_SYNC_ERROR vxwErrNo 3997697 File S:\sy_bld_main_Bld_sn/stargate2/gantry/source/mp/MPShell/src/MPShell.cpp Line 336 Param1 36700168 Param2 3997697 Info MP Latest Fatal Error Data. EalTime 8/23/2021 2:38:41 AM Repair Action : Life Ping between GHOST and GMP lost for some time (>2 seconds). Check laser cabling to center laser. Can be shorting out causing the Ghost to restart. Customer stated the couch was popping at the same time. Not sure they are related. Suggest checking for debris in the horizontal movement as well as covers binding. If the couch is hard to move at various points in horizontal travel, very possible the Kerk Screw and bearings are going bad and need to be replaced. Disconnect the linkage between the Kerk Screw drive and the table assembly to halve the problem. If it is the lead screw it could be the Kerk nut, the brake/clutch dragging/ or the thrust bearings needing replacement. If the table trolley is the issue it could be the linear bearings needing lubrication or replacement. Also check the Horizontal encoders for proper operation, tight mechanical connections and calibrations. Check the motor coupling between the encoders. Lubrication of Horizontal drive per PM manual. Check 5v switching power for at least 4.85v measured on the Couch Control Board. The power supply is adjustable. Refer to the Couch troubleshooting manual for further assistance. Technician Name : Miller, John M Internal Comments : Error # and/or description of error : TmStamp 106389 ERR_TYPE NORMAL ErrNum S_MASTER_SEQUENCER_SHOT_OUT_OF_SEQUENCE_OR_ERR vxwErrNo 0 File S:\sy_bld_main_Bld_sn/stargate2/gantry/source/ghost/Sequencer/src/ShotTrans.cpp Line 774 Param1 128 Param2 0 Info EalTime 8During Shot state Param1 has a response bit map: Bit 0 (1) 0x01 XOFF_GUARD - if 1 the RHOST did not respond on a request from GHOST to Switch off X-ray. Bit 1 (2) 0x02 GRANT_GUARD - if 1 the RHOST did not respond on a request from GHOST to go to Grant State Bit 2 (4) 0x04 ARMED_GUARD - if 1 the RHOST did not respond on a request from GHOST to go to Armed State. Bit 3 (8) 0x08 FILAMENT_GUARD - if 1 the RHOST did not respond on a request from GHOST to preheat the filament or Filament timer was set twice. Bit 4 (16) 0x10 EXECUTING_GUARD. Bit 5 (32) 0x20 PRE_SHOT_DELAY_GUARD - If 1 Could mean that the RHOST did not responded for GEN_GRANT or GEN ARM. Bit 6 (64) 0x40 XON_GUARD - If 1 RHOST did not detect Xray ON Bit 7(128) 0x80 XARM_GUARD - If 1 the RHOST ARMED state timed out. Bit 8 (256) 0x100 DC_TAKESHOT_GUARD - If 1 the DMC did not start the shot. Bit 9 (512) 0x200 COUCH_GO_HOR_MOVE_GUARD - If 1 the couch did not move (correctly) during surview or Helix scan. 8/21/2021 11:27:55 PM TmStamp 204122 ERR_TYPE FATAL ErrNum S_MPL_VX_SEND_ERROR vxwErrNo 3997697 File S:\sy_bld_main_Bld_sn/stargate2/gantry/source/gaf/mpl/src/MPL.cpp Line 598 Param1 -1 Param2 429062 Info 0 EalTime 8/23/2021 2:38:40 AM ERR_TYPE NORMAL ErrNum S_MP_MPSHELL_SERIAL_LIFE_SYNC_ERROR vxwErrNo 3997697 File S:\sy_bld_main_Bld_sn/stargate2/gantry/source/mp/MPShell/src/MPShell.cpp Line 336 Param1 36700168 Param2 3997697 Info MP Latest Fatal Error Data. EalTime 8/23/2021 2:38:41 AM Repair Action : Life Ping between GHOST and GMP lost for some time (>2 seconds). Check laser cabling to center laser. Can be shorting out causing the Ghost to restart. Customer stated the couch was popping at the same time. Not sure they are related. Suggest checking for debris in the horizontal movement as well as covers binding. If the couch is hard to move at various points in horizontal travel, very possible the Kerk Screw and bearings are going bad and need to be replaced. Disconnect the linkage between the Kerk Screw drive and the table assembly to halve the problem. If it is the lead screw it could be the Kerk nut, the brake/clutch dragging/ or the thrust bearings needing replacement. If the table trolley is the issue it could be the linear bearings needing lubrication or replacement. Also check the Horizontal encoders for proper operation, tight mechanical connections and calibrations. Check the motor coupling between the encoders. Lubrication of Horizontal drive per PM manual. Check 5v switching power for at least 4.85v measured on the Couch Control Board. The power supply is adjustable. Refer to the Couch troubleshooting manual for further assistance. Diagnostic performed by Engineer : Collect information from customer and view log files. ----------------------------------------------------------------------------- Product/Part - FRU,ASSY, VERTICLE MOTOR/BRAKE 20:1 Part number - 459801403271 Actual Qty - 1.0000 Parts Consumed Description: 1.00000,459801402611,BALL SCREW ASSEMBLY SERVICE KIT | 3 |  | Vertical drive check with field to get bugreport | [12 empty] | 118182951 |  | 1244897 |  |  | Investigation |  | NotEngineering | [5 empty] | Engineering | [6 empty] | Vertical movement | [9 empty] | System | Couch-HW | VerticalDriver | Couch-HW-VerticalDriver | Couch-HW-VerticalDriver | NotEngineering |  | Couch-HW | Couch-HW-VerticalDriver | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | n | y | y | 0 | y | #N/A
1331 |  |  | United States | AdventHealth Zephyrhills | 300106 | 4443300% |  | Applaction issue with Brain Perfusion program not pulling up some of windows | Internal Remarks : Interaction Type: New Case Caller: James Golden IP & Issue: 54648842 Applaction issue with Brain Perfusion program not pulling up some of windows Event Type: Incident Service Request: Remote Clinical Did Customer Decline Remote Support? No Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Created Clinical W/O routed to Clinical Next Step Action: N/A Verbal PO Number: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- Work Order #: WO-04203601 26-Aug-2021 Garcia Marrero, Jorge The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, no follow-up requested Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: Applaction issue with Brain Perfusion program not pulling up some of windows User Impact: None Patient Impact: None Current Software Version: Diagnostic performed by Engineer : Collecting data from Brain Perfusion study to create a Customer Feedback case about Brain Perfusion application that fail to create a report correctly. Follow Up Required : Follow up to monitor performance of application ----------------------------------------------------------------------------- Work Order #: WO-04203560 Boyd Craig The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: Applaction issue with Brain Perfusion program not pulling up some of windows User Impact: None Patient Impact: None Current Software Version: Diagnostic performed by Engineer : Applications spoke with customer regarding issue and covered all bases related to it. PRS accessed. ULV: no normal errors seen. Follow Up Required : FSE needed on site to evaluate issue. Recommendations: Reload software. ----------------------------------------------------------------------------- Work Order #: WO-04203451 Mujica, David The system meets the specification for the performed service and is returned to use Problem description by Clinical Support Specialist : Customer Function/Role: Technologist Device Use: In use Expected and Unexpected Results: Color Maps Issue for BP User Impact: none Patient Impact: none Current Software Version: Diagnostic performed by Engineer : Russ indicated that the color maps is intermittently failing when post processing a CT Jog Brain Perfusion study. No errors were given on the scan portion, no errors given when loading the scan into the BP post processing software. The CT system's local is optimal at 71% availability and daily shutdown's with weekly complete power cycles are performed per Russ. RSE group involved, bug report captured. Case transferred to Tech for dispatch for remedy onsite. No further clinical support needed. No rescan, no reinjection per Russ. ----------------------------------------------------------------------------- | 5 |  | Collecting data from Brain Perfusion study to create a Customer Feedback case about Brain Perfusion application that fail to create a report correctly. TW PR ID 1254071 collect bugreport and detail information for engineering analysis | [12 empty] | 118222828 |  | 1254071 |  |  | Evaluation |  | CTUDT00633806 | [5 empty] | Engineering | [6 empty] | Brain Perfusion | [5 empty] | Refresh |  |  |  | System | Console-SW | Perfusion | Console-SW-Perfusion | Console-SW-Perfusion | CTUDT00633806 |  | Console-SW | Console-SW-Perfusion | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | n | y | y | 0 | y | #N/A
1332 |  |  | United States | AdventHealth Zephyrhills | 300099 | 4443100% |  | GANTRY WAS TILTED AND E STOP OPENED - NOW UNABLE TO CLOSE E STOP | Internal Remarks : Interaction Type: Customer Call Back Caller: vinny IP & Issue: calling for FSE Dispatch tonight 8/23 Event Type: Incident Service Request: Onsite Did Customer Decline Remote Support? Yes Preferred Service Time: today 8/23 Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: created remote wo Next Step Action: N/A Verbal PO Number: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- Work Order #: WO-04197124 25-Aug-2021 Stillwell Philip The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: Technician Device Use: Clinical Expected and Unexpected Results: None User Impact: None Patient Impact: None Current Software Version: Resolution : Did not realize Jorge Garcia made it to site and got scanner working. Ran T/C, axial, & helical. ----------------------------------------------------------------------------- Work Order #: WO-04177610 24-Aug-2021 Garcia Marrero, Jorge The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: Diagnostics Expected and Unexpected Results: Tilted gantry to clean bottom and E Stop opened. Now unable to close E Stop even after reboot from wall. User Impact: No Patient Impact: No Current Software Version:4.0 Diagnostic performed by Engineer : Tilt limit switch in the right leg of the gantry was activated. Gantry was move out of the tilt limit switch manualy. System is fully operational. ----------------------------------------------------------------------------- Work Order #: WO-04177552 Dangos John The system does not meet the specification for the performed service and is not returned to use Diagnostic performed by Engineer : Tilted gantry to clean bottom and E Stop opened. Now unable to close E Stop even after reboot from wall. Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: Diagnostics Expected and Unexpected Results: Tilted gantry to clean bottom and E Stop opened. Now unable to close E Stop even after reboot from wall. User Impact: No Patient Impact: No Current Software Version:4.0 Follow Up Required : Left/Right Tilt Safety switch is activated and now unable to close E Stop. Release safety switch in service mode first. ----------------------------------------------------------------------------- | 2 |  | Tilt limit switch in the right leg of the gantry was activated. Gantry was move out of the tilt limit switch manualy. System is fully operational. | [12 empty] | 118193757 |  | 1245727 |  |  | Investigation | [14 empty] | Tilt limit switch | [5 empty] | Non-Refresh | bar: No; No
1333 |  |  | United States | Lakeland Regional Medical Center | 300143 | 4443100% |  | The gantry is clicking when rotating | Integration : Confirmation of order 54061215 saved Integration : Confirmation of order 54061215 saved ----------------------------------------------------------------------------- Work Order #: WO-04175076 26-Aug-2021 Garcia Marrero, Jorge The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Problem description by engineer : The gantry is clicking when rotating When was the first occurrence? Within... : 23-Aug-2021 How often is issue occurring? : Continuously Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: No further information is available as per the customer Patient Impact/Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer Current Software Version: 4.1.7.25030 Diagnostic performed by Engineer : FOV mylar was creating tic noise inside the gantry. The FOV mylar was adjusted. System is fully operational. ----------------------------------------------------------------------------- Work Order #: WO-04175059 Atz, Michael The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, no follow-up requested Problem Description by Engineer : Problem description by engineer : The gantry is clicking when rotating When was the first occurrence? Within... : 23-Aug-2021 How often is issue occurring? : Continuously Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: No further information is available as per the customer Patient Impact/Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer Current Software Version: 4.1.7.25030 Diagnostic performed by Engineer : The gantry is clicking when rotating Internal Comments : Error # and/or description of error : none Repair Action : The gantry is clicking when rotating Follow Up Required : Error # and/or description of error : none Repair Action : The gantry is clicking when rotating Technician Name : Atz, Michael ----------------------------------------------------------------------------- | 3 |  | FOV mylar was creating tic noise inside the gantry. The FOV mylar was adjusted. System is fully operational. | [12 empty] | 118192581 |  | 1245021 |  |  | Evaluation | [14 empty] | Mylar | [5 empty] | Non-Refresh | bar: No; No
1334 |  |  | United States | Select Specialty Hospitals | 300037 | 4446800% |  | warning message about cracked compensator | Internal Remarks : Title: Internal Remarks Initials: Twila Actions Taken: I called and left FSE a VM inquiring if he's able to accept WO Outcome: N/A Next Steps: 9/29/21 @ 11:55am EST. Please call FSE/Beard Michael back in 20 min (12:15pm). Internal Remarks : FADB: (FSE Aware of Double Booking) Title: Internal Remarks Initials: Twila Actions Taken: Per message sent to me in Teams from FSE/Beard, Michael - Dispatch for 10/7/21. I Dispatched for 9am EST. Outcome: Dispatched WO per FSE's Request IP: 53340864 Next Steps: None Internal Remarks : Phantom cals stopped with error. Had FSe reboot everything and it is working. Integration : Confirmation of order 54230698 saved ----------------------------------------------------------------------------- Work Order #: WO-04446472 05-Oct-2021 Ostapchenko Patrick The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Problem description by engineer : warning message about cracked compensator during air calibration Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: warning message about cracked compensator during air calibration Patient Impact/Outcome: none User Impact: none Current Software Version: 4.1.7.25037 Diagnostic performed by Engineer : moved adm mod 22 to slot 26, and 41 to 26. Perform cals, clean cone ring and test ok. Tools & Equipment : MULTIMETER,2022-03-31 00:00:00 ----------------------------------------------------------------------------- Work Order #: WO-04445475 Sheffield, Scott The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Follow Up Required : Error # and/or description of error : cracked compensator. warning Repair Action : There are 2 previous cases for this problem 0117903846 7/16/2021 0117779333 6/24/2021 Recommendation: The Auto bad det Signal test plot shows that detetor module 22 is low on the plot compaired to modules 21 and 23. I will attach the plots to the case. Use the compensator analizer tool and check compensator for crack when onsite. Swap detector module 26 with detector module 22 and do necessary calibrations. Refer to calibration matrix. Consult with NSS Sam Pierce for more information on repair. I was consulting with NSS Sam Pierce for the recommendations. Technician Name : Sheffield, Scott Internal Comments : Error # and/or description of error : cracked compensator. warning Repair Action : There are 2 previous cases for this problem 0117903846 7/16/2021 0117779333 6/24/2021 Recommendation: The Auto bad det Signal test plot shows that detetor module 22 is low on the plot compaired to modules 21 and 23. I will attach the plots to the case. Swap detector module 26 with detector module 22 and do necessary calibrations. Refer to calibration matrix. Consult with NSS Sam Pierce for more information on repair. I was consulting with NSS Sam Pierce for the recommendations. Diagnostic performed by Engineer : Brian said that during air cal he got message about cracked compensator. He is not seeing any artifacts on phantom images. They have not had any patients to scan, he is scanning the phantom with Clinical Applications support that is onsite. System is working at this time. I downloaded logs and plots and reviewed them. Problem Description by Engineer : Problem description by engineer : warning message about cracked compensator during air calibration Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: warning message about cracked compensator during air calibration Patient Impact/Outcome: none User Impact: none Current Software Version: 4.1.7.25037 ----------------------------------------------------------------------------- |  |  | may related to Beam and DMS check | [12 empty] | 118453627 |  | 1368962 |  |  | Evaluation |  | ECR-076862 | [5 empty] | Engineering | [6 empty] | Compensator | [5 empty] | Refresh |  |  | Upgrade | System | DMS-SW | Beam-DMSCheck | DMS-SW-Beam-DMSCheck | DMS-SW-Beam-DMSCheck | ECR-076862 |  | DMS-SW | DMS-SW-Beam-DMSCheck | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | #N/A
1335 |  |  | United States | Christ Hospital | 300205 | 4446000% |  | 0798CDY72 Calcium Score Case - Intermittently Cannot Initialize | Work Order #: WO-04388474 22-Sep-2021 Hanlin Christopher The system meets the specification for the performed service and is returned to use Problem Description by Engineer : DESCRIPTION OF PROBLEM (EXPECTED/UNEXPECTED RESULTS) Calcium Score Case - Intermittently Cannot Initialize CALLER FUNCTION/ROLE Biomed DEVICE INTENDED USE Diagnostic WHEN WAS THE FIRST OCCURRENCE? WITHIN . . . <1 Week HOW OFTEN IS ISSUE OCCURRING? Intermittent WHERE IN THE END-USER WORKFLOW IS THE ISSUE OCCURRING? Post acquisition/scan WHAT RECENT EVENTS HAVE OCCURRED? During normal use NEGATIVE OPERATOR IMPACT Delay in scanning NEGATIVE PATIENT IMPACT/PROCEDURE OUTCOME Delay in being scanned HAS CUSTOMER CREATED A SPECIFIC SYSTEM EVENT REPORT OF ISSUE? Yes DEVICE SOFTWARE VERSION 4.1.10 Resolution : Examined error logger and parsed BUGREP. Issue caused by intermittent horizontal motor relay errors. Resubmitting plan works to work around issue. Issue traced to K103 Horizontal Motor Relay. ----------------------------------------------------------------------------- Work Order #: WO-04388306 Weaver David The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : DESCRIPTION OF PROBLEM (EXPECTED/UNEXPECTED RESULTS) Calcium Score Case - Intermittently Cannot Initialize CALLER FUNCTION/ROLE Biomed DEVICE INTENDED USE Diagnostic WHEN WAS THE FIRST OCCURRENCE? WITHIN . . . <1 Week HOW OFTEN IS ISSUE OCCURRING? Intermittent WHERE IN THE END-USER WORKFLOW IS THE ISSUE OCCURRING? Post acquisition/scan WHAT RECENT EVENTS HAVE OCCURRED? During normal use NEGATIVE OPERATOR IMPACT Delay in scanning NEGATIVE PATIENT IMPACT/PROCEDURE OUTCOME Delay in being scanned HAS CUSTOMER CREATED A SPECIFIC SYSTEM EVENT REPORT OF ISSUE? Yes DEVICE SOFTWARE VERSION 4.1.10 Diagnostic performed by Engineer : Calcium Score Case - Intermittently Cannot Initialize Follow Up Required : Calcium Score Case - Intermittently Cannot Initialize Cannot access system via PRS, must page out to FSE. Encouraged FSE to reach out to NSS for advice. Possible CT Host software reload. ----------------------------------------------------------------------------- |  |  | Resolution : Examined error logger and parsed BUGREP. Issue caused by intermittent horizontal motor relay errors. Resubmitting plan works to work around issue. Issue traced to K103 Horizontal Motor Relay. | [12 empty] | 118366055 |  | 1340567 |  |  | Investigation |  | NotEngineering-PartReplacement | [5 empty] | Engineering | [6 empty] | Hori Motor Relay | [5 empty] | Refresh |  |  | Upgrade | System | Couch-HW | Relay | Couch-HW-Relay | Couch-HW-Relay | NotEngineering-PartReplacement | Relay | Couch-HW | Couch-HW-Relay | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1336 |  |  | United States | Christ Hospital | 300205 | 4445900% |  | 0699CFXXX Calcium Score Case - No Initialize [ PRS Fail ] | Integration : Confirmation of order 54186844 saved ----------------------------------------------------------------------------- Work Order #: WO-04381465 Carpenter, Kevin The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: Untrained biomed Cory Device Use: Diag Expected and Unexpected Results: Gantry would not Init for CS exam, until power cycled User Impact: None Patient Impact: None Current Software Version: IDK Diagnostic performed by Engineer : PRS/ FTP - problem - unable to get BugRep All ports scan green / open, and the FTP client does launch but the CT side of the FTP dialog never populates. Called biomed back ~ offered to dispatch, but he declined. Resolution : Untrained biomed declined to have FSE onsite to review bugrep as CT is now working. Also there are other biomed who are trained who could likely review the bugrep. ----------------------------------------------------------------------------- |  |  | Resolution : Examined error logger and parsed BUGREP. Issue caused by intermittent horizontal motor relay errors. Resubmitting plan works to work around issue. Issue traced to K103 Horizontal Motor Relay. | [12 empty] | 118360983 |  | 1337625 |  |  | Closed - No Further Investigation |  | NotEngineering-PartReplacement | [5 empty] | Engineering | [6 empty] | Hori Motor Relay | [5 empty] | Refresh |  |  | Upgrade | System | Couch-HW | Relay | Couch-HW-Relay | Couch-HW-Relay | NotEngineering-PartReplacement | Relay | Couch-HW | Couch-HW-Relay | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1337 |  |  | United States | AdventHealth Zephyrhills | 300106 | 4444800% |  | Unable to reset password/ logged in correctly | Integration : Confirmation of order 54139504 saved Internal Remarks : Interaction Type: New Case Caller: Jessie Ros IP & Issue: 54648842 Unable to reset password/ log in correctly Event Type: Incident Service Request: Remote Clinical Did Customer Decline Remote Support? No Preferred Service Time: Regular Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: CASE Next Step Action: N/A Verbal PO Number: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- Work Order #: WO-04311218 Curry, Dyan The system meets the specification for the performed service and is returned to use Problem description by Clinical Support Specialist : Customer Function/ Role: technologist Device Use: in use name Expected and Unexpected Results: locked out of scanner User Impact: none Patient Impact: none Current Software version: 4.1. Diagnostic performed by Engineer : advised customer to wait 15 mins . Scanner will unlock and then enter correct password. Also advised to use Emergency Customer Function/ Role: technologist Device Use: in use name Expected and Unexpected Results: locked out of scanner User Impact: none Patient Impact: none Current Software version: 4.1. password to scan if needed until it unlocks ----------------------------------------------------------------------------- |  |  | Password | [12 empty] | 118304183 |  | 1302199 |  |  | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Password | [5 empty] | Refresh |  |  | Upgrade | bar: No; No
1338 |  |  | United States | AdventHealth Zephyrhills | 300099 | 4444200% |  | 0711CDY73 Stuck in init. | Integration : Confirmation of order 54117038 saved ----------------------------------------------------------------------------- Work Order #: WO-04276721 03-Sep-2021 Sher Mark The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Problem description by engineer : Customer states the system was stuck in init state and wouldn't scan. When was the first occurrence? Within... : 03-Sep-2021 How often is issue occurring? : First occurrence Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Stuck in init Patient Impact/Outcome: Delay in having Scan User Impact: Delay in Scan Current Software Version: 4.1.7.25030 Resolution : moved dvd ram option from lanconfig ----------------------------------------------------------------------------- Work Order #: WO-04276521 Miller, John M The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Problem description by engineer : Customer states the system was stuck in init state and wouldn't scan. When was the first occurrence? Within... : 03-Sep-2021 How often is issue occurring? : First occurrence Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Stuck in init Patient Impact/Outcome: Delay in having Scan User Impact: Delay in Scan Current Software Version: 4.1.7.25030 Diagnostic performed by Engineer : Collect information from customer and view log files. Internal Comments : Error # and/or description of error : Type : Critical Date : 9/2/2021 Time : 10:26:53 PM Event : 41 Source : Microsoft-Windows-Kernel-Power Category : (63) User : \SYSTEM Computer : 128-CT-ED Description: The description for Event ID ( 41 ) in Source ( Microsoft-Windows-Kernel-Power ) could not be found. Either the component that raises this event is not installed on the computer or the installation is corrupted.You can install or repair the component or try to change Description Server. The following information was included with the event (insertion strings): 0 00000000 00000000 00000000 00000000 0 132751131434068928 0 0 0 0 0 No Gantry errors Repair Action : Suggest running the Dell or HP Diagnostic (Dell press F12 during bootup, HP press Esc then F2) to determine the health of the PC. Also check the health of the local drive. Also suggest re-seating memory modules. Suggest replacing hard drive and reload software on new drive. 459801652541 - HP TurboDrive 512GB TLC Z4 SSD w/adapter 1- Long Time to initialize: Is this problem fixed? How? If not, have you checked LAN config and removed DVDRAM from Local devices? Follow Up Required : Error # and/or description of error : Type : Critical Date : 9/2/2021 Time : 10:26:53 PM Event : 41 Source : Microsoft-Windows-Kernel-Power Category : (63) User : \SYSTEM Computer : 128-CT-ED Description: The description for Event ID ( 41 ) in Source ( Microsoft-Windows-Kernel-Power ) could not be found. Either the component that raises this event is not installed on the computer or the installation is corrupted.You can install or repair the component or try to change Description Server. The following information was included with the event (insertion strings): 0 00000000 00000000 00000000 00000000 0 132751131434068928 0 0 0 0 0 No Gantry errors Repair Action : Suggest running the Dell or HP Diagnostic (Dell press F12 during bootup, HP press Esc then F2) to determine the health of the PC. Also check the health of the local drive. Also suggest re-seating memory modules. Suggest replacing hard drive and reload software on new drive. 459801652541 - HP TurboDrive 512GB TLC Z4 SSD w/adapter 1- Long Time to initialize: Is this problem fixed? How? If not, have you checked LAN config and removed DVDRAM from Local devices? Technician Name : Miller, John M ----------------------------------------------------------------------------- |  |  | Long Time to initialize | [12 empty] | 118274392 |  | 1286066 |  |  | Closed - No Further Investigation |  | CTUDT00633734 | [5 empty] | Engineering | [6 empty] | Long Time to initialize | [5 empty] | Refresh |  |  | Upgrade | System | Console-SW | DVDRAM | Console-SW-DVDRAM | Console-SW-DVDRAM | CTUDT00633734 |  | Console-SW | Console-SW-DVDRAM | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | #N/A
1339 |  |  | China | Jiangsu Province Yangzhou City No. 1 People's Hospital, West District Hospital | 345022 | 4446900% |  | 02N/Ingenuity Core128/Elite China/key fault | T2 Activities:Problem Description: Summary of customer's problems: Control panel button is not working. Detailed description of Customer's problems or malfunctions: CTBOX buttons are not responsive. It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: NA Remote troubleshoot and test result: The customer said Down and Exit bed buttons on ctbox are not working properly. Down and Exit bed buttons on Gantry panel are normal PRS TR status: it's not through, ask engineer for maintenance on site. Repair Action: Suggest onsite service procedures: Check the CT BOX. We recommend spare parts 12NC and names: " 459801705022 CT BOX ASSEMBLY_BEST" Internal Remarks : 02.10.2021 04:53:01 UTC (86036024) "RDI Accuracy:NA" $Q7: (Confirmed reported problem: CTbox can't move couh down and back.) $Q8: (No error report. The down key and back key of ctbox are defective. Need to be changed.) $Q9: (Changed CTbox assembly_best and 12NC is 459801705022.) | 3 |  | Changed CTbox assembly_best and 12NC is 459801705022 | [12 empty] | 118462384 |  | 1375863 |  |  | Opened |  | CTUDT00632185 | [5 empty] | Engineering | [6 empty] | CT Box | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | CT Box | Gantry-HW-CT Box | Gantry-HW-CT Box | CTUDT00632185 | CT Box | Gantry-HW | Gantry-HW-CT Box | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1340 |  |  | China | Xifeng TCM Hospital | 345014 | 4446900% |  | CT box of the CT5000 is of poor quality. | CT box of the CT5000 is of poor quality. 12NC:459801705022 A month ago,the down button of the CT box is out of order, so FSE replaced the CT box . But now,the old problem popped up again. At the moment ,we had to replace it again. So the customer complained about the poor quality of the product. The first broken CT box SN :107 The second broken CT box SN:313 |  |  | CT box of the CT5000 is of poor quality | [14 empty] | 1373289 |  |  | Investigation |  | CTUDT00632185 | [5 empty] | Engineering | [6 empty] | CT Box | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | CT Box | Gantry-HW-CT Box | Gantry-HW-CT Box | CTUDT00632185 | CT Box | Gantry-HW | Gantry-HW-CT Box | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1341 |  |  | China | Sichuan Province Yibin City No.2 T.C.M. Hospital | 345035 | 4446800% |  | 02N/Ingenuity Core128/Elite China/ | T2 Activities : Problem Description: Scan interrupted by ARC Troubleshooting Action: Check logger APM ARC 5:0. No more repeat, do filament calibration and take monitoring. Repair Action: No more repeat, do filament calibration and take monitoring. | 3 |  | Perform filament calibration to resolve tube arc issue and keep monitoring. | [12 empty] | 118451184 |  | 1371429 |  |  | Evaluation |  | CTUDT00630085 | [5 empty] | Engineering | [6 empty] | Tube Arc | [5 empty] | Refresh |  |  | System issue | System | IC-HW | Tub Arc | IC-HW-Tub Arc | IC-HW-Tub Arc | CTUDT00630085 | Tub Arc | IC-HW | IC-HW-Tub Arc | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1342 |  |  | China | Laiyang City People's Hospital | 345005 | 4446800% |  | 05L/Ingenuity Core128/Elite China/worklist fault | External Remarks: 30.09.2021 01:31:08 UTC (86038250) Service details: worklist fail to refresh, reconfigure server node. service result: device returns to normal working condition after service Internal Remarks : 30.09.2021 01:31:11 UTC (86038250) "RDI Accuracy:yes" $Q7: (Confirmed reported problem & failed to refresh worklist) $Q8: ( check lan configuration & found worklist node lost) $Q9: ( reconfig His/Ris server node by lanconfig app ) $Q10: (Test and inspection data refer to: PA & Device returned to full functionality) | 3 |  | reconfig His/Ris server node via lanconfig | [12 empty] | 118447052 |  | 1370663 |  |  | Opened | [7 empty] | Others | [6 empty] | HIS/RIS Configuration | [5 empty] | Refresh |  |  | System issue | bar: No; No
1343 |  |  | Germany | Gesundheit Nord gGmbH Klinikverbund | 40002 | 4446800% |  | The system sporadically makes a strange USB sound like Connect / disconnect | Problem Description by Engineer: intermittent disconnect USB can be heard. Probably IBOX CT console USB connection. Resolution: Voltage on the power supply for the I-Box was increased to 5.33VDC. In the tests performed, the system corresponds to the manufacturer's specifications and is ready for clinical use. Tools & Equipment : MULTIMETER,2022-05-25 00:00:00 | 4 |  | intermittent disconnect USB | [12 empty] | 118451431 |  | 1367341 |  |  | Evaluation |  | TBD-NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Ibox USB | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | TBD-NoEnoughInfo | Gantry-HW-TBD-NoEnoughInfo | Gantry-HW-TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo | Gantry-HW | Gantry-HW-TBD-NoEnoughInfo | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1344 |  |  | Japan | Social Medical Corporation BOKOI, Nikko Memorial Hospital | 341011 | 4446800% |  | 01CS: RS: When creating an MPR from an Axail image of the head, it is recognized that there are 2 series of data. | Problem Description by Engineer : When MPR view is reconstructed from head axial images, the data is recognized as being two series. Resolution : I asked him to use the Merge function to process the data that was recognized as being split into two series. I asked him to change the scan type to 2D axial or perform a helical scan. Internal Comments: [Subject] 01CS:RS:Mr. Sato: When MPR view is reconstructed from head axial images, the data is recognized as being two series. [Internal Comments] None [4/8 Elements] Q5. How did the examination end? None. Q6. Did any mistaken findings (mistaken diagnosis or treatment) result from the problem? No. Q7. Was there any unintended unnecessary exposure? No Q8 Artifacts? None | 4 |  | Q&A | [12 empty] | 118453258 |  | 1372503 |  |  | NAC | [7 empty] | Application | [6 empty] | Q&A - MPR | [5 empty] | Refresh |  |  | System issue | bar: No; No
1345 |  |  | Japan | Social Medical Corporation BOKOI, Nikko Memorial Hospital | 341011 | 4446800% |  | 02CS: KKI: E-Stop open occurs when the bed is horizontally operated >> S_HMC_TAPE_SWITCH_PRESSED> Restored by system OFF / ON | Problem Description by Engineer: The E-stop opened when he was operating the bed horizontally. Resolution : Response Turning the system off/on achieved recovery. Internal Comments: [Subject] 02CS:KKI:Mr. Ando: The E-stop opened when he was operating the bed horizontally>>S_HMC_TAPE_SWITCH_PRESSED> It recovered when he turned the system off/on. [Internal Comments] [Results of checking the log] The following error occurred. S_HMC_TAPE_SWITCH_PRESSED It seems that the tape switch, Down button and Up button were pressed at the same time, and the emergency stop happened. I think it was a temporary problem, so I asked him to monitor the situation. [4/8 Elements] Q5. How did the examination end? None. Q6. Did any mistaken findings (mistaken diagnosis or treatment) result from the problem? No. Q7. Was there any unintended unnecessary exposure? No Q8 Artifacts? None | 4 |  | t seems that the tape switch, Down button and Up button were pressed at the same time, and the emergency stop happened. | [12 empty] | 118447376 |  | 1369140 |  |  | Investigation | [7 empty] | Others | [6 empty] | Misoperation | [5 empty] | Refresh |  |  | System issue | bar: No; No
1346 |  |  | Portugal | Centro Hospitalar de Vila Nova | 40030 | 4446800% |  | Faulty southern philips TV equipment does not perform topograms or exams | Work Order #: WO-04443053 29-Sep-2021 Jaime Granja The system does not meet the specification for the performed service and is not returned to use Follow Up Required: Replacement of the optical transmitter on the rotating part of the gantry. Eventually the problem may be in the associated optical fibers. Diagnostic performed by Engineer: Detected malfunction in the detector data communication circuit (DMS) for the stationary part of the gantry. It is necessary to replace the transmitter and/or associated optical fibers. Problem Description by Engineer : Customer Function/Role: Maintenance Department Device Use: Diagnostics Expected and Unexpected Results: Unable to acquire images. User Impact: Examinations to be done on another system on the same hospital or postponed Patient Impact: Examinations to be done on another system on the same hospital or postponed Current Software Version: 4.1.10 ----------------------------------------------------------------------------- | 3 |  | Product/Part - Module, Optical RX Stator Moog Velocity Part number - 453567538322 Product/Part - Harness-Rotor TX Fiber Assy Part number - 453567538381 Product/Part - Stator Optics Assembly Part number - 453567538341 Product/Part - MODULE, ROTOR ELECTRONICS TX Part number - 453567538334 | [12 empty] | 118451509 |  | 1369206 |  |  | Evaluation |  | TBD_OBS | [5 empty] | Engineering | [6 empty] | TX/RX | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Other | Gantry-HW-Other | Gantry-HW-Other | TBD_OBS | TBD-NoEnoughInfo | Gantry-HW | Gantry-HW-Other | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1347 |  |  | China | DaKuai Town Health, Fengquan District, Xinxiang City, Henan Province | 345009 | 4446700% |  | 02N/Ingenuity Core128/Elite China/Blue screen | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure T2 Activities:Problem Description: Blue screen Troubleshooting Action: Remotely instruct the customer to restart the system. Repair Action: Remotely instruct the customer to restart the system. Test/Inspection: Functional Check PASSED. The device returned to full functionality. Test Equipment: N/A SAP Transfer:Case Transferred to SAP | 3 |  | Blue screen reported from customer and system worked after restart the system. AI: Check with TSE and collect the bluescreen photo and bugreport for eng analysis | [12 empty] | 118439242 |  | 1366085 |  |  | Closed - No Further Investigation |  | CTUDT00632553 | [5 empty] | Engineering | [6 empty] | Blue Screen | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Blue-Screen | Console-SW-Blue-Screen | Console-SW-Blue-Screen | CTUDT00632553 |  | Console-SW | Console-SW-Blue-Screen | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | #N/A
1348 |  |  | Japan | Social Welfare Organization Saiseikai Imperial Gift Foundation, Inc. | 343000 | 4446700% |  | 01DS: KKI: Contrast-enhanced tracker CT host freezes and scan does not proceed >> S_MASTER_SEQUENCER_ACQ_NOT_OK / Error # 12 --Unknown error … | Integration : Confirmation of order 54220738 saved ----------------------------------------------------------------------------- Work Order #: WO-04429715 Kimura, Kazumasa The system meets the specification for the performed service and is returned to use. Internal Comments: [Subject] 01DS:KKI:Mr. Moriwaki: After a contrast trucker scan, the CT host froze and the scan did not progress.>>S_MASTER_SEQUENCER_ACQ_NOT_OK/ Error #12 - Unknown error 0x12> Restarting the CT host restored the scan. [Internal Comments] [Results of checking the log] The following error occurred. S_MASTER_SEQUENCER_ACQ_NOT_OK From S1 Error #12 - Unknown error 0x12 I did not find an error that allowed me to identify the cause of the problem. Dispatched: 1 RC FSE Registration Information: Name of FSE 1 who visited the site: Yasuhiro Kawashige FSE 1 visit date and time: 2021/09/2816:00 FSE 1 hours worked: 2 Diagnostic performed by Engineer : Response: Logging out/in to the CT host solved the problem. Problem Description by Engineer: After a contrast trucker scan, the CT host froze and the scan did not progress. ----------------------------------------------------------------------------- Work Order #: WO-04431295 28-Sep-2021 Kawashige, Yasuhiro The system meets the specification for the performed service and is returned to use. Internal Comments: [CX related] Did you agree on how to complete the report with the person in charge of devices? <Select one of the following> Yes. I completed normal reporting that day. [Work report (paper) submission] Did you submit a work report on paper? NO Resolution : [Work details] I investigated the problem of a contrast scan examination not proceeding, but it was not reproduced. I checked the log entries for both systems (HOST PC, CIRS) when the problem occurred, but there was no particular clear error. However, the CIRS had not been restarted for more than a month, so it is possible that the CIRS operation had become sluggish. Therefore, I restarted the Host PC and CIRS onsite, and I carried out an IQ test and a contrast scan test. I carried out the above and confirmed normal system operation. I apologize for the inconvenience caused. | 4 |  | the CIRS had not been restarted for more than a month, so it is possible that the CIRS operation had become sluggish. Therefore, I restarted the Host PC and CIRS onsite, and I carried out an IQ test and a contrast scan test. | [12 empty] | 118439293 |  | 1362781 |  |  | Investigation | [7 empty] | Others | [6 empty] | Reboot - CIRS | [5 empty] | Refresh |  |  | System issue | bar: No; No
1349 |  |  | Japan | 独立行政法人国立病院機構 | 341013 | 4446700% |  | 01CS: RS: Scan application crashed while shooting >> Scan was running as planned | Problem Description by Engineer: The scanning application crashed while shooting Resolution: The scan itself ran with the planned content. No noticeable hardware errors were found in the various logs. From the situation, it is assumed that the scanning application has crashed temporarily. Request a restart of the host computer and complete. Internal Comments: [Subject] 01CS: RS: Mr. Mamamoto: The scanning application crashed during shooting >> The scan was running as planned [Internal Comments] none [4/8 Elements] Q5 How did the test end ?: None Q6 Presence or absence of incorrect findings (misdiagnosis, mistreatment) due to problems: None Q7 Presence or absence of unintended unnecessary exposure: None Q8 Presence / absence of artifacts: None | 4 |  | "The scan itself ran with the planned content. No noticeable hardware errors were found in the various logs. From the situation, it is assumed that the scanning application has crashed temporarily. Request a restart of the host computer and complete | [12 empty] | 118440496 |  | 1362177 |  |  | Closed - No Further Investigation |  | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | App crash | [5 empty] | Refresh |  |  | System issue | System | TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | NoEnoughInfo |  | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | #N/A
1350 |  |  | United States | Philips Medical Capital | 40053 | 4446700% |  | 0711CDN8A FSE req'd for DICOM problem | Internal Remarks : FADB: (FSE Aware of Double Booking) Title: Internal Remarks Initials: Twila Actions Taken: Dispatched to FSE/Fairchild, Andrew - (9/28/21 @ 3pm EST). Outcome: Dispatched WO per PFSE/McDowell Preston to FSE/Fairchild IP: 88993323 Next Steps: None **************************** Work Order #: WO-04433863 28-Sep-2021 Fairchild, Andrew The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Diagnostic performed by Engineer : Cust says FSE Andrew Fairchild is aware They decline T2 Remote Tech Support ~ FSE req'ed. Diagnostic performed by Engineer : Troubleshooted issues with PACS, can query and send but not retrieve. Customer is checking with IT to confirm their side is good. If issue persists, will contact tech support for assistance. Follow Up Required : Additional PACS troubleshooting ----------------------------------------------------------------------------- | 3 |  | Customer is checking with IT to confirm their side is good. If issue persists, will contact tech support for assistance. | [12 empty] | 118443199 |  | 1365766 |  |  | Evaluation | [7 empty] | TBD | [6 empty] | PACS Retrieve | [5 empty] | Refresh |  |  | System issue | bar: No; No
1351 |  |  | China | Yangfangdang Hospital | 40035 | 4446600% |  | Wechat:02N/Ingenuity CT 728326/transfer fault | T2 Activities:Problem Description: Summary of customer's problems: ISP cannot be used to transfer back and forth with CT Detailed description of Customer's problems or malfunctions: ISP cannot be used to transfer back and forth with CT It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: Test CT connection in ISP, but configuration test failed Parameter record is ISP 172.20.110.25 CT 172.20.110.149 AE MXV Remote troubleshoot and test result: 1. confirm CT and ISP parameter setup, see if it's correct 2. Check network, check node type, re-add node SAP Transfer:Dispatched for On site Support External Remarks: 28.09.2021 07:53:59 UTC (86039460) Service details: ISP cannot transfer images with CT host to each other Service Result: The device returned to working normally. Internal Remarks : 28.09.2021 07:54:02 UTC (86039460) "RDI Accuracy:Yes" $Q7: (Confirmation & Description of problem) Confirmed reported problem: ISP can't deliver image to host. $Q8: (Troubleshooting & Result Check it, found have lots of image delivered with failure, $Q9: (Repair steps) Do reset spooler, and add ISP network node again, device work normal, | 3 |  | Do reset spooler, and add ISP network node again, device work normal, ISP can deliver image to host. AI: Check with TSE if it's CT or ISP issue. | [12 empty] | 118431322 |  | 1361845 |  |  | Closed - No Further Investigation | [7 empty] | TBD | [6 empty] | ISP Transfer | [5 empty] | Refresh |  |  | System issue | bar: No; No
1352 |  |  | Poland | Government Agency. | 341008 | 4446600% |  | Turn off gantry when tilting. In addition, the need to configure the PACS server. | The gantry turns off during tilting. Additionally, it is necessary to configure the PACS server. |  |  | Duplicate | [14 empty] | 1384552 |  |  | Duplicate | [7 empty] | Duplicate | [6 empty] | Duplicate | [5 empty] | Refresh |  |  | System issue | bar: No; No
1353 |  |  | Poland | Government Agency. | 341008 | 4446600% |  | Turn off gantry when tilting. In addition, the need to configure the PACS server. | PACS server configuration - correct. Tilting the gantry - the fuse was off. Enabled. Works properly. | 3 |  | unexpect fuse disconnect/E-stop open when tilting; config PACS | [12 empty] | 118435526 |  | 1363518 |  |  | Investigation |  | NotEngineering | [5 empty] | Engineering | [6 empty] | E-Stop - Tilting | [5 empty] | Refresh |  |  | System issue | System | NotEngineering | ShouldbeCustomer | NotEngineering-ShouldbeCustomer | NotEngineering-ShouldbeCustomer | NotEngineering | NotEngineering-ShouldbeCustomer | NotEngineering | NotEngineering-ShouldbeCustomer | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1354 |  |  | Puerto Rico | METRO MAYAGUEZ INC | 40009 | 4446600% |  | When performing calibration of the machine indicates that there is a strange artifact, it may be that the contrast... | Problem Description by Engineer : Problem description by engineer : "When performing calibration of the machine indicates that there is a strange artifact, it may be that the contrast medium" When was the first occurrence? Within... : 27-Sep-2021 09:00:00 Diagnostic performed by Engineer : The customer reported artifacts when doing patient studies. The image calibration failed to complete. The gantry cover was open. I visually inspected the area. I found contras found under the plastic center cover. This contrast is what is causing the image artifact. The contrast has to be removed Resolution : The contrast spots were removed from the Millar. The calibration was completed successfully. The CT was returned to the customer. | 3 |  | The contrast spots were removed from the Millar | [12 empty] | 118436926 |  | 1360478 |  |  | Investigation | [7 empty] | Others | [6 empty] | Contrast - Mylar | [5 empty] | Refresh |  |  | System issue | bar: No; No
1355 |  |  | China | 江苏省第二中医院 | 345027 | 4446500% |  | Wechat:02N/Ingenuity Core128/Elite China/Couch fault | BU Q&R Feedback: No Reason: Corrective Maintenance Safety Questions : Case: 0118428950 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Problem Reported by Customer : Couch fault T2 Activities : Problem Description: 客户问题概述： 其他类型故障。 客户问题详细描述或故障现象： 床卡死 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 床卡死动不了 远程诊断与测试结果： 客户说断电情况下床也推不动，应该是被撞的 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 检查床水平运动 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 26.09.2021 11:26:17 UTC (86036024)服务内容：调整床的外壳机械结构。 服务结果：维修后，设备恢复正常工作。 Internal Remarks : 26.09.2021 11:26:20 UTC (86036024) "RDI Accuracy:Yes" $Q7: (Confirmed reported problem. The couch can't move in or out.) $Q8: (Mannually move couch in or out and found it is hard to move. The cover of couch was out of shape, due to be knocked by patient table.) $Q9: (Adjusted the mechanism of couch cover.) $Q10: (Test and inspection data refer to PA. Device returned to full functionality.) $Q11: Calibrated Tools (NA). ----------------------------------------------------------------------------- | 3 |  | $Q8: (Mannually move couch in or out and found it is hard to move. The cover of couch was out of shape, due to be knocked by patient table.) | [12 empty] | 118428950 |  | 1358336 |  |  | Investigation |  | TBD-CouchCover | [5 empty] | Engineering | [6 empty] | Couch cover | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | Noise | Couch-HW-Noise | Couch-HW-Noise | TBD-CouchCover | Noise | Couch-HW | Couch-HW-Noise | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1356 |  |  | Japan | Social Medical Corporation BOKOI, Nikko Memorial Hospital | 341011 | 4446500% |  | 04DS: KK: The scan does not run after the rolling blackouts are restored. >> S_DMC_SHOT_COMM_COLLAPSE | Work Order #: WO-04417624 Kojima, Kenichi The system does not meet the specification for the performed service and is not returned to use Problem Description by Engineer: Unable to scan following restoration of power after a planned power outage.>> S_DMC_SHOT_COMM_COLLAPSE Diagnostic performed by Engineer: A S_DMC_SHOT_COMM_COLLAPSE error occurred. Physical channel order is 0 x 0. Internal Comments: [Subject] 04DS: KK: Mr. Ando: Unable to scan following restoration of power after a planned power outage.>> S_DMC_SHOT_COMM_COLLAPSE Shipped part: 453567538334: MODULE, ROTOR ELECTRONICS TX QTY: 1 ZREP/Traceable --None-- 453567538322: Module, Optical RX Stator Moog Velocity QTY: 1 ZREP/Traceable --None-- 453567057167:FRC ASSEMBLY QTY:1 ZREP/Traceable --None-- Dispatched: 1 RC FSE Registration Information: Name of FSE 1 who visited the site: Hodaka Hashimoto FSE 1 visit date and time: 2021/09/2616:00 FSE 1 hours worked: 4 ----------------------------------------------------------------------------- | 2 |  | [Details of work] After the rolling blackout, X-rays could not be emitted, so support was provided by telephone. Since it was completed by telephone support, we are returning during the visit. | [12 empty] | 118429061 |  | 1357741 |  |  | Closed - No Further Investigation | [7 empty] | TBD | [6 empty] | Unable to Scan | [5 empty] | Refresh |  |  | System issue | bar: No; No
1357 |  |  | China | 苏州广慈肿瘤医院有限公司 | 345019 | 4446400% |  | 02N/Ingenuity Core128/Elite China/equipment error | BU Q&R Feedback: No SAP Customer Symptom Code: HANG Customer Symptom Code: Hang/Crash Reason: Corrective Maintenance Engineer Failure Code: Wear-out/Breakage/Mechanical Failure Repair Activity Code: Instructed Customer Problem Reported by Customer : 02N/Ingenuity Core128/Elite China T2 Activities : Problem Description: 客户报修思创工作站电脑故障。 Troubleshooting Action: 客户报修的工作站不属于CT机电脑或星云工作站维修范围内，已联系客户去找思创工作站部门报修。 Repair Action: 客户报修的工作站不属于CT机电脑或星云工作站维修范围内，已联系客户去找思创工作站部门报修。 Test/Inspection: customer confirm can scan patient and image ok.functional Check PASSED. Device returned to full functionality. Test Equipment: N/A Safety Questions : Case: 0118427823 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Bo Last updated by: Wang Bo 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 3 |  | 客户报修的工作站不属于CT机电脑或星云工作站维修范围内，已联系客户去找思创工作站部门报修。 | [12 empty] | 118427823 |  | 1357795 |  |  | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | ISP | [5 empty] | Refresh |  |  | System issue | bar: No; No
1358 |  |  | China | 江西省儿童医院 | 40041 | 4446300% |  | Wechat:02N/Ingenuity CT 728326/Remove PRS | BU Q&R Feedback: No SAP Customer Symptom Code: CENV Customer Symptom Code: Customer Environmental Concern Reason: Corrective Maintenance Engineer Failure Code: Environmental Cause Repair Activity Code: Configuration Change Problem Reported by Customer : Remove PRS Safety Questions : Case: 0118421861 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Bai Jingwen Last updated by: Bai Jingwen 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Troubleshooting Action: answer customer question Repair Action: answer customer question Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: N/A ----------------------------------------------------------------------------- | [15 empty] | 118421861 |  | 1355688 |  |  | Closed - Non-Complaint | [20 empty] | Refresh |  |  | System issue | bar: No; No
1359 |  |  | China | 林甸县医院 | 345038 | 4446300% |  | 02N/Ingenuity Core128/Elite China/Printing failed | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ Safety Questions : Case: 0118421668 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingsong Last updated by: Lu Mingsong 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: Printing failed Troubleshooting Action: Remotely instruct the customer to Select printer settings. Repair Action: Remotely instruct the customer to Select printer settings. Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: N/A ----------------------------------------------------------------------------- | 3 |  | Remotely instruct the customer to Select printer settings. | [12 empty] | 118421668 |  | 1355556 |  |  | Closed - No Further Investigation | [7 empty] | Application | [6 empty] | Print setting | [5 empty] | Refresh |  |  | System issue | bar: No; No
1360 |  |  | Japan | 社会医療法人 母恋 日鋼記念病院 | 341011 | 4446300% |  | 02DS: KK: Last night, when I tried to shoot, E-Stop … | Work Order #: WO-04408495 Konishi, Hideo Regulatory Questions : N/A Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- Work Order #: WO-04408430 Kojima, Kenichi The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : 昨晩、撮影しようとしたらE-Stop Openが発生し、その後Closeできなかった。>>TEALからの再起動にて復帰。>>S_MC_DEVICE_VERT_POT_OPEN、S_VMC_AXIS_UNALLOWED_MOVE Diagnostic performed by Engineer : S_MC_DEVICE_VERT_POT_OPEN、S_VMC_AXIS_UNALLOWED_MOVEが発生しE-Stop Openしていました。 再Closeしようとしていましたが_MPS_MAIN_DRIVE_NOT_CLOSINGとなりCloseできなかったようです。 TEALからの再起動にて復旧。本日検査終了後対応。 Internal Comments : [Subject] 02DS:KK:ｻﾄｳ様:昨晩、撮影しようとしたらE-Stop Openが発生し、その後Closeできなかった。>>TEALからの再起動にて復帰。>>S_MC_DEVICE_VERT_POT_OPEN、S_VMC_AXIS_UNALLOWED_MOVE [出庫部品] 459800623331:Vertical Transducer QTY:1 ZSID --None-- [ディスパッチ] 1 RC -----------------FSE登録情報----------------- -訪問FSE1氏名 小西 英生 -FSE1訪問日時 2021/09/2416:00 -FSE1作業時間 4 [出荷指示] -----------------出荷指示#1----------------- -配送条件 CPU(倉庫引取り・輸送手配無し) -納期 2021/09/24 13:00 必着 -配送先名 札幌倉庫止め・FSEが引取り -受取者名 小西 英生 -電話番号 080-1605-0255 -その他 -依頼者名 小島健一 ----------------------------------------------------------------------------- | 4 |  | Part/Product - Vertical Transducer Part Number - 459800623331 | [12 empty] | 118421645 |  | 1351234 |  |  | Closed - No Further Investigation |  | NotEngineering-partreplace | [5 empty] | Engineering | [6 empty] | Vertical Transducer | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | VerticalTransducer | Couch-HW-VerticalTransducer | Couch-HW-VerticalTransducer | NotEngineering-partreplace |  | Couch-HW | Couch-HW-VerticalTransducer | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | #N/A
1361 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 337000 | 4446300% |  | Gantry dönerken sürtme sesi geliyor | Integration: Confirmation of order 54206872 has been saved. ----------------------------------------------------------------------------- Work Order #: WO-04409960 24-Sep-2021 Can Ozgenlik The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Rubbing noise in the gantry. Resolution : We visited the site for the gantry noise and we found that one of the fan cables was a little higher than where it should be. The cable was fixed in its position. No other problems were found. Test scans with phantom were completed successfully. Device is up. ----------------------------------------------------------------------------- Work Order #: WO-04409565 Soner Saglam The system does not meet the specification for the performed service and is not returned to use Problem Description by Engineer : Rubbing noise in the gantry. Diagnostic performed by Engineer : In the video sent, rubbing noise was found. We asked the user to stop using the device in this condition. Site visit is required. Follow Up Required : In the video sent, rubbing noise was found. We asked the user to stop using the device in this condition. Site visit is required. ----------------------------------------------------------------------------- | 2 |  | The system was examined for the problem of sound coming from the gantry, and it was seen that one of the fan cables was a little higher than it should be. The cable is fixed in place. No other adverse events were observed. Test shots were made with the Phantom, no problem was seen, the device is in active condition. Check Ali to collect picture of issue | [12 empty] | 118422825 |  | 1352051 |  |  | Closed - No Further Investigation | [7 empty] | Operation | [6 empty] | Fan cable | [5 empty] | Refresh |  |  | System issue | bar: No; No
1362 |  |  | China | 四川省宜宾市第二中医院 | 345035 | 4446200% |  | 02N/Ingenuity Core128/Elite China/information consultation | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer : information consultation Next Step Action : 邮件ASP chengduyunjia跟进 维护保养体系的相关规章制度要求 Safety Questions : Case: 0118416849 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Zhou Xiaohong Last updated by: Zhou Xiaohong 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | [15 empty] | 118416849 |  | 1352299 |  |  | Closed - Non-Complaint | [20 empty] | Refresh |  |  | System issue | bar: No; No
1363 |  |  | United States | Philips | 40021 | 4446200% |  | Ingenuity 5000 has had continued issues with Tube heat | •Feedback Type: Technical Issue •Other Details: Other •Submitter Email: rhonda.swain@philips.com •CFM Form Received Date (Outlook email received date): 09/25/2021 Ingenuity 5000 installed October 2021 has had continued issues with Tube heat. The customer has placed multiple calls to the CCSC and me for help. We have adjusted the exam cards, and the issue continues. This week, I was onsite for a follow-up visit and learned that the tube heat issues had not been resolved. The hospital is a level two trauma center, and the exam cards can be complex. We see no reason in the exam cards that would be causing the tube heat issues. This does cause a delay in patient care. The staff sometimes get messages that they have to wait three to four minutes for the tube to cool before they can scan. All pertinent details: The reported issue did not cause harm or injury to a patient,but does sometimes cause a delay in patient care. |  |  | We have adjusted the exam cards, and the issue continues | [14 empty] | 1355116 |  |  | Investigation |  | CTD00005656 | [5 empty] | Engineering | [6 empty] | Tube Overheat | [5 empty] | Refresh |  |  | System issue | System | Gantry-SW | Tube over heat | Gantry-SW-Tube over heat | Gantry-SW-Tube over heat | CTD00005656 | Tube Heat | Gantry-SW | Gantry-SW-Tube over heat | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1364 |  |  | United States | McLaren Macomb | 40021 | 4446200% |  | 4.1.10 Tube heat again | Integration : Confirmation of order 54202832 saved ----------------------------------------------------------------------------- Work Order #: WO-04404140 29-Sep-2021 Koren Moshe The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: Diagnostic Expected and Unexpected Results: Ongoing Tube Heat Warnings User Impact: No Patient Impact: No Current Software Version: 4.1.10 Diagnostic performed by Engineer : Issue will be resolved by chief TECH and APPS. ----------------------------------------------------------------------------- Work Order #: WO-04412476 27-Sep-2021 Mona Tony The system meets the specification for the performed service and is returned to use Resolution : Ran Filament Cals, copied Exam-card on external drive & Bug-report. Send information for analysis. David Gaffney the Hospital Radiology manager will provide a technologist contact information to work with Philips applications to tweak the Exam-Cards. Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: ----------------------------------------------------------------------------- Work Order #: WO-04403795 Maier, Donnie The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Follow Up Required : Several calls with tube heat issues in last 6 months. Create escalation, customer complaint, as necessary. Clin Apps is onsite and Remote Clin Apps has been involved. Protocols have been adjusted as far as lowering DRI, etc. already and cu is still getting tube heat warnings. May want to look into if protocols were converted for example as part of an upgrade. Diagnostic performed by Engineer : Clinical Applications is onsite and remote clin apps has also assisted on this site. Protocols have been adjusted (Bolus Tracker cycle times decreased, lower DRI, etc.) according to earlier cases. Dispatching for FSE to escalate and create Customer Complaint if necessary, etc. Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: Diagnostic Expected and Unexpected Results: Ongoing Tube Heat Warnings User Impact: No Patient Impact: No Current Software Version: 4.1.10 ----------------------------------------------------------------------------- Work Order #: WO-04403695 Hirchak, Michele The system meets the specification for the performed service and is returned to use Resolution : Called Rhonda Swain, apps onsite, to let her know the FSE will be there tomorrow at 2pm, according to the notes in the case. Closing clinical work order. MBH Problem description by Clinical Support Specialist : Customer Function/Role: Apps onsite. Device USE: IN USE Expected and Unexpected Results:Tube heat, happening very often, protocols are set very low. User Impact: NONE Patient Impact: NONE Current Software Version: 4.1.10 No rescan, no re-injection, no patient harm. Diagnostic performed by Engineer : Rhonda Swain, CES, is onsite. She observed multiple tube heat messages yesterday. Called today to review protocol changes from previous customer calls. PE Abd/Pelvis reviewed. Auto Scan time off, set RT to .75. DRI settings are low. Protocols should not be giving tube heat. Customer requesting FSE to come and troubleshoot. I will have an RSE look at logs. Case yesterday 9/22/21 @ 3:30pm EST. MBH ----------------------------------------------------------------------------- | 5 |  | We have adjusted the exam cards, and the issue continues | [12 empty] | 118418154 |  | 1349067 |  |  | Closed - No Further Investigation |  | Tube Heat | [5 empty] | Engineering | [6 empty] | Tube Overheat | [5 empty] | Refresh |  |  | System issue | System | Gantry-SW | Tube over heat | Gantry-SW-Tube over heat | Gantry-SW-Tube over heat | Tube Heat | Tube Heat | Gantry-SW | Gantry-SW-Tube over heat | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1365 |  |  | Chile | BANCO DE CHILE | 341005 | 4446100% |  | CL1688-SOLICITA APOYO PARA REALIZAR AJUSTE DE HORA EN EL EQUIPO | Integration:Confirmation of order 54198333 saved ----------------------------------------------------------------------------- Work Order #: WO-04397274 [REDACTED] The system meets the specification for the performed service and is returned to use. Problem Description by Engineer : Problem description by engineer : Set time in the equipment. Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected results: No time problems in the system. Patient Impact/Outcome: No impact. User Impact: No impact. Current Software Version: 4.1.10 Resolution: The time is adjusted in the reconstructor. Test is performed with an exam and it is verified that now the images have the correct time. Internal Comments : Error # and/or description of error : Incorrect time in reconstructor. Repair Action: The time is adjusted in the reconstructor. Test is performed with an exam and it is verified that now the images have the correct time. Diagnostic performed by Engineer: There is a problem with the time of the reconstructor. The images contain at a different time than the one in the main computer. ----------------------------------------------------------------------------- | 3 |  | Resolution: The time is adjusted in the reconstructor. Test is performed with an exam and it is verified that now the images have the correct time. | [12 empty] | 118412909 |  | 1349710 |  |  | Evaluation |  | TBD-TimeSync | [5 empty] | Engineering | [6 empty] | Time sync | [5 empty] | Refresh |  |  | System issue | System | Console-SW | TimeSync | Console-SW-TimeSync | Console-SW-TimeSync | TBD-TimeSync |  | Console-SW | Console-SW-TimeSync | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | #N/A
1366 |  |  | Japan | Maki Hospital | 341017 | 4446100% |  | 03CS:KK:検査しようとしたらエラーが出て停止。>>Anode Arc発生。>>再スキャンで検査可。 | Integration : Confirmation of order 54194725 saved Integration:According to factory calendar 20210923 (Posting date) is not a workday (check entry) ----------------------------------------------------------------------------- Work Order #: WO-04392955 Kojima, Kenichi The system meets the specification for the performed service and is returned to use. Resolution: Anode arc had occurred. It seemed to be a sporadic problem, so he agreed to keep an eye on it. I reported the issue to an FSE who is scheduled to visit the site for inspection. Problem Description by Engineer : An error occurred when he tried to run a scan, and it stopped.>> Anode arcing occurred.>> He was able to perform the examination with a rescan. Internal Comments: [Subject] 03CS:KK: Mr. Togawa: An error occurred when he tried to run a scan, and it stopped.>> Anode arcing occurred.>> He was able to perform the examination with a rescan. [4/8 Elements] Q5. How did the examination end? None. Q6. Did any mistaken findings (mistaken diagnosis or treatment) result from the problem? No. Q7. Was there any unintended unnecessary exposure? No Q8 Artifacts? None ----------------------------------------------------------------------------- | 4 |  | wait and see. | [12 empty] | 118369358 |  | 1342796 |  |  | Investigation | [7 empty] | TBD | [6 empty] | TBD | [5 empty] | Refresh |  |  | System issue | bar: No; No
1367 |  |  | China | 息烽县中医医院 | 345014 | 4446000% |  | Wechat：02N/Core128/The key of couch fault | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : The key of couch fault Safety Questions : Case: 0118363153 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Pan Chuchu Last updated by: Pan Chuchu 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action : 客户：CT BOX降床键故障 已通知值班工程师 刘国平18066898085 先跟进客户判断是否需要派工，短信已发 Next Step Action : Site1 FSE 唐广林 18084189777 正在通话中 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : 唐工反馈他明天去正安的单子维修，这个场地去不了 联系SITE2 FSE 胡克 18585068636 告知可以派给他，已发派工短信通知到工程师 ----------------------------------------------------------------------------- | 3 |  | CT Box button | [12 empty] | 118363153 |  | 1342767 |  |  | Closed - No Further Investigation |  | CTUDT00632185 | [5 empty] | Engineering | [6 empty] | CT Box | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | CT Box | Gantry-HW-CT Box | Gantry-HW-CT Box | CTUDT00632185 | CT Box | Gantry-HW | Gantry-HW-CT Box | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1368 |  |  | Japan | Independent administrative institution National Hospital Organization | 341013 | 4446000% |  | 04CS:KKI:撮影開始時に、エラー発生＞＞S_DMC_SHOT_READY_STATE_TIMEOUT | Work Order #: WO-04385479 Kimura, Kazumasa Problem Description by Engineer: Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer: N/A ----------------------------------------------------------------------------- | 4 |  | TBD | [12 empty] | 118363767 |  | 1343092 |  |  | Closed - No Further Investigation | [7 empty] | TBD | [6 empty] | TBD | [5 empty] | Refresh |  |  | System issue | bar: No; No
1369 |  |  | Puerto Rico | HOSPITAL DR SUSONI INC | 40013 | 4446000% |  | Table noise when going up and down | Integration : Confirmation of order 54193468 saved ----------------------------------------------------------------------------- Work Order #: WO-04390999 28-Sep-2021 Ramos Rivera, Pablo The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Problem description by engineer : noise in vertical movement When was the first occurrence? Within... : 21-Sep-2021 13:00:00 How often is issue occurring? : Intermittent Malfunction area : Table Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: abnormal noise Patient Impact/Outcome: n/a User Impact: abnormal noise Current Software Version: unknown Resolution : Lubricated the vertical movement mechanism of the patient support. Now there is no noise when it moves vertically. ----------------------------------------------------------------------------- Work Order #: WO-04390455 Caldelas Aldana, Carlos Manuel The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Problem description by engineer : noise in vertical movement When was the first occurrence? Within... : 21-Sep-2021 13:00:00 How often is issue occurring? : Intermittent Malfunction area : Table Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: abnormal noise Patient Impact/Outcome: n/a User Impact: abnormal noise Current Software Version: unknown Diagnostic performed by Engineer : abnomral noise detected in vertical movement of patient support. no errors reported while in use. Follow Up Required : Error # and/or description of error : abnormal noise Repair Action : n/a What are the next Troubleshooting steps : inspect, clean and lubricate inf necesary vertical elements of patient suport. Technician Name : Caldelas Aldana, Carlos Manuel Resolution : FSE Visit Internal Comments : Error # and/or description of error : abnormal noise Repair Action : n/a ----------------------------------------------------------------------------- | 3 |  | inspect, clean and lubricate inf necesary vertical elements of patient suport. | [12 empty] | 118367512 |  | 1342002 |  |  | Closed - No Further Investigation |  | E-125059 | [5 empty] | Engineering | [6 empty] | Table Noise | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | Lubricant | Couch-HW-Lubricant | Couch-HW-Lubricant | E-125059 | Lubricant | Couch-HW | Couch-HW-Lubricant | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1370 |  |  | United States | Philips Medical Capital | 40053 | 4446000% |  | 1301CDN - Customer requested a new destination added to the system | Integration : Confirmation of order 54193152 saved ----------------------------------------------------------------------------- Work Order #: WO-04390599 22-Sep-2021 Fairchild, Andrew The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: Diagnostics Expected and Unexpected Results: Add a new destination User Impact: No Patient Impact: No Current Software Version:4.1 Resolution : Customer was unable to communicate to a newly installed site. Changed AE-Title in PACS to allow for transmission of data. System restored to full functionality without restriction. ----------------------------------------------------------------------------- Work Order #: WO-04390532 Dangos John The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: Diagnostics Expected and Unexpected Results: Add a new destination User Impact: No Patient Impact: No Current Software Version:4.1 Diagnostic performed by Engineer : Customer requested a new destination added but does not have all the network information needed to add remotely. Customer will get information from IT dept and will have it ready for local FSE when he arrives Follow Up Required : Customer requested a new destination added but does not have all the network information needed to add remotely. Customer will get information from IT dept and will have it ready for local FSE when he arrives. The only Network info provided is the AE Title: MVODC-CT Customer will still need to provide IP address:? Port:?, Name:? ----------------------------------------------------------------------------- | 5 |  | Customer requested a new destination added | [12 empty] | 118367566 |  | 1341765 |  |  | NAC | [7 empty] | NAC | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | bar: No; No
1371 |  |  | United States | Westgate Hospital | 40001 | 4445900% |  | 82488278 FAILED CALIBRATION DUE TO TUBE ARCING | Internal Remarks : **CUST REQUESTED ONSITE SUPPORT Interaction Type: New Case Caller: jim dodson - IPCR 7300 N 99TH AVE GLENDALE AZ IP & Issue: 82488278 FAILED CALIBRATION DUE TO TUBE ARCING Event Type: Incident Service Request: Onsite Did Customer Decline Remote Support? Yes Preferred Service Time: N/A Entitlement Options: Quote Reason for Entitlement Claim: N/A Action Taken: CREATED CASE AND QUOTE Next Step Action: Awaiting Quote Verbal PO Number: 51-982518 Verbal PO Approver: JIM DODSON Verbal PO CAP: Integration : VBAKKOM has been processed successfully-VBAPKOM has been processed successfully-VBAPKOM has been processed successfully-KONVKOM has been processed successfully-KONVKOM has been processed successfully-KONVKOM has been processed successfully-KONVKOM has been processed successfully-KONVKOM has been processed successfully-KONVKOM has been processed successfully-KONVKOM has been processed successfully-KONVKOM has been processed successfully-REQ2RES DMR 0018218618 has been saved-MPG is present on the product code : P9-MPG is present on the product code : P9 Integration : Confirmation of order 54187107 saved ----------------------------------------------------------------------------- Work Order #: WO-04381399 27-Sep-2021 Leon Peter The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : Performed long tube conditioning. Performed filament calibrations. Performed short tube conditioning. Ran test scans on system, all tests passed￼. ----------------------------------------------------------------------------- | 4 |  | Actions Taken: Dispatched per fse request Outcome: scheduled for 9/22 | [12 empty] | 118360881 |  | 1341234 |  |  | Closed - No Further Investigation |  | CTUDT00630085 | [5 empty] | Engineering | [6 empty] | Tube Arc | [5 empty] | Refresh |  |  | System issue | System | IC-HW | Tub Arc | IC-HW-Tub Arc | IC-HW-Tub Arc | CTUDT00630085 | Tub Arc | IC-HW | IC-HW-Tub Arc | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1372 |  |  | China | 四川省宜宾市第二中医院 | 345035 | 4445800% |  | 02N/Ingenuity Core128/Elite China/ | BU Q&R Feedback: No Reason: Corrective Maintenance Next Step Action : 故障：添加打印机 RSE:LUMINGSONG Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ Safety Questions : Case: 0118356269 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Gao Luqi Last updated by: Gao Luqi 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : Xu Zehong ----------------------------------------------------------------------------- | [15 empty] | 118356269 |  | 1336226 |  |  | Closed - Non-Complaint | [20 empty] | Refresh |  |  | System issue | bar: No; No
1373 |  |  | China | Shandong Province Chest Hospital | 40025 | 4445700% |  | 02N/Ingenuity CT 728326/Software login failure | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 02N/Ingenuity CT 728326 Safety Questions:Case: 0118355210 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingsong Last updated by: Lu Mingsong 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer's problems: Can't log on to the scan interface. Detailed description of Customer's problems or malfunctions: Can't log on to the scan interface. It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: N/A Remote troubleshoot and test result: N/A PRS TR status: Only M2M is not connected, ask FSE check onsite. Repair Action: Suggest onsite service procedures: Can't log in to the scan interface. It's a software bug; use the "smart card" to reactivate the user. We recommend spare parts 12NC and names: N/A Memo of Customer request: None SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Next Step Action: huang zhaokun External Remarks: 19.09.2021 01:21:26 UTC (86038405) Fault Description: The CT account can't log in to the software. It has a blue screen after logging in. Service Details: Used IST privileges to log in to the configuration and acquire the Bug report.Replaced the Hosp-Lan network cable. Swapped the switch interface position where the network cable is located. Set the network card's speed and changed the camera Association time. Service result: device returns to normal working condition after service.Need follow-up to monitor usage. Internal Remarks : 23.09.2021 13:35:47 UTC (86038405) "RDI Accuracy: YES." $Q7: ( Confirmed reported problem: Software can't login) ) $Q8: ( Login in Software use smart card and configuration. & L3 support reply this is Windows 10 Bug and need to Fixed future.) $Q9: ( Login Software use smart card and configuration.) $Q10: ( Test and inspection data refer to: PA . Device returned to full functionality.) $Q11: Calibrated Tools name (N/A) ^Current Tube installed date: N/A ^Current Tube usage: N/A ----------------------------------------------------------------------------- | 3 |  | Blue Screen | [12 empty] | 118355210 |  | 1335501 |  |  | Closed - No Further Investigation |  | CTUDT00632553 | [5 empty] | Engineering | [6 empty] | Blue Screen | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Blue-Screen | Console-SW-Blue-Screen | Console-SW-Blue-Screen | CTUDT00632553 | Blue-Screen | Console-SW | Console-SW-Blue-Screen | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1374 |  |  | China | Fujian Provincial Hospital | 345011 | 4445700% |  | 02N/Ingenuity Core128/Elite China//system performance low | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Calibration/Adjustment Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ SAP Transfer:Dispatched for On site Support Safety Questions:Case: 0118354932 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Gaoge Last updated by: Wang Gaoge 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer's problems: system is running slow or crashed. Detailed description of Customer's problems or malfunctions: The system is slow and is stuck and needs debugging It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: System is slow. Remote troubleshoot and test result: Consider mcafee causing slow operation PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Replace virus database + close Mcafee update Test/Inspection: N/A Test Equipment: N/A SAP Transfer:Case Transferred to SAP Next Step Action: LI FENG Internal Remarks : 19.09.2021 02:15:48 UTC (86036224) RDI Accuracy: Yes $Q7: (Confirmed reported problem. The CT App run is slowly.) $Q8: (On site check and find the McAfee version is ok.) $Q9: (Delete the rubbish file and do disk clean up.) $Q10: (Test and inspection data refer to: PA.+Device returned to full functionality.) $Q11: Calibrated Tools name (FLUKE 287C Multimeter) S/N (16490128) NextCalibration date (10/20/2021) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date: 09/13/2020 ^Current Tube usage: 17832 External Remarks: 19.09.2021 02:15:38 UTC (86036224) Service details: CT software runs slowly and is stuck.Check on the site to make sure that the virus database of the Mcafee anti-virus software is normal.Clean junk files and clean disk. service result: device returns to normal working condition after service ----------------------------------------------------------------------------- | 3 |  | System is slow. Delete the rubbish file and do disk clean up. | [12 empty] | 118354932 |  | 1335350 |  |  | Closed - No Further Investigation |  | CTUDT00632261 | [5 empty] | Engineering | [6 empty] | System slow-DISK FULL | [5 empty] | Refresh |  |  | System issue | System | Console-SW | DiskFull | Console-SW-DiskFull | Console-SW-DiskFull | CTUDT00632261 | DiskFull | Console-SW | Console-SW-DiskFull | No | #REF! | #REF! | #REF! | #REF! | #REF! | No | #REF! | #REF! | #REF! | #REF! | #REF! | y | y | y | #REF! | System slow | #REF!
1375 |  |  | China | Yangfangdang Hospital | 40035 | 4445600% |  | Wechat:02N/Ingenuity CT 728326/print information display setting | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Incompatible Use Repair Activity Code: Instructed Customer Problem Reported by Customer: wechat:02N/Ingenuity CT 728326 Safety Questions:Case: 0118349098 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Bai Jingwen Last updated by: Bai Jingwen 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: print information display setting Troubleshooting Action: the time and date not display on head images print. instruct customer change the information setting. system is ok. Repair Action: Functional Check PASSED. The device returned to full functionality. Test/Inspection: Functional Check PASSED. The device returned to full functionality. Test Equipment: N/A SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | the time and date not display on head images print. instruct customer change the information setting. system is ok. | [12 empty] | 118349098 |  | 1333533 |  |  | Closed - No Further Investigation | [7 empty] | Application | [6 empty] | Print setting | [5 empty] | Refresh |  |  | System issue | bar: No; No
1376 |  |  | China | Medical center (contagious disease, He nan, Xu chang) | 40047 | 4445600% |  | System prompted e-stop problem, replaced front control panel | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 02n/Ingenuity CT 728326 Safety Questions:Case: 0118348031 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer's problems: Other types of error. Detailed description of Customer's problems or malfunctions: The customer reported for the system prompting the emergency switch problem It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: The customer reported for the system prompting the emergency switch problem Remote troubleshoot and test result: The customer rotated emergency switch and it restore. The panel has an emergency switch problem PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Replace the gantry left side front and back panels 0 "459800484501 CONTROL PANEL ASSEMBLY-FRONT LH 459800484531 OPERATOR PANEL ASSEMBLY-REAR RH" SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Next Step Action: PPD 459800484501*1 459800484531*1, 24H, FSE: Huang yong -86-18039573296 arranged shipment External Remarks: 18.09.2021 08:16:23 UTC (86038368) service details: [emergency switch on the panel is faulty. Replace the front panel.] Service result: [The device returns to normal working condition after maintenance.] Other important issues: [None] Internal Remarks : 18.09.2021 08:16:26 UTC (86038368) RDI Accuracy: Yes $Q7: (Confirm reported problem,estop button broken ) $Q8: (Performed visual inspection, found the estop button broken, need to replace the front control panel ) $Q9: (Replace the front control panel ) $Q10: (Test and inspection data refer to: PA. Device returned to full functionality.) $Q11:Calibrated Tools(N/A). ----------------------------------------------------------------------------- | 3 |  | estop button broken, Replace the front control panel 459800484501*1 | [12 empty] | 118348031 |  | 1333173 |  |  | Investigation |  | NotEngineering | [5 empty] | Engineering | [6 empty] | E-Stop button | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Control Panel | Gantry-HW-Control Panel | Gantry-HW-Control Panel | NotEngineering |  | Gantry-HW | Gantry-HW-Control Panel | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | #N/A
1377 |  |  | China | Sichuan Province Yibin City No.2 T.C.M. Hospital | 345035 | 4445600% |  | 02N/Ingenuity Core128/Elite China/artifact | BU Q&R Feedback: No SAP Customer Symptom Code: QUAL Customer Symptom Code: Quality Problem with Image or Signal Reason: Corrective Maintenance Engineer Failure Code: Performance or Calibration Drift/Out-of-Specification Repair Activity Code: Calibration/Adjustment Problem Reported by Customer: 02N/Ingenuity Core128/Elite China Safety Questions:Case: 0118347991 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Case Transferred to SAP T2 Activities:Problem Description: artifact Troubleshooting Action: customer call service, after air calibration, scan with artifact images. instruct customer perform air calibration again, system is ok. Repair Action: perform air calibration. system is ok. Test/Inspection: Functional Check PASSED. The device returned to full functionality. Test Equipment: N/A ----------------------------------------------------------------------------- | 3 |  | scan with artifact images. instruct customer perform air calibration again, system is ok. | [12 empty] | 118347991 |  | 1333135 |  |  | Investigation | [7 empty] | Application | [6 empty] | Air Cal. | [5 empty] | Refresh |  |  | System issue | bar: No; No
1378 |  |  | China | Medical center (contagious disease, He nan, Xu chang) | 40047 | 4445500% |  | 02N/Ingenuity CT 728326/gantry cannot power up | BU Q&R Feedback: No SAP Customer Symptom Code: BOOT Customer Symptom Code: Boot up/Power-up Problem Reason: Corrective Maintenance Engineer Failure Code: Electrical/Electronic Failure Repair Activity Code: Instructed Customer Problem Reported by Customer: 02N/Ingenuity CT 728326 Safety Questions:Case: 0118343291 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Gaoge Last updated by: Wang Gaoge 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities: Problem Description: The gantry can't start up. Troubleshooting Action: Guided the customer to check the E-stops on the front and back of the gantry. The equipment is back to normal use. Repair Action: Guided the customer to check the emergency switch. The E-stop on the back of the gantry was pressed by mistake. After restoring the emergency switch, the equipment is back to normal. Test/Inspection: N/A Test Equipment: N/A SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | Guided the customer to check the emergency switch. The E-stop on the back of the gantry was pressed by mistake. After restoring the emergency switch, the equipment is back to normal. | [12 empty] | 118343291 |  | 1329507 |  |  | Closed - No Further Investigation | [7 empty] | Application | [6 empty] | False alert | [5 empty] | Refresh |  |  | System issue | bar: No; No
1379 |  |  | China | Ning'an City No. 2 People's Hospital | 345039 | 4445500% |  | Wechat:02N/Ingenuity Core128/Elite China/Required Operation Instructions | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: Instructed Customer Safety Questions:Case: 0118343310 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Zixuan Last updated by: Wang Zixuan 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Problem Reported by Customer : Required Operation Instructions T2 Activities : Problem Description: Required Operation Instructions Troubleshooting Action: Required Operation Instructions Repair Action: Give customer IFU and TRG file Test/Inspection: Functional Check PASSED. The device returned to full functionality. Test Equipment: NA SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | [15 empty] | 118343310 |  | 1329523 |  |  | Closed - Non-Complaint | [20 empty] | Refresh |  |  | System issue | bar: No; No
1380 |  |  | Turkey | ADIYAMAN OZEL SAGINSTURTAS SAN VE TIC A.S | 341006 | 4445500% |  | Hava kalibrasyonu yapılamıyor, konstrat uyarısı veriyor | Integration: Confirmation of order 54172012 has been saved. ----------------------------------------------------------------------------- Work Order #: WO-04360531 23-Sep-2021 Turgay Baylan The system meets the specification for the performed service and is returned to use Resolution : We opened the gantry covers to clean any contrasting agents. The system was returned to use in functional condition after testing with a patient. Problem Description by Engineer : Error displayed during air calibration. ----------------------------------------------------------------------------- Work Order #: WO-04360428 Soner Saglam The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, no follow-up requested Diagnostic performed by Engineer : Contrasting agent leak. There are artifacts in images. The leak should be cleaned with a site visit. Problem Description by Engineer : Error displayed during air calibration. Follow Up Required : Contrasting agent leak. There are artifacts in images. The leak should be cleaned with a site visit. ----------------------------------------------------------------------------- | 3 |  | Contrasting agent leak. There are artifacts in images. The leak should be cleaned with a site visit. | [12 empty] | 118344743 |  | 1325909 |  |  | Evaluation | [7 empty] | Customer | [6 empty] | Contrast | [5 empty] | Refresh |  |  | System issue | bar: No; No
1381 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 337000 | 4445500% |  | CT Box Arızası | Integration: Confirmation of order 54170625 has been saved. ----------------------------------------------------------------------------- Work Order #: WO-04358387 27-Sep-2021 Can Ozgenlik The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Some of the keys on the CTBox are out of use. Parts Consumed : 1.00000,459801705022,CT BOX ASSEMBLY_BEST Resolution : The problem was solved after replacing the faulty CTbox. Device is operational. ----------------------------------------------------------------------------- Work Order #: WO-04358334 Soner Saglam The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, no follow-up requested Resolution : CTBox replacement. Problem Description by Engineer : Some of the keys on the CTBox are out of use. Follow Up Required : CTBox replacement. ----------------------------------------------------------------------------- | 3 |  | Product/Part - CT BOX ASSEMBLY_BEST Part number - 459801705022 | [12 empty] | 118342763 |  | 1324559 |  |  | Closed - No Further Investigation |  | CTUDT00632185 | [5 empty] | Engineering | [6 empty] | CT Box | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | CT Box | Gantry-HW-CT Box | Gantry-HW-CT Box | CTUDT00632185 | CT Box | Gantry-HW | Gantry-HW-CT Box | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1382 |  |  | United States | Philips Medical Capital | 40055 | 4445500% |  | Backup media damaged, replaced USB | Internal Remarks : FADB: (FSE Aware of Double Booking) Title: Internal Remarks Initials: Twila Actions Taken: Per message sent to me in Teams from FSE/Turley, Rich - Create Case and Dispatch WO for 9/17/21 (afternoon). I Dispatched for 1pm. Outcome: Created Case and Dispatched WO per FSE's Request IP: 88192637 Next Steps: None Integration : Confirmation of order 54174928 saved Integration : According to factory calendar 20210919 (Posting date) is not a workday (check entry) ----------------------------------------------------------------------------- Work Order #: WO-04363444 17-Sep-2021 Turley Rich The system meets the specification for the performed service and is returned to use Resolution : Replaced bad missing backup media Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Parts Consumed : 1.00000,459801128162,64GB USB BACKUP AND RESTORE ----------------------------------------------------------------------------- | 5 |  | 1.00000,459801128162,64GB USB BACKUP AND RESTORE | [12 empty] | 118346771 |  | 1332491 |  |  | Closed - Done |  | NotEngineering | [5 empty] | Engineering | [6 empty] | USB | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | USBExtended | Gantry-HW-USBExtended | Gantry-HW-USBExtended | NotEngineering |  | Gantry-HW | Gantry-HW-USBExtended | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | #N/A
1383 |  |  | China | 河南许昌市中心医院(传染病) | 40047 | 4445400% |  | 02N/Ingenuity CT 728326 | BU Q&R Feedback: No SAP Customer Symptom Code: BOOT Customer Symptom Code: Boot up/Power-up Problem Reason: Corrective Maintenance Engineer Failure Code: Software Defect or Virus Repair Activity Code: Instructed Customer Problem Reported by Customer : 02N/Ingenuity CT 728326 Safety Questions : Case: 0118336170 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Guoping Last updated by: Liu Guoping 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: cannot login Troubleshooting Action: need change password Repair Action: guide customer change password Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: N/A ----------------------------------------------------------------------------- | 3 |  | guide customer change password | [12 empty] | 118336170 |  | 1324704 |  |  | Closed - No Further Investigation | [7 empty] | Application | [6 empty] | Password | [5 empty] | Refresh |  |  | System issue | bar: No; No
1384 |  |  | China | Zhejiang Province Yueqing No. 1 People’s Hospital | 40054 | 4445400% |  | 02n/Ingenuity CT 728326/ECG wave transfer to host failed | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: Chip monitor is not connected to CT Safety Questions:Case: 0118336105 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Lili Last updated by: Lu Lili 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: ECG wave transfer to host failed Troubleshooting Action: this is intermittent issue, ECG monitor display wave is normal, but transfer to host failed. maybe cables connection problem. nemo ECG monitor to gantry's cable. Repair Action: nemo cable Test/Inspection: N/A Test Equipment: N/A T2 Activities:Problem Description: Summary of customer's problems: Other types of error. Detailed description of Customer's problems or malfunctions: ECG signal cannot be transferred to host It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: ECG signal cannot be transferred to host Remote troubleshoot and test result: ECG display is normal, host is frequently disconnected, it's possible that the connection of the middle wire causing signal transmission failure PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Replace ECG to rack connection cable to test 0 CBL, ECG Monitor trigger 4M Cable 455012301901 SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Next Step Action: PPD455012301901*1 have arranged for 18h delivery, please dispatch FSE: Huang Yi-86-15958725282 Next Step Action:huang yi External Remarks: 16.09.2021 13:03:41 UTC (86036249) Service details: ECG signal loss. Checked and tested, found cable from IBOX to HOST is faulty. Re plug in and adjust, then restore. service result: device returns to normal working condition after service Other significant matters: none Internal Remarks : 16.09.2021 13:03:44 UTC (86036249) RDI Accuracy:Yes $Q7: (Confirmed reported problem, ECG signal lost) $Q8: (check the connection found ibox disconnect with host) $Q9: (replug and insert the related cables) $Q10: (Test and inspection data refer to: PA. Device returned to full functionality.) $Q11: Calibrated tools name (N/A). S/N (N/A). Next calibration date (N/A); For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date:N/A ^Current Tube usage: second:N/A ----------------------------------------------------------------------------- | 3 |  | ECG signal loss. Checked and tested, found cable from IBOX to HOST is faulty. Re plug in and adjust, then restore. | [12 empty] | 118336105 |  | 1324651 |  |  | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Cable loose | [5 empty] | Refresh |  |  | System issue | bar: No; No
1385 |  |  | China | Dongguan Donghua Hospital Co., Ltd. | 40059 | 4445400% |  | Disk space was insufficient, replaced hard disk, reinstalled software | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: Out of memory Safety Questions:Case: 0118335062 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Zixuan Last updated by: Wang Zixuan 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer's problems: failure related to computer software. Detailed description of Customer's problems or malfunctions: Space is insufficient It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: NA Remote troubleshoot and test result: When scanning, it prompts that local space is insufficient. Checked the local and there is still 20% space left. The customer informed that it has also occurred before. PRS TR status: No PRS. Repair Action: Suggest onsite service procedures: Check and clean the C disk and D disk space, whether there are junk files; delete patient data and do a full rebuild; check the software. SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Next Step Action:LIN ZHITAO External Remarks : 17.09.2021 08:43:35 UTC (86034034) Service Details: The scan is intermittently interrupted. Prompted that there was a hard disk buffer space issue. Replaced the hard disk and reinstalled the software. Service Result: The device returned to working normally. Other significant matters: [none] Internal Remarks : 17.09.2021 08:43:38 UTC (86034034) RDI Accuracy:YES $Q7: (Confirmed reported problem, Out of memory ) $Q8: (out of disk memory when write data, the hard disk was not well) $Q9: (replace a new SSD) $Q10: (Test and inspection data refer to: PA. Device returned to full functionality) Q11: Calibrated Tools name ( NA ) S/N ( NA ) Next Calibration date ( NA ) For Mainland China CV / XR, CT engineers only. Please fill in: Current Tube installed date:NA Current Tube usage: NA ----------------------------------------------------------------------------- | 3 |  | The scan is intermittently interrupted. Prompted that there was a hard disk buffer space issue. Replaced the hard disk and reinstalled the software. | [12 empty] | 118335062 |  | 1324092 |  |  | Closed - Done |  | CTUDT00632261 | [5 empty] | Engineering | [6 empty] | Disk full | [5 empty] | Refresh |  |  | System issue | System | Console-SW | DiskFull | Console-SW-DiskFull | Console-SW-DiskFull | CTUDT00632261 | DiskFull | Console-SW | Console-SW-DiskFull | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1386 |  |  | China | 莱阳市人民医院 | 345005 | 4445400% |  | Blue screen and "reboot the device" error occurred during Brain scanning | Blue screen and "reboot the device" error occurred during Brain scanning |  |  | System is out of warranty, Link to "Blue screen" defect CTUDT00632553 | [14 empty] | 1324012 |  |  | Investigation |  | CTUDT00632553 | [5 empty] | Engineering | [6 empty] | Blue Screen | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Blue-Screen | Console-SW-Blue-Screen | Console-SW-Blue-Screen | CTUDT00632553 | Blue-Screen | Console-SW | Console-SW-Blue-Screen | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | #REF! | y | #N/A
1387 |  |  | United States | Unity Health Systems | 40037 | 4445400% |  | Monthly QA | Integration : Confirmation of order 54169063 saved ----------------------------------------------------------------------------- Work Order #: WO-04355570 16-Sep-2021 Taylor, Durrance The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: tech Device Use: scheduled activity Expected and Unexpected Results: Monthly QA User Impact: none Patient Impact: none Current Software Version: 4.1.10 Resolution : Monthly QA completed ----------------------------------------------------------------------------- | [15 empty] | 118340430 |  | 1326668 |  |  | Closed - Non-Complaint | [20 empty] | Refresh |  |  | System issue | bar: No; No
1388 |  |  | China | Shaanxi Province Ankang T.C.M. Hospital | 345041 | 4445300% |  | WeChat：02N/Ingenuity Core128/Elite China/reconstruct image | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: WeChat：02N/Ingenuity Core128/Elite China Safety Questions:Case: 0118330314 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: mo xueke Last updated by: mo xueke 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer's problems: For a certain patient, there is a problem with the reconstructed image It's a technology problem with the equipment itself: Yes. Troubleshooting Action: remote diagnosis and testing result: Suspect software, raw data problem PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Suggest engineer to check raw data and reconstruction card on the field.。。。 SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Next Step Action: zhang yifei Internal Remarks : 22.09.2021 03:37:56 UTC (86039086) $Q7: (Confirmed reported problem. & The system IRS have error problem ) $Q8: ( 1.Check IRS system. We need to restart system ) $Q9: ( 1. change system time). 2. Test scan for CT was OK. $Q10: (Test and inspection data refer to: PA. & Device returned to full functionality.) $Q11: Calibrated Tools name (N.A.) External Remarks: 22.09.2021 03:37:53 UTC (86039086) service reason: rebuild failed Service details: remote instruct customer to reboot CIRS, restore Sevice result: Device work normal. ----------------------------------------------------------------------------- | 3 |  | $Q7: (Confirmed reported problem. & The system IRS have error problem ) $Q8: (1.Check IRS system& We need to restart system ) $Q9: (1. change system time) . 2. Test scan for CT was OK. ) | [12 empty] | 118330314 |  | 1320767 |  |  | Investigation |  | CTUDT00632116 | [5 empty] | Engineering | [6 empty] | CIRS error | [5 empty] | Refresh |  |  | System issue | System | CIRS-SW | MPR Recon | CIRS-SW-MPR Recon | CIRS-SW-MPR Recon | CTUDT00632116 | MPR Recon | CIRS-SW | CIRS-SW-MPR Recon | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1389 |  |  | Chile | Scotiabank Chile | 40008 | 4445200% |  | CL1675-HUELE A QUEMADO | Integration:Confirmation of order 54153944 saved ----------------------------------------------------------------------------- Work Order #: WO-04351965 28-Sep-2021 [REDACTED] The system meets the specification for the performed service and is returned to use. Problem Description by Engineer : Problem description by engineer : Burning smell in examination room. Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected Results: No problems with burning smell. Patient Impact/Outcome: No impact. User Impact: No impact. Current Software Version: 4.1.10 Diagnostic performed by Engineer: Visual and functional check-ups are performed in the equipment. - Isolation transformer is opened to check its components. - Internal components (power supplies, transformers, fuses and electronic cards) are checked in the equipment. In both cases, no overheating signs or visually affected components are detected. Functional tests are performed with X-rays and image quality is assessed without problems. Equipment is left under observation, event has not occurred again and it has been working normally for over a week so there are doubts that the problem comes from the scanner. Resolution: Equipment is under monitoring, no damage is detected in components and the equipment works normally. ----------------------------------------------------------------------------- Work Order #: WO-04336506 [REDACTED] The system meets the specification for the performed service and is returned to use. Diagnostic performed by Engineer: Intermittent burning smell in the examination room. The room is checked with the cusstomer's help, but we are unable to identify the cause of the problem. It is asked to check the electrical board and air conditioning, but the customer reports that no problems were found. Case is sent to Planning to schedule a visit. Internal Comments : Error # and/or description of error : Burning smell in examination room. Repair Action: - Check system due to burning smell in examination room. Follow Up Required: Error Number and/or Description: Burning smell in examination room. Repair Action: - Check system due to burning smell in examination room. Technician's Name: [REDACTED] Problem Description by Engineer : Problem description by engineer : Burning smell in examination room. Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected Results: No problems with burning smell. Patient Impact/Outcome: No impact. User Impact: No impact. Current Software Version: 4.1.10 ----------------------------------------------------------------------------- | 2 |  | Burning smell in examination room, unable to identify the cause of the problem remotely, onsite visit is planned at 28 Sep (Formal PR962707 at April, After confirmed, it's not from CT, but from air condition) | [12 empty] | 118324751 |  | 1318359 |  |  | Investigation | [7 empty] | TBD | [6 empty] | Burning | [5 empty] | Refresh |  |  | System issue | bar: No; No
1390 |  |  | China | Xifeng County Traditional Chinese Medicine Hospital | 345014 | 4445200% |  | 05R/Ingenuity Core128/Elite China/cannot scan | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 05R/Ingenuity Core128/Elite China SAP Transfer:Dispatched for On site Support T2 Activities:Problem Description: Summary of customer's problems: Scanning failed. Detailed description of Customer's problems or malfunctions: Gantry acquisition system cannot comply。 It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: "event logger display Get DoseMODrefref, did not retrieve valid data " Remote troubleshoot and test result: Normal after restart PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Analyze reference detector data and confirm if it is the problem of converting or reference detector We recommend spare parts 12NC and names: BB Reference Detector FRU 459801702431 459801531261 DMS Ref. Converter FRU SAP Transfer:Case Transferred to SAP External Remarks: 27.09.2021 16:17:38 UTC (86039448) failure: scan failure. Failure analysis: Check event logger, display no data resend.Optical path testing is normal. Occasional failure. Failure handling: observe during use. Internal Remarks : 27.09.2021 16:17:41 UTC (86039448) $Q1: () $Q2: () $Q3: () $Q4: () "RDI Accuracy:YES" $Q7: (Confirmed reported problem.the scan failed. $Q8: (Troubleshooting,On site, Check evnt loger for no data resend. The optical path test is normal, and the occasional failure occurs. $Q9: (Repair steps. Using observation.12nc: ) $Q10: (Test and inspection data refer to: PA. Device returned to full functionality._ $Q11: Calibrated Tools name:N/A For Mainland China CV / XR, CT engineers only. Please fill in: tangguanglin Current Tube installed date: N/A Current Tube usage: scan second_N/A Safety Questions:Case: 0118321509 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Guoping Last updated by: Liu Guoping 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action: PPD 459801702431 ordered 459801702431 Sub no. 459800258961 available, FSE confirm when to ship. FSE: Tang Guanglin-86-18084189777 ----------------------------------------------------------------------------- | 3 |  | PPD 459801702431 已经订货 459801702431 替代号459800258961 有货 等FSE 确定何时发货 FSE:唐广林-86-18084189777 | [12 empty] | 118321509 |  | 1315760 |  |  | Closed - No Further Investigation |  | NotEngineering | [5 empty] | Engineering | [6 empty] | Ref. Converte | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Reference Converter | Gantry-HW-Reference Converter | Gantry-HW-Reference Converter | NotEngineering |  | Gantry-HW | Gantry-HW-Reference Converter | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | #N/A
1391 |  |  | Japan | Independent administrative institution National Hospital Organization | 341013 | 4445200% |  | 01DS:TM:撮影途中でCT-Hostがフリーズ>>9秒ほどしてから回復するが撮影タイミングが遅れてしまった>>エラーコード無し | Integration : Confirmation of order 54148695 saved ----------------------------------------------------------------------------- Work Order #: WO-04329894 13-Sep-2021 Yamamoto, Yuko The system meets the specification for the performed service and is returned to use. Problem Description by Engineer: Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Regulatory Questions : N/A Resolution: There may have been a delay due to a software malfunction. A new patch (SP1) has been released, so I installed it.(V4.1.10.xxx210 to V4.1.10.xxx220) Internal Comments: [CX related] Did you agree on how to complete the report with the person in charge of devices? <Select one of the following> Yes. I completed normal reporting that day. [Work report (paper) submission] Did you submit a work report on paper? <YES/ Q1. How did the examination finish? Select from the following: -Only the scan in question was retaken ・The customer resumed the examination from the point at which it had stopped. Q2. Were there any mistaken findings (misdiagnosis, mistaken treatment) due to the problem? None. Q3. Were there any artifacts? None. Q4. Was there any unintended unnecessary exposure? None. ----------------------------------------------------------------------------- Work Order #: WO-04329886 Miyakawa, Takeyuki The system meets the specification for the performed service and is returned to use. Internal Comments: [Subject] 01DS:TM: Mr. Yamamoto: The CT-Host froze during a scan>>It recovered after about 9 seconds, but the timing of the scan was delayed>>No error codes. [Internal Comments] Results of checking the log: There were no hardware errors on logger.mdb. There was no record of any events on the CT-Host or the CIRS side. There were no problems with the amount of disk usage, etc.Both the CT and CIRS restarted this morning. Dispatched: 1 RC FSE Registration Information: Name of FSE 1 who visited the site: Yutaka Yamamoto FSE 1 visit date and time: 2021/09/1316:30 FSE 1 hours worked: 2 Diagnostic performed by Engineer: $version = "v4.1.10.25210" 9:24 Call description: After doing Prep, the CT-Host froze and stopped working. After about 9 seconds, the operation started but the timing of the scan was delayed. The similar problem occurred before, and we have reinstalled the CT-Host software. Problem Description by Engineer: The CT-Host froze during a scan>>It recovered after about 9 seconds, but the timing of the scan was delayed>>No error codes. ----------------------------------------------------------------------------- | 4 |  | The CT-Host froze during a scan>>It recovered after about 9 seconds, The similar problem occurred before, A new patch (SP1) has been released, so I installed it.(V4.1.10.xxx210 to V4.1.10.xxx220) | [12 empty] | 118318777 |  | 1312698 |  |  | Investigation |  | CTUDT00632520 | [5 empty] | Engineering | [6 empty] | Yellow Screen | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1392 |  |  | United States | Lackey Memorial Hospital | 40015 | 4445200% |  | Q/O could not get ECG to work on CT | Integration : Confirmation of order 54153532 saved Integration : Confirmation of order 54153532 saved ----------------------------------------------------------------------------- Work Order #: WO-04336542 14-Sep-2021 Denison Chuck The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: C.T. tech. Device Use: Philips C.T. Expected and Unexpected Results: ECG. not working on host monitor. User Impact: C.T. still worked. Patient Impact: C.T. ECG. not showing. Current Software Version:4.1.10 Diagnostic performed by Engineer : Checked all connections for the ECG monitor. Not showing on the Host console. Found connection on i-box was loose. Tested the performance of the unit. Worked as should. ----------------------------------------------------------------------------- | 3 |  | Checked all connections for the ECG monitor. Not showing on the Host console. Found connection on i-box was loose. Tested the performance of the unit. Worked as should | [12 empty] | 118324775 |  | 1314504 |  |  | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Cable loose | [5 empty] | Refresh |  |  | System issue | bar: No; No
1393 |  |  | United States | Philips Medical Capital | 40053 | 4445200% |  | 4.1 Reformats on aixials are being cut off. | Integration : Confirmation of order 54152659 saved ----------------------------------------------------------------------------- Work Order #: WO-04334832 Giles Brian The system meets the specification for the performed service and is returned to use Problem description by Clinical Support Specialist : Customer Function/Role: Tech Device Use: In Use Expected and Unexpected Results: Cut off reformats User Impact: None Patient Impact: None Current Software Version: 4.1 Diagnostic performed by Engineer : Brought David Mujica (RAS) onto the call call to help her out ----------------------------------------------------------------------------- Work Order #: WO-04334814 Mujica, David The system meets the specification for the performed service and is returned to use Problem description by Clinical Support Specialist : Customer Function/Role: Technologist Device Use: In use Expected and Unexpected Results: User Impact: none Patient Impact: none Current Software Version: 4.1 Diagnostic performed by Engineer : Alyssa was aided to merge images in CT viewer to create coronal and sagittal images on a CT head exam. Merge was explored in CT viewer. All images were reconstructed and saved successfully. No further clinical support needed. No rescan, no loss of data, no reinjection, per Alyssa. ----------------------------------------------------------------------------- | 5 |  | Alyssa was aided to merge images in CT viewer to create coronal and sagittal images on a CT head exam. Merge was explored in CT viewer. All images were reconstructed and saved successfully. | [12 empty] | 118323508 |  | 1313180 |  |  | Closed - No Further Investigation | [7 empty] | Application | [6 empty] | Merge image | [5 empty] | Refresh |  |  | System issue | bar: No; No
1394 |  |  | China | Zhalantun City People's Hospital | 40045 | 4445100% |  | wechat:02N/Ingenuity CT 728326/Scan, image transmission and reconstruction failure | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: Scan, image transmission and reconstruction failure Safety Questions:Case: 0118318000 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer's problems: Other types of error. Detailed description of Customer's problems or malfunctions: Continuous scan can cause overheating. Image transmission can crash. Reconstructed images are not recognized in the clinic It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: Continuous scan can cause overheating. Image transmission can crash. Reconstructed images are not recognized in the clinic Remote troubleshoot and test result: Scan protocol setup need to be adjusted PRS TR status: it's not through, ask engineer for maintenance on site. Repair Action: Suggest onsite service procedures: Adjust scan protocol. Communicated with trainer. Setup again and check image transmission issue SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support External Remarks: 15.09.2021 04:33:31 UTC (86038219) service details: scan software freezes, restore host software. service result: device returns to normal working condition after repairing Internal Remarks : 15.09.2021 04:33:34 UTC (86038219) RDI Accuracy（Yes）Q7: (Confirmation & Description of problem. Host software error. )$Q8: (On site check Host software failed, Need to install Host software.) $Q9: (Install Host software.) $Q10: (Test and inspection data refer to: PA. Device returned to full functionality.) $Q11: Calibrated tools name (FLUKE287) S/N (12470209) Next calibration date (06/30/2021) For Mainland China CV / XR, CT engineers only. Please fill in: Current Tube installed date Tube installed date: ##5/ 19/ 2021 ^Current Tube usage: scan second_##35126 ----------------------------------------------------------------------------- | 3 |  | Q7: (Confirmation & Description of problem. Host software error. )$Q8: (On site check Host software failed, Need to install Host software.) $Q9: (Install Host software.) $Q10: (Test and inspection data refer to: PA. Device returned to full functionality.) | [12 empty] | 118318000 |  | 1312034 |  |  | Investigation |  | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Recon failure | [5 empty] | Refresh |  |  | System issue | System | CIRS-SW | NoEnoughInfo | CIRS-SW-NoEnoughInfo | CIRS-SW-NoEnoughInfo | NoEnoughInfo |  | CIRS-SW | CIRS-SW-NoEnoughInfo | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | #N/A
1395 |  |  | China | Zhalantun City People's Hospital | 40045 | 4445100% |  | 02N/Ingenuity CT 728326/account locked | BU Q&R Feedback: No SAP Customer Symptom Code: MARF Customer Symptom Code: Monitor Alarm Failure Reason: Corrective Maintenance Engineer Failure Code: Incompatible Use Repair Activity Code: Instructed Customer Problem Reported by Customer: 02N/Ingenuity CT 728326 Safety Questions:Case: 0118317948 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: CT account locked Troubleshooting Action: wait half hour , then reuse CT account , input password, system is ok. Repair Action: wait half hour , then reuse CT account , input password, system is ok. Test/Inspection: N/A Test Equipment: N/A SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | wait half hour , then reuse CT account , input password, system is ok. | [12 empty] | 118317948 |  | 1311866 |  |  | Closed - No Further Investigation | [7 empty] | Application | [6 empty] | Password | [5 empty] | Refresh |  |  | System issue | bar: No; No
1396 |  |  | China | Caofeidian District Hospital | 345016 | 4444900% |  | Wechat:02N/Ingenuity Core128/Elite China/Estop open | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer:Estop open Safety Questions:Case: 0118308389 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Bai Jingwen Last updated by: Bai Jingwen 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Case Transferred to SAP T2 Activities:Problem Description: Summary of customer's problems: ESTOP disconnected. Detailed description of Customer's problems or malfunctions: The gantry loses power continuously. It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: No PRS. Remote troubleshoot and test result: Restart but didn't work. PRS TR status: No PRS. Repair Action: Suggest onsite service procedures: Suggest troubleshooting based on log on the site. We recommend spare parts 12NC and names: N/A SAP Transfer:Dispatched for On site Support External Remarks: 15.09.2021 02:40:08 UTC (86039063) service details: replace R2D Service Result: The device returned to normal. Internal Remarks : 15.09.2021 02:40:11 UTC (86039063) "RDI Accuracy:Yes" $Q7: (Confirmation & Description of problem) Confirmed reported problem, sometimes gantry power down $Q8: (Troubleshooting & Results) Check the error log, need to replace R2D $Q9: (Repair steps) Replace the R2D. $Q10: (Test and Inspection data & Device status) Test and inspection data refer to: PA. Device returned to full functionality. $Q11: Calibrated Tools name to PA ----------------------------------------------------------------------------- | 3 |  | Confirmed reported problem, sometimes gantry power down $Q8: (Troubleshooting & Results) Check the error log, need to replace R2D $Q9: (Repair steps) Replace the R2D. | [12 empty] | 118308389 |  | 1309146 |  |  | Closed - No Further Investigation |  | CTUDT00630864 | [5 empty] | Engineering | [6 empty] | R2D | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | R2D | Gantry-HW-R2D | Gantry-HW-R2D | CTUDT00630864 | R2D | Gantry-HW | Gantry-HW-R2D | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1397 |  |  | China | 黑龙江省桦川县人民医院 | 345018 | 4444900% |  | 02N/Ingenuity Core128/Elite China/Mouse and keyboard faults | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Replace Part Problem Reported by Customer : Mouse and keyboard faults Safety Questions : Case: 0118307510 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Gaoge Last updated by: Wang Gaoge 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 键盘鼠标泡水了，坏了 Troubleshooting Action: nemo键盘鼠标给客户，客户自行更换 Repair Action: nemo 459800918431 Cherry USB keyboard English Silver 459800918441 Microsoft Optical Mouse 200 Test/Inspection: N/A Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | nemo键盘鼠标给客户，客户自行更换 | [12 empty] | 118307510 |  | 1308883 |  |  | NAC | [7 empty] | NEMO | [6 empty] | NEMO | [5 empty] | Refresh |  |  | System issue | bar: No; No
1398 |  |  | United States | Unity Health Systems | 40037 | 4444900% |  | Errors during study with large patient; Inspected system and confirmed that system's operating normally without issues. | Integration : Confirmation of order 54146605 saved ----------------------------------------------------------------------------- Work Order #: WO-04324718 16-Sep-2021 Taylor, Durrance The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: Diagnostic Expected and Unexpected Results: Errors during study with large patient User Impact: No Patient Impact: No Current Software Version: 4.1.10 Resolution : This all happened 9/10/21 00:39- 00:49 hours involving the patient support. Technician reported the patient 520 lbs rubbing against the sides of the gantry during the scan. Technician stated, "The AP scout was completed and there was an error after and the lateral scout was not completed. They moved the table in and zeroed again and the same thing happened with the second attempt". Customer also stated that after the patient was taken to another hospital that, "They completed the QA and IQ check after they removed the patient from the room. Everything went fine with that and has been working consistently since". Error logs showed that the table was not reaching the requested position in time. While I was doing the patient support PM I inspected the tables horizontal resistance and no issue was found there. I did perform a horizontal calibration as well. They also reported they could not run a bug report. I had them shut down the system and try again and they still could not run one. During the PM the system was blue button'd and all computers powered down. After power down the scanner was able to perform a bug report. This event occurred this one time and has not presented itself since 9/10. ----------------------------------------------------------------------------- Work Order #: WO-04324574 Maier, Donnie The system meets the specification for the performed service and is returned to use Follow Up Required : Customer explained they had some errors on 9/10 just after midnight during a study with a heavy patient. System is functioning properly now. She has already talked to PFSE Taylor Durrance and asked to have a case created and dispatched for 9/15 @1:00PM local time as he was already going to be onsite. CHA only had 3 cases with this particular soft limit error. States to calibrate couch horizontal and if continues to replace limit switches. Customer also stated that this was a heavy patient and only occurred this one time so may be a result of heavy patient. May calibrate and try to reproduce issue. Diagnostic performed by Engineer : PFSE Taylor Durrance is aware and asked for customer to call in and create a case. Also would like it scheduled for 9/15 1:00PM according to customer. System doesn't appear to be in radar yet. Cu states that they received an error just after midnight this morning 9/10 on this system during a study with a larger patient. I will attempt to pull logs and add case notes if possible. Dispatching as requested. I pulled logs/analyzed and seen S_HMC_AXIS_WRONG_DEST_REACHED normal error @00:41:36 and then S_MC_DEVICE_IN_SOFT_LIMIT_ON Normal Error followed by S_MANMOT_HORIZONTAL_START_ERROR. Only 3 cases in CHA and they elude to performing Couch Horizontal Calibration first and then proceed to checking/replacing limit switches. However cu indicated this was a large patient so I would run thru couch calibration and try to reproduce error as it only occurred once according to cu. I cannot verify however as system doesn't appear to be in Radar yet. Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: Diagnostic Expected and Unexpected Results: Errors during study with large patient User Impact: No Patient Impact: No Current Software Version: 4.1.10 ----------------------------------------------------------------------------- | 3 |  | After power down the scanner was able to perform a bug report. This event occurred this one time and has not presented itself since 9/10. | [12 empty] | 118315169 |  | 1308351 |  |  | Closed - Done | [7 empty] | Customer | [6 empty] | Over load | [5 empty] | Refresh |  |  | System issue | bar: No; No
1399 |  |  | Poland | University Clinical Hospital in | 341001 | 4444800% |  | UPS do CT | Integration: Confirmation of order 54140370 saved. ----------------------------------------------------------------------------- Work Order #: WO-04312462 22-Sep-2021 Marcin Knyps The system meets the specification for the performed service and is returned to use Problem Description by Engineer: Parts were collected for diagnostics. Resolution: Collection and writing of spare parts. Correct operation of the device. Parts Consumed : 1.00000,459801652561,HP Z Turbo Drive 2TB TLC Z8 SSD Kit 1.00000,459801652551,HP Z Turbo Drive 256GB TLC Z8 SSD Module 2.00000,459801583761,8GB DDR4-2666 RegRAM 1.00000,459800306392,EXTERNAL BATT PACK, 3KVA UPS, 72VDC 1.00000,459800306382,UPS,NA,230V 50/60HZ,3KVA ----------------------------------------------------------------------------- | 5 |  | 1.00000,459801652561,HP Z Turbo Drive 2TB TLC Z8 SSD Kit 1.00000,459801652551,HP Z Turbo Drive 256GB TLC Z8 SSD Module 2.00000,459801583761,8GB DDR4-2666 RegRAM 1.00000,459800306392,EXTERNAL BATT PACK, 3KVA UPS, 72VDC 1.00000,459800306382,UPS,NA,230V 50/60HZ,3KVA | [12 empty] | 118305028 |  | 1307285 |  |  | Closed - No Further Investigation | [7 empty] | TBD | [6 empty] | TBD | [5 empty] | Refresh |  |  | System issue | bar: No; No
1400 |  |  | United States | Odessa Regional Medical Center | 40029 | 4444800% |  | True Coronal/Sagittal on Scoliosis Patient | Work Order #: WO-04313269 Mujica, David The system meets the specification for the performed service and is returned to use Problem description by Clinical Support Specialist : Customer Function/Role: Technologist Device Use: In use Expected and Unexpected Results: Scoliosis patient reconstruction User Impact: none Patient Impact: none Current Software Version: 4.1 Diagnostic performed by Engineer : Maya indicated that she may need help straightening a CT scan in post processing of a scoliosis patient. It was confirmed that this patient was already diagnosed with a scoliosis condition, therefore the orthopod/radiologist would wish to see the condition of the scoliosis and therefore straightening the spine would defeat the conditional diagnosis. Maya will inquire from the radiologist if additional straightening views are needed. Maya said she could construct true coronal and sagittal by herself with the data given. Maya had no further questions or need for clinical support. No rescan, no loss of data, no reinjection per Maya. ----------------------------------------------------------------------------- | [15 empty] | 118305606 |  | 1303468 |  |  | Closed - Non-Complaint | [20 empty] | Refresh |  |  | System issue | bar: No; No
1401 |  |  | Japan | Maki Hospital | 341017 | 4444800% |  | 01CS:RS: Inquiry about changing a patient name. >> I explained the "Change patient details" function. | Work Order #: WO-04309241 Satake, Rui The system meets the specification for the performed service and is returned to use. Problem Description by Engineer : Inquiry about changing a patient name. Resolution : I explained the "Change patient details" function. Internal Comments: [Subject] 01CS:RS:Mr. Togawa: Inquiry about changing a patient name. >> I explained the "Change patient details" function. [Internal Comments] None [4/8 Elements] Q5. How did the examination end? None. Q6. Did any mistaken findings (mistaken diagnosis or treatment) result from the problem? No. Q7. Was there any unintended unnecessary exposure? No Q8 Artifacts? None ----------------------------------------------------------------------------- | [15 empty] | 118300918 |  | 1300880 |  |  | Closed - Non-Complaint | [20 empty] | Refresh |  |  | System issue | bar: No; No
1402 |  |  | China | 山东安丘市中医医院 | 345017 | 4444700% |  | 02N/Ingenuity Core128/Elite China/gantry noise | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity Core128/Elite China SAP Transfer : Case Transferred to SAP Safety Questions : Case: 0118296659 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Bo Last updated by: Wang Bo 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 设备异响。 客户问题详细描述或故障现象： 机架旋转时异响。 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： log无报错。 远程诊断与测试结果： 重启无效。已和工程师确认rotor皮带裂损。 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议检查异响来源，必要时更换皮带。 建议备件12NC及名称： PPD 453566492401 SAP Transfer : Dispatched for Onsite Support Next Step Action : PPD 453566492401*1 24H FSE:杨兴波-86-13583235930 已安排出货 External Remarks : 09.09.2021 08:02:48 UTC Xingbo Yang (86038250)服务内容：旋转噪声，更换Rotor驱动皮带。 服务结果：维修后，设备恢复正常工作。 Internal Remarks : 09.09.2021 08:02:50 UTC Xingbo Yang (86038250) "RDI Accuracy:yes" $Q7: (Confirmed reported problem & make noise when the rotor totation) $Q8: ( check rotor drive belt & found belt broken ) $Q9: ( replace poly_belt 453566492401 ) $Q10: (Test and inspection data refer to: PA & Device returned to full functionality) $Q11: Calibrated tools name(NA) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date: 9/ 30/ 2020 ^Current Tube usage: scan second_250394 ----------------------------------------------------------------------------- | 3 |  | $Q7: (Confirmed reported problem & make noise when the rotor totation) $Q8: ( check rotor drive belt & found belt broken ) $Q9: ( replace poly_belt 453566492401 ) | [12 empty] | 118296659 |  | 1301514 |  |  | Closed - No Further Investigation |  | ECR-077012 | [5 empty] | Engineering | [6 empty] | Belt | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Belt | Gantry-HW-Belt | Gantry-HW-Belt | ECR-077012 | Belt | Gantry-HW | Gantry-HW-Belt | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1403 |  |  | China | Shaanxi Province Ankang T.C.M. Hospital | 345041 | 4444700% |  | WeChat：02N/Ingenuity Core128/Elite China | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: WeChat: 02N/Ingenuity Core128 Safety Questions:Case: 0118296354 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: mo xueke Last updated by: mo xueke 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Case Transferred to SAP T2 Activities:Problem Description: Summary of customer's problems: system is running slow or crashed. Detailed description of Customer's problems or malfunctions: The computer host system crashing. It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: The log has no errors. Remote troubleshoot and test result: Restart but didn't work. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Suggest to add memory and observe usage according to TSE suggestion. We recommend spare parts 12NC and names: PPD 459801633881 Internal Remarks : 26.09.2021 02:25:58 UTC (86039086) $Q7: (Confirmed reported problem. & The system software have error ) $Q8: ( 1.Check system software. We need to replace RAM) $Q9: ( 1. replace RAM). 2. Test scan for CT was OK. $Q10: (Test and inspection data refer to: PA. The device returned to full functionality. $Q11: Calibrated Tools name (N.A.) SAP Transfer:Dispatched for On site Support External Remarks: 26.09.2021 02:25:56 UTC (86039086) service reason: the system is not working smoothly Service details: add 64g ram Sevice result: Device work normal. ----------------------------------------------------------------------------- | 5 |  | 服务原因：系统运行卡顿 服务内容：加装64g内存条 服务结果：设备工作正常 | [12 empty] | 118296354 |  | 1301293 |  |  | Closed - No Further Investigation |  | AddRAM | [5 empty] | Engineering | [6 empty] | System slow disk full | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Performance | Console-SW-Performance | Console-SW-Performance | AddRAM | Performance | Console-SW | Console-SW-Performance | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | System slow | #N/A
1404 |  |  | China | Medical center (contagious disease, He nan, Xu chang) | 40047 | 4444700% |  | 02n/Ingenuity CT 728326/images transfer failed | BU Q&R Feedback: No SAP Customer Symptom Code: DIDM Customer Symptom Code: Patient or Image Data Management/Post Processing Data Reason: Corrective Maintenance Engineer Failure Code: Software Defect or Virus Repair Activity Code: Install Update Problem Reported by Customer: 02n/Ingenuity CT 728326 SAP Transfer:Case Transferred to SAP Next Step Action: 15637449981 repeat call in, transfer yang Safety Questions:Case: 0118296297 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: images transfer from CT to ISP failed Troubleshooting Action: customer call service, images transfer from CT to ISP failed. FSE has perform FCO72800740, this issue not solved. FSE call back to customer, will perform FCO72800758. close this case. Repair Action: FSE will perform FCO72800758,. Test/Inspection: N/A Test Equipment: N/A ----------------------------------------------------------------------------- | 3 |  | FSE will perform FCO72800758 to resolve image transfer issue | [12 empty] | 118296297 |  | 1301237 |  |  | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS transfer | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1405 |  |  | Poland | University Clinical Hospital in | 341001 | 4444700% |  | Assistance during installation of Stellant CT DUAL injector. | Integration: Confirmation of order 54133709 saved. ----------------------------------------------------------------------------- Work Order #: WO-04303316 26-Sep-2021 Gerard Dabrowka The system meets the specification for the performed service and is returned to use Problem Description by Engineer: integration of the contrast injector. Diagnostic performed by Engineer: integration of the contrast injector Stellant Bayer. ----------------------------------------------------------------------------- | [15 empty] | 118297498 |  | 1302123 |  |  | Closed - Non-Complaint | [20 empty] | Refresh |  |  | System issue | bar: No; No
1406 |  |  | Germany | Gesundheit Nord gGmbH Klinikverbund | 40002 | 4444600% |  | Probleme bei der Konstanzprüfung | Internal Remarks: 459801415161 Customer comment: # Arranged appointment with Ms. Schichtl for 9/16, 2 PM # Integration:Confirmation of order 54127560 saved ----------------------------------------------------------------------------- Work Order #: WO-04294491 16-Sep-2021 Halama, Thomas The system meets the specification for the performed service and is returned to use Problem Description by Engineer: System Phantom Head defective. Parts Consumed: 1.00000,459801415161,HARMONIZED PHANTOM HEAD SECTION Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution: Performed constancy test. In the tests performed, the system corresponds to the manufacturer's specifications and is ready for clinical use. ----------------------------------------------------------------------------- | 3 |  | System Phantom Head defective. Parts Consumed: 1.00000,459801415161,HARMONIZED PHANTOM HEAD SECTION | [12 empty] | 118289659 |  | 1297009 |  |  | Evaluation |  | NotEngineering-PartReplace | [5 empty] | Engineering | [6 empty] | Head Phantom | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | Phantom | Couch-HW-Phantom | Couch-HW-Phantom | NotEngineering-PartReplace |  | Couch-HW | Couch-HW-Phantom | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | #N/A
1407 |  |  | China | Medical center (contagious disease, He nan, Xu chang) | 40047 | 4444600% |  | 02N/Ingenuity CT 728326/change dicom setting | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 02N/Ingenuity CT 728326 Safety Questions:Case: 0118289121 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Dispatched for On site Support SAP Transfer:Case Transferred to SAP External Remarks: 07.09.2021 10:32:18 UTC Yong Huang (86038368) Service details: [customer request to add dicom node] Service result: [The device returns to normal working condition after maintenance.] Other important issues: [None] Internal Remarks: 07.09.2021 10:32:20 UTC Yong Huang (86038368) RDI Accuracy: Yes $Q7: (Confirm reported problem, change dicom set ) $Q8: (Performed visual inspection, need to change dicom set) $Q9: (change dicom set ) $Q10: (Test and inspection data refer to: PA. Device returned to full functionality.) $Q11: Calibrated Tools (N/A) T2 Activities:Problem Description: Summary of customer's problems: Other types of error. Detailed description of Customer's problems or malfunctions: Modify DICOM setting It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: Modify DICOM setting Remote troubleshoot and test result: FSE modify dicom setting onsite PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Modify DICOM setting ----------------------------------------------------------------------------- | [15 empty] | 118289121 |  | 1296651 |  |  | Closed - Non-Complaint | [20 empty] | Refresh |  |  | System issue | bar: No; No
1408 |  |  | United States | Clay County Medical Center | 40003 | 4444600% |  | Front Left Control Panel Up Button is intermittent. | Next Step Action : left panel doesn't work right , FSE approved to order. ----------------------------------------------------------------------------- Work Order #: WO-04297146 10-Sep-2021 Stauffer Jeffrey The system meets the specification for the performed service and is returned to use Resolution : Replaced front left control panel. Diagnostic performed by Engineer : Looked at front left control panel and saw that it was damaged. Pressed up button and it worked but had to be pressed hard to get it to work. Problem Description by Engineer : Customer Function/Role: Philips FSE Device Use: Diagnostic CT exams Expected and Unexpected Results: Expected - up button to be easy to press. Unexpected - up button on left front control panel is damaged and hard to press. User Impact: None Patient Impact: None Current Software Version: 4.1.10 Parts Consumed : 1.00000,459801484991,CONTROL PANEL ASSEMBLY-FRONT LEFT Tools & Equipment : MULTIMETER,2023-06-30 00:00:00 ----------------------------------------------------------------------------- | 3 |  | Replaced front left control panel. 1.00000,459801484991,CONTROL PANEL ASSEMBLY-FRONT LEFT | [12 empty] | 118292376 |  | 1299156 |  |  | Closed - No Further Investigation |  | tbd | [5 empty] | Engineering | [6 empty] | Control panel | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Control Panel | Gantry-HW-Control Panel | Gantry-HW-Control Panel | Not a Defect | Control Panel | Gantry-HW | Gantry-HW-Control Panel | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1409 |  |  | Chile | BANCO DE CHILE | 341005 | 4444500% |  | CL1688 - TIME SET IS REQUIRED. | Integration:Confirmation of order 54123296 saved ----------------------------------------------------------------------------- Work Order #: WO-04288476 Valenzuela Velázquez, Eduardo Ariel The system meets the specification for the performed service and is returned to use. Problem Description by Engineer : Problem description by engineer : Adjust time. Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected Results: The time is set. Patient Impact/Outcome: No impact. User Impact: No impact. Current Software Version: Not provided. Diagnostic performed by Engineer: Due to a new time in the country, a time set is requested. Internal Comments : Error # and/or description of error : Adjust time. Repair Action: - The time is updated. Resolution: The equipment's time is updated. ----------------------------------------------------------------------------- | 4 |  | Customer request setting | [12 empty] | 118285065 |  | 1290972 |  |  | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Time setting | [5 empty] | Refresh |  |  | System issue | bar: No; No
1410 |  |  | China | Henan Province Puyang City People’s Hospital | 40031 | 4444500% |  | 02N/Ingenuity CT 728326/software hang up | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 02N/Ingenuity CT 728326 Safety Questions:Case: 0118284008 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Dispatched for On site Support SAP Transfer:Case Transferred to SAP External Remarks : 07.09.2021 06:48:56 UTC Yongping Luo (86038950) Service details: Software crash, slow response. Service result: Reset the software, after cleaned the disk, the equipment returned to normal. Continue to monitor. Other significant matters: none Internal Remarks : 07.09.2021 14:29:08 UTC Yongping Luo (86038950) $Q7:Confirmed report problem: the APP no response. Q8:Perform visual check on site, find the software running slowly, reset the software and clean the disk device return to full function. $Q9:reset the software and clean the disk, Device return to full function. $Q10: test and inspection data refer to PA. Device returned to full functionality. $Q11:Calibrated Tools refer to PA. Next Step Action: Luo Yongping 015226131133 call in and disp him Informed that the customer got drunk, and could not report for repair T2 Activities:Problem Description: Summary of customer's problems: Other types of error. Detailed description of Customer's problems or malfunctions: The software is crashing. It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: The software is crashing. Remote troubleshoot and test result: FSE reports for repair on behalf of the customer that software crash, yellow screen PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Re-install the system, and conduct FCO ----------------------------------------------------------------------------- | 3 |  | $Q7:Confirmed report problem: the APP no response. Q8:Perform visual check on site, find the software running slowly, reset the software and clean the disk device return to full function. FCO72800758 performed | [12 empty] | 118284008 |  | 1292924 |  |  | Investigation |  | CTUDT00632520 | [5 empty] | Engineering | [6 empty] | SP1 FCO | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1411 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 337000 | 4444500% |  | spill guard order | Integration: Entry 1 in table OPERATION (method CHANGE) was not found. - Error occurred during the processing of the BAPI methods. Integration: Confirmation of order 54122643 has been saved. ----------------------------------------------------------------------------- Work Order #: WO-04288143 Parts Consumed : 1.00000,453566136722,APRON PACKING-SPILL GUARD ----------------------------------------------------------------------------- Work Order #: WO-04287967 Can Ozgenlik The system meets the specification for the performed service and is returned to use Resolution : The spill guard which is laid on the gantry cover to prevent opaque material leak was ordered. Device is operational. Problem Description by Engineer : Spill guard order. Follow Up Required : Spill guard order. ----------------------------------------------------------------------------- | 3 |  | 1.00000,453566136722,APRON PACKING-SPILL GUARD @William, need check what is the kit include, and the propose of the kit. | [12 empty] | 118284526 |  | 1290640 |  |  | Evaluation | [7 empty] | Others | [6 empty] | NEMO | [5 empty] | Refresh |  |  | System issue | bar: No; No
1412 |  |  | Portugal | Centro Hospitalar de Vila Nova | 40030 | 4444500% |  | Repair of the couch top of the Emergency CT that is broken | [16 empty] | 118286513 |  | 1294456 |  |  | Closed - Duplicate | [20 empty] | Refresh |  |  | System issue | bar: No; No
1413 |  |  | China | Juancheng Ma Aiyun WCH | 345044 | 4444500% |  | 02N/Ingenuity Core128/Elite China/config exam data | BU Q&R Feedback: No SAP Customer Symptom Code: MARF Customer Symptom Code: Monitor Alarm Failure Reason: Corrective Maintenance Engineer Failure Code: Incompatible Use Repair Activity Code: Instructed Customer Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ Safety Questions:Case: 0118284164 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Jianqiang Last updated by: Wang Jianqiang 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: consulting the exam data how to config Troubleshooting Action: information customer can config in preference Repair Action: instruct to customer Test/Inspection: Functional Check PASSED. The device returned to full functionality. Test Equipment: NA SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | [15 empty] | 118284164 |  | 1293042 |  |  | Closed - Non-Complaint | [20 empty] | Refresh |  |  | System issue | bar: No; No
1414 |  |  | United States | Philips Medical Capital | 40028 | 4444500% |  | 0599CDN83 EStop Opens when couch motion fails | Next Step Action : Cust to call back if Holiday service wanted. If they do call back then either a PO# will be provided, of the name of the HCF Mgr approving billable holiday OT. Also agreed if they did not call back by 5pm, that case would be dispatched HTM Tuesday AM. Customer has not called back, but Mike Blair says to use Jerry Prendergast ' as PO# Integration : Confirmation of order 54123438 saved ----------------------------------------------------------------------------- Work Order #: WO-04290212 06-Sep-2021 Plewa Michael The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: CTTech Device Use: Diag Expected and Unexpected Results: Estop opens when couch motion fails User Impact: None Patient Impact: None Current Software Version:IDK Diagnostic performed by Engineer : Reset gantry. Checked error logs and estop circuit. Ran test scans . Tested OK. ----------------------------------------------------------------------------- Work Order #: WO-04288565 Carpenter, Kevin The system does not meet the specification for the performed service and is not returned to use Problem Description by Engineer : Customer Function/Role: CTTech Device Use: Diag Expected and Unexpected Results: Estop opens when couch motion fails User Impact: None Patient Impact: None Current Software Version:IDK Diagnostic performed by Engineer : Site PRS / RSN not working ~ possible misconfig - CT RSE unable to get logs. Cust to call back if they want billable onsite service ~ if they don't call by 5PM, then they've agree RSE is to dispatch for HTM Tuesday 9/7 Diagnostic performed by Engineer : Cust never called back for onsite service ~ but Mike Blair says to use "Jerry Prendergast" as PO# and to dispatch. Follow Up Required : PRS / RSN to site is not working properly. Cust was to call back if Holiday service was wanted ~ CT RSE did not receive call from site. FSE please review logs onsite as customer reports EStop opens each time couch motion is attemoted. ----------------------------------------------------------------------------- | 2 |  | Reset gantry. Checked error logs and estop circuit. Ran test scans . Tested OK. | [12 empty] | 118284998 |  | 1290809 |  |  | Investigation |  | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Gantry reboot | [5 empty] | Refresh |  |  | System issue | System | TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | NoEnoughInfo |  | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | #N/A
1415 |  |  | United States | Odessa Regional Medical Center | 40029 | 4444400% |  | Spiral artifact, unknown for customer cancelled onsite service | Integration : Confirmation of order 54119947 saved ----------------------------------------------------------------------------- Work Order #: WO-04281297 Carpenter, Kevin The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Diagnostic performed by Engineer : CT RSE advised customer to LOCK patients scan which have this artifact ( she says there are 3 cases ) ~ She says she has already wiped the bore to clear any specks of contrast and has then run Air Cals but the artifact persists. Site only has Remote Console on PRS / PRS - CT RSE unable to pull logs, images nor cal data. Problem Description by Engineer : Customer Function/Role: CTTech Device Use: Diag Expected and Unexpected Results: Spiral Artifact User Impact: None Patient Impact: None Current Software Version: 4 Follow Up Required : Cust reports Spiral Artifact on Helical scans ~ even after wipe bore and running air cals. Service is requested Tuesday AM. FSE run Bad Dets and inspect inner DMS cover to contrast, etc. ----------------------------------------------------------------------------- | 3 |  | cx stated no servie is needed issue is resolved. | [12 empty] | 118277997 |  | 1289541 |  |  | Closed - Done | [7 empty] | Others | [6 empty] | Artifact | [5 empty] | Refresh |  |  | System issue | bar: No; No
1416 |  |  | China | 江西省儿童医院 | 40041 | 4444300% |  | 02N/Ingenuity CT/Gantry error | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: Instructed Customer Problem Reported by Customer : 02N/Ingenuity CT/ Next Step Action : CT Application 刘俊强 15314108698 来电 Safety Questions : Case: 0118276017 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wennian Han Last updated by: Wennian Han 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action : 应用老师在现场培训时，演示重启过程，后开机发现机器报错。建议应用老师关机长些时间，再重启，后反馈机器扫描正常。 Next Step Action : 也通知了FSE相关情况 T2 Activities : Problem Description: Gantry error Troubleshooting Action: Application specialist onsite training, demonstrated the reboot process to customer, and then power on and found that can not scan. Repair Action: Advise application specialist power down for a longer time, then reboot, and then feedback that the scan normally. Inform FSE related situation. Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: NA SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | Advise application specialist power down for a longer time, then reboot, and then feedback that the scan normally. Inform FSE related situation. | [12 empty] | 118276017 |  | 1289199 |  |  | Closed - No Further Investigation | [7 empty] | Application | [6 empty] | Training | [5 empty] | Refresh |  |  | System issue | bar: No; No
1417 |  |  | China | 河南省新乡市凤泉区大块镇卫生院 | 345009 | 4444300% |  | 02N/Core128/need reset password | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Software Defect or Virus Repair Activity Code: Configuration Change Problem Reported by Customer : 02N/Core128/ Safety Questions : Case: 0118276372 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Guoping Last updated by: Liu Guoping 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: need reset password Troubleshooting Action: need reset password Repair Action: guide customer reset password Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | [15 empty] | 118276372 |  | 1289297 |  |  | Closed - Non-Complaint | [20 empty] | Refresh |  |  | System issue | bar: No; No
1418 |  |  | China | 鄄城马爱云妇儿医院 | 345044 | 4444200% |  | 02N/Ingenuity Core128/Elite China/add printer | BU Q&R Feedback: No SAP Customer Symptom Code: MARF Customer Symptom Code: Monitor Alarm Failure Reason: Corrective Maintenance Engineer Failure Code: Incompatible Use Repair Activity Code: Instructed Customer Problem Reported by Customer : 02N/Ingenuity Core128/Elite China SAP Transfer : Case Transferred to SAP Safety Questions : Case: 0118270896 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Jianqiang Last updated by: Wang Jianqiang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: add printer Troubleshooting Action: information customer can add the printer under olevel modl Repair Action: instruct to customer add the printer by themselvs Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: NA ----------------------------------------------------------------------------- | [15 empty] | 118270896 |  | 1287404 |  |  | Closed - Non-Complaint | [20 empty] | Refresh |  |  | System issue | bar: No; No
1419 |  |  | United States | Leesburg Regional Medical Center | 40022 | 4444200% |  | Screen froze; Rebooted host and CIRS to resolve. | Integration : Confirmation of order 54117367 saved Integration : According to factory calendar 20210905 (Posting date) is not a workday (check entry) ----------------------------------------------------------------------------- Work Order #: WO-04276916 Morrow, Dan The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Problem description by engineer : Screen froze and customer rebooted system. Raw data missing. When was the first occurrence? Within... : 03-Sep-2021 How often is issue occurring? : First occurrence Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Missing raw data Patient Impact/Outcome: none User Impact: missing raw data Current Software Version: 4.1.10 Diagnostic performed by Engineer : Consulted RAS Dyan Curry. Raw data was lost because system was rebooted before raw data was complete. Reviewed logs and no errors reported. Customer preformed a test scan and system is functional after Host/CIRS reboot. Internal Comments : Error # and/or description of error : No errors Repair Action : Consulted RAS Dyan Curry. Raw data was lost because system was rebooted before raw data was complete. Reviewed logs and no errors reported. Customer preformed a test scan and system is functional after Host/CIRS reboot. Resolution : Consulted RAS Dyan Curry. Raw data was lost because system was rebooted before raw data was complete. Reviewed logs and no errors reported. Customer preformed a test scan and system is functional after Host/CIRS reboot. OK to close. ----------------------------------------------------------------------------- | 3 |  | Raw data was lost because system was rebooted before raw data was complete. Reviewed logs and no errors reported. Customer preformed a test scan and system is functional after Host/CIRS reboot. | [12 empty] | 118274663 |  | 1286204 |  |  | Closed - Done | [7 empty] | Others | [6 empty] | Reboot - Host&CIRS | [5 empty] | Refresh |  |  | System issue | bar: No; No
1420 |  |  | Australia |  | 341000 | 4444100% |  | Cant complete short tube condtioning | Integration : Confirmation of order 54112963 saved ----------------------------------------------------------------------------- Work Order #: WO-04300038 09-Sep-2021 Ruben MARQUEZ The system meets the specification for the performed service and is returned to use Diagnostic performed by Engineer : Checked the CIRS - unable to connect remotely but functional. Checked system logs and carried out visual diagnostics. Resolution : Assisted David. Replaced the rotor transmitter. Cold started the system and carried out tube warm up and test scans - system ok. Carried out constancy, IQ and stand by during procedures- system ok. Problem Description by Engineer : Customer Function/Role: Radiographer Device Use: not in use Expected and Unexpected Results: Gantry/CIRS error. User Impact: NA Patient Impact: NA Current Software Version:4.1.10 ----------------------------------------------------------------------------- Work Order #: WO-04271467 27-Sep-2021 David BROWN The system meets the specification for the performed service and is returned to use Problem Description by Engineer : System unable to carry out short Tube conditioning. Resolution : Carry out system inspection. Carry full diagnostics. Reseat all PCB’s in CIRS. Carry out functional testing. Return to service. ----------------------------------------------------------------------------- Work Order #: WO-04300042 09-Sep-2021 David BROWN The system does not meet the specification for the performed service and is not returned to use Problem Description by Engineer : System not working , failing on tube conditioning. Diagnostic performed by Engineer : Check over system RSE reported problem with CIRS AQU board. Send Bug reports to Engineer for checking. ----------------------------------------------------------------------------- Work Order #: WO-04300901 05-Sep-2021 Dan Quoc Le The system does not meet the specification for the performed service and is not returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : Could be due to the optical TX/RX in the gantry ----------------------------------------------------------------------------- Work Order #: WO-04281235 05-Sep-2021 David BROWN The system does not meet the specification for the performed service and is not returned to use Problem Description by Engineer : System will not complete a short tube conditioning. Diagnostic performed by Engineer : Check over system. Check all boards on CIRS and reseat. Reset system. Problem with CIRS computer and boards. Order Aquisitor PCB and fit ASAP ----------------------------------------------------------------------------- Work Order #: WO-04270951 Brian Donnelly The system does not meet the specification for the performed service and is not returned to use Follow Up Required : Error # and/or description of error : vBufferBytes = 547356672 AcqBufferByteSize=268435456 NumBuffers=2 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 Preparing for test execution... AcqGem_Diag::DiagRegisterTest Acq board#: 0 PASSED Diagnostic Data Path Info after DIAG_REGISTER_TEST : Real-time stream Physical channel order: 0x0 CommOK Status: 0 Loss of Signal: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 Line Fault Indicator: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 AcqGem_Diag::OdlMemoryDataBusTest Acq board#: 0, Gem board#: 0 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp945:AcqGem_Diag::PollOnAcqRecvBytes: AcqPutPointer(0x0), acqGetPtr(0x0), Bytes Received(0x0) != Bytes Expected(0x5580) ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp1457: AcqGem_Diag::OdlMemoryDataBusTest PollOnAcqRecvBytes FAILED on indx: 0 AcqGem_Diag::OdlMemoryDataBusTest FAILED. Diagnostic Data Path Info after ODL_MEM_DATA_BUS_TEST : Real-time stream Physical channel order: 0x0 CommOK Status: 0 Loss of Signal: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 Line Fault Indicator: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 AcqGem_Diag::SpiFlashIdVerifyTest PASSED Diagnostic Data Path Info after FPGA_FLASH_ID_CODE_VERIFY_TEST : Real-time stream Physical channel order: 0x0 CommOK Status: 0 Loss of Signal: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 Line Fault Indicator: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp945:AcqGem_Diag::PollOnAcqRecvBytes: AcqPutPointer(0x0), acqGetPtr(0x0), Bytes Received(0x0) != Bytes Expected(0xffd500) ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp1259: AcqGem_Diag::DmaGem2Acq PollOnAcqRecvBytes FAILED. AcqGem_Diag::DmaTest FAILED. Diagnostic Data Path Info after DMA_TEST : Real-time stream Physical channel order: 0x0 CommOK Status: 0 Loss of Signal: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 Line Fault Indicator: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp945:AcqGem_Diag::PollOnAcqRecvBytes: AcqPutPointer(0x0), acqGetPtr(0x0), Bytes Received(0x0) != Bytes Expected(0xffd500) ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp1259: AcqGem_Diag::DmaGem2Acq PollOnAcqRecvBytes FAILED. AcqGem_Diag::DmaTest FAILED. Diagnostic Data Path Info after DMA_TEST : Real-time stream Physical channel order: 0x0 CommOK Status: 0 Loss of Signal: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 Line Fault Indicator: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp945:AcqGem_Diag::PollOnAcqRecvBytes: AcqPutPointer(0x0), acqGetPtr(0x0), Bytes Received(0x0) != Bytes Expected(0xffd500) ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp1259: AcqGem_Diag::DmaGem2Acq PollOnAcqRecvBytes FAILED. AcqGem_Diag::DmaTest FAILED. Diagnostic Data Path Info after DMA_TEST : Real-time stream Physical channel order: 0x0 CommOK Status: 0 Loss of Signal: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 Line Fault Indicator: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp945:AcqGem_Diag::PollOnAcqRecvBytes: AcqPutPointer(0x0), acqGetPtr(0x0), Bytes Received(0x0) != Bytes Expected(0xffd500) ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp1259: AcqGem_Diag::DmaGem2Acq PollOnAcqRecvBytes FAILED. AcqGem_Diag::DmaTest FAILED. Diagnostic Data Path Info after DMA_TEST : Real-time stream Physical channel order: 0x0 CommOK Status: 0 Loss of Signal: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 Line Fault Indicator: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiagThread.cpp55:Acquire error: WaitForStartOfSeries Gantry data error found while wait for start of series. ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp2609: Acquire GetThreadStatus(1) unsuccessful. AcqGem_Diag::GantryDataChecksTest : subtest(1) : Normal Acquisition Test FAILED. GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiagThread.cpp55:Acquire error: WaitForStartOfSeries Gantry data error found while wait for start of series. ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp2609: Unexpected WorkItem WorkType Error bit(s) set: ErrorBitsFound: 0x1, ErrorBitsExpected: 0x10 AcqGem_Diag::GantryDataChecksTest : subtest(2) : Invalid Reading Number within a shot Test FAILED. GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiagThread.cpp55:Acquire error: WaitForStartOfSeries Gantry data error found while wait for start of series. ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp2609: Unexpected WorkItem WorkType Error bit(s) set: ErrorBitsFound: 0x1, ErrorBitsExpected: 0x400 AcqGem_Diag::GantryDataChecksTest : subtest(3) : Measurement Reading not in a shot Test FAILED. GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiagThread.cpp55:Acquire error: WaitForStartOfSeries Gantry data error found while wait for start of series. ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp2609: Unexpected WorkItem WorkType Error bit(s) set: ErrorBitsFound: 0x1, ErrorBitsExpected: 0x200 AcqGem_Diag::GantryDataChecksTest : subtest(4) : Idle Reading within a shot Test FAILED. GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiagThread.cpp55:Acquire error: WaitForStartOfSeries Gantry data error found while wait for start of series. ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp2609: Unexpected WorkItem WorkType Error bit(s) set: ErrorBitsFound: 0x1, ErrorBitsExpected: 0x40 AcqGem_Diag::GantryDataChecksTest : subtest(5) : Reading UID changed in a shot Test FAILED. GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiagThread.cpp55:Acquire error: WaitForStartOfSeries Gantry data error found while wait for start of series. ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp2609: Unexpected WorkItem WorkType Error bit(s) set: ErrorBitsFound: 0x1, ErrorBitsExpected: 0x20 AcqGem_Diag::GantryDataChecksTest : subtest(6) : Invalid DFS Setting In Reading Header Test FAILED. GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiagThread.cpp55:Acquire error: WaitForStartOfSeries Gantry data error found while wait for start of series. ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp2609: Unexpected WorkItem WorkType Error bit(s) set: ErrorBitsFound: 0x1, ErrorBitsExpected: 0x8 AcqGem_Diag::GantryDataChecksTest : subtest(7) : Non-zero Reading number at start of shot Test FAILED. GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiagThread.cpp55:Acquire error: WaitForStartOfSeries Gantry data error found while wait for start of series. ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp2609: Unexpected WorkItem WorkType Error bit(s) set: ErrorBitsFound: 0x1, ErrorBitsExpected: 0x2 AcqGem_Diag::GantryDataChecksTest : subtest(8) : Non-zero Shot number at start of Series Test FAILED. GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiagThread.cpp55:Acquire error: WaitForStartOfSeries Gantry data error found while wait for start of series. ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp2609: Unexpected WorkItem WorkType Error bit(s) set: ErrorBitsFound: 0x1, ErrorBitsExpected: 0x800 AcqGem_Diag::GantryDataChecksTest : subtest(9) : Reading slice size changed in series Test FAILED. GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiagThread.cpp55:Acquire error: WaitForStartOfSeries Gantry data error found while wait for start of series. Diagnostic Data Path Info after TIMEOUT_WAITING_FOR_NEXT_READING_TEST : Real-time stream Physical channel order: 0x0 CommOK Status: 0 Loss of Signal: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 Line Fault Indicator: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 AcqGem_Diag::GantryDataChecksTest : subtest(10) : Timeout waiting for next reading Test PASSED GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiagThread.cpp55:Acquire error: WaitForStartOfSeries Gantry data error found while wait for start of series. ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp2609: Unexpected WorkItem WorkType Error bit(s) set: ErrorBitsFound: 0x1, ErrorBitsExpected: 0x80 AcqGem_Diag::GantryDataChecksTest : subtest(11) : Maximum consecutive bad readings Test FAILED. GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiagThread.cpp55:Acquire error: WaitForStartOfSeries Gantry data error found while wait for start of series. ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp2609: Unexpected WorkItem WorkType Error bit(s) set: ErrorBitsFound: 0x1, ErrorBitsExpected: 0x1000 AcqGem_Diag::GantryDataChecksTest : subtest(12) : System memory overflow Test FAILED. GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiagThread.cpp55:Acquire error: WaitForStartOfSeries Gantry data error found while wait for start of series. ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp2609: Unexpected WorkItem WorkType Error bit(s) set: ErrorBitsFound: 0x1, ErrorBitsExpected: 0x2000 AcqGem_Diag::GantryDataChecksTest : subtest(13) : WorkQueue overflow Test FAILED. GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 Diagnostic Cycle Report ------------------------ Tests completed: 20 Tests remaining: 0 Total tests queued: 20 Cycles completed: 1 Cycles remaining: 0 Tests passed this cycle: 3 Tests failed this cycle: 17 Total tests passed: 3 Total tests failed: 17 Repair Action : Downloaded bugreport and could see h2R errors. Opened CIRS logs and could see that the physical channel for the ACQ was 0x0 Logged into the CIRS remotely and ran ACQ diagnostics. 17 of 20 tests failed. Technician Name : Brian Donnelly Internal Comments : Error # and/or description of error : vBufferBytes = 547356672 AcqBufferByteSize=268435456 NumBuffers=2 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 Preparing for test execution... AcqGem_Diag::DiagRegisterTest Acq board#: 0 PASSED Diagnostic Data Path Info after DIAG_REGISTER_TEST : Real-time stream Physical channel order: 0x0 CommOK Status: 0 Loss of Signal: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 Line Fault Indicator: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 AcqGem_Diag::OdlMemoryDataBusTest Acq board#: 0, Gem board#: 0 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp945:AcqGem_Diag::PollOnAcqRecvBytes: AcqPutPointer(0x0), acqGetPtr(0x0), Bytes Received(0x0) != Bytes Expected(0x5580) ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp1457: AcqGem_Diag::OdlMemoryDataBusTest PollOnAcqRecvBytes FAILED on indx: 0 AcqGem_Diag::OdlMemoryDataBusTest FAILED. Diagnostic Data Path Info after ODL_MEM_DATA_BUS_TEST : Real-time stream Physical channel order: 0x0 CommOK Status: 0 Loss of Signal: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 Line Fault Indicator: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 AcqGem_Diag::SpiFlashIdVerifyTest PASSED Diagnostic Data Path Info after FPGA_FLASH_ID_CODE_VERIFY_TEST : Real-time stream Physical channel order: 0x0 CommOK Status: 0 Loss of Signal: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 Line Fault Indicator: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp945:AcqGem_Diag::PollOnAcqRecvBytes: AcqPutPointer(0x0), acqGetPtr(0x0), Bytes Received(0x0) != Bytes Expected(0xffd500) ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp1259: AcqGem_Diag::DmaGem2Acq PollOnAcqRecvBytes FAILED. AcqGem_Diag::DmaTest FAILED. Diagnostic Data Path Info after DMA_TEST : Real-time stream Physical channel order: 0x0 CommOK Status: 0 Loss of Signal: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 Line Fault Indicator: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp945:AcqGem_Diag::PollOnAcqRecvBytes: AcqPutPointer(0x0), acqGetPtr(0x0), Bytes Received(0x0) != Bytes Expected(0xffd500) ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp1259: AcqGem_Diag::DmaGem2Acq PollOnAcqRecvBytes FAILED. AcqGem_Diag::DmaTest FAILED. Diagnostic Data Path Info after DMA_TEST : Real-time stream Physical channel order: 0x0 CommOK Status: 0 Loss of Signal: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 Line Fault Indicator: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp945:AcqGem_Diag::PollOnAcqRecvBytes: AcqPutPointer(0x0), acqGetPtr(0x0), Bytes Received(0x0) != Bytes Expected(0xffd500) ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp1259: AcqGem_Diag::DmaGem2Acq PollOnAcqRecvBytes FAILED. AcqGem_Diag::DmaTest FAILED. Diagnostic Data Path Info after DMA_TEST : Real-time stream Physical channel order: 0x0 CommOK Status: 0 Loss of Signal: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 Line Fault Indicator: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp945:AcqGem_Diag::PollOnAcqRecvBytes: AcqPutPointer(0x0), acqGetPtr(0x0), Bytes Received(0x0) != Bytes Expected(0xffd500) ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp1259: AcqGem_Diag::DmaGem2Acq PollOnAcqRecvBytes FAILED. AcqGem_Diag::DmaTest FAILED. Diagnostic Data Path Info after DMA_TEST : Real-time stream Physical channel order: 0x0 CommOK Status: 0 Loss of Signal: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 Line Fault Indicator: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiagThread.cpp55:Acquire error: WaitForStartOfSeries Gantry data error found while wait for start of series. ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp2609: Acquire GetThreadStatus(1) unsuccessful. AcqGem_Diag::GantryDataChecksTest : subtest(1) : Normal Acquisition Test FAILED. GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiagThread.cpp55:Acquire error: WaitForStartOfSeries Gantry data error found while wait for start of series. ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp2609: Unexpected WorkItem WorkType Error bit(s) set: ErrorBitsFound: 0x1, ErrorBitsExpected: 0x10 AcqGem_Diag::GantryDataChecksTest : subtest(2) : Invalid Reading Number within a shot Test FAILED. GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiagThread.cpp55:Acquire error: WaitForStartOfSeries Gantry data error found while wait for start of series. ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp2609: Unexpected WorkItem WorkType Error bit(s) set: ErrorBitsFound: 0x1, ErrorBitsExpected: 0x400 AcqGem_Diag::GantryDataChecksTest : subtest(3) : Measurement Reading not in a shot Test FAILED. GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiagThread.cpp55:Acquire error: WaitForStartOfSeries Gantry data error found while wait for start of series. ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp2609: Unexpected WorkItem WorkType Error bit(s) set: ErrorBitsFound: 0x1, ErrorBitsExpected: 0x200 AcqGem_Diag::GantryDataChecksTest : subtest(4) : Idle Reading within a shot Test FAILED. GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiagThread.cpp55:Acquire error: WaitForStartOfSeries Gantry data error found while wait for start of series. ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp2609: Unexpected WorkItem WorkType Error bit(s) set: ErrorBitsFound: 0x1, ErrorBitsExpected: 0x40 AcqGem_Diag::GantryDataChecksTest : subtest(5) : Reading UID changed in a shot Test FAILED. GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiagThread.cpp55:Acquire error: WaitForStartOfSeries Gantry data error found while wait for start of series. ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp2609: Unexpected WorkItem WorkType Error bit(s) set: ErrorBitsFound: 0x1, ErrorBitsExpected: 0x20 AcqGem_Diag::GantryDataChecksTest : subtest(6) : Invalid DFS Setting In Reading Header Test FAILED. GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiagThread.cpp55:Acquire error: WaitForStartOfSeries Gantry data error found while wait for start of series. ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp2609: Unexpected WorkItem WorkType Error bit(s) set: ErrorBitsFound: 0x1, ErrorBitsExpected: 0x8 AcqGem_Diag::GantryDataChecksTest : subtest(7) : Non-zero Reading number at start of shot Test FAILED. GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiagThread.cpp55:Acquire error: WaitForStartOfSeries Gantry data error found while wait for start of series. ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp2609: Unexpected WorkItem WorkType Error bit(s) set: ErrorBitsFound: 0x1, ErrorBitsExpected: 0x2 AcqGem_Diag::GantryDataChecksTest : subtest(8) : Non-zero Shot number at start of Series Test FAILED. GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiagThread.cpp55:Acquire error: WaitForStartOfSeries Gantry data error found while wait for start of series. ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp2609: Unexpected WorkItem WorkType Error bit(s) set: ErrorBitsFound: 0x1, ErrorBitsExpected: 0x800 AcqGem_Diag::GantryDataChecksTest : subtest(9) : Reading slice size changed in series Test FAILED. GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiagThread.cpp55:Acquire error: WaitForStartOfSeries Gantry data error found while wait for start of series. Diagnostic Data Path Info after TIMEOUT_WAITING_FOR_NEXT_READING_TEST : Real-time stream Physical channel order: 0x0 CommOK Status: 0 Loss of Signal: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 Line Fault Indicator: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 AcqGem_Diag::GantryDataChecksTest : subtest(10) : Timeout waiting for next reading Test PASSED GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiagThread.cpp55:Acquire error: WaitForStartOfSeries Gantry data error found while wait for start of series. ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp2609: Unexpected WorkItem WorkType Error bit(s) set: ErrorBitsFound: 0x1, ErrorBitsExpected: 0x80 AcqGem_Diag::GantryDataChecksTest : subtest(11) : Maximum consecutive bad readings Test FAILED. GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiagThread.cpp55:Acquire error: WaitForStartOfSeries Gantry data error found while wait for start of series. ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp2609: Unexpected WorkItem WorkType Error bit(s) set: ErrorBitsFound: 0x1, ErrorBitsExpected: 0x1000 AcqGem_Diag::GantryDataChecksTest : subtest(12) : System memory overflow Test FAILED. GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiagThread.cpp55:Acquire error: WaitForStartOfSeries Gantry data error found while wait for start of series. ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp2609: Unexpected WorkItem WorkType Error bit(s) set: ErrorBitsFound: 0x1, ErrorBitsExpected: 0x2000 AcqGem_Diag::GantryDataChecksTest : subtest(13) : WorkQueue overflow Test FAILED. GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 Diagnostic Cycle Report ------------------------ Tests completed: 20 Tests remaining: 0 Total tests queu | 2 |  | Replaced the rotor transmitter. Cold started the system and carried out tube warm up and test scans - system ok. | [12 empty] | 118269593 |  | 1282954 |  |  | Closed - No Further Investigation |  | TBD_OBS | [5 empty] | Engineering | [6 empty] | TX | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Other | Gantry-HW-Other | Gantry-HW-Other | TBD_OBS | TBD-NoEnoughInfo | Gantry-HW | Gantry-HW-Other | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1421 |  |  | China | 鄄城马爱云妇儿医院 | 345044 | 4444100% |  | 02N/Ingenuity Core128/Elite China//consult | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ SAP Transfer : Case Transferred to SAP Safety Questions : Case: 0118266107 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Cao Last updated by: Liu Cao 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: Customer consult issue Troubleshooting Action: Customer not in working time. may call back tomorrow Repair Action: Customer not in working time. may call back tomorrow Test/Inspection: N/A Test Equipment: N/A ----------------------------------------------------------------------------- | [15 empty] | 118266107 |  | 1284519 |  |  | Closed - Non-Complaint | [20 empty] | Refresh |  |  | System issue | bar: No; No
1422 |  |  | Australia | Maia Financial Pty Limited | 341000 | 4444000% |  | Bad Detector | Work Order #: WO-04262865 David BROWN Diagnostic performed by Engineer : N/A Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: ----------------------------------------------------------------------------- | 3 |  | Product/Part - TDMS V2 M.Det Class A FRU Part number - 459800911271 | [12 empty] | 118262778 |  | 1282716 |  |  | Evaluation |  | CTUDT00632110 | [5 empty] | Engineering | [6 empty] | DMS Module | [5 empty] | Refresh |  |  | System issue | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | DMS Module | DMS-HW | DMS-HW-DMS Module | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1423 |  |  | China | Medical center (contagious disease, He nan, Xu chang) | 40047 | 4444000% |  | 02N/Ingenuity CT 728326/Add node | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer: 02N/Ingenuity CT 728326 Safety Questions:Case: 0118259071 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Gao Luqi Last updated by: Gao Luqi 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: Add node and test network Troubleshooting Action: Customer add node but test failed Delete Retrieve and query service, then test pass Repair Action: Instruct customer. Test/Inspection: N/A Test Equipment: N/A SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | Add node and test network | [12 empty] | 118259071 |  | 1279895 |  |  | NAC | [7 empty] | NAC | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | bar: No; No
1424 |  |  | United States | Mount Grant General Hospital | 40012 | 4444000% |  | 0698CDN - Gantry noise and also one image will not transfer to PACS | Gantry noise and also one image will not transfer to PACS |  |  | Need CHU to collect more probelm detail for investigation@Miley | [14 empty] | 1326147 |  |  | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS transfer | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1425 |  |  | United States | Mount Grant General Hospital | 40012 | 4444000% |  | 0698CDN - Gantry noise and also one image will not transfer to PACS | Integration : Confirmation of order 54105234 saved ----------------------------------------------------------------------------- Work Order #: WO-04315202 13-Sep-2021 Schuler Bruce The system meets the specification for the performed service and is returned to use Resolution : Replaced belt and checked image transfer status. No other issues found Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: Diagnostics Expected and Unexpected Results: Gantry noise and image will not transfer to PACS- FSE is aware User Impact: Patient Impact: Current Software Version: Parts Consumed : 1.00000,453566492401,POLY-VEE BELT ----------------------------------------------------------------------------- Work Order #: WO-04260727 09-Sep-2021 Schuler Bruce The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Diagnostic performed by Engineer : Found belt to squeak every revolution. Will Replace belt Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: Diagnostics Expected and Unexpected Results: Gantry noise and image will not transfer to PACS- FSE is aware User Impact: Patient Impact: Current Software Version: Follow Up Required : Replace Belt ----------------------------------------------------------------------------- Work Order #: WO-04260647 Dangos John The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: Diagnostics Expected and Unexpected Results: Gantry noise and image will not transfer to PACS- FSE is aware User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : Gantry noise and image will not transfer to PACS- FSE is aware Follow Up Required : Gantry noise and image will not transfer to PACS- FSE is aware ----------------------------------------------------------------------------- | 3 |  | 1.00000,453566492401,POLY-VEE BELT | [12 empty] | 118261133 |  | 1277051 |  |  | Closed - No Further Investigation |  | ECR-077012 | [5 empty] | Engineering | [6 empty] | Belt | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Belt | Gantry-HW-Belt | Gantry-HW-Belt | ECR-077012 | Belt | Gantry-HW | Gantry-HW-Belt | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1426 |  |  | United States | CENTRAL VERMONT MEDICAL CENTER | 40024 | 4444000% |  | Ring artifact on images; Replaced detector module,cleaned contrast out of ultra high Res filter, adjusted/calibrated DMS | Internal Remarks : Provided approval for a detector module. Also recommended checking the UHR comb to see if there is contrast in it along with on the Collimator. ----------------------------------------------------------------------------- Work Order #: WO-04263504 02-Sep-2021 Sangster, Robert The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Problem description by engineer : ring artifacts on temporal bone head scans Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: ring artifacts on temporal bone head scans Patient Impact/Outcome: delay, moved patient to another scanner and rescanned patient. User Impact: delay Current Software Version: Not provided Parts Consumed : 1.00000,459800911271,TDMS V2 M.Det Class A FRU Resolution : Replace detector module 23. Clean contrast out of ultra high Res filter. Adjust and calibrate the DMS as needed. ----------------------------------------------------------------------------- Work Order #: WO-04260353 01-Sep-2021 Sangster, Robert The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, no follow-up requested Problem Description by Engineer : Problem description by engineer : ring artifacts on temporal bone head scans Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: ring artifacts on temporal bone head scans Patient Impact/Outcome: delay, moved patient to another scanner and rescanned patient. User Impact: delay Current Software Version: Not provided Follow Up Required : Replace detector module Diagnostic performed by Engineer : Trouble shot issue to bad detector ordered parts ----------------------------------------------------------------------------- Work Order #: WO-04259603 Sheffield, Scott The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Problem description by engineer : ring artifacts on temporal bone head scans Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: ring artifacts on temporal bone head scans Patient Impact/Outcome: delay, moved patient to another scanner and rescanned patient. User Impact: delay Current Software Version: Not provided Follow Up Required : Error # and/or description of error : none Repair Action : I downloaded bad dets, air cal preview images and delta air. No bad dets were found during air cals Air cal images were black I did not see anything in the scan field logs attached to case for NSS if needed Recommended: Open gantry and check dms and collimator mylar covers for contrast or other fluids that may be causing artifacts. Also check the gantry mylar that it is clean Run bad detector test Consult with NSS for further assistance Technician Name : Sheffield, Scott Internal Comments : Error # and/or description of error : none Repair Action : I downloaded bad dets, air cal preview images and delta air. No bad dets were found during air cals Air cal images were black I did not see anything in the scan field logs attached to case for NSS if needed Recommended: Open gantry and check dms and collimator mylar covers for contrast or other fluids that may be causing artifacts. Also check the gantry mylar that it is clean Run bad detector test Consult with NSS for further assistance Diagnostic performed by Engineer : Chad said that he got the ring artifacts on the temporal head scan only on all images. He did a regular scan and did not have artifacts. He scanned the phantom with the temporal protocol and it had the same artifacts. I had him copy the phantom images to cdr phantom and I downloaded the images and saw the artifacts. I will attached them and logs to case. I had Chad run air cals, it gave him a message saying something was in scan field. He cleaned gantry mylar then air cals ran. After air cals he scanned the phantom again and still has the artifacts for temporal heads. He is not seeing the artifacts on any other type scans at this time. He answered yes to safety question # 2 he said because he had to re-scan the patient and the patient received extra radiation because of the re-scan. I'm dispatching FSE. ----------------------------------------------------------------------------- | 3 |  | Product/Part - TDMS V2 M.Det Class A FRU Part number - 459800911271 | [12 empty] | 118260423 |  | 1276409 |  |  | Closed - Done |  | CTUDT00632110 | [5 empty] | Engineering | [6 empty] | DMS Module | [5 empty] | Refresh |  |  | System issue | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | DMS Module | DMS-HW | DMS-HW-DMS Module | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1427 |  |  | China | Xinjiang Province Yili Prefecture Friendship Hospital | 40048 | 4443900% |  | 02N/Ingenuity CT 728326/ESTOP opened | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer:ESTOP opened SAP Transfer:Dispatched for On site Support SAP Transfer:Case Transferred to SAP External Remarks: 01.09.2021 03:44:12 UTC Jianjun Sang (86038298) Service Details: The gantry couldn't power on. Instructed the customer to check all the emergency switches, and they found that an emergency switch was pressed in. They reset the emergency switch, powered the gantry back on and the system is working properly. service result: device returns to normal working condition after repairing Internal Remarks:01.09.2021 03:44:15 UTC Jianjun Sang (86038298) $Q7: (Confirmed reported problem,the system e_stop open, can not powere on) $Q8: ( ask the customer to check e_stop switch find one switch was pressed ) $Q9: (ask the customer to reset the e_stop switch and restat the system , the system works normal ) $Q10: (Test and inspection data refer to: PA. Device returned to full functionality) $Q11: N/A Next Step Action:SANG Jianjun Safety Questions:Case: 0118254164 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Gao Luqi Last updated by: Gao Luqi 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- |  |  | ask the customer to check e_stop switch find one switch was pressed | [12 empty] | 118254164 |  | 1276688 |  |  | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Misoperation | [5 empty] | Refresh |  |  | System issue | bar: No; No
1428 |  |  | Spain | S.C.S. HOSP. GRAL. DE FUERTEVENTURA | 40020 | 4443900% |  | Images are not sent to PACS, RIS or local. | Integration:Confirmation of order 54097057 saved ----------------------------------------------------------------------------- Work Order #: WO-04248201 Carlos Saumell Lasala The system meets the specification for the performed service and is returned to use. Diagnostic performed by Engineer: The transfer queue is deleted, but the studies continue to transfer very slowly to the PACS. Problem Description by Engineer: They cannot send studies to the PACS. Regulatory Questions : N/A Resolution: Finally, the IT technicians of the Hospital found that the cause of the problem was in the PACS itself, and they have already solved it. ----------------------------------------------------------------------------- |  |  | IT technicians of the Hospital found that the cause of the problem was in the PACS itself, and they have already solved it. | [12 empty] | 118251247 |  | 1275013 |  |  | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | PACS Vendor | [5 empty] | Refresh |  |  | System issue | bar: No; No
1429 |  |  | China | Inner Mongolia Hinggan League People's Hospital | 345030 | 4443900% |  | 02N/Ingenuity Core128/Elite China/software stuck | BU Q&R Feedback: No Reason: Corrective Maintenance Safety Questions:Case: 0118249760 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Dai Suyang Last updated by: Dai Suyang 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Case Transferred to SAP External Remarks: 01.09.2021 08:33:00 UTC Zhen Liang (86039057) Service Result: [The customer said the equipment can crash when in use. Implement the database replacement and standby according to the customer's request.device returns to normal working condition.] Other significant matters: none Internal Remarks: 01.09.2021 08:33:03 UTC Zhen Liang (86039057) $Q7: Confirmed reported problem,the system was slowly and system crashed。 $Q8: on sit service，data base might be damaged. $Q9: do backup and restore .system is ok! $Q10:Test and inspection data refer to: PA. Device returned to not full functionality. $Q11: Calibrated Toolsnamen/a Current Tube usage: scan second 47825. SAP Transfer:Dispatched for On site Support Problem Reported by Customer: 02N/Ingenuity Core128/Elite China T2 Activities:Problem Description: Summary of customer's problems: system is running slow or crashed. Detailed description of Customer's problems or malfunctions: The software crashed. It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: The log has no errors. Remote troubleshoot and test result: Restart but didn't work. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: It is recommended to first reinstall the software to resolve the fault. If it is still not resolved, it is recommended to collect the bug report and upload to the technical support engineer (TSE). We recommend spare parts 12NC and names: N/A ----------------------------------------------------------------------------- |  |  | $Q7: Confirmed reported problem,the system was slowly and system crashed。 $Q8: on sit service，data base might be damaged. $Q9: do backup and restore .system is ok! | [12 empty] | 118249760 |  | 1274263 |  |  | Closed - No Further Investigation |  | NoEnoughInfo | [5 empty] | engineering | [6 empty] | System slow | [5 empty] | Refresh |  |  | System issue | System | TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | NoEnoughInfo |  | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | #N/A
1430 |  |  | China | The Second Hospital of Haibei Prefecture in Qinghai Province | 345029 | 4443900% |  | 02N/Ingenuity Core128/Elite China/need onsite support | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Clean / Lubricate Problem Reported by Customer: 02N/Ingenuity Core128/Elite China Safety Questions:Case: 0118247962 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Bai Jingwen Last updated by: Bai Jingwen 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer's problems: Other types of error. Detailed description of Customer's problems or malfunctions: The top of the machine room is leaking water onto the machine. Need to go to the field to check if it can be turned on It's a technology problem with the equipment itself: No. Troubleshooting Action: key fault code or symptom: The machine room is leaking water. The equipment needs to be checked to see if it can be turned on normally Remote troubleshoot and test result: need on site support PRS TR status: it's not through, ask engineer for maintenance on site. Repair Action: Suggest onsite service procedures: 1. Check the water leaking in the frame. Clean and dry it, and try to start the machine Test/Inspection: N/A Test Equipment: N/A SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Internal Remarks: 31.08.2021 08:14:41 UTC Yonghua Xu (86001334) RDI Accuracy: no $Q7:Confirmation reported problem,The CT scanning room is leaking, and the gantry is flooded. It has been dried, and an engineer is required to check and start the machine. $Q8:The CT scanning room is leaking, and the gantry is flooded. It has been dried, and an engineer is required to check and start the machine. Check that the equipment gantry, couch, TEAL and UPS are dry properly. $Q9:Check that the equipment gantry, couch, TEAL and UPS are dry properly.Power on test equipment tube warming and scan normal. $Q10:Test and Inspection data refer to PA , Do a few patient, Observation of a certain period of time. The system is ok. Device returned to full functionality. $Calibrated Tools name (NA) S/N (NA) Next Calibration date (NA) External Remarks: 31.08.2021 08:14:38 UTC Yonghua Xu (86001334) Department: Radiology Department's CT Room. Contact Person: Guo Rongshou. Contact Phone: 15309708576 Fault Description: The CT machine room leaks water. The main frame was in water, now it's dried. Ask the field service engineer to check and start the machine. Solution/Maintenance measures: The CT machine room leaks water. The main frame was in water, now it's dried. Ask the field service engineer to check and start the machine.Checked the device's gantry, couch, Teal cabinet and UPS and they are normal. Powered on to test the device, preheating and scanning are normal. Complete corresponding calibration, complete backing up. Device returns to normal working condition after service. ----------------------------------------------------------------------------- |  |  | $Q9:Check that the equipment gantry, couch, TEAL and UPS are dry properly.Power on test equipment tube warming and scan normal. | [12 empty] | 118247962 |  | 1273769 |  |  | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Water | [5 empty] | Refresh |  |  | System issue | [26 empty] | #N/A | bar: No to NAC
1431 |  |  | China | Jiangsu Province Yangzhou City No. 1 People's Hospital, West District Hospital | 345022 | 4443800% |  | 06J: Relocation/Ingenuity Core128/Elite China | BU Q&R Feedback: No Reason: Lifecycle Solution Problem Reported by Customer: Relocation Field Service Engineer Notes:This PO is a Reference Number only SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Safety Questions:Case: 0118241254 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Tao Fengming Last updated by: Tao Fengming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | [15 empty] | 118241254 |  | 1269145 |  |  | Closed - Non-Complaint | [20 empty] | Refresh |  |  | System issue | [26 empty] | #N/A | bar: No to NAC
1432 |  |  | China | Medical center (contagious disease, He nan, Xu chang) | 40047 | 4443800% |  | 02N/Ingenuity CT 728326/need edit examcard | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 02N/Ingenuity CT 728326 Safety Questions:Case: 0118241119 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Guoping Last updated by: Liu Guoping 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | [15 empty] | 118241119 |  | 1269083 |  |  | Closed - Non-Complaint | [20 empty] | Refresh |  |  | System issue | [26 empty] | #N/A | bar: No to NAC
1433 |  |  | China | Zhalantun City People's Hospital | 40045 | 4443800% |  | WeChat：02N/Ingenuity CT 728326/The tube overheating | BU Q&R Feedback: No SAP Customer Symptom Code: BOOT Customer Symptom Code: Boot up/Power-up Problem Reason: Corrective Maintenance Engineer Failure Code: Electrical/Electronic Failure Repair Activity Code: Instructed Customer Problem Reported by Customer:The tube overheating Safety Questions:Case: 0118241109 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Zixuan Last updated by: Wang Zixuan 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Troubleshooting Action: answer customer question Repair Action: answer customer question Test/Inspection: Functional Check PASSED. The device returned to full functionality. Test Equipment: N/A SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- |  |  | answer customer question | [12 empty] | 118241109 |  | 1269070 |  |  | Closed - No Further Investigation | [7 empty] | Application | [6 empty] | Q&A - Tube Overheat | [5 empty] | Refresh |  |  | System issue | [26 empty] | #N/A | bar: No to NAC
1434 |  |  | China | Dongguan Donghua Hospital Co., Ltd. | 40059 | 4443800% |  | WeChat: 02N/Ingenuity CT 728326/scan fault | BU Q&R Feedback: No Reason: Corrective Maintenance Safety Questions:Case: 0118241024 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Zixuan Last updated by: Wang Zixuan 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Problem Reported by Customer: scan fault T2 Activities:Problem Description: Summary of customer's problems: Scan malfunction Detailed description of Customer's problems or malfunctions: Scan malfunction It's a technology problem with the equipment itself: Yes. Troubleshooting Action: summary of customer's problem: Scan malfunction Detailed description of Customer's problems or malfunctions: Scan malfunction It's a technology problem with the equipment itself: Yes. Repair Action: Suggest onsite service procedures: Check the system log file and service further according to the log file. We recommend spare parts 12NC and names: N/A Memo of Customer request: None SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support External Remarks:07.09.2021 05:51:52 UTC Zhitao Lin (86034034) Fault description: Scan interruption, prompted error Service details: Checked the equipment, analyzed error file, did not find hardware error. Continue to monitor Service result: After service, the machine is back to normal. Other important issues: None Internal Remarks:07.09.2021 05:51:55 UTC Zhitao Lin (86034034) $Q7: (Confirmed reported problem, scan abort ) $Q8: (check log and system hardware, find nothing serious. ) $Q9: (inspect the system and restart the system) $Q10: (Test and inspection data refer to PA . Device returned to full functionality ) $Q11: Calibrated Tools name ( refer to PA ) S/N ( ) Next Calibration date ( refer to PA ) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date: N/A ^Current Tube usage: N/A ----------------------------------------------------------------------------- |  |  | inspect the system and restart the system, did not find hardware error. Continue to monitor | [12 empty] | 118241024 |  | 1268985 |  |  | Investigation | [7 empty] | Others | [6 empty] | Reboot - System | [5 empty] | Refresh |  |  | System issue | [26 empty] | #N/A | bar: No to NAC
1435 |  |  | China | 东莞东华医院有限公司 | 40059 | 4450000% |  | Wechat:02N/Ingenuity CT 728326/Connection monitoring | T2 Activities : Problem Description: 客户问题概述： 连接MP5 客户问题详细描述或故障现象： 连接MP5监护 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 机架后侧ECG口连接不上 远程诊断与测试结果： 确认心电接口连接是否正常 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 确认心电接口连接是否正常 Test Equipment: N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : FSE林志桃 13570009796 T2 Activities : Problem Description: PPD 24H: 455012301901 ECG Monitor Trigger 7M Ca 1 补发备件 Next Step Action : PPD ECG Monitor Trigger 7M Ca 455012301901 1 已安排24H到医院，FSE：林志桃-86-13570009796 | 3 |  | 服务内容：心电图仪连接不到机架，线和接头有问题，更换两根连接线 服务结果：设备恢复正常工作 | [12 empty] | 118633745 |  | 1475840 |  |  | Opened - Pending Translation | [7 empty] | Others | [6 empty] | ECG | [5 empty] | Refresh |  |  | System issue | [26 empty] | #N/A | bar: No to NAC
1436 |  |  | China | 山东省胸科医院 | 40025 | 4449900% |  | Wechat:01/Ingenuity CT 728326/Blue screen | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer : Blue screen Safety Questions : Case: 0118631654 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: han susu Last updated by: han susu 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: Cannot login in and stuck at desktop Troubleshooting Action: Cannot login in and stuck at desktop Repair Action: Login Philips account and run workflow, than issue fixed. Consider as Win10 bug. Test Equipment: N/A ----------------------------------------------------------------------------- | 3 |  | Login Philips account and run workflow, than issue fixed. Consider as Win10 bug. Link to Blue screen defect which in SP2 scope, | [12 empty] | 118631654 |  | 1475209 |  |  | Opened - Pending Translation |  | CTUDT00632553 | [5 empty] | Engineering | [6 empty] | Blue Screen | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Blue-Screen | Console-SW-Blue-Screen | Console-SW-Blue-Screen | CTUDT00632553 | Blue-Screen | Console-SW | Console-SW-Blue-Screen | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1437 |  |  | China | 黑龙江省桦川县人民医院 | 345018 | 4449800% |  | 02N/Ingenuity Core128/nemo couch cushion | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Replace Part Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ Safety Questions : Case: 0118624895 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Cao Last updated by: Liu Cao 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: Need couch cushion Troubleshooting Action: Need couch cushion Repair Action: Nemo 72h 453567517311 Cushion, Patient Support iCT 1 Test/Inspection: N/A Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | NEMO | [12 empty] | 118624895 |  | 1473233 |  |  | Opened - Pending Translation | [7 empty] | TBD | [6 empty] | NEMO | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | n | y | y | y | y | #N/A
1438 |  |  | China | Wuxi Jiangyuan Hospital | 345020 | 4449800% |  | Oil pump has abnormal noise，anode not up to speed. | Oil pump has abnormal noise，anode not up to speed. |  |  | $Q7: (Confirm reported problem. Gantry make noise.) $Q8: (Open gantry cover find Anode make noise.) $Q9: (Replace the tube12NC:989000086352.) | [14 empty] | 1469113 |  |  | Evaluation |  | PartReplacement | [5 empty] | Engineering | [6 empty] | Tube | [5 empty] | Refresh |  |  | System issue | System | IC-HW | TubeReplacement | IC-HW-TubeReplacement | IC-HW-TubeReplacement | PartReplacement |  | IC-HW | IC-HW-TubeReplacement | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | #N/A
1439 |  |  | China | Jilin Province Zhaonan City TCM Hospital | 345006 | 4449700% |  | 01/Ingenuity Core128/Elite China/The bed makes a strange noise | Next Step Action:Transfer to sales handling, this case is valid within 90 days starts from this case created, once beyond of this period without any further steps, this case is considered invalid, customer refuses our service T2 Activities:Problem Description: Summary of customer's problems: The device makes a strange sound. Detailed description of Customer's problems or malfunctions: The bed is noisy when moving horizontally It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: The bed is noisy when moving Remote troubleshoot and test result: Need an engineer to go to the site to check. PRS TR status: it's not through, ask engineer for maintenance on site. Repair Action: Suggest onsite service procedures: 1. need to check if it's related to coupling replacement | 3 |  | related to coupling replacement, customer refuses our service | [12 empty] | 118619750 |  | 1469974 |  |  | Opened |  | E-125059 | [5 empty] | Engineering | [6 empty] | Couch Noise | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | Lubricant | Couch-HW-Lubricant | Couch-HW-Lubricant | E-125059 | Lubricant | Couch-HW | Couch-HW-Lubricant | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1440 |  |  | China | Sichuan Province Yibin City No.2 T.C.M. Hospital | 345035 | 4449700% |  | Application/Ingenuity Core128/Elite China//Enhanced scanning issues | Next Step Action: newly installed device, respond, during enhanced scan, encounter trigger and dosing issue. Need to transfer to app teacher to follow up ASAP, confirm if protocol setup or operation issue.The result needs a receipt, thanks. | 5 |  | Application Support 新装机客户，反应增强扫描，触发和打药出问题，需转应用老师尽快跟进，确认是否协议设置或者操作问题。 | [12 empty] | 118618420 |  | 1469146 |  |  | Opened | [7 empty] | Application | [6 empty] | Application Support | [5 empty] | Refresh |  |  | System issue | [26 empty] | #N/A | bar: No to NAC
1441 |  |  | United States | Lackey Memorial Hospital | 40015 | 4449600% |  | Table Noises | Integration : Confirmation of order 54360580 saved Integration : According to factory calendar 20211030 (Posting date) is not a workday (check entry) ----------------------------------------------------------------------------- Work Order #: WO-04644426 28-Oct-2021 Denison Chuck The system meets the specification for the performed service and is returned to use Diagnostic performed by Engineer : The Customer said that sometimes there is a noise when lowering the patient support. Checked the couch covers. Made small adjustment to the covers. will monitor the performance of the couch support. Problem Description by Engineer : Customer Function/Role: C.T. Tech Device Use: Philips C.T. Expected and Unexpected Results: None System Up. User Impact: None System up. Patient Impact: None System up. Current Software Version: 4.1.8 | 5 |  | Made small adjustment to the covers. will monitor the performance of the couch support. | [12 empty] | 118616928 |  | 1466558 |  |  | Evaluation | [7 empty] | Others | [6 empty] | Couch Cover | [5 empty] | Refresh |  |  | System issue | [26 empty] | #N/A | bar: No to NAC
1442 |  |  | United States | Unity Health Systems | 40037 | 4449600% |  | 0306CFX84 Gantry cannot comply | The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Problem description by engineer : Gantry cannot comply When was the first occurrence? Within... : 27-Oct-2021 How often is issue occurring? : First occurrence Malfunction area : Table Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Not provided Patient Impact/Outcome: None User Impact: Customer had to restart Current Software Version: Not provided Diagnostic performed by Engineer : System went down and customer power cycled and did a bugreport now. Biomed will come in the morning and run checks. Br`uce want us to check the bugreport and let him know what may have caused this issue. Internal Comments : Error # and/or description of error : 2021/10/27 02:54:56.520__001 S_HMC_CTRL_A_POSITION_ERROR_LIMIT_EXCEEDED ERROR_NORMAL Series Exit Logger.mdb Checked ACS Code 25 (P1=25) - External problem is preventing the controller to track the position correctly. Not enough (relative) encoder pulses reached the ACS in time Repair Action : Check for an obstruction or mechanical load problem, check for any obstructions that might have caused this matter (Couch friction, Caught mattress, etc.). Resolution : This may have been due to an obstruction on the couch or a hardware glitch. Customer power cycled and system is running as normal now | 2 |  | This may have been due to an obstruction on the couch or a hardware glitch. Customer power cycled and system is running as normal now | [12 empty] | 118610869 |  | 1460632 |  |  | Investigation |  | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Reboot | [5 empty] | Refresh |  |  | System issue | System | Gantry-SW | NoEnoughInfo | Gantry-SW-NoEnoughInfo | Gantry-SW-NoEnoughInfo | NoEnoughInfo | TBD-NoEnoughInfo | Gantry-SW | Gantry-SW-NoEnoughInfo | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1443 |  |  | China | 江苏省徐州医学院附属第三医院 | 40039 | 4449600% |  | 02G/Ingenuity CT 728326/add dicom node | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Connection/Cable/Interconnect Repair Activity Code: Instructed Customer Problem Reported by Customer : 02G/Ingenuity CT 728326 Safety Questions : Case: 0118610545 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: add dicom node Troubleshooting Action: instruct customer add dicom node, modify exam card. system is ok, close case. Repair Action: Functional Check PASSED. Device returned to full functionality. Test/Inspection: N/A Test Equipment: N/A ----------------------------------------------------------------------------- |  |  | 2021-10-27 | [12 empty] | 118610545 |  | 1464913 |  |  | Closed - Non-Complaint | [7 empty] | 2021-10-27 | [6 empty] | 2021-10-28 14:53 | [5 empty] | Refresh |  |  | System issue | [26 empty] | #N/A | bar: No to NAC
1444 |  |  | United States | Leesburg Regional Medical Center | 40022 | 4449600% |  | HIS RIS issue | Expected and Unexpected results: Unable to connect to HIS/RIS Patient Impact/Outcome: none User Impact: Manual patient data entry Current Software Version: 4.1.10 Diagnostic performed by Engineer : Able to ping system, Able to connect to CIRS, Customer able to send to other destinations. Unable to log into olevel due to password change. Checked worklist setting = ok. Advised customer to contact PACS admin. Will leave case open if customer needs further assistance. Internal Comments : Error # and/or description of error : No error Repair Action : Able to ping system, Able to connect to CIRS, Customer able to send to other destinations. Unable to log into olevel due to password change. Checked worklist setting = ok. Advised customer to contact PACS admin. Will leave case open if customer needs further assistance. Resolution : No further needs from customer. Closing case. | 3 |  | No further needs from customer. customer to contact PACS admin | [12 empty] | 118609634 |  | 1459632 |  |  | Evaluation | [7 empty] | Others | [6 empty] | PACS Vendor | [5 empty] | Refresh |  |  | System issue | [26 empty] | #N/A | bar: No to NAC
1445 |  |  | China | Dongguan Donghua Hospital Co., Ltd. | 40059 | 4449500% |  | wechat：02N/Ingenuity CT 728326/Panel fault | T2 Activities:Problem Description: Summary of customer's problems: Other types of error. Detailed description of Customer's problems or malfunctions: Gantry Right Panel button is not working. It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: Confirmed the problem with the engineer. Remote troubleshoot and test result: Restart but didn't work. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Suggest to replace with a new right panel of the frame to eliminate the problem. We recommend spare parts 12NC and names: PPD 459800484511 Next Step Action:PPD CONTROL PANEL ASSEMBLY-FRONT RH 459800484511 1 parts are shipped, FSE: Lin Zhitao-86-13570009796 24H SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support ----------------------------------------------------------------------------- | 3 |  | 服务内容：机架右侧面板按键破损，更换新的面板 服务结果：设备恢复正常工作 | [12 empty] | 118597517 |  | 1460488 |  |  | Opened - Pending Translation |  | CTUDT00632181 | [5 empty] | Engineering | [6 empty] | GCP | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Gantry Panel | Gantry-HW-Gantry Panel | Gantry-HW-Gantry Panel | CTUDT00632181 | Control Panel | Gantry-HW | Gantry-HW-Gantry Panel | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1446 |  |  | Spain | Comercial Centro Medico | 40042 | 4449400% |  | Unable to turn off the computer in the daily shutdown | Work Order #: WO-04615355 [REDACTED] The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled. Regulatory Questions : N/A Problem Description by Engineer: We talk with Vanessa, the CT technician. In the mornings, when trying to do tube heating, or start acquisitions, the equipment displays a message indicating a reconstructor failure (CIRS). They have to reset the equipment completely to be able to work. Diagnostic performed by Engineer: The host error logs do not show any detailed information about the cause of the problem; simply the gantry cancels the acquisitions due to a host order. But it is not possible to access the CIRS Event Viewer since there is no direct connection to this computer. Follow Up Required: We should go to the site and check the CIRS event file when acquisitions or tube heating are blocked (see first thing on 10/25). See if the communication between the gantry and CIRS shows then any problem. ----------------------------------------------------------------------------- On Monday 10/25/2021, when launching the tube heating, the process did not start. They rebooted the system and after the acquisition, it gave a CIRS error. In the system error log, there is a main contactor error, which was resolved with the first reboot. Then the IMR rebuilder did not start correctly. They rebooted again and it already worked properly. During my diagnosis, they performed several patients, and the system worked correctly. It is not necessary to intervene in the system because they are specific errors that have not been reproduced the rest of the week. | 3 |  | In the system error log, there is a main contactor error, which was resolved with the first reboot. Then the IMR rebuilder did not start correctly. They rebooted again and it already worked properly. | [12 empty] | 118589840 |  | 1455731 |  |  | Evaluation - Pending Translation |  | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Reboot | [5 empty] | Refresh |  |  | System issue | System | CIRS-SW | NoEnoughInfo | CIRS-SW-NoEnoughInfo | CIRS-SW-NoEnoughInfo |  |  |  |  | No | 0 | 0 | 0 | 0 | 0 | No | y | y | y | y | y | y | y | y | 0 | y | #N/A
1447 |  |  | China | Wuxi City Jiangyuan Hospital | 345020 | 4449400% |  | 02N/Ingenuity Core128/Elite China/have noise | T2 Activities:Problem Description: Summary of customer's problems: Other types of error. Detailed description of Customer's problems or malfunctions: There is a loud noise in the rack It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: There is a loud noise in the rack Remote troubleshoot and test result: It should be a loud fan noise PRS TR status: it's not through, ask engineer for maintenance on site. Repair Action: Suggest onsite service procedures: Check the fans, SAP Transfer:Dispatched for On site Support ----------------------------------------------------------------------------- | 3 |  | $Q7: (Confirm reported problem. Gantry make noise.) $Q8: (Open gantry cover find Anode make noise.) $Q9: (Replace the tube12NC:989000086352.) | [12 empty] | 118587501 |  | 1455637 |  |  | Opened - Pending Translation |  | PartReplacement | [5 empty] | Engineering | [6 empty] | Tube | [5 empty] | Refresh |  |  | System issue | System | IC-HW | TubeReplacement | IC-HW-TubeReplacement | IC-HW-TubeReplacement | PartReplacement |  | IC-HW | IC-HW-TubeReplacement | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | #N/A
1448 |  |  | Japan | Independent administrative institution National Hospital Organization | 341013 | 4449400% |  | HY: Memory replacement work | Work Order #: WO-04610607 27-Oct-2021 TE 005 The system meets the specification for the performed service and is returned to use. Memory replacement operation check Reported to the on-duty technician for approval Parts shipping Part number (12nc) Part 1: 459801633881 Part name Part 1: 64GB DDR4-2666 (8x8GB) RegRAM Number of parts 1: 1 | 5 |  | Replace Memory Set | [12 empty] | 118585611 |  | 1455379 |  |  | Evaluation |  | MemoryReplacement | [5 empty] | Engineering | [6 empty] | Memory | [5 empty] | Refresh |  |  | System issue | System | Console-HW | MemoryReplacement | Console-HW-MemoryReplacement | Console-HW-MemoryReplacement | MemoryReplacement | MemoryReplacement | Console-HW | Console-HW-MemoryReplacement | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1449 |  |  | United States | McLaren Macomb | 40021 | 4449400% |  | Have a question about a study | Internal Remarks : CSA INTERACTION TEMPLATE : Interaction Type: New Case Caller: Nadia Lewis - phone # 586 945 4530 IP & Issue: 85380634 - Have a question about a study Event Type: Incident Service Request: Remote Clinical Did Customer Decline Remote Support? No Preferred Service Time: ASAP Entitlement Options: Claims Contract (with reason) Reason for Entitlement Claim: N/A Action Taken: Created case & locked it for entitlement verification Next Step Action: Awaiting Entitlement Verbal PO Number: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- Work Order #: WO-04616049 Maier, Donnie The system meets the specification for the performed service and is returned to use Diagnostic performed by Engineer : Called and left VM @11:30AM EST for contact in case to call back in for assistance. Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: Diagnostic Expected and Unexpected Results: Question about a study User Impact: No Patient Impact: No Current Software Version: 4.1.10 | 5 |  | No clinical applications support is needed | [12 empty] | 118589569 |  | 1453484 |  |  | Evaluation | [7 empty] | Application | [6 empty] | Application Support | [5 empty] | Refresh |  |  | System issue | [26 empty] | #N/A | bar: No to NAC
1450 |  |  | Puerto Rico | HOSPITAL DR SUSONI INC | 40013 | 4449400% |  | Device does not accepts Login and password | Work Order #: WO-04615609 25-Oct-2021 Ramos Rivera, Pablo The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Problem description by engineer : User is disabled. Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected results: User CT enabled. Patient Impact/Outcome: System is down User Impact: System is down Current Software Version: 4.1.10 Resolution : CT user was disabled. Reseted the CT user password. Now it is working properly. Ran abdomen scans, system is working properly. ----------------------------------------------------------------------------- Work Order #: WO-04614430 Valenzuela Velasquez, Eduardo Ariel The system does not meet the specification for the performed service and is not returned to use Internal Comments : Error # and/or description of error : CT user disabled. Repair Action : - Reset password. - Enable emergency user. - GIve access to customer for the IT user Diagnostic performed by Engineer : CT user is disabled, no login possible. Remote access is limited and reset password is not available. Problem Description by Engineer : Problem description by engineer : User is disabled. Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected results: User CT enabled. Patient Impact/Outcome: System is down User Impact: System is down Current Software Version: 4.1.10 Follow Up Required : Error # and/or description of error : CT user disabled. Repair Action : - Reset password. - Enable emergency user. - GIve access to customer for the IT user | 2 |  | - Reset password. - Enable emergency user. - GIve access to customer for the IT user | [12 empty] | 118589165 |  | 1453243 |  |  | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Password | [5 empty] | Refresh |  |  | System issue | [26 empty] | #N/A | bar: No to NAC
1451 |  |  | China | Henan Province Puyang City People’s Hospital | 40031 | 4449300% |  | wechat:02N/Ingenuity CT 728326/system crash | T2 Activities:Problem Description: Summary of customer's problems: Other types of error. Detailed description of Customer's problems or malfunctions: Computer is crashing. It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: Computer is crashing. Remote troubleshoot and test result: It returned to normal after restart. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Collect bug report, escalation SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Internal Remarks : 27.10.2021 12:57:39 UTC (86038950) Q7:Confirmed report problem: the host crash. Q8:Perform visual check on site, clean the disk and continue to watch. $Q9:clean the disk and continue to watch, Device return to full function. $Q10: test and inspection data refer to PA. Device returned to full functionality. $Q11:Calibrated Tools refer to PA. External Remarks: 27.10.2021 12:57:36 UTC (86038950) service details: host machine crash. Click on any area without response. Clean the disc and junk on the site. Continue to monitor. Service outcome: clean the disc and junk on the field. The equipment is back to normal. Continue to monitor. Other significant matters: none | 5 |  | 服务内容： 主机死机，点击任何地方无响应，到场清理磁盘和垃圾，继续观察。 服务结果： 到场清理磁盘和垃圾，设备恢复正常，继续观察。 Need bugreport for investigation | [12 empty] | 118582823 |  | 1451593 |  |  | Closed - No Further Investigation | [7 empty] | Engineering | [6 empty] | SW Crash | [5 empty] | Refresh |  |  | System issue | System | Console-SW | DiskFull | Console-SW-DiskFull | Console-SW-DiskFull |  |  |  |  | No | 0 | 0 | #REF! | 0 | 0 | No | y | y | y | y | y | y | y | y | y | y | #N/A
1452 |  |  | Poland | Independent Public Health Maintenance Organization of the Ministry of Interior Affairs and Administration in Katowice. | 341008 | 4449100% |  | DICOM configuration and inspection of sanitary protection fixed guards. | DICOM configuration and inspection of sanitary protection fixed guards. Work related to the preparation of the tomograph for commissioning in a new location: 20/10/2021 - visit, DICOM configuration + assistance for Sanepid 28/10/2021 - remote service configuration visit. | 4 |  | NAC_DICOM Configuration | [12 empty] | 118578700 |  | 1450273 |  |  | Evaluation - Pending Translation | [7 empty] | NAC | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [26 empty] | #N/A | bar: No to NAC
1453 |  |  | China | Jilin Province Zhaonan City TCM Hospital | 345006 | 4449100% |  | 01/Ingenuity Core128/Elite China/Estop opened | Service details: The horizontal movement of the bed loses power. The coupling position of the encoder is damaged, replace. Service outcome: device returns to normal working condition after service Other significant matters: None. Internal Remarks : 24.10.2021 07:15:53 UTC (86038446) RDI Accuracy ：No $Q7: (Confirmed reported problem: E-stop opened when the couch Horizontal motion) $Q8: (Review logger and check the couch, The couch Horizontal coupling damaged) $Q9: (Replaced the part and calibration) $Q10: (Test and inspection data refer to: PA. The device returned to full functionality. $Q11 Calibrated Tools name (FLUKE287C) S/N (27420003) Next Calibration date (03/17/2022) | 3 |  | Replace Horizontal coupling, Adjust the cover and test for the noise issue | [12 empty] | 118576982 |  | 1449651 |  |  | Investigation - Pending Translation |  | CTUDT00632179 | [5 empty] | Engineering | [6 empty] | Coupling | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | Motor Coupling | Couch-HW-Motor Coupling | Couch-HW-Motor Coupling | CTUDT00632179 | Motor Coupling | Couch-HW | Couch-HW-Motor Coupling | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1454 |  |  | China | Jiangxi Province Children’s Hospital | 40041 | 4449100% |  | 02N/Ingenuity CT 728326/Gantry nosie | SAP Transfer:Dispatched for On site Support T2 Activities:Problem Description: Summary of customer's problems: The device makes a strange sound. Detailed description of Customer's problems or malfunctions: There is a problem with the couch's cover. It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: NA Remote troubleshoot and test result: The gantry's rotation is noisy when scanning but not when it's at rest. It is recommended for the customer to consider not using it for now.Contacted the field service engineer, who says that it's a problem with the position of the slide rail pulley that happened before. PRS TR status: it's not through, ask engineer for maintenance on site. Repair Action: Suggest onsite service procedures: Check the rail pulley components. Check for scraping, large belt, etc. We recommend spare parts 12NC and names: " 459800445141 SPLIT TAPER BUSHING； 453566563321 293,76POLY-V-SHEAVE (T84A-1007)" Next Step Action: PPD have arranged for 24h delivery, please dispatch labor. Field service engineer: Li Jiawei - 86-18579112310 External Remarks: 24.10.2021 15:53:30 UTC (86091272)"SUAT. Ready for Clinical Use local time" : "2021-10-23 16:30" Service description: The customer reported that the machine is very noisy. Checked and found that the motor wheel is moving around, causing friction. Replaced the roller cone block and adjusted the mechanical position. Service result: The device returns to normal working condition after service. External Remarks: 25.10.2021 02:18:10 UTC (86091296) Service details: The customer reported that the machine is very noisy. Checked and found that the motor wheel is moving around, causing friction. Replaced the roller cone block and adjusted the mechanical position. Service result: The device returns to normal working condition after service. Internal Remarks : 27.10.2021 10:39:34 UTC (86091272) "RDI Accuracy:Yes" $Q7:(Confirmed reported problem. gantry noise.) $Q8:( Split Taper bushing position moved, The belt is friction with the machine. ) $Q9:( Replace POLY-V-SHEAVE solved the problem.) $Q10:(Test and inspection data refer to PA. Device returned to full functionality.) $Q11:Calibrated Tools: No tools used. Internal Remarks : 27.10.2021 14:51:05 UTC (86091296)"RDI Accuracy:Yes" $Q7:(Confirmed reported problem. gantry noise.) $Q8:( Split Taper bushing position moved, The belt is friction with the machine. ) $Q9:( Replace POLY-V-SHEAVE solved the problem.) $Q10:(Test and inspection data refer to PA. Device returned to full functionality.) $Q11:Calibrated Tools: No tools used. | 2 |  | Split Taper bushing position adjusted. Replace POLY-V-SHEAVE 453566563321 solved the problem. | [12 empty] | 118575835 |  | 1449198 |  |  | Closed - No Further Investigation |  | ECR-076490 | [5 empty] | Engineering | [6 empty] | Sheave | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Sheave | Gantry-HW-Sheave | Gantry-HW-Sheave | ECR-076490 |  | Gantry-HW | Gantry-HW-Sheave | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | #N/A
1455 |  |  | Japan | Independent administrative institution National Hospital Organization | 341013 | 4449100% |  | 03DS: RS: Tracker behavior becomes heavy and then E-STOP OPEN >> Reboot and return >> S_GEN_FAULT_ANODE_TUBE_ARC / S_GEN_FAULT_CATHODE_TUBE_ARC | Work Order #: WO-04598981 27-Oct-2021 Yamamoto, Yuko The system meets the specification for the performed service and is returned to use. Resolution: Performed witnessing. Internal Comments: [CX related] Did you agree on how to complete the report with the person in charge of devices? <Select one of the following> Yes. I completed normal reporting that day. [Work report (paper) submission] Did you submit a work report on paper? Yes. Work Order #: WO-04598349 22-Oct-2021 Kawazoe, Junichi The system meets the specification for the performed service and is returned to use. Regulatory Questions : N/A Resolution: E tc [gibberish] Internal Comments: [CX related] Did you agree on how to complete the report with the person in charge of devices? YES On-site the next day [Work report (paper) submission] Did you submit a work report on paper? YES Resolution : Work details: I carried out Long TC and Filament Cal.I acquired the bug report.There is no problem with the operation of the equipment.I reported to Technician Yamamoto that the work was complete. ----------------------------------------------------------------------------- Work Order #: WO-04598090 Satake, Rui The system meets the specification for the performed service and is returned to use. Diagnostic performed by Engineer: After he had switched the key switch on/off and restarted the host, it recovered and scanning was possible. I confirmed a record of an X-ray tube discharge from the log. S_GEN_FAULT_ANODE_TUBE_ARC/S_GEN_FAULT_CATHODE_TUBE_ARC However, I could not determine from the log what had slowed down the host immediately before. I asked the customer to create a bug report.An FSE will respond. Problem Description by Engineer: During a tracker scan, the host's behavior became sluggish and then the E-Stop opened. >>> "S_GEN_FAULT_ANODE_TUBE_ARC/S_GEN_FAULT_CATHODE_TUBE_ARC" Internal Comments: [Subject] 03DS:RS:Mr. Yamamoto:During a tracker scan, the host's behavior became sluggish and then the E-Stop opened. >> Following a restart, it recovered. >> "S_GEN_FAULT_ANODE_TUBE_ARC/S_GEN_FAULT_CATHODE_TUBE_ARC" [Internal Comments] None Dispatched: 1 RC FSE Registration Information: Name of FSE 1 who visited the site: Junichi Kawazoe [L2 confirmation] FSE 1 visit date and time: 2021/10/2216:00 FSE 1 hours worked: 4 ----------------------------------------------------------------------------- | 4 |  | I carried out Long TC and Filament Cal. I acquired the bug report.There is no problem with the operation of the equipment. | [12 empty] | 118575342 |  | 1444759 |  |  | Evaluation |  | CTUDT00630085 | [5 empty] | Engineering | [6 empty] | Tube Arc | [5 empty] | Refresh |  |  | System issue | System | IC-HW | Tub Arc | IC-HW-Tub Arc | IC-HW-Tub Arc | CTUDT00630085 | Tub Arc | IC-HW | IC-HW-Tub Arc | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1456 |  |  | Turkey | Republic of Turkey Ministry of Health, Bagcilar | 40010 | 4449000% |  | 0115747107/WO-01714930 Torque tool calibration process | Integration: VBAKKOM has been processed successfully-VBAPKOM has been processed successfully-KONVKOM has been processed successfully-KONVKOM has been processed successfully-KONVKOM has been processed successfully-KONVKOM has been processed successfully-REQ2RES DMR 0018257989 has been saved-MPG is present on the product code: P9. ----------------------------------------------------------------------------- | 3 |  | NAC_Torque tool calibration | [12 empty] | 118571566 |  | 1446313 |  |  | Evaluation | [7 empty] | NAC | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [26 empty] | #N/A | bar: No to NAC
1457 |  |  | China | Shaanxi Province Xianyang Central Hospital (Xianyang City No. 2 People's Hospital) | 40050 | 4449000% |  | 02N/Ingenuity CT 728326/Rack power down | T2 Activities:Problem Description: Summary of customer's problems: Frequent power failure of rack Detailed description of Customer's problems or malfunctions: Frequent power failure of rack It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: S_MC_DEVICE_VERT_POT_OPEN Remote troubleshoot and test result: N/A PRS TR status: it's not through, ask engineer for maintenance on site. Repair Action: Suggest onsite service procedures: On site, check the couch's vertical encoder. If necessary, replace it. We recommend spare parts 12NC and names: Vertical Transducer : 459800623331 Memo of Customer request: None SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Next Step Action: PPD parts 459800623331 arranged for 24H to the hospital, Guo Yujie - 86-15158180650 External Remarks: 24.10.2021 02:52:01 UTC (86038990) service details: the machine frequently loses power. Replace the bed's height POT and it restore. service result: device returns to normal working condition after service ----------------------------------------------------------------------------- | 3 |  | Vertical Transducer : 459800623331 | [12 empty] | 118571505 |  | 1446274 |  |  | Investigation |  | NotEngineering-partreplace | [5 empty] | Engineering | [6 empty] | Vertical Transducer | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | VerticalTransducer | Couch-HW-VerticalTransducer | Couch-HW-VerticalTransducer | NotEngineering-partreplace |  | Couch-HW | Couch-HW-VerticalTransducer | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | #N/A
1458 |  |  | China | Fujian Provincial Hospital | 345011 | 4449000% |  | 02N/Ingenuity Core128/Worklist and PACS | Unable to connect PACS and WORKLIST. Troubleshooting found network cable error. Replace the network cable with The Hospital's IT dept. Service result (required): Device returns to normal working condition after service Other significant matters: none Internal Remarks : 22.10.2021 03:56:29 UTC (86036497) “RDI Accuracy: Yes” $Q7: confirmed the reported problem, Pacs and worklist are disconnected. $Q8: on site Reviewed the error log, Performed visual inspection and found Network cables are faulty. $Q9: Cooperate with hospital information department to replace network cables. $Q10: Test and Inspection data refer to:PA.Device returned to full functionality. $Q11: Calibrated Tools refer to:PA ^Current Tube installed date: NA ^Current Tube usage: NA | 3 |  | Network cables are faulty. | [12 empty] | 118569398 |  | 1444853 |  |  | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Net cable | [5 empty] | Refresh |  |  | System issue | [26 empty] | #N/A | bar: No to NAC
1459 |  |  | Japan | Maki Hospital | 341017 | 4448900% |  | HY: Arc countermeasure work | Work Order #: WO-04583810 27-Oct-2021 Yamamoto, Yoshiyuki The system meets the specification for the performed service and is returned to use. Regulatory Questions : N/A Problem Description by Engineer: Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : Work details: Arching has been occurring occasionally on the X-ray tube while it is warming up (Short TC), and given the frequency of the problem, I thought it could be due to accumulated impurities (causing arcing) in the X-ray tube, so I performed aging (Long TC). Internal Comments: [CX related] Did you agree on how to complete the report with the person in charge of devices? Yes. I completed normal reporting that day. [Work report (paper) submission] Did you submit a work report on paper? <Yes> Q1. How did the examination finish? ・As it stopped before X-ray emission, no extra irradiation Q2. Were there any mistaken findings (misdiagnosis, mistaken treatment) due to the problem? None. Q3. Were there any artifacts? None. Q4. Was there any unintended unnecessary exposure? None. ----------------------------------------------------------------------------- | 5 |  | it could be due to accumulated impurities (causing arcing) in the X-ray tube, so I performed aging (Long TC). Also replace Part/Product - ECG Monitor Trigger 7M Ca Part Number - 455012301901 | [12 empty] | 118564450 |  | 1441305 |  |  | Closed - No Further Investigation |  | CTUDT00630085 | [5 empty] | Engineering | [6 empty] | Tube Arc | [5 empty] | Refresh |  |  | System issue | System | IC-HW | Tub Arc | IC-HW-Tub Arc | IC-HW-Tub Arc | CTUDT00630085 | Tub Arc | IC-HW | IC-HW-Tub Arc | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1460 |  |  | China | 北京市中西医结合医院 | 40032 | 4448900% |  | Wechat:02N/Ingenuity CT 728326/Screen fault | T2 Activities : Problem Description: 客户问题概述： 其他类型故障。 客户问题详细描述或故障现象： 显示器不亮 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 显示器不亮 远程诊断与测试结果： 数据线连接问题 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 检查显示器连线 SAP Transfer : Dispatched for Onsite Support | 3 |  | found DVI cable broken,after change backup cable,device work normal. | [12 empty] | 118562409 |  | 1439956 |  |  | Closed - No Further Investigation |  | CTUDT00632173 | [5 empty] | Engineering | [6 empty] | DVI Cable | [5 empty] | Refresh |  |  | System issue | System | Console-HW | Adaptor | Console-HW-Adaptor | Console-HW-Adaptor | CTUDT00632173 | Adaptor | Console-HW | Console-HW-Adaptor | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1461 |  |  | China | 陕西省咸阳中心医院 (咸阳市第二人民医院） | 40050 | 4448900% |  | wechat:02N/Ingenuity CT 728326/tube can't expose | SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Problem Reported by Customer : tube can't expose T2 Activities : Problem Description: 客户问题概述： 扫描失败。 客户问题详细描述或故障现象： 机架无法扫描。 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： PRS不通。 远程诊断与测试结果： 重启无效。 PRS TR状态： 不通，请工程师到场检查。 Repair Action: 建议现场维修步骤： 建议现场根据log或event信息排查机器故障。 建议备件12NC及名称： N/A External Remarks : 20.10.2021 03:07:37 UTC (86038990)"SUAT. Ready for Clinical Use local time" : "2021-10-20 11:00" 服务内容：不能曝光，重启后恢复。 服务结果：服务后，设备正常工作。 ----------------------------------------------------------------------------- | 2 |  | 服务内容：不能曝光，重启后恢复。 服务结果：服务后，设备正常工作。 review error logs ,found CIRS event log report error ,ACQ not ok, check the CIRS optical fiber is connect, perform the VD test, not found error, so to perform the restart the CIRS computer, device runing normal, so think customer observation and use. $Q9:maybe is CIRS computer not normal start, restart CIRS device return normal., | [12 empty] | 118561943 |  | 1439686 |  |  | Closed - No Further Investigation |  | CTUDT00630408 | [5 empty] | Engineering | [6 empty] | ACQ_NOT_OK | [5 empty] | Refresh |  |  | System issue | System | CIRS-SW | ACQ_NOT_OK | CIRS-SW-ACQ_NOT_OK | CIRS-SW-ACQ_NOT_OK | ACQ_NOT_OK | ACQ_NOT_OK | CIRS-SW | CIRS-SW-ACQ_NOT_OK | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | Reboot | #N/A
1462 |  |  | Japan | Maki Hospital | 341017 | 4448900% |  | 作業日程に関して>>FSEより回答 |  |  |  | 2021-10-20 | [12 empty] | 118561986 |  | 1439708 |  |  | Closed - Non-Complaint | [7 empty] | 2021-10-20 | [6 empty] | 2021-10-21 09:18 | [5 empty] | Refresh |  |  | System issue | [26 empty] | #N/A | bar: No to NAC
1463 |  |  | Australia | Maia Financial Pty Limited | 341000 | 4449000% |  | Quality Control Computer Network Configuration | Integration : Confirmation of order 54329236 saved Integration : According to factory calendar 20211023 (Posting date) is not a workday (check entry) ----------------------------------------------------------------------------- Work Order #: WO-04589670 Steve, Teixeira The system meets the specification for the performed service and is returned to use Problem Description by Engineer : ________________________________________________________________ CUSTOMER REPORTED PROBLEM OR EVENT DESCRIPTION: 21 Oct 2021 Customer called requesting assistance to change the AE Titles for the Quality Control Computers, in the CT network configuration. ________________________________________________________________ Diagnostic performed by Engineer : ________________________________________________________________ CUSTOMER REPORTED PROBLEM OR EVENT DESCRIPTION: 21 Oct 2021 Customer called requesting assistance to change the AE Titles for the Quality Control Computers, in the CT network configuration. ________________________________________________________________ CONFIGURATION: 21 Oct 2021 Assisted the Customer with the configuration of the RAD600 and RAD600Vidar network nodes. Customer to test new configuration. ________________________________________________________________ FOLLOW UP: 22 Oct 2021 Contacted site and confirmed that transfers successful. Customer had to do further configuring, as IP Addresses also required amending. The AE Title and IP address changed to: RA600 - AE Title: NC_MUR_RA_02 - 10.64.65.143 RA600Vidar - AE Title : NC_MUR_RA_01 - 10.64.65.142. ________________________________________________________________ Resolution : ________________________________________________________________ CASE CLOSED: As the transfers to the Quality Control computers are successful after the CT network configuration, as agreed with the Customer, no further service related action will be taken on this case. ________________________________________________________________ ----------------------------------------------------------------------------- |  |  | 2021-10-21 | [12 empty] | 118568654 |  | 1439684 |  |  | Closed - Non-Complaint | [7 empty] | 2021-10-21 | [6 empty] | 2021-10-21 09:08 | [5 empty] | Refresh |  |  | System issue | [26 empty] | #N/A | bar: No to NAC
1464 |  |  | Japan | 医療法人清翠会 牧病院 | 341017 | 4448800% |  | 03CS: KKI: Short TC stopped at 76% >> S_GEN_FAULT_ANODE_TUBE_ARC / S_RHOST_HVG_HSS_GENERATOR_ARC_FAUL | Work Order #: WO-04576935 Kimura, Kazumasa The system meets the specification for the performed service and is returned to use Problem Description by Engineer : ShortTCが76%で止まっていた＞＞S_GEN_FAULT_ANODE_TUBE_ARC/S_RHOST_HVG_HSS_GENERATOR_ARC_FAUL Resolution : 【対応】 継続して検査可 Internal Comments : [Subject] 03CM:KKI:ﾄｶﾞﾜ様:ShortTCが76%で止まっていた＞＞S_GEN_FAULT_ANODE_TUBE_ARC/S_RHOST_HVG_HSS_GENERATOR_ARC_FAUL [Internal Comments] 【ログ確認結果】 S_GEN_FAULT_ANODE_TUBE_ARC S_RHOST_HVG_HSS_GENERATOR_ARC_FAUL 陽極側にてアークが発生しておりました。 様子見とさせて頂いております。 [4/8 Elements] Q5 その検査はどの様に終了しましたか?:無しQ6 問題により間違った所見（誤診、誤治療）の有無:無し Q7 意図しない不要な曝射の有無:無し Q8 アーチファクトの有無:無し ----------------------------------------------------------------------------- | 4 |  | Duplicate with 118564450 | [12 empty] | 118558926 |  | 1436681 |  |  | Closed - No Further Investigation | [7 empty] | TBD | [6 empty] | Tube Arc | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | n | y | y | y | y | #N/A
1465 |  |  | China | 山东省泰安市中医二院 | 345037 | 4448800% |  | 02N/Ingenuity Core128/Elite China/report error | T2 Activities : Problem Description: 客户问题概述： 扫描失败。 客户问题详细描述或故障现象： 无法扫描 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： S_COUCHMGR_HOR_MOTION_STOP_ERROR；S_HMC_VELOCITY_ENCODER_FAILURE ；S_MPS_WAITING_FOR_MAIN_DRIVE_TO_CLOSE 远程诊断与测试结果： 床水平运动掉电一次，后main contactor不吸合。客户告知大关过一次无效。晚上没人上班，无法尝试再重启，沟通后先派工。 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 尝试大关机器稍久再开机；检查gullwing switch等main contactor相关；检查床水平阻力，考虑观察掉电故障是否偶发一次 SAP Transfer : Dispatched for Onsite Support Next Step Action : JIN XIAO External Remarks : 19.10.2021 10:53:12 UTC (86038373)"SUAT. Ready for Clinical Use local time" : "2021-10-19 18:30" External Remarks : 19.10.2021 11:19:59 UTC (86038373)故障现象: CT扫描失败 服务内容: 按压前盖，重启机架 服务结果: 维修后,设备恢复正常. Internal Remarks : 19.10.2021 11:54:05 UTC (86038373) “RDI Accuracy: YES" $Q7: (Confirmed reported problem, CT scanning failed) $Q8: ( Check gantry log, found "S_MPS_MAIN_DRIVE_NOT_CLOSING" error) $Q9: ( Press gantry front cover and reboot CT gantry) $Q10: ( Test and inspection data refer to: PA. Device returned to full functionality) $Q11: Calibrated Tools name ( N/A) ^Current Tube installed date: N/A ^Current Tube usage: N/A ----------------------------------------------------------------------------- | 2 |  | Press gantry front cover and reboot CT gantry, system back to normal | [12 empty] | 118557515 |  | 1436279 |  |  | Investigation | [7 empty] | Engineering | [6 empty] | Front Cover | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Misalignment | Gantry-HW-Misalignment | Gantry-HW-Misalignment | 0 | Noise | Gantry-HW | Gantry-HW-Misalignment | No | 0 | 0 | 0 | 0 | 0 | No | y | y | 0 | 0 | 0 | y | y | y | y | y | #N/A
1466 |  |  | China | 广东省广州市第八人民医院（嘉禾院区） | 345032 | 4448700% |  | Wechat:02N/Ingenuity Core128/Elite China/air conditioning temperature is too high | BU Q&R Feedback: No SAP Customer Symptom Code: CENV Customer Symptom Code: Customer Environmental Concern Reason: Corrective Maintenance Engineer Failure Code: Environmental Cause Repair Activity Code: Instructed Customer Problem Reported by Customer : Wechat:02N/Ingenuity Core128/Elite China Safety Questions : Case: 0118550159 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Bai Jingwen Last updated by: Bai Jingwen T2 Activities : Troubleshooting Action: air conditioning，contact third party Repair Action: air conditioning，contact third party Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: N/A | 5 |  | NAC_air conditioning problem | [12 empty] | 118550159 |  | 1430864 |  |  | Closed - No Further Investigation | [7 empty] | NAC | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [26 empty] | #N/A | bar: No to NAC
1467 |  |  | United Kingdom | Philips | 40016 | 4448700% |  | adaptor was faulty, replaced the adaptor cable | Whilst installing the monitors for the first time, one of the minitors would not come up correctly. After swapping ports on the back of the Host PC and changing DP adaptors we found that the adaptor was faulty. This is also the case with the iQon on the same site. I believe other sites are also having the same issue during install. Not able to change the date in the boxes below. Retrospective complaint. 14th Oct 2021. FOR UKI REFERENCE ONLY: Feedback type: Technical issue UKI KMO: Paul Rothwell Submitter email: warren.thorpe@philips.com |  |  | DVI adapter | [14 empty] | 1428290 |  |  | Closed - Done |  | CTUDT00632173 | [5 empty] | Engineering | [6 empty] | DVI adapter | [5 empty] | Refresh |  |  | System issue | System | Console-HW | Adaptor | Console-HW-Adaptor | Console-HW-Adaptor | CTUDT00632173 | Adaptor | Console-HW | Console-HW-Adaptor | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1468 |  |  | Japan | Maki Hospital | 341017 | 4448600% |  | 04CS: KKI: Bed does not work >> S_MPS_MAIN_DRIVE_NOT_CLOSING szInfo: MP: Waiting for MDC to close ..> Restored by TEAL OFF / ON | 04CS: KKI: Bed does not work >> S_MPS_MAIN_DRIVE_NOT_CLOSING szInfo: MP: Waiting for MDC to close ..> Restored by TEAL OFF / ON . [Subject] 04CS: KKI: Noda: Bed does not work >> S_MPS_MAIN_DRIVE_NOT_CLOSING szInfo: MP: Waiting for MDC to close ..> Restored by TEAL OFF / ON [Internal Comments] [Log confirmation result] The following error has occurred. S_MPS_MAIN_DRIVE_NOT_CLOSING szInfo: MP: Waiting for MDC to close .. Since the key switch was turned off twice, the gantry power supply did not turn off, so I turned the key switch twice. It is probable that the gantry side was hung because the gantry did not start up normally after that. It seems to be a temporary phenomenon, so I will wait and see. [4/8 Elements] Q5 How did the test end ?: None Q6 Presence or absence of incorrect findings (misdiagnosis, mistreatment) due to problems: None Q7 Presence or absence of unintended unnecessary exposure: None Q8 Presence / absence of artifacts: None | 4 |  | Since the key switch was turned off twice, the gantry power supply did not turn off, so I turned the key switch twice. It is probable that the gantry side was hung because the gantry did not start up normally after that. It seems to be a temporary phenomenon, so I will wait and see. | [12 empty] | 118548042 |  | 1429673 |  |  | Closed - No Further Investigation |  | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Reboot | [5 empty] | Refresh |  |  | System issue | System | Gantry-SW | NoEnoughInfo | Gantry-SW-NoEnoughInfo | Gantry-SW-NoEnoughInfo |  |  |  |  | No | 0 | 0 | 0 | 0 | 0 | No | y | y | y | y | y | y | y | y | 0 | y | #N/A
1469 |  |  | China | Jilin Province Zhaonan City TCM Hospital | 345006 | 4448400% |  | 01/Ingenuity Core128/Elite China/Mouse fault | N/A Next Step Action: ASP jilinyiyang follow Transfer to sales handling, this case is valid within 90 days starts from this case created, once beyond of this period without any further steps, this case is considered invalid, customer refuses our service | 3 |  | out of warranty, customer refuses our service | [12 empty] | 118541561 |  | 1424965 |  |  | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Mouse | [5 empty] | Refresh |  |  | System issue | [26 empty] | #N/A | bar: No to NAC
1470 |  |  | China | 河南省新乡市凤泉区大块镇卫生院 | 345009 | 4448300% |  | 02N/Ingenuity Core128/Elite China/gantry noise | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity Core128/Elite China T2 Activities : Problem Description: 客户问题概述： 设备异响。 客户问题详细描述或故障现象： 设备异响，静止时也响 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 设备异响，静止时也响 远程诊断与测试结果： 怀疑探测器风扇 PRS TR状态： 不通，请工程师到场检查。 Repair Action: 建议现场维修步骤： 检查噪音来源，根据发现做相关的troubleshooting 建议备件12NC及名称： "BR64 Air Fan Replacement Kit 459800980091 Replaces BR AIR FAN REPLACEMENT KIT 455011509231; 000130" SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : PPD placement Kit 459800980091 1 已安排 36H到医院，FSE：郭海滨-86-18738187087 Safety Questions : Case: 0118534604 | 3 |  | $Q8:perform visiual and conform the noise was from the fan of the DMS. and the fan was wore out $Q9:replace the defective fans | [12 empty] | 118534604 |  | 1421695 |  |  | Closed - No Further Investigation |  | CTUDT00632110 | [5 empty] | Engineering | [6 empty] | DMS Fan | [5 empty] | Refresh |  |  | System issue | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | DMS Module | DMS-HW | DMS-HW-DMS Module | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1471 |  |  | China | 曹妃甸区医院 | 345016 | 4448300% |  | Wechat:02N/Ingenuity Core128/Elite China/CT box bed lowering button cannot be used | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : CT box bed lowering button cannot be used Safety Questions : Case: 0118534137 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: mo xueke Last updated by: mo xueke 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 其他类型故障。 客户问题详细描述或故障现象： CT box的降床的键不好用 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： CT box的降床的键不好用 远程诊断与测试结果： CT box的降床的键不好用 PRS TR状态： 无PRS。 Repair Action: 建议现场维修步骤： 检查CT box，如有必要进行更换 建议备件12NC及名称： 453567027195 CT box T2 Activities : Problem Description: 客户问题概述： 其他类型故障。 客户问题详细描述或故障现象： CT box的降床的键不好用 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： CT box的降床的键不好用 远程诊断与测试结果： CT box的降床的键不好用 PRS TR状态： 无PRS。 Repair Action: 建议现场维修步骤： 检查CT box，如有必要进行更换 建议备件12NC及名称： "CT BOX ASSEMBLY_BEST 459801705022" SAP Transfer : Dispatched for Onsite Support Next Step Action : PPD CT BOX ASSEMBLY_BEST 459801705022 1 已安排24H到医院，FSE：张翔-86-18600120460 SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | CT BOX ASSEMBLY_BEST 459801705022 | [12 empty] | 118534137 |  | 1421355 |  |  | Closed - No Further Investigation |  | CTUDT00632185 | [5 empty] | Engineering | [6 empty] | CT Box | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | CT Box | Gantry-HW-CT Box | Gantry-HW-CT Box | CTUDT00632185 | CT Box | Gantry-HW | Gantry-HW-CT Box | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1472 |  |  | China | 东南大学附属中大医院江北院区 | 345021 | 4448300% |  | 02n/Ingenuity Core128/Elite China/exposure failed | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02n/Ingenuity Core128/Elite China Safety Questions : Case: 0118532093 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 其他类型故障。 客户问题详细描述或故障现象： 不能曝光 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 不能曝光 远程诊断与测试结果： 不能曝光，机架报错，重启后还是不行 PRS TR状态： 不通，请工程师到场检查。 Repair Action: 建议现场维修步骤： 查看报错信息，指导医院正常使用注意事项 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support ----------------------------------------------------------------------------- | 2 |  | $Q7: (Confirmed reported problem,CIRS boot up failed and system report CIRS connection error) $Q8: (on site check CIRS statues，found during boot up，system can't detect CPU fans ) $Q9: (reconnect CPU fans and computer cover) | [12 empty] | 118532093 |  | 1420172 |  |  | Investigation |  | ShouldbeService | [5 empty] | TBD | [6 empty] | CPU Fan | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | n | y | y | 0 | y | #N/A
1473 |  |  | Germany | Gesundheit Nord gGmbH Klinikverbund | 40002 | 4448300% |  | In the case of the thorax, some of the lung windows are not sent to the PACS | Internal Remarks: Mr. Deerberg, question on voicemail, can he pull the error log here? Customer Comment: Did not reach the customer. Customer comment: Appointment discussed with Ms. Schichtl. Friday, November 5, 2021, approx. 9:00 am Problem Description by Engineer: Error description: Sniffer from Mr. Hallermann was attached to the customer During the thorax examinations, in part, the lung windows are not sent to the PACS Error has already been known for approx. 1 year. Please check remotely Error message: - Error occurred on: Yesterday and today 14.10.21 Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- Work Order #: WO-04546390 Sandner, Michael The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Follow-up required: Request for a technical service to copy the data from the laptop; no system time required; colleague should only copy the data from the laptop since that does not work remotely; then consultation with R. Kraschweski and send him the data; copy data regarding the following periods: 12.10. Bollhagen, Doris 11:09 am 14.10. Böttcher, Gisela 10:03 am Mureia, Correia 10:35 am Problem Description by Engineer: Error description: Sniffer from Mr. Hallermann was attached to the customer During the thorax examinations, in part, the lung windows are not sent to the PACS Error has already been known for approx. 1 year. Please check remotely Error message: - Error occurred on: Yesterday and today 14.10.21 Diagnostic performed by Engineer: return call Ms. Schichtl, discussion of the problem; call to Th. Halama - only voicemail; 3 cases since the day before yesterday, sometimes it works automatically, sometimes not, in approx. 30% of the case, there is no automatic transmission to the PACS. ----------------------------------------------------------------------------- | 4 |  | In the case of the thorax, some of the lung windows are not sent to the PACS Error has been known for about 1 year. 3 cases since the day before yesterday, sometimes it works automatically, sometimes not, about 30% does not go into PACS automatically | [12 empty] | 118534170 |  | 1417111 |  |  | Evaluation |  | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | PACS Transfer_Lung | [5 empty] | Refresh |  |  | System issue | System | Console-SW | NoEnoughInfo | Console-SW-NoEnoughInfo | Console-SW-NoEnoughInfo | CTUDT00632185 |  | Console-SW | Console-SW-NoEnoughInfo | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | #N/A
1474 |  |  | China | Jilin Province Songyuan City Central Hospital | 345008 | 4448200% |  | 02n/Ingenuity Core128/TheGantry power down | BU Q&R Feedback: No Reason: Corrective Maintenance Internal Remarks : 16.10.2021 15:28:30 UTC (86038306) "RDI Accuracy:Yes" $Q7: (Confirmed reported problem&estop open) $Q8: (review the system logger，the R2D error，the R2D damaged） $Q9: (replaced the R2D） $Q10: (Test and inspection data refer to: PA. Device returned to full functionality) $Q11 Calibrated Tools name (FLUKE287) S/N (16490094) Next Calibration date (10/19/2021) SAP Transfer:Case Transferred to SAP Problem Reported by Customer: 02n/Ingenuity Core128/Elite China SAP Transfer:Dispatched for On site Support T2 Activities:Problem Description: Summary of customer's problems: Gantry intermittently power off. It's a technology problem with the equipment itself: Yes. Troubleshooting Action: remote diagnosis and testing result: It's very difficult to determine the problem. We suspect the couch and signal brush.。。problem PRS TR status: it's not through, ask engineer for maintenance on site. Repair Action: Suggest onsite service procedures: It is recommended to have a field service engineer go on site to check the error log and determine the problem. Memo of Customer request: contact hospitalm when arrive Safety Questions:Case: 0118526792 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Weng Jinming Last updated by: Weng Jinming Next Step Action: PPD called to contact the field service engineer to confirm the alternate number 459800467341*1, shipped 24H, Chi Xiaodong - 86-18943690010 External Remarks: 14.10.2021 11:31:59 UTC (86038306) Service details: The device intermittently powers off, R2D malfunction, replace. Service outcome: device returns to normal working condition after service Other significant matters: None. ----------------------------------------------------------------------------- | 3 |  | replaced the R2D 459800467341 | [12 empty] | 118526792 |  | 1416266 |  |  | Investigation |  | CTUDT00630864 | [5 empty] | Engineering | [6 empty] | R2D | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | R2D | Gantry-HW-R2D | Gantry-HW-R2D | CTUDT00630864 | R2D | Gantry-HW | Gantry-HW-R2D | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1475 |  |  | China | Dongguan Donghua Hospital Co., Ltd. | 40059 | 4448100% |  | Image trasferring gets stuck; Had the customer reset the spooler to resolve. | BU Q&R Feedback: No SAP Customer Symptom Code: HANG Customer Symptom Code: Hang/Crash Reason: Corrective Maintenance Engineer Failure Code: Configuration Error Repair Activity Code: Configuration Change Problem Rep T2 Activities: Problem Description: Automatic sending of images gets stuck. Troubleshooting Action: Tried to have the customer reset spooler to reset. The customer said he will reset it at work on October 13th, then observe again. Repair Action: Tried to have the customer reset spooler to reset. The customer said he will reset it at work on October 13th, then observe. Test/Inspection: customer confirm can scan patient and image ok.functional Check PASSED. The device returned to full functionality. Test Equipment: N/A SAP Transfer:Case Transferred to SAP | 3 |  | reset spooler | [12 empty] | 118520957 |  | 1412285 |  |  | Closed - Done |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS Transfer | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1476 |  |  | China | Shandong Province Chest Hospital | 40025 | 4448100% |  | WeChat：02N/Ingenuity CT 728326/software hang up | T2 Activities:Problem Description: Summary of customer's problems: Other types of error. Detailed description of Customer's problems or malfunctions: SW stuck. It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: SW stuck. Remote troubleshoot and test result: There is still not enough memory. FSE requires more memory PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Increase memory 0 459801813461 16GB DDR4-2933 (2x8GB) RegRAM SAP Transfer:Case Transferred to SAP External Remarks: 13.10.2021 07:50:51 UTC (86038405) error: blue screen when turned on to login, software is crashing. Service details: add memory to host Z4. Service result: device returns to normal working condition after service. Next Step Action: huang zhaokun SAP Transfer:Dispatched for On site Support Next Step Action: PPD 459801813461*1 shipped, 24H Huang Zhaokun-86-18654536122 | 3 |  | Increase memory 459801813461 16GB DDR4-2933 (2x8GB) | [12 empty] | 118519183 |  | 1411148 |  |  | Investigation |  | Performance Slow | [5 empty] | Engineering | [6 empty] | Performance Slow | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Performance Slow | Console-SW-Performance Slow | Console-SW-Performance Slow | Performance Slow | Performance Slow | Console-SW | Console-SW-Performance Slow | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1477 |  |  | China | No.8 People's Hospital of Guangzhou City, Guangdong Province (Jiahe Branch) | 345032 | 4448100% |  | 02N/Ingenuity Core128/Elite China/IRS disconnected | T2 Activities:Problem Description: Customer consult the machine life span. Troubleshooting Action: inform customer the machine life span. Device is normal. Repair Action: Inform customer the machine life span, and the machine is normal. Test/Inspection: customer confirm can scan patient and image ok.functional Check PASSED. The device returned to full functionality. Test Equipment: N/A T2 Activities:Problem Description: Summary of customer's problems: The host cannot be connected to CIRS Detailed description of Customer's problems or malfunctions: Instruct customer, unplug and replug in hard disc, still the same It's a technology problem with the equipment itself: Yes. Troubleshooting Action: remote diagnosis and testing result: Suspect CIRS, hardware issue PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Suggest engineer to go to the field to check CIRS, and reinstall the software.Teach customer to use Philips WeChat to request service SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support External Remarks: 12.10.2021 05:28:32 UTC (86034061)"SUAT. Ready for Clinical Use local time" : "2021-10-12 14:30" Service content: CIRS can not start. Memory failure. Replace the memory and test, back to normal Service outcome: Device returns to normal working condition after service. ----------------------------------------------------------------------------- | 2 |  | replace memory and finish recommended test 服务内容: 重建计算机无法开机，内存故障，更换内存，测试恢复正常 服务结果: 服务完成后,设备恢复正常工作 | [12 empty] | 118518739 |  | 1410844 |  |  | Closed - No Further Investigation |  | MemoryReplacement | [5 empty] | Engineering | [6 empty] | Memory | [5 empty] | Refresh |  |  | System issue | System | Console-HW | MemoryReplacement | Console-HW-MemoryReplacement | Console-HW-MemoryReplacement | MemoryReplacement | MemoryReplacement | Console-HW | Console-HW-MemoryReplacement | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1478 |  |  | China | Fujian Provincial Hospital | 345011 | 4448100% |  | 02N/Ingenuity Core128/Elite China | BU Q&R Feedback: No Reason: Corrective Maintenance External Remarks: 15.10.2021 02:31:02 UTC (86036497) Service details: CT software occasionally lagging.Check on the field to confirm it's the application software which is abnormal.Replace hard disc, install OS and app, import backup file, and complete corresponding configuration service result: device returns to normal working condition after service Internal Remarks : 15.10.2021 02:31:05 UTC (86036497)RDI Accuracy: Yes $Q7: (Confirmed reported problem. The Host Apps run slowly.) $Q8: (On site check and find the OS and Apps are abnormal.) $Q9: (Replace the Host hard disk and re-install OS & Apps. And then do restore.) $Q10: (Test and inspection data refer to: PA.+Device returned to full functionality.) $Q11: Calibrated Tools name (FLUKE 287C Multimeter) S/N (16490128) NextCalibration date (10/08/2022) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date: NA ^Current Tube usage: NA Problem Reported by Customer: 02N/Ingenuity Core128/Elite China Safety Questions:Case: 0118518507 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Guoping Last updated by: Liu Guoping 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer's problems: SW stuck. Detailed description of Customer's problems or malfunctions: SW stuck. It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: SW stuck. " Remote troubleshoot and test result: Reboot no avail. Need to replace the hard disc and re install the software PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Replace the hard disk and reinstall the software. We recommend spare parts 12NC and names: HP TurboDrive 512GB TLC Z4 SSD w/adapter 459801652541 SAP Transfer:Case Transferred to SAP External Remarks: 15.10.2021 00:49:15 UTC (86036224) Service details: CT software occasionally lagging.Check on the field to confirm it's the application software which is abnormal.Replace hard disc, install OS and app, import backup file, and complete corresponding configuration service result: device returns to normal working condition after service Internal Remarks : 15.10.2021 00:49:18 UTC (86036224) RDI Accuracy: Yes $Q7: (Confirmed reported problem. The Host Apps run slowly.) $Q8: (On site check and find the OS and Apps are abnormal.) $Q9: (Replace the Host hard disk and re-install OS & Apps. And then do restore.) $Q10: (Test and inspection data refer to: PA.+Device returned to full functionality.) $Q11: Calibrated Tools name (FLUKE 287C Multimeter) S/N (16490128) NextCalibration date (10/08/2022) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date: NA ^Current Tube usage: NA SAP Transfer:Dispatched for On site Support Next Step Action: PPD 459801652541*1 shipped, 24H Ding Shi duo-86-13509388496 Next Step Action: Ding shiduo | 3 |  | Replace the Host hard disk(459801652541) and re-install OS & Apps. And then do restore. | [12 empty] | 118518507 |  | 1410696 |  |  | Closed - No Further Investigation |  | HardDiskReplacement | [5 empty] | Engineering | [6 empty] | SW slow | [5 empty] | Refresh |  |  | System issue | System | Console-HW | HardDiskReplacement | Console-HW-HardDiskReplacement | Console-HW-HardDiskReplacement | HardDiskReplacement | HardDiskReplacement | Console-HW | Console-HW-HardDiskReplacement | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1479 |  |  | Peru | BANCO BBVA PERU | 40034 | 4448100% |  | Boton de CTbox no funciona | Work Order #: WO-04534498 [REDACTED] The system meets the specification for the performed service and is returned to use. Problem Description by Engineer: Customer Function/Role: Field Service Engineer Device Use: Diagnostic equipment Expected and Unexpected Results: CTBox button does not allow to move the table down. User Impact: NA Patient Impact: NA Current Software Version: 4.1.10 Diagnostic performed by Engineer: CTBox button to move the table down does not work. CTBox needs to be replaced. Follow Up Required: [REDACTED] Contact: [REDACTED] ----------------------------------------------------------------------------- | 5 |  | Product/Part - CT BOX ASSEMBLY_BEST Part number - 459801705022 | [12 empty] | 118524225 |  | 1410393 |  |  | Closed - No Further Investigation |  | CTUDT00632185 | [5 empty] | Engineering | [6 empty] | CT Box | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | CT Box | Gantry-HW-CT Box | Gantry-HW-CT Box | CTUDT00632185 | CT Box | Gantry-HW | Gantry-HW-CT Box | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1480 |  |  | Japan | Maki Hospital | 341017 | 4447900% |  | 03DM: RS: Error occurred during test scan >> Rescan is possible >> S_GEN_FAULT_ANODE_TUBE_ARC | Work Order #: WO-04520283 Satake, Rui Problem Description by Engineer: Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer: N/A ----------------------------------------------------------------------------- | 4 |  | LTC is required | [12 empty] | 118512793 |  | 1406421 |  |  | Closed - No Further Investigation |  | CTUDT00630085 | [5 empty] | Engineering | [6 empty] | Tube Arc | [5 empty] | Refresh |  |  | System issue | System | IC-HW | Tub Arc | IC-HW-Tub Arc | IC-HW-Tub Arc | CTUDT00630085 | Tub Arc | IC-HW | IC-HW-Tub Arc | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1481 |  |  | United States | Odessa Regional Medical Center | 40029 | 4447900% |  | 1001CDN81 Issues with QC. | The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: System will not initiate User Impact: System unavailable for use Patient Impact: Current Software Version: Diagnostic performed by Engineer : checked system and verified issue. ----------------------------------------------------------------------------- Work Order #: WO-04519987 Mark Finnigan The system meets the specification for the performed service and is returned to use Follow Up Required : Spoke with Mia on the customer callback. She is not reporting any issues with artifacts but she is stating that she is having s tough time getting QC to run. It sounds like the phantom holder needs to be adjusted. Make sure that it is level and centered in bith axis and not skewed from the top perspective. Internal Comments : The customer originally called in Sunday with a complaint about a possible ring. I asked that they do an air calibration but they were too busy at that time. I was not able to get a hold of them latter in the day. Called back on Monday and spoke with Mia. She believes that an air calibration was never done and that she was too busy to do one today. She added that they have been having issues running QC as it fails. No other details are provided. I would start by checking the phantom holder alignment and leveling. Their PRS is not working so no further remore assistance can be rendered. The site is connected but all ports are blocked preventing any access. Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: System will not initiate User Impact: System unavailable for use Patient Impact: Current Software Version: Diagnostic performed by Engineer : Spoke with the CT tech. They could not get the system to initiate. Had them power cycle the machine and it worked properly after the restart. Diagnostic performed by Engineer : Spoke with Mia on the customer callback. She is not reporting any issues with artifacts but she is stating that she is having s tough time getting QC to run. It sounds like the phantom holder needs to be adjusted. Make sure that it is level and centered in bith axis and not skewed from the top perspective. ----------------------------------------------------------------------------- | 2 |  | power cycle the machine and it worked properly after the restart. | [12 empty] | 118512628 |  | 1406026 |  |  | Closed - No Further Investigation |  | ShouldbeCustomer | [5 empty] | Engineering | [6 empty] | Reboot | [5 empty] | Refresh |  |  | System issue | System | Console-SW | ShouldbeCustomer | Console-SW-ShouldbeCustomer | Console-SW-ShouldbeCustomer | ShouldbeCustomer |  | Console-SW | Console-SW-ShouldbeCustomer | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | #N/A
1482 |  |  | China | 北京市中西医结合医院 | 40032 | 4447800% |  | 02N/Ingenuity CT 728326/network change | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity CT 728326 SAP Transfer : Dispatched for Onsite Support SAP Transfer : Case Transferred to SAP Safety Questions : Case: 0118510577 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 其他类型故障。 客户问题详细描述或故障现象： worklist 修改参数 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： worklist 修改参数 远程诊断与测试结果： 需要工程师到场帮忙修改worklist参数 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 修改worklist参数 ----------------------------------------------------------------------------- | 3 |  | NAC_Change network configuration | [12 empty] | 118510577 |  | 1403271 |  |  | Closed - No Further Investigation | [7 empty] | NAC | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [26 empty] | #N/A | bar: No to NAC
1483 |  |  | China | Wuhan Province Wuhan Ren’ai Hospital | 345004 | 4447800% |  | 05K FOR IB/Ingenuity Core128/Elite China | T2 Activities:Problem Description: Summary of customer's problems: Images transmission failure. Detailed description of Customer's problems or malfunctions: The images can't be transmitted to Shenfei Yun and PACS It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: "The images can't be transmitted to Shenfei Yun and PACS " Remote troubleshoot and test result: PACS is not through, reset invalid PRS TR status: it's not through, ask engineer for maintenance on site. Repair Action: Suggest onsite service procedures: Check network, reset spooler SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Next Step Action: Customer called in again to push, informed FSE Hu to respond to customer External Remarks: 10.10.2021 15:01:36 UTC (86036472) Service details: [Image transmission to Shenfei Yun and PACS failed. Re-setup IP, restore after connecting to hospital network. After service, device is restored] service result: [device returns to normal working condition after service] Internal Remarks : 10.10.2021 15:11:48 UTC (86036472) RDI Accuracy: Yes $Q7: (Confirm reported problem & Image cannot be transmitted to shenfeiyun and PACS） $Q8: (The scene inspection found that the image could not be transmitted. Due to the doctor's feedback that the film could be printed, the data was suspected to be stuck. Spooler reset was executed, but the transmission still failed after the restart, and the LAN was connected The config test network was also abnormal. The network IP was detected and it was found that the second IP was set by the hospital, and the second IP was also lost. The information department was asked to re-assign the address, and the Settings were restored） $Q9: ( Reset the IP address to connect to the hospital network then the system recovery) $Q10: (Test and inspection data refer to PA ,Device return to full functionality) $Q11: Calibrated Tools name :N/A For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date:N/A ^Current Tube usage: scan second_N/A ----------------------------------------------------------------------------- | 3 |  | Reset the IP address to connect to the hospital network then the system recovery | [12 empty] | 118510268 |  | 1403172 |  |  | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | PACS Config | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | n | y | y | y | y | #N/A
1484 |  |  | China | Harbin City Sen Hai Hospital | 345007 | 4447600% |  | Wechat:05K for IB/Ingenuity Core128/Elite China/High pressure injector failure | Problem Reported by Customer : Fan inoperative Next Step Action: The console enhanced small computer's fan is not working. Remote Service Engineer (RSE) Weng Jinming Next Step Action: Remote Service Engineer (RSE) feedback: need to dispatch labor Next Step Action: Remote service engineer (RSE) feedback: Sorry, the hospital's description is wrong, it's a high-pressure injector issue, contact the third party. Next Step Action: Called Teacher Zhu back at the hospital to consult about the manufacturer's information. Teacher Zhu informed that the manufacturer has been contacted, and now it has been fixed. | 3 |  | NAC_Injector problem | [12 empty] | 118496929 |  | 1396890 |  |  | Evaluation | [7 empty] | NAC | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [26 empty] | #N/A | bar: No to NAC
1485 |  |  | Poland | Government Agency. | 341002 | 4447700% |  | Przeniesienie aparatu z kontenera do budynku głównego | Integration:Check the settlement rule-Entry 1 in table OPERATION (method CHANGE) was not found-Error during processing of BAPI methods-MPG is present on the product code : P7 Integration: Confirmation of order 54273802 saved. Integration:According to factory calendar 20211023 (Posting date) is not a workday (check entry) ----------------------------------------------------------------------------- Work Order #: WO-04511999 15-Oct-2021 Mr. Mariusz Pachla The system meets the specification for the performed service and is returned to use Tools & Equipment : MULTIMETER,2022-07-24 00:00:00 Problem Description by Engineer: Moving the device from the container to the main building. Resolution: Moving the device from the container to the main building. ----------------------------------------------------------------------------- Work Order #: WO-04511990 22-Oct-2021 Gerard Dabrowka The system meets the specification for the performed service and is returned to use Tools & Equipment : TORQUE WRENCH,2021-10-26 00:00:00 TORQUE WRENCH,2021-10-26 00:00:00 Problem Description by Engineer: order to transfer the device. Diagnostic performed by Engineer: Transfer supervision, disassembly of the device, installation in a new location. ----------------------------------------------------------------------------- |  |  | 2021-10-08 | [12 empty] | 118506501 |  | 1402011 |  |  | Closed - Non-Complaint | [7 empty] | 2021-10-08 | [6 empty] | 2021-10-09 18:22 | [5 empty] | Refresh |  |  | System issue | [26 empty] | #N/A | bar: No to NAC
1486 |  |  | China | Wuhan Province Wuhan Ren’ai Hospital | 345004 | 4447700% |  | 05K for IB/Ingenuity Core128/Elite China/ | BU Q&R Feedback: No Reason: Corrective Maintenance External Remarks: 15.10.2021 06:57:03 UTC (86036493) Service details: [Doctor forgot password, failed to log in to scan software. Restored to normal after resetting ct account password. After service, device is restored] service result: [device returns to normal working condition after service] Internal Remarks : 15.10.2021 06:57:06 UTC (86036493)RDI Accuracy: Yes $Q7: (Confirm reported problem: The doctor forgot his password and couldn't log into the scanning software) $Q8: (When I tried to log in with IT account, I found that the password was also wrong, so I went in to reset the password in maintenance mode. However, I happened to enter the account incorrectly for many times, and the account was locked. After unlocking, I reset the CT account password） $Q9: ( Reset the CT account password then the system recovery) $Q10: (Test and inspection data refer to PA ,Device return to full functionality) $Q11: Calibrated Tools name :N/A For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date:N/A ^Current Tube usage: scan second_N/A Problem Reported by Customer: 05K for IB/Ingenuity Core128/Elite China Safety Questions:Case: 0118505272 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Guoping Last updated by: Liu Guoping 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer's problems: Other types of error. Detailed description of Customer's problems or malfunctions: Forgot password and failed to log in It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: Forgot password and failed to log in " Remote troubleshoot and test result: PRS is not through, customer doesn't know IT password PRS TR status: it's not through, ask engineer for maintenance on site. Repair Action: Suggest onsite service procedures: 1. reset password; 2. replace IT password, and ask customer to keep the account safe SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Next Step Action: hu fangtao External Remarks: 09.10.2021 09:43:01 UTC (86036472)"SUAT. Ready for Clinical Use local time" : "2021-10-08 23:00" External Remarks: 10.10.2021 15:15:18 UTC (86036472) Service details: [Doctor forgot password, failed to log in to scan software. Restored to normal after resetting ct account password. After service, device is restored] service result: [device returns to normal working condition after service] Internal Remarks : 10.10.2021 15:18:21 UTC (86036472) RDI Accuracy: Yes $Q7: (Confirm reported problem & The doctor forgot his password and couldn't log into the scanning software） $Q8: (When I tried to log in with IT account, I found that the password was also wrong, so I went in to reset the password in maintenance mode. However, I happened to enter the account incorrectly for many times, and the account was locked. After unlocking, I reset the CT account password） $Q9: ( Reset the CT account password then the system recovery) $Q10: (Test and inspection data refer to PA ,Device return to full functionality) $Q11: Calibrated Tools name :N/A For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date:N/A ^Current Tube usage: scan second_N/A ----------------------------------------------------------------------------- |  |  | 2021-10-08 | [12 empty] | 118505272 |  | 1401508 |  |  | Closed - Non-Complaint | [7 empty] | 2021-10-08 | [6 empty] | 2021-10-09 14:47 | [5 empty] | Refresh |  |  | System issue | [26 empty] | #N/A | bar: No to NAC
1487 |  |  | Australia | Maia Financial Pty Limited | 341000 | 4447600% |  | PACS Changes | Work Order #: WO-04522710 20-Oct-2021 Ruben MARQUEZ The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: Radiographer Device Use: in use Expected and Unexpected Results: Change Network DICOM nodes. User Impact: NA Patient Impact: NA Current Software Version: 4.1.10 Resolution : Assisted David in creating new DICOM nodes and amending local AET changes on the system. Tested all destinations - all passed. ----------------------------------------------------------------------------- |  |  | 2021-10-07 | [12 empty] | 118499748 |  | 1398575 |  |  | Closed - Non-Complaint | [7 empty] | 2021-10-07 | [6 empty] | 2021-10-08 20:37 | [5 empty] | Refresh |  |  | System issue | [26 empty] | #N/A | bar: No to NAC
1488 |  |  | Turkey | ADIYAMAN OZEL SAGINSTURTAS SAN VE TIC A.S | 341006 | 4447500% |  | Image artifact, cleaned the UHR filter | Resolution : The system was returned to the use in functional condition after cleaning the UHR filter and air calibration. ----------------------------------------------------------------------------- Work Order #: WO-04495201 Salih Mendi The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Problem description by engineer : Artefact on images Information to support the complaint handling process Customer Function/Role: Not provided Device Use: Not provided Expected and Unexpected results: Not provided Patient Impact/Outcome: Not provided User Impact: Not provided Current Software Version: Not provided. Diagnostic performed by Engineer : There are artefacts on the images using the UHR filter due to contrasting agent leak/contamination. Internal Comments : Error # and/or description of error : Poor image quality caused from contrast on UHR filter. Repair Action : The problem occurs only with UHR filter imaging. The contrasting agent should be removed from the UHR filter. Resolution : The user was remotely supported for the image quality problem. In diagnosis, we found that the artefact was caused by contrasting agent leak. The site should be visited to clean the UHR filter. If replacement doesn't solve the problem, UHR filter should be ordered for replacement. The system is partially up. Follow Up Required : Error # and/or description of error : Poor image quality caused from contrast on UHR filter. Repair Action : The problem occurs only with UHR filter imaging. The contrasting agent should be removed from the UHR filter. | 3 |  | artefact was caused by contrasting agent leak, Poor image quality caused from contrast on UHR filter. | [12 empty] | 118493421 |  | 1389887 |  |  | Closed - Done | [7 empty] | Others | [6 empty] | Contrast | [5 empty] | Refresh |  |  | System issue | [26 empty] | #N/A | bar: No to NAC
1489 |  |  | Japan | Maki Hospital | 341017 | 4447500% |  | 02CS:KK:固定帯劣化。 | Work Order #: WO-04491841 ----------------------------------------------------------------------------- Work Order #: WO-04491836 Kojima, Kenichi The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Resolution : Restraint strap is deteriorated and he asks for a new one. Problem Description by Engineer : Deteriorated restraint strap. ----------------------------------------------------------------------------- | 4 |  | Parts Consumed Description: 2.00000,453567009951,L-6" PATIENT RESTRAINT STRAP/ 42" | [12 empty] | 118490725 |  | 1387963 |  |  | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Strap | [5 empty] | Refresh |  |  | System issue | [26 empty] | #N/A | bar: No to NAC
1490 |  |  | China | Shandong Province Tai'an City T.C.M. No. 2 Hospital | 345037 | 4447400% |  | 02N/Ingenuity Core128/Elite China/Can't rebuild | Problem Reported by Customer: can't rebuild Next Step Action: Remote Service Engineer (RSE) Cao Liu Next Step Action: Calling back was successful. SAP Transfer:Case Transferred to SAP External Remarks : 06.10.2021 13:44:58 UTC (86038373) Fault Description: There are fewer images reconstructed Service content: Clean and unplug GPU graphic card Service outcome: Device returns to normal working condition after servicing. Internal Remarks : 07.10.2021 13:43:33 UTC (86038373) “RDI Accuracy: YES" $Q7: (Confirmed reported problem, CT image reconstruction failed ) $Q8: (Do CT image reconstruction manually, found it was also failed) $Q9: ( Clean and reseat GPU card) $Q10: (Test and inspection data refer to: PA. Device returned to full functionality) $Q11: Calibrated Tools name ( N/A) ^Current Tube installed date: N/A ^Current Tube usage: N/A Next Step Action:FSE Tao Chen 018560023293 call in and disp him Unable to rebuild. ----------------------------------------------------------------------------- | 3 |  | Clean and reseat GPU card | [12 empty] | 118484594 |  | 1388065 |  |  | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | GPU | [5 empty] | Refresh |  |  | System issue | [26 empty] | #N/A | bar: No to NAC
1491 |  |  | Japan | Maki Hospital | 341017 | 4447300% |  | 02CS:KK:固定帯劣化。 | Work Order #: WO-04474261 ----------------------------------------------------------------------------- Work Order #: WO-04474233 Kojima, Kenichi The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Deteriorated restraint strap. Resolution: The restraint strap is deteriorated, so he wants us to send a new one. ----------------------------------------------------------------------------- | 4 |  | 453566451921,14 INCH RESTRAINT STRAP | [12 empty] | 118477364 |  | 1381004 |  |  | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Strap | [5 empty] | Refresh |  |  | System issue | [26 empty] | #N/A | bar: No to NAC
1492 |  |  | United States | Unity Health Systems | 40037 | 4447400% |  | Monthly QA | Integration : Confirmation of order 54257308 saved ----------------------------------------------------------------------------- Work Order #: WO-04487915 05-Oct-2021 Taylor, Durrance The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: tech Device Use: scheduled event Expected and Unexpected Results: monthly qa User Impact: none Patient Impact: none Current Software Version: 4.1.10 Resolution : Monthly QA completed ----------------------------------------------------------------------------- |  |  | 2021-10-05 | [12 empty] | 118487789 |  | 1387210 |  |  | Closed - Non-Complaint | [7 empty] | 2021-10-05 | [6 empty] | 2021-10-06 03:17 | [5 empty] | Refresh |  |  | System issue | [26 empty] | #N/A | bar: No to NAC
1493 |  |  | United States | Leesburg Regional Medical Center | 40022 | 4447200% |  | 0602CDN84 Left side UP/Down button hard to push. | Work Order #: WO-04473095 05-Oct-2021 Mayhew Bruce The system meets the specification for the performed service and is returned to use Resolution : Left UP/Down buttons hard to push. No Patient Involved Expected and Unexpected Results:Left UP/Down buttons hard to push. User Impact: Can't scan Patient Impact: None Current Software Version: 4.1.10 Resolution : Replaced left panel on front of gantry and tested successfully. Problem Description by Engineer : Front left panel vertical up button is difficult to press. Parts Consumed : 1.00000,459801484991,CONTROL PANEL ASSEMBLY-FRONT LEFT ----------------------------------------------------------------------------- Work Order #: WO-04490158 05-Oct-2021 Denning, Dale The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : Helped replace the controller and tested out the system. ----------------------------------------------------------------------------- Work Order #: WO-04473086 Miller, John M The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Follow Up Required : Errors: Control Panel Right side UP/DN hard to press. Suggest checking cabling connections and/or replace Control Panel. 459800484511 - CONTROL PANEL ASSEMBLY-FRONT RH Internal Comments : Errors: Control Panel Right side UP/DN hard to press. Suggest checking cabling connections and/or replace Control Panel. 459800484511 - CONTROL PANEL ASSEMBLY-FRONT RH Diagnostic performed by Engineer : Collect information from customer and view log files. Problem Description by Engineer : Left UP/Down buttons hard to push. No Patient Involved Expected and Unexpected Results:Left UP/Down buttons hard to push. User Impact: Can't scan Patient Impact: None Current Software Version: 4.1.10 ----------------------------------------------------------------------------- | 3 |  | 459800484511 - CONTROL PANEL ASSEMBLY-FRONT RH | [12 empty] | 118476542 |  | 1380784 |  |  | Closed - No Further Investigation |  | CTUDT00632181 | [5 empty] | Engineering | [6 empty] | GCP | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Gantry Panel | Gantry-HW-Gantry Panel | Gantry-HW-Gantry Panel | CTUDT00632181 | Control Panel | Gantry-HW | Gantry-HW-Gantry Panel | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1494 |  |  | Chile | BANCO DE CHILE | 341005 | 4447300% |  | CL1688-FALTA CABLE SG PARA MONITOR | Integration:Confirmation of order 54251519 saved ----------------------------------------------------------------------------- Work Order #: WO-04480200 Parts Consumed : 1.00000,455012301901,ECG Monitor Trigger 7M Ca ----------------------------------------------------------------------------- Work Order #: WO-04479707 [REDACTED] The system meets the specification for the performed service and is returned to use. Internal Comments : Error # and/or description of error : NA Repair Action: NA Follow Up Required: Error Number and/or Description: NA Repair Action: NA Technician's Name: [REDACTED] Dispatch to: CLINICA AVANSALUD PROVIDENCIA SA Address: Salvador 100 Providencia Contact Person: [REDACTED] Diagnostic performed by Engineer: The customer does not find the interconnection cable between the ECG monitor and the Gantry. Problem Description by Engineer : Problem description by engineer : The customer does not find the connection cable between the ECG monitor and the Gantry. When was the first occurrence? Within... : 04-Oct-2021 09:30:00 How often is issue occurring? : Continuously Malfunction area : Accessories Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected results: Cardiology studies cannot be performed. Patient Impact/Outcome: Cardiology studies cannot be performed. User Impact: Cardiology studies cannot be performed. Current Software Version: 4 ----------------------------------------------------------------------------- |  |  | 2021-10-04 | [12 empty] | 118481906 |  | 1382414 |  |  | Closed - Non-Complaint | [7 empty] | 2021-10-04 | [6 empty] | 2021-10-04 23:12 | [5 empty] | Refresh |  |  | System issue | [26 empty] | #N/A | bar: No to NAC
1495 |  |  | United States | Philips Medical Capital | 40053 | 4447000% |  | 0699CDN80 Squeak noise during gantry rotation | Internal Remarks : FADB: (FSE Aware of Double Booking) Title: Internal Remarks Initials: Twila Actions Taken: Dispatched WO to FSE/Fairchild, Andrew for 10/4/21 @ 1pm EST. Outcome: Dispatched WO per PFSE's Request (McDowell Preston) IP: 88993323 Next Steps: None **************************** Note: I called PFSE/McDowell Preston - He stated to Dispatch WO to FSE/Fairchild, Andrew - Being they're already aware of Customer's issue Integration : Confirmation of order 54243944 saved Integration : Confirmation of order 54254068 saved ----------------------------------------------------------------------------- Work Order #: WO-04483938 05-Oct-2021 Parts Consumed : 1.00000,453566492401,POLY-VEE BELT ----------------------------------------------------------------------------- Work Order #: WO-04466982 05-Oct-2021 Fairchild, Andrew The system meets the specification for the performed service and is returned to use Follow Up Required : Belt replacement next PM Resolution : Replaced tensioner, recommend a new belt to be installed next PM Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: 4.1.10 Parts Consumed : 1.00000,453567003451,TENSIONER ASSY ----------------------------------------------------------------------------- Work Order #: WO-04483594 05-Oct-2021 McDowell Preston The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : replaced the belt tensioner and adjusted the bracket for smooth operation. verified no noise, and also performed gantry rotor balance. System fully operational at this time. ----------------------------------------------------------------------------- | 3 |  | replaced the belt tensioner 453567003451,TENSIONER ASSY and adjusted the bracket for smooth operation | [12 empty] | 118471860 |  | 1379945 |  |  | Closed - No Further Investigation |  | ECR-077012 | [5 empty] | Engineering | [6 empty] | Belt Tensioner | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Belt | Gantry-HW-Belt | Gantry-HW-Belt | ECR-077012 | Belt | Gantry-HW | Gantry-HW-Belt | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1496 |  |  | Japan | Maki Hospital | 341017 | 4447000% |  | 03CM:KK:Short TC中に停止。再スタートで完了。>>S_GEN_FAULT_ANODE_TUBE_ARC | Integration : Confirmation of order 54242329 saved Integration:According to factory calendar 20211009 (Posting date) is not a workday (check entry) ----------------------------------------------------------------------------- Work Order #: WO-04465091 Kojima, Kenichi The system meets the specification for the performed service and is returned to use. Resolution: S_GEN_FAULT_ANODE_TUBE_ARC had occurred. He restarted it and it completed. Problem Description by Engineer : Short TC stopped partway through.Completed after restart.>> S_GEN_FAULT_ANODE_TUBE_ARC Internal Comments: [Subject] 03CM:KK:Mr. Hoshi: Short TC stopped partway through.Completed after restart.>> S_GEN_FAULT_ANODE_TUBE_ARC [4/8 Elements] Q5. How did the examination end? None. Q6. Did any mistaken findings (mistaken diagnosis or treatment) result from the problem? No. Q7. Was there any unintended unnecessary exposure? No Q8 Artifacts? None ----------------------------------------------------------------------------- | 4 |  | Short TC stopped partway through.Completed after restart. | [12 empty] | 118470518 |  | 1379514 |  |  | Closed - No Further Investigation |  | CTUDT00630085 | [5 empty] | Engineering | [6 empty] | Tube Arc | [5 empty] | Refresh |  |  | System issue | System | IC-HW | Tub Arc | IC-HW-Tub Arc | IC-HW-Tub Arc | CTUDT00630085 | Tub Arc | IC-HW | IC-HW-Tub Arc | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1497 |  |  | Puerto Rico | HOSPITAL METROPOLITANO SAN GERMAN | 40017 | 4447000% |  | It does not perform studies | 07-Oct-2021 Hernandez, Alvin The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Problem description by engineer : Gantry error. Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: No gantry errors. Patient Impact/Outcome: System is down. User Impact: System is down. Current Software Version: 4.1.10 Diagnostic performed by Engineer : The filament circuit on the cathode power module is defective. This is a complete assembly and has to be replaced. Parts Consumed : 1.00000,453567560002,Spellman 80 KW RoHS CPM Assy Resolution : The cathode power module was delivered to the hospital facilities. The CPM module was installed and calibrated. All the calibrations concerning the generator replacement were performed. A new system backup was made. The CT was tested and left fully operational. ----------------------------------------------------------------------------- Work Order #: WO-04480516 06-Oct-2021 Pinto-Rodriguez, Felipe The system meets the specification for the performed service and is returned to use Resolution : The Cathode Power Module has been replaced. Problem Description by Engineer : Filament errors Diagnostic performed by Engineer : There are Cathode Power Module filament errors. ----------------------------------------------------------------------------- Work Order #: WO-04468193 01-Oct-2021 Hernandez, Alvin The system does not meet the specification for the performed service and is not returned to use Tools & Equipment : MULTIMETER,2022-04-30 00:00:00 Follow Up Required : The Cathode Power Module needs to be replaced. The Cathode has a filament problem. Diagnostic performed by Engineer : The following actions were done: 1. I create a Bug report from the error events data. 2. Using the Can-Trace tool I evaluate the problem in real-time. Can-Trace reported filament fault. 3. The ATP reports three errors: - Filament fault. - No filament current. - No filament feedback. 4. I did an X-ray tube inspection: -The Cathode module had an electric spike mark. - The high tension cable was visually verified. The Cables are ok. 5. The X-ray tube filament was measured. - The Filament is not open. - Filament measures 0.29 ohms. The filament is OK. - The four-pin connections on each side of the high tension cable are not internally open. They pass the conductivity test. 6. The No-Load Test was done. The test failed. This means that the X-ray tube is not the problem. The conclusion is that the Cathode Power Module is defective. This error is very likely to be caused by many incoming power failures during the week. The Cathode Power Module needs to be replaced. Recommendations: It is greatly advised to connect the Brilliance Ingenuity CT to a UPS system. These will prevent the CT from damage and downtimes caused by the power failures disruptions. Resolution : The Power Cathode Module was ordered. It will be delivered to the hospital facilities. The Part information is the following: Part Desc: Spellman 80 KW RoHS CPM Assy. Part Num: 453567560002. Deliver to Hospital Metro San German. As soon as the part arrives I will be there to install it. Problem Description by Engineer : Problem description by engineer : Gantry error. Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: No gantry errors. Patient Impact/Outcome: System is down. User Impact: System is down. Current Software Version: 4.1.10 ----------------------------------------------------------------------------- Work Order #: WO-04467269 Valenzuela Velasquez, Eduardo Ariel The system does not meet the specification for the performed service and is not returned to use Follow Up Required : Error # and/or description of error : S_RHOST_HVG_HSS_GENERATOR_NON_ARC_FAULT Repair Action : - Check log file in detail Technician Name : Valenzuela Velasquez, Eduardo Ariel Problem Description by Engineer : Problem description by engineer : Gantry error. Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: No gantry errors. Patient Impact/Outcome: System is down. User Impact: System is down. Current Software Version: 4.1.10 Diagnostic performed by Engineer : A gantry error is present and the tube warm up is not possible. Remote access is not working properly. The case is sent to planning. Internal Comments : Error # and/or description of error : S_RHOST_HVG_HSS_GENERATOR_NON_ARC_FAULT Repair Action : - Check log file in detail ----------------------------------------------------------------------------- | 2 |  | The filament circuit on the cathode power module is defective. This is a complete assembly and has to be replaced. Parts Consumed : 1.00000,453567560002,Spellman 80 KW RoHS CPM Assy | [12 empty] | 118472055 |  | 1378117 |  |  | Closed - No Further Investigation |  | TBD-APM/CPM | [5 empty] | Engineering | [6 empty] | CPM | [5 empty] | Refresh |  |  | System issue | System | IC-HW | APM/CPM | IC-HW-APM/CPM | IC-HW-APM/CPM | TBD-APM/CPM | APM/CPM | IC-HW | IC-HW-APM/CPM | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1498 |  |  | Japan | Independent administrative institution National Hospital Organization | 341013 | 4447000% |  | AF:固定帯交換 | Integration : Confirmation of order 54241525 saved ----------------------------------------------------------------------------- Work Order #: WO-04464183 06-Oct-2021 Yamamoto, Yuko The system meets the specification for the performed service and is returned to use. Regulatory Questions : N/A Parts Consumed: 2.00000,453567009961,L-14" PATIENT RESTRAINT STRAP/ 42" 2.00000,453567009951,L-6" PATIENT RESTRAINT STRAP/ 42" Internal Comments: [CX related] Did you agree on how to complete the report with the person in charge of devices? <Select one of the following> Yes. I completed normal reporting that day. [Work report (paper) submission] Did you submit a work report on paper? Yes. Q1. How did the examination finish? Select from the following: *Because the answer is No for Q1, not applicable Q2. Were there any mistaken findings (misdiagnosis, mistaken treatment) due to the problem? None. Q3. Were there any artifacts? None. Q4. Was there any unintended unnecessary exposure? None. Problem Description by Engineer: Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : I replaced the patient restraint strap. October 6, 2021 / I checked with the person in charge Mr. Yamamoto. ----------------------------------------------------------------------------- |  |  | 2021-10-01 | [12 empty] | 118469809 |  | 1379329 |  |  | Closed - Non-Complaint | [7 empty] | 2021-10-01 | [6 empty] | 2021-10-02 16:22 | [5 empty] | Refresh |  |  | System issue | [26 empty] | #N/A | bar: No to NAC
1499 |  |  | United States | Adventhealth Zephyrhills | 300106 | 4449800% |  | ECM password expired | CT SWare Ver = 4.1.10 Clin Apps req'ed RSE support ~ Exam Card Mgr password expired and cust cant reset password. RSE advised to login with HCF define Admin password, then reset ECMgr password. Cust says they do not have Admin password. FSE please assist customer reseting ECMgr password ~ educate and assist customer so they understand password system and have their HCF specific password noted where they can be found. Problem Description by Engineer : Customer Function/Role: Case from CT Apps Device Use: Diag Expected and Unexpected Results: Reset Exam Card Mgr password User Impact: None Patient Impact: None Current Software Version: 4.1.10 Diagnostic performed by Engineer : Accepted case from CT Clin Apps ~ As customer did not know HCF defined Admin password ~ we were not able to assist them resetting their Exam Card Mgr password. FSE dispatched. | 4 |  | Reset password according Refresh policy | [12 empty] | 118629203 |  | 1471670 |  |  | Evaluation | [7 empty] | Customer | [6 empty] | Password | [5 empty] | Refresh |  |  | Upgrade | [26 empty] | #N/A | bar: No to NAC
1500 |  |  | United States | BIBB MEDICAL CENTER | 310175 | 4449700% |  | Can CT system perform DEXA scans | Laura indicated that her old DEXA system is down and was inquiring if the CT system would be capable of gathering Bone Mineral Density images. After some investigation their CT system is capable of gathering this information, however a Bone Mineral Post processing packaged must be purchased to take these measurements after a CT scan. I had inquired if she needed me to get a quote but Laura said she would reach out to her Account Manager. A follow up call was performed to make sure this quote was generated for Laura. No rescan, no loss of data, no reinjection Per Laura. Laura had no further questions or needed for clinical support. | 5 |  | NAC_Q&A | [12 empty] | 118621855 |  | 1466633 |  |  | Evaluation | [7 empty] | NAC | [6 empty] | NAC | [5 empty] | Non-Refresh |  |  | Upgrade | [26 empty] | #N/A | bar: No to NAC
1501 |  |  | United States | Christ Hospital | 300205 | 4449500% |  | Intermittent problem with cardiac gating. | Integration : Confirmation of order 54353188 saved Integration : According to factory calendar 20211030 (Posting date) is not a workday (check entry) ----------------------------------------------------------------------------- Work Order #: WO-04634273 Mark Finnigan The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: Biomed Device Use: No patient involved. Expected and Unexpected Results: The cardiac gating waveform has had intermittent disconnects. User Impact: The customer has to power cycle the gantry when this happens. Patient Impact: Current Software Version: Diagnostic performed by Engineer : Spoke with Corey in biomed. He staes that the cardiac signal has experienced intermittent drop-outs requiring a power cycle to recover. The new CAT 6 cable has already been installed in place of the USB extenders. The cable between the ECG monitor and gantry has already been replaced, I recommended that Corey check and adjust the iBox 5VDC power supply to read 5.2 at the iBox. The next step would be to flash the iBox. If that does not help then the iBox should be replaced. Corey will take the first two steps and let us know if that helped. Diagnostic performed by Engineer : Called the customer back looking for an update. Call went to VM. Will try again later today. Diagnostic performed by Engineer : Corey called back in . The PS was at 5.185 volts. He flashed the iBox. He is going to proactively order a iBox board and needed the part number. The PN is 459800166652 Resolution : Provided the customer with te part number requested. He will complete the repair. The customer does not report to Philips so this case is closed as a resolution and complete. ----------------------------------------------------------------------------- | 4 |  | The PS was at 5.185 volts. He flashed the iBox. He is going to proactively order a iBox board and needed the part number. The PN is 459800166652 | [12 empty] | 118606839 |  | 1457901 |  |  | Investigation | [20 empty] | Non-Refresh |  |  | Upgrade | [26 empty] | #N/A | bar: No to NAC
1502 |  |  | United States | Philips Medical Capital | 300057 | 4449500% |  | 0999CFXXX Error message: Artifact stating a crack compensator | Work Order #: WO-04632649 Powell, Jamie The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Problem description by engineer : Error message: Artifact stating a crack compensator while running air cals Information to support the complaint handling process Customer Function/Role: Not provided Device Use: Not provided Expected and Unexpected results: Not provided Patient Impact/Outcome: Not provided User Impact: Not provided Current Software Version: Not provided Diagnostic performed by Engineer : Spoke to customer who states that she saw the message yesterday but she's not seeing any artifacts or any issues. Internal Comments : Error # and/or description of error : Down loading calibration files and no issues found. Ran cracked compensator tool and no indication found. Could be a false alarm or contrast in the field. Repair Action : Down loading calibration files and no issues found. Ran cracked compensator tool and no indication found. Could be a false alarm or contrast in the field. Resolution : No issues found with image quality and no indication of cracked comp. ----------------------------------------------------------------------------- | 4 |  | No issues found with image quality and no indication of cracked comp. | [12 empty] | 118605667 |  | 1457221 |  |  | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | False Alarm | [5 empty] | Non-Refresh |  |  | Upgrade | [26 empty] | #N/A | bar: No to NAC
1503 |  |  | United States | Select Medical Corporation | 32261 | 4449000% |  | 0599CFXXX E-stop not closing | Integration : Confirmation of order 54332955 saved ----------------------------------------------------------------------------- Work Order #: WO-04594414 Boyd Craig The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: E-stop not closing User Impact: None Patient Impact: None Current Software Version: Diagnostic performed by Engineer : E-stop not closing on bootup. Power cycle system PRS accessed. ULV: no normal errors seen. System operational. Test scans are good. Resolution : Power cycle system PRS accessed. ULV: no normal errors seen. System operational. Test scans are good. ----------------------------------------------------------------------------- | 2 | [14 empty] | 118572701 |  | 1442695 |  |  | Investigation | [20 empty] | Non-Refresh |  |  | Upgrade | [26 empty] | #N/A | bar: No to NAC
1504 |  |  | United States | Adventhealth Zephyrhills | 300106 | 4448900% |  | Q/O L3 Support requested a video from Brain Perfusion application, color map issue | Integration : Confirmation of order 54326617 saved ----------------------------------------------------------------------------- Work Order #: WO-04585496 20-Oct-2021 Garcia Marrero, Jorge The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : Capture video of Brain Perfusion application from host computer monitor. ----------------------------------------------------------------------------- | 5 |  | Make a video from Brain Perfusion application, color map issue for engineering investigation | [12 empty] | 118565695 |  | 1439474 |  |  | Investigation |  | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Brain Perfusion | [5 empty] | Refresh |  |  | Upgrade | System | Console-SW | NoEnoughInfo | Console-SW-NoEnoughInfo | Console-SW-NoEnoughInfo | NoEnoughInfo | NoEnoughInfo | Console-SW | Console-SW-NoEnoughInfo | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1505 |  |  | United States | AdventHealth Zephyrhills | 300099 | 4448800% |  | Black ring artifat on images; Performed air calibraiton to resolve the issue. | Integration : Confirmation of order 54320397 saved ----------------------------------------------------------------------------- Work Order #: WO-04578554 20-Oct-2021 Sher Mark The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: Ring artifact User Impact: None Patient Impact: None Current Software Version: Diagnostic performed by Engineer : customer cleared issue wiyh aircals. ran bad detectors and fee test.. checked system ----------------------------------------------------------------------------- Work Order #: WO-04577520 Boyd Craig The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: Ring artifact User Impact: None Patient Impact: None Current Software Version: Diagnostic performed by Engineer : Black ring artifact. Air calibrations performed. PRS accessed. Physics Analysis shows M26 bad. Possible contrast on window. Follow Up Required : FSE needed on site to replace detector module 26. Recommendations: Clean gantry window. Check for anything in the field of view. Perform air calibrations. Perform Test scans. Analyze images. Run bad detector test and FEE Detector Tests. Replace M26 if needed. ----------------------------------------------------------------------------- | 3 | [14 empty] | 118559312 |  | 1433008 |  |  | Closed - Done | [20 empty] | non-Refresh |  |  | Upgrade | [26 empty] | #N/A | bar: No to NAC
1506 |  |  | Netherlands | Stichting Rijnstate Ziekenhuis locatie Arnhem | 52101 | 4448800% |  | Additional CT imaging due to occurring interference. | "Submitter name: Frank Nieuwenhuijse Submitter email: frank.nieuwenhuijse@philips.com Other contact information: Tel.: 0648109620 Best method of contact: e-mail Form submit date: October 19 2021 Attachments: 1 Reason for delay: There were 2 days delay between the occurrence of the event and reporting to philips. This report was made to an email address (philips employee works in-house in Rijnstate) of the Rijnstate organization. This email was only available on October 18. (this email is only available in the Rijnstate building) Response required for submitter/customer: No Additional notes: No Safety Impact: Other details: The patient received an extra dose of X-ray radiation (272.5 mGy) Feedback type: Technical Issue Equipment ID: ---- -------------------------------------------------- ----------------- Problem definition: A malfunction has been reported with the CT scanner. This malfunction has made it necessary to repeat part of the scan protocol. This has added extra radiation. ng dose delivered to patient. This patient also received an extra dose of intravenous contrast agent. Description of incident The patient was referred to radiology for CT examination. Indication AAA size and indications for endoleak. An adapted version of scan protocol 4.2 was protocolled by the radiologist for this (in which, in addition to the abdomen, the thorax had to be scanned in both acquisitions and a blank CT thorax also had to be made). The patient was scanned on an Ingenuity CT scanner (NZC109, Philips, 2018) in Zevenaar. While performing bolus tracking for correct contrast timing, the speed of reconstructions visible on scanner was so low that an additional delay occurred causing the arterial phase acquisition to start too late. MBB aborted protocol after this scan so that the 2nd acquisition was cancelled. After this consultation took place with the intervention radiologist and the examination with regard to the contrast series was repeated. What now appears to be facts: - The patient has received an extra dose. From dose report in PACS it can be concluded that the total dose of this study yields a dose length product of 840.6 mGy.cm. Of this, 272.5 mGy.cm is due to additional CT imaging due to interference occurring. - The cause of the bolus tracking problem appears to be a stagnant reconstruction process while performing bolus track. Event/occurence date: October 11, 2021 Philips aware date: October 13, 2021 Ability to obtain the information: October 18, 2021" |  |  | Under Engineering investigation, need raise defect to tracking | [14 empty] | 1431481 |  |  | Investigation |  | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Bolus Tracking | [5 empty] | Refresh |  |  | Upgrade | System | Console-SW | NoEnoughInfo | Console-SW-NoEnoughInfo | Console-SW-NoEnoughInfo | NoEnoughInfo | NoEnoughInfo | Console-SW | Console-SW-NoEnoughInfo | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1507 |  |  | United States | Select Medical Corporation | 32261 | 4448700% |  | 1205CFXXX Gantry not coming on. | Integration : Confirmation of order 54314443 saved ----------------------------------------------------------------------------- Work Order #: WO-04569963 Morrow, Dan The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Problem description by engineer : System left off over the weekend. Gantry will not turn on. Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: Unable to use Patient Impact/Outcome: none User Impact: Unable to use Current Software Version: 4.1.7 Diagnostic performed by Engineer : Reviewed logs. Rhost not communicating. I walked the customer through a power cycle of the system. After power cycle system is up. Customer able to run STC and test images. OK to close. Internal Comments : Error # and/or description of error : S_RTDH_HC_TIMEOUT_ERROR Repair Action : Reviewed logs. Rhost not communicating. I walked the customer through a power cycle of the system. After power cycle system is up. Customer able to run STC and test images. OK to close. Resolution : Reviewed logs. Rhost not communicating. I walked the customer through a power cycle of the system. After power cycle system is up. Customer able to run STC and test images. OK to close. ----------------------------------------------------------------------------- | 2 | [14 empty] | 118553179 |  | 1428814 |  |  | Closed - No Further Investigation | [20 empty] | Non-Refresh |  |  | Upgrade | [26 empty] | #N/A | bar: No to NAC
1508 |  |  | United States | Upson Regional Medical Center | 310115 | 4448000% |  | Gantry error, replaced DMS fan and heaters, GPU, cleaned slip ring | Integration : Confirmation of order 54284159 saved ----------------------------------------------------------------------------- Work Order #: WO-04526142 12-Oct-2021 Lee John The system meets the specification for the performed service and is returned to use Resolution : Replaced dms heater fan assemblies, gpu in CRC, and cleaned the slip rings. Problem Description by Engineer : Customer Function/Role: CTTech Device Use: Diag Expected and Unexpected Results: Gantry Error User Impact: None Patient Impact: None Current Software Version: IDK Parts Consumed : 1.00000,459800550141,BR64T RoHS HEATER REPLACEMENT KIT 1.00000,459800550141,BR64T RoHS HEATER REPLACEMENT KIT 1.00000,459800980091,BR64 Air Fan Replacement Kit 1.00000,459801622381,NVidia P2000 GPU Card ----------------------------------------------------------------------------- Work Order #: WO-04531921 Sheffield, Scott The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Internal Comments : Error # and/or description of error : 2021/10/12 09:15:52.000__010 CirsEvt_GPUMonitor_177 CIRS_Error Pre Series GPUMonitor177GPUMonitor::RunConfigTests() FAILED (device=0), error details: CudaDevice::CheckConfig failed: GPU current PCIe link width is 8 which is less than the maximum value of 16, which may result in slower performance. Failed device info: Name: Quadro P2000 Serial: 1560420600473 PCI bus ID: 0000:03:00.0 2021/10/12 09:15:57.000__010 CirsEvt_ProxyServer_90 CIRS_Error Series Prep ProxyServer90Exception occurred (type: class Cirs::BasicExceptStream<class Cirs::CirsExcept>) at SeriesProcess.cpp(1442): Error #12 - CheckGantryConnection: ACQ board or Gantry-to-CIRS datapath failure. CIRS Event Log Checked 2021/10/12 09:15:57.961__002 S_DMC_SHOT_COMM_COLLAPSE ERROR_NORMAL Series Exit CStateStudyExit Logger.mdb Checked Type : Error Date : 10/12/2021 Time : 9:33:45 AM Event : 0 Source : Supervisor Category : Error User : N/A Computer : CIRSPHI-27FF54D Description: ImageStudyElmt(342): Error #5 - #61658D43:000 (ACQ #3696) - Failed to start properly 2021/10/12 09:33:46.036 S_DMC_SHOT_COMM_COLLAPSE ERROR_NORMAL Series Exit CStateStudyExit Logger.mdb Checked 2021/10/12 09:43:56.000__001 CirsEvt_GPUMonitor_177 CIRS_Error Series Exit GPUMonitor177GPUMonitor::RunConfigTests() FAILED (device=0), error details: CudaDevice::CheckConfig failed: GPU current PCIe link width is 8 which is less than the maximum value of 16, which may result in slower performance. Failed device info: Name: Quadro P2000 2021/10/12 09:43:56.000__079 CirsEvt_ProxyServer_90 CIRS_Error Series Exit ProxyServer90Exception occurred (type: class Cirs::BasicExceptStream<class Cirs::Os::SEExcept>) at OsSETranslator.cpp(215): Error #12 - ACCESS_VIOLATION occurred at 0x00007FF787C341B1 attempting to read from address 0x0000000000000010 Repair Action : I called FSE John and spoke to him about the GPU errors. John is about to be onsite to work on system. John said that he saw GPU errors in the bug report he reviewed yesterday and he ordered a GPU board already. Diagnostic performed by Engineer : Katy said that they had a power outage this morning and now system is down. They have rebooted a couple times. I was able to log into cirs with remote desktop and troubleshoot problem. Problem Description by Engineer : Problem description by engineer : cirs can't initialize Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: cirs can't initialize Patient Impact/Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer Current Software Version: Not provided Resolution : FSE John is going to resolve the problem for onsite wo WO-04526142. ----------------------------------------------------------------------------- Work Order #: WO-04526121 Carpenter, Kevin The system does not meet the specification for the performed service and is not returned to use Problem Description by Engineer : Customer Function/Role: CTTech Device Use: Diag Expected and Unexpected Results: Gantry Error User Impact: None Patient Impact: None Current Software Version: IDK Diagnostic performed by Engineer : CT RSE remotely collected error logs and reviewed same. Added details and possible CHA parts correlations to dispatch follow-up notes for FSE. Follow Up Required : 2021/10/11 10:28:45.733 S_MPROC_DEVICE_NO_DATA_FROM_GMP_SERIAL P1=5 P2=5 xMany 10:28:45.217 S_MP_CTRL_NO_MORE_AT_SPEED P1=8000 MP-Ctrl No-More-AT-Speed. IP: 8045 10:28:44.328 S_RMC_INVALID_POSITION 10:28:44.286__08 S_CTRL_LINK_GHOST_GOT_CRC_FRAMING_ERROR Last gantry power cycle ~ 14 days ago CHA not working at this time ~ no specific parts recommendations. FSE consult NSS as needed. ----------------------------------------------------------------------------- | 3 | [14 empty] | 118517645 |  | 1406037 |  |  | Closed - Done | [20 empty] | Non-Refresh |  |  | Upgrade | [26 empty] | #N/A | bar: No to NAC
1509 |  |  | United States | AdventHealth Zephyrhills | 300099 | 4447800% |  | Estop not closing; Replaced couch relay and tilt actuators, performed tilt calibrations. | Integration : Confirmation of order 54278587 saved ----------------------------------------------------------------------------- Work Order #: WO-04518522 13-Oct-2021 Garcia Marrero, Jorge The system meets the specification for the performed service and is returned to use Diagnostic performed by Engineer : Installation of left and right Linear Actuator on gantry legs and calibration of gantry Tilt. System is fully operational. Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: E-stop not closing User Impact: None Patient Impact: None Current Software Version: Parts Consumed : 1.00000,453566502541,LINEAR ACTUATOR SKF LH 1.00000,453567114552,ACTUATOR RH SKF 1.00000,453566496941,RELAY,24VDC PLUGIN W/COVER ----------------------------------------------------------------------------- Work Order #: WO-04533050 12-Oct-2021 Watson Adam The system meets the specification for the performed service and is returned to use Resolution : Replaced gantry tilt actuators, performed tilt calibrations, and backed up NVRAM. Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: ----------------------------------------------------------------------------- Work Order #: WO-04518316 09-Oct-2021 Watson Adam The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Diagnostic performed by Engineer : Troubleshot system and found the right side tilt actuator has lifted off of its switches causing estop to open when the customer tilted the gantry. Engineer engaged the actuator switches to move the gantry to the 0 degree position and instructed the customers not to tilt the gantry until he new actuators arrive. Also found horizontal motor relay open errors when the gantry would attempt to scan and move the table. Replaced the horizontal motor relay. System has been turned back over to the customer for clinical use. Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: E-stop not closing User Impact: None Patient Impact: None Current Software Version: Follow Up Required : Replace both tilt actuators ----------------------------------------------------------------------------- Work Order #: WO-04518271 Boyd Craig The system does not meet the specification for the performed service and is not returned to use Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: E-stop not closing User Impact: None Patient Impact: None Current Software Version: Diagnostic performed by Engineer : Patient vomited in gantry. When gantry tilted the E-stop tripped. System power cycled. All E-stops cycled. E-stop not closing. Unable to retrieve error logs. Follow Up Required : PO WC001-0001651531. Customer requesting service today. FSE needed on site to evaluate E-stop circuit and error logs. ----------------------------------------------------------------------------- | 2 | [14 empty] | 118511359 |  | 1402522 |  |  | Closed - Done | [20 empty] | Non-Refresh |  |  | Upgrade | [26 empty] | #N/A | bar: No to NAC
1510 |  |  | United States | Upson Regional Medical Center | 333009 | 4447500% |  | 4.1.10 EC Manager password expired | Integration : Confirmation of order 54274523 saved ----------------------------------------------------------------------------- Work Order #: WO-04512872 Hirchak, Michele The system meets the specification for the performed service and is returned to use Problem description by Clinical Support Specialist : Customer Function/Role: TECH Device USE: IN USE Expected and Unexpected Results: Reset EC Manager password. User Impact: NONE Patient Impact: NONE Current Software Version: 4.1.10 No rescan, no re-injection, no patient harm. Diagnostic performed by Engineer : Customer unable to get in EC Manager. Message tells her password has expired. I told her their IT Admin should be able to reset it. However, their IT Admin is out of the office. This is not an emergency, she will let me know if she wants the FSE to come out to reset. MBH Resolution : Customer was in contact with FSE. John Lee will be going there for another issue, he will reset the passwords when he is onsite. Spoke to John to confirm. MBH Problem description by Clinical Support Specialist : Called to f/u. No answer, no voicemail.MBH ----------------------------------------------------------------------------- | 5 |  | Reset password according Refresh policy | [12 empty] | 118496068 |  | 1396246 |  |  | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Password | [5 empty] | Refresh |  |  | Upgrade | [26 empty] | #N/A | bar: No to NAC
1511 |  |  | United States | Upson Regional Medical Center | 310115 | 4447400% |  | Temperature control error. | duplicate case to 0118517645 | 5 | [14 empty] | 118489673 |  | 1391704 |  |  | Investigation | [20 empty] | Non-Refresh |  |  | Upgrade | [26 empty] | #N/A | bar: No to NAC
1512 |  |  | United States | Philips Medical Capital | 300057 | 4447400% |  | Image artifact, inreased Idose level and centrered patient | Work Order #: WO-04489209 Curry, Dyan The system meets the specification for the performed service and is returned to use Problem description by Clinical Support Specialist : Customer Function/ Role: technologist Device Use: in use name Expected and Unexpected Results: artifact questions User Impact: none Patient Impact: none Current Software version: 4.1 Diagnostic performed by Engineer : FSE emailed images to me. MAS is low and patient is off centered. Advised FSE to increase Idose level and center patient. check MAS profile before scanning | 4 | [14 empty] | 118488817 |  | 1386699 |  |  | Closed - Done | [20 empty] | Non-Refresh |  |  | Upgrade | [26 empty] | #N/A | bar: No to NAC
1513 |  |  | United States | Philips Medical Capital | 300057 | 4447300% |  | 0999CDN83 Star Artifact on Thins | Work Order #: WO-04496328 08-Oct-2021 Williams John The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: CTTech Device Use: Diag Expected and Unexpected Results: Star Artifact User Impact: None Patient Impact: None Current Software Version: IDK Resolution : Review images with technologist . Passed on information applications gave me. ----------------------------------------------------------------------------- Work Order #: WO-04482035 05-Oct-2021 Williams John The system meets the specification for the performed service and is returned to use Follow Up Required : Observe and convey applications recommendations regarding image quality. Will work with apps while onsite. Problem Description by Engineer : Customer Function/Role: CTTech Device Use: Diag Expected and Unexpected Results: Star Artifact User Impact: None Patient Impact: None Current Software Version: IDK Diagnostic performed by Engineer : Download and send images to applications and NSS. Work via phone with applications. ----------------------------------------------------------------------------- Work Order #: WO-04481960 Carpenter, Kevin The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Follow Up Required : Star Artifact on Thins ~ FSE please review artifact images onsite. RSE asked cust to copy artifact to CDR for easy download ~ but images not found there. RSE also had cust ' lock ' Pt exam contain artifact ~ see series with ' thins' Problem Description by Engineer : Customer Function/Role: CTTech Device Use: Diag Expected and Unexpected Results: Star Artifact User Impact: None Patient Impact: None Current Software Version: IDK Diagnostic performed by Engineer : Star Artifact on Thins RSE asked cust to copy artifact to CDR for easy download ~ but images not found there. RSE also had cust ' lock ' Pt exam contain artifact ~ see series with ' thins' ----------------------------------------------------------------------------- | 3 | [14 empty] | 118483434 |  | 1383083 |  |  | Investigation | [20 empty] | Non-Refresh |  |  | Upgrade | [26 empty] | #N/A | bar: No to NAC
1514 |  |  | Spain | POLICLINICA NTRA SRA DEL ROSARIO SL | 336190 | 4447300% |  | The UPS has failed. | Work Order #: WO-04486742 05-Oct-2021 Parts Consumed : 1.00000,459800306392,EXTERNAL BATT PACK, 3KVA UPS, 72VDC ----------------------------------------------------------------------------- Work Order #: WO-04477685 05-Oct-2021 [REDACTED] The system meets the specification for the performed service and is returned to use. Resolution: Battery is requested. Problem Description by Engineer: The UPS is damaged. Regulatory Questions : N/A Follow Up Required: Order battery to send directly to the customer. ----------------------------------------------------------------------------- Work Order #: WO-04477048 [REDACTED] The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled. Diagnostic performed by Engineer: I talk to [REDACTED]. The UPS has problems with the battery, the equipment is in bypass mode, need to replace the battery. Problem Description by Engineer: The UPS is damaged. Regulatory Questions : N/A Follow Up Required: Replace the battery. ----------------------------------------------------------------------------- | 3 | [14 empty] | 118479908 |  | 1381591 |  |  | Closed - No Further Investigation | [20 empty] | Non-Refresh |  |  | Upgrade | [26 empty] | #N/A | bar: No to NAC
1515 |  |  | Australia | Maia Financial Pty Limited | 341000 | 4452900% | 48 | Recording Voice Messages Assistance | Work Order #: WO-04866036 Steve, Teixeira Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 5 |  | Application Support | [12 empty] | 118802178 | In Progress | 1568366 |  |  | Opened | Opened | [6 empty] | Application | [6 empty] | Application Support | [5 empty] | Refresh |  |  | System issue | bar: No; No
1516 |  |  | Australia | Maia Financial Pty Limited | 341000 | 4452500% | 47 | Surview scan stopped <CT> | Work Order #: WO-04849511 Steve, Teixeira The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : CUSTOMER REPORTED PROBLEM OR EVENT DESCRIPTION: 26 Nov 2021 The Customer called to report on behalf of a colleague that during a Chest Abdomen study the surview scan stopped 3/4 of the way through. Spoke to the Radiographer scanning at the time and informed that the first surview stopped 3/4 of the way through and produced a message asking if the User wanted to proceed to the 2nd surview. On the 2nd atempt the scan appeared to complete but no image appeared. Whilst on the phone, on the 3rd attempt the User was able to complete a dual durview and the next scans successfully. Study completed. Bug report generated. ________________________________________________________________ Diagnostic performed by Engineer : CUSTOMER REPORTED PROBLEM OR EVENT DESCRIPTION: 26 Nov 2021 The Customer called to report on behalf of a colleague that during a Chest Abdomen study the surview scan stopped 3/4 of the way through. Spoke to the Radiographer scanning at the time and was informed that the first surview stopped 3/4 of the way through and produced a message asking if the User wanted to proceed to the 2nd surview. On the 2nd attempt, the scan appeared to complete but no image appeared. Whilst on the phone, on the 3rd attempt the User was able to complete a dual durview and the next scans successfully. Study completed. Bug report generated. SOFTWARE REVISION OF THE SYSTEM: Ver. 4.1.10 WAS A BUG REPORT SAVED? Yes. LOG REVIEW: A review of the logs confirms that during the 2 attempts the surview scans stopped (at 10:09:05 AM and 10:11:19 AM 26 Nov 2021) due to arcing (anode and cathode). TROUBLESHOOTING: 26 Nov 2021 Assisted the Customer with running Filament calibration and starting the long tube conditioning. CUSTOMER UPDATE: 26 Nov 2021 Sent an email update to the Customer summarising the events. Email extract reads: " HI XXXXX, As discussed today the surview scans stopped due to arcing. ZZZZ successfully completed the Filament calibration and started a long tube conditioning which I believe is complete by now. ZZZZ (hoping I have the spelling correct) wanted to know where to find the tube heat status. The CT Application Specialist YYYY confirmed that It is found under Utilities (Spanner icon) -> Show tube heat. Generally we want the tube heat to be around 10%, and if the system is idle for extended periods, a short tube conditioning prior to the study is good practice. Given the arcing today please be mindful of the tube heat and running a short tube conditioning when the system is idle for more than 2 hours, or when called out to do after hour scans. This is purely a recommendation as appreciate that different Users are involved and there are other work demands involved. I will keep the case open until next week to ensure that there are no further signs of arcing, as we may need to run another long tube conditioning should the arcing persist. As a gentle reminder please restart the Host computer daily and the CIRS Server/s weekly, although not related to the arcing it does minimize the chance of lockups, performance issues etc. Please feel free to contact me should you have any further queries in this regard. ..." FOLLOW UP: 2 Dec 2021 Called the Customer and confirmed that no further reports of arcing-related messages or scans stopping prematurely. ________________________________________________________________ Internal Comments : LOG REVIEW: *********** ________________________________________________________________ Filament cal last done 16 Oct 2020. ________________________________________________________________ LOGGER: * 2021/11/26 10:19:34.226__01 ERROR_BLOB CANMS2 ------ Series 24 Load Structure ----- 2021/11/26 10:19:34.137__01 ERROR_BLOB CANMS2 ------ Series 23 Load Structure ----- 2021/11/26 10:19:34.058__05 ERROR_BLOB CANMS2 ------ Series 22 Load Structure ----- 2021/11/26 10:19:00.245 EVENT_MAJOR GSE ----------------- Series 21 Complete ----------------- 2021/11/26 10:18:58.655__01 EVENT_MAJOR GSE - - - - - - Shot 0 Starts - - - - - - 2021/11/26 10:18:46.763__05 ERROR_BLOB CANMS2 ------ Series 21 Load Structure ----- 2021/11/26 10:17:22.136 EVENT_MAJOR GSE ----------------- Series 20 Complete ----------------- 2021/11/26 10:17:13.604__01 EVENT_MAJOR GSE - - - - - - Shot 0 Starts - - - - - - 2021/11/26 10:16:53.503 EVENT_MAJOR GSE ----------------- Series 19 Complete ----------------- 2021/11/26 10:16:43.973__01 EVENT_MAJOR GSE - - - - - - Shot 0 Starts - - - - - - 2021/11/26 10:16:01.178__01 ERROR_BLOB CANMS2 ------ Series 20 Load Structure ----- 2021/11/26 10:16:01.113__09 ERROR_BLOB CANMS2 ------ Series 19 Load Structure ----- 2021/11/26 10:15:49.450__02 EVENT_MINOR GSE GantryStartStudyBkst was sent to Gantry 2021/11/26 10:15:49.450 EVENT_MINOR GSE GSE::CState::StateBkst::STUDY_STARTED was sent to GPC 2021/11/26 10:15:30.288 EVENT_MINOR GSE Series 18 cancelled, due to GANTRY_CANCEL_SOURCE_HOST, in state CStateSeriesReady 2021/11/26 10:15:30.269 EVENT_MINOR GSE GSE::CState::StateBkst::STUDY_ENDED was sent to GPC 2021/11/26 10:15:19.183 ERROR_WARNING Gantry CStateSeriesReady 0 ERROR 165295 WARN S_MASTER_SEQUENCER_SERIES_DELAY_LESS_THAN_MIN 0 4123 2021/11/26 10:15:12.994 ERROR_BLOB CANMS2 ------ Series 18 Load Structure ----- 2021/11/26 10:14:36.903 EVENT_MINOR GSE Series 17 cancelled, due to GANTRY_CANCEL_SOURCE_CTBOX, in state CStateSeriesReady 2021/11/26 10:14:34.271 ERROR_WARNING Gantry CStateSeriesReady 0 ERROR 165250 WARN S_MASTER_SEQUENCER_SERIES_DELAY_LESS_THAN_MIN 0 4123 2021/11/26 10:14:13.695__06 ERROR_BLOB CANMS2 ------ Series 17 Load Structure ----- 2021/11/26 10:11:19.982 ERROR_WARNING Gantry CStateStudyExit 1 ERROR 165056 WARN S_GEN_MGR_SYNC_OFF_OUT_OF_SEQUENCE 0 0 2021/11/26 10:11:19.976 ERROR_WARNING Gantry CStateStudyExit 1 ERROR 165056 WARN S_GEN_FAULT_CATHODE_TUBE_ARC 5 33 5:5 CPM Arc 2021/11/26 10:11:19.934 ERROR_WARNING Gantry CStateStudyExit 1 ERROR 165056 WARN S_GEN_FAULT_ANODE_TUBE_ARC 5 33 5:0 APM Arc 2021/11/26 10:11:19.932 ERROR_WARNING Gantry CStateStudyExit 1 ERROR 165056 WARN S_GEN_FAULT_ANALOG_FAULT 2 32 2:5 Analog Fault 2021/11/26 10:11:19.887__15 ERROR_NORMAL Gantry CStateStudyExit 0 ERROR 165056 NORMAL S_MASTER_SEQUENCER_NO_XRAY 0 1 2021/11/26 10:11:19.885__03 EVENT_MINOR GSE CStateShot:Series 15 cancelled, due to GANTRY_CANCEL_SOURCE_NORMAL_ERROR, in state CStateShot 2021/11/26 10:11:18.102__01 EVENT_MAJOR GSE - - - - - - Shot 0 Starts - - - - - - 2021/11/26 10:10:58.646__01 ERROR_BLOB CANMS2 ------ Series 16 Load Structure ----- 2021/11/26 10:10:58.576__09 ERROR_BLOB CANMS2 ------ Series 15 Load Structure ----- 2021/11/26 10:10:11.632 EVENT_MINOR GSE Series 14 cancelled, due to GANTRY_CANCEL_SOURCE_CTBOX, in state CStateSeriesPrep 2021/11/26 10:09:35.927__04 ERROR_BLOB CANMS2 ------ Series 14 Load Structure ----- 2021/11/26 10:09:19.417 EVENT_MINOR GSE Series 13 cancelled, due to GANTRY_CANCEL_SOURCE_CTBOX, in state CStateSeriesPrep 2021/11/26 10:09:15.693__06 ERROR_BLOB CANMS2 ------ Series 13 Load Structure ----- 2021/11/26 10:09:07.087__01 ERROR_NORMAL Gantry CStateStudyExit 0 ERROR 164923 NORMAL S_MASTER_SEQUENCER_EXIT_STATE_ERR 4096 8369 2021/11/26 10:09:05.071 ERROR_WARNING Gantry CStateStudyExit 1 ERROR 164921 WARN S_GEN_MGR_SYNC_OFF_OUT_OF_SEQUENCE 0 0 2021/11/26 10:09:05.068 ERROR_WARNING Gantry CStateStudyExit 1 ERROR 164921 WARN S_GEN_FAULT_CATHODE_TUBE_ARC 5 33 5:5 CPM Arc 2021/11/26 10:09:05.032__06 ERROR_WARNING Gantry CStateStudyExit 1 ERROR 164921 WARN S_GEN_FAULT_ANODE_TUBE_ARC 5 33 5:0 APM Arc 2021/11/26 10:09:05.029 EVENT_MINOR GSE CStateShot:Series 11 cancelled, due to GANTRY_CANCEL_SOURCE_NORMAL_ERROR, in state CStateShot 2021/11/26 10:09:05.027 ERROR_WARNING Gantry CStateShot 1 ERROR 164921 WARN S_GEN_FAULT_ANALOG_FAULT 2 32 2:5 Analog Fault 2021/11/26 10:09:05.021 ERROR_NORMAL Gantry CStateShot 1 ERROR 164921 NORMAL S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT 0 0 2021/11/26 10:09:00.335 EVENT_MAJOR GSE - - - - - - Shot 0 Starts - - - - - - 2021/11/26 10:08:38.732 ERROR_BLOB CANMS2 ------ Series 12 Load Structure ----- 2021/11/26 10:08:38.651__07 ERROR_BLOB CANMS2 ------ Series 11 Load Structure ----- 2021/11/26 10:08:27.554__02 EVENT_MINOR GSE GantryStartStudyBkst was sent to Gantry 2021/11/26 10:08:27.554 EVENT_MINOR GSE GSE::CState::StateBkst::STUDY_STARTED was sent to GPC * ________________________________________________________________ Resolution : CASE CLOSED: As the system is operational after running the relevant calibrations and monitoring, as agreed with the Customer, no further service-related action will be taken on this case. This issue will continue to be monitored via Customer calls and alerts over time, to identify trends if any, that may warrant further action. ________________________________________________________________ ----------------------------------------------------------------------------- | 3 |  | As the system is operational after running the relevant calibrations and monitoring, as agreed with the Customer, no further service-related action will be taken on this case. This issue will continue to be monitored via Customer calls and alerts over time, to identify trends if any, that may warrant further action. | [12 empty] | 118788208 | Incomplete | 1558982 |  |  | Investigation | Investigation | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Scan error | [5 empty] | Refresh |  |  | System issue | System | Console-SW | NoEnoughInfo | Console-SW-NoEnoughInfo | Console-SW-NoEnoughInfo | NoEnoughInfo | NoEnoughInfo | bar: No; No
1517 |  |  | Australia | Murwillumbah Hospital | 341000 | 4452500% | 47 | Phantom is hard to use when performing IQ test on an Ingenuity system; |  | #N/A |  | Phantom is hard to use when performing IQ test on an Ingenuity system; | [14 empty] | 1554916 |  |  | Investigation | Investigation | H-Phantom | [5 empty] | Engineering | [6 empty] | H-Phantom RFC | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | H-Phantom | Couch-HW-H-Phantom | Couch-HW-H-Phantom | H-Phantom | H-Phantom | bar: No; No
1518 |  |  | China | 江苏大学附属医院 | 345023 | 4452900% | 48 | wechat：02N/Ingenuity Core128/Elite China/Deleting Network Node | BU Q&R Feedback: No SAP Customer Symptom Code: HANG Customer Symptom Code: Hang/Crash Reason: Corrective Maintenance Engineer Failure Code: Configuration Error Repair Activity Code: Configuration Change Problem Reported by Customer : wechat：02N/Ingenuity Core128/Elite China/ Safety Questions : Case: 0118802918 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Bo Last updated by: Wang Bo 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户需要删除节点。 Troubleshooting Action: 远程协助客户删除添加的节点。 Repair Action: 远程协助客户删除添加的节点。机器正常。 Test/Inspection: customer confirm can scan patient and image ok.functional Check PASSED. Device returned to full functionality. Test Equipment: N/A ----------------------------------------------------------------------------- | 3 | [14 empty] | 118802918 | Fixed | 1568804 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [19 empty] | Refresh |  |  | System issue | bar: No; No
1519 |  |  | Chile | BANCO DE CHILE | 341005 | 4452300% | 47 | CL1688 EL GANTRY SE APAGA | Work Order #: WO-04840371 26-Nov-2021 [REDACTED] The system meets the specification for the performed service and is returned to use. Diagnostic performed by Engineer: Equipment stops once with errors: S_ROTOR_MCO S_RMC_Error_MCO Problem Description by Engineer : Problem description by engineer : Gantry shuts off Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected Results: No gantry errors. Patient Impact/Outcome: Delay in studies. User Impact: Delay in studies. Current Software Version: 4.1.10 Resolution: Optical and electrical slip ring inspection is done, firmware is loaded to GMP. ----------------------------------------------------------------------------- Work Order #: WO-04829286 [REDACTED] The system meets the specification for the performed service and is returned to use. Internal Comments : Error # and/or description of error : S_MPROC_DEVICE_NO_DATA_FROM_GMP_SERIAL S_MPS_WAITING_FOR_MAIN_DRIVE_TO_CLOSE Repair Action: - Check GMP connections and cables. - Reconfigure GMP. Diagnostic performed by Engineer: The gantry turns off. In the error log we observed an interruption in the communication of the gantry and an error in a controller. Gantry check-up is needed. Problem Description by Engineer : Problem description by engineer : Gantry shuts off Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected Results: No gantry errors. Patient Impact/Outcome: Delay in studies. User Impact: Delay in studies. Current Software Version: 4.1.10 Follow Up Required: Error Number and/or Description: S_MPROC_DEVICE_NO_DATA_FROM_GMP_SERIAL S_MPS_WAITING_FOR_MAIN_DRIVE_TO_CLOSE Repair Action: - Check GMP connections and cables. - Reconfigure GMP. Technician's Name: [REDACTED] ----------------------------------------------------------------------------- | 3 |  | - Check GMP connections and cables. - Reconfigure GMP | [12 empty] | 118772061 | Closed | 1553335 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | TBD-GMP | [5 empty] | Engineering | [6 empty] | GMP | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | GMP | Gantry-HW-GMP | Gantry-HW-GMP | TBD-GMP | GMP | Gantry-HW | Gantry-HW-GMP | bar: No; No
1520 |  |  | China | Medical center (contagious disease, He nan, Xu chang) | 40047 | 4450700% | 44 | 02N/Ingenuity CT 728326/host have no free disk space | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Calibration/Adjustment Problem Reported by Customer: 02N/Ingenuity CT 728326 Safety Questions:Case: 0118675828 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Gaoge Last updated by: Wang Gaoge 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer's problems: failure related to computer software. Detailed description of Customer's problems or malfunctions: Host system failed to start It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: The host computer failed to start software Remote troubleshoot and test result: There is not enough storage space for the host PRS TR status: it's not through, ask engineer for maintenance on site. Repair Action: Suggest onsite service procedures: 1. re-format and re-install; 2. delete redundant data Test/Inspection: N/A Test Equipment: N/A SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support ----------------------------------------------------------------------------- | 2 |  | $Q7: (Confirmed reported problem ,low disk space ) $Q8:( low disk space ,software can not run ) Q9 (delete log files and other no used files , then issue solved !) | [12 empty] | 118675828 | In Process | 1498106 |  |  | Evaluation - Pending Translation | Evaluation - Pending Translation | CTUDT00632261 | [5 empty] | Engineering | [6 empty] | Disk Full | [5 empty] | Refresh |  |  | System issue | System | Console-SW | DiskFull | Console-SW-DiskFull | Console-SW-DiskFull | CTUDT00632261 | DiskFull | Console-SW | Console-SW-DiskFull | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1521 |  |  | Chile | BANCO DE CHILE | 341005 | 4451500% | 46 | CL1688 ERROR EN EL ESTUDIO POST-CARDIACO | Work Order #: WO-04766671 [REDACTED] The system meets the specification for the performed service and is returned to use. Problem Description by Engineer: Customer Function/Role: Technician Device Use: Diagnostic Expected and Unexpected Results: No error messages. User Impact: Delay in studies. Patient Impact: Delay in studies. Current Software Version: 11.1.3 Internal Comments: RSW not available Diagnostic performed by Engineer: A warning error is displayed for access to reserved programs. After accepting the message, they can continue working. The equipment software needs to be reviewed. Follow Up Required: Configuration check-up and software check-up. Evaluate possible software reloading. ----------------------------------------------------------------------------- | 4 |  | CL1688 ERROR IN POST-CARDIAC STUDY I get a warning error for accessing reserved programs. After accepting the message, you can continue working. Equipment software revision is required. | [12 empty] | 118718057 | In Progress | 1526933 |  |  | Evaluation | Evaluation | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | SW error | [5 empty] | Refresh |  |  | System issue | System | Console-SW | NoEnoughInfo | Console-SW-NoEnoughInfo | Console-SW-NoEnoughInfo | NoEnoughInfo | NoEnoughInfo | bar: No; No
1522 |  |  | China | Sichuan Province Yibin City No.2 T.C.M. Hospital | 345035 | 4453000% | 48 | 02N/Ingenuity Core128/Elite China/couch noise | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 02N/Ingenuity Core128/Elite China Safety Questions:Case: 0118811414 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang T2 Activities:Problem Description: Summary of customer's problems: Other types of error. Detailed description of Customer's problems or malfunctions: noise with the bed It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: noise with the bed Remote troubleshoot and test result: Not sure about the specifics. Need to go to the field to check PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Check to see if it's caused by the shell, then add lube and try SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support ----------------------------------------------------------------------------- | 3 |  | noise with the bed $Q7: (Confirmation the report problem.Couch noise when the couch up from low position.) $Q8: (Couch noise when the couch up from low position.check and found the cover was rub.adjust and clean the cover,lubricate the cover.) $Q9: (Couch noise when the couch up from low position.check and found the cover was rub.adjust and clean the cover,lubricate the cover.) | [12 empty] | 118811414 | Closed | 1573362 |  |  | Opened - Pending Translation | Opened - Pending Translation | E-125059 | [5 empty] | Engineering | [6 empty] | Couch Noise | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | Lubricant | Couch-HW-Lubricant | Couch-HW-Lubricant | E-125059 | Lubricant | Couch-HW | Couch-HW-Lubricant | bar: No; No
1523 |  |  | China | Lindian County Hospital | 345038 | 4452900% | 48 | 02N/Ingenuity Core128/Elite China/Yellow screen | T2 Activities:Problem Description: Summary of customer's problems: The host has a yellow screen. Detailed description of Customer's problems or malfunctions: The host has a yellow screen. It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: N/A Remote troubleshoot and test result: N/A PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: It is recommended to do a workaround upgrade. We recommend spare parts 12NC and names: N/A Memo of Customer request: None Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support ----------------------------------------------------------------------------- | 3 |  | $Q7: (Confirmed reported problem & host down) $Q8: (host down & host down) $Q9: (restart host) $Q10: (functional test pass. & Device returned to full functionality) | [12 empty] | 118802673 | Closed | 1568658 |  |  | Investigation - Pending Translation | Investigation - Pending Translation | CTUDT00632520 | [5 empty] | Engineering | [6 empty] | Yellow Screen | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | bar: No; No
1524 |  |  | China | 扎兰屯市人民医院 | 40045 | 4452900% | 48 | 02N/Ingenuity CT 728326/Ist authorization | BU Q&R Feedback: No SAP Customer Symptom Code: HANG Customer Symptom Code: Hang/Crash Reason: Corrective Maintenance Engineer Failure Code: Configuration Error Repair Activity Code: Configuration Change Problem Reported by Customer : 2N/Ingenuity CT 728326/ Next Step Action : ASP yiyuanji 015604188333 call in T2 Activities : Problem Description: 保养工程师需要IST授权。 Troubleshooting Action: 远程协助保养工程师进行授权。 Repair Action: 远程协助保养工程师进行授权。授权失败，PRS报错连接失败。无法进行授权。 Test/Inspection: customer confirm can scan patient and image ok.functional Check PASSED. Device returned to full functionality. Test Equipment: N/A Safety Questions : Case: 0118802526 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Bo Last updated by: Wang Bo 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 3 |  | 保养工程师需要IST授权。 | [12 empty] | 118802526 | Fixed | 1568572 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Service | [6 empty] | IST Push | [5 empty] | Refresh |  |  | System issue | bar: No; No
1525 |  |  | China | Shaanxi Province Xianyang Central Hospital (Xianyang City No. 2 People's Hospital) | 40050 | 4452900% | 48 | 02N/Ingenuity CT 728326/gantry noise when rotating | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Replace Part Safety Questions:Case: 0118801707 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Gaoge Last updated by: Wang Gaoge 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer's problems: The device makes a strange sound. Detailed description of Customer's problems or malfunctions: Abnormal sound can be heard when the gantry rotates. It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: An abnormal sound occurs on the tensioner. Remote troubleshoot and test result: Consider that the belt is causing the sound or that the tensioner is causing the sound. PRS TR status: No PRS. Repair Action: Suggest onsite service procedures: 1. Replace the belt or tensioner. We recommend spare parts 12NC and names: "453567003451 TENSIONER ASSY 453566492401 POLY-VEE BELT" Test/Inspection: N/A Test Equipment: N/A Next Step Action: PPD 453566492401*1 453567003451*1 all have been shipped. field service engineer: Guo Yujie - 86-15158180650 24H Problem Reported by Customer: 02N/Ingenuity CT 728326 SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support ----------------------------------------------------------------------------- | 3 |  | $Q7:confirmed report problem,ct device has noise when gantry rotation, customer can not use device $Q8:review error logs ,not found any error ,to check gantry rotor motor,belt and belt Tensioner , found is Tensioner has noise, belt is normal, so replace the new tensioner, gantry rotation not has noise, device runing normal, so think customer observation and use. $Q9:I replace the new tensioner, gantry rotation not has noise, device return normal., $Q10:Test and inspection data refer to PA ，Device returned to full functionality 453567003451 TENSIONER ASSY | [12 empty] | 118801707 | Closed | 1568109 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | ECR-077012 | [5 empty] | Engineering | [6 empty] | TENSIONER ASSY | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Belt | Gantry-HW-Belt | Gantry-HW-Belt | ECR-077012 | Belt | Gantry-HW | Gantry-HW-Belt | bar: No; No
1526 |  |  | China | Inner Mongolia Alxa Right Banner People's Hospital | 345031 | 4452900% | 48 | 02N/Ingenuity Core128/Elite China/power on failed | BU Q&R Feedback: No SAP Customer Symptom Code: BOOT Customer Symptom Code: Boot up/Power-up Problem Reason: Corrective Maintenance Engineer Failure Code: Incompatible Use Repair Activity Code: Instructed Customer Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ T2 Activities:Problem Description: power on failed Troubleshooting Action: customer call service, system can't power on. instruct customer check power on, customer not power on teal transform, instruct customer power on it, then power on UPS, system is ok. Repair Action: nothing to do, close case. Test/Inspection: Functional Check PASSED. The device returned to full functionality. Test Equipment: N/A SAP Transfer:Case Transferred to SAP Safety Questions:Case: 0118800659 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 3 |  | customer call service, system can't power on. instruct customer check power on, customer not power on teal transform, instruct customer power on it, then power on UPS, system is ok. | [12 empty] | 118800659 | Closed | 1567498 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Others | [6 empty] | System Reboot | [5 empty] | Refresh |  |  | System issue | bar: No; No
1527 |  |  | China | Jinan City No. 4 People's Hospital | 40038 | 4452900% | 48 | Intermittent power failure, closed keyswitch and rebooted CT gantry | BU Q&R Feedback: No Reason: Corrective Maintenance T2 Activities:Problem Description: Summary of customer's problems: The gantry intermittently loses power. It's a technology problem with the equipment itself: Yes. Troubleshooting Action: remote diagnosis and testing result: Engineer Chen Tao is reporting repair for the customer. It's very difficult to determine the problem PRS TR status: it's not through, ask engineer for maintenance on site. Repair Action: Suggest onsite service procedures: Suggest engineer to go to the hospital to check the error log and determine the problem. Teach the customer to use Philips WeChat to report for repair Safety Questions:Case: 0118800493 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Weng Jinming Last updated by: Weng Jinming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Problem Reported by Customer:Intermittent power failure SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support ----------------------------------------------------------------------------- | 3 |  | $Q7: (Confirmed reported problem, E-stop open intermittently) $Q8: (Close keyswitch and it was normal) $Q9: ( Reboot CT gantry) | [12 empty] | 118800493 | Fixed | 1567319 |  |  | Closed - Done | Closed - Done | [6 empty] | Others | [6 empty] | System Reboot | [5 empty] | Refresh |  |  | System issue | bar: No
1528 |  |  | China | Xuancheng City Orthopedic Hospital | 345042 | 4452800% | 47 | WeChat：02N/Ingenuity Core128/Elite China/instal patch | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Calibration/Adjustment Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ Safety Questions:Case: 0118799473 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Cao Last updated by: Liu Cao 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer's problems: running stuck Detailed description of Customer's problems or malfunctions: running stuck It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: Need to install patch Remote troubleshoot and test result: Install patch PRS TR status: NA Repair Action: Suggest onsite service procedures: Install patch SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Next Step Action:Zhou Bing ----------------------------------------------------------------------------- | 5 |  | $Q7: (Confirmed reported problem,CT host software crash） $Q8: ( check the system logger，found the software issue ) $Q9: (reinstall the host software , trouble disppeared.) | [12 empty] | 118799473 | Closed | 1565023 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | incorrect installation | [5 empty] | Others | [6 empty] | SW Crash | [5 empty] | Refresh |  |  | System issue |  |  |  | - | - | incorrect installation | 0 | bar: No
1529 |  |  | China | Medical center (contagious disease, He nan, Xu chang) | 40047 | 4452800% | 47 | 02N/Ingenuity CT 728326/ | BU Q&R Feedback: No Reason: Corrective Maintenance Safety Questions:Case: 0118799206 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Guoping Last updated by: Liu Guoping 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer's problems: ESTOP disconnected. Detailed description of Customer's problems or malfunctions: Twice Estop disconnected It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: Twice Estop disconnected Remote troubleshoot and test result: It's normal after re-keying PRS TR status: No PRS. Repair Action: Suggest onsite service procedures: Check the logger and confirm the problem. Problem Reported by Customer: 02N/Ingenuity CT 728326 Next Step Action: HUANG YONG SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support ----------------------------------------------------------------------------- | 3 |  | $Q7: (Confirm reported problem, estop open) $Q8: (On site checked the logger ,found relay open error， need to replace the relay) $Q9: (Replace the relay ) | [12 empty] | 118799206 | Closed | 1564625 |  |  | Investigation - Pending Translation | Investigation - Pending Translation | [6 empty] | Engineering | [6 empty] | Relay | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | Relay | Couch-HW-Relay | Couch-HW-Relay | NotEngineering-PartReplacement | Relay | Couch-HW | Couch-HW-Relay | bar: No
1530 |  |  | China | 济南市第四人民医院 | 40038 | 4452600% | 47 | 02N/Ingenuity CT 728326/add node | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: Instructed Customer Problem Reported by Customer : 02N/Ingenuity CT 728326/ Safety Questions : Case: 0118789101 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wennian Han Last updated by: Wennian Han 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: add node Troubleshooting Action: add workstation node Repair Action: guide customer to add workstation node, feedback transmission is normal Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: NA ----------------------------------------------------------------------------- | 5 | [14 empty] | 118789101 | Fixed | 1562265 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [19 empty] | Refresh |  |  | System issue | bar: No
1531 |  |  | China | Zhejiang Province Yueqing No. 1 People’s Hospital | 40054 | 4452500% | 47 | Application/Ingenuity CT 728326 | BU Q&R Feedback: No Reason: Application Support Safety Questions:Case: 0118782324 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingsong Last updated by: Lu Mingsong 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Problem Reported by Customer: 02N/Ingenuity CT 728326 ----------------------------------------------------------------------------- | 5 | [14 empty] | 118782324 | In Process | 1559229 |  |  | Evaluation | Evaluation | [6 empty] | Application | [6 empty] | Application Support | [5 empty] | Refresh |  |  | System issue | bar: No
1532 |  |  | China | Jiangsu University Affiliated Hospital | 345023 | 4450600% | 44 | 02N/Ingenuity Core128/Elite China/Can not boot | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : Can not boot SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Safety Questions:Case: 0118675308 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Zhao Shixin Last updated by: Zhao Shixin 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A External Remarks: 07.11.2021 13:50:04 UTC (86036528)"SUAT. Ready for Clinical Use local time" : "2021-11-07 10:45" External Remarks: 07.11.2021 14:18:05 UTC (86036528) Service details: Device fail to boot up normally. Onsite find CIRS is not turned on. Re do switch on and off procedure and train the operator on machine start up procedure, then restore. Service Result: The device returned to normal. Internal Remarks : 07.11.2021 14:21:15 UTC (86036528) RDI Accuracy: Yes $Q7: (Confirm reported problem, the gantry can't power on.) $Q8: (check the console found CIRS did not turned on.) $Q9:Turn on the CIRS. $Q10: (Test and inspection data refer to PA. Device returned to full functionality.) $Q11: Calibrated tools refer to PA. ----------------------------------------------------------------------------- | 2 |  | (86036528) Service details: Device fail to boot up normally. Onsite find CIRS is not turned on. Re do switch on and off procedure and train the operator on machine start up procedure, then restore. | [12 empty] | 118675308 | New | 1497474 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Others | [6 empty] | CIRS Power on | [5 empty] | Refresh |  |  | System issue | bar: No
1533 |  |  | China | Heilongjiang Province Huachuan County People's Hospital | 345018 | 4452400% | 47 | Application/Core128 | BU Q&R Feedback: No Reason: Application Support Problem Reported by Customer : Application Safety Questions:Case: 0118775224 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Pan Chuchu Last updated by: Pan Chuchu Next Step Action: Customer request second clinical training. Yang Jun 13817647801 yang.yj.jun@philips.com Guo Fan 13501285722 fan.guo@philips.com, follow up. ----------------------------------------------------------------------------- | 5 |  | Application Training 客户申请第二次临床培训， 杨俊 13817647801 yang.yj.jun@philips.com 郭凡 13501285722 fan.guo@philips.com 跟进 | [12 empty] | 118775224 | New | 1555527 |  |  | NAC | Evaluation | [6 empty] | Application | [6 empty] | APP Training | [5 empty] | Refresh |  |  | System issue | bar: No
1534 |  |  | China | Jiangsu Xuzhou Medical School Affiliated No. 3 Hospital | 40039 | 4452300% | 47 | 02G/Ingenuity CT 728326/Adapter failure | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer: 02G/Ingenuity CT 728326/Adapter failure Safety Questions:Case: 0118772508 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: shao jianan Last updated by: shao jianan 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 3 |  | 02G/Ingenuity CT 728326/Adapter failure | [12 empty] | 118772508 | Closed | 1553677 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632173 | [5 empty] | Engineering | [6 empty] | Adapter | [5 empty] | Refresh |  |  | System issue | System | Console-HW | Adaptor | Console-HW-Adaptor | Console-HW-Adaptor | CTUDT00632173 | Adaptor | Console-HW | Console-HW-Adaptor | bar: No
1535 |  |  | China | Jiangsu Province No. 2 Chinese Medicine Hospital | 345027 | 4452300% | 47 | 02N/Ingenuity Core128/Elite China/tube overheating | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ Next Step Action: Caller: Dr Fu: 013914749786 T2 Activities:Problem Description: Summary of customer's problems: Other types of error. Detailed description of Customer's problems or malfunctions: The customer said the tube is frequently running in overheat protection. It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: Log has no error. Oil cooling is normal, around 40 degrees all day. Remote troubleshoot and test result: Restart but didn't work. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: It is recommended to check if the customer set up too many slices in scan protocol. Verify if the protocol dose setting is too high in mAs or scan time is too long causing the tube heat capacity to reach the limit. We recommend spare parts 12NC and names: N/A Safety Questions:Case: 0118763700 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Bo Last updated by: Wang Bo 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Next Step Action:Qi liangdong ----------------------------------------------------------------------------- | 3 |  | $Q7: (Confirmed reported problem: Heat capacity warnimg sometimes.) $Q8: (Checked the error log, the temperature of oil flow is normal. Checked scan protocol, found mA value of one scan protocol is too higher.) $Q9: (Adjusted the scan parameter, keep monitoring.) | [12 empty] | 118763700 | In Process | 1552109 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Customer | [6 empty] | Tube Overheat | [5 empty] | Refresh |  |  | System issue | bar: No
1536 |  |  | China | Henan Province Puyang City People’s Hospital | 40031 | 4452300% | 47 | 02N/Ingenuity CT 728326/The tube overheating | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Safety Questions:Case: 0118762961 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingsong Last updated by: Lu Mingsong 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: The tube overheating Troubleshooting Action: The room temperature is too high. Advise customer to lower room temperature. Repair Action: The room temperature is too high.Advise customer to lower room temperature. Test/Inspection: Functional Check PASSED. The device returned to full functionality. Test Equipment: N/A Problem Reported by Customer: 02N/Ingenuity CT 728326 SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | Troubleshooting Action: The room temperature is too high. Advise customer to lower room temperature. Repair Action: The room temperature is too high.Advise customer to lower room temperature. Test/Inspection: Functional Check PASSED. The device returned to full functionality. | [12 empty] | 118762961 | Closed | 1551612 |  |  | Investigation | Investigation | [6 empty] | Customer | [6 empty] | Tube Overheat | [5 empty] | Refresh |  |  | System issue | bar: No
1537 |  |  | China | Inner Mongolia Autonomous Region IV Hospital | 40023 | 4452100% | 46 | 05K for IB/Ingenuity CT 728326/password expired | BU Q&R Feedback: No SAP Customer Symptom Code: MARF Customer Symptom Code: Monitor Alarm Failure Reason: Corrective Maintenance Engineer Failure Code: Incompatible Use Repair Activity Code: Instructed Customer Problem Reported by Customer: 05K for IB/Ingenuity CT 728326/ Field Service Engineer Notes:This PO is a Reference Number only Safety Questions:Case: 0118754238 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Password expired Troubleshooting Action: instruct customer change password, close case. Repair Action: instruct customer change password, close case. Test/Inspection: Functional Check PASSED. The device returned to full functionality. Test Equipment: N/A ----------------------------------------------------------------------------- | 3 |  | Repair Action: instruct customer change password, close case. | [12 empty] | 118754238 | Fixed | 1544053 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Others | [6 empty] | Password | [5 empty] | Refresh |  |  | System issue | bar: No
1538 |  |  | China | Jilin Province Zhaonan City TCM Hospital | 345006 | 4452100% | 46 | 05L/Ingenuity Core128/Elite China/artifact | BU Q&R Feedback: No Reason: Corrective Maintenance Safety Questions:Case: 0118754020 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer's problems: Other types of error. Detailed description of Customer's problems or malfunctions: Artifact It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: Ring artifact. Remote troubleshoot and test result: Detector channel failure PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Analyze data, find artifact location, mask channel SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support External Remarks: 22.11.2021 00:48:45 UTC (86039473) Service Details: There are artifacts in the images. Work on the bad detector and seal the bad channel. service result: device returns to normal working condition after service Problem Reported by Customer : 05L/Ingenuity Core128/Elite China/ Internal Remarks : 22.11.2021 00:48:48 UTC (86039473) RDI Accuracy: Yes $Q7:(Confirmed reported problem & the image has ring artifacts) $Q8: (Do the bad detector and analyze data,have bad channel) $Q9: (Shielding bad channels,cperform relative test,system is ok) $Q10: (Test and inspection data refer to: PA. Device returned to full functionality.) $Q11: Calibrated Tools name: refer to PA Next Step Action: Liu Yadong ----------------------------------------------------------------------------- | 3 |  | $Q7:(Confirmed reported problem & the image has ring artifacts) $Q8: (Do the bad detector and analyze data,have bad channel) $Q9: (Shielding bad channels,cperform relative test,system is ok) | [12 empty] | 118754020 | Closed | 1543949 |  |  | Investigation | Investigation | CTUDT00632110 | [5 empty] | Engineering | [6 empty] | Ring Artifact | [5 empty] | Refresh |  |  | System issue | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | DMS Module | DMS-HW | DMS-HW-DMS Module | bar: No
1539 |  |  | China | Heilongjiang Province Huachuan County People's Hospital | 345018 | 4452400% | 47 | 05K For IB/Core128/Consult warranty | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer:Consult warranty Safety Questions:Case: 0118775223 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Pan Chuchu Last updated by: Pan Chuchu 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action: Case#0118775223 Heilongjiang Province Huachuan County People's Hospital, Ingenuity Core128 EQ#87552930, customer consult warranty expiration date Teacher Zhang 15246304333 Next Step Action: Screenshot of warranty period has been sent to Dark Blue Trading Ltd. Co to call back customer to inform Next Step Action: Case# 0118775224 follow up, Yang, Jun <yang.yj.jun@philips.com> feedback: have arranged two trainings. I asked the trainer to explain to the hospital. Next Step Action: Case# 0118776997 followed up, Yang, Jun <yang.yj.jun@philips.com> already arranged for trainer to contact DEALER, follow up with 2nd follow up visit ----------------------------------------------------------------------------- | 5 | [14 empty] | 118775223 | Closed | 1555529 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [19 empty] | Refresh |  |  | System issue | bar: No
1540 |  |  | China | Jiangxi Province Children’s Hospital | 40041 | 4452000% | 46 | 02N/Ingenuity CT 728326/rotation noise | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 02N/Ingenuity CT 728326 SAP Transfer:Case Transferred to SAP T2 Activities:Problem Description: Summary of customer's problems: The device makes a strange sound. Detailed description of Customer's problems or malfunctions: NOISE IN ROTATION It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: NA Remote troubleshoot and test result: The customer informed the rack scan when rotating has noise, not when stationary, contacted the engineer, suspected tension wheel problems PRS TR status: it's not through, ask engineer for maintenance on site. Repair Action: Suggest onsite service procedures: Check the tension wheel, belt-related mechanical structure We recommend spare parts 12NC and names: " 453567003451 TENSIONER ASSY: 453566502101 SGL MTR SPINDLE ALIGN TOOL" Next Step Action: PPD 453566502101 1 & 453567003451 1 scheduled 72H to hospital, FSE Li Jiawei -86-18579112310 Safety Questions:Case: 0118752935 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created by: Wennian Han Last updated by: Wennian Han 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Dispatched for On site Support ----------------------------------------------------------------------------- | 3 |  | 453566502101 1 & 453567003451 $Q7:(Confirmed reported problem，gantry noise) $Q8:( Check rotation system, found the noise from tension pulley ) $Q9:( Replace the tension pulley solved the problem.) | [12 empty] | 118752935 | Closed | 1543530 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | ECR-077012 | [5 empty] | Engineering | [6 empty] | Tensioner | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Belt | Gantry-HW-Belt | Gantry-HW-Belt | ECR-077012 | Belt | Gantry-HW | Gantry-HW-Belt | bar: No
1541 |  |  | China | Jiangsu Xuzhou Medical School Affiliated No. 3 Hospital | 40039 | 4452000% | 46 | 02G/Ingenuity CT 728326/Gantry nosie | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 01/Ingenuity CT 728326 Safety Questions:Case: 0118752470 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created by: Wennian Han Last updated by: Wennian Han 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Gantry noise Troubleshooting Action: The gantry is noisy when rotating. Repair Action: Check for scraping: the belt, tension pulley, etc. ----------------------------------------------------------------------------- | 3 |  | T2 Activities:Problem Description: Gantry noise Troubleshooting Action: The gantry is noisy when rotating. Repair Action: Check for scraping: the belt, tension pulley, etc. | [12 empty] | 118752470 | In Process | 1543427 |  |  | Evaluation | Evaluation | ECR-077012 | [5 empty] | Engineering | [6 empty] | Tensioner | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Belt | Gantry-HW-Belt | Gantry-HW-Belt | ECR-077012 | Belt | Gantry-HW | Gantry-HW-Belt | bar: No
1542 |  |  | China | Heilongjiang Province Huachuan County People's Hospital | 345018 | 4451900% | 46 | Cannot hear operator voice in Gantry room, | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Replace Part SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support External Remarks: 21.11.2021 07:09:55 UTC (86039442) Service details: check ct-box cable, adjust sound volume Service Result: The device returned to normal. Internal Remarks : 21.11.2021 07:09:57 UTC (86039442) "RDI Accuracy:Yes" $Q7: (Confirmed reported problem: The horn doesn't sound ) $Q8: (check cable ok) $Q9:(adjust voice) $Q10: (Test and inspection data refer to: PA. The device returned to full functionality. $Q11 Calibrated Tools name (no) Problem Reported by Customer: The horn doesn't sound Next Step Action: PPD451213440782*1 have arranged for 24h delivery, please dispatch FSE: Jin Jun-86-13796075213 Safety Questions:Case: 0118746499 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Gaoge Last updated by: Wang Gaoge 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer's problems: Other types of error. Detailed description of Customer's problems or malfunctions: There was no sound in the rack when speaking It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: Use CT BOX to speak, no sound. Auto audio playback is normal Remote troubleshoot and test result: Consider CT BOX failure. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: 1. check audio playback; 2. replace CT box We recommend spare parts 12NC and names: "459801705022 CT BOX ASSEMBLY_BEST " Test/Inspection: N/A Test Equipment: N/A ----------------------------------------------------------------------------- | 3 |  | $Q7: (Confirmed reported problem&The horn doesn't sound ) $Q8: (check cable ok） $Q9:(adjust voice) 459801705022 CT BOX ASSEMBLY_BEST | [12 empty] | 118746499 | Closed | 1541375 |  |  | Investigation | Investigation | CTUDT00632185 | [5 empty] | Engineering | [6 empty] | CT Box | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | CT Box | Gantry-HW-CT Box | Gantry-HW-CT Box | CTUDT00632185 | CT Box | Gantry-HW | Gantry-HW-CT Box | No: No; y: 0; y: 0; y: 0; y: 0; y: 0; No: No
1543 |  |  | China | Shandong Province Tai'an City T.C.M. No. 2 Hospital | 345037 | 4451900% | 46 | System opens estop intermittently; Adjusted the gantry gullwing switch position to resolve. | BU Q&R Feedback: No Reason: Corrective Maintenance Safety Questions:Case: 0118746354 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingsong Last updated by: Lu Mingsong 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer's problems: The gantry lost power. Detailed description of Customer's problems or malfunctions: The gantry lost power. It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: MCO open Remote troubleshoot and test result: N/A PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Adjust Gullwing switch position. We recommend spare parts 12NC and names: N/A Memo of Customer request: None Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ Next Step Action: Customer said to call back Teacher Chen: 18560023293 Next Step Action: power on failure intermittently SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Internal Remarks : 19.11.2021 11:57:21 UTC (86038373) “RDI Accuracy: YES" $Q7: (Confirmed reported problem, close E-stop failed intermittently $Q8: ( Check gantry log, found "MCO open" error. Push gantry front cover and it can be closed normally) $Q9: ( Adjust the position of gullwing) $Q10: (Test and inspection data refer to: PA. Device returned to full functionality) $Q11: Calibrated Tools name ( N/A) ^Current Tube installed date: N/A ^Current Tube usage: N/A External Remarks : 19.11.2021 11:30:32 UTC (86038373) Fault Description: Power on by turning key intermittently failed Service content: Adjust gull wing switch position Service outcome: Device returns to normal working condition after servicing. ----------------------------------------------------------------------------- | 3 |  | $Q7: (Confirmed reported problem, close E-stop failed intermittently $Q8: ( Check gantry log, found "MCO open" error. Push gantry front cover and it can be closed normally) $Q9: ( Adjust the position of gullwing) | [12 empty] | 118746354 | Closed | 1541213 |  |  | Closed - Done | Closed - Done | AssemblyAdjust | [5 empty] | Others | [6 empty] | Gullwing | [5 empty] | Refresh |  |  | System issue |  |  |  | - | - | AssemblyAdjust | 0 | bar: No
1544 |  |  | China | Guangxi Province Teng County People’s Hospital | 40036 | 4451800% | 46 | 01/Ingenuity CT 728326/The image is not very clear | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : The image is not very clear Next Step Action: OB Zhangyuanyuan emailed to inform that Teacher Tan's feedback during the follow-up visit is as follows: 1. After the scan is finished, there is no one-button Couch In operation. 2. The diagnosing doctor said the images are not very clear and a bit coarse; they are not as bright as Siemens. 3. We need a clinical teacher to come out and train them again. There was no medical case when I came. I still don't understand it after saying that. Please help with the feedback, and please ask the online engineer to check if it is possible to instruct the customer remotely.Thanks! Next Step Action: I called and texted CT Application Specialist Deng Mingming at 18929200533 to follow up and reply to the customer. Next Step Action: Regarding the third point, I emailed ASP YingLian to follow up with Business. Safety Questions:Case: 0118739639 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Zhou Xiaohong Last updated by: Zhou Xiaohong 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action: Deng Mingming (18929200533) said that the customer rejected Teacher Deng's callback to handle it remotely and requested to have a clinical teacher on site. Teacher Deng has confirmed with the training manager, the site has had on-site training twice and there is no need to go on site again. Next Step Action: ASP YingLian replied: This device is still within the Philips warranty period. Already followed up by Manager Lin Ruixin. T2 Activities:Problem Description: Summary of customer's problems: Other types of error. Detailed description of Customer's problems or malfunctions: "During the follow-up visit, Teacher Tan said the problem is as follows: 1. After the scan is finished, there is no one-button Couch In operation. 2. The diagnosing doctor said the images are not very clear and a bit coarse; they are not as bright as Siemens. 3. We need a clinical teacher to come out and train them again. There was no medical case when I came. I still don't understand it after saying that." It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: Follow-up feedback Remote troubleshoot and test result: "1. Inform the customer that for the CT box from the last century, the one Couch In button doesn't exist, but on both sides of the gantry there is one Couch In button. 2. Regarding the feedback's second item, transfer to operations, the application engineer will call back to process it. 3. For the third item, please transfer to Business to follow up." PRS TR status: it's not through, ask engineer for maintenance on site. ----------------------------------------------------------------------------- | 3 | [14 empty] | 118739639 | In Process | 1538261 |  |  | Evaluation | Evaluation | [6 empty] | Application | [6 empty] | Application Support | [5 empty] | Refresh |  |  | System issue | bar: No
1545 |  |  | Spain | Comercial Centro Medico | 40042 | 4450200% | 44 | Error de reconstructor | Work Order #: WO-04686646 05-Nov-2021 [REDACTED] The system meets the specification for the performed service and is returned to use. Problem Description by Engineer: Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Regulatory Questions : N/A Diagnostic performed by Engineer: Reconstructor error message is reproduced. It stops sending images during reconstruction. Resolution: After checking the error log of the different components, I do not find an element in the hardware that could be causing the error, so I re-installed the software on both servers of the reconstruction system. After reloading the software, I check the correct operation of the system. The system works correctly and is ready for use. ----------------------------------------------------------------------------- Work Order #: WO-04683286 02-Nov-2021 [REDACTED] The system meets the specification for the performed service and is returned to use. Regulatory Questions : N/A Problem Description by Engineer: Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Follow Up Required: Check operation in S2 IMR Diagnostic performed by Engineer: When launching a patient study, the error "restart reconstruction system" is displayed. The S2 reconstructor of the IMR crashes. I reset it and works correctly. After performing a patient it crashes again. There is no network connection from the HOST. It is restarted and fails again. I do not find any evident error in the error logs of the system or the reconstructors. The monitor and keyboard are connected, awaiting for the error to reappear, but it does not. The system is tested several times and does not fail. ----------------------------------------------------------------------------- | 2 |  | Diagnostic performed by Engineer: When launching a patient study, the error "restart reconstruction system" is displayed. The S2 reconstructor of the IMR crashes. I reset it and works correctly. After performing a patient it crashes again. There is no network connection from the HOST. It is restarted and fails again. I do not find any evident error in the error logs of the system or the reconstructors. The monitor and keyboard are connected, awaiting for the error to reappear, but it does not. The system is tested several times and does not fail. | [12 empty] | 118650957 | Closed | 1480937 |  |  | Investigation | Investigation | part replace | [5 empty] | Others | Net Switch | [5 empty] | Net Switch | [5 empty] | Refresh |  |  | System issue |  |  |  | - | - | part replace | 0 | bar: No
1546 |  |  | China | 浙江省乐清第一人民医院 | 40054 | 4451700% | 46 | wechat：02N/Ingenuity CT 728326/software running slowly | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : wechat：02N/Ingenuity CT 728326 Safety Questions : Case: 0118730781 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: software running slowly Troubleshooting Action: software running slowly ,after customer add menory for test, speed will change fast. so customer request upgrade memory. nemo it . Repair Action: nemo memory . Test/Inspection: N/A Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | software running slowly ,after customer add menory for test, speed will change fast. so customer request upgrade memory. nemo it . | [12 empty] | 118730781 | In Process | 1534771 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | MemoryReplacement | [5 empty] | Engineering | [6 empty] | Memory | [5 empty] | Refresh |  |  | System issue | System | Console-HW | MemoryReplacement | Console-HW-MemoryReplacement | Console-HW-MemoryReplacement | MemoryReplacement | MemoryReplacement | Console-HW | Console-HW-MemoryReplacement | bar: No
1547 |  |  | China | Nanjing Medical University Affiliated Yifu Hospital | 345028 | 4451600% | 46 | 02N/Ingenuity Core128/Elite China/couch noise | BU Q&R Feedback: No Reason: Corrective Maintenance Safety Questions:Case: 0118720842 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingsong Last updated by: Lu Mingsong 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer's problems: The vertical movement of the couch is noisy. Detailed description of Customer's problems or malfunctions: The feedback: sounded like friction It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: N/A Remote troubleshoot and test result: N/A PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Check the screw and horizontal glide on the site. If necessary add lube. We recommend spare parts 12NC and names: N/A Memo of Customer request: None SAP Transfer:Dispatched for On site Support SAP Transfer:Case Transferred to SAP Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ Next Step Action:YAO Liu Bin ----------------------------------------------------------------------------- | 3 |  | $Q7: (Confirm reported problem. Patient table has noise when moves vertically.) $Q8: (Reviewed the error log and found the ball screw and slide rails are leak of oil.) $Q9: (Lubricate the ball screw and slide rails.) | [12 empty] | 118720842 | In Process | 1528788 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | E-125059 | [5 empty] | Engineering | [6 empty] | Lubrication | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | Lubricant | Couch-HW-Lubricant | Couch-HW-Lubricant | E-125059 | Lubricant | Couch-HW | Couch-HW-Lubricant | bar: No
1548 |  |  | China | Heilongjiang Province Huachuan County People's Hospital | 345018 | 4452200% | 47 | praise/Ingenuity Core128/Elite China/engineer | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer: praise Safety Questions:Case: 0118755415 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Qian Wang Last updated by: Qian Wang 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action: Praise engineer Jin Jun, each time when report for repair, he arrive on site timely, and can solve the problem. he is very professional with good attitude. Email jingwen to follow up ----------------------------------------------------------------------------- | 3 | [14 empty] | 118755415 | Closed | 1546758 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [19 empty] | Refresh |  |  | System issue | bar: No
1549 |  |  | China | Jiangxi Province Yugan Chu East Hospital (New Branch) | 337500 | 4451600% | 46 | Wechat:02N/Ingenuity Core128/Elite China/Unable to log in after restarting the system | BU Q&R Feedback: No SAP Customer Symptom Code: BOOT Customer Symptom Code: Boot up/Power-up Problem Reason: Corrective Maintenance Engineer Failure Code: Electrical/Electronic Failure Repair Activity Code: Reboot/Restart Problem Reported by Customer: Unable to log in after restarting the system Safety Questions:Case: 0118720501 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Weng Jinming Last updated by: Weng Jinming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Troubleshooting Action: instruct customer to restart the system Repair Action: instruct customer to restart the system Test/Inspection: Functional Check PASSED. The device returned to full functionality. Test Equipment: N/A ----------------------------------------------------------------------------- | 3 |  | Troubleshooting Action: instruct customer to restart the system Repair Action: instruct customer to restart the system | [12 empty] | 118720501 | Fixed | 1528562 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632553 | [5 empty] | Engineering | [6 empty] | System Reboot | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Blue-Screen | Console-SW-Blue-Screen | Console-SW-Blue-Screen |  | Blue-Screen | Console-SW | Console-SW-Blue-Screen | bar: No
1550 |  |  | China | Heilongjiang Province Huachuan County People's Hospital | 345018 | 4451500% | 46 | 05K FOR IB/Ingenuity Core128/Elite China/can not burn | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 05K FOR IB/Ingenuity Core128/Elite China Next Step Action: Wait for customer to confirm host model in the afternoon Safety Questions:Case: 0118712859 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created by: Wennian Han Last updated by: Wennian Han 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | 05K FOR IB/Ingenuity Core128/Elite China/can not burn | [12 empty] | 118712859 | Fixed | 1523118 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Service | [6 empty] | Burn | [5 empty] | Refresh |  |  | System issue |  |  |  | - | - | 0 | 0 | bar: No
1551 |  |  | China | Guangxi Province Teng County People’s Hospital | 40036 | 4451500% | 46 | 02N/Ingenuity CT/has prodigious sound | BU Q&R Feedback: No Reason: Corrective Maintenance SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Next Step Action: Manual Case#PCC2020111411 Caller: CT Teacher Tan. 15878407675 Hospital: Guangxi Teng County People's Hospital Device model: Ingenuity CT 728326 EQ# 87185445 SN#40036 Fault Description: Q1&Q2: Follow up contact: CT Teacher Tan. 15878407675 Warranty info: 02N SITE1:FSE Huang Yao pei 18128388600 Dispatch FSE/ASP: FSE Huang Yao pei 18128388600 RSE Wang Jianqiang: Customer issues: After the maintenance, ding dong noise is often heard. Detailed description of Customer's problems or malfunctions: After the maintenance, ding dong noise is often heard. It's a technology problem with the equipment itself: Yes. Key error code or symptom: After the maintenance, ding dong noise is often heard. Remote troubleshoot and test result: Suspect a cable is poorly in contact. Check cause of the abnormal sound. The cable needs to be plugin tight. Also the cable from the back of the rack to the monitor is not plugged in securely. PRS TR status: it's not through, ask engineer for maintenance on site. Suggest to go to the site to repair: Check the wiring and unplug and re plug tightly. Problem Reported by Customer: Has prodigious sound Safety Questions:Case: 0118712610 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Pan Chuchu Last updated by: Pan Chuchu 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 3 |  | $Q6: (Gantry noise and ECG signal lost) $Q7: (Confirm reported problem & Found ibox ECG signal problem.) $Q8: (Found ibox convert box connected problem & Need to adjust ibox cable.) $Q9: (Readjust ibox cable.) | [12 empty] | 118712610 | New | 1522889 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | IBOX | [5 empty] | Engineering | [6 empty] | IBOX | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | IBOX | Gantry-HW-IBOX | Gantry-HW-IBOX | IBOX | IBOX | Gantry-HW | Gantry-HW-IBOX | bar: No
1552 |  |  | China | Jiangsu University Affiliated Hospital | 345023 | 4451100% | 45 | 02N/Ingenuity Core128/Elite China/spooler reset | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ T2 Activities:Problem Description: Need reset spooler Troubleshooting Action: Need reset spooler Repair Action: Remotely run script Safety Questions:Case: 0118697950 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Cao Last updated by: Liu Cao 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 3 |  | Troubleshooting Action: Need reset spooler Repair Action: Remotely run script | [12 empty] | 118697950 | Fixed | 1513961 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Reset Spooler | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | Relay | Couch-HW-Relay | Couch-HW-Relay | NotEngineering-PartReplacement | Relay | Couch-HW | Couch-HW-Relay | bar: No
1553 |  |  | China | Jiangsu University Affiliated Hospital | 345023 | 4451100% | 45 | Wechat:02N/Ingenuity Core128/Elite China/Unable to transfer | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: Instructed Customer Problem Reported by Customer: Unable to transfer Safety Questions:Case: 0118697789 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Bai Jingwen Last updated by: Bai Jingwen 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities: Problem Description: Unable to transfer Troubleshooting Action: Unable to transfer image Repair Action: reboot HOST computer, then transmission is normal Test/Inspection: Functional Check PASSED. The device returned to full functionality. Test Equipment: NA SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | Troubleshooting Action: Unable to transfer image Repair Action: reboot HOST computer, then transmission is normal | [12 empty] | 118697789 | Fixed | 1513926 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | Image Transfer | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | bar: No
1554 |  |  | China | Inner Mongolia Alxa Right Banner People's Hospital | 345031 | 4451100% | 45 | 02N/Ingenuity Core128/Elite China/estop close failed | BU Q&R Feedback: No SAP Customer Symptom Code: BOOT Customer Symptom Code: Boot up/Power-up Problem Reason: Corrective Maintenance Engineer Failure Code: Electrical/Electronic Failure Repair Activity Code: Instructed Customer Problem Reported by Customer :Ingenuity Core128/Elite China Safety Questions:Case: 0118697661 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: estop close failed Troubleshooting Action: check errorlogger, no error found, instruct customer check emergency button, and close estop, system is ok. close case. Repair Action: nothing to do, close case. Test/Inspection: N/A Test Equipment: N/A ----------------------------------------------------------------------------- | 3 |  | T2 Activities:Problem Description: estop close failed Troubleshooting Action: check errorlogger, no error found, instruct customer check emergency button, and close estop, system is ok. close case. Repair Action: nothing to do, close case. | [12 empty] | 118697661 | Fixed | 1513877 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Others | [6 empty] | System Reboot | [5 empty] | Refresh |  |  | System issue | bar: No
1555 |  |  | China | Caofeidian District Hospital | 345016 | 4451000% | 45 | 05K For IB/Ingenuity Core128/Elite China/Unable to burn | BU Q&R Feedback: No Reason: Corrective Maintenance T2 Activities:Problem Description: Summary of customer's problems: Other types of error. Detailed description of Customer's problems or malfunctions: Customer problem It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: NA Remote troubleshoot and test result: The customer refused to transfer and did not answer the phone.Contacted the engineer to tell him that there is a disc burning recognition issue. The customer needs on site assistance. PRS TR status: No PRS. Repair Action: Suggest onsite service procedures: Confirm if the customer's disc burning procedure is correct; check Rom drive, software; rule out disc issue Safety Questions:Case: 0118692235 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created by: Wennian Han Last updated by: Wennian Han 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Internal Remarks : 11.11.2021 15:07:20 UTC (86038892) "RDI Accuracy:Yes" $Q7: Confirm reported problems:system need to change configuration. $Q8: Trouble shooting steps:change the film configuration. $Q9: Function returns to normal after the of configuration is changed. $Q10: Test and Inspection data refer to PA. Device status:Device returned to full functionally $Q11: N/A Problem Reported by Customer: Unable to burn External Remarks: 11.11.2021 10:22:45 UTC (86038892) service details: machine burn failure. After the process, it is back to normal. service result: device returned to normal working condition after repairing ----------------------------------------------------------------------------- | 3 |  | $Q7: Confirm reported problems:system need to change configuration. $Q8: Trouble shooting steps:change the film configuration. $Q9: Function returns to normal after the of configuration is changed. | [12 empty] | 118692235 | Closed | 1510753 |  |  | Evaluation | Evaluation | [6 empty] | Service | [6 empty] | Burn | [5 empty] | Refresh |  |  | System issue |  |  |  | - | - | 0 | 0 | bar: No
1556 |  |  | China | Jiangsu Xuzhou Medical School Affiliated No. 3 Hospital | 40039 | 4451800% | 46 | 02G/Ingenuity CT 728326/add dicom node | BU Q&R Feedback: No SAP Customer Symptom Code: CENV Customer Symptom Code: Customer Environmental Concern Reason: Corrective Maintenance Engineer Failure Code: Environmental Cause Repair Activity Code: Instructed Customer Problem Reported by Customer: 02G/Ingenuity CT 728326 Safety Questions:Case: 0118740003 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Gaoge Last updated by: Wang Gaoge 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Customer need to add a printer Troubleshooting Action: Instructed the customer to add printers. Need to find the device administrator account and password. Repair Action: Instructed the customer to add the printer. Test/Inspection: N/A Test Equipment: N/A ----------------------------------------------------------------------------- | 3 | [14 empty] | 118740003 | Fixed | 1538350 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [19 empty] | Refresh |  |  | System issue | bar: No
1557 |  |  | China | Medical center (contagious disease, He nan, Xu chang) | 40047 | 4451000% | 45 | 02N/Ingenuity CT 728326/application hanged up | BU Q&R Feedback: No SAP Customer Symptom Code: HANG Customer Symptom Code: Hang/Crash Reason: Corrective Maintenance Engineer Failure Code: Environmental Cause Repair Activity Code: Instructed Customer SAP Transfer:Case Transferred to SAP Safety Questions:Case: 0118690732 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Gaoge Last updated by: Wang Gaoge 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Problem Reported by Customer: 02N/Ingenuity CT 728326 T2 Activities:Problem Description: Software stuck when boot up, fail to enter interface Troubleshooting Action: Instruct customer to log off current user, log in again, then restore Repair Action: Instruct customer to log in again Test/Inspection: N/A Test Equipment: N/A ----------------------------------------------------------------------------- | 3 |  | T2 Activities:Problem Description: Software stuck when boot up, fail to enter interface Troubleshooting Action: Instruct customer to log off current user, log in again | [12 empty] | 118690732 | Fixed | 1509836 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632553 | [5 empty] | Engineering | [6 empty] | SW Stuck | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Blue-Screen | Console-SW-Blue-Screen | Console-SW-Blue-Screen |  | Blue-Screen | Console-SW | Console-SW-Blue-Screen | bar: No
1558 |  |  | China | Fujian Province Medical University Affiliated No. 1 Hospital | 345012 | 4450900% | 45 | 02N/Ingenuity Core128/Elite China/right panel key fault | BU Q&R Feedback: No Reason: Corrective Maintenance T2 Activities:Problem Description: Summary of customer's problems: Control panel button is not working. Detailed description of Customer's problems or malfunctions: Right control panel buttons are not good It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: Right control panel buttons are not good Remote troubleshoot and test result: Buttons are not usable. The customer needs to replace it PRS TR status: No PRS. Repair Action: Suggest onsite service procedures: 1. Check the right control panel.2. Replace the right control panel according to the manual We recommend spare parts 12NC and names: "CONTROL PANEL ASSEMBLY-FRONT RH 459800484511" Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ Next Step Action: FSE Li Feng 018650369579 call in Next Step Action: Calling back was successful. Safety Questions:Case: 0118685700 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Jianqiang Last updated by: Wang Jianqiang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Next Step Action : PPD 459800484511 *1, arranged 24H delivery, Field Service Engineer Li Feng - 86-18650369579 ----------------------------------------------------------------------------- | 3 |  | CONTROL PANEL ASSEMBLY-FRONT RH 459800484511 | [12 empty] | 118685700 | In Process | 1506266 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632181 | [5 empty] | Engineering | [6 empty] | GCP | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Gantry Panel | Gantry-HW-Gantry Panel | Gantry-HW-Gantry Panel | CTUDT00632181 | Gantry Panel | Gantry-HW | Gantry-HW-Gantry Panel | No: No; y: 0; y: 0; y: 0; y: 0; y: 0; No: No
1559 |  |  | China | Suzhou Guang ci Tumor Hospital Co., Ltd | 345019 | 4450900% | 45 | The images has artifacts. Performed the air calibration again to resolve. | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Calibration/Adjustment Internal Remarks : 11.11.2021 16:01:02 UTC (86090280) RDI Accuracy: Yes $Q7: (Confirm reported problem,the images has artifacts.) $Q8: (Check the air calibration table and found has error.) $Q9: (perform the air calibration according to service manual) $Q10: (Test and inspection data refer to PA. Device returned to full functionality.) $Q11: Calibrated tools refer to PA. SAP Transfer:Case Transferred to SAP Problem Reported by Customer:image fault Safety Questions:Case: 0118684258 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Zixuan Last updated by: Wang Zixuan 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Dispatched for On site Support T2 Activities:Problem Description: Summary of customer's problems: Image density is nonuniform. Detailed description of Customer's problems or malfunctions: present in all parts, air calibration has been done this morning It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: The compensator tool did not prompt for cracks Remote troubleshoot and test result: Confirm raw data and calibrate PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Confirm raw data and calibrate Next Step Action: Xu Haotian External Remarks : 09.11.2021 11:23:38 UTC (86090280) Service details: There are artifacts to the images. Re-run air calibration and it restore. Service Result: The device returned to normal. ----------------------------------------------------------------------------- | 3 |  | $Q7: (Confirm reported problem,the images has artifacts.) $Q8: (Check the air calibration table and found has error.) $Q9: (perform the air calibration according to service manual) | [12 empty] | 118684258 | In Process | 1505450 |  |  | Closed - Done | Closed - Done | ShouldbeCustomer | [5 empty] | Engineering | [6 empty] | Image Artifact | [5 empty] | Refresh |  |  | System issue | System | Console-SW | ShouldbeCustomer | Console-SW-ShouldbeCustomer | Console-SW-ShouldbeCustomer | ShouldbeCustomer | ShouldbeCustomer | bar: No
1560 |  |  | United States | Westgate Hospital | 40001 | 4451800% | 46 | 0699CFXXX customer requesting support - part id - standoff | Internal Remarks : Interaction Type: New Case Caller: Justin Bertino IP & Issue: 82488278 Event Type: Part ID Service Request: Remote Technical Did Customer Decline Remote Support? No, not entitled Preferred Service Time: n/a Entitlement Options: Free Remote Reason for Entitlement Claim: part id Action Taken: created case, wo and routed call to rtech for part id callback Next Step Action: N/A Verbal PO Number: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- Work Order #: WO-04798258 Miller, John M The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Problem description by engineer : Customer needs part number of the stand off for the Footswitch I/F board. When was the first occurrence? Within... : 18-Nov-2021 How often is issue occurring? : First occurrence Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: BioMed Device Use: No patient involved Expected and Unexpected results: None Patient Impact/Outcome: Not provided User Impact: None Current Software Version: N/A Diagnostic performed by Engineer : Collect information from customer and view log files. Internal Comments : Error # and/or description of error : Research part number. Repair Action : Researched parts number of the Stand off for the Foot Switch Interface board. The customer is okay with closing the case. 459800088051 - Standoff-Hex M/F 8-32 X2 Resolution : Researched parts number of the Stand off for the Foot Switch Interface board. The customer is okay with closing the case. 459800088051 - Standoff-Hex M/F 8-32 X2 ----------------------------------------------------------------------------- | 5 | [14 empty] | 118744103 | Incomplete | 1536147 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [19 empty] | Refresh |  |  | System issue | bar: No
1561 |  |  | China | Guangxi Province Teng County People’s Hospital | 40036 | 4450700% | 44 | 02N/Ingenuity CT 728326/power off | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer:power off Safety Questions:Case: 0118675996 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Qian Wang Last updated by: Qian Wang 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action: huang yaopei SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support ----------------------------------------------------------------------------- | 3 |  | $Q6: (Power up problem) $Q7: (Confirm reported problem & Found local power problem.) $Q8: (Found power up problem . & Need to check power resource.) $Q9: (Check power input and restart system.) | [12 empty] | 118675996 | New | 1498428 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Others | [6 empty] | System Reboot | [5 empty] | Refresh |  |  | System issue | bar: No
1562 |  |  | United States | Philips Medical Capital | 40028 | 4450400% | 44 | 0199CDN84 Recon box error | Work Order #: WO-04699813 10-Nov-2021 Abbas Mohamed Abbas The system meets the specification for the performed service and is returned to use Tools & Equipment : MULTIMETER,2023-05-31 00:00:00 Resolution : Reset all boards for recon. Verify recon booting up. verify acquisition. verify QC Problem Description by Engineer : Problem description by engineer : Recon box error When was the first occurrence? Within... : 04-Nov-2021 How often is issue occurring? : Continuously Malfunction area : Recon Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: Not provided Patient Impact/Outcome: No further information is available as per the customer User Impact: Cannot use system to scan Current Software Version: Not provided ----------------------------------------------------------------------------- Work Order #: WO-04699406 Ali, Adam The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Internal Comments : Error # and/or description of error : RECON box error Repair Action : Check / repair power button on CIRS Run CIRS diagnostics Diagnostic performed by Engineer : Recon box error on the system. Customer power cycled the system, found as soon as the power button on CIRS is released it will turn itself off. Problem Description by Engineer : Problem description by engineer : Recon box error When was the first occurrence? Within... : 04-Nov-2021 How often is issue occurring? : Continuously Malfunction area : Recon Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: Not provided Patient Impact/Outcome: No further information is available as per the customer User Impact: Cannot use system to scan Current Software Version: Not provided Follow Up Required : Error # and/or description of error : RECON box error Repair Action : Check / repair power button on CIRS Run CIRS diagnostics Technician Name : Ali, Adam ----------------------------------------------------------------------------- | 2 |  | Reset all boards for recon. Verify recon booting up. verify acquisition. verify QC | [12 empty] | 118663947 | Incomplete | 1487783 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | ReconBox | [5 empty] | Engineering | [6 empty] | CIRS HW | [5 empty] | Refresh |  |  | System issue | System | Console-HW | ReconBox | Console-HW-ReconBox | Console-HW-ReconBox | ReconBox | ReconBox | bar: No
1563 |  |  | China | Shaanxi Province Ankang T.C.M. Hospital | 345041 | 4450700% | 44 | 02N/Ingenuity Core128/Elite China/estop open | BU Q&R Feedback: No Reason: Corrective Maintenance Internal Remarks : 15.11.2021 04:27:45 UTC (86039086) Q7: (Confirmed reported problem. & The Gantry shut down. $Q8: (Check the error log was not found. Check the table & We need to calibrate table) $Q9: (1. calibrate table2. Test use for CT was OK. $Q10: (Test and inspection data refer to: PA. & Device returned to full functionality.) $Q11: NA External Remarks: 15.11.2021 04:27:42 UTC (86039086) Service Reason: The gantry powered off. Service Details: Instructed the customer to calibrate. Result of service: device operating normally Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ T2 Activities:Problem Description: Summary of customer's problems: ESTOP disconnected. Detailed description of Customer's problems or malfunctions: The frame lost power when moving the bed board horizontally. It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: Confirmed the fault with the customer.PRS Password expiredPlease reset SSH2 FTP password onsite Remote troubleshoot and test result: Restart but didn't work. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: It is recommended to check scan bed horizontal position readings. Re perform horizontal position calibration. Based on the error report, determine if clutch is suction closed well. If necessary replace the absolute and relative encoder to eliminate the problem. We recommend spare parts 12NC and names: PPD 453566500911, 45356500901 SAP Transfer:Case Transferred to SAP Next Step Action: PPD 453566500901 *1 +453566500911* pending to confirm the replacement number, Zhang Yifei-86-17621702391 SAP Transfer:Dispatched for On site Support Safety Questions:Case: 0118675529 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: han susu Last updated by: han susu 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 3 |  | The frame lost power when moving the bed board horizontally. Q7: (Confirmed reported problem. & The Gantry shut down. ) $Q8: ( Check the error log was not found. Check the table & We need to calibrate table) 453566500901 *1+453566500911* 待确认替待号 | [12 empty] | 118675529 | In Process | 1497632 |  |  | Investigation | Investigation | PartReplacement | [5 empty] | Others | [6 empty] | E-Stop | [5 empty] | Refresh |  |  | System issue |  |  |  | - | - | PartReplacement | 0 | bar: No
1564 |  |  | Japan | Social Medical Corporation BOKOI, Nikko Memorial Hospital | 341011 | 4451100% | 45 | 04DS: KK: Ring artifact occurred during coronary shooting. | Work Order #: WO-04741237 11-Nov-2021 Konishi, Hideo The system meets the specification for the performed service and is returned to use. Resolution: [Work details] </br> <br> signals were not received at all from the 2 channels of Detector Module 25. I replaced Detector Module 25 and that improved the situation. After the replacement, I ran a test scan and confirmed that there were no problems. Internal Comments: [CX related] Did you agree on how to complete the report with the person in charge of devices? <Select one of the following> Yes. I completed normal reporting that day. [Work report (paper) submission] Did you submit a work report on paper? Yes Q1. How did the examination finish? Select from the following: *Because the answer is No for Q1, not applicable Q2. Were there any mistaken findings (misdiagnosis, mistaken treatment) due to the problem? None. Q3. Were there any artifacts? None. Q4. Was there any unintended unnecessary exposure? None. Resolution: "Dead channel" was confirmed in Mod25 during a bad detector test. Parts Consumed : 1.00000,459800911271,TDMS V2 M.Det Class A FRU Problem Description by Engineer: A coronary scan produced a ring artifact. Regulatory Questions : N/A ----------------------------------------------------------------------------- Work Order #: WO-04741195 Kojima, Kenichi The system does not meet the specification for the performed service and is not returned to use Internal Comments: [Subject] 04DS: KK: Mr. Nishiya: A coronary scan produced a ring artifact. Shipped part: 459800911271:TDMS V2 M.Det Class A FRU QTY:1 ZREP/Traceable --None-- Dispatched: 1 RC FSE Registration Information: Name of FSE 1 who visited the site: Hideo Konishi FSE 1 visit date and time: 2021/11/1112:00 FSE 1 hours worked: 4 Dispatch instructions: Dispatch instruction #1: Shipping conditions: CPU (remove from warehouse, no shipping arrangements) Delivery date: Must arrive by 10:00 on November 11, 2021. Shipping destination: Sapporo warehouse. FSE will collect. To be collected by: Hideo Konishi Phone number: 080-1605-0255 Other: Person making request: Kenichi Kojima Problem Description by Engineer: A coronary scan produced a ring artifact. Diagnostic performed by Engineer: The ring artifact appears to have been caused by a problem with the module. An FSE will respond. Follow Up Required : * ----------------------------------------------------------------------------- | 2 |  | 459800911271:TDMS V2 M.Det Class A FRU QTY:1 | [12 empty] | 118697409 | Closed | 1509826 |  |  | Investigation | Investigation | CTUDT00632110 | [5 empty] | Engineering | [6 empty] | Ring Artifact | [5 empty] | Refresh |  |  | System issue | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | DMS Module | DMS-HW | DMS-HW-DMS Module | bar: No
1565 |  |  | China | Suzhou Guang ci Tumor Hospital Co., Ltd | 345019 | 4450500% | 44 | 02N/Ingenuity Core128/Elite China/Image contrast | BU Q&R Feedback: No Reason: Corrective Maintenance Safety Questions:Case: 0118668075 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Jianqiang Last updated by: Wang Jianqiang 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Problem Reported by Customer: Image contrast T2 Activities:Problem Description: Summary of customer's problems: Other types of error. Detailed description of Customer's problems or malfunctions: Image quality is low and contrast test exceeds range. Need an engineer to calibrate It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: Image contrast is low, failed. Remote troubleshoot and test result: There is a small problem with image quality PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: On-site calibration of images, air, and module. SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support ----------------------------------------------------------------------------- | 3 |  | 服务内容：设备低分测试没过，重新水膜校正 服务结果: 设备正常使用 | [12 empty] | 118668075 | In Process | 1494923 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Others | [6 empty] | Calibration | [5 empty] | Refresh |  |  | System issue | bar: No
1566 |  |  | China | Guangxi Province Teng County People’s Hospital | 40036 | 4450400% | 44 | 02N/Ingenuity CT 728326/Replace the head mop pad | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : Replace the head mop pad Safety Questions:Case: 0118664106 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: mo xueke Last updated by: mo xueke 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | 5 |  | head mop pad | [12 empty] | 118664106 | Closed | 1491986 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Others | [6 empty] | Head pad | [5 empty] | Refresh |  |  | System issue | bar: No
1567 |  |  | China | Laiyang City People's Hospital | 345005 | 4451500% | 46 | 05/Ingenuity Core128/Elite China/Add a node | BU Q&R Feedback: No SAP Customer Symptom Code: MARF Customer Symptom Code: Monitor Alarm Failure Reason: Corrective Maintenance Engineer Failure Code: Incompatible Use Repair Activity Code: Instructed Customer Problem Reported by Customer : 05/Ingenuity Core128/Elite China/ Next Step Action: 20211114sjn15 Safety Questions:Case: 0118713549 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: liu xiaohu Last updated by: liu xiaohu 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: ask to add a dicom node Troubleshooting Action: instruct to customer to add the node under advanced account Repair Action: instruct the customer to add the node by themselves Test/Inspection: Functional Check PASSED. The device returned to full functionality. Test Equipment: NA ----------------------------------------------------------------------------- | 3 | [14 empty] | 118713549 | Fixed | 1523602 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [19 empty] | Refresh |  |  | System issue | bar: No
1568 |  |  | China | Shandong Province Chest Hospital | 40025 | 4450200% | 44 | Login failed | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Calibration/Adjustment Safety Questions:Case: 0118651053 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Gaoge Last updated by: Wang Gaoge 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Dispatched for On site Support Problem Reported by Customer: 05L/Ingenuity CT 728326 T2 Activities:Problem Description: Summary of customer's problems: failure related to computer software. Detailed description of Customer's problems or malfunctions: System shows blue screen It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: The Windows blue screen requires IST privileges before it can be used again. Remote troubleshoot and test result: Software fault PRS TR status: it's not through, ask engineer for maintenance on site. Repair Action: Suggest onsite service procedures: Need to reinstall the Windows system and software. Test/Inspection: N/A Test Equipment: N/A SAP Transfer:Case Transferred to SAP Next Step Action: huang zhaokun External Remarks: 02.11.2021 10:01:12 UTC (86038405) Service Details: Used IST privileges to log in to the software, then rebooted the Host. Service result: The device returned to normal working conditions after being repaired. Internal Remarks : 02.11.2021 10:01:15 UTC (86038405) "RDI Accuracy: YES." $Q7: ( Confirmed reported problem: Software can't login) ) $Q8: ( Login in Software use smart card and configuration. & L3 support reply this is Windows 10 Bug and need to Fixed future.) $Q9: ( Login Software use smart card and configuration.) $Q10: ( Test and inspection data refer to: PA . Device returned to fully functionality.) $Q11: Calibrated Tools name (N/A) ^Current Tube installed date: N/A ^Current Tube usage: N/A ----------------------------------------------------------------------------- | 3 |  | $Q7: ( Confirmed reported problem: Software can't login) ) $Q8: ( Login in Software use smart card and configuration. & L3 support reply this is Windows 10 Bug and need to Fixed future.) $Q9: ( Login Software use smart card and configuration.) | [12 empty] | 118651053 | Closed | 1483386 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632553 | [5 empty] | Engineering | [6 empty] | Blue screen login | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Blue-Screen | Console-SW-Blue-Screen | Console-SW-Blue-Screen | CTUDT00632553 | Blue-Screen | bar: No
1569 |  |  | China | Zhejiang Province Yueqing No. 1 People’s Hospital | 40054 | 4450200% | 44 | 02N/Ingenuity CT/ECG signal lost sometimes | BU Q&R Feedback: No Reason: Corrective Maintenance Next Step Action NTD: Customer has contacted site1, refused to transfer to RSE Problem Reported by Customer: 02N/Ingenuity CT/ SAP Transfer:Case Transferred to SAP Safety Questions:Case: 0118649526 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer's problems: Other types of error. Detailed description of Customer's problems or malfunctions: Sometimes the ECG signal is lost It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: Sometimes the ECG signal is lost Remote troubleshoot and test result: ECG signal transmission problems. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Replace ECG signal's direct connection, no adapter. Lower to the fault point 0 459801731761 USB EXTENDER CABLE FROM IBOX TO CONSOLE SAP Transfer:Dispatched for On site Support Next Step Action:huang yi Next Step Action: PPD part USB EXTENDER CABLE FROM IBOX TO CONSOLE 459801731761 1, has arranged 12H to the site, FSE: Huang Yi -86-15958725282 Internal Remarks : 03.11.2021 05:12:04 UTC (86036249) RDI Accuracy:Yes $Q7: (Confirmed reported problem: ECG signal intermittent lost) $Q8: (check found usb often disconnect from ibox to host) $Q9: (replace the USB cable from ibox to host) $Q10: (Test and inspection data refer to: PA, Device returned to full functionality.) $Q11: Calibrated tools name(N/A) S/N (N/A) Next calibration date (N/A) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date:N/A ^Current Tube usage: second:N/A External Remarks: 03.11.2021 05:12:02 UTC (86036249) service details: ECG signal is disconnected intermittently. Suspect the cable from IBOX to HOST is unstable. Replace the cable. service result: device returns to normal working condition after service Other significant matters: none ----------------------------------------------------------------------------- | 3 |  | $Q7: (Confirmed reported problem: ECG signal intermittent lost) $Q8: (check found usb often disconnect from ibox to host) $Q9: (replace the USB cable from ibox to host) | [12 empty] | 118649526 | Closed | 1482588 |  |  | Investigation | Investigation | IBOX | [5 empty] | Engineering | [6 empty] | IBOX USB | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | IBOX | Gantry-HW-IBOX | Gantry-HW-IBOX | IBOX | IBOX | Gantry-HW | Gantry-HW-IBOX | bar: No
1570 |  |  | China | Guangxi Province Teng County People’s Hospital | 40036 | 4450100% | 44 | 02N/Ingenuity CT 728326/pacs received images quality is poor | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 02N/Ingenuity CT 728326 Safety Questions:Case: 0118638063 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer's problems: Other types of error. Detailed description of Customer's problems or malfunctions: The image quality received by pacs is reduced It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: The image quality received by pacs is reduced Remote troubleshoot and test result: Pacs compressed the image after receiving it PRS TR status: Everything passed through. Repair Action: Suggest onsite service procedures: Check PACS node setup. Work with PACS to find cause SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support ----------------------------------------------------------------------------- | 5 |  | $Q6: (Image problem) $Q7: (Confirm reported problem & Found image problem.) $Q8: (Found image problem . & Need to do image calibration.) $Q9: (Do QA calibration.) | [12 empty] | 118638063 | Closed | 1479208 |  |  | Investigation | Investigation | [6 empty] | Others | [6 empty] | Calibration | [5 empty] | Refresh |  |  | System issue | bar: No
1571 |  |  | China | Shaanxi Province Ankang T.C.M. Hospital | 345041 | 4450100% | 44 | WeChat：02N/Ingenuity Core128/Elite China/couch fault | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Clean / Lubricate Problem Reported by Customer: couch fault SAP Transfer:Case Transferred to SAP Safety Questions:Case: 0118637504 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Gaoge Last updated by: Wang Gaoge 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer's problems: There is a fault with the horizontal movement of the scan bed. Detailed description of Customer's problems or malfunctions: There is an abnormal sound when the couch is moving horizontally. It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: There is an abnormal sound when the couch is moving horizontally. Remote troubleshoot and test result: Consider drive shaft friction or a different shaft causing the abnormal sound. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: 1. Check the source of the abnormal sound in the horizontal drive. Lubricate or adjust to eliminate the abnormal sound. Test/Inspection: N/A Test Equipment: N/A SAP Transfer:Dispatched for On site Support T2 Activities : Repair Action: replace Motor Coupling 453567099401 External Remarks: 08.11.2021 02:48:32 UTC (86039086) service reason: patient bed is noisy. Service Contents: replace coupling Result of service: device operating normally Internal Remarks : 08.11.2021 02:48:35 UTC (86039086) Q7: (Confirmed reported problem. & The Gantry couple have problem. $Q8: (Check the error log was not found. Check the couple & We need to replace the couple) $Q9: ( 1. Replace the couple2. Test use for CT was OK. $Q10: (Test and inspection data refer to: PA. The device returned to full functionality. $Q11: NA ----------------------------------------------------------------------------- | 3 |  | replace Motor Coupling 453567099401 | [12 empty] | 118637504 | Closed | 1478839 |  |  | Investigation | Investigation | CTUDT00632179 | [5 empty] | Engineering | [6 empty] | Coupling | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | Motor Coupling | Couch-HW-Motor Coupling | Couch-HW-Motor Coupling | CTUDT00632179 | Motor Coupling | Couch-HW | Couch-HW-Motor Coupling | bar: No
1572 |  |  | Germany | Gesundheit Nord gGmbH Klinikverbund | 40002 | 4452900% | 48 | Kunde bemängelt die Bildqualität des System bei Axialen Scans, Biopsien und Kopfaufnahmen | Internal Remarks: See email attachments => Ms. Bösemann /Application will still clarify ----------------------------------------------------------------------------- | 3 |  | Customer complains about the image quality of the system for axial scans, biopsies and head images | [12 empty] | 118803627 | New | 1569285 |  |  | Investigation | Investigation | [6 empty] | Application | [6 empty] | Application | [5 empty] | Refresh |  |  | System issue | bar: No
1573 |  |  | Japan | Independent administrative institution National Hospital Organization | 341015 | 4451900% | 46 | CT:APP:FN | ①冠動脈の後にAoを撮影したい。冠動脈はBTでAoはTimedで組みsaveできますが、大丈夫でしょうか？②胸服部単純+CS+BT+冠動脈+胸腹部2相を撮影するとTube_hotが出ます。#_of_cycleを少なくしたり、Cycle_Timeを長くすることで効果ありますか？③最近2件ほど起きた現象ですが、CTViewerで作成したMPRが真っ黒になりました。他のサイトでもありますか？ |  |  | Split CFM for MPR and Bolus Track issue Case Number# 0118748208 Description ① I want to take a picture of Ao after the coronary artery. The coronary arteries can be assembled with BT and Ao with Timed, but is that okay? ② (2) Tube_hot appears when you take a picture of Chest Hattori Simple + CS + BT + Coronary Arteries + Chest and Abdomen Phase 2. Is it effective to reduce #_of_cycle or increase Cycle_Time? ③ (3) This is a phenomenon that has occurred about two times recently, but the MPR created by CTViewer has turned black. Is it another site? Response from Remote Clinical Support Nakagawa, Futoshi: ① Basically, BT and Timed cannot be used together. Even if you can save, some settings may have been changed automatically. Please do a test scan once. ② I think that there will be some effect by reviewing the BT conditions. The number of times is 30 → 20, the cycle time is 1.3sec → 1.7sec, etc. ③ I have hardly heard of it. If the phenomenon occurs repeatedly, please leave a history such as taking a picture. | [14 empty] | 1567977 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632116 | [5 empty] | Engineering | Application | [5 empty] | MPR | [5 empty] | Refresh |  |  | System issue | System | CIRS-SW | MPR Recon | CIRS-SW-MPR Recon | CIRS-SW-MPR Recon | CTUDT00632116 | MPR Recon | CIRS-SW | CIRS-SW-MPR Recon | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1574 |  |  | Japan | Independent administrative institution National Hospital Organization | 341015 | 4451900% | 46 | CT:APP:FN | ①冠動脈の後にAoを撮影したい。冠動脈はBTでAoはTimedで組みsaveできますが、大丈夫でしょうか？②胸服部単純+CS+BT+冠動脈+胸腹部2相を撮影するとTube_hotが出ます。#_of_cycleを少なくしたり、Cycle_Timeを長くすることで効果ありますか？③最近2件ほど起きた現象ですが、CTViewerで作成したMPRが真っ黒になりました。他のサイトでもありますか？ |  |  | Split CFM for MPR and Bolus Track issue | [14 empty] | 1567972 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | Application | [5 empty] | Bolus Track | [5 empty] | IB Exist |  |  | System issue | system | cosnole-sw | bolustracking | cosnole-sw-bolustracking | cosnole-sw-bolustracking | 0 | bolustracking | bar: No
1575 |  |  | Japan | 医療法人清翠会 牧病院 | 341017 | 4452800% | 47 | 03CS:KKI:ShortTC時に、アーク発生＞＞継続してShortTC実施＞szInfo: SoftArcReg 0x211*/S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT | Work Order #: WO-04862714 Kimura, Kazumasa Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 4 |  | Continue to implement Short TC | [12 empty] | 118798971 | In Progress | 1564481 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00630085 | [5 empty] | Engineering | [6 empty] | Generator Arc | [5 empty] | Refresh |  |  | System issue | System | IC-HW | Tub Arc | IC-HW-Tub Arc | IC-HW-Tub Arc | CTUDT00630085 | Tub Arc | IC-HW | IC-HW-Tub Arc | bar: No
1576 |  |  | Japan | Tomakomai City Hospital | 341014 | 4452500% | 47 | I will cancel because I made a mistake | Work Order #: WO-04843034 Miyakawa, Takeyuki Problem Description by Engineer: Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer: N/A ----------------------------------------------------------------------------- | 4 |  | I will cancel because I made a mistake Therefore, it cannot be canceled. Thank you for the closing process. | [12 empty] | 118782466 | In Progress | 1559261 |  |  | Evaluation | Evaluation | [6 empty] | Others | [6 empty] | Mistake | [5 empty] | Refresh |  |  | System issue | bar: No
1577 |  |  | Japan | Tomakomai City Hospital | 341014 | 4452500% | 47 | 01DS: TM: When burning a CD, the data can be written, but failed is displayed. | Work Order #: WO-04843151 Miyakawa, Takeyuki The system meets the specification for the performed service and is returned to use. Diagnostic performed by Engineer: When he burns a CD, data can be written, but a "Failed" message was displayed Problem Description by Engineer: When he burns a CD, data can be written, but a "Failed" message was displayed Internal Comments: [Subject] 01DS:TM:Mr. Hanada: When he burns a CD, data can be written, but a "Failed" message was displayed [Internal Comments] FSE will respond. Shipped part: 459801631381:HPE HP 9.5mm Slim SuperMulti DVD Writer QTY:1 ZSID BackOrder Dispatched: 1 RC FSE Registration Information: Name of FSE 1 who visited the site: Mototsugu Sugahara FSE 1 visit date and time: 2021/11/3018:00 FSE 1 hours worked: 2 Dispatch instructions: Dispatch instruction #1: Shipping conditions: NBD (mixed loading, next day or thereafter) Delivery date: November 26, 2021, no time specified Shipping destination: Sapporo Office To be collected by: Mototsugu Sugahara Phone number: Other: If there is no stock in Japan, please order it from overseas. Person making request: Follow Up Required : * ----------------------------------------------------------------------------- Work Order #: WO-04843316 Sugawara, Mototsugu Problem Description by Engineer: When he burns a CD, data can be written, but a "Failed" message was displayed Regulatory Questions : N/A Diagnostic performed by Engineer: N/A ----------------------------------------------------------------------------- | 4 |  | 1.00000,459801631381,HPE HP 9.5mm Slim SuperMulti DVD Writer | [12 empty] | 118782728 | In Progress | 1556046 |  |  | Evaluation | Evaluation | TBD-DVDDrive | [5 empty] | Engineering | [6 empty] | ShouldnotEngineering | [5 empty] | Refresh |  |  | System issue | System | Console-HW | ShouldnotEngineering | Console-HW-DVDDrive | Console-HW-DVDDrive | TBD-DVDDrive | DVDDrive | Console-HW | Console-HW-DVDDrive | bar: No
1578 |  |  | Japan | Social Welfare Organization Saiseikai Imperial Gift Foundation, Inc. | 343000 | 4452400% | 47 | 01CS: TM: Message "Missing ..." was displayed after shooting >> All images were created by closing the window >> No hardware error | Work Order #: WO-04836280 Miyakawa, Takeyuki Problem Description by Engineer: Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer: N/A ----------------------------------------------------------------------------- | 4 |  | Description: $ version = "v4.1.10.25210" 9:49 Call content The message "Missing ..." was displayed twice after the Axial shooting of the head. Click OK to complete all the images. I am creating a Bug Report. correspondence Confirm that there are no errors in the Event log for both CIRS and CT-Host. It was judged that there was an error on the CT-Host side, and we asked you to restart the CT-Host when the time was up, and we decided to continue using it. https://philipsb2bsc.my.salesf | [12 empty] | 118775297 | In Progress | 1555574 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | SW Error | [5 empty] | Refresh |  |  | System issue | System | Console-SW | NoEnoughInfo | Console-SW-NoEnoughInfo | Console-SW-NoEnoughInfo | NoEnoughInfo | NoEnoughInfo | bar: No
1579 |  |  | Japan | Social Medical Corporation BOKOI, Nikko Memorial Hospital | 341011 | 4452300% | 47 | 04CS: KKI: 11/23 An error occurred when the key switch was turned on >> Continued use | Work Order #: WO-04833729 Kimura, Kazumasa Problem Description by Engineer: Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer: N/A ----------------------------------------------------------------------------- | 4 |  | 04CS: KKI: Tsukakoshi: 11/23 An error occurred when the key switch was turned on >> Continued use [Internal Comments] [Log confirmation result] It is probably caused by the following error. S_TILT_C_MANUAL_TILT_OFF It was said that a tilt error occurred. It seems that it was used continuously because the tilt operation was confirmed and there was no problem. After that, when the key switch was turned off / on, the same error did not occur. We have decided that this is a temporary phenomenon and will wait and see. | [12 empty] | 118774968 | In Progress | 1555475 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Tilt | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | NoEnoughInfo | Gantry-HW-NoEnoughInfo | Gantry-HW-NoEnoughInfo | NoEnoughInfo | NoEnoughInfo | bar: No
1580 |  |  | Japan | Independent administrative institution National Hospital Organization | 341015 | 4451900% | 46 | CT:APP:FN | Work Order #: WO-04802999 Nakagawa, Futoshi The system meets the specification for the performed service and is returned to use. Resolution: (1) Basically speaking, BT and Timed cannot be used together.Even if saving can be done, it is possible that some settings have been automatically changed.I asked him to try a test scan.(2) I think it will be somewhat effective if the BT conditions are reviewed.Reduce the number of cycles from 30 to 20, and then change the cycle time from 1.3sec to 1.7sec, etc.(3) I've hardly ever heard of it.If the problem occurs repeatedly, please take photos and leave a record. Problem description by Clinical Support Specialist : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer: NA ----------------------------------------------------------------------------- | 5 | [14 empty] | 118748208 | Closed | 1538262 |  |  | Investigation | Investigation | [6 empty] | Application | [6 empty] | Application | [5 empty] | Refresh |  |  | System issue | bar: No
1581 |  |  | Japan | Maki Hospital | 341017 | 4451000% | 45 | JS:warranty期間中のフォローアップ | Work Order #: WO-04732805 Ootani, Natsuki Regulatory Questions : N/A Problem Description by Engineer: Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer: N/A ----------------------------------------------------------------------------- | 5 | [14 empty] | 118690608 | In Progress | 1509798 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [19 empty] | Refresh |  |  | System issue | bar: No
1582 |  |  | Japan | Maki Hospital | 341017 | 4451800% | 46 | An error occurred during head shooting, performed long tube calibration | Work Order #: WO-04788786 Kimura, Kazumasa The system meets the specification for the performed service and is returned to use. Diagnostic performed by Engineer: Examinations can continue Problem Description by Engineer: An error occurred during a head scan. Internal Comments: [Subject] 03DM:KKI:Mr. Noda: An error occurred during a head scan. >> SoftArcReg 0x211*/S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT >> He was able to continue the examination. [Internal Comments] [Results of checking the log] An arc occurred. SoftArcReg 0x211* S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT I think it was probably the initial arcing after X-ray tube replacement. I think it was a temporary problem, so I asked him to monitor the situation. Dispatched: 1 RC FSE Registration Information: Name of FSE 1 who visited the site: Oyama Tetsuo FSE 1 visit date and time: 2021/11/1822:00 FSE 1 hours worked: 2 Follow Up Required : * ----------------------------------------------------------------------------- Work Order #: WO-04789955 Ohyama, Tetsuo Regulatory Questions : N/A Problem Description by Engineer: An error occurred during a head scan. Diagnostic performed by Engineer: N/A ----------------------------------------------------------------------------- | 4 |  | I think it was probably the initial arcing after X-ray tube replacement. I think it was a temporary problem, so I asked him to monitor the situation. | [12 empty] | 118735894 | In Progress | 1533185 |  |  | Closed - Done | Closed - Done | CTUDT00630085 | [5 empty] | Engineering | [6 empty] | Generator Arc | [5 empty] | Refresh |  |  | System issue | System | IC-HW | Tub Arc | IC-HW-Tub Arc | IC-HW-Tub Arc | CTUDT00630085 | Tub Arc | IC-HW | IC-HW-Tub Arc | bar: No
1583 |  |  | Japan | 医療法人清翠会 牧病院 | 341017 | 4451600% | 46 | 01CS: RS: CT indicator at the bottom left of the screen is blinking >> Remove and eliminate the SR file that was transferred by mistake | Work Order #: WO-04768640 Satake, Rui The system meets the specification for the performed service and is returned to use Problem Description by Engineer : 画面左下のCTインジケータが点滅している Internal Comments : [Subject] 01CS:RS:ﾄｶﾞﾜ様:画面左下のCTインジケータが点滅している>>誤って転送したSRファイルをRemoveし解消 [Internal Comments] なし [4/8 Elements] Q5 その検査はどの様に終了しましたか?:無しQ6 問題により間違った所見（誤診、誤治療）の有無:無し Q7 意図しない不要な曝射の有無:無し Q8 アーチファクトの有無:無し Diagnostic performed by Engineer : 誤って転送したSRファイルをRemoveし解消 ----------------------------------------------------------------------------- | 4 |  | Diagnostic performed by Engineer: Removes and resolves erroneously transferred SR files | [12 empty] | 118719656 | In Progress | 1522964 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | TBD-DoseReport | [5 empty] | Engineering | [6 empty] | DSR | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Dose SR | Console-SW-Dose SR | Console-SW-Dose SR | TBD-DoseReport | Dose SR | Console-SW | Console-SW-Dose SR | bar: No
1584 |  |  | Japan | Tomakomai City Hospital | 341014 | 4451100% | 45 | 02DS: KKI: Abnormal noise is generated when the bed is horizontally operated. | Work Order #: WO-04743178 Sugawara, Mototsugu Regulatory Questions : N/A Problem Description by Engineer : The couch makes a rattling noise when it is moved horizontally. Diagnostic performed by Engineer: N/A ----------------------------------------------------------------------------- Work Order #: WO-04742802 Kimura, Kazumasa The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Internal Comments: [Subject] 02DS:KKI:Mr. Hisano: The couch makes a rattling noise when it is moved horizontally. [Internal Comments] None Shipped part: 453566489521:MCG /PITTMAN MOTOR ASSEMBLY QTY:1 ZSID Sapporo (U612) 453566494823:CLUTCH - HORIZONTAL DRIVE QTY:1 ZSID/Traceable Sapporo (U612) Dispatched: 1 RC FSE Registration Information: Name of FSE 1 who visited the site: Mototsugu Sugahara FSE 1 visit date and time: 2021/11/1220:00 FSE 1 hours worked: 2 Dispatch instructions: Dispatch instruction #1: Shipping conditions: ASAP (Arrange shipment) Delivery date: ASAP Shipping destination: The FSE will pick it up at Shin-Chitose Airport. To be collected by: Mototsugu Sugahara Phone number: 080-1053-9342 Other: Person making request: Kazumasa Kimura Follow Up Required : * Problem Description by Engineer : The couch makes a rattling noise when it is moved horizontally. Diagnostic performed by Engineer: Examinations can be performed without any problems.No error. ----------------------------------------------------------------------------- | 3 |  | Description: [Details of work] There was a misalignment of the coupling (elastic metal cylinder) that connects the horizontal drive shaft of the sleeper and the encoder for position detection, causing noise. The alignment has been readjusted to eliminate the abnormal noise. November 11, 2021 Approved by Mr. Fujiki. | [12 empty] | 118699016 | In Progress | 1511119 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632179 | [5 empty] | Engineering | [6 empty] | Coupling | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | Motor Coupling | Couch-HW-Motor Coupling | Couch-HW-Motor Coupling | CTUDT00632179 | Motor Coupling | Couch-HW | Couch-HW-Motor Coupling | bar: No
1585 |  |  | United States | Lackey Memorial Hospital | 40015 | 4450900% | 45 | Clinical support | Internal Remarks : Interaction Type: New Case Caller: Michea McLemore IP & Issue: 85117259 & Protocol Event Type: Incident Service Request: Remote Clinical Did Customer Decline Remote Support? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Created case & xfer Clinapps Next Step Action: N/A Verbal PO Number: Verbal PO Approver: Verbal PO CAP: Integration : Confirmation of order 54417555 saved ----------------------------------------------------------------------------- Work Order #: WO-04731324 Curry, Dyan The system meets the specification for the performed service and is returned to use Problem description by Clinical Support Specialist : Customer Function/ Role: technologist Device Use: in use Expected and Unexpected Results: ped EC questions User Impact: none Patient Impact: none Current Software version: 4.1 Diagnostic performed by Engineer : answered dose tools questions for pediatric ecs. advised customer dose tools should not be used of head and spine work for infants. ----------------------------------------------------------------------------- | 5 | [14 empty] | 118689619 | Closed | 1504587 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [19 empty] | Refresh |  |  | System issue | bar: No
1586 |  |  | Japan | 医療法人清翠会 牧病院 | 341017 | 4451000% | 45 | Recon error, confirmed system could work properly | Integration : Confirmation of order 54419806 saved ----------------------------------------------------------------------------- Work Order #: WO-04734347 Kimura, Kazumasa The system meets the specification for the performed service and is returned to use Internal Comments : [Subject] 01CS:KKI:ﾄｶﾞﾜ様:”Not all images for (Patient Name) were created..”のポップアップ発生＞＞継続して検査可 [Internal Comments] 【ログ確認結果】 ハード的なエラーは確認できておりません。 撮影途中でキャンセルを行い撮影を終了しておりました。 おそらく、その影響でポップアップが発生したものと思われます。 一時的な現象だと思われますので、様子見とさせて頂いております。 [4/8 Elements] Q5 その検査はどの様に終了しましたか?:無しQ6 問題により間違った所見（誤診、誤治療）の有無:無し Q7 意図しない不要な曝射の有無:無し Q8 アーチファクトの有無:無し Resolution : 継続して検査可 Problem Description by Engineer : ”Not all images for (Patient Name) were created..”のポップアップ発生 ----------------------------------------------------------------------------- | 4 |  | Work Order #: WO-04734347 Kimura, Kazumasa The system meets the specification for the performed service and is returned to use Internal Comments: [Subject] 01CS: KKI: Togawa: "Not all images for (Patient Name) were created .." pop-up occurred >> Continued inspection possible [Internal Comments] [Log confirmation result] We have not confirmed any hardware errors. I canceled the shooting in the middle of shooting and finished shooting. Perhaps that's what caused the pop-up. It seems to be a temporary phenomenon, so I will wait and see. [4/8 Elements] Q5 How did the test end ?: None Q6 Presence or absence of incorrect findings (misdiagnosis, mistreatment) due to the problem: None Q7 Presence or absence of unintended unnecessary exposure: None Q8 Presence / absence of artifacts: None Resolution: Can be continuously inspected Problem Description by Engineer: "Not all images for (Patient Name) were created .." pop-up occurred | [12 empty] | 118692096 | Closed | 1510375 |  |  | Closed - Done | Closed - Done | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Popup message | [5 empty] | Refresh |  |  | System issue | System | CIRS-SW | NoEnoughInfo | CIRS-SW-NoEnoughInfo | CIRS-SW-NoEnoughInfo | NoEnoughInfo | NoEnoughInfo | bar: No
1587 |  |  | Peru | BANCO BBVA PERU | 40034 | 4450800% | 45 | 40034-THERE IS AN ERROR IN THE SYSTEM, IT DOES NOT ALLOW TO DO CALIBRATION OR HEATING THE TUBE | Work Order #: WO-04720573 [REDACTED] The system does not meet the specification for the performed service and is not returned to use. Problem Description by Engineer : Problem description by engineer : Gantry error Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected Results: No gantry errors. Patient Impact/Outcome: Equipment is down. User Impact: Equipment stopped. Current Software Version: 4.1.10 Diagnostic performed by Engineer: There is an error in the gantry and the device does not allow tube heating. The remote access is not working. They were able to extract the error file and the customer sent them for review. Cathode arcing error occurs in the high voltage stage. The cathode tank will be ordered. Internal Comments : Error # and/or description of error : S_GEN_FAULT_ANALOG_FAULT 2:5 Analog Fault S_GEN_FAULT_CATHODE_TUBE_ARC 5:5 CPM Arc Repair Action: - Check high voltage contacts. - Perform HV diagnostics - If necessary, change the cathode module. Follow Up Required: Error Number and/or Description: S_GEN_FAULT_ANALOG_FAULT 2:5 Analog Fault S_GEN_FAULT_CATHODE_TUBE_ARC 5:5 CPM Arc Repair Action: - Check high voltage contacts. - Perform HV diagnostics - If necessary, change the cathode module. Technician's Name: [REDACTED] ----------------------------------------------------------------------------- Work Order #: WO-04730012 [REDACTED] Problem Description by Engineer : Problem description by engineer : Gantry error Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected Results: No gantry errors. Patient Impact/Outcome: Equipment is down. User Impact: Equipment stopped. Current Software Version: 4.1.10 Diagnostic performed by Engineer: N/A ----------------------------------------------------------------------------- | 2 |  | Resolution Description: Inspection of high voltage connectors in X-ray tube and generators was carried out, finding evidence of arcs in cathode, cleaning and lubrication of connectors was carried out, long conditioning tests of X-ray tube were continued without inconveniences . Satisfactory functional tests were performed. | [12 empty] | 118680961 | Closed | 1503168 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | HV Connector | [5 empty] | Engineering | [6 empty] | HV Connector | [5 empty] | Refresh |  |  | System issue | System | IC-HW | HV Connector | IC-HW-HV Connector | IC-HW-HV Connector | HV Connector | HV Connector | bar: No
1588 |  |  | Japan | Nikko Memorial Hospital | 341011 | 4450500% | 44 | Unable to recon coronal/sagittal with tilt angle during non-helical scan, CTUDT00632116 addressed investigation |  | #N/A |  | CTUDT00632116 | [14 empty] | 1491718 |  |  | Closed - Done | Closed - Done | CTUDT00632116 | [5 empty] | Engineering | [6 empty] | coronal/sagittal | [5 empty] | Refresh |  |  | System issue | System | CIRS-SW | MPR Recon | CIRS-SW-MPR Recon | CIRS-SW-MPR Recon | CTUDT00632116 | MPR Recon | CIRS-SW | CIRS-SW-MPR Recon | bar: No
1589 |  |  | Japan | 医療法人清翠会 牧病院 | 341017 | 4450200% | 44 | 03DM: RS: X-ray system cannot comply at the time of Survuew and stop scanning >> Continue scanning after replanning >> S_GEN_FAULT_ANODE_TUBE_ARC | Work Order #: WO-04768072 16-Nov-2021 Ohyama, Tetsuo The system meets the specification for the performed service and is returned to use Internal Comments : [CX関連] 機器担当者と報告完了方法を合意できましたか？ ＜以下のいずれかを選択＞ YES 当日通常報告完了 [作業報告書（紙）提出] 作業報告書は紙で提出しましたか？ YES Q1その検査はどの様に終了しましたか?: ＜以下から選択＞ ・中断した箇所から再撮影 Q2 問題により間違った所見（誤診、誤治療）の有無: ・無し Q3 アーチファクトの有無: ・無し Q4 意図しない不要な曝射の有無: ・無し Regulatory Questions : N/A Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : WO-04705797にて対応、報告済み。翌日、立会い分WOになります。 ----------------------------------------------------------------------------- Work Order #: WO-04716475 16-Nov-2021 Kawazoe, Junichi The system meets the specification for the performed service and is returned to use Resolution : Case#0118649575報告済み Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Regulatory Questions : N/A ----------------------------------------------------------------------------- Work Order #: WO-04705797 15-Nov-2021 Ono, Junpei The system meets the specification for the performed service and is returned to use Resolution : [使用部品] 部品番号:453567560022 /部品名:Spellman 80 KW RoHS CPM Assy /個数:1 Tools & Equipment : TORQUE WRENCH,2022-05-11 00:00:00 Parts Consumed : 1.00000,453567560002,Spellman 80 KW RoHS CPM Assy 1.00000,989000086352,PHILIPS MRC 880 Follow Up Required : N/A Resolution : [作業内容詳細] アークが頻発する件につきまして、Ｘ線管球の交換および高圧発生装置の 交換作業を実施致しました。 作業後、装置の動作確認を行い正常に動作する旨確認致しました。 X線管球の交換後、焦点のキャリブレーションにて一部の項目が基準値内に 収まらないために、調査およびＸ線管球の再交換により予定時間を大幅に 超過しご迷惑をおかけし大変申し訳ありませんでした。 Problem Description by Engineer : Survuewの際にX-ray system cannot complyとでてスキャン停止 Regulatory Questions : N/A ----------------------------------------------------------------------------- Work Order #: WO-04682011 Satake, Rui The system meets the specification for the performed service and is returned to use Diagnostic performed by Engineer : 再プラン後スキャン継続. S_GEN_FAULT_ANODE_TUBE_ARCの発生を確認。 FSE対応。 Problem Description by Engineer : Survuewの際にX-ray system cannot complyとでてスキャン停止 Follow Up Required : * Internal Comments : [Subject] 03DM:RS:ﾄｶﾞﾜ様:Survuewの際にX-ray system cannot complyとでてスキャン停止>>再プラン後スキャン継続>>S_GEN_FAULT_ANODE_TUBE_ARC [Internal Comments] なし [出庫部品] 989000086352:PHILIPS MRC 880 QTY:1 ZREP/Traceable 大阪(U615) 453566495091:1XRT SERVICE TOOL ASSY EXP QTY:1 ZSID 大阪(U615) 459800089031:Tool, XRT Strap Pad Assy QTY:1 ZSID 大阪(U615) 459800911271:TDMS V2 M.Det Class A FRU QTY:1 ZREP/Traceable 大阪(U615) [ディスパッチ] 1 RC -----------------FSE登録情報----------------- -訪問FSE1氏名 小野 純平 -FSE1訪問日時 2021/11/1410:00 -FSE1作業時間 10 -訪問FSE2氏名 川添 純一 L1 Engineer -FSE2訪問日時 2021/11/1410:00 -FSE2作業時間 10 [出荷指示] -----------------出荷指示#1----------------- -配送条件 TDD(日時指定) -納期 2021/11/14 10:00 指定 -配送先名 現地直送 -受取者名 小野 純平 -電話番号 080-9880-9929 -その他 同時引き取り有り:待機5h -依頼者名 佐竹類 ----------------------------------------------------------------------------- | 4 |  | [Details of work] Regarding the frequent occurrence of arcs, we replaced the X-ray tube and the high-voltage generator. After the work, we checked the operation of the device and confirmed that it works normally. After replacing the X-ray tube, some items did not fit within the standard values ​​in the focus calibration, so the scheduled time was significantly exceeded due to the investigation and re-replacement of the X-ray tube, which caused inconvenience. I'm very sorry. 989000086352:PHILIPS MRC 880 QTY:1 ZREP/Traceable 大阪(U615) 453566495091:1XRT SERVICE TOOL ASSY EXP QTY:1 ZSID 大阪(U615) 459800089031:Tool, XRT Strap Pad Assy QTY:1 ZSID 大阪(U615) 459800911271:TDMS V2 M.Det Class A FRU QTY:1 ZREP/Traceable 大阪(U615) | [12 empty] | 118649575 | Fixed | 1482624 |  |  | Investigation | Investigation | CTUDT00630085 | [5 empty] | Engineering | [6 empty] | Tube Arc | [5 empty] | Refresh |  |  | System issue | System | IC-HW | Tub Arc | IC-HW-Tub Arc | IC-HW-Tub Arc | CTUDT00630085 | Tub Arc | IC-HW | IC-HW-Tub Arc | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y
1590 |  |  | Japan | Independent administrative institution National Hospital Organization | 341013 | 4450100% | 44 | 01DS: KKI: Freezes when shooting Prep >> Recovers after 5 seconds and can shoot | Work Order #: WO-04698358 02-Nov-2021 Yamamoto, Yuko Regulatory Questions : N/A Problem Description by Engineer: Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer: N/A ----------------------------------------------------------------------------- Work Order #: WO-04667805 02-Nov-2021 Yamamoto, Yuko The system meets the specification for the performed service and is returned to use. Regulatory Questions : N/A Problem Description by Engineer: Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution: I will retrieve and analyze the log. November 1, 2021 I reported to the person in charge Mr. Takumi Yamamoto. Internal Comments: [CX related] Did you agree on how to complete the report with the person in charge of devices? <Select one of the following> Yes. I completed normal reporting that day. [Work report (paper) submission] Did you submit a work report on paper? <YES/ Q1. How did the examination finish? Select from the following: *Because the answer is No for Q1, not applicable Q2. Were there any mistaken findings (misdiagnosis, mistaken treatment) due to the problem? None. Q3. Were there any artifacts? None. Q4. Was there any unintended unnecessary exposure? None. ----------------------------------------------------------------------------- Work Order #: WO-04667773 Kimura, Kazumasa The system meets the specification for the performed service and is returned to use. Diagnostic performed by Engineer: It recovered after 5 seconds and a scan could be performed. Problem Description by Engineer: The system froze during a Prep scan. Internal Comments: [Subject] 01DS:KKI:Mr. Yamamoto: The system froze during a Prep scan >> 5 seconds later it recovered and he was able to run a scan. [Internal Comments] The log showed no hardware errors. The CT host was restarted first thing in the morning, but the cause could not be identified. Dispatched: 1 RC FSE Registration Information: Name of FSE 1 who visited the site: Yutaka Yamamoto FSE 1 visit date and time: 2021/11/0117:00 FSE 1 hours worked: 2 ----------------------------------------------------------------------------- | 4 |  | CT application was freeze about 5sec after SAS injection. | [12 empty] | 118636379 | In Progress | 1475816 |  |  | Investigation | Investigation | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | SW Freeze | [5 empty] | Refresh |  |  | System issue | System | Console-SW | NoEnoughInfo | Console-SW-NoEnoughInfo | Console-SW-NoEnoughInfo | NoEnoughInfo | NoEnoughInfo | bar: No
1591 |  |  | Peru | BANCO BBVA PERU | 40034 | 4452300% | 47 | 40034 WHEN A SECOND IP IS CONFIGURED FOR REMOTE CONNECTION PURPOSES, YOU CANNOT ENTER THE LAN CONFIG, NOR … |  | 5 |  | Known issue, link to LAN Config can not open defect | [12 empty] | 118773692 | New | 1554584 |  |  | Opened | Opened | CTUDT00631055 | [5 empty] | Engineering | [6 empty] | LAN Config | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Lan Config | Console-SW-Lan Config | Console-SW-Lan Config | CTUDT00631055 | Lan Config | Console-SW | Console-SW-Lan Config | bar: No
1592 |  |  | Chile | Scotiabank Chile | 40008 | 4451900% | 46 | CL1675-ERROR DE TUBO | Work Order #: WO-04815858 23-Nov-2021 [REDACTED] The system meets the specification for the performed service and is returned to use. Resolution: 30A/600V fuse replacement since it is open in cathode. Functional tests are conducted with no issues. Diagnostic performed by Engineer: Remote diagnosis is performed and a blown cathode fuse is found. Problem Description by Engineer : Problem description by engineer : Gantry error. Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected Results: No gantry errors. Patient Impact/Outcome: Equipment is down. User Impact: Equipment stopped. Current Software Version: 4.1.10 Parts Consumed : 2.00000,459800066152,FUSE, 30A/600V AJT30 ----------------------------------------------------------------------------- Work Order #: WO-04806757 [REDACTED] The system does not meet the specification for the performed service and is not returned to use. Problem Description by Engineer : Problem description by engineer : Gantry error. Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected Results: No gantry errors. Patient Impact/Outcome: Equipment is down. User Impact: Equipment stopped. Current Software Version: 4.1.10 Diagnostic performed by Engineer: An error appears in phase voltage in the gantry. With the customer's help, a blown fuse is identified in the gantry. Fuse must be replaced. Internal Comments : Error # and/or description of error : S_GEN_FAULT_KPM_480_LINE_PHASES_MISSING Repair Action: - Replace fuse. - Set input voltage for the gantry. Follow Up Required: Error Number and/or Description: S_GEN_FAULT_KPM_480_LINE_PHASES_MISSING Repair Action: - Replace fuse. - Set input voltage for the gantry. Technician's Name: [REDACTED] ----------------------------------------------------------------------------- | 2 |  | - Replace fuse. - Set input voltage for the gantry. 2.00000,459800066152,FUSE, 30A/600V AJT30 | [12 empty] | 118751009 | Closed | 1543055 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | Not a Defect | [5 empty] | Engineering | [6 empty] | FUSE | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Other | Gantry-HW-Other | Gantry-HW-Other | Not a Defect | TBD-NoEnoughInfo | Gantry-HW | Gantry-HW-Other | bar: No
1593 |  |  | China | 河南许昌市中心医院(传染病) | 40047 | 4450500% | 44 | 02N/Ingenuity CT 728326/add the printer | BU Q&R Feedback: No SAP Customer Symptom Code: CENV Customer Symptom Code: Customer Environmental Concern Reason: Corrective Maintenance Engineer Failure Code: Environmental Cause Repair Activity Code: Configuration Change Problem Reported by Customer : 02N/Ingenuity CT 728326 Safety Questions : Case: 0118670402 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Weng Jinming Last updated by: Weng Jinming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Troubleshooting Action: add the printer Repair Action: add the printer Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: N/A ----------------------------------------------------------------------------- | 5 | [14 empty] | 118670402 | Fixed | 1495869 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [19 empty] | Refresh |  |  | System issue | bar: No
1594 |  |  | China | 广西藤县人民医院 | 40036 | 4450500% | 44 | Application/Ingenuity CT 728326/ | BU Q&R Feedback: No Reason: Application Support Problem Reported by Customer : Application/Ingenuity CT 728326/ Next Step Action : 邓老师015177135531 来电 Next Step Action : 已联系应用丰川13004507022 跟进 Safety Questions : Case: 0118669635 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Gao Luqi Last updated by: Gao Luqi 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 3 | [14 empty] | 118669635 | New | 1495565 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [19 empty] | Refresh |  |  | System issue | bar: No
1595 |  |  | Poland | University Clinical Hospital in | 341001 | 4452300% | 47 | problem with the appearance of research on the worklist | Work Order #: WO-04827835 Gerard Dabrowka The system meets the specification for the performed service and is returned to use Diagnostic performed by Engineer: In consultation with an IT specialist and an IT specialist from the hospital, we found that despite the computer was turned off, the host's IP address responded to Ping. It means that there is another computer with the same IP in the hospital network that has been assigned to the Hospital LAN card. Problem Description by Engineer : Problem description by engineer : The worklist does not appear, patient data was entered manually. After performing the exam, it is not sent to PACS or is sent partially. When was the first occurrence? Within... : 22-Nov-2021 18:00:00 How often is issue occurring? : Intermittent Information to support the complaint handling process Customer Function/Role: Not provided Device Use: Not provided Expected and Unexpected results: Not provided Patient Impact/Outcome: Not provided User Impact: Not provided Current Software Version: Not provided Internal Comments : Error # and/or description of error : There is no evident error, the system is not connected to PRS. Repair Action : During the next PM visit planned for 30 Nov 2021, the IT specialist will assign a new IP for the CT scanner from the restricted list of addresses. Extra Internal comment field : Continuation during the PM visit. Resolution: IT technician from the hospital will locate and turn the computer off with duplicate IP address. ----------------------------------------------------------------------------- | 4 |  | IT technician from the hospital will locate and turn the computer off with duplicate IP address. | [12 empty] | 118769322 | Closed | 1548856 |  |  | Evaluation | Evaluation | [6 empty] | Others | [6 empty] | Worklist | [5 empty] | Refresh |  |  | System issue | bar: No
1596 |  |  | Poland | Government Agency. | 341010 | 4450300% | 44 | Recognizable artifact on images due to faulty detector module;Replaced detector module and calibrated | Work Order #: WO-04719373 09-Nov-2021 Mr. Robert Kamola The system meets the specification for the performed service and is returned to use Problem Description by Engineer: Replacement of the damaged module in the detector. Resolution: Replacement of the damaged module in detector no. 34. Calibrations and tests after the repair. The fault was eliminated, the device works correctly. ----------------------------------------------------------------------------- Work Order #: WO-04700546 09-Nov-2021 Mr. Robert Kamola The system does not meet the specification for the performed service and is not returned to use Problem Description by Engineer: The report concerns image artifacts. Resolution: The report concerns artifacts on the image. Test scans and detector quality tests were performed - incorrect signal was detected on module no. 34. The detector module needs to be replaced. The device is out of order. ----------------------------------------------------------------------------- | 4 |  | Resolution: The report concerns artifacts on the image. Test scans and detector quality tests were performed - incorrect signal was detected on module no. 34. The detector module needs to be replaced. The device is out of order. 1.00000,453567528582,Kit, Torque Wrench (0-100FT-LBS) 1.00000,459800911271,TDMS V2 M.Det Class A FRU 1.00000,452250066171,TOOL KIT, COMPACT | [12 empty] | 118658881 | In Progress | 1488513 |  |  | Closed - Done | Closed - Done | CTUDT00632110 | [5 empty] | Engineering | [6 empty] | dms module | [5 empty] | Refresh |  |  | System issue | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | DMS Module | DMS-HW | DMS-HW-DMS Module | bar: No
1597 |  |  | Puerto Rico | HOSPITAL DR SUSONI INC | 40013 | 4453000% | 48 | Patient support lubrication | Work Order #: WO-04881573 Ramos Rivera, Pablo The system meets the specification for the performed service and is returned to use Follow Up Required : lubricate patient support Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: patient support lubrication needed User Impact: none Patient Impact: none Current Software Version:4.1.10 Diagnostic performed by Engineer : Order lubrication kit, lubricate patient support ----------------------------------------------------------------------------- | 5 | [14 empty] | 118815263 | In Progress | 1571447 |  |  | Evaluation | Evaluation | E-125059 | [5 empty] | Engineering | [6 empty] | Lubrication | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | Lubricant | Couch-HW-Lubricant | Couch-HW-Lubricant | E-125059 | Lubricant | Couch-HW | Couch-HW-Lubricant | bar: No
1598 |  |  | Spain | Comercial Centro Medico | 40042 | 4451100% | 45 | Rebuilder error appears again | Work Order #: WO-04743628 12-Nov-2021 [REDACTED] The system meets the specification for the performed service and is returned to use. Diagnostic performed by Engineer: Intermittent connection error with reconstructors. Problem Description by Engineer: Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Regulatory Questions : N/A Resolution: The connection switch is changed between the 3 computers of the CT console. I perform an operation test and the system is working. Parts Consumed: 1.00000,459800457501,SMC SMCGS803 GB Network Switch 8 port ----------------------------------------------------------------------------- | 3 |  | Resolution: The connection switch is changed between the 3 computers of the CT console. I perform an operation test and the system is working. Parts Consumed: 1.00000,459800457501,SMC SMCGS803 GB Network Switch 8 port | [12 empty] | 118699798 | Fixed | 1515241 |  |  | Closed - Done | Closed - Done | part replace | [5 empty] | Others | [6 empty] | Net Switch | [5 empty] | Refresh |  |  | System issue |  |  |  | - | - | part replace | 0 | bar: No
1599 |  |  | Spain | S.C.S. HOSP. GRAL. DE FUERTEVENTURA | 40020 | 4450300% | 44 | Streaks appear in images when making skulls | Work Order #: WO-04691426 [REDACTED] The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled. Regulatory Questions : N/A Problem Description by Engineer: In some cranial protocol, a horizontal line shaped artifact is showing up across the image. Follow Up Required: We would need to go by and check with the phantom and by using the same cranial protocol if the artifacts are repeating. Check the Raw Data for these images to find out if it the issue is coming from the gantry. Diagnostic performed by Engineer: The images of the relevant patient are checked, and the artifact is observed in one of the series; this artifact is moving vertically along the series images. It seems to be caused by something external to the reconstruction circle. ----------------------------------------------------------------------------- | 3 |  | Check the Raw Data for these images to find out if it the issue is coming from the gantry. Diagnostic performed by Engineer: The images of the relevant patient are checked, and the artifact is observed in one of the series; this artifact is moving vertically along the series images. It seems to be caused by something external to the reconstruction circle. | [12 empty] | 118657714 | Closed | 1484487 |  |  | Investigation | Investigation | TBD-NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Image Artifact | [5 empty] | Refresh |  |  | System issue | System | TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo | bar: No
1600 |  |  | China | Fujian Provincial Hospital | 345011 | 4451800% | 46 | 02N/Ingenuity Core128/stop 0x00000ed | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 02N/Ingenuity Core128/Elite China SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Safety Questions:Case: 0118734289 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Weng Jinming Last updated by: Weng Jinming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer's problems: Host computer blue screen Detailed description of Customer's problems or malfunctions: Service requester is not on site. It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: Stop 0x00000ed Remote troubleshoot and test result: Contact engineer Ding Shi duo, suspect software, memory.。。problem, hard disc has been replaced before PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Suggest engineer to check host computer memory, graphic card, reinstall software on site, teach customer to use Philips WeChat to report for repair Next Step Action: Ding shiduo ----------------------------------------------------------------------------- | 2 |  | $Q7:Confirmed reported problem. the host blue screen. $Q8: on site Reviewed the error log , Performed visual inspection and During the BIOS hardware test, found the memory is abnormal. $Q9: Clean Computer host, unplug Computer host memory. | [12 empty] | 118734289 | In Process | 1537109 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Others | [6 empty] | Memory | [5 empty] | Refresh |  |  | System issue |  |  |  | - | - | 0 | 0 | bar: No
1601 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 337000 | 4451000% | 45 | Hava kalibrasyonunu geçmiyor | Work Order #: WO-04734527 16-Nov-2021 Osman Alpay The system meets the specification for the performed service and is returned to use Diagnostic performed by Engineer : Air Calibration fails with error. Problem Description by Engineer : Problem description by engineer : Air cal could be completed. Information to support the complaint handling process Customer Function/Role: Not provided Device Use: Not provided Expected and Unexpected results: Not provided Patient Impact/Outcome: Not provided User Impact: Not provided Current Software Version: Not provided. Resolution : We opened the gantry cover to check DMS. We found contrasting agent deposits on the UHR Grid and detectors. System was returned to use in functional condition after cleaning the contrasting agent. Also, gantry room ar conditioners were faulty and the ambient temperature was high. We advised the user that they should have the a/c unit repaired. Tools & Equipment : MULTIMETER,2022-02-05 00:00:00 ----------------------------------------------------------------------------- Work Order #: WO-04734382 Salih Mendi The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Internal Comments : Error # and/or description of error : Air cal could be completed. Repair Action : Air cal could be completed. site visit needed for further inspection. There is no artefact on images also. Diagnostic performed by Engineer : Air cal could be completed. Problem Description by Engineer : Problem description by engineer : Air cal could be completed. Information to support the complaint handling process Customer Function/Role: Not provided Device Use: Not provided Expected and Unexpected results: Not provided Patient Impact/Outcome: Not provided User Impact: Not provided Current Software Version: Not provided. Resolution : We remotely supported the user for the air calibration problem. Mylar was cleaned. When starting the air calibration, object warning inside the imaging area error occurred. Site visit is needed for a detailed inspection. The user reported that there were no artefacts and the system was up and running. Follow Up Required : Error # and/or description of error : Air cal could be completed. Repair Action : Air cal could be completed. site visit needed for further inspection. There is no artefact on images also. Technician Name : Salih Mendi ----------------------------------------------------------------------------- | 3 |  | Resolution : We remotely supported the user for the air calibration problem. Mylar was cleaned. When starting the air calibration, object warning inside the imaging area error occurred. Site visit is needed for a detailed inspection. The user reported that there were no artefacts and the system was up and running. | [12 empty] | 118692132 | Closed | 1506220 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Others | [6 empty] | Contrast Intrusion | [5 empty] | Refresh |  |  | System issue | bar: No
1602 |  |  | United Kingdom | STOCKPORT NHS FOUNDATION TRUST | 40016 | 4451400% | 45 | Replace harmonized phantom with the old blue style system phantom. | Internal Remarks : Email from Warren Thorpe dated the 12th Nov 2021 Hospital Name: Stepping Hill Hospital, Stockport Customer Contact Name (not FSE or Philips Employee) : Karen Andrew Contact Telephone number:01614194782 Serial/Pin/Tech ID: 40016 System out of action? No System Priority: 4 - Intermittent Problem Detail of work/Problem Description: Replace harmonized phantom with the old blue style system phantom after the customer complained about alignment issues with the new ones. Replace phantom and run calibrations. Customer Temperature: Green No. of FSE/s to fit: 1 Hours to fit (excluding travel time): 3 Required FSE skill level: 3 - Adequate/Experienced Q1 Was the equipment in ‘Clinical use’ at the time the issue was discovered?: No Q2 Was any patient or user harmed?: No Parts TPR Review needed: Choose an item. TPR Reviewer: Click here to enter text. Part number: 453566502022 Description: Phantom Holder Assy, Carbon Top Quantity: 1 FCO part: Choose an item. PUDO: Click here to enter text. TPR Review needed: Choose an item. TPR Reviewer: Click here to enter text. Part number: 453567135961 Description: EXTENDED PHANTOM PERFORMAN KIT Quantity: 1 FCO part: Choose an item. PUDO:Click here to enter text. Delivery detail if to site: Click here to enter text. Any other relevant info: Customer has complained, so knows that we are ordering a new system phantom. Date and time needs to be arranged. Best Regards, Warren Thorpe Customer comment : Warren to attend on 7/12 - Confirmed with Karen Andrew ----------------------------------------------------------------------------- Work Order #: WO-04758834 Warren Thorpe Regulatory Questions : N/A Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 4 | [14 empty] | 118711553 | In Progress | 1521512 |  |  | Evaluation | Evaluation | H-Phantom | [5 empty] | Engineering | [6 empty] | H-Phantom | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | H-Phantom | Couch-HW-H-Phantom | Couch-HW-H-Phantom | H-Phantom | H-Phantom | bar: No
1603 |  |  | United States | Central Vermont Medical Center | 40024 | 4452400% | 47 | Images not sending to PACs | Work Order #: WO-04882078 02-Dec-2021 Sangster, Robert The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Diagnostic performed by Engineer : Made adjustments to the network settings, set two of the queue nodes (AMBRA and iSite) to series level instead of image. Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : Working with site and national support to resolve network issues. ----------------------------------------------------------------------------- Work Order #: WO-04841061 29-Nov-2021 Sangster, Robert The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : Issue with pacs reviewed. Bug report gathered and sent for review. Follow Up Required : Adjust dicom settings to send series to pacs ----------------------------------------------------------------------------- | 4 |  | Diagnostic performed by Engineer : Issue with pacs reviewed. Bug report gathered and sent for review. Follow Up Required : Adjust dicom settings to send series to pacs under engineering investigation | [12 empty] | 118781092 | In Progress | 1558724 |  |  | Evaluation | Evaluation | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS Transfer | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y
1604 |  |  | China | Dongguan City Integrated Traditional Chinese and Western medicine Hospital | 345013 | 4450200% | 44 | 02N/Ingenuity Core128/Elite China/Check equipment | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ Safety Questions:Case: 0118651564 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingsong Last updated by: Lu Mingsong 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer's problems: The equipment's warranty is about to expire. The customer asked the field service engineer to go on site to check the equipment's condition. Detailed description of Customer's problems or malfunctions: The equipment's warranty is about to expire. The customer asked the field service engineer to go on site to check the equipment's condition. It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: N/A Remote troubleshoot and test result: N/A PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Go on site and check the equipment's condition. We recommend spare parts 12NC and names: N/A Memo of Customer request: None SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support External Remarks: 03.11.2021 11:53:23 UTC (86034125) Service details: inspection before machine warranty expires. Service outcome: the machine is checked and found no abnormality. The machine is working normally. Other important issues: None Internal Remarks : 05.11.2021 15:42:16 UTC (86034125) "RDI Accuracy:Yes" $Q1: () $Q2: () $Q3: () $Q4: () $Q7: (Confirmation & Description of problem) Confirmed reported problem, customer requires check of the machine situation before the warranty expires. $Q9: (Repair steps) check the machine as the customer required, no abnormality is found. $Q10: (Test and inspection data refer to PA . Device returned to full functionality ) $Q11: Calibrated Tools name ( refer to PA ) S/N ( ) Next Calibration date ( refer to PA ) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date: N/A ^Current Tube usage: N/A External Remarks: 08.11.2021 02:25:39 UTC (86001504) Service details: inspection before machine warranty expires. Service outcome: the machine is checked and found no abnormality. The machine is working normally. Other important issues: None Internal Remarks: 08.11.2021 02:25:42 UTC (86001504)RDI Accuracy:Yes $Q7: (Confirmed the reported problem, the sales asked to checked the device status) $Q8: (no fault found by other FSE, double confirmed) $Q9: (no action required) $Q10: (Test and inspection data refer to N/A, PA by other FSE. Device returned to full functionality.) $Q11: $Calibrated Tools name (N/A) S/N (N/A) Next Calibration date (N/A) ----------------------------------------------------------------------------- | 5 | [14 empty] | 118651564 | In Process | 1483697 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [19 empty] | Refresh |  |  | System issue | bar: No
1605 |  |  | United States | UF Health Leesburg Hospital | 40022 | 4450400% | 44 | Result would jump when planning a scan with multiple results |  | #N/A |  | Known issue: result jump | [14 empty] | 1524653 |  |  | Investigation | Investigation | CTUDT00632986 | [5 empty] | Engineering | [6 empty] | Result Jump | [5 empty] | Refresh |  |  | System issue | System | Console-SW | ResultSelectionJump | Console-SW-ResultSelectionJump | Console-SW-ResultSelectionJump | CTUDT00632986 | ResultSelectionJump | Console-SW | Console-SW-ResultSelectionJump | bar: No
1606 |  |  | China | Hebei Province Geriatric Hospital | 345015 | 4450200% | 44 | 05K For IB/Core128/Consult Manufacture date | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer:consult Manufacture date Safety Questions:Case: 0118649876 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Pan Chuchu Last updated by: Pan Chuchu 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action: Case# 0118649876 Hebei Province Geriatrics Hospital, Ingenuity Core128, EQ# 84722645, customer want to query the device's manufacturer date The caller is teacher Zhang at equipment dept 031189892511 Next Step Action: Notified ASP Jiunuo Business to follow up with customer ----------------------------------------------------------------------------- | 3 | [14 empty] | 118649876 | Closed | 1482740 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [19 empty] | Refresh |  |  | System issue | bar: No
1607 |  |  | United States | Philips Medical Capital | 40053 | 4450800% | 45 | PACS troubleshooting | Work Order #: WO-04725430 16-Nov-2021 Fairchild, Andrew The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Diagnostic performed by Engineer : Unable to log into Portal system with service user names/passwords. Follow-up created with Portal site ID. Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: 4.1.10 ----------------------------------------------------------------------------- Work Order #: WO-04723289 08-Nov-2021 Fairchild, Andrew The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: 4.1.10 Diagnostic performed by Engineer : Troubleshooted multiple configurations and ran tests with logging turned up to debug layer. Data and logs gathered to be forwarded to other support. Equipment downtime needs to be scheduled for additional troubleshooting when support available. Follow Up Required : Continue troubleshooting with all involved parties that work with the PACS system. ----------------------------------------------------------------------------- | 5 |  | Troubleshooted multiple configurations and ran tests with logging turned up to debug layer. Data and logs gathered to be forwarded to other support. Equipment downtime needs to be scheduled for additional troubleshooting when support available. Continue troubleshooting with all involved parties that work with the PACS system. | [12 empty] | 118682975 | Closed | 1501066 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS Transfer | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y
1608 |  |  | United States | Leesburg Regional Medical Center | 40022 | 4450600% | 44 | 0711CFXXX HIS/RIS not populating | Integration : Confirmation of order 54403986 saved ----------------------------------------------------------------------------- Work Order #: WO-04714094 Mcginnis, Mike The system meets the specification for the performed service and is returned to use Internal Comments : Error # and/or description of error : his/ris only showing yesterdays patients Repair Action : I had him check the by date filter and it was set for yesterday. He cleared that filter and it is now pulling current scheduled patients Diagnostic performed by Engineer : Angelo said the his/ris only shows yesterdays patient list. I had him check the by date filter and it was set for yesterday. He cleared that filter and it is now pulling current scheduled patients Problem Description by Engineer : Problem description by engineer : HIS/RIS only showing yesterdays patients When was the first occurrence? Within... : 06-Nov-2021 How often is issue occurring? : First occurrence Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: HIS/RIS only showing yesterdays patients Patient Impact/Outcome: none User Impact: Could not use his/ris Current Software Version: Not provided Resolution : I had him check the by date filter and it was set for yesterday. He cleared that filter and it is now pulling current scheduled patients ----------------------------------------------------------------------------- | 3 |  | Repair Action : I had him check the by date filter and it was set for yesterday. He cleared that filter and it is now pulling current scheduled patients Diagnostic performed by Engineer : Angelo said the his/ris only shows yesterdays patient list. I had him check the by date filter and it was set for yesterday. He cleared that filter and it is now pulling current scheduled patients Problem Description by Engineer : Problem description by engineer : | [12 empty] | 118675011 | Closed | 1496538 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Application | [6 empty] | HIS/RIS | [5 empty] | Refresh |  |  | System issue | bar: No
1609 |  |  | China | Panjin City DaWa District People's Hospital | 345040 | 4450200% | 44 | 02N/Ingenuity Core128/Elite China/add node | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 02N/Ingenuity Core128/Elite China Safety Questions:Case: 0118649455 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Dispatched for On site Support T2 Activities:Problem Description: Summary of customer's problems: Other types of error. Detailed description of Customer's problems or malfunctions: Add node It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: Add node Remote troubleshoot and test result: Add nodes. Need an engineer to go to the field PRS TR status: it's not through, ask engineer for maintenance on site. Repair Action: Suggest onsite service procedures: Add node SAP Transfer:Case Transferred to SAP External Remarks: 02.11.2021 11:33:43 UTC (86038273) service details: need to add network nodes. Add nodes. service result: device returns to normal working condition after service Internal Remarks : 02.11.2021 11:33:45 UTC (86038273) RDI Accuracy:Yes $Q7: (Confirmed reported problem: add dicom device) $Q8: (add dicom device) $Q9:(add dicom device) $Q10: (Test and Inspection data refer to pa & Device returned to full functionality) $Q11:Calibrated Tools name (NA) ----------------------------------------------------------------------------- | 3 | [14 empty] | 118649455 | Closed | 1482538 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [19 empty] | Refresh |  |  | System issue | bar: No
1610 |  |  | United States | Clay County Medical Center | 40003 | 4450300% | 44 | Physics phantom wires have lost tension. | Work Order #: WO-04695066 05-Nov-2021 Stauffer Jeffrey The system meets the specification for the performed service and is returned to use Resolution : Ordered new phantom and picked up at UPS. Phantom will be delivered to customer and used on next PM in December 2021. Diagnostic performed by Engineer : Looked at phantom during last service call and saw that tension on wires in physics phantom are loose. Problem Description by Engineer : Customer Function/Role: Philips FSE Device Use: General CT Exams Expected and Unexpected Results: Unexpected - tension wire in physics phantom is not tight. User Impact: None Patient Impact: None Current Software Version:4.10 Parts Consumed : 1.00000,459801415171,HARMONIZED PHANTOM PHYSICS SECTION ----------------------------------------------------------------------------- | 3 |  | Resolution : Ordered new phantom and picked up at UPS. Phantom will be delivered to customer and used on next PM in December 2021. Diagnostic performed by Engineer : Looked at phantom during last service call and saw that tension on wires in physics phantom are loose. 1.00000,459801415171,HARMONIZED PHANTOM PHYSICS SECTION | [12 empty] | 118660267 | Closed | 1489660 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | Physics Phantom | [5 empty] | Engineering | [6 empty] | Physics Phantom | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | Physics Phantom | Couch-HW-Physics Phantom | Physics Phantom-Couch-HW-Physics Phantom | Physics Phantom | Physics Phantom | bar: No
1611 |  |  | China | Beijing City Hospital of Traditional Chinese and Western Medicine | 40032 | 4450100% | 44 | 02N/Ingenuity CT 728326/stand by | BU Q&R Feedback: No Reason: Customer Relations SAP Transfer:Dispatched for On site Support Problem Reported by Customer: stand by Next Step Action: 013811972406 Yu called Safety Questions:Case: 0118640561 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Qian Wang Last updated by: Qian Wang 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action : fan guochao SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | 5 | [14 empty] | 118640561 | New | 1479550 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [19 empty] | Refresh |  |  | System issue | bar: No
1612 |  |  | China | DaKuai Town Health, Fengquan District, Xinxiang City, Henan Province | 345009 | 4452400% | 47 | 02N/Ingenuity Core128/Elite China/Gantry Unable to start | BU Q&R Feedback: No Reason: Corrective Maintenance SAP Transfer:Dispatched for On site Support Problem Reported by Customer: Gantry Unable to start Safety Questions:Case: 0118779357 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Zhang Caiqin Last updated by: Zhang Caiqin 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Case Transferred to SAP External Remarks: 25.11.2021 11:28:11 UTC (86038950)"SUAT. Ready for Clinical Use local time" : "2021-11-25 15:00" External Remarks: 26.11.2021 12:26:25 UTC (86038950) Service details: No key box when turned on, scan failed. On-site check found MKD's NC port open. Service outcome: after replacing MKD contactor, the equipment is back to normal. Other significant matters: none ----------------------------------------------------------------------------- | 2 |  | Service details: No key box when turned on, scan failed. On-site check found MKD's NC port open. Service outcome: after replacing MKD contactor, the equipment is back to normal. | [12 empty] | 118779357 | Closed | 1557847 |  |  | Investigation - Pending Translation | Investigation - Pending Translation | PartReplacement | [5 empty] | Others | [6 empty] | MKD Contactor | [5 empty] | Refresh |  |  | System issue |  |  |  | - |  | PartReplacement | 0 | bar: No
1613 |  |  | United States | Clay County Medical Center | 40003 | 4450300% | 44 | Left Control Panel on Gantry needs adjustment. | Integration : Confirmation of order 54389280 saved ----------------------------------------------------------------------------- Work Order #: WO-04693042 03-Nov-2021 Stauffer Jeffrey The system meets the specification for the performed service and is returned to use Diagnostic performed by Engineer : Pushed up button and it did not work. Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: General CT exams Expected and Unexpected Results: Left control panel does not work properly. User Impact: None Patient Impact: None Current Software Version:4 Resolution : Replaced left control panel. ----------------------------------------------------------------------------- | 3 | [14 empty] | 118658916 | Closed | 1486176 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632181 | [5 empty] | Engineering | [6 empty] | GCP | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Gantry Panel | Gantry-HW-Gantry Panel | Gantry-HW-Gantry Panel | CTUDT00632181 | Gantry Panel | Gantry-HW | Gantry-HW-Gantry Panel | No: No; y: 0; y: 0; y: 0; y: 0; y: 0; No: No
1614 |  |  | China | 河南省新乡市凤泉区大块镇卫生院 | 345009 | 4450100% | 44 | 02N/Ingenuity Core128/Elite China/Customer consultation | BU Q&R Feedback: No SAP Customer Symptom Code: HANG Customer Symptom Code: Hang/Crash Reason: Corrective Maintenance Engineer Failure Code: Wear-out/Breakage/Mechanical Failure Repair Activity Code: Instructed Customer Problem Reported by Customer : 02N/Ingenuity Core128/Elite China T2 Activities : Problem Description: 客户咨询是否能够把64排物理结构探测器暂时锁住两排。 Troubleshooting Action: 告诉客户探测器物理结构无法锁住，并通知工程师郭海滨跟进。 Repair Action: 告诉客户探测器物理结构无法锁住，并通知工程师郭海滨跟进。机器无故障。 Test/Inspection: customer confirm can scan patient and image ok.functional Check PASSED. Device returned to full functionality. Test Equipment: N/A Safety Questions : Case: 0118637059 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Bo Last updated by: Wang Bo 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 3 | [14 empty] | 118637059 | Closed | 1478558 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [19 empty] | Refresh |  |  | System issue | bar: No
1615 |  |  | Argentina | INSTITUTO NACIONAL DE | 40027 | 4450100% | 44 | SE BLOQUEO USUARIO DEBIDO A INTRODUCIR MAL LA CONTRASEÑA | Integration:Confirmation of order 54378031 saved ----------------------------------------------------------------------------- Work Order #: WO-04667653 [REDACTED] The system meets the specification for the performed service and is returned to use. Problem Description by Engineer : Problem description by engineer : THE USER GOT BLOCKED BECAUSE THE PASSWORD WAS ENTERED WRONGLY. When was the first occurrence? Within... : 01-Nov-2021 08:00:00 How often is issue occurring? : First occurrence Malfunction area : Console Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected Results: It is working properly. Patient Impact/Outcome: Delay in attention User Impact: Delay in work flow. Current Software Version: 4.1.10 Diagnostic performed by Engineer: We contacted the user, they said that a field service engineer visited the site and helped them to restore the password. Equipment functioning correctly. Internal Comments : Error # and/or description of error : Blocked password. Repair Action: We contacted the user, and they said that a field service engineer visited the site and helped them to restore the password. Equipment functioning correctly. Resolution: Serviced remotely. ----------------------------------------------------------------------------- | 2 | [14 empty] | 118636312 | Closed | 1476930 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [19 empty] | Refresh |  |  | System issue | bar: No
1616 |  |  | United States | Mount Grant General Hospital | 40012 | 4450300% | 44 | surview image questions | Integration : Confirmation of order 54391354 saved ----------------------------------------------------------------------------- Work Order #: WO-04696344 Curry, Dyan The system meets the specification for the performed service and is returned to use Problem description by Clinical Support Specialist : Customer Function/ Role: technologist Device Use: in use Expected and Unexpected Results: surview image labels User Impact: none Patient Impact: none Current Software version: 4.1.10 Diagnostic performed by Engineer : FSE calling for assistance. Dual surview labeling both images as 100-1. PACS not accepting 2 images labeled the same. Did test scan on KIVM 4.1. Surviews are both labeled 100-1. Advised FSE that is the way system is set up. There are no options to change from user /application side. Advised FSE to check with PACS admin , there could be a setting. Also to contact his NSS for advice. ----------------------------------------------------------------------------- | 4 | [14 empty] | 118661233 | Closed | 1486082 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Application | [6 empty] | Q&A | [5 empty] | Refresh |  |  | System issue | bar: No
1617 |  |  | United States | Adventhealth Zephyrhills | 300106 | 4453000% | 48 | 1302CDY72 slow transferring images to pacs | Work Order #: WO-04881472 02-Dec-2021 Watson Adam The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Problem description by engineer : slow transferring images to pacs Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: slow transferring images to pacs Patient Impact/Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer Current Software Version: 4.1.7.25030 Resolution : When engineer arrived onsite, images were transferring normally. Customer said the problem happens on larger studies. Checked bug report that customer has made, and nothing showed in the logs. Ran Tracert on both hospital CT’s and both had identical issues. When tracert was sent to the default gateway, the trace would happen immediately. When the trace was ran to the PACS and BackUp PACS, there were time out errors before successfully completing the trace. Engineer talked with next level support and they agreed that the problem appears to be a hospital switch. Instructed customer reach out to their IT department to troubleshoot network switches. No confirmed problem with the Philips CT. ----------------------------------------------------------------------------- Work Order #: WO-04879925 Sheffield, Scott The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Problem description by engineer : slow transferring images to pacs Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: slow transferring images to pacs Patient Impact/Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer Current Software Version: 4.1.7.25030 Diagnostic performed by Engineer : Russ said that this has been an ongoing problem for about 4 months. It is intermittent problem does not happen all the time. I guided Russ thru olevel to reset spooler and reboot host computer. The images transferred fine afterwards. I reviewed the bug report and did not find any problems. Follow Up Required : Error # and/or description of error : none Repair Action : Previous cases the FSE has noted problem is with the customer network. Russ told me the pacs ip address is 10.189.15.12 I compared the lanconfig settings for pacs for this scanner and the other scanner that Mark told me and they looked the same. system1 site # 5468842 transfers slow C:\>ping 10.189.15.12 ping 10.189.15.12 Pinging 10.189.15.12 with 32 bytes of data: Reply from 10.189.15.12: bytes=32 time<1ms TTL=254 Reply from 10.189.15.12: bytes=32 time<1ms TTL=254 Reply from 10.189.15.12: bytes=32 time<1ms TTL=254 Reply from 10.189.15.12: bytes=32 time<1ms TTL=254 Ping statistics for 10.189.15.12: Packets: Sent = 4, Received = 4, Lost = 0 (0% loss), Approximate round trip times in milli-seconds: Minimum = 0ms, Maximum = 0ms, Average = 0ms C:\>tracert 10.189.15.12 tracert 10.189.15.12 Tracing route to 10.189.15.12 over a maximum of 30 hops 1 * * * Request timed out. 2 <1 ms <1 ms <1 ms 10.189.15.12 Trace complete. scanner 54648530 transfers fine C:\>tracert 10.189.15.12 tracert 10.189.15.12 Tracing route to 10.189.15.12 over a maximum of 30 hops 1 * * * Request timed out. 2 <1 ms <1 ms <1 ms 10.189.15.12 Trace complete. Technician Name : Sheffield, Scott Internal Comments : Error # and/or description of error : none Repair Action : Previous cases the FSE has noted problem is with the customer network. Russ told me the pacs ip address is 10.189.15.12 I compared the lanconfig settings for pacs for this scanner and the other scanner that Mark told me and they looked the same. system1 site # 5468842 transfers slow C:\>ping 10.189.15.12 ping 10.189.15.12 Pinging 10.189.15.12 with 32 bytes of data: Reply from 10.189.15.12: bytes=32 time<1ms TTL=254 Reply from 10.189.15.12: bytes=32 time<1ms TTL=254 Reply from 10.189.15.12: bytes=32 time<1ms TTL=254 Reply from 10.189.15.12: bytes=32 time<1ms TTL=254 Ping statistics for 10.189.15.12: Packets: Sent = 4, Received = 4, Lost = 0 (0% loss), Approximate round trip times in milli-seconds: Minimum = 0ms, Maximum = 0ms, Average = 0ms C:\>tracert 10.189.15.12 tracert 10.189.15.12 Tracing route to 10.189.15.12 over a maximum of 30 hops 1 * * * Request timed out. 2 <1 ms <1 ms <1 ms 10.189.15.12 Trace complete. scanner 54648530 transfers fine C:\>tracert 10.189.15.12 tracert 10.189.15.12 Tracing route to 10.189.15.12 over a maximum of 30 hops 1 * * * Request timed out. 2 <1 ms <1 ms <1 ms 10.189.15.12 Trace complete. ----------------------------------------------------------------------------- | 3 |  | Customer said the problem happens on larger studies SW Version: SP1 FCO72800754 Russ said that this has been an ongoing problem for about 4 months. It is intermittent problem does not happen all the time. I guided Russ thru olevel to reset spooler and reboot host computer. The images transferred fine afterwards. I reviewed the bug report and did not find any problems. Engineer talked with next level support and they agreed that the problem appears to be a hospital switch. Instructed customer reach out to their IT department to troubleshoot network switches. No confirmed problem with the Philips CT. | [12 empty] | 118814165 | Fixed | 1571314 |  |  | Evaluation |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS Transfer | [5 empty] | Refresh |  |  | Upgrade | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y
1618 |  |  | United States | Lakeland Regional Medical Center | 300143 | 4452000% | 46 | CT Box flashing. | [16 empty] | 118753770 | Fixed | 1543896 |  |  | Closed - No Further Investigation | [20 empty] | Non-Refresh |  |  | Upgrade | bar: No
1619 |  |  | United States | Lakeland Regional Medical Center | 300143 | 4452000% | 46 | CT box flashing; Replaced CT box buttons and enabled button. | Work Order #: WO-04811947 22-Nov-2021 Watson Adam The system meets the specification for the performed service and is returned to use Parts Consumed : 1.00000,453567027202,Keypad assy, CT Box 1.00000,453566496941,RELAY,24VDC PLUGIN W/COVER Resolution : Replaced CT box buttons and enable button issue cleared. Attempted to perform a scan and logs showed horizontal motor relay errors. Replaced relay and performed test scans with no further issues. Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: ----------------------------------------------------------------------------- Work Order #: WO-04811517 21-Nov-2021 Sher Mark The system does not meet specification and is out of use Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: CT Box flashing User Impact: None Patient Impact: None Current Software Version: Diagnostic performed by Engineer : parts delay expedited Adam Watson will be taking over call Follow Up Required : replace button board ----------------------------------------------------------------------------- Work Order #: WO-04810589 21-Nov-2021 Sher Mark The system does not meet specification and is out of use Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: CT Box flashing User Impact: None Patient Impact: None Current Software Version: Diagnostic performed by Engineer : found system reports stuck command button ordered button board Follow Up Required : Install ordered parts ----------------------------------------------------------------------------- Work Order #: WO-04810495 Boyd Craig The system does not meet the specification for the performed service and is not returned to use Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: CT Box flashing User Impact: None Patient Impact: None Current Software Version: Diagnostic performed by Engineer : PRS accessed. ULV: S_COMMAND_BUTTON_STUCK_ERROR, AEnb Follow Up Required : FSE needed on site to evaluate CT Box Buttons. S_COMMAND_BUTTON_STUCK_ERROR , AEnb. Recommendations; Try to unstick Scan Enable button on CT Box. Replace parts if needed. CT Box 453567027195, CRC- Power Distribution Module 459801032622 ----------------------------------------------------------------------------- | [15 empty] | 118753804 | Closed | 1543400 |  |  | Closed - Done | [20 empty] | Non-Refresh |  |  | Upgrade | bar: No
1620 |  |  | Netherlands | Stichting Rijnstate Ziekenhuis | 52101 | 4450900% | 45 | Systeem wil niet scannen ganrtry can not initialize. | Work Order #: WO-04726691 09-Nov-2021 EDDY ELFRING The system meets the specifications for the performed service and was returned to use. Follow Up Required: send parts attn Frank Nieuwenhuijs Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution: Relay 24 DC replaced. ----------------------------------------------------------------------------- Work Order #: WO-04726599 JOHNNY BASTIAENS The system does not meet the full specifications for the performed service and was returned to use with limitations as accepted by the Customer, Follow-up work scheduled. Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnosis performed by Engineer : A restart did not solve the problem. A field service engineer will go on site. Follow-up Required : Diagnose. ----------------------------------------------------------------------------- Work Order #: WO-04734867 09-Nov-2021 Parts Consumed : 2.00000,453566496941,RELAY,24VDC PLUGIN W/COVER 10.00000,300006737161,ECG adaptor FRU ----------------------------------------------------------------------------- | [15 empty] | 118686034 | Closed | 1501932 |  |  | Evaluation | [20 empty] | Non-Refresh |  |  | Upgrade | bar: No
1621 |  |  | Netherlands | Stichting Rijnstate Ziekenhuis | 52101 | 4450900% | 45 | Gantry can not initialize | Integration:Confirmation of order 54413401 saved ----------------------------------------------------------------------------- Work Order #: WO-04726624 JOHNNY BASTIAENS The system meets the specifications for the performed service and was returned to use. Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer: Duplicate case 0118686034 ----------------------------------------------------------------------------- | [15 empty] | 118686038 | Closed | 1501912 |  |  | Closed - Duplicate | [20 empty] | Non-Refresh |  |  | Upgrade | bar: No
1622 |  |  | United States | Lakeland Regional Medical Center | 320474 | 4450400% | 44 | Head Arm Rest needs replacing | Internal Remarks : approved: per the attachment added to the case the customers contract should include accessory/consumables. ----------------------------------------------------------------------------- Work Order #: WO-04702339 08-Nov-2021 Watson Adam The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : Delivered pad to customer Parts Consumed : 1.00000,459800042241,Cushion-Head/ Armrest ----------------------------------------------------------------------------- | [15 empty] | 118665982 | Incomplete | 1493402 |  |  | Closed - Non-Complaint | [20 empty] | Non-Refresh |  |  | Upgrade | bar: No
1623 |  |  | United States | Philips Medical Capital | 40055 | 4453600% |  | Needs cable for in room monitor replaced | Work Order #: WO-04917752 Maier, Donnie The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Follow Up Required : Customer explained that the trigger cable for in room physiological monitor needs replaced (intermittently working). Looked up trigger cable part number in CHA 455012301901 - ECG Monitor Trigger 7M Ca (10 on Shelf in Parts-O-Pedia) after confirming with PFSE Rich Turley that is indeed what he is needing ordered to the site. I will order to site (verified address and POC with PFSE) and create onsite for PFSE Turley as requested. Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: Diagnostic Expected and Unexpected Results: Trigger cable for in room monitor needs replaced User Impact: No Patient Impact: No Current Software Version: 4.1.10.25210 Diagnostic performed by Engineer : Customer explained that the trigger cable for in room physiological monitor needs replaced (intermittently working). Looked up trigger cable part number in CHA 455012301901 - ECG Monitor Trigger 7M Ca after confirming with PFSE Rich Turley that is indeed what he is needing ordered to the site. I will order to site (verified address and POC with PFSE) and create onsite for PFSE Turley as requested. ----------------------------------------------------------------------------- Work Order #: WO-04917909 Turley Rich Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: Diagnostic Expected and Unexpected Results: Trigger cable for in room monitor needs replaced User Impact: No Patient Impact: No Current Software Version: 4.1.10.25210 Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 4 | [14 empty] | 118844487 | Closed | 1590033 |  |  | Opened | Opened | [6 empty] | Engineering | [6 empty] | ECG trigger cable | [5 empty] | IB Exist | 455012301901 | ECG Monitor Trigger 7M Ca | bar: No
1624 |  |  | China | 息烽县中医医院 | 345014 | 4453300% |  | 05k for IB/Ingenuity Core128/Elite China/CT box | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 05k for IB/Ingenuity Core128/Elite China/ Next Step Action : FSE Hu Ke 018585068636 call in, SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : 已回访成功 Safety Questions : Case: 0118832558 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Weng Jinming Last updated by: Weng Jinming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 工程师胡克替客户报修，需要更换CT box 设备本身技术问题： 是。 Troubleshooting Action: 远程诊断与测试结果： 工程师胡克替客户报修，需要更换CT box PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议工程师到场检查更换CT box 建议备件12NC及名称： CT box：459801704602 ----------------------------------------------------------------------------- | 3 | [14 empty] | 118832558 | Closed | 1585998 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | Engineering | [6 empty] | CT-Box | [5 empty] | IB Exist | 459801705022 | CT BOX ASSEMBLY_BEST | bar: No
1625 |  |  | United States | Philips Medical Capital | 40005 | 4453300% |  | Error Gantry cant comply. | Work Order #: WO-04906131 Greer Daniel The system does not meet the specification for the performed service and is not returned to use Follow Up Required : Error # and/or description of error : gantry can't comply error. Repair Action : Check tape switch's and foot switch's. Then slip ring data flow. Technician Name : Greer Daniel Internal Comments : Error # and/or description of error : gantry can't comply error. Repair Action : Check tape switch's and foot switch's. Then slip ring data flow. Problem Description by Engineer : Problem description by engineer : Gantry cannot comply errors. When was the first occurrence? Within... : 03-Dec-2021 How often is issue occurring? : First occurrence Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Gantry cannot comply errors. Patient Impact/Outcome: Not provided User Impact: System down Current Software Version: 4.1.10 Diagnostic performed by Engineer : Could not get logs to download. Customer went through shut down and power up. System still has gantry can't comply error. ----------------------------------------------------------------------------- Work Order #: WO-04906385 Womack Stephen Problem Description by Engineer : Problem description by engineer : Gantry cannot comply errors. When was the first occurrence? Within... : 03-Dec-2021 How often is issue occurring? : First occurrence Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Gantry cannot comply errors. Patient Impact/Outcome: Not provided User Impact: System down Current Software Version: 4.1.10 Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- Work Order #: WO-04907443 Dhaher Zaid Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 2 |  | Check tape switch's and foot switch's. Then slip ring data flow. Could not get logs to download. Customer went through shut down and power up. System still has gantry can't comply error. | [5 empty] | Gantry_HW | TR/RX | [5 empty] | 118835316 | In Progress | 1584225 |  |  | Opened | Opened | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Gantry Can't comply | [5 empty] | Refresh |  |  | System issue | System | Gantry-SW | NoEnoughInfo | Gantry-SW-NoEnoughInfo | Gantry-SW-NoEnoughInfo | NoEnoughInfo | NoEnoughInfo | Gantry-SW | Gantry-SW-NoEnoughInfo | bar: No
1626 |  |  | China | Wuhan Province Wuhan Ren’ai Hospital | 345004 | 4453200% |  | 05K for IB/Ingenuity Core128/Elite China/Slow image reconstruction | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 05K for IB/Ingenuity Core128/Elite China Safety Questions:Case: 0118826256 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingsong Last updated by: Lu Mingsong 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer's problems: Image reconstruction is slow. Image transmission often fails Detailed description of Customer's problems or malfunctions: Image reconstruction is slow. Image transmission often fails It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: N/A Remote troubleshoot and test result: N/A PRS TR status: No PRS. Repair Action: Suggest onsite service procedures: Clean the reconstruction cabinet filter mesh. Check the reconstruction card cooling fan and software system. We recommend spare parts 12NC and names: N/A Memo of Customer request: The customer wants an engineer to go on site to do maintenance SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support ----------------------------------------------------------------------------- | 3 | [14 empty] | 118826256 | In Process | 1582414 |  |  | Opened | Opened | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | Recon slow, image transfer fail | [5 empty] | Refresh |  |  |  | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | bar: No
1627 |  |  | Japan | Independent administrative institution National Hospital Organization | 341015 | 4453300% |  | X-ray system cannot complyと出てスキャンが途中で停止した。>>再撮影でスキャン可。>>S_GEN_FAULT_ANODE_TUBE_ARC | Work Order #: WO-04901733 Kojima, Kenichi The system meets the specification for the performed service and is returned to use. Problem Description by Engineer : The error message "X-ray system cannot comply" appeared and a scan stopped partway through.>> Another attempt was made and scanning was possible.>> S_GEN_FAULT_ANODE_TUBE_ARC Resolution : [Subject] 03CS:KK:Mr. Hisaki: The error message "X-ray system cannot comply" appeared and a scan stopped partway through.>> Another attempt was made and scanning was possible.>> S_GEN_FAULT_ANODE_TUBE_ARC [4/8 Elements] Q5. How did the examination end? Rescanned from where scan had stopped. Q6. Did any mistaken findings (mistaken diagnosis or treatment) result from the problem? No. Q7. Was there any unintended unnecessary exposure? No Q8 Artifacts? None Resolution: Arcing occurred.It seems to have been a once-off problem. Scanning was possible after a rescan.Suggested the client to wait-and-see. From the logger, 2021/12/03 11:53:27.320__02 [--o] ----------------- Series 76 Ready ----------------- 2021/12/03 11:53:30.090__01 [-oo] - - - - - - Shot 0 Starts - - - - - - 2021/12/03 11:53:38.481 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_ANALOG_FAULT szInfo:2:5 Analog Fault P1:2/P2:32 2021/12/03 11:53:38.523 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_ANODE_TUBE_ARC 2021/12/03 11:53:38.527 <*HV*> Type: NORMAL ErrNum: S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT szInfo: P1:0/P2: 2021/12/03 11:53:38.533__02 [CANCEL] CStateShot:Series 76 cancelled, due to GANTRY_CANCEL_SOURCE_NORMAL_ERROR, in state CStateShot 2021/12/03 11:53:38.536__02 [***] - - - - - - Study Exit - - - - - - 2021/12/03 11:53:38.536__04 <WARN> Type: WARN ErrNum: S_GEN_MGR_SYNC_OFF_OUT_OF_SEQUENCE P1:0/P2:0 2021/12/03 11:53:40.581 <WARN> Type: WARN ErrNum: S_FDOM_PROFILE_FRAME_ISR_TBL_END_DEFAULT_MA_USED P1:0/P2:0 2021/12/03 11:53:40.586__02 <*ERROR*> Type: NORMAL ErrNum: S_MASTER_SEQUENCER_EXIT_STATE_ERR P1:4096/P2:8369 2021/12/03 11:53:40.605__01 <WARN> Type: WARN ErrNum: S_FDOM_PROFILE_FRAME_ISR_TBL_END_DEFAULT_MA_USED P1:0/P2:0 ----------------------------------------------------------------------------- | 4 | [14 empty] | 118831795 | Closed | 1581622 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Arc | [5 empty] | IB Exist | bar: No
1628 |  |  | Japan | 医療法人清翠会 牧病院 | 341017 | 4453200% |  | PACS側へ画像転送が止まっている＞＞CTホストの再起動にて復旧 | Work Order #: WO-04892792 Kimura, Kazumasa Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 4 | [14 empty] | 118824423 | Closed | 1581418 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS transfer stop, back to normal after Host reboot | [5 empty] | Refresh |  |  |  | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | bar: No
1629 |  |  | United States | Odessa Regional Medical Center | 40029 | 4453100% |  | Changing Table Top | Work Order #: WO-04890835 Brian, Kluever, Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- Work Order #: WO-04890855 Zeitouni Fadi Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 5 |  | Adjust the table cover | [12 empty] | 118822892 | Closed | 1580458 |  |  | Evaluation | Evaluation | [6 empty] | Other | [6 empty] | Noise with table movement | [5 empty] | IB Exist | bar: No
1630 |  |  | China | 内蒙古阿拉善右旗人民医院 | 345031 | 4453100% |  | 02N/Core128/artifact | BU Q&R Feedback: No Reason: Corrective Maintenance Next Step Action : PPD 459800911281*1 24H FSE:梁震-86-15849133876 备件已安排出货 Safety Questions : Case: 0118820007 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wennian Han Last updated by: Wennian Han 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 图像伪影。 客户问题详细描述或故障现象： 环形伪影 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： NA 远程诊断与测试结果： 环形伪影，大关重启无效 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 检查模块，屏蔽通道，结合场地实际情况，必要时更换模块 建议备件12NC及名称： " 459800911281 TDMS V2 M.Det Class B FRU" Next Step Action : 派工先联系su ruijie SAP Transfer : Dispatched for Onsite Support SAP Transfer : Case Transferred to SAP Problem Reported by Customer : 02N/Core128/ Next Step Action : NTD:备件组接到客户报修电话 报修2 ：345031 内蒙古阿拉善右旗人民医院 Ingenuity Core128/Elite China（客户给的SN查到的是128排，但她报修说是64排，请确认） 故障：图像有环形伪影 联系人： 王老师- 13948833710 已转RSE：hanwennian Next Step Action : FSE 苏瑞杰 13304719335 电话没接，稍后继续联系 Next Step Action : FSE 苏瑞杰 13304719335告知可以派给他，已发派工短信通知到工程师 ----------------------------------------------------------------------------- | 3 |  | Replace M30 | [12 empty] | 118820007 | Closed | 1578394 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | Engineering | [6 empty] | Artifact | [5 empty] | IB Exist | 459800911281 | TDMS V2.X M.Det Class B FRU | bar: No
1631 |  |  | Japan | 医療法人清翠会 牧病院 | 341017 | 4453200% |  | sub_selectionの使用方法ですが、2つのシリーズを同時に読み込むと転送時にWW/WLが同じになってしまいます。 | Work Order #: WO-04894151 Nakagawa, Futoshi The system meets the specification for the performed service and is returned to use Resolution : どちらもシリーズ番号が小さい方のWW/WLになってしまいます。基本的には2つ読み込んで処理するアプリケーションではないため、ご希望に添えない仕様となっています。 Problem description by Clinical Support Specialist : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : NA ----------------------------------------------------------------------------- | 5 | [14 empty] | 118825781 | Closed | 1577665 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Others | [6 empty] | WW/WL | [5 empty] | Refresh | bar: No
1632 |  |  | China | 中国医科大学附属盛京医院大连分院 | 345001 | 4453100% |  | 05R/Ingenuity Core128/Elite China/estop open | BU Q&R Feedback: No Reason: Corrective Maintenance Next Step Action : PPD POWER SUPPLY 120 VAC DUAL OUTPUT ROHS 453567374153 1，24H场地，FSE于志龙-86-18640955881 Problem Reported by Customer : 05R/Ingenuity Core128/Elite China Safety Questions : Case: 0118817570 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Bo Last updated by: Wang Bo 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Dispatched for Onsite Support T2 Activities : Problem Description: 客户问题概述： ESTOP断开。 客户问题详细描述或故障现象： 机架掉电。 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： LIFESYNC P1=2&4 远程诊断与测试结果： 重启无效。 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议测量vicor和FRC供电，必要时先更换VICOR排除掉电故障。 建议备件12NC及名称： PPD POWER SUPPLY 120 VAC DUAL OUTPUT 453567374153 SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | Replace the Dual Output Vicor PS | [12 empty] | 118817570 | Closed | 1577280 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | Engineering | [6 empty] | E-stop open | [5 empty] | IB Exist | 453567374153 | POWER SUPPLY 120 VAC DUAL OUTPUT ROHS | bar: No
1633 |  |  | China | 中国医科大学附属盛京医院大连分院 | 345001 | 4453100% |  | 05R/Ingenuity Core128/Elite China/scan fault | BU Q&R Feedback: No Reason: Corrective Maintenance Next Step Action : PPD MODULE, ROTOR ELECTRONICS TX 453567538334 1，24H场地，FSE于志龙-86-18640955881 Problem Reported by Customer : 05R/Ingenuity Core128/Elite China T2 Activities : Problem Description: 客户问题概述： 扫描失败。 客户问题详细描述或故障现象： 机架扫描中断。 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： S_DMC_SHOT_COMM_COLLAPSE physical address 0x0 远程诊断与测试结果： 重启无效。工程师已到场诊断。 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议测量TX输入，必要时更换TX。 建议备件12NC及名称： "PPD MODULE, ROTOR ELECTRONICS TX 453567538334 " Safety Questions : Case: 0118816819 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Bo Last updated by: Wang Bo 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support ----------------------------------------------------------------------------- | 3 |  | Scan cannot initialize, TX no output, replace the TX | [12 empty] | 118816819 | Closed | 1576930 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | Engineering | [6 empty] | Scan cannot initialize | [5 empty] | IB Exist | 453567538334 | ROTOR ELECTRONICS TX | bar: No
1634 |  |  | China | Xifeng County Traditional Chinese Medicine Hospital | 345014 | 4453100% |  | 05K For IB/Ingenuity Core128/Elite China/Belt cracking | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 05K for IB/Ingenuity Core128/Elite China Safety Questions:Case: 0118816542 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Bai Jingwen Last updated by: Bai Jingwen 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action: Tang Guanglin found the belt is broken during maintenance and needs to be replaced. Dispatch hu ke to RSE, ask RSE PPD part T2 Activities:Problem Description: Summary of customer's problems: BELT BROKEN Detailed description of Customer's problems or malfunctions: The rotating motor belt of the frame cracked It's a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: N/A Remote troubleshoot and test result: N/A PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Replace the gantry's rotation motor belt. We recommend spare parts 12NC and names: POLY-VEE BELT: 453566492401 Memo of Customer request: None Next Step Action:PPD POLY-VEE BELT 453566492401 1.24H on site, FSE Tang Guanglin -86-18084189777 SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support ----------------------------------------------------------------------------- | 3 |  | Belet crack found during PM, replace the belt | [12 empty] | 118816542 | Closed | 1576707 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Belt crack | [5 empty] | IB Exist | 453566492401 | POLY-VEE BELT | bar: No
1635 |  |  | Japan | 社会医療法人 母恋 日鋼記念病院 | 341011 | 4454300% |  | ShortTC時にアーク発生＞＞継続してShortTCを行い正常終了＞S_GEN_FAULT_ANODE_TUBE_ARC/S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT | Work Order #: WO-04970689 Kimura, Kazumasa The system meets the specification for the performed service and is returned to use Follow Up Required : * Internal Comments : [Subject] 03DS:KKI:アンドウ:ShortTC時にアーク発生＞＞継続してShortTCを行い正常終了＞S_GEN_FAULT_ANODE_TUBE_ARC/S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT [Internal Comments] ログ確認結果 以下のエラーが発生しております。 S_GEN_FAULT_ANODE_TUBE_ARC S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT [出庫部品] 453567560012:Spellman 80 KW RoHS APM Assy# QTY:1 ZREP/Traceable 東京(U611) 989000005011:HV-ANODECABLE PHILIPS MRC880 CT 1,84m QTY:1 ZSID 札幌(U612) 453566495091:1XRT SERVICE TOOL ASSY EXP QTY:1 ZSID 札幌(U612) 453566495101:GENERATOR SERVICE TOOL QTY:1 ZSID 札幌(U612) 459800089031:Tool, XRT Strap Pad Assy QTY:1 ZSID 札幌(U612) 989000086352:PHILIPS MRC 880 QTY:1 ZREP/Traceable 札幌(U612) [ディスパッチ] 1 RC -----------------FSE登録情報----------------- -訪問FSE1氏名 橋本 穂高 -FSE1訪問日時 2021/12/141:00 -FSE1作業時間 2 [出荷指示] -----------------出荷指示#1----------------- -配送条件 ASAP(輸送手配) -納期 ASAP -配送先名 札幌倉庫止め・FSEが引取り -受取者名 橋本 穂高 -電話番号 070-3876-3812 -その他 -依頼者名 木村 和正 Problem Description by Engineer : ShortTC時にアーク発生 Diagnostic performed by Engineer : 継続してShortTCを行い正常終了 ----------------------------------------------------------------------------- Work Order #: WO-04971129 Hashimoto, Hodaka Regulatory Questions : N/A Problem Description by Engineer : ShortTC時にアーク発生 Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 4 | [14 empty] | 118887945 | Closed | 1612666 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | Engineering | [6 empty] | Arc | [5 empty] | IB Exist | 453567560012 | Spellman 80 KW RoHS APM Assy | bar: No
1636 |  |  | Turkey | ADIYAMANOZELSAGINSTURTASSANVETICA.S | 341006 | 4454300% |  | CT belt siparişi | Work Order #: WO-04969812 Salih Mendi The system meets the specification for the performed service and is returned to use Diagnostic performed by Engineer : CT belt siparişi Problem Description by Engineer : Problem description by engineer : Main belt noise issue Information to support the complaint handling process Customer Function/Role: Not provided Device Use: Not provided Expected and Unexpected results: Not provided Patient Impact/Outcome: Not provided User Impact: Not provided Current Software Version: Not provided Internal Comments : Error # and/or description of error : Main belt noise issue Repair Action : Main belt noise issue Extra Internal comment field : Belt değişiminden sonra Tool "unused- good and sealed" olarak raporlanmalıdır. Follow Up Required : Error # and/or description of error : Main belt noise issue Repair Action : Main belt noise issue Extra Internal comment field : Belt değişiminden sonra Tool "unused- good and sealed" olarak raporlanmalıdır. Technician Name : Salih Mendi Resolution : Main belt'in rotasyon sırasında gürültü çıkardığı şikayeti üzerine, main belt siparişi yapılacaktır. Sistem faal durumdadır. ----------------------------------------------------------------------------- | 5 | [14 empty] | 118887293 | Closed | 1612234 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | Engineering | [6 empty] | Belt noise | [5 empty] | IB Exist | 453566492401 | POLY-VEE BELT | bar: No
1637 |  |  | Japan | 独立行政法人国立病院機構 | 341013 | 4454300% |  | 撮影時にカタンカタンという異音が発生 | Work Order #: WO-04967241 Miyakawa, Takeyuki The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Follow Up Required : * Diagnostic performed by Engineer : a Problem Description by Engineer : a ----------------------------------------------------------------------------- Work Order #: WO-04967265 Takada, Ushio Problem Description by Engineer : a Regulatory Questions : N/A Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | [15 empty] | 118884943 | Closed | 1611472 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | Engineering | [6 empty] | DMS Cover noise | [5 empty] | IB Exist | 459800624062 | TDMS Cover RoHS FRU | bar: No
1638 |  |  | China | 河南省濮阳市人民医院 | 40031 | 4454100% |  | Wechat:05K for IB/Ingenuity CT 728326/Stop scanning | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : Stop scanning Safety Questions : Case: 0118882346 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Zhou Xiaohong Last updated by: Zhou Xiaohong 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 扫描中断 客户问题详细描述或故障现象： 扫描中断 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： S_MASTER_SEQUENCER_ACQ_NOT_OK；S_DMC_SHOT_COMM_COLLAPSE 远程诊断与测试结果： 不能远程连接重建主机，eventlog不能下载 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 现场检查数据通路，做可视化诊断，查看eventlog，检查TX,RX,需要时更换 建议备件12NC及名称： N/A 客户需求备注： 无 Next Step Action : Luo Yongping SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support ----------------------------------------------------------------------------- |  |  | ACQ not ok, resend, replace both TX and RX | [12 empty] | 118882346 | Closed | 1610864 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | Engineering | [6 empty] | ACQ not ok | [5 empty] | IB Exist | Not provided | TX and RX | bar: No
1639 |  |  | United States | Philips Medical Capital | 40053 | 4454000% |  | EKG isn't connecting to the CT | Internal Remarks : FSE did PM lastnight and now ekg not connecting to the CT. ----------------------------------------------------------------------------- |  |  | Conducted IBOX test, checked all cable connections and full reboot of the gantry. IBOX is functioning normally. | [12 empty] | 118881127 | Closed | 1610581 |  |  | Opened | Opened | [6 empty] | Other | [6 empty] | ECG isn't connecting to CT | [5 empty] | IB Exist | bar: No
1640 |  |  | Japan | Maki Hospital | 341017 | 4454000% |  | PACS側へ画像転送が止まっている | Work Order #: WO-04955542 Kawazoe, Junichi Regulatory Questions : N/A Diagnostic performed by Engineer: N/A Problem Description by Engineer: Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: ----------------------------------------------------------------------------- | 5 | [14 empty] | 118876751 | In Progress | 1609141 |  |  | Evaluation | Evaluation | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | Manual PACS transfer stuck | [5 empty] | Refresh |  |  |  | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | bar: No
1641 |  |  | United States | Odessa Regional Medical Center | 40029 | 4454000% |  | IQ is failing temperature control malfunction errors | Internal Remarks : Planner Initials: JJ Actions Taken: FSE-Brian Kluever called in to have WO-04958358 dispatch to him. FSE said he will contact the cx I put FADB in the planning notes I dispatched the WO per FSE request Next Steps: N/A Last Updated: 12:40PM EST 12/10/21 ----------------------------------------------------------------------------- Work Order #: WO-04958263 Maier, Donnie The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Follow Up Required : Customer states they are unable to perform the scheduled IQ test they are trying to run. They are getting temperature control malfunction errors. I asked if gantry room feels overly warm and she states no, normal temperatures. I asked about any weird noises like fans going bad and she stated no. system doesn't appear to have telnet of FTP ports open to pull logs. I asked if customer has tried to perform a normal scan to see if there are artifacts and they have not. I let them know that maybe they can try putting phantom up on table and running belly scan to check for artifacts. I will dispatch FSE to try and duplicate and perform further troubleshooting onsite as we don't have access remotely. Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: No Patient Involved Expected and Unexpected Results: Cant pass the normal IQ Check giving Temperatuere Control Malfunction Error User Impact: No Patient Impact: No Current Software Version: 4.1.10 Diagnostic performed by Engineer : Customer states they are unable to perform the scheduled IQ test they are trying to run. They are getting temperature control malfunction errors. I asked if gantry room feels overly warm and she states no, normal temperatures. I asked about any weird noises like fans going bad and she stated no. system doesn't appear to have telnet of FTP ports open to pull logs. I asked if customer has tried to perform a normal scan to see if there are artifacts and they have not. I let them know that maybe they can try putting phantom up on table and running belly scan to check for artifacts. I will dispatch FSE to try and duplicate and perform further troubleshooting onsite as we don't have access remotely. ----------------------------------------------------------------------------- Work Order #: WO-04958358 Brian, Kluever, Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: No Patient Involved Expected and Unexpected Results: Cant pass the normal IQ Check giving Temperatuere Control Malfunction Error User Impact: No Patient Impact: No Current Software Version: 4.1.10 Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- |  |  | With occasion heating error on IQ only, can pass the IQ test, ask customer to check if will happen again. | [12 empty] | 118879015 | Closed | 1606802 |  |  | Opened | Opened | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Heating error | [5 empty] | Refresh |  |  |  | System | Console-SW | NoEnoughInfo | Console-SW-NoEnoughInfo | Console-SW-NoEnoughInfo | NoEnoughInfo | NoEnoughInfo | Console-SW | Console-SW-NoEnoughInfo | bar: No
1642 |  |  | China | Jinan City No. 4 People's Hospital | 40038 | 4453900% |  | 02N/Ingenuity CT 728326/Delete data | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: Database Recreation / Recovery Problem Reported by Customer: 02N/Ingenuity CT 728326 Next Step Action:FSE Tao Chen 018560023293 call in and disp him Safety Questions:Case: 0118859120 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Guoping Last updated by: Liu Guoping 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: need delete data Troubleshooting Action: need delete date via PRS Repair Action: delete date via PRS Test/Inspection: Functional Check PASSED. The device returned to full functionality. Test Equipment: N/A SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- |  |  | Delete data via PRS | [12 empty] | 118859120 | Closed | 1604616 |  |  | Opened | Opened | [6 empty] | TBD | [6 empty] | Delete data | [5 empty] | Refresh | bar: No
1643 |  |  | China | Hebei Province Tangshan City Workers Hospital | 40044 | 4453900% |  | 02N/Ingenuity CT 728326/barcode reader missing | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Replace Part SAP Transfer:Case Transferred to SAP Problem Reported by Customer: 02N/Ingenuity CT 728326 Safety Questions:Case: 0118858991 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Gaoge Last updated by: Wang Gaoge 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: The person who reported the repair said the scan code gun that came with the machine was lost. Need to send another code scan gun Resolver Action: nemo code scan gun to the customer, the customer replaced it by themselves Repair Action: nemo 455018005052 Test/Inspection: N/A Test Equipment: N/A ----------------------------------------------------------------------------- | [15 empty] | 118858991 | Closed | 1604511 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Customer | [6 empty] | Barcoder scanner missing | [5 empty] | Non-refresh | 455018005052 | BARCODE SCANNER LS2208 USB | bar: No
1644 |  |  | United States | Odessa Regional Medical Center | 40029 | 4453900% |  | Need to remove the word 'draft' off of the calcium scoring report. | Internal Remarks : Interaction Type: New Case Caller: Liliana Echeverria IP & Issue: 84812213; Customer requesting clinical support: Need to remove the word 'draft' off of the calcium scoring report. Event Type: Incident Service Request: Remote Clinical Did Customer Decline Remote Support? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Created Case & Work Order Next Step Action: N/A Verbal PO Number: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- Work Order #: WO-04950411 Mujica, David The system meets the specification for the performed service and is returned to use Problem description by Clinical Support Specialist : Customer Function/Role: Technologist Device Use: In use Expected and Unexpected Results: Draft on Report Removal User Impact: none Patient Impact: none Current Software Version: 4.1 Diagnostic performed by Engineer : Liliana required support for removing the word "draft" off of their Calcium Score reports. After final approval of the the reporting process the system prompts you where you would like to send the final report to. At this phase you must select printer and then the final report will not say "draft". Liliana had no further questions or need for clinical support. No rescan, no loss of data, no reinjection per Liliana. ----------------------------------------------------------------------------- | 5 |  | After final approval of the the reporting process the system prompts you where you would like to send the final report to. At this phase you must select printer and then the final report will not say "draft". | [12 empty] | 118872043 | Closed | 1603038 |  |  | Opened | Opened | [6 empty] | Customer | [6 empty] | Remove "draft" in the calcium scoring report | [5 empty] | Non-refresh | bar: No
1645 |  |  | Puerto Rico | METRO MAYAGUEZ INC | 40009 | 4453900% |  | Throw message Gantry cannot comply and the table gets stuck | Work Order #: WO-04950634 09-Dec-2021 Valenzuela Velasquez, Eduardo Ariel The system does not meet the specification for the performed service and is not returned to use Problem Description by Engineer : Problem description by engineer : Gantry cannot comply Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: NO gantry errors. Patient Impact/Outcome: System is down User Impact: System is down Current Software Version: 4.1.10 Internal Comments : Error # and/or description of error : S_DMC_REFERENCE_CONV_DATA Repair Action : - Power supply replacement. Follow Up Required : Error # and/or description of error : S_DMC_REFERENCE_CONV_DATA Repair Action : - Power supply replacement. Technician Name : Valenzuela Velasquez, Eduardo Ariel Diagnostic performed by Engineer : Gantry was checked and a power supply is bad. The part will be ordered. ----------------------------------------------------------------------------- Work Order #: WO-04948963 09-Dec-2021 Hernandez, Alvin The system does not meet the specification for the performed service and is not returned to use Follow Up Required : The Reference detector Power Supply for the Ingenuity CT will be Arrived tomorrow by FedEx. The part was ordered by Ing Edwardo Valenzuela. I will be here to install it. Resolution : The power supply was ordered by FSE Edwardo Valenzuela. It will be installed as soon as I received it. Problem Description by Engineer : Problem description by engineer : Gantry cannnot comply. Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: NO gantry error. Patient Impact/Outcome: System is down. User Impact: System is down Current Software Version: 4.1.10 Diagnostic performed by Engineer : The problem was caused by the -5.0/+5.0 Vdc defectie Reference Detector power supply. this power supply needs to be replaced. ----------------------------------------------------------------------------- Work Order #: WO-04948201 Valenzuela Velasquez, Eduardo Ariel The system does not meet the specification for the performed service and is not returned to use Problem Description by Engineer : Problem description by engineer : Gantry cannnot comply. Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: NO gantry error. Patient Impact/Outcome: System is down. User Impact: System is down Current Software Version: 4.1.10 Diagnostic performed by Engineer : System is showing an error with the reference converter. Gantry mus be checked for power supply and reference converter. Internal Comments : Error # and/or description of error : S_DMC_REFERENCE_CONV_DATA Repair Action : Check POWER SUPPLY, TRIPLE OUTPUT 459800359381 - Check power supply - Check reference converter and cables - Check DMS fuses - Check contrast media Follow Up Required : Error # and/or description of error : S_DMC_REFERENCE_CONV_DATA Repair Action : Check POWER SUPPLY, TRIPLE OUTPUT 459800359381 - Check power supply - Check reference converter and cables - Check DMS fuses - Check contrast media Technician Name : Valenzuela Velasquez, Eduardo Ariel | 2 | [14 empty] | 118870499 | Closed | 1602219 |  |  | Opened | Opened | [6 empty] | Engineering | [6 empty] | Gantry cannot comply | [5 empty] | IB Exist | 459800723941 | PSA5V SUPPLY ASS'Y | bar: No
1646 |  |  | Turkey | ADIYAMANOZELSAGINSTURTASSANVETICA.S | 341006 | 4453900% |  | Kontrat madde temizliği (Contrast matter cleanup ) | Work Order #: WO-04948632 Turgay Baylan The system meets the specification for the performed service and is returned to use Resolution : CT cihazı bakım için program yapılması gerekmektedir, bakım sırasında kontrast temizliğide yapılacaktır. Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: ----------------------------------------------------------------------------- | 3 | [14 empty] | 118870833 | Closed | 1602064 |  |  | Evaluation - Pending Translation | Evaluation - Pending Translation | [6 empty] | TBD | [6 empty] | Contrast matter cleanup | [5 empty] | IB Exist | bar: No
1647 |  |  | Spain | S.C.S. HOSP. GRAL. DE FUERTEVENTURA | 40020 | 4453900% |  | Salen las imágenes con rayas El fallo ha comenzado ayer En uso clínico Ubicación: Planta 1era | Work Order #: WO-04937851 [REDACTED] The system does not meet the specification for the performed service and is not returned to use. Follow Up Required: We need to go and check the central detector modules. Exchange with external module to see if the artifact moves. Replace module as needed. Regulatory Questions : N/A Problem Description by Engineer: I talk to [REDACTED]. A very noticeable artifact can be observed in all the acquired images. Straight lines focused in the center. Diagnostic performed by Engineer: Images from different protocols (cranial and abdomen studies) are checked, and we observe what seems to be a failure of either a detector or the central detectors. They cause lines in the shape of stretch marks. ----------------------------------------------------------------------------- | 2 |  | Se sustituye el módulo 21 del detector que no pasaba los test de detectores. Una vez sustituido se realizan todos los test indicados en el manual de usuario tras sustitución de módulos incluyendo comprobaciones de calidad de imagen para comprobar que no se reproducen los artefactos. | [12 empty] | 118860492 | Closed | 1601184 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Artifact | [5 empty] | IB Exist | 459800911271 | TDMS V2 M.Det Class A FRU | bar: No
1648 |  |  | China | The Second Hospital of Haibei Prefecture in Qinghai Province | 345029 | 4453800% |  | 02N/Ingenuity Core128/Elite China/Replace worklist parameters | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : Replace worklist parameters Safety Questions:Case: 0118854890 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Zixuan Last updated by: Wang Zixuan 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support T2 Activities:Problem Description: Summary of customer's problems: Change the PACS system, it's not in a network segment, a field service engineer is needed on site to work with them. It's a technology problem with the equipment itself: Yes. Troubleshooting Action: remote diagnosis and testing result: Change the PACS system, it's not in a network segment, a field service engineer is needed on site to work with them. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Change the PACS system, it's not in a network segment, a field service engineer is needed on site to work with them. Memo of Customer request: Contact the person who requested service and confirm the time. Next Step Action: Site1 Field Service Engineer Xu Yonghua 13327661334 Inform OK to dispatch him, a dispatch text message has been sent to the engineer ----------------------------------------------------------------------------- | 5 | [14 empty] | 118854890 | Closed | 1601133 |  |  | Closed - Non-Complaint | Opened | [6 empty] | Customer | [6 empty] | Support for hospital PACS installation | [5 empty] | IB Existing | bar: No
1649 |  |  | China | Anhui Tongling City Bo'ai hospital new campus | 345034 | 4453800% |  | 02N/Ingenuity Core128/Elite China/nstall the high-pressure syringe | BU Q&R Feedback: No Reason: Corrective Maintenance Safety Questions:Case: 0118854624 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Bo Last updated by: Wang Bo 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ T2 Activities:Problem Description: Customer need to dispatch labor to install high pressure injector. Ask C1 to follow up. Troubleshooting Action: Customer need to dispatch labor to install high pressure injector. Ask C1 to follow up. Repair Action: Customer need to dispatch labor to install high pressure injector. Ask C1 help to follow up. ----------------------------------------------------------------------------- | [15 empty] | 118854624 | In Process | 1600931 |  |  | Closed - Non-Complaint | Opened | [6 empty] | Customer | [6 empty] | Do injector installation | [5 empty] | IB Existing | bar: No
1650 |  |  | Japan | Independent administrative institution National Hospital Organization | 341015 | 4453900% |  | X-Ray system cannot complyとメッセージが出てスキャンが止まった。>>ログを確認しましたがエラーは見受けられず。 | Work Order #: WO-04937298 Kojima, Kenichi The system meets the specification for the performed service and is returned to use. Problem Description by Engineer : The message "X-ray system cannot comply" appeared and a scan stopped.>> I checked the log, but there was no error. Resolution: I checked the log but did not find any error, such as arcing. It seems "X-ray system cannot comply" appeared, so a high-voltage system error may have occurred. I had the FSE contact the customer, and he agreed to keep an eye on it. Internal Comments: [Subject] 03CS:KK:Mr. Hisaki: "X-ray system cannot comply" appeared and a scan stopped.>> I checked the log, but there was no error. [4/8 Elements] Q5. How did the examination end? Rescanned from where scan had stopped. Q6. Did any mistaken findings (mistaken diagnosis or treatment) result from the problem? No. Q7. Was there any unintended unnecessary exposure? No Q8 Artifacts? None ----------------------------------------------------------------------------- | 4 |  | ログを確認しましたがアークなどのエラーは確認できず。 X-Ray system cannot complyと出たらしいので何かしら高圧系のエラーが起きたと思われます。 FSEより連絡を入れていただき様子見とさせていただきました。 | [12 empty] | 118859965 | Closed | 1600567 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | XRAY system cannot comply | [5 empty] | IB Exist |  |  |  | System | Gantry-SW | bar: No
1651 |  |  | Japan | Maki Hospital | 341017 | 4453700% |  | Long TC作業 | Work Order #: WO-04922289 08-Dec-2021 Kawazoe, Junichi The system meets the specification for the performed service and is returned to use. Internal Comments: [CX related] Did you agree on how to complete the report with the person in charge of devices? <Select one of the following> Yes. I completed normal reporting that day. [Work report (paper) submission] Did you submit a work report on paper? Yes Q1. How did the examination finish? Select from the following: *Because the answer is No for Q1, not applicable Q2. Were there any mistaken findings (misdiagnosis, mistaken treatment) due to the problem? None. Q3. Were there any artifacts? None. Q4. Was there any unintended unnecessary exposure? None. Resolution : Work details: I performed a tube adjustment (Long Tube Conditioning, Filament Calibration). There is no problem with the operation of the equipment. Problem Description by Engineer: Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Regulatory Questions : N/A ----------------------------------------------------------------------------- | 5 |  | Long Tube condition and filament calibration | [12 empty] | 118848277 | Closed | 1596654 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Engineering | [6 empty] | PM to do LTC and Filament | [5 empty] | IB Existing | bar: No
1652 |  |  | China | Hebei Province Geriatric Hospital | 345015 | 4453600% |  | 01/Ingenuity Core128/Elite China/add dicom node | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Calibration/Adjustment Next Step Action: Dispatched Engineer Liang Longse to follow 159 3113 3573 up Problem Reported by Customer: 01/Ingenuity Core128/Elite China Next Step Action : Transfer to sales handling, this case is valid within 90 days starts from this case created, once beyond of this period without any further steps, this case is considered invalid, customer refuses our service Next Step Action: contacted JiuNuo ASP to follow up with Business Safety Questions:Case: 0118840146 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Gaoge Last updated by: Wang Gaoge 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer's problems: Other types of error. Detailed description of Customer's problems or malfunctions: Need to add nodes. Admin password doesn't help It's a technology problem with the equipment itself: No. Troubleshooting Action: key fault code or symptom: Admin password is missing Remote troubleshoot and test result: Need an engineer to reset admin password on the site. PRS reset failed PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: 1. Reset admin password Test/Inspection: N/A Test Equipment: N/A ----------------------------------------------------------------------------- | [15 empty] | 118840146 | In Process | 1590950 |  |  | Opened | Opened | [6 empty] | Customer | [6 empty] | Forget IT password | [5 empty] | Refresh | bar: No
1653 |  |  | China | 河北省唐山市工人医院 | 40044 | 4454900% |  | 02N/Ingenuity CT 728326/cable failed |  | 3 | [14 empty] | 118924822 | Closed | 1635285 |  |  | Opened - Pending Translation | Opened - Pending Translation | CTD00006847 | [5 empty] | Service | [6 empty] | Video cable | [5 empty] | IB Exist |  |  |  | System | CIRS-SW | ACQ_NOT_OK | CIRS-SW-ACQ_NOT_OK | CIRS-SW-ACQ_NOT_OK | CTD00006847 | ACQ_NOT_OK | CIRS-SW | CIRS-SW-ACQ_NOT_OK | bar: No
1654 |  |  | China | 河北省唐山市工人医院 | 40044 | 4454900% |  | 02N/Ingenuity CT 728326/dispaly problem |  |  |  | 机器无法使用，处理视频线后正常。 服务结果：服务后设备恢复正常。 | [12 empty] | 118924795 | Closed | 1635248 |  |  | Opened - Pending Translation | Opened - Pending Translation | CTD00006847 | [5 empty] | Service | [6 empty] | Video cable | [5 empty] | IB Exist |  |  |  | System | CIRS-SW | ACQ_NOT_OK | CIRS-SW-ACQ_NOT_OK | CIRS-SW-ACQ_NOT_OK | CTD00006847 | ACQ_NOT_OK | =BE1656 | CIRS-SW-ACQ_NOT_OK | bar: No
1655 |  |  | China | HEBEI Tangshan Staff Hosp. | 40044 | 4455000% |  | System can not scan. Err code ACQ NOT OK. |  | CFM | [14 empty] | CFM |  | 1635224 |  |  | Opened | Opened | CTUDT00630408 | [5 empty] | Engineering | [6 empty] | ACQ NOT OK | [5 empty] | Refresh |  |  |  | System | CIRS-SW | ACQ_NOT_OK | CIRS-SW-ACQ_NOT_OK | CIRS-SW-ACQ_NOT_OK | CTUDT00630408 | ACQ_NOT_OK | CIRS-SW | CIRS-SW-ACQ_NOT_OK | bar: No
1656 |  |  | China | 扎兰屯市人民医院 | 40045 | 4454900% |  | 02N/Ingenuity CT/cannot scan |  | 2 | [14 empty] | 118924630 | In Process | 1635172 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | Engineering | [6 empty] | Gantry initialization | [5 empty] | IB Exist | bar: No
1657 |  |  | China | 内蒙古兴安盟人民医院 | 345030 | 4454800% |  | 02N/Ingenuity Core128/Elite China/ring artifact |  | 3 |  | Image artifact, M12 replaced. | [12 empty] | 118923872 | Closed | 1634692 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | Engineering | [6 empty] | Artifact | [5 empty] | IB Exist | 459800911281 | TDMS V2 M.Det Class B FRU | bar: No
1658 |  |  | China | 河南省新乡市凤泉区大块镇卫生院 | 345009 | 4454700% |  | 05K for IB/Ingenuity Core128/Elite China/Customer consultation | [16 empty] | 118918764 | Closed | 1633115 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | Customer | [6 empty] | Consult for Service Contract and report for the workstation issue | [5 empty] | Non-refresh | bar: No
1659 |  |  | Puerto Rico | DORADO HEALTH INC DBA | 40007 | 4454700% |  | It throws an error that asks to delete data since there is not enough space but there is only 5% occupation |  |  |  | Customer store the patient data to local hard disk | [12 empty] | 118921581 | Closed | 1631997 |  |  | Opened | Opened | CTUDT00632261 | [5 empty] | Engineering | [6 empty] | Disk full | [5 empty] | Refresh |  |  |  | System | Console-SW | DiskFull | Console-SW-DiskFull | Console-SW-DiskFull | CTUDT00632261 | DiskFull | Console-SW | Console-SW-DiskFull | bar: No
1660 |  |  | Japan | 独立行政法人国立病院機構 | 341015 | 4454700% |  | 撮影したデータが短くExtendしました。シンスライスも結合されていますが、MPRを作成するためにCTViewerで開くと片方しか表示されません。 |  | 5 |  | Start up the CView and open the siries tab and split into two. If you select two, right-click, Marge will appear, so choose and combine | [12 empty] | 118919064 | Closed | 1629595 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Application | [6 empty] | Clinical consult: marge combine | [5 empty] | Refresh | bar: No
1661 |  |  | United Kingdom | STEPPING HILL HOSPITAL | 40016 | 4454500% |  | at 13:30pm on 14/12/21 during a diagnostic colonoscopy the scanner reported at reconstruction error Customer did not try an off-line recon nor did they do a bugreport as reported that they were very busy The customer rescanned the patientok but resulting in an addition DLP and dose Since the incident the scanner has been in use without any reported errors after FSE Warren Thorpe spoke with customer a bugreport was generated today at 12:15 on 15/12/21 FOR UKI REFERENCE ONLY: Feedback type: potential incident UKI KMO: LisaHollis Submitter email: doug.anderson_2@philips.com |  | CFM | [14 empty] | CFM |  | 1625818 |  |  | Investigation | Investigation | [6 empty] | TBD | [6 empty] | Recon error | [5 empty] | Refresh | bar: No
1662 |  |  | Japan | Maki Hospital | 341017 | 4454600% |  | Quick viewerを開こうとしたらBugrep作成するようポップアップが出た。>>Host再起動後正常動作。ハード的なエラーは無し。 |  | 4 |  | Restart | [12 empty] | 118905403 | Closed | 1624861 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Popup create Bugreport request, after click the quick view. | [5 empty] | Refresh |  |  |  | System | Console-SW | NoEnoughInfo | Console-SW-NoEnoughInfo | Console-SW-NoEnoughInfo | NoEnoughInfo | NoEnoughInfo | Console-SW | Console-SW-NoEnoughInfo | bar: No
1663 |  |  | China | 息烽县中医医院 | 345014 | 4454500% |  | WeChat：05K for IB/Ingenuity Core128/Elite China/Mat breakage |  | Zhai, Lisa | [14 empty] | 118899967 | Closed | 1624597 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Mat breakage | [5 empty] | IB Exist | bar: No
1664 |  |  | China | 河北省唐山市工人医院 | 40044 | 4454500% |  | 02N/Ingenuity CT 728326/ACQ issue |  |  |  | Reload CIRS software, re-plug the ACQ board | [12 empty] | 118898185 | In Process | 1623633 |  |  | Opened - Pending Translation | Opened - Pending Translation | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | ACQ error | [5 empty] | Refresh |  |  |  | System | Console-HW | NoEnoughInfo | Console-HW-NoEnoughInfo | Console-HW-NoEnoughInfo | NoEnoughInfo | NoEnoughInfo | Console-HW | Console-HW-NoEnoughInfo | bar: No
1665 |  |  | United Kingdom | STOCKPORT NHS FOUNDATION TRUST | 40016 | 4454400% |  | came up with system error during reconstruction - occurred at 13:30 |  | 4 | [14 empty] | 118895711 | In Progress | 1621949 |  |  | Investigation | Investigation | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | reconstruction failed during a colonoscopy | [5 empty] | Refresh |  |  |  | System | CIRS-SW | NoEnoughInfo | CIRS-SW-NoEnoughInfo | CIRS-SW-NoEnoughInfo | NoEnoughInfo | NoEnoughInfo | CIRS-SW | CIRS-SW-NoEnoughInfo | bar: No
1666 |  |  | China | Lindian County Hospital | 345038 | 4454400% |  | 02N/Ingenuity Core128/Elite/yellow China |  | 5 |  | Reload Host software | [12 empty] | 118892511 | In Process | 1619715 |  |  | Opened | Opened | CTUDT00632520 | [5 empty] | Engineering | [6 empty] | Console stuck and yellow screen after scan complete. | [5 empty] | Refresh |  |  |  | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow-Screen | Console-SW | Console-SW-Yellow-Screen | bar: No
1667 |  |  | Japan | Maki Hospital | 341017 | 4454400% |  | ガントリー回転時に異音が発生する |  | 5 |  | Replace the main belt | [12 empty] | 118891082 | Closed | 1618885 |  |  | Evaluation | Evaluation | [6 empty] | Engineering | [6 empty] | Noise when gantry rotation | [5 empty] | IB Exist | 453566492401 | POLY-VEE BELT | bar: No
1668 |  |  | Japan | Tomakomai City Hospital | 341014 | 4454400% |  | 画面がワンモニターになってしまった （The screen has become one monitor ） |  | Zhu, Chun |  | One of the DisplayPort cabinets on the console side was missing from the signal. Cable replacement | [12 empty] | 118897534 | In Progress | 1618706 |  |  | Evaluation - Pending Translation | Evaluation - Pending Translation | ReplaceCable | [5 empty] | Engineering | [6 empty] | 1 Video signal missing | [5 empty] | Refresh |  | Cable |  | System | Console-HW | ReplaceCable | Console-HW-ReplaceCable | Console-HW-ReplaceCable | ReplaceCable | ReplaceCable | Console-HW | Console-HW-ReplaceCable | bar: No
1669 |  |  | United States | Lackey Memorial Hospital | 40015 | 4454400% |  | Gantry wont turn on. |  | 2 | [14 empty] | 118892135 | Fixed | 1616223 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Gantry cannot initialize | [5 empty] | IB Exist | 453567017751 | RHOST SG164 ASSY | bar: No
1670 |  |  | Portugal | Centro Hospitalar de Vila Nova | 40030 | 4454300% |  | O TC do SU aquando a programação de duas aquisições sequenciadas arterial e venosa, dá erro na altura de fazer a aquisição arterial e é abortada pelo (The SU CT when programming two sequential acquisitions in blood and vein, gives error at the time of making the arterial acquisition and is aborted by the equipment) equipamento. ao fazer ok ao erro e retirando a contagem do tempo é possível fazer as aquisições. Problema tem vindo a ser frequente radiologia Convencional - U1 SN: 40030 PO: OT 2021120349 | [16 empty] | 118886493 | Closed | 1615521 |  |  | Evaluation | Evaluation | [6 empty] | TBD | [6 empty] | System error inbetween two acquisition series | [5 empty] | TBD | bar: No
1671 |  |  | China | Hebei Province Tangshan City Workers Hospital | 40044 | 4454300% |  | WeChat：02N/Ingenuity CT 728326/report error |  | Cai, Cathy |  | Replace GPU | [12 empty] | 118884797 | Closed | 1614437 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | Replace-GPU | [5 empty] | Engineering | [6 empty] | System rebuild slowly-GPU | [5 empty] | Refresh | 459801622381 | NVidia P2000 GPU Card |  | System | Console-HW | GPU | Console-HW-GPU | Console-HW-GPU | Replace-GPU | GPU | Console-HW | Console-HW-GPU | bar: No
1672 |  |  | United States | Philips Medical Capital | 300057 | 4453300% |  | Fan is noisy, replaced 453566503142,Fan, Tube Axial | PR1584749 investigates the ring artifact issue, PR1644755 investigates the noisy fan issue. Ring artifact. #40 replacement. 459800594811,TDMS Module Class A | Kanzeg, Mike | [14 empty] | 118836197 | Closed | 1644755 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | Fan noise | [5 empty] | IB Exist | 453566503142 | Fan, Tube Axial | bar: No
1673 |  |  | Mexico | DE LAGE LANDEN SA DE CV | 10007 | 4454900% |  | EL EQUIPO ESTA BLOQUEADO. | Work Order #: WO-05015036 [REDACTED] The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled. Problem Description by Engineer : Problem description by engineer : The system opens the e-Stop. When was the first occurrence? Within... : December 18 2021, 6:00:00 PM How often is issue occurring? : Intermittent Malfunction area : Table Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: The gantry shuts down. Patient Impact/Outcome: Scan is aborted. User Impact: Scans cannot be done and starting device is hard. Current Software Version: 4.1.10 Diagnostic performed by Engineer: The system is checked because it has been failing related to aborted acquisitions or e-Stop opening and being unable to be restored until several full resets. Bug report is checked remotely and communication failure logs are found in canbus because of the table controller. Case is derived to Planning area with some suggestions for parts for visit. Internal Comments : Error # and/or description of error : S_LIFESYNCMGRTASK_LIFESYNC_FAIL P1=8 CRC RESENDS Repair Action: Resends were also found in channels A and C mainly. Check. Follow Up Required: Error Number and/or Description: S_LIFESYNCMGRTASK_LIFESYNC_FAIL P1=8 CRC RESENDS Repair Action: Resends were also found in channels A and C mainly. Check. What are the next Troubleshooting steps : DC power supply and DC-Couch Control board cable proactive replacement. Check data path to check resends detected mainly in channel A and C. Technician's Name: [REDACTED] ----------------------------------------------------------------------------- Work Order #: WO-05022597 [REDACTED] Problem Description by Engineer : Problem description by engineer : The system opens the e-Stop. When was the first occurrence? Within... : December 18 2021, 6:00:00 PM How often is issue occurring? : Intermittent Malfunction area : Table Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: The gantry shuts down. Patient Impact/Outcome: Scan is aborted. User Impact: Scans cannot be done and starting device is hard. Current Software Version: 4.1.10 Diagnostic performed by Engineer: N/A ----------------------------------------------------------------------------- | 2 | [14 empty] | 118925159 | Fixed | 1635717 |  |  | Opened | Opened | [6 empty] | Engineering | [6 empty] | E-stop open | [5 empty] | IB Exist | 453566503742 and 453567133673 | ASSY, 4 OUTPUT DC SUPPLY and CBL, DC POWER | bar: No
1674 |  |  | United States | Upson Regional Medical Center | 310115 | 4454600% |  | System error during scan, e-stop opened - Replaced faulty 24V couch relay | Work Order #: WO-05005844 18-Dec-2021 Davis Paul The system meets the specification for the performed service and is returned to use Resolution : 24V couch relay is faulty. Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: Diagnostic Expected and Unexpected Results: System Error Please recover sequence error User Impact: No Patient Impact: No Current Software Version: 4.1.10.25220 Parts Consumed : 1.00000,453566496941,RELAY,24VDC PLUGIN W/COVER 1.00000,453566496941,RELAY,24VDC PLUGIN W/COVER ----------------------------------------------------------------------------- Work Order #: WO-05000975 18-Dec-2021 Davis Paul The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: Diagnostic Expected and Unexpected Results: System Error Please recover sequence error User Impact: No Patient Impact: No Current Software Version: 4.1.10.25220 Resolution : 24V couch relay faulty. Follow Up Required : Ran bug report. Error not recurring. ----------------------------------------------------------------------------- Work Order #: WO-05000882 Maier, Donnie The system does not meet the specification for the performed service and is not returned to use Follow Up Required : Customer states system is giving multiple errors. Most recent was a system error during scan please reconstruct. Customer states they have already power cycled completely down to wall breaker multiple times. She believes they were having other errors as well last night as far as couch free floating as if estop opened. Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: Diagnostic Expected and Unexpected Results: System Error Please recover sequence error User Impact: No Patient Impact: No Current Software Version: 4.1.10.25220 Diagnostic performed by Engineer : Customer states system is giving multiple errors. Most recent was a system error during scan please reconstruct. Customer states they have already power cycled completely down to wall breaker multiple times. She believes they were having other errors as well last night as far as couch free floating as if estop opened. I will try to pull logs, analyze and add case notes if possible. ----------------------------------------------------------------------------- | Mulhern, Melissa |  | Replace the couch 24V relay | [12 empty] | 118914162 | Closed | 1626503 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Gantry wouldn't initialize Couch Free Floating | [5 empty] | IB Exist | 453566496941 | RELAY, 24VDC PLUG-IN W/COVER | bar: No
1675 |  |  | United States | Christ Hospital | 300205 | 4454400% |  | 0903CFXXX parts ID request | Internal Remarks : Interaction Type: New Case Caller: Corey Claxton IP & Issue: 65842604 & parts id// Event Type: Incident Service Request: Remote Technical Did Customer Decline Remote Support? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: created case and WO and routed customer to tech support Next Step Action: N/A Verbal PO Number: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- Work Order #: WO-04982116 Morrow, Dan The system meets the specification for the performed service and is returned to use Resolution : Customer wants part number and part ordered for class A V2 module. 459800911271 - TDMS V2 M.Det Class A FRU. NEMO ordered part for customer. Internal Comments : Error # and/or description of error : No error. Repair Action : Customer wants part number and part ordered for class A V2 module. 459800911271 - TDMS V2 M.Det Class A FRU. NEMO ordered part for customer. Diagnostic performed by Engineer : Customer wants part number and part ordered for class A V2 module. 459800911271 - TDMS V2 M.Det Class A FRU. NEMO ordered part for customer. Follow Up Required : Please ship requested parts to the following location for Next Day Arrival (NDA): The Christ Hospital-Ft Wright Attention: CT/Cory Claxton 1955 Dixie Hwy FT WRIGHT KY 41011-2792 Problem Description by Engineer : Problem description by engineer : Parts ID request When was the first occurrence? Within... : 14-Dec-2021 How often is issue occurring? : First occurrence Malfunction area : Detector Sys Information to support the complaint handling process Customer Function/Role: BioMed Device Use: No patient involved Expected and Unexpected results: Part ID Patient Impact/Outcome: none User Impact: Artifact Current Software Version: Not provided ----------------------------------------------------------------------------- Work Order #: WO-04982201 ----------------------------------------------------------------------------- | Shai, Gal | [14 empty] | 118897024 | In Progress | 1618436 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Customer order the part | [5 empty] | IB Exist | 459800911271 | TDMS V2 M.Det Class A FRU | bar: No
1676 |  |  | United States | Christ Hospital | 300205 | 4454400% |  | 65842604 RING ARTIFACTS | Internal Remarks : **CUST REQUESTED TECH SUPPORT Interaction Type: New Case Caller: CORI CLAXTON IP & Issue: 65842604 RING ARTIFACTS Event Type: Incident Service Request: Remote Technical Did Customer Decline Remote Support? No Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: CREATED CASE AND REMOTE WO; RTAC Next Step Action: N/A Verbal PO Number: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- Work Order #: WO-05003253 22-Dec-2021 Hanlin Christopher The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: Technician Device Use: during normal scanning Expected and Unexpected Results: Customer reports artifacts in image; not quite a complete circle at times. User Impact: Affecting image quality, but scans are diagnostic. Customer is continuing scanning Patient Impact: NA except for potential image quality issue. Current Software Version: Resolution : Unable to replace detector module in question ; part was not ordered. There is no artifact in images, but module is scheduled for replacement due to noise issue found during calibrations. Customer will order module and reschedule replacement at their convenience. ----------------------------------------------------------------------------- Work Order #: WO-04981237 16-Dec-2021 Hanlin Christopher The system meets the specification for the performed service and is returned to use Tools & Equipment : MULTIMETER,2022-06-30 00:00:00 Problem Description by Engineer : Customer Function/Role: Technician Device Use: during normal scanning Expected and Unexpected Results: Customer reports artifacts in image; not quite a complete circle at times. User Impact: Affecting image quality, but scans are diagnostic. Customer is continuing scanning Patient Impact: NA except for potential image quality issue. Current Software Version: Follow Up Required : Replace detector module 24 due to marginal noise spikes in module. Resolution : Ran Ring ID to locate potential artifact; verified bad detector elements and module by running Bad Detector diagnostics. Replaced detector module and recalibrated system per recommendation. ----------------------------------------------------------------------------- Work Order #: WO-04981056 Powell, Jamie The system does not meet the specification for the performed service and is not returned to use Diagnostic performed by Engineer : customer states that they're having ring artifacts during scans.. Customer ran air cals bust still having the issue.. No contrast that customer knows of. No ftp access to download logs. Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Follow Up Required : No ftp access to download logs. customer ran air cals but still having issues. Check for contrast. Run bad detector test. Run fee test. Internal Comments : No ftp access to download logs. customer ran air cals but still having issues. Check for contrast. Run bad detector test. Run fee test. ----------------------------------------------------------------------------- | Shai, Gal |  | Replace module 24 | [12 empty] | 118896346 | Closed | 1617837 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Artifact | [5 empty] | IB Exist | No infor | bar: No
1677 |  |  | United States | Select Medical Corporation | 32261 | 4453600% |  | reference account is currently locked | Internal Remarks : Interaction Type: New Case Caller: Jesse Nieves IP & Issue: 7688812 currently locked out of system Event Type: Incident Service Request: Remote Technical Did Customer Decline Remote Support? No Preferred Service Time: N/A Entitlement Options: Quote Reason for Entitlement Claim: N/A Action Taken: Created case and quote Next Step Action: N/A Verbal PO Number: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | [15 empty] | 118843629 | New | 1593453 |  |  | Evaluation | Evaluation | [6 empty] | Customer | [6 empty] | Customer account lock | [5 empty] | Refresh | bar: No
1678 |  |  | Australia | CITISCAN RADIOLOGY PTY LTD | 10491 | 4453600% |  | Gantry turned of using the UNLOAD option | Work Order #: WO-04935704 17-Dec-2021 Richard Wade The system meets the specification for the performed service and is returned to use Problem Description by Engineer : CUSTOMER REPORTED PROBLEM OR EVENT DESCRIPTION: 6 Dec 2021 The Customer called to report on behalf of a colleague that the Gantry turned off using the UNLOAD option. The issue occurred at approx. 10:05 AM and was resolved by turning the gantry back on. Resolution : Performed patient support cable checks - all ok. Performed numerous load/ unload motions all without error. Checked and tested system ok. ----------------------------------------------------------------------------- Work Order #: WO-04912325 Steve, Teixeira The system meets the specification for the performed service and is returned to use Problem Description by Engineer : CUSTOMER REPORTED PROBLEM OR EVENT DESCRIPTION: 6 Dec 2021 The Customer called to report on behalf of a colleague that the Gantry turned off using the UNLOAD option. The issue occurred at approx. 10:05 AM and was resolved by turning the gantry back on. Diagnostic performed by Engineer : ________________________________________________________________ CUSTOMER REPORTED PROBLEM OR EVENT DESCRIPTION: 6 Dec 2021 The Customer called to report on behalf of a colleague that the Gantry turned off using the UNLOAD option. The issue occurred at approx. 10:05 AM and was resolved by turning the gantry back on. SOFTWARE REVISION OF THE SYSTEM: Ver. 4.1.10 WAS A BUG REPORT SAVED? No TROUBLESHOOTING - 6 Dec 2021: Informed the Customer that would download and review the logs. Logs confirm that the UNLOAD button was pressed at 09:46:17.145300 AM 6 Dec 2021 and released at 09:46:24.975473 AM. The logs indicate that the Gantry turned off shortly after at 09:46:26.996509 AM due to ESTOP error "S_COUCHMGR_HOR_MOTION_STOP_ERROR " at 09:46:27.183 AM followed by Normal type Error "S_MANMOT_HORIZONTAL_STOP_ERROR" at 09:46:30.087 AM. The errors appear to indicate that the Couch did not stop in time. Informed the Customer of the findings. The Customer confirmed with the colleague unloading the patient, that the patient was not moving unexpectedly at the time. CUSTOMER UPDATE: 9 Dec 2021 Sent an email update to the Customer advising of the action plan. Email extract reads: "Hi XXXX, Regarding the case logged on the 6th Dec 2021 for the gantry turning off whilst using the unload – Case 0118839759. As discussed, a review of the logs shows the User released the UNLOAD button just before the bed reached the out limit, it then appears that the bed although requested to stop continued to move slightly which was not expected, and as consequence turned the gantry off due to an unexpected movement. A Planned Maintenance service is scheduled for Thursday the 16th at 12:30 PM, so the attending Engineer as part of the service will inspect the integrity of the Horizontal couch assemblies and action where required. Please feel free to contact me should you have any further queries in this regard. ..." ONSITE WORK ORDER: Please refer to the onsite work order WO-04935704 for further details of the actions taken to address this issue. ________________________________________________________________ Internal Comments : LOG REVIEW: *********** ________________________________________________________________ LOGGER: * 2021/12/06 10:37:38.038 EVENT_MINOR GSE GSE::CState::StateBkst::STUDY_ENDED was sent to GPC 2021/12/06 10:35:39.042 EVENT_MAJOR GSE ----------------- Series 12 Complete ----------------- 2021/12/06 10:29:46.870 EVENT_MAJOR GSE ----------------- Series 11 Complete ----------------- 2021/12/06 10:29:20.504 EVENT_MAJOR GSE ----------------- Series 10 Complete ----------------- 2021/12/06 10:28:40.889 EVENT_MINOR GSE GSE::CState::StateBkst::STUDY_STARTED was sent to GPC 2021/12/06 09:47:18.789 ERROR_WARNING Gantry CStateConfigured 3 ERROR 522593 WARN S_HMC_COUCH_24V_IS_ON 3 1073938000 2021/12/06 09:47:18.740__02 EVENT_MINOR GSE GantryEstopClosedBkst in CStateConfigured 2021/12/06 09:47:18.646 ERROR_WARNING Gantry CStateConfigured 0 ERROR 522593 WARN S_MPS_WAITING_FOR_MAIN_DRIVE_TO_CLOSE 0 0 MP: Waiting for MDC to close.. 2021/12/06 09:46:30.087 ERROR_NORMAL Gantry CStateConfigured 0 ERROR 522544 NORMAL S_MANMOT_HORIZONTAL_STOP_ERROR 0 0 2021/12/06 09:46:27.183 ERROR_FATAL Gantry CStateConfigured 3 ERROR 522541 ESTOP S_COUCHMGR_HOR_MOTION_STOP_ERROR 0 0 2021/12/06 09:46:27.039 ERROR_FATAL Gantry GantryEstopOpenedBkst(0) in CStateSeriesNotLoaded 2021/12/06 09:46:09.711 EVENT_MINOR GSE GSE::CState::StateBkst::STUDY_ENDED was sent to GPC 2021/12/06 09:45:27.792 EVENT_MAJOR GSE ----------------- Series 9 Complete ----------------- * ________________________________________________________________ CANTRACE: * 2021/12/06 09:46:17.145300 [PANEL] PANEL_UNLOAD_PANEL_REQ [PRESS] 2021/12/06 09:46:17.201735 COUCH_START_HOR_JOG_REQ 2021/12/06 09:46:17.259700 COUCH_REPORT_HOR_POS_BKST 5136 1(OUT) 0 0 2021/12/06 09:46:17.512603 COUCH_REPORT_HOR_POS_BKST 5193 1(OUT) 4276 0 2021/12/06 09:46:17.763763 COUCH_REPORT_HOR_POS_BKST 5376 1(OUT) 9269 0 2021/12/06 09:46:18.015409 COUCH_REPORT_HOR_POS_BKST 5683 1(OUT) 14245 0 2021/12/06 09:46:18.265495 COUCH_REPORT_HOR_POS_BKST 6111 1(OUT) 18433 0 2021/12/06 09:46:18.516199 COUCH_REPORT_HOR_POS_BKST 6575 1(OUT) 18539 0 2021/12/06 09:46:18.768578 COUCH_REPORT_HOR_POS_BKST 7041 1(OUT) 18488 0 2021/12/06 09:46:19.019183 COUCH_REPORT_HOR_POS_BKST 7505 1(OUT) 18506 0 2021/12/06 09:46:19.269510 COUCH_REPORT_HOR_POS_BKST 7968 1(OUT) 18510 0 2021/12/06 09:46:19.520522 COUCH_REPORT_HOR_POS_BKST 8432 1(OUT) 18499 0 2021/12/06 09:46:19.772574 COUCH_REPORT_HOR_POS_BKST 8898 1(OUT) 18491 0 2021/12/06 09:46:20.022802 COUCH_REPORT_HOR_POS_BKST 9362 1(OUT) 18476 0 2021/12/06 09:46:20.274589 COUCH_REPORT_HOR_POS_BKST 9826 1(OUT) 18493 0 2021/12/06 09:46:20.525580 COUCH_REPORT_HOR_POS_BKST 10292 1(OUT) 18493 0 2021/12/06 09:46:20.777598 COUCH_REPORT_HOR_POS_BKST 10757 1(OUT) 18499 0 2021/12/06 09:46:21.028266 COUCH_REPORT_HOR_POS_BKST 11222 1(OUT) 18495 0 2021/12/06 09:46:21.278869 COUCH_REPORT_HOR_POS_BKST 11686 1(OUT) 18494 0 2021/12/06 09:46:21.530088 COUCH_REPORT_HOR_POS_BKST 12150 1(OUT) 18492 0 2021/12/06 09:46:21.781585 COUCH_REPORT_HOR_POS_BKST 12614 1(OUT) 18498 0 2021/12/06 09:46:22.032298 COUCH_REPORT_HOR_POS_BKST 13080 1(OUT) 18551 0 2021/12/06 09:46:22.283489 COUCH_REPORT_HOR_POS_BKST 13544 1(OUT) 18497 0 2021/12/06 09:46:22.534097 COUCH_REPORT_HOR_POS_BKST 14008 1(OUT) 18508 0 2021/12/06 09:46:22.786107 COUCH_REPORT_HOR_POS_BKST 14473 1(OUT) 18501 0 2021/12/06 09:46:23.036215 COUCH_REPORT_HOR_POS_BKST 14937 1(OUT) 18516 0 2021/12/06 09:46:23.287252 COUCH_REPORT_HOR_POS_BKST 15400 1(OUT) 18495 0 2021/12/06 09:46:23.538756 COUCH_REPORT_HOR_POS_BKST 15865 1(OUT) 18480 0 2021/12/06 09:46:23.789504 COUCH_REPORT_HOR_POS_BKST 16330 1(OUT) 18472 0 2021/12/06 09:46:24.040424 COUCH_REPORT_HOR_POS_BKST 16794 1(OUT) 18493 0 2021/12/06 09:46:24.291803 COUCH_REPORT_HOR_POS_BKST 17258 1(OUT) 18508 0 2021/12/06 09:46:24.543007 COUCH_REPORT_HOR_POS_BKST 17724 1(OUT) 18490 0 2021/12/06 09:46:24.792994 COUCH_REPORT_HOR_POS_BKST 18184 1(OUT) 17834 129 2021/12/06 09:46:24.975473 [PANEL] PANEL_UNLOAD_PANEL_REQ [RELEASE] 2021/12/06 09:46:24.984781 COUCH_STOP_HOR_JOG_REQ 2021/12/06 09:46:25.002961 COUCH_STOP_VERT_JOG_REQ 2021/12/06 09:46:25.044888 COUCH_REPORT_HOR_POS_BKST 18561 1(OUT) 12924 129 2021/12/06 09:46:25.296010 COUCH_REPORT_HOR_POS_BKST 18811 1(OUT) 7962 129 2021/12/06 09:46:25.547028 COUCH_REPORT_HOR_POS_BKST 18936 1(OUT) 2988 129 2021/12/06 09:46:25.798474 COUCH_REPORT_HOR_POS_BKST 18952 1(OUT) 3 133 2021/12/06 09:46:26.049129 COUCH_REPORT_HOR_POS_BKST 18952 1(OUT) 3 133 2021/12/06 09:46:26.300214 COUCH_REPORT_HOR_POS_BKST 18952 1(OUT) 6 133 2021/12/06 09:46:26.551292 COUCH_REPORT_HOR_POS_BKST 18952 1(OUT) 3 133 2021/12/06 09:46:26.802750 COUCH_REPORT_HOR_POS_BKST 18952 1(OUT) 9 133 2021/12/06 09:46:26.996509 **RHOST_ESTOP_OPENED_BKST** 2021/12/06 09:46:26.996799 **GHOST_ESTOP_OPENED_BKST** 2021/12/06 09:46:26.997096 **DMC_ESTOP_OPENED_BKST** 2021/12/06 09:46:26.998104 **GANTRY_ESTOP_OPENED_BKST** 2021/12/06 09:46:27.006350 **COUCH_ESTOP_OPENED_BKST** 2021/12/06 09:46:27.008788 COUCH_REPORT_HOR_POS_BKST 18952 1(OUT) 9 389 2021/12/06 09:46:27.023419 COUCH_REPORT_VERT_POS_BKST 1066 1(DOWN) 1 2021/12/06 09:46:27.163575 ==COUCH_ESTOP_ERROR_BKST== S_COUCHMGR_HOR_MOTION_STOP_ERROR 2021/12/06 09:46:27.176639 ==GANTRY_ESTOP_ERROR_BKST== S_COUCHMGR_HOR_MOTION_STOP_ERROR 2021/12/06 09:46:30.051295 <<NORMAL_ERROR>> S_MANMOT_HORIZONTAL_STOP_ERROR 2021/12/06 09:46:30.064170 <<NORMAL_ERROR>> S_MANMOT_HORIZONTAL_STOP_ERROR 2021/12/06 09:46:30.350342 <<GANTRY_EAL_REPORT_ANS>> GHOST EAL_NORMAL S_MANMOT_HORIZONTAL_STOP_ERROR 2021/12/06 09:47:06.703867 COUCH_REPORT_HOR_POS_BKST 18951 0(IN) 0 260 2021/12/06 09:47:06.955585 COUCH_REPORT_HOR_POS_BKST 18948 0(IN) 0 260 2021/12/06 09:47:07.207334 COUCH_REPORT_HOR_POS_BKST 18779 0(IN) 0 260 ________________________________________________________________ Follow Up Required : CUSTOMER REPORTED PROBLEM OR EVENT DESCRIPTION: 6 Dec 2021 The Customer called to report on behalf of a colleague that the Gantry turned off using the UNLOAD option. The issue occurred at approx. 10:05 AM and was resolved by turning the gantry back on. TROUBLESHOOTING - 6 Dec 2021: Informed the Customer that would download and review the logs. Logs confirm that the UNLOAD button was pressed at 09:46:17.145300 AM 6 Dec 2021 and released at 09:46:24.975473 AM. The logs indicate that the Gantry turned off shortly after at 09:46:26.996509 AM due to ESTOP error "S_COUCHMGR_HOR_MOTION_STOP_ERROR " at 09:46:27.183 AM followed by Normal type Error "S_MANMOT_HORIZONTAL_STOP_ERROR" at 09:46:30.087 AM. The errors appear to indicate that the Couch did not stop in time. Informed the Customer of the findings. The Customer confirmed with the colleague unloading the patient, that the patient was not moving unexpectedly at the time. NEXT STEP ACTION: An onsite work order is to be dispatched to coincide with PM work order WO-02519912 at Citiscan at 12:30 16th Dec 2021. Requested the attending Engineer to please inspect the couch mechanics in terms of horizontal travel as mentioned in emails. Included the AEL recommendations associated with error S_COUCHMGR_HOR_MOTION_STOP_ERROR that may help the Engineer focus on the areas of concern. The Customer has been made aware of the action plan via email. ----------------------------------------------------------------------------- | Howard, Amanda |  | Performed patient support cable checks - all ok | [12 empty] | 118839759 | Closed | 1588236 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Gantry off when upload | [5 empty] | IB Exist | bar: No
1679 |  |  | United States | Philips Medical Capital | 300057 | 4453300% |  | Ring artifacts appear in images, replaced 459800594811 TDMS Module Class A FRU | Internal Remarks : approved: onsite is covered by risk sharing model. Internal Remarks : Interaction New Case Caller: John Williams IP & Issue: 54851181 FSE callback for NSS Event Type: Incident Service Request: Onsite Did Customer Decline Remote Support? Yes Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: warm transferred to NSS Next Step Action: N/A Verbal PO Number: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- Work Order #: WO-04927213 14-Dec-2021 Williams John The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Problem description by engineer : Ring artifacts on scans Information to support the complaint handling process Customer Function/Role: Not provided Device Use: Not provided Expected and Unexpected results: Not provided Patient Impact/Outcome: Not provided User Impact: Not provided Current Software Version: Not provided Parts Consumed : 1.00000,459800594811,TDMS Module Class A FRU Resolution : Replace detector module ----------------------------------------------------------------------------- Work Order #: WO-04910786 08-Dec-2021 Williams John The system meets the specification for the performed service and is returned to use Resolution : noisy fan Problem Description by Engineer : Problem description by engineer : Ring artifacts on scans Information to support the complaint handling process Customer Function/Role: Not provided Device Use: Not provided Expected and Unexpected results: Not provided Patient Impact/Outcome: Not provided User Impact: Not provided Current Software Version: Not provided Parts Consumed : 1.00000,453566503142,Fan, Tube Axial ----------------------------------------------------------------------------- Work Order #: WO-04910787 07-Dec-2021 Williams John The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Follow Up Required : Replace module 40 Problem Description by Engineer : Problem description by engineer : Ring artifacts on scans Information to support the complaint handling process Customer Function/Role: Not provided Device Use: Not provided Expected and Unexpected results: Not provided Patient Impact/Outcome: Not provided User Impact: Not provided Current Software Version: Not provided Diagnostic performed by Engineer : Bad detector module ----------------------------------------------------------------------------- Work Order #: WO-04907471 03-Dec-2021 Williams John The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Problem description by engineer : Ring artifacts on scans Information to support the complaint handling process Customer Function/Role: Not provided Device Use: Not provided Expected and Unexpected results: Not provided Patient Impact/Outcome: Not provided User Impact: Not provided Current Software Version: Not provided Follow Up Required : Replace module 37 Follow Up Required : Replace Noisy Fan Diagnostic performed by Engineer : Rings in image ----------------------------------------------------------------------------- Work Order #: WO-04907434 Powell, Jamie The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Internal Comments : Looked at image and used ring ID.. Looks like M6 or M37 is bad. Internal Comments : Error # and/or description of error : Ring artifacts on scans Repair Action : Ring artifacts on scans Follow Up Required : Error # and/or description of error : Ring artifacts on scans Repair Action : Ring artifacts on scans Looked at image and used ring ID.. Looks like M6 or M37 is bad. Technician Name : Powell, Jamie Problem Description by Engineer : Problem description by engineer : Ring artifacts on scans Information to support the complaint handling process Customer Function/Role: Not provided Device Use: Not provided Expected and Unexpected results: Not provided Patient Impact/Outcome: Not provided User Impact: Not provided Current Software Version: Not provided Diagnostic performed by Engineer : Ring artifacts on scans.. Customer states that the last air cal was ran on monday.. Suggested customer run another but she has scans coming up and will run one when she can. ----------------------------------------------------------------------------- | Kanzeg, Mike |  | Rplace module | [12 empty] | 118836197 | Closed | 1584749 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | Artifact | [5 empty] | IB Exist | 459800594811 | TDMS Module Class A FRU | bar: No
1680 |  |  | China | 阳高县医疗集团(阳高县人民医院 | 345025 | 4455600% |  | 05K For IB/Core128/Gantry noise |  | 3 |  | Replace the main belt | [12 empty] | 118957042 | Closed | 1654215 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | Engineering | [6 empty] | Gantry noise when rotation | [5 empty] | IB Exist | 453566492401 | POLY-VEE BELT | bar: No
1681 |  |  | China | 四川省宜宾市第二中医院 | 345035 | 4455600% |  | 02N/Ingenuity Core128/Elite China/couch vertical problem |  | 3 |  | Replace vertical transducer, Couch vertical re-calibration | [12 empty] | 118957019 | Closed | 1654142 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | Engineering | [6 empty] | S_Master_sequencer_vert_couch_pos_err | [5 empty] | IB Exist | 459800623331 | FRU, Vertical Transducer | bar: No
1682 |  |  | China | 四川省宜宾市第二中医院 | 345035 | 4455600% |  | 02N/Ingenuity Core128/Elite China/noise with rotation |  | 3 |  | Replae the main belt | [12 empty] | 118956787 | Closed | 1654084 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | Engineering | [6 empty] | Gantry noise when rotation | [5 empty] | IB Exist | 453566492401 | POLY-VEE BELT | bar: No
1683 |  |  | United States | Willis Knighton Hospital | 40004 | 4455600% |  | Local drive is full |  | 2 | [14 empty] | 118957431 | Fixed | 1653957 |  |  | Opened | Opened | CTUDT00632261 | [5 empty] | Engineering | [6 empty] | Disk full | [5 empty] | Refresh |  |  |  | System | Console-SW | DiskFull | Console-SW-DiskFull | Console-SW-DiskFull | CTUDT00632261 | DiskFull | Console-SW | Console-SW-DiskFull | bar: No
1684 |  |  | Portugal | Centro Hospitalar de Vila Nova | 40030 | 4455200% |  | TAC do SU está a dar erros de reconstrução. Está a trabalhar mas tem dado erros. (It's making reconstruction errors. ) Radiologia Convencional- U1 |  | 3 |  | Reconstructor software error when preparing the data receipt. Indication to the operator to restart the CIRS server. Trouble disappeared after CIRS restarted. There's no report of a new occurrence. No follow-up required. | [12 empty] | 118944241 | Closed | 1649021 |  |  | Opened - Pending Translation | Opened - Pending Translation | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | CIRS Reconstruction error | [5 empty] | Refresh |  |  |  | System | CIRS-SW | NoEnoughInfo | CIRS-SW-NoEnoughInfo | CIRS-SW-NoEnoughInfo | NoEnoughInfo | NoEnoughInfo | CIRS-SW | CIRS-SW-NoEnoughInfo | bar: No
1685 |  |  | Argentina | INSTITUTO NACIONAL DE | 40027 | 4455200% |  | NO SE PUEDE INICIAR EL CUADRO DE RECONSTRUCCIÓN (cirs). CAMBIE LOS PARÁMETROS DE EXPLORACIÓN |  | 2 | [14 empty] | 118945072 | Closed | 1648159 |  |  | Opened - Pending Translation | Opened - Pending Translation | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | CIRS rebuild table fail | [5 empty] | Refresh |  |  |  | System | CIRS-SW | NoEnoughInfo | CIRS-SW-NoEnoughInfo | CIRS-SW-NoEnoughInfo | NoEnoughInfo | NoEnoughInfo | CIRS-SW | CIRS-SW-NoEnoughInfo | bar: No
1686 |  |  | China | 莱阳市人民医院 | 345005 | 4455200% |  | WeChat:05L/Ingenuity Core128/Elite China/No heart rate display |  | 3 |  | instruct to customer the rate can see in scan | [12 empty] | 118941802 | Closed | 1647343 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | Customer | [6 empty] | no heart rate display | [5 empty] | Refresh | bar: No
1687 |  |  | Japan | 社会福祉法人恩賜財団済生会 | 343000 | 4455300% |  | Loadフットスイッチを操作した瞬間E-STOP OPENとなった |  | 4 WANG LIN, LV |  | Suspect cables block the Moog light | [12 empty] | 118946610 | Closed | 1646610 |  |  | Investigation - Pending Translation | Investigation - Pending Translation | [6 empty] | Engineering | [6 empty] | S_CTRL_LINK_DMC_GOT_CRC_FRAMING_ERROR | [5 empty] | Non-Refresh | bar: No
1688 |  |  | China | 河南省濮阳市人民医院 | 40031 | 4455200% |  | WeChat:05K for IB/Ingenuity CT 728326/Scan module damaged |  | 3 |  | Replace module | [12 empty] | 118940433 | Closed | 1646592 |  |  | Opened - Pending Translation | Opened - Pending Translation | CTUDT00632110 | [5 empty] | Engineering | [6 empty] | dms module | [5 empty] | Refresh | 459800911281 | TDMS V2.X M.Det Class B FRU |  | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | DMS Module | DMS-HW | DMS-HW-DMS Module | bar: No
1689 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 337000 | 4455200% |  | Masa hareketlerinde ses geliyor (Sounds on the desk movements) |  | 3 |  | When the mechanical parts of the table were checked, the sound was detected from the bed motion brake. Replacing brake cleanup sets the problem on the device | [12 empty] | 118941867 | Fixed | 1643701 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | Engineering | [6 empty] | Noise on couch vertical movement | [5 empty] | Non-Refresh | bar: No
1690 |  |  | Poland | Government Agency. | 341009 | 4455100% |  | Integracja TK z systemem HIS |  | 4 | [14 empty] | 118935452 | In Progress | 1643293 |  |  | Opened | Opened | [6 empty] | TBD | [6 empty] | TBD | [5 empty] | TBD | bar: No
1691 |  |  | Japan | Maki Hospital | 341017 | 4455100% |  | 本日午前中にパスワードがロックされた件に関して(With the password locked in the morning ) |  | 4 Yang, Cancan | [14 empty] | 118934845 | Closed | 1642194 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Customer | [6 empty] | Password locked | [5 empty] | Refresh | [6 empty] | - | - | 0 | 0 | 0 | - | bar: No
1692 |  |  | United States | Philips Medical Capital | 40053 | 4455000% |  | slow application when doing coronary scans. |  | 3 WANG LIN, LV | [14 empty] | 118932562 | In Progress | 1641742 |  |  | Investigation | Investigation | [6 empty] | TBD | [6 empty] | TBD | [5 empty] | TBD | bar: No
1693 |  |  | Puerto Rico | METRO MAYAGUEZ INC | 40009 | 4455100% |  | Table Joy stick does not work |  | 3 |  | The Joystick was not the problem. The problem was caused by the cable JOYSTICK PANEL MUT CONN. HARNESS. PN: 453567542042. | [12 empty] | 118937100 | Closed | 1641475 |  |  | Evaluation | Evaluation | [6 empty] | Engineering | [6 empty] | Joystick not work | [5 empty] | IB Exist | 453567542042 | CABLE, JOYSTICK PANEL MNT CONN. HARNESS | bar: No
1694 |  |  | Poland | Rzadowa Agencja | 341008 | 4455000% |  | Integracja TK z systemem HIS |  | 4 | [14 empty] | 118929865 | Closed | 1639653 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Other | [6 empty] | Other | [5 empty] | Non-Refresh | bar: No
1695 |  |  | France | CLINIQUE DE FLANDRE | 40019 | 4455100% |  | problème de mot de passe, quand il faut entrer le nouveau, ça a bloqué au bout de plusieurs essais. (Password problem. When a new password needs to be entered, it is blocked after multiple attempts ) |  | 3 | [14 empty] | 118934988 | Closed | 1639157 |  |  | Opened | Opened | [6 empty] | Customer | [6 empty] | Password locked | [5 empty] | Refresh | [6 empty] | - | - | 0 | 0 | 0 | - | bar: No
1696 |  |  | Turkey | ADIYAMAN OZEL SAGINSTURTAS SAN VE TIC A.S | 341006 | 4455100% |  | Mouse siparişi |  | 3 | [14 empty] | 118935314 | Closed | 1638834 |  |  | Evaluation | Evaluation | [6 empty] | Customer | [6 empty] | Customer order the part | [5 empty] | IB Exist | 459801378751 | White 1000 DPI 3 Button Mouse | bar: No
1697 |  |  | China | 江苏大学附属医院 | 345023 | 4455000% |  | PRS Connection/02N/Ingenuity Core128/Elite China |  | 5 Wang, Elaine |  | Add dual IP adress | [12 empty] | 118926310 | Closed | 1637996 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Service | [6 empty] | PRS not connect | [5 empty] | Refresh | [6 empty] | - | - | 0 | 0 | 0 | - | bar: No
1698 |  |  | Japan | 医療法人清翠会 牧病院 | 341017 | 4455000% |  | パスワードを間違えロックされてしまった。>>ユーザーitでログインいただきロック解除していただきました。(The password has been mistaken.) > > I locked it with the user it) |  | 4 Yang, Cancan | [14 empty] | 118932927 | Closed | 1637976 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Customer | [6 empty] | Password locked | [5 empty] | Refresh | [6 empty] | - | - | 0 | 0 | 0 | - | bar: No
1699 |  |  | China | DaKuai Town Health, Fengquan District, Xinxiang City, Henan Province | 345009 | 4455000% |  | 05K for IB/Ingenuity Core128/Elite China/consult information display |  | 3 Wang, Elaine |  | customer consult print information display issue, answer customer's question | [12 empty] | 118926167 | Closed | 1637881 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Customer consult printer setting | [5 empty] | Non-Refresh | bar: No
1700 | Date0103 |  | China | 陕西省咸阳中心医院 (咸阳市第二人民医院） | 40050 | 4456200% |  | wechat:05R/Ingenuity CT 728326 |  | 3 | [14 empty] | 118990920 | New | 1669486 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | TBD | bar: No
1701 |  |  | United States | Philips Medical Capital | 40055 | 4456100% |  | 0711CDN83Gantry Initialize Errors/Please Restart the Gantry |  | 2 | No bugreport yet | None | [12 empty] | 118990606 | Closed | 1669474 |  |  | Opened | Opened | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Intermittent gantry initial error | [5 empty] | Refresh |  |  |  | System | Gantry-SW | NoEnoughInfo | Gantry-SW-NoEnoughInfo | Gantry-SW-NoEnoughInfo | NoEnoughInfo | NoEnoughInfo | Gantry-SW | Gantry-SW-NoEnoughInfo | bar: No
1702 |  |  | United States | Philips Medical Capital | 40028 | 4456200% |  | 0505CFXXX Gantry not coming on |  | 2 |  | Recycle the E-stop button on CT BOX | [12 empty] | 118991931 | Closed | 1669309 |  |  | Evaluation | Evaluation | [6 empty] | Customer | [6 empty] | E-stop closing | [5 empty] | IB Existing | bar: No
1703 |  |  | France | CLINIQUE DE FLANDRE | 40019 | 4456100% |  | Le Host ne redémarre plus | Display failure when host restarts. After stopping the host on the "stop" button on the front panel, the host will not restart | 2 |  | FSE identified the display displaying issue when restar the host, no further information about what action done for the display. | [12 empty] | 118987426 | Closed | 1667428 |  |  | Opened | Opened | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Didplaying | [5 empty] | Refresh |  |  |  | System | Console-SW | NoEnoughInfo | Console-SW-NoEnoughInfo | Console-SW-NoEnoughInfo | NoEnoughInfo | NoEnoughInfo | Console-SW | Console-SW-NoEnoughInfo | bar: No
1704 |  |  | China | 宁安市第二人民医院 | 345039 | 4456000% |  | 02N/Ingenuity Core128/add the PACS |  | 5 |  | FSE to add the PACS as requested | [12 empty] | 118975961 | Closed | 1667101 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Add PACS | [5 empty] | Non-refresh | bar: No
1705 |  |  | United States | Philips Medical Capital | 40005 | 4455900% |  | Q/O Thumping and squeaking noise generated by drive belt |  | 5 |  | Replace the main belt and tensioner | [12 empty] | 118972284 | Closed | 1665316 |  |  | Evaluation | Evaluation | [6 empty] | Engineering | [6 empty] | Main belt noise | [5 empty] | IB Existing | 453567003451, 453566492401 | TENSIONER ASSY, POLY-VEE BELT | bar: No
1706 |  |  | China | 河南许昌市中心医院(传染病) | 40047 | 4455900% |  | 02N/Ingenuity CT 728326/disk space problem |  | 3 |  | Performed visual inspection ,found useless files, delete the useless file | [12 empty] | 118970175 | Closed | 1663990 |  |  | Opened - Pending Translation | Opened - Pending Translation | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Disk full | [5 empty] | Refresh |  |  |  | System | TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | bar: No
1707 |  |  | China | DaKuai Town Health, Fengquan District, Xinxiang City, Henan Province | 345009 | 4455900% |  | 02N/Ingenuity Core128/Elite China |  | 3 Wang, Elaine |  | N/A | [12 empty] | 118969785 | Closed | 1663825 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Customer | [6 empty] | 神飞云工作站 | [5 empty] | Non-refresh | bar: No
1708 |  |  | United States | Unity Health Systems | 40037 | 4455800% |  | MONTHLY QA |  | 5 Dantha, Rajashekar |  | N/A | [12 empty] | 118968008 | Closed | 1662882 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Monthly QA | [5 empty] | Non-refresh | bar: No
1709 |  |  | United States | Odessa Regional Medical Center | 40029 | 4455800% |  | infant EC questions |  | 4 WANG LIN, LV |  | advised customer to hand inject and scan with 5 sec delay. Follow up with customer. Scan went well, but she will scan immediately after injection next time | [12 empty] | 118966789 | Closed | 1662008 |  |  | Closed - Done | Closed - Done | [6 empty] | Customer | [6 empty] | Clinical consult of the injection delay for infant | [5 empty] | Non-refresh | bar: No
1710 |  |  | China | 苏州广慈肿瘤医院有限公司 | 345019 | 4455800% |  | 02N/Ingenuity Core128/Elite China/Quality Control |  | 5 |  | Do the image calibration | [12 empty] | 118965430 | Closed | 1661057 |  |  | Opened - Pending Translation | Opened - Pending Translation | Calibration | [5 empty] | Engineering | [6 empty] | Low contract resolution failure | [5 empty] | Refresh |  |  |  | System | NotEngineering | Calibration | NotEngineering-Calibration | NotEngineering-Calibration | Calibration | Calibration | NotEngineering | NotEngineering-Calibration | bar: No
1711 |  |  | China | Sichuan Province Yibin City No.2 T.C.M. Hospital | 345035 | 4455800% |  | 02N/Ingenuity Core128/Elite China |  | 5 Wang, Elaine |  | N/A | [12 empty] | 118964516 | Closed | 1660536 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Add the node | [5 empty] | Non-refresh | bar: No
1712 |  |  | China | Shaanxi Province Ankang T.C.M. Hospital | 345041 | 4455800% |  | WeChat:02N/Ingenuity Core128/Elite China/Scan fault |  | 3 WANG LIN, LV |  | delete patient data, and do full rebuild | [12 empty] | 118963411 | Closed | 1660062 |  |  | Evaluation | Evaluation | SystemStuck | [5 empty] | Engineering | [6 empty] | Scan fault, software stuck | [5 empty] | Refresh |  |  |  | System | Console-SW | SystemStuck | Console-SW-SystemStuck | Console-SW-SystemStuck | SystemStuck | SystemStuck | Console-SW | Console-SW-SystemStuck | bar: No
1713 | Date0110 |  | China | 内蒙古兴安盟人民医院 | 345030 | 4457000% |  | 05K for IB/Ingenuity Core128/Elite China/gantry initialize failed |  | 3 |  | On site check the scan parameter, change the scan parameter | [12 empty] | 119025216 | Closed | 1690655 |  |  | Opened - Pending Translation | Opened - Pending Translation | NotEngineering | [5 empty] | Engineering | [6 empty] | Gantry initialize failure | [5 empty] | Refresh |  |  |  | System | NotEngineering | ParameterSetting | NotEngineering-ParameterSetting | NotEngineering-ParameterSetting | NotEngineering | ParameterSetting | NotEngineering | NotEngineering-ParameterSetting | bar: No
1714 |  |  | Australia | Maia Financial Pty Limited | 341000 | 4457000% |  | Gantry Acquisition cannot comply failure, system is down. |  | 2 |  | Replace 453567538334,MODULE, ROTOR ELECTRONICS TX replacement | [12 empty] | 119026083 | Closed | 1690554 |  |  | Opened | Opened | [6 empty] | Engineering | [6 empty] | Gantry acquisition cannot comply | [5 empty] | IB Existing | 453567538334 | MODULE, ROTOR ELECTRONICS TX | bar: No
1715 |  |  | China | 哈尔滨市森海医院 | 345007 | 4456900% |  | 01/Ingenuity Core128/Elite China/password was locked |  | 3 |  | instruct customer wait more than 20 min to log in again. make a change password | [12 empty] | 119023787 | Closed | 1690098 |  |  | Opened - Pending Translation | Opened - Pending Translation | CTUDT00632236 | [5 empty] | Engineering | [6 empty] | Password locked | [5 empty] | Refresh |  |  |  | System | Console-SW | User Account | Console-SW-User Account | Console-SW-User Account | CTUDT00632236 | User Account | Console-SW | Console-SW-User Account | bar: No
1716 |  |  | Japan | 医療法人清翠会 牧病院 | 341017 | 4456900% |  | データ転送が途中で停止。Logout/Loginでは解消されずコンピュータの再起動で復旧した |  | 4 Zhu, Chun |  | Reboot the computer | [12 empty] | 119023916 | Closed | 1688848 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | Data transfer stop | [5 empty] | Refresh |  |  |  | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | bar: No
1717 |  |  | Germany | Gesundheit Nord gGmbH Klinikverbund | 40002 | 4456800% |  | FM: Fehlfunktion an der Temperaturregelung , bitte an FSE wenden (Error message with Air Calibration ) |  | 3 |  | FRU replaced | [12 empty] | 119019245 | Closed | 1688749 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | Engineering | [6 empty] | DMS Temperature | [5 empty] | IB Existing | 459800325761, 453567488921, 459800624062, | Brush, Signal Block Assembly; Power Supply Assy / Single; TDMS Cover RoHS FRU | bar: No
1718 |  |  | China | 四川省宜宾市第二中医院 | 345035 | 4456800% |  | 02N/Ingenuity Core128/Elite China/couch vertical error |  | 4 |  | Replace the couch vertical moter brake | [12 empty] | 119018734 | Closed | 1688529 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | Engineering | [12 empty] | IB Existing | bar: No
1719 |  |  | China | 无锡市江原医院 | 345020 | 4456700% |  | 05K for IB/Ingenuity Core128/Elite China |  | 3 |  | None | [12 empty] | 119013713 | Closed | 1685160 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | TBD | [6 empty] | TBD | [5 empty] | TBD | bar: No
1720 |  |  | United States | Central Vermont Medical Center | 40024 | 4456700% |  | Error stating the recon is not initialized. |  | 3 |  | Flashed the DMC | [12 empty] | 119014761 | Closed | 1682144 |  |  | Evaluation | Evaluation | NotEngineering | [5 empty] | Engineering | [6 empty] | Recon is not initialized | [5 empty] | Refresh |  |  |  | System | NotEngineering | DMC | NotEngineering-DMC | NotEngineering-DMC | NotEngineering | DMC | NotEngineering | NotEngineering-DMC | bar: No
1721 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 337000 | 4456700% |  | E- stop ilerlemiyor (E-stop not progressing) |  | 2 |  | Slip ring clean up | [12 empty] | 119014667 | Closed | 1682141 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | Engineering | [6 empty] | E-stop cannot be closed | [5 empty] | IB Existing | bar: No
1722 |  |  | China | Xinjiang Province Yili Prefecture Friendship Hospital | 40048 | 4456600% |  | 02N/Ingenuity CT 728326/gantry fault |  | 3 |  | Reboot gantry and host | [12 empty] | 119006027 | Closed | 1680588 |  |  | Opened | Opened | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Gantry cannot scan | [5 empty] | Refresh |  |  |  | System | Gantry-SW | NoEnoughInfo | Gantry-SW-NoEnoughInfo | Gantry-SW-NoEnoughInfo | NoEnoughInfo | NoEnoughInfo | Gantry-SW | Gantry-SW-NoEnoughInfo | bar: No
1723 |  |  | Japan | 独立行政法人国立病院機構 | 341015 | 4456700% |  | 04DS:KK:リングアーチファクト発生。 |  | 3 Yang, Cancan |  | Replace the module | [12 empty] | 119012031 | Closed | 1680466 |  |  | Evaluation - Pending Translation | Evaluation - Pending Translation | [6 empty] | Engineering | [6 empty] | Ring artifact | [5 empty] | IB Existing | 459800074772 | BR.64 TDMS V2 (RoHS FRU) | bar: No
1724 |  |  | China | Shaanxi Province Ankang T.C.M. Hospital | 345041 | 4456500% |  | Application/Ingenuity Core128/Elite China |  | 5 |  | N/A | [12 empty] | 118998684 | New | 1675936 |  |  | Opened | Opened | [6 empty] | TBD | [6 empty] | TBD | [5 empty] | TBD | bar: No
1725 |  |  | United States | Philips Medical Capital | 40055 | 4456400% |  | Abnormal Sound from Gantry |  | 3 Dantha, Rajashekar |  | Abnormal noise at high speed/ discolored track, found 3 foot section on belt track of rotor (V-Groove) has foreign material, cleaned 6 times, tested unable to hear afterwaards | [12 empty] | 118997843 | Closed | 1671626 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Gantry rotation noise | [5 empty] | IB Existing | bar: No
1726 |  |  | Peru | BANCO BBVA PERU | 40034 | 4456400% |  | 40034-ERROR DE CALIBRACION Y DE IQ COLOCANDO TIEMPO DE ESPERA |  | 3 |  | The IQ check error message was due to the detector temperature. We provide instructions to perform this procedure in the future. | [12 empty] | 118996183 | Closed | 1671002 |  |  | Opened | Opened | NotEngineering | [5 empty] | Engineering | [6 empty] | DMS Temp error during IQ check | [5 empty] | Refresh |  |  |  | System | DMS-HW | Instruction | DMS-HW-Instruction | DMS-HW-Instruction | NotEngineering | Instruction | DMS-HW | DMS-HW-Instruction | bar: No
1727 | Date0117 |  | China | 河南省新乡市凤泉区大块镇卫生院 | 345009 | 4457500% |  | 05K For IB/Ingenuity Core128/Elite China/Crash |  | 5 |  | just remote service . instruct the customer to delete patient data and restart computer | [12 empty] | 119059752 | Closed | 1711914 |  |  | Opened - Pending Translation | Opened - Pending Translation | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Host software stuck | [5 empty] | Refresh |  |  |  | System | Console-SW | NoEnoughInfo | Console-SW-NoEnoughInfo | Console-SW-NoEnoughInfo | NoEnoughInfo | NoEnoughInfo | Console-SW | Console-SW-NoEnoughInfo | bar: No
1728 |  |  | China | 中国医科大学附属盛京医院大连分院 | 345001 | 4457500% |  | 05R/Ingenuity Core128/Elite China/xray error |  | 3 |  | Clean the HV connector, filament cal, LTC | [12 empty] | 119059781 | Closed | 1711928 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | Engineering | [6 empty] | ARC | [5 empty] | IB Existing | bar: No
1729 |  |  | China | 羊坊店医院 | 40035 | 4457500% |  | 05K for IB/Ingenuity CT 728326/ |  | 3 |  | delete data and restart host | [12 empty] | 119059347 | Fixed | 1711720 |  |  | Opened - Pending Translation | Opened - Pending Translation | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Software system run slowly | [5 empty] | Refresh |  |  |  | System | Console-SW | NoEnoughInfo | Console-SW-NoEnoughInfo | Console-SW-NoEnoughInfo | NoEnoughInfo | NoEnoughInfo | Console-SW | Console-SW-NoEnoughInfo | bar: No
1730 |  |  | United States | Clay County Medical Center | 40003 | 4457500% |  | Research FCO requirements for system. |  | 5 |  | Researched current software level (is V4.1.10.25202) and future implementation of FCO72800737 and FCO72800740 | [12 empty] | 119065701 | Closed | 1710741 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Engineering | [6 empty] | Implement FCO | [5 empty] | Refresh | bar: No
1731 |  |  | United States | Unity Health Systems | 40037 | 4457500% |  | Gantry speaker not working |  | 3 |  | Two cables for intercom and speaker were disconnected. Reconnected them and verified normal operation of system | [12 empty] | 119064442 | Closed | 1709721 |  |  | Opened | Opened | [6 empty] | Other | [6 empty] | Gantry audio cables not connected | [5 empty] | IB Existing | bar: No
1732 |  |  | China | 林甸县医院 | 345038 | 4457400% |  | 02N/Ingenuity Core128/Elite China/Change IP |  | 3 |  | New PACS/RIS has different IP, need change HOST and ISP IP | [12 empty] | 119051002 | Closed | 1707110 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Customer | [6 empty] | Change IP | [5 empty] | IB Existing | bar: No
1733 |  |  | China | Harbin City Sen Hai Hospital | 345007 | 4457400% |  | 01/Ingenuity Core128/Elite China/Consulting the warranty |  | 3 |  | N/A | [12 empty] | 119050862 | Closed | 1706990 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Consolut Warranty | [5 empty] | IB Existing | bar: No
1734 |  |  | United States | Leesburg Regional Medical Center | 40022 | 4457300% |  | 0705CDY73 Water spilled on keyboard. |  | 2 |  | Replace the keyboard | [12 empty] | 119049434 | Closed | 1705871 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Water spilled on the keyboard | [5 empty] | Non-refresh | bar: No
1735 |  |  | United States | Leesburg Regional Medical Center | 40022 | 4457400% |  | Need XR-29 certificate |  | 5 |  | created PDF certificate . emailed to customer. Uploaded to TEAMS files | [12 empty] | 119055716 | Closed | 1704854 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Engineering | [6 empty] | XR 29 Request | [5 empty] | Refresh | [6 empty] | - | - | 0 | 0 | 0 | - | bar: No
1736 |  |  | United Kingdom | STOCKPORT NHS FOUNDATION TRUST | 40016 | 4457400% |  | The table is not coming on, scanner down: error message saying that the system is not in a working state |  | 2 Suggest to close, not Refresh SW issue. |  | Reboot gantry, Host, CIRS | [12 empty] | 119054406 | Closed | 1704206 |  |  | Opened | Opened | NotEngineering | [5 empty] | Engineering | [6 empty] | System is not in a working state | [5 empty] | Refresh |  |  |  | System | NotEngineering | SeemsHWProblem | NotEngineering-SeemsHWProblem | NotEngineering-SeemsHWProblem | NotEngineering | SeemsHWProblem | NotEngineering | NotEngineering-SeemsHWProblem | bar: No
1737 |  |  | China | 济南市第四人民医院 | 40038 | 4457300% |  | Appliaction/Ingenuity CT 728326 |  | 5 |  | TBD | [12 empty] | 119039999 | New | 1702553 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | TBD | [6 empty] | TBD | [5 empty] | TBD | bar: No
1738 |  |  | China | Nanchang University First Affiliated hospital (Jinggangshan Hospital) | 345002 | 4457200% |  | 01/Ingenuity Core128/Elite China/add printer |  | 3 Zhai, Lisa |  | None | [12 empty] | 119033110 | Closed | 1698344 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Add printer | [5 empty] | IB Existing | bar: No
1739 |  |  | China | Hebei Province Tangshan City Workers Hospital | 40044 | 4457200% |  | 02N/Ingenuity CT 728326/Maintenance work order content |  | 5 |  | TBD | [12 empty] | 119032969 | New | 1698263 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | TBD | [6 empty] | TBD | [5 empty] | TBD | bar: No
1740 |  |  | Australia | Maia Financial Pty Limited | 341000 | 4457300% |  | Ring artifact |  | 2 WANG LIN, LV |  | Replace Module 31 | [12 empty] | 119039779 | Closed | 1698250 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Ring Artifact | [5 empty] | IB Existing | 459800911271 | TDMS V2 M.Det Class A FRU | bar: No
1741 |  |  | China | Shaanxi Province Ankang T.C.M. Hospital | 345041 | 4457200% |  | 02N/Ingenuity Core128/Elite China/change print header |  | 3 Zhai, Lisa |  | N/A | [12 empty] | 119032918 | Closed | 1698226 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Change printer header | [5 empty] | IB Existing | bar: No
1742 |  |  | Japan | Maki Hospital | 341017 | 4456900% |  | DICOM image transfer stopped |  | Lu, Xueying | [14 empty] | CTUDT00632127, CTUDT00631602. Wait for SP2 deploy.Ask for the detail SW version of the customer, if HF3 and SP1 have been implemented. |  | 1698189 |  |  | Investigation | Investigation | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS transfer | [5 empty] | Refresh |  |  |  | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | bar: No
1743 |  |  | United States | Philips Medical Capital | 40028 | 4457200% |  | Artifact in the field of view. Customer requesting onsite service in the morning. |  | 2 De la Zerda, Yotam |  | Replace Module 17 | [12 empty] | 119032657 | Closed | 1698133 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Artifact | [5 empty] | IB Existing | 459800911271 | TDMS V2 M.Det Class A FRU | bar: No
1744 |  |  | China | Hebei Province Tangshan City Workers Hospital | 40044 | 4457200% |  | Wechat:02N/Ingenuity CT 728326/intercom fault |  | 3 |  | Reconnect the intercom cables | [12 empty] | 119032464 | Closed | 1698003 |  |  | Opened | Opened | [6 empty] | Others | [6 empty] | Gantry audio cables not connected | [5 empty] | IB Existing | bar: No
1745 |  |  | China | Zhalantun City People's Hospital | 40045 | 4457100% |  | 02N/Ingenuity CT/gantry rotor brake fault |  | 3 |  | REPLACE BRAKE ASSY FOR AC | [12 empty] | 119027577 | Closed | 1694155 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | Engineering | [6 empty] | Main motor problem | [5 empty] | IB Existing | 453566502551 | SPARE BRAKE ASSY FOR AC MOTOR | bar: No
1746 |  |  | China | Beijing City Hospital of Traditional Chinese and Western Medicine | 40032 | 4457100% |  | 05K for IB/Ingenuity CT 728326/Gantry noise |  | 3 |  | Performed visual inspection, DMS fan | [12 empty] | 119026617 | Fixed | 1693379 |  |  | Opened | Opened | [6 empty] | Engineering | [6 empty] | Gantry noise | [5 empty] | IB Existing |  | DMS Fan | bar: No
1747 |  |  | United States | Philips Medical Capital | 40028 | 4457100% |  | Partial Ring Artifact |  | 3 Smith-Kifus, Vikki |  | None | [12 empty] | 119031824 | Closed | 1692909 |  |  | Investigation | Investigation | [6 empty] | TBD | [6 empty] | Partial ring artifact | [5 empty] | IB Existing | bar: No
1748 |  |  | China | Jiangsu University Affiliated Hospital | 345023 | 4457000% |  | Wechat:05K for IB/Ingenuity Core128/Elite China/Images cannot be automatically transferred to PACS |  | 3 Zhai, Lisa |  | check the auto storage configuration ，found the auto storage configuration changed, re-config the auto storage | [12 empty] | 119025632 | Closed | 1691281 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | Image cannot auto transferred to PACS | [5 empty] | Refresh |  |  |  | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | bar: No
1749 |  |  | United States | Ashland Hospital Corporation | 40011 | 4453500% |  | E-stop can not be closed due to vertical potentiometer connection issue, FSE cleaned potentiometer connection |  | Suggest to closed this case. 2 Cohen, Shmuel |  | Cleaned the contacts and the string of the potentiometer. Ran the couch up and down multiple times and performed tests scans | [12 empty] | 118839314 | Closed | 1588094 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | E-stop | [5 empty] | IB Existing | bar: No
1750 | Date0124 |  | United States | Philips Medical Capital | 40055 | 4458400% |  | Streak Artifact on Heads and Face studies |  | 2 |  | Cleaned all surfaces within the FOV and ran air calibration. The image artifact was eliminated and the system returned to the customer ready for use | [12 empty] | 119105972 | Closed | 1735649 |  |  | Opened | Opened | [6 empty] | Others | [6 empty] | Streak Artifact | [5 empty] | IB Existing | bar: No
1751 |  |  | China | 东南大学附属中大医院江北院区 | 345021 | 4458300% |  | 05K/Ingenuity Core128/Elite China/report errors |  | 3 |  | on site check the couch statues，found the top-right cover was distorted，need to adjust | [12 empty] | 119104489 | Closed | 1735612 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | Engineering | [6 empty] | Couch movement stop | [5 empty] | IB Existing | bar: No
1752 |  |  | China | Shandong Province Tai'an City T.C.M. No. 2 Hospital | 345037 | 4458300% |  | 02N/Ingenuity Core128/Elite China/ct box key pressed not ok |  | 3 |  | Replace CT Box | [12 empty] | 119103684 | Closed | 1735508 |  |  | Opened | Opened | [6 empty] | Engineering | [6 empty] | CT Box Button | [5 empty] | IB Existing | 459801705022 | CT BOX ASSEMBLY_BEST | bar: No
1753 |  |  | Japan | Independent administrative institution National Hospital Organization | 341015 | 4458100% |  | WY:ディテクター交換作業 |  | 5 Xu, Xiaoting |  | Replace module | [12 empty] | 119090993 | Closed | 1730053 |  |  | Closed - Cancelled | Closed - Cancelled | [6 empty] | Engineering | [6 empty] | Rplace module | [5 empty] | IB Existing | 459800074772 | BR.64 TDMS V2 (RoHS FRU) | bar: No
1754 |  |  | China | 林甸县医院 | 345038 | 4458100% |  | WeChat：application/Ingenuity Core128/Elite China |  | 3 |  | Clinical engineer feedback the customer | [12 empty] | 119089517 | In Progress | 1729229 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Customer sonsult for the head scan protocol | [5 empty] | IB Existing | bar: No
1755 |  |  | United States | Clay County Medical Center | 40003 | 4458000% |  | Knee wedge pad and head/neck pad are coming apart. |  | 5 | [14 empty] | 119088067 | Closed | 1728208 |  |  | Evaluation | Evaluation | CushionBreakdown | [5 empty] | Engineering | [6 empty] | Head / Armreset Cushion break down | [5 empty] | Refresh | 453566502591, 459800042241 | Cushion-Head/ Armrest, PATIENT CUSHIONS EXP |  | System | Couch-HW | CushionBreakdown | Couch-HW-CushionBreakdown | Couch-HW-CushionBreakdown | CushionBreakdown | CushionBreakdown | Couch-HW | Couch-HW-CushionBreakdown | bar: No
1756 |  |  | United States | Clay County Medical Center | 40003 | 4458000% |  | Replace CT Slicker on Couch |  | 5 Wedemeier, Tracy |  | Replace the couch cushion | [12 empty] | 119087728 | Closed | 1727994 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Table cushion | [5 empty] | IB Existing | 459801350221 | STANDARD CUSHION-47.24 INCH MATTRESS | bar: No
1757 |  |  | United States | Clay County Medical Center | 40003 | 4458000% |  | Pads/cushions are already failing |  | CFM SHU, Tim | [14 empty] | CFM |  | 1726643 |  |  | Investigation | Investigation | CTD00006748 | [5 empty] | Engineering | [6 empty] | pad/ cushion | [5 empty] | IB Existing |  |  |  | System | Couch-HW | Accessories | bar: No
1758 |  |  | China | 陕西省安康中医医院 | 345041 | 4458000% |  | wechat：02N/Ingenuity Core128/Elite China/report errors |  | 3 Cai, Cathy |  | Reboot gantry | [12 empty] | 119083073 | Closed | 1724894 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | X-ray system cannot comply | [5 empty] | Refresh |  |  |  | System | Gantry-HW | NoEnoughINfo | Gantry-HW-NoEnoughINfo | Gantry-HW-NoEnoughINfo | NoEnoughInfo | NoEnoughINfo | Gantry-HW | Gantry-HW-NoEnoughINfo | bar: No
1759 |  |  | United States | Central Vermont Medical Center | 40024 | 4458000% |  | Q/O: Error stating the recon is not initialized. |  | 3 Wedemeier, Tracy |  | None, 1 time occur, no available bug report | [12 empty] | 119082878 | Closed | 1724656 |  |  | Investigation | Investigation | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Recon is not initialized | [5 empty] | Refresh |  |  |  | System | CIRS-SW | NoEnoughInfo | CIRS-SW-NoEnoughInfo | CIRS-SW-NoEnoughInfo | NoEnoughInfo | NoEnoughInfo | CIRS-SW | CIRS-SW-NoEnoughInfo | bar: No
1760 |  |  | United States | Clay County Medical Center | 40003 | 4458000% |  | Operation Manual is not installed on system. |  | 5 |  | Installed Operation Manual on application software from CD disk | [12 empty] | 119088210 | Closed | 1724191 |  |  | Evaluation | Evaluation | [6 empty] | Service | [6 empty] | Operation Manual is not installed on system. | [5 empty] | Refresh | bar: No
1761 |  |  | United States | Unity Health Systems | 40037 | 4457900% |  | 0304CDN8A Table making a rubbing sound when moving in |  | 3 |  | Located the noise coming from the clutch on the rear of the table. When the table moves in. the clutch seems to be wobbling. Since the system is not down and the noise can barely be heard, it will be kept monitored for now | [12 empty] | 119080646 | Closed | 1723393 |  |  | Evaluation | Evaluation | [6 empty] | Engineering | [6 empty] | Couch movement noise | [5 empty] | IB Existing | bar: No
1762 |  |  | China | 福建省立医院 | 345011 | 4457900% |  | 05K For IB/Ingenuity Core128/Elite China/Check the machine |  | 5 Wang, Elaine |  | N/A | [12 empty] | 119075426 | Closed | 1720326 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Support customer for the quality check | [5 empty] | IB Existing | bar: No
1763 |  |  | China | Inner Mongolia Hinggan League People's Hospital | 345030 | 4457900% |  | 05K for IB/Ingenuity Core128/Elite China/image artifact |  | 3 |  | Add M12, channel 181 into the bad detector list | [12 empty] | 119075313 | Closed | 1720265 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | Engineering | [6 empty] | Ring artifact | [5 empty] | IB Existing | bar: No
1764 |  |  | China | Xinjiang Province Yili Prefecture Friendship Hospital | 40048 | 4457800% |  | 02N/Ingenuity CT 728326/change HIS/RIS config |  | 3 Zhai, Lisa |  | None | [12 empty] | 119071417 | Closed | 1716866 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Change HIS/RIS config | [5 empty] | IB Existing | bar: No
1765 |  |  | Japan | Tomakomai City Hospital | 341014 | 4457800% |  | SF:ガントリー回転時に異音 |  | 5 Xu, Xiaoting |  | Replace main belt and tensioner bracket | [12 empty] | 119070533 | Imcomplete | 1716355 |  |  | Evaluation | Evaluation | [6 empty] | Engineering | [6 empty] | Gantry rotation noise | [5 empty] | IB Existing | 453567178301, 453566492401, | Bracket, Motor to Tensioner; POLY-VEE BELT | bar: No
1766 |  |  | Japan | Maki Hospital | 341017 | 4457800% |  | 03CS:KKI:撮影中にエラー発生＞＞szInfo: SoftArcReg 0x211* /S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT |  | 4 WANG LIN, LV |  | Longtc (removal of inert gas in X-ray tube) was carried out for the error caused by the discharge in the X-ray tube. | [12 empty] | 119069836 | Closed | 1715961 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Arc | [5 empty] | IB Existing | bar: No
1767 |  |  | China | Le qing City People's Hospital | 40054 | 4457800% |  | 02N/Ingenuity CT 728326/Dual display |  | 4 |  | Replace video adapter cable, clean software junk files. | [12 empty] | 119069836 |  | 1715767 |  |  | Opened | Opened | CTUDT00632173 | [5 empty] | Engineering | [6 empty] | Dual monitor didplay | [5 empty] | Refresh |  |  |  | System | Console-HW | Monitor | Console-HW-Monitor | Console-HW-Monitor | CTUDT00632173 | Monitor | Console-HW | Console-HW-Monitor | bar: No
1768 |  |  | United States | Central Vermont Medical Center | 40024 | 4456700% |  | Error stating the recon is not initialized. CFM and bug report are required to check the detail of the recon error. |  | CFM E, Sravanthi |  | TBD | [12 empty] | CFM |  | 1715357 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | TBD | [6 empty] | TBD | [5 empty] | TBD | bar: No
1769 |  |  | Portugal | Centro Hospitalar de Vila Nova | 40030 | 4457800% |  | O TAC da Urgência sobreaquece sem motivo aparente, sendo que este tipo de anomalia foi já identificado várias vezes |  | 3 |  | User messages stating that the acquisition parameters exceeded the generator and/or tube capacity. Customer has to decrease acquisition parameters in order to continue | [12 empty] | 119072657 | Closed | 1715261 |  |  | Evaluation | Evaluation | [6 empty] | Engineering | [6 empty] | User message: exceed the tube capacity | [5 empty] | IB Existing | bar: No
1770 |  |  | China | DaKuai Town Health, Fengquan District, Xinxiang City, Henan Province | 345009 | 4457700% |  | 05K For IB/Core128/ESTOP opened |  | 3 Cai, Cathy |  | Ask customer to check the foot switch | [12 empty] | 119068537 | Closed | 1714527 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | E-stop open | [5 empty] | IB Existing | bar: No
1771 |  |  | Japan | 医療法人清翠会 牧病院 | 341017 | 4457800% |  | 【CT:APP:NO】 |  | 5 Zhu, Chun |  | Since the parameter cannot be changed, replace the protocol | [12 empty] | 119071342 | Closed | 1714420 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Consolut for protocol parameter change | [5 empty] | IB Existing | bar: No
1772 |  |  | China | Shandong Province Chest Hospital | 40025 | 4457700% |  | 05L/Ingenuity CT 728326/blue screen Silva will check for the similar blue screen during CT login case. CTUDT00632553 |  | 2 |  | Login in Software use smart card and configuration | [12 empty] | 119068325 | Closed | 1714155 |  |  | Opened | Opened | CTUDT00632553 | [5 empty] | Engineering | [6 empty] | Blue screen when login CT. | [5 empty] | Refresh |  |  |  | System | Console-SW | Blue-Screen | Console-SW-Blue-Screen | Console-SW-Blue-Screen | CTUDT00632553 | Blue-Screen | Console-SW | Console-SW-Blue-Screen | bar: No
1773 |  |  | China | Inner Mongolia Hinggan League People's Hospital | 345030 | 4457700% |  | 05K for IB/Ingenuity Core128/Elite China/wrong password |  | 3 Zhai, Lisa |  | Notify customer wait for half hour, then re-enter password. | [12 empty] | 119067917 | Closed | 1713864 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Customer | [6 empty] | Password lock | [5 empty] | Refresh | bar: No
1774 |  |  | United States | Unity Health Systems | 40037 | 4457500% |  | MONTHLY QA |  | 5 Dantha, Rajashekar |  | N/A | [12 empty] | 119066415 | Closed | 1713419 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Monthly QA | [5 empty] | Non-refresh | bar: No
1775 |  |  | China | China Medical University Affiliated Shengjing Hospital Dalian Branch | 345001 | 4457500% |  | 05R/Ingenuity Core128/Elite China/cushion |  | 3 |  | None | [12 empty] | 119062453 | Closed | 1712212 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Customer | [6 empty] | Customer need cushion | [5 empty] | Non-refresh | 453567517311 | Cushion, Patient Support iCT 2 | bar: No
1776 | Date0103U |  | United States | Select Specialty Hospitals | 310051 | 4455900% |  | in the middle of procedure , needing clinical support ASAP |  | 1 Mohanta, Deepranjan |  | N/A | [12 empty] | 118972702 | Closed | 1662461 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Need clinical support | [5 empty] | Non-refresh | bar: No
1777 | Date0110U |  | United States | Adventhealth Zephyrhills | 300106 | 4456800% |  | Issue with perfusion mapping software |  | 5 |  | Russ indicated that post processing absolute value graph was not acquiring the necessary data per the Radiologist on CT Brain Perfusion studies. It was explained to Russ that the interscan delay is too long to acquire these values due to the size of the detector array and the interscan delay of the jog cycle. | [12 empty] | 119021901 | Closed | 1687244 |  |  | Opened | Opened | NotEngineering | [5 empty] | Engineering | [6 empty] | CT Brain Perfusion study post processing absolute value graph was not acquiring the necessary data per the Radiologist | [5 empty] | Refresh |  |  |  | System | NotEngineering | ParameterSetting | NotEngineering-ParameterSetting | NotEngineering-ParameterSetting | NotEngineering | ParameterSetting | NotEngineering | NotEngineering-ParameterSetting | bar: No
1778 |  |  | United States | Lakeland Regional Medical Center | 320474 | 4456600% |  | Q/O Customer Visit |  | 5 |  | Customer visit. No problems noted | [12 empty] | 119011183 | Closed | 1680115 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | None | [6 empty] | None | [5 empty] | TBD | bar: No
1779 |  |  | United States | Adventhealth Zephyrhills | 300106 | 4456400% |  | Needs password reset - to change protocols |  | 5 E, Sravanthi | [14 empty] | 118996944 | Closed | 1671122 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Customer request for password reset | [5 empty] | Non-refresh | bar: No
1780 | Date0117U |  | Mexico | DE LAGE LANDEN SA DE CV | 10007 | 4457500% |  | EMPEZO MARCAR LEYENDA: EL SISTEMA DE ADQUISICION DEL GANTRY NO PUEDE CUMPLIR CON LOS REQUISITOS. (Legend of start mark: gantry procurement system cannot meet the requirements) |  | Wait for case close, then close this one. 2 |  | TBD | [12 empty] | 119058058 | In Progress | 1711239 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | TBD | [6 empty] | TBD | [5 empty] | TBD | bar: No
1781 |  |  | Netherlands | Stichting Rijnstate Ziekenhuis | 52101 | 4457500% |  | NZC109 Terugkomende ringartefacten met CT hersenen |  | 3 |  | Reboot gantry | [12 empty] | 119062380 | Closed | 1708651 |  |  | Opened | Opened | [6 empty] | Engineering | [6 empty] | On 12-1 and 14-1 there has been an error with the aircal | [5 empty] | IB Existing | bar: No
1782 |  |  | United States | Select Specialty Hospitals | 310051 | 4457200% |  | Unable to burn a CD |  | 5 Reja, Afroj |  | None | [12 empty] | 119037460 | Closed | 1701199 |  |  | Evaluation | Evaluation | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Unable to burn CD | [9 empty] | System | TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | bar: No
1783 | Date0124U |  | United States | Adventhealth Zephyrhills | 300099 | 4458200% |  | Detector Noise |  | 3 |  | Replace the module | [12 empty] | 119102132 | Imcomplete | 1734864 |  |  | Opened | Opened | [6 empty] | Engineering | [6 empty] | Contrast inside, image artifact | [5 empty] | IB Existing | 459800594811 | TDMS Module Class A FRU | bar: No
1784 |  |  | United States | Christ Hospital | 300205 | 4458200% |  | 65842604 IP System will not power on |  | 2 |  | Replace tape switch cable | [12 empty] | 119103253 | Closed | 1733242 |  |  | Opened | Opened | [6 empty] | Engineering | [6 empty] | E-stop cannot be closed | [5 empty] | IB Existing | 459800623252 | FRU, CBL, HANDLE TAPE SWITCH | bar: No
1785 |  |  | Netherlands | Stichting Rijnstate Ziekenhuis | 52101 | 4458000% |  | Geen ECG signaal op host computer |  | 3 |  | Replace the ECG cable | [12 empty] | 119086967 | Closed | 1722972 |  |  | Opened | Opened | [6 empty] | Engineering | [6 empty] | ECG | [5 empty] | IB Existing | 455012301901 | ECG Monitor Trigger 7M Ca | bar: No
1786 |  |  | United States | Select Medical Corporation | 32261 | 4457800% |  | error code and unit is not working properly |  | 2 |  | Ran air calibrations on system and found no problems with the system. | [12 empty] | 119074698 | Closed | 1718985 |  |  | Evaluation | Evaluation | [6 empty] | TBD | [6 empty] | TBD | [5 empty] | TBD | bar: No
1787 | Date0214 |  | China | 新疆伊犁州友谊医院 | 40048 | 4460500% |  | 02N/Ingenuity CT 728326/add node | CT has no problem. Customer forget OLEVEL password, can not add node now. Customer feedback no need add it now. Inform FSE that set password when next time onsite service | 5 |  | Reset olevel password in next visit | [12 empty] | 119220768 | Closed | 1802238 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Customer | [6 empty] | Forget olevel password | [5 empty] | Refresh | [12 empty] | No: Yes
1788 |  |  | China | 四川省宜宾市第二中医院 | 345035 | 4460500% |  | 05K For IB/Ingenuity Core128/Elite China/software crash | An error was been found when open the application.restore the Backup,the problem was same as before | 3 |  | Reload the SW | [12 empty] | 119220747 | Closed | 1802200 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Error when open the application | [5 empty] | Refresh |  |  |  | System | Console-SW | NoEnoughInfo | Console-SW-NoEnoughInfo | Console-SW-NoEnoughInfo | NoEnoughInfo | NoEnoughInfo | Console-SW | Console-SW-NoEnoughInfo | bar: No
1789 |  |  | China | Fujian Provincial Hospital | 345011 | 4460300% |  | 05K for IB/Ingenuity Core128/Elite China/image artifact |  | 3 |  | Bad detector test, add the bad detector in the list | [12 empty] | 119212899 | Closed | 1799824 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Ring artifact | [5 empty] | IB Existing | [12 empty] | No: Yes
1790 |  |  | United States | Philips Medical Capital | 40055 | 4460200% |  | 0699CDN84 can't hear auto voice commands at gantry since last service |  | 3 |  | Removed the gantry cover and reconnected the gantry speaker and microphone. Tested the auto voice and everything was working as it should. | [12 empty] | 119209686 | Closed | 1797588 |  |  | Investigation | Investigation | [6 empty] | Service | [6 empty] | Gantry rear intercom cable not connected. | [5 empty] | IB Existing | [12 empty] | No: Yes
1791 |  |  | China | Jiangxi Province Children’s Hospital | 40041 | 4460200% |  | 02N/Ingenuity CT 728326/artifact |  | 3 |  | FSE onsite and check compenstor pass. After air calibration, no more artifact | [12 empty] | 119206501 | Closed | 1795803 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Artifact | [5 empty] | IB Existing | [12 empty] | No: Yes
1792 |  |  | Peru | BANCO BBVA PERU | 40034 | 4460100% |  | CADA SEMANA PIDE CALIBRACION Y ENVIA MENSAJE DE ERROR DE TEMPERATURA | Temperature error when doing calibration | 3 |  | None | [12 empty] | 119205728 | Closed | 1795273 |  |  | Evaluation | Evaluation | [6 empty] | Customer | [6 empty] | Temperature error | [5 empty] | Refresh | [6 empty] | - | - | 0 | 0 | 0 | - | No: Yes
1793 |  |  | United States | Lackey Memorial Hospital | 40015 | 4460100% |  | Calcium Scoring off/ not communicating - not tracking the heartrate |  | 3 ZHANG, Yue Xing |  | None | [12 empty] | 119205544 | Closed | 1795081 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | No track of heart rate | [5 empty] | IB Existing | [12 empty] | No: Yes
1794 |  |  | Germany | Gesundheit Nord gGmbH Klinikverbund | 40002 | 4460100% |  | Artefakte bei sequentiellen Scans untersuchen |  | 3 |  | A-plane aligned and necessary calibrations carried out. Artifacts still present. Data compiled and forwarded to the factory. System operational. | [12 empty] | 119203651 | Closed | 1793755 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Artifact | [5 empty] | IB Existing | [12 empty] | No: Yes
1795 |  |  | China | 江西省儿童医院 | 40041 | 4460000% |  | 02N/Ingenuity CT 728326/ Artifact |  | 3 ZHANG, Yue Xing |  | has ring artifact, customer did no do air calibration for a long time Repair Action: FSE feedback that no need dispatch, do air calibration and keep monitoring. Test/Inspection: Functional Check PASSE | [12 empty] | 119193681 | Closed | 1787314 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Artifact | [5 empty] | IB Existing | [12 empty] | No: Yes
1796 |  |  | United States | Philips Medical Capital | 40053 | 4459900% |  | 0802CDN85 Recurring issue EKG monitor not connecting to CT |  | 3 Cohen, Shmuel |  | Per RSE checked 5V power supply to iBox. Voltage slightly under 5V, adjusted to between 5 and 5.1 as recommended | [12 empty] | 119191364 | Closed | 1785655 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | EKG Monitoring disconnect | [5 empty] | IB Existing | [12 empty] | No: Yes
1797 |  |  | Japan | 医療法人清翠会 牧病院 | 341017 | 4460000% |  | 【CT:APP:RK】 |  | 5 Zhu, Chun |  | None | [12 empty] | 119195776 | Closed | 1784109 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | N/A | [5 empty] | Non-refresh | [12 empty] | No: Yes
1798 |  |  | Japan | Maki Hospital | 341017 | 4460000% |  | If the Patient Orientation set when entering patient information is incorrectly set and scanned, it cannot be corrected. RFC |  | CFM Lu, Xueying |  | N/A | [12 empty] | CFM |  | 1783058 |  |  | Closed - Done | Closed - Done | [6 empty] | Customer | [6 empty] | Patient Orientation set incorrectly | [5 empty] | Non-refresh | [12 empty] | No: Yes
1799 |  |  | China | Laiyang City People's Hospital | 345005 | 4459100% |  | ADR: Error: Cantry acquisition system cannot comply NMPA, Sean to check with Chinal Local TSE |  | CFM Lu, Xueying |  | None | [12 empty] | NMPA |  | 1782945 |  |  | Closed - Done | Closed - Done | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Gantry acquisition system cannot comply | [5 empty] | Refresh |  |  |  | System | Gantry-SW | NoEnoughINfo | Gantry-SW-NoEnoughINfo | Gantry-SW-NoEnoughINfo | NoEnoughInfo | NoEnoughINfo | Gantry-SW | Gantry-SW-NoEnoughINfo | bar: No
1800 |  |  | United States | Philips Medical Capital | 40053 | 4459900% |  | Scanner locks up during gated studies | EC has more than 15 results. also has Cal Score series with addition results. Did test scan and hour glass appeared with a wait until customer can continue planning. Did another test scan with reference EC, no hourglass or delay. | 4 Mohanta, Deepranjan |  | Advised customer to remove some functional results , leaving 0 and 40% or to build new EC from refence set. Also local disc at 25% remaining, deleted old exams from local. also advised customer to keep disk space at 50-60% available | [12 empty] | 119192014 | Closed | 1781846 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Customer | [6 empty] | Scanner lock up. | [5 empty] | Refresh | [6 empty] | - | - | 0 | 0 | 0 | - | No: Yes
1801 |  |  | Puerto Rico | METRO MAYAGUEZ INC | 40009 | 4459900% |  | Throws message: gantry can not comply |  | 2 ZHANG, Yue Xing |  | The TX was replaced to resolve the issue | [12 empty] | 119190214 | Imcomplete | 1781674 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Gantry cannot comply | [5 empty] | IB Existing | 453567560151 | Kit, Moog Velocity Laser Power | [10 empty] | No: Yes
1802 |  |  | United States | Lackey Memorial Hospital | 40015 | 4459900% |  | Getting clicking noise and error message during scan; Could not find any events caused problems. |  | 5 XU, Miley |  | Checked the CIRS Event Logs. Nothing showing as errors. Could not find any events caused problems | [12 empty] | 119191523 | Closed | 1781469 |  |  | Closed - Done | Closed - Done | Should be assembly issue. so change to Others | [5 empty] | Others | [6 empty] | Clicking noise and error message during scan | [5 empty] | Refresh | [6 empty] | - | - | Should be assembly issue. so change to Others | 0 | 0 | - | No: Yes
1803 |  |  | United States | Central Vermont Medical Center | 40024 | 4459900% |  | 0699CFXXX IQ Check and Air Cal Failed |  | 4 |  | Reboot Console and gantry | [12 empty] | 119188616 | Closed | 1780192 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | NoEnoughINfo | [5 empty] | Engineering | [6 empty] | IQ check and Air cal fail | [5 empty] | Refresh |  |  |  | System | Console-SW | NoEnoughINfo | Console-SW-NoEnoughINfo | Console-SW-NoEnoughINfo | NoEnoughINfo | NoEnoughINfo | Console-SW | Console-SW-NoEnoughINfo | bar: No
1804 |  |  | China | Lindian County Hospital | 345038 | 4459700% |  | Wechat:02N/Ingenuity Core128/Elite China/There is noise |  | 3 |  | Power on the UPS | [12 empty] | 119184777 | Closed | 1779223 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Customer | [6 empty] | Noise | [5 empty] | Non-refresh | [12 empty] | No: Yes
1805 |  |  | United States | Philips Medical Capital | 40055 | 4459700% |  | 0304CDN84 Unable to run table in/out |  | 2 Wedemeier, Tracy |  | R2D or CGMP were replaced to resolve the issue | [12 empty] | 119184325 | Closed | 1777502 |  |  | Closed - Duplicate | Closed - Duplicate | [6 empty] | Engineering | [6 empty] | Unable to run table in/out | [5 empty] | IB Existing |  | R2D, CGMP | [10 empty] | No: Yes
1806 |  |  | Argentina | INSTITUTO NACIONAL DE | 40027 | 4459600% |  | LA PRIMERA ES QUE EN EL MOVIMIENTO VERTICAL DE LA CAMILLA SE GENERA UN SONIDO FUERA DE LO HABITUAL | Lubrication of vertical mechanism | 3 |  | Vertical sliding lubrication mechanism | [12 empty] | 119182792 | Closed | 1776768 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Lubrication | [5 empty] | IB Existing | [12 empty] | No: Yes
1807 |  |  | Argentina | INSTITUTO NACIONAL DE | 40027 | 4459600% |  | SE REALIZAN LOS ESTUDIOS CARDIACOS HAY QUE REINICIAR TODO EL EQUIPO PORQUE NO RECONOCE EL MONITOR MULTIPAREMETRICO | Cardiac examination must restart the whole device because it cannot recognize the multi area monitor | 3 |  | Replace IBOX-2 | [12 empty] | 119182814 | In Progress | 1776772 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Cardiac examination | [5 empty] | IB Existing | 459800166651 | I-BOX-2 BASE BOARD ASSY | [10 empty] | No: Yes
1808 |  |  | United States | Leesburg Regional Medical Center | 40022 | 4459500% |  | Customer Wants To Update the Worklist and Store |  | 5 Mohanta, Deepranjan | [14 empty] | 119176029 | Closed | 1773219 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Customer want to update the worklist and store | [5 empty] | Non-refresh | [12 empty] | No: Yes
1809 |  |  | United States | Unity Health Systems | 40037 | 4459500% |  | Couch is not moving horizontally Gantry Cant Init |  | 2 Temzin, Maria |  | None. System itself back to normal. | [12 empty] | 119175579 | Closed | 1772876 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Couch can move horizontal, suspect horizontal motor relay | [5 empty] | IB Existing | [12 empty] | No: Yes
1810 |  |  | China | Jinan City No. 4 People's Hospital | 40038 | 4459400% |  | 05K For IB/Ingenuity CT 728326/gantry power down | intermittent E-Stop open | 3 WANG, Peter |  | Replace R2D | [12 empty] | 119160348 | Closed | 1769792 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | E-stop open | [5 empty] | IB Existing | 459800467343 | R2D | [10 empty] | No: Yes
1811 |  |  | United States | Philips Medical Capital | 40055 | 4459400% |  | Gantry acquisition system can't comply error due to R2D failure; FSE replaced the R2D |  | 4 Shai, Gal |  | Replace R2D and CGMP | [12 empty] | 119170019 | Closed | 1769485 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | Gantry acquisition system cannot comply | [5 empty] | IB Existing | 459800467343, 459800440404 | R2D, CGMP | [10 empty] | No: Yes
1812 |  |  | United States | Philips Medical Capital | 40053 | 4459300% |  | Log collection for slow coronary scans | Obtained log information on two scans using different protocols for review and comparison | 3 |  | N/A | [12 empty] | 119159286 | Closed | 1769168 |  |  | Evaluation | Evaluation | [6 empty] | Service | [6 empty] | Collect Log | [5 empty] | Non-refresh | [12 empty] | No: Yes
1813 |  |  | Chile | Scotiabank Chile | 40008 | 4459400% |  | TIENE FISURADA LA CORREA DEL GANTRY POLI 5 |  | 4 |  | Replace main belt | [12 empty] | 119166236 | Incomplete | 1768688 |  |  | Evaluation | Evaluation | [6 empty] | Engineering | [6 empty] | Main belt has broken | [5 empty] | IB Existing | 453566492401 | POLY-VEE BELT | [10 empty] | No: Yes
1814 |  |  | Chile | Scotiabank Chile | 40008 | 4459200% |  | CL1675-PROBLEMA CON LA HORA |  | 4 |  | Se ingresa por via remota al reconstructor y se ajusta la hora para coincidir con la hora del computador principal. Se comprueba que las imágenes ahora aparecen con la hora correcta | [12 empty] | 119152878 | Incomplete | 1763938 |  |  | Investigation | Investigation | time not sync, belongs to service | [5 empty] | Service | [6 empty] | Image display the incorrect time | [5 empty] | Refresh | [6 empty] | - | - | time not sync, belongs to service | 0 | 0 | - | bar: No
1815 |  |  | United States | Philips Medical Capital | 40055 | 4459200% |  | QO: material for install |  | 5 |  | None | [12 empty] | 119150998 | Closed | 1763841 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | N/A | [5 empty] | Non-refresh | bar: No
1816 |  |  | Chile | BANCO DE CHILE | 341005 | 4459300% |  | CL1688-FALLA DE SOFTWARE, NO LEVANTA EL SOFTWARE |  | 2 |  | Reload host SW | [12 empty] | 119156334 | Imcomplete | 1763457 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632553 | [5 empty] | Engineering | [6 empty] | Login CT, cannot open application, blue screen | [5 empty] | Refresh |  |  |  | System | Console-SW | Blue-Screen | Console-SW-Blue-Screen | Console-SW-Blue-Screen | CTUDT00632553 | Blue-Screen | Console-SW | Console-SW-Blue-Screen | bar: No
1817 |  |  | Turkey | ADIYAMAN OZEL SAGINSTURTAS SAN VE TIC A.S | 341006 | 4459200% |  | bazı sekans isimlerinde yazım yanlışlığı var. | There's a mistake in writing in some random names. | 3 Dogan, Meltem |  | It was described to the user how to change the filming protocols. System is running | [12 empty] | 119151455 | Closed | 1761966 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Customer | [6 empty] | Write in some random names | [5 empty] | Non-refresh | bar: No
1818 |  |  | Italy | ASST SPEDALI CIVILI | 40026 | 4459200% |  | ARTEFATTO SULLE IMMAGINI | Image artifact | 3 Lu, Xueying |  | This problem is already under management with case 0119150388 . This report is closed as double | [12 empty] | 119150706 | Closed | 1760129 |  |  | N/A, Duplicated | Investigation | [6 empty] | N/A, Duplicated | bar: No
1819 |  |  | Italy | ASST SPEDALI CIVILI | 40026 | 4459200% |  | Artefatto concentrico | Concentric artefact | 3 Lu, Xueying | [14 empty] | 119150388 | Closed | 1760109 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Image Artifact | [5 empty] | IB Existing | 459800911271 | TDMS V2 M.Det Class A FRU | bar: No
1820 |  |  | United States | Lackey Memorial Hospital | 40015 | 4459100% |  | 0308CDN84 3-System Restricted |  | 3 XU, Miley |  | Replace vertical transducer | [12 empty] | 119143848 | Closed | 1758669 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | E-stop open | [5 empty] | IB Existing | 459800623331 | Vertical Transducer | bar: No
1821 |  |  | Japan | Maki Hospital | 341017 | 4459000% |  | 01CS:RS:朝のホスト再起動後画面がフリーズしている>>リモートから再起動し解消>>顕著なハードのエラーなし | The screen is frozen after the morning restart | 4 Zhu, Chun |  | Restart and reboot from remote. The event log shows no marked hardware error. Judging from malfunction on temporary applications. Follow-up | [12 empty] | 119142546 | Closed | 1758190 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Screen frozen after restart monitor | [5 empty] | Refresh |  |  |  | System | TBD | NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo | NoEnoughInfo | NoEnoughInfo | TBD | TBD-NoEnoughInfo | bar: No
1822 |  |  | Poland | Government Agency. | 341010 | 4458900% |  | Unstable temperature in the detector, replaced heater modules, power supply cable and common DMS power |  | 3 Lu, Xueying |  | Replace the heater, 48V power supply, CDMP | [12 empty] | 119134833 | Closed | 1757509 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | DMS temperature is not stable | [5 empty] | IB Existing | 459800559353, 459800550141, 453567434781, | CDMP RoHS with pack; BR64T RoHS HEATER REPLACEMENT KIT; CBL, Heater 48VDC Power Supply | bar: No
1823 |  |  | Poland | Government Agency. | 341002 | 4458900% |  | Pękaja zaczepy | Hooks are breaking | 4 |  | Replace the strap | [12 empty] | 119134588 | Fixed | 1757423 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Hooks are breaking | [5 empty] | IB Existing | 453566451921 | 14 INCH RESTRAINT STRAP | bar: No
1824 |  |  | Japan | Maki Hospital | 341017 | 4459000% |  | 01DS:RS:キーボードが反応しない>>延長ケーブルを介さず、HOSTに直接接続し復帰 |  | 4 Zhu, Chun |  | The keyboard doesn't react to me Confirmed reproducibility of the phenomenon I can't. From the situation that occurred, USB Extended cable and I decided to exchange extension cable. After work, confirm the operation of the device I confirmed that it works. | [12 empty] | 119141839 | Closed | 1756825 |  |  | Evaluation | Evaluation | [6 empty] | Engineering | [6 empty] | Keyboard not working | [5 empty] | IB Existing | 452213250062 | USB EXTENSION CABLE A-A 3 METERS | bar: No
1825 |  |  | China | Laiyang City People's Hospital | 345005 | 4458900% |  | 01/Ingenuity Core128/Elite China/scan interrupt |  | 3 WANG, Peter |  | None | [12 empty] | 119131954 | In Process | 1756475 |  |  | Investigation | Investigation | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | ACQ not OK | [5 empty] | Refresh |  |  |  | System | Gantry-SW | NoEnoughInfo | Gantry-SW-NoEnoughInfo | Gantry-SW-NoEnoughInfo | NoEnoughInfo | NoEnoughInfo | Gantry-SW | Gantry-SW-NoEnoughInfo | bar: No
1826 |  |  | China | Jiangxi Province Children’s Hospital | 40041 | 4458900% |  | 02N/Ingenuity CT 728326/Login account locked |  | 3 Wang, Elaine |  | 客户多次输错密码后，主机登录账号被锁，告知账户被锁半个小时后可以解锁，等待半个小时后登录正常。 | [12 empty] | 119131739 | Fixed | 1756388 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Customer | [6 empty] | Account locked | [5 empty] | Refresh | [6 empty] | - | - | 0 | 0 | 0 | - | bar: No
1827 |  |  | Poland | Government Agency. | 341002 | 4458800% |  | sprawdzenie paska rotora gantry | Loud gantry rotor belt operation | 4 |  | Replace the main belt | [12 empty] | 119127889 | Closed | 1753636 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Gantry rotate main belt loud | [5 empty] | IB Existing | 453566492401 | POLY-VEE BELT | bar: No
1828 |  |  | China | ZJ Yueqing No.1 people Hosp | 40054 | 4458700% |  | Duplicate to PR1738861_Intermittent ring artifact in head images. |  | CFM XU, Miley |  | Swap module position, and re-calibrate | [12 empty] | CFM |  | 1752690 |  |  | Closed - Duplicate | Closed - Duplicate | [6 empty] | Engineering | [6 empty] | Intermittent ring artifact | [5 empty] | IB Existing | bar: No
1829 |  |  | United States | Philips Medical Capital | 40055 | 4458800% |  | Ring on image due to faulty detector modules; |  | 4 XU, Miley |  | Replaced detector modules and calibrated DMS | [12 empty] | 119129904 | Incomplete | 1751910 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | Image Artifact | [5 empty] | IB Existing | 459800911271 | TDMS V2 M.Det Class A FRU | bar: No
1830 |  |  | Portugal | Centro Hospitalar de Vila Nova | 40030 | 4458700% |  | Paragens sucessivas do software entre exames no tc do SU sendo necessário fazer shutdown do equipamento | Successive software stops between exams in the CT of the SU requiring shutdown of the equipment | 3 |  | The reported problem could not be reproduced and the analysis of the bug reports was inconclusive. System software update expected to be available soon | [12 empty] | 119123395 | Closed | 1750826 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | PendingBugRep | [5 empty] | Engineering | [6 empty] | Successive software stops between exams | [5 empty] | Refresh |  |  |  | System | TBD | NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo | PendingBugRep | NoEnoughInfo | TBD | TBD-NoEnoughInfo | bar: No
1831 |  |  | Puerto Rico | DORADO HEALTH INC DBA | 40007 | 4458800% |  | No video on monitor, replaced host computer |  | 3 SHU, Tim |  | Replae Host | [12 empty] | 119128366 | Fixed | 1749805 |  |  | Closed - Done | Closed - Done | DEFOA | [5 empty] | Engineering | [6 empty] | No video on the monitor | [5 empty] | Refresh | 459801493151 | CT Gen5.3 Host, Z4 |  | System | Console-HW | Monitor | Console-HW-Monitor | Console-HW-Monitor | DEFOA | Monitor | Console-HW | Console-HW-Monitor | bar: No
1832 |  |  | China | Zhalantun City People's Hospital | 40045 | 4458700% |  | 02N/Ingenuity CT 728326/Transport node change |  | 3 Wang, Elaine | [14 empty] | 119120122 | Closed | 1748217 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | N/A | [5 empty] | Non refresh | bar: No
1833 |  |  | China | Zhalantun City People's Hospital | 40045 | 4458700% |  | 02N/Ingenuity CT 728326/Image artifacts |  | 2 WANG, Peter |  | Module M39 and M40 were replaced | [12 empty] | 119120114 | Closed | 1748210 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Image Artifact | [5 empty] | IB Existing | 459800911281 | TDMS Module V2 | bar: No
1834 |  |  | United States | Leesburg Regional Medical Center | 40022 | 4458700% |  | Trying fill out the ACR paperwork to get accredited. Needs info on how to change scan field of view, number of... |  | 3 E, Sravanthi |  | N/A | [12 empty] | 119121843 | Closed | 1745330 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | N/A | [5 empty] | Non refresh | bar: No
1835 |  |  | Japan | 苫小牧市立病院 | 341014 | 4458700% |  | 【CT:APP:NO】 | A tube hot can be seen by a protocol to shoot the chest after taking a neck (surview hand grater). Would you like to cut DRI? | 5 Zhu, Chun |  | I think it is one countermeasure to switch to fixed MAS. If the maximum MAS is input to the protocol, the generation is reduced. | [12 empty] | 119120771 | Fixed | 1744325 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | NoEnoughINfo | [5 empty] | Engineering | [6 empty] | Tube is too hot | [5 empty] | Refresh |  |  |  | System | Gantry-SW | TubTooHot | Gantry-SW-TubTooHot | Gantry-SW-TubTooHot | NoEnoughINfo | TubTooHot | Gantry-SW | Gantry-SW-TubTooHot | bar: No
1836 |  |  | China | 四川省宜宾市第二中医院 | 345035 | 4458600% |  | application/Ingenuity Core128/Elite China/ |  | 3 Zhai, Lisa |  | N/A | [12 empty] | 119112938 | New | 1743398 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | N/A | [5 empty] | Non refresh | bar: No
1837 |  |  | China | Shaanxi Province Xianyang Central Hospital (Xianyang City No. 2 People's Hospital) | 40050 | 4458600% |  | wechat:02N/Ingenuity CT 728326/disk space release failed |  | 3 Wang, Elaine |  | Clean D disk | [12 empty] | 119112742 | In Process | 1743261 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632261 | [5 empty] | Engineering | [6 empty] | Disk full | [5 empty] | Refresh |  |  |  | System | Console-SW | DiskFull | Console-SW-DiskFull | Console-SW-DiskFull | CTUDT00632261 | DiskFull | Console-SW | Console-SW-DiskFull | bar: No
1838 |  |  | China | Le qing City People's Hospital | 40054 | 4458500% |  | Image artifact - Swapped module locations and calibrated system |  | 3 Mulhern, Melissa |  | Swap module position, and re-calibrate | [12 empty] | 119106843 | Closed | 1738861 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | Intermittent ring artifact | [5 empty] | IB Existing | bar: No
1839 |  |  | China | Inner Mongolia Hinggan League People's Hospital | 345030 | 4458500% |  | 05K for IB/Ingenuity Core128/Elite China/artifact |  | 3 WANG, Peter |  | Replace module | [12 empty] | 119106582 | Closed | 1738702 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Image Artifact | [5 empty] | IB Existing | 459800911281 | TDMS Module V2 | bar: No
1840 |  |  | China | Zhalantun City People's Hospital | 40045 | 4458500% |  | Wechat:02N/Ingenuity CT 728326/Images fault |  | 3 Zhai, Lisa |  | Replace module | [12 empty] | 119106242 | Closed | 1738506 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Image Artifact | [5 empty] | IB Existing |  | Module | bar: No
1841 |  |  | Turkey | Yıldırım Beyazıt Universitesi Yenimahalle Eğitim Ve Araştırma Hastanesi | 337000 |  |  | The device is too slow and fails quickly |  | CFM WANG, Peter |  | N/A | [12 empty] | CFM |  | 1736840 |  |  | Investigation | Investigation | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Device is too slow, and fails quickly | [5 empty] | Refresh |  |  |  | System | Console-SW | Performance | Console-SW-Performance | Console-SW-Performance | NoEnoughInfo | Performance | Console-SW | Console-SW-Performance | bar: No
1842 |  |  | China | Fujian Province Medical University Affiliated No. 1 Hospital | 345012 | 4458400% |  | E-stops will not close, adjusted left and right tilt switch assembly |  | 3 Kanzeg, Mike |  | Re-adjust the tilt switch | [12 empty] | 119105682 | Closed | 1736474 |  |  | Closed - Done | Closed - Done | [6 empty] | 车载CT | [6 empty] | E-stop open | [5 empty] | IB Existing | bar: No
1843 |  |  | Portugal | Centro Hospitalar de Vila Nova | 40030 | 4457800% |  | Iberia-PT-Centro Hospitalar de Vila Nova-Ingenuity CT 728326 |  | CFM ZHANG, Yue Xing |  | Modify scan parameter | [12 empty] | CFM |  | 1736433 |  |  | Investigation | Investigation | CTD00007125 | [5 empty] | Engineering | [6 empty] | Tube is too hot, increase KV or decrease mAs | [5 empty] | Refresh |  |  |  | System | Gantry-SW | TubTooHeat | Gantry-SW-TubTooHeat | Gantry-SW-TubTooHeat | CTD00007125 | TubTooHeat | Gantry-SW | Gantry-SW-TubTooHeat | bar: No
1844 | Date0221 |  | China | 济南市第四人民医院 | 40038 | 4461000% |  | 05K For IB/Ingenuity CT/Can not scan |  | 2 |  | Replace 2pcs of 453566503142 tube axial fan. | [12 empty] | 119250866 | Closed | 1824106 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | CT Cannot scan | [5 empty] | IB Existing | 453566503142 | tube axial fan | [10 empty] | No: Yes
1845 |  |  | Japan | 苫小牧市立病院 | 341014 | 4461000% |  | WY:モニタケーブル交換 | Monitor cable exchange | 5 |  | Replace the DVI-Mini DP adaptor | [12 empty] | 119249023 | In Progress | 1823424 |  |  | Evaluation | Evaluation | DEFOA | [5 empty] | Engineering | [6 empty] | Monitor cable exchange | [5 empty] | Refresh | 459801495951 | Mini-DisplayPort To DVI Adapter Cable |  | System | Console-HW | Monitor | Console-HW-Monitor | Console-HW-Monitor | DEFOA | Monitor | Console-HW | Console-HW-Monitor | bar: No
1846 |  |  | United States | Philips Medical Capital | 40028 | 4461000% |  | 0705CDY70 Keyboard numbers not working |  | 3 |  | Replace the keyboard | [12 empty] | 119253236 | In Progress | 1822728 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Keyboard not working | [5 empty] | IB Existing | 459800918431 | Cherry USB keyboard English Silver | [10 empty] | No: Yes
1847 |  |  | Turkey | ADIYAMANOZELSAGINSTURTASSANVETICA.S | 341006 | 4460900% |  | Gantry çekimi durdurdu | Recon performance is bad. Recon is slow | 3 |  | User is connected to the system by remote access, request for slow access to the recon process. Run disk testing done. The result was successfully completed. Raw data disc formatted. After the work, the problem was confirmed. The system is active | [12 empty] | 119245308 | Closed | 1820022 |  |  | Investigation | Investigation | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Recon is slow | [5 empty] | Refresh |  |  |  | System | TBD | NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo | NoEnoughInfo | NoEnoughInfo | TBD | TBD-NoEnoughInfo | bar: No
1848 |  |  | Turkey | ADIYAMANOZELSAGINSTURTASSANVETICA.S | 341006 | 4461000% |  | görüntüde çizgi çıkıyor | Poor image quality | 3 |  | UHR filter cleanup done, air calibration bad detector and beam and dms check and pass. The device has been successfully ejected | [12 empty] | 119249529 | In Progress | 1819889 |  |  | Investigation | Investigation | E-125059 | [5 empty] | Engineering | [6 empty] | Poor IQ/ Flurian | [5 empty] | Refresh |  |  |  | System | Couch-HW | Cleaning | Couch-HW-Cleaning | Couch-HW-Cleaning | E-125059 | Cleaning | Couch-HW | Couch-HW-Cleaning | bar: No
1849 |  |  | China | 北京市中西医结合医院 | 40032 | 4460900% |  | 05K for IB/Ingenuity CT 728326/add node |  | 5 |  | N/A | [12 empty] | 119241757 | Fixed | 1818896 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | N/A | [5 empty] | Non refresh | [12 empty] | No: Yes
1850 |  |  | Japan | 独立行政法人国立病院機構 | 341015 | 4461000% |  | 画面のCTのインジケータが黄色で点滅を繰り返している | Screen CT indicator repeats yellow blinking | 4 Yang, Cancan |  | Delete the left of the quota on the screen | [12 empty] | 119247922 | Closed | 1818850 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632520 | [5 empty] | Engineering | [6 empty] | Yellow screen blinking | [5 empty] | Refresh |  |  |  | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow-Screen | Console-SW | Console-SW-Yellow-Screen | bar: No
1851 |  |  | United States | Leesburg Regional Medical Center | 40022 | 4460900% |  | ACR question |  | 4 |  | N/A | [12 empty] | 119245183 | Closed | 1816755 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | N/A | [5 empty] | Non refresh | [12 empty] | No: Yes
1852 |  |  | Chile | Scotiabank Chile | 40008 | 4460700% |  | CL1675-CUANDO SE TOMA UN ANGIOGRAMA, SE PIERDE EL ROI | When you do angiography, you lose ROI | 4 |  | He checked the incident record of submitting questions during the exam. It is verified that the pause key is pressed in the ROI positioning stage in the workflow. The detailed work done during the review will be reviewed in the application. Steps to enhance the customer's angiographic procedures. After verification, the problem was not raised subsequently. | [12 empty] | 119234542 | In Progress | 1813872 |  |  | Opened | Opened | [6 empty] | Customer | [6 empty] | ROI loose during angiography | [5 empty] | Non refresh | [12 empty] | No: Yes
1853 |  |  | Poland | Government Agency. | 341009 | 4460700% |  | Ustawienie automatycznego restaru |  | 4 |  | N/A | [12 empty] | 119232729 | Closed | 1812501 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | N/A | [5 empty] | Non refresh | [12 empty] | No: Yes
1854 |  |  | China | Jiangxi Province Children’s Hospital | 40041 | 4460700% |  | 02N/Ingenuity CT 728326/artifact |  | 3 |  | Replace the module | [12 empty] | 119229022 | In Process | 1810251 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Ring Artifact | [5 empty] | IB Existing | [12 empty] | No: Yes
1855 |  |  | China | Inner Mongolia Hinggan League People's Hospital | 345030 | 4460600% |  | 05K for IB/Ingenuity Core128/Gantry belt | Gantry main belt worn out. | 3 Zhai, Lisa |  | Replace the main belt. | [12 empty] | 119222032 | Closed | 1805144 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Main belt worn out | [5 empty] | IB Existing | 453566492401 | POLY-VEE BELT | [10 empty] | No: Yes
1856 |  |  | United States | Philips Medical Capital | 40055 | 4460600% |  | No ECG signal on Host; RSE instructed customer to restart the Host and Gantry |  | 3 Shai, Gal |  | Reboot Console and Gantry | [12 empty] | 119227201 | Closed | 1804528 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | No ECG Signal in Host | [5 empty] | IB Existing | [12 empty] | No: Yes
1857 |  |  | Chile | BANCO DE CHILE | 341005 | 4460600% |  | FALLA DE SOFTWARE EN EL ANGIOTAC | After reinstalling the software, the device is not allowed to have a cardiac examination | 3 |  | Check the USB connection of the electrocardiograph (Ibox), detect that the host does not detect the card, restart the Ibox card and re-establish communication. Use ECG monitor for function test | [12 empty] | 119225449 | In Progress | 1804145 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | IBOX | [5 empty] | Engineering | [6 empty] | device is not allowed to have a cardiac examination | [5 empty] | Refresh |  |  |  | System | Gantry-HW | IBOX | Gantry-HW-IBOX | Gantry-HW-IBOX | IBOX | IBOX | Gantry-HW | Gantry-HW-IBOX | bar: No
1858 |  |  | France | CLINIQUE DE FLANDRE | 40019 | 4460600% |  | Probleme de table |  | 3 |  | Replace the table switch cable | [12 empty] | 119225913 | Closed | 1803910 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Table problem | [5 empty] | IB Existing | 459800623252 | FRU, CBL, HANDLE TAPE SWITCH | [10 empty] | No: Yes
1859 | Date0228 |  | Chile | Scotiabank Chile | 40008 | 4461800% |  | CORTA LOS EXAMENES |  | 4 |  | TBD | [12 empty] | 119297882 | New | 1849697 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | TBD | [6 empty] | TBD | [5 empty] | TBD | [12 empty] | No: Yes
1860 |  |  | China | 内蒙古自治区第四医院 | 40023 | 4461800% |  | 05L/Ingenuity CT 728326/have noise |  | 3 |  | TBD | [12 empty] | 119297746 | New | 1849628 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | TBD | [6 empty] | TBD | [5 empty] | TBD | [12 empty] | No: Yes
1861 |  |  | China | Hebei Province Tangshan Workers Hospital | 40044 | 4461700% |  | wechat:02n/Ingenuity CT 728326/system report error |  | 3 |  | TBD | [12 empty] | 119294102 | New | 1848864 |  |  | Opened | Opened | [6 empty] | TBD | [6 empty] | TBD | [5 empty] | TBD | [12 empty] | No: Yes
1862 |  |  | China | Laiyang City People's Hospital | 345005 | 4461700% |  | pay job：01/Ingenuity Core128/Elite China/SWO#54874484/Buy parts |  | 3 |  | Replace R2D | [12 empty] | 119291069 | New | 1847697 |  |  | Opened | Opened | [6 empty] | Engineering | [6 empty] | ACQ not OK | [5 empty] | IB Existing | 459800467343 | R2D | [10 empty] | No: Yes
1863 |  |  | Japan | National Hospital Organization | 341013 | 4461700% |  | 03CS:KK:10時30分ごろスキャン停止。再撮影でスキャン可。>> S_GEN_FAULT_ANODE_TUBE_ARC |  | 4 WANG LIN, LV |  | None | [12 empty] | 119291623 | Closed | 1844720 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Tube ARC | [5 empty] | IB Existing | [12 empty] | No: Yes
1864 |  |  | China | 铜陵博爱医院有限公司 | 345034 | 4461600% |  | 05K for IB/Ingenuity Core128/Elite China/couch horizontal movement error |  | 3 |  | Replace the encoder | [12 empty] | 119283825 | In Process | 1844335 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | Engineering | [6 empty] | couch horizontal movement error | [5 empty] | IB Existing |  | Couch horizontal encoders | [10 empty] | No: Yes
1865 |  |  | China | 福建省立医院 | 345011 | 4461600% |  | 05K For IB/Core128/Need move CT |  | 5 |  | 车载CT需要到场配合调整输入电。到场完成配合工作。 | [12 empty] | 119283034 | In Process | 1843954 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | Customer | [6 empty] | 配合车载CT调整输入电 | [5 empty] | Non refresh | [12 empty] | No: Yes
1866 |  |  | China | 广西藤县人民医院 | 40036 | 4461600% |  | 01/Ingenuity CT/Customer consultation |  | 3 |  | 客户咨询此报修的设备是否是能谱CT。告知客户Ingenuity CT 不属于能谱CT。 | [12 empty] | 119282074 | Fixed | 1843710 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | Customer | [6 empty] | 咨询ING是否是能谱CT | [5 empty] | Non refresh | [12 empty] | No: Yes
1867 |  |  | United States | Unity Health Systems | 40037 | 4461500% |  | Monthly QA |  | 5 |  | N/A | [12 empty] | 119278798 | In Progress | 1842888 |  |  | Opened | Opened | [6 empty] | Customer | [6 empty] | Monthly QA | [5 empty] | Non refresh | [12 empty] | No: Yes
1868 |  |  | Poland | Rzadowa Agencja | 341002 | 4461500% |  | Głośna praca w czasie podnoszenia stołu | Loud work while lifting the table | 4 |  | Lubrication of the running elements in the base of the table | [12 empty] | 119273732 | Closed | 1840540 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | Engineering | [6 empty] | Couch vertical movement noise | [5 empty] | IB Existing | [12 empty] | No: Yes
1869 |  |  | China | 陕西省咸阳中心医院 (咸阳市第二人民医院） | 40050 | 4461500% |  | 02N/Ingenuity CT 728326/artifact | Image artifact | 3 |  | TBD | [12 empty] | 119273253 | In Process | 1840352 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | TBD | [6 empty] | TBD | [5 empty] | TBD | [12 empty] | No: Yes
1870 |  |  | United States | Odessa Regional Medical Center | 40029 | 4461400% |  | CT couch is making a creaking noise. FSE required to examine. |  | 4 |  | tested the table just found some normal table noise | [12 empty] | 119266879 | Incomplete | 1835864 |  |  | Evaluation | Evaluation | [6 empty] | Engineering | [6 empty] | Couch movement noise | [5 empty] | IB Existing | [12 empty] | No: Yes
1871 |  |  | China | 陕西省咸阳中心医院 (咸阳市第二人民医院） | 40050 | 4461400% |  | 02N/Ingenuity CT 72832/Image quality issue |  | 3 |  | 客户对图像质量不满意，建议联系应用培训专家跟进 | [12 empty] | 119265591 | In Process | 1834798 |  |  | Opened - Pending Translation | Opened - Pending Translation | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | IQ satisfaction | [5 empty] | Refresh |  |  |  | System | Console-SW | NoEnoughInfo | Console-SW-NoEnoughInfo | Console-SW-NoEnoughInfo | NoEnoughInfo | NoEnoughInfo | Console-SW | Console-SW-NoEnoughInfo | bar: No
1872 |  |  | China | 江苏省徐州医学院附属第三医院 | 40039 | 4461400% |  | 02N/Ingenuity CT 728326/acq not ok |  | 2 Zhai, Lisa |  | Replace TX Box | [12 empty] | 119263416 | Closed | 1833485 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | ACQ not OK | [5 empty] | IB Existing |  | TX Box | [10 empty] | No: Yes
1873 |  |  | China | 江苏省扬州市第一人民医院西区医院 | 345022 | 4461300% |  | 05K/Ingenuity Core128/Elite China/reset password |  | 3 Zhai, Lisa |  | N/A | [12 empty] | 119259328 | In Process | 1830248 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | N/A | [5 empty] | Non refresh | [12 empty] | No: Yes
1874 |  |  | China | Laiyang City People's Hospital | 345005 | 4461300% |  | 05L/Ingenuity Core128/Elite China/report error | 冠脉扫描失败 | 3 |  | Replace R2D | [12 empty] | 119258068 | Closed | 1829389 |  |  | Opened | Opened | [6 empty] | Engineering | [6 empty] | Scan error | [5 empty] | IB Existing | 459800467343 | R2D | [10 empty] | No: Yes
1875 |  |  | China | 宁安市第二人民医院 | 345039 | 4461300% |  | 02N/Ingenuity Core128/Elite China/Yellow screen |  | 3 |  | 升级软件至4.1.10.25220 | [12 empty] | 119258026 | Closed | 1829353 |  |  | Opened - Pending Translation | Opened - Pending Translation | CTUDT00632520 | [5 empty] | Engineering | [6 empty] | Yellow screen | [5 empty] | Refresh |  |  |  | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow-Screen | Console-SW | Console-SW-Yellow-Screen | bar: No
1876 |  |  | China | DaKuai Town Health, Fengquan District, Xinxiang City, Henan Province | 345009 | 4461300% |  | Application/Ingenuity Core128/Elite China | 客户咨询应用问题 | 5 |  | 已联系Application工程师跟进 | [12 empty] | 119257970 | New | 1829296 |  |  | Opened | Opened | [6 empty] | Customer | [6 empty] | consolut for applicable usage | [5 empty] | Non refresh | [12 empty] | No: Yes
1877 | Date0214U |  | Mexico | DE LAGE LANDEN SA DE CV | 10007 | 4460300% |  | HACE RUIDO AL GIRAR EL ROTOR | The equipment made noise on the gantry | 3 |  | Replace the main belt. | [12 empty] | 119218089 | In Progress | 1799517 |  |  | Evaluation | Evaluation | [6 empty] | Engineering | [6 empty] | Gantry rotation noise | [5 empty] | IB Existing | 453566492401 | POLY-VEE BELT | [10 empty] | No: Yes
1878 |  |  | United States | Wake Med Cary Hospital | 9591 | 4460200% |  | NEMO |  | 3 |  | N/A | [12 empty] | 119210589 | Closed | 1797366 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | N/A | [5 empty] | Non refresh | [12 empty] | No: Yes
1879 |  |  | United States | Wake Med Cary Hospital | 9591 | 4460100% |  | NEMO |  | 3 Mukherjee, Susmita |  | N/A | [12 empty] | 119205537 | Closed | 1793202 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | N/A | [5 empty] | Non refresh | [12 empty] | No: Yes
1880 |  |  | United States | Adventhealth Zephyrhills | 300106 | 4460000% |  | Cirs errors; Customer performed a reboot | Reboot CIRS | 3 Squires, Ciera | [14 empty] | 119199380 | Fixed | 1791295 |  |  | Closed - Done | Closed - Done | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | CIRS Error | [5 empty] | Refresh |  |  |  | System | CIRS-SW | NoEnoughInfo | CIRS-SW-NoEnoughInfo | CIRS-SW-NoEnoughInfo | NoEnoughInfo | NoEnoughInfo | CIRS-SW | CIRS-SW-NoEnoughInfo | bar: No
1881 |  |  | United States | Wake Med Cary Hospital | 9591 | 4460100% |  | DMC shot comm collapse errors - Reboot to resolve |  | 3 Mulhern, Melissa |  | Reboot the system | [12 empty] | 119205423 | Closed | 1790791 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | DMC shot comm collapse error | [5 empty] | IB Existing | [12 empty] | No: Yes
1882 |  |  | United States | Adventhealth Zephyrhills | 300106 | 4459900% |  | 1203CDY70 x-ray system can't comply errors |  | 2 Wedemeier, Tracy |  | CPM replaced | [12 empty] | 119190362 | Closed | 1780811 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | X-ray system cannot comply | [5 empty] | IB Existing |  | CPM Generator | [10 empty] | No: Yes
1883 |  |  | United States | Adventhealth Zephyrhills | 300099 | 4459800% |  | 0199CDN84 Vertical lines through images. |  | 3 LI, Xue |  | Clean the some contrast on the mylar ring | [12 empty] | 119186133 | Closed | 1780145 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | some contrast on the mylar ring | [5 empty] | IB Existing | [12 empty] | No: Yes
1884 |  |  | United States | Christ Hospital | 300205 | 4459200% |  | LIFECYCLE |  | 2 E, Sravanthi |  | N/A | [12 empty] | 119152147 | Closed | 1763052 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | N/A | [5 empty] | Non refresh | bar: No
1885 |  |  | United States | Lakeland Regional Medical Center | 320474 | 4459200% |  | Q/O |  | E, Sravanthi |  | N/A | [12 empty] | 119152117 | Closed | 1761010 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | N/A | [5 empty] | Non refresh | bar: No
1886 | Date0228U |  | United States | Select Medical Corporation | 32261 | 4461600% |  | Linear artifacts showing up when doing CT's of head |  | 3 |  | None | [12 empty] | 119287186 | Incomplete | 1846310 |  |  | Opened | Opened | [6 empty] | Engineering | [6 empty] | Line artifact | [5 empty] | IB Existing | [12 empty] | No: Yes
1887 |  |  | Australia | CITISCAN RADIOLOGY PTY LTD | 10491 | 4461400% |  | Left Hand monitor power switch broken, monitor is currently on, system operational. |  | 3 |  | Replace the monitor and the Graphic card | [12 empty] | 119269085 | In Progress | 1833267 |  |  | Evaluation | Evaluation | [6 empty] | Engineering | [6 empty] | Monittor broken | [5 empty] | IB Existing | 459801140821, 459801536801 | 19" LCD display monitor EA193Mi, NVIDIA Quadro P400 2GB (3)mDP GFX | [10 empty] | No: Yes
1888 |  |  | United States | Adventhealth Zephyrhills | 300106 | 4461400% |  | 4.1.10 laser QA failed 2mm, and no SMPTE pattern. |  | 5 Das, Rituparna |  | TBD | [12 empty] | 119268772 | In Progress | 1832961 |  |  | Evaluation | Evaluation | [6 empty] | TBD | [6 empty] | TBD | [5 empty] | TBD | [12 empty] | No: Yes
1889 |  |  | United States | Lakeland Regional Health Systems | 300143 | 4461300% |  | Q/O: Right panel touch panel in switch works but is getting harder to press |  | 5 |  | Replace the control panel | [12 empty] | 119260732 | Closed | 1831229 |  |  | Evaluation | Evaluation | [6 empty] | Engineering | [6 empty] | Panel button hard to press | [5 empty] | IB Existing | 459800484511 | CONTROL PANEL ASSEMBLY-FRONT RH | [10 empty] | No: Yes
| Date0307 |  | China | Zhalut Banner Mein medical hospital | 345036 | 4462400% |  | 05k for IB/ Ingenuity Core128/Elite China/Microphone interference |  | 3 |  | information customer to adjust the button of control voice, adjust the voice button | [12 empty] | 119332804 | Closed | 1871894 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Customer | [6 empty] | CT Box noise too loud | [5 empty] | IB Existing | [12 empty] | No: Yes
|  |  | Poland | Rare Agency. | 341010 | 4462400% |  | Artefakty |  | 4 |  | M11, channel 161 low signal, do air calibration, systm test | [12 empty] | 119332181 | Closed | 1871648 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | Artifact | [5 empty] | IB Existing | [12 empty] | No: Yes
|  |  | United States | Philips Medical Capital | 40055 | 4462500% |  | 1302CFXXX Images not going to Pacs |  | 3 |  | Customer states they can archive to other locations. Suggested they check with Pacs admin/IT department to determine if Pacs is having problems. Contacted customer and archiving is working after Pacs Admin helped. | [12 empty] | 119336448 | Closed | 1871610 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Customer | [6 empty] | PACS | [5 empty] | Non refresh | [12 empty] | No: Yes
|  |  | United States | Odessa Regional Medical Center | 40029 | 4462400% |  | Voice Manager |  | 4 |  | The Voice Manager was selected under the Wrench - the English language was selected - Add Phrase was selected - Maya recorded a new breathing instruction and save the new instruction successfully. Maya had no further questions or need for clinical support | [12 empty] | 119333815 | Closed | 1869393 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Engineering | [6 empty] | Clinical support - Voice Manager | [5 empty] | Refresh | [12 empty] | No: Yes
|  |  | China | Zhalut Banner Mein medical hospital | 345036 | 4462300% |  | 05K for IB/Ingenuity Core128/Elite China/tube condition failed |  | 3 |  | None | [12 empty] | 119324806 | Closed | 1867415 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Rotor seek reference time out | [5 empty] | IB Existing | [12 empty] | No: Yes
|  |  | China | Fujian Provincial Hospital | 345011 | 4462300% |  | 05K for IB/Ingenuity Core128/Elite China/panel failed |  | 3 |  | Replace the control panel | [12 empty] | 119324547 | Closed | 1867232 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Control panel malfunction | [5 empty] | IB Existing | 459800484501 | CONTROL PANEL ASSEMBLY-FRONT LH | [10 empty] | No: Yes
|  |  | China | Laiyang City People's Hospital | 345005 | 4462100% |  | Scan abort due to Gantry acquisition system cannot comply error reported |  | WANG LIN, LV |  | None. No bug report for this case, FSE was not required to visit this site. | [14 empty] | 1866490 |  |  | Closed - Done | Closed - Done | CTUDT00630408 | [5 empty] | Engineering | [6 empty] | ACQ not OK | [5 empty] | Refresh |  |  |  | System | CIRS-SW | ACQ_NOT_OK | CIRS-SW-ACQ_NOT_OK | CIRS-SW-ACQ_NOT_OK | CTUDT00630408 | ACQ_NOT_OK | CIRS-SW | CIRS-SW-ACQ_NOT_OK | No: Yes
|  |  | China | Lindian County Hospital | 345038 | 4462200% |  | 02N/Core128/The belt fault |  | 3 |  | FSE cannot on site for inspect, since the covid 19. Suspect belt or tension Assy | [12 empty] | 119322503 | In Process | 1866164 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Gantry rotation noise | [5 empty] | IB Existing | [12 empty] | No: Yes
|  |  | United States | Philips Medical Capital | 40005 | 4462200% |  | LS Quote |  | 5 |  | New case, the customer ask for quote | [12 empty] | 119321776 | New | 1865692 |  |  | Evaluation | Evaluation | [6 empty] | Customer | [6 empty] | Quote | [5 empty] | Non refresh | [12 empty] | No: Yes
|  |  | United States | Lackey Memorial Hospital | 40015 | 4462200% |  | Q/O unit is making gantry noise |  | 5 |  | FSE onsite check, no more noise of the gantry. | [12 empty] | 119321317 | Closed | 1864385 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Gantry noise | [5 empty] | IB Existing | [12 empty] | No: Yes
|  |  | China | Inner Mongolia Autonomous Region No. 4 Hospital | 40023 | 4462200% |  | Appliaction/Ingenuity CT 728326 |  | 5 |  | Clinical training for enhanced scan | [12 empty] | 119319786 | New | 1864097 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Ehanced scan | [5 empty] | Non refresh | [12 empty] | No: Yes
|  |  | China | Inner Mongolia Autonomous Region No. 4 Hospital | 40023 | 4462200% |  | pay job：01/Ingenuity CT 728326/swo#54899233 buy parts |  | 3 |  | Replace main belt | [12 empty] | 119316866 | Closed | 1862246 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Engineering | [6 empty] | Gantry rotation noise | [5 empty] | IB Existing | 453566492401 | POLY-VEE BELT | [10 empty] | No: Yes
|  |  | China | Beijing City Hospital of Integrated Chinese and Western Medicine | 40032 | 4462200% |  | 05K for IB/Ingenuity CT 728326/ECG issue |  | 3 |  | Replace IBOX and IBOX PS | [12 empty] | 119316239 | Closed | 1861940 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | ECG issue | [5 empty] | IB Existing | 459800166651 | I-BOX-2 BASE BOARD ASSY | [10 empty] | No: Yes
|  |  | China | Sxi feng County T.C.M. Hospital | 345014 | 4462100% |  | 01/Ingenuity Core128/Elite China/add node |  | 3 Zhai, Lisa |  | Provide account and password, instruct customer to add node | [12 empty] | 119311089 | Closed | 1858161 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Add node | [5 empty] | Non refresh | [12 empty] | No: Yes
|  |  | China | Sxi feng County T.C.M. Hospital | 345014 | 4462100% |  | 01/Ingenuity Core128/Elite China/ADD dicom node |  | 3 Cai, Cathy | [14 empty] | 119310704 | Closed | 1857875 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Add Dicom node | [5 empty] | Non refresh | [12 empty] | No: Yes
|  |  | China | Shandong Province Chest Hospital | 40025 | 4462100% |  | 05L/Ingenuity CT 728326/add the printer |  | 5 Zhai, Lisa | [14 empty] | 119310649 | Fixed | 1857812 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Add printer | [5 empty] | Non refresh | [12 empty] | No: Yes
|  |  | United States | Philips Medical Capital | 40055 | 4462100% |  | QO: Add new ct to scanner |  | 5 |  | Added new ct as network node verified reachable | [12 empty] | 119315231 | Closed | 1857317 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Engineering | [6 empty] | Add new CT as network node | [5 empty] | Non refresh | [12 empty] | No: Yes
|  |  | China | Xinjiang YiliuZhou Friendship Hospital | 40048 | 4462100% |  | wechat：02N/Ingenuity CT 728326/The administrator account and password cannot be logged in |  | 3 Zhai, Lisa | [14 empty] | 119309814 | Closed | 1857286 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Engineering | [6 empty] | Administrator account cannot login | [5 empty] | Refresh | [12 empty] | No: Yes
|  |  | China | Inner Mongolia Alasan Right Banner People's Hospital | 345031 | 4462000% |  | 01/Ingenuity Core128/Elite China/consult for password |  | 3 Zhai, Lisa | [14 empty] | 119304377 | Fixed | 1854070 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Consult for password | [5 empty] | Non refresh | [12 empty] | No: Yes
|  |  | United States | Ashland Hospital Corporation | 40011 | 4462000% |  | PM WAS DONE ON THURSDAY 2.24.22 AND NOW ITS RUNNING SLOW |  | 4 Dantha, Rajashekar |  | Recycled the power to the HOST and CIRS computers. Performed Visual Diagnostics on the CIRS and Database Cleanup. | [12 empty] | 119308478 | Closed | 1853132 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | Performance Slow | [5 empty] | Service | [6 empty] | System running slow | [5 empty] | Refresh |  |  |  | System |  |  | - | - | Performance Slow |  |  |  | No: Yes
|  |  | China | Jinan City No. 4 People's Hospital | 40038 | 4462000% |  | Application/05K for IB/Ingenuity CT 728326 |  | 5 |  | None | [12 empty] | 119301166 | Closed | 1852848 |  |  | Investigation | Investigation | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Vessel enhance scan protocol issue | [5 empty] | Refresh |  |  |  | System | TBD | NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo | bar: No
|  |  | China | Guangdong Province Guangzhou City No.8 People's Hospital (Jiahe Hospital District) | 345032 | 4462000% |  | 05K for IB/Ingenuity Core128/Elite China/Scan error |  | 3 |  | Replace the TX box | [12 empty] | 119300552 | Closed | 1852646 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | Scan abort | [5 empty] | IB Existing |  | TX Box | [10 empty] | No: Yes
|  |  | United States | Philips Medical Capital | 40053 | 4462000% |  | 0699CFXXX Received a temperature control malfunction error when running QC |  | 4 Mohanta, Deepranjan |  | Power cycle the system | [12 empty] | 119306348 | Closed | 1851328 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | DMStemperature error | [5 empty] | IB Existing | [12 empty] | No: Yes
| Date0307U |  | United States | Upson Regional Medical Center | 310115 | 4462200% |  | 84628582 TABLE WILL NOT COME OUT WITH FOOT SWITCH OR GANTRY SWITCH - HAS TO BE DONE MANUALLY |  | 5 |  | performed horizontal table cals | [12 empty] | 119320599 | Closed | 1864840 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | Couch can only move manually | [5 empty] | IB Existing | [12 empty] | No: Yes
|  |  | Mexico | DE LAGE LANDEN SA DE CV | 10007 | 4462100% |  | EL EQUIPO MUESTRA ERROR NO SE PUEDE INICIAR LA RECONTRUCCION |  | 3 |  | Replace the FRC | [12 empty] | 119314864 | Closed | 1856898 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | No power to TX | [5 empty] | IB Existing | 453567057167 | FRC ASSEMBLY | [10 empty] | No: Yes
|  |  | Netherlands | Stichting Rijnstate Ziekenhuis | 52101 | 4462100% |  | Bolus tracking scan start niet vanwege trage recon delay. |  | 3 |  | Reload server 1 software | [12 empty] | 119313407 | Closed | 1855542 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Slow recon delay | [5 empty] | Refresh |  |  |  | System | Console-SW | BolusTracking | Console-SW-BolusTracking | Console-SW-BolusTracking | 0 |  |  |  | No: Yes
|  |  | United States | Wake Med Cary Hospital | 9591 | 4462100% |  | 0399CFXXX Need part number for Mattress Pad |  | 5 Mohanta, Deepranjan | [14 empty] | 119313234 | Closed | 1855472 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Customer consult for the pad PN | [5 empty] | Non refresh | [12 empty] | No: Yes
| Date0314U |  | United States | BIBB MEDICAL CENTER | 310175 | 4463000% |  | QO: cracked lexan ring |  | 5 |  | Replace the Lexan ring | [12 empty] | 119365623 | Closed | 1889691 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | Lexan ring crack | [5 empty] | IB Existing | 459800048742 | RING-LEXAN GN5 | [10 empty] | No: Yes
|  |  | United States | Adventhealth Zephyrhills | 300099 | 4462800% |  | Q/O |  | 3 |  | None | [12 empty] | 119351998 | Closed | 1880130 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | TBD | [6 empty] | TBD | [5 empty] | TBD | [12 empty] | No: Yes
| Date0314 |  | China | Jiangxi Province Yugan Chudong Hospital (New Branch) | 337500 | 4463300% |  | wechat:05K/Ingenuity Core128/Elite China/Button failure |  | 3 |  | Replace CT Box | [12 empty] | 119374121 | Closed | 1897118 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632185 | [5 empty] | Engineering | [6 empty] | CT Box button malfunction | [5 empty] | TBD |  | CT Box |  | System | Gantry-HW | CT Box | Gantry-HW-CT Box | Gantry-HW-CT Box | CTUDT00632185 | CT Box | Gantry-HW | Gantry-HW-CT Box | No: Yes
|  |  | United States | Leesburg Regional Medical Center | 40022 | 4463200% |  | 85536292 the A-STOP WILL NOT BOOT |  | 2 |  | Replace the left control panel | [12 empty] | 119373696 | Closed | 1896221 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | E-stop cannot close | [5 empty] | IB Existing | 459800484501 | CONTROL PANEL ASSEMBLY-FRONT LH | [10 empty] | No: Yes
|  |  | China | Lindian County Hospital | 345038 | 4463100% |  | 02N/Ingenuity Core128/Elite China/ECG signal lost |  | 3 |  | None, access constraint due to Covid19 | [12 empty] | 119366570 | In Process | 1894532 |  |  | Investigation | Investigation | [6 empty] | TBD | [6 empty] | TBD | [5 empty] | TBD | [12 empty] | No: Yes
|  |  | China | Lindian County Hospital | 345038 | 4463100% |  | 02N/Ingenuity Core128/Elite China/ECG fault |  | 3 |  | Replace the usb cable between IBOX and Host | [12 empty] | 119367244 | Closed | 1894802 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Engineering | [6 empty] | ECG issue | [5 empty] | IB Existing | 459801731761 | USB EXTENDER CABLE FROM IBOX TO CONSOLE | [10 empty] | No: Yes
|  |  | China | Inner Mongolia Xing'an People's Hospital | 345030 | 4463100% |  | 05k for IB/ Ingenuity Core128/Elite China/Artifact |  | 3 |  | Air calibration | [12 empty] | 119367368 | Closed | 1894884 |  |  | Investigation | Investigation | air calibration | [5 empty] | Customer | [6 empty] | Artifact | [5 empty] | Refresh | [12 empty] | No: Yes
|  |  | China | Inner Mongolia Xing'an People's Hospital | 345030 | 4463100% |  | 05k for IB/ Ingenuity Core128/couch cushion |  | 5 |  | Provide couch curshion | [12 empty] | 119368315 | Closed | 1895326 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Customer need couch curshion | [5 empty] | Non refresh | [12 empty] | No: Yes
|  |  | Portugal | Centro Hospitalar de Vila Nova | 40030 | 4463100% |  | Substituição de colchão da TC do SU que se encontra danificado e embebido em sangue |  | 3 |  | TBD | [12 empty] | 119370548 | New | 1896165 |  |  | Investigation | Investigation | [6 empty] | TBD | [6 empty] | TBD | [5 empty] | TBD | [12 empty] | No: Yes
|  |  | Chile | BANCO DE CHILE | 341005 | 4463000% |  | SE CORTO DISPARO EN EXAMEN CAT |  | 3 |  | None | [12 empty] | 119365128 | Closed | 1889756 |  |  | Investigation | Investigation | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Short shot in CAT | [5 empty] | Refresh |  |  |  | System | TBD | NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo |  |  |  |  | No: Yes
|  |  | Japan | Kinan Hospital | 340000 | 4463000% |  | CX:ガントリベアリンググリスアップ |  | 5 Yang, Cancan |  | Apply grease | [12 empty] | 119359819 | Closed | 1890185 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Grease grill up | [5 empty] | IB Existing | [12 empty] | No: Yes
|  |  | China | Suzhou Guang ci Tumor Hospital Co., Ltd | 345019 | 4463000% |  | 05K for IB/Ingenuity Core128/Elite China/gantry noise |  | 3 |  | Replace the tensioner | [12 empty] | 119361326 | Closed | 1891110 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Gantry noise | [5 empty] | IB Existing | 453567003451 | TENSIONER ASSY | [10 empty] | No: Yes
|  |  | China | Shaanxi Ankang T.C.M | 345041 | 4462900% |  | WeChat：02N/Ingenuity Core128/Recon cann't initialize |  | 3 Gao, Mia |  | Restart the system | [12 empty] | 119353376 | Closed | 1885516 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | ReconNotInitialize | [5 empty] | Engineering | [6 empty] | Recon cannot initialize | [5 empty] | Refresh |  |  |  | System | CIRS-SW | ReconNotInitialize | CIRS-SW-ReconNotInitialize | CIRS-SW-ReconNotInitialize |  |  |  |  | No: Yes
|  |  | China | Shandong Anchu City T.C.M | 345017 | 4462900% |  | 02N/Ingenuity Core128/Elite China/keyboard failure |  | 3 Zhai, Lisa |  | Replace the keyboard | [12 empty] | 119354837 | Closed | 1886511 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Keyboard not working | [5 empty] | IB Existing | 459800918431 | Cherry USB keyboard English Silver | [10 empty] | No: Yes
|  |  | United States | Unity Health Systems | 40037 | 4462900% |  | Monthly QA |  | 5 |  | N/A | [12 empty] | 119357528 | Fixed | 1888805 |  |  | Evaluation | Evaluation | [6 empty] | Customer | [6 empty] | Monthly QA | [5 empty] | Non refresh | [12 empty] | No: Yes
|  |  | United States | Philips Medical Capital | 40028 | 4462800% |  | 1202CDN84 Maximum number of hard arcs |  | 2 |  | Ran a Long Tube Conditioning and Filament calibrations before and after to eliminate arcing in the system | [12 empty] | 119348890 | Closed | 1877958 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | Tube Arc | [5 empty] | IB Existing | [12 empty] | No: Yes
|  |  | United States | Philips Medical Capital | 40055 | 4462800% |  | 0704CFXXX Local full |  | 3 Boddolla, Ramu |  | Deleted orphaned files. Customer checked local space and it's back to 70% free | [12 empty] | 119351031 | Closed | 1879783 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632261 | [5 empty] | Engineering | [6 empty] | Local disk is full | [5 empty] | Refresh |  |  |  | System | Console-SW | DiskFull | Console-SW-DiskFull | Console-SW-DiskFull | CTUDT00632261 | DiskFull | Console-SW | Console-SW-DiskFull | No: Yes
|  |  | China | Zhalantun City People's Hospital | 40045 | 4462800% |  | 02N/Ingenuity CT 728326/IST authorize |  | 3 Gao, Mia |  | N/A | [12 empty] | 119348087 | Closed | 1881830 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | Non refresh | [12 empty] | No: Yes
|  |  | United States | Philips Medical Capital | 40053 | 4462800% |  | QO: Blood pressure cuff order |  | 5 |  | Unable to order blood pressure cuff directly, provided customer with contact to order | [12 empty] | 119351842 | Closed | 1884627 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Order | [5 empty] | Non refresh | [12 empty] | No: Yes
|  |  | Japan | National Hospital Organization | 341015 | 4462700% |  | SF:ディテクターモジュール交換 |  | 5 Xu, Xiaoting |  | Replace the module | [12 empty] | 119339688 | Closed | 1876826 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Artifact | [5 empty] | IB Existing | 459800911271 | TDMS V2 M.Det Class A FRU | [10 empty] | No: Yes
|  |  | China | Suzhou Guang ci Tumor Hospital Co., Ltd | 345019 | 4462700% |  | 05K for IB/Ingenuity Core128/Elite China/Control panel key failure |  | 3 |  | Replace the control panel | [12 empty] | 119340409 | Closed | 1877313 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Control panel malfunction | [5 empty] | IB Existing | 459801484981 | CONTROL PANEL ASSEMBLY-FRONT RIGHT | [10 empty] | No: Yes
| Date0321U |  | Spain | POLICLINICA NTRA SRA DEL ROSARIO SL | 336190 | 4463600% |  | No han podido arrancarlo, se queda bloqueado al intentar encenderlo The password expire date will be in 90 days. |  | 2 |  | By entering with the administrator password, finally we were able to check and update the usual user passwords, so that the client enters the application of the host computer. | [12 empty] | 119396901 | Closed | 1910144 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632236 | [5 empty] | Engineering | [6 empty] | Password | [5 empty] | Refresh |  |  |  | System | Console-SW | User Account | Console-SW-User Account | Console-SW-User Account | CTUDT00632236 | User Account | Console-SW | Console-SW-User Account | No: Yes
| Date0321 |  | China | 河南许昌市中心医院(传染病) | 40047 | 4464000% |  | 05K for IB/Ingenuity CT 728326/ECG problem |  | 3 |  | Reboot | [12 empty] | 119417707 | Fixed | 1919941 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | ECG issue | [5 empty] | Refresh |  |  |  | System | TBD | NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo | NoEnoughInfo |  |  |  | No: Yes
|  |  | China | Henan Xuchang City Central Hospital (Infectious Diseases) | 40047 | 4463800% |  | 05k for IB/ Ingenuity CT 728326/Head rest failure |  | 3 |  | None | [12 empty] | 119409162 | Closed | 1917647 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | HeadRest | [5 empty] | Engineering | [6 empty] | Head rest failure | [5 empty] | Refresh |  |  |  | System | Couch-HW | HeadRest | Couch-HW-HeadRest | Couch-HW-HeadRest | HeadRest |  |  |  | No: Yes
|  |  | China | Suzhou Guang ci Tumor Hospital Co., Ltd | 345019 | 4463800% |  | 01/Ingenuity Core128/Elite China/High pressure syringe product fault |  | 3 |  | TBD | [12 empty] | 119409498 | In Process | 1917737 |  |  | Evaluation | Evaluation | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | High pressure syring fail | [5 empty] | Refresh |  |  |  | System | TBD | NotEngineering | TBD-NotEngineering | TBD-NotEngineering |  |  |  |  | No: Yes
|  |  | Spain | S.C.S. HOSP. GRAL. DE FUERTEVENTURA | 40020 | 4463800% |  | Defectos estéticos encontrados durante el mantenimiento preventivo |  | 5 |  | It is required to use consumable materials during the next preventive maintenance. The customer was informed that although the patient pad is considered to be an accessory not covered by the maintenance contract, in this case, it is ordered and provided free of charge | [12 empty] | 119412660 | Closed | 1918478 |  |  | Evaluation | Evaluation | [6 empty] | Engineering | [6 empty] | Crushion | [5 empty] | Non refresh | 459801350221 | STANDARD CUSHION-47.24 INCH MATTRESS PAD | [10 empty] | No: Yes
|  |  | Poland | Rzadowa Agencja | 341009 | 4463800% |  | W trakcie skanowania pojawiają się pierścienie pochodzące najprawdopodobniej z detektora. |  | 3 |  | Replace the module | [12 empty] | 119413147 | Closed | 1918663 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Ring Artifact | [5 empty] | IB Existing | 459800911271 | TDMS V2 M.Det Class A FRU | [10 empty] | No: Yes
|  |  | United States | Odessa Regional Medical Center | 40029 | 4463800% |  | recon images |  | 4 |  | N/A | [12 empty] | 119414279 | Canceled | 1919097 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A | [12 empty] | No: Yes
|  |  | Chile | BANCO DE CHILE | 341005 | 4463800% |  | NO HAY SEÑAL ECG EN ANGIOTAC |  | 3 |  | TBD | [12 empty] | 119415052 | In Progress | 1919280 |  |  | Evaluation | Evaluation | [6 empty] | Engineering | [6 empty] | No ECG Signal | [5 empty] | IB Existing | [12 empty] | No: Yes
|  |  | China | Henan Xuchang City Central Hospital (Infectious Diseases) | 40047 | 4463700% |  | 05K/Ingenuity CT 728326/Check warranty |  | 3 |  | N/A | [12 empty] | 119405527 | Closed | 1915270 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A | [12 empty] | No: Yes
|  |  | United States | Unity Health Systems | 40037 | 4463700% |  | Monthly QA |  | 5 |  | N/A | [12 empty] | 119407021 | Closed | 1916477 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A | [12 empty] | No: Yes
|  |  | China | Zhalantun City People's Hospital | 40045 | 4463600% |  | 02N/Ingenuity CT/CT box&noise |  | 3 |  | Replace the main belt | [12 empty] | 119395480 | Closed | 1909316 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Gantry rotation noise | [5 empty] | IB Existing | 453566492401 | BELT, 5-RIB POLY-VEE, 198" | [10 empty] | No: Yes
|  |  | China | Inner Mongolia Xing'an People's Hospital | 345030 | 4463600% |  | 05K for IB/Core128/couch cushion |  | 3 Zhai, Lisa |  | None | [12 empty] | 119395938 | Closed | 1909506 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Engineering | [6 empty] | Crushion | [5 empty] | Non refresh | [12 empty] | No: Yes
|  |  | China | Zhalantun City People's Hospital | 40045 | 4463600% |  | 02N/Ingenuity CT/CT box |  | 3 |  | Replace CT Box | [12 empty] | 119395988 | Closed | 1909533 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | CT Box button malfunction | [5 empty] | IB Existing | 453567027195 | CTBOX ASSEMBLY RAFI/TILTED (RoHS) | [10 empty] | No: Yes
|  |  | China | Le qing City People's Hospital | 40054 | 4463600% |  | 02N/Ingenuity CT 728326/scan complete and image not recon |  | 3 |  | Replace GPU P2000 | [12 empty] | 119396185 | Closed | 1909676 |  |  | Investigation | Investigation | GPU | [5 empty] | Engineering | [6 empty] | Not recon | [5 empty] | Refresh | 459801622381 | NVidia P2000 GPU Card |  | System | Console-HW | GPU | Console-HW-GPU | Console-HW-GPU |  |  |  |  | No: Yes
|  |  | China | Shandong Anchu City T.C.M | 345017 | 4463600% |  | 02N/Ingenuity Core128/Password |  | 5 Zhai, Lisa |  | N/A | [12 empty] | 119396596 | Closed | 1909951 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A | [12 empty] | No: Yes
|  |  | China | Jiangsu Province No. 2 T.M. Hospital | 345027 | 4463600% |  | 05K/Ingenuity Core128/The time was not correct |  | 5 |  | Remote guide customer correct the time | [12 empty] | 119397038 | Closed | 1910246 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | ShouldbeService | [5 empty] | Service | [6 empty] | Host time incorrect | [5 empty] | Refresh | [6 empty] | - | - | ShouldbeService |  |  |  | No: Yes
|  |  | China | Henan Xuchang City Central Hospital (Infectious Diseases) | 40047 | 4463600% |  | 02N/Ingenuity CT 728326/Hung up |  | 3 |  | N/A | [12 empty] | 119397426 | Closed | 1910487 |  |  | Closed - Done | Closed - Done | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Software hung up | [5 empty] | Refresh |  |  |  | System | TBD | NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo | NoEnoughInfo |  |  |  | No: Yes
|  |  | China | Fujian Provincial Hospital | 345011 | 4463500% |  | 05K for IB/Ingenuity Core128/Elite China/ACCU ERROR |  | 3 |  | Replace ACCU | [12 empty] | 119385119 | Closed | 1905156 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | F301 burn | [5 empty] | IB Existing | 459800298762 | AC CONTROL UNIT PCB | [10 empty] | No: Yes
|  |  | Japan | Kinan Hospital | 340000 | 4463500% |  | MN:寝台部グリスアップ |  | 5 Xu, Xiaoting |  | Bed section grease up was implemented. There is no problem with the operation of the device. | [12 empty] | 119385423 | Fixed | 1905400 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Other | [6 empty] | Do lubrication | [5 empty] | Non refresh | [12 empty] | No: Yes
|  |  | China | Zhalut Banner Mein medical hospital | 345036 | 4463500% |  | 05k for ib/Ingenuity Core128/Elite China/add dicom node |  | 3 |  | N/A | [12 empty] | 119385732 | Closed | 1905560 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A | [12 empty] | No: Yes
|  |  | China | Yang gao County Medical Group (Yang gao County People's Hospital). | 345025 | 4463500% |  | Application/Ingenuity Core128/Elite China |  | 5 |  | N/A | [12 empty] | 119386003 | New | 1905747 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A | [12 empty] | No: Yes
|  |  | China | Jiangxi Province Children's Hospital | 40041 | 4463500% |  | 02N/Ingenuity CT 728326/gantry noise |  | 3 |  | Adjust the POLY-V SHEAVE position | [12 empty] | 119386538 | Closed | 1906082 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Gangry rotation noise | [5 empty] | IB Existing | [12 empty] | No: Yes
|  |  | Turkey | ADIYAMANOZELSAGINSTURTASSANVETICA.S | 341006 | 4463400% |  | Görüntülerde artefak var |  | 3 WANG LIN, LV |  | the App Support Team investigated the imagination of artefact and reported that the images were normal with the exact residue of the shooting. Specifies whether the lines in the word view will not appear when maximising the remainder. The hospital will be informed about this. No field visits will be performed. The system is active. | [12 empty] | 119380685 | Closed | 1898350 |  |  | Investigation | Investigation | [6 empty] | Customer | [6 empty] | Artifact with residue shooting | [5 empty] | Non refresh | [12 empty] | No: Yes
|  |  | United States | Philips Medical Capital | 40028 | 4463400% |  | 0902CDY73 x6558 QA Low contrast failed |  | 3 Mukherjee, Susmita |  | Customer resolved the issue themselves, cancel the call | [12 empty] | 119381869 | Closed | 1898820 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Customer | [6 empty] | Low contrast failed | [5 empty] | Non refresh | [12 empty] | No: Yes
|  |  | Chile | BANCO DE CHILE | 341005 | 4463400% |  | EL GANTRY SE APAGA |  | 2 |  | Checked the equipment, error log but could not find a specific cause of the problem, the problem could be caused by a power drop in the power supply of the equipment, equipment is operational and under observation | [12 empty] | 119381360 | Closed | 1898857 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Customer | [6 empty] | System Power Off | [5 empty] | Non refresh | [12 empty] | No: Yes
|  |  | China | China Medical University Affiliated Shengjing Hospital Dalian Campus | 345001 | 4463400% |  | 05R/Ingenuity Core128/Elite China/Unable to scan |  | 2 |  | check error log, no filament current, test tube filament&filament open. Replace the tube. | [12 empty] | 119376368 | Closed | 1900666 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Filament fault | [5 empty] | IB Existing |  | Tube | [10 empty] | No: Yes
|  |  | China | Henan Xuchang City Central Hospital (Infectious Diseases) | 40047 | 4463300% |  | 02N/Ingenuity CT 728326/couch rear cover fault |  | 3 WANG LIN, LV |  | Replace the broken cover | [12 empty] | 119374636 | Closed | 1897687 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Couch cover broken | [5 empty] | IB Existing | 459800808451 | COVER ASSY, REAR SUBFRAME, GREY | [10 empty] | No: Yes
|  |  | Chile | Clinica Redsalud Rancagua | 40008 | 4459400% |  | Noise detected when gantry was rotating |  | CFM |  | TBD | [14 empty] | 1916699 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Gantry rotation noise | [5 empty] | IB Existing | [12 empty] | No: Yes
| Date0328U |  | United States | Adventhealth Zephyrhills | 300099 | 4464500% |  | ring artifact |  | 3 |  | Air calibration | [12 empty] | 119450770 | Closed | 1939430 |  |  | Closed - Done | Closed - Done | air calibration | [5 empty] | Customer | [6 empty] | Ring Artifact | [5 empty] | Refresh | [6 empty] | - | - | air calibration |  |  |  | No: Yes
| Date0328 |  | United States | Mclaren Macomb | 40021 | 4464500% |  | Injector not connecting |  | 3 |  | System reboot | [12 empty] | 119450002 | Closed | 1938803 |  |  | Closed - Done | Closed - Done | CTD00007194 | [5 empty] | Engineering | [6 empty] | Injector not connecting | [5 empty] | Refresh |  |  |  | System | Console-SW | InjectorLose | Console-SW-InjectorLose | Console-SW-InjectorLose | CTD00007194 | Console-SW | InjectorLose | Console-SW-InjectorLose | No: Yes
|  |  | Argentina | INSTITUTO NACIONAL DE | 40027 | 4464500% |  | NO DEJA REALIZAR CALIBRACIONES Y NO NO NOS DEJA HACER ESTUDIOS |  | 3 |  | TBD, reboot can restore the system, but will occur again, ask FSE to check on site | [12 empty] | 119449156 | In Progress | 1939031 |  |  | Opened | Opened | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | System cannot do calibration | [5 empty] | Refresh |  |  |  | System | TBD | NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo |  |  |  |  | No: Yes
|  |  | Japan | Social Welfare Corporation 3-beekai | 343000 | 4464400% |  | 01CS:RS:問い合わせ:画面左下のCTインジケータでのImage Transfer failedというメッセージついて |  | 5 Yang, Cancan |  | None | [12 empty] | 119438961 | Fixed | 1931674 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | Image transfer fail | [5 empty] | Refresh |  |  |  | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No: Yes
|  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 337000 | 4464400% |  | Pacs sistemine bağlanmasını istiyorlar |  | 3 |  | N/A | [12 empty] | 119441298 | In Progress | 1933474 |  |  | Evaluation | Evaluation | [6 empty] | Customer | [6 empty] | customer want to connect to the PACS | [5 empty] | Non refresh | [12 empty] | No: Yes
|  |  | Japan | Maki Hospital | 341017 | 4464400% |  | WY:warranty期間中のフォローアップ |  | 5 Yang, Cancan |  | N/A | [12 empty] | 119439229 | Closed | 1936163 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A | [12 empty] | No: Yes
|  |  | Portugal | Centro Hospitalar de Vila Nova | 40030 | 4464400% |  | Faz ruído ao deslizar a mesa e ao angular a gantry |  | 3 |  | Lubrication | [12 empty] | 119442567 | Closed | 1937305 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Noise when couch movement and gantry tilt | [5 empty] | IB Existing | [12 empty] | No: Yes
|  |  | Japan | Maki Hospital | 341017 | 4464300% |  | 01CS:KK:AZEへの画像転送が止まっている。>>Unable to Connect with the target device.となっていましたがPingを確認するとPingは通りました。 |  | 4 Zhu, Chun |  | Hospital Lan issue | [12 empty] | 119432861 | Closed | 1928195 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Customer | [6 empty] | Image transfer fail | [5 empty] | Non refresh | [12 empty] | No: Yes
|  |  | Spain | S.C.S. HOSP. GRAL. DE FUERTEVENTURA | 40020 | 4464300% |  | Las imagenes salen mal, se piensa que es el detector. |  | 2 |  | Replace the M15 | [12 empty] | 119434326 | Incomplete | 1928833 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Artifact | [5 empty] | IB Existing | 459800911271 | TDMS V2 M.Det Class A FRU | [10 empty] | No: Yes
|  |  | United States | Philips Medical Capital | 40028 | 4464300% |  | failed quality control test |  | 3 |  | Replace HP phantom head section | [12 empty] | 119435041 | Closed | 1929017 |  |  | Evaluation | Evaluation | Hphantom Head | [5 empty] | Engineering | [6 empty] | IQ fail | [5 empty] | Refresh | 459801415161 | HARMONIZED PHANTOM HEAD SECTION |  | System | Couch-HW | H-phantom | Couch-HW-H-phantom | Couch-HW-H-phantom |  |  |  |  | No: Yes
|  |  | China | Zhalantun City People's Hospital | 40045 | 4464300% |  | 02N/Ingenuity CT 728326/Automatic voice volume is high |  | 3 Zhai, Lisa |  | Remote guide customer to adjust the voice volumn | [12 empty] | 119433356 | Closed | 1932264 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Customer | [6 empty] | Voice volume high | [5 empty] | Non refresh | [12 empty] | No: Yes
|  |  | China | Henan Xuchang City Central Hospital (Infectious Diseases) | 40047 | 4464300% |  | 05K For IB/Ingenuity CT 728326/consult node information |  | 5 Cai, Cathy |  | N/A | [12 empty] | 119433727 | Closed | 1932493 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A | [12 empty] | No: Yes
|  |  | Chile | BANCO DE CHILE | 341005 | 4464200% |  | PEDIR PARTES POSTERIOR EVALUACION |  | 5 SHU, Tim |  | TBD | [12 empty] | 119429129 | In Progress | 1926494 |  |  | Investigation | Investigation | [6 empty] | TBD | [6 empty] | TBD | [5 empty] | TBD | [12 empty] | No: Yes
|  |  | China | Le qing City People's Hospital | 40054 | 4464200% |  | 02N/Ingenuity CT 728326/report errors |  | 3 WANG LIN, LV |  | Reinstall Host and CIR SW | [12 empty] | 119426425 | Closed | 1927824 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | SW intermittent hung up doing CTA scan | [5 empty] | Refresh |  |  |  | System | Console-SW | Installation | Console-SW-Installation | Console-SW-Installation | 0 | Installation | Console-SW | Installation-Console-SW | No: Yes
|  |  | Puerto Rico | METRO MAYAGUEZ INC | 40009 | 4464200% |  | Device is showing lines at imaging on Pilot |  | 3 |  | Clean the contrast | [12 empty] | 119429197 | Closed | 1929795 |  |  | Closed - Done | Closed - Done | [6 empty] | Customer | [6 empty] | Image Artifact | [5 empty] | Non refresh | [12 empty] | No: Yes
|  |  | United States | Clay County Medical Center | 40003 | 4464100% |  | Patient Database on Local drive only shows last 2 days |  | 3 Boddolla, Ramu |  | Performed a Full Rebuild of the patient Local database | [12 empty] | 119423356 | Closed | 1921974 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | Bugrep | [5 empty] | Engineering | [6 empty] | Only last 2 days patient data show in Local | [5 empty] | Refresh |  |  |  | System | Console-SW | LocalDBLost | Console-SW-LocalDBLost | Console-SW-LocalDBLost |  |  |  |  | No: Yes
|  |  | China | Suzhou Guang ci Tumor Hospital Co., Ltd | 345019 | 4464100% |  | 05K for IB/Core128/Blue screen |  | 3 |  | Close the back programs | [12 empty] | 119419075 | Closed | 1922520 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632553 | [5 empty] | Engineering | [6 empty] | Intermitten blue screen after reboot Host | [5 empty] | Refresh |  |  |  | System | Console-SW | Blue-Screen | Console-SW-Blue-Screen | Console-SW-Blue-Screen | CTUDT00632553 | Blue-Screen | Console-SW | Console-SW-Blue-Screen | bar: No
|  |  | China | Fujian Province Medical University No. 1 Affiliated Hospital | 345012 | 4464100% |  | SA:05K for IB/Ingenuity Core128/Elite China/recon speed |  | 3 Cai, Cathy |  | 车载CT，delete the stuck queries to AI terminal | [12 empty] | 119419492 | Closed | 1922791 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Recon speed slow | [5 empty] | IB existing |  |  |  | System | CIRS-SW | [7 empty] | No: Yes
|  |  | China | Shaanxi Ankang T.C.M | 345041 | 4464100% |  | Wechat:02N/Ingenuity Core128/Elite China/time not accuray |  | 5 |  | Modify the host time | [12 empty] | 119419569 | Closed | 1922832 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | TBD-TimeSync | [5 empty] | Engineering | [6 empty] | Host time is not accuracy | [5 empty] | TBD |  |  |  | System | Console-SW | Time sync | Console-SW-Time sync | Console-SW-Time sync | TBD-TimeSync |  | Console-SW | Console-SW-Time sync | No: Yes
|  |  | Poland | Uniwersytecki Szpital Kliniczny w | 341001 | 4464100% |  | Awaria aparatu |  | 4 WANG LIN, LV |  | Correct the FRC input plug | [12 empty] | 119422936 | Closed | 1925481 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Fail to active scan after power outage | [5 empty] | IB Existing | [12 empty] | No: Yes
|  |  | United States | Westgate Hospital | 40001 | 4464100% |  | Have a question about longitude condition |  | 5 |  | N/A | [12 empty] | 119423646 | Closed | 1926199 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A | [12 empty] | No: Yes
|  |  | United States | Clay County Medical Center | 40003 | 4464000% |  | During scans, a gantry error was populated. |  | 5 WANG LIN, LV | [14 empty] | 119418853 | Closed | 1922262 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | Gantry intermittent error | [5 empty] | IB Existing | 453566496941 | RELAY,24VDC PLUGIN W/COVER | [10 empty] | No: Yes
| Date0406 |  | China | Xinjiang YiliuZhou Friendship Hospital | 40048 | 4464800% |  | 05K for IB/Ingenuity CT 728326/RH panel |  | 3 Ma, Julita |  | Replace the right control panel | [12 empty] | 119455567 | Closed | 1945840 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Right control panel malfunction | [5 empty] | Refresh | 459800484511 | CONTROL PANEL ASSEMBLY-FRONT RIGHT |  | System | Gantry-HW | Control Panel | Gantry-HW-Control Panel | Gantry-HW-Control Panel |  | Control Panel | Gantry-HW | Gantry-HW-Control Panel | No: Yes
|  |  | China | Heilongjiang Province Huachuan County People's Hospital | 345018 | 4464800% |  | 01/Ingenuity Core128/Elite China/Image failure |  | 3 |  | TBD, suspect the reference detector and converter | [12 empty] | 119456018 | In Process | 1946071 |  |  | Evaluation | Evaluation | [6 empty] | Engineering | [6 empty] | Image contrast artifact | [5 empty] | IB Existing | [12 empty] | No: Yes
|  |  | China | Xinjiang YiliuZhou Friendship Hospital | 40048 | 4464800% |  | 05K for IB/Ingenuity CT 728326/transfer slow |  | 3 Zhai, Lisa |  | Guide customer change the IP | [12 empty] | 119457234 | Closed | 1946882 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Customer | [6 empty] | Guide customer change IP | [5 empty] | Non refresh | bar: No to y; y to y
|  |  | United States | Mount Grant General Hospital | 40012 | 4465000% |  | 0711CDY72 not able to login or change the user//customer requested tech support |  | 3 |  | None. Customer cancelled the call, issue resolved by themselves | [12 empty] | 119474726 | Closed | 1953285 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Cannot login or change user | [5 empty] | Refresh |  |  |  | System | Console-SW | Not a defect | Console-SW-Not a defect | Console-SW-Not a defect |  | Not a defect | Console-SW | Console-SW-Not a defect | No: Yes
|  |  | United States | Philips Medical Capital | 40028 | 4465000% |  | phantom was dropped |  | 3 |  | None | [12 empty] | 119472930 | Closed | 1956133 |  |  | Closed - Duplicate | Closed - Duplicate | [6 empty] | N/A | [6 empty] | Phantom drop down | [5 empty] | N/A | bar: No to y; y to y
|  |  | United States | Philips Medical Capital | 40028 | 4465000% |  | someone dropped fantam bar broke off inside |  | 5 |  | N/A | [12 empty] | 119472976 | Closed | 1956164 |  |  | Closed - Duplicate | Closed - Duplicate | [6 empty] | N/A | [6 empty] | Phantom drop down | [5 empty] | N/A | bar: No to y; y to y
|  |  | China | Shaanxi Province Xianyang Central Hospital (Xianyang City No.2 People's Hospital) | 40050 | 4465000% |  | 02N/Ingenuity CT 728326/consult |  | 5 Zhai, Lisa |  | N/A | [12 empty] | 119473389 | Closed | 1956502 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A | bar: No to y; y to y
|  |  | Poland | Uniwersytecki Szpital Kliniczny w | 341001 | 4465100% |  | Konfiguracja nowego węzła |  | 4 |  | N/A | [12 empty] | 119484334 | Closed | 1960992 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A | bar: No to y; y to y
|  |  | United States | Philips Medical Capital | 40055 | 4465300% |  | 0711CFXXX Local Full |  | 3 |  | Full Rebuild, deleting, and restarting host | [12 empty] | 119496938 | Closed | 1963391 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632261 | [5 empty] | Engineering | [6 empty] | Local disk is full | [5 empty] | Refresh |  |  |  | System | Console-SW | DiskFull | Console-SW-DiskFull | Console-SW-DiskFull | CTUDT00632261 | DiskFull | Console-SW | Console-SW-DiskFull | No: Yes
|  |  | China | Suzhou Guang ci Tumor Hospital Co., Ltd | 345019 | 4465200% |  | 01/Ingenuity Core128/Elite China/transfer image fault |  | 3 |  | PACS SW issue, customer contact with PACS service | [12 empty] | 119488761 | Closed | 1963428 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Customer | [6 empty] | Image transfer fail | [5 empty] | Non refresh | bar: No to y; y to y
|  |  | United States | Unity Health Systems | 40037 | 4465300% |  | Couch would not move in and out. |  | 3 |  | None. Customer did not request the service, since the error did not happen again. | [12 empty] | 119498086 | Closed | 1964804 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | Couch cannot move in/out | [5 empty] | IB Existing | [12 empty] | No: Yes
|  |  | United States | Philips Medical Capital | 40053 | 4465200% |  | Left side cover popping when tilting |  | 3 |  | Replace the gantry cover | [12 empty] | 119495228 | Closed | 1965222 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | Left side cover poping when tilting | [5 empty] | IB Existing | 459801485211 | ASSEMBLY-UPPER SIDE COVER LH | [10 empty] | No: Yes
|  |  | Spain | S.C.S. HOSP. GRAL. DE FUERTEVENTURA | 40020 | 4465500% |  | Ha salido un artefacto en la radiografía. |  | 3 |  | Replace defective module, and re-calibration. | [12 empty] | 119502367 | Closed | 1967257 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Image Artifact | [5 empty] | IB Existing | 459800911271 | TDMS V2 M.Det Class A FRU | [10 empty] | No: Yes
|  |  | Germany | Gesundheit Nord gGmbH Klinikverbund | 40002 | 4465500% |  | Koronare und Sagittale Rekonstruktion vom Kopf werden nicht richtig berechnet. |  | 3 |  | None | [12 empty] | 119501530 | Closed | 1969415 |  |  | Investigation | Investigation | CTUDT00632116 | [5 empty] | Engineering | [6 empty] | Coronary and sagittal reconstruction of the head are not calculated correctly. | [5 empty] | Refresh |  |  |  | System | CIRS-SW | MPR Recon | CIRS-SW-MPR Recon | CIRS-SW-MPR Recon | CTUDT00632116 | MPR Recon | CIRS-SW | CIRS-SW-MPR Recon | No: Yes
| Date0407U |  | Mexico | DE LAGE LANDEN SA DE CV | 10007 | 4465000% |  | VERIFICACION DE CALIBRACION DE AIRE Y MONITOR SIGNOS VITALES. |  | 5 |  | Replace the trigger cable | [12 empty] | 119469722 | Closed | 1954700 |  |  | Evaluation | Evaluation | [6 empty] | Engineering | [6 empty] | Check monitor vita signal | [5 empty] | IB Existing | 455012301901 | ECG Monitor Trigger 7M Cable | [10 empty] | No: Yes
|  |  | United States | Select Medical Corporation | 32261 | 4464900% |  | Service Call |  | 5 Butler, Daria |  | replaced the Mylar ring | [12 empty] | 119467155 | Closed | 1953028 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | N/A | [5 empty] | IB Existing | 459800048742 | RING-LEXAN GN5 | [10 empty] | No: Yes
|  |  | United States | Philips Medical Capital | 300057 | 4464800% |  | 0799CFXXX too many login attempts |  | 3 Datta, Avinaba |  | Call sustomer, and ask for wait for 30min, then try again. | [12 empty] | 119458842 | Closed | 1944077 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632236 | [5 empty] | Engineering | [6 empty] | System locked due to wrong password login too many times | [5 empty] | Refresh |  |  |  | System | Console-SW | User Account | Console-SW-User Account | Console-SW-User Account | CTUDT00632236 | User Account | Console-SW | Console-SW-User Account | No: Yes
| Date0411 |  | China | Jiangxi Province Children's Hospital | 40041 | 4466000% |  | 05K for IB/Ingenuity CT 728326/ left monitor display fault |  | 3 |  | Replace the left DP-DVI adaptor | [12 empty] | 119533401 | Closed | 1987919 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | DEFOA | [5 empty] | Engineering | [6 empty] | Left monitor display fault | [5 empty] | Refresh |  |  |  | System | Console-HW | Monitor | Console-HW-Monitor | Console-HW-Monitor | DEFOA | Monitor | Console-HW | Console-HW-Monitor | No: Yes
|  |  | United States | UNITY HEALTH SYSTEMS | 40037 | 4465900% |  | Couch would not move in and out. |  | 3 |  | Lubricant the horizontal table rails | [12 empty] | 119529894 | Closed | 1986880 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | Couch cannot move in/out | [5 empty] | IB Existing | [12 empty] | No: Yes
|  |  | Argentina | INSTITUTO NACIONAL DE | 40027 | 4465900% |  | TIRA DISTINTOS ERRORES EN CONSOLA DE COMANDOS |  | 4 |  | TBD | [12 empty] | 119529955 | New | 1986893 |  |  | Evaluation | Evaluation | [6 empty] | TBD | [6 empty] | TBD | [5 empty] | TBD | bar: No to y; y to y
|  |  | Poland | Rzadowa Agencja | 341008 | 4465900% |  | Uruchomienie aparatu |  | 4 |  | None | [12 empty] | 119529421 | Closed | 1986698 |  |  | Evaluation | Evaluation | [6 empty] | N/A | [6 empty] | Support for perform CT test | [5 empty] | N/A | bar: No to y; y to y
|  |  | United States | Mclaren Macomb | 40021 | 4465900% |  | 0711CFXXX protocol password expired. |  | 3 |  | Introduce to customer to reset the password via IT admin or call local service to reset the password. | [12 empty] | 119532290 | Closed | 1985569 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Engineering | [6 empty] | Protocol password expired | [5 empty] | Refresh |  |  |  | System | Console-SW | User Account | Console-SW-User Account | Console-SW-User Account |  | User Account | Console-SW | Console-SW-User Account | No: Yes
|  |  | Chile | BANCO DE CHILE | 341005 | 4465900% |  | LOS EXAMENES APARECEN CON HORARIO CAMBIADO |  | 5 |  | Rebuilding time. | [12 empty] | 119530859 | Closed | 1984528 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Wrong schedule appers in the patient image. | [5 empty] | Refresh |  |  |  | System | Console-SW | NotEnoughInformation | Console-SW-NotEnoughInformation | Console-SW-NotEnoughInformation |  | NotEnoughInformation | Console-SW | Console-SW-NotEnoughInformation | No: Yes
|  |  | Spain | S.C.S. HOSP. GRAL. DE FUERTEVENTURA | 40020 | 4463800% |  | Defectos estéticos encontrados durante el mantenimiento preventivo |  |  |  | N/A | [12 empty] | CFM |  | 1983711 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A | bar: No to y; y to y
|  |  | Chile | BANCO DE CHILE | 341005 | 4465800% |  | SOLICITA EL CAMBIO DE HORA DEL EQUIPO |  | 3 |  | Update time in computer. | [12 empty] | 119524895 | Closed | 1980925 |  |  | Evaluation | Evaluation | [6 empty] | Customer | [6 empty] | Request to change device time | [5 empty] | Non refresh | bar: No to y; y to y
|  |  | United States | Central Vermont Medical Center | 40024 | 4465800% |  | There is a bad detector. The FSE asked the customer to place a call so that he can order a part. |  | 4 |  | Replace the defective module | [12 empty] | 119524843 | Closed | 1980820 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | Image artifact | [5 empty] | IB Existing | 459800911281 | TDMS V2 M.Det Class B FRU | [10 empty] | No: Yes
|  |  | United States | Philips Medical Capital LLC | 40028 | 4465700% |  | the vertical motion buttons on the ct control box do not work |  | 5 |  | Replace the CT Box | [12 empty] | 119519238 | Closed | 1980773 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | Vertical button on CT Box malfunction. | [5 empty] | IB Existing | 453567027195 | CTBOX ASSEMBLY RAFI/TILTED (RoHS) | [10 empty] | No: Yes
|  |  | China | Jiangxi Province Children's Hospital | 40041 | 4465700% |  | 05k for IB/Ingenuity CT 728326/Maintenance record |  | 3 |  | TBD | [12 empty] | 119514884 | New | 1978387 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | TBD | [6 empty] | TBD | [5 empty] | TBD | bar: No to y; y to y
|  |  | China | Heilongjiang Province Huachuan County People's Hospital | 345018 | 4465700% |  | Wechat：01/Ingenuity Core128/Elite China/Double screen to single screen |  | 3 Zhai, Lisa |  | None | [12 empty] | 119513280 | Closed | 1977326 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | N/A | [6 empty] | Double screen to single screen | [5 empty] | N/A | bar: No to y; y to y
|  |  | Germany | Gesundheit Nord gGmbH Klinikverbund Bremen Klinikum Bremen-Ost Chirurgische Klinik | 40002 | 4465500% |  | MPR reconstruction incomplete |  | WANG, Peter |  | None | [12 empty] | CFM |  | 1976231 |  |  | Investigation | Investigation | CTUDT00632116 | [5 empty] | Engineering | [6 empty] | MPR reconstruction incomplete | [5 empty] | Refresh |  |  |  | System | CIRS-SW | MPR Recon | CIRS-SW-MPR Recon | CIRS-SW-MPR Recon | CTUDT00632116 | MPR Recon | CIRS-SW | CIRS-SW-MPR Recon | No: Yes
|  |  | Viet Nam | Công ty TNHH Thương mại và | 40046 | 4465600% |  | CT BOX got defective |  | 3 |  | Replace the CT Box | [12 empty] | 119508962 | Closed | 1974461 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | CT Box defective | [5 empty] | IB Existing | 459801705022 | CT BOX ASSEMBLY_BEST | [10 empty] | No: Yes
| Date0411U |  | Mexico | DE LAGE LANDEN SA DE CV | 10007 | 4465900% |  | EL EQUIPO ESTA ESTA MUY LENTO PARA LA ADQUISICION DE IMAGEN, PRESENTA CUENTA REGRESIVA PARA EXPLORACIÓN. SE TIENE... |  | 3 |  | Clean the filter, and re-cal EP | [12 empty] | 119532439 | Closed | 1987775 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Image acq speed low | [5 empty] | Refresh |  |  |  | System | CIRS-SW | ReconSlow | CIRS-SW-ReconSlow | CIRS-SW-ReconSlow | TBD-ReconSlow | ReconSlow | CIRS-SW | CIRS-SW-ReconSlow | No: Yes
|  |  | United States | Adventhealth Zephyrhills | 300106 | 4465900% |  | 0611CDY70 Gantry shutoff |  | 3 |  | Replace the dual output PS | [12 empty] | 119531173 | Closed | 1984817 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | gantry shut down | [5 empty] | IB Existing | 453567374153 | POWER SUPPLY 120 VAC DUAL OUTPUT | [10 empty] | No: Yes
|  |  | United States | Lakeland Regional Health Systems | 300143 | 4465800% |  | Parts Needed |  | 3 |  | Deliver new pads and straps to customer. | [12 empty] | 119524606 | Incomplete | 1984484 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Customer order parts. | [5 empty] | IB Existing | 455300002861, 459800033821, 459800042241, 453567009961, 453567111051 | Mouse pad, Flat Cushion, Head/armrest cushion, L-14"PATIENT RESTRAINT STRAP/ 42", HEADREST STRAP ASSY | bar: No to y; y to y
|  |  | United States | Lakeland Regional Health Systems | 320474 | 4465800% |  | Parts Needed |  | 3 |  | Deliver new pads and straps to customer. | [12 empty] | 119524552 | Closed | 1984449 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Customer order parts. | [5 empty] | IB Existing | 455300002861, 459800033821, 459800042241, 453567009961, 453567111051 | Mouse pad, Flat Cushion, Head/armrest cushion, L-14"PATIENT RESTRAINT STRAP/ 42", HEADREST STRAP ASSY | bar: No to y; y to y
| Date0418 |  | United States | UNITY HEALTH SYSTEMS | 40037 | 4466500% |  | Monthly QA |  | 5 |  | N/A | [12 empty] | 119559615 | Closed | 2006339 |  |  | Evaluation | Evaluation | [6 empty] | N/A | [6 empty] | Monthly QA | [5 empty] | IB Existing | bar: No to y; y to y
|  |  | China | Zhalut Banner Mein medical hospital | 345036 | 4466500% |  | 05k for ib/Ingenuity Core128/Elite China/monitor |  | 5 |  | Cancelled, system no error. | [12 empty] | 119558576 | Fixed | 2005650 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A | bar: No to y; y to y
|  |  | China | Zhalut Banner Mein medical hospital | 345036 | 4466500% |  | 05k for ib/Ingenuity Core128/Elite China/Stand by |  | 5 |  | Cancelled, system no error. | [12 empty] | 119558405 | Fixed | 2005525 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A | bar: No to y; y to y
|  |  | China | Shandong Province Chest Hospital | 40025 | 4466500% |  | 05L/Ingenuity CT 728326/blue screen |  | 3 |  | N/A | [12 empty] | 119557528 | Closed | 2004906 |  |  | Closed - Done | Closed - Done | CTUDT00632553 | [5 empty] | Engineering | [6 empty] | Bule screen | [5 empty] | Refresh |  |  |  | System | Console-SW | Blue-Screen | Console-SW-Blue-Screen | Console-SW-Blue-Screen | CTUDT00632553 | Blue-Screen | Console-SW | Console-SW-Blue-Screen | No: Yes
|  |  | Japan | Nikko Memorial Hospital | 341011 | 4466600% |  | CT:APP:AK |  | 5 Xu, Xiaoting |  | None | [12 empty] | 119562703 | Closed | 2004742 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A | bar: No to y; y to y
|  |  | Japan | Nikko Memorial Hospital | 341011 | 4466500% |  | 01DS:KKI:心臓検査にて、バンディングアーチファクト発生 |  | 4 Yang, Cancan |  | Operation mistake, thickness were changed. | [12 empty] | 119561669 | Closed | 2004570 |  |  | Investigation | Investigation | [6 empty] | Customer | [6 empty] | Cardiac testing results in banding artifacts | [5 empty] | Non refresh | bar: No to y; y to y
|  |  | China | Shaanxi Province Xianyang Central Hospital (Xianyang City No.2 People's Hospital) | 40050 | 4466500% |  | Wechat:02N/Ingenuity CT 728326/Controller key failure |  | 3 |  | Replace the CT Box | [12 empty] | 119556787 | Closed | 2004465 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632185 | [5 empty] | Engineering | [6 empty] | CT Box cannot control the patient support movement | [5 empty] | TBD | 459801705022 | CT BOX ASSEMBLY_BEST |  | System | Gantry-HW | CT Box | Gantry-HW-CT Box | Gantry-HW-CT Box | CTUDT00632185 | CT Box | Gantry-HW | Gantry-HW-CT Box | No: Yes
|  |  | United States | UNITY HEALTH SYSTEMS | 40037 | 4466500% |  | Q/O: Foot switch needs adjustment |  | 5 |  | Adjust the footswitch | [12 empty] | 119560819 | Closed | 2003678 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Service | [6 empty] | Footswitch need to be adjusted | [5 empty] | Non refresh | bar: No to y; y to y
|  |  | Japan | Yonago Medical Center | 341015 | 4466500% |  | Slight ring artifact, performed air calibration |  | 4 SHU, Tim |  | Air calibration | [12 empty] | 119557344 | Closed | 2001278 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | Ring artifact | [5 empty] | IB Existing | [12 empty] | No: Yes
|  |  | China | Henan Xuchang City Central Hospital (Infectious Diseases) | 40047 | 4466400% |  | 05K for IB/Ingenuity CT 728326/auto storage setting |  | 3 Cai, Cathy |  | Instructure customer to set the auto storage device. | [12 empty] | 119550210 | Closed | 2000495 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Customer reqest for set the auto storage device | [5 empty] | Non refresh | bar: No to y; y to y
|  |  | United States | Philips Medical Capital LLC | 40028 | 4466300% |  | ext 4203* iq check fail and air calibration (foreign object was found error) |  | 3 |  | Clean the foreign object | [12 empty] | 119545758 | Closed | 1997580 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | IQ Check fail | [5 empty] | IB Existing | [12 empty] | No: Yes
|  |  | China | Henan Xuchang City Central Hospital (Infectious Diseases) | 40047 | 4466300% |  | 05K for IB/Ingenuity CT 728326/ADD PACS |  | 3 Ma, Julita |  | Add node of HIS and PACS | [12 empty] | 119545428 | Closed | 1997331 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | Support connect HIS and PACS | [5 empty] | Non refresh | bar: No to y; y to y
|  |  | United States | Odessa Regional Medical Center | 40029 | 4466300% |  | 1302CDY73 Images not transferring to PACS |  | 3 |  | checked network configurations, pacs and all destinations, no problem found no issues detected | [12 empty] | 119548945 | Closed | 1995673 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | Image transfer to PACS fail | [5 empty] | Refresh |  |  |  | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No: Yes
|  |  | United States | Philips Medical Capital LLC | 40053 | 4466200% |  | equipment is making a squeaking noise when gantry is titled/please refer to 0119495228 |  | 5 |  | Adjusted CT side covers to eliminate noise when tilting | [12 empty] | 119541962 | Closed | 1995243 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | Squeaking noise when gantry tilting | [5 empty] | Refresh |  |  |  | System | Gantry-HW | Noise | Gantry-HW-Noise | Gantry-HW-Noise | TBD-GantryNoise | Noise | Gantry-HW | Gantry-HW-Noise | No: Yes
|  |  | Japan | Tomakomai City Hospital | 341014 | 4466200% |  | SF:寝台カバーの歪み |  | 5 Xu, Xiaoting |  | Replace the BRACKET-PEDESTAL | [12 empty] | 119538282 | Closed | 1992463 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Couch noise | [5 empty] | IB Existing | 459800721902 | BRACKET-PEDESTAL | [10 empty] | No: Yes
|  |  | United States | Mclaren Macomb | 40021 | 4466200% |  | 9799CFXXX need to chg protocol |  | 3 |  | Reset IT password, and introduce how to reset the CT user password. | [12 empty] | 119539383 | Closed | 1989961 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Cannot login the exam card management | [5 empty] | Refresh | bar: No to y; y to y
| Date0418U |  | United States | Lakeland Regional Health Systems | 320474 | 4466500% |  | MP5 monitor mounting bracket needs to be replaced |  | 3 |  | Replace the MP5 bracket | [12 empty] | 119560751 | Closed | 2007115 |  |  | Evaluation | Evaluation | [6 empty] | Engineering | [6 empty] | Need to replace the MP5 install bracket | [5 empty] | Non refresh | 451261019401 | IV-MP5 MET Mounting Adapter | [10 empty] | No: Yes
|  |  | United States | Upson Regional Medical Center | 310115 | 4466500% |  | Remote Alert: CT-PRO-GANTRY_ControlLink |  | 5 |  | Clean the slip ring | [12 empty] | 119560813 | Closed | 2003574 |  |  | Evaluation | Evaluation | [6 empty] | Engineering | [6 empty] | gantry control link alert | [5 empty] | IB Existing | [12 empty] | No: Yes
|  |  | United States | Adventhealth Zephyrhills | 300099 | 4466300% |  | No able to send images to PACS |  | 3 |  | Customer properly seating CAT5 Cable | [12 empty] | 119547377 | Closed | 1997732 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Cannot transfor images to PACS | [5 empty] | Non refresh | [12 empty] | No: Yes
|  |  | United States | Adventhealth Zephyrhills | 300099 | 4466300% |  | Scanner not sending to PACS |  | 3 |  | Reseated Nic cable. And cold started system | [12 empty] | 119546605 | Closed | 1994977 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Cannot transfor images to PACS | [5 empty] | Non refresh | [12 empty] | No: Yes
|  |  | Spain | POLICLINICA NTRA SRA DEL ROSARIO SL | 336190 | 4466300% |  | No deja disparar, error intermitente |  | 3 |  | Replace the anode module | [12 empty] | 119544364 | Closed | 1992306 |  |  | Evaluation | Evaluation | [6 empty] | Engineering | [6 empty] | Anode fail | [5 empty] | IB Existing | 453567560012 | Spellman 80 KW RoHS APM Assy | [10 empty] | No: Yes
| Date0425 |  | Japan | Yonago Medical Center | 341015 | 4467600% |  | 01CS:KK:1件目の患者様撮影開始時に患者情報登録画面が出てこなかった。 |  | 4 |  | None | [12 empty] | 119602286 | Closed | 2030664 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632553 | [5 empty] | Engineering | [6 empty] | When logout / logged in the customer judgment, it seems to have frozen in the blue screen | [5 empty] | Refresh |  |  |  | System | Console-SW | Blue-Screen | Console-SW-Blue-Screen | Console-SW-Blue-Screen | CTUDT00632553 | Blue-Screen | Console-SW | Console-SW-Blue-Screen | No: Yes
|  |  | China | Shandong Province Chest Hospital | 40025 | 4467500% |  | Wechat:05L/Ingenuity CT /Couch cann't up/down |  | 3 |  | 重新连接绝对编码器接线，进行扫描床水平校正 | [12 empty] | 119600669 | Closed | 2030650 |  |  | Opened | Opened | TBD | [5 empty] | Engineering | [6 empty] | Couch cannot move up/down | [5 empty] | Refresh |  |  |  | System | Couch-HW | VerticalDriver | Couch-HW-VerticalDriver | Couch-HW-VerticalDriver | TBD | VerticalDriver | Couch-HW | Couch-HW-VerticalDriver | No: Yes
|  |  | China | Shandong Province Chest Hospital | 40025 | 4466500% |  | System software can't log in. |  | Cai, Cathy |  | None | [12 empty] | CFM |  | 2030325 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632553 | [5 empty] | Engineering | [6 empty] | System software can't log in. blue screen | [5 empty] | Refresh |  |  |  | System | Console-SW | Blue-Screen | Console-SW-Blue-Screen | Console-SW-Blue-Screen | CTUDT00632553 | Blue-Screen | Console-SW | Console-SW-Blue-Screen | No: Yes
|  |  | Poland | Rzadowa Agencja | 341002 | 4467300% |  | Wielokrotne wyłączanie się gantry. |  | 4 |  | Replace the dual output PS and the signal brush | [12 empty] | 119596804 | Closed | 2029435 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Gantry shut down automatically | [5 empty] | IB Existing | POWER SUPPLY 120 VAC DUAL OUTPUT ROHS, Brush, Signal Block Assembly | [11 empty] | No: Yes
|  |  | China | Shaanxi Province Xianyang Central Hospital (Xianyang City No.2 People's Hospital) | 40050 | 4467300% |  | 02N/Ingenuity CT 728326/The image has artifacts |  | 5 |  | TBD | [12 empty] | 119594347 | In Process | 2028495 |  |  | Opened | Opened | [6 empty] | TBD | [6 empty] | Image artifact | [5 empty] | TBD | bar: No to y; y to y
|  |  | China | Shaanxi Ankang T.C.M | 345041 | 4467100% |  | Wechat：02N/Ingenuity Core128/Elite China/Transfer work order |  | 5 |  | N/A | [12 empty] | 119582264 | New | 2020990 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A | bar: No to y; y to y
|  |  | United States | Ascension Seton Medical Center | 40006 | 4467100% |  | 0304CFXXX Parts ID Encoder failure |  | 5 |  | Provided customer with part numbers for the horizontal couch encoders and the gantry speake | [12 empty] | 119585643 | Closed | 2019538 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Ask for the part part number | [5 empty] | Non Refresh | bar: No to y; y to y
|  |  | Chile | Scotiabank Chile | 40008 | 4467100% |  | EL GANTRY NO INICIA |  | 2 |  | Replace the R2D | [12 empty] | 119583336 | In Progress | 2018170 |  |  | Opened | Opened | [6 empty] | Engineering | [6 empty] | System start-up error | [5 empty] | IB Existing | 459800467343 | R2D ASSEMBLY | [10 empty] | No: Yes
|  |  | China | Shandong Province Chest Hospital | 40025 | 4467000% |  | WeChat：05L/Ingenuity CT 728326/Blue screen cannot start |  | 3 |  | None | [12 empty] | 119575391 | Closed | 2016595 |  |  | Evaluation | Evaluation | CTUDT00632553 | [5 empty] | Engineering | [6 empty] | Blue screen after restart host | [5 empty] | Refresh |  |  |  | System | Console-SW | Blue-Screen | Console-SW-Blue-Screen | Console-SW-Blue-Screen | CTUDT00632553 | Blue-Screen | Console-SW | Console-SW-Blue-Screen | No: Yes
|  |  | Japan | Maki Hospital | 341017 | 4467000% |  | Scan aborted |  | 4 ZHANG, Yue Xing |  | None | [12 empty] | 119574541 | Closed | 2016040 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Tube soft arc | [5 empty] | IB Existing | [12 empty] | No: Yes
|  |  | United States | Lackey Memorial Hospital | 40015 | 4466900% |  | 0711CDN85 unit stating that a critical error has occurred |  | 2 Butler, Daria |  | restored the backup and exam cards | [12 empty] | 119572105 | Closed | 2010830 |  |  | Evaluation | Evaluation | [6 empty] | Engineering | [6 empty] | system errors after closing exam cards | [5 empty] | Refresh |  |  |  | System | Console-SW | System-Error | Console-SW-System-Error | Console-SW-System-Error | 0 | Blue-Screen | Console-SW | Console-SW-System-Error | No: Yes
|  |  | Puerto Rico | DORADO HEALTH INC DBA | 40007 | 4466900% |  | Ring artifact in images; Cleaned contrast on the DMS detectors and ran calibrations. |  | 3 XU, Miley |  | contrast liquid was found in multiple detectors, do the deep clean. | [12 empty] | 119571019 | Closed | 2010713 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | Image artifact | [5 empty] | IB Existing | [12 empty] | No: Yes
| Date0425U |  | United States | Adventhealth Zephyrhills | 300099 | 4467522% |  | IT password reset needed |  | 3 |  | The password has no issue when FSE on site. | [12 empty] | 119598751 | Closed | 2030080 |  |  | Opened | Opened | [6 empty] | Customer | [6 empty] | Password need reset | [5 empty] | Refresh | bar: No to y; y to y
|  |  | United States | Adventhealth Zephyrhills | 300099 | 4467511% |  | Add PACS node |  | 3 |  | Add PACS node | [12 empty] | 119598242 | Closed | 2029911 |  |  | Opened | Opened | [6 empty] | Customer | [6 empty] | Add the PACS node | [5 empty] | Non Refresh | bar: No to y; y to y
|  |  | United States | Adventhealth Zephyrhills | 300099 | 4467283% |  | 1301CDN8A Need to add multiple nodes |  | 5 |  | Add the nodes | [12 empty] | 119583154 | Closed | 2021621 |  |  | Evaluation | Evaluation | [6 empty] | Customer | [6 empty] | Need to add multi nodes | [5 empty] | Non Refresh | bar: No to y; y to y
|  |  | United States | Adventhealth Zephyrhills | 300106 | 4467283% |  | 1301CDN8A Need to add multiple nodes |  | 5 |  | Dicom node add remotely. | [12 empty] | 119583138 | Closed | 2021605 |  |  | Evaluation | Evaluation | [6 empty] | Customer | [6 empty] | Need to add multi nodes | [5 empty] | Non Refresh | bar: No to y; y to y
|  |  | United States | Select Medical Corporation | 32261 | 4467191% |  | Yesterday the device had a ring artifact. Today they were unable to perform QA testing. Device won't power on and... |  | 2 Miller, Steve |  | Add bad detector in the compliance list | [12 empty] | 119577235 | Closed | 2017931 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Ring artifact | [5 empty] | IB Existing | [12 empty] | No: Yes
|  |  | United States | Philips Medical Capital LLC | 300057 | 4467002% |  | has a question about info not crossing over to the scanner Requested tech support |  | 5 Dantha, Rajashekar |  | RIS/PACS system is the rootcause, not the CT system. | [12 empty] | 119571779 | Closed | 2010682 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | RIS/PACS | [6 empty] | Infor not corss over to the scanner | [5 empty] | Non Refresh | bar: No to y; y to y
| Date0516 |  | China | Hebei Province Geriatric Hospital | 345015 | 4469400% |  | 01/Ingenuity Core128/Elite China/ct value shift |  |  |  | TBD | [12 empty] | 119705678 | In Process | 2093780 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | CT number not accuracy | [5 empty] | Refresh |  |  |  | System | DMS-SW | CT number | DMS-SW-CT number | CT number-DMS-SW-CT number | 0 | Cover | Couch-HW |  | No: Yes
|  |  | China | Dongguan City Hospital of Integrated Chinese and Western Medicine | 345013 | 4469300% |  | Wechat:05R/Ingenuity Core128/Elite China/Gantry noise |  |  |  | Replace the main driver belt | [12 empty] | 119694749 | Closed | 2089138 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Gantry rotation noise | [5 empty] | IB Existing | 453566492401 | BELT, 5-RIB POLY-VEE, 198"
|  |  | China | Nanchang University No. 1 Affiliated Hospital (Jinggangshan Hospital) | 345002 | 4469200% |  | 01/Ingenuity Core128/Elite China/Query tube exposure record |  | Zhai, Lisa |  | N/A | [12 empty] | 119687327 | Closed | 2084269 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A
|  |  | China | Shaanxi Ankang T.C.M | 345041 | 4469200% |  | 05k for IB/ Ingenuity Core128/Elite China/Repair and maintenance work order |  |  |  | TBD | [12 empty] | 119686970 | New | 2084065 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Engineering | [6 empty] | Routine PM | [5 empty] | Non Refresh
|  |  | China | Shandong Anchu City T.C.M | 345017 | 4469200% |  | 05K For IB/Core128/ct box key pressed error |  |  |  | Replace the CT Box | [12 empty] | 119686924 | In Process | 2084023 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632185 | [5 empty] | Engineering | [6 empty] | CT Box button malfunction | [5 empty] | Refresh | 459801705022 | CT BOX ASSEMBLY_BEST |  | System | Gantry-HW | CT Box | Gantry-HW-CT Box | Gantry-HW-CT Box | CTUDT00632185 | CT Box | Gantry-HW | Gantry-HW-CT Box | No: Yes
|  |  | United States | Lackey Memorial Hospital | 40015 | 4469200% |  | Doesn't know how to Backup Protocols on Exam Cards |  | Das, Sayan |  | Work with customer on loading the EC onto the Thumb Drive to do the store. | [12 empty] | 119692201 | Closed | 2083728 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Customer do not know how to do the EC protocol backup. | [5 empty] | Non Refresh
|  |  | China | Sxi feng County T.C.M. Hospital | 345014 | 4469100% |  | 01/Ingenuity Core128/Elite China/CT box damage |  |  |  | TBD | [12 empty] | 119682291 | In Process | 2080786 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632185 | [5 empty] | Engineering | [6 empty] | CT Box button malfunction | [5 empty] | Refresh |  |  |  | System | Gantry-HW | CT Box | Gantry-HW-CT Box | Gantry-HW-CT Box | CTUDT00632185 | CT Box | Gantry-HW | Gantry-HW-CT Box | No: Yes
|  |  | China | Harbin City Senhai Hospital | 345007 | 4469100% |  | 01/Core128/consulting |  | Zhai, Lisa |  | N/A | [12 empty] | 119680749 | Closed | 2079879 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A
|  |  | Puerto Rico | METRO MAYAGUEZ INC | 40009 | 4469100% |  | It does not expose, gantry can not comply |  | Howard, Amanda |  | N/A. When FSE got to the site, the system was working already. | [12 empty] | 119683764 | Closed | 2077715 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | Arc | [5 empty] | IB Existing
|  |  | Japan | YAMAGATA SAISEI HOSPITAL | 343000 | 4469000% |  | AF:寝台下降時に異音がする |  | Yang, Cancan |  | Adjust the couch cover | [12 empty] | 119673828 | Closed | 2075696 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | NA | [5 empty] | Engineering | [6 empty] | Couch noise | [5 empty] | Refresh |  |  |  | System | Couch-HW | Cover | Couch-HW-Cover | Cover-Couch-HW-Cover | NA | Cover | Couch-HW |  | No: Yes
|  |  | China | Hebei Province Tangshan Workers Hospital | 40044 | 4469000% |  | Wechat：01/Ingenuity CT 728326/gantry abnormal noise |  |  |  | TBD | [12 empty] | 119673359 | In Process | 2075407 |  |  | Evaluation | Evaluation | [6 empty] | Engineering | [6 empty] | Gantry rotation noise | [5 empty] | IB Existing
|  |  | Chile | Scotiabank Chile | 40008 | 4469000% |  | EL GANTRY NO PUEDE INICIAR, LA FALLA REINCIDIO |  |  |  | Reload CGMP software, do calibration and image test | [12 empty] | 119678645 | Closed | 2074804 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | CGMP software corrupted due to sysetm unexpected shut down | [5 empty] | IB Existing
|  |  | China | Harbin City Senhai Hospital | 345007 | 4468800% |  | 01/Ingenuity Core128/Elite China/artifact |  |  |  | TBD | [12 empty] | 119669631 | In Process | 2071983 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Artifact | [5 empty] | IB Existing
|  |  | United States | UNITY HEALTH SYSTEMS | 40037 | 4468700% |  | Monthly QA |  |  |  | None | [12 empty] | 119666980 | In Progress | 2071270 |  |  | Evaluation | Evaluation | [6 empty] | Customer | [6 empty] | Monthly QA | [5 empty] | Non Refresh
|  |  | Japan | Maki Hospital | 341017 | 4468700% |  | SF:Long TC作業 |  | Zhu, Chun |  | Long tube condition | [12 empty] | 119664625 | Closed | 2070323 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | N/A | [6 empty] | None | [5 empty] | IB Existing
|  |  | Germany | Gesundheit Nord gGmbH Klinikverbund Bremen Klinikum Bremen-Ost | 40002 |  |  | Coronal and Sagital MPR were not complet reconstructed |  | Lu, Xueying |  | CTUDT00632116 MPR, in SP2 scope | [14 empty] | 2069481 |  |  | Investigation | Investigation | CTUDT00632116 | [5 empty] | Engineering | [6 empty] | MPR incomplete reconstructed | [5 empty] | Refresh |  |  |  | System | CIRS-SW | MPR Recon | CIRS-SW-MPR Recon | CIRS-SW-MPR Recon | CTUDT00632116 | MPR Recon | CIRS-SW | CIRS-SW-MPR Recon | No: Yes
|  |  | United Kingdom | STOCKPORT NHS FOUNDATION TRUST | 40016 | 4468400% |  | *FSE* Replace bottom front cover. Damage was picked up during a routine service. |  |  |  | Replace the lower front cover | [12 empty] | 119649287 | Closed | 2060215 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Cover damage | [5 empty] | IB Existing | 459801483431 | COVER-FRONT LOWER
|  |  | China | Jilin Province Songwong City Central Hospital | 345008 | 4468400% |  | 01/Ingenuity Core128/Elite China |  | Cai, Cathy |  | None (Customer contact the sales themselves) | [12 empty] | 119646638 | Closed | 2058588 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Artifact | [5 empty] | IB Existing
|  |  | China | Zhalantun City People's Hospital | 40045 | 4468400% |  | wechat:01/Ingenuity CT 728326 |  | Zhai, Lisa |  | None | [12 empty] | 119646627 | Closed | 2058578 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | Account locked | [5 empty] | Refresh |  |  |  | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No: Yes
|  |  | United States | Clay County Medical Center | 40003 | 4468400% |  | Change Philips VPN Encryption to SHA256 |  | Mohanta, Deepranjan |  | N/A | [12 empty] | 119649010 | Closed | 2057325 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A
|  |  | Japan | Kinan Hospital | 340000 | 4468300% |  | SF:システムファントム破損 |  | Zhu, Chun |  | Replace the H phantom | [12 empty] | 119641966 | Closed | 2055336 |  |  | Evaluation | Evaluation | [6 empty] | Engineering | [6 empty] | H phantom broken | [5 empty] | Refresh | 459801608821 | HARMONIZED PHANTOM BODY SECTION-40MM |  | System | Couch-HW | H-Phantom | Couch-HW-H-Phantom | Couch-HW-H-Phantom |  | H-Phantom | Couch-HW | Couch-HW-H-Phantom | No: Yes
|  |  | United States | Clay County Medical Center | 40003 | 4468300% |  | need to change the Encryption of the vpn tunnel |  | Dantha, Rajashekar |  | N/A | [12 empty] | 119646122 | Closed | 2054930 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A
|  |  | China | Jiangxi Province Children's Hospital | 40041 | 4468200% |  | 05K for IB/Ingenuity CT 728326/Reconstruction tank failure |  | Cai, Cathy |  | Ajust the CIRS time | [12 empty] | 119640240 | Closed | 2053025 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | CT image time is incorrect | [5 empty] | Refresh |  |  |  | System | Console-SW | TimeSync | Console-SW-TimeSync | Console-SW-TimeSync |  | TimeSync | Console-SW | Console-SW-TimeSync | No: Yes
|  |  | China | Sxi feng County T.C.M. Hospital | 345014 | 4468000% |  | WeChat：01/Ingenuity Core128/Elite China/CT BOX Key failure |  | Cai, Cathy |  | TBD | [12 empty] | 119633696 | In Process | 2051531 |  |  | Evaluation | Evaluation | CTUDT00632185 | [5 empty] | Engineering | [6 empty] | CT Box couch down button malfunction | [5 empty] | Refresh |  |  |  | System | Gantry-HW | CT Box | Gantry-HW-CT Box | Gantry-HW-CT Box | CTUDT00632185 | CT Box | Gantry-HW | Gantry-HW-CT Box | No: Yes
|  |  | China | Qinghai Province HaibeiZhou No. 2 Hospital | 345029 | 4468000% |  | 01/Ingenuity Core128/Elite China/Abnormal noise of machine |  | Zhai, Lisa |  | TBD | [12 empty] | 119630840 | In Process | 2050604 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Abnormal noise in system | [5 empty] | IB Existing
|  |  | United States | Odessa Regional Medical Center | 40029 | 4468000% |  | Vertical Noisy Couch |  |  |  | Lubricant the couch mechanical system | [12 empty] | 119635055 | Closed | 2049862 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Couch vertical movement noise | [5 empty] | IB Existing
|  |  | Portugal | Centro Hospitalar de Vila Nova | 40030 | 4467900% |  | Erro no reconstrutor |  |  |  | None. FSE on site, the system working fine. | [12 empty] | 119627324 | Closed | 2048272 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Rebuilder problem | [5 empty] | Refresh |  |  |  | System | Console-SW | Rebuilder | Console-SW-Rebuilder | Console-SW-Rebuilder |  | Rebuilder | Console-SW | Console-SW-Rebuilder | No: Yes
|  |  | China | Inner Mongolia Xing'an People's Hospital | 345030 | 4467900% |  | Appliaction/Ingenuity Core128/Elite China |  | Zhai, Lisa |  | N/A | [12 empty] | 119624588 | New | 2046983 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A
|  |  | United States | Philips Medical Capital LLC | 40055 | 4467800% |  | 0711CDN89 Phantom files are full. FSE Vic Donald to be onsite for other issue. |  | Mohanta, Deepranjan |  | TBD | [12 empty] | 119622593 | Incomplete | 2044590 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Host HD full | [5 empty] | Refresh |  |  |  | System | Console-SW | Dispk full | Console-SW-Dispk full | Console-SW-Dispk full |  | Dispk full | Console-SW | Console-SW-Dispk full | No: Yes
|  |  | Japan | Yonago Medical Center | 341015 | 4467900% |  | CT:APP:AK |  | Zhu, Chun |  | Introduce the correct protocol to scan | [12 empty] | 119626458 | Closed | 2043699 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Customer | [6 empty] | None | [5 empty] | Non Refresh
|  |  | Japan | Maki Hospital | 341017 | 4467800% |  | 01CS:KKI:PACSへの画像転送が止まる |  | Xu, Xiaoting |  | None | [12 empty] | 119617884 | Closed | 2042952 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Image transfer to PACS stop | [5 empty] | Refresh |  |  |  | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection |  | Remote Connection | Console-SW | Console-SW-Remote Connection | No: Yes
|  |  | China | Jiangsu University Affiliated Hospital | 345023 | 4467800% |  | 05k/Ingenuity Core128/Elite China/ password missing |  | Zhai, Lisa |  | Guide customer modify the password and add printer | [12 empty] | 119617411 | Closed | 2042670 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Consult for password and add printer | [5 empty] | Refresh |  |  |  | System | NA | NA | NA-NA | NA-NA |  | NA | NA | NA-NA | No: Yes
|  |  | Spain | S.C.S. HOSP. GRAL. DE FUERTEVENTURA | 40020 | 4467800% |  | Problemas en el envio de estudios. Juan Pedro 607941742 - Wendy 928862092 |  |  |  | N/A | [12 empty] | 119619331 | Closed | 2040818 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A
|  |  | China | Jiangxi Province Children's Hospital | 40041 | 4467700% |  | 05K FOR IB/Ingenuity CT 728326/estop open |  | Zhai, Lisa |  | Reboot the CT system | [12 empty] | 119608972 | Closed | 2037869 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Estop Open | [5 empty] | IB Existing
|  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 337000 | 4467700% |  | Gantry fan arızası - yerinde tespit |  |  |  | Replace the fan | [12 empty] | 119613139 | Closed | 2037770 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Gantry fan error | [5 empty] | IB Existing | 459800980091 | BR64 Air Fan Replacement Kit
|  |  | United States | Leesburg Regional Medical Center | 40022 | 4467600% |  | Line/Ring artifacts on images during clinical use; Detector module replaced |  | Sanchez, Josiah |  | Replace detector module | [12 empty] | 119605142 | Closed | 2035217 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | Ring artifact | [5 empty] | IB Existing
|  |  | Japan | Tomakomai City Hospital | 341014 | 4467700% |  | 01CS:KK:17時45分頃HostからTSHSTSRV02に画像転送ができなかった。>>自然復旧したようですがログ確認依頼。関連は不明ですがFaulting application name:... |  | Zhu, Chun |  | None | [12 empty] | 119611808 | Closed | 2035048 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Host image transfer issue | [5 empty] | IB Existing
| Date0516U |  | United States | Wake Med Cary Hospital | 9591 | 4469200% |  | NEMO |  | 3 Dantha, Rajashekar |  | N/A | [12 empty] | 119690502 | Closed | 2085763 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A
|  |  | United States | Wake Med Cary Hospital | 9591 | 4469200% |  | X-ray system can not comply |  | 3 Reiner, Eyal |  | None | [12 empty] | 119691968 | Closed | 2083280 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | XRAY system can not comply | [5 empty] | IB Existing
|  |  | United States | Wake Med Cary Hospital | 9591 | 4468700% |  | NEMO |  | 3 Boddolla, Ramu |  | N/A | [12 empty] | 119668360 | Closed | 2071720 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A
|  |  | United States | Adventhealth Zephyrhills | 300099 | 4468700% |  | Gantry turned off itself and error message occurred, re-seated and cleaned all affected connections |  | 2 Lu, Xueying |  | Constrast inside gantry, clean and reset rotor connectors. | [12 empty] | 119669106 | Closed | 2069794 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | Constrast leakage | [5 empty] | IB Existing
|  |  | United States | Wake Med Cary Hospital | 9591 | 4468700% |  | 0308CFXXX Customer Requesting Support - Parts ID |  | 5 Mohanta, Deepranjan |  | N/A | [12 empty] | 119668177 | Closed | 2069214 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A
|  |  | United States | Adventhealth Zephyrhills | 300106 | 4468600% |  | Customer may have deleted the some DICOM node information accidentally and needs help restore those nodes. FSE required. |  | 1900-01-03 |  | Looks like a policy in the hospital network isn’t allowing to send studies between them | [12 empty] | 119661608 | Closed | 2065670 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Customer | [6 empty] | Dicom node communications | [5 empty] | Refresh
|  |  | United States | Wake Med Cary Hospital | 9591 | 4468400% |  | Table making knocking noise when moving vertically; No correction performed |  | 3 Squires, Ciera |  | None | [12 empty] | 119650530 | Closed | 2057885 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | Couch horizontal movement knock noise | [5 empty] | IB Existing
|  |  | United States | Wake Med Cary Hospital | 9591 | 4468300% |  | NEMO |  | 3 Boddolla, Ramu |  | N/A | [12 empty] | 119643824 | Closed | 2056374 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A
|  |  | United States | Wake Med Cary Hospital | 9591 | 4468300% |  | High Voltage Errors - NEMO - Cathode Power Module & Control PCA |  | 3 Furtado, Peter | [14 empty] | 119643586 | Closed | 2053605 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | Grid feedback imbalance | [5 empty] | IB Existing | 453567028521; 453566502911 | - CH/* CATHODE POWER MODULE; - CHPRINTED CIRCUIT CNTRL BD
| Date0523 |  | China | Nanjing Medical University Affiliated Yif Hospital | 345028 | 4470200% |  | WeChat：05K/Ingenuity Core128/Elite China/Image artifact |  | 3 |  | Isolate the bad detector and do air-cal | [12 empty] | 119750638 | Closed | 2116980 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Image artifact | [5 empty] | IB Existing
|  |  | United States | UNITY HEALTH SYSTEMS | 40037 | 4470100% |  | cluncking noise during rotation |  | 2 |  | Re-align the inner ring cover | [12 empty] | 119750421 | Fixed | 2116931 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Cluncking noise during rotation | [5 empty] | IB Existing
|  |  | China | 鄄 City Ma'aiyun Women and Children's Hospital | 345044 | 4470100% |  | 05K for IB/Ingenuity Core128/Elite China/mouse damage |  | 3 |  | Replace the mouse | [12 empty] | 119746432 | Closed | 2115795 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Mouse broken | [5 empty] | IB Existing
|  |  | China | Fujian Province Medical University Affiliated United Hospital | 345010 | 4470100% |  | 05K for IB/Ingenuity Core128/Elite China/panel damage |  | 3 |  | Replace the control panel | [12 empty] | 119746266 | Closed | 2115726 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Panel damage | [5 empty] | IB Existing | 459800484511 | CONTROL PANEL ASSEMBLY-FRONT RH
|  |  | China | Fujian Province Medical University Affiliated United Hospital | 345010 | 4470100% |  | 05k for IB/Ingenuity Core128/Elite China/To connect to the Internet |  | 3 |  | N/A | [12 empty] | 119745619 | Closed | 2115470 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A
|  |  | United States | Philips Medical Capital LLC | 40028 | 4470000% |  | Need help exporting files |  | 5 |  | N/A | [12 empty] | 119742823 | Closed | 2113803 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A
|  |  | Japan | Tomakomai City Hospital | 341014 | 4470000% |  | 01CS:KKI:右側の画面にてDirectoryが表示されない＞＞CTホストの再起動にて復旧 |  | 4 Yang, Cancan |  | Restart the host computer | [12 empty] | 119737839 | Closed | 2110967 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Directory is not displayed on the right screen | [5 empty] | Refresh |  |  |  | System | Console-HW | Mornitor | Console-HW-Mornitor | Console-HW-Mornitor |  | Mornitor | Console-HW | Console-HW-Mornitor | No: Yes
|  |  | China | Jiangsu Province No. 2 T.M. Hospital | 345027 | 4470000% |  | Wechat：05K/Core128/image fault |  | 3 |  | Re-install the SW | [12 empty] | 119737788 | Closed | 2110922 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | Quick film malfunction | [5 empty] | Refresh |  |  |  | System | Console-SW | Film | Console-SW-Film | Console-SW-Film |  | Film | Console-SW | Console-SW-Film | No: Yes
|  |  | China | Jiangsu Province No. 2 T.M. Hospital | 345027 | 4469900% |  | Wechat：05K/Core128/No image |  | 3 WANG LIN, LV |  | Guide customer use ISP or execute 1 serial once for film | [12 empty] | 119732920 | Closed | 2107756 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | No image in film view | [5 empty] | Refresh |  |  |  | System | Console-SW | Film | Console-SW-Film | Console-SW-Film |  | Film | Console-SW | Console-SW-Film | No: Yes
|  |  | China | Zhalantun City People's Hospital | 40045 | 4469900% |  | Wechat：01/Ingenuity CT/Consult |  | 3 Ma, Julita |  | N/A | [12 empty] | 119732077 | Closed | 2107338 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A
|  |  | China | Beijing City Hospital of Integrated Chinese and Western Medicine | 40032 | 4469900% |  | 05K for IB/Ingenuity CT 728326/problem not confirm |  | 3 Zhai, Lisa |  | FSE contact with customer, state the system is workin gnormal. | [12 empty] | 119731038 | Closed | 2106637 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A
|  |  | Spain | S.C.S. HOSP. GRAL. DE FUERTEVENTURA | 40020 | 4469900% |  | Tiene la colchoneta y almohadilla rotas |  | 3 | [14 empty] | 119732364 | Closed | 2103292 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Cushion broken | [5 empty] | Refresh | 459800126661 | Cushion-47.24 Inch Mattress Pad |  | System | Couch-HW | Cushion broken | Couch-HW-Cushion broken | Couch-HW-Cushion broken |  | Cushion broken | Couch-HW | Couch-HW-Cushion broken | No: Yes
|  |  | China | Fujian Province Medical University No. 1 Affiliated Hospital | 345012 | 4469800% |  | SA:05K for IB/Ingenuity Core128/Elite China |  | 3 Zhai, Lisa |  | Guide customer to modify the password. | [12 empty] | 119724396 | Closed | 2102120 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | Account locked | [5 empty] | Refresh |  |  |  | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No: Yes
|  |  | China | Shaanxi Ankang T.C.M | 345041 | 4469800% |  | 05k for IB/Ingenuity Core128/Elite China/Does not automatically rebuild the image |  | 3 |  | None | [12 empty] | 119724336 | Closed | 2102069 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | Does not rebuild the image | [5 empty] | Refresh |  |  |  | System | Console-SW | Image recon | Console-SW-Image recon | Console-SW-Image recon |  | Image recon | Console-SW | Console-SW-Image recon | No: Yes
|  |  | China | Hebei Province Tangshan Workers Hospital | 40044 | 4469700% |  | 01/Ingenuity CT 728326/buy parts |  | 3 |  | TBD | [12 empty] | 119719634 | New | 2098933 |  |  | Evaluation | Evaluation | [6 empty] | TBD | [6 empty] | TBD | [5 empty] | TBD
|  |  | China | Shandong Province Chest Hospital | 40025 | 4469700% |  | Wechat：01/Ingenuity CT 728326/equipment failure |  | 3 |  | TBD | [12 empty] | 119719131 | In Process | 2098624 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Tube too hot, and cannot cool down | [5 empty] | IB Existing
|  |  | China | Henan Xuchang City Central Hospital (Infectious Diseases) | 40047 | 4469700% |  | 05K For IB/ Ingenuity CT/Image fault |  | 5 |  | TBD | [12 empty] | 119718022 | New | 2097894 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | first 2 slice show left half image, last 2 slice show right half image. | [5 empty] | TBD |  |  |  | System | Console-SW | Image recon | Console-SW-Image recon | Console-SW-Image recon |  | Image recon | Console-SW | Console-SW-Image recon | No: Yes
|  |  | China | Beijing City Hospital of Integrated Chinese and Western Medicine | 40032 | 4469700% |  | Scan failed with gantry cannot comply error, replaced Resolver to Digital Assembly |  | 3 ZHANG, Yue Xing |  | Replace R2D. | [12 empty] | 119717557 | Closed | 2097550 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | ACQ not OK | [5 empty] | IB Existing |  | R2D
| Date0523U |  | Spain | POLICLINICA NTRA SRA DEL ROSARIO SL | 336190 | 4470000% |  | Uno de los laser no se enciende |  | 3 |  | Replace Laser | [12 empty] | 119740345 | Closed | 2108759 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Laser not on | [5 empty] | IB Existing | 453567033441 | LASER ASSY, RH FRONT COVER
|  |  | United States | Wake Med Cary Hospital | 9591 | 4469700% |  | NEMO |  | 3 |  | N/A | [12 empty] | 119723011 | Closed | 2101400 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A
|  |  | United States | Adventhealth Zephyrhills | 300099 | 4469700% |  | 0599CDY72 Unable to turn gantry on |  | 2 Boddolla, Ramu |  | Replace the victor PS and the brush | [12 empty] | 119723253 | Closed | 2096961 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [13 empty] | Unable to turn gantry on | [5 empty] | IB Existing | 453567374153 | Victor PS, Brush block
|  |  | United States | Adventhealth Zephyrhills | 300099 | 4469700% |  | Q/O |  | 3 |  | N/A | [12 empty] | 119723118 | Fixed | 2096877 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A
| Date0530 |  | China | Jiangsu Southeast University Affiliated China University Hospital | 345021 | 4470900% |  | SA:05K/Ingenuity Core128/Elite China/Add network node |  | 3 |  | Add the network node | [12 empty] | 119796474 | Closed | 2139615 |  |  | Opened | Opened | [6 empty] | Customer | [6 empty] | Customer ask for add network node | [5 empty] | Non Refresh
|  |  | United Kingdom | STOCKPORT NHS FOUNDATION TRUST | 40016 | 4470800% |  | images haven’t reconstructed on, we have done them from the raw data |  | 3 |  | TBD | [12 empty] | 119790751 | In Progress | 2138101 |  |  | Evaluation | Evaluation | [6 empty] | Engineering | [6 empty] | Image not reconstructed | [5 empty] | Refresh |  |  |  | System | Console-SW | Image recon | Console-SW-Image recon | Console-SW-Image recon |  | Image recon | Console-SW | Console-SW-Image recon | No: Yes
|  |  | China | Dongguan Donghua Hospital Co., Ltd | 40059 | 4470800% |  | WeChat：SA:05K FOR IB/Ingenuity CT /Reconstruct image |  | 3 |  | Add RAM | [12 empty] | 119789082 | Closed | 2137518 |  |  | Opened | Opened | [6 empty] | Engineering | [6 empty] | Image reconstruction slow | [5 empty] | Refresh |  |  |  | System | Console-SW | Image recon | Console-SW-Image recon | Console-SW-Image recon |  | Image recon | Console-SW | Console-SW-Image recon | No: Yes
|  |  | United States | Philips Medical Capital LLC | 40005 | 4470800% |  | advising on product alert letter 9599CFXXX |  | 5 |  | N/A | [12 empty] | 119794227 | Closed | 2137110 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A
|  |  | Japan | Yonago Medical Center | 341013 | 4470800% |  | 05CS:KK:検査室内の除湿器が排水されず満水になっている。>>装置には影響ないことから、そのままにしていただき、後日FSEが確認となりました。 |  | 4 Xu, Xiaoting |  | N/A | [12 empty] | 119791065 | Closed | 2134847 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A
|  |  | China | 山东省胸科医院 | 40025 | 4470700% |  | 01/Ingenuity CT 728326/add printer |  | 3 |  | N/A | [12 empty] | 119777665 | Closed | 2134230 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A
|  |  | China | Ning'an City No. 2 People's Hospital | 345039 | 4470700% |  | 01/Ingenuity Core128/add the printer |  | 3 |  | TBD | [12 empty] | 119776203 | In Process | 2133996 |  |  | Opened | Opened | [6 empty] | Customer | [6 empty] | Add the printer | [5 empty] | Non Refresh
|  |  | Portugal | Centro Hospitalar de Vila Nova | 40030 | 4470400% |  | Com barulhos anormais vindos da gantry |  | 3 |  | TBD | [12 empty] | 119756950 | New | 2122651 |  |  | Evaluation | Evaluation | [6 empty] | Engineering | [6 empty] | Gantry abnormal noise | [5 empty] | IB Existing
|  |  | Spain | S.C.S. HOSP. GRAL. DE FUERTEVENTURA | 40020 | 4470400% |  | Se cambió uno de los sensores pero el problema continua, no funciona |  | 2 WANG LIN, LV |  | Replace the module, do the cal. | [12 empty] | 119755457 | In Progress | 2118388 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Artifact | [5 empty] | IB Existing
| Date0530U |  | United States | Upson Regional Medical Center | 333009 | 4470900% |  | After Power Outage/Unable to Scan |  | 2 |  | Customer restart the entire system | [12 empty] | 119795620 | Closed | 2137815 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Unable to scan | [5 empty] | IB Existing
|  |  | United States | Adventhealth Zephyrhills | 300099 | 4470700% |  | 1301CDY73 Network changes needed |  | 3 |  | N/A | [12 empty] | 119782475 | Closed | 2132768 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Need to change network | [5 empty] | IB Existing
|  |  | United States | Adventhealth Zephyrhills | 300099 | 4470400% |  | Gantry will not turn on. |  | 2 |  | Repalce the main fused, inspection brush block, test syttem. | [12 empty] | 119759920 | Closed | 2120325 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Gantry cannot turn on | [5 empty] | IB Existing
|  |  | United States | Wellstar Kennestone Hospital | 68117 | 4470400% |  | CIRS isn't powering on; Customer resolved |  | 2 |  | Customer had a power outage. Ask customer press the start contactor on the wall. | [12 empty] | 119756967 | Closed | 2118801 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Customer | [6 empty] | Cannot power on the system | [5 empty] | IB Existing
Formulas:
Check-Level3-Type = [Level3]
Note on K2: Shi, Brook: 1. Frequency 2. Effect 3. Safety related
Note on AG2: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: blank cells with light yellow
Note on BN2: Wang, Jing : gap in light yellow and marked with 'n' Fiill the cell with the value extracted from Field Issue-Service
Note on BO2: Wang, Jing : gap in light yellow and marked with 'n'
Note on BP2: Wang, Jing : gap in light yellow and marked with 'n'
Note on BQ2: Wang, Jing : gap in light yellow and marked with 'n'
Note on BR2: Wang, Jing : gap in light yellow and marked with 'n'
Note on BS2: Wang, Jing : refresh 3-month rolling manually by marking June, July, August as Yes (do not refer to service data)
Note on BT2: Wang, Jing : compared BS and BM and mark the gap in light yellow and n
Note on BY2: Wang, Jing : COMPARED WITH AM
Note on BZ2: Wang, Jing : Pay attention to the value '#N/A': the gap is special that serviice provided a formula to get the value, yet the SE directly presents the value in text format.

## Engineer-Pivot
PR Status | (Multiple Items)
IB Exist/New Found | (Multiple Items)
3MonthRolling | (All)
CHECK
Row Labels | Count of PR Status | Count of PR Status2
Engineering | 898 | 66.77%
Others | 123 | 75.91%
Customer | 112 | 84.24%
Application | 64 | 89%
TBD | 63 | 93.68%
Service | 56 | 97.84%
Operation | 24 | 99.63%
NAC | 5 | 100%
Grand Total | 1345
all pivot data is from column AE to BM in sheet 'Field IssueSE'. Previous Version (Ingenuity CT 5000 FOK tracker_Final version-Jul2021-005) showed that the data is about 716 in total for level 0 with same filter attributes.
[18 empty] | Console-SW | 24
PR Status | (Multiple Items) | [16 empty] | TBD | 7
IB Exist/New Found | (Multiple Items) | [16 empty] | CIRS-SW | 6
3MonthRolling | (All) | [16 empty] | Gantry-HW | 5
Funct. Owner | Engineering | [16 empty] | Couch-HW | 4
Level1 | (All) | [16 empty] | Console-HW | 3
[18 empty] | Grand Total | 49
Row Labels | Count of PR Status | Count of PR Status2
Console-SW | 282 | 31.4%
Gantry-HW | 182 | 51.67%
Couch-HW | 97 | 62.47%
IC-HW | 49 | 67.93%
DMS-HW | 46 | 73.05%
Console-HW | 40 | 77.51%
NotEngineering | 39 | 81.85% | [15 empty] | TB Confirm | 7 | 0.14286 | [6 empty] | Console-SW | 21 |  | HW | 5 | Console-SW | 21
Couch-SW | 34 | 85.63% | [15 empty] | SW | 30 | 0.61224 | [6 empty] | Gantry-SW | 5 |  | SW | 30 | TBD | 9
CIRS-SW | 34 | 89.42% | [15 empty] | HW | 12 | 0.2449 | [6 empty] | Console-HW | 4 |  | NOTENGINEERING | 4 | Gantry-SW | 5
Gantry-SW | 26 | 92.32% | [16 empty] | 49 | [7 empty] | Couch-HW | 4 |  | TBCONFIRM | #REF! | NotEngineering | 4
TBD-Bugreport | 23 | 94.88% | [24 empty] | CIRS-SW | 4 |  | TOTAL | #REF! | Console-HW | 4
TBD-NoEnoughInfo | 21 | 97.22% | [24 empty] | Gantry-HW | 3 |  |  |  | Couch-HW | 4
DMS-SW | 15 | 98.89% | [24 empty] | DMS-HW | 1 |  |  |  | CIRS-SW | 4
TBD | 9 | 99.89% | [29 empty] | Gantry-HW | 3
NA | 1 | 1 | [29 empty] | TBD-NoEnoughInfo | 1
Grand Total | 898 | [30 empty] | DMS-HW | 1
PR Status | (Multiple Items) | [20 empty] | Resolve | Ongoing | Out |  |  |  | Resolve | Ongoing | NoEnoughInfo | Out |  |  |  |  | Resolve | Ongoing | NoEnoughInfo | Out
IB Exist/New Found | (Multiple Items) | [19 empty] | Console-SW | 14 | 10 | 0 | 24 | 34 | Gantry-HW | 23 | 5 | 3 | 2 | 33 |  | Gantry-HW-Belt | 6 | 6
Funct. Owner | Engineering | [19 empty] | CIRS-SW |  | 4 | 2 | 6 | 8 | Couch-HW | 15 |  |  | 4 | 19 |  | Gantry-HW-CT Box | 7 | 7
3MonthRolling | Yes | [19 empty] | Gantry-SW |  |  |  | 0 | 3 | IC-HW | 9 |  |  | 4 | 13 |  | Gantry-HW-R2D | 4 | 4
Level1 | system | [19 empty] | Couch-SW |  |  |  | 0 |  | DMS-HW |  |  |  | 6 | 6 |  | Gantry-HW-Noise | 4 | 4
Level2 | (All) | [19 empty] | DMS-SW |  |  |  | 0 |  | console-hw | 3 | 6 |  |  | 9 |  | Gantry-HW-TBD-NoEnoughInfo | 2 |  |  | 2
[21 empty] | Total | 14 | 14 | 2 | 30 | [8 empty] | Gantry-HW-SlipRingClean | 2 |  |  |  | 2
Row Labels | Count of PR Status | Count of PR Status2 | [19 empty] | 46.67% | 46.67% | 6.667% | [9 empty] | Gantry-HW-Lexan Ring | 2 | 2
TBD-NoEnoughInfo | 6 | 12.24% | [19 empty] | 47 | 47 | 7 | [9 empty] | Gantry-HW-Sheave | 2 |  |  | 2
Console-SW-Blue-Screen | 5 | 22.45% | [14 empty] | 1200% |  |  |  | Rolling 3 months SW Issue Status (Total 30; 47% resolved; 47% no enough info; 7% out of Refresh) | [12 empty] | Gantry-HW-Other | 2 |  |  | 2
Console-SW-NoEnoughInfo | 3 | 28.57% | [14 empty] | ss12 | [18 empty] | 23 | 0 | 6 | 2
Console-SW-User Account | 2 | 32.65%
Console-SW-Remote Connection | 2 | 36.73% | [33 empty] | Resolve | Ongoing | NoEnoughInfo | Out
Console-SW-Yellow-Screen | 2 | 40.82% | [31 empty] | Console-SW-Yellow-Screen | 6 | 6
Console-HW-Monitor | 2 | 44.9% |  |  | Remote Connection | Yellow-Screen | DiskFull | Blue-Screen | Performance | Service Tool | Lan Config | Exam Card | Dose SR | HF3 | ResultSelectionJump | Imagemissng-printer | [17 empty] | Console-SW-Remote Connection | 5 | 5
CIRS-SW-MPR Recon | 2 | 48.98% |  |  | 84 | 37 | 10 | 9 | 7 | 4 | 3 | 3 | 3 | 2 | 2 | 2 | [17 empty] | Console-SW-Blue-Screen | 5 | 5
Console-SW-DiskFull | 2 | 53.06% | [31 empty] | Console-SW-Performance | 3 |  | 3
Gantry-HW-CT Box | 2 | 57.14% | [31 empty] | Console-SW-Imagemissing-Printer | 2 |  | 2
TBD-NotEngineering | 1 | 59.18% |  |  | Belt | 26 |  | Yellow-Screen | 16 |  | Lubricant | 11 |  | Lubricant | 14 | [18 empty] | Console-SW-DiskFull | 2 | 2
Couch-HW-HeadRest | 1 | 61.22% |  |  | CT Box | 24 |  | Remote Connection | 14 |  | Vertical spring | 6 |  | Motor Coupling | 8 | [18 empty] | Console-SW-WorklistUsage | 1 |  | 1
Console-SW-System-Error | 1 | 0.63265 |  |  | Gantry Panel | 11 |  | Bugrep | 4 |  | Motor Coupling | 6 |  | Relay | 7 | [18 empty] | Console-SW-NotDefect | 1 |  |  |  | 1
Couch-HW-H-Phantom | 1 | 65.31% |  |  | Lexan Ring | 18 |  | HF3 | 2 |  | AssemblyAdjust | 3 |  | Vertical spring | 6 | [18 empty] | Console-SW-Bugrep | 1 |  |  | 1
Console-SW-Installation | 1 | 67.35% |  |  | GantryLaser | 9 |  | DiskFull | 2 |  | VerticalTransducer | 3 |  | VerticalTransducer | 5 | [20 empty] | 18 | 6 | 1 | 1
Couch-HW-VerticalDriver | 1 | 69.39% |  |  | SlipRingClean | 8 |  | ResultSelectionJump | 2 |  | Horizontal linkage | 3 |  | TBD-bugrep | 3
Console-SW-Not a defect | 1 | 0.71429 |  |  | Noise | 8 |  | TBD-BlueScreen | 1 |  | Phantom | 3 |  | Horizontal linkage | 3
Gantry-HW-Control Panel | 1 | 73.47% |  |  | R2D | 7 |  | SerielsJump | 1 |  | Relay | 3 |  | Phantom | 3
Console-SW-BolusTracking | 1 | 75.51% |  |  | GMP | 5 |  | CTA | 1 |  | TBD-bugrep | 2 |  | Encoder | 2
Console-SW-Time sync | 1 | 0.77551 |  |  | FlatSpring | 4 |  | TBD-Performance | 1 |  | Arm Rest | 2 |  | TabSwitch | 2
Console-SW-InjectorLose | 1 | 79.59% |  |  | PowerOff | 3 |  | NotDefect | 1 |  | TabSwitch | 2 |  | AssemblyAdjust | 2
CIRS-SW-ACQ_NOT_OK | 1 | 81.63% |  |  | Injector | 3 |  | EC Upgrade | 1 |  | Mattress scan line | 2 |  | Horizontal motor | 2
CIRS-SW-ReconNotInitialize | 1 | 83.67% |  |  | USBExtended | 3 |  | FilmStuck | 1 |  | Horizontal motor | 2 |  | Noise | 2
Gantry-HW-Noise | 1 | 85.71% |  |  | UPS | 2 |  | Biopsy | 1 |  | Cover | 2 |  | Mattress scan line | 2
Console-SW-LocalDBLost | 1 | 87.76% |  |  | CCT pedal | 2 |  | AuditRepository | 1 |  | RFC | 2 |  | Arm Rest | 2
CIRS-SW-ReconSlow | 1 | 89.8% | [8 empty] | Noise | 2 |  | Cover | 2
Console-SW-NotEnoughInformation | 1 | 0.91837 | [11 empty] | RFC | 2
Console-HW-GPU | 1 | 93.88%
CIRS-SW-NoEnoughInfo | 1 | 95.92%
Couch-HW-Cleaning | 1 | 97.96%
Gantry-HW-IBox | 1 | 100%
Grand Total | 49 |  |  |  | Belt | 6 |  | Yellow-Screen | 6 |  | VerticalTransducer | 2
[5 empty] | CT Box | 5 |  | Remote Connection | 5 |  | AssemblyAdjust | 2
[5 empty] | R2D | 4 |  | Blue-Screen | 5 |  | Lubricant | 2
[5 empty] | Noise | 4 |  | Performance | 3 |  | Phantom | 1
[5 empty] | TBD-NoEnoughInfo | 2 |  | Imagemissing-Printer | 2 |  | Vertical spring | 1
[5 empty] | SlipRingClean | 2 |  | DiskFull | 2 |  | Table Cover | 1
[5 empty] | Lexan Ring | 2 |  | WorklistUsage | 1 |  | VelocityEncoder | 1
[5 empty] | Sheave | 2 |  | NotDefect | 1 |  | VerticalEncoder | 1
[5 empty] | Other | 2 |  | Bugrep | 1 |  | AMC | 1
[11 empty] | TabSwitch | 1
[11 empty] | TBD-bugrep | 1
[5 empty] | Remote Connection | 83 |  | Remote Connection | 88
[5 empty] | Yellow-Screen | 37 |  | Yellow-Screen | 38
[5 empty] | DiskFull | 9 |  | Blue-Screen | 11
[5 empty] | Blue-Screen | 8 |  | DiskFull | 11
[5 empty] | Performance | 7 |  | NoEnoughInfo | 7
[5 empty] | Service Tool | 4 |  | Performance | 7
[5 empty] | ExamCard | 3 |  | Lan Config | 5
[5 empty] | Lan Config | 3 |  | Dose SR | 4
[5 empty] | Dose SR | 3 |  | Service Tool | 4
[5 empty] | HF3 | 2 |  | ResultSelectionJump | 3
[5 empty] | ResultSelectionJump | 2 |  | ExamCard | 3
[5 empty] | TBD-NoEnoughInfo | 2 |  | ShouldbeCustomer | 3
[5 empty] | Console-SW-TBD-bugrep | 2 |  | Console-SW-TBD-NoEnoughInfo | 2
[5 empty] | Console-SW-Imagemissing-Printer | 2 |  | Console-SW-HF3 | 2
[8 empty] | Console-SW-Imagemissing-Printer | 2
[8 empty] | Console-SW-TBD-bugrep | 2
[5 empty] | Console-SW-Blue-Screen | 5 |  | CIRS-SW-MPR Recon | 2
[5 empty] | Console-SW-NoEnoughInfo | 3 |  | CIRS-SW-ReconNotInitialize | 1
[5 empty] | Console-SW-Yellow-Screen | 2 |  | CIRS-SW-ReconSlow | 1
[5 empty] | Console-SW-User Account | 2 |  | CIRS-SW-ACQ_NOT_OK | 1
[5 empty] | Console-SW-Remote Connection | 2 |  | CIRS-SW-NoEnoughInfo | 1
[5 empty] | Console-SW-DiskFull | 2
[5 empty] | Console-SW-LocalDBLost | 1
[5 empty] | Console-SW-NotEnoughInformation | 1
[5 empty] | Console-SW-Not a defect | 1
[5 empty] | Console-SW-InjectorLose | 1
[5 empty] | Console-SW-System-Error | 1
[5 empty] | Console-SW-Installation | 1
[5 empty] | Console-SW-Time sync | 1
[5 empty] | Console-SW-BolusTracking | 1
TBD-NoEnoughInfo | 8
Console-SW-NoEnoughInfo | 5
Console-SW-Remote Connection | 4
Console-HW-Monitor | 3
Gantry-SW-NoEnoughInfo | 3
Console-SW-Yellow-Screen | 2
NotEngineering-ParameterSetting | 2
CIRS-SW-NoEnoughInfo | 2
Console-SW-DiskFull | 2
Console-SW-User Account | 2
Console-SW-Blue-Screen | 2
NotEngineering-DMC | 1
Gantry-SW-TubTooHot | 1
NotEngineering-SeemsHWProblem | 1
Console-SW-BolusTracking | 1
Gantry-HW-IBox | 1
Couch-HW-HeadRest | 1
DMS-HW-Instruction | 1
Couch-HW-Cleaning | 1
Couch-HW-CushionBreakdown | 1
Couch-HW-H-Phantom | 1
Gantry-SW-TubTooHeat | 1
Console-HW-GPU | 1
Gantry-HW-CT Box | 1
Console-SW-Installation | 1
Console-SW-InjectorLose | 1
TBD-NoEnoughInfo-TBD-NoEnoughInfo | 1
CIRS-SW-ReconNotInitialize | 1
TBD-NotEngineering | 1
Console-SW-LocalDBLost | 1
CIRS-SW-ACQ_NOT_OK | 1
Gantry-HW-NoEnoughInfo | 1
Rolling 3 Month Analysis Repeated Issues (40 of 56) | 40 | 56 | 0.71429
type | PR#
TBD-NoEnoughInfo | 8 |  | TBD-NoEnoughInfo | 8 | 14.29%
Console-SW-NoEnoughInfo | 5 |  | Console-SW-NoEnoughInfo | 5 | 23.21%
Console-SW-Remote Connection | 4 |  | Console-SW-Remote Connection | 4 | 30.36%
Console-HW-Monitor | 3 |  | Console-HW-Monitor | 3 | 35.71%
Gantry-SW-NoEnoughInfo | 3 |  | Gantry-SW-NoEnoughInfo | 3 | 41.07%
Console-SW-Yellow-Screen | 2 |  | Console-SW-Yellow-Screen | 2 | 44.64%
NotEngineering-ParameterSetting | 2 |  | NotEngineering-ParameterSetting | 2 | 48.21%
CIRS-SW-NoEnoughInfo | 2 |  | CIRS-SW-NoEnoughInfo | 2 | 51.79%
Console-SW-DiskFull | 2 |  | Console-SW-DiskFull | 2 | 55.36%
Console-SW-User Account | 2 |  | Console-SW-User Account | 2 | 58.93%
Console-SW-Blue-Screen | 2 |  | Console-SW-Blue-Screen | 2 | 62.5%
Gantry-SW-TubTooHot | 2 |  | NotEngineering-DMC | 1 | 64.29%
Console-SW-BolusTracking | 1 |  | Gantry-SW-TubTooHot | 1 | 66.07%
Gantry-HW-CT Box | 1 |  | NotEngineering-SeemsHWProblem | 1 | 67.86%
Console-SW-InjectorLose | 1 |  | Console-SW-BolusTracking | 1 | 69.64%
|  |  | Gantry-HW-IBox | 1 | 71.43%
|  |  | Couch-HW-HeadRest | 1 | 73.21%
|  |  | DMS-HW-Instruction | 1 | 75%
|  |  | Couch-HW-Cleaning | 1 | 76.79%
|  |  | Couch-HW-CushionBreakdown | 1 | 78.57%
|  |  | Couch-HW-H-Phantom | 1 | 80.36%
|  |  | Gantry-SW-TubTooHeat | 1 | 82.14%
|  |  | Console-HW-GPU | 1 | 83.93%
|  |  | Gantry-HW-CT Box | 1 | 85.71%
|  |  | Console-SW-Installation | 1 | 87.5%
|  |  | Console-SW-InjectorLose | 1 | 89.29%
|  |  | TBD-NoEnoughInfo-TBD-NoEnoughInfo | 1 | 91.07%
|  |  | CIRS-SW-ReconNotInitialize | 1 | 92.86%
|  |  | TBD-NotEngineering | 1 | 94.64%
|  |  | Console-SW-LocalDBLost | 1 | 96.43%
|  |  | CIRS-SW-ACQ_NOT_OK | 1 | 98.21%
|  |  | Gantry-HW-NoEnoughInfo | 1 | 100%
Funct. Owner | Engineering
Count of # | Column Labels
Row Labels | Remote Connection | Remote Connection-1602 | Remote Connection-1602TBD | Remote Connection-Broken | Remote Connection-Not a Defect | Remote Connection-Worklist | Grand Total
40020 | 1 |  | 3 |  | 3 |  | 7
345005 | 3 |  |  |  |  | 1 | 4
40007 | 4 | [5 empty] | 4
40010 |  |  | 1 |  | 2 | 1 | 4
345025 | 3 | [5 empty] | 3
40024 | 3 | [5 empty] | 3
345001 | 2 | [5 empty] | 2
40019 |  |  |  | 2 |  |  | 2
40015 | 2 | [5 empty] | 2
345017 |  |  | 2 |  |  |  | 2
40002 | 1 | [5 empty] | 1
40021 | 1 | [5 empty] | 1
341000 | 1 | [5 empty] | 1
345014 | 1 | [5 empty] | 1
40006 | 1 | [5 empty] | 1
40023 |  |  |  |  | 1 |  | 1
40031 |  | 1 |  |  |  |  | 1
40029 | 1 | [5 empty] | 1
40005 | [5 empty] | 1 | 1
40030 | 1 | [5 empty] | 1
40025 | 1 | [5 empty] | 1
Grand Total | 26 | 1 | 6 | 2 | 6 | 3 | 44
IB Exist/New Found | (Multiple Items)
PR Status | (Multiple Items)
Level1 | system
Funct. Owner | (Multiple Items)
Row Labels | Count of PR Status
CTUDT00632127 | 81
NoEnoughInfo | 53
CTUDT00632110 | 38
CTUDT00630085 | 33
NotEngineering | 31
Not A Defect | 28
ECR-077012 | 25
CTUDT00631367 | 24
CTUDT00632185 | 24
TBD_Bugreport | 22
CTUDT00632553 | 19
ECR-075975 | 16
TBD-Translation | 16
(blank) | 16
CTUDT00632520 | 16
TBD-NoEnoughInfo | 15
CTUDT00630138/630160 | 14
E-125059 | 14
CTUDT00632261 | 13
CTUDT00632173 | 13
CTUDT00632116 | 12
NotEngineering-partreplace | 9
TBD_OBS | 8
NotEngineering-SlipringClean | 8
CTUDT00632179 | 8
CTUDT00632181 | 8
TBD_Tube hot_bugreport | 7
CTUDT00631602 | 7
TBD-APM/CPM | 7
TBD-GantryNoise | 7
TBD-Bugrep | 7
CTUDT00630408 | 7
CTUDT00630864 | 7
TBD_CouchEstop | 6
ShouldbeCustomer | 5
TBD-NoEnoughInfo-TBD-NoEnoughInfo | 5
CTUDT00632986 | 4
TBD-GMP | 4
CTUDT00630107 | 4
CTUDT00632236 | 4
CTUDT00631055 | 4
CTUDT00632833 | 4
CTUDT00631058 | 4
Tube Heat | 4
NotEngineering-PartReplacement | 4
TBD-SW-Slow | 4
TBD-Encoder | 3
CTUDT00631671 | 3
CTUDT00631634 | 3
DEFOA | 3
AssemblyAdjust | 3
TBD_Outer laser adjustment | 3
CTUDT00631156 | 3
CTUDT00631545/CTUDT00630971 | 3
CTUDT00631669 | 3
TBD-Service | 3
TBD-RFC | 3
CTUDT00632187 | 3
TBD-DoseReport | 3
ECR-076490 | 3
MemoryReplacement | 3
IBOX | 3
CTUDT00631143 | 2
TBD-SAS | 2
H-Phantom | 2
RFC | 2
PartReplacement | 2
CTUDT00632298 | 2
ECR-077066 | 2
TBD | 2
TBD-PowerOff | 2
BlackScreen | 2
CTUDT00631165 | 2
CTUDT00632114 | 2
Performance Slow | 2
CTUDT00632125 | 2
CTUDT00630969 | 2
CTUDT00633269 | 2
TBD-Initilization | 2
CTUDT00633684 | 2
CTUDT00631638 | 2
TBD_UPS | 2
ECR-075975 | 2
CTUDT00632177 | 2
CTUDT00631545 | 2
Duplicate | 2
TBD-USBExtended | 2
TBD-CouchCover | 2
CTUDT00631136 | 2
CTUDT00631283 | 2
TBD-DVDDrive | 2
HF3 | 2
ShouldbeService | 2
AddRAM | 1
GPU | 1
CTUDT00632803 | 1
NotEngineering-PartOrder | 1
CTUDT00632175 | 1
CTUDT00631717 | 1
TBD-TubGenerate | 1
CTUDT00632584 | 1
Calibration | 1
NotEngineering-Service | 1
TBD-Ghost | 1
CTUDT00632610 | 1
CTUDT00632455 | 1
CTUDT00632618 | 1
TBD-SW-IQ | 1
CTUDT00632793 | 1
TBD-WW | 1
Power Issue | 1
HV Connector | 1
CIRS-SW-CannotRecon | 1
PendingBugRep | 1
ShouldbeApplication | 1
NotEngineering-HorizontalCalibration | 1
ShouldbeClinical | 1
TBD-HeadCushion | 1
CTUDT00632916 | 1
HardDiskReplacement | 1
CTUDT00631140 | 1
TBD-ReconSlow | 1
CTUDT00632121 | 1
CTUDT00630971 | 1
CTUDT00633122 | 1
TBD-TimeSync | 1
TBD_Couch Relay | 1
Not Engineering | 1
CTUDT00632183 | 1
NotDefect | 1
TBD_HR_MP5 | 1
NotEngineering_Hwconnection | 1
TBD_IQ | 1
ReplaceCable | 1
TBD_KV_Perf_Bugrep | 1
CushionBreakdown | 1
TBD_NoEnoughInfo | 1
ReconNotInitialize | 1
TBD_Noise | 1
Hphantom Head | 1
ACQCardReplacement | 1
TBD-GasPring | 1
CTUDT00633734 | 1
CTUDT00628691 | 1
TBD_Restart | 1
TBD-HostNoise | 1
CTUDT00633806 | 1
FRC-Replace | 1
CustomerNetIssue | 1
TBD-Performance | 1
TBD_Vertical Transducer | 1
TBD-PRS | 1
TBD_Workstation | 1
Image Artifact | 1
TBD-ACQ Module | 1
TBD-ScanParameter | 1
TBD-Adapter | 1
TBD-ServiceTool | 1
TBD-AMC | 1
CTUDT00630991 | 1
CTUDT00631142 | 1
CTUDT00631134 | 1
TBD-Artifact | 1
TBD-UHR | 1
DMSMotherboard | 1
TBD-WorklistUsage | 1
TBD-BugRepError | 1
Timeout | 1
TBD-Button | 1
CTUDT00631039 | 1
TBD-Calibration | 1
NotEngineering_CouchEstop_Redocalibration | 1
TBD-CanCard | 1
ReconBox | 1
TBD-CardiacShootIssue | 1
Physics Phantom | 1
TBD-CCT Performance_Bugrep | 1
Replace-GPU | 1
CTUDT00632189 | 1
SystemStuck | 1
TBD-CTDUT | 1
CTD00005656 | 1
CTUDT00632190 | 1
CTD00007125 | 1
CTUDT00630958 | 1
HeadRest | 1
CollimatorTimerOout | 1
CTD00007194 | 1
TBD-ExamSummary | 1
Bugrep | 1
TBD-GantryError | 1
ECR-076862 | 1
Grand Total | 878
PR Status | (Multiple Items)
Funct. Owner | Engineering
IB Exist/New Found | (Multiple Items)
3MonthRolling | Yes
Level1 | (All)
Row Labels | Count of # | Count of #2
Console-SW | 61 | 0.22846
Gantry-HW | 43 | 0.38951
(blank) | 40 | 0.53933
Couch-HW | 19 | 0.61049
NotEngineering | 18 | 0.6779
IC-HW | 14 | 0.73034
DMS-HW | 14 | 0.78277
Gantry-SW | 12 | 0.82772
TBD-Bugreport | 12 | 0.87266
Couch-SW | 10 | 0.91011
Console-HW | 9 | 0.94382
TBD-NoEnoughInfo | 7 | 0.97004
DMS-SW | 5 | 0.98876
CIRS-SW | 3 | 1
Grand Total | 267
Note on D5: Silva: check counts in 'Field IssueSE' if no gap, mark with green, instead red

## FOKSumary
| 270 | 195 | 0 | 1
Type | Complaint number | Human interface or workflow | IQ Known | IQ NEW
Console-SW-Remote Connection | 94 | y | N | N
Console-SW-Yellow-Screen | 39 | Y | N | N
Console-SW-Blue-Screen | 13 | Y | N | N
Console-SW-DiskFull | 13 | Y | N | N
Console-SW-NoEnoughInfo | 12 | N | N | N
Console-SW-Performance | 7 | N | N | N
Console-SW-Lan Config | 5 | Y | N | N
Console-SW-Service Tool | 4 | Y | N | N
Console-SW-Dose SR | 4 | N | N | N
Console-SW-ExamCard | 3 | N | N | N
Console-SW-ShouldbeCustomer | 3 | N | N | N
Console-SW-ResultSelectionJump | 3 | Y | N | N
Console-SW-SystemStuck | 2 | N | N | N
Console-SW-TBD-bugrep | 2 | N | N | N
Console-SW-HF3 | 2 | Y | N | N
Console-SW-Imagemissing-Printer | 2 | N | N | N
Console-SW-TBD-NoEnoughInfo | 2 | N | N | N
Console-SW-User Account | 2 | Y | N | N
Console-SW-RFC | 1 | N | N | N
Console-SW-InstallationVersion | 1 | N | N | N
Console-SW-Spectral Error Message | 1 | N | N | N
Console-SW-Login | 1 | Y | N | N
Console-SW-BugRepError | 1 | N | N | N
Console-SW-MIP | 1 | N | N | N
Console-SW-ExamSummary | 1 | N | N | N
Console-SW-MPR Recon | 1 | Y | N | N
Console-SW-TBD_Bugreport | 1 | N | N | N
Console-SW-Bugrep | 1 | N | N | N
Console-SW-DVDRAM | 1 | N | N | N
Console-SW-NotDefect | 1 | N | N | N
Console-SW-DeviceConfig | 1 | N | N | N
Console-SW-TimeSync | 1 | N | N | N
Console-SW-SerielsJump | 1 | Y | N | N
Console-SW-WorklistUsage | 1 | Y | N | N
Console-SW-FilmStuck | 1 | Y | N | N
Console-SW-CannotRecon | 1 | Y | N | N
Console-SW-Biopsy | 1 | N | N | N
Console-SW-CCT Performance | 1 | N | N | N
Console-SW-CTA | 1 | Y | N | N
Console-SW-PreviewImage-WW-WC | 1 | Y | N | N
Console-SW-TBD-Performance | 1 | N | N | N
Console-SW-PRSPort | 1 | N | N | N
Console-SW-RealTimeSurview | 1 | N | N | N
Console-SW-AuditRepository | 1 | N | N | N
Console-SW-WW | 1 | N | N | N
Console-SW-Performance Slow | 1 | Y | N | N
Console-SW-TBD-OBS | 1 | N | N | N
Console-SW-Perfusion | 1 | N | N | N
Console-SW-PoorIQ | 1 | Y | N | Y
CIRS-SW-MPR Recon | 9 | Y | N | N
CIRS-SW-NoEnoughInfo | 7 | N | N | N
CIRS-SW-ACQ_NOT_OK | 6 | N | N | N
CIRS-SW-TBD_NoEnoughInfo | 1 | N | N | N
CIRS-SW-ReconSlow | 1 | N | N | N
CIRS-SW-TBD-NoEnoughInfo | 1 | N | N | N
CIRS-SW-Installation | 1 | N | N | N
CIRS-SW-CannotRecon | 1 | Y | N | N

## Jul2021CompliantARlIST
CTUDT | ARNumber | Row Labels | Sum of ARNumber | Row Labels | Sum of ARNumber | checknew
CTUDT00632127 | 69 | CTUDT00632127 | 69 | CTUDT00632127 | 69 | CTUDT00632127
CTUDT00632110 | 34 | CTUDT00632110 | 34 | CTUDT00632110 | 34 | CTUDT00632110
CTUDT00631367 | 24 | CTUDT00630085 | 30 | CTUDT00630085 | 30 | CTUDT00630085
CTUDT00630085 | 30 | CTUDT00631367 | 24 | CTUDT00631367 | 24 | CTUDT00631367
ECR-077012 | 22 | CTUDT00632185 | 23 | CTUDT00632185 | 23 | CTUDT00632185
CTUDT00632185 | 23 | ECR-077012 | 22 | ECR-077012 | 22 | #N/A
ECR-075975 | 18 | ECR-075975 | 18 | ECR-075975 | 16 | #N/A
CTUDT00630138 | 14 | CTUDT00630138 | 14 | CTUDT00630138/630160 | 14 | #N/A
CTUDT00630160 | 14 | CTUDT00630160 | 14 | CTUDT00632520 | 12 | CTUDT00632520
CTUDT00632520 | 12 | CTUDT00632520 | 12 | CTUDT00632173 | 11 | CTUDT00632173
CTUDT00632173 | 11 | CTUDT00632173 | 11 | E-125059 | 10 | #N/A
CTUDT00632116 | 8 | E-125059 | 10 | CTUDT00632553 | 9 | CTUDT00632553
CTUDT00631602 | 7 | CTUDT00632553 | 9 | CTUDT00632116 | 8 | CTUDT00632116
E-125059 | 10 | CTUDT00630971 | 8 | CTUDT00631602 | 7 | CTUDT00631602
CTUDT00632261 | 7 | CTUDT00632116 | 8 | CTUDT00632261 | 7 | CTUDT00632261
CTUDT00632179 | 6 | CTUDT00631602 | 7 | CTUDT00630864 | 7 | CTUDT00630864
CTUDT00630864 | 7 | CTUDT00632261 | 7 | CTUDT00632179 | 6 | CTUDT00632179
CTUDT00632553 | 9 | CTUDT00630864 | 7 | CTUDT00632181 | 6 | CTUDT00632181
CTUDT00632833 | 4 | CTUDT00632181 | 6 | CTUDT00630408 | 4 | CTUDT00630408
CTUDT00631058 | 4 | CTUDT00632179 | 6 | CTUDT00630107 | 4 | CTUDT00630107
CTUDT00630107 | 4 | CTUDT00631545 | 5 | CTUDT00632833 | 4 | CTUDT00632833
CTUDT00632181 | 6 | CTUDT00631058 | 4 | CTUDT00631058 | 4 | CTUDT00631058
CTUDT00630408 | 4 | CTUDT00632833 | 4 | CTUDT00631634 | 3 | CTUDT00631634
CTUDT00631156 | 3 | CTUDT00630408 | 4 | CTUDT00631669 | 3 | CTUDT00631669
CTUDT00631634 | 3 | CTUDT00630107 | 4 | CTUDT00632187 | 3 | CTUDT00632187
CTUDT00632187 | 3 | CTUDT00631634 | 3 | ECR-076490 | 3 | #N/A
CTUDT00631669 | 3 | CTUDT00632986 | 3 | CTUDT00631671 | 3 | CTUDT00631671
CTUDT00631545 | 5 | CTUDT00631156 | 3 | CTUDT00632986 | 3 | CTUDT00632986
CTUDT00630971 | 4 | CTUDT00632187 | 3 | CTUDT00631545/CTUDT00630971 | 3 | #N/A
CTUDT00632986 | 3 | CTUDT00631669 | 3 | CTUDT00631156 | 3 | CTUDT00631156
CTUDT00631671 | 3 | CTUDT00631671 | 3 | CTUDT00630969 | 2 | CTUDT00630969
CTUDT00632298 | 2 | CTUDT00631165 | 2 | CTUDT00631143 | 2 | CTUDT00631143
CTUDT00631165 | 2 | CTUDT00631055 | 2 | CTUDT00631638 | 2 | CTUDT00631638
CTUDT00631283 | 2 | CTUDT00633684 | 2 | CTUDT00632114 | 2 | CTUDT00632114
CTUDT00632177 | 2 | HF3 | 2 | CTUDT00633269 | 2 | CTUDT00633269
CTUDT00633269 | 2 | CTUDT00631638 | 2 | CTUDT00631545 | 2 | CTUDT00631545
CTUDT00633684 | 2 | CTUDT00631283 | 2 | CTUDT00632298 | 2 | CTUDT00632298
CTUDT00632114 | 2 | CTUDT00633269 | 2 | CTUDT00633684 | 2 | CTUDT00633684
CTUDT00632125 | 2 | CTUDT00632298 | 2 | CTUDT00632125 | 2 | CTUDT00632125
ECR-076490 | 2 | CTUDT00632177 | 2 | CTUDT00631165 | 2 | CTUDT00631165
CTUDT00631055 | 2 | CTUDT00632114 | 2 | CTUDT00631055 | 2 | CTUDT00631055
ECR-077066 | 2 | ECR-076490 | 2 | CTUDT00631283 | 2 | CTUDT00631283
HF3 | 2 | CTUDT00630969 | 2 | CTUDT00632177 | 2 | CTUDT00632177
CTUDT00631136 | 2 | ECR-077066 | 2 | ECR-077066 | 2 | #N/A
CTUDT00631638 | 2 | CTUDT00631136 | 2 | ECR-075975 | 2 | #N/A
CTUDT00630969 | 2 | CTUDT00631143 | 2 | CTUDT00630971 | 1 | CTUDT00630971
CTUDT00631143 | 2 | CTUDT00632125 | 2 | ECR-076862 | 1 | #N/A
CTUDT00633122 | 1 | CTUDT00631142 | 1 | CTUDT00631140 | 1 | CTUDT00631140
CTUDT00632584 | 1 | CTUDT00632121 | 1 | CTUDT00632236 | 1 | CTUDT00632236
CTUDT00632183 | 1 | CTUDT00632189 | 1 | CTUDT00632175 | 1 | CTUDT00632175
CTUDT00632610 | 1 | CTUDT00628691 | 1 | CTUDT00630991 | 1 | CTUDT00630991
CTUDT00632175 | 1 | CTUDT00633122 | 1 | CTUDT00631717 | 1 | CTUDT00631717
CTUDT00630958 | 1 | CTUDT00632455 | 1 | CTUDT00632121 | 1 | CTUDT00632121
CTUDT00632803 | 1 | CTUDT00630958 | 1 | CTUDT00632189 | 1 | CTUDT00632189
CTUDT00630971 | 4 | CTUDT00631134 | 1 | CTUDT00632455 | 1 | CTUDT00632455
CTUDT00631717 | 1 | CTUDT00632236 | 1 | CTUDT00632803 | 1 | CTUDT00632803
CTUDT00632189 | 1 | CTUDT00631039 | 1 | CTUDT00631039 | 1 | CTUDT00631039
CTUDT00632236 | 1 | CTUDT00631140 | 1 | CTUDT00633122 | 1 | CTUDT00633122
CTUDT00630991 | 1 | CTUDT00632584 | 1 | CTUDT00628691 | 1 | CTUDT00628691
CTUDT00632618 | 1 | CTUDT00631717 | 1 | CTUDT00631134 | 1 | CTUDT00631134
CTUDT00632793 | 1 | CTUDT00632610 | 1 | CTUDT00632584 | 1 | CTUDT00632584
CTUDT00631039 | 1 | CTUDT00632183 | 1 | CTUDT00632183 | 1 | CTUDT00632183
CTUDT00631134 | 1 | CTUDT00632618 | 1 | CTUDT00632610 | 1 | CTUDT00632610
CTUDT00631140 | 1 | CTUDT00630991 | 1 | CTUDT00631142 | 1 | CTUDT00631142
CTUDT00632121 | 1 | CTUDT00632793 | 1 | CTUDT00632618 | 1 | CTUDT00632618
CTUDT00632455 | 1 | CTUDT00632190 | 1 | CTUDT00633806 | 1 | #N/A
CTUDT00631142 | 1 | CTUDT00632803 | 1 | CTUDT00632793 | 1 | CTUDT00632793
CTUDT00628691 | 1 | CTUDT00632175 | 1 | CTUDT00633734 | 1 | #N/A
CTUDT00632190 | 1 | CTUDT00632916 | 1 | CTUDT00630958 | 1 | CTUDT00630958
CTUDT00632916 | 1 | (blank) |  | CTUDT00632190 | 1
|  | Grand Total | 432 | CTUDT00632916 | 1

## Engineering-SlidesTable
IDX | Sequence | 2nd level Pareto | 3rd level Pareto | CTUDT | 2nd level Pareto | 3rd level Pareto | Problem Description | Root causes | Countermeasures | Owner | RD Timing | Stage2 Timing | Complaint # | Mitigate Complaint # | Impact on Gap | Status (%) | Percentage | 100% | Complaint | 8.99 | Complaints | 810
108 | 71 | Console-SW | Console-SW-Remote Connection | CTUDT00632127 | Console-SW | Console-SW-Remote Connection | Several FOK site reply back unsuccessful sending images to PACS again. The failure frequency of this issue is very high. | 1. There is no retry mechanism when transfer failure happens. 2. Image transfer loaded which lead spooler stuck, real root cause still under investigation. | Temporary: Add additional logs in HF3 and keep observ and continue trigger investigation Permanent: Code change to fix it in SP1. | Sankari S / iCAP | 2021-06-30 | 2021-06-30 | 64 | 59 | 0.65483 | 75% | 64 | 64
108 | 71 | Console-SW | Console-SW-Remote Connection | CTUDT00631602 | Console-SW | Console-SW-Remote Connection | Several FOK site reply back unsuccessful sending images to PACS again. The failure frequency of this issue is very high. | 1. There is no retry mechanism when transfer failure happens. 2. Image transfer loaded which lead spooler stuck, real root cause still under investigation. | Temporary: Add additional logs in HF3 and keep observ and continue trigger investigation Permanent: Code change to fix it in SP1. | Sankari S / iCAP | 2021-03-30 | 2021-03-30 | 7 | 7 | 0.077691 | 75% | 7 | 7
108 | 71 | Console-SW | Console-SW-Remote Connection | CTUDT00632114 | Console-SW | Console-SW-Remote Connection | Image transfer stuck related to network connection | Network is not stable enough between host and third party remote device. | Temporary: Need user check network and reset spooler. Permanent: Trigger further investigation and consider adding possible retry mechanism in SP1 | Sankari S / iCAP | 2021-06-30 | 2021-06-30 | 2 | 0 | 0 | 100% | 2 | 2
108 | 71 | Console-SW | Console-SW-Remote Connection | CTUDT00631283 | Console-SW | Console-SW-Remote Connection-Worklist | Several FOK site reply back remote device retrieve or transfer failure(Work list populating issue) | User use wrong worklist filter. | Temporary: Give clinical application training to user. Permanent: Give clinical application training to user. | Caroline Burke / Clinical Application | 2020-10-27 | 2020-10-27 | 2 | 0 | 0 | 100% | 2 | 2
108 | 71 | Console-SW | Console-SW-Remote Connection | CTUDT00630958 | Console-SW | Console-SW-Remote Connection | Several FOK site reply back unsuccessful sending images to PACS again. The failure frequency of this issue is very high. | 1. There is no retry mechanism when transfer failure happens. 2. Image transfer loaded which lead spooler stuck, real root cause still under investigation. | Temporary: Add additional logs in HF3 and keep observ and continue trigger investigation Permanent: Code change to fix it in SP1. | Sankari S / iCAP | 2021-02-15 | 2021-02-15 | 1 | 1 | 0.011099 | 100% | 1 | 1
108 | 27 | Console-SW | Console-SW-Yellow-Screen | CTUDT00632520/CTUDT00631367 | Console-SW | Console-SW-Yellow-Screen | FOK-Image transfer issue and yellow screen issue happened | TBD | Temporary: Reboot system Permanent: code fix. First round of sw patch had been deliverred to team. | Shubdeep / Console | 2021-05-30 | 2021-05-30 | 25 | 25 | 0.27747 | 25% | #N/A | #N/A
108 | 27 | Console-SW | Console-SW-Yellow-Screen | CTUDT00631367 | Console-SW | Console-SW-Yellow-Screen | Yellow screen issue happens when use workflow, user have to relogin system or restart host computer. | Patient Directory .cfg was using a mutex mechanism. Two threads access PatientDirectory.cfg at same time to trigger conflict. | Temporary: Need user relogin system Permanent: Trigger code change and merged in SP1 WIP34 | Sankari S / iCAP | 2021-03-30 | 2021-03-30 | 6 | 6 | 0.066593 | 75% | 6 | 25
108 | 27 | Console-SW | Console-SW-Yellow-Screen | CTUDT00632520/CTUDT00631367 | Console-SW | Console-SW-Yellow-Screen | Yellow screen issue happens when use workflow, user have to restart host computer. | Patient Directory .cfg was using a mutex mechanism. Two threads access PatientDirectory.cfg at same time to trigger conflict. | Temporary: Need user relogin system Permanent: Trigger code change and merged in SP1 WIP34 | Sankari S / iCAP | 2021-03-30 | 2021-03-30 | 6 | 6 | 0.066593 | 100% | 1 | #N/A
108 | 6 | Console-SW | Console-SW-MPR Recon | CTUDT00632116 | Console-SW | Console-SW-MPR Recon | MPR Recon with wrong size for Brain/Head Axial EC, but only happens in one site. | Rootcause is not clear, need further investigation. | Temporary: Reload SW in customer site. Permanent: Trigger further investigation | Shubdeep / Console | 2021-06-30 | 2021-06-30 | 6 | 6 | 0.066593 | 50% | 5 | 6
108 | 6 | Console-SW | Console-SW-MPR Recon | CTUDT00632183 | Console-SW | Console-SW-MPR Recon | MPR cannot be created from Axial image. | Rootcause is not clear, need further investigation. | Temporary: Reload SW in customer site. Permanent: Trigger further investigation | Console4 | 2021-06-30 | 2021-06-30 | 1 | 1 | 0.011099 | 50% | 1 | 1
108 | 4 | Console-SW | Console-SW-DiskFull | CTUDT00632261 | Console-SW | Console-SW-Diskfull | The dump files are created in D drive filling the disk space. | Rootcause is not clear, need further investigation. | Temporary: Restart system. Permanent: Trigger further investigation, may need code change | Console4 | 2021-03-30 | 2021-03-30 | 5 | 5 | 0.055494 | 25% | 5 | 5
108 | 4 | Console-SW | Console-SW-Service Tool | CTUDT00631143 | Console-SW | Console-SW-Service Tool | Enable Remote Access not present in Home Menu. | Service click DoD cofig tool as mistake, then remote connect tool is remobe from list. | Temporary: Reinstalll the software in customer site. Permanent: Deliver tips to service team and high light the DoD config tool usage. | Brook Shi / Service | 2020-09-29 | TBD | 2 | 0 | 0 | 100% | 2 | 2
108 | 4 | Console-SW | Console-SW-Service Tool | TBD-ServiceTool | Console-SW | Console-SW-Service tool |  |  |  |  | NA | NA | 1 | 0 | 0 |  | 1 | #N/A
108 | 2 | Console-SW | Console-SW-ResultSelectionJump | CTUDT00632986 | Console-SW | Console-SW-ResultSelectionJump | When user select one seriels, it jump to another with random. | under investigation. | Temporary: annoying issue to user, IB eisting issue . Permanent: code fix. | Console4 | TBD | TBD | 3 | 0 | 0 |  | 2 | 3
108 | 2 | Console-SW | Console-SW-HF3 | HF3 | Console-SW | Console-SW-HF3 | HF3 Workaround | performing Fco workaround | Temporary: FCO wokraround Permanent: fix the issue in SP1 | Console4 | 2021-06-30 | 2021-06-30 | 2 | 2 | 0.022198 |  | 2 | #N/A
108 | 2 | Console-SW | Console-SW-Biopsy | CTUDT00632916 | Console-SW | Console-SW-Biopsy | CCT coverage not correct | Upgrade CCT protocol convert incorrectly | Temporary: Update EC coverage Permanent: Support EC converting and then convert EC with right parameters | Console4 | 2021-08-30 | 2021-08-30 | 1 | 1 | 0.011099 |  | 1 | 1
108 | 2 | Console-SW | Console-SW-AuditRepository | CTUDT00633122 | Console-SW | Console-SW-AuditRepository | Audit log too huge to fill D disk | under investigation. | Temporary: Remove audit log from D disk. Permanent: check root cause and trigger code fix. | Console4 | TBD | TBD | 1 | 0 | 0 |  | 1 | 1
60 | 14 | Gantry-HW | Gantry-HW-Lexan Ring | ECR-075975 | Gantry-HW | Gantry-HW-Lexan Ring | 1. Contrast media or other fluid ingressed inside the gantry (~8 cases). 2. Image Artifact (~4 cases). 3. Noise isuue (~1 case). 4. Gantry-HW-SlipRingClean (~1 case). | Fitting dimension and tolerance between the Lexan ring and the front cover was not set properly leading to the fitting between Lexan ring and gantry front cover is not good. | To implement updated Lexan ring design in refresh (E-128610). | Lucas Peng Brook Shi | 2020-08-05 | 2020-08-05 | 16 | 14 | 0.15538 | 75% | 12 | 16
60 | 10 | Gantry-HW | Gantry-HW-Belt | ECR-077012 | Gantry-HW | Gantry-HW-Belt | 1. There is noise coming from the location where the belt was installed during gantry rotation (~7 cases). 2. Dual pot malfunction/broken leadind to gantry tilting problem. (~3 cases). | 1.1. Belt was stretched longer/loose without enough tension and the position of the belt tensioner was not adjusted to make the belt tensioned. 1.2. Not enough adjustment position on belt tensioner bracket. 2. Tension force applied on the dual pot belt is too big. | 1.1. Regular check the position of the belt tensioner during PM and adjust it if the tensioner was found hanging down (ECT314279). 1.2. Design change to add more adjustment position on the tensioner bracket for tensioner position adjustment (ECT303629). 2. Design change to define the tension force of the dual pot belt as 20N~30N and control it on every produciton and field unit (ECT293665). | Alan Li Brook Shi | 2021-04-12 | 2021-04-12 | 16 | 16 | 0.17758 | 50% | 14 | 16
60 | 10 | Gantry-HW | Gantry-HW-CT Box | CTUDT00632185 | Gantry-HW | Gantry-HW-CT Box | 1. Button not full function (~7 cases). 2. CRC power supply issue (~2 cases). 3. Unclear replacement (~1 case). | Pending on failure part return back. | Pending on failure part return back. | Alan Li / Gantry Brook Shi / Service | 2021-12-20 | TBD | 16 | 13 | 0.14428 | 25% | 13 | 16
60 | 9 | Gantry-HW | Gantry-HW-GantryLaser | TBD_Outer laser adjustment | Gantry-HW | Gantry-HW-Laser | 1. The position of the external laser mounted on the front cover was changed and not adjusted well after field installation (~2 cases). 2. Unclear (~1 case). | 1.1. The front cover was not secured strongly leading to the position change during shipment. 1.2. The front cover was not closed properly by FSE after PM. | 1.1.1. Increase front cover mounting screw torque to give more design margin for holding the cover in position (E-127321). 1.1.2. Service manual update to add check points of the 8 front cover mounting screws loose and re-tighten if necessary (ECT314279). 1.2. Update current Ingenuity & BigBore Gantry R&R manual to add additional steps with photographs to improve gantry front cover closure instructions and train FSE (E-126517). | Andy Shao Brook Shi | 2020-06-23 | 2020-06-23 | 3 | 3 | 0.033296 | 25% | 3 | #N/A
60 | 9 | Gantry-HW | Gantry-HW-GantryLaser | CTUDT00631634 | Gantry-HW | Gantry-HW-Laser | 1. The position of the external laser mounted on the front cover was changed after installation. | 1. The front cover was not secured strongly leading to the position change during shipment. | 1. Increase front cover mounting screw torque to give more design margin for holding the cover in position (E-127321). | Andy Shao / Gantry | 2020-09-02 | 2020-09-02 | 3 | 3 | 0.033296 | 75% | 3 | 3
60 | 9 | Gantry-HW | Gantry-HW-GantryLaser | ECR-077066 | Gantry-HW | Gantry-HW-Laser | 1. The position of the inner laser mounted on the rotor was changed after certain period. | 1. There is no loose prevention method for the inner laser mounting screws after they are adjusted. | 1. Changed the mechanical structure design of the inner laser mounting to introduce the loose prevention nuts (ECT299595). | Andy Shao | 2021-04-28 | 2021-04-28 | 2 | 2 | 0.022198 | 75% | 2 | 2
60 | 5 | Gantry-HW | Gantry-HW-SlipRingClean | NotEngineering-SlipringClean | Gantry-HW | Gantry-HW-SlipRingClean | 1. The slipring and brushes cleanness issue causing problem of intermittent interruption of acquisitions. | Dust/debris/foreign material enter into the lubricant of the slip ring leading to the intermittent contact between the brush and slip ring base. | Increase the frequency of PM to clean the slip ring system (Current frequency is half year). | Brook Shi | 2021-09-30 | TBD | 6 | 5 | 0.055494 | 25% | 6 | #N/A
60 | 4 | Gantry-HW | Gantry-HW-Noise | TBD-GantryNoise | Gantry-HW-Noise | Gantry-HW-Noise | 1. Noise come from the gantry. | 1. Front cover’ laser cable routing is not placed well so that the cable gets in touch with the rotation frame occasionally. 2. The front cover was not secured strongly leading to the Lexan ring position change during shipment, making the Lexan ring gets in touch with the collimator occasionally during gantry rotation. | 1. Modify cable routing on front cover to enlarge the distance between the cable and the rotational frame to avoid getting in touch with each other (E-127321). 2. Increase front cover mounting screw torque to give more design margin for holding the cover in position (E-127321). 3. Add PM checkpoints to check if the torque seal applied on the front cover mounting screws/nuts is cracked or broken, if yes, adjust the cover position, retighten the screws/nuts with specified torque, rotate the gantry and make sure there is not any abnormal noise coming from the gantry (ECT314279). | Andy Shao Brook Shi | 2021-04-28 | 2021-04-28 | 3 | 3 | 0.033296 | 25% | 3 | #N/A
60 | 4 | Gantry-HW | Gantry-HW-FlatSpring | CTUDT00632833 | Gantry-HW-FlatSpring | Gantry-HW-FlatSpring |  |  |  | NA |  |  | 4 | 0 | 0 |  | 4 | 4
60 | 4 | Gantry-HW | Gantry-HW-Gantry Panel | CTUDT00632181 | Gantry-HW | Gantry-HW-Gantry Panel | 1. Control panel button broken while covered with the protection film (~1 case). 2. Button not full function (~2 cases). 3. Unclear (~1 case). | 1. Some chemical elements in the protection film chemically reacted with the control panel button of silicone rubber material. 2. Pending on failure part return back for investigation. | 1. Remove the Raucodrape® Surgical Incise Drape from the control panel. Suggest use other plastic film which is not sensitive to rubber to cover the control panel. 2. Pending on failure part return back for investigation. | Alan Li / Gantry Brook Shi / Service | 2021-11-30 | TBD | 4 | 4 | 0.044395 | 25% | 4 | 4
26 | 20 | IC-HW | IC-HW-Tub Arc | CTUDT00630085 | IC-HW | IC-HW-Tub Arc | Tube conditioning aging not sufficient before shipment | Tube conditioning aging not sufficient before shipment | Temporary: short tube conditioning by customer Permanent: Improve tube aging through at Philips with software update (TC Boost), released in WIP33 | Michael Yu / Gantry | 2021-03-30 | 2021-03-30 | 23 | 20 | 0.22198 | 75% | 20 | 23
26 | 20 | IC-HW | IC-HW-Tub Arc | CTUDT00632177 | IC-HW | IC-HW-Hard Arcs | Tube Arc | 1. HV connection/insulation not good. 2. Tube conditioning aging not sufficient before shipment. | Temporary: Inspect HV connections, candlesticks, silicone grease, etc. Perform long tube conditioning. Perform filament calibrations. Inspect Power Brush Blocks / slip ring for poor connections. Perform no-load test. Permanent: Optimize the current Tube conditioning process. Tube conditioning boost to replace Long tube conditioning (LTC). TC Boost will have a conditioning algorithm that is more effective in reconditioning (saving) tubes but also the time needed for this procedure is shorter and much more predictable. | Dragon Zhou / Gantry | 2021-06-30 | 2021-06-30 | 2 | 2 | 0.022198 | 50% | 2 | 2
26 | 6 | IC-HW | IC-HW-APM/CPM | TBD-APM/CPM | IC-HW | IC-HW-APM/CPM | Generator HW issue. TBD | TBD | TBD | Michael Yu / Gantry | 2021-12-30 | 2021-12-30 | 4 | 0 | 0 | 25% | 4 | #N/A
26 | 6 | IC-HW | IC-HW-APM/CPM | TBD_KV_Perf_Bugrep | IC-HW | IC-HW-APM/CPM |  |  |  |  | NA | NA | 1 | 0 | 0 |  | 1 | #N/A
26 | 6 | IC-HW | IC-HW-APM/CPM | TBD-TubGenerate | IC-HW | IC-HW-APM/CPM |  |  |  |  | NA | NA | 1 | 0 | 0 |  | 1 | #N/A
25 | 25 | DMS-HW | DMS-HW-DMS Module | CTUDT00632110 | DMS-HW | DMS-HW-DMS Module | Artifact happens in customer site because DMS module malfunction | DMS module malfunction | Temporary: Replace DMS Module Permanent: Request further investigation and confirm DMS Module quality improvement | Limor Yatiz / DMS | 2021-12-31 | 2021-12-31 | 26 | 0 | 0 | 100% | 24 | 26
25 | 25 | DMS-HW | DMS-HW-DMS Module | CTUDT00631717 | DMS-HW | DMS-HW-DMS Module | Ring artifact happens in customer site because DMS module malfunction | DMS module malfunction | Temporary: Replace DMS Module Permanent: Request further investigation and confirm DMS Module quality improvement | Limor Yatiz / DMS | 2021-12-31 | 2021-12-31 | 1 | 0 | 0 | 100% | 1 | 1
22 | 22 | Couch-SW | Couch-SW-Horizontal Move Block | CTUDT00630138/630160 | Couch-SW | Couch-SW-Horizontal Move Block | Couch horizontal motion stop by mechanical stop | Couch horizontal motion stop by mechanical stop | Temporary: Refresh hotfix 1 SW patch released to field Permanent: Couch SW update for manual motion stop improvement, released in WIP33 Release with SP2 | Chang Xian Zhang / Couch | 2020-10-30 | 2020-10-30 | 13 | 13 | 0.14428 | 100% | 13 | #N/A
22 | 22 | Couch-SW | Couch-SW-Horizontal Move Block | CTUDT00631156 | Couch-SW | Couch-SW-Horizontal Move Block | Couch horizontal motion stop by mechanical stop | Couch horizontal motion stop by mechanical stop | Temporary: Refresh hotfix 1 SW patch released to field Permanent: Couch SW update for manual motion stop improvement, released in WIP33 Release with SP2 | Chang Xian Zhang / Couch | 2020-10-30 | 2020-10-30 | 3 | 3 | 0.033296 | 100% | 3 | 3
22 | 22 | Couch-SW | Couch-SW-Horizontal Move Block | CTUDT00631669 | Couch-SW | Couch-SW-Horizontal Move Block | Couch horizontal motion stop by mechanical stop | Couch horizontal motion stop by mechanical stop | Temporary: Refresh hotfix 1 SW patch released to field Permanent: Couch SW update for manual motion stop improvement, released in WIP33 Release with SP2 | Chang Xian Zhang / Couch | 2020-10-30 | 2020-10-30 | 3 | 3 | 0.033296 | 75% | 3 | 3
22 | 22 | Couch-SW | Couch-SW-Horizontal Move Block | CTUDT00631039 | Couch-SW | Couch-SW-Horizontal Move Block | Couch horizontal motion stop by mechanical stop | Couch horizontal motion stop by mechanical stop | Temporary: Refresh hotfix 1 SW patch released to field Permanent: Couch SW update for manual motion stop improvement, released in WIP33 Release with SP2 | Chang Xian Zhang / Couch | 2020-10-30 | 2020-10-30 | 1 | 1 | 0.011099 | 100% | 1 | 1
22 | 22 | Couch-SW | Couch-SW-Horizontal Move Block | CTUDT00632189 | Couch-SW | Couch-SW-Estop open | User specific operation lead couch Estop open. | Software timeout | Temporary: Add 2s between manual horizontal and vertical motion Permanent: Update couch and gantry SW Release with SP2 | Chang Xian Zhang / Couch | 2021-01-23 | 2021-01-23 | 1 | 1 | 0.011099 | 50% | 1 | 1
13 | 5 | Couch-HW | Couch-HW-Motor Coupling | CTUDT00632179 | Couch-HW | Couch-HW-Motor Coupling | Repeatable Couch Horizontal Motor Coupling | Investigation in process | Temporary: Check coupling bending during assembly Permanent: Coupling mouting improvement. Update assembly WI | Tommy Sun / Couch | 2021-06-30 | 2021-06-30 | 5 | 5 | 0.055494 | 25% | 5 | 5
13 | 4 | Couch-HW | Couch-HW-Vertical spring | CTUDT00631058 | Couch-HW | Couch-HW-VerticalTransducer | No sufficient info, need part return |  |  |  | NA | NA | 6 | 3 | 0.033296 |  | 3 | 6
13 | 4 | Couch-HW | Couch-HW-Lubricant | E-125059 | Couch-HW | Couch-HW-Lubricant | Couch vertical movement noise from side rail. | short of oil selection guideline. This lubricating oil is not suitable for couch long time operation usage. | qualify new lubricating oil. | Tommy Sun / Couch | 2020-12-15 | 2020-12-15 | 4 | 4 | 0.044395 |  | 4 | #N/A
13 | 1899-12-31 | Couch-HW | Couch-HW-Matress | CTUDT00632793 | Couch-HW | Couch-HW-Matress | Artifact ahppens when uer put patient head directly on couch and perfomring scan, then check image with incorrect WC,WW. | Incorrect usage of scan | Temporary: Info customer Permanent: add note in CRN to mention it. | Couch-HW | 2021-08-30 | 2021-08-30 | 1 | 0 | 0 |  | 1 | 1
8 | 8 | DMS-SW | DMS-SW-Phantom Detect | CTUDT00630107 | DMS-SW | DMS-SW-Phantom Detect |  |  |  |  | NA | NA | 5 | 0 | 0 |  | 4 | 5
8 | 8 | DMS-SW | DMS-SW-Phantom Detect | CTUDT00631545 | DMS-SW | DMS-SW-AutoCentering | [Refresh FOK] Phantom not seen during IQ check | Auto Centering detect phantom algorithm has known limitation on some couch position. | Temporaty: User can redo the calibration. Or perform calibration manually. Permanent: Improve algorithm to support all couch position. | DMS | NA | NA | 2 | 0 | 0 | Postponed | 2 | 2
8 | 8 | DMS-SW | DMS-SW-Phantom Detect | CTUDT00630971 | DMS-SW | DMS-SW-AutoCentering | FOK: During IQ check, scanner stated no phantom was detected, but actually the phantom is on the scout. | Auto Centering detect phantom algorithm has known limitation on some couch position. | Temporaty: User can redo the calibration. Or perform calibration manually. Permanent: Improve algorithm to support all couch position. | DMS | NA | NA | 2 | 0 | 0 | Closed | 1 | 2
7 | 7 | Gantry-SW | Gantry-SW-TubHot | TBD_Tube hot_bugreport | Gantry-SW | Gantry-SW-TubHot | [6 empty] | 9 | 0 | 0 |  | 9 | #N/A
7 | 7 | Console-HW | Console-HW-Adaptor | CTUDT00632173 | Console-HW | Console-HW-Adaptor | Random red artifacts happens when performing surview or clinical scan in customer sites. | DVI Adaptor failure with poor quality, multi site have report this adaptor failure during installation. | Temporary: Replace DVI adaptor of monitor Permanent: Further investigation of DVI quality status and looking for new supplier of DVI adaptor | Shelvin Yuan / ConsoleHW | 2021-09-01 | 2021-09-01 | 8 | 8 | 0.08879 | 50% | 8 | 8

## tmp
Defect | Number | 2nd level Pareto | 3rd level Pareto | xx | yy | zz
CTUDT00632127 | 54 | Console-SW |  | Console-SW-Remote Connection | Console-SW | 71 | Row Labels | Sum of zz
CTUDT00632110 | 23 | CTUDT00632110 |  | Console-SW-Yellow-Screen | Console-SW | 27 | Console-SW | 111
CTUDT00632641 | 20 | CTUDT00632520 |  | DMS-HW-DMS Module | DMS-HW | 25 | Gantry-HW | 63
CTUDT00630085 | 18 | CTUDT00630085 |  | Couch-SW-Horizontal Move Block | Couch-SW | 22 | IC-HW | 26
ECR-075975 | 14 | ECR-075975 |  | IC-HW-Tub Arc | IC-HW | 20 | DMS-HW | 25
CTUDT00630138/630160 | 13 | CTUDT00630138/630160 |  | Gantry-HW-Lexan Ring | Gantry-HW | 14 | Couch-SW | 22
ECR-077012 | 10 | ECR-077012 |  | Gantry-HW-CT Box | Gantry-HW | 10 | Couch-HW | 17
CTUDT00632185 | 10 | CTUDT00632185 |  | Gantry-HW-Belt | Gantry-HW | 10 | DMS-SW | 8
CTUDT00631602 | 7 | CTUDT00631602 |  | Gantry-HW-GantryLaser | Gantry-HW | 9 | Console-HW | 6
CTUDT00632173 | 7 | CTUDT00632173 |  | DMS-SW-Phantom Detect | DMS-SW | 8 | Grand Total | 278
CTUDT00631367 | 6 | CTUDT00631367 |  | Gantry-SW-TubHot | Gantry-HW | 7
NotEngineering-SlipringClean | 5 | NotEngineering-SlipringClean |  | Console-HW-Adaptor | Console-HW | 6
CTUDT00632116 | 5 | CTUDT00632116 |  | IC-HW-APM/CPM | IC-HW | 6
TBD_CouchEstop | 5 | Couch-SW | Couch-SW-TBD_CouchEstop | Gantry-HW-SlipRingClean | Gantry-HW | 5
CTUDT00632179 | 5 | CTUDT00632179 |  | Console-SW-MPR Recon | Console-SW | 5
TBD-APM/CPM | 4 | TBD-APM/CPM |  | Couch-HW-Motor Coupling | Couch-HW | 5
CTUDT00632181 | 4 | CTUDT00632181 |  | Console-SW-DiskFull | Console-SW | 4
CTUDT00630107 | 4 | DMS-SW | DMS-SW-Phantom Detect | Couch-HW-Vertical spring | Couch-HW | 4
CTUDT00632833 | 4 | Gantry-HW | Gantry-HW-FlatSpring | Couch-HW-Lubricant | Couch-HW | 4
E-125059 | 4 | Couch-HW | Couch-HW-Lubricant | Couch-HW-FlatSpring | Couch-HW | 4
CTUDT00632261 | 4 | CTUDT00632261 |  | Console-SW-Service Tool | Console-SW | 4
TBD_Outer laser adjustment | 3 | TBD_Outer laser adjustment |  | Gantry-HW-Noise | Gantry-HW | 4
CTUDT00631058 | 3 | CTUDT00631058 |  | Gantry-HW-Gantry Panel | Gantry-HW | 4
TBD-GMP | 3 | TBD-GMP
CTUDT00632298 | 3 | CTUDT00632298
CTUDT00631156 | 3 | CTUDT00631156
CTUDT00631634 | 3 | CTUDT00631634
TBD-GantryNoise | 3 | TBD-GantryNoise
CTUDT00631669 | 3 | CTUDT00631669
CTUDT00631671 | 3 | CTUDT00631671
TBD-PowerOff | 3 | TBD-PowerOff
CTUDT00632803 | 2 | Couch-HW | Couch-HW-Phantom
CTUDT00631283 | 2 | CTUDT00631283
CTUDT00632177 | 2 | CTUDT00632177
TBD-USBExtended | 2 | TBD-USBExtended
TBD-DoseReport | 2 | Console-SW | Console-SW-Dose SR
CTUDT00631545 | 2 | CTUDT00631545 |  | Row Labels | Count of # | Count of #2
TBD_UPS | 2 | TBD_UPS |  | Console-SW-Remote Connection | 10 | 11.76%
TBD-SAS | 2 | TBD-SAS |  | NotEngineering | 7 | 20%
CTUDT00630969 | 2 | CTUDT00630969 |  | Couch-SW-TBD-bugrep | 5 | 25.88%
CTUDT00632187 | 2 | CTUDT00632187 |  | Gantry-HW-Belt | 5 | 31.76%
CTUDT00631055 | 2 | CTUDT00631055 |  | TBD-TBD-NoEnoughInfo | 4 | 36.47%
CTUDT00632986 | 2 | Console-SW | Console-SW-ResultSelectionJump | Console-SW-Yellow-Screen | 3 | 40%
TBD-Initilization | 2 | Console-SW | Console-SW-Performance | Gantry-HW-CT Box | 3 | 43.53%
CTUDT00631165 | 2 | CTUDT00631165 |  | Console-HW-BlackScreen | 3 | 47.06%
CTUDT00631638 | 2 | CTUDT00631638 |  | Gantry-HW-Noise | 2 | 49.41%
CTUDT00632114 | 2 | CTUDT00632114 |  | Gantry-SW-Tube Heat | 2 | 51.76%
CTUDT00631143 | 2 | CTUDT00631143 |  | IC-HW-Tub Arc | 2 | 54.12%
CTUDT00632125 | 2 | CTUDT00632125 |  | DMS-SW-Phantom Detect | 2 | 56.47%
ECR-077066 | 2 | ECR-077066 |  | Couch-HW-ShouldbeAssembly | 2 | 58.82%
TBD-ExamSummary | 1 | Console-SW | Console-SW-ExamSummary
TBD-ServiceTool | 1 | Console-SW | Console-SW-Service tool
CTUDT00632183 | 1 | CTUDT00632183
TBD_Restart | 1 | TBD_Restart
TBD-Ghost | 1 | TBD-Ghost
CTUDT00631717 | 1 | CTUDT00631717
TBD-PRS | 1 | Console-SW | Console-SW-PRSPort
CTUDT00630958 | 1 | CTUDT00630958
TBD-UHR | 1 | TBD-UHR
CTUDT00632189 | 1 | CTUDT00632189
TBD-DVDDrive | 1 | Console-HW | Console-HW-DVDDrive
CTUDT00632190 | 1 | CTUDT00632190
CTUDT00631140 | 1 | CTUDT00631140
CTUDT00632236 | 1 | CTUDT00632236
TBD-HeadCushion | 1 | TBD-HeadCushion
TBD_KV_Perf_Bugrep | 1 | IC-HW | IC-HW-APM/CPM
TBD-Performance | 1 | Console-SW | Console-SW-TBD-Performance
TBD_NoEnoughInfo | 1 | CIRS-SW | CIRS-SW-TBD_NoEnoughInfo
CTUDT00630991 | 1 | CTUDT00630991
TBD_Noise | 1 | Gantry-HW | Gantry-HW-TBD_Noise
TBD-SW-Slow | 1 | Console-SW | Console-SW-Performance
CTUDT00630864 | 1 | CTUDT00630864
TBD-WW | 1 | Console-SW | Console-SW-WW
CTUDT00631134 | 1 | CTUDT00631134
TBD-Couch Relay | 1 | TBD-Couch Relay
CTUDT00632584 | 1 | CTUDT00632584
TBD-Encoder | 1 | TBD-Encoder
CTUDT00632618 | 1 | CTUDT00632618
TBD-GantryError | 1 | TBD-GantryError
TBD_Vertical Transducer | 1 | Couch-HW | Couch-HW-VerticalTransducer
TBD-GasPring | 1 | TBD-GasPring
TBD_Workstation | 1 | Console-SW | Console-SW-DeviceConfig
CTUDT00631142 | 1 | CTUDT00631142
TBD-ACQ Module | 1 | Console-HW | Console-HW-ACQ-Module
TBD-HostNoise | 1 | Console-HW | Console-HW-Noise
TBD-ACQ-Not-OK | 1 | TBD-ACQ-Not-OK
CTUDT00632175 | 1 | CTUDT00632175
TBD-AMC | 1 | Couch-HW | Couch-HW-AMC
CTUDT00630971 | 1 | CTUDT00630971
CTUDT00632121 | 1 | CTUDT00632121
TBD-ReconSlow | 1 | CIRS-SW | CIRS-SW-ReconSlow
TBD-Artifact | 1 | DMS-SW | DMS-SW-Artifact
CTUDT00628691 | 1 | CTUDT00628691
CTUDT00632793 | 1 | Couch-HW | Couch-HW-Matress
TBD-SW-IQ | 1 | DMS-SW | DMS-SW-IQ
TBD-BugRepError | 1 | Console-SW | Console-SW-BugRepError
TBD-TubGenerate | 1 | IC-HW | IC-HW-APM/CPM
TBD-Button | 1 | Couch-HW | Couch-HW-UnlockButton
TBD-Calibration | 1 | TBD-Calibration
TBD_IQ | 1 | Couch-HW | Couch-HW-H-Phantom
TBD-CanCard | 1 | TBD-CanCard
TBD-CardiacShootIssue | 1 | Gantry-SW | Gantry-SW-IBOX2
TBD-CCT Performance_Bugrep | 1 | Console-SW | Console-SW-CCT Performance
CTUDT00631039 | 1 | CTUDT00631039
CTUDT00631545/CTUDT00630971 | 1 | #N/A
CTUDT00632121 | 1 | CTUDT00632121

## Engineer-SlidesTable
Leve1 | Level2 | Level3 | Level2+3 | Root causes | Countermeasures | Countermeasures Short Term | Action (PA) Long Term | Owner | Timing | Count of Complaint (PR#) No. | Percentage | Progress | AR list
System | Console-SW | Remote Connection | Console-SW-Remote Connection | Workflow and one scenario of image transfer stuck was fixed. Others Still not clear. USP is doing investigation. | Temporay: Restart Host and retry. Permanent: Trigger code change and merge to SP1. | Restart Host and retry. | Trigger code change and merge to SP1. | Sankari S | SP1 | 32 | 3.271 | 50% | CTUDT00632127(23) - new open, track 23 complaints, similar case of 1602. CTUDT00632114(2) - new open to track image transfer issue related to network connection CTUDT00631602(7) - original AR, root cause not clear CTUDT00631283 (3)- worklist configuration issue, closed. CTUDT00630958 (1) - once scenario of image transfer stuck issue, fixed in WIP34.
System | Couch-SW | Horizontal Move Block | Couch-SW-Horizontal Move Block | Couch horizontal motion stop by mechanical stop | Temporay: Refresh hotfix 1 SW patch released to field Permanent: Couch SW update for manual motion stop improvement, released in WIP33 | Refresh hotfix 1 SW patch released to field | Couch SW update for manual motion stop improvement, released in WIP33 | Chang Xian Zhang | Aug 2020 | 19 | 2.1733 | 100% | CTUDT00630138/630160 CTUDT00631156 CTUDT00631669 CTUDT00631039
System | Gantry-HW | Tub Arc | Gantry-HW-Tub Arc | Tube conditioning aging not sufficient before shipment | Temporay: short tube conditioning by customer Permanent: Improve tube aging through at Philips with software update, released in WIP33 | short tube conditioning by customer | Improve tube aging through at Philips with software update, released in WIP33 | Michael Yu | SP1 | 7 | 0.80049 | 100%
System | Console-SW | WF-Yellow Screen | Console-SW-WF-Yellow Screen | two application access PatientDirectory.cfg at same time. | Temporay: Need user Relogin system Permanent: Trigger code change and merged in SP1 WIP34 | Need user Relogin system | Trigger code change and merged in SP1 WIP34 | Sankari S | SP1 | 6 | 0.68614 | 75%
System | Gantry-HW | Gantry laser | Gantry-HW-gantry laser | clearance between laser cable and rotor not under well control | Temporay: FSE to replace the laser and ensure enough clearance between cable and rotor Permanent: redesign cable routing to ensure enough clearance to rotor (E-127321) | FSE to replace the laser and ensure enough clearance between cable and rotor | redesign cable routing to ensure enough clearance to rotor (E-127321) | Andy Shao | ECO SCE | 3 | 0.34307 | 75%
System | Couch-HW | Horizontal linkage | Couch-HW-Horizontal linkage | Couch horizontal Linkage damage | Temporay: Replaced with new designed linakge FRU Permanent: New linkage design improvement done. Forward production cut in in Q2,2020 | Replaced with new designed linakge FRU | New linkage design improvement done. Forward production cut in in Q2,2020 | Tommy Sun | ECO SCE | 3 | 0.34307 | 100%
System | Couch-HW | Phantom | Couch-HW-Phantom | Phantom broken | Temporay: Replace with new Hardware. Permanent: Design change shall be initiated by ECR-074469 and implemented by E-127916. | Replace with new Hardware. | Design change shall be initiated by ECR-074469 and implemented by E-127916. | Ada Wang | ECO SCE | 2 | 0.22871 | 75%
System | Gantry-HW | Lexan Ring | Gantry-HW-Lexan Ring | Lexan ring size is big than bore R&D planned the Lexan Ring change in refresh but rejected by refresh team | Temporay: FSE to adjust Lexan ring in field Permanent: To implement updated Lexan ring design in refresh | FSE to adjust Lexan ring in field | To implement updated Lexan ring design in refresh | Lucas | ECO SCE | 2 | 0.22871 | 75%
System | Couch-HW | Vertical spring | Couch-HW-Vertical spring | Spring is not adjust well | Temporay: OPS WI updated for help spring alignment Permanent: Update spring installation design to avoide assembly and noise issue | OPS WI updated for help spring alignment | Update spring installation design to avoide assembly and noise issue | Tommy Sun | ECO SCE | 2 | 0.22871 | 50%
System | DMS-HW | Compensator | DMS-HW-Compensator | Air Cal Failure Cracked Compensator message | Temporay: Replace cracked compensator Permanent: Do Beam and DMS check in OPS site to capture the issue in early phase | Replace cracked compensator | Do Beam and DMS check in OPS site to capture the issue in early phase | Silva | SP1 | 2 | 0.22871 | 75% | New Added
System | Console-SW | Slab Viewer | Console-SW-Slab Viewer | User use wrong operation to do image roation. | Temporay: Training user with correct operation. Permanent: Add training steps in clinical application training. | Training user with correct operation. | Add training steps in clinical application training. | Caroline | SP1 | 2 | 0.22871 | 75%
System | Gantry-HW | R2D | Gantry-HW-R2D | R2D design change (U1 EOL triggered)verificatio nnot sufficient | Temporay: verify the impct of current R2D and confirmed no impact on performance, communicated to gloabel service already Permanent: Update R2D design | verify the impct of current R2D and confirmed no impact on performance, communicated to gloabel service already | Update R2D design | Yaping | ECO SCE | 1 | 0.11436 | 100%
System | Gantry-HW | gantry cover | Gantry-HW-Gantry cover | Cabel not fixed | Temporay: Adjust and fix the cable Permanent: Drawing update and supplier quality control ECR-075587 submitted to fix drawing update | Adjust and fix the cable | Drawing update and supplier quality control ECR-075587 submitted to fix drawing update | Andy Shao | ECO SCE | 1 | 0.11436 | 75%
System | Couch-HW | Scan range |  | Couch scan line definition is not defined | Temporay: Create TRR to define scanline value. Update mattress and service manual Permanent: Define scan line value in accessory SSDS with SE | Create TRR to define scanline value. Update mattress and service manual | Define scan line value in accessory SSDS with SE | Ada Wang | TRR | 1 | 0.11436 | 25% | Removed, it is Application
System | DMS-SW | Help Page | DMS-SW-Help Page | wrong instruction in help instruction | Temporay: Uncheck the Infant phantom in preference setting Permanent: Infant phantom always been used and tested in OPS, service calibration manual inlucde step to the uncheck step according to customer order. | Uncheck the Infant phantom in preference setting | Infant phantom always been used and tested in OPS, service calibration manual inlucde step to the uncheck step according to customer order. | Ariel | SP1 | 1 | 0.11436 | 75%
System | Console-SW | LanConfig | Console-SW-LanConfig | Multiple IPs configured the Lanconfig is not able to load more than 2 IPs. | Temporay: Uncheck IPv6 when the connection is no longer supported. Permanent: Trigger code changed and merge in WIP33. | Uncheck IPv6 when the connection is no longer supported. | Trigger code changed and merge in WIP33. | Sankari S | SP1 | 1 | 0.11436 | 75%
System | Console-SW | Spectral Error Message | Console-SW-Spectral Error Message | Some Image DICOM tags contains Spectral related annotations which trigger the warning message. | Temporay: User click okay, and can continue the image transfer. Permanent: Trigger code change and merge in SP1. | User click okay, and can continue the image transfer. | Trigger code change and merge in SP1. | Shubdeep | SP1 | 1 | 0.11436 | 50%
System | Console-SW | InstallationVersion | Console-SW-InstallationVersion | Adobe keep update lead non standard installation | Temporay: Give feedback to user. Permanent: Trigger code changed and merged in SP1 | Give feedback to user. | Trigger code changed and merged in SP1 | Shubdeep | SP1 | 1 | 0.11436 | 50%

## Engineer-ComplaintARFullLi
IDX | Defect (CTUDT#) | Count of PR# | Func.Owner | Description | RDState | Score/Pre-score | Hazard | Subsystem | Owner | Resolution | In HF3 Scope | Resolution | Refresh/IB | PR(y/n) | OneEMS | DEFOA | AllFieldFeedback | SW/HW | Status | Issue description | Corrective actions | Score | Subsystem | Due Date | Actual Closure | Note | Product | Level | Level1 | Level3 | Level2- | Level3 | Root causes | Countermeasures Short Term | Action (PA) Long Term | Owner | Timing | Impact on Gap | Current Status | Submit Date | Submit Date wk | DefectIn | RDFixDate | RDFixDate WK | Defect out | Complaint Contribution | SS | Check
1 | CTUDT00632127 | 22 | Engineering | FOK: More Sites meet Image transfer stuck issue need to restart host to recover | Investigated | 2 | NO | ICAP | Sankari | SP1 | Should fix, GAP | No solution now. Need transfer image to remote deivce from other remote device. | Refresh | Y | Y | N | PRCall | SW | 0 | Image transfer stuck | [7 empty] | ING-Console-SW-Remote Connection-Image transfer stuck | ING | Remote Connection | Console-SW | Remote Connection | Still not clear. Pending bug report collection. | Need user Restart Host and retry. | Should trigger code change and merge to SP1. | Sankari S | SP1 |  |  | December 24, 2020 12:25:28 AM GMT+08:00 | WK52 | 52 | NotFix yet | NA | NA | 0 | y | 22
2 | CTUDT00630160 | 12 | Duplicate | FOK - Couch horizontal motion stop by mechanical stop | Closed | 2 | NO | Couch | cx_zhang | Fix-HF1 | NoNeed | Fixed with FCO72800735 | Refresh | Y | Y | N | PRCall | SW | 1 | Couch FCO7280035 | [7 empty] | ING-Couch-SW-Horizontal Move Block | ING | Horizontal Move Block | Couch-SW | Horizontal Move Block | Couch horizontal motion stop by mechanical stop | Refresh hotfix 1 SW patch released to field | Couch SW update for manual motion stop improvement, released in WIP33 | Chang Xian Zhang | 4422700% |  |  | June 25, 2020 6:41:00 PM GMT+08:00 | WK26 | 26 | Jul 20 2020 | WK30 | 30 | 0 | y | 12
3 | CTUDT00630138/630160 | 12 | Engineering | FOK - Couch horizontal motion stop by mechanical stop | Closed | 2 | YES | Couch | cx_zhang | Fix-HF1 | HF1 fixed | Fixed with FCO72800735 | Refresh | Y | Y | N | PRCall | SW | 1 | Couch FCO7280035 | [7 empty] | ING-Couch-SW-Horizontal Move Block | ING | Horizontal Move Block | Couch-SW | Horizontal Move Block | Couch horizontal motion stop by mechanical stop | Refresh hotfix 1 SW patch released to field | Couch SW update for manual motion stop improvement, released in WIP33 | Chang Xian Zhang | 4422700% |  |  | 4400400% | WK26 | 26 | Jul 20 2020 | WK30 | 30 | 12 | #N/A | #N/A
4 | CTUDT00632110 | 7 | Engineering | FOK: Bad Detector calibration failed and ring artifacts occurred | Closed |  | NO | DMS |  | WAS | N | [18 empty] | DMS Module | DMS-HW | DMS | [7 empty] | December 23, 2020 11:09:54 AM GMT+08:00 | WK52 | 52 | NotFix | NA | NA |  | y | 7
5 | CTUDT00630085 | 7 | Engineering | The tube can not warm up and light TBC-TubAcr | Closed | 4 | NO | GCX | Andy | WIP33 | N |  | IB | Y | Y | Y | PR | HW | 0 | Tub Arch | [7 empty] | ING-Gantry-HW-Tub Arc | ING | Tub Arc | IC-HW | Tub Arc | Tube conditioning aging not sufficient before shipment | short tube conditioning by customer | Improve tube aging through at Philips with software update | Silva/Goblin | TBD |  |  | June 16, 2020 2:57:07 PM GMT+08:00 | WK25 | 25 | October 26, 2020 1:21:02 PM GMT+08:00 | WK44 | 44 | 7 | y | 7
6 | CTUDT00631602 | 7 | Engineering | FOK: Several FOK site reply back unsuccessful sending images to PACS again. | Accepted | 2 | NO | ICAP | Sankari | SP1 | Should fix, GAP | No solution now. Need transfer image to remote deivce from other remote device. | Refresh | Y | Y | N | PRCall | SW | 0.25 | PACS Issue | [7 empty] | ING-Console-SW-Remote Connection | ING | Remote Connection-1602TBD | Console-SW | Remote Connection | Still not clear. USP is doing investigation. | Need user Restart Host and retry. | Should trigger code change and merge to SP1. | Sankari S | SP1 |  |  | October 22, 2020 1:40:48 PM GMT+08:00 | WK43 | 43 | AUG 30 2021 | WK88 | 88 | 7 | n | 7
7 | CTUDT00631367/CTUDT00628691 | 6 | Engineering | FOK: Yellow screen issue happens when use workflow 8691 combined with this item together. | Resolved | 2 | NO | Console4 | modi_s | WIP34 | Y | 0 | Refresh |  |  |  |  | SW | 0 | Yellow screen | [7 empty] | ING-Console-SW-Work flow | ING | #N/A | #N/A | WF-Yellow Screen | two application access PatientDirectory.cfg at same time. | Need user Relogin system | Trigger code change and merged in SP1 WIP34 | Sankari S | 4446600% |  | 90% | 4410300% | WK40 | 40 | May 13 2020 | WK20 | 20 |  | y | 6
8 | CTUDT00632116 | 5 | Engineering | Axial scan MPR Recon- thin slice, not fully cover recon range | Accepted |  | YES | Console4 |  | Open-SW | N | [6 empty] | SW | [11 empty] | MPR Recon | Console-SW | MPR Recon | [7 empty] | December 23, 2020 2:22:29 PM GMT+08:00 | WK52 | 52 | NotFix YET | TBD | TBD | 0 | y | 5
9 | CTUDT00632181 | 4 | Engineering | FOK - GCP control panel issues are reported from 4 FOK sites. | Accepted |  | NO | GCX | [21 empty] | Gantry Panel | Gantry-HW | [8 empty] | January 5, 2021 9:55:24 PM GMT+08:00 | WK54 | 54 | AUG 30 2021 | WK88 | 88 | 4 | n | FOK
10 | CTUDT00631156 | 3 | Engineering | FOK: Couch horizontal movement issue happens. | Closed | 4 | NO | couch | Robin Zhou | Fix-HW | N | Do mechanical adjust for field machine. Will fix completely in SP | Refresh | Y | Y | N | PRCall | HW | 0.25 | Couch Travel Range - HW limit | [7 empty] | ING-Couch-HW-Mechanism Range | ING | Horizontal Move Block | Couch-SW | Horizontal Move Block | Couch horizontal motion stop by mechanical stop | Refresh hotfix 1 SW patch released to field | Couch SW update for manual motion stop improvement, released in WIP33 | Chang Xian Zhang | 4422700% |  |  | September 16, 2020 2:18:52 PM GMT+08:00 | WK38 | 38 | Oct 27 2020 | WK44 | 44 | 3 | y | 3
11 | CTUDT00631634 | 3 | Engineering | FOK - Gantry laser cable short circuit or damage can cause Ghost reboot to trigger EStop | Closed | 4 | YES | GCX |  | Fix-ECR | N | 0 | IB | Y | Y | N | PRCall | HW | 0 | TBC-Front cover improvement-lasercable | [7 empty] | ING-Gantry-HW-Laser Cabel | ING | gantry laser | Gantry-HW | gantry laser | clearance between laser cable and rotor not under well control | FSE to replace the laser and ensure enough clearance between cable and rotor | redesign cable routing to ensure enough clearance to rotor | Andy Shao | E-127321 released, now in supplier chain enrichment |  | 100% | October 26, 2020 10:29:00 AM GMT+08:00 | WK44 | 44 | Sep 1 2020 | WK36 | 36 | 1 | y | 3
12 | CTUDT00631669 | 3 | Engineering | FOK: Couch manual control moving cause ESTOP in the hospital | Closed | 3 | NO | Couch | cx_zhang | Open-HW | N | under investigation. | IB | Y | Y | N | PRCall | HW | 0 | CouchHW-ESTOP | [7 empty] | ING-Couch-HW-Couch move trigger ESTOP | ING | Horizontal Move Block | Couch-SW | Horizontal Move Block | Couch horizontal motion stop by mechanical stop | Refresh hotfix 1 SW patch released to field | Couch SW update for manual motion stop improvement, released in WIP33 | Chang Xian Zhang | 4422700% |  |  | October 28, 2020 10:19:16 AM GMT+08:00 | WK44 | 44 | Dec 27 2020 | WK52 | 52 |  | y | 3
13 | CTUDT00631671 | 3 | Engineering | Couch linkage issue | Resolved | 2 | YES | Couch | Robin Zhou | Open-HW | N | 0 | IB | Y | Y | N | PRCall | HW | 0 | Couch-linkage | [7 empty] | ING-Couch-HW-Linkage | ING | Horizontal linkage | Couch-HW | Horizontal linkage | Couch horizontal Linkage damage | Replaced with new designed linakge FRU | New linkage design improvement done. Forward production cut in in Q2,2020 | Tommy Sun | 4395100% |  |  | October 28, 2020 1:52:18 PM GMT+08:00 | WK44 | 44 | January 4, 2021 3:45:44 PM GMT+08:00 | WK54 | 54 | 3 | n | 3
14 | CTUDT00632173 | 3 | Engineering | FOK: Several FOK site reply back remote device retrieve or transfer failure. | Accepted | 4 | NO | CIRS | Veer | Close-WAS | N | WAS | Refresh | Y | Y | N | PRCall | Not Defect | 0.5 | Usr wrong setting | [7 empty] | ING-Console-SW-Remote Connection | ING | Adaptor | Console-HW | Remote Connection | Customer Wrong operation | Give feedback to user | N/A, closed. | Shubdeep | 4409500% |  | 100% | January 5, 2021 2:40:48 PM GMT+08:00 | WK54 | 54 | AUG 30 2021 | WK88 | 88 | 3 | n | FOK
15 | CTUDT00632179 | 3 | Engineering | FOK - Horizontal movement coupling happens in customer site. three sites report similar issue. | Accepted |  | YES | Couch | [21 empty] | Motor Coupling | Couch-HW | [8 empty] | January 5, 2021 9:33:30 PM GMT+08:00 | WK54 | 54 | AUG 30 2021 | WK88 | 88 | 3 | n | FOK
16 | CTUDT00632185 | 3 | Engineering | FOK - CT Box buttons malfunction, reported in three FOK sites. | Accepted |  | NO | GCX | [21 empty] | CT Box | Gantry-HW | [8 empty] | January 6, 2021 12:02:41 AM GMT+08:00 | WK54 | 54 | AUG 30 2021 | WK88 | 88 | 3 | n | FOK
17 | CTUDT00631545 | 3 | Engineering | FOK: Phantom not seen message during IQ check | Postponed | 3 | NO | DMS |  | Close-phantom | N | 0 | Refresh | Y | Y | N | PRCall | SW | 0 | Phantom Detect | [7 empty] | ING-DMS-SW-Auto Centering | ING | Phantom Detect | DMS-SW | Auto Centering | The Phantom detection algorithm failed to recognize the thin Body Phantom. | Workaround provided: move the phantom or the couch a little bit and re-try, and it pass. | Trigger code changes to replace current tolerance based methodology can fix this, but postpone it and keep observation in FOK period. | Ariel |  |  |  | October 20, 2020 9:48:16 AM GMT+08:00 | WK43 | 43 | Not fix , 29 Oct 2020 | NA | NA | 0 | n | 3
18 | CTUDT00631136 | 2 | OPS | FOK: Recon - Slab Tool-Image Rotation not working properly | Closed | 3 | NO | Console4 | Gaurav | Close-WAS | NA | WAS | IB | Y | Y | N | PRCall | Not Defect | 0.25 | Image Rotation in Slab tool | [7 empty] | ING-Console-SW-Work flow | ING |  | Console-SW | Work flow | User use wrong operation to do image roation. | Give feedback to user | N/A, closed. | Shubdeep | 4408900% |  | 100% | September 16, 2020 12:02:40 AM GMT+08:00 | WK38 | 38 | Fix by clinical training Nov 2 2020 | WK45 | 45 | 2 | #N/A | #N/A
19 | CTUDT00631283 | 2 | Engineering | FOK: Several FOK site reply back remote device retrieve or transfer failure(Work list populating issue) | Closed |  | NO | Console4 | [11 empty] | Worklist config issue | [9 empty] | Remote Connection-Worklist | Console-SW | Remote Connection-worklist | worklist configuration issue | WAS | [5 empty] | September 21, 2020 9:31:32 PM GMT+08:00 | WK39 | 39 | Oct 27 2020 | WK44 | 44 | 2 | n | 2
20 | CTUDT00630136 ECR-074469 | 2 | Engineering | FOK - Adjusting knob on phantom holder is bent. some phantom damage. And phantom only include system phantom label, no head body phantom diff. | RD Close |  | NO | Couch |  | Fix-ECR | N | [6 empty] | HW | [11 empty] | #N/A | #N/A | Phantom | Phantom damage by user or during transport phase | Replace broken phatnom with new one. | Do design improvement to add Knob, labels according to customer feedback. | Ada Wang | 4386000% |  |  | 4400400% | WK26 | 26 | Aug 25 2020 | WK35 | 35 | 2 | #N/A | #N/A
21 | CTUDT00630969 | 2 | Engineering | FOK: Mt Grant Hospital Routine Air Cal Failure Cracked Compensator message | Resolved | 2 | NO | Support-System | Yeshan_M | Fix-ECR | N | Trigger MTS update and WI update. | NA | Y | Y | N | PRCall | SW | 0.75 | Beam+DMS check | [7 empty] | ING-DMS-HW-Compensator | ING | Compensator | DMS-HW | DMS Module | Air Cal Failure Cracked Compensator message | Temporary: Replace cracked compensator Permanent: Do Beam and DMS check in OPS site to capture the issue in early phase |  | Silva |  |  |  | September 1, 2020 7:24:26 PM GMT+08:00 | WK36 | 36 | January 4, 2021 1:49:34 PM GMT+08:00 | WK54 | 54 | 2 | y | 2
22 | CTUDT00631058 | 2 | Engineering | Headline: FOK: Design improvement considertion: Table base is creaking loudly when 12” from the completely down position and 12” away in the upward movement | Assigned | 4 | NO | Couch | Ada Wang | Fix-ECR | N | Adjust srping. Further investigation. | IB | Y | N | N | PR | HW | 0.5 | couch noise | [7 empty] | ING-Couch-HW-Spring | ING | Vertical spring | Couch-HW | Vertical spring | Spring is not adjust well | OPS WI updated for help spring alignment | Update spring installation design to avoide assembly and noise issue | Tommy Sun | 4386000% |  |  | September 9, 2020 2:46:22 PM GMT+08:00 | WK37 | 37 | DEC 31 2021 | WK105 | 105 |  | y | 2
23 | CTUDT00631143 | 2 | Engineering | FOK: Enable Remote Access not present in Home Menu. | Closed | 3 | NO | Console4 | Gaurav | Close-WAS | N | Impact servicability after service use enable US DoD config. RFC to add confirmation message. | Refresh | Y | N | N | PR | Not Defect | 0.5 | US DOD Tool | [7 empty] | ING-Console-SW-Service Tool | ING | Service Tool | Console-SW | Service Tool | Remote access window keep displaying on top of workflow. | Give feedback to user. | N/A, closed. | Shubdeep | 4408900% |  | 100% | September 16, 2020 11:54:21 AM GMT+08:00 | WK38 | 38 | NO PLAN TO FIX | NA | NA | 0 | y | 2
24 | CTUDT00631165 | 2 | Engineering | FOK: Customer Feedback Willis Knighton EStop on Couch Zero | Closed | 3 | NO | GCX | Jerry Tan | Fix-HW | N | hw Replace ment. | Refresh | Y | Y | N | PRCall | SW | 0.25 | Estop couch zero | [7 empty] | ING-Gantry-HW-Control Panel | ING | Horizontal motor | Couch-HW | Horizontal motor | not clear, not reproducable, log reviewed and found Ghost reset. | mornitor this site to track if this issue will happen again. | NA | Jerry Tan | TBD |  |  | September 16, 2020 5:43:13 PM GMT+08:00 | WK38 | 38 | nov 3 2020 | WK45 | 45 | 2 | y | 2
25 | CTUDT00632114 | 2 | Engineering | Image transfer to Pacs stuck due to broken connection | Accepted |  | NO | Console4 |  | TBD | N | [6 empty] | SW | [11 empty] | Remote Connection-Broken | Console-SW | Remote Connection | image transfer stuck due to net work connection issue | User retart Host to recover |  |  | SP1 |  |  | December 23, 2020 1:56:29 PM GMT+08:00 | WK52 | 52 | AUG 30 2021 | WK88 | 88 | 2 | y | 2
26 | CTUDT00632125 | 2 | Engineering | Impossible to use Arm Rest with current design. | Accepted |  | NO | Couch |  | Open-HW | N | [6 empty] | HW | [11 empty] | Arm Rest | Couch-HW | Arm Rest | Velcro design make it's impossible to use Arm Rest. | Put Arm Rest between pad and top. | Trigger design investigation and make final decision how to improve. | Ada Wang | 4386000% |  |  | December 23, 2020 11:27:59 PM GMT+08:00 | WK52 | 52 | NO PLAN TO FIX | TBD | TBD | 0 | y | 2
27 | CTUDT00632187 | 2 | Engineering | FOK - E-Stop open during unload patient. | Investigated |  | NO | Couch | [21 empty] | Estop | Couch-SW | [8 empty] | January 6, 2021 9:40:40 AM GMT+08:00 | WK54 | 54 | 4422600% | WK57 | 57 |  | y | 2
28 | TBC-Lexan Ring ECR-075975 | 2 | Engineering | Lexan Ring size is bigger than bore size | ECR Track | 2 | NO | GCX | WILLAIM | Open-ECR | N | Need check design tolerance from RD level. | Refresh | Y | Y | Y | PR | HW | 0.25 | Lexan Ring | [7 empty] | ING-Gantry-HW-Lexan Ring | ING | Lexan Ring | Gantry-HW | Lexan Ring | updated Lexan Ring design not implemented in Refresh cover installation is not well adjusted (adjustment is not easy) R&D planned the Lexan Ring change in refresh but rejected by refresh team | FSE to adjust Lexan ring in field | to implement updated Lexan ring design in refresh | Lucas | E-128610 released, now in supplier chain enrichment |  |  | 4402100% | WK28 | 28 | Jul 209 2020 | WK31 | 31 | 2 | #N/A | #N/A
29 | CTUDT00628691 | 1 | Duplicate | Val:Refresh: Patient registration page blank after end exam and Patient Directory stopped working while executing Performance Test | Resolved | 3 | NO | ICAP | 0 | WIP30 | Y | 0 | Refresh | Y | Y | N | PRCall | SW | 0.75 | patient regist page blank | [7 empty] | ING-Console-SW-Work flow | ING | #N/A | Console-SW | Work flow | two application access PatientDirectory.cfg at same time. | Need user Relogin system | Trigger code change and merged in SP1 WIP34 | Sankari S | 4412500% |  | 90% | February 13, 2020 9:04:03 PM GMT+08:00 | WK7 | 7 | May 13, 2020 6:41:20 PM GMT+08:00 | WK20 | 20 | 0 | y | 1
30 | TBC-washer | 1 | OPS | Washer in installation kit is not fit field installation size | OPS | 2 | NO | GCX | William | Open-HW | NoNeed |  | IB | Y | Y | Y | PR | HW | 0 | Washer | [7 empty] | ING-Gantry-HW-Washer | ING | #N/A | Gantry-HW | Washer | Supplier delivered wrong (smaller) washer to Suzhou Philips | Local purchase qualified washer for installation | to be updated by SQ and MQ | Alan-->Tianxing Chen | 4422600% |  |  | #N/A | #N/A | #N/A | #N/A | #N/A | NA | 1 | #N/A | #N/A
31 | CTUDT00629920 | 1 | Engineering | Gaps of Instruction, Help Page and Physics Spec for Calibration with H-Phantom | Closed | 1900-01-03 | NO | DMS |  | WIP33 | N | [6 empty] | SW | [11 empty] | Help Page | DMS-SW | Help Page | Some error message list in Help page of phantom calibration instruction | Send information to FSE and performing correponding calibration | Trigger code fix to fix it in WIP33. | Limor | 4414700% |  |  | May 27, 2020 10:42:21 PM GMT+08:00 | WK22 | 22 | October 11, 2020 4:25:36 PM GMT+08:00 | wk41 | 41 | 1 | #N/A | #N/A
32 | CTUDT00630134 ECR-075587 | 1 | Engineering | FOK - Cable Harness Rear Cover Close Up been pressed by screw hole | Closed |  | NO | GCX |  | Fix-ECR | N | [6 empty] | HW | [11 empty] | gantry cover | Gantry-HW | gantry cover | Supplier didn't follow design to make the part. | Remove cable from the screw hole. | Trigger design improvement to avoid supplier wrong instruction. | Andy Shao | 4386000% |  |  | 4400400% | WK26 | 26 | June 23 2020 | WK26 | 26 |  | #N/A | #N/A
33 | CTUDT00630864 | 1 | Engineering | FOK: System rotation issue when first installation on site. | Closed | 2 | NO | GCX | wang_c | Fix-HW | N | had fixed in field with R2D replacement. Further check need after getting the HW | IB | Y | N | N | PR | HW | NA | R2D broken | [7 empty] | ING-Gantry-HW-R2D | ING | R2D | Gantry-HW | R2D | R2D design change (U1 EOL triggered)verificatio nnot sufficient | verify the impct of current R2D and confirmed no impact on performance, communicated to gloabel service already | update R2D design | Yaping | E-128870 released, now in supplier chain enrichment (all pproved) |  |  | August 25, 2020 10:41:23 AM GMT+08:00 | WK35 | 35 | NO PLAN TO FIX | NA | NA | 0 | y | 1
34 | CTUDT00630958 | 1 | Engineering | Image transfer to PACS problem | Resolved | 2 | NO | ICAP | Sankari | WIP33 | Y | Fixed in SP1 | Refresh | Y | Y | N |  | SW | [11 empty] | Remote Connection | Console-SW | Remote Connection | the transfer is waiting to complete the read endlessly and showing the status Waiting in Queue | As a workaround , perform spooler reset\restart or restart the workflow | Code fix and merge to SP1. Not completed fix, need full fix in 1602. | Sankari S | WIP33 |  |  | August 31, 2020 1:46:48 PM GMT+08:00 | WK36 | 36 | October 23, 2020 4:40:57 PM GMT+08:00 | wk43 | 43 | 2 | y | 1
35 | CTUDT00630971 | 1 | Engineering | FOK: During IQ check, scanner stated no phantom was detected, but actually the phantom is on the scout. | Closed | 4 | NO | Support-Physics | braunstein_d | Close-phantom | N | DMS investigation: The described flow might occur when the phantom is positioned in a way that the holder screw is at it's out most position. In such case, it is recommended to reposition the phantom and start the process over. | Refresh | Y | N | N | PR | SW | 1 | Phantom Detect | [7 empty] | ING-DMS-SW-Auto Centering | ING | Phantom Detect | DMS-SW | Auto Centering | The Phantom detection algorithm failed to recognize the thin Body Phantom. | Workaround provided: move the phantom or the couch a little bit and re-try, and it pass. | N/A, closed. | Ariel |  |  |  | September 1, 2020 7:42:17 PM GMT+08:00 | WK36 | 36 | NO PLAN TO FIX | NA | NA | 0 | y | 1
36 | CTUDT00630991 | 1 | Engineering | FoK: Installation is non standard shown in installation validation tool. | Opened | 4 | NO | Console4 | Gaurav | SP1 | N | need code fix. Current, give feedback to user that ignore the gap | Refresh | Y | N | N | PR | SW | 0.5 | Adobe auto upgrade | [7 empty] | ING-Console-SW-Adobe | ING | InstallationVersion | Console-SW | InstallationVersion | Adobe keep update lead non standard installation | Give feedback to user. | Trigger code changed and merged in SP1 | Shubdeep | 4407700% |  |  | September 3, 2020 1:46:45 PM GMT+08:00 | WK36 | 36 | 4443800% | WK88 | 88 | 1 | y | 1
37 | CTUDT00630796 | 1 | OPS | FOK: R2D LED display wrong status in field. | Closed | 4 | NO | GCX | yaping_l | Fix-ECR | N | Design change to avoid wrong instruction of LED | IB | Y | Y | N | PRCall | HW | 0.75 | R2D LED | [7 empty] | ING-Gantry-HW-R2D | ING | R2D | Gantry-HW | R2D | R2D design change (U1 EOL triggered)verificatio nnot sufficient | verify the impct of current R2D and confirmed no impact on performance, communicated to gloabel service already | update R2D design | Yaping | E-128870 released, now in supplier chain enrichment (all pproved) |  |  | August 19, 2020 12:34:42 PM GMT+08:00 | WK34 | 34 | 4407400% | WK36 | 36 | 1 | #N/A | #N/A
38 | CTUDT00631640 | 1 | OPS | FOK - Missing handle for legacy phantom of eye phantom | Closed | 2 | NO | Couch | Ada Wang | Fix-HW | NA | ECR-077026 | Refresh |  | Y | Y | Y | HW | 0 | Eye phantom holder | [7 empty] | ING-Couch-HW-Eye phantom holder | ING | 0 | Couch-HW | Eye phantom holder | BOM error for phantom group for different configuration | CBOM updated for BOM issue | Review with CM,product manager for CBOM update. Create check list for CM update change | Ada Wang | 4386000% |  |  | October 26, 2020 12:56:10 PM GMT+08:00 | WK44 | 44 | Oct 29 2020 | WK44 | 44 | 1 | #N/A | #N/A
39 | CTUDT00631039 | 1 | Engineering | FOK: Gantry not ready message after scanner sat idle for a few | Closed | 4 | YES | GCX | Jerry Tan | Fix-HF1 | HF1 fixed | close as duplicate and merge to HF1 FCO | TBD | Y | N | N | PR | SW | 0.25 | Gantry not ready | [7 empty] | ING-Couch-HW-Horizontal Move Block | ING | Horizontal Move Block | Couch-SW | Horizontal Move Block | Couch horizontal motion stop by mechanical stop | Refresh hotfix 1 SW patch released to field | Couch SW update for manual motion stop improvement, released in WIP33 | Chang Xian Zhang | 4416500% |  |  | September 8, 2020 10:41:51 AM GMT+08:00 | WK37 | 37 | nov 3 2020 | WK45 | 45 | 1 | y | 1
40 | CTUDT00631055 | 1 | Engineering | FOK:Refresh - Lan Config open with error if Network IPV4 add more than one static IP address | Closed | 3 | NO | ICAP | Gaurav | WIP33 | Y | disale IPV6 from customer site. Need code change | Refresh | Y | N | N | PR | SW | 0.5 | IPV6 | [7 empty] | ING-Console-SW-Lan Config | ING | Lan Config | Console-SW | Lan Config | Multiple IPs configured the Lanconfig is not able to load more than 2 IPs. | Uncheck IPv6 when the connection is no longer supported. | Trigger code changed and merge in WIP33. | Sankari S | 4412700% |  |  | September 9, 2020 9:23:09 AM GMT+08:00 | WK37 | 37 | October 23, 2020 4:34:10 PM GMT+08:00 | wk43 | 43 | 3 | y | 1
41 | CTUDT00631134 | 1 | Engineering | FOK: RFC: Customer feedback for 3 items which need be improved in ING system. | Closed | 4 | NO | Couch | silva | Close-RFC | N | RFC from customer | IB | Y | Y | N | PRCall | Not Defect | 0.25 | RFC | [7 empty] | ING-Couch-HW-RFC | ING | RFC | Couch-HW | RFC | Current couch design does not including the requirement for use gurney for patient transfer | DRB reviewed all the content and investigation info and agree to close it as Change Proposal and consider in next Gen Couch. | DRB reviewed all the content and investigation info and agree to close it as Change Proposal and consider in next Gen Couch. | N/A | N/A |  |  | September 15, 2020 11:47:52 PM GMT+08:00 | WK38 | 38 | Oct 29 2020 | WK44 | 44 | 1 | y | 1
42 | CTUDT00631140 | 1 | Engineering | FOK: The message box of remote control cannot be hidden. | Closed | 4 | NO | Console4 | Gaurav | Close-RFC | N | RFC from customer to hide remote control window. But from security perspective, it should not permit to hide. | IB | Y | N | N | PR | Not Defect | 0.5 | Remote Control | [7 empty] | ING-Console-SW-Service Tool | ING | Service Tool | Console-SW | Service Tool | This is the current design implemented in Refresh. This Banner (Message) doesn't not have the minimize button . It supports only movable to different position and always stays on top of the screen. | WAS and closed. | N/A, closed. | Shubdeep |  |  |  | September 16, 2020 9:25:19 AM GMT+08:00 | WK38 | 38 | sep 25 2020, not need fix | NA | NA | 0 | y | 1
43 | CTUDT00631142 | 1 | Engineering | FOK-Warning message "Spectral annotations check" pop up when adding ING images into ISP film. | Accepted | 3 | NO | Console4 | Gaurav | SP2 | N | need code fix. Current, give feedback to user that ignore the message | Refresh | Y | N | N | PR | SW | 0.25 | DICOM Spetral MSG | [7 empty] | ING-Console-SW-Spectral Error Message | ING | Spectral Error Message | Console-SW | Spectral Error Message | Some Image DICOM tags contains Spectral related annotations which trigger the warning message. | User click okay, and can continue the image transfer. | Trigger code change and merge in SP1. | Shubdeep |  |  |  | September 16, 2020 11:50:06 AM GMT+08:00 | WK38 | 38 | DEC 31 2021 | WK105 | 105 | 1 | n | FOK; SP2_Candidate
44 | CTUDT00631636 | 1 | Engineering | FOK - RFC: Customer wanted a long cross hair on the axials. Explained that does not exist on 4.1.7 | Accepted | 4 | NO | Console4 |  | Close-RFC | N | 0 | IB | Y | Y | N | PRCall | Not Defect | 0 | RFC-Cross hair | [7 empty] | ING-Console-SW-RFC | ING | RFC | Console-SW | RFC | No such function in Refresh Ing, this is RFC. | N/A | [5 empty] | October 26, 2020 10:37:38 AM GMT+08:00 | WK44 | 44 | DEC 31 2021 | WK105 | 105 | 1 | n | FOK; SP2_Score2Postpone
45 | CTUDT00631638 | 1 | Engineering | FOK - scannable range line on pad doesn’t match table line | Assigned | 2 | YES | Couch | Ada Wang | Fix-ECR | N | 0 | IB | Y | Y | N | PRCall | HW | 0 | scanrange line | [7 empty] | ING-Couch-HW-Scan range | ING | Mattress scan line | Couch-HW | Scan range | Couch scan line definition is not defined | Create TRR to define scanline value. Update mattress and service manual | Define scan line value in accessory SSDS with SE | Ada Wang | 4386000% |  |  | October 26, 2020 10:59:24 AM GMT+08:00 | WK44 | 44 | 4443800% | WK88 | 88 | 1 | y | 1
46 | CTUDT00631717 | 1 | Engineering | Artifact captured even pass Beam and DMS check | Closed | 3 | NO | DMS | Dmtry | Fix-HW | N | DMS Module need replace from HW level. But also need consider why Beam + DMS check not captured. | Refresh | Y | Y | N | PRCall | HW | 0.25 | Artifact | [7 empty] | ING-DMS-HW-Artifact | ING | DMS Module | DMS-HW | Artifact | Artifact Captured during clinical scan DMS Modue Problem, Need replace. | Temporary: Replace DMS Module Permanent: Send back to HFA DMS team and do trouble shooting | N/A, closed. | Ariel | 4413700% |  | 100% | November 2, 2020 2:40:54 PM GMT+08:00 | WK45 | 45 | Nov 4 2020, part replacement | WK45 | 45 | 1 | y | 1
47 | CTUDT00632121 | 1 | Engineering | FOK: Exam Summary does not show the dose for the scan unless has the SR report turned on & Can not uncheck PACS as a send to Destination if PACS is listed as a Default Destination. | Accepted |  | NO | Console4 |  | Open-SW | N | [6 empty] | SW | [11 empty] | Dose SR | Console-SW | Dose SR | [7 empty] | December 23, 2020 10:35:54 PM GMT+08:00 | WK52 | 52 | AUG 30 2021 | WK88 | 88 | 1 | n | FOK
48 | CTUDT00632175 | 1 | Engineering | FOK - Customer complaint: Couch pad arrived with extreme wrinkles in the Velcro. | Accepted |  | NO | Couch | [21 empty] | Mattress | Couch-HW | [8 empty] | January 5, 2021 3:46:07 PM GMT+08:00 | WK54 | 54 | AUG 30 2021 | WK88 | 88 | 1 | n | FOK
49 | CTUDT00632177 | 1 | Engineering | FOK - Customer complaint: 1202CDOD1 Maximum Number Of Hard Arcs Errors | Accepted |  | NO | GCX | [21 empty] | #N/A | Gantry-HW | [8 empty] | January 5, 2021 5:26:28 PM GMT+08:00 | WK54 | 54 | AUG 30 2021 | WK88 | 88 | 1 | n | FOK
50 | CTUDT00632183 | 1 | Engineering | FOK - MPR cannot be created from Axial image. | Accepted |  | YES | CIRS | [21 empty] | CannotRecon | Console-SW | [8 empty] | January 5, 2021 11:42:41 PM GMT+08:00 | WK54 | 54 | AUG 30 2021 | WK88 | 88 | 1 | n | FOK
51 | CTUDT00632189 | 1 | Engineering | FOK- couch horizontal and vertical control motion timeout problem,inner 20ms send both start and end command to couch | Investigated |  | NO | Couch | [21 empty] | Estop open | Couch-HW | [8 empty] | January 6, 2021 10:40:51 AM GMT+08:00 | WK54 | 54 | 4443800% | WK88 | 88 | 1 | y | 1
52 | CTUDT00632190 | 1 | Engineering | FOK - E-stop open during unload patient with tilt function | Investigated |  | NO | GCX | [21 empty] | Estop | Gantry-HW | [8 empty] | January 6, 2021 11:06:20 AM GMT+08:00 | WK54 | 54 | 4443800% | WK88 | 88 | 1 | y | 1
53 | ECR-077012 | 1 | Engineering |  | ECR Track |  | NO | GCX | [21 empty] | Belt | Gantry-HW | [8 empty] | 4411500% | WK41 | 41 | Dec 2 2020 | WK49 | 49 |  | y | 1
54 | TBD_Disk-Bugrep | 1 | Engineering | Although the patient database is empty in the device, the hard disk seems full | TBR | 3 | NO | Console4 | Shubndeep | TBR | N | Need further investigation. After performing disk repair, issue resolved. | Refresh | Y | Y | Y | PR | HW | 0 | Hard disk issue. | [7 empty] | ING-Console-SW-Disk | ING | #N/A | Console-HW | Disk | [7 empty] | TBD | #N/A | #N/A |  | TBD | TBD |  | y | 1
55 | TBD-Cannot Recon | 1 | Engineering | MPR scan cannot recon due to CIRS S2 service running | Installation instruction |  | YES | CIRS |  | TBD | N | [6 empty] | Doc | [11 empty] | Cannot Recon | Console-SW | Cannot Recon | Service forgot to set start up sequence of service mode, trigger IP flouting. | Set service mode and restart system. | Enhance service training to avoid missing this step. | Brook Shi | 4386000% |  |  | TBD | #N/A | #N/A |  | TBD | TBD |  | #N/A | #N/A
56 | TBD-CCT Performance | 1 | Engineering | Pending Bug report | TBSubmit |  | NO | Console4 |  | TBD | N | [6 empty] | SW | [11 empty] | Cannot Recon | Console-SW | CCT Performance | [7 empty] | TBD | #N/A | #N/A |  | TBD | TBD |  | #N/A | #N/A
Formulas:
Level3 = VLOOKUP([Defect (CTUDT#)],'Field IssueSE'!$AG2:BG287,26,FALSE)

## Pivot-mitigate
Func.Owner | (Multiple Items)
Row Labels | Count of Defect (CTUDT#) | Sum of Complaint Contribution
NA | 7 | 0
TBD | 5 | 0
WK105 | 3 | 2
WK20 | 1
WK26 | 1
WK30 | 1 | 12
WK31 | 1 | 2
WK35 | 1 | 2
WK36 | 2 | 2
wk41 | 1 | 1
wk43 | 2 | 5
WK44 | 5 | 14
WK45 | 4 | 6
WK49 | 1
WK52 | 1
WK54 | 2 | 5
WK57 | 1
WK88 | 14 | 30
#N/A | 1 | 1
(blank)
Grand Total | 54 | 82
WK | UniqDefectIn |  | WK | UniqDefectOut | Sum of Complaint Contribution
WK22 | 1 |  | NA | 7 | 0
WK25 | 1 |  | TBD | 5 | 0
WK26 | 3 |  | WK105 | 3 | 2
WK28 | 1 |  | WK20 | 1
WK34 | 1 |  | WK26 | 1
WK35 | 1 |  | WK30 | 1 | 12 | WK36
WK36 | 4 |  | WK31 | 1 | 2
WK37 | 3 |  | WK35 | 1 | 2
WK38 | 7 |  | WK36 | 2 | 2
WK39 | 1 |  | wk41 | 1 | 1
WK40 | 1 |  | wk43 | 2 | 5
WK41 | 1 |  | WK44 | 5 | 14
WK43 | 2 |  | WK45 | 4 | 6
WK44 | 6 |  | WK49 | 1
WK45 | 1 |  | WK52 | 1
WK52 | 6 |  | WK54 | 2 | 5
#N/A | 4 |  | WK57 | 1
WK54 | 10 |  | WK88 | 14 | 30
|  |  | #N/A | 1 | 1
(blank)

## Sheet2
Row Labels | Count of #
ECR-075975 | 2
CTUDT00628691 | 1
CTUDT00630085 | 7
CTUDT00630138 | 12
CTUDT00630160 | 12
CTUDT00630864 | 1
CTUDT00630958 | 1
CTUDT00630969 | 2
CTUDT00630971 | 1
CTUDT00630991 | 1
CTUDT00631039 | 1
CTUDT00631055 | 1
CTUDT00631058 | 2
CTUDT00631134 | 1
CTUDT00631140 | 1
CTUDT00631142 | 1
CTUDT00631143 | 2
CTUDT00631156 | 3
CTUDT00631165 | 2
CTUDT00631283 | 2
CTUDT00631367 | 5
CTUDT00631545 | 3
CTUDT00631602 | 7
CTUDT00631634 | 3
CTUDT00631636 | 1
CTUDT00631638 | 1
CTUDT00631669 | 3
CTUDT00631671 | 3
CTUDT00631717 | 1
CTUDT00632110 | 7
CTUDT00632114 | 2
CTUDT00632116 | 5
CTUDT00632121 | 1
CTUDT00632125 | 2
CTUDT00632127 | 22
CTUDT00632173 | 3
CTUDT00632175 | 1
CTUDT00632177 | 1
CTUDT00632179 | 3
CTUDT00632181 | 4
CTUDT00632183 | 1
CTUDT00632185 | 3
CTUDT00632187 | 2
CTUDT00632189 | 1
CTUDT00632190 | 1
ECR-077012 | 1
ECR-XXXX | 2
Not A Defect | 26
RFC | 2
TBD_Bugreport | 4
TBD_Disk-Bugrep | 1
TBD_Ganty error | 1
TBD_KV_Perf_Bugrep | 1
TBD_OBS | 4
TBD_PACS | 1
TBD_Slipring | 2
TBD_Vertical Transducer | 1
TBD_Workstation | 1
TBD-ACQ Module | 1
TBD-Bugrep | 1
TBD-CCT Performance_Bugrep | 1

## 7-Jan-2021FullList
id | Project | State | Submitter | Submit_Date | Headline | VersionFoundIn | VersionsFixedIn | Workaround | WhoVerifies | ReferenceID | RequirementID | Resolution | Owner | ToBeRepeated | Score | Module | Subsystem | Hazard | State | VersionFixPlannedFor | Products | CFR_ID | VersionsFixedIn | DateLastMerged
CTUDT00632199 | Refresh | Submitted | bingxin_lin | January 6, 2021 4:08:23 PM GMT+08:00 | [BB Refresh]Needle position did not display decimal part when the decimal part is 0 | WIP35 |  |  | Subsystem |  |  |  |  | 0 |  |  |  |  | Submitted |  | Brilliance Big Bore
CTUDT00632197 | Refresh | Submitted | zhong_x | January 6, 2021 3:56:29 PM GMT+08:00 | “gantry cannot initialize. Please change scan parameters or restart the Gantry” always pop up if move couch vertical up/down after initialize surview | WIP35 |  |  | Subsystem |  |  |  |  | 0 |  |  |  |  | Submitted |  | Brilliance Big Bore
CTUDT00632197 | Refresh | Submitted | zhong_x | January 6, 2021 3:56:29 PM GMT+08:00 | “gantry cannot initialize. Please change scan parameters or restart the Gantry” always pop up if move couch vertical up/down after initialize surview | WIP35 |  |  | Subsystem |  |  |  |  | 0 |  |  |  |  | Submitted |  | Ingenuity 64
CTUDT00632195 | Refresh | Submitted | bingxin_lin | January 6, 2021 3:34:33 PM GMT+08:00 | [BB Refresh][IVC]Horizonal , Vertical , Tilt position on IVC control box did not update when press corresponding buttons on it | WIP35 |  |  | Subsystem |  |  |  |  | 0 |  |  |  |  | Submitted |  | Brilliance Big Bore
CTUDT00632190 | Refresh | Investigated | lin_s | January 6, 2021 11:06:20 AM GMT+08:00 | FOK - E-Stop open during unload patient with tilt function. | FOK |  | Design change ECR-077604 will be executed to define tension force. | Subsystem | FOK | NA |  | li_a | 0 |  |  | GCX | NO | Investigated |  | Ingenuity 64 | PR501607
CTUDT00632189 | Refresh | Investigated | cx_zhang | January 6, 2021 10:40:51 AM GMT+08:00 | FOK- couch horizontal and vertical control motion timeout problem,inner 20ms send both start and end command to couch | WIP30.1 |  |  | System | FOK | N/A |  | cx_zhang | 0 |  |  | Couch | NO | Investigated |  | Ingenuity 64 | PR485771
CTUDT00632187 | Refresh | Investigated | lin_s | January 6, 2021 9:40:40 AM GMT+08:00 | FOK - E-Stop open during unload patient. | FOK |  |  | Subsystem | FOK | N/A |  | cx_zhang | 0 |  |  | Couch | NO | Investigated |  | Ingenuity 64 | PR501147 PR503538
CTUDT00632185 | Refresh | Accepted | lin_s | January 6, 2021 12:02:41 AM GMT+08:00 | FOK - CT Box buttons malfunction, reported in three FOK sites. | FOK |  |  | Subsystem | FOK |  |  | wang_f | 0 |  |  | GCX | NO | Accepted |  | Ingenuity 64 | PR387967 PR427944 PR450087
CTUDT00632183 | Refresh | Accepted | lin_s | January 5, 2021 11:42:41 PM GMT+08:00 | FOK - MPR cannot be created from Axial image. | FOK |  |  | Subsystem | FOK |  |  | rakic_k | 0 |  |  | CIRS | YES | Accepted |  | Ingenuity 64 | PR394655
CTUDT00632181 | Refresh | Accepted | lin_s | January 5, 2021 9:55:24 PM GMT+08:00 | FOK - GCP control panel issues are reported from 4 FOK sites. | FOK |  |  | Subsystem | FOK |  |  | wang_f | 0 |  |  | GCX | NO | Accepted |  | Ingenuity 64 | PR435982 PR438049 PR443925 PR487099
CTUDT00632179 | Refresh | Accepted | lin_s | January 5, 2021 9:33:30 PM GMT+08:00 | FOK - Horizontal movement coupling happens in customer site. three sites report similar issue. | FOK |  |  | Subsystem | FOK |  |  | zhou_r | 0 |  |  | Couch | YES | Accepted |  | Ingenuity 64 | PR239416 PR345758 PR395546
CTUDT00632177 | Refresh | Accepted | lin_s | January 5, 2021 5:26:28 PM GMT+08:00 | FOK - Customer complaint: 1202CDOD1 Maximum Number Of Hard Arcs Errors | FOK |  |  | Subsystem | FOK |  |  | zhou_d | 0 |  |  | GCX | NO | Accepted |  | Ingenuity 64 | PR507320
CTUDT00632175 | Refresh | Accepted | lin_s | January 5, 2021 3:46:07 PM GMT+08:00 | FOK - Customer complaint: Couch pad arrived with extreme wrinkles in the Velcro. | FOK |  |  | Subsystem | FOK |  |  | zhou_r | 0 |  |  | Couch | NO | Accepted |  | Ingenuity 64 | PR315368
CTUDT00632173 | Refresh | Accepted | lin_s | January 5, 2021 2:40:48 PM GMT+08:00 | FOK - Random Red Artifacts happens after performing surview or clinical scan | FOK |  |  | Subsystem | FOK |  |  | yuan_s | 0 |  |  | CIRS | NO | Accepted |  | Ingenuity 64 | PR347941 PR402457 PR414927
CTUDT00632172 | Refresh | Assigned | altshuller_mark | January 4, 2021 4:07:30 PM GMT+08:00 | The design of the CTDI free air test for the 128x collimation mode is prone to causing unnecessary failures | WIP30.1 |  |  | Subsystem |  | N/A |  | Lior_Embon | 0 | 2 |  | Support-Physics | NO | Assigned | HF3 | Ovation
CTUDT00632168 | Refresh | Accepted | zhong_x | December 31, 2020 2:33:50 PM GMT+08:00 | CIRS Logs are missing if not restart Host after do CIRS configuration. | WIP25.9 |  |  | Subsystem |  |  |  | rakic_k | 0 |  |  | CIRS | NO | Accepted |  | Ingenuity 64
CTUDT00632157 | Refresh | Accepted | wang_l | December 30, 2020 5:49:41 PM GMT+08:00 | Refresh: "Not all images were created.." when recon with specific scan parameters | WIP25.8 |  |  | Subsystem |  |  |  | mukherjee_s | 0 |  |  | Console4 | YES | Accepted |  | Ingenuity 64
CTUDT00632153 | Refresh | Accepted | lei_yang | December 29, 2020 1:29:07 PM GMT+08:00 | Refresh: The DoseRight message display differs slightly from the host requirements(sSource ID: 412) | WIP35 |  |  | Subsystem |  |  |  | mukherjee_s | 0 |  |  | Console4 | NO | Accepted |  | Brilliance Big Bore
CTUDT00632153 | Refresh | Accepted | lei_yang | December 29, 2020 1:29:07 PM GMT+08:00 | Refresh: The DoseRight message display differs slightly from the host requirements(sSource ID: 412) | WIP35 |  |  | Subsystem |  |  |  | mukherjee_s | 0 |  |  | Console4 | NO | Accepted |  | Ingenuity 64
CTUDT00632150 | Refresh | Accepted | lei_yang | December 29, 2020 11:01:19 AM GMT+08:00 | Refresh: No DoseRight warning message appears after enabled and disabled DoseRight option. | WIP35 |  |  | Subsystem |  |  |  | mukherjee_s | 0 |  |  | Console4 | NO | Accepted |  | Ingenuity 64
CTUDT00632141 | Refresh | Accepted | ding_f | December 28, 2020 4:37:48 PM GMT+08:00 | With "Fast Out", “Tilt Forward” or “Tilt Backward” button pressed, Couch/Gantry motion continues when Load button is pressed, but stops immediately if releasing the Load button. | ING_WIP35 |  |  | Subsystem |  |  |  | yang_h | 0 |  |  | Console4 | NO | Accepted |  | Brilliance 64
CTUDT00632141 | Refresh | Accepted | ding_f | December 28, 2020 4:37:48 PM GMT+08:00 | With "Fast Out", “Tilt Forward” or “Tilt Backward” button pressed, Couch/Gantry motion continues when Load button is pressed, but stops immediately if releasing the Load button. | ING_WIP35 |  |  | Subsystem |  |  |  | yang_h | 0 |  |  | Console4 | NO | Accepted |  | Brilliance Big Bore
CTUDT00632139 | Refresh | Accepted | zhong_x | December 28, 2020 3:57:53 PM GMT+08:00 | Only upper part displayed in Patient tab after click end exam | WIP35 |  |  | Subsystem |  |  |  | mukherjee_s | 0 |  |  | Console4 | NO | Accepted |  | Brilliance Big Bore
CTUDT00632136 | Refresh | Accepted | cai_t | December 28, 2020 9:10:54 AM GMT+08:00 | Console status is not right after powering off Gantry from “Press Manual to start scan” status | BB_WIP33 |  |  | Subsystem |  |  |  | yang_h | 0 |  |  | GCX | YES | Accepted |  | Brilliance Big Bore
CTUDT00632127 | Refresh | Investigated | zhang_t | December 24, 2020 12:25:28 AM GMT+08:00 | FOK: More Sites meet Image transfer stuck issue need to restart host to recover | ING 4.1.10 |  | NA | Subsystem | FOK; SP1_PM_Pending | NA |  | s_sa | 0 |  |  | iCAP | NO | Investigated |  | Ingenuity 64 | 267626 339755 341665 359739 360185 357261 356232 371889 378368 398673 423095 423104 427597 456455 462639 462732 470009 473938 279232 460334 501165 493887 499505
CTUDT00632125 | Refresh | Accepted | zhang_t | December 23, 2020 11:27:59 PM GMT+08:00 | FOK: The new Ingenuity 5000 table pad now has Velcro the entire length of the pad which made it impossible to use the arm rest | ING4.1.10 |  |  | Subsystem | FOK |  |  | wang_ada | 0 |  |  | Couch | NO | Accepted |  | Ingenuity 64 | PR393742
CTUDT00632121 | Refresh | Accepted | zhang_t | December 23, 2020 10:35:54 PM GMT+08:00 | FOK: Exam Summary does not show the dose for the scan unless has the SR report turned on & Can not uncheck PACS as a send to Destination if PACS is listed as a Default Destination. | Ing 4.1.10 |  |  | Subsystem | FOK |  |  | mukherjee_s | 0 |  |  | Console4 | NO | Accepted |  | Ingenuity 64 | PR467713
CTUDT00632116 | Refresh | Accepted | zhang_t | December 23, 2020 2:22:29 PM GMT+08:00 | FOK: MPR Recon with wrong size for Brain/Head Axial EC | Ing HF2 |  |  | Subsystem | FOK |  |  | mukherjee_s | 0 |  |  | Console4 | YES | Accepted |  | Ingenuity 64 | PR444629 PR444106
CTUDT00632114 | Refresh | Accepted | zhang_t | December 23, 2020 1:56:29 PM GMT+08:00 | FOK: Image transfer stuck related to network connection | Ing HF2 |  |  | Subsystem | FOK |  |  | mukherjee_s | 0 |  |  | Console4 | NO | Accepted |  | Ingenuity 64 | PR307590 PR309138
CTUDT00632110 | Refresh | Closed | wang_l | December 23, 2020 11:09:54 AM GMT+08:00 | FOK: Bad Detector calibration failed and ring artifacts occurred | WIP25.2 |  |  | Subsystem | FOK |  | Works As Specified / Not A Defect | yatziv_l | 0 |  |  | DMS | NO | Closed |  | Ingenuity 64 | 479060
CTUDT00632108 | Refresh | Assigned | hilas | December 23, 2020 12:57:20 AM GMT+08:00 | FOK: TDMC reset during air calibration and gave a lifesync failure followed by opening estop. | 4.7.7 |  | Required gantry restart & Re-execute air (as user done, and it passed) | System | SP1_PM_Approved | n/a |  | frider_o | 0 | 3 |  | DMS | NO | Assigned | SP1 | IQon1 | PR487143 PR487155
CTUDT00632105 | Refresh | Accepted | lei_yang | December 22, 2020 9:12:32 AM GMT+08:00 | Refresh: No ECG Gating option displayed after enabled Ecg option in Body Helical EC. | WIP35 |  |  | Subsystem |  |  |  | mukherjee_s | 0 |  |  | Console4 | NO | Accepted |  | Brilliance Big Bore
CTUDT00632105 | Refresh | Accepted | lei_yang | December 22, 2020 9:12:32 AM GMT+08:00 | Refresh: No ECG Gating option displayed after enabled Ecg option in Body Helical EC. | WIP35 |  |  | Subsystem |  |  |  | mukherjee_s | 0 |  |  | Console4 | NO | Accepted |  | Ingenuity 64
CTUDT00632103 | Refresh | Assigned | Lior_Embon | December 22, 2020 5:22:28 AM GMT+08:00 | The design of the CTDI free air test for the 128x collimation mode is prone to causing unnecessary failures | WIP30.1 |  |  | Subsystem |  | n/a |  | Lior_Embon | 0 | 2 |  | Support-Physics | NO | Assigned | SP1 | Ovation
CTUDT00632096 | Refresh | Accepted | wang_l | December 21, 2020 5:08:55 PM GMT+08:00 | Refresh - Transfer ICON not displayed but images transferred to PACS completed | WIP35 |  |  | Subsystem |  |  |  | s_sa | 0 |  |  | iCAP | NO | Accepted |  | Brilliance Big Bore
CTUDT00632094 | Refresh | Accepted | wang_l | December 21, 2020 4:38:42 PM GMT+08:00 | Refresh BB - I2P Configuration file restore failed if I2P config was default value | WIP35 |  |  | Subsystem |  |  |  | mukherjee_s | 0 |  |  | Service Tools | NO | Accepted |  | Brilliance Big Bore
CTUDT00632088 | Refresh | NotAccepted | wang_pete | December 18, 2020 2:28:09 PM GMT+08:00 | FOK Refresh BB: No warning about acquisition out of Surview , and warning - sometimes - when reconstruction inside the surview | WIP25.8 HF2 |  |  | Subsystem |  |  |  |  | 0 |  |  |  |  | NotAccepted |  | Ingenuity BigBore | 293722
CTUDT00632084 | Refresh | Accepted | wang_pete | December 18, 2020 10:46:49 AM GMT+08:00 | FOK Refresh BB: In oncology mode, the images are send to simulation software. issue when changing the position X or Y. | WIP25.8 HF2 |  |  | Subsystem | FOK |  |  | zhong_x | 0 |  |  | Console4 | YES | Accepted |  | Brilliance Big Bore | 293719
CTUDT00632080 | Refresh | Accepted | lakshmikanth_s | December 17, 2020 5:50:42 PM GMT+08:00 | "Move to outer laser" button was functional while scan was in progress. | WIP35 |  |  | Subsystem |  |  |  | mukherjee_s | 0 |  |  | Console4 | YES | Accepted |  | Brilliance Big Bore
CTUDT00632078 | Refresh | Accepted | lakshmikanth_s | December 17, 2020 5:43:47 PM GMT+08:00 | Incremental and Save Work Position buttons are functional during other motion is in progress. | WIP35 |  |  | Subsystem |  |  |  | mukherjee_s | 0 |  |  | Console4 | YES | Accepted |  | Brilliance Big Bore
CTUDT00632075 | Refresh | Accepted | zhong_x | December 17, 2020 3:37:48 PM GMT+08:00 | Scan Region entirely outside of surview image while no notify message | WIP35 |  |  | Subsystem |  |  |  | mukherjee_s | 0 |  |  | Console4 | NO | Accepted |  | Brilliance Big Bore
CTUDT00632073 | Refresh | Accepted | bingxin_lin | December 17, 2020 10:09:26 AM GMT+08:00 | The Move to inner and outer laser LED lights do not behave the same between the workflow is opened and exited | WIP33 |  |  | Subsystem |  |  |  | mukherjee_s | 0 |  |  | Console4 | NO | Accepted |  | Brilliance Big Bore
CTUDT00632071 | Refresh | Scored | hilas | December 16, 2020 11:16:48 PM GMT+08:00 | FOK: Lost Images - CT profusion patient shows in studies, but it does not show any scans done. | 4.7.7 |  |  | Subsystem | SP1_PM_Pending | N/A |  | mukherjee_s | 0 | 3 |  | Console4 | YES | Scored | SP2 | IQon1 | PR459013 PR476143
CTUDT00632069 | Refresh | Closed | hilas | December 16, 2020 9:52:41 PM GMT+08:00 | FOK: recon slowness and transfer to remote device issue | 4.7.7 |  |  | Subsystem |  | N/A | Works As Specified / Not A Defect | janakiraman_j | 0 |  |  | Console4 | NO | Closed |  | IQon1 | PR#464118
CTUDT00632067 | Refresh | Accepted | bingxin_lin | December 16, 2020 4:23:37 PM GMT+08:00 | Couch stopped moving and Move to inner and outer laser button did not work after press the Load Foot Pedal and release it during the move to Outer Laser motion | WIP33 |  |  | Subsystem |  |  |  | dongping_w | 0 |  |  | Support-System | NO | Accepted |  | Brilliance Big Bore
CTUDT00632062 | Refresh | Accepted | madankumar_s | December 15, 2020 3:11:31 PM GMT+08:00 | Non Ascii Exam is not opening after Upgrade and backup restore | WIP35 |  |  | Subsystem |  |  |  | mukherjee_s | 0 |  |  | Service Tools | NO | Accepted |  | Ingenuity 64
CTUDT00632060 | Refresh | Closed | bingxin_lin | December 15, 2020 11:03:44 AM GMT+08:00 | Couch vertical position directly from 56 to 111, did not stop at around 85 vertical position | WIP33 |  |  | Subsystem |  | N/A | Works As Specified / Not A Defect | cx_zhang | 0 |  |  | Couch | NO | Closed |  | Brilliance Big Bore
CTUDT00632058 | Refresh | Scored | portal_l | December 15, 2020 5:00:30 AM GMT+08:00 | Ict-SP WIP 33- Beam Quality Check in Ops user failed yet notes as pass | WIP 33 |  |  | Subsystem | PM_Pending | n/a |  | frider_o | 0 | 4 |  | DMS | NO | Scored | WIP37SegReq | Ovation
CTUDT00632056 | Refresh | Assigned | allaberdiev_v | December 14, 2020 7:39:09 PM GMT+08:00 | IQon:DHF302480:Bad Detectors DC User Regions on the plots are not according to spec. | WIP 34 |  |  | Subsystem | SP1_PM_Approved | n/a |  | frider_o | 0 | 2 |  | DMS | NO | Assigned | WIP37SegReq | IQon1
CTUDT00632048 | Refresh | Closed | barnea_d | December 14, 2020 3:22:22 AM GMT+08:00 | Refresh: Sanity: Air Calibration in CT User stops abruptly | WIP 33 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | frider_o | 0 |  |  | DMS | NO | Closed |  | Ovation
CTUDT00632046 | Refresh | Accepted | zhang_t | December 11, 2020 1:59:35 PM GMT+08:00 | Fok: User experience waiting cursor during plan operations like drag, resizing | iQon 4.7.7 |  |  | Subsystem |  |  |  | mukherjee_s | 0 | 2 |  | Console4 | NO | Accepted |  | IQon1
CTUDT00632042 | Refresh | Assigned | portal_l | December 11, 2020 2:51:19 AM GMT+08:00 | IQon WIP 35- BD does not run in DC\ops user, and pops up Error message of:"Failed to retrieve Nvram information from all 42 DMS modules" even tough the SN&PN exsits. | WIP 35 |  |  | Subsystem | SP1_PM_Approved | n/a |  | frider_o | 0 | 2 |  | DMS | NO | Assigned | WIP37SegReq | Ovation
CTUDT00632042 | Refresh | Assigned | portal_l | December 11, 2020 2:51:19 AM GMT+08:00 | IQon WIP 35- BD does not run in DC\ops user, and pops up Error message of:"Failed to retrieve Nvram information from all 42 DMS modules" even tough the SN&PN exsits. | WIP 35 |  |  | Subsystem | SP1_PM_Approved | n/a |  | frider_o | 0 | 2 |  | DMS | NO | Assigned | WIP37SegReq | IQon1
CTUDT00632039 | Refresh | Accepted | jeffr | December 10, 2020 9:44:26 PM GMT+08:00 | Raw Data Utility description in IFU needs further information | 4.7.7 |  |  | Subsystem |  |  |  | rajalakshmi_j | 0 |  |  | Support-System | NO | Accepted |  | IQon1 | PR281325
CTUDT00632027 | Refresh | Assigned | allaberdiev_v | December 9, 2020 10:12:53 PM GMT+08:00 | ICT:DHF302480:Bad Detectors Service User Regions on the plots are not according to spec. | WIP 35 |  |  | Subsystem | SP1_PM_Approved | n/a |  | frider_o | 0 | 2 |  | DMS | NO | Assigned | WIP37SegReq | Ovation
CTUDT00632020 | Refresh | Assigned | altshuller_mark | December 9, 2020 3:23:27 PM GMT+08:00 | Gantry Touch Panel - Not color indication during appeared during the button pressing. | Hawk |  |  | Subsystem | SP1_PM_Approved | N/A |  | Shai_Goldberg | 0 | 3 | Gantry SW | GCX | NO | Assigned | WIP37SegReq | Hawk
CTUDT00632019 | Refresh | Assigned | altshuller_mark | December 9, 2020 3:17:18 PM GMT+08:00 | Gantry Touch Display buttons respond only from double click | Hawk |  |  | Subsystem | SP1_PM_Approved | N/A |  | Shai_Goldberg | 0 | 3 | Gantry SW | GCX | NO | Assigned | WIP37SegReq | Hawk
CTUDT00632013 | Refresh | Closed | allaberdiev_v | December 8, 2020 8:50:15 PM GMT+08:00 | ICT:DHF302480:BadDetectors "Normalized Signal" high and low limits are not according to spec. | WIP 35 |  |  | Subsystem |  | n/a | Duplicate | frider_o | 0 |  |  | DMS | NO | Closed |  | Ovation
CTUDT00632011 | Refresh | Accepted | altshuller_mark | December 7, 2020 8:27:23 PM GMT+08:00 | Performance degradation due to whitelisting/anti virus | Hawk |  | Perform offline recon | Subsystem | SP1_PM_Approved |  |  | chandran_a | 0 | 2 |  | Support-Security | YES | Accepted | WIP37SegReq | IQon1
CTUDT00632007 | Refresh | Resolved | rakic_k | December 4, 2020 11:50:36 PM GMT+08:00 | CIRS Cannot initialize when scanning Perfusion with more than 7 groups | WIP35 | WIP36 |  | Subsystem |  | N/A | Fixed | richards_b | 0 | 3 | CIRS_SW_ImageProcess | CIRS | NO | Resolved | WIP36SubCom | Ovation |  | WIP36 | December 15, 2020 8:09:08 PM GMT+08:00
CTUDT00632007 | Refresh | Resolved | rakic_k | December 4, 2020 11:50:36 PM GMT+08:00 | CIRS Cannot initialize when scanning Perfusion with more than 7 groups | WIP35 | WIP36 |  | Subsystem |  | N/A | Fixed | richards_b | 0 | 3 | CIRS_SW_ImageProcess | CIRS | NO | Resolved | WIP36SubCom | Brilliance 64 |  | WIP36 | December 15, 2020 8:09:08 PM GMT+08:00
CTUDT00632007 | Refresh | Resolved | rakic_k | December 4, 2020 11:50:36 PM GMT+08:00 | CIRS Cannot initialize when scanning Perfusion with more than 7 groups | WIP35 | WIP36 |  | Subsystem |  | N/A | Fixed | richards_b | 0 | 3 | CIRS_SW_ImageProcess | CIRS | NO | Resolved | WIP36SubCom | Brilliance Big Bore |  | WIP36 | December 15, 2020 8:09:08 PM GMT+08:00
CTUDT00632007 | Refresh | Resolved | rakic_k | December 4, 2020 11:50:36 PM GMT+08:00 | CIRS Cannot initialize when scanning Perfusion with more than 7 groups | WIP35 | WIP36 |  | Subsystem |  | N/A | Fixed | richards_b | 0 | 3 | CIRS_SW_ImageProcess | CIRS | NO | Resolved | WIP36SubCom | Ingenuity 64 |  | WIP36 | December 15, 2020 8:09:08 PM GMT+08:00
CTUDT00632007 | Refresh | Resolved | rakic_k | December 4, 2020 11:50:36 PM GMT+08:00 | CIRS Cannot initialize when scanning Perfusion with more than 7 groups | WIP35 | WIP36 |  | Subsystem |  | N/A | Fixed | richards_b | 0 | 3 | CIRS_SW_ImageProcess | CIRS | NO | Resolved | WIP36SubCom | IQon1 |  | WIP36 | December 15, 2020 8:09:08 PM GMT+08:00
CTUDT00632005 | Refresh | Closed | modi_s | December 4, 2020 5:35:29 PM GMT+08:00 | Two RDSR data was generated for one study. | 4.7.5 |  |  | Subsystem | failing req |  | Forwarded to Different Database | periasamy_g | 0 |  |  | Console4 | NO | Closed |  | IQon1 | N/A
CTUDT00631997 | Refresh | Resolved | cx_zhang | December 3, 2020 4:20:43 PM GMT+08:00 | For Gantry-couch unload fuction communication interface change | WIP24 | WIP36 | release unload button and press again | System | SP1_PM_Approved | N/A | Fixed | yang_h | 0 | 4 | Gantry Software | Couch | NO | Resolved | WIP36SubCom | Ovation | N/A | WIP36 | December 21, 2020 5:33:10 PM GMT+08:00
CTUDT00631997 | Refresh | Resolved | cx_zhang | December 3, 2020 4:20:43 PM GMT+08:00 | For Gantry-couch unload fuction communication interface change | WIP24 | WIP36 | release unload button and press again | System | SP1_PM_Approved | N/A | Fixed | yang_h | 0 | 4 | Gantry Software | Couch | NO | Resolved | WIP36SubCom | Brilliance 64 | N/A | WIP36 | December 21, 2020 5:33:10 PM GMT+08:00
CTUDT00631997 | Refresh | Resolved | cx_zhang | December 3, 2020 4:20:43 PM GMT+08:00 | For Gantry-couch unload fuction communication interface change | WIP24 | WIP36 | release unload button and press again | System | SP1_PM_Approved | N/A | Fixed | yang_h | 0 | 4 | Gantry Software | Couch | NO | Resolved | WIP36SubCom | Brilliance Big Bore | N/A | WIP36 | December 21, 2020 5:33:10 PM GMT+08:00
CTUDT00631997 | Refresh | Resolved | cx_zhang | December 3, 2020 4:20:43 PM GMT+08:00 | For Gantry-couch unload fuction communication interface change | WIP24 | WIP36 | release unload button and press again | System | SP1_PM_Approved | N/A | Fixed | yang_h | 0 | 4 | Gantry Software | Couch | NO | Resolved | WIP36SubCom | Ingenuity 64 | N/A | WIP36 | December 21, 2020 5:33:10 PM GMT+08:00
CTUDT00631997 | Refresh | Resolved | cx_zhang | December 3, 2020 4:20:43 PM GMT+08:00 | For Gantry-couch unload fuction communication interface change | WIP24 | WIP36 | release unload button and press again | System | SP1_PM_Approved | N/A | Fixed | yang_h | 0 | 4 | Gantry Software | Couch | NO | Resolved | WIP36SubCom | IQon1 | N/A | WIP36 | December 21, 2020 5:33:10 PM GMT+08:00
CTUDT00631989 | Refresh | Resolved | altshuller_mark | December 2, 2020 3:32:50 PM GMT+08:00 | Hawk WIP25.2: Workflow was stuck - 100% CPU by process “Philips.SI.GLP” | Hawk | WIP37 | N/A | Subsystem | SP1_PM_Approved | N/A | Fixed | venkatachalam_p | 0 | 2 | Service Tools_Software | Service Tools | YES | Resolved | WIP37SubCom | Ovation |  | WIP37 | December 17, 2020 7:25:25 PM GMT+08:00
CTUDT00631989 | Refresh | Resolved | altshuller_mark | December 2, 2020 3:32:50 PM GMT+08:00 | Hawk WIP25.2: Workflow was stuck - 100% CPU by process “Philips.SI.GLP” | Hawk | WIP37 | N/A | Subsystem | SP1_PM_Approved | N/A | Fixed | venkatachalam_p | 0 | 2 | Service Tools_Software | Service Tools | YES | Resolved | WIP37SubCom | Brilliance 64 |  | WIP37 | December 17, 2020 7:25:25 PM GMT+08:00
CTUDT00631989 | Refresh | Resolved | altshuller_mark | December 2, 2020 3:32:50 PM GMT+08:00 | Hawk WIP25.2: Workflow was stuck - 100% CPU by process “Philips.SI.GLP” | Hawk | WIP37 | N/A | Subsystem | SP1_PM_Approved | N/A | Fixed | venkatachalam_p | 0 | 2 | Service Tools_Software | Service Tools | YES | Resolved | WIP37SubCom | Brilliance Big Bore |  | WIP37 | December 17, 2020 7:25:25 PM GMT+08:00
CTUDT00631989 | Refresh | Resolved | altshuller_mark | December 2, 2020 3:32:50 PM GMT+08:00 | Hawk WIP25.2: Workflow was stuck - 100% CPU by process “Philips.SI.GLP” | Hawk | WIP37 | N/A | Subsystem | SP1_PM_Approved | N/A | Fixed | venkatachalam_p | 0 | 2 | Service Tools_Software | Service Tools | YES | Resolved | WIP37SubCom | Ingenuity 64 |  | WIP37 | December 17, 2020 7:25:25 PM GMT+08:00
CTUDT00631989 | Refresh | Resolved | altshuller_mark | December 2, 2020 3:32:50 PM GMT+08:00 | Hawk WIP25.2: Workflow was stuck - 100% CPU by process “Philips.SI.GLP” | Hawk | WIP37 | N/A | Subsystem | SP1_PM_Approved | N/A | Fixed | venkatachalam_p | 0 | 2 | Service Tools_Software | Service Tools | YES | Resolved | WIP37SubCom | IQon1 |  | WIP37 | December 17, 2020 7:25:25 PM GMT+08:00
CTUDT00631987 | Refresh | Closed | dongping_w | December 2, 2020 3:01:49 PM GMT+08:00 | FOK Refresh BB:When changing the IMR Image Definition, the system does not populate when you select "Apply to All Results". | WIP12.6 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | dongping_w | 0 |  |  |  |  | Closed |  | Brilliance Big Bore
CTUDT00631985 | Refresh | Assigned | cai_t | December 2, 2020 10:29:49 AM GMT+08:00 | Filament Calibration has additional scan parameter: Small 450 and Large 450, 550 than the SSRS(GCX.SSRS.iCT.18577) required. | WIP33 |  |  | Subsystem |  | GCX.SSRS.iCT.18577 |  | wang_k | 0 | 2 |  | GCX | NO | Assigned | SSRS | Brilliance Big Bore
CTUDT00631980 | Refresh | Resolved | 320053769 | December 1, 2020 9:14:35 PM GMT+08:00 | WIP35 – delivered serviceability doesn’t match the installed version. | WIP35 | WIP36 |  | Subsystem | SP1_PM_Approved | CIRS.SSRS.5362 | Fixed | richards_b | 0 | 2 | CIRS_SW_ServiceNode | CIRS | NO | Resolved | WIP36SubCom | Ovation |  | WIP36 | December 15, 2020 8:09:32 PM GMT+08:00
CTUDT00631980 | Refresh | Resolved | 320053769 | December 1, 2020 9:14:35 PM GMT+08:00 | WIP35 – delivered serviceability doesn’t match the installed version. | WIP35 | WIP36 |  | Subsystem | SP1_PM_Approved | CIRS.SSRS.5362 | Fixed | richards_b | 0 | 2 | CIRS_SW_ServiceNode | CIRS | NO | Resolved | WIP36SubCom | Brilliance 64 |  | WIP36 | December 15, 2020 8:09:32 PM GMT+08:00
CTUDT00631980 | Refresh | Resolved | 320053769 | December 1, 2020 9:14:35 PM GMT+08:00 | WIP35 – delivered serviceability doesn’t match the installed version. | WIP35 | WIP36 |  | Subsystem | SP1_PM_Approved | CIRS.SSRS.5362 | Fixed | richards_b | 0 | 2 | CIRS_SW_ServiceNode | CIRS | NO | Resolved | WIP36SubCom | Brilliance Big Bore |  | WIP36 | December 15, 2020 8:09:32 PM GMT+08:00
CTUDT00631980 | Refresh | Resolved | 320053769 | December 1, 2020 9:14:35 PM GMT+08:00 | WIP35 – delivered serviceability doesn’t match the installed version. | WIP35 | WIP36 |  | Subsystem | SP1_PM_Approved | CIRS.SSRS.5362 | Fixed | richards_b | 0 | 2 | CIRS_SW_ServiceNode | CIRS | NO | Resolved | WIP36SubCom | Ingenuity 64 |  | WIP36 | December 15, 2020 8:09:32 PM GMT+08:00
CTUDT00631980 | Refresh | Resolved | 320053769 | December 1, 2020 9:14:35 PM GMT+08:00 | WIP35 – delivered serviceability doesn’t match the installed version. | WIP35 | WIP36 |  | Subsystem | SP1_PM_Approved | CIRS.SSRS.5362 | Fixed | richards_b | 0 | 2 | CIRS_SW_ServiceNode | CIRS | NO | Resolved | WIP36SubCom | IQon1 |  | WIP36 | December 15, 2020 8:09:32 PM GMT+08:00
CTUDT00631973 | Refresh | Closed | zhang_t | December 1, 2020 1:44:59 PM GMT+08:00 | FOK: System time synch is not correct | FOK |  | No Workaround. | Subsystem | SP1_PM_Pending | NA | Works As Specified / Not A Defect | s_srivastava | 0 | 4 |  | Console4 | NO | Closed |  | IQon1 | PR342057
CTUDT00631966 | Refresh | Accepted | zhong_x | November 30, 2020 5:44:34 PM GMT+08:00 | For (QRS) not detected notification, criteria defined in System Requirement is not clear. | WIP33 |  |  | Subsystem |  |  |  | lin_s | 0 |  |  | Support-System | NO | Accepted |  | Brilliance Big Bore
CTUDT00631966 | Refresh | Accepted | zhong_x | November 30, 2020 5:44:34 PM GMT+08:00 | For (QRS) not detected notification, criteria defined in System Requirement is not clear. | WIP33 |  |  | Subsystem |  |  |  | lin_s | 0 |  |  | Support-System | NO | Accepted |  | Ingenuity 64
CTUDT00631964 | Refresh | Closed | portal_l | November 30, 2020 3:49:04 PM GMT+08:00 | ICt WIP 33- A gap between Spec (DHF198870) to Help Page regarding Steps Calibration For required Phantom. | WIP 33 |  | Use WI | Subsystem | PM_Close_NoFix | N/A | No Plan to Fix | yatziv_l | 0 | 2 |  | DMS | NO | Closed | SubNC | Ovation
CTUDT00631955 | Refresh | Closed | barnea_d | November 29, 2020 11:10:11 PM GMT+08:00 | Refresh: DMS-SRS: The word “direction” isn’t translated in all of the Swedish help page parts | WIP 33 |  |  | Subsystem |  | DMS-SRS-1206 | No Plan to Fix | yatziv_l | 0 | 3 |  | DMS | NO | Closed |  | Ovation
CTUDT00631955 | Refresh | Closed | barnea_d | November 29, 2020 11:10:11 PM GMT+08:00 | Refresh: DMS-SRS: The word “direction” isn’t translated in all of the Swedish help page parts | WIP 33 |  |  | Subsystem |  | DMS-SRS-1206 | No Plan to Fix | yatziv_l | 0 | 3 |  | DMS | NO | Closed |  | Brilliance 64
CTUDT00631955 | Refresh | Closed | barnea_d | November 29, 2020 11:10:11 PM GMT+08:00 | Refresh: DMS-SRS: The word “direction” isn’t translated in all of the Swedish help page parts | WIP 33 |  |  | Subsystem |  | DMS-SRS-1206 | No Plan to Fix | yatziv_l | 0 | 3 |  | DMS | NO | Closed |  | Brilliance Big Bore
CTUDT00631955 | Refresh | Closed | barnea_d | November 29, 2020 11:10:11 PM GMT+08:00 | Refresh: DMS-SRS: The word “direction” isn’t translated in all of the Swedish help page parts | WIP 33 |  |  | Subsystem |  | DMS-SRS-1206 | No Plan to Fix | yatziv_l | 0 | 3 |  | DMS | NO | Closed |  | Ingenuity 64
CTUDT00631955 | Refresh | Closed | barnea_d | November 29, 2020 11:10:11 PM GMT+08:00 | Refresh: DMS-SRS: The word “direction” isn’t translated in all of the Swedish help page parts | WIP 33 |  |  | Subsystem |  | DMS-SRS-1206 | No Plan to Fix | yatziv_l | 0 | 3 |  | DMS | NO | Closed |  | IQon1
CTUDT00631950 | Refresh | Accepted | zhong_x | November 29, 2020 3:44:49 PM GMT+08:00 | From X-ray on to first preview image display exceed 4 seconds sometimes. | WIP33 |  |  | Subsystem |  |  |  | zhong_x | 1 |  |  | Support-System | NO | Accepted |  | Brilliance Big Bore
CTUDT00631942 | Refresh | NotAccepted | lakshmikanth_s | November 26, 2020 6:20:16 PM GMT+08:00 | Not able load Cardiac and Onco Gating EC to Plan and in ECM | WIP 34 |  |  | Subsystem |  |  |  |  | 0 |  |  |  |  | NotAccepted |  | IQon1
CTUDT00631940 | Refresh | Closed | lakshmikanth_s | November 26, 2020 6:03:19 PM GMT+08:00 | Results are loaded to Spectral CT Viewer through auto launch, number of images displayed in status is incorrect | WIP34 |  |  | Subsystem |  | ?? | Works As Specified / Not A Defect | aviw | 0 |  |  | Support_Spectral | NO | Closed |  | IQon1
CTUDT00631938 | Refresh | Closed | lakshmikanth_s | November 26, 2020 5:41:12 PM GMT+08:00 | Spectral CT viewer was crashed after scan and recon when auto launch was set for 13 spectral results. | WIP34 |  | Rerun sCTV and load images manually | Subsystem |  | ?? | No Plan to Fix | aviw | 0 | 4 |  | Support_Spectral | NO | Closed |  | IQon1
CTUDT00631934 | Refresh | Closed | hilas | November 26, 2020 12:08:11 AM GMT+08:00 | FOK: CTVs Maps wrong translation in French | 4.7.7 |  | load pdf spactral page and re-check | System | FOK; SP2_Postpone | NA | No Plan to Fix | aviw | 0 | 4 |  | Support_Spectral | NO | Closed |  | IQon1 | PR286041
CTUDT00631922 | Refresh | Opened | bingxin_lin | November 25, 2020 4:16:01 PM GMT+08:00 | The phantom is difficult to be detected without Phantom's Adult Body section at Constancy test calibration when using Therapy Top | WIP33 |  |  | Subsystem | CRN | n/a |  | rajalakshmi_j | 0 | 3 |  | Support-System | NO | Opened | SP1-Mar-21 | Ovation
CTUDT00631922 | Refresh | Opened | bingxin_lin | November 25, 2020 4:16:01 PM GMT+08:00 | The phantom is difficult to be detected without Phantom's Adult Body section at Constancy test calibration when using Therapy Top | WIP33 |  |  | Subsystem | CRN | n/a |  | rajalakshmi_j | 0 | 3 |  | Support-System | NO | Opened | SP1-Mar-21 | Brilliance Big Bore
CTUDT00631922 | Refresh | Opened | bingxin_lin | November 25, 2020 4:16:01 PM GMT+08:00 | The phantom is difficult to be detected without Phantom's Adult Body section at Constancy test calibration when using Therapy Top | WIP33 |  |  | Subsystem | CRN | n/a |  | rajalakshmi_j | 0 | 3 |  | Support-System | NO | Opened | SP1-Mar-21 | Ingenuity 64
CTUDT00631919 | Refresh | Accepted | lei_yang | November 25, 2020 11:17:15 AM GMT+08:00 | Refresh: Step&shoot first clinical scan no wave display. | WIP33 |  |  | Subsystem | SP2_Candidate |  |  | mukherjee_s | 0 |  |  | Console4 | YES | Accepted |  | Ingenuity 64
CTUDT00631917 | Refresh | Accepted | bingxin_lin | November 25, 2020 11:04:37 AM GMT+08:00 | "Net Command has stopped working" message popped up after canceling the collection of budreport | WIP33 |  |  | Subsystem | SP2_Candidate |  |  | mukherjee_s | 0 |  |  | Console4 | NO | Accepted |  | Brilliance Big Bore
CTUDT00631915 | Refresh | Investigated | bingxin_lin | November 25, 2020 10:41:08 AM GMT+08:00 | "System error" message displayed after zoom in of the surview image | WIP33 |  |  | Subsystem | TBR, SP1_TBR | NA |  | aswani_v | 1 |  |  | Console4 | NO | Investigated |  | Brilliance Big Bore
CTUDT00631908 | Refresh | NotAccepted | lakshmikanth_s | November 24, 2020 6:41:28 PM GMT+08:00 | System Error occurred after clicking on End Exam. | WIP 33 |  |  | Subsystem |  |  |  |  | 0 |  |  |  |  | NotAccepted |  | Brilliance Big Bore
CTUDT00631905 | Refresh | Accepted | Revital_Cohen | November 24, 2020 3:26:27 AM GMT+08:00 | Refresh Service pack O-level User in Service Tools is able to adjust the time in Internet tab and have the ability to perform DMS diagnostics although he should not be able to do so | WIP33 |  |  | Subsystem | SP2_Candidate |  |  | periasamy_g | 0 |  |  | Service Tools | NO | Accepted | WIP37SegReq | Ovation
CTUDT00631905 | Refresh | Accepted | Revital_Cohen | November 24, 2020 3:26:27 AM GMT+08:00 | Refresh Service pack O-level User in Service Tools is able to adjust the time in Internet tab and have the ability to perform DMS diagnostics although he should not be able to do so | WIP33 |  |  | Subsystem | SP2_Candidate |  |  | periasamy_g | 0 |  |  | Service Tools | NO | Accepted | WIP37SegReq | IQon1
CTUDT00631898 | Refresh | Closed | zhang_l | November 23, 2020 3:53:02 PM GMT+08:00 | Long Tube Conditioning Sequences in QA are inconsistent with Gantry SSRS. | WIP33 | DHF317353 rev09 DHF317363 rev10 |  | System |  | GCX.SSRS.iCT.16213; GCX.SSRS.iCT.18533 | Fixed | wang_k | 0 | 2 | Design Documentation | GCX | NO | Closed | SSRS new version | Brilliance 64 |  | DHF317353 rev09 DHF317363 rev10 | December 15, 2020 2:58:13 PM GMT+08:00
CTUDT00631898 | Refresh | Closed | zhang_l | November 23, 2020 3:53:02 PM GMT+08:00 | Long Tube Conditioning Sequences in QA are inconsistent with Gantry SSRS. | WIP33 | DHF317353 rev09 DHF317363 rev10 |  | System |  | GCX.SSRS.iCT.16213; GCX.SSRS.iCT.18533 | Fixed | wang_k | 0 | 2 | Design Documentation | GCX | NO | Closed | SSRS new version | Brilliance Big Bore |  | DHF317353 rev09 DHF317363 rev10 | December 15, 2020 2:58:13 PM GMT+08:00
CTUDT00631898 | Refresh | Closed | zhang_l | November 23, 2020 3:53:02 PM GMT+08:00 | Long Tube Conditioning Sequences in QA are inconsistent with Gantry SSRS. | WIP33 | DHF317353 rev09 DHF317363 rev10 |  | System |  | GCX.SSRS.iCT.16213; GCX.SSRS.iCT.18533 | Fixed | wang_k | 0 | 2 | Design Documentation | GCX | NO | Closed | SSRS new version | Ingenuity 64 |  | DHF317353 rev09 DHF317363 rev10 | December 15, 2020 2:58:13 PM GMT+08:00
CTUDT00631895 | Refresh | Resolved | periasamy_g | November 23, 2020 12:28:15 PM GMT+08:00 | Requirements update with Boundary conditions | 4.1.10 | WIP36 |  | Subsystem | SP1_PM_Approved ; Doc_Update | Console-SRS.10089, Console-SRS.10090, Console-SRS,10099 | Fixed | periasamy_g | 0 | 2 | Requirements Documentation | Console4 | NO | Resolved | SSRS | Ovation |  | WIP36 | December 11, 2020 12:36:02 AM GMT+08:00
CTUDT00631895 | Refresh | Resolved | periasamy_g | November 23, 2020 12:28:15 PM GMT+08:00 | Requirements update with Boundary conditions | 4.1.10 | WIP36 |  | Subsystem | SP1_PM_Approved ; Doc_Update | Console-SRS.10089, Console-SRS.10090, Console-SRS,10099 | Fixed | periasamy_g | 0 | 2 | Requirements Documentation | Console4 | NO | Resolved | SSRS | Brilliance 64 |  | WIP36 | December 11, 2020 12:36:02 AM GMT+08:00
CTUDT00631895 | Refresh | Resolved | periasamy_g | November 23, 2020 12:28:15 PM GMT+08:00 | Requirements update with Boundary conditions | 4.1.10 | WIP36 |  | Subsystem | SP1_PM_Approved ; Doc_Update | Console-SRS.10089, Console-SRS.10090, Console-SRS,10099 | Fixed | periasamy_g | 0 | 2 | Requirements Documentation | Console4 | NO | Resolved | SSRS | Brilliance Big Bore |  | WIP36 | December 11, 2020 12:36:02 AM GMT+08:00
CTUDT00631895 | Refresh | Resolved | periasamy_g | November 23, 2020 12:28:15 PM GMT+08:00 | Requirements update with Boundary conditions | 4.1.10 | WIP36 |  | Subsystem | SP1_PM_Approved ; Doc_Update | Console-SRS.10089, Console-SRS.10090, Console-SRS,10099 | Fixed | periasamy_g | 0 | 2 | Requirements Documentation | Console4 | NO | Resolved | SSRS | Ingenuity 64 |  | WIP36 | December 11, 2020 12:36:02 AM GMT+08:00
CTUDT00631895 | Refresh | Resolved | periasamy_g | November 23, 2020 12:28:15 PM GMT+08:00 | Requirements update with Boundary conditions | 4.1.10 | WIP36 |  | Subsystem | SP1_PM_Approved ; Doc_Update | Console-SRS.10089, Console-SRS.10090, Console-SRS,10099 | Fixed | periasamy_g | 0 | 2 | Requirements Documentation | Console4 | NO | Resolved | SSRS | IQon1 |  | WIP36 | December 11, 2020 12:36:02 AM GMT+08:00
CTUDT00631891 | Refresh | Accepted | wang_l | November 20, 2020 5:01:23 PM GMT+08:00 | Refresh - Bolus Tracker: The Threshold textbox got stucked when trying to edit in Tracker scene | WIP33 |  |  | Subsystem | SP2_Candidate |  |  | periasamy_g | 0 |  |  | Console4 | NO | Accepted |  | Ingenuity 64
CTUDT00631885 | Refresh | Resolved | lakshmikanth_s | November 19, 2020 8:31:10 PM GMT+08:00 | Few tools are missing in Service tools for iQon and BigBore | WIP33 | WIP36 |  | Subsystem | SP1_PM_Approved | NA | Fixed | chakraborty_d | 0 | 2 | Service Tools_Software | Service Tools | NO | Resolved | WIP36SubCom | Ovation |  | WIP36 | December 8, 2020 3:37:46 PM GMT+08:00
CTUDT00631885 | Refresh | Resolved | lakshmikanth_s | November 19, 2020 8:31:10 PM GMT+08:00 | Few tools are missing in Service tools for iQon and BigBore | WIP33 | WIP36 |  | Subsystem | SP1_PM_Approved | NA | Fixed | chakraborty_d | 0 | 2 | Service Tools_Software | Service Tools | NO | Resolved | WIP36SubCom | Brilliance 64 |  | WIP36 | December 8, 2020 3:37:46 PM GMT+08:00
CTUDT00631885 | Refresh | Resolved | lakshmikanth_s | November 19, 2020 8:31:10 PM GMT+08:00 | Few tools are missing in Service tools for iQon and BigBore | WIP33 | WIP36 |  | Subsystem | SP1_PM_Approved | NA | Fixed | chakraborty_d | 0 | 2 | Service Tools_Software | Service Tools | NO | Resolved | WIP36SubCom | Brilliance Big Bore |  | WIP36 | December 8, 2020 3:37:46 PM GMT+08:00
CTUDT00631885 | Refresh | Resolved | lakshmikanth_s | November 19, 2020 8:31:10 PM GMT+08:00 | Few tools are missing in Service tools for iQon and BigBore | WIP33 | WIP36 |  | Subsystem | SP1_PM_Approved | NA | Fixed | chakraborty_d | 0 | 2 | Service Tools_Software | Service Tools | NO | Resolved | WIP36SubCom | Ingenuity 64 |  | WIP36 | December 8, 2020 3:37:46 PM GMT+08:00
CTUDT00631885 | Refresh | Resolved | lakshmikanth_s | November 19, 2020 8:31:10 PM GMT+08:00 | Few tools are missing in Service tools for iQon and BigBore | WIP33 | WIP36 |  | Subsystem | SP1_PM_Approved | NA | Fixed | chakraborty_d | 0 | 2 | Service Tools_Software | Service Tools | NO | Resolved | WIP36SubCom | IQon1 |  | WIP36 | December 8, 2020 3:37:46 PM GMT+08:00
CTUDT00631875 | Refresh | Resolved | rakic_k | November 17, 2020 10:37:35 PM GMT+08:00 | CIRS’s collected records are to large – Causing RADAR monitoring to fail | WIP33 | WIP35 |  | Subsystem | SP1_PM_Approved | NA | Fixed | richards_b | 0 | 3 | CIRS_SW_Installation | CIRS | NO | Resolved | WIP35SubCom | Ovation |  | WIP35 | December 5, 2020 6:00:29 AM GMT+08:00
CTUDT00631875 | Refresh | Resolved | rakic_k | November 17, 2020 10:37:35 PM GMT+08:00 | CIRS’s collected records are to large – Causing RADAR monitoring to fail | WIP33 | WIP35 |  | Subsystem | SP1_PM_Approved | NA | Fixed | richards_b | 0 | 3 | CIRS_SW_Installation | CIRS | NO | Resolved | WIP35SubCom | Brilliance 64 |  | WIP35 | December 5, 2020 6:00:29 AM GMT+08:00
CTUDT00631875 | Refresh | Resolved | rakic_k | November 17, 2020 10:37:35 PM GMT+08:00 | CIRS’s collected records are to large – Causing RADAR monitoring to fail | WIP33 | WIP35 |  | Subsystem | SP1_PM_Approved | NA | Fixed | richards_b | 0 | 3 | CIRS_SW_Installation | CIRS | NO | Resolved | WIP35SubCom | Brilliance Big Bore |  | WIP35 | December 5, 2020 6:00:29 AM GMT+08:00
CTUDT00631875 | Refresh | Resolved | rakic_k | November 17, 2020 10:37:35 PM GMT+08:00 | CIRS’s collected records are to large – Causing RADAR monitoring to fail | WIP33 | WIP35 |  | Subsystem | SP1_PM_Approved | NA | Fixed | richards_b | 0 | 3 | CIRS_SW_Installation | CIRS | NO | Resolved | WIP35SubCom | Ingenuity 64 |  | WIP35 | December 5, 2020 6:00:29 AM GMT+08:00
CTUDT00631875 | Refresh | Resolved | rakic_k | November 17, 2020 10:37:35 PM GMT+08:00 | CIRS’s collected records are to large – Causing RADAR monitoring to fail | WIP33 | WIP35 |  | Subsystem | SP1_PM_Approved | NA | Fixed | richards_b | 0 | 3 | CIRS_SW_Installation | CIRS | NO | Resolved | WIP35SubCom | IQon1 |  | WIP35 | December 5, 2020 6:00:29 AM GMT+08:00
CTUDT00631868 | Refresh | Accepted | hilas | November 17, 2020 7:34:09 PM GMT+08:00 | FoK: Host stuck and preview image not reconstructed | 4.7.7 |  |  | Subsystem | SP1_PM_Pending | N/A |  | mukherjee_s | 0 |  |  | Console4 | YES | Accepted |  | IQon1 | PR405360 PR497763
CTUDT00631866 | Refresh | Accepted | hilas | November 17, 2020 5:27:31 PM GMT+08:00 | FoK: The GPC experienced BSOD (blue screen of death) | 4.7.7 |  |  | Subsystem | FoK |  |  | dongping_w | 0 |  |  | Support-System | NO | Accepted |  | IQon1 | PR224015 PR415059
CTUDT00631862 | Refresh | Accepted | lei_yang | November 17, 2020 9:19:34 AM GMT+08:00 | When the EC in ECM is right-clicked and then click the other group, the group will be offset, and the EC cannot be selected. | WIP33 |  |  | Subsystem | SP2_Postpone |  |  | mukherjee_s | 0 |  |  | Console4 | NO | Accepted |  | Brilliance 64
CTUDT00631862 | Refresh | Accepted | lei_yang | November 17, 2020 9:19:34 AM GMT+08:00 | When the EC in ECM is right-clicked and then click the other group, the group will be offset, and the EC cannot be selected. | WIP33 |  |  | Subsystem | SP2_Postpone |  |  | mukherjee_s | 0 |  |  | Console4 | NO | Accepted |  | Brilliance Big Bore
CTUDT00631862 | Refresh | Accepted | lei_yang | November 17, 2020 9:19:34 AM GMT+08:00 | When the EC in ECM is right-clicked and then click the other group, the group will be offset, and the EC cannot be selected. | WIP33 |  |  | Subsystem | SP2_Postpone |  |  | mukherjee_s | 0 |  |  | Console4 | NO | Accepted |  | Ingenuity 64
CTUDT00631860 | Refresh | Opened | lei_yang | November 17, 2020 9:13:31 AM GMT+08:00 | Workflow suddenly disappears in the process of collecting bug report. | WIP33 |  |  | Subsystem | CRN, SP2_Score2Postpone | N/A |  | rajalakshmi_j | 0 | 3 |  | Console4 | NO | Opened | SP1-Mar-21-CRN | Ingenuity 64
CTUDT00631856 | Refresh | Accepted | 320085947 | November 16, 2020 8:34:16 PM GMT+08:00 | a display issue in acquisition scene | WIP34 |  |  | Subsystem | SP2_Candidate |  |  | mukherjee_s | 0 |  |  | Console4 | NO | Accepted |  | iCT SDCT
CTUDT00631853 | Refresh | Scored | prabhan_p | November 16, 2020 4:50:22 PM GMT+08:00 | Film tab screen turns yellow. | WIP33 |  | NA | Subsystem | FOK; PM_Pending | NA |  | srushti_k | 0 | 2 |  | iCAP | NO | Scored |  | Ovation
CTUDT00631853 | Refresh | Scored | prabhan_p | November 16, 2020 4:50:22 PM GMT+08:00 | Film tab screen turns yellow. | WIP33 |  | NA | Subsystem | FOK; PM_Pending | NA |  | srushti_k | 0 | 2 |  | iCAP | NO | Scored |  | Brilliance 64
CTUDT00631853 | Refresh | Scored | prabhan_p | November 16, 2020 4:50:22 PM GMT+08:00 | Film tab screen turns yellow. | WIP33 |  | NA | Subsystem | FOK; PM_Pending | NA |  | srushti_k | 0 | 2 |  | iCAP | NO | Scored |  | Brilliance Big Bore
CTUDT00631853 | Refresh | Scored | prabhan_p | November 16, 2020 4:50:22 PM GMT+08:00 | Film tab screen turns yellow. | WIP33 |  | NA | Subsystem | FOK; PM_Pending | NA |  | srushti_k | 0 | 2 |  | iCAP | NO | Scored |  | Ingenuity 64
CTUDT00631853 | Refresh | Scored | prabhan_p | November 16, 2020 4:50:22 PM GMT+08:00 | Film tab screen turns yellow. | WIP33 |  | NA | Subsystem | FOK; PM_Pending | NA |  | srushti_k | 0 | 2 |  | iCAP | NO | Scored |  | IQon1
CTUDT00631852 | Refresh | Resolved | cai_t | November 16, 2020 4:49:16 PM GMT+08:00 | [Old GCP]Couch stopped moving when Press and Hold Horizontal In button then Press Vertical Up button. | WIP33 | WIP36 |  | Subsystem | SP1_PM_Approved | NA/ | Change Proposal | tan_j | 0 | 2 | Gantry Software | GCX | NO | Resolved | WIP36SegReq | Brilliance 64 |  | WIP36 | December 8, 2020 3:45:02 PM GMT+08:00
CTUDT00631852 | Refresh | Resolved | cai_t | November 16, 2020 4:49:16 PM GMT+08:00 | [Old GCP]Couch stopped moving when Press and Hold Horizontal In button then Press Vertical Up button. | WIP33 | WIP36 |  | Subsystem | SP1_PM_Approved | NA/ | Change Proposal | tan_j | 0 | 2 | Gantry Software | GCX | NO | Resolved | WIP36SegReq | Brilliance Big Bore |  | WIP36 | December 8, 2020 3:45:02 PM GMT+08:00
CTUDT00631852 | Refresh | Resolved | cai_t | November 16, 2020 4:49:16 PM GMT+08:00 | [Old GCP]Couch stopped moving when Press and Hold Horizontal In button then Press Vertical Up button. | WIP33 | WIP36 |  | Subsystem | SP1_PM_Approved | NA/ | Change Proposal | tan_j | 0 | 2 | Gantry Software | GCX | NO | Resolved | WIP36SegReq | Ingenuity 64 |  | WIP36 | December 8, 2020 3:45:02 PM GMT+08:00
CTUDT00631850 | Refresh | Closed | allaberdiev_v | November 15, 2020 10:28:08 PM GMT+08:00 | Ingenuity: DHF170395 : Spatial Resolution 5 Head UHR Impulse Response Constancy limits don't appear in the spec DHF170395 table 8.7.6 Constancy limits for Harmonized phantom | WIP 33 |  |  | Subsystem |  | N/A | N/A - Test Design | gehtin_e | 0 | 2 | Test Design | DMS | NO | Closed | SP1_Doc | Ovation |  |  | November 25, 2020 9:28:43 PM GMT+08:00
CTUDT00631850 | Refresh | Closed | allaberdiev_v | November 15, 2020 10:28:08 PM GMT+08:00 | Ingenuity: DHF170395 : Spatial Resolution 5 Head UHR Impulse Response Constancy limits don't appear in the spec DHF170395 table 8.7.6 Constancy limits for Harmonized phantom | WIP 33 |  |  | Subsystem |  | N/A | N/A - Test Design | gehtin_e | 0 | 2 | Test Design | DMS | NO | Closed | SP1_Doc | Brilliance Big Bore |  |  | November 25, 2020 9:28:43 PM GMT+08:00
CTUDT00631850 | Refresh | Closed | allaberdiev_v | November 15, 2020 10:28:08 PM GMT+08:00 | Ingenuity: DHF170395 : Spatial Resolution 5 Head UHR Impulse Response Constancy limits don't appear in the spec DHF170395 table 8.7.6 Constancy limits for Harmonized phantom | WIP 33 |  |  | Subsystem |  | N/A | N/A - Test Design | gehtin_e | 0 | 2 | Test Design | DMS | NO | Closed | SP1_Doc | Ingenuity 64 |  |  | November 25, 2020 9:28:43 PM GMT+08:00
CTUDT00631850 | Refresh | Closed | allaberdiev_v | November 15, 2020 10:28:08 PM GMT+08:00 | Ingenuity: DHF170395 : Spatial Resolution 5 Head UHR Impulse Response Constancy limits don't appear in the spec DHF170395 table 8.7.6 Constancy limits for Harmonized phantom | WIP 33 |  |  | Subsystem |  | N/A | N/A - Test Design | gehtin_e | 0 | 2 | Test Design | DMS | NO | Closed | SP1_Doc | IQon1 |  |  | November 25, 2020 9:28:43 PM GMT+08:00
CTUDT00631847 | Refresh | Accepted | elavarasan_s | November 13, 2020 3:24:37 PM GMT+08:00 | When user clicked on Start final Recon it is taking more than 30 secs to load the images in view scene. In View 2 Status line message displayed as Reconstructing even after recon is completed | WIP 33 |  |  | System | SP2_Candidate |  |  | mukherjee_s | 0 |  |  | Console4 | YES | Accepted |  | Brilliance Big Bore
CTUDT00631847 | Refresh | Accepted | elavarasan_s | November 13, 2020 3:24:37 PM GMT+08:00 | When user clicked on Start final Recon it is taking more than 30 secs to load the images in view scene. In View 2 Status line message displayed as Reconstructing even after recon is completed | WIP 33 |  |  | System | SP2_Candidate |  |  | mukherjee_s | 0 |  |  | Console4 | YES | Accepted |  | IQon1
CTUDT00631845 | Refresh | Scored | hilas | November 13, 2020 1:16:04 AM GMT+08:00 | FOK: Film application slowness | 4.7.7 |  | NA | Subsystem | FOK; PM_Pending | NA |  | srushti_k | 0 | 2 |  | iCAP | NO | Scored |  | Ovation | PR399302 PR399356 PR402396
CTUDT00631845 | Refresh | Scored | hilas | November 13, 2020 1:16:04 AM GMT+08:00 | FOK: Film application slowness | 4.7.7 |  | NA | Subsystem | FOK; PM_Pending | NA |  | srushti_k | 0 | 2 |  | iCAP | NO | Scored |  | Brilliance 64 | PR399302 PR399356 PR402396
CTUDT00631845 | Refresh | Scored | hilas | November 13, 2020 1:16:04 AM GMT+08:00 | FOK: Film application slowness | 4.7.7 |  | NA | Subsystem | FOK; PM_Pending | NA |  | srushti_k | 0 | 2 |  | iCAP | NO | Scored |  | Brilliance Big Bore | PR399302 PR399356 PR402396
CTUDT00631845 | Refresh | Scored | hilas | November 13, 2020 1:16:04 AM GMT+08:00 | FOK: Film application slowness | 4.7.7 |  | NA | Subsystem | FOK; PM_Pending | NA |  | srushti_k | 0 | 2 |  | iCAP | NO | Scored |  | Ingenuity 64 | PR399302 PR399356 PR402396
CTUDT00631845 | Refresh | Scored | hilas | November 13, 2020 1:16:04 AM GMT+08:00 | FOK: Film application slowness | 4.7.7 |  | NA | Subsystem | FOK; PM_Pending | NA |  | srushti_k | 0 | 2 |  | iCAP | NO | Scored |  | IQon1 | PR399302 PR399356 PR402396
CTUDT00631841 | Refresh | Closed | hilas | November 12, 2020 10:40:34 PM GMT+08:00 | FOK: End Exam slowness | 4.7.7 |  |  | Subsystem | SP2_Candidate, FOK,HISRIS | NA | Does Not Apply | s_sa | 0 |  |  | iCAP | NO | Closed |  | IQon1
CTUDT00631837 | Refresh | Resolved | hilas | November 12, 2020 9:52:02 PM GMT+08:00 | FOK: Recon error with not all image created msg | 4.7.7 | WIP36 | User can ignore the message and continue with their workflow. | Subsystem | FOK; HF3; SP1_PM_Approved | N/A | Fixed | richards_b | 0 | 4 | CIRS_SW_Data_Management | CIRS | NO | Resolved | SP1 | Ovation | 405410 | WIP36 | December 23, 2020 9:18:47 PM GMT+08:00
CTUDT00631837 | Refresh | Resolved | hilas | November 12, 2020 9:52:02 PM GMT+08:00 | FOK: Recon error with not all image created msg | 4.7.7 | WIP36 | User can ignore the message and continue with their workflow. | Subsystem | FOK; HF3; SP1_PM_Approved | N/A | Fixed | richards_b | 0 | 4 | CIRS_SW_Data_Management | CIRS | NO | Resolved | SP1 | Brilliance 64 | 405410 | WIP36 | December 23, 2020 9:18:47 PM GMT+08:00
CTUDT00631837 | Refresh | Resolved | hilas | November 12, 2020 9:52:02 PM GMT+08:00 | FOK: Recon error with not all image created msg | 4.7.7 | WIP36 | User can ignore the message and continue with their workflow. | Subsystem | FOK; HF3; SP1_PM_Approved | N/A | Fixed | richards_b | 0 | 4 | CIRS_SW_Data_Management | CIRS | NO | Resolved | SP1 | Brilliance Big Bore | 405410 | WIP36 | December 23, 2020 9:18:47 PM GMT+08:00
CTUDT00631837 | Refresh | Resolved | hilas | November 12, 2020 9:52:02 PM GMT+08:00 | FOK: Recon error with not all image created msg | 4.7.7 | WIP36 | User can ignore the message and continue with their workflow. | Subsystem | FOK; HF3; SP1_PM_Approved | N/A | Fixed | richards_b | 0 | 4 | CIRS_SW_Data_Management | CIRS | NO | Resolved | SP1 | Ingenuity 64 | 405410 | WIP36 | December 23, 2020 9:18:47 PM GMT+08:00
CTUDT00631837 | Refresh | Resolved | hilas | November 12, 2020 9:52:02 PM GMT+08:00 | FOK: Recon error with not all image created msg | 4.7.7 | WIP36 | User can ignore the message and continue with their workflow. | Subsystem | FOK; HF3; SP1_PM_Approved | N/A | Fixed | richards_b | 0 | 4 | CIRS_SW_Data_Management | CIRS | NO | Resolved | SP1 | IQon1 | 405410 | WIP36 | December 23, 2020 9:18:47 PM GMT+08:00
CTUDT00631835 | Refresh | Closed | hilas | November 12, 2020 9:31:54 PM GMT+08:00 | FOK: after system upgrade Auto voice Language changed to Arabic instead of Mandarin Chinese | 4.7.7 |  |  | Subsystem | SP2_Candidate | Console-SRS.15464 | Not Enough Information | venkatachalam_p | 0 |  |  | Console4 | NO | Closed |  | IQon1 | PR405382
CTUDT00631832 | Refresh | Accepted | dongping_w | November 12, 2020 2:48:35 PM GMT+08:00 | FOK BB- FOV change back to 600mm after setting to 350mm when using off line recon | WIP12.6 |  |  | Subsystem | FOK |  |  | dongping_w | 0 |  |  | Support-System | NO | Accepted |  | Brilliance Big Bore
CTUDT00631831 | Refresh | Resolved | rakic_k | November 12, 2020 7:09:24 AM GMT+08:00 | Hawk Performance: Not all images created error during 200 Patients Test | WIP33 | WIP34 | Yes, user can perform offline recon and all images will appear | Subsystem | SP1_PM_Approved | N/A | Fixed | rakic_k | 0 | 3 | CIRS_SW_ImgNodeSvr | CIRS | YES | Resolved | WIP34SubCom | Ovation |  | WIP34 | November 19, 2020 11:57:31 PM GMT+08:00
CTUDT00631831 | Refresh | Resolved | rakic_k | November 12, 2020 7:09:24 AM GMT+08:00 | Hawk Performance: Not all images created error during 200 Patients Test | WIP33 | WIP34 | Yes, user can perform offline recon and all images will appear | Subsystem | SP1_PM_Approved | N/A | Fixed | rakic_k | 0 | 3 | CIRS_SW_ImgNodeSvr | CIRS | YES | Resolved | WIP34SubCom | Brilliance 64 |  | WIP34 | November 19, 2020 11:57:31 PM GMT+08:00
CTUDT00631831 | Refresh | Resolved | rakic_k | November 12, 2020 7:09:24 AM GMT+08:00 | Hawk Performance: Not all images created error during 200 Patients Test | WIP33 | WIP34 | Yes, user can perform offline recon and all images will appear | Subsystem | SP1_PM_Approved | N/A | Fixed | rakic_k | 0 | 3 | CIRS_SW_ImgNodeSvr | CIRS | YES | Resolved | WIP34SubCom | Brilliance Big Bore |  | WIP34 | November 19, 2020 11:57:31 PM GMT+08:00
CTUDT00631831 | Refresh | Resolved | rakic_k | November 12, 2020 7:09:24 AM GMT+08:00 | Hawk Performance: Not all images created error during 200 Patients Test | WIP33 | WIP34 | Yes, user can perform offline recon and all images will appear | Subsystem | SP1_PM_Approved | N/A | Fixed | rakic_k | 0 | 3 | CIRS_SW_ImgNodeSvr | CIRS | YES | Resolved | WIP34SubCom | Ingenuity 64 |  | WIP34 | November 19, 2020 11:57:31 PM GMT+08:00
CTUDT00631831 | Refresh | Resolved | rakic_k | November 12, 2020 7:09:24 AM GMT+08:00 | Hawk Performance: Not all images created error during 200 Patients Test | WIP33 | WIP34 | Yes, user can perform offline recon and all images will appear | Subsystem | SP1_PM_Approved | N/A | Fixed | rakic_k | 0 | 3 | CIRS_SW_ImgNodeSvr | CIRS | YES | Resolved | WIP34SubCom | IQon1 |  | WIP34 | November 19, 2020 11:57:31 PM GMT+08:00
CTUDT00631827 | Refresh | Closed | hilas | November 9, 2020 11:37:26 PM GMT+08:00 | FOK: Host stuck and status monitor flickering | 4.7.7 |  | NA | Subsystem |  | NA | Duplicate | tiwari_g | 0 |  |  | Console4 | NO | Closed |  | Ovation | PR388983
CTUDT00631827 | Refresh | Closed | hilas | November 9, 2020 11:37:26 PM GMT+08:00 | FOK: Host stuck and status monitor flickering | 4.7.7 |  | NA | Subsystem |  | NA | Duplicate | tiwari_g | 0 |  |  | Console4 | NO | Closed |  | Brilliance 64 | PR388983
CTUDT00631827 | Refresh | Closed | hilas | November 9, 2020 11:37:26 PM GMT+08:00 | FOK: Host stuck and status monitor flickering | 4.7.7 |  | NA | Subsystem |  | NA | Duplicate | tiwari_g | 0 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore | PR388983
CTUDT00631827 | Refresh | Closed | hilas | November 9, 2020 11:37:26 PM GMT+08:00 | FOK: Host stuck and status monitor flickering | 4.7.7 |  | NA | Subsystem |  | NA | Duplicate | tiwari_g | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64 | PR388983
CTUDT00631827 | Refresh | Closed | hilas | November 9, 2020 11:37:26 PM GMT+08:00 | FOK: Host stuck and status monitor flickering | 4.7.7 |  | NA | Subsystem |  | NA | Duplicate | tiwari_g | 0 |  |  | Console4 | NO | Closed |  | IQon1 | PR388983
CTUDT00631826 | Refresh | Assigned | altshuller_mark | November 9, 2020 9:37:08 PM GMT+08:00 | Restrict USB boot by modifying BIOS Boot order. | Hawk |  |  | Subsystem | SP1_PM_Approved | N/A |  | hilas | 0 | 2 | Documentation | Support-System | NO | Assigned | SP1_InstallationSpec | Ovation
CTUDT00631826 | Refresh | Assigned | altshuller_mark | November 9, 2020 9:37:08 PM GMT+08:00 | Restrict USB boot by modifying BIOS Boot order. | Hawk |  |  | Subsystem | SP1_PM_Approved | N/A |  | hilas | 0 | 2 | Documentation | Support-System | NO | Assigned | SP1_InstallationSpec | Brilliance 64
CTUDT00631826 | Refresh | Assigned | altshuller_mark | November 9, 2020 9:37:08 PM GMT+08:00 | Restrict USB boot by modifying BIOS Boot order. | Hawk |  |  | Subsystem | SP1_PM_Approved | N/A |  | hilas | 0 | 2 | Documentation | Support-System | NO | Assigned | SP1_InstallationSpec | Ingenuity 64
CTUDT00631826 | Refresh | Assigned | altshuller_mark | November 9, 2020 9:37:08 PM GMT+08:00 | Restrict USB boot by modifying BIOS Boot order. | Hawk |  |  | Subsystem | SP1_PM_Approved | N/A |  | hilas | 0 | 2 | Documentation | Support-System | NO | Assigned | SP1_InstallationSpec | Ingenuity BigBore
CTUDT00631826 | Refresh | Assigned | altshuller_mark | November 9, 2020 9:37:08 PM GMT+08:00 | Restrict USB boot by modifying BIOS Boot order. | Hawk |  |  | Subsystem | SP1_PM_Approved | N/A |  | hilas | 0 | 2 | Documentation | Support-System | NO | Assigned | SP1_InstallationSpec | IQon1
CTUDT00631824 | Refresh | Assigned | lin_s | November 9, 2020 8:21:30 PM GMT+08:00 | [Refresh]: Cardiac Viewer and CT Reporting are in option list of KGS for Br64 and also same as marketing strategy, but the SRS had been removed | RFD |  |  | Subsystem |  | ProductConfig |  | lin_s | 0 | 3 |  | Support-System | NO | Assigned | SP1_Doc | Brilliance 64
CTUDT00631822 | Refresh | Closed | hilas | November 9, 2020 5:01:02 PM GMT+08:00 | FOK: Refresh Hotfix2 installation failed. | 4.7.7 |  |  | Subsystem | FOK | NA | Works As Specified / Not A Defect | venugopal_s | 0 |  |  | DI | NO | Closed |  | IQon1 | PR389055
CTUDT00631819 | Refresh | Postponed | bingxin_lin | November 9, 2020 4:30:45 PM GMT+08:00 | Error message popped up after remove the X2 Base Unit device | WIP33 |  |  | Subsystem |  | n/a |  | frider_o | 0 | 4 |  | DMS | NO | Postponed | SubNC | Brilliance Big Bore
CTUDT00631818 | Refresh | Accepted | altshuller_mark | November 9, 2020 1:52:11 PM GMT+08:00 | Network ports for internal system communication are configured for two network ranges in the firewall instead of a single network range | Hawk |  |  | Subsystem | SP2_Candidate |  |  | periasamy_g | 0 |  |  | Console4 | YES | Accepted |  | Ovation
CTUDT00631818 | Refresh | Accepted | altshuller_mark | November 9, 2020 1:52:11 PM GMT+08:00 | Network ports for internal system communication are configured for two network ranges in the firewall instead of a single network range | Hawk |  |  | Subsystem | SP2_Candidate |  |  | periasamy_g | 0 |  |  | Console4 | YES | Accepted |  | Brilliance 64
CTUDT00631818 | Refresh | Accepted | altshuller_mark | November 9, 2020 1:52:11 PM GMT+08:00 | Network ports for internal system communication are configured for two network ranges in the firewall instead of a single network range | Hawk |  |  | Subsystem | SP2_Candidate |  |  | periasamy_g | 0 |  |  | Console4 | YES | Accepted |  | Brilliance Big Bore
CTUDT00631818 | Refresh | Accepted | altshuller_mark | November 9, 2020 1:52:11 PM GMT+08:00 | Network ports for internal system communication are configured for two network ranges in the firewall instead of a single network range | Hawk |  |  | Subsystem | SP2_Candidate |  |  | periasamy_g | 0 |  |  | Console4 | YES | Accepted |  | Ingenuity 64
CTUDT00631818 | Refresh | Accepted | altshuller_mark | November 9, 2020 1:52:11 PM GMT+08:00 | Network ports for internal system communication are configured for two network ranges in the firewall instead of a single network range | Hawk |  |  | Subsystem | SP2_Candidate |  |  | periasamy_g | 0 |  |  | Console4 | YES | Accepted |  | IQon1
CTUDT00631818 | Refresh | Accepted | altshuller_mark | November 9, 2020 1:52:11 PM GMT+08:00 | Network ports for internal system communication are configured for two network ranges in the firewall instead of a single network range | Hawk |  |  | Subsystem | SP2_Candidate |  |  | periasamy_g | 0 |  |  | Console4 | YES | Accepted |  | CT 7500
CTUDT00631817 | Refresh | Accepted | altshuller_mark | November 9, 2020 1:47:04 PM GMT+08:00 | Whitelisting is not restricting unauthorized software execution through WMI commands | Hawk |  |  | Subsystem | SP2_Candidate |  |  | periasamy_g | 0 |  |  | Console4 | YES | Accepted |  | Ovation
CTUDT00631817 | Refresh | Accepted | altshuller_mark | November 9, 2020 1:47:04 PM GMT+08:00 | Whitelisting is not restricting unauthorized software execution through WMI commands | Hawk |  |  | Subsystem | SP2_Candidate |  |  | periasamy_g | 0 |  |  | Console4 | YES | Accepted |  | Brilliance 64
CTUDT00631817 | Refresh | Accepted | altshuller_mark | November 9, 2020 1:47:04 PM GMT+08:00 | Whitelisting is not restricting unauthorized software execution through WMI commands | Hawk |  |  | Subsystem | SP2_Candidate |  |  | periasamy_g | 0 |  |  | Console4 | YES | Accepted |  | Brilliance Big Bore
CTUDT00631817 | Refresh | Accepted | altshuller_mark | November 9, 2020 1:47:04 PM GMT+08:00 | Whitelisting is not restricting unauthorized software execution through WMI commands | Hawk |  |  | Subsystem | SP2_Candidate |  |  | periasamy_g | 0 |  |  | Console4 | YES | Accepted |  | Ingenuity 64
CTUDT00631817 | Refresh | Accepted | altshuller_mark | November 9, 2020 1:47:04 PM GMT+08:00 | Whitelisting is not restricting unauthorized software execution through WMI commands | Hawk |  |  | Subsystem | SP2_Candidate |  |  | periasamy_g | 0 |  |  | Console4 | YES | Accepted |  | IQon1
CTUDT00631817 | Refresh | Accepted | altshuller_mark | November 9, 2020 1:47:04 PM GMT+08:00 | Whitelisting is not restricting unauthorized software execution through WMI commands | Hawk |  |  | Subsystem | SP2_Candidate |  |  | periasamy_g | 0 |  |  | Console4 | YES | Accepted |  | CT 7500
CTUDT00631816 | Refresh | Accepted | altshuller_mark | November 9, 2020 1:41:50 PM GMT+08:00 | SSH port is open on the console for network connections from unauthorized sources | Hawk |  |  | Subsystem | SP2_Candidate |  |  | periasamy_g | 0 |  |  | Console4 | YES | Accepted |  | Ovation
CTUDT00631816 | Refresh | Accepted | altshuller_mark | November 9, 2020 1:41:50 PM GMT+08:00 | SSH port is open on the console for network connections from unauthorized sources | Hawk |  |  | Subsystem | SP2_Candidate |  |  | periasamy_g | 0 |  |  | Console4 | YES | Accepted |  | Brilliance 64
CTUDT00631816 | Refresh | Accepted | altshuller_mark | November 9, 2020 1:41:50 PM GMT+08:00 | SSH port is open on the console for network connections from unauthorized sources | Hawk |  |  | Subsystem | SP2_Candidate |  |  | periasamy_g | 0 |  |  | Console4 | YES | Accepted |  | Brilliance Big Bore
CTUDT00631816 | Refresh | Accepted | altshuller_mark | November 9, 2020 1:41:50 PM GMT+08:00 | SSH port is open on the console for network connections from unauthorized sources | Hawk |  |  | Subsystem | SP2_Candidate |  |  | periasamy_g | 0 |  |  | Console4 | YES | Accepted |  | Ingenuity 64
CTUDT00631816 | Refresh | Accepted | altshuller_mark | November 9, 2020 1:41:50 PM GMT+08:00 | SSH port is open on the console for network connections from unauthorized sources | Hawk |  |  | Subsystem | SP2_Candidate |  |  | periasamy_g | 0 |  |  | Console4 | YES | Accepted |  | IQon1
CTUDT00631816 | Refresh | Accepted | altshuller_mark | November 9, 2020 1:41:50 PM GMT+08:00 | SSH port is open on the console for network connections from unauthorized sources | Hawk |  |  | Subsystem | SP2_Candidate |  |  | periasamy_g | 0 |  |  | Console4 | YES | Accepted |  | CT 7500
CTUDT00631815 | Refresh | Accepted | altshuller_mark | November 9, 2020 1:33:55 PM GMT+08:00 | Content can be downloaded from unauthorized external sources | Hawk |  |  | Subsystem | SP2_Candidate |  |  | periasamy_g | 0 |  |  | Console4 | YES | Accepted |  | Ovation
CTUDT00631815 | Refresh | Accepted | altshuller_mark | November 9, 2020 1:33:55 PM GMT+08:00 | Content can be downloaded from unauthorized external sources | Hawk |  |  | Subsystem | SP2_Candidate |  |  | periasamy_g | 0 |  |  | Console4 | YES | Accepted |  | Brilliance 64
CTUDT00631815 | Refresh | Accepted | altshuller_mark | November 9, 2020 1:33:55 PM GMT+08:00 | Content can be downloaded from unauthorized external sources | Hawk |  |  | Subsystem | SP2_Candidate |  |  | periasamy_g | 0 |  |  | Console4 | YES | Accepted |  | Brilliance Big Bore
CTUDT00631815 | Refresh | Accepted | altshuller_mark | November 9, 2020 1:33:55 PM GMT+08:00 | Content can be downloaded from unauthorized external sources | Hawk |  |  | Subsystem | SP2_Candidate |  |  | periasamy_g | 0 |  |  | Console4 | YES | Accepted |  | Ingenuity 64
CTUDT00631815 | Refresh | Accepted | altshuller_mark | November 9, 2020 1:33:55 PM GMT+08:00 | Content can be downloaded from unauthorized external sources | Hawk |  |  | Subsystem | SP2_Candidate |  |  | periasamy_g | 0 |  |  | Console4 | YES | Accepted |  | IQon1
CTUDT00631815 | Refresh | Accepted | altshuller_mark | November 9, 2020 1:33:55 PM GMT+08:00 | Content can be downloaded from unauthorized external sources | Hawk |  |  | Subsystem | SP2_Candidate |  |  | periasamy_g | 0 |  |  | Console4 | YES | Accepted |  | CT 7500
CTUDT00631806 | Refresh | Accepted | hilas | November 8, 2020 9:29:34 PM GMT+08:00 | Refresh SP1: New Branding requirements for iCT scanners need to be implemented in Refresh for China systems only | 4.1.10 |  |  | Subsystem | SP1_PM_Approved |  |  | hilas | 0 |  |  | Support-System | NO | Accepted | SP1_Doc | Ovation
CTUDT00631802 | Refresh | Resolved | hilas | November 8, 2020 9:15:18 PM GMT+08:00 | Refresh SP: New Branding requirements for iCT scanners need to be implemented in Refresh for China systems only | 4.1.10 | WIP36 | NA | Subsystem | SP1_PM_Approved | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Resolved | WIP35SubCom;WIP36SubCom | Ovation |  | WIP36 | November 27, 2020 1:35:09 PM GMT+08:00
CTUDT00631800 | Refresh | Resolved | altshuller_mark | November 8, 2020 9:11:50 PM GMT+08:00 | Hawk:When End exam button is pressed Exam Summary Series and Dose structured series (5000) are not generated in PD in Dutch language. | Hawk | WIP35 | NA | Subsystem | SP1_PM_Approved | N/A | Fixed | challa_s | 0 | 2 | Console Software | Console4 | NO | Resolved | WIP35SubCom | Ovation |  | WIP35 | November 25, 2020 1:35:28 PM GMT+08:00
CTUDT00631800 | Refresh | Resolved | altshuller_mark | November 8, 2020 9:11:50 PM GMT+08:00 | Hawk:When End exam button is pressed Exam Summary Series and Dose structured series (5000) are not generated in PD in Dutch language. | Hawk | WIP35 | NA | Subsystem | SP1_PM_Approved | N/A | Fixed | challa_s | 0 | 2 | Console Software | Console4 | NO | Resolved | WIP35SubCom | Brilliance 64 |  | WIP35 | November 25, 2020 1:35:28 PM GMT+08:00
CTUDT00631800 | Refresh | Resolved | altshuller_mark | November 8, 2020 9:11:50 PM GMT+08:00 | Hawk:When End exam button is pressed Exam Summary Series and Dose structured series (5000) are not generated in PD in Dutch language. | Hawk | WIP35 | NA | Subsystem | SP1_PM_Approved | N/A | Fixed | challa_s | 0 | 2 | Console Software | Console4 | NO | Resolved | WIP35SubCom | Ingenuity 64 |  | WIP35 | November 25, 2020 1:35:28 PM GMT+08:00
CTUDT00631800 | Refresh | Resolved | altshuller_mark | November 8, 2020 9:11:50 PM GMT+08:00 | Hawk:When End exam button is pressed Exam Summary Series and Dose structured series (5000) are not generated in PD in Dutch language. | Hawk | WIP35 | NA | Subsystem | SP1_PM_Approved | N/A | Fixed | challa_s | 0 | 2 | Console Software | Console4 | NO | Resolved | WIP35SubCom | IQon1 |  | WIP35 | November 25, 2020 1:35:28 PM GMT+08:00
CTUDT00631799 | Refresh | Closed | altshuller_mark | November 8, 2020 3:03:51 PM GMT+08:00 | Excessive privileges in file system application configuration files | Hawk |  |  | Subsystem |  | N/A | No Plan to Fix | tiwari_g | 0 | 3 |  | Console4 | NO | Closed |  | Ovation
CTUDT00631799 | Refresh | Closed | altshuller_mark | November 8, 2020 3:03:51 PM GMT+08:00 | Excessive privileges in file system application configuration files | Hawk |  |  | Subsystem |  | N/A | No Plan to Fix | tiwari_g | 0 | 3 |  | Console4 | NO | Closed |  | IQon1
CTUDT00631799 | Refresh | Closed | altshuller_mark | November 8, 2020 3:03:51 PM GMT+08:00 | Excessive privileges in file system application configuration files | Hawk |  |  | Subsystem |  | N/A | No Plan to Fix | tiwari_g | 0 | 3 |  | Console4 | NO | Closed |  | Hawk
CTUDT00631795 | Refresh | Closed | prabhan_p | November 6, 2020 10:25:03 PM GMT+08:00 | Parameters are not set as default after Select “CT Simulation Exam Card” option. | Wip 33 |  | NA | Subsystem | SP2_Postpone | NA | Works As Specified / Not A Defect | suresh_b | 0 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00631793 | Refresh | Assigned | prabhan_p | November 6, 2020 8:51:07 PM GMT+08:00 | Head Dual Energy or Body Dual Energy was not present in Scan type drop down. | WIP 33 |  |  | Subsystem | SP1_PM_Approved; Test Update | NA |  | jindal_k | 0 | 3 |  | Console4 | NO | Assigned | Test Design | IQon1
CTUDT00631791 | Refresh | Accepted | prabhan_p | November 6, 2020 8:33:00 PM GMT+08:00 | Error displayed while login as CT user first time. | Wip 33 |  |  | Subsystem | SP2_Candidate |  |  | mukherjee_s | 1 |  |  | Console4 | NO | Accepted |  | Brilliance Big Bore
CTUDT00631789 | Refresh | Closed | barnea_d | November 6, 2020 1:07:48 AM GMT+08:00 | Refresh: QA: Percentage Bar shows 100% from beginning of scan for in Service User- Intermittently | WIP 33 |  |  | Subsystem |  |  | Duplicate | frider_o | 0 |  |  | DMS | NO | Closed |  | Ovation
CTUDT00631769 | Refresh | Assigned | lin_s | November 3, 2020 8:41:27 PM GMT+08:00 | RADAR is unable to collect Huge USP logs due to missing UspLog Extractor tool in Refresh release | 4.7.7 |  |  | Subsystem |  | CT-NM.SRS.Service.3137 |  | lin_s | 0 | 3 | Service Tools_Software | Support-System | NO | Assigned | SP1_Doc | Ovation
CTUDT00631769 | Refresh | Assigned | lin_s | November 3, 2020 8:41:27 PM GMT+08:00 | RADAR is unable to collect Huge USP logs due to missing UspLog Extractor tool in Refresh release | 4.7.7 |  |  | Subsystem |  | CT-NM.SRS.Service.3137 |  | lin_s | 0 | 3 | Service Tools_Software | Support-System | NO | Assigned | SP1_Doc | Brilliance 64
CTUDT00631769 | Refresh | Assigned | lin_s | November 3, 2020 8:41:27 PM GMT+08:00 | RADAR is unable to collect Huge USP logs due to missing UspLog Extractor tool in Refresh release | 4.7.7 |  |  | Subsystem |  | CT-NM.SRS.Service.3137 |  | lin_s | 0 | 3 | Service Tools_Software | Support-System | NO | Assigned | SP1_Doc | Brilliance Big Bore
CTUDT00631769 | Refresh | Assigned | lin_s | November 3, 2020 8:41:27 PM GMT+08:00 | RADAR is unable to collect Huge USP logs due to missing UspLog Extractor tool in Refresh release | 4.7.7 |  |  | Subsystem |  | CT-NM.SRS.Service.3137 |  | lin_s | 0 | 3 | Service Tools_Software | Support-System | NO | Assigned | SP1_Doc | Ingenuity 64
CTUDT00631769 | Refresh | Assigned | lin_s | November 3, 2020 8:41:27 PM GMT+08:00 | RADAR is unable to collect Huge USP logs due to missing UspLog Extractor tool in Refresh release | 4.7.7 |  |  | Subsystem |  | CT-NM.SRS.Service.3137 |  | lin_s | 0 | 3 | Service Tools_Software | Support-System | NO | Assigned | SP1_Doc | IQon1
CTUDT00631768 | Refresh | Accepted | shterner_a | November 3, 2020 8:27:22 PM GMT+08:00 | Refresh: critical error when cancel sign out of CT or Service user | WIP33 |  |  | Subsystem | SP2_Candidate |  |  | mukherjee_s | 0 |  |  | Console4 | NO | Accepted |  | Ovation
CTUDT00631768 | Refresh | Accepted | shterner_a | November 3, 2020 8:27:22 PM GMT+08:00 | Refresh: critical error when cancel sign out of CT or Service user | WIP33 |  |  | Subsystem | SP2_Candidate |  |  | mukherjee_s | 0 |  |  | Console4 | NO | Accepted |  | Hawk
CTUDT00631764 | Refresh | Accepted | bingxin_lin | November 3, 2020 6:46:21 PM GMT+08:00 | [FOK][BB Refresh][IVC]Need a method to track multiple needles (at least 4), have a real time measurement of the distance between each needle to allow the user to know they are parallel | RFD |  |  | System | FOK, IVC SW; SP2_Score2Postpone |  |  | mukherjee_s | 0 |  |  | Console4 | NO | Accepted |  | Brilliance Big Bore
CTUDT00631762 | Refresh | NotAccepted | bingxin_lin | November 3, 2020 6:42:12 PM GMT+08:00 | [FOK][BB Refresh][IVC]Measurements & tools need to be streamlines | RFD |  |  | System | FOK |  |  |  | 0 |  |  |  |  | NotAccepted |  | Brilliance Big Bore
CTUDT00631760 | Refresh | NotAccepted | bingxin_lin | November 3, 2020 6:38:35 PM GMT+08:00 | [FOK][BB Refresh][IVC]When a scan or needle location has been saved, & the table moved to that location, the center image of a 3 display will be at that location, but not in volume | RFC |  |  | System | FOK |  |  |  | 0 |  |  |  |  | NotAccepted |  | Brilliance Big Bore
CTUDT00631758 | Refresh | Accepted | bingxin_lin | November 3, 2020 6:34:12 PM GMT+08:00 | [FOK][BB Refresh][IVC]Although selected saved needle position ,we still have to select the radio button before moving to that location. | RFD |  |  | System | FOK, IVC SW; SP2_Score2Postpone |  |  | mukherjee_s | 0 |  |  | Console4 | NO | Accepted |  | Brilliance Big Bore
CTUDT00631756 | Refresh | Accepted | bingxin_lin | November 3, 2020 6:28:24 PM GMT+08:00 | [FOK][BB Refresh][IVC]The incorrect couch position indicated on Surview and during Recon mode for CCT procedure | RFD |  |  | Subsystem | FOK; SP2_Candidate |  |  | mukherjee_s | 0 |  |  | Console4 | YES | Accepted |  | Brilliance Big Bore
CTUDT00631753 | Refresh | NotAccepted | bingxin_lin | November 3, 2020 5:37:16 PM GMT+08:00 | [FOK][BB Refresh][IVC]Metal reduction for small metals eg needles required to adequately see surrounding tissue & structures. | RFD |  |  | System | FOK |  |  |  | 0 |  |  |  |  | NotAccepted |  | Brilliance Big Bore
CTUDT00631749 | Refresh | NotAccepted | bingxin_lin | November 3, 2020 4:54:29 PM GMT+08:00 | [FOK][BB Refresh][IVC]The technologists would like to have the option of typing in a scan pos. at the console during scan planning on console. | RFD |  |  | System | FOK |  |  |  | 0 |  |  |  |  | NotAccepted |  | Brilliance Big Bore
CTUDT00631746 | Refresh | Accepted | bingxin_lin | November 3, 2020 4:46:38 PM GMT+08:00 | [FOK][BB Refresh][IVC] Request to be able to copy and paste a graphic to all images in the scan viewer during volume CCT mode | RFD |  |  | System | FOK; SP2_Score2Postpone |  |  | mukherjee_s | 0 |  |  | Console4 | NO | Accepted |  | Brilliance Big Bore
CTUDT00631744 | Refresh | Closed | bingxin_lin | November 3, 2020 4:40:30 PM GMT+08:00 | [FOK][BB Refresh][IVC] Gantry Initialization time is way too long | RFD |  |  | Subsystem | FOK | na | Works As Specified / Not A Defect | zhong_x | 0 |  |  | GCX | NO | Closed |  | Brilliance Big Bore
CTUDT00631742 | Refresh | Accepted | bingxin_lin | November 3, 2020 4:30:44 PM GMT+08:00 | [FOK][BB Refresh][IVC] Save a scan or needle position is available after adding CCT EC | RFD |  |  | System | FOK; SP2_Score2Postpone |  |  | mukherjee_s | 0 |  |  | Console4 | NO | Accepted |  | Brilliance Big Bore
CTUDT00631739 | Refresh | Accepted | bingxin_lin | November 3, 2020 4:15:30 PM GMT+08:00 | [FOK][BB Refresh][IVC]To be able to start X-Rays from the control room (CT-Box or SW) even when in CCT | RFD |  |  | System | FOK; SP2_Score2Postpone |  |  | mukherjee_s | 0 |  |  | Console4 | NO | Accepted |  | Brilliance Big Bore
CTUDT00631737 | Refresh | NotAccepted | bingxin_lin | November 3, 2020 4:09:22 PM GMT+08:00 | [FOK][BB Refresh][IVC] Boot of the touchscreen is very long, but fortunately HW bed controls can already be used | RFD |  |  | Subsystem | FOK |  |  |  | 0 |  |  |  |  | NotAccepted |  | Brilliance Big Bore
CTUDT00631735 | Refresh | NotAccepted | bingxin_lin | November 3, 2020 3:44:38 PM GMT+08:00 | [FOK][BB Refresh][IVC] The order of options in the context menu needs to be changed | RFD |  |  | System | FOK |  |  |  | 0 |  |  |  |  | NotAccepted |  | Brilliance Big Bore
CTUDT00631733 | Refresh | NotAccepted | bingxin_lin | November 3, 2020 3:41:47 PM GMT+08:00 | [FOK][BB Refresh][IVC]Scrolling through images awkward | RFD |  |  | System | FOK |  |  |  | 0 |  |  |  |  | NotAccepted |  | Brilliance Big Bore
CTUDT00631730 | Refresh | Accepted | bingxin_lin | November 3, 2020 3:21:35 PM GMT+08:00 | [FOK][BB Refresh][IVC]Function improvement and function increase on IVC | RFD |  |  | System | FOK, RFC; SP2_Score2Postpone |  |  | mukherjee_s | 0 |  |  | Console4 | NO | Accepted |  | Brilliance Big Bore
CTUDT00631728 | Refresh | Accepted | bingxin_lin | November 3, 2020 2:58:47 PM GMT+08:00 | [FOK][BB Refresh][IVC]Some features on touch panel should be taken to Big Bore IVC | RFD |  |  | System | FOK, RFC; SP2_Score2Postpone |  |  | mukherjee_s | 0 |  |  | Console4 | NO | Accepted |  | Brilliance Big Bore
CTUDT00631725 | Refresh | Accepted | bingxin_lin | November 3, 2020 2:35:00 PM GMT+08:00 | [FOK][BB Refresh][IVC]The saved locations are not reflected on the scan plan | RFD |  |  | Subsystem | FOK RFC ; SP2_Score2Postpone |  |  | mukherjee_s | 0 |  |  | Console4 | NO | Accepted |  | Brilliance Big Bore
CTUDT00631722 | Refresh | Closed | dongping_w | November 3, 2020 1:34:41 PM GMT+08:00 | FOK Refresh_BB_ING: New passwords could cause Radar issues | WIP12.6 |  |  | Subsystem | FOK | NA | Works As Specified / Not A Defect | lin_s | 0 |  |  | Support-System | NO | Closed | WIP36SegReq | Ovation | na
CTUDT00631722 | Refresh | Closed | dongping_w | November 3, 2020 1:34:41 PM GMT+08:00 | FOK Refresh_BB_ING: New passwords could cause Radar issues | WIP12.6 |  |  | Subsystem | FOK | NA | Works As Specified / Not A Defect | lin_s | 0 |  |  | Support-System | NO | Closed | WIP36SegReq | Brilliance 64 | na
CTUDT00631722 | Refresh | Closed | dongping_w | November 3, 2020 1:34:41 PM GMT+08:00 | FOK Refresh_BB_ING: New passwords could cause Radar issues | WIP12.6 |  |  | Subsystem | FOK | NA | Works As Specified / Not A Defect | lin_s | 0 |  |  | Support-System | NO | Closed | WIP36SegReq | Brilliance Big Bore | na
CTUDT00631722 | Refresh | Closed | dongping_w | November 3, 2020 1:34:41 PM GMT+08:00 | FOK Refresh_BB_ING: New passwords could cause Radar issues | WIP12.6 |  |  | Subsystem | FOK | NA | Works As Specified / Not A Defect | lin_s | 0 |  |  | Support-System | NO | Closed | WIP36SegReq | Ingenuity 64 | na
CTUDT00631722 | Refresh | Closed | dongping_w | November 3, 2020 1:34:41 PM GMT+08:00 | FOK Refresh_BB_ING: New passwords could cause Radar issues | WIP12.6 |  |  | Subsystem | FOK | NA | Works As Specified / Not A Defect | lin_s | 0 |  |  | Support-System | NO | Closed | WIP36SegReq | IQon1 | na
CTUDT00631717 | Refresh | Closed | wang_l | November 2, 2020 2:40:54 PM GMT+08:00 | Refresh - ring-shaped artifact was shown in images | WIP22.5 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | yatziv_l | 0 |  |  | DMS | NO | Closed |  | Ingenuity 64 | PR 339255 PR 356828
CTUDT00631714 | Refresh | Accepted | lin_s | November 2, 2020 9:29:21 AM GMT+08:00 | Loading an EC to plan scene takes too long. | WIP25.5 |  |  | Subsystem | SP1_PM_Approved |  |  | mukherjee_s | 0 | 2 | Requirements Documentation | Console4 | NO | Accepted | WIP37SegReq | Ovation
CTUDT00631714 | Refresh | Accepted | lin_s | November 2, 2020 9:29:21 AM GMT+08:00 | Loading an EC to plan scene takes too long. | WIP25.5 |  |  | Subsystem | SP1_PM_Approved |  |  | mukherjee_s | 0 | 2 | Requirements Documentation | Console4 | NO | Accepted | WIP37SegReq | Brilliance 64
CTUDT00631714 | Refresh | Accepted | lin_s | November 2, 2020 9:29:21 AM GMT+08:00 | Loading an EC to plan scene takes too long. | WIP25.5 |  |  | Subsystem | SP1_PM_Approved |  |  | mukherjee_s | 0 | 2 | Requirements Documentation | Console4 | NO | Accepted | WIP37SegReq | Brilliance Big Bore
CTUDT00631714 | Refresh | Accepted | lin_s | November 2, 2020 9:29:21 AM GMT+08:00 | Loading an EC to plan scene takes too long. | WIP25.5 |  |  | Subsystem | SP1_PM_Approved |  |  | mukherjee_s | 0 | 2 | Requirements Documentation | Console4 | NO | Accepted | WIP37SegReq | Ingenuity 64
CTUDT00631714 | Refresh | Accepted | lin_s | November 2, 2020 9:29:21 AM GMT+08:00 | Loading an EC to plan scene takes too long. | WIP25.5 |  |  | Subsystem | SP1_PM_Approved |  |  | mukherjee_s | 0 | 2 | Requirements Documentation | Console4 | NO | Accepted | WIP37SegReq | IQon1
CTUDT00631713 | Refresh | Closed | altshuller_mark | November 1, 2020 9:19:16 PM GMT+08:00 | Attempt to reassign objsPartType Error during S&S TAVI protocol | V4.1.7 | WIP32 | User can perform offline reconstruction with the available data. | Subsystem | SP1_PM_Approved | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_Data_Management | CIRS | YES | Closed | WIP32SubCom | Ovation | 0112321625 0113008764 | WIP32 | November 4, 2020 10:18:39 PM GMT+08:00
CTUDT00631713 | Refresh | Closed | altshuller_mark | November 1, 2020 9:19:16 PM GMT+08:00 | Attempt to reassign objsPartType Error during S&S TAVI protocol | V4.1.7 | WIP32 | User can perform offline reconstruction with the available data. | Subsystem | SP1_PM_Approved | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_Data_Management | CIRS | YES | Closed | WIP32SubCom | Brilliance 64 | 0112321625 0113008764 | WIP32 | November 4, 2020 10:18:39 PM GMT+08:00
CTUDT00631713 | Refresh | Closed | altshuller_mark | November 1, 2020 9:19:16 PM GMT+08:00 | Attempt to reassign objsPartType Error during S&S TAVI protocol | V4.1.7 | WIP32 | User can perform offline reconstruction with the available data. | Subsystem | SP1_PM_Approved | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_Data_Management | CIRS | YES | Closed | WIP32SubCom | Ingenuity 64 | 0112321625 0113008764 | WIP32 | November 4, 2020 10:18:39 PM GMT+08:00
CTUDT00631713 | Refresh | Closed | altshuller_mark | November 1, 2020 9:19:16 PM GMT+08:00 | Attempt to reassign objsPartType Error during S&S TAVI protocol | V4.1.7 | WIP32 | User can perform offline reconstruction with the available data. | Subsystem | SP1_PM_Approved | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_Data_Management | CIRS | YES | Closed | WIP32SubCom | IQon1 | 0112321625 0113008764 | WIP32 | November 4, 2020 10:18:39 PM GMT+08:00
CTUDT00631712 | Refresh | Assigned | altshuller_mark | November 1, 2020 7:14:23 PM GMT+08:00 | HAWK : The offline ECG waves did not appear for first result of cardiac scan in view2 | Hawk |  |  | Subsystem | SP2_Candidate | N/A |  | tiwari_g | 0 | 2 | Console software | Console4 | NO | Assigned | SP2 | IQon1
CTUDT00631704 | Refresh | Opened | wang_l | October 30, 2020 9:05:59 AM GMT+08:00 | Refresh - Left Panel Stucked when load EC to Plan | WIP32 | WIP36 |  | Subsystem | SP1_PM_Approved | NA | Fixed | s_srivastava | 0 | 3 | Console Software | Console4 | NO | Opened | WIP37SegReq | Brilliance Big Bore |  | WIP36
CTUDT00631702 | Refresh | Closed | portal_l | October 29, 2020 10:12:06 PM GMT+08:00 | Ict WIP 31- A gap between Spec Of Reasons to Workflow under RND user for Harmonized phantom and for Legacy for UHR Focal Spot X Position Checkbox . | WIP 31 | WIP34 |  | Subsystem | SP1_PM_Approved | n/a | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP34SubCom | Ovation |  | WIP34 | November 12, 2020 7:22:57 PM GMT+08:00
CTUDT00631700 | Refresh | Closed | portal_l | October 29, 2020 9:52:12 PM GMT+08:00 | IQon Refresh WIP 31-SNR test per half tile limits Under BD calibration are not as spec requires. | WIP 31 | WIP34 |  | Subsystem | PM_Approved | n/a | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP34SubCom | IQon1 |  | WIP34 | November 12, 2020 7:20:01 PM GMT+08:00
CTUDT00631698 | Refresh | Accepted | s_namburi | October 29, 2020 9:09:39 PM GMT+08:00 | RFC: Need to update latest OS patches in Refresh product software | 4.1.10/4.7.7/4.8 |  |  | Subsystem | SP2_Candidate |  |  | mukherjee_s | 0 |  |  | Support-Security | NO | Accepted | SP2 | Ovation
CTUDT00631698 | Refresh | Accepted | s_namburi | October 29, 2020 9:09:39 PM GMT+08:00 | RFC: Need to update latest OS patches in Refresh product software | 4.1.10/4.7.7/4.8 |  |  | Subsystem | SP2_Candidate |  |  | mukherjee_s | 0 |  |  | Support-Security | NO | Accepted | SP2 | Brilliance 64
CTUDT00631698 | Refresh | Accepted | s_namburi | October 29, 2020 9:09:39 PM GMT+08:00 | RFC: Need to update latest OS patches in Refresh product software | 4.1.10/4.7.7/4.8 |  |  | Subsystem | SP2_Candidate |  |  | mukherjee_s | 0 |  |  | Support-Security | NO | Accepted | SP2 | Brilliance Big Bore
CTUDT00631698 | Refresh | Accepted | s_namburi | October 29, 2020 9:09:39 PM GMT+08:00 | RFC: Need to update latest OS patches in Refresh product software | 4.1.10/4.7.7/4.8 |  |  | Subsystem | SP2_Candidate |  |  | mukherjee_s | 0 |  |  | Support-Security | NO | Accepted | SP2 | Ingenuity 64
CTUDT00631698 | Refresh | Accepted | s_namburi | October 29, 2020 9:09:39 PM GMT+08:00 | RFC: Need to update latest OS patches in Refresh product software | 4.1.10/4.7.7/4.8 |  |  | Subsystem | SP2_Candidate |  |  | mukherjee_s | 0 |  |  | Support-Security | NO | Accepted | SP2 | IQon1
CTUDT00631694 | Refresh | Closed | 320085947 | October 29, 2020 8:20:02 PM GMT+08:00 | The Clinical scan starts automatically after the tracker scan completes with no indication. | WIP25.8 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | tiwari_g | 0 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00631692 | Refresh | Closed | hilas | October 29, 2020 7:14:33 PM GMT+08:00 | FOK: The HEAD Cushion is too thick and make Patient uncomfortable | 4.7.7 |  | Customer use the cushion from competitor. | Subsystem | FoK | N/A | Change Proposal | wang_ada | 0 |  |  | Couch | NO | Closed |  | Ovation | PR203334
CTUDT00631692 | Refresh | Closed | hilas | October 29, 2020 7:14:33 PM GMT+08:00 | FOK: The HEAD Cushion is too thick and make Patient uncomfortable | 4.7.7 |  | Customer use the cushion from competitor. | Subsystem | FoK | N/A | Change Proposal | wang_ada | 0 |  |  | Couch | NO | Closed |  | Brilliance 64 | PR203334
CTUDT00631692 | Refresh | Closed | hilas | October 29, 2020 7:14:33 PM GMT+08:00 | FOK: The HEAD Cushion is too thick and make Patient uncomfortable | 4.7.7 |  | Customer use the cushion from competitor. | Subsystem | FoK | N/A | Change Proposal | wang_ada | 0 |  |  | Couch | NO | Closed |  | Brilliance Big Bore | PR203334
CTUDT00631692 | Refresh | Closed | hilas | October 29, 2020 7:14:33 PM GMT+08:00 | FOK: The HEAD Cushion is too thick and make Patient uncomfortable | 4.7.7 |  | Customer use the cushion from competitor. | Subsystem | FoK | N/A | Change Proposal | wang_ada | 0 |  |  | Couch | NO | Closed |  | Ingenuity 64 | PR203334
CTUDT00631692 | Refresh | Closed | hilas | October 29, 2020 7:14:33 PM GMT+08:00 | FOK: The HEAD Cushion is too thick and make Patient uncomfortable | 4.7.7 |  | Customer use the cushion from competitor. | Subsystem | FoK | N/A | Change Proposal | wang_ada | 0 |  |  | Couch | NO | Closed |  | IQon1 | PR203334
CTUDT00631690 | Refresh | Closed | wang_l | October 29, 2020 4:45:00 PM GMT+08:00 | Refresh - Remote Configuration/Tight VNC/Enable Remote Access/ULV did not work | WIP32 | WIP35 | N/A | Subsystem | SP1_PM_Approved | NA | Fixed | venkatachalam_p | 0 | 2 | Service Tools_Software | Console4 | NO | Closed | WIP35SubCom | Ovation |  | WIP35 | December 1, 2020 4:24:41 PM GMT+08:00
CTUDT00631690 | Refresh | Closed | wang_l | October 29, 2020 4:45:00 PM GMT+08:00 | Refresh - Remote Configuration/Tight VNC/Enable Remote Access/ULV did not work | WIP32 | WIP35 | N/A | Subsystem | SP1_PM_Approved | NA | Fixed | venkatachalam_p | 0 | 2 | Service Tools_Software | Console4 | NO | Closed | WIP35SubCom | Brilliance 64 |  | WIP35 | December 1, 2020 4:24:41 PM GMT+08:00
CTUDT00631690 | Refresh | Closed | wang_l | October 29, 2020 4:45:00 PM GMT+08:00 | Refresh - Remote Configuration/Tight VNC/Enable Remote Access/ULV did not work | WIP32 | WIP35 | N/A | Subsystem | SP1_PM_Approved | NA | Fixed | venkatachalam_p | 0 | 2 | Service Tools_Software | Console4 | NO | Closed | WIP35SubCom | Brilliance Big Bore |  | WIP35 | December 1, 2020 4:24:41 PM GMT+08:00
CTUDT00631690 | Refresh | Closed | wang_l | October 29, 2020 4:45:00 PM GMT+08:00 | Refresh - Remote Configuration/Tight VNC/Enable Remote Access/ULV did not work | WIP32 | WIP35 | N/A | Subsystem | SP1_PM_Approved | NA | Fixed | venkatachalam_p | 0 | 2 | Service Tools_Software | Console4 | NO | Closed | WIP35SubCom | Ingenuity 64 |  | WIP35 | December 1, 2020 4:24:41 PM GMT+08:00
CTUDT00631690 | Refresh | Closed | wang_l | October 29, 2020 4:45:00 PM GMT+08:00 | Refresh - Remote Configuration/Tight VNC/Enable Remote Access/ULV did not work | WIP32 | WIP35 | N/A | Subsystem | SP1_PM_Approved | NA | Fixed | venkatachalam_p | 0 | 2 | Service Tools_Software | Console4 | NO | Closed | WIP35SubCom | IQon1 |  | WIP35 | December 1, 2020 4:24:41 PM GMT+08:00
CTUDT00631671 | Refresh | Resolved | zhou_r | October 28, 2020 1:52:18 PM GMT+08:00 | FOK- The customer reported a problem with the horizontal table movement | FOK | E-111607 |  | Subsystem | SP2_Candidate; HW | N/A | Fixed | zhou_r | 0 | 2 | Couch Mechanical | Couch | YES | Resolved | HW Replace | Ingenuity 64 | PR#265518 & PR#272532 PR360975 | E-111607 | January 4, 2021 3:45:44 PM GMT+08:00
CTUDT00631669 | Refresh | Closed | cx_zhang | October 28, 2020 10:19:16 AM GMT+08:00 | FOK: Couch manual control moving cause ESTOP in the hospital | WIP30.1 |  |  | System | FOK | N/A | Works As Specified / Not A Defect | cx_zhang | 0 |  |  | Couch | NO | Closed |  | Ingenuity 64 | PR324680 PR305910 PR348881
CTUDT00631663 | Refresh | Resolved | hilas | October 27, 2020 10:07:46 PM GMT+08:00 | Refresh: New Branding requirements for iCT scanners need to be implemented in Refresh for China systems only | 4.1.10 | WIP36 |  | Subsystem | SP1_PM_Approved | Console-SRS.18409 | Fixed | s_srivastava | 0 | 2 | Requirements Documentation | Console4 | NO | Resolved | WIP36SubCom | Ovation |  | WIP36 | December 12, 2020 1:59:41 AM GMT+08:00
CTUDT00631645 | Refresh | Accepted | Kevin_ZHAO | October 26, 2020 3:08:32 PM GMT+08:00 | Second BT scan can start by pressing Space bar in keyboard | WIP32 |  |  | Subsystem | CRN; SP2_Close; WAS |  |  | rajalakshmi_j | 0 |  |  | Console4 | YES | Accepted | SP1-Mar-21-CRN | Brilliance Big Bore
CTUDT00631645 | Refresh | Accepted | Kevin_ZHAO | October 26, 2020 3:08:32 PM GMT+08:00 | Second BT scan can start by pressing Space bar in keyboard | WIP32 |  |  | Subsystem | CRN; SP2_Close; WAS |  |  | rajalakshmi_j | 0 |  |  | Console4 | YES | Accepted | SP1-Mar-21-CRN | Ingenuity 64
CTUDT00631640 | Refresh | Closed | lin_s | October 26, 2020 12:56:10 PM GMT+08:00 | FOK - Missing handle for legacy phantom of eye phantom | FOK |  | Workaroud: The issues were tracked by DEFOA, production sent correct eye phantom handle to customer. For long term solution, add proper eye phantom handle into system configuration (CE, PCM and C-BOM). ECR refer to ECR-077026. | Subsystem |  | N/A | Forwarded to Different Database | wang_ada | 0 | 2 |  | Couch | NO | Closed |  | Ovation
CTUDT00631640 | Refresh | Closed | lin_s | October 26, 2020 12:56:10 PM GMT+08:00 | FOK - Missing handle for legacy phantom of eye phantom | FOK |  | Workaroud: The issues were tracked by DEFOA, production sent correct eye phantom handle to customer. For long term solution, add proper eye phantom handle into system configuration (CE, PCM and C-BOM). ECR refer to ECR-077026. | Subsystem |  | N/A | Forwarded to Different Database | wang_ada | 0 | 2 |  | Couch | NO | Closed |  | Brilliance Big Bore
CTUDT00631640 | Refresh | Closed | lin_s | October 26, 2020 12:56:10 PM GMT+08:00 | FOK - Missing handle for legacy phantom of eye phantom | FOK |  | Workaroud: The issues were tracked by DEFOA, production sent correct eye phantom handle to customer. For long term solution, add proper eye phantom handle into system configuration (CE, PCM and C-BOM). ECR refer to ECR-077026. | Subsystem |  | N/A | Forwarded to Different Database | wang_ada | 0 | 2 |  | Couch | NO | Closed |  | Ingenuity 64
CTUDT00631640 | Refresh | Closed | lin_s | October 26, 2020 12:56:10 PM GMT+08:00 | FOK - Missing handle for legacy phantom of eye phantom | FOK |  | Workaroud: The issues were tracked by DEFOA, production sent correct eye phantom handle to customer. For long term solution, add proper eye phantom handle into system configuration (CE, PCM and C-BOM). ECR refer to ECR-077026. | Subsystem |  | N/A | Forwarded to Different Database | wang_ada | 0 | 2 |  | Couch | NO | Closed |  | IQon1
CTUDT00631638 | Refresh | Assigned | lin_s | October 26, 2020 10:59:24 AM GMT+08:00 | FOK - scannable range line on pad doesn’t match table line | FOK |  |  | Subsystem |  | CT-NM.SRS.SysPerf.540/541/542/543/687/567 |  | wang_ada | 0 | 2 |  | Couch | YES | Assigned | DMR new revision | Ovation | PR249304 PR371217
CTUDT00631638 | Refresh | Assigned | lin_s | October 26, 2020 10:59:24 AM GMT+08:00 | FOK - scannable range line on pad doesn’t match table line | FOK |  |  | Subsystem |  | CT-NM.SRS.SysPerf.540/541/542/543/687/567 |  | wang_ada | 0 | 2 |  | Couch | YES | Assigned | DMR new revision | Brilliance 64 | PR249304 PR371217
CTUDT00631638 | Refresh | Assigned | lin_s | October 26, 2020 10:59:24 AM GMT+08:00 | FOK - scannable range line on pad doesn’t match table line | FOK |  |  | Subsystem |  | CT-NM.SRS.SysPerf.540/541/542/543/687/567 |  | wang_ada | 0 | 2 |  | Couch | YES | Assigned | DMR new revision | Ingenuity 64 | PR249304 PR371217
CTUDT00631638 | Refresh | Assigned | lin_s | October 26, 2020 10:59:24 AM GMT+08:00 | FOK - scannable range line on pad doesn’t match table line | FOK |  |  | Subsystem |  | CT-NM.SRS.SysPerf.540/541/542/543/687/567 |  | wang_ada | 0 | 2 |  | Couch | YES | Assigned | DMR new revision | IQon1 | PR249304 PR371217
CTUDT00631636 | Refresh | Accepted | lin_s | October 26, 2020 10:37:38 AM GMT+08:00 | FOK - RFC: Customer wanted a long cross hair on the axials. Explained that does not exist on 4.1.7 | 4.1.10 |  |  | System | FOK; SP2_Score2Postpone |  |  | mukherjee_s | 0 |  |  | Console4 | NO | Accepted |  | Ovation | PR256112
CTUDT00631636 | Refresh | Accepted | lin_s | October 26, 2020 10:37:38 AM GMT+08:00 | FOK - RFC: Customer wanted a long cross hair on the axials. Explained that does not exist on 4.1.7 | 4.1.10 |  |  | System | FOK; SP2_Score2Postpone |  |  | mukherjee_s | 0 |  |  | Console4 | NO | Accepted |  | Brilliance 64 | PR256112
CTUDT00631636 | Refresh | Accepted | lin_s | October 26, 2020 10:37:38 AM GMT+08:00 | FOK - RFC: Customer wanted a long cross hair on the axials. Explained that does not exist on 4.1.7 | 4.1.10 |  |  | System | FOK; SP2_Score2Postpone |  |  | mukherjee_s | 0 |  |  | Console4 | NO | Accepted |  | Brilliance Big Bore | PR256112
CTUDT00631636 | Refresh | Accepted | lin_s | October 26, 2020 10:37:38 AM GMT+08:00 | FOK - RFC: Customer wanted a long cross hair on the axials. Explained that does not exist on 4.1.7 | 4.1.10 |  |  | System | FOK; SP2_Score2Postpone |  |  | mukherjee_s | 0 |  |  | Console4 | NO | Accepted |  | Ingenuity 64 | PR256112
CTUDT00631636 | Refresh | Accepted | lin_s | October 26, 2020 10:37:38 AM GMT+08:00 | FOK - RFC: Customer wanted a long cross hair on the axials. Explained that does not exist on 4.1.7 | 4.1.10 |  |  | System | FOK; SP2_Score2Postpone |  |  | mukherjee_s | 0 |  |  | Console4 | NO | Accepted |  | IQon1 | PR256112
CTUDT00631634 | Refresh | Closed | lin_s | October 26, 2020 10:29:00 AM GMT+08:00 | FOK - Gantry laser cable short circuit or damage can cause Ghost reboot to trigger EStop | FOK |  | Design change has been done for front cover drawings to add the cable tie 453567274081 by E-127321. | Subsystem | FOK | N/A | Forwarded to Different Database | shao_a | 0 |  |  | GCX | YES | Closed |  | Ingenuity 64 | PR257297 PR261207
CTUDT00631631 | Refresh | Closed | hilas | October 26, 2020 1:14:27 AM GMT+08:00 | FOK: Recon got stuck and slow transfer to ISP spectral | 4.7.7 |  | None | Subsystem |  | NA | System Repaired | richards_b | 0 |  |  | CIRS | YES | Closed |  | IQon1 | PR359083
CTUDT00631629 | Refresh | Scored | hilas | October 26, 2020 12:00:24 AM GMT+08:00 | FOK: System stay in init state | 4.7.7 |  | No Workaround. | Subsystem | SP1_PM_Pending | NA |  | janakiraman_j | 0 | 4 |  | Console4 | NO | Scored |  | IQon1 | PR342057
CTUDT00631624 | Refresh | Assigned | guraddi_r | October 23, 2020 9:09:59 PM GMT+08:00 | FoK: Application locking up while doing interventional studies | v4.8 |  |  | System | FoK; SP2_Candidate | NA |  | janakiraman_j | 0 | 3 |  | Console4 | YES | Assigned | SP2 | Ovation | NA
CTUDT00631624 | Refresh | Assigned | guraddi_r | October 23, 2020 9:09:59 PM GMT+08:00 | FoK: Application locking up while doing interventional studies | v4.8 |  |  | System | FoK; SP2_Candidate | NA |  | janakiraman_j | 0 | 3 |  | Console4 | YES | Assigned | SP2 | Brilliance 64 | NA
CTUDT00631624 | Refresh | Assigned | guraddi_r | October 23, 2020 9:09:59 PM GMT+08:00 | FoK: Application locking up while doing interventional studies | v4.8 |  |  | System | FoK; SP2_Candidate | NA |  | janakiraman_j | 0 | 3 |  | Console4 | YES | Assigned | SP2 | Brilliance Big Bore | NA
CTUDT00631624 | Refresh | Assigned | guraddi_r | October 23, 2020 9:09:59 PM GMT+08:00 | FoK: Application locking up while doing interventional studies | v4.8 |  |  | System | FoK; SP2_Candidate | NA |  | janakiraman_j | 0 | 3 |  | Console4 | YES | Assigned | SP2 | Ingenuity 64 | NA
CTUDT00631624 | Refresh | Assigned | guraddi_r | October 23, 2020 9:09:59 PM GMT+08:00 | FoK: Application locking up while doing interventional studies | v4.8 |  |  | System | FoK; SP2_Candidate | NA |  | janakiraman_j | 0 | 3 |  | Console4 | YES | Assigned | SP2 | IQon1 | NA
CTUDT00631614 | Refresh | Postponed | rakic_k | October 22, 2020 10:30:07 PM GMT+08:00 | MsgPipe Watchdog Reliability Issue | WIP30.1 |  |  | Subsystem | SP2 | N/A |  | richards_b | 0 | 3 |  | CIRS | YES | Postponed | SP2 | Ovation
CTUDT00631614 | Refresh | Postponed | rakic_k | October 22, 2020 10:30:07 PM GMT+08:00 | MsgPipe Watchdog Reliability Issue | WIP30.1 |  |  | Subsystem | SP2 | N/A |  | richards_b | 0 | 3 |  | CIRS | YES | Postponed | SP2 | Brilliance 64
CTUDT00631614 | Refresh | Postponed | rakic_k | October 22, 2020 10:30:07 PM GMT+08:00 | MsgPipe Watchdog Reliability Issue | WIP30.1 |  |  | Subsystem | SP2 | N/A |  | richards_b | 0 | 3 |  | CIRS | YES | Postponed | SP2 | Brilliance Big Bore
CTUDT00631614 | Refresh | Postponed | rakic_k | October 22, 2020 10:30:07 PM GMT+08:00 | MsgPipe Watchdog Reliability Issue | WIP30.1 |  |  | Subsystem | SP2 | N/A |  | richards_b | 0 | 3 |  | CIRS | YES | Postponed | SP2 | Ingenuity 64
CTUDT00631614 | Refresh | Postponed | rakic_k | October 22, 2020 10:30:07 PM GMT+08:00 | MsgPipe Watchdog Reliability Issue | WIP30.1 |  |  | Subsystem | SP2 | N/A |  | richards_b | 0 | 3 |  | CIRS | YES | Postponed | SP2 | IQon1
CTUDT00631612 | Refresh | Closed | rakic_k | October 22, 2020 10:21:53 PM GMT+08:00 | Reconstruction failure from ImgNode Creation | WIP30.1 | WIP33 |  | Subsystem | SP1_PM_Approved | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_SuprNodeSvr | CIRS | YES | Closed | WIP33SubCom | Ovation |  | WIP33 | November 4, 2020 3:46:23 AM GMT+08:00
CTUDT00631612 | Refresh | Closed | rakic_k | October 22, 2020 10:21:53 PM GMT+08:00 | Reconstruction failure from ImgNode Creation | WIP30.1 | WIP33 |  | Subsystem | SP1_PM_Approved | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_SuprNodeSvr | CIRS | YES | Closed | WIP33SubCom | Brilliance 64 |  | WIP33 | November 4, 2020 3:46:23 AM GMT+08:00
CTUDT00631612 | Refresh | Closed | rakic_k | October 22, 2020 10:21:53 PM GMT+08:00 | Reconstruction failure from ImgNode Creation | WIP30.1 | WIP33 |  | Subsystem | SP1_PM_Approved | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_SuprNodeSvr | CIRS | YES | Closed | WIP33SubCom | Brilliance Big Bore |  | WIP33 | November 4, 2020 3:46:23 AM GMT+08:00
CTUDT00631612 | Refresh | Closed | rakic_k | October 22, 2020 10:21:53 PM GMT+08:00 | Reconstruction failure from ImgNode Creation | WIP30.1 | WIP33 |  | Subsystem | SP1_PM_Approved | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_SuprNodeSvr | CIRS | YES | Closed | WIP33SubCom | Ingenuity 64 |  | WIP33 | November 4, 2020 3:46:23 AM GMT+08:00
CTUDT00631612 | Refresh | Closed | rakic_k | October 22, 2020 10:21:53 PM GMT+08:00 | Reconstruction failure from ImgNode Creation | WIP30.1 | WIP33 |  | Subsystem | SP1_PM_Approved | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_SuprNodeSvr | CIRS | YES | Closed | WIP33SubCom | IQon1 |  | WIP33 | November 4, 2020 3:46:23 AM GMT+08:00
CTUDT00631609 | Refresh | Closed | allaberdiev_v | October 22, 2020 9:37:17 PM GMT+08:00 | IQon: DHF259320 : Deviation between spec mA parameter and mA in the results | WIP 31 | DHF259320 Rev09 |  | Subsystem | SP1_PM_Approved | N/A | Fixed | richards_b | 0 | 2 | Design Documentation | Support-Physics | NO | Closed | WIP34SubCom | IQon1 |  | DHF259320 Rev09 | November 24, 2020 9:58:06 PM GMT+08:00
CTUDT00631604 | Refresh | Opened | 320053767 | October 22, 2020 2:48:32 PM GMT+08:00 | Grayed out field “number of images” available for user input. | WIP 25.8 |  |  | Subsystem | SP2_Postpone; CRN | Console-SRS.12805 |  | rajalakshmi_j | 0 | 3 |  | Console4 | YES | Opened | SP1-Mar-21-CRN | Ovation
CTUDT00631604 | Refresh | Opened | 320053767 | October 22, 2020 2:48:32 PM GMT+08:00 | Grayed out field “number of images” available for user input. | WIP 25.8 |  |  | Subsystem | SP2_Postpone; CRN | Console-SRS.12805 |  | rajalakshmi_j | 0 | 3 |  | Console4 | YES | Opened | SP1-Mar-21-CRN | Brilliance 64
CTUDT00631604 | Refresh | Opened | 320053767 | October 22, 2020 2:48:32 PM GMT+08:00 | Grayed out field “number of images” available for user input. | WIP 25.8 |  |  | Subsystem | SP2_Postpone; CRN | Console-SRS.12805 |  | rajalakshmi_j | 0 | 3 |  | Console4 | YES | Opened | SP1-Mar-21-CRN | Brilliance Big Bore
CTUDT00631604 | Refresh | Opened | 320053767 | October 22, 2020 2:48:32 PM GMT+08:00 | Grayed out field “number of images” available for user input. | WIP 25.8 |  |  | Subsystem | SP2_Postpone; CRN | Console-SRS.12805 |  | rajalakshmi_j | 0 | 3 |  | Console4 | YES | Opened | SP1-Mar-21-CRN | Ingenuity 64
CTUDT00631604 | Refresh | Opened | 320053767 | October 22, 2020 2:48:32 PM GMT+08:00 | Grayed out field “number of images” available for user input. | WIP 25.8 |  |  | Subsystem | SP2_Postpone; CRN | Console-SRS.12805 |  | rajalakshmi_j | 0 | 3 |  | Console4 | YES | Opened | SP1-Mar-21-CRN | IQon1
CTUDT00631602 | Refresh | Accepted | lin_s | October 22, 2020 1:40:48 PM GMT+08:00 | FOK: Several FOK site reply back unsuccessful sending images to PACS again. | FOK |  |  | Subsystem | FOK; SP1_PM_Approved |  |  | s_sa | 0 |  |  | iCAP | NO | Accepted | WIP37SegReq | Brilliance Big Bore | PR309138 PR226030 PR267191 PR279732 PR354667 PR267626 PR319756 PR308737 PR339755 PR332249 PR341665 PR359739 PR348670 PR354900 PR360185 PR357261 PR356232 PR348693 PR371889 PR378368 PR392810 PR439686 PR466844 PR465836 PR437631 PR457075 PR307590 PR307569 PR460583
CTUDT00631602 | Refresh | Accepted | lin_s | October 22, 2020 1:40:48 PM GMT+08:00 | FOK: Several FOK site reply back unsuccessful sending images to PACS again. | FOK |  |  | Subsystem | FOK; SP1_PM_Approved |  |  | s_sa | 0 |  |  | iCAP | NO | Accepted | WIP37SegReq | Ingenuity 64 | PR309138 PR226030 PR267191 PR279732 PR354667 PR267626 PR319756 PR308737 PR339755 PR332249 PR341665 PR359739 PR348670 PR354900 PR360185 PR357261 PR356232 PR348693 PR371889 PR378368 PR392810 PR439686 PR466844 PR465836 PR437631 PR457075 PR307590 PR307569 PR460583
CTUDT00631600 | Refresh | Resolved | richards_b | October 22, 2020 10:19:56 AM GMT+08:00 | Race condition in Resource Manager using Disk Handler Service can possible cause missing views and degrade image quality | WIP30.1 | WIP33 |  | Subsystem | SP1_PM_Approved | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_DataManager | CIRS | YES | Resolved | WIP33SubCom | IQon1 |  | WIP33 | November 4, 2020 3:46:52 AM GMT+08:00
CTUDT00631598 | Refresh | Resolved | richards_b | October 22, 2020 10:05:02 AM GMT+08:00 | CIRS Refresh: Address Recon Stall Reli issues | WIP30 | WIP33 | User can perform offline recon | Subsystem | SP1_PM_Approved | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_ImgNodeSvr | CIRS | YES | Resolved | WIP33SubCom | Ovation |  | WIP33 | November 4, 2020 3:47:25 AM GMT+08:00
CTUDT00631598 | Refresh | Resolved | richards_b | October 22, 2020 10:05:02 AM GMT+08:00 | CIRS Refresh: Address Recon Stall Reli issues | WIP30 | WIP33 | User can perform offline recon | Subsystem | SP1_PM_Approved | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_ImgNodeSvr | CIRS | YES | Resolved | WIP33SubCom | Brilliance 64 |  | WIP33 | November 4, 2020 3:47:25 AM GMT+08:00
CTUDT00631598 | Refresh | Resolved | richards_b | October 22, 2020 10:05:02 AM GMT+08:00 | CIRS Refresh: Address Recon Stall Reli issues | WIP30 | WIP33 | User can perform offline recon | Subsystem | SP1_PM_Approved | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_ImgNodeSvr | CIRS | YES | Resolved | WIP33SubCom | Brilliance Big Bore |  | WIP33 | November 4, 2020 3:47:25 AM GMT+08:00
CTUDT00631598 | Refresh | Resolved | richards_b | October 22, 2020 10:05:02 AM GMT+08:00 | CIRS Refresh: Address Recon Stall Reli issues | WIP30 | WIP33 | User can perform offline recon | Subsystem | SP1_PM_Approved | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_ImgNodeSvr | CIRS | YES | Resolved | WIP33SubCom | Ingenuity 64 |  | WIP33 | November 4, 2020 3:47:25 AM GMT+08:00
CTUDT00631598 | Refresh | Resolved | richards_b | October 22, 2020 10:05:02 AM GMT+08:00 | CIRS Refresh: Address Recon Stall Reli issues | WIP30 | WIP33 | User can perform offline recon | Subsystem | SP1_PM_Approved | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_ImgNodeSvr | CIRS | YES | Resolved | WIP33SubCom | IQon1 |  | WIP33 | November 4, 2020 3:47:25 AM GMT+08:00
CTUDT00631592 | Refresh | Closed | altshuller_mark | October 21, 2020 8:26:17 PM GMT+08:00 | Hawk Validation (WIP 21): DDC Calibration in O-Level( O-Level user) failed exception is “access denied” to C:\DDC_ON.txt. | Hawk | WIP33 |  | Subsystem | SP1_PM_Approved | n/a | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP33SubCom | IQon1 |  | WIP33 | October 29, 2020 8:07:14 PM GMT+08:00
CTUDT00631585 | Refresh | Assigned | bingxin_lin | October 21, 2020 4:21:32 PM GMT+08:00 | [FOK][BB Refresh][IVC]The wired connection is a tripping hazard when mount on the cart. | RFD | BB_WIP13.4 |  | Subsystem | FOK; SP2_Candidate; HW | 459801623171 | Change Proposal | qin_c | 0 | 4 | CIRS_HW_Projector | GCX | YES | Assigned | HW improvement | Brilliance Big Bore | Cable routing | BB_WIP13.4
CTUDT00631582 | Refresh | Opened | bingxin_lin | October 21, 2020 4:09:55 PM GMT+08:00 | [FOK][BB Refresh][IVC]The IFU is not match the wording printed on the IVC unit | RFD |  | N/A | Subsystem | FOK | N/A |  | rajalakshmi_j | 0 | 3 |  | Support-System | NO | Opened | SP1-Mar-21 | Brilliance Big Bore
CTUDT00631576 | Refresh | Assigned | bingxin_lin | October 21, 2020 3:14:19 PM GMT+08:00 | [FOK][BB Refresh][IVC]Save as Needle Position and Save as Scan Position option displayed as the same at right click menu under French language | RFD |  |  | Subsystem | SP1_PM_Approved ; FOK | NA |  | mukherjee_s | 0 | 2 |  | iCAP | YES | Assigned | WIP37SegReq | Brilliance Big Bore
CTUDT00631575 | Refresh | Resolved | altshuller_mark | October 21, 2020 1:33:48 PM GMT+08:00 | Hawk Performance, Refresh FOK: "Not all images were created" message appeared during Performance Automation | FoK | WIP34 |  | Subsystem | SP1_PM_Approved, FOK | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_Installation | CIRS | YES | Resolved | WIP34SubCom | Ovation | TBD | WIP34 | November 17, 2020 5:19:13 AM GMT+08:00
CTUDT00631575 | Refresh | Resolved | altshuller_mark | October 21, 2020 1:33:48 PM GMT+08:00 | Hawk Performance, Refresh FOK: "Not all images were created" message appeared during Performance Automation | FoK | WIP34 |  | Subsystem | SP1_PM_Approved, FOK | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_Installation | CIRS | YES | Resolved | WIP34SubCom | Brilliance 64 | TBD | WIP34 | November 17, 2020 5:19:13 AM GMT+08:00
CTUDT00631575 | Refresh | Resolved | altshuller_mark | October 21, 2020 1:33:48 PM GMT+08:00 | Hawk Performance, Refresh FOK: "Not all images were created" message appeared during Performance Automation | FoK | WIP34 |  | Subsystem | SP1_PM_Approved, FOK | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_Installation | CIRS | YES | Resolved | WIP34SubCom | Brilliance Big Bore | TBD | WIP34 | November 17, 2020 5:19:13 AM GMT+08:00
CTUDT00631575 | Refresh | Resolved | altshuller_mark | October 21, 2020 1:33:48 PM GMT+08:00 | Hawk Performance, Refresh FOK: "Not all images were created" message appeared during Performance Automation | FoK | WIP34 |  | Subsystem | SP1_PM_Approved, FOK | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_Installation | CIRS | YES | Resolved | WIP34SubCom | Ingenuity 64 | TBD | WIP34 | November 17, 2020 5:19:13 AM GMT+08:00
CTUDT00631575 | Refresh | Resolved | altshuller_mark | October 21, 2020 1:33:48 PM GMT+08:00 | Hawk Performance, Refresh FOK: "Not all images were created" message appeared during Performance Automation | FoK | WIP34 |  | Subsystem | SP1_PM_Approved, FOK | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_Installation | CIRS | YES | Resolved | WIP34SubCom | IQon1 | TBD | WIP34 | November 17, 2020 5:19:13 AM GMT+08:00
CTUDT00631573 | Refresh | Assigned | bingxin_lin | October 21, 2020 12:54:33 PM GMT+08:00 | [FOK][BB Refresh][IVC]Hardware issue : Rocker switch not completely vertical and the cable cover on interface was broken after installation. | RFD | BB_WIP13.4 BB_WIP13.3 BB_WIP13.2 | For rocker switch not vertical: Defective cap is assembled to the rocker switch and withou well check. Suggest to update assembly WI in supplier to check this point and update purchasing spec in Philips to add this vertical requirement. For cable wrapper was broken After investigate in Suzhou Bay and photo check from field, we find the cable fixing ring is very sharp it will very easily to broken the cable wrapper, at the same time the cable bending requirements is high. Proposal A: Add bend relief on cable + Connector fix ring , avoid sharp edge Bend relief boot can release bending force, move sharp edge can solve broken risk. Proposal B: Changed the connector type from vertical into elbow plug (90 degree)+Connector fix ring , avoid sharp edge elbow plug has risk of cable direction will be fixed, can not realize left and right both direction ok. Proposal C: Change raw cable into more flexible cable Risk of this is EMC impact, This raw cable needs signal and power combine together, EMC requirement is high, new type and flexible cable selection is not easy. Normally flexible cable need special materials and process, Color keeps white is not easy, The new selected cable color may be not matches with the controller side cable color. | Subsystem | FOK; SP2_Candidate; HW | 459801623171 | Change Proposal | qin_c | 0 | 2 | CIRS_HW_CIRS_Server | GCX | NO | Assigned | Design Change | Brilliance Big Bore | Drawing | BB_WIP13.4 BB_WIP13.3 BB_WIP13.2
CTUDT00631571 | Refresh | Accepted | bingxin_lin | October 21, 2020 12:34:19 PM GMT+08:00 | [BB Refresh][IVC]Estop opened during pressing the move to select position by joystick mouse | WIP25.8 |  |  | Subsystem | SP2_Candidate |  |  | mukherjee_s | 0 |  |  | Console4 | YES | Accepted | SP2 | Brilliance Big Bore
CTUDT00631569 | Refresh | Resolved | bingxin_lin | October 21, 2020 12:23:20 PM GMT+08:00 | [BB Refresh][IVC]Quickly move the IVC rocker switch to maximum angle 20,couch moving very slowly, Pop up error message box "Patient table cannot comply" | WIP25.8 | WIP35 |  | Subsystem | SP1_PM_Approved | NA | Fixed | 320058084 | 0 | 3 | Console Software | Console4 | NO | Resolved | WIP35SubCom | Brilliance Big Bore |  | WIP35 | November 20, 2020 8:22:55 PM GMT+08:00
CTUDT00631567 | Refresh | Accepted | bingxin_lin | October 21, 2020 12:07:09 PM GMT+08:00 | [BB Refresh][IVC]The connection cannot be established after plug in the network cable if the IVC is started without network cable | WIP25.8 |  |  | Subsystem | SP2_Score2Postpone |  |  | mukherjee_s | 0 |  |  | Console4 | NO | Accepted |  | Brilliance Big Bore
CTUDT00631558 | Refresh | Accepted | bingxin_lin | October 20, 2020 6:18:21 PM GMT+08:00 | [BB Refresh][IVC]The IVC is hung after pull out the network cable during the couch moving and plug in the cable again | WIP25.8 |  |  | Subsystem | SP2_Score2Postpone |  |  | mukherjee_s | 0 |  |  | Console4 | YES | Accepted |  | Brilliance Big Bore
CTUDT00631551 | Refresh | Accepted | ding_f | October 20, 2020 3:40:00 PM GMT+08:00 | Missing images during transferring to ISP when scanning a coronal result with setting ThinSlice and AutoStorage to ISP. | BB_WIP25.8 |  | Manually copy the study. | Subsystem | SP2_Close | NA |  | mukherjee_s | 0 | 4 |  | Console4 | NO | Accepted |  | Ovation
CTUDT00631551 | Refresh | Accepted | ding_f | October 20, 2020 3:40:00 PM GMT+08:00 | Missing images during transferring to ISP when scanning a coronal result with setting ThinSlice and AutoStorage to ISP. | BB_WIP25.8 |  | Manually copy the study. | Subsystem | SP2_Close | NA |  | mukherjee_s | 0 | 4 |  | Console4 | NO | Accepted |  | Brilliance 64
CTUDT00631551 | Refresh | Accepted | ding_f | October 20, 2020 3:40:00 PM GMT+08:00 | Missing images during transferring to ISP when scanning a coronal result with setting ThinSlice and AutoStorage to ISP. | BB_WIP25.8 |  | Manually copy the study. | Subsystem | SP2_Close | NA |  | mukherjee_s | 0 | 4 |  | Console4 | NO | Accepted |  | Brilliance Big Bore
CTUDT00631551 | Refresh | Accepted | ding_f | October 20, 2020 3:40:00 PM GMT+08:00 | Missing images during transferring to ISP when scanning a coronal result with setting ThinSlice and AutoStorage to ISP. | BB_WIP25.8 |  | Manually copy the study. | Subsystem | SP2_Close | NA |  | mukherjee_s | 0 | 4 |  | Console4 | NO | Accepted |  | Ingenuity 64
CTUDT00631551 | Refresh | Accepted | ding_f | October 20, 2020 3:40:00 PM GMT+08:00 | Missing images during transferring to ISP when scanning a coronal result with setting ThinSlice and AutoStorage to ISP. | BB_WIP25.8 |  | Manually copy the study. | Subsystem | SP2_Close | NA |  | mukherjee_s | 0 | 4 |  | Console4 | NO | Accepted |  | IQon1
CTUDT00631548 | Refresh | Assigned | cai_t | October 20, 2020 1:09:05 PM GMT+08:00 | View2 application popups an ErrorViewTwo window when Overwriting view2 by clicking End exam and continue working button | WIP25.8 |  |  | Subsystem | SP2_Candidate | NA |  | janakiraman_j | 0 | 3 |  | Console4 | NO | Assigned | SP2 | Ingenuity 64
CTUDT00631545 | Refresh | Postponed | zhang_t | October 20, 2020 9:48:16 AM GMT+08:00 | [Refresh FOK] Phantom not seen during IQ check | WIP 25.2 |  |  | Subsystem | FOK; | n/a |  | frider_o | 0 | 3 |  | DMS | NO | Postponed | SubNC | Brilliance Big Bore | PR315698
CTUDT00631545 | Refresh | Postponed | zhang_t | October 20, 2020 9:48:16 AM GMT+08:00 | [Refresh FOK] Phantom not seen during IQ check | WIP 25.2 |  |  | Subsystem | FOK; | n/a |  | frider_o | 0 | 3 |  | DMS | NO | Postponed | SubNC | Ingenuity 64 | PR315698
CTUDT00631536 | Refresh | Closed | zhong_x | October 19, 2020 3:21:03 PM GMT+08:00 | “System error” pop up when cancel Edit Results in Slab Viewer of RBS result. | WIP25.8 |  |  | Subsystem | SP2_Postpone | NA | No Plan to Fix | s_srivastava | 0 | 4 |  | Console4 | NO | Closed |  | Ovation
CTUDT00631536 | Refresh | Closed | zhong_x | October 19, 2020 3:21:03 PM GMT+08:00 | “System error” pop up when cancel Edit Results in Slab Viewer of RBS result. | WIP25.8 |  |  | Subsystem | SP2_Postpone | NA | No Plan to Fix | s_srivastava | 0 | 4 |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00631536 | Refresh | Closed | zhong_x | October 19, 2020 3:21:03 PM GMT+08:00 | “System error” pop up when cancel Edit Results in Slab Viewer of RBS result. | WIP25.8 |  |  | Subsystem | SP2_Postpone | NA | No Plan to Fix | s_srivastava | 0 | 4 |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00631536 | Refresh | Closed | zhong_x | October 19, 2020 3:21:03 PM GMT+08:00 | “System error” pop up when cancel Edit Results in Slab Viewer of RBS result. | WIP25.8 |  |  | Subsystem | SP2_Postpone | NA | No Plan to Fix | s_srivastava | 0 | 4 |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00631536 | Refresh | Closed | zhong_x | October 19, 2020 3:21:03 PM GMT+08:00 | “System error” pop up when cancel Edit Results in Slab Viewer of RBS result. | WIP25.8 |  |  | Subsystem | SP2_Postpone | NA | No Plan to Fix | s_srivastava | 0 | 4 |  | Console4 | NO | Closed |  | IQon1
CTUDT00631521 | Refresh | Closed | altshuller_mark | October 18, 2020 4:24:57 PM GMT+08:00 | "end exam" and "log out" buttons grayed out after aborting timed scans | IQon HB1 | WIP34 |  | System | SP1_PM_Approved | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_SeriesProcess | CIRS | YES | Closed | WIP34SubCom | IQon1 | PR165531 | WIP34 | November 17, 2020 5:18:44 AM GMT+08:00
CTUDT00631520 | Refresh | Accepted | altshuller_mark | October 18, 2020 4:20:45 PM GMT+08:00 | Popup "Not all images were created" occurred for a conventional result with Save SBI checked | IQon HB1 |  |  | Subsystem | SP2_Candidate |  |  | periasamy_g | 0 |  |  | Console4 | NO | Accepted |  | IQon1 | PR 9430129 Esc# 0114460042
CTUDT00631519 | Refresh | Accepted | altshuller_mark | October 18, 2020 4:16:27 PM GMT+08:00 | Scanner application closed after Locator scan completed, study had to be reopened, used previous surviews and repeated the Locator, then study planned and completed without error. | IQon HB1 |  |  | Subsystem | InvestigateP1 ; SP2_Candidate | N/A |  | mukherjee_s | 0 |  |  | Console4 | YES | Accepted | SP2 | IQon1 | 8539578
CTUDT00631518 | Refresh | Assigned | altshuller_mark | October 18, 2020 4:12:06 PM GMT+08:00 | Dual Surviews completed but then System Error appeared | IQon HB1 |  | No workaround | Subsystem | SP2_Candidate | N/A |  | modi_s | 0 | 4 |  | Console4 | NO | Assigned | SP2 | Ovation
CTUDT00631518 | Refresh | Assigned | altshuller_mark | October 18, 2020 4:12:06 PM GMT+08:00 | Dual Surviews completed but then System Error appeared | IQon HB1 |  | No workaround | Subsystem | SP2_Candidate | N/A |  | modi_s | 0 | 4 |  | Console4 | NO | Assigned | SP2 | Brilliance 64
CTUDT00631518 | Refresh | Assigned | altshuller_mark | October 18, 2020 4:12:06 PM GMT+08:00 | Dual Surviews completed but then System Error appeared | IQon HB1 |  | No workaround | Subsystem | SP2_Candidate | N/A |  | modi_s | 0 | 4 |  | Console4 | NO | Assigned | SP2 | IQon1
CTUDT00631517 | Refresh | Accepted | altshuller_mark | October 18, 2020 4:08:15 PM GMT+08:00 | Scan start GUI did not displayed | IQon HB1 |  | NO | Subsystem | InvestigateP1; SP2_Candidate | N/A |  | mukherjee_s | 0 | 4 |  | Console4 | NO | Accepted |  | IQon1 | CFR# 9472339
CTUDT00631516 | Refresh | Closed | altshuller_mark | October 18, 2020 4:05:06 PM GMT+08:00 | Bolus tracking fail due to ACQ_NOT_OK after skipping to clinical scan | IQon HB1 |  |  | System |  | N/A | No Plan to Fix | ivanf | 0 | 3 |  | DMS | YES | Closed | SubNC | IQon1 | N/A
CTUDT00631515 | Refresh | Assigned | altshuller_mark | October 18, 2020 4:00:09 PM GMT+08:00 | INQ: [Scan fail due to "Unknown Error" message (S_MASTER_SEQUENCER_GRANT_ERR)] [Osaka Minato Centeral Hospital] [iCT SP] [Escalation Case 0115097780] | Ovation |  |  | System | SP2 | N/A |  | luo_h | 0 | 3 |  | GCX | YES | Assigned | SP2 | Ovation | 0115097780
CTUDT00631515 | Refresh | Assigned | altshuller_mark | October 18, 2020 4:00:09 PM GMT+08:00 | INQ: [Scan fail due to "Unknown Error" message (S_MASTER_SEQUENCER_GRANT_ERR)] [Osaka Minato Centeral Hospital] [iCT SP] [Escalation Case 0115097780] | Ovation |  |  | System | SP2 | N/A |  | luo_h | 0 | 3 |  | GCX | YES | Assigned | SP2 | iCT2 | 0115097780
CTUDT00631514 | Refresh | Accepted | altshuller_mark | October 18, 2020 3:56:38 PM GMT+08:00 | Host slowness is observed over a duration of time | IQon HB1 |  |  | Subsystem | SP2_Candidate |  |  | mukherjee_s | 0 |  |  | Console4 | NO | Accepted | SP2 | IQon1 | PR 10066519
CTUDT00631513 | Refresh | Resolved | altshuller_mark | October 18, 2020 3:48:38 PM GMT+08:00 | No preview image is displayed before the end of a scan. | IQon HB1 | WIP33 | N/A | Subsystem | SP1_PM_Approved | N/A | Fixed | richards_b | 0 | 4 | CIRS_SW_ImgNodeSvr | CIRS | YES | Resolved | WIP33SubCom | Ovation | PR 10006313 escalation 0114358315 | WIP33 | November 3, 2020 6:08:44 AM GMT+08:00
CTUDT00631513 | Refresh | Resolved | altshuller_mark | October 18, 2020 3:48:38 PM GMT+08:00 | No preview image is displayed before the end of a scan. | IQon HB1 | WIP33 | N/A | Subsystem | SP1_PM_Approved | N/A | Fixed | richards_b | 0 | 4 | CIRS_SW_ImgNodeSvr | CIRS | YES | Resolved | WIP33SubCom | Brilliance 64 | PR 10006313 escalation 0114358315 | WIP33 | November 3, 2020 6:08:44 AM GMT+08:00
CTUDT00631513 | Refresh | Resolved | altshuller_mark | October 18, 2020 3:48:38 PM GMT+08:00 | No preview image is displayed before the end of a scan. | IQon HB1 | WIP33 | N/A | Subsystem | SP1_PM_Approved | N/A | Fixed | richards_b | 0 | 4 | CIRS_SW_ImgNodeSvr | CIRS | YES | Resolved | WIP33SubCom | Brilliance Big Bore | PR 10006313 escalation 0114358315 | WIP33 | November 3, 2020 6:08:44 AM GMT+08:00
CTUDT00631513 | Refresh | Resolved | altshuller_mark | October 18, 2020 3:48:38 PM GMT+08:00 | No preview image is displayed before the end of a scan. | IQon HB1 | WIP33 | N/A | Subsystem | SP1_PM_Approved | N/A | Fixed | richards_b | 0 | 4 | CIRS_SW_ImgNodeSvr | CIRS | YES | Resolved | WIP33SubCom | Ingenuity 64 | PR 10006313 escalation 0114358315 | WIP33 | November 3, 2020 6:08:44 AM GMT+08:00
CTUDT00631513 | Refresh | Resolved | altshuller_mark | October 18, 2020 3:48:38 PM GMT+08:00 | No preview image is displayed before the end of a scan. | IQon HB1 | WIP33 | N/A | Subsystem | SP1_PM_Approved | N/A | Fixed | richards_b | 0 | 4 | CIRS_SW_ImgNodeSvr | CIRS | YES | Resolved | WIP33SubCom | IQon1 | PR 10006313 escalation 0114358315 | WIP33 | November 3, 2020 6:08:44 AM GMT+08:00
CTUDT00631510 | Refresh | Closed | dongping_w | October 16, 2020 4:51:17 PM GMT+08:00 | FOK Refresh BB:The main computer console is running very slow | WIP12.6 |  |  | Subsystem | FOK |  | Duplicate | dongping_w | 0 |  |  |  |  | Closed |  | Brilliance Big Bore | 323385
CTUDT00631508 | Refresh | Closed | dongping_w | October 16, 2020 3:19:39 PM GMT+08:00 | FOK Refresh BB:Transfer stuck to PACS need to restart Host | WIP12.6 |  |  | Subsystem | FOK |  | Duplicate | dongping_w | 0 |  |  |  |  | Closed |  | Brilliance Big Bore | 304653
CTUDT00631506 | Refresh | Accepted | lin_s | October 16, 2020 1:48:12 PM GMT+08:00 | FOK RFC: Customer feedback RFC: Customer would like to have a means to move the Tabletop into the exact required position during a biopsy or interventional procedure without having to move the tabletop manually | FOK |  |  | System | FOK; SP2_Score2Postpone |  |  | periasamy_g | 0 |  |  | Console4 | NO | Accepted |  | Ingenuity 64 | 266893
CTUDT00631504 | Refresh | Accepted | lin_s | October 16, 2020 1:43:36 PM GMT+08:00 | FOK RFC: Customer feedback that would like to see the Surview display as it is being acquired in the event they feel they have enough coverage and prevent unnecessary exposure | FOK |  |  | System | FOK; SP2_Score2Postpone |  |  | periasamy_g | 0 |  |  | Console4 | YES | Accepted |  | Ingenuity 64 | 266893
CTUDT00631498 | Refresh | Accepted | s_namburi | October 15, 2020 10:30:02 PM GMT+08:00 | Fok: Sometimes no image for more than 30S after the Surview scan was done | 4.7.7 |  |  | Subsystem | SP2_Candidate |  |  | janakiraman_j | 0 |  |  | Console4 | NO | Accepted | SP2 | IQon1 | PR324710
CTUDT00631496 | Refresh | Assigned | s_namburi | October 15, 2020 10:11:33 PM GMT+08:00 | Fok: Host computer is slow and laggy | 4.7.7 |  |  | Subsystem | SP2_Candidate ; | NA |  | janakiraman_j | 0 | 2 |  | Console4 | NO | Assigned | SP2 | IQon1
CTUDT00631491 | Refresh | Opened | aloni_e | October 15, 2020 4:12:06 PM GMT+08:00 | Misalignment between iCT IFU and TDS regarding spatial resolution | 4.1.7 |  | N/A | Subsystem |  | N/A |  | rajalakshmi_j | 0 | 2 |  | Support-System | NO | Opened | SP1-Mar-21 | Ovation
CTUDT00631474 | Refresh | Closed | lei_yang | October 14, 2020 3:32:17 PM GMT+08:00 | “Slab viewer editing is not available…” pop up when clicking on Edit Results in Slab Viewer. | WIP25.8 |  |  | Subsystem | SP2_Postpone | NA | No Plan to Fix | s_srivastava | 0 | 3 |  | Console4 | NO | Closed |  | Ovation
CTUDT00631474 | Refresh | Closed | lei_yang | October 14, 2020 3:32:17 PM GMT+08:00 | “Slab viewer editing is not available…” pop up when clicking on Edit Results in Slab Viewer. | WIP25.8 |  |  | Subsystem | SP2_Postpone | NA | No Plan to Fix | s_srivastava | 0 | 3 |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00631474 | Refresh | Closed | lei_yang | October 14, 2020 3:32:17 PM GMT+08:00 | “Slab viewer editing is not available…” pop up when clicking on Edit Results in Slab Viewer. | WIP25.8 |  |  | Subsystem | SP2_Postpone | NA | No Plan to Fix | s_srivastava | 0 | 3 |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00631474 | Refresh | Closed | lei_yang | October 14, 2020 3:32:17 PM GMT+08:00 | “Slab viewer editing is not available…” pop up when clicking on Edit Results in Slab Viewer. | WIP25.8 |  |  | Subsystem | SP2_Postpone | NA | No Plan to Fix | s_srivastava | 0 | 3 |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00631474 | Refresh | Closed | lei_yang | October 14, 2020 3:32:17 PM GMT+08:00 | “Slab viewer editing is not available…” pop up when clicking on Edit Results in Slab Viewer. | WIP25.8 |  |  | Subsystem | SP2_Postpone | NA | No Plan to Fix | s_srivastava | 0 | 3 |  | Console4 | NO | Closed |  | IQon1
CTUDT00631469 | Refresh | Closed | dunitza_m | October 14, 2020 3:07:13 AM GMT+08:00 | FOK Lille: Gantry touch panel | v.4.8 |  |  | System | FOK |  | Works As Specified / Not A Defect | dongping_w | 0 |  |  | Support-System | NO | Closed |  | Brilliance Big Bore
CTUDT00631467 | Refresh | Accepted | dunitza_m | October 14, 2020 3:05:00 AM GMT+08:00 | FOK Lille: IVC Rocker Switch reliability | v4.8 |  | Update Curve | System | FOK; SP2_Candidate | Update curve WI |  | qin_c | 0 |  |  | GCX | NO | Accepted |  | Brilliance Big Bore
CTUDT00631467 | Refresh | Accepted | dunitza_m | October 14, 2020 3:05:00 AM GMT+08:00 | FOK Lille: IVC Rocker Switch reliability | v4.8 |  | Update Curve | System | FOK; SP2_Candidate | Update curve WI |  | qin_c | 0 |  |  | GCX | NO | Accepted |  | Hawk
CTUDT00631465 | Refresh | Closed | dunitza_m | October 14, 2020 3:02:44 AM GMT+08:00 | FOK Lille: IVC Rocker Switch performance | v4.8 |  |  | System | FOK; SP1_PM_Approved | NA | Duplicate | 320058084 | 0 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00631465 | Refresh | Closed | dunitza_m | October 14, 2020 3:02:44 AM GMT+08:00 | FOK Lille: IVC Rocker Switch performance | v4.8 |  |  | System | FOK; SP1_PM_Approved | NA | Duplicate | 320058084 | 0 |  |  | Console4 | NO | Closed |  | Hawk
CTUDT00631463 | Refresh | Closed | dunitza_m | October 14, 2020 3:00:37 AM GMT+08:00 | FOK Lille: SW IVC right click menu localization | v.4.8 |  |  | System | FOK |  | Duplicate | mukherjee_s | 0 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00631463 | Refresh | Closed | dunitza_m | October 14, 2020 3:00:37 AM GMT+08:00 | FOK Lille: SW IVC right click menu localization | v.4.8 |  |  | System | FOK |  | Duplicate | mukherjee_s | 0 |  |  | Console4 | NO | Closed |  | Hawk
CTUDT00631461 | Refresh | Opened | dunitza_m | October 14, 2020 2:57:53 AM GMT+08:00 | FOK MCVI: IFU discrepancy | BB v4.8 |  | N/A | System | FOK; Doc_Update | N/A |  | rajalakshmi_j | 0 | 3 |  | Support-System | NO | Opened | SP1-Mar-21 | Brilliance Big Bore
CTUDT00631459 | Refresh | Closed | dunitza_m | October 14, 2020 2:53:50 AM GMT+08:00 | FOK MCVI: IVC Cart footprint is too big | v4.8 |  | A: Fix on ceiling directly--Don’t occupy ground space and is easy to move. Risk 1. Site doesn't support ceiling mounting 2. Interference with current structure (interference with Monitors) 3. unstable during operation because the IVC is suspended 4. Suzhou CT bay ceiling structure needs to be modified 5.Cable needs to be hidden in the wall. Wall needs to be modified. Bay and Site doesn‘t support this modification. It’s also difficult to replace cable for service when cable is hidden in wall 6. Cable routing needs to be hidden in the tube and this needs to be customized and prototype delivery takes long time than expected Efforts:30 weeks B Fix on steel structure--Don’t occupy ground space and is easy to move. 1.. Site doesn't support steel mounting (Wall strength and space ) 2. Interference with current structure (interference with Monitors) 3. unstable during operation because the IVC is suspended 4. Suzhou CT bay wall structure needs to be modified 5.Cable needs to be hidden in the wall. Wall needs to be modified. Bay and Site doesn't support this modification. It's difficult to replace cable for service in this situation 6. Cable routing needs to be hidden in the tube and this needs to be customized and prototype delivery takes long time than expected | System | FOK | 459801630321 | Change Proposal | qin_c | 0 |  |  | GCX | NO | Closed |  | Brilliance Big Bore | DMR347478
CTUDT00631459 | Refresh | Closed | dunitza_m | October 14, 2020 2:53:50 AM GMT+08:00 | FOK MCVI: IVC Cart footprint is too big | v4.8 |  | A: Fix on ceiling directly--Don’t occupy ground space and is easy to move. Risk 1. Site doesn't support ceiling mounting 2. Interference with current structure (interference with Monitors) 3. unstable during operation because the IVC is suspended 4. Suzhou CT bay ceiling structure needs to be modified 5.Cable needs to be hidden in the wall. Wall needs to be modified. Bay and Site doesn‘t support this modification. It’s also difficult to replace cable for service when cable is hidden in wall 6. Cable routing needs to be hidden in the tube and this needs to be customized and prototype delivery takes long time than expected Efforts:30 weeks B Fix on steel structure--Don’t occupy ground space and is easy to move. 1.. Site doesn't support steel mounting (Wall strength and space ) 2. Interference with current structure (interference with Monitors) 3. unstable during operation because the IVC is suspended 4. Suzhou CT bay wall structure needs to be modified 5.Cable needs to be hidden in the wall. Wall needs to be modified. Bay and Site doesn't support this modification. It's difficult to replace cable for service in this situation 6. Cable routing needs to be hidden in the tube and this needs to be customized and prototype delivery takes long time than expected | System | FOK | 459801630321 | Change Proposal | qin_c | 0 |  |  | GCX | NO | Closed |  | Hawk | DMR347478
CTUDT00631457 | Refresh | Assigned | dunitza_m | October 14, 2020 2:49:03 AM GMT+08:00 | FOK MCVI US and Lille France: IVC Joystick mouse performance | Big Bore v4.8 |  | There is IVC Curve Update Instruction could already fix this issue | System | FOK; | CTUDT00631457 |  | qin_c | 0 | 2 |  | GCX | NO | Assigned | SP2 | Brilliance Big Bore | MOUSE
CTUDT00631457 | Refresh | Assigned | dunitza_m | October 14, 2020 2:49:03 AM GMT+08:00 | FOK MCVI US and Lille France: IVC Joystick mouse performance | Big Bore v4.8 |  | There is IVC Curve Update Instruction could already fix this issue | System | FOK; | CTUDT00631457 |  | qin_c | 0 | 2 |  | GCX | NO | Assigned | SP2 | Hawk | MOUSE
CTUDT00631455 | Refresh | Postponed | dunitza_m | October 14, 2020 2:33:48 AM GMT+08:00 | FOK MCVI: IVC icons for Scan and Work position confusion | Refresh BB v4.8 |  |  | System | FOK | N/A |  | qin_c | 0 | 3 |  | GCX | NO | Postponed |  | Brilliance Big Bore | n/a
CTUDT00631449 | Refresh | Closed | bingxin_lin | October 13, 2020 6:15:29 PM GMT+08:00 | "There is an error in FilmView application..." message popped up after logout from service user and login with CT user | WIP25.8 |  | Relaunch of workflow or logout and login | Subsystem |  | NA | Duplicate | s_sa | 0 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00631449 | Refresh | Closed | bingxin_lin | October 13, 2020 6:15:29 PM GMT+08:00 | "There is an error in FilmView application..." message popped up after logout from service user and login with CT user | WIP25.8 |  | Relaunch of workflow or logout and login | Subsystem |  | NA | Duplicate | s_sa | 0 |  |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00631449 | Refresh | Closed | bingxin_lin | October 13, 2020 6:15:29 PM GMT+08:00 | "There is an error in FilmView application..." message popped up after logout from service user and login with CT user | WIP25.8 |  | Relaunch of workflow or logout and login | Subsystem |  | NA | Duplicate | s_sa | 0 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00631449 | Refresh | Closed | bingxin_lin | October 13, 2020 6:15:29 PM GMT+08:00 | "There is an error in FilmView application..." message popped up after logout from service user and login with CT user | WIP25.8 |  | Relaunch of workflow or logout and login | Subsystem |  | NA | Duplicate | s_sa | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00631449 | Refresh | Closed | bingxin_lin | October 13, 2020 6:15:29 PM GMT+08:00 | "There is an error in FilmView application..." message popped up after logout from service user and login with CT user | WIP25.8 |  | Relaunch of workflow or logout and login | Subsystem |  | NA | Duplicate | s_sa | 0 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00631447 | Refresh | Accepted | wang_l | October 13, 2020 3:50:22 PM GMT+08:00 | Refresh - Scan barcode cannot retrieve patient data for the first time when end one active exam | WIP32 |  |  | Subsystem | SP2_Candidate |  |  | mukherjee_s | 0 |  |  | Console4 | NO | Accepted | SP2 | Brilliance Big Bore
CTUDT00631447 | Refresh | Accepted | wang_l | October 13, 2020 3:50:22 PM GMT+08:00 | Refresh - Scan barcode cannot retrieve patient data for the first time when end one active exam | WIP32 |  |  | Subsystem | SP2_Candidate |  |  | mukherjee_s | 0 |  |  | Console4 | NO | Accepted | SP2 | Ingenuity 64
CTUDT00631437 | Refresh | Resolved | altshuller_mark | October 13, 2020 12:32:19 PM GMT+08:00 | Hawk Validation (WIP 19.5): 55% phase in Step and Shoot EC | Hawk | WIP36 |  | Subsystem | SP1_PM_Approved | N/A | Fixed | sannihith_r | 0 | 2 | Console Generic | Console4 | NO | Resolved | WIP36SubCom | Ovation |  | WIP36 | December 1, 2020 3:44:05 PM GMT+08:00
CTUDT00631437 | Refresh | Resolved | altshuller_mark | October 13, 2020 12:32:19 PM GMT+08:00 | Hawk Validation (WIP 19.5): 55% phase in Step and Shoot EC | Hawk | WIP36 |  | Subsystem | SP1_PM_Approved | N/A | Fixed | sannihith_r | 0 | 2 | Console Generic | Console4 | NO | Resolved | WIP36SubCom | Brilliance 64 |  | WIP36 | December 1, 2020 3:44:05 PM GMT+08:00
CTUDT00631437 | Refresh | Resolved | altshuller_mark | October 13, 2020 12:32:19 PM GMT+08:00 | Hawk Validation (WIP 19.5): 55% phase in Step and Shoot EC | Hawk | WIP36 |  | Subsystem | SP1_PM_Approved | N/A | Fixed | sannihith_r | 0 | 2 | Console Generic | Console4 | NO | Resolved | WIP36SubCom | Ingenuity 64 |  | WIP36 | December 1, 2020 3:44:05 PM GMT+08:00
CTUDT00631437 | Refresh | Resolved | altshuller_mark | October 13, 2020 12:32:19 PM GMT+08:00 | Hawk Validation (WIP 19.5): 55% phase in Step and Shoot EC | Hawk | WIP36 |  | Subsystem | SP1_PM_Approved | N/A | Fixed | sannihith_r | 0 | 2 | Console Generic | Console4 | NO | Resolved | WIP36SubCom | IQon1 |  | WIP36 | December 1, 2020 3:44:05 PM GMT+08:00
CTUDT00631435 | Refresh | Accepted | zhong_x | October 13, 2020 11:46:03 AM GMT+08:00 | IVC HW Control got stuck during perform non-interventional Test. | WIP25.8 |  |  | Subsystem | SP2_Candidate |  |  | mukherjee_s | 0 |  |  | Console4 | NO | Accepted |  | Brilliance Big Bore
CTUDT00631431 | Refresh | Opened | bingxin_lin | October 12, 2020 11:00:43 AM GMT+08:00 | [BB Refresh][IVC]The couch or gantry are keep moving when slide on the move up/down button or tilt button on IVC box and release it | WIP25.8 |  |  | Subsystem | SP1_PM_Approved | NA |  | rajalakshmi_j | 0 | 4 |  | Console4 | YES | Opened | SP1-Mar-21-CRN | Brilliance Big Bore
CTUDT00631412 | Refresh | Closed | levi_a | October 8, 2020 6:34:32 PM GMT+08:00 | Refresh - iCT-SP - Missing BugRep Directories | WIP32 | WIP33 | N/A | Subsystem | SP1_PM_Approved ; Blocker | STPS-SSRS-1102 | Fixed | venkatachalam_p | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP33SubCom | Ovation |  | WIP33 | October 27, 2020 11:16:16 PM GMT+08:00
CTUDT00631409 | Refresh | Closed | yatziv_l | October 8, 2020 2:36:12 PM GMT+08:00 | Refresh SP1: CTDI: The 16th "AirCTDI_15 Pediatric Body" of "CTDI Free Air" test failed | WIP31 | WIP33 | N/A | Subsystem | SP1_PM_Approved | N/A | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP33SubCom | Ovation |  | WIP33 | October 29, 2020 8:10:09 PM GMT+08:00
CTUDT00631409 | Refresh | Closed | yatziv_l | October 8, 2020 2:36:12 PM GMT+08:00 | Refresh SP1: CTDI: The 16th "AirCTDI_15 Pediatric Body" of "CTDI Free Air" test failed | WIP31 | WIP33 | N/A | Subsystem | SP1_PM_Approved | N/A | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP33SubCom | Brilliance 64 |  | WIP33 | October 29, 2020 8:10:09 PM GMT+08:00
CTUDT00631409 | Refresh | Closed | yatziv_l | October 8, 2020 2:36:12 PM GMT+08:00 | Refresh SP1: CTDI: The 16th "AirCTDI_15 Pediatric Body" of "CTDI Free Air" test failed | WIP31 | WIP33 | N/A | Subsystem | SP1_PM_Approved | N/A | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP33SubCom | Brilliance Big Bore |  | WIP33 | October 29, 2020 8:10:09 PM GMT+08:00
CTUDT00631409 | Refresh | Closed | yatziv_l | October 8, 2020 2:36:12 PM GMT+08:00 | Refresh SP1: CTDI: The 16th "AirCTDI_15 Pediatric Body" of "CTDI Free Air" test failed | WIP31 | WIP33 | N/A | Subsystem | SP1_PM_Approved | N/A | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP33SubCom | Ingenuity 64 |  | WIP33 | October 29, 2020 8:10:09 PM GMT+08:00
CTUDT00631409 | Refresh | Closed | yatziv_l | October 8, 2020 2:36:12 PM GMT+08:00 | Refresh SP1: CTDI: The 16th "AirCTDI_15 Pediatric Body" of "CTDI Free Air" test failed | WIP31 | WIP33 | N/A | Subsystem | SP1_PM_Approved | N/A | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP33SubCom | IQon1 |  | WIP33 | October 29, 2020 8:10:09 PM GMT+08:00
CTUDT00631401 | Refresh | Accepted | hilas | October 6, 2020 3:03:03 AM GMT+08:00 | FoK: The IQon at PCH experienced a BSOD (blue screen of death) | 4.7.7 installed with Hotfix 1 |  | User has to reboot the system. | System | FOK; SP2_Score2Postpone | NA |  | janakiraman_j | 0 |  |  | Console4 | NO | Accepted |  | IQon1 | PR296882
CTUDT00631397 | Refresh | Closed | zhong_x | October 5, 2020 3:26:19 PM GMT+08:00 | Nonstandard Installation for WIP25.8 | WIP25.8 |  |  | Subsystem | InvestigateP1 | None | No Plan to Fix | veeraswamy_k | 0 | 4 |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00631393 | Refresh | Opened | zhong_x | October 5, 2020 3:21:55 PM GMT+08:00 | "We can't sign into your account...message popped up during the clinical user login" when User Name contains special sign as % | WIP25.8 |  |  | Subsystem | CRN; Test Design | NA |  | rajalakshmi_j | 0 | 3 |  | Support-System | NO | Opened | SP1-Mar-21-CRN | Brilliance Big Bore
CTUDT00631378 | Refresh | Closed | prabhan_p | October 1, 2020 12:24:01 AM GMT+08:00 | When we loaded result in View to slab viewer Not all images were reconstructed image popup is coming and Save icon is also disappearing. | WIP25.8 |  | Edit the result on Surview instead of Slab Viewer. | Subsystem | InvestigateP1 | NA | No Plan to Fix | janakiraman_j | 0 | 4 |  | Console4 | NO | Closed |  | IQon1
CTUDT00631374 | Refresh | Assigned | bagdady_b | September 30, 2020 4:28:51 PM GMT+08:00 | Refresh- Review before save result takes more than 4 minutes to be enabled for edit | WIP25.5 |  |  | Subsystem | SP2_Postpone | NA |  | 320057847 | 0 | 2 |  | Console4 | NO | Assigned | SP2 | IQon1
CTUDT00631369 | Refresh | Closed | altshuller_mark | September 30, 2020 12:33:33 PM GMT+08:00 | [HAWK]-Dots artifact in QA Axial Head 3D and Brain Perfusion EC | Hawk | WIP33 |  | Subsystem | SP1_PM_Approved | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_ImgNodeSvr | CIRS | YES | Closed | WIP33SubCom | Ovation |  | WIP33 | November 3, 2020 6:08:55 AM GMT+08:00
CTUDT00631367 | Refresh | Closed | lin_s | September 30, 2020 9:40:09 AM GMT+08:00 | FOK: Yellow screen issue happens when use workflow | 4.1.10 | WIP34 |  | Subsystem | FOK; SP1_PM_Approved | N/A | Duplicate | s_srivastava | 0 | 2 | Console USP_EBW | Console4 | NO | Closed | WIP34SubCom | Ovation | PR293042 PR295524 | WIP34 | November 4, 2020 2:50:40 PM GMT+08:00
CTUDT00631367 | Refresh | Closed | lin_s | September 30, 2020 9:40:09 AM GMT+08:00 | FOK: Yellow screen issue happens when use workflow | 4.1.10 | WIP34 |  | Subsystem | FOK; SP1_PM_Approved | N/A | Duplicate | s_srivastava | 0 | 2 | Console USP_EBW | Console4 | NO | Closed | WIP34SubCom | Brilliance 64 | PR293042 PR295524 | WIP34 | November 4, 2020 2:50:40 PM GMT+08:00
CTUDT00631367 | Refresh | Closed | lin_s | September 30, 2020 9:40:09 AM GMT+08:00 | FOK: Yellow screen issue happens when use workflow | 4.1.10 | WIP34 |  | Subsystem | FOK; SP1_PM_Approved | N/A | Duplicate | s_srivastava | 0 | 2 | Console USP_EBW | Console4 | NO | Closed | WIP34SubCom | Brilliance Big Bore | PR293042 PR295524 | WIP34 | November 4, 2020 2:50:40 PM GMT+08:00
CTUDT00631367 | Refresh | Closed | lin_s | September 30, 2020 9:40:09 AM GMT+08:00 | FOK: Yellow screen issue happens when use workflow | 4.1.10 | WIP34 |  | Subsystem | FOK; SP1_PM_Approved | N/A | Duplicate | s_srivastava | 0 | 2 | Console USP_EBW | Console4 | NO | Closed | WIP34SubCom | Ingenuity 64 | PR293042 PR295524 | WIP34 | November 4, 2020 2:50:40 PM GMT+08:00
CTUDT00631367 | Refresh | Closed | lin_s | September 30, 2020 9:40:09 AM GMT+08:00 | FOK: Yellow screen issue happens when use workflow | 4.1.10 | WIP34 |  | Subsystem | FOK; SP1_PM_Approved | N/A | Duplicate | s_srivastava | 0 | 2 | Console USP_EBW | Console4 | NO | Closed | WIP34SubCom | IQon1 | PR293042 PR295524 | WIP34 | November 4, 2020 2:50:40 PM GMT+08:00
CTUDT00631355 | Refresh | Closed | elavarasan_s | September 28, 2020 9:25:53 PM GMT+08:00 | MPR result with non primary accession number is not loading to slab viewer. | BB_WIP25.7 |  |  | Subsystem | SP2_Postpone | NA | No Plan to Fix | s_srivastava | 0 | 3 |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00631349 | Refresh | Closed | lei_yang | September 25, 2020 5:04:11 PM GMT+08:00 | “The specified path is invalid” is always displayed when unplug the USB drive and click the “Export” button again. | WIP25.6 |  |  | Subsystem | SP2_Score2Postpone | N/A | No Plan to Fix | guraddi_r | 0 | 4 |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00631349 | Refresh | Closed | lei_yang | September 25, 2020 5:04:11 PM GMT+08:00 | “The specified path is invalid” is always displayed when unplug the USB drive and click the “Export” button again. | WIP25.6 |  |  | Subsystem | SP2_Score2Postpone | N/A | No Plan to Fix | guraddi_r | 0 | 4 |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00631347 | Refresh | Closed | lei_yang | September 25, 2020 4:59:07 PM GMT+08:00 | “ExamCardManager has stopped working” pop up when the USB drive is pulled out and continue to export EC. | WIP25.6 |  | Do not remove USB If user want to export exam card to USB. | Subsystem |  | None | No Plan to Fix | veeraswamy_k | 0 | 3 |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00631347 | Refresh | Closed | lei_yang | September 25, 2020 4:59:07 PM GMT+08:00 | “ExamCardManager has stopped working” pop up when the USB drive is pulled out and continue to export EC. | WIP25.6 |  | Do not remove USB If user want to export exam card to USB. | Subsystem |  | None | No Plan to Fix | veeraswamy_k | 0 | 3 |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00631340 | Refresh | Opened | lin_s | September 24, 2020 5:14:56 PM GMT+08:00 | Not able to login with Chinese credentials | V4.7.7 |  | No WorkAround | Subsystem | SP1_PM_Approved | STPS-SSRS-1201 |  | rajalakshmi_j | 0 | 2 |  | Support-System | NO | Opened | SP1-Mar-21-CRN | Ovation
CTUDT00631340 | Refresh | Opened | lin_s | September 24, 2020 5:14:56 PM GMT+08:00 | Not able to login with Chinese credentials | V4.7.7 |  | No WorkAround | Subsystem | SP1_PM_Approved | STPS-SSRS-1201 |  | rajalakshmi_j | 0 | 2 |  | Support-System | NO | Opened | SP1-Mar-21-CRN | Brilliance 64
CTUDT00631340 | Refresh | Opened | lin_s | September 24, 2020 5:14:56 PM GMT+08:00 | Not able to login with Chinese credentials | V4.7.7 |  | No WorkAround | Subsystem | SP1_PM_Approved | STPS-SSRS-1201 |  | rajalakshmi_j | 0 | 2 |  | Support-System | NO | Opened | SP1-Mar-21-CRN | Brilliance Big Bore
CTUDT00631340 | Refresh | Opened | lin_s | September 24, 2020 5:14:56 PM GMT+08:00 | Not able to login with Chinese credentials | V4.7.7 |  | No WorkAround | Subsystem | SP1_PM_Approved | STPS-SSRS-1201 |  | rajalakshmi_j | 0 | 2 |  | Support-System | NO | Opened | SP1-Mar-21-CRN | Ingenuity 64
CTUDT00631340 | Refresh | Opened | lin_s | September 24, 2020 5:14:56 PM GMT+08:00 | Not able to login with Chinese credentials | V4.7.7 |  | No WorkAround | Subsystem | SP1_PM_Approved | STPS-SSRS-1201 |  | rajalakshmi_j | 0 | 2 |  | Support-System | NO | Opened | SP1-Mar-21-CRN | IQon1
CTUDT00631337 | Refresh | Scored | cai_t | September 24, 2020 1:44:57 PM GMT+08:00 | [Refresh] :It occurs a system error while using Snagit to get screenshot - PmsRemoting service crash | WIP25.5 |  | logout and re-login as CT user \perform system restart. | Subsystem | PM_Pending | NA |  | s_sa | 0 | 3 |  | iCAP | NO | Scored |  | Ovation
CTUDT00631337 | Refresh | Scored | cai_t | September 24, 2020 1:44:57 PM GMT+08:00 | [Refresh] :It occurs a system error while using Snagit to get screenshot - PmsRemoting service crash | WIP25.5 |  | logout and re-login as CT user \perform system restart. | Subsystem | PM_Pending | NA |  | s_sa | 0 | 3 |  | iCAP | NO | Scored |  | Brilliance 64
CTUDT00631337 | Refresh | Scored | cai_t | September 24, 2020 1:44:57 PM GMT+08:00 | [Refresh] :It occurs a system error while using Snagit to get screenshot - PmsRemoting service crash | WIP25.5 |  | logout and re-login as CT user \perform system restart. | Subsystem | PM_Pending | NA |  | s_sa | 0 | 3 |  | iCAP | NO | Scored |  | Brilliance Big Bore
CTUDT00631337 | Refresh | Scored | cai_t | September 24, 2020 1:44:57 PM GMT+08:00 | [Refresh] :It occurs a system error while using Snagit to get screenshot - PmsRemoting service crash | WIP25.5 |  | logout and re-login as CT user \perform system restart. | Subsystem | PM_Pending | NA |  | s_sa | 0 | 3 |  | iCAP | NO | Scored |  | Ingenuity 64
CTUDT00631337 | Refresh | Scored | cai_t | September 24, 2020 1:44:57 PM GMT+08:00 | [Refresh] :It occurs a system error while using Snagit to get screenshot - PmsRemoting service crash | WIP25.5 |  | logout and re-login as CT user \perform system restart. | Subsystem | PM_Pending | NA |  | s_sa | 0 | 3 |  | iCAP | NO | Scored |  | IQon1
CTUDT00631334 | Refresh | Closed | dongping_w | September 24, 2020 12:42:52 PM GMT+08:00 | FOK Refresh BB: the doctor want to have dicom info : performin physician | WIP12.6 |  |  | Subsystem | FOK | RFC | Change Proposal | guraddi_r | 0 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore | 290264
CTUDT00631332 | Refresh | Closed | dongping_w | September 24, 2020 12:38:42 PM GMT+08:00 | FOK Refresh BB: mAS is not normal when there is a mental part in the planned area | WIP12.6 |  |  | Subsystem | FOK |  | Works As Specified / Not A Defect | dongping_w | 0 |  |  | Support-System | NO | Closed | com-24-oct-2020 | Brilliance Big Bore | 290246
CTUDT00631331 | Refresh | Assigned | ding_f | September 24, 2020 12:36:08 PM GMT+08:00 | The time taken to View2 exceeds 15s when clicked "End Exam and Continue working". | WIP25.5 |  |  | Subsystem | SP2_Postpone; Req_update | NA |  | tiwari_g | 0 | 2 |  | Console4 | NO | Assigned | SP2_Doc | Ovation
CTUDT00631331 | Refresh | Assigned | ding_f | September 24, 2020 12:36:08 PM GMT+08:00 | The time taken to View2 exceeds 15s when clicked "End Exam and Continue working". | WIP25.5 |  |  | Subsystem | SP2_Postpone; Req_update | NA |  | tiwari_g | 0 | 2 |  | Console4 | NO | Assigned | SP2_Doc | Brilliance 64
CTUDT00631331 | Refresh | Assigned | ding_f | September 24, 2020 12:36:08 PM GMT+08:00 | The time taken to View2 exceeds 15s when clicked "End Exam and Continue working". | WIP25.5 |  |  | Subsystem | SP2_Postpone; Req_update | NA |  | tiwari_g | 0 | 2 |  | Console4 | NO | Assigned | SP2_Doc | Brilliance Big Bore
CTUDT00631331 | Refresh | Assigned | ding_f | September 24, 2020 12:36:08 PM GMT+08:00 | The time taken to View2 exceeds 15s when clicked "End Exam and Continue working". | WIP25.5 |  |  | Subsystem | SP2_Postpone; Req_update | NA |  | tiwari_g | 0 | 2 |  | Console4 | NO | Assigned | SP2_Doc | Ingenuity 64
CTUDT00631331 | Refresh | Assigned | ding_f | September 24, 2020 12:36:08 PM GMT+08:00 | The time taken to View2 exceeds 15s when clicked "End Exam and Continue working". | WIP25.5 |  |  | Subsystem | SP2_Postpone; Req_update | NA |  | tiwari_g | 0 | 2 |  | Console4 | NO | Assigned | SP2_Doc | IQon1
CTUDT00631331 | Refresh | Assigned | ding_f | September 24, 2020 12:36:08 PM GMT+08:00 | The time taken to View2 exceeds 15s when clicked "End Exam and Continue working". | WIP25.5 |  |  | Subsystem | SP2_Postpone; Req_update | NA |  | tiwari_g | 0 | 2 |  | Console4 | NO | Assigned | SP2_Doc | Hawk
CTUDT00631329 | Refresh | Closed | dongping_w | September 24, 2020 12:32:20 PM GMT+08:00 | FOK Refresh BB: Dissatisfaction the avalaible pitchs in V4.8 and with the option 32, are the same than in V2.3, 3.5, 4.2 : very limited. | WIP12.6 |  |  | Subsystem | FOK | N/A | Change Proposal | dongping_w | 0 |  |  | Support-System | NO | Closed |  | Brilliance Big Bore | 290028
CTUDT00631327 | Refresh | Closed | dongping_w | September 24, 2020 12:26:42 PM GMT+08:00 | FOK Refresh BB:Exam Card Manager: Dissatisfaction news in version Window 10 worst than before; when an ECM we select exam cards it is less visible than before and not easy for users color of the selected lines when exam cards are selected. | WIP12.6 |  |  | Subsystem | FOK, LCM_Candidate | NA | Change Proposal | guraddi_r | 0 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore | 290005
CTUDT00631325 | Refresh | Closed | dongping_w | September 24, 2020 11:24:05 AM GMT+08:00 | FOK Refresh BB:Dissatisfaction news in version Window 10 worst than before; translation review as "revoir" in French | WIP12.6 | WIP34 | NA | System | FOK, SP1_PM_Approved | NA | Fixed | srushti_k | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP34SubCom | Ovation | 290003 | WIP34 | November 11, 2020 3:43:03 PM GMT+08:00
CTUDT00631325 | Refresh | Closed | dongping_w | September 24, 2020 11:24:05 AM GMT+08:00 | FOK Refresh BB:Dissatisfaction news in version Window 10 worst than before; translation review as "revoir" in French | WIP12.6 | WIP34 | NA | System | FOK, SP1_PM_Approved | NA | Fixed | srushti_k | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP34SubCom | Brilliance 64 | 290003 | WIP34 | November 11, 2020 3:43:03 PM GMT+08:00
CTUDT00631325 | Refresh | Closed | dongping_w | September 24, 2020 11:24:05 AM GMT+08:00 | FOK Refresh BB:Dissatisfaction news in version Window 10 worst than before; translation review as "revoir" in French | WIP12.6 | WIP34 | NA | System | FOK, SP1_PM_Approved | NA | Fixed | srushti_k | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP34SubCom | Brilliance Big Bore | 290003 | WIP34 | November 11, 2020 3:43:03 PM GMT+08:00
CTUDT00631325 | Refresh | Closed | dongping_w | September 24, 2020 11:24:05 AM GMT+08:00 | FOK Refresh BB:Dissatisfaction news in version Window 10 worst than before; translation review as "revoir" in French | WIP12.6 | WIP34 | NA | System | FOK, SP1_PM_Approved | NA | Fixed | srushti_k | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP34SubCom | Ingenuity 64 | 290003 | WIP34 | November 11, 2020 3:43:03 PM GMT+08:00
CTUDT00631325 | Refresh | Closed | dongping_w | September 24, 2020 11:24:05 AM GMT+08:00 | FOK Refresh BB:Dissatisfaction news in version Window 10 worst than before; translation review as "revoir" in French | WIP12.6 | WIP34 | NA | System | FOK, SP1_PM_Approved | NA | Fixed | srushti_k | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP34SubCom | IQon1 | 290003 | WIP34 | November 11, 2020 3:43:03 PM GMT+08:00
CTUDT00631321 | Refresh | Investigated | lei_yang | September 24, 2020 9:06:57 AM GMT+08:00 | The host computer suddenly appears blue screen while editing an EC in ECM. | BB_WIP25.5 |  |  | Subsystem | InvestigateP2 | NA |  | tan_j | 1 |  |  | GCX | NO | Investigated |  | Brilliance Big Bore
CTUDT00631318 | Refresh | Closed | 320085947 | September 23, 2020 9:16:33 PM GMT+08:00 | When changing the language from Chinese to Dutch (or any other language) and performing a few cycles of login and logout the Chinese and Dutch are both appearing in the login screen. | WIP25.5 |  | User can login to philips_service and change the Keyboard language from chinese to non-English. | Subsystem |  | NA | No Plan to Fix | s_sa | 0 | 3 |  | iCAP | NO | Closed |  | Ovation
CTUDT00631318 | Refresh | Closed | 320085947 | September 23, 2020 9:16:33 PM GMT+08:00 | When changing the language from Chinese to Dutch (or any other language) and performing a few cycles of login and logout the Chinese and Dutch are both appearing in the login screen. | WIP25.5 |  | User can login to philips_service and change the Keyboard language from chinese to non-English. | Subsystem |  | NA | No Plan to Fix | s_sa | 0 | 3 |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00631318 | Refresh | Closed | 320085947 | September 23, 2020 9:16:33 PM GMT+08:00 | When changing the language from Chinese to Dutch (or any other language) and performing a few cycles of login and logout the Chinese and Dutch are both appearing in the login screen. | WIP25.5 |  | User can login to philips_service and change the Keyboard language from chinese to non-English. | Subsystem |  | NA | No Plan to Fix | s_sa | 0 | 3 |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00631318 | Refresh | Closed | 320085947 | September 23, 2020 9:16:33 PM GMT+08:00 | When changing the language from Chinese to Dutch (or any other language) and performing a few cycles of login and logout the Chinese and Dutch are both appearing in the login screen. | WIP25.5 |  | User can login to philips_service and change the Keyboard language from chinese to non-English. | Subsystem |  | NA | No Plan to Fix | s_sa | 0 | 3 |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00631318 | Refresh | Closed | 320085947 | September 23, 2020 9:16:33 PM GMT+08:00 | When changing the language from Chinese to Dutch (or any other language) and performing a few cycles of login and logout the Chinese and Dutch are both appearing in the login screen. | WIP25.5 |  | User can login to philips_service and change the Keyboard language from chinese to non-English. | Subsystem |  | NA | No Plan to Fix | s_sa | 0 | 3 |  | iCAP | NO | Closed |  | IQon1
CTUDT00631316 | Refresh | Closed | 320085947 | September 23, 2020 8:52:12 PM GMT+08:00 | Loading an EC to plan scene takes too long. | WIP25.5 | DHF298554 Rev17 DHF298555 Rev13 DHF298556 Rev.17 DHF298553 Rev.18 |  | System | SP1-PM_Approved | CT-NM.SRS.SysPerf.131 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | SP1 | Ovation |  | DHF298554 Rev17 DHF298555 Rev13 DHF298556 Rev.17 DHF298553 Rev.18 | October 23, 2020 8:48:14 AM GMT+08:00
CTUDT00631316 | Refresh | Closed | 320085947 | September 23, 2020 8:52:12 PM GMT+08:00 | Loading an EC to plan scene takes too long. | WIP25.5 | DHF298554 Rev17 DHF298555 Rev13 DHF298556 Rev.17 DHF298553 Rev.18 |  | System | SP1-PM_Approved | CT-NM.SRS.SysPerf.131 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | SP1 | Brilliance 64 |  | DHF298554 Rev17 DHF298555 Rev13 DHF298556 Rev.17 DHF298553 Rev.18 | October 23, 2020 8:48:14 AM GMT+08:00
CTUDT00631316 | Refresh | Closed | 320085947 | September 23, 2020 8:52:12 PM GMT+08:00 | Loading an EC to plan scene takes too long. | WIP25.5 | DHF298554 Rev17 DHF298555 Rev13 DHF298556 Rev.17 DHF298553 Rev.18 |  | System | SP1-PM_Approved | CT-NM.SRS.SysPerf.131 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | SP1 | Brilliance Big Bore |  | DHF298554 Rev17 DHF298555 Rev13 DHF298556 Rev.17 DHF298553 Rev.18 | October 23, 2020 8:48:14 AM GMT+08:00
CTUDT00631316 | Refresh | Closed | 320085947 | September 23, 2020 8:52:12 PM GMT+08:00 | Loading an EC to plan scene takes too long. | WIP25.5 | DHF298554 Rev17 DHF298555 Rev13 DHF298556 Rev.17 DHF298553 Rev.18 |  | System | SP1-PM_Approved | CT-NM.SRS.SysPerf.131 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | SP1 | Ingenuity 64 |  | DHF298554 Rev17 DHF298555 Rev13 DHF298556 Rev.17 DHF298553 Rev.18 | October 23, 2020 8:48:14 AM GMT+08:00
CTUDT00631316 | Refresh | Closed | 320085947 | September 23, 2020 8:52:12 PM GMT+08:00 | Loading an EC to plan scene takes too long. | WIP25.5 | DHF298554 Rev17 DHF298555 Rev13 DHF298556 Rev.17 DHF298553 Rev.18 |  | System | SP1-PM_Approved | CT-NM.SRS.SysPerf.131 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | SP1 | IQon1 |  | DHF298554 Rev17 DHF298555 Rev13 DHF298556 Rev.17 DHF298553 Rev.18 | October 23, 2020 8:48:14 AM GMT+08:00
CTUDT00631312 | Refresh | Closed | gT_s | September 23, 2020 4:58:46 PM GMT+08:00 | Refresh : Results assigned with different accession numbers were shown under one accession number | WIP25.5 |  | Use Change Patient Details on the series level | Subsystem | PM_SP2_candidate, CRN | NA | No Plan to Fix | s_srivastava | 0 | 3 |  | Console4 | NO | Closed |  | IQon1
CTUDT00631309 | Refresh | Closed | zhong_x | September 23, 2020 3:18:02 PM GMT+08:00 | System Error occurred when trying to edit result in Slab Viewer for ThinSlice results with different procedures. | WIP25.5 | WIP25.6 |  | Subsystem | PM_Approved | NA | Duplicate | s_srivastava | 0 | 2 | Console Software | Console | NO | Closed | NextWIP | Brilliance Big Bore |  | WIP25.6
CTUDT00631304 | Refresh | Opened | vinoliya_j | September 23, 2020 1:01:51 PM GMT+08:00 | Refresh : Information about Short tube conditioning should be corrected in different user documents | WIP30.1 |  | N/A | Subsystem |  | N/A |  | rajalakshmi_j | 0 | 2 |  | Support-System | NO | Opened | SP1-Mar-21-CRN | Brilliance 64 | NA
CTUDT00631304 | Refresh | Opened | vinoliya_j | September 23, 2020 1:01:51 PM GMT+08:00 | Refresh : Information about Short tube conditioning should be corrected in different user documents | WIP30.1 |  | N/A | Subsystem |  | N/A |  | rajalakshmi_j | 0 | 2 |  | Support-System | NO | Opened | SP1-Mar-21-CRN | Brilliance Big Bore | NA
CTUDT00631304 | Refresh | Opened | vinoliya_j | September 23, 2020 1:01:51 PM GMT+08:00 | Refresh : Information about Short tube conditioning should be corrected in different user documents | WIP30.1 |  | N/A | Subsystem |  | N/A |  | rajalakshmi_j | 0 | 2 |  | Support-System | NO | Opened | SP1-Mar-21-CRN | iCT4 | NA
CTUDT00631304 | Refresh | Opened | vinoliya_j | September 23, 2020 1:01:51 PM GMT+08:00 | Refresh : Information about Short tube conditioning should be corrected in different user documents | WIP30.1 |  | N/A | Subsystem |  | N/A |  | rajalakshmi_j | 0 | 2 |  | Support-System | NO | Opened | SP1-Mar-21-CRN | Ingenuity 64 | NA
CTUDT00631304 | Refresh | Opened | vinoliya_j | September 23, 2020 1:01:51 PM GMT+08:00 | Refresh : Information about Short tube conditioning should be corrected in different user documents | WIP30.1 |  | N/A | Subsystem |  | N/A |  | rajalakshmi_j | 0 | 2 |  | Support-System | NO | Opened | SP1-Mar-21-CRN | IQon1 | NA
CTUDT00631295 | Refresh | Assigned | herip_yapoujian | September 22, 2020 7:52:16 PM GMT+08:00 | Rerfresh WIP25.5: Slow response and delay in lateral surview image appearance. | WIP25.5 |  |  | Subsystem | SP1_PM_Approved | N/A |  | atin_s | 0 | 2 |  | Console4 | YES | Assigned | WIP37SegReq | Ovation
CTUDT00631295 | Refresh | Assigned | herip_yapoujian | September 22, 2020 7:52:16 PM GMT+08:00 | Rerfresh WIP25.5: Slow response and delay in lateral surview image appearance. | WIP25.5 |  |  | Subsystem | SP1_PM_Approved | N/A |  | atin_s | 0 | 2 |  | Console4 | YES | Assigned | WIP37SegReq | Brilliance 64
CTUDT00631295 | Refresh | Assigned | herip_yapoujian | September 22, 2020 7:52:16 PM GMT+08:00 | Rerfresh WIP25.5: Slow response and delay in lateral surview image appearance. | WIP25.5 |  |  | Subsystem | SP1_PM_Approved | N/A |  | atin_s | 0 | 2 |  | Console4 | YES | Assigned | WIP37SegReq | Brilliance Big Bore
CTUDT00631295 | Refresh | Assigned | herip_yapoujian | September 22, 2020 7:52:16 PM GMT+08:00 | Rerfresh WIP25.5: Slow response and delay in lateral surview image appearance. | WIP25.5 |  |  | Subsystem | SP1_PM_Approved | N/A |  | atin_s | 0 | 2 |  | Console4 | YES | Assigned | WIP37SegReq | Ingenuity 64
CTUDT00631295 | Refresh | Assigned | herip_yapoujian | September 22, 2020 7:52:16 PM GMT+08:00 | Rerfresh WIP25.5: Slow response and delay in lateral surview image appearance. | WIP25.5 |  |  | Subsystem | SP1_PM_Approved | N/A |  | atin_s | 0 | 2 |  | Console4 | YES | Assigned | WIP37SegReq | IQon1
CTUDT00631293 | Refresh | Closed | herip_yapoujian | September 22, 2020 7:34:07 PM GMT+08:00 | Rerfresh WIP25.5: Exam application is not responding message while uploading EC | WIP25.5 |  |  | Subsystem |  | NA | No Plan to Fix | modi_s | 0 | 4 |  | Console4 | NO | Closed |  | IQon1
CTUDT00631291 | Refresh | Resolved | herip_yapoujian | September 22, 2020 5:44:25 PM GMT+08:00 | Rerfresh WIP25.5: CIRS cannot initialize | WIP 25.5 | WIP35 | User can hit go again | Subsystem | SP1_PM_Approved | N/A | Fixed | richards_b | 0 | 3 | CIRS_SW_Data_Management | CIRS | YES | Resolved | WIP35SubCom | Ovation |  | WIP35 | December 5, 2020 6:00:54 AM GMT+08:00
CTUDT00631291 | Refresh | Resolved | herip_yapoujian | September 22, 2020 5:44:25 PM GMT+08:00 | Rerfresh WIP25.5: CIRS cannot initialize | WIP 25.5 | WIP35 | User can hit go again | Subsystem | SP1_PM_Approved | N/A | Fixed | richards_b | 0 | 3 | CIRS_SW_Data_Management | CIRS | YES | Resolved | WIP35SubCom | Brilliance 64 |  | WIP35 | December 5, 2020 6:00:54 AM GMT+08:00
CTUDT00631291 | Refresh | Resolved | herip_yapoujian | September 22, 2020 5:44:25 PM GMT+08:00 | Rerfresh WIP25.5: CIRS cannot initialize | WIP 25.5 | WIP35 | User can hit go again | Subsystem | SP1_PM_Approved | N/A | Fixed | richards_b | 0 | 3 | CIRS_SW_Data_Management | CIRS | YES | Resolved | WIP35SubCom | Brilliance Big Bore |  | WIP35 | December 5, 2020 6:00:54 AM GMT+08:00
CTUDT00631291 | Refresh | Resolved | herip_yapoujian | September 22, 2020 5:44:25 PM GMT+08:00 | Rerfresh WIP25.5: CIRS cannot initialize | WIP 25.5 | WIP35 | User can hit go again | Subsystem | SP1_PM_Approved | N/A | Fixed | richards_b | 0 | 3 | CIRS_SW_Data_Management | CIRS | YES | Resolved | WIP35SubCom | Ingenuity 64 |  | WIP35 | December 5, 2020 6:00:54 AM GMT+08:00
CTUDT00631291 | Refresh | Resolved | herip_yapoujian | September 22, 2020 5:44:25 PM GMT+08:00 | Rerfresh WIP25.5: CIRS cannot initialize | WIP 25.5 | WIP35 | User can hit go again | Subsystem | SP1_PM_Approved | N/A | Fixed | richards_b | 0 | 3 | CIRS_SW_Data_Management | CIRS | YES | Resolved | WIP35SubCom | IQon1 |  | WIP35 | December 5, 2020 6:00:54 AM GMT+08:00
CTUDT00631285 | Refresh | NotAccepted | prabhan_p | September 21, 2020 11:31:41 PM GMT+08:00 | In End Exam & Continue scenario sometimes some results are marked completed(Save icon displayed) while the recon is still progress in RQM. | Wip25.5 |  |  | Subsystem |  |  |  |  | 0 |  |  |  |  | NotAccepted |  | Brilliance Big Bore
CTUDT00631283 | Refresh | Closed | lin_s | September 21, 2020 9:31:32 PM GMT+08:00 | FOK: Several FOK site reply back remote device retrieve or transfer failure(Work list populating issue) | 4.1.10 |  |  | Subsystem | FOK; InvestigateP1 | none | Works As Specified / Not A Defect | veeraswamy_k | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64 | PR189748 PR307590 PR309138
CTUDT00631275 | Refresh | Accepted | bingxin_lin | September 21, 2020 7:24:06 PM GMT+08:00 | [BB Refresh][IVC]The IVC lights show confusion after the first installation | WIP25.2 |  |  | Subsystem | SP2_Score2Postpone | STPS-SSRS-1351 |  | 320021835 | 1 |  |  | Console4 | YES | Accepted |  | Brilliance Big Bore
CTUDT00631274 | Refresh | Closed | altshuller_mark | September 21, 2020 6:34:59 PM GMT+08:00 | In split study scenario, Surview of second and third study is not transferred to configured remote device and Transfer icon is not displayed in Patient Directory. | WIP25.5 | WIP32 |  | Subsystem | SP1, SP1_PM_Approved | N/A | Fixed | challa_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP32SubCom | Ovation |  | WIP32 | October 5, 2020 2:01:56 PM GMT+08:00
CTUDT00631274 | Refresh | Closed | altshuller_mark | September 21, 2020 6:34:59 PM GMT+08:00 | In split study scenario, Surview of second and third study is not transferred to configured remote device and Transfer icon is not displayed in Patient Directory. | WIP25.5 | WIP32 |  | Subsystem | SP1, SP1_PM_Approved | N/A | Fixed | challa_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP32SubCom | Brilliance 64 |  | WIP32 | October 5, 2020 2:01:56 PM GMT+08:00
CTUDT00631274 | Refresh | Closed | altshuller_mark | September 21, 2020 6:34:59 PM GMT+08:00 | In split study scenario, Surview of second and third study is not transferred to configured remote device and Transfer icon is not displayed in Patient Directory. | WIP25.5 | WIP32 |  | Subsystem | SP1, SP1_PM_Approved | N/A | Fixed | challa_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP32SubCom | Brilliance Big Bore |  | WIP32 | October 5, 2020 2:01:56 PM GMT+08:00
CTUDT00631274 | Refresh | Closed | altshuller_mark | September 21, 2020 6:34:59 PM GMT+08:00 | In split study scenario, Surview of second and third study is not transferred to configured remote device and Transfer icon is not displayed in Patient Directory. | WIP25.5 | WIP32 |  | Subsystem | SP1, SP1_PM_Approved | N/A | Fixed | challa_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP32SubCom | Ingenuity 64 |  | WIP32 | October 5, 2020 2:01:56 PM GMT+08:00
CTUDT00631274 | Refresh | Closed | altshuller_mark | September 21, 2020 6:34:59 PM GMT+08:00 | In split study scenario, Surview of second and third study is not transferred to configured remote device and Transfer icon is not displayed in Patient Directory. | WIP25.5 | WIP32 |  | Subsystem | SP1, SP1_PM_Approved | N/A | Fixed | challa_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP32SubCom | IQon1 |  | WIP32 | October 5, 2020 2:01:56 PM GMT+08:00
CTUDT00631273 | Refresh | Closed | altshuller_mark | September 21, 2020 6:25:27 PM GMT+08:00 | result of scan is not selected by default on view2 | HotFix2 | WIP32 |  | Subsystem | SP1, SP1_PM_Approved | N/A | Fixed | s_srivastava | 0 | 2 | Console Generic | Console4 | NO | Closed | WIP32SubCom | Ovation |  | WIP32 | October 6, 2020 12:18:00 AM GMT+08:00
CTUDT00631273 | Refresh | Closed | altshuller_mark | September 21, 2020 6:25:27 PM GMT+08:00 | result of scan is not selected by default on view2 | HotFix2 | WIP32 |  | Subsystem | SP1, SP1_PM_Approved | N/A | Fixed | s_srivastava | 0 | 2 | Console Generic | Console4 | NO | Closed | WIP32SubCom | Brilliance 64 |  | WIP32 | October 6, 2020 12:18:00 AM GMT+08:00
CTUDT00631273 | Refresh | Closed | altshuller_mark | September 21, 2020 6:25:27 PM GMT+08:00 | result of scan is not selected by default on view2 | HotFix2 | WIP32 |  | Subsystem | SP1, SP1_PM_Approved | N/A | Fixed | s_srivastava | 0 | 2 | Console Generic | Console4 | NO | Closed | WIP32SubCom | Brilliance Big Bore |  | WIP32 | October 6, 2020 12:18:00 AM GMT+08:00
CTUDT00631273 | Refresh | Closed | altshuller_mark | September 21, 2020 6:25:27 PM GMT+08:00 | result of scan is not selected by default on view2 | HotFix2 | WIP32 |  | Subsystem | SP1, SP1_PM_Approved | N/A | Fixed | s_srivastava | 0 | 2 | Console Generic | Console4 | NO | Closed | WIP32SubCom | Ingenuity 64 |  | WIP32 | October 6, 2020 12:18:00 AM GMT+08:00
CTUDT00631273 | Refresh | Closed | altshuller_mark | September 21, 2020 6:25:27 PM GMT+08:00 | result of scan is not selected by default on view2 | HotFix2 | WIP32 |  | Subsystem | SP1, SP1_PM_Approved | N/A | Fixed | s_srivastava | 0 | 2 | Console Generic | Console4 | NO | Closed | WIP32SubCom | IQon1 |  | WIP32 | October 6, 2020 12:18:00 AM GMT+08:00
CTUDT00631268 | Refresh | Closed | dongping_w | September 21, 2020 5:14:32 PM GMT+08:00 | FOK Refresh BB: engineer's arc description at the pipe level resulted in estop cut-off during lubrication testing at some resolutions | WIP12.6 |  |  | Subsystem | FOK |  | System Repaired | tan_j | 0 |  |  | Support-System | NO | Closed |  | Brilliance Big Bore | 277763
CTUDT00631266 | Refresh | Closed | dongping_w | September 21, 2020 5:08:18 PM GMT+08:00 | FOK Refresh BB: Once import into ARIA/Eclipse, the ROI has a Hounsfield Unit override applied. | WIP12.6 |  |  | Subsystem | FOK |  | Does Not Apply | dongping_w | 0 |  |  |  |  | Closed |  | Brilliance Big Bore | 280446
CTUDT00631263 | Refresh | Closed | dongping_w | September 21, 2020 4:54:36 PM GMT+08:00 | FOK Refresh BB: The customer wants to be able to import into Aria the Exam Summary (CT dose record), as was previously possible | WIP12.6 |  |  | Subsystem | FOK |  | Not Enough Information | dongping_w | 0 |  |  |  |  | Closed |  | Brilliance Big Bore | 280440
CTUDT00631260 | Refresh | Closed | dongping_w | September 21, 2020 4:45:58 PM GMT+08:00 | FOK Refresh BB: Changement APM | WIP12.6 |  |  | Subsystem | FOK |  | Not Enough Information | dongping_w | 0 |  |  |  |  | Closed |  | Brilliance Big Bore | 286822
CTUDT00631257 | Refresh | Closed | dongping_w | September 21, 2020 4:36:50 PM GMT+08:00 | FOK Refresh BB: RGSC does not have wave form on CT scanner | FOK |  |  | Subsystem | FOK |  | Not Enough Information | dongping_w | 0 |  |  |  |  | Closed |  | Brilliance Big Bore | 286681
CTUDT00631255 | Refresh | Assigned | dongping_w | September 21, 2020 4:26:50 PM GMT+08:00 | FOK Refresh BB: The system does not label the image or dataset when IMR is used | FOK |  |  | Subsystem | FOK, SP2_Candidate | N/A |  | 320057847 | 0 | 4 |  | Console4 | NO | Assigned | SP2 | Ovation | 285040
CTUDT00631255 | Refresh | Assigned | dongping_w | September 21, 2020 4:26:50 PM GMT+08:00 | FOK Refresh BB: The system does not label the image or dataset when IMR is used | FOK |  |  | Subsystem | FOK, SP2_Candidate | N/A |  | 320057847 | 0 | 4 |  | Console4 | NO | Assigned | SP2 | Brilliance Big Bore | 285040
CTUDT00631255 | Refresh | Assigned | dongping_w | September 21, 2020 4:26:50 PM GMT+08:00 | FOK Refresh BB: The system does not label the image or dataset when IMR is used | FOK |  |  | Subsystem | FOK, SP2_Candidate | N/A |  | 320057847 | 0 | 4 |  | Console4 | NO | Assigned | SP2 | Ingenuity 64 | 285040
CTUDT00631255 | Refresh | Assigned | dongping_w | September 21, 2020 4:26:50 PM GMT+08:00 | FOK Refresh BB: The system does not label the image or dataset when IMR is used | FOK |  |  | Subsystem | FOK, SP2_Candidate | N/A |  | 320057847 | 0 | 4 |  | Console4 | NO | Assigned | SP2 | IQon1 | 285040
CTUDT00631252 | Refresh | Assigned | lV_d | September 21, 2020 1:27:42 PM GMT+08:00 | When building or editing a Step & Shoot CTA Exam Card, the user increment can't be 50% of the slice thickness | 4.1.7 |  |  | System | FOK;SP2_Postpone | N/A |  | atin_s | 0 | 2 |  | Console4 | NO | Assigned | SP2 | Ingenuity 64 | 169992
CTUDT00631251 | Refresh | Resolved | altshuller_mark | September 21, 2020 12:39:05 PM GMT+08:00 | Visual diagnostic CIRS Recon SWBP test is faling | IQON GPU 4.7.5.43533 | WIP34 | Documentation for operations and service would need to be updated to not select this test when testing on a single rack. If the other tests are passing, it is an indication that the recon server is working as expected. This test was just provided to show test coverage for 2D reconstructions. | Subsystem | SP1, SP1_PM_Approved | N/A | Fixed | richards_b | 0 | 2 | Service Documentation | CIRS | NO | Resolved | WIP34SubCom | IQon1 |  | WIP34 | November 17, 2020 5:37:33 AM GMT+08:00
CTUDT00631247 | Refresh | Assigned | zhong_x | September 20, 2020 1:39:06 PM GMT+08:00 | After transfer coordinate from Console to LAP Laser System, "The deformation of the CT image is not allowed" message pop up at LAP System. | BB_WIP25.2 |  | None | Subsystem | SP1_PM_Approved | NA |  | Snehashis_Dey | 0 | 2 |  | Console4 | NO | Assigned | WIP37SegReq | Brilliance Big Bore
CTUDT00631243 | Refresh | Closed | bingxin_lin | September 18, 2020 8:26:47 PM GMT+08:00 | [BB Refresh][IVC]Quickly move the IVC rocker switch to the left or right limit angle, the couch can't move at 150mm/s high speed | WIP25.2 |  |  | Subsystem | InvestigateP1 |  | Duplicate | 320021835 | 0 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00631241 | Refresh | Resolved | bingxin_lin | September 18, 2020 11:04:13 AM GMT+08:00 | [BB Refresh][IVC]Need to update IVC SW about the FW files | WIP25.5 | WIP35 |  | Subsystem | SP1_PM_Approved | Console-SRS.16018 | Fixed | guraddi_r | 0 | 2 | IVC Software | Console4 | NO | Resolved | WIP35SubCom | Brilliance Big Bore |  | WIP35 | November 18, 2020 3:10:09 PM GMT+08:00
CTUDT00631239 | Refresh | Closed | bingxin_lin | September 18, 2020 10:25:34 AM GMT+08:00 | [BB Refresh][IVC]"IvHwControlsApplication has stopped working "message popup after logout from workflow and login again. | WIP25.2 | WIP35 |  | System | SP1_PM_Approved | Console-SRS.16462 | Fixed | guraddi_r | 0 | 4 | IVC Software | Console4 | NO | Closed | WIP35SubCom | Brilliance Big Bore |  | WIP35 | November 29, 2020 3:52:37 PM GMT+08:00
CTUDT00631237 | Refresh | Opened | bingxin_lin | September 18, 2020 10:03:28 AM GMT+08:00 | [BB Refresh][IVC]"Getting Windows ready Don't turn off your computer"message displayed during the IVC shutdown | WIP25.2 |  |  | Subsystem | CRN | Console-SRS.16547 |  | rajalakshmi_j | 0 | 4 |  | Support-System | NO | Opened | SP1-Mar-21-CRN | Brilliance Big Bore
CTUDT00631198 | Refresh | Closed | krisher_i | September 18, 2020 1:14:51 AM GMT+08:00 | High Priority Result is not prioritized in View2 | WIP25.5 |  |  | Subsystem |  | N/A | No Plan to Fix | challa_s | 0 | 4 |  | Console4 | NO | Closed |  | Ovation
CTUDT00631198 | Refresh | Closed | krisher_i | September 18, 2020 1:14:51 AM GMT+08:00 | High Priority Result is not prioritized in View2 | WIP25.5 |  |  | Subsystem |  | N/A | No Plan to Fix | challa_s | 0 | 4 |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00631198 | Refresh | Closed | krisher_i | September 18, 2020 1:14:51 AM GMT+08:00 | High Priority Result is not prioritized in View2 | WIP25.5 |  |  | Subsystem |  | N/A | No Plan to Fix | challa_s | 0 | 4 |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00631198 | Refresh | Closed | krisher_i | September 18, 2020 1:14:51 AM GMT+08:00 | High Priority Result is not prioritized in View2 | WIP25.5 |  |  | Subsystem |  | N/A | No Plan to Fix | challa_s | 0 | 4 |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00631198 | Refresh | Closed | krisher_i | September 18, 2020 1:14:51 AM GMT+08:00 | High Priority Result is not prioritized in View2 | WIP25.5 |  |  | Subsystem |  | N/A | No Plan to Fix | challa_s | 0 | 4 |  | Console4 | NO | Closed |  | IQon1
CTUDT00631184 | Refresh | Assigned | bagdady_b | September 17, 2020 7:41:30 PM GMT+08:00 | Refresh - RBS results are not enabled for edit, Incomplete data in view2 | WIP25.5 |  |  | System | SP2_Postpone | NA |  | challa_s | 0 | 2 |  | Console4 | NO | Assigned | SP2 | Ovation
CTUDT00631184 | Refresh | Assigned | bagdady_b | September 17, 2020 7:41:30 PM GMT+08:00 | Refresh - RBS results are not enabled for edit, Incomplete data in view2 | WIP25.5 |  |  | System | SP2_Postpone | NA |  | challa_s | 0 | 2 |  | Console4 | NO | Assigned | SP2 | Brilliance 64
CTUDT00631184 | Refresh | Assigned | bagdady_b | September 17, 2020 7:41:30 PM GMT+08:00 | Refresh - RBS results are not enabled for edit, Incomplete data in view2 | WIP25.5 |  |  | System | SP2_Postpone | NA |  | challa_s | 0 | 2 |  | Console4 | NO | Assigned | SP2 | Brilliance Big Bore
CTUDT00631184 | Refresh | Assigned | bagdady_b | September 17, 2020 7:41:30 PM GMT+08:00 | Refresh - RBS results are not enabled for edit, Incomplete data in view2 | WIP25.5 |  |  | System | SP2_Postpone | NA |  | challa_s | 0 | 2 |  | Console4 | NO | Assigned | SP2 | Ingenuity 64
CTUDT00631184 | Refresh | Assigned | bagdady_b | September 17, 2020 7:41:30 PM GMT+08:00 | Refresh - RBS results are not enabled for edit, Incomplete data in view2 | WIP25.5 |  |  | System | SP2_Postpone | NA |  | challa_s | 0 | 2 |  | Console4 | NO | Assigned | SP2 | IQon1
CTUDT00631180 | Refresh | Closed | bagdady_b | September 17, 2020 6:36:14 PM GMT+08:00 | Refresh- Transfer icon doesn't appear even though the Series was transferred to remote devices | WIP25.5 |  | User can perform manual copy for the respective study | Subsystem |  | NA | Duplicate | s_sa | 1 |  |  | iCAP | NO | Closed |  | Ovation
CTUDT00631180 | Refresh | Closed | bagdady_b | September 17, 2020 6:36:14 PM GMT+08:00 | Refresh- Transfer icon doesn't appear even though the Series was transferred to remote devices | WIP25.5 |  | User can perform manual copy for the respective study | Subsystem |  | NA | Duplicate | s_sa | 1 |  |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00631180 | Refresh | Closed | bagdady_b | September 17, 2020 6:36:14 PM GMT+08:00 | Refresh- Transfer icon doesn't appear even though the Series was transferred to remote devices | WIP25.5 |  | User can perform manual copy for the respective study | Subsystem |  | NA | Duplicate | s_sa | 1 |  |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00631180 | Refresh | Closed | bagdady_b | September 17, 2020 6:36:14 PM GMT+08:00 | Refresh- Transfer icon doesn't appear even though the Series was transferred to remote devices | WIP25.5 |  | User can perform manual copy for the respective study | Subsystem |  | NA | Duplicate | s_sa | 1 |  |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00631180 | Refresh | Closed | bagdady_b | September 17, 2020 6:36:14 PM GMT+08:00 | Refresh- Transfer icon doesn't appear even though the Series was transferred to remote devices | WIP25.5 |  | User can perform manual copy for the respective study | Subsystem |  | NA | Duplicate | s_sa | 1 |  |  | iCAP | NO | Closed |  | IQon1
CTUDT00631173 | Refresh | Closed | lei_yang | September 17, 2020 9:15:48 AM GMT+08:00 | Install WIP25.2 and WIP25.5 patch failed. | WIP25.2/WIP25.5 | Hotfix2 | user should start re-installation and update to Hot fix1 and hot fix2 | Subsystem |  | None | Fixed | hilas | 0 | 2 | Service Documentation | Support-System | NO | Closed | installation spec new version | Ovation |  | Hotfix2 | October 29, 2020 12:56:25 AM GMT+08:00
CTUDT00631173 | Refresh | Closed | lei_yang | September 17, 2020 9:15:48 AM GMT+08:00 | Install WIP25.2 and WIP25.5 patch failed. | WIP25.2/WIP25.5 | Hotfix2 | user should start re-installation and update to Hot fix1 and hot fix2 | Subsystem |  | None | Fixed | hilas | 0 | 2 | Service Documentation | Support-System | NO | Closed | installation spec new version | Brilliance 64 |  | Hotfix2 | October 29, 2020 12:56:25 AM GMT+08:00
CTUDT00631173 | Refresh | Closed | lei_yang | September 17, 2020 9:15:48 AM GMT+08:00 | Install WIP25.2 and WIP25.5 patch failed. | WIP25.2/WIP25.5 | Hotfix2 | user should start re-installation and update to Hot fix1 and hot fix2 | Subsystem |  | None | Fixed | hilas | 0 | 2 | Service Documentation | Support-System | NO | Closed | installation spec new version | Brilliance Big Bore |  | Hotfix2 | October 29, 2020 12:56:25 AM GMT+08:00
CTUDT00631173 | Refresh | Closed | lei_yang | September 17, 2020 9:15:48 AM GMT+08:00 | Install WIP25.2 and WIP25.5 patch failed. | WIP25.2/WIP25.5 | Hotfix2 | user should start re-installation and update to Hot fix1 and hot fix2 | Subsystem |  | None | Fixed | hilas | 0 | 2 | Service Documentation | Support-System | NO | Closed | installation spec new version | Ingenuity 64 |  | Hotfix2 | October 29, 2020 12:56:25 AM GMT+08:00
CTUDT00631173 | Refresh | Closed | lei_yang | September 17, 2020 9:15:48 AM GMT+08:00 | Install WIP25.2 and WIP25.5 patch failed. | WIP25.2/WIP25.5 | Hotfix2 | user should start re-installation and update to Hot fix1 and hot fix2 | Subsystem |  | None | Fixed | hilas | 0 | 2 | Service Documentation | Support-System | NO | Closed | installation spec new version | IQon1 |  | Hotfix2 | October 29, 2020 12:56:25 AM GMT+08:00
CTUDT00631165 | Refresh | Closed | lin_s | September 16, 2020 5:43:13 PM GMT+08:00 | FOK: Customer Feedback Willis Knighton EStop on Couch Zero. | FOK |  |  | Subsystem | FOK | NA | System Repaired | yaping_l | 0 |  |  | GCX | NO | Closed |  | Ingenuity 64 | PR278069
CTUDT00631164 | Refresh | Closed | prabhan_p | September 16, 2020 5:42:22 PM GMT+08:00 | Right monitor turns yellow during offline recon. | wip 25.5 |  |  | Subsystem |  |  | Duplicate | tiwari_g | 0 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00631160 | Refresh | Resolved | bingxin_lin | September 16, 2020 3:13:52 PM GMT+08:00 | The IVC power on button turned to off after IVC finished started | WIP25.2 | WIP35 |  | Subsystem | SP1_PM_Approved | Console-SRS.16547, Console-SRS.16549 | Fixed | guraddi_r | 0 | 2 | IVC Software | Console4 | NO | Resolved | WIP35SubCom | Brilliance Big Bore |  | WIP35 | November 29, 2020 3:53:10 PM GMT+08:00
CTUDT00631156 | Refresh | Closed | lin_s | September 16, 2020 2:18:52 PM GMT+08:00 | FOK: Couch horizontal movement issue happens. | 4.1.10 |  |  | Subsystem | FOK; SP1_PM_Approved | NA | System Repaired | cx_zhang | 0 |  |  | Couch | NO | Closed | WIP33SubCom | Ingenuity 64 | PR264536 PR265579 PR266795 PR269690
CTUDT00631155 | Refresh | Closed | altshuller_mark | September 16, 2020 1:57:55 PM GMT+08:00 | Hawk Pre-Validation (WIP 18.4): Couldn't trigger clinical scan manually during bolus tracker sequences | Hawk |  |  | System |  | N/A | Duplicate | tiwari_g | 0 | 2 |  | Console4 | YES | Closed |  | IQon1
CTUDT00631153 | Refresh | Closed | altshuller_mark | September 16, 2020 1:54:30 PM GMT+08:00 | Workflow freeze while Trauma scanning | Hawk |  | NA | Subsystem |  | N/A | Duplicate | tiwari_g | 0 | 2 |  | Console4 | NO | Closed |  | IQon1
CTUDT00631151 | Refresh | Assigned | altshuller_mark | September 16, 2020 1:42:03 PM GMT+08:00 | WIP13.1 – after selecting ‘Abdomen’ EC to load to plan scene the popups “System is still initializing and there seems to be a problem” appeared. | Hawk |  | Relaunch the application. | System | SP2_Candidate | N/A |  | tiwari_g | 0 | 3 | Console Generic | Console4 | NO | Assigned | SP2 | Ovation
CTUDT00631151 | Refresh | Assigned | altshuller_mark | September 16, 2020 1:42:03 PM GMT+08:00 | WIP13.1 – after selecting ‘Abdomen’ EC to load to plan scene the popups “System is still initializing and there seems to be a problem” appeared. | Hawk |  | Relaunch the application. | System | SP2_Candidate | N/A |  | tiwari_g | 0 | 3 | Console Generic | Console4 | NO | Assigned | SP2 | Brilliance 64
CTUDT00631151 | Refresh | Assigned | altshuller_mark | September 16, 2020 1:42:03 PM GMT+08:00 | WIP13.1 – after selecting ‘Abdomen’ EC to load to plan scene the popups “System is still initializing and there seems to be a problem” appeared. | Hawk |  | Relaunch the application. | System | SP2_Candidate | N/A |  | tiwari_g | 0 | 3 | Console Generic | Console4 | NO | Assigned | SP2 | Brilliance Big Bore
CTUDT00631151 | Refresh | Assigned | altshuller_mark | September 16, 2020 1:42:03 PM GMT+08:00 | WIP13.1 – after selecting ‘Abdomen’ EC to load to plan scene the popups “System is still initializing and there seems to be a problem” appeared. | Hawk |  | Relaunch the application. | System | SP2_Candidate | N/A |  | tiwari_g | 0 | 3 | Console Generic | Console4 | NO | Assigned | SP2 | Ingenuity 64
CTUDT00631151 | Refresh | Assigned | altshuller_mark | September 16, 2020 1:42:03 PM GMT+08:00 | WIP13.1 – after selecting ‘Abdomen’ EC to load to plan scene the popups “System is still initializing and there seems to be a problem” appeared. | Hawk |  | Relaunch the application. | System | SP2_Candidate | N/A |  | tiwari_g | 0 | 3 | Console Generic | Console4 | NO | Assigned | SP2 | IQon1
CTUDT00631143 | Refresh | Closed | lin_s | September 16, 2020 11:54:21 AM GMT+08:00 | FOK: Enable Remote Access not present in Home Menu. | 4.1.10 |  | The FSE should be educated not to execute the tool for Non-DOD sites and there is no way to revert these changes | Subsystem | FOK | N/A | Works As Specified / Not A Defect | Prabhakar_Priyank_Muppidi | 0 |  |  | Console4 | NO | Closed |  | Ovation | PR275190
CTUDT00631143 | Refresh | Closed | lin_s | September 16, 2020 11:54:21 AM GMT+08:00 | FOK: Enable Remote Access not present in Home Menu. | 4.1.10 |  | The FSE should be educated not to execute the tool for Non-DOD sites and there is no way to revert these changes | Subsystem | FOK | N/A | Works As Specified / Not A Defect | Prabhakar_Priyank_Muppidi | 0 |  |  | Console4 | NO | Closed |  | Brilliance 64 | PR275190
CTUDT00631143 | Refresh | Closed | lin_s | September 16, 2020 11:54:21 AM GMT+08:00 | FOK: Enable Remote Access not present in Home Menu. | 4.1.10 |  | The FSE should be educated not to execute the tool for Non-DOD sites and there is no way to revert these changes | Subsystem | FOK | N/A | Works As Specified / Not A Defect | Prabhakar_Priyank_Muppidi | 0 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore | PR275190
CTUDT00631143 | Refresh | Closed | lin_s | September 16, 2020 11:54:21 AM GMT+08:00 | FOK: Enable Remote Access not present in Home Menu. | 4.1.10 |  | The FSE should be educated not to execute the tool for Non-DOD sites and there is no way to revert these changes | Subsystem | FOK | N/A | Works As Specified / Not A Defect | Prabhakar_Priyank_Muppidi | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64 | PR275190
CTUDT00631143 | Refresh | Closed | lin_s | September 16, 2020 11:54:21 AM GMT+08:00 | FOK: Enable Remote Access not present in Home Menu. | 4.1.10 |  | The FSE should be educated not to execute the tool for Non-DOD sites and there is no way to revert these changes | Subsystem | FOK | N/A | Works As Specified / Not A Defect | Prabhakar_Priyank_Muppidi | 0 |  |  | Console4 | NO | Closed |  | IQon1 | PR275190
CTUDT00631143 | Refresh | Closed | lin_s | September 16, 2020 11:54:21 AM GMT+08:00 | FOK: Enable Remote Access not present in Home Menu. | 4.1.10 |  | The FSE should be educated not to execute the tool for Non-DOD sites and there is no way to revert these changes | Subsystem | FOK | N/A | Works As Specified / Not A Defect | Prabhakar_Priyank_Muppidi | 0 |  |  | Console4 | NO | Closed |  | Hawk | PR275190
CTUDT00631142 | Refresh | Accepted | zhong_x | September 16, 2020 11:50:06 AM GMT+08:00 | FOK-Warning message "Spectral annotations check" pop up when adding ING images into ISP film. | WIP25.2 |  |  | Subsystem | FOK; SP2_Candidate |  |  | 320057847 | 0 |  |  | Console4 | NO | Accepted | SP2 | Ingenuity 64 | 277227
CTUDT00631140 | Refresh | Closed | lin_s | September 16, 2020 9:25:19 AM GMT+08:00 | FOK: The message box of remote control cannot be hidden. | 4.1.10 |  | No Workaround | Subsystem | FOK | STPS-SSRS-98 | Works As Specified / Not A Defect | venkatachalam_p | 0 |  |  | Console4 | NO | Closed |  | Ovation | PR243621
CTUDT00631140 | Refresh | Closed | lin_s | September 16, 2020 9:25:19 AM GMT+08:00 | FOK: The message box of remote control cannot be hidden. | 4.1.10 |  | No Workaround | Subsystem | FOK | STPS-SSRS-98 | Works As Specified / Not A Defect | venkatachalam_p | 0 |  |  | Console4 | NO | Closed |  | Brilliance 64 | PR243621
CTUDT00631140 | Refresh | Closed | lin_s | September 16, 2020 9:25:19 AM GMT+08:00 | FOK: The message box of remote control cannot be hidden. | 4.1.10 |  | No Workaround | Subsystem | FOK | STPS-SSRS-98 | Works As Specified / Not A Defect | venkatachalam_p | 0 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore | PR243621
CTUDT00631140 | Refresh | Closed | lin_s | September 16, 2020 9:25:19 AM GMT+08:00 | FOK: The message box of remote control cannot be hidden. | 4.1.10 |  | No Workaround | Subsystem | FOK | STPS-SSRS-98 | Works As Specified / Not A Defect | venkatachalam_p | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64 | PR243621
CTUDT00631140 | Refresh | Closed | lin_s | September 16, 2020 9:25:19 AM GMT+08:00 | FOK: The message box of remote control cannot be hidden. | 4.1.10 |  | No Workaround | Subsystem | FOK | STPS-SSRS-98 | Works As Specified / Not A Defect | venkatachalam_p | 0 |  |  | Console4 | NO | Closed |  | IQon1 | PR243621
CTUDT00631136 | Refresh | Closed | lin_s | September 16, 2020 12:02:40 AM GMT+08:00 | FOK: Recon - Slab Tool-Image Rotation not working properly | 4.1.10 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | atin_s | 0 | 3 |  | Console4 | NO | Closed | SP1 | Ingenuity 64 | 274937 278078
CTUDT00631134 | Refresh | Closed | lin_s | September 15, 2020 11:47:52 PM GMT+08:00 | FOK: RFC: Customer feedback for 3 items which need be improved in ING system. | RFD |  |  | System | FOK | N/A | Change Proposal | zhou_r | 0 |  |  | Couch | NO | Closed |  | Ingenuity 64 | 266893
CTUDT00631132 | Refresh | Closed | periasamy_g | September 15, 2020 11:01:48 PM GMT+08:00 | User management tool is allowing to create the username with special characters (unsupported) and later while logging its popping error as invalid username | WIP 25.4 |  |  | Subsystem |  | NA | Duplicate | kumar_gau | 0 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00631132 | Refresh | Closed | periasamy_g | September 15, 2020 11:01:48 PM GMT+08:00 | User management tool is allowing to create the username with special characters (unsupported) and later while logging its popping error as invalid username | WIP 25.4 |  |  | Subsystem |  | NA | Duplicate | kumar_gau | 0 |  |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00631132 | Refresh | Closed | periasamy_g | September 15, 2020 11:01:48 PM GMT+08:00 | User management tool is allowing to create the username with special characters (unsupported) and later while logging its popping error as invalid username | WIP 25.4 |  |  | Subsystem |  | NA | Duplicate | kumar_gau | 0 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00631132 | Refresh | Closed | periasamy_g | September 15, 2020 11:01:48 PM GMT+08:00 | User management tool is allowing to create the username with special characters (unsupported) and later while logging its popping error as invalid username | WIP 25.4 |  |  | Subsystem |  | NA | Duplicate | kumar_gau | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00631132 | Refresh | Closed | periasamy_g | September 15, 2020 11:01:48 PM GMT+08:00 | User management tool is allowing to create the username with special characters (unsupported) and later while logging its popping error as invalid username | WIP 25.4 |  |  | Subsystem |  | NA | Duplicate | kumar_gau | 0 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00631130 | Refresh | Closed | periasamy_g | September 15, 2020 10:43:14 PM GMT+08:00 | Two requirements are capturing about same behavior with different performance index numbers | 4.7.5 | DHF303211_Rev24 DHF316705 Rev23 | NA | Subsystem |  | Console-SRS.11582 | Fixed | periasamy_g | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | SSRS new version | Ovation |  | DHF303211_Rev24 DHF316705 Rev23 | October 9, 2020 9:47:27 PM GMT+08:00
CTUDT00631130 | Refresh | Closed | periasamy_g | September 15, 2020 10:43:14 PM GMT+08:00 | Two requirements are capturing about same behavior with different performance index numbers | 4.7.5 | DHF303211_Rev24 DHF316705 Rev23 | NA | Subsystem |  | Console-SRS.11582 | Fixed | periasamy_g | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | SSRS new version | Brilliance 64 |  | DHF303211_Rev24 DHF316705 Rev23 | October 9, 2020 9:47:27 PM GMT+08:00
CTUDT00631130 | Refresh | Closed | periasamy_g | September 15, 2020 10:43:14 PM GMT+08:00 | Two requirements are capturing about same behavior with different performance index numbers | 4.7.5 | DHF303211_Rev24 DHF316705 Rev23 | NA | Subsystem |  | Console-SRS.11582 | Fixed | periasamy_g | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | SSRS new version | Brilliance Big Bore |  | DHF303211_Rev24 DHF316705 Rev23 | October 9, 2020 9:47:27 PM GMT+08:00
CTUDT00631130 | Refresh | Closed | periasamy_g | September 15, 2020 10:43:14 PM GMT+08:00 | Two requirements are capturing about same behavior with different performance index numbers | 4.7.5 | DHF303211_Rev24 DHF316705 Rev23 | NA | Subsystem |  | Console-SRS.11582 | Fixed | periasamy_g | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | SSRS new version | Ingenuity 64 |  | DHF303211_Rev24 DHF316705 Rev23 | October 9, 2020 9:47:27 PM GMT+08:00
CTUDT00631130 | Refresh | Closed | periasamy_g | September 15, 2020 10:43:14 PM GMT+08:00 | Two requirements are capturing about same behavior with different performance index numbers | 4.7.5 | DHF303211_Rev24 DHF316705 Rev23 | NA | Subsystem |  | Console-SRS.11582 | Fixed | periasamy_g | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | SSRS new version | IQon1 |  | DHF303211_Rev24 DHF316705 Rev23 | October 9, 2020 9:47:27 PM GMT+08:00
CTUDT00631130 | Refresh | Closed | periasamy_g | September 15, 2020 10:43:14 PM GMT+08:00 | Two requirements are capturing about same behavior with different performance index numbers | 4.7.5 | DHF303211_Rev24 DHF316705 Rev23 | NA | Subsystem |  | Console-SRS.11582 | Fixed | periasamy_g | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | SSRS new version | Hawk |  | DHF303211_Rev24 DHF316705 Rev23 | October 9, 2020 9:47:27 PM GMT+08:00
CTUDT00631122 | Refresh | Closed | elavarasan_s | September 15, 2020 4:21:58 PM GMT+08:00 | In split study scenario, Surview of second and third study is not transferred to configured remote device and Transfer icon is not displayed in Patient Directory. | WIP25.5 | WIP25.6 |  | Subsystem | SP1_PM_Approved | N/A | Fixed | challa_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP25.6SegReq | Ovation |  | WIP25.6 | September 21, 2020 6:37:10 PM GMT+08:00
CTUDT00631122 | Refresh | Closed | elavarasan_s | September 15, 2020 4:21:58 PM GMT+08:00 | In split study scenario, Surview of second and third study is not transferred to configured remote device and Transfer icon is not displayed in Patient Directory. | WIP25.5 | WIP25.6 |  | Subsystem | SP1_PM_Approved | N/A | Fixed | challa_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP25.6SegReq | Brilliance 64 |  | WIP25.6 | September 21, 2020 6:37:10 PM GMT+08:00
CTUDT00631122 | Refresh | Closed | elavarasan_s | September 15, 2020 4:21:58 PM GMT+08:00 | In split study scenario, Surview of second and third study is not transferred to configured remote device and Transfer icon is not displayed in Patient Directory. | WIP25.5 | WIP25.6 |  | Subsystem | SP1_PM_Approved | N/A | Fixed | challa_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP25.6SegReq | Brilliance Big Bore |  | WIP25.6 | September 21, 2020 6:37:10 PM GMT+08:00
CTUDT00631122 | Refresh | Closed | elavarasan_s | September 15, 2020 4:21:58 PM GMT+08:00 | In split study scenario, Surview of second and third study is not transferred to configured remote device and Transfer icon is not displayed in Patient Directory. | WIP25.5 | WIP25.6 |  | Subsystem | SP1_PM_Approved | N/A | Fixed | challa_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP25.6SegReq | Ingenuity 64 |  | WIP25.6 | September 21, 2020 6:37:10 PM GMT+08:00
CTUDT00631122 | Refresh | Closed | elavarasan_s | September 15, 2020 4:21:58 PM GMT+08:00 | In split study scenario, Surview of second and third study is not transferred to configured remote device and Transfer icon is not displayed in Patient Directory. | WIP25.5 | WIP25.6 |  | Subsystem | SP1_PM_Approved | N/A | Fixed | challa_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP25.6SegReq | IQon1 |  | WIP25.6 | September 21, 2020 6:37:10 PM GMT+08:00
CTUDT00631117 | Refresh | Closed | prabhan_p | September 15, 2020 2:31:46 PM GMT+08:00 | Right monitor turns yellow during offline recon. | wip 25.5 |  | user performs any action in the left monitor, | Subsystem | InvestigateP2 | N/A | No Plan to Fix | modi_s | 0 | 4 |  | Console4 | NO | Closed |  | IQon1
CTUDT00631115 | Refresh | Closed | lakshmikanth_s | September 15, 2020 2:17:32 PM GMT+08:00 | System or Bugrep error when clicked on Volume icon in View2 tools | WIP25.5 |  |  | Subsystem | SP1_PM_Approved | NA | Duplicate | s_srivastava | 0 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00631115 | Refresh | Closed | lakshmikanth_s | September 15, 2020 2:17:32 PM GMT+08:00 | System or Bugrep error when clicked on Volume icon in View2 tools | WIP25.5 |  |  | Subsystem | SP1_PM_Approved | NA | Duplicate | s_srivastava | 0 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00631113 | Refresh | Closed | wang_l | September 15, 2020 1:19:36 PM GMT+08:00 | Refresh - The duplicate localized language options were shown when clicking Win+Space in Workflow for Dutch/Spanish | WIP25.5 |  | Do not press Win+ Space as the application and the Login window are already allowing the respective keyboard selected in preferences. NOTE: The usage of Win+Space and SHIFT key is only for Chinese keyboard as the user has to select between the English and Chinese. | Subsystem |  | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | Ovation
CTUDT00631113 | Refresh | Closed | wang_l | September 15, 2020 1:19:36 PM GMT+08:00 | Refresh - The duplicate localized language options were shown when clicking Win+Space in Workflow for Dutch/Spanish | WIP25.5 |  | Do not press Win+ Space as the application and the Login window are already allowing the respective keyboard selected in preferences. NOTE: The usage of Win+Space and SHIFT key is only for Chinese keyboard as the user has to select between the English and Chinese. | Subsystem |  | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00631113 | Refresh | Closed | wang_l | September 15, 2020 1:19:36 PM GMT+08:00 | Refresh - The duplicate localized language options were shown when clicking Win+Space in Workflow for Dutch/Spanish | WIP25.5 |  | Do not press Win+ Space as the application and the Login window are already allowing the respective keyboard selected in preferences. NOTE: The usage of Win+Space and SHIFT key is only for Chinese keyboard as the user has to select between the English and Chinese. | Subsystem |  | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00631113 | Refresh | Closed | wang_l | September 15, 2020 1:19:36 PM GMT+08:00 | Refresh - The duplicate localized language options were shown when clicking Win+Space in Workflow for Dutch/Spanish | WIP25.5 |  | Do not press Win+ Space as the application and the Login window are already allowing the respective keyboard selected in preferences. NOTE: The usage of Win+Space and SHIFT key is only for Chinese keyboard as the user has to select between the English and Chinese. | Subsystem |  | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00631113 | Refresh | Closed | wang_l | September 15, 2020 1:19:36 PM GMT+08:00 | Refresh - The duplicate localized language options were shown when clicking Win+Space in Workflow for Dutch/Spanish | WIP25.5 |  | Do not press Win+ Space as the application and the Login window are already allowing the respective keyboard selected in preferences. NOTE: The usage of Win+Space and SHIFT key is only for Chinese keyboard as the user has to select between the English and Chinese. | Subsystem |  | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | IQon1
CTUDT00631105 | Refresh | Assigned | 320085947 | September 14, 2020 7:24:00 PM GMT+08:00 | FoK: Study date and time filter is unfunctional in "History" tab | WIP25.5 |  |  | System | FoK,SP2_Candidate | Console-SRS.17643 |  | s_srivastava | 0 | 3 |  | Console4 | NO | Assigned | SP2 | Ovation | PR271983
CTUDT00631105 | Refresh | Assigned | 320085947 | September 14, 2020 7:24:00 PM GMT+08:00 | FoK: Study date and time filter is unfunctional in "History" tab | WIP25.5 |  |  | System | FoK,SP2_Candidate | Console-SRS.17643 |  | s_srivastava | 0 | 3 |  | Console4 | NO | Assigned | SP2 | Brilliance 64 | PR271983
CTUDT00631105 | Refresh | Assigned | 320085947 | September 14, 2020 7:24:00 PM GMT+08:00 | FoK: Study date and time filter is unfunctional in "History" tab | WIP25.5 |  |  | System | FoK,SP2_Candidate | Console-SRS.17643 |  | s_srivastava | 0 | 3 |  | Console4 | NO | Assigned | SP2 | Brilliance Big Bore | PR271983
CTUDT00631105 | Refresh | Assigned | 320085947 | September 14, 2020 7:24:00 PM GMT+08:00 | FoK: Study date and time filter is unfunctional in "History" tab | WIP25.5 |  |  | System | FoK,SP2_Candidate | Console-SRS.17643 |  | s_srivastava | 0 | 3 |  | Console4 | NO | Assigned | SP2 | Ingenuity 64 | PR271983
CTUDT00631105 | Refresh | Assigned | 320085947 | September 14, 2020 7:24:00 PM GMT+08:00 | FoK: Study date and time filter is unfunctional in "History" tab | WIP25.5 |  |  | System | FoK,SP2_Candidate | Console-SRS.17643 |  | s_srivastava | 0 | 3 |  | Console4 | NO | Assigned | SP2 | IQon1 | PR271983
CTUDT00631103 | Refresh | Closed | altshuller_mark | September 14, 2020 5:14:45 PM GMT+08:00 | result of scan is not selected by default on view2 | HotFix2 | WIP25.6 |  | Subsystem | PM_Approved | N/A | Fixed | s_srivastava | 0 | 2 | Console Generic | Console4 | NO | Closed | HotFix2 | Hawk |  | WIP25.6 | September 21, 2020 6:29:21 PM GMT+08:00
CTUDT00631096 | Refresh | Closed | 320085947 | September 13, 2020 9:10:14 PM GMT+08:00 | Wrong serviceability version number appears in IQon WIP25.5 | WIP25.5 |  |  | Subsystem |  | N/A | Works As Specified / Not A Defect | richards_b | 0 |  |  | CIRS | NO | Closed |  | IQon1
CTUDT00631095 | Refresh | Assigned | gT_s | September 13, 2020 9:09:11 PM GMT+08:00 | Refresh : One of thin slice image is not getting transferred to the remote devices | WIP 25.5 |  |  | Subsystem | SP2_Candidate | NA |  | challa_s | 0 | 2 |  | Console4 | NO | Assigned | SP2 | Ovation
CTUDT00631095 | Refresh | Assigned | gT_s | September 13, 2020 9:09:11 PM GMT+08:00 | Refresh : One of thin slice image is not getting transferred to the remote devices | WIP 25.5 |  |  | Subsystem | SP2_Candidate | NA |  | challa_s | 0 | 2 |  | Console4 | NO | Assigned | SP2 | Brilliance 64
CTUDT00631095 | Refresh | Assigned | gT_s | September 13, 2020 9:09:11 PM GMT+08:00 | Refresh : One of thin slice image is not getting transferred to the remote devices | WIP 25.5 |  |  | Subsystem | SP2_Candidate | NA |  | challa_s | 0 | 2 |  | Console4 | NO | Assigned | SP2 | Brilliance Big Bore
CTUDT00631095 | Refresh | Assigned | gT_s | September 13, 2020 9:09:11 PM GMT+08:00 | Refresh : One of thin slice image is not getting transferred to the remote devices | WIP 25.5 |  |  | Subsystem | SP2_Candidate | NA |  | challa_s | 0 | 2 |  | Console4 | NO | Assigned | SP2 | Ingenuity 64
CTUDT00631095 | Refresh | Assigned | gT_s | September 13, 2020 9:09:11 PM GMT+08:00 | Refresh : One of thin slice image is not getting transferred to the remote devices | WIP 25.5 |  |  | Subsystem | SP2_Candidate | NA |  | challa_s | 0 | 2 |  | Console4 | NO | Assigned | SP2 | IQon1
CTUDT00631086 | Refresh | Assigned | cai_t | September 11, 2020 7:45:11 PM GMT+08:00 | [Refresh] :Language setting cannot be restored by Optional Backup/Restore. | WIP25.4 |  | N/A | Subsystem | SP2_Candidate | STPS-SSRS-159 |  | venkatachalam_p | 0 | 3 |  | Service Tools | NO | Assigned | SP2 | Brilliance Big Bore
CTUDT00631075 | Refresh | Assigned | wang_l | September 11, 2020 9:13:09 AM GMT+08:00 | Refresh: Lan Config Stop Working when do Restore | WIP31 |  | Since upgrade and back and restore was done , some window corruption happened and needed a restart of the system. | Subsystem | SP1_PM_Approved | NA |  | hilas | 0 | 4 |  | Support-System | NO | Assigned | SP1_Doc | Ovation
CTUDT00631075 | Refresh | Assigned | wang_l | September 11, 2020 9:13:09 AM GMT+08:00 | Refresh: Lan Config Stop Working when do Restore | WIP31 |  | Since upgrade and back and restore was done , some window corruption happened and needed a restart of the system. | Subsystem | SP1_PM_Approved | NA |  | hilas | 0 | 4 |  | Support-System | NO | Assigned | SP1_Doc | Brilliance 64
CTUDT00631075 | Refresh | Assigned | wang_l | September 11, 2020 9:13:09 AM GMT+08:00 | Refresh: Lan Config Stop Working when do Restore | WIP31 |  | Since upgrade and back and restore was done , some window corruption happened and needed a restart of the system. | Subsystem | SP1_PM_Approved | NA |  | hilas | 0 | 4 |  | Support-System | NO | Assigned | SP1_Doc | Brilliance Big Bore
CTUDT00631075 | Refresh | Assigned | wang_l | September 11, 2020 9:13:09 AM GMT+08:00 | Refresh: Lan Config Stop Working when do Restore | WIP31 |  | Since upgrade and back and restore was done , some window corruption happened and needed a restart of the system. | Subsystem | SP1_PM_Approved | NA |  | hilas | 0 | 4 |  | Support-System | NO | Assigned | SP1_Doc | Ingenuity 64
CTUDT00631075 | Refresh | Assigned | wang_l | September 11, 2020 9:13:09 AM GMT+08:00 | Refresh: Lan Config Stop Working when do Restore | WIP31 |  | Since upgrade and back and restore was done , some window corruption happened and needed a restart of the system. | Subsystem | SP1_PM_Approved | NA |  | hilas | 0 | 4 |  | Support-System | NO | Assigned | SP1_Doc | IQon1
CTUDT00631060 | Refresh | Closed | dongping_w | September 9, 2020 3:44:01 PM GMT+08:00 | FOK Refresh BB: A ding sound made with every keystrokes sometimes | WIP12.6 |  |  | Subsystem | FOK; InvestigateP2 |  | Does Not Apply | atin_s | 0 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore | 261233
CTUDT00631058 | Refresh | Assigned | lin_s | September 9, 2020 2:46:22 PM GMT+08:00 | Headline: FOK: Design improvement considertion: Table base is creaking loudly when 12” from the completely down position and 12” away in the upward movement | 4.1.10.25065 |  | Disassembly hte spring and redo the spring alignment will fix the issue. | Subsystem | FOK; LCM_Candidate ; SP2_Candidate | N/A |  | zhou_r | 0 | 4 |  | Couch | NO | Assigned | Design Improvement Next Version | Ovation | 250901
CTUDT00631058 | Refresh | Assigned | lin_s | September 9, 2020 2:46:22 PM GMT+08:00 | Headline: FOK: Design improvement considertion: Table base is creaking loudly when 12” from the completely down position and 12” away in the upward movement | 4.1.10.25065 |  | Disassembly hte spring and redo the spring alignment will fix the issue. | Subsystem | FOK; LCM_Candidate ; SP2_Candidate | N/A |  | zhou_r | 0 | 4 |  | Couch | NO | Assigned | Design Improvement Next Version | Brilliance 64 | 250901
CTUDT00631058 | Refresh | Assigned | lin_s | September 9, 2020 2:46:22 PM GMT+08:00 | Headline: FOK: Design improvement considertion: Table base is creaking loudly when 12” from the completely down position and 12” away in the upward movement | 4.1.10.25065 |  | Disassembly hte spring and redo the spring alignment will fix the issue. | Subsystem | FOK; LCM_Candidate ; SP2_Candidate | N/A |  | zhou_r | 0 | 4 |  | Couch | NO | Assigned | Design Improvement Next Version | Brilliance Big Bore | 250901
CTUDT00631058 | Refresh | Assigned | lin_s | September 9, 2020 2:46:22 PM GMT+08:00 | Headline: FOK: Design improvement considertion: Table base is creaking loudly when 12” from the completely down position and 12” away in the upward movement | 4.1.10.25065 |  | Disassembly hte spring and redo the spring alignment will fix the issue. | Subsystem | FOK; LCM_Candidate ; SP2_Candidate | N/A |  | zhou_r | 0 | 4 |  | Couch | NO | Assigned | Design Improvement Next Version | Ingenuity 64 | 250901
CTUDT00631058 | Refresh | Assigned | lin_s | September 9, 2020 2:46:22 PM GMT+08:00 | Headline: FOK: Design improvement considertion: Table base is creaking loudly when 12” from the completely down position and 12” away in the upward movement | 4.1.10.25065 |  | Disassembly hte spring and redo the spring alignment will fix the issue. | Subsystem | FOK; LCM_Candidate ; SP2_Candidate | N/A |  | zhou_r | 0 | 4 |  | Couch | NO | Assigned | Design Improvement Next Version | IQon1 | 250901
CTUDT00631057 | Refresh | Closed | altshuller_mark | September 9, 2020 1:26:28 PM GMT+08:00 | Hawk rotor balance false positive calibration | Hawk | WIP34 | No Workaround | System | SP1_PM_Approved | N/A | Fixed | morton_t | 0 | 2 | Physics | Support-Physics | NO | Closed | WIP34SubCom | IQon1 |  | WIP34 | December 2, 2020 3:11:02 AM GMT+08:00
CTUDT00631055 | Refresh | Closed | wang_l | September 9, 2020 9:23:09 AM GMT+08:00 | FOK:Refresh - Lan Config open with error if Network IPV4 add more than one static IP address | WIP31 | WIP33 | As a workaround , in such cases please uncheck IPv6 when the connection is no longer supported. | Subsystem | FOK; SP1_PM_Approved | NA | Fixed | srushti_k | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP33SubCom | Ovation |  | WIP33 | October 23, 2020 4:34:10 PM GMT+08:00
CTUDT00631055 | Refresh | Closed | wang_l | September 9, 2020 9:23:09 AM GMT+08:00 | FOK:Refresh - Lan Config open with error if Network IPV4 add more than one static IP address | WIP31 | WIP33 | As a workaround , in such cases please uncheck IPv6 when the connection is no longer supported. | Subsystem | FOK; SP1_PM_Approved | NA | Fixed | srushti_k | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP33SubCom | Brilliance 64 |  | WIP33 | October 23, 2020 4:34:10 PM GMT+08:00
CTUDT00631055 | Refresh | Closed | wang_l | September 9, 2020 9:23:09 AM GMT+08:00 | FOK:Refresh - Lan Config open with error if Network IPV4 add more than one static IP address | WIP31 | WIP33 | As a workaround , in such cases please uncheck IPv6 when the connection is no longer supported. | Subsystem | FOK; SP1_PM_Approved | NA | Fixed | srushti_k | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP33SubCom | Brilliance Big Bore |  | WIP33 | October 23, 2020 4:34:10 PM GMT+08:00
CTUDT00631055 | Refresh | Closed | wang_l | September 9, 2020 9:23:09 AM GMT+08:00 | FOK:Refresh - Lan Config open with error if Network IPV4 add more than one static IP address | WIP31 | WIP33 | As a workaround , in such cases please uncheck IPv6 when the connection is no longer supported. | Subsystem | FOK; SP1_PM_Approved | NA | Fixed | srushti_k | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP33SubCom | Ingenuity 64 |  | WIP33 | October 23, 2020 4:34:10 PM GMT+08:00
CTUDT00631055 | Refresh | Closed | wang_l | September 9, 2020 9:23:09 AM GMT+08:00 | FOK:Refresh - Lan Config open with error if Network IPV4 add more than one static IP address | WIP31 | WIP33 | As a workaround , in such cases please uncheck IPv6 when the connection is no longer supported. | Subsystem | FOK; SP1_PM_Approved | NA | Fixed | srushti_k | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP33SubCom | IQon1 |  | WIP33 | October 23, 2020 4:34:10 PM GMT+08:00
CTUDT00631055 | Refresh | Closed | wang_l | September 9, 2020 9:23:09 AM GMT+08:00 | FOK:Refresh - Lan Config open with error if Network IPV4 add more than one static IP address | WIP31 | WIP33 | As a workaround , in such cases please uncheck IPv6 when the connection is no longer supported. | Subsystem | FOK; SP1_PM_Approved | NA | Fixed | srushti_k | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP33SubCom | Hawk |  | WIP33 | October 23, 2020 4:34:10 PM GMT+08:00
CTUDT00631053 | Refresh | Resolved | shaham_t | September 9, 2020 1:43:18 AM GMT+08:00 | Login page is switching to Windows login and login is getting failed when new user added with "mu" special character | WIP25 | WIP35 |  | Subsystem | SP1_PM_Approved | NA | Fixed | kumar_gau | 0 | 2 | Console Software | Console4 | NO | Resolved | WIP35SubCom | Ovation |  | WIP35 | November 27, 2020 6:40:49 PM GMT+08:00
CTUDT00631053 | Refresh | Resolved | shaham_t | September 9, 2020 1:43:18 AM GMT+08:00 | Login page is switching to Windows login and login is getting failed when new user added with "mu" special character | WIP25 | WIP35 |  | Subsystem | SP1_PM_Approved | NA | Fixed | kumar_gau | 0 | 2 | Console Software | Console4 | NO | Resolved | WIP35SubCom | Brilliance 64 |  | WIP35 | November 27, 2020 6:40:49 PM GMT+08:00
CTUDT00631053 | Refresh | Resolved | shaham_t | September 9, 2020 1:43:18 AM GMT+08:00 | Login page is switching to Windows login and login is getting failed when new user added with "mu" special character | WIP25 | WIP35 |  | Subsystem | SP1_PM_Approved | NA | Fixed | kumar_gau | 0 | 2 | Console Software | Console4 | NO | Resolved | WIP35SubCom | Brilliance Big Bore |  | WIP35 | November 27, 2020 6:40:49 PM GMT+08:00
CTUDT00631053 | Refresh | Resolved | shaham_t | September 9, 2020 1:43:18 AM GMT+08:00 | Login page is switching to Windows login and login is getting failed when new user added with "mu" special character | WIP25 | WIP35 |  | Subsystem | SP1_PM_Approved | NA | Fixed | kumar_gau | 0 | 2 | Console Software | Console4 | NO | Resolved | WIP35SubCom | Ingenuity 64 |  | WIP35 | November 27, 2020 6:40:49 PM GMT+08:00
CTUDT00631053 | Refresh | Resolved | shaham_t | September 9, 2020 1:43:18 AM GMT+08:00 | Login page is switching to Windows login and login is getting failed when new user added with "mu" special character | WIP25 | WIP35 |  | Subsystem | SP1_PM_Approved | NA | Fixed | kumar_gau | 0 | 2 | Console Software | Console4 | NO | Resolved | WIP35SubCom | IQon1 |  | WIP35 | November 27, 2020 6:40:49 PM GMT+08:00
CTUDT00631052 | Refresh | Resolved | shaham_t | September 8, 2020 10:00:09 PM GMT+08:00 | FoK: Customer reported with the issue in Application, look like slow process and with Not recon image message | 4.7.7 | WIP32 |  | Subsystem | SP1_PM_Approved | NA | Fixed | s_srivastava | 0 | 2 | Console Software | Console4 | NO | Resolved | WIP32SubCom | Ovation | PR200433 PR203331 PR217226 | WIP32 | October 6, 2020 12:17:13 AM GMT+08:00
CTUDT00631052 | Refresh | Resolved | shaham_t | September 8, 2020 10:00:09 PM GMT+08:00 | FoK: Customer reported with the issue in Application, look like slow process and with Not recon image message | 4.7.7 | WIP32 |  | Subsystem | SP1_PM_Approved | NA | Fixed | s_srivastava | 0 | 2 | Console Software | Console4 | NO | Resolved | WIP32SubCom | Brilliance 64 | PR200433 PR203331 PR217226 | WIP32 | October 6, 2020 12:17:13 AM GMT+08:00
CTUDT00631052 | Refresh | Resolved | shaham_t | September 8, 2020 10:00:09 PM GMT+08:00 | FoK: Customer reported with the issue in Application, look like slow process and with Not recon image message | 4.7.7 | WIP32 |  | Subsystem | SP1_PM_Approved | NA | Fixed | s_srivastava | 0 | 2 | Console Software | Console4 | NO | Resolved | WIP32SubCom | Brilliance Big Bore | PR200433 PR203331 PR217226 | WIP32 | October 6, 2020 12:17:13 AM GMT+08:00
CTUDT00631052 | Refresh | Resolved | shaham_t | September 8, 2020 10:00:09 PM GMT+08:00 | FoK: Customer reported with the issue in Application, look like slow process and with Not recon image message | 4.7.7 | WIP32 |  | Subsystem | SP1_PM_Approved | NA | Fixed | s_srivastava | 0 | 2 | Console Software | Console4 | NO | Resolved | WIP32SubCom | Ingenuity 64 | PR200433 PR203331 PR217226 | WIP32 | October 6, 2020 12:17:13 AM GMT+08:00
CTUDT00631052 | Refresh | Resolved | shaham_t | September 8, 2020 10:00:09 PM GMT+08:00 | FoK: Customer reported with the issue in Application, look like slow process and with Not recon image message | 4.7.7 | WIP32 |  | Subsystem | SP1_PM_Approved | NA | Fixed | s_srivastava | 0 | 2 | Console Software | Console4 | NO | Resolved | WIP32SubCom | IQon1 | PR200433 PR203331 PR217226 | WIP32 | October 6, 2020 12:17:13 AM GMT+08:00
CTUDT00631048 | Refresh | Postponed | lakshmikanth_s | September 8, 2020 2:23:49 PM GMT+08:00 | Not able to login with Chinese credentials | WIP25.4 |  | No WorkAround | Subsystem |  | STPS-SSRS-1201 |  | kumar_gau | 0 | 2 |  | Console4 | NO | Postponed |  | Brilliance Big Bore
CTUDT00631048 | Refresh | Postponed | lakshmikanth_s | September 8, 2020 2:23:49 PM GMT+08:00 | Not able to login with Chinese credentials | WIP25.4 |  | No WorkAround | Subsystem |  | STPS-SSRS-1201 |  | kumar_gau | 0 | 2 |  | Console4 | NO | Postponed |  | IQon1
CTUDT00631047 | Refresh | Postponed | lei_yang | September 8, 2020 2:23:18 PM GMT+08:00 | “ResultsEngineExe has stopped working” pops up when clicking log out from workflow. | WIP31 |  |  | Subsystem | InvestigateP2 ; SP2_Candidate | NA |  | janakiraman_j | 0 | 3 |  | Console4 | NO | Postponed |  | Ingenuity 64
CTUDT00631045 | Refresh | Resolved | lei_yang | September 8, 2020 2:05:11 PM GMT+08:00 | After push remote certificate, the application Visual Diagnostics and Remote Configuration will not be displayed in service tool. | WIP31 | WIP34 |  | Subsystem | SP1_PM_Approved | NA | Fixed | chakraborty_d | 0 | 2 | Service Tools_Software | Service Tools | NO | Resolved | WIP34SubCom | Ovation |  | WIP34 | November 5, 2020 3:13:59 PM GMT+08:00
CTUDT00631045 | Refresh | Resolved | lei_yang | September 8, 2020 2:05:11 PM GMT+08:00 | After push remote certificate, the application Visual Diagnostics and Remote Configuration will not be displayed in service tool. | WIP31 | WIP34 |  | Subsystem | SP1_PM_Approved | NA | Fixed | chakraborty_d | 0 | 2 | Service Tools_Software | Service Tools | NO | Resolved | WIP34SubCom | Brilliance 64 |  | WIP34 | November 5, 2020 3:13:59 PM GMT+08:00
CTUDT00631045 | Refresh | Resolved | lei_yang | September 8, 2020 2:05:11 PM GMT+08:00 | After push remote certificate, the application Visual Diagnostics and Remote Configuration will not be displayed in service tool. | WIP31 | WIP34 |  | Subsystem | SP1_PM_Approved | NA | Fixed | chakraborty_d | 0 | 2 | Service Tools_Software | Service Tools | NO | Resolved | WIP34SubCom | Brilliance Big Bore |  | WIP34 | November 5, 2020 3:13:59 PM GMT+08:00
CTUDT00631045 | Refresh | Resolved | lei_yang | September 8, 2020 2:05:11 PM GMT+08:00 | After push remote certificate, the application Visual Diagnostics and Remote Configuration will not be displayed in service tool. | WIP31 | WIP34 |  | Subsystem | SP1_PM_Approved | NA | Fixed | chakraborty_d | 0 | 2 | Service Tools_Software | Service Tools | NO | Resolved | WIP34SubCom | Ingenuity 64 |  | WIP34 | November 5, 2020 3:13:59 PM GMT+08:00
CTUDT00631045 | Refresh | Resolved | lei_yang | September 8, 2020 2:05:11 PM GMT+08:00 | After push remote certificate, the application Visual Diagnostics and Remote Configuration will not be displayed in service tool. | WIP31 | WIP34 |  | Subsystem | SP1_PM_Approved | NA | Fixed | chakraborty_d | 0 | 2 | Service Tools_Software | Service Tools | NO | Resolved | WIP34SubCom | IQon1 |  | WIP34 | November 5, 2020 3:13:59 PM GMT+08:00
CTUDT00631039 | Refresh | Closed | lin_s | September 8, 2020 10:41:51 AM GMT+08:00 | FOK: Gantry not ready message after scanner sat idle for a few | 4.1.10 |  |  | Subsystem | FOK | NA | Duplicate | tan_j | 0 |  |  | GCX | YES | Closed |  | Ingenuity 64 | 261069
CTUDT00631037 | Refresh | Closed | lei_yang | September 8, 2020 8:37:10 AM GMT+08:00 | “CardiacServer has stopped working” pops up when clicking log out from workflow. | WIP31 |  |  | Subsystem | InvestigateP2 | NA | No Plan to Fix | janakiraman_j | 0 | 4 |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00631026 | Refresh | Resolved | lei_yang | September 7, 2020 11:04:54 AM GMT+08:00 | There is no response and no pop-up box appears when clicking the “Print” button in EC Sharing with CT user. | WIP31 | WIP37 |  | Subsystem | InvestigateP1; SP1_PM_Approved | NA | Fixed | challa_s | 0 | 2 | Console Software | Console4 | NO | Resolved | WIP37SegReq | Ovation |  | WIP37 | December 16, 2020 7:39:38 PM GMT+08:00
CTUDT00631026 | Refresh | Resolved | lei_yang | September 7, 2020 11:04:54 AM GMT+08:00 | There is no response and no pop-up box appears when clicking the “Print” button in EC Sharing with CT user. | WIP31 | WIP37 |  | Subsystem | InvestigateP1; SP1_PM_Approved | NA | Fixed | challa_s | 0 | 2 | Console Software | Console4 | NO | Resolved | WIP37SegReq | Brilliance 64 |  | WIP37 | December 16, 2020 7:39:38 PM GMT+08:00
CTUDT00631026 | Refresh | Resolved | lei_yang | September 7, 2020 11:04:54 AM GMT+08:00 | There is no response and no pop-up box appears when clicking the “Print” button in EC Sharing with CT user. | WIP31 | WIP37 |  | Subsystem | InvestigateP1; SP1_PM_Approved | NA | Fixed | challa_s | 0 | 2 | Console Software | Console4 | NO | Resolved | WIP37SegReq | Brilliance Big Bore |  | WIP37 | December 16, 2020 7:39:38 PM GMT+08:00
CTUDT00631026 | Refresh | Resolved | lei_yang | September 7, 2020 11:04:54 AM GMT+08:00 | There is no response and no pop-up box appears when clicking the “Print” button in EC Sharing with CT user. | WIP31 | WIP37 |  | Subsystem | InvestigateP1; SP1_PM_Approved | NA | Fixed | challa_s | 0 | 2 | Console Software | Console4 | NO | Resolved | WIP37SegReq | Ingenuity 64 |  | WIP37 | December 16, 2020 7:39:38 PM GMT+08:00
CTUDT00631026 | Refresh | Resolved | lei_yang | September 7, 2020 11:04:54 AM GMT+08:00 | There is no response and no pop-up box appears when clicking the “Print” button in EC Sharing with CT user. | WIP31 | WIP37 |  | Subsystem | InvestigateP1; SP1_PM_Approved | NA | Fixed | challa_s | 0 | 2 | Console Software | Console4 | NO | Resolved | WIP37SegReq | IQon1 |  | WIP37 | December 16, 2020 7:39:38 PM GMT+08:00
CTUDT00631024 | Refresh | Assigned | elavarasan_s | September 4, 2020 11:08:27 PM GMT+08:00 | ePHI data is not scrubbed in Bugrep for CIRS Mdu and ExamRecordsRepository XML.gz file 1.Patient birth time in CIRS mdu didn't round off to nearest 10 minutes time 2.Study date in ExamRecordsRepository XML.gz didn’t modify the “day” to “1”. | WIP25.4 |  |  | Subsystem | SP2_Candidate | Console-SRS.18077 |  | veeraswamy_k | 0 | 3 |  | Service Tools | NO | Assigned | SP2 | Ovation
CTUDT00631024 | Refresh | Assigned | elavarasan_s | September 4, 2020 11:08:27 PM GMT+08:00 | ePHI data is not scrubbed in Bugrep for CIRS Mdu and ExamRecordsRepository XML.gz file 1.Patient birth time in CIRS mdu didn't round off to nearest 10 minutes time 2.Study date in ExamRecordsRepository XML.gz didn’t modify the “day” to “1”. | WIP25.4 |  |  | Subsystem | SP2_Candidate | Console-SRS.18077 |  | veeraswamy_k | 0 | 3 |  | Service Tools | NO | Assigned | SP2 | Brilliance 64
CTUDT00631024 | Refresh | Assigned | elavarasan_s | September 4, 2020 11:08:27 PM GMT+08:00 | ePHI data is not scrubbed in Bugrep for CIRS Mdu and ExamRecordsRepository XML.gz file 1.Patient birth time in CIRS mdu didn't round off to nearest 10 minutes time 2.Study date in ExamRecordsRepository XML.gz didn’t modify the “day” to “1”. | WIP25.4 |  |  | Subsystem | SP2_Candidate | Console-SRS.18077 |  | veeraswamy_k | 0 | 3 |  | Service Tools | NO | Assigned | SP2 | Brilliance Big Bore
CTUDT00631024 | Refresh | Assigned | elavarasan_s | September 4, 2020 11:08:27 PM GMT+08:00 | ePHI data is not scrubbed in Bugrep for CIRS Mdu and ExamRecordsRepository XML.gz file 1.Patient birth time in CIRS mdu didn't round off to nearest 10 minutes time 2.Study date in ExamRecordsRepository XML.gz didn’t modify the “day” to “1”. | WIP25.4 |  |  | Subsystem | SP2_Candidate | Console-SRS.18077 |  | veeraswamy_k | 0 | 3 |  | Service Tools | NO | Assigned | SP2 | Ingenuity 64
CTUDT00631024 | Refresh | Assigned | elavarasan_s | September 4, 2020 11:08:27 PM GMT+08:00 | ePHI data is not scrubbed in Bugrep for CIRS Mdu and ExamRecordsRepository XML.gz file 1.Patient birth time in CIRS mdu didn't round off to nearest 10 minutes time 2.Study date in ExamRecordsRepository XML.gz didn’t modify the “day” to “1”. | WIP25.4 |  |  | Subsystem | SP2_Candidate | Console-SRS.18077 |  | veeraswamy_k | 0 | 3 |  | Service Tools | NO | Assigned | SP2 | IQon1
CTUDT00631022 | Refresh | Closed | gT_s | September 4, 2020 10:28:16 PM GMT+08:00 | Refresh : The patient details are getting duplicated multiple times in the Patient directory | WIP 25.4 |  | Performing Rebuild operation | Subsystem |  | NA | No Plan to Fix | srushti_k | 0 | 3 |  | iCAP | NO | Closed |  | Ovation
CTUDT00631022 | Refresh | Closed | gT_s | September 4, 2020 10:28:16 PM GMT+08:00 | Refresh : The patient details are getting duplicated multiple times in the Patient directory | WIP 25.4 |  | Performing Rebuild operation | Subsystem |  | NA | No Plan to Fix | srushti_k | 0 | 3 |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00631022 | Refresh | Closed | gT_s | September 4, 2020 10:28:16 PM GMT+08:00 | Refresh : The patient details are getting duplicated multiple times in the Patient directory | WIP 25.4 |  | Performing Rebuild operation | Subsystem |  | NA | No Plan to Fix | srushti_k | 0 | 3 |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00631022 | Refresh | Closed | gT_s | September 4, 2020 10:28:16 PM GMT+08:00 | Refresh : The patient details are getting duplicated multiple times in the Patient directory | WIP 25.4 |  | Performing Rebuild operation | Subsystem |  | NA | No Plan to Fix | srushti_k | 0 | 3 |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00631022 | Refresh | Closed | gT_s | September 4, 2020 10:28:16 PM GMT+08:00 | Refresh : The patient details are getting duplicated multiple times in the Patient directory | WIP 25.4 |  | Performing Rebuild operation | Subsystem |  | NA | No Plan to Fix | srushti_k | 0 | 3 |  | iCAP | NO | Closed |  | IQon1
CTUDT00631018 | Refresh | Opened | vinoliya_j | September 4, 2020 2:54:56 PM GMT+08:00 | Refresh : Quick reference guide needs an update under section "Scan procedure" | WIP30.1 |  | N/A | Subsystem |  | N/A |  | rajalakshmi_j | 0 | 3 |  | Support-System | NO | Opened | SP1-Mar-21 | Ovation
CTUDT00631018 | Refresh | Opened | vinoliya_j | September 4, 2020 2:54:56 PM GMT+08:00 | Refresh : Quick reference guide needs an update under section "Scan procedure" | WIP30.1 |  | N/A | Subsystem |  | N/A |  | rajalakshmi_j | 0 | 3 |  | Support-System | NO | Opened | SP1-Mar-21 | Brilliance 64
CTUDT00631018 | Refresh | Opened | vinoliya_j | September 4, 2020 2:54:56 PM GMT+08:00 | Refresh : Quick reference guide needs an update under section "Scan procedure" | WIP30.1 |  | N/A | Subsystem |  | N/A |  | rajalakshmi_j | 0 | 3 |  | Support-System | NO | Opened | SP1-Mar-21 | Brilliance Big Bore
CTUDT00631018 | Refresh | Opened | vinoliya_j | September 4, 2020 2:54:56 PM GMT+08:00 | Refresh : Quick reference guide needs an update under section "Scan procedure" | WIP30.1 |  | N/A | Subsystem |  | N/A |  | rajalakshmi_j | 0 | 3 |  | Support-System | NO | Opened | SP1-Mar-21 | Ingenuity 64
CTUDT00631018 | Refresh | Opened | vinoliya_j | September 4, 2020 2:54:56 PM GMT+08:00 | Refresh : Quick reference guide needs an update under section "Scan procedure" | WIP30.1 |  | N/A | Subsystem |  | N/A |  | rajalakshmi_j | 0 | 3 |  | Support-System | NO | Opened | SP1-Mar-21 | IQon1
CTUDT00631016 | Refresh | Resolved | bingxin_lin | September 4, 2020 1:23:27 PM GMT+08:00 | Some service tools items are not displayed correctly under Olevel user | WIP31 | WIP34 |  | Subsystem | SP1_PM_Approved | NA | Fixed | chakraborty_d | 0 | 2 | Service Tools_Software | Console4 | NO | Resolved | WIP34SubCom | Brilliance Big Bore |  | WIP34 | November 5, 2020 2:53:14 PM GMT+08:00
CTUDT00631013 | Refresh | Accepted | cai_t | September 4, 2020 9:59:00 AM GMT+08:00 | [Refresh] :The system always popups ErrorExamApplication warning after loading an ExamCard. | WIP25.4 |  |  | Subsystem | SP2_Candidate | N/A |  | modi_s | 0 |  |  | Console4 | NO | Accepted |  | Brilliance Big Bore
CTUDT00631010 | Refresh | Accepted | zhong_x | September 4, 2020 9:03:06 AM GMT+08:00 | PatientDirectory.exe is not responding message pop up when save display transferred Icon Setting | WIP31 |  |  | Subsystem | SP2_Candidate |  |  | periasamy_g | 0 |  |  | Console4 | NO | Accepted |  | Brilliance Big Bore
CTUDT00630991 | Refresh | Opened | s_namburi | September 3, 2020 1:46:45 PM GMT+08:00 | FoK: Installation is non standard shown in installation validation tool. | 4.1.10 |  |  | Subsystem | FoK; SP1_PM_Approved | NIL | Fixed | chandran_a | 0 | 4 | Console Software | Console4 | NO | Opened | WIP37SegReq | Ovation | PR250834 PR268516
CTUDT00630991 | Refresh | Opened | s_namburi | September 3, 2020 1:46:45 PM GMT+08:00 | FoK: Installation is non standard shown in installation validation tool. | 4.1.10 |  |  | Subsystem | FoK; SP1_PM_Approved | NIL | Fixed | chandran_a | 0 | 4 | Console Software | Console4 | NO | Opened | WIP37SegReq | Ingenuity 64 | PR250834 PR268516
CTUDT00630989 | Refresh | Closed | lakshmikanth_s | September 3, 2020 12:32:34 PM GMT+08:00 | Change Password string is not displayed | WIP24.2 |  |  | Subsystem | SP2_Score2Postpone | na | Works As Specified / Not A Defect | kumar_gau | 1 |  |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00630989 | Refresh | Closed | lakshmikanth_s | September 3, 2020 12:32:34 PM GMT+08:00 | Change Password string is not displayed | WIP24.2 |  |  | Subsystem | SP2_Score2Postpone | na | Works As Specified / Not A Defect | kumar_gau | 1 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00630989 | Refresh | Closed | lakshmikanth_s | September 3, 2020 12:32:34 PM GMT+08:00 | Change Password string is not displayed | WIP24.2 |  |  | Subsystem | SP2_Score2Postpone | na | Works As Specified / Not A Defect | kumar_gau | 1 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00630985 | Refresh | Closed | gT_s | September 2, 2020 11:43:52 PM GMT+08:00 | Refresh : The € euro symbol is displayed as question mark ? in the application | WIP25.2 |  |  | Subsystem | InvestigateP2 | Console-SRS: 15526 | Duplicate | sannihith_r | 0 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00630985 | Refresh | Closed | gT_s | September 2, 2020 11:43:52 PM GMT+08:00 | Refresh : The € euro symbol is displayed as question mark ? in the application | WIP25.2 |  |  | Subsystem | InvestigateP2 | Console-SRS: 15526 | Duplicate | sannihith_r | 0 |  |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00630985 | Refresh | Closed | gT_s | September 2, 2020 11:43:52 PM GMT+08:00 | Refresh : The € euro symbol is displayed as question mark ? in the application | WIP25.2 |  |  | Subsystem | InvestigateP2 | Console-SRS: 15526 | Duplicate | sannihith_r | 0 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00630985 | Refresh | Closed | gT_s | September 2, 2020 11:43:52 PM GMT+08:00 | Refresh : The € euro symbol is displayed as question mark ? in the application | WIP25.2 |  |  | Subsystem | InvestigateP2 | Console-SRS: 15526 | Duplicate | sannihith_r | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00630985 | Refresh | Closed | gT_s | September 2, 2020 11:43:52 PM GMT+08:00 | Refresh : The € euro symbol is displayed as question mark ? in the application | WIP25.2 |  |  | Subsystem | InvestigateP2 | Console-SRS: 15526 | Duplicate | sannihith_r | 0 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00630982 | Refresh | Closed | portal_l | September 2, 2020 6:19:22 PM GMT+08:00 | Ict-SP- Service Tolerance CTDIfree-air % for case 16 in Spec which is 15 in the machine are different, the machine has limit 30% while the Spec requires limit of 25% | WIP 31 | WIP33 |  | Subsystem | SP1_PM_Approved | n/a | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP33SubCom | Ovation |  | WIP33 | October 29, 2020 8:04:23 PM GMT+08:00
CTUDT00630977 | Refresh | Closed | swack_g | September 2, 2020 3:02:11 PM GMT+08:00 | Refresh: iCT: Request for change: remove the checking of cracks on small and medium wedges to prevent the false error messages to the CT user. | WIP31 | WIP33 |  | Subsystem | SP1_PM_Approved | n/a | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP33SubCom | Ovation |  | WIP33 | October 29, 2020 4:58:24 PM GMT+08:00
CTUDT00630971 | Refresh | Closed | lin_s | September 1, 2020 7:42:17 PM GMT+08:00 | FOK: During IQ check, scanner stated no phantom was detected, but actually the phantom is on the scout. | Rfd |  |  | System | FOK |  | Works As Specified / Not A Defect | braunstein_d | 0 |  |  | Support-Physics | NO | Closed |  | Ingenuity 64 | 249299
CTUDT00630969 | Refresh | Resolved | lin_s | September 1, 2020 7:24:26 PM GMT+08:00 | FOK: Mt Grant Hospital Routine Air Cal Failure Cracked Compensator message | 4.1.10 |  |  | System | FOK | N/A | Change Proposal | yeshan_m | 0 | 2 | DMS Calibrations | Support-System | NO | Resolved | MTS, WI new vesion | Ovation | 247224 |  | January 4, 2021 1:49:34 PM GMT+08:00
CTUDT00630969 | Refresh | Resolved | lin_s | September 1, 2020 7:24:26 PM GMT+08:00 | FOK: Mt Grant Hospital Routine Air Cal Failure Cracked Compensator message | 4.1.10 |  |  | System | FOK | N/A | Change Proposal | yeshan_m | 0 | 2 | DMS Calibrations | Support-System | NO | Resolved | MTS, WI new vesion | Brilliance 64 | 247224 |  | January 4, 2021 1:49:34 PM GMT+08:00
CTUDT00630969 | Refresh | Resolved | lin_s | September 1, 2020 7:24:26 PM GMT+08:00 | FOK: Mt Grant Hospital Routine Air Cal Failure Cracked Compensator message | 4.1.10 |  |  | System | FOK | N/A | Change Proposal | yeshan_m | 0 | 2 | DMS Calibrations | Support-System | NO | Resolved | MTS, WI new vesion | Brilliance Big Bore | 247224 |  | January 4, 2021 1:49:34 PM GMT+08:00
CTUDT00630969 | Refresh | Resolved | lin_s | September 1, 2020 7:24:26 PM GMT+08:00 | FOK: Mt Grant Hospital Routine Air Cal Failure Cracked Compensator message | 4.1.10 |  |  | System | FOK | N/A | Change Proposal | yeshan_m | 0 | 2 | DMS Calibrations | Support-System | NO | Resolved | MTS, WI new vesion | Ingenuity 64 | 247224 |  | January 4, 2021 1:49:34 PM GMT+08:00
CTUDT00630969 | Refresh | Resolved | lin_s | September 1, 2020 7:24:26 PM GMT+08:00 | FOK: Mt Grant Hospital Routine Air Cal Failure Cracked Compensator message | 4.1.10 |  |  | System | FOK | N/A | Change Proposal | yeshan_m | 0 | 2 | DMS Calibrations | Support-System | NO | Resolved | MTS, WI new vesion | IQon1 | 247224 |  | January 4, 2021 1:49:34 PM GMT+08:00
CTUDT00630964 | Refresh | Closed | altshuller_mark | September 1, 2020 3:30:06 PM GMT+08:00 | FoK: Transfer icon didn't showed on some series that already send to PACS | 4.7.7 | WIP32 |  | Subsystem | SP1_PM_Approved | N/A | Fixed | pandey_a | 0 | 4 | Console Software | Console4 | NO | Closed | WIP32SubCom | Ovation | PR217800 | WIP32 | October 7, 2020 1:36:50 PM GMT+08:00
CTUDT00630964 | Refresh | Closed | altshuller_mark | September 1, 2020 3:30:06 PM GMT+08:00 | FoK: Transfer icon didn't showed on some series that already send to PACS | 4.7.7 | WIP32 |  | Subsystem | SP1_PM_Approved | N/A | Fixed | pandey_a | 0 | 4 | Console Software | Console4 | NO | Closed | WIP32SubCom | Brilliance 64 | PR217800 | WIP32 | October 7, 2020 1:36:50 PM GMT+08:00
CTUDT00630964 | Refresh | Closed | altshuller_mark | September 1, 2020 3:30:06 PM GMT+08:00 | FoK: Transfer icon didn't showed on some series that already send to PACS | 4.7.7 | WIP32 |  | Subsystem | SP1_PM_Approved | N/A | Fixed | pandey_a | 0 | 4 | Console Software | Console4 | NO | Closed | WIP32SubCom | Brilliance Big Bore | PR217800 | WIP32 | October 7, 2020 1:36:50 PM GMT+08:00
CTUDT00630964 | Refresh | Closed | altshuller_mark | September 1, 2020 3:30:06 PM GMT+08:00 | FoK: Transfer icon didn't showed on some series that already send to PACS | 4.7.7 | WIP32 |  | Subsystem | SP1_PM_Approved | N/A | Fixed | pandey_a | 0 | 4 | Console Software | Console4 | NO | Closed | WIP32SubCom | Ingenuity 64 | PR217800 | WIP32 | October 7, 2020 1:36:50 PM GMT+08:00
CTUDT00630964 | Refresh | Closed | altshuller_mark | September 1, 2020 3:30:06 PM GMT+08:00 | FoK: Transfer icon didn't showed on some series that already send to PACS | 4.7.7 | WIP32 |  | Subsystem | SP1_PM_Approved | N/A | Fixed | pandey_a | 0 | 4 | Console Software | Console4 | NO | Closed | WIP32SubCom | IQon1 | PR217800 | WIP32 | October 7, 2020 1:36:50 PM GMT+08:00
CTUDT00630964 | Refresh | Closed | altshuller_mark | September 1, 2020 3:30:06 PM GMT+08:00 | FoK: Transfer icon didn't showed on some series that already send to PACS | 4.7.7 | WIP32 |  | Subsystem | SP1_PM_Approved | N/A | Fixed | pandey_a | 0 | 4 | Console Software | Console4 | NO | Closed | WIP32SubCom | Hawk | PR217800 | WIP32 | October 7, 2020 1:36:50 PM GMT+08:00
CTUDT00630962 | Refresh | Closed | bingxin_lin | September 1, 2020 2:54:26 PM GMT+08:00 | CTDI Free Air(AirCTDI_8) test failed during H phantom calibration | WIP31 |  |  | Subsystem |  |  | Invalid/Uncontrolled Configuration | lin_s | 1 |  |  | Support-System | NO | Closed | TBR | Brilliance Big Bore
CTUDT00630961 | Refresh | Closed | altshuller_mark | August 31, 2020 7:42:15 PM GMT+08:00 | IQon High Frequency Suspend/Resume Improvement | CIRS 208/209 version | CIRS-5.1.0.0210 | 1. Re-boot the computer in order to recover the missing images 2. Site could upgrade to a 6 server configuration 3. Reduce the number of SBIs or adjust workflow to reduce the number of suspends/resumes (waiting for recons to finish before moving to next patient). | Subsystem | PM_Approved | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_SuprNodeSvr | CIRS | YES | Closed | WIP32SubCom | IQon1 | PR191940 | CIRS-5.1.0.0210 | September 12, 2020 3:53:21 AM GMT+08:00
CTUDT00630960 | Refresh | Closed | shaham_t | August 31, 2020 5:56:19 PM GMT+08:00 | FoK: Exam Manager has stopped working error ( Auto recon image didn't send to PACS and create duplicate series) | 4.7.7 | WIP32 |  | Subsystem | SP1_PM_Approved | N/A | Fixed | pandey_a | 0 | 3 | Console Software | Console4 | NO | Closed | WIP32SubCom | Ovation | PR217257 | WIP32 | October 7, 2020 1:35:56 PM GMT+08:00
CTUDT00630960 | Refresh | Closed | shaham_t | August 31, 2020 5:56:19 PM GMT+08:00 | FoK: Exam Manager has stopped working error ( Auto recon image didn't send to PACS and create duplicate series) | 4.7.7 | WIP32 |  | Subsystem | SP1_PM_Approved | N/A | Fixed | pandey_a | 0 | 3 | Console Software | Console4 | NO | Closed | WIP32SubCom | Brilliance 64 | PR217257 | WIP32 | October 7, 2020 1:35:56 PM GMT+08:00
CTUDT00630960 | Refresh | Closed | shaham_t | August 31, 2020 5:56:19 PM GMT+08:00 | FoK: Exam Manager has stopped working error ( Auto recon image didn't send to PACS and create duplicate series) | 4.7.7 | WIP32 |  | Subsystem | SP1_PM_Approved | N/A | Fixed | pandey_a | 0 | 3 | Console Software | Console4 | NO | Closed | WIP32SubCom | Brilliance Big Bore | PR217257 | WIP32 | October 7, 2020 1:35:56 PM GMT+08:00
CTUDT00630960 | Refresh | Closed | shaham_t | August 31, 2020 5:56:19 PM GMT+08:00 | FoK: Exam Manager has stopped working error ( Auto recon image didn't send to PACS and create duplicate series) | 4.7.7 | WIP32 |  | Subsystem | SP1_PM_Approved | N/A | Fixed | pandey_a | 0 | 3 | Console Software | Console4 | NO | Closed | WIP32SubCom | Ingenuity 64 | PR217257 | WIP32 | October 7, 2020 1:35:56 PM GMT+08:00
CTUDT00630960 | Refresh | Closed | shaham_t | August 31, 2020 5:56:19 PM GMT+08:00 | FoK: Exam Manager has stopped working error ( Auto recon image didn't send to PACS and create duplicate series) | 4.7.7 | WIP32 |  | Subsystem | SP1_PM_Approved | N/A | Fixed | pandey_a | 0 | 3 | Console Software | Console4 | NO | Closed | WIP32SubCom | IQon1 | PR217257 | WIP32 | October 7, 2020 1:35:56 PM GMT+08:00
CTUDT00630958 | Refresh | Resolved | altshuller_mark | August 31, 2020 1:46:48 PM GMT+08:00 | Image transfer to PACS problem | IQon HB1 | WIP33 | As a workaround , perform spooler reset\restart or restart the workflow | Subsystem | SP1_PM_Approved | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Resolved | WIP33SubCom | Ovation | Escalation #0115421125 | WIP33 | October 23, 2020 4:40:57 PM GMT+08:00
CTUDT00630958 | Refresh | Resolved | altshuller_mark | August 31, 2020 1:46:48 PM GMT+08:00 | Image transfer to PACS problem | IQon HB1 | WIP33 | As a workaround , perform spooler reset\restart or restart the workflow | Subsystem | SP1_PM_Approved | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Resolved | WIP33SubCom | Brilliance 64 | Escalation #0115421125 | WIP33 | October 23, 2020 4:40:57 PM GMT+08:00
CTUDT00630958 | Refresh | Resolved | altshuller_mark | August 31, 2020 1:46:48 PM GMT+08:00 | Image transfer to PACS problem | IQon HB1 | WIP33 | As a workaround , perform spooler reset\restart or restart the workflow | Subsystem | SP1_PM_Approved | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Resolved | WIP33SubCom | Brilliance Big Bore | Escalation #0115421125 | WIP33 | October 23, 2020 4:40:57 PM GMT+08:00
CTUDT00630958 | Refresh | Resolved | altshuller_mark | August 31, 2020 1:46:48 PM GMT+08:00 | Image transfer to PACS problem | IQon HB1 | WIP33 | As a workaround , perform spooler reset\restart or restart the workflow | Subsystem | SP1_PM_Approved | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Resolved | WIP33SubCom | Ingenuity 64 | Escalation #0115421125 | WIP33 | October 23, 2020 4:40:57 PM GMT+08:00
CTUDT00630958 | Refresh | Resolved | altshuller_mark | August 31, 2020 1:46:48 PM GMT+08:00 | Image transfer to PACS problem | IQon HB1 | WIP33 | As a workaround , perform spooler reset\restart or restart the workflow | Subsystem | SP1_PM_Approved | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Resolved | WIP33SubCom | IQon1 | Escalation #0115421125 | WIP33 | October 23, 2020 4:40:57 PM GMT+08:00
CTUDT00630926 | Refresh | Closed | dongping_w | August 27, 2020 3:16:46 PM GMT+08:00 | FOK:BB The customer is pleased the buttons are illuminated and have annotations describing function under the buttons. | WIP12.6 |  |  | System | FOK |  | Works As Specified / Not A Defect | tiwari_g | 0 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore | 241627
CTUDT00630923 | Refresh | Closed | dongping_w | August 27, 2020 3:08:16 PM GMT+08:00 | FOK:BB The customer would like to add two features to the Motion section of the illuminated screen portion of the control panel on the Interventional Control Control Console (ICC) | WIP12.6 |  |  | System | FOK |  | Change Proposal | tiwari_g | 0 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore | 241611
CTUDT00630922 | Refresh | Closed | lakshmikanth_s | August 27, 2020 3:05:07 PM GMT+08:00 | “&” character is allowed for user name. | WIP24.1 |  |  | Subsystem |  | N/A | Works As Specified / Not A Defect | juster_d | 0 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00630919 | Refresh | Assigned | shaham_t | August 27, 2020 2:30:45 PM GMT+08:00 | Refresh Fok: Gap between the new foot pedal and the gantry | 4.1.10 |  | 1.Move couch and gantry closer to each other if possible to decrease the distance between Couch and Gantry; 2.Add additional spacer or sponge to fill the gap. (Preferred) | Subsystem | FoK; LCM_Candidate ; SP2_Candidate | N/A |  | wang_ada | 0 | 2 |  | Couch | YES | Assigned | Design change | Ovation | PR243477 PR250832
CTUDT00630919 | Refresh | Assigned | shaham_t | August 27, 2020 2:30:45 PM GMT+08:00 | Refresh Fok: Gap between the new foot pedal and the gantry | 4.1.10 |  | 1.Move couch and gantry closer to each other if possible to decrease the distance between Couch and Gantry; 2.Add additional spacer or sponge to fill the gap. (Preferred) | Subsystem | FoK; LCM_Candidate ; SP2_Candidate | N/A |  | wang_ada | 0 | 2 |  | Couch | YES | Assigned | Design change | Brilliance 64 | PR243477 PR250832
CTUDT00630919 | Refresh | Assigned | shaham_t | August 27, 2020 2:30:45 PM GMT+08:00 | Refresh Fok: Gap between the new foot pedal and the gantry | 4.1.10 |  | 1.Move couch and gantry closer to each other if possible to decrease the distance between Couch and Gantry; 2.Add additional spacer or sponge to fill the gap. (Preferred) | Subsystem | FoK; LCM_Candidate ; SP2_Candidate | N/A |  | wang_ada | 0 | 2 |  | Couch | YES | Assigned | Design change | Brilliance Big Bore | PR243477 PR250832
CTUDT00630919 | Refresh | Assigned | shaham_t | August 27, 2020 2:30:45 PM GMT+08:00 | Refresh Fok: Gap between the new foot pedal and the gantry | 4.1.10 |  | 1.Move couch and gantry closer to each other if possible to decrease the distance between Couch and Gantry; 2.Add additional spacer or sponge to fill the gap. (Preferred) | Subsystem | FoK; LCM_Candidate ; SP2_Candidate | N/A |  | wang_ada | 0 | 2 |  | Couch | YES | Assigned | Design change | IQon1 | PR243477 PR250832
CTUDT00630918 | Refresh | Accepted | lin_s | August 27, 2020 2:24:16 PM GMT+08:00 | Automatic Acceptance does not comply with 61223-3-5, edition 2.0. | WIP24.1 |  |  | Subsystem |  |  |  | rajalakshmi_j | 0 | 2 | DMS Calibrations | Support-System | NO | Accepted | SP1-Mar-21 | Ovation
CTUDT00630918 | Refresh | Accepted | lin_s | August 27, 2020 2:24:16 PM GMT+08:00 | Automatic Acceptance does not comply with 61223-3-5, edition 2.0. | WIP24.1 |  |  | Subsystem |  |  |  | rajalakshmi_j | 0 | 2 | DMS Calibrations | Support-System | NO | Accepted | SP1-Mar-21 | Brilliance Big Bore
CTUDT00630918 | Refresh | Accepted | lin_s | August 27, 2020 2:24:16 PM GMT+08:00 | Automatic Acceptance does not comply with 61223-3-5, edition 2.0. | WIP24.1 |  |  | Subsystem |  |  |  | rajalakshmi_j | 0 | 2 | DMS Calibrations | Support-System | NO | Accepted | SP1-Mar-21 | Ingenuity 64
CTUDT00630918 | Refresh | Accepted | lin_s | August 27, 2020 2:24:16 PM GMT+08:00 | Automatic Acceptance does not comply with 61223-3-5, edition 2.0. | WIP24.1 |  |  | Subsystem |  |  |  | rajalakshmi_j | 0 | 2 | DMS Calibrations | Support-System | NO | Accepted | SP1-Mar-21 | IQon1
CTUDT00630917 | Refresh | Closed | dongping_w | August 27, 2020 2:23:44 PM GMT+08:00 | FOK: BB The customer would like the middle mouse button of IVC on the right and left sides of the mouse. | WIP12.6 |  |  | System | FOK |  | Change Proposal | tiwari_g | 0 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore | 239270
CTUDT00630912 | Refresh | Closed | dongping_w | August 27, 2020 2:14:30 PM GMT+08:00 | FOK: BB Customer would like to turn the laser on and off within the interventional scan ruler using mouse click | WIP12.6 |  |  | System | FOK |  | Change Proposal | tiwari_g | 0 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore | 241653
CTUDT00630910 | Refresh | Closed | allaberdiev_v | August 27, 2020 1:16:03 PM GMT+08:00 | Refresh: DHF170395_74: Harmonized Acceptance and Constancy HTML "According to" value was not according to spec. | WIP 31 | WIP33 | N/A | Subsystem | SP1_PM_Approved | N/A | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP33SubCom | Ovation |  | WIP33 | October 29, 2020 4:53:58 PM GMT+08:00
CTUDT00630910 | Refresh | Closed | allaberdiev_v | August 27, 2020 1:16:03 PM GMT+08:00 | Refresh: DHF170395_74: Harmonized Acceptance and Constancy HTML "According to" value was not according to spec. | WIP 31 | WIP33 | N/A | Subsystem | SP1_PM_Approved | N/A | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP33SubCom | Brilliance 64 |  | WIP33 | October 29, 2020 4:53:58 PM GMT+08:00
CTUDT00630910 | Refresh | Closed | allaberdiev_v | August 27, 2020 1:16:03 PM GMT+08:00 | Refresh: DHF170395_74: Harmonized Acceptance and Constancy HTML "According to" value was not according to spec. | WIP 31 | WIP33 | N/A | Subsystem | SP1_PM_Approved | N/A | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP33SubCom | Brilliance Big Bore |  | WIP33 | October 29, 2020 4:53:58 PM GMT+08:00
CTUDT00630910 | Refresh | Closed | allaberdiev_v | August 27, 2020 1:16:03 PM GMT+08:00 | Refresh: DHF170395_74: Harmonized Acceptance and Constancy HTML "According to" value was not according to spec. | WIP 31 | WIP33 | N/A | Subsystem | SP1_PM_Approved | N/A | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP33SubCom | Ingenuity 64 |  | WIP33 | October 29, 2020 4:53:58 PM GMT+08:00
CTUDT00630910 | Refresh | Closed | allaberdiev_v | August 27, 2020 1:16:03 PM GMT+08:00 | Refresh: DHF170395_74: Harmonized Acceptance and Constancy HTML "According to" value was not according to spec. | WIP 31 | WIP33 | N/A | Subsystem | SP1_PM_Approved | N/A | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP33SubCom | IQon1 |  | WIP33 | October 29, 2020 4:53:58 PM GMT+08:00
CTUDT00630908 | Refresh | Closed | lei_yang | August 27, 2020 12:23:09 PM GMT+08:00 | The 16th "AirCTDI_15 Pediatric Body" of "CTDI Free Air" test failed. | WIP31 | WIP34 |  | System | SP1_PM_Approved | N/A | Fixed | s_srivastava | 0 | 2 | Console Software | Console4 | NO | Closed | WIP34SubCom | Ingenuity 64 |  | WIP34 | November 12, 2020 8:56:47 PM GMT+08:00
CTUDT00630906 | Refresh | Closed | bingxin_lin | August 27, 2020 12:04:45 PM GMT+08:00 | ”Not all images for ‘Hcor Basic’ were created….”message popped up during the Hcor Basic test calibration | WIP31 | WIP33 | User can perform offline recon or re-start calibration. | Subsystem | SP1_PM_Approved | n/a | Fixed | richards_b | 0 | 3 | CIRS_SW_ImgNodeSvr | CIRS | NO | Closed | WIP33SubCom | Ovation |  | WIP33 | November 4, 2020 3:45:59 AM GMT+08:00
CTUDT00630906 | Refresh | Closed | bingxin_lin | August 27, 2020 12:04:45 PM GMT+08:00 | ”Not all images for ‘Hcor Basic’ were created….”message popped up during the Hcor Basic test calibration | WIP31 | WIP33 | User can perform offline recon or re-start calibration. | Subsystem | SP1_PM_Approved | n/a | Fixed | richards_b | 0 | 3 | CIRS_SW_ImgNodeSvr | CIRS | NO | Closed | WIP33SubCom | Brilliance 64 |  | WIP33 | November 4, 2020 3:45:59 AM GMT+08:00
CTUDT00630906 | Refresh | Closed | bingxin_lin | August 27, 2020 12:04:45 PM GMT+08:00 | ”Not all images for ‘Hcor Basic’ were created….”message popped up during the Hcor Basic test calibration | WIP31 | WIP33 | User can perform offline recon or re-start calibration. | Subsystem | SP1_PM_Approved | n/a | Fixed | richards_b | 0 | 3 | CIRS_SW_ImgNodeSvr | CIRS | NO | Closed | WIP33SubCom | Brilliance Big Bore |  | WIP33 | November 4, 2020 3:45:59 AM GMT+08:00
CTUDT00630906 | Refresh | Closed | bingxin_lin | August 27, 2020 12:04:45 PM GMT+08:00 | ”Not all images for ‘Hcor Basic’ were created….”message popped up during the Hcor Basic test calibration | WIP31 | WIP33 | User can perform offline recon or re-start calibration. | Subsystem | SP1_PM_Approved | n/a | Fixed | richards_b | 0 | 3 | CIRS_SW_ImgNodeSvr | CIRS | NO | Closed | WIP33SubCom | Ingenuity 64 |  | WIP33 | November 4, 2020 3:45:59 AM GMT+08:00
CTUDT00630906 | Refresh | Closed | bingxin_lin | August 27, 2020 12:04:45 PM GMT+08:00 | ”Not all images for ‘Hcor Basic’ were created….”message popped up during the Hcor Basic test calibration | WIP31 | WIP33 | User can perform offline recon or re-start calibration. | Subsystem | SP1_PM_Approved | n/a | Fixed | richards_b | 0 | 3 | CIRS_SW_ImgNodeSvr | CIRS | NO | Closed | WIP33SubCom | IQon1 |  | WIP33 | November 4, 2020 3:45:59 AM GMT+08:00
CTUDT00630900 | Refresh | Closed | portal_l | August 26, 2020 9:23:38 PM GMT+08:00 | Ingenuity 64 -CT number maximum limit of scan Body CT13c Air was different than Spec required. | WIP 31 | WIP33 | N/A | Subsystem | SP1_PM_Approved | N/A | Fixed | frider_o | 0 | 2 | DMS Software | DMS | YES | Closed | WIP33SubCom | Brilliance 64 |  | WIP33 | October 29, 2020 4:02:49 PM GMT+08:00
CTUDT00630900 | Refresh | Closed | portal_l | August 26, 2020 9:23:38 PM GMT+08:00 | Ingenuity 64 -CT number maximum limit of scan Body CT13c Air was different than Spec required. | WIP 31 | WIP33 | N/A | Subsystem | SP1_PM_Approved | N/A | Fixed | frider_o | 0 | 2 | DMS Software | DMS | YES | Closed | WIP33SubCom | Ingenuity 64 |  | WIP33 | October 29, 2020 4:02:49 PM GMT+08:00
CTUDT00630890 | Refresh | Resolved | cai_t | August 26, 2020 10:47:41 AM GMT+08:00 | [ACS]Couch jog1 horizontal moving exceeding electrical limit hit the mechanical limit, and then press the unload button Couch start vertical moving down | WIP25 | WIP36 | • How is the issue manifested? duiring the verification check the couch horizontal moving trajectory the trigger deceleration position not as designed check the software code find one type parameter from GINT32 to GUINT32 , cause a value compare error. because GINT32 value should change to GUINT32 when compare with GUINT32 value • What is the impact on the user, patient, service personnel, bystander or environment? Couch manual motion horizontal motion last moving trajectory not as design, but no impact to patient. stopped position normally as design before electrical limit. • What system/products are impacted? Sub-system? Software version? Other Software versions impacted? • What is the frequency of occurrence? How reproducible is the issue? N/A • Is there a possible workaround, can or did the user perform any action to complete the desired activity? If so, describe it. no need action, next version software will solve this little bug • Which requirements are being reviewed for this assessment? Are they being met? (Consider service, manufacturing, security and privacy requirements as well). DHF173447-1824: Couch shall include travel limit switches to stop over travel and uncommented motions. The requirement is met. Manual motion Unload function requirement:GCX.SSRS.iCT.16185, The requirement is not met. • Additionally, provide Software modules, Part numbers, or other design detail available: change the parameter type back to GINT32 the moving as designed. would solve this error. | Subsystem | SP1_PM_Approved | NA | Fixed | cx_zhang | 0 | 2 | Couch Software | Couch | NO | Resolved | WIP33SubComWIP36SubCom | Ovation |  | WIP36 | December 16, 2020 1:25:05 PM GMT+08:00
CTUDT00630890 | Refresh | Resolved | cai_t | August 26, 2020 10:47:41 AM GMT+08:00 | [ACS]Couch jog1 horizontal moving exceeding electrical limit hit the mechanical limit, and then press the unload button Couch start vertical moving down | WIP25 | WIP36 | • How is the issue manifested? duiring the verification check the couch horizontal moving trajectory the trigger deceleration position not as designed check the software code find one type parameter from GINT32 to GUINT32 , cause a value compare error. because GINT32 value should change to GUINT32 when compare with GUINT32 value • What is the impact on the user, patient, service personnel, bystander or environment? Couch manual motion horizontal motion last moving trajectory not as design, but no impact to patient. stopped position normally as design before electrical limit. • What system/products are impacted? Sub-system? Software version? Other Software versions impacted? • What is the frequency of occurrence? How reproducible is the issue? N/A • Is there a possible workaround, can or did the user perform any action to complete the desired activity? If so, describe it. no need action, next version software will solve this little bug • Which requirements are being reviewed for this assessment? Are they being met? (Consider service, manufacturing, security and privacy requirements as well). DHF173447-1824: Couch shall include travel limit switches to stop over travel and uncommented motions. The requirement is met. Manual motion Unload function requirement:GCX.SSRS.iCT.16185, The requirement is not met. • Additionally, provide Software modules, Part numbers, or other design detail available: change the parameter type back to GINT32 the moving as designed. would solve this error. | Subsystem | SP1_PM_Approved | NA | Fixed | cx_zhang | 0 | 2 | Couch Software | Couch | NO | Resolved | WIP33SubComWIP36SubCom | Brilliance 64 |  | WIP36 | December 16, 2020 1:25:05 PM GMT+08:00
CTUDT00630890 | Refresh | Resolved | cai_t | August 26, 2020 10:47:41 AM GMT+08:00 | [ACS]Couch jog1 horizontal moving exceeding electrical limit hit the mechanical limit, and then press the unload button Couch start vertical moving down | WIP25 | WIP36 | • How is the issue manifested? duiring the verification check the couch horizontal moving trajectory the trigger deceleration position not as designed check the software code find one type parameter from GINT32 to GUINT32 , cause a value compare error. because GINT32 value should change to GUINT32 when compare with GUINT32 value • What is the impact on the user, patient, service personnel, bystander or environment? Couch manual motion horizontal motion last moving trajectory not as design, but no impact to patient. stopped position normally as design before electrical limit. • What system/products are impacted? Sub-system? Software version? Other Software versions impacted? • What is the frequency of occurrence? How reproducible is the issue? N/A • Is there a possible workaround, can or did the user perform any action to complete the desired activity? If so, describe it. no need action, next version software will solve this little bug • Which requirements are being reviewed for this assessment? Are they being met? (Consider service, manufacturing, security and privacy requirements as well). DHF173447-1824: Couch shall include travel limit switches to stop over travel and uncommented motions. The requirement is met. Manual motion Unload function requirement:GCX.SSRS.iCT.16185, The requirement is not met. • Additionally, provide Software modules, Part numbers, or other design detail available: change the parameter type back to GINT32 the moving as designed. would solve this error. | Subsystem | SP1_PM_Approved | NA | Fixed | cx_zhang | 0 | 2 | Couch Software | Couch | NO | Resolved | WIP33SubComWIP36SubCom | Brilliance Big Bore |  | WIP36 | December 16, 2020 1:25:05 PM GMT+08:00
CTUDT00630890 | Refresh | Resolved | cai_t | August 26, 2020 10:47:41 AM GMT+08:00 | [ACS]Couch jog1 horizontal moving exceeding electrical limit hit the mechanical limit, and then press the unload button Couch start vertical moving down | WIP25 | WIP36 | • How is the issue manifested? duiring the verification check the couch horizontal moving trajectory the trigger deceleration position not as designed check the software code find one type parameter from GINT32 to GUINT32 , cause a value compare error. because GINT32 value should change to GUINT32 when compare with GUINT32 value • What is the impact on the user, patient, service personnel, bystander or environment? Couch manual motion horizontal motion last moving trajectory not as design, but no impact to patient. stopped position normally as design before electrical limit. • What system/products are impacted? Sub-system? Software version? Other Software versions impacted? • What is the frequency of occurrence? How reproducible is the issue? N/A • Is there a possible workaround, can or did the user perform any action to complete the desired activity? If so, describe it. no need action, next version software will solve this little bug • Which requirements are being reviewed for this assessment? Are they being met? (Consider service, manufacturing, security and privacy requirements as well). DHF173447-1824: Couch shall include travel limit switches to stop over travel and uncommented motions. The requirement is met. Manual motion Unload function requirement:GCX.SSRS.iCT.16185, The requirement is not met. • Additionally, provide Software modules, Part numbers, or other design detail available: change the parameter type back to GINT32 the moving as designed. would solve this error. | Subsystem | SP1_PM_Approved | NA | Fixed | cx_zhang | 0 | 2 | Couch Software | Couch | NO | Resolved | WIP33SubComWIP36SubCom | Ingenuity 64 |  | WIP36 | December 16, 2020 1:25:05 PM GMT+08:00
CTUDT00630890 | Refresh | Resolved | cai_t | August 26, 2020 10:47:41 AM GMT+08:00 | [ACS]Couch jog1 horizontal moving exceeding electrical limit hit the mechanical limit, and then press the unload button Couch start vertical moving down | WIP25 | WIP36 | • How is the issue manifested? duiring the verification check the couch horizontal moving trajectory the trigger deceleration position not as designed check the software code find one type parameter from GINT32 to GUINT32 , cause a value compare error. because GINT32 value should change to GUINT32 when compare with GUINT32 value • What is the impact on the user, patient, service personnel, bystander or environment? Couch manual motion horizontal motion last moving trajectory not as design, but no impact to patient. stopped position normally as design before electrical limit. • What system/products are impacted? Sub-system? Software version? Other Software versions impacted? • What is the frequency of occurrence? How reproducible is the issue? N/A • Is there a possible workaround, can or did the user perform any action to complete the desired activity? If so, describe it. no need action, next version software will solve this little bug • Which requirements are being reviewed for this assessment? Are they being met? (Consider service, manufacturing, security and privacy requirements as well). DHF173447-1824: Couch shall include travel limit switches to stop over travel and uncommented motions. The requirement is met. Manual motion Unload function requirement:GCX.SSRS.iCT.16185, The requirement is not met. • Additionally, provide Software modules, Part numbers, or other design detail available: change the parameter type back to GINT32 the moving as designed. would solve this error. | Subsystem | SP1_PM_Approved | NA | Fixed | cx_zhang | 0 | 2 | Couch Software | Couch | NO | Resolved | WIP33SubComWIP36SubCom | IQon1 |  | WIP36 | December 16, 2020 1:25:05 PM GMT+08:00
CTUDT00630885 | Refresh | Closed | shaham_t | August 25, 2020 11:36:26 PM GMT+08:00 | SRS update required based on new CTDI values | WIP31 | DHF298556 Rev.16 DHF298553 Rev.17 |  | Subsystem | SP1_PM_Approved | CT-NM.SRS.IQ.1728; CT-NM.SRS.IQ.3137; CT-NM.SRS.IQ.3104 | Fixed | shaham_t | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | SRS | Ovation |  | DHF298556 Rev.16 DHF298553 Rev.17 | October 6, 2020 7:22:41 PM GMT+08:00
CTUDT00630885 | Refresh | Closed | shaham_t | August 25, 2020 11:36:26 PM GMT+08:00 | SRS update required based on new CTDI values | WIP31 | DHF298556 Rev.16 DHF298553 Rev.17 |  | Subsystem | SP1_PM_Approved | CT-NM.SRS.IQ.1728; CT-NM.SRS.IQ.3137; CT-NM.SRS.IQ.3104 | Fixed | shaham_t | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | SRS | IQon1 |  | DHF298556 Rev.16 DHF298553 Rev.17 | October 6, 2020 7:22:41 PM GMT+08:00
CTUDT00630881 | Refresh | Accepted | lakshmikanth_s | August 25, 2020 9:17:44 PM GMT+08:00 | On first logout from OLevel, new keyboard language is coming in login screen on a fresh installed system. It should be English | WIP24.2 |  |  | Subsystem | SP2_Score2Postpone |  |  | kumar_gau | 0 |  |  | Console4 | NO | Accepted |  | IQon1
CTUDT00630872 | Refresh | Postponed | swack_g | August 25, 2020 6:16:08 PM GMT+08:00 | Refresh: Help Pages not fully translated | WIP 31 |  |  | Subsystem |  | N/A |  | yatziv_l | 0 | 3 |  | DMS | NO | Postponed | SubNC | Ovation
CTUDT00630872 | Refresh | Postponed | swack_g | August 25, 2020 6:16:08 PM GMT+08:00 | Refresh: Help Pages not fully translated | WIP 31 |  |  | Subsystem |  | N/A |  | yatziv_l | 0 | 3 |  | DMS | NO | Postponed | SubNC | Brilliance 64
CTUDT00630872 | Refresh | Postponed | swack_g | August 25, 2020 6:16:08 PM GMT+08:00 | Refresh: Help Pages not fully translated | WIP 31 |  |  | Subsystem |  | N/A |  | yatziv_l | 0 | 3 |  | DMS | NO | Postponed | SubNC | Brilliance Big Bore
CTUDT00630872 | Refresh | Postponed | swack_g | August 25, 2020 6:16:08 PM GMT+08:00 | Refresh: Help Pages not fully translated | WIP 31 |  |  | Subsystem |  | N/A |  | yatziv_l | 0 | 3 |  | DMS | NO | Postponed | SubNC | Ingenuity 64
CTUDT00630872 | Refresh | Postponed | swack_g | August 25, 2020 6:16:08 PM GMT+08:00 | Refresh: Help Pages not fully translated | WIP 31 |  |  | Subsystem |  | N/A |  | yatziv_l | 0 | 3 |  | DMS | NO | Postponed | SubNC | IQon1
CTUDT00630871 | Refresh | Resolved | shaham_t | August 25, 2020 4:19:42 PM GMT+08:00 | FoK: Both the New Factory Exam Cards and User restored Exam Cards are present in User Group after Upgrade | 4.7.7 | WIP34 |  | Subsystem | SP1_PM_Approved | N/A | Fixed | venkatachalam_p | 0 | 2 | Service Documentation | Console4 | NO | Resolved | WIP34SubCom | Ovation | PR192844 PR203329 | WIP34 | November 12, 2020 6:29:47 PM GMT+08:00
CTUDT00630871 | Refresh | Resolved | shaham_t | August 25, 2020 4:19:42 PM GMT+08:00 | FoK: Both the New Factory Exam Cards and User restored Exam Cards are present in User Group after Upgrade | 4.7.7 | WIP34 |  | Subsystem | SP1_PM_Approved | N/A | Fixed | venkatachalam_p | 0 | 2 | Service Documentation | Console4 | NO | Resolved | WIP34SubCom | Brilliance 64 | PR192844 PR203329 | WIP34 | November 12, 2020 6:29:47 PM GMT+08:00
CTUDT00630871 | Refresh | Resolved | shaham_t | August 25, 2020 4:19:42 PM GMT+08:00 | FoK: Both the New Factory Exam Cards and User restored Exam Cards are present in User Group after Upgrade | 4.7.7 | WIP34 |  | Subsystem | SP1_PM_Approved | N/A | Fixed | venkatachalam_p | 0 | 2 | Service Documentation | Console4 | NO | Resolved | WIP34SubCom | Brilliance Big Bore | PR192844 PR203329 | WIP34 | November 12, 2020 6:29:47 PM GMT+08:00
CTUDT00630871 | Refresh | Resolved | shaham_t | August 25, 2020 4:19:42 PM GMT+08:00 | FoK: Both the New Factory Exam Cards and User restored Exam Cards are present in User Group after Upgrade | 4.7.7 | WIP34 |  | Subsystem | SP1_PM_Approved | N/A | Fixed | venkatachalam_p | 0 | 2 | Service Documentation | Console4 | NO | Resolved | WIP34SubCom | Ingenuity 64 | PR192844 PR203329 | WIP34 | November 12, 2020 6:29:47 PM GMT+08:00
CTUDT00630871 | Refresh | Resolved | shaham_t | August 25, 2020 4:19:42 PM GMT+08:00 | FoK: Both the New Factory Exam Cards and User restored Exam Cards are present in User Group after Upgrade | 4.7.7 | WIP34 |  | Subsystem | SP1_PM_Approved | N/A | Fixed | venkatachalam_p | 0 | 2 | Service Documentation | Console4 | NO | Resolved | WIP34SubCom | IQon1 | PR192844 PR203329 | WIP34 | November 12, 2020 6:29:47 PM GMT+08:00
CTUDT00630868 | Refresh | Closed | lin_s | August 25, 2020 12:18:36 PM GMT+08:00 | FOK: Unable to connect to worklist on customer site. | FOK |  |  | Subsystem |  |  | Works As Specified / Not A Defect | lin_s | 0 |  |  |  |  | Closed |  | Ingenuity 64 | 234535
CTUDT00630866 | Refresh | Accepted | lin_s | August 25, 2020 12:12:42 PM GMT+08:00 | FOK: Right rear gantry panel tilt LED cannot lit | FOK |  |  | Subsystem | FOK |  |  | ickes_d | 0 |  |  | Support-System | NO | Accepted |  | Ingenuity 64 | 152146
CTUDT00630864 | Refresh | Closed | lin_s | August 25, 2020 10:41:23 AM GMT+08:00 | FOK: System rotation issue when first installation on site. | FOK |  |  | System | FoK | NA | System Repaired | yaping_l | 0 |  |  | GCX | NO | Closed |  | Ingenuity 64 | 202722
CTUDT00630845 | Refresh | Accepted | lei_yang | August 24, 2020 4:16:03 PM GMT+08:00 | “ScannerFrameworkWinApp has stopped working” error message pops up after launching workflow app. | WIP25.2 |  |  | Subsystem | SP2_Candidate | NA |  | mukherjee_s | 0 | 4 |  | Console4 | NO | Accepted |  | Ovation
CTUDT00630845 | Refresh | Accepted | lei_yang | August 24, 2020 4:16:03 PM GMT+08:00 | “ScannerFrameworkWinApp has stopped working” error message pops up after launching workflow app. | WIP25.2 |  |  | Subsystem | SP2_Candidate | NA |  | mukherjee_s | 0 | 4 |  | Console4 | NO | Accepted |  | Brilliance 64
CTUDT00630845 | Refresh | Accepted | lei_yang | August 24, 2020 4:16:03 PM GMT+08:00 | “ScannerFrameworkWinApp has stopped working” error message pops up after launching workflow app. | WIP25.2 |  |  | Subsystem | SP2_Candidate | NA |  | mukherjee_s | 0 | 4 |  | Console4 | NO | Accepted |  | Brilliance Big Bore
CTUDT00630845 | Refresh | Accepted | lei_yang | August 24, 2020 4:16:03 PM GMT+08:00 | “ScannerFrameworkWinApp has stopped working” error message pops up after launching workflow app. | WIP25.2 |  |  | Subsystem | SP2_Candidate | NA |  | mukherjee_s | 0 | 4 |  | Console4 | NO | Accepted |  | Ingenuity 64
CTUDT00630845 | Refresh | Accepted | lei_yang | August 24, 2020 4:16:03 PM GMT+08:00 | “ScannerFrameworkWinApp has stopped working” error message pops up after launching workflow app. | WIP25.2 |  |  | Subsystem | SP2_Candidate | NA |  | mukherjee_s | 0 | 4 |  | Console4 | NO | Accepted |  | IQon1
CTUDT00630843 | Refresh | Closed | dongping_w | August 24, 2020 4:01:05 PM GMT+08:00 | FOk BB Bellows not working because of i-Box issue | WIP12.6 | WIP33 | For the bellows function abnormal issue, please power down/up the Gantry, reset the IBOX2 board and re-start the work flow. | Subsystem | FOK; SP1_PM_Approved | NA | Fixed | luo_h | 0 | 3 | Gantry Software | GCX | NO | Closed | WIP33SubCom | Brilliance 64 | 222240 | WIP33 | November 2, 2020 3:57:31 PM GMT+08:00
CTUDT00630843 | Refresh | Closed | dongping_w | August 24, 2020 4:01:05 PM GMT+08:00 | FOk BB Bellows not working because of i-Box issue | WIP12.6 | WIP33 | For the bellows function abnormal issue, please power down/up the Gantry, reset the IBOX2 board and re-start the work flow. | Subsystem | FOK; SP1_PM_Approved | NA | Fixed | luo_h | 0 | 3 | Gantry Software | GCX | NO | Closed | WIP33SubCom | Brilliance Big Bore | 222240 | WIP33 | November 2, 2020 3:57:31 PM GMT+08:00
CTUDT00630843 | Refresh | Closed | dongping_w | August 24, 2020 4:01:05 PM GMT+08:00 | FOk BB Bellows not working because of i-Box issue | WIP12.6 | WIP33 | For the bellows function abnormal issue, please power down/up the Gantry, reset the IBOX2 board and re-start the work flow. | Subsystem | FOK; SP1_PM_Approved | NA | Fixed | luo_h | 0 | 3 | Gantry Software | GCX | NO | Closed | WIP33SubCom | Ingenuity 64 | 222240 | WIP33 | November 2, 2020 3:57:31 PM GMT+08:00
CTUDT00630839 | Refresh | Closed | dongping_w | August 24, 2020 3:40:27 PM GMT+08:00 | FOK-BB 4D importing into MIM and Eclipse can not be completed automatically | WIP12.6 |  |  | Subsystem | FOK; InvestigateP2 | None | Change Proposal | sannihith_r | 0 | 3 |  | Console4 | NO | Closed |  | Brilliance Big Bore | 224675
CTUDT00630836 | Refresh | Closed | dongping_w | August 24, 2020 3:23:51 PM GMT+08:00 | FOK-BB Customer requests the ability to have more than 1 mapped drive when logged in as user | WIP12.6 |  |  | System | FOK | N/A | Change Proposal | atin_s | 0 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore | 224673
CTUDT00630833 | Refresh | Opened | s_namburi | August 24, 2020 3:14:47 PM GMT+08:00 | FoK: Screen display was freeze for a while. (time not counting) | 4.7.7 | WIP34 |  | System | SP1_PM_Approved; InvestigateP1 | N/A |  | s_srivastava | 0 | 2 | Console Software | Console4 | NO | Opened | WIP37SegReq | Ovation | PR217799 | WIP34 | November 5, 2020 8:43:53 PM GMT+08:00
CTUDT00630833 | Refresh | Opened | s_namburi | August 24, 2020 3:14:47 PM GMT+08:00 | FoK: Screen display was freeze for a while. (time not counting) | 4.7.7 | WIP34 |  | System | SP1_PM_Approved; InvestigateP1 | N/A |  | s_srivastava | 0 | 2 | Console Software | Console4 | NO | Opened | WIP37SegReq | Brilliance 64 | PR217799 | WIP34 | November 5, 2020 8:43:53 PM GMT+08:00
CTUDT00630833 | Refresh | Opened | s_namburi | August 24, 2020 3:14:47 PM GMT+08:00 | FoK: Screen display was freeze for a while. (time not counting) | 4.7.7 | WIP34 |  | System | SP1_PM_Approved; InvestigateP1 | N/A |  | s_srivastava | 0 | 2 | Console Software | Console4 | NO | Opened | WIP37SegReq | Brilliance Big Bore | PR217799 | WIP34 | November 5, 2020 8:43:53 PM GMT+08:00
CTUDT00630833 | Refresh | Opened | s_namburi | August 24, 2020 3:14:47 PM GMT+08:00 | FoK: Screen display was freeze for a while. (time not counting) | 4.7.7 | WIP34 |  | System | SP1_PM_Approved; InvestigateP1 | N/A |  | s_srivastava | 0 | 2 | Console Software | Console4 | NO | Opened | WIP37SegReq | Ingenuity 64 | PR217799 | WIP34 | November 5, 2020 8:43:53 PM GMT+08:00
CTUDT00630833 | Refresh | Opened | s_namburi | August 24, 2020 3:14:47 PM GMT+08:00 | FoK: Screen display was freeze for a while. (time not counting) | 4.7.7 | WIP34 |  | System | SP1_PM_Approved; InvestigateP1 | N/A |  | s_srivastava | 0 | 2 | Console Software | Console4 | NO | Opened | WIP37SegReq | IQon1 | PR217799 | WIP34 | November 5, 2020 8:43:53 PM GMT+08:00
CTUDT00630830 | Refresh | Assigned | dongping_w | August 24, 2020 2:47:59 PM GMT+08:00 | FOK-BB holder is very difficult to slide onto therapy table top | BB_WIP12.6 |  | No workaround is recommended. The issue does not affect the function and intended use. The design improvement will be implemented. Then the user will easily install the Phantom Holder in place. The design change will be initiated by ECR-076387 and implemented by E-129050. | Subsystem | FOK; LCM_Candidate; SP2_Candidate | N/A |  | wang_ada | 0 | 4 |  | Couch | NO | Assigned | Next Design | Ovation | 222254
CTUDT00630830 | Refresh | Assigned | dongping_w | August 24, 2020 2:47:59 PM GMT+08:00 | FOK-BB holder is very difficult to slide onto therapy table top | BB_WIP12.6 |  | No workaround is recommended. The issue does not affect the function and intended use. The design improvement will be implemented. Then the user will easily install the Phantom Holder in place. The design change will be initiated by ECR-076387 and implemented by E-129050. | Subsystem | FOK; LCM_Candidate; SP2_Candidate | N/A |  | wang_ada | 0 | 4 |  | Couch | NO | Assigned | Next Design | Brilliance 64 | 222254
CTUDT00630830 | Refresh | Assigned | dongping_w | August 24, 2020 2:47:59 PM GMT+08:00 | FOK-BB holder is very difficult to slide onto therapy table top | BB_WIP12.6 |  | No workaround is recommended. The issue does not affect the function and intended use. The design improvement will be implemented. Then the user will easily install the Phantom Holder in place. The design change will be initiated by ECR-076387 and implemented by E-129050. | Subsystem | FOK; LCM_Candidate; SP2_Candidate | N/A |  | wang_ada | 0 | 4 |  | Couch | NO | Assigned | Next Design | Brilliance Big Bore | 222254
CTUDT00630830 | Refresh | Assigned | dongping_w | August 24, 2020 2:47:59 PM GMT+08:00 | FOK-BB holder is very difficult to slide onto therapy table top | BB_WIP12.6 |  | No workaround is recommended. The issue does not affect the function and intended use. The design improvement will be implemented. Then the user will easily install the Phantom Holder in place. The design change will be initiated by ECR-076387 and implemented by E-129050. | Subsystem | FOK; LCM_Candidate; SP2_Candidate | N/A |  | wang_ada | 0 | 4 |  | Couch | NO | Assigned | Next Design | Ingenuity 64 | 222254
CTUDT00630830 | Refresh | Assigned | dongping_w | August 24, 2020 2:47:59 PM GMT+08:00 | FOK-BB holder is very difficult to slide onto therapy table top | BB_WIP12.6 |  | No workaround is recommended. The issue does not affect the function and intended use. The design improvement will be implemented. Then the user will easily install the Phantom Holder in place. The design change will be initiated by ECR-076387 and implemented by E-129050. | Subsystem | FOK; LCM_Candidate; SP2_Candidate | N/A |  | wang_ada | 0 | 4 |  | Couch | NO | Assigned | Next Design | IQon1 | 222254
CTUDT00630827 | Refresh | Opened | lei_yang | August 24, 2020 10:06:48 AM GMT+08:00 | "Zero button became disabled after closing estop" in case open Estop before end exam. | WIP25.2 |  |  | Subsystem | CRN | NA |  | rajalakshmi_j | 0 | 4 |  | Support-System | NO | Opened | SP1-Mar-21-CRN | Brilliance 64
CTUDT00630827 | Refresh | Opened | lei_yang | August 24, 2020 10:06:48 AM GMT+08:00 | "Zero button became disabled after closing estop" in case open Estop before end exam. | WIP25.2 |  |  | Subsystem | CRN | NA |  | rajalakshmi_j | 0 | 4 |  | Support-System | NO | Opened | SP1-Mar-21-CRN | Brilliance Big Bore
CTUDT00630827 | Refresh | Opened | lei_yang | August 24, 2020 10:06:48 AM GMT+08:00 | "Zero button became disabled after closing estop" in case open Estop before end exam. | WIP25.2 |  |  | Subsystem | CRN | NA |  | rajalakshmi_j | 0 | 4 |  | Support-System | NO | Opened | SP1-Mar-21-CRN | Ingenuity 64
CTUDT00630818 | Refresh | Closed | wang_l | August 21, 2020 7:02:58 PM GMT+08:00 | Refresh: Conflict rule of Exam Card Import and Export for identify the EC name | WIP25.2 |  |  | Subsystem |  | NA | No Plan to Fix | guraddi_r | 0 | 4 |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00630818 | Refresh | Closed | wang_l | August 21, 2020 7:02:58 PM GMT+08:00 | Refresh: Conflict rule of Exam Card Import and Export for identify the EC name | WIP25.2 |  |  | Subsystem |  | NA | No Plan to Fix | guraddi_r | 0 | 4 |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00630802 | Refresh | Opened | shaham_t | August 19, 2020 8:14:12 PM GMT+08:00 | FoK: Keyboard language setting not retained as per preference setting language. Changes back to English keyboard in Login page after every restart. | 4.7.7 | WIP36 | NA | System | FoK; SP1_PM_Approved | N/A | Fixed | kumar_gau | 0 | 2 | Console Software | Console4 | NO | Opened | WIP34SubCom | Ovation | PR197817 | WIP36 | November 12, 2020 7:52:16 PM GMT+08:00
CTUDT00630802 | Refresh | Opened | shaham_t | August 19, 2020 8:14:12 PM GMT+08:00 | FoK: Keyboard language setting not retained as per preference setting language. Changes back to English keyboard in Login page after every restart. | 4.7.7 | WIP36 | NA | System | FoK; SP1_PM_Approved | N/A | Fixed | kumar_gau | 0 | 2 | Console Software | Console4 | NO | Opened | WIP34SubCom | Brilliance 64 | PR197817 | WIP36 | November 12, 2020 7:52:16 PM GMT+08:00
CTUDT00630802 | Refresh | Opened | shaham_t | August 19, 2020 8:14:12 PM GMT+08:00 | FoK: Keyboard language setting not retained as per preference setting language. Changes back to English keyboard in Login page after every restart. | 4.7.7 | WIP36 | NA | System | FoK; SP1_PM_Approved | N/A | Fixed | kumar_gau | 0 | 2 | Console Software | Console4 | NO | Opened | WIP34SubCom | Brilliance Big Bore | PR197817 | WIP36 | November 12, 2020 7:52:16 PM GMT+08:00
CTUDT00630802 | Refresh | Opened | shaham_t | August 19, 2020 8:14:12 PM GMT+08:00 | FoK: Keyboard language setting not retained as per preference setting language. Changes back to English keyboard in Login page after every restart. | 4.7.7 | WIP36 | NA | System | FoK; SP1_PM_Approved | N/A | Fixed | kumar_gau | 0 | 2 | Console Software | Console4 | NO | Opened | WIP34SubCom | Ingenuity 64 | PR197817 | WIP36 | November 12, 2020 7:52:16 PM GMT+08:00
CTUDT00630802 | Refresh | Opened | shaham_t | August 19, 2020 8:14:12 PM GMT+08:00 | FoK: Keyboard language setting not retained as per preference setting language. Changes back to English keyboard in Login page after every restart. | 4.7.7 | WIP36 | NA | System | FoK; SP1_PM_Approved | N/A | Fixed | kumar_gau | 0 | 2 | Console Software | Console4 | NO | Opened | WIP34SubCom | IQon1 | PR197817 | WIP36 | November 12, 2020 7:52:16 PM GMT+08:00
CTUDT00630796 | Refresh | Closed | lin_s | August 19, 2020 12:34:42 PM GMT+08:00 | FOK: R2D LED display wrong status in field. | wip22.5 |  | Provide a method to help FSE distinguish the true or false error. | System | FOK | NA | Forwarded to Different Database | yaping_l | 0 |  |  | GCX | NO | Closed |  | Ingenuity 64
CTUDT00630793 | Refresh | Closed | cai_t | August 19, 2020 9:43:33 AM GMT+08:00 | The value of directionM of last 7 COUCH_REPORT_VERT_POS_BKST is incorrect when completing unload action. | WIP22.5 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | cx_zhang | 0 |  |  | Couch | NO | Closed |  | Brilliance 64
CTUDT00630793 | Refresh | Closed | cai_t | August 19, 2020 9:43:33 AM GMT+08:00 | The value of directionM of last 7 COUCH_REPORT_VERT_POS_BKST is incorrect when completing unload action. | WIP22.5 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | cx_zhang | 0 |  |  | Couch | NO | Closed |  | Brilliance Big Bore
CTUDT00630793 | Refresh | Closed | cai_t | August 19, 2020 9:43:33 AM GMT+08:00 | The value of directionM of last 7 COUCH_REPORT_VERT_POS_BKST is incorrect when completing unload action. | WIP22.5 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | cx_zhang | 0 |  |  | Couch | NO | Closed |  | Ingenuity 64
CTUDT00630793 | Refresh | Closed | cai_t | August 19, 2020 9:43:33 AM GMT+08:00 | The value of directionM of last 7 COUCH_REPORT_VERT_POS_BKST is incorrect when completing unload action. | WIP22.5 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | cx_zhang | 0 |  |  | Couch | NO | Closed |  | IQon1
CTUDT00630789 | Refresh | Closed | Noy_Seelig | August 18, 2020 9:38:56 PM GMT+08:00 | Refresh: CTDI Free Air wrong Window Width value for Pediatric Body | WIP 31 | WIP33 | N/A | Subsystem | SP1_PM_Approved | N/A | Fixed | frider_o | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP33SubCom | IQon1 |  | WIP33 | October 29, 2020 4:39:37 PM GMT+08:00
CTUDT00630783 | Refresh | Closed | elavarasan_s | August 18, 2020 6:30:11 PM GMT+08:00 | Login page is switching to Windows login and login is getting failed when new user added with "mu" special character | WIP25 | WIP 25.4 |  | Subsystem |  | NA | Fixed | kumar_gau | 0 | 2 | Console Software | Console4 | NO | Closed | WIP25.4_SegReq | Ovation |  | WIP 25.4 | August 31, 2020 6:41:56 PM GMT+08:00
CTUDT00630783 | Refresh | Closed | elavarasan_s | August 18, 2020 6:30:11 PM GMT+08:00 | Login page is switching to Windows login and login is getting failed when new user added with "mu" special character | WIP25 | WIP 25.4 |  | Subsystem |  | NA | Fixed | kumar_gau | 0 | 2 | Console Software | Console4 | NO | Closed | WIP25.4_SegReq | Brilliance 64 |  | WIP 25.4 | August 31, 2020 6:41:56 PM GMT+08:00
CTUDT00630783 | Refresh | Closed | elavarasan_s | August 18, 2020 6:30:11 PM GMT+08:00 | Login page is switching to Windows login and login is getting failed when new user added with "mu" special character | WIP25 | WIP 25.4 |  | Subsystem |  | NA | Fixed | kumar_gau | 0 | 2 | Console Software | Console4 | NO | Closed | WIP25.4_SegReq | Brilliance Big Bore |  | WIP 25.4 | August 31, 2020 6:41:56 PM GMT+08:00
CTUDT00630783 | Refresh | Closed | elavarasan_s | August 18, 2020 6:30:11 PM GMT+08:00 | Login page is switching to Windows login and login is getting failed when new user added with "mu" special character | WIP25 | WIP 25.4 |  | Subsystem |  | NA | Fixed | kumar_gau | 0 | 2 | Console Software | Console4 | NO | Closed | WIP25.4_SegReq | Ingenuity 64 |  | WIP 25.4 | August 31, 2020 6:41:56 PM GMT+08:00
CTUDT00630783 | Refresh | Closed | elavarasan_s | August 18, 2020 6:30:11 PM GMT+08:00 | Login page is switching to Windows login and login is getting failed when new user added with "mu" special character | WIP25 | WIP 25.4 |  | Subsystem |  | NA | Fixed | kumar_gau | 0 | 2 | Console Software | Console4 | NO | Closed | WIP25.4_SegReq | IQon1 |  | WIP 25.4 | August 31, 2020 6:41:56 PM GMT+08:00
CTUDT00630733 | Refresh | Closed | lin_s | August 12, 2020 7:10:02 PM GMT+08:00 | Fok: After Refresh Upgrade : cannot connect for diag the Site from PRS portal | 4.7.7 | WIP33 | M2M can be used to access these applications from PRS | Subsystem | SP1_PM_Approved | N/A | Fixed | Prabhakar_Priyank_Muppidi | 0 | 3 | Console Software | Service Tools | NO | Closed | WIP33SubCom | Ovation | PR197818 | WIP33 | October 17, 2020 4:13:16 PM GMT+08:00
CTUDT00630733 | Refresh | Closed | lin_s | August 12, 2020 7:10:02 PM GMT+08:00 | Fok: After Refresh Upgrade : cannot connect for diag the Site from PRS portal | 4.7.7 | WIP33 | M2M can be used to access these applications from PRS | Subsystem | SP1_PM_Approved | N/A | Fixed | Prabhakar_Priyank_Muppidi | 0 | 3 | Console Software | Service Tools | NO | Closed | WIP33SubCom | Brilliance 64 | PR197818 | WIP33 | October 17, 2020 4:13:16 PM GMT+08:00
CTUDT00630733 | Refresh | Closed | lin_s | August 12, 2020 7:10:02 PM GMT+08:00 | Fok: After Refresh Upgrade : cannot connect for diag the Site from PRS portal | 4.7.7 | WIP33 | M2M can be used to access these applications from PRS | Subsystem | SP1_PM_Approved | N/A | Fixed | Prabhakar_Priyank_Muppidi | 0 | 3 | Console Software | Service Tools | NO | Closed | WIP33SubCom | Brilliance Big Bore | PR197818 | WIP33 | October 17, 2020 4:13:16 PM GMT+08:00
CTUDT00630733 | Refresh | Closed | lin_s | August 12, 2020 7:10:02 PM GMT+08:00 | Fok: After Refresh Upgrade : cannot connect for diag the Site from PRS portal | 4.7.7 | WIP33 | M2M can be used to access these applications from PRS | Subsystem | SP1_PM_Approved | N/A | Fixed | Prabhakar_Priyank_Muppidi | 0 | 3 | Console Software | Service Tools | NO | Closed | WIP33SubCom | Ingenuity 64 | PR197818 | WIP33 | October 17, 2020 4:13:16 PM GMT+08:00
CTUDT00630733 | Refresh | Closed | lin_s | August 12, 2020 7:10:02 PM GMT+08:00 | Fok: After Refresh Upgrade : cannot connect for diag the Site from PRS portal | 4.7.7 | WIP33 | M2M can be used to access these applications from PRS | Subsystem | SP1_PM_Approved | N/A | Fixed | Prabhakar_Priyank_Muppidi | 0 | 3 | Console Software | Service Tools | NO | Closed | WIP33SubCom | IQon1 | PR197818 | WIP33 | October 17, 2020 4:13:16 PM GMT+08:00
CTUDT00630709 | Refresh | Closed | s_namburi | August 12, 2020 3:54:22 PM GMT+08:00 | FoK: Incorrect date format even after change in preference | 4.7.7 |  |  | Subsystem |  | NA | Change Proposal | guraddi_r | 0 |  |  | Console4 | NO | Closed |  | IQon1 | PR203431
CTUDT00630706 | Refresh | Closed | s_namburi | August 12, 2020 3:43:28 PM GMT+08:00 | FoK: 'The data to reconstruct the results of this scan is corrupted.’ message displayed during offline recon | 4.7.7 | WIP 35 | NA | Subsystem | FoK; SP1_PM_Approved | NA | Fixed | 320057847 | 0 | 2 | Console Software | Console4 | YES | Closed | WIP35SubCom | Ovation | PR203354 | WIP 35 | November 19, 2020 3:32:33 PM GMT+08:00
CTUDT00630706 | Refresh | Closed | s_namburi | August 12, 2020 3:43:28 PM GMT+08:00 | FoK: 'The data to reconstruct the results of this scan is corrupted.’ message displayed during offline recon | 4.7.7 | WIP 35 | NA | Subsystem | FoK; SP1_PM_Approved | NA | Fixed | 320057847 | 0 | 2 | Console Software | Console4 | YES | Closed | WIP35SubCom | Brilliance 64 | PR203354 | WIP 35 | November 19, 2020 3:32:33 PM GMT+08:00
CTUDT00630706 | Refresh | Closed | s_namburi | August 12, 2020 3:43:28 PM GMT+08:00 | FoK: 'The data to reconstruct the results of this scan is corrupted.’ message displayed during offline recon | 4.7.7 | WIP 35 | NA | Subsystem | FoK; SP1_PM_Approved | NA | Fixed | 320057847 | 0 | 2 | Console Software | Console4 | YES | Closed | WIP35SubCom | Ingenuity 64 | PR203354 | WIP 35 | November 19, 2020 3:32:33 PM GMT+08:00
CTUDT00630706 | Refresh | Closed | s_namburi | August 12, 2020 3:43:28 PM GMT+08:00 | FoK: 'The data to reconstruct the results of this scan is corrupted.’ message displayed during offline recon | 4.7.7 | WIP 35 | NA | Subsystem | FoK; SP1_PM_Approved | NA | Fixed | 320057847 | 0 | 2 | Console Software | Console4 | YES | Closed | WIP35SubCom | IQon1 | PR203354 | WIP 35 | November 19, 2020 3:32:33 PM GMT+08:00
CTUDT00630696 | Refresh | Closed | s_namburi | August 12, 2020 2:10:06 PM GMT+08:00 | FoK: Error " System Error, Please restart the system" during Save BugRep | 4.7.7 |  |  | Subsystem |  |  | Duplicate | s_namburi | 0 |  |  | Support-System | NO | Closed |  | IQon1 | PR217798
CTUDT00630689 | Refresh | Closed | s_namburi | August 12, 2020 1:23:51 PM GMT+08:00 | FoK: Transfer icon didn't showed on some series that already send to PACS | 4.7.7 | WIP25.5 |  | Subsystem | PM_Approved | N/A | Fixed | pandey_a | 0 | 4 | Console Software | Console4 | NO | Closed | HotFix2 | Ovation | PR217800 | WIP25.5 | September 5, 2020 1:53:52 AM GMT+08:00
CTUDT00630689 | Refresh | Closed | s_namburi | August 12, 2020 1:23:51 PM GMT+08:00 | FoK: Transfer icon didn't showed on some series that already send to PACS | 4.7.7 | WIP25.5 |  | Subsystem | PM_Approved | N/A | Fixed | pandey_a | 0 | 4 | Console Software | Console4 | NO | Closed | HotFix2 | Brilliance 64 | PR217800 | WIP25.5 | September 5, 2020 1:53:52 AM GMT+08:00
CTUDT00630689 | Refresh | Closed | s_namburi | August 12, 2020 1:23:51 PM GMT+08:00 | FoK: Transfer icon didn't showed on some series that already send to PACS | 4.7.7 | WIP25.5 |  | Subsystem | PM_Approved | N/A | Fixed | pandey_a | 0 | 4 | Console Software | Console4 | NO | Closed | HotFix2 | Brilliance Big Bore | PR217800 | WIP25.5 | September 5, 2020 1:53:52 AM GMT+08:00
CTUDT00630689 | Refresh | Closed | s_namburi | August 12, 2020 1:23:51 PM GMT+08:00 | FoK: Transfer icon didn't showed on some series that already send to PACS | 4.7.7 | WIP25.5 |  | Subsystem | PM_Approved | N/A | Fixed | pandey_a | 0 | 4 | Console Software | Console4 | NO | Closed | HotFix2 | Ingenuity 64 | PR217800 | WIP25.5 | September 5, 2020 1:53:52 AM GMT+08:00
CTUDT00630689 | Refresh | Closed | s_namburi | August 12, 2020 1:23:51 PM GMT+08:00 | FoK: Transfer icon didn't showed on some series that already send to PACS | 4.7.7 | WIP25.5 |  | Subsystem | PM_Approved | N/A | Fixed | pandey_a | 0 | 4 | Console Software | Console4 | NO | Closed | HotFix2 | IQon1 | PR217800 | WIP25.5 | September 5, 2020 1:53:52 AM GMT+08:00
CTUDT00630689 | Refresh | Closed | s_namburi | August 12, 2020 1:23:51 PM GMT+08:00 | FoK: Transfer icon didn't showed on some series that already send to PACS | 4.7.7 | WIP25.5 |  | Subsystem | PM_Approved | N/A | Fixed | pandey_a | 0 | 4 | Console Software | Console4 | NO | Closed | HotFix2 | Hawk | PR217800 | WIP25.5 | September 5, 2020 1:53:52 AM GMT+08:00
CTUDT00630644 | Refresh | Resolved | richards_b | August 10, 2020 7:49:33 PM GMT+08:00 | FoK: Poitiers: CIRS Get Ready Reli issue | 4.7.7 | WIP34 | Perform offline recon to recover images | Subsystem | FoK; SP1_PM_Approved | N/A | Fixed | richards_b | 0 | 3 | CIRS_SW_SuprNodeSvr | CIRS | YES | Resolved | WIP34SubCom | Ovation | PR215950 | WIP34 | November 17, 2020 5:16:37 AM GMT+08:00
CTUDT00630644 | Refresh | Resolved | richards_b | August 10, 2020 7:49:33 PM GMT+08:00 | FoK: Poitiers: CIRS Get Ready Reli issue | 4.7.7 | WIP34 | Perform offline recon to recover images | Subsystem | FoK; SP1_PM_Approved | N/A | Fixed | richards_b | 0 | 3 | CIRS_SW_SuprNodeSvr | CIRS | YES | Resolved | WIP34SubCom | Brilliance 64 | PR215950 | WIP34 | November 17, 2020 5:16:37 AM GMT+08:00
CTUDT00630644 | Refresh | Resolved | richards_b | August 10, 2020 7:49:33 PM GMT+08:00 | FoK: Poitiers: CIRS Get Ready Reli issue | 4.7.7 | WIP34 | Perform offline recon to recover images | Subsystem | FoK; SP1_PM_Approved | N/A | Fixed | richards_b | 0 | 3 | CIRS_SW_SuprNodeSvr | CIRS | YES | Resolved | WIP34SubCom | Brilliance Big Bore | PR215950 | WIP34 | November 17, 2020 5:16:37 AM GMT+08:00
CTUDT00630644 | Refresh | Resolved | richards_b | August 10, 2020 7:49:33 PM GMT+08:00 | FoK: Poitiers: CIRS Get Ready Reli issue | 4.7.7 | WIP34 | Perform offline recon to recover images | Subsystem | FoK; SP1_PM_Approved | N/A | Fixed | richards_b | 0 | 3 | CIRS_SW_SuprNodeSvr | CIRS | YES | Resolved | WIP34SubCom | Ingenuity 64 | PR215950 | WIP34 | November 17, 2020 5:16:37 AM GMT+08:00
CTUDT00630644 | Refresh | Resolved | richards_b | August 10, 2020 7:49:33 PM GMT+08:00 | FoK: Poitiers: CIRS Get Ready Reli issue | 4.7.7 | WIP34 | Perform offline recon to recover images | Subsystem | FoK; SP1_PM_Approved | N/A | Fixed | richards_b | 0 | 3 | CIRS_SW_SuprNodeSvr | CIRS | YES | Resolved | WIP34SubCom | IQon1 | PR215950 | WIP34 | November 17, 2020 5:16:37 AM GMT+08:00
CTUDT00630642 | Refresh | Resolved | s_namburi | August 10, 2020 5:11:46 PM GMT+08:00 | FoK: Operation console application process slower | 4.7.7 | WIP34 |  | Subsystem | FoK; SP1_PM_Approved | NA | Fixed | challa_s | 0 | 2 | Console Software | Console4 | NO | Resolved | WIP34SubCom | Ovation | PR200433 | WIP34 | November 10, 2020 1:09:43 PM GMT+08:00
CTUDT00630642 | Refresh | Resolved | s_namburi | August 10, 2020 5:11:46 PM GMT+08:00 | FoK: Operation console application process slower | 4.7.7 | WIP34 |  | Subsystem | FoK; SP1_PM_Approved | NA | Fixed | challa_s | 0 | 2 | Console Software | Console4 | NO | Resolved | WIP34SubCom | Brilliance 64 | PR200433 | WIP34 | November 10, 2020 1:09:43 PM GMT+08:00
CTUDT00630642 | Refresh | Resolved | s_namburi | August 10, 2020 5:11:46 PM GMT+08:00 | FoK: Operation console application process slower | 4.7.7 | WIP34 |  | Subsystem | FoK; SP1_PM_Approved | NA | Fixed | challa_s | 0 | 2 | Console Software | Console4 | NO | Resolved | WIP34SubCom | Brilliance Big Bore | PR200433 | WIP34 | November 10, 2020 1:09:43 PM GMT+08:00
CTUDT00630642 | Refresh | Resolved | s_namburi | August 10, 2020 5:11:46 PM GMT+08:00 | FoK: Operation console application process slower | 4.7.7 | WIP34 |  | Subsystem | FoK; SP1_PM_Approved | NA | Fixed | challa_s | 0 | 2 | Console Software | Console4 | NO | Resolved | WIP34SubCom | Ingenuity 64 | PR200433 | WIP34 | November 10, 2020 1:09:43 PM GMT+08:00
CTUDT00630642 | Refresh | Resolved | s_namburi | August 10, 2020 5:11:46 PM GMT+08:00 | FoK: Operation console application process slower | 4.7.7 | WIP34 |  | Subsystem | FoK; SP1_PM_Approved | NA | Fixed | challa_s | 0 | 2 | Console Software | Console4 | NO | Resolved | WIP34SubCom | IQon1 | PR200433 | WIP34 | November 10, 2020 1:09:43 PM GMT+08:00
CTUDT00630640 | Refresh | Closed | s_namburi | August 10, 2020 5:07:21 PM GMT+08:00 | FoK: Exam Manager has stopped working error ( Auto recon image didn't send to PACS and create duplicate series) | 4.7.7 | WIP25.5 |  | Subsystem | SP1_PM_Approved | N/A | Fixed | pandey_a | 0 | 3 | Console Software | Console4 | NO | Closed | WIP25.4_SegReq | Ovation | PR217257 | WIP25.5 | September 5, 2020 1:53:19 AM GMT+08:00
CTUDT00630640 | Refresh | Closed | s_namburi | August 10, 2020 5:07:21 PM GMT+08:00 | FoK: Exam Manager has stopped working error ( Auto recon image didn't send to PACS and create duplicate series) | 4.7.7 | WIP25.5 |  | Subsystem | SP1_PM_Approved | N/A | Fixed | pandey_a | 0 | 3 | Console Software | Console4 | NO | Closed | WIP25.4_SegReq | Brilliance 64 | PR217257 | WIP25.5 | September 5, 2020 1:53:19 AM GMT+08:00
CTUDT00630640 | Refresh | Closed | s_namburi | August 10, 2020 5:07:21 PM GMT+08:00 | FoK: Exam Manager has stopped working error ( Auto recon image didn't send to PACS and create duplicate series) | 4.7.7 | WIP25.5 |  | Subsystem | SP1_PM_Approved | N/A | Fixed | pandey_a | 0 | 3 | Console Software | Console4 | NO | Closed | WIP25.4_SegReq | Brilliance Big Bore | PR217257 | WIP25.5 | September 5, 2020 1:53:19 AM GMT+08:00
CTUDT00630640 | Refresh | Closed | s_namburi | August 10, 2020 5:07:21 PM GMT+08:00 | FoK: Exam Manager has stopped working error ( Auto recon image didn't send to PACS and create duplicate series) | 4.7.7 | WIP25.5 |  | Subsystem | SP1_PM_Approved | N/A | Fixed | pandey_a | 0 | 3 | Console Software | Console4 | NO | Closed | WIP25.4_SegReq | Ingenuity 64 | PR217257 | WIP25.5 | September 5, 2020 1:53:19 AM GMT+08:00
CTUDT00630640 | Refresh | Closed | s_namburi | August 10, 2020 5:07:21 PM GMT+08:00 | FoK: Exam Manager has stopped working error ( Auto recon image didn't send to PACS and create duplicate series) | 4.7.7 | WIP25.5 |  | Subsystem | SP1_PM_Approved | N/A | Fixed | pandey_a | 0 | 3 | Console Software | Console4 | NO | Closed | WIP25.4_SegReq | IQon1 | PR217257 | WIP25.5 | September 5, 2020 1:53:19 AM GMT+08:00
CTUDT00630638 | Refresh | Closed | s_namburi | August 10, 2020 4:53:54 PM GMT+08:00 | FoK: Error " System Error, Please restart the system" during Save BugRep | 4.7.7 |  |  | Subsystem |  | NA | Duplicate | srushti_k | 0 |  |  | iCAP | NO | Closed |  | Ovation | PR217798
CTUDT00630638 | Refresh | Closed | s_namburi | August 10, 2020 4:53:54 PM GMT+08:00 | FoK: Error " System Error, Please restart the system" during Save BugRep | 4.7.7 |  |  | Subsystem |  | NA | Duplicate | srushti_k | 0 |  |  | iCAP | NO | Closed |  | Brilliance 64 | PR217798
CTUDT00630638 | Refresh | Closed | s_namburi | August 10, 2020 4:53:54 PM GMT+08:00 | FoK: Error " System Error, Please restart the system" during Save BugRep | 4.7.7 |  |  | Subsystem |  | NA | Duplicate | srushti_k | 0 |  |  | iCAP | NO | Closed |  | Brilliance Big Bore | PR217798
CTUDT00630638 | Refresh | Closed | s_namburi | August 10, 2020 4:53:54 PM GMT+08:00 | FoK: Error " System Error, Please restart the system" during Save BugRep | 4.7.7 |  |  | Subsystem |  | NA | Duplicate | srushti_k | 0 |  |  | iCAP | NO | Closed |  | Ingenuity 64 | PR217798
CTUDT00630638 | Refresh | Closed | s_namburi | August 10, 2020 4:53:54 PM GMT+08:00 | FoK: Error " System Error, Please restart the system" during Save BugRep | 4.7.7 |  |  | Subsystem |  | NA | Duplicate | srushti_k | 0 |  |  | iCAP | NO | Closed |  | IQon1 | PR217798
CTUDT00630633 | Refresh | Closed | s_namburi | August 10, 2020 3:57:37 PM GMT+08:00 | FoK: Auto Voice delay - appears once | 4.7.7 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | Wang_Dexin | 0 |  |  | GCX | NO | Closed |  | IQon1 | PR195130
CTUDT00630602 | Refresh | Closed | periasamy_g | August 4, 2020 9:47:46 PM GMT+08:00 | FoK: Customer reported with the issue in Application, look like slow process and with Not recon image message | 4.7.7 | WIP25.7 |  | Subsystem | FoK ; SP1_PM_Approved | NA | Fixed | s_srivastava | 0 | 2 | Console Software | Console4 | NO | Closed | WIP25.7_SegReq | Ovation | PR200433 PR203331 PR217226 | WIP25.7 | September 29, 2020 4:18:07 PM GMT+08:00
CTUDT00630602 | Refresh | Closed | periasamy_g | August 4, 2020 9:47:46 PM GMT+08:00 | FoK: Customer reported with the issue in Application, look like slow process and with Not recon image message | 4.7.7 | WIP25.7 |  | Subsystem | FoK ; SP1_PM_Approved | NA | Fixed | s_srivastava | 0 | 2 | Console Software | Console4 | NO | Closed | WIP25.7_SegReq | Brilliance 64 | PR200433 PR203331 PR217226 | WIP25.7 | September 29, 2020 4:18:07 PM GMT+08:00
CTUDT00630602 | Refresh | Closed | periasamy_g | August 4, 2020 9:47:46 PM GMT+08:00 | FoK: Customer reported with the issue in Application, look like slow process and with Not recon image message | 4.7.7 | WIP25.7 |  | Subsystem | FoK ; SP1_PM_Approved | NA | Fixed | s_srivastava | 0 | 2 | Console Software | Console4 | NO | Closed | WIP25.7_SegReq | Brilliance Big Bore | PR200433 PR203331 PR217226 | WIP25.7 | September 29, 2020 4:18:07 PM GMT+08:00
CTUDT00630602 | Refresh | Closed | periasamy_g | August 4, 2020 9:47:46 PM GMT+08:00 | FoK: Customer reported with the issue in Application, look like slow process and with Not recon image message | 4.7.7 | WIP25.7 |  | Subsystem | FoK ; SP1_PM_Approved | NA | Fixed | s_srivastava | 0 | 2 | Console Software | Console4 | NO | Closed | WIP25.7_SegReq | Ingenuity 64 | PR200433 PR203331 PR217226 | WIP25.7 | September 29, 2020 4:18:07 PM GMT+08:00
CTUDT00630602 | Refresh | Closed | periasamy_g | August 4, 2020 9:47:46 PM GMT+08:00 | FoK: Customer reported with the issue in Application, look like slow process and with Not recon image message | 4.7.7 | WIP25.7 |  | Subsystem | FoK ; SP1_PM_Approved | NA | Fixed | s_srivastava | 0 | 2 | Console Software | Console4 | NO | Closed | WIP25.7_SegReq | IQon1 | PR200433 PR203331 PR217226 | WIP25.7 | September 29, 2020 4:18:07 PM GMT+08:00
CTUDT00630594 | Refresh | Resolved | richards_b | August 4, 2020 8:05:43 PM GMT+08:00 | Tracker TTFI Improvement with background recons | 4.7.5.43530 | WIP33 | The user would need to trigger the clinical scan manually if they saw the tracker was delayed. The IFU also asks the user specifically to "Monitor the bolus values on the images during the Tracker scan to ensure the scan is proceeding as planned." | Subsystem | SP1_PM_Approved | N/A | Fixed | richards_b | 0 | 3 | CIRS_SW_ImgNodeSvr | CIRS | NO | Resolved | WIP33SubCom | Ovation | PR101705 PR125930 PR321611 | WIP33 | November 4, 2020 3:45:24 AM GMT+08:00
CTUDT00630594 | Refresh | Resolved | richards_b | August 4, 2020 8:05:43 PM GMT+08:00 | Tracker TTFI Improvement with background recons | 4.7.5.43530 | WIP33 | The user would need to trigger the clinical scan manually if they saw the tracker was delayed. The IFU also asks the user specifically to "Monitor the bolus values on the images during the Tracker scan to ensure the scan is proceeding as planned." | Subsystem | SP1_PM_Approved | N/A | Fixed | richards_b | 0 | 3 | CIRS_SW_ImgNodeSvr | CIRS | NO | Resolved | WIP33SubCom | Brilliance 64 | PR101705 PR125930 PR321611 | WIP33 | November 4, 2020 3:45:24 AM GMT+08:00
CTUDT00630594 | Refresh | Resolved | richards_b | August 4, 2020 8:05:43 PM GMT+08:00 | Tracker TTFI Improvement with background recons | 4.7.5.43530 | WIP33 | The user would need to trigger the clinical scan manually if they saw the tracker was delayed. The IFU also asks the user specifically to "Monitor the bolus values on the images during the Tracker scan to ensure the scan is proceeding as planned." | Subsystem | SP1_PM_Approved | N/A | Fixed | richards_b | 0 | 3 | CIRS_SW_ImgNodeSvr | CIRS | NO | Resolved | WIP33SubCom | Brilliance Big Bore | PR101705 PR125930 PR321611 | WIP33 | November 4, 2020 3:45:24 AM GMT+08:00
CTUDT00630594 | Refresh | Resolved | richards_b | August 4, 2020 8:05:43 PM GMT+08:00 | Tracker TTFI Improvement with background recons | 4.7.5.43530 | WIP33 | The user would need to trigger the clinical scan manually if they saw the tracker was delayed. The IFU also asks the user specifically to "Monitor the bolus values on the images during the Tracker scan to ensure the scan is proceeding as planned." | Subsystem | SP1_PM_Approved | N/A | Fixed | richards_b | 0 | 3 | CIRS_SW_ImgNodeSvr | CIRS | NO | Resolved | WIP33SubCom | Ingenuity 64 | PR101705 PR125930 PR321611 | WIP33 | November 4, 2020 3:45:24 AM GMT+08:00
CTUDT00630594 | Refresh | Resolved | richards_b | August 4, 2020 8:05:43 PM GMT+08:00 | Tracker TTFI Improvement with background recons | 4.7.5.43530 | WIP33 | The user would need to trigger the clinical scan manually if they saw the tracker was delayed. The IFU also asks the user specifically to "Monitor the bolus values on the images during the Tracker scan to ensure the scan is proceeding as planned." | Subsystem | SP1_PM_Approved | N/A | Fixed | richards_b | 0 | 3 | CIRS_SW_ImgNodeSvr | CIRS | NO | Resolved | WIP33SubCom | IQon1 | PR101705 PR125930 PR321611 | WIP33 | November 4, 2020 3:45:24 AM GMT+08:00
CTUDT00630592 | Refresh | Closed | richards_b | August 4, 2020 7:44:08 PM GMT+08:00 | IQon High Frequency Suspend/Resume Improvement | CIRS 208/209 version | CIRS-5.1.0.0210 | 1. Re-boot the computer in order to recover the missing images 2. Site could upgrade to a 6 server configuration 3. Reduce the number of SBIs or adjust workflow to reduce the number of suspends/resumes (waiting for recons to finish before moving to next patient). | Subsystem | SP1_PM_Approved | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_SuprNodeSvr | CIRS | YES | Closed | HotFix2 | IQon1 | PR191940 | CIRS-5.1.0.0210 | September 14, 2020 8:55:16 PM GMT+08:00
CTUDT00630582 | Refresh | Closed | Noy_Seelig | August 3, 2020 9:11:42 PM GMT+08:00 | Spec update- Acceptance and Constancy wrong FOV for Ingenuity HP | Spec Rev18 | WIP33 |  | Subsystem | SP1_PM_Approved | N/A | Fixed | richards_b | 0 | 2 | Design Documentation | Support-Physics | NO | Closed | spec new version | Ingenuity 64 |  | WIP33 | November 4, 2020 6:16:04 AM GMT+08:00
CTUDT00630525 | Refresh | Closed | cai_t | July 27, 2020 2:57:01 PM GMT+08:00 | [Old GCP]Couch can go up to 4 position in vertical direction when Pressing “Horizontal In” and “Load”button | WIP25 | WIP32 | don't press "horizontal in" button with load button at same time | Subsystem | SP1_PM_Approved | NA | Fixed | yang_h | 0 | 3 | Gantry Software | GCX | YES | Closed | WIP32SubCom | Ovation |  | WIP32 | September 14, 2020 11:00:11 AM GMT+08:00
CTUDT00630525 | Refresh | Closed | cai_t | July 27, 2020 2:57:01 PM GMT+08:00 | [Old GCP]Couch can go up to 4 position in vertical direction when Pressing “Horizontal In” and “Load”button | WIP25 | WIP32 | don't press "horizontal in" button with load button at same time | Subsystem | SP1_PM_Approved | NA | Fixed | yang_h | 0 | 3 | Gantry Software | GCX | YES | Closed | WIP32SubCom | Brilliance 64 |  | WIP32 | September 14, 2020 11:00:11 AM GMT+08:00
CTUDT00630525 | Refresh | Closed | cai_t | July 27, 2020 2:57:01 PM GMT+08:00 | [Old GCP]Couch can go up to 4 position in vertical direction when Pressing “Horizontal In” and “Load”button | WIP25 | WIP32 | don't press "horizontal in" button with load button at same time | Subsystem | SP1_PM_Approved | NA | Fixed | yang_h | 0 | 3 | Gantry Software | GCX | YES | Closed | WIP32SubCom | Brilliance Big Bore |  | WIP32 | September 14, 2020 11:00:11 AM GMT+08:00
CTUDT00630525 | Refresh | Closed | cai_t | July 27, 2020 2:57:01 PM GMT+08:00 | [Old GCP]Couch can go up to 4 position in vertical direction when Pressing “Horizontal In” and “Load”button | WIP25 | WIP32 | don't press "horizontal in" button with load button at same time | Subsystem | SP1_PM_Approved | NA | Fixed | yang_h | 0 | 3 | Gantry Software | GCX | YES | Closed | WIP32SubCom | Ingenuity 64 |  | WIP32 | September 14, 2020 11:00:11 AM GMT+08:00
CTUDT00630525 | Refresh | Closed | cai_t | July 27, 2020 2:57:01 PM GMT+08:00 | [Old GCP]Couch can go up to 4 position in vertical direction when Pressing “Horizontal In” and “Load”button | WIP25 | WIP32 | don't press "horizontal in" button with load button at same time | Subsystem | SP1_PM_Approved | NA | Fixed | yang_h | 0 | 3 | Gantry Software | GCX | YES | Closed | WIP32SubCom | IQon1 |  | WIP32 | September 14, 2020 11:00:11 AM GMT+08:00
CTUDT00630523 | Refresh | Closed | lin_s | July 27, 2020 9:24:18 AM GMT+08:00 | Refresh: CT-NM.SRS.Modes.1863 didn't show right behavior when pulmo signal disconnect for 4D CT scan. | 4.1.10 | DHF298554 Rev15 |  | System |  | Modes.1863 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | SRS revision | Brilliance 64 |  | DHF298554 Rev15 | August 6, 2020 8:22:32 PM GMT+08:00
CTUDT00630523 | Refresh | Closed | lin_s | July 27, 2020 9:24:18 AM GMT+08:00 | Refresh: CT-NM.SRS.Modes.1863 didn't show right behavior when pulmo signal disconnect for 4D CT scan. | 4.1.10 | DHF298554 Rev15 |  | System |  | Modes.1863 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | SRS revision | Brilliance Big Bore |  | DHF298554 Rev15 | August 6, 2020 8:22:32 PM GMT+08:00
CTUDT00630523 | Refresh | Closed | lin_s | July 27, 2020 9:24:18 AM GMT+08:00 | Refresh: CT-NM.SRS.Modes.1863 didn't show right behavior when pulmo signal disconnect for 4D CT scan. | 4.1.10 | DHF298554 Rev15 |  | System |  | Modes.1863 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | SRS revision | Ingenuity 64 |  | DHF298554 Rev15 | August 6, 2020 8:22:32 PM GMT+08:00
CTUDT00630521 | Refresh | Assigned | cai_t | July 27, 2020 9:17:45 AM GMT+08:00 | The Slice Increment displayed a negative number in image parameter page after offline recon. | WIP25 |  |  | Subsystem | SP2_Postpone | NA |  | s_srivastava | 0 | 2 |  | Console4 | NO | Assigned | SP2 | Ovation
CTUDT00630521 | Refresh | Assigned | cai_t | July 27, 2020 9:17:45 AM GMT+08:00 | The Slice Increment displayed a negative number in image parameter page after offline recon. | WIP25 |  |  | Subsystem | SP2_Postpone | NA |  | s_srivastava | 0 | 2 |  | Console4 | NO | Assigned | SP2 | Brilliance 64
CTUDT00630521 | Refresh | Assigned | cai_t | July 27, 2020 9:17:45 AM GMT+08:00 | The Slice Increment displayed a negative number in image parameter page after offline recon. | WIP25 |  |  | Subsystem | SP2_Postpone | NA |  | s_srivastava | 0 | 2 |  | Console4 | NO | Assigned | SP2 | Brilliance Big Bore
CTUDT00630521 | Refresh | Assigned | cai_t | July 27, 2020 9:17:45 AM GMT+08:00 | The Slice Increment displayed a negative number in image parameter page after offline recon. | WIP25 |  |  | Subsystem | SP2_Postpone | NA |  | s_srivastava | 0 | 2 |  | Console4 | NO | Assigned | SP2 | Ingenuity 64
CTUDT00630521 | Refresh | Assigned | cai_t | July 27, 2020 9:17:45 AM GMT+08:00 | The Slice Increment displayed a negative number in image parameter page after offline recon. | WIP25 |  |  | Subsystem | SP2_Postpone | NA |  | s_srivastava | 0 | 2 |  | Console4 | NO | Assigned | SP2 | IQon1
CTUDT00630519 | Refresh | Closed | cai_t | July 27, 2020 8:52:22 AM GMT+08:00 | The Cardiac Phase displayed 75A% after adjusting ECG wave. | WIP25 |  |  | Subsystem |  | NA | No Plan to Fix | janakiraman_j | 0 | 4 |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00630519 | Refresh | Closed | cai_t | July 27, 2020 8:52:22 AM GMT+08:00 | The Cardiac Phase displayed 75A% after adjusting ECG wave. | WIP25 |  |  | Subsystem |  | NA | No Plan to Fix | janakiraman_j | 0 | 4 |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00630519 | Refresh | Closed | cai_t | July 27, 2020 8:52:22 AM GMT+08:00 | The Cardiac Phase displayed 75A% after adjusting ECG wave. | WIP25 |  |  | Subsystem |  | NA | No Plan to Fix | janakiraman_j | 0 | 4 |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00630512 | Refresh | Closed | dongping_w | July 26, 2020 8:50:46 PM GMT+08:00 | Refresh-BB 4D only can use with a minimum breathing rate of 6bpm and a max breathing rate of 30bpm | 4.8 |  |  | System |  |  | Works As Specified / Not A Defect | dongping_w | 0 |  |  |  |  | Closed |  | Brilliance Big Bore
CTUDT00630506 | Refresh | Closed | yatziv_l | July 26, 2020 2:20:46 PM GMT+08:00 | IB: High rates of XRT failure in the CTDI free air test (FDA test 307) | 4.1.7 | WIP31 |  | System | SP1_PM_Approved | NA | Fixed | challa_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP31SegReq | Ovation |  | WIP31 | August 19, 2020 3:01:25 PM GMT+08:00
CTUDT00630506 | Refresh | Closed | yatziv_l | July 26, 2020 2:20:46 PM GMT+08:00 | IB: High rates of XRT failure in the CTDI free air test (FDA test 307) | 4.1.7 | WIP31 |  | System | SP1_PM_Approved | NA | Fixed | challa_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP31SegReq | IQon1 |  | WIP31 | August 19, 2020 3:01:25 PM GMT+08:00
CTUDT00630495 | Refresh | Closed | s_namburi | July 23, 2020 10:00:09 PM GMT+08:00 | FoK: Keyboard language setting not retained as per preference setting language. Changes back to English keyboard in Login page after every restart. | 4.7.7 | WIP 25.4 | NA | Subsystem | FoK; PM_Approved | NA | Fixed | kumar_gau | 0 | 2 | Console Software | Console4 | NO | Closed | WIP25.4_SegReq | Ovation | PR197817 | WIP 25.4 | September 3, 2020 7:08:21 PM GMT+08:00
CTUDT00630495 | Refresh | Closed | s_namburi | July 23, 2020 10:00:09 PM GMT+08:00 | FoK: Keyboard language setting not retained as per preference setting language. Changes back to English keyboard in Login page after every restart. | 4.7.7 | WIP 25.4 | NA | Subsystem | FoK; PM_Approved | NA | Fixed | kumar_gau | 0 | 2 | Console Software | Console4 | NO | Closed | WIP25.4_SegReq | Brilliance 64 | PR197817 | WIP 25.4 | September 3, 2020 7:08:21 PM GMT+08:00
CTUDT00630495 | Refresh | Closed | s_namburi | July 23, 2020 10:00:09 PM GMT+08:00 | FoK: Keyboard language setting not retained as per preference setting language. Changes back to English keyboard in Login page after every restart. | 4.7.7 | WIP 25.4 | NA | Subsystem | FoK; PM_Approved | NA | Fixed | kumar_gau | 0 | 2 | Console Software | Console4 | NO | Closed | WIP25.4_SegReq | Brilliance Big Bore | PR197817 | WIP 25.4 | September 3, 2020 7:08:21 PM GMT+08:00
CTUDT00630495 | Refresh | Closed | s_namburi | July 23, 2020 10:00:09 PM GMT+08:00 | FoK: Keyboard language setting not retained as per preference setting language. Changes back to English keyboard in Login page after every restart. | 4.7.7 | WIP 25.4 | NA | Subsystem | FoK; PM_Approved | NA | Fixed | kumar_gau | 0 | 2 | Console Software | Console4 | NO | Closed | WIP25.4_SegReq | Ingenuity 64 | PR197817 | WIP 25.4 | September 3, 2020 7:08:21 PM GMT+08:00
CTUDT00630495 | Refresh | Closed | s_namburi | July 23, 2020 10:00:09 PM GMT+08:00 | FoK: Keyboard language setting not retained as per preference setting language. Changes back to English keyboard in Login page after every restart. | 4.7.7 | WIP 25.4 | NA | Subsystem | FoK; PM_Approved | NA | Fixed | kumar_gau | 0 | 2 | Console Software | Console4 | NO | Closed | WIP25.4_SegReq | IQon1 | PR197817 | WIP 25.4 | September 3, 2020 7:08:21 PM GMT+08:00
CTUDT00630493 | Refresh | Closed | s_namburi | July 23, 2020 9:47:41 PM GMT+08:00 | FoK: Creating new user need HOST Restart, if not cannot login from new user created | 4.7.7 |  | Not Applicable | Subsystem | FoK | N/A | Works As Specified / Not A Defect | Prabhakar_Priyank_Muppidi | 1 |  |  | Support-Security | NO | Closed |  | IQon1 | PR197810
CTUDT00630491 | Refresh | Closed | lei_yang | July 23, 2020 2:35:33 PM GMT+08:00 | Couch start vertical moving down before reaching the horizontal outermost position by pressing unload button. | WIP25 | WIP25.2 |  | Subsystem | Related to FoK | NA | Fixed | cx_zhang | 0 | 2 | Couch Software | Couch | NO | Closed | WIP25.1 | Ovation |  | WIP25.2 | August 10, 2020 6:58:13 PM GMT+08:00
CTUDT00630491 | Refresh | Closed | lei_yang | July 23, 2020 2:35:33 PM GMT+08:00 | Couch start vertical moving down before reaching the horizontal outermost position by pressing unload button. | WIP25 | WIP25.2 |  | Subsystem | Related to FoK | NA | Fixed | cx_zhang | 0 | 2 | Couch Software | Couch | NO | Closed | WIP25.1 | Brilliance 64 |  | WIP25.2 | August 10, 2020 6:58:13 PM GMT+08:00
CTUDT00630491 | Refresh | Closed | lei_yang | July 23, 2020 2:35:33 PM GMT+08:00 | Couch start vertical moving down before reaching the horizontal outermost position by pressing unload button. | WIP25 | WIP25.2 |  | Subsystem | Related to FoK | NA | Fixed | cx_zhang | 0 | 2 | Couch Software | Couch | NO | Closed | WIP25.1 | Brilliance Big Bore |  | WIP25.2 | August 10, 2020 6:58:13 PM GMT+08:00
CTUDT00630491 | Refresh | Closed | lei_yang | July 23, 2020 2:35:33 PM GMT+08:00 | Couch start vertical moving down before reaching the horizontal outermost position by pressing unload button. | WIP25 | WIP25.2 |  | Subsystem | Related to FoK | NA | Fixed | cx_zhang | 0 | 2 | Couch Software | Couch | NO | Closed | WIP25.1 | Ingenuity 64 |  | WIP25.2 | August 10, 2020 6:58:13 PM GMT+08:00
CTUDT00630491 | Refresh | Closed | lei_yang | July 23, 2020 2:35:33 PM GMT+08:00 | Couch start vertical moving down before reaching the horizontal outermost position by pressing unload button. | WIP25 | WIP25.2 |  | Subsystem | Related to FoK | NA | Fixed | cx_zhang | 0 | 2 | Couch Software | Couch | NO | Closed | WIP25.1 | IQon1 |  | WIP25.2 | August 10, 2020 6:58:13 PM GMT+08:00
CTUDT00630480 | Refresh | Closed | shaham_t | July 22, 2020 11:48:27 PM GMT+08:00 | SRS Updates are required to support IEC61223-3-5 Ed 2.0 | 4.7.7, 4.1.10, 4.8 | DHF298556 Rev.15 DHF298553 Rev.16 DHF298554 Rev.14 DHF298555 Rev.10 |  | Subsystem | SP1_PM_Approved | CT-NM.SRS.RegStand.324; CT-NM.SRS.IQ.2558; CT-NM.SRS.IQ.2559; CT-NM.SRS.IQ.1523; CT-NM.SRS.IQ.1525 | Fixed | shaham_t | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | SRS New version | Ovation |  | DHF298556 Rev.15 DHF298553 Rev.16 DHF298554 Rev.14 DHF298555 Rev.10 | October 6, 2020 7:15:10 PM GMT+08:00
CTUDT00630480 | Refresh | Closed | shaham_t | July 22, 2020 11:48:27 PM GMT+08:00 | SRS Updates are required to support IEC61223-3-5 Ed 2.0 | 4.7.7, 4.1.10, 4.8 | DHF298556 Rev.15 DHF298553 Rev.16 DHF298554 Rev.14 DHF298555 Rev.10 |  | Subsystem | SP1_PM_Approved | CT-NM.SRS.RegStand.324; CT-NM.SRS.IQ.2558; CT-NM.SRS.IQ.2559; CT-NM.SRS.IQ.1523; CT-NM.SRS.IQ.1525 | Fixed | shaham_t | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | SRS New version | Brilliance 64 |  | DHF298556 Rev.15 DHF298553 Rev.16 DHF298554 Rev.14 DHF298555 Rev.10 | October 6, 2020 7:15:10 PM GMT+08:00
CTUDT00630480 | Refresh | Closed | shaham_t | July 22, 2020 11:48:27 PM GMT+08:00 | SRS Updates are required to support IEC61223-3-5 Ed 2.0 | 4.7.7, 4.1.10, 4.8 | DHF298556 Rev.15 DHF298553 Rev.16 DHF298554 Rev.14 DHF298555 Rev.10 |  | Subsystem | SP1_PM_Approved | CT-NM.SRS.RegStand.324; CT-NM.SRS.IQ.2558; CT-NM.SRS.IQ.2559; CT-NM.SRS.IQ.1523; CT-NM.SRS.IQ.1525 | Fixed | shaham_t | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | SRS New version | Brilliance Big Bore |  | DHF298556 Rev.15 DHF298553 Rev.16 DHF298554 Rev.14 DHF298555 Rev.10 | October 6, 2020 7:15:10 PM GMT+08:00
CTUDT00630480 | Refresh | Closed | shaham_t | July 22, 2020 11:48:27 PM GMT+08:00 | SRS Updates are required to support IEC61223-3-5 Ed 2.0 | 4.7.7, 4.1.10, 4.8 | DHF298556 Rev.15 DHF298553 Rev.16 DHF298554 Rev.14 DHF298555 Rev.10 |  | Subsystem | SP1_PM_Approved | CT-NM.SRS.RegStand.324; CT-NM.SRS.IQ.2558; CT-NM.SRS.IQ.2559; CT-NM.SRS.IQ.1523; CT-NM.SRS.IQ.1525 | Fixed | shaham_t | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | SRS New version | Ingenuity 64 |  | DHF298556 Rev.15 DHF298553 Rev.16 DHF298554 Rev.14 DHF298555 Rev.10 | October 6, 2020 7:15:10 PM GMT+08:00
CTUDT00630480 | Refresh | Closed | shaham_t | July 22, 2020 11:48:27 PM GMT+08:00 | SRS Updates are required to support IEC61223-3-5 Ed 2.0 | 4.7.7, 4.1.10, 4.8 | DHF298556 Rev.15 DHF298553 Rev.16 DHF298554 Rev.14 DHF298555 Rev.10 |  | Subsystem | SP1_PM_Approved | CT-NM.SRS.RegStand.324; CT-NM.SRS.IQ.2558; CT-NM.SRS.IQ.2559; CT-NM.SRS.IQ.1523; CT-NM.SRS.IQ.1525 | Fixed | shaham_t | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | SRS New version | IQon1 |  | DHF298556 Rev.15 DHF298553 Rev.16 DHF298554 Rev.14 DHF298555 Rev.10 | October 6, 2020 7:15:10 PM GMT+08:00
CTUDT00630474 | Refresh | Closed | s_namburi | July 22, 2020 8:43:29 PM GMT+08:00 | Fok: After Refresh Upgrade : cannot connect for diag the Site from PRS portal | 4.7.7 | 459801587371 Rev H 459801587381 Rev F 459801587391 Rev E 459801587361 Rev G 459801587331 Rev H 459801829121 Rev A | M2M can be used to access these applications from PRS | System | FoK; SP1_PM_Approved; Doc_Update | N/A | Fixed | shi_b | 0 | 3 | Console Software | Support-Service | NO | Closed | Service Manual | Ovation | PR197818 | 459801587371 Rev H 459801587381 Rev F 459801587391 Rev E 459801587361 Rev G 459801587331 Rev H 459801829121 Rev A | November 9, 2020 11:07:04 AM GMT+08:00
CTUDT00630474 | Refresh | Closed | s_namburi | July 22, 2020 8:43:29 PM GMT+08:00 | Fok: After Refresh Upgrade : cannot connect for diag the Site from PRS portal | 4.7.7 | 459801587371 Rev H 459801587381 Rev F 459801587391 Rev E 459801587361 Rev G 459801587331 Rev H 459801829121 Rev A | M2M can be used to access these applications from PRS | System | FoK; SP1_PM_Approved; Doc_Update | N/A | Fixed | shi_b | 0 | 3 | Console Software | Support-Service | NO | Closed | Service Manual | Brilliance 64 | PR197818 | 459801587371 Rev H 459801587381 Rev F 459801587391 Rev E 459801587361 Rev G 459801587331 Rev H 459801829121 Rev A | November 9, 2020 11:07:04 AM GMT+08:00
CTUDT00630474 | Refresh | Closed | s_namburi | July 22, 2020 8:43:29 PM GMT+08:00 | Fok: After Refresh Upgrade : cannot connect for diag the Site from PRS portal | 4.7.7 | 459801587371 Rev H 459801587381 Rev F 459801587391 Rev E 459801587361 Rev G 459801587331 Rev H 459801829121 Rev A | M2M can be used to access these applications from PRS | System | FoK; SP1_PM_Approved; Doc_Update | N/A | Fixed | shi_b | 0 | 3 | Console Software | Support-Service | NO | Closed | Service Manual | Brilliance Big Bore | PR197818 | 459801587371 Rev H 459801587381 Rev F 459801587391 Rev E 459801587361 Rev G 459801587331 Rev H 459801829121 Rev A | November 9, 2020 11:07:04 AM GMT+08:00
CTUDT00630474 | Refresh | Closed | s_namburi | July 22, 2020 8:43:29 PM GMT+08:00 | Fok: After Refresh Upgrade : cannot connect for diag the Site from PRS portal | 4.7.7 | 459801587371 Rev H 459801587381 Rev F 459801587391 Rev E 459801587361 Rev G 459801587331 Rev H 459801829121 Rev A | M2M can be used to access these applications from PRS | System | FoK; SP1_PM_Approved; Doc_Update | N/A | Fixed | shi_b | 0 | 3 | Console Software | Support-Service | NO | Closed | Service Manual | Ingenuity 64 | PR197818 | 459801587371 Rev H 459801587381 Rev F 459801587391 Rev E 459801587361 Rev G 459801587331 Rev H 459801829121 Rev A | November 9, 2020 11:07:04 AM GMT+08:00
CTUDT00630474 | Refresh | Closed | s_namburi | July 22, 2020 8:43:29 PM GMT+08:00 | Fok: After Refresh Upgrade : cannot connect for diag the Site from PRS portal | 4.7.7 | 459801587371 Rev H 459801587381 Rev F 459801587391 Rev E 459801587361 Rev G 459801587331 Rev H 459801829121 Rev A | M2M can be used to access these applications from PRS | System | FoK; SP1_PM_Approved; Doc_Update | N/A | Fixed | shi_b | 0 | 3 | Console Software | Support-Service | NO | Closed | Service Manual | IQon1 | PR197818 | 459801587371 Rev H 459801587381 Rev F 459801587391 Rev E 459801587361 Rev G 459801587331 Rev H 459801829121 Rev A | November 9, 2020 11:07:04 AM GMT+08:00
CTUDT00630463 | Refresh | Closed | zhong_x | July 22, 2020 10:59:05 AM GMT+08:00 | Intend of horizontal in/out fast in CT-NM.SRS.SysPerf.574 is not clear | WIP25 | DHF298554 Rev16 |  | Subsystem | SP1_PM_Approved | CT-NM.SRS.SysPerf.574 | Fixed | lin_s | 0 | 3 | Requirements Documentation | Support-System | NO | Closed | SRS Next version | Brilliance 64 |  | DHF298554 Rev16 | September 15, 2020 11:38:44 AM GMT+08:00
CTUDT00630463 | Refresh | Closed | zhong_x | July 22, 2020 10:59:05 AM GMT+08:00 | Intend of horizontal in/out fast in CT-NM.SRS.SysPerf.574 is not clear | WIP25 | DHF298554 Rev16 |  | Subsystem | SP1_PM_Approved | CT-NM.SRS.SysPerf.574 | Fixed | lin_s | 0 | 3 | Requirements Documentation | Support-System | NO | Closed | SRS Next version | Ingenuity 64 |  | DHF298554 Rev16 | September 15, 2020 11:38:44 AM GMT+08:00
CTUDT00630422 | Refresh | Closed | s_namburi | July 20, 2020 5:15:07 PM GMT+08:00 | FoK: After Service Reboot - Philips_Service User Account was not Accessible - Had to Reset Philips_Service user using IT_admin | 4.7.7 |  | N/A | Subsystem | FoK | N/A | Works As Specified / Not A Defect | Prabhakar_Priyank_Muppidi | 1 |  |  | Console4 | NO | Closed |  | IQon1 | PR192813
CTUDT00630417 | Refresh | Closed | s_namburi | July 20, 2020 4:51:34 PM GMT+08:00 | FoK: Both the New Factory Exam Cards and User restored Exam Cards are present in User Group after Upgrade | 4.7.7 | 459801693571 Rev. C 459801693541 Rev. C |  | Subsystem | FoK | N/A | Fixed | sela_a | 0 | 2 | Service Documentation | Support-Service | NO | Closed | Service Documents | Ovation | PR192844 PR203329 | 459801693571 Rev. C 459801693541 Rev. C | August 25, 2020 2:48:33 PM GMT+08:00
CTUDT00630417 | Refresh | Closed | s_namburi | July 20, 2020 4:51:34 PM GMT+08:00 | FoK: Both the New Factory Exam Cards and User restored Exam Cards are present in User Group after Upgrade | 4.7.7 | 459801693571 Rev. C 459801693541 Rev. C |  | Subsystem | FoK | N/A | Fixed | sela_a | 0 | 2 | Service Documentation | Support-Service | NO | Closed | Service Documents | Brilliance 64 | PR192844 PR203329 | 459801693571 Rev. C 459801693541 Rev. C | August 25, 2020 2:48:33 PM GMT+08:00
CTUDT00630417 | Refresh | Closed | s_namburi | July 20, 2020 4:51:34 PM GMT+08:00 | FoK: Both the New Factory Exam Cards and User restored Exam Cards are present in User Group after Upgrade | 4.7.7 | 459801693571 Rev. C 459801693541 Rev. C |  | Subsystem | FoK | N/A | Fixed | sela_a | 0 | 2 | Service Documentation | Support-Service | NO | Closed | Service Documents | Brilliance Big Bore | PR192844 PR203329 | 459801693571 Rev. C 459801693541 Rev. C | August 25, 2020 2:48:33 PM GMT+08:00
CTUDT00630417 | Refresh | Closed | s_namburi | July 20, 2020 4:51:34 PM GMT+08:00 | FoK: Both the New Factory Exam Cards and User restored Exam Cards are present in User Group after Upgrade | 4.7.7 | 459801693571 Rev. C 459801693541 Rev. C |  | Subsystem | FoK | N/A | Fixed | sela_a | 0 | 2 | Service Documentation | Support-Service | NO | Closed | Service Documents | Ingenuity 64 | PR192844 PR203329 | 459801693571 Rev. C 459801693541 Rev. C | August 25, 2020 2:48:33 PM GMT+08:00
CTUDT00630417 | Refresh | Closed | s_namburi | July 20, 2020 4:51:34 PM GMT+08:00 | FoK: Both the New Factory Exam Cards and User restored Exam Cards are present in User Group after Upgrade | 4.7.7 | 459801693571 Rev. C 459801693541 Rev. C |  | Subsystem | FoK | N/A | Fixed | sela_a | 0 | 2 | Service Documentation | Support-Service | NO | Closed | Service Documents | IQon1 | PR192844 PR203329 | 459801693571 Rev. C 459801693541 Rev. C | August 25, 2020 2:48:33 PM GMT+08:00
CTUDT00630415 | Refresh | Closed | s_namburi | July 20, 2020 3:35:10 PM GMT+08:00 | FoK: Duplicate Exam Cards Found during Exam Card Manager Export | 4.7.7 | 459801693541 Rev. C 459801693571 Rev. C |  | Subsystem | FoK | N/A | Fixed | sela_a | 0 | 4 | Service Documentation | Support-Service | NO | Closed | Service Documents | Ovation | PR192828 | 459801693541 Rev. C 459801693571 Rev. C | August 25, 2020 2:47:50 PM GMT+08:00
CTUDT00630415 | Refresh | Closed | s_namburi | July 20, 2020 3:35:10 PM GMT+08:00 | FoK: Duplicate Exam Cards Found during Exam Card Manager Export | 4.7.7 | 459801693541 Rev. C 459801693571 Rev. C |  | Subsystem | FoK | N/A | Fixed | sela_a | 0 | 4 | Service Documentation | Support-Service | NO | Closed | Service Documents | Brilliance 64 | PR192828 | 459801693541 Rev. C 459801693571 Rev. C | August 25, 2020 2:47:50 PM GMT+08:00
CTUDT00630415 | Refresh | Closed | s_namburi | July 20, 2020 3:35:10 PM GMT+08:00 | FoK: Duplicate Exam Cards Found during Exam Card Manager Export | 4.7.7 | 459801693541 Rev. C 459801693571 Rev. C |  | Subsystem | FoK | N/A | Fixed | sela_a | 0 | 4 | Service Documentation | Support-Service | NO | Closed | Service Documents | Brilliance Big Bore | PR192828 | 459801693541 Rev. C 459801693571 Rev. C | August 25, 2020 2:47:50 PM GMT+08:00
CTUDT00630415 | Refresh | Closed | s_namburi | July 20, 2020 3:35:10 PM GMT+08:00 | FoK: Duplicate Exam Cards Found during Exam Card Manager Export | 4.7.7 | 459801693541 Rev. C 459801693571 Rev. C |  | Subsystem | FoK | N/A | Fixed | sela_a | 0 | 4 | Service Documentation | Support-Service | NO | Closed | Service Documents | Ingenuity 64 | PR192828 | 459801693541 Rev. C 459801693571 Rev. C | August 25, 2020 2:47:50 PM GMT+08:00
CTUDT00630415 | Refresh | Closed | s_namburi | July 20, 2020 3:35:10 PM GMT+08:00 | FoK: Duplicate Exam Cards Found during Exam Card Manager Export | 4.7.7 | 459801693541 Rev. C 459801693571 Rev. C |  | Subsystem | FoK | N/A | Fixed | sela_a | 0 | 4 | Service Documentation | Support-Service | NO | Closed | Service Documents | IQon1 | PR192828 | 459801693541 Rev. C 459801693571 Rev. C | August 25, 2020 2:47:50 PM GMT+08:00
CTUDT00630411 | Refresh | Closed | shaham_t | July 20, 2020 1:34:33 PM GMT+08:00 | Refresh ORE: Acceptance and performance tests fail on iCT | WIP24.1 | WIP31 | N/A | Subsystem | SP1_PM_Approved | N/A | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP31SegReq | Ovation |  | WIP31 | July 22, 2020 7:22:00 PM GMT+08:00
CTUDT00630408 | Refresh | Resolved | lin_s | July 20, 2020 1:03:57 PM GMT+08:00 | FOK - Gantry Initialization error happens in FOK site: The Regional Medical Center | 4.1.10 | WIP33 | The user can acknowledge the error and re-submit the scan request. | Subsystem | FOK; SP1_PM_Approved | NA | Fixed | richards_b | 0 | 3 | CIRS_SW_SeriesProcess | CIRS | YES | Resolved | WIP33SubCom | Ovation |  | WIP33 | November 4, 2020 3:44:17 AM GMT+08:00
CTUDT00630408 | Refresh | Resolved | lin_s | July 20, 2020 1:03:57 PM GMT+08:00 | FOK - Gantry Initialization error happens in FOK site: The Regional Medical Center | 4.1.10 | WIP33 | The user can acknowledge the error and re-submit the scan request. | Subsystem | FOK; SP1_PM_Approved | NA | Fixed | richards_b | 0 | 3 | CIRS_SW_SeriesProcess | CIRS | YES | Resolved | WIP33SubCom | Brilliance 64 |  | WIP33 | November 4, 2020 3:44:17 AM GMT+08:00
CTUDT00630408 | Refresh | Resolved | lin_s | July 20, 2020 1:03:57 PM GMT+08:00 | FOK - Gantry Initialization error happens in FOK site: The Regional Medical Center | 4.1.10 | WIP33 | The user can acknowledge the error and re-submit the scan request. | Subsystem | FOK; SP1_PM_Approved | NA | Fixed | richards_b | 0 | 3 | CIRS_SW_SeriesProcess | CIRS | YES | Resolved | WIP33SubCom | Brilliance Big Bore |  | WIP33 | November 4, 2020 3:44:17 AM GMT+08:00
CTUDT00630408 | Refresh | Resolved | lin_s | July 20, 2020 1:03:57 PM GMT+08:00 | FOK - Gantry Initialization error happens in FOK site: The Regional Medical Center | 4.1.10 | WIP33 | The user can acknowledge the error and re-submit the scan request. | Subsystem | FOK; SP1_PM_Approved | NA | Fixed | richards_b | 0 | 3 | CIRS_SW_SeriesProcess | CIRS | YES | Resolved | WIP33SubCom | Ingenuity 64 |  | WIP33 | November 4, 2020 3:44:17 AM GMT+08:00
CTUDT00630408 | Refresh | Resolved | lin_s | July 20, 2020 1:03:57 PM GMT+08:00 | FOK - Gantry Initialization error happens in FOK site: The Regional Medical Center | 4.1.10 | WIP33 | The user can acknowledge the error and re-submit the scan request. | Subsystem | FOK; SP1_PM_Approved | NA | Fixed | richards_b | 0 | 3 | CIRS_SW_SeriesProcess | CIRS | YES | Resolved | WIP33SubCom | IQon1 |  | WIP33 | November 4, 2020 3:44:17 AM GMT+08:00
CTUDT00630396 | Refresh | Resolved | wang_l | July 16, 2020 12:58:51 PM GMT+08:00 | Refresh: "CTA" was missed in Preference>Windows Preset after upgrade by Upgrade utility | WIP30.1 | WIP36 | User can manually add the CTA row in preferences after upgrade | System | Doc_Update; SP1_PM_Approved | Console-SRS.17124 | Fixed | periasamy_g | 0 | 3 | Requirements Documentation | Console4 | NO | Resolved | SSRS | Ovation |  | WIP36 | December 11, 2020 12:36:56 AM GMT+08:00
CTUDT00630396 | Refresh | Resolved | wang_l | July 16, 2020 12:58:51 PM GMT+08:00 | Refresh: "CTA" was missed in Preference>Windows Preset after upgrade by Upgrade utility | WIP30.1 | WIP36 | User can manually add the CTA row in preferences after upgrade | System | Doc_Update; SP1_PM_Approved | Console-SRS.17124 | Fixed | periasamy_g | 0 | 3 | Requirements Documentation | Console4 | NO | Resolved | SSRS | Brilliance 64 |  | WIP36 | December 11, 2020 12:36:56 AM GMT+08:00
CTUDT00630396 | Refresh | Resolved | wang_l | July 16, 2020 12:58:51 PM GMT+08:00 | Refresh: "CTA" was missed in Preference>Windows Preset after upgrade by Upgrade utility | WIP30.1 | WIP36 | User can manually add the CTA row in preferences after upgrade | System | Doc_Update; SP1_PM_Approved | Console-SRS.17124 | Fixed | periasamy_g | 0 | 3 | Requirements Documentation | Console4 | NO | Resolved | SSRS | Brilliance Big Bore |  | WIP36 | December 11, 2020 12:36:56 AM GMT+08:00
CTUDT00630396 | Refresh | Resolved | wang_l | July 16, 2020 12:58:51 PM GMT+08:00 | Refresh: "CTA" was missed in Preference>Windows Preset after upgrade by Upgrade utility | WIP30.1 | WIP36 | User can manually add the CTA row in preferences after upgrade | System | Doc_Update; SP1_PM_Approved | Console-SRS.17124 | Fixed | periasamy_g | 0 | 3 | Requirements Documentation | Console4 | NO | Resolved | SSRS | Ingenuity 64 |  | WIP36 | December 11, 2020 12:36:56 AM GMT+08:00
CTUDT00630394 | Refresh | Opened | zhong_x | July 16, 2020 10:28:59 AM GMT+08:00 | White cover symbols were not mentioned in Symbol Glossary document | WIP12.6 |  | N/A | Subsystem |  | N/A |  | rajalakshmi_j | 0 | 3 |  | Support-System | NO | Opened | SP1-Mar-21 | Ovation
CTUDT00630394 | Refresh | Opened | zhong_x | July 16, 2020 10:28:59 AM GMT+08:00 | White cover symbols were not mentioned in Symbol Glossary document | WIP12.6 |  | N/A | Subsystem |  | N/A |  | rajalakshmi_j | 0 | 3 |  | Support-System | NO | Opened | SP1-Mar-21 | Brilliance Big Bore
CTUDT00630394 | Refresh | Opened | zhong_x | July 16, 2020 10:28:59 AM GMT+08:00 | White cover symbols were not mentioned in Symbol Glossary document | WIP12.6 |  | N/A | Subsystem |  | N/A |  | rajalakshmi_j | 0 | 3 |  | Support-System | NO | Opened | SP1-Mar-21 | Ingenuity 64
CTUDT00630394 | Refresh | Opened | zhong_x | July 16, 2020 10:28:59 AM GMT+08:00 | White cover symbols were not mentioned in Symbol Glossary document | WIP12.6 |  | N/A | Subsystem |  | N/A |  | rajalakshmi_j | 0 | 3 |  | Support-System | NO | Opened | SP1-Mar-21 | IQon1
CTUDT00630365 | Refresh | Postponed | lei_yang | July 14, 2020 11:28:16 AM GMT+08:00 | Load patient to extreme position and press release load button for several times, ESTOP will Open. | WIP22.5 |  |  | Subsystem |  | NA |  | tan_j | 0 | 3 |  | Couch | NO | Postponed |  | Ovation
CTUDT00630365 | Refresh | Postponed | lei_yang | July 14, 2020 11:28:16 AM GMT+08:00 | Load patient to extreme position and press release load button for several times, ESTOP will Open. | WIP22.5 |  |  | Subsystem |  | NA |  | tan_j | 0 | 3 |  | Couch | NO | Postponed |  | Brilliance 64
CTUDT00630365 | Refresh | Postponed | lei_yang | July 14, 2020 11:28:16 AM GMT+08:00 | Load patient to extreme position and press release load button for several times, ESTOP will Open. | WIP22.5 |  |  | Subsystem |  | NA |  | tan_j | 0 | 3 |  | Couch | NO | Postponed |  | Brilliance Big Bore
CTUDT00630365 | Refresh | Postponed | lei_yang | July 14, 2020 11:28:16 AM GMT+08:00 | Load patient to extreme position and press release load button for several times, ESTOP will Open. | WIP22.5 |  |  | Subsystem |  | NA |  | tan_j | 0 | 3 |  | Couch | NO | Postponed |  | Ingenuity 64
CTUDT00630365 | Refresh | Postponed | lei_yang | July 14, 2020 11:28:16 AM GMT+08:00 | Load patient to extreme position and press release load button for several times, ESTOP will Open. | WIP22.5 |  |  | Subsystem |  | NA |  | tan_j | 0 | 3 |  | Couch | NO | Postponed |  | IQon1
CTUDT00630361 | Refresh | Closed | zhong_x | July 13, 2020 8:41:43 PM GMT+08:00 | FOK-Patient Size missing from surview image | BB_WIP12.6 | WIP34 |  | Subsystem | FOK; SP1_PM_Approved | Console-SRS.4746 | Fixed | s_srivastava | 0 | 2 | Console Software | Console4 | NO | Closed | WIP34SubCom | Ovation |  | WIP34 | November 10, 2020 3:30:51 PM GMT+08:00
CTUDT00630361 | Refresh | Closed | zhong_x | July 13, 2020 8:41:43 PM GMT+08:00 | FOK-Patient Size missing from surview image | BB_WIP12.6 | WIP34 |  | Subsystem | FOK; SP1_PM_Approved | Console-SRS.4746 | Fixed | s_srivastava | 0 | 2 | Console Software | Console4 | NO | Closed | WIP34SubCom | Brilliance 64 |  | WIP34 | November 10, 2020 3:30:51 PM GMT+08:00
CTUDT00630361 | Refresh | Closed | zhong_x | July 13, 2020 8:41:43 PM GMT+08:00 | FOK-Patient Size missing from surview image | BB_WIP12.6 | WIP34 |  | Subsystem | FOK; SP1_PM_Approved | Console-SRS.4746 | Fixed | s_srivastava | 0 | 2 | Console Software | Console4 | NO | Closed | WIP34SubCom | Brilliance Big Bore |  | WIP34 | November 10, 2020 3:30:51 PM GMT+08:00
CTUDT00630361 | Refresh | Closed | zhong_x | July 13, 2020 8:41:43 PM GMT+08:00 | FOK-Patient Size missing from surview image | BB_WIP12.6 | WIP34 |  | Subsystem | FOK; SP1_PM_Approved | Console-SRS.4746 | Fixed | s_srivastava | 0 | 2 | Console Software | Console4 | NO | Closed | WIP34SubCom | Ingenuity 64 |  | WIP34 | November 10, 2020 3:30:51 PM GMT+08:00
CTUDT00630361 | Refresh | Closed | zhong_x | July 13, 2020 8:41:43 PM GMT+08:00 | FOK-Patient Size missing from surview image | BB_WIP12.6 | WIP34 |  | Subsystem | FOK; SP1_PM_Approved | Console-SRS.4746 | Fixed | s_srivastava | 0 | 2 | Console Software | Console4 | NO | Closed | WIP34SubCom | IQon1 |  | WIP34 | November 10, 2020 3:30:51 PM GMT+08:00
CTUDT00630350 | Refresh | Closed | zhang_l | July 12, 2020 8:38:26 PM GMT+08:00 | DMS HW12 deviation may occurred on WIP22.5 after restart gantry | WIP22.5 | 459801587361 RevG 459801587351 RevG 459801587381 RevF 459801587391 RevE 459801853871 RevB |  | Subsystem |  | N/A | Fixed | shi_b | 0 | 4 | DMS Software | Support-Service | NO | Closed | com-24-oct-2020 | Brilliance 64 |  | 459801587361 RevG 459801587351 RevG 459801587381 RevF 459801587391 RevE 459801853871 RevB | October 29, 2020 2:40:42 PM GMT+08:00
CTUDT00630350 | Refresh | Closed | zhang_l | July 12, 2020 8:38:26 PM GMT+08:00 | DMS HW12 deviation may occurred on WIP22.5 after restart gantry | WIP22.5 | 459801587361 RevG 459801587351 RevG 459801587381 RevF 459801587391 RevE 459801853871 RevB |  | Subsystem |  | N/A | Fixed | shi_b | 0 | 4 | DMS Software | Support-Service | NO | Closed | com-24-oct-2020 | Ingenuity 64 |  | 459801587361 RevG 459801587351 RevG 459801587381 RevF 459801587391 RevE 459801853871 RevB | October 29, 2020 2:40:42 PM GMT+08:00
CTUDT00630346 | Refresh | Accepted | zhang_l | July 12, 2020 8:18:51 PM GMT+08:00 | “Snagit has stopped working” message popup during installation process. | WIP22.5 |  |  | Subsystem | SP2_Candidate |  |  | Prabhakar_Priyank_Muppidi | 0 |  |  | Console4 | NO | Accepted |  | Brilliance 64
CTUDT00630325 | Refresh | Resolved | lin_s | July 9, 2020 10:13:44 AM GMT+08:00 | FOK - Couch horizontal motion stop by mechanical stop | RFD version | WIP32 |  | Subsystem | fok; SP1_PM_Approved | NA | Fixed | cx_zhang | 0 | 2 | Couch Electrical | Couch | YES | Resolved | WIP32SubCom | Ovation | PR# 159883 | WIP32 | September 28, 2020 9:05:47 AM GMT+08:00
CTUDT00630325 | Refresh | Resolved | lin_s | July 9, 2020 10:13:44 AM GMT+08:00 | FOK - Couch horizontal motion stop by mechanical stop | RFD version | WIP32 |  | Subsystem | fok; SP1_PM_Approved | NA | Fixed | cx_zhang | 0 | 2 | Couch Electrical | Couch | YES | Resolved | WIP32SubCom | Brilliance 64 | PR# 159883 | WIP32 | September 28, 2020 9:05:47 AM GMT+08:00
CTUDT00630325 | Refresh | Resolved | lin_s | July 9, 2020 10:13:44 AM GMT+08:00 | FOK - Couch horizontal motion stop by mechanical stop | RFD version | WIP32 |  | Subsystem | fok; SP1_PM_Approved | NA | Fixed | cx_zhang | 0 | 2 | Couch Electrical | Couch | YES | Resolved | WIP32SubCom | Brilliance Big Bore | PR# 159883 | WIP32 | September 28, 2020 9:05:47 AM GMT+08:00
CTUDT00630325 | Refresh | Resolved | lin_s | July 9, 2020 10:13:44 AM GMT+08:00 | FOK - Couch horizontal motion stop by mechanical stop | RFD version | WIP32 |  | Subsystem | fok; SP1_PM_Approved | NA | Fixed | cx_zhang | 0 | 2 | Couch Electrical | Couch | YES | Resolved | WIP32SubCom | Ingenuity 64 | PR# 159883 | WIP32 | September 28, 2020 9:05:47 AM GMT+08:00
CTUDT00630325 | Refresh | Resolved | lin_s | July 9, 2020 10:13:44 AM GMT+08:00 | FOK - Couch horizontal motion stop by mechanical stop | RFD version | WIP32 |  | Subsystem | fok; SP1_PM_Approved | NA | Fixed | cx_zhang | 0 | 2 | Couch Electrical | Couch | YES | Resolved | WIP32SubCom | IQon1 | PR# 159883 | WIP32 | September 28, 2020 9:05:47 AM GMT+08:00
CTUDT00630324 | Refresh | Closed | shaham_t | July 8, 2020 11:59:48 PM GMT+08:00 | RFC: SSDE is required to be optional for the Chinese market | 4.7.7, 4.1.10 | WIP31 |  | Subsystem | SP1_PM_Approved | NA | Fixed | challa_s | 0 | 2 | Console-Console4 | Console4 | NO | Closed | WIP31SubCom | Ovation |  | WIP31 | July 21, 2020 12:25:14 PM GMT+08:00
CTUDT00630324 | Refresh | Closed | shaham_t | July 8, 2020 11:59:48 PM GMT+08:00 | RFC: SSDE is required to be optional for the Chinese market | 4.7.7, 4.1.10 | WIP31 |  | Subsystem | SP1_PM_Approved | NA | Fixed | challa_s | 0 | 2 | Console-Console4 | Console4 | NO | Closed | WIP31SubCom | Brilliance 64 |  | WIP31 | July 21, 2020 12:25:14 PM GMT+08:00
CTUDT00630324 | Refresh | Closed | shaham_t | July 8, 2020 11:59:48 PM GMT+08:00 | RFC: SSDE is required to be optional for the Chinese market | 4.7.7, 4.1.10 | WIP31 |  | Subsystem | SP1_PM_Approved | NA | Fixed | challa_s | 0 | 2 | Console-Console4 | Console4 | NO | Closed | WIP31SubCom | Ingenuity 64 |  | WIP31 | July 21, 2020 12:25:14 PM GMT+08:00
CTUDT00630324 | Refresh | Closed | shaham_t | July 8, 2020 11:59:48 PM GMT+08:00 | RFC: SSDE is required to be optional for the Chinese market | 4.7.7, 4.1.10 | WIP31 |  | Subsystem | SP1_PM_Approved | NA | Fixed | challa_s | 0 | 2 | Console-Console4 | Console4 | NO | Closed | WIP31SubCom | IQon1 |  | WIP31 | July 21, 2020 12:25:14 PM GMT+08:00
CTUDT00630252 | Refresh | Closed | allaberdiev_v | July 2, 2020 6:57:33 PM GMT+08:00 | Ingenuity and iCT: Typo in 3 cases of Performance Specification DHF170396: Tolerance is ±25 instead of ±30. | WIP30 | WIP31 |  | Subsystem | SP1_PM_Approved | n/a | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP31SubCom | Ovation |  | WIP31 | July 13, 2020 10:42:56 PM GMT+08:00
CTUDT00630252 | Refresh | Closed | allaberdiev_v | July 2, 2020 6:57:33 PM GMT+08:00 | Ingenuity and iCT: Typo in 3 cases of Performance Specification DHF170396: Tolerance is ±25 instead of ±30. | WIP30 | WIP31 |  | Subsystem | SP1_PM_Approved | n/a | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP31SubCom | Ingenuity 64 |  | WIP31 | July 13, 2020 10:42:56 PM GMT+08:00
CTUDT00630248 | Refresh | Opened | jeffr | July 2, 2020 3:04:12 PM GMT+08:00 | IB: High rates of XRT failure in the CTDI free air test (FDA test 307) | 4.1.10 and 4.7.7 | WIP31 wip25.10 |  | Subsystem | SP1_PM_Approved; HF3 | N/A | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | NO | Opened | WIP31SegReq | Ovation |  | WIP31 wip25.10 | July 27, 2020 11:51:49 PM GMT+08:00
CTUDT00630248 | Refresh | Opened | jeffr | July 2, 2020 3:04:12 PM GMT+08:00 | IB: High rates of XRT failure in the CTDI free air test (FDA test 307) | 4.1.10 and 4.7.7 | WIP31 wip25.10 |  | Subsystem | SP1_PM_Approved; HF3 | N/A | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | NO | Opened | WIP31SegReq | IQon1 |  | WIP31 wip25.10 | July 27, 2020 11:51:49 PM GMT+08:00
CTUDT00630245 | Refresh | Assigned | lin_s | July 2, 2020 2:19:05 PM GMT+08:00 | Ingenuity collimator failed in the opening calibration at Suzhou and Haifa site. | 4.1.7 |  |  | Subsystem | SP1_PM_Approved; Doc_update | NA |  | peng_lucas | 0 | 3 |  | GCX | NO | Assigned | New DMR | Ingenuity 64
CTUDT00630198 | Refresh | Closed | s_namburi | June 29, 2020 9:35:08 PM GMT+08:00 | Failed to Import few exam cards using EC import option in ECM | 4.1.10 | WIP31 |  | Subsystem | SP1_PM_Approved | NA | Fixed | challa_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP31SubCom | Brilliance 64 |  | WIP31 | July 16, 2020 8:50:20 PM GMT+08:00
CTUDT00630198 | Refresh | Closed | s_namburi | June 29, 2020 9:35:08 PM GMT+08:00 | Failed to Import few exam cards using EC import option in ECM | 4.1.10 | WIP31 |  | Subsystem | SP1_PM_Approved | NA | Fixed | challa_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP31SubCom | Ingenuity 64 |  | WIP31 | July 16, 2020 8:50:20 PM GMT+08:00
CTUDT00630190 | Refresh | Closed | lin_s | June 29, 2020 8:29:16 PM GMT+08:00 | FOK - Clinical got a critical error when exiting out the Exam Card Manager. | 4.1.10 | WIP31 | Logout and Login as CT user or system restart. | Subsystem | FOK; SP1_PM_Approved | NA | Fixed | srushti_k | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP31SegReq | Ovation |  | WIP31 | July 17, 2020 4:21:50 PM GMT+08:00
CTUDT00630190 | Refresh | Closed | lin_s | June 29, 2020 8:29:16 PM GMT+08:00 | FOK - Clinical got a critical error when exiting out the Exam Card Manager. | 4.1.10 | WIP31 | Logout and Login as CT user or system restart. | Subsystem | FOK; SP1_PM_Approved | NA | Fixed | srushti_k | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP31SegReq | Brilliance 64 |  | WIP31 | July 17, 2020 4:21:50 PM GMT+08:00
CTUDT00630190 | Refresh | Closed | lin_s | June 29, 2020 8:29:16 PM GMT+08:00 | FOK - Clinical got a critical error when exiting out the Exam Card Manager. | 4.1.10 | WIP31 | Logout and Login as CT user or system restart. | Subsystem | FOK; SP1_PM_Approved | NA | Fixed | srushti_k | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP31SegReq | Brilliance Big Bore |  | WIP31 | July 17, 2020 4:21:50 PM GMT+08:00
CTUDT00630190 | Refresh | Closed | lin_s | June 29, 2020 8:29:16 PM GMT+08:00 | FOK - Clinical got a critical error when exiting out the Exam Card Manager. | 4.1.10 | WIP31 | Logout and Login as CT user or system restart. | Subsystem | FOK; SP1_PM_Approved | NA | Fixed | srushti_k | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP31SegReq | Ingenuity 64 |  | WIP31 | July 17, 2020 4:21:50 PM GMT+08:00
CTUDT00630190 | Refresh | Closed | lin_s | June 29, 2020 8:29:16 PM GMT+08:00 | FOK - Clinical got a critical error when exiting out the Exam Card Manager. | 4.1.10 | WIP31 | Logout and Login as CT user or system restart. | Subsystem | FOK; SP1_PM_Approved | NA | Fixed | srushti_k | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP31SegReq | IQon1 |  | WIP31 | July 17, 2020 4:21:50 PM GMT+08:00
CTUDT00630186 | Refresh | Resolved | PADMAKUMAR_BS | June 29, 2020 4:44:32 PM GMT+08:00 | RADAR is unable to collect Huge USP logs due to missing UspLog Extractor tool in Refresh release | 4.7.7 | WIP32 |  | System | SP1_PM_Approved | NA | Fixed | yellapurkar_r | 0 | 3 | Service Tools_Software | Service Tools | NO | Resolved | WIP32SubCom | Ovation |  | WIP32 | September 21, 2020 7:26:51 PM GMT+08:00
CTUDT00630186 | Refresh | Resolved | PADMAKUMAR_BS | June 29, 2020 4:44:32 PM GMT+08:00 | RADAR is unable to collect Huge USP logs due to missing UspLog Extractor tool in Refresh release | 4.7.7 | WIP32 |  | System | SP1_PM_Approved | NA | Fixed | yellapurkar_r | 0 | 3 | Service Tools_Software | Service Tools | NO | Resolved | WIP32SubCom | Brilliance 64 |  | WIP32 | September 21, 2020 7:26:51 PM GMT+08:00
CTUDT00630186 | Refresh | Resolved | PADMAKUMAR_BS | June 29, 2020 4:44:32 PM GMT+08:00 | RADAR is unable to collect Huge USP logs due to missing UspLog Extractor tool in Refresh release | 4.7.7 | WIP32 |  | System | SP1_PM_Approved | NA | Fixed | yellapurkar_r | 0 | 3 | Service Tools_Software | Service Tools | NO | Resolved | WIP32SubCom | Brilliance Big Bore |  | WIP32 | September 21, 2020 7:26:51 PM GMT+08:00
CTUDT00630186 | Refresh | Resolved | PADMAKUMAR_BS | June 29, 2020 4:44:32 PM GMT+08:00 | RADAR is unable to collect Huge USP logs due to missing UspLog Extractor tool in Refresh release | 4.7.7 | WIP32 |  | System | SP1_PM_Approved | NA | Fixed | yellapurkar_r | 0 | 3 | Service Tools_Software | Service Tools | NO | Resolved | WIP32SubCom | Ingenuity 64 |  | WIP32 | September 21, 2020 7:26:51 PM GMT+08:00
CTUDT00630186 | Refresh | Resolved | PADMAKUMAR_BS | June 29, 2020 4:44:32 PM GMT+08:00 | RADAR is unable to collect Huge USP logs due to missing UspLog Extractor tool in Refresh release | 4.7.7 | WIP32 |  | System | SP1_PM_Approved | NA | Fixed | yellapurkar_r | 0 | 3 | Service Tools_Software | Service Tools | NO | Resolved | WIP32SubCom | IQon1 |  | WIP32 | September 21, 2020 7:26:51 PM GMT+08:00
CTUDT00630179 | Refresh | Closed | altshuller_mark | June 28, 2020 5:04:20 PM GMT+08:00 | AutoStoreAdapter Failed to update file name in Memory manager.Exception. | 4.7.5 | WIP12 |  | System |  | N/A | Fixed | janakiraman_j | 0 | 2 | Console Software | Console4 | YES | Closed | WIP32SegReq | IQon1 | PR9887200 | WIP12 | July 24, 2020 7:10:49 PM GMT+08:00
CTUDT00630178 | Refresh | Closed | altshuller_mark | June 28, 2020 4:58:36 PM GMT+08:00 | Application crash happens during Functional CT analysis. | 4.1.0 | WIP31 |  | System |  | N/A | Fixed | aswani_v | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP31SegReq | Ovation | TBD | WIP31 | July 13, 2020 2:27:53 PM GMT+08:00
CTUDT00630160 | Refresh | Closed | shaham_t | June 25, 2020 6:41:00 PM GMT+08:00 | FoK: Couch horizontal movement error casing the couch to stop moving | 4.1.10 |  |  | Subsystem | FOK | N/A | Duplicate | jt_zhao | 0 |  |  | Couch | NO | Closed |  | Ovation
CTUDT00630160 | Refresh | Closed | shaham_t | June 25, 2020 6:41:00 PM GMT+08:00 | FoK: Couch horizontal movement error casing the couch to stop moving | 4.1.10 |  |  | Subsystem | FOK | N/A | Duplicate | jt_zhao | 0 |  |  | Couch | NO | Closed |  | Brilliance 64
CTUDT00630160 | Refresh | Closed | shaham_t | June 25, 2020 6:41:00 PM GMT+08:00 | FoK: Couch horizontal movement error casing the couch to stop moving | 4.1.10 |  |  | Subsystem | FOK | N/A | Duplicate | jt_zhao | 0 |  |  | Couch | NO | Closed |  | Brilliance Big Bore
CTUDT00630160 | Refresh | Closed | shaham_t | June 25, 2020 6:41:00 PM GMT+08:00 | FoK: Couch horizontal movement error casing the couch to stop moving | 4.1.10 |  |  | Subsystem | FOK | N/A | Duplicate | jt_zhao | 0 |  |  | Couch | NO | Closed |  | Ingenuity 64
CTUDT00630160 | Refresh | Closed | shaham_t | June 25, 2020 6:41:00 PM GMT+08:00 | FoK: Couch horizontal movement error casing the couch to stop moving | 4.1.10 |  |  | Subsystem | FOK | N/A | Duplicate | jt_zhao | 0 |  |  | Couch | NO | Closed |  | IQon1
CTUDT00630142 | Refresh | Closed | zhang_t | June 24, 2020 10:34:26 AM GMT+08:00 | RFC: Improve Grid Space and Grid View type | BB_WIP13.4 |  |  | Subsystem |  | NA | Change Proposal | dongping_w | 0 | 5 |  | Support-System | NO | Closed |  | Brilliance Big Bore
CTUDT00630141 | Refresh | Resolved | karjagi_a | June 22, 2020 6:35:02 PM GMT+08:00 | RFC : Pre-Val: HAWK : (WIP15.1): Wrong direction of incremental movement button on IVC SW | HAWK | WIP36 | None | System | SP1_PM_Approved | N/A | Fixed | suresh_b | 0 | 2 | Console Generic | Console4 | YES | Resolved | WIP36SubCom | Brilliance Big Bore |  | WIP36 | December 7, 2020 2:02:11 PM GMT+08:00
CTUDT00630138 | Refresh | Closed | lin_s | June 22, 2020 4:40:46 PM GMT+08:00 | FOK - Couch horizontal motion stop by mechanical stop | 4.1.10 | WIP25 |  | Subsystem | FoK | N/A | Fixed | cx_zhang | 0 | 2 | Couch Software | Couch | YES | Closed | WIP25_SegReq | Ovation | PR# 159883 PR# 915108 | WIP25 | July 20, 2020 8:01:44 PM GMT+08:00
CTUDT00630138 | Refresh | Closed | lin_s | June 22, 2020 4:40:46 PM GMT+08:00 | FOK - Couch horizontal motion stop by mechanical stop | 4.1.10 | WIP25 |  | Subsystem | FoK | N/A | Fixed | cx_zhang | 0 | 2 | Couch Software | Couch | YES | Closed | WIP25_SegReq | Brilliance Big Bore | PR# 159883 PR# 915108 | WIP25 | July 20, 2020 8:01:44 PM GMT+08:00
CTUDT00630138 | Refresh | Closed | lin_s | June 22, 2020 4:40:46 PM GMT+08:00 | FOK - Couch horizontal motion stop by mechanical stop | 4.1.10 | WIP25 |  | Subsystem | FoK | N/A | Fixed | cx_zhang | 0 | 2 | Couch Software | Couch | YES | Closed | WIP25_SegReq | Ingenuity 64 | PR# 159883 PR# 915108 | WIP25 | July 20, 2020 8:01:44 PM GMT+08:00
CTUDT00630138 | Refresh | Closed | lin_s | June 22, 2020 4:40:46 PM GMT+08:00 | FOK - Couch horizontal motion stop by mechanical stop | 4.1.10 | WIP25 |  | Subsystem | FoK | N/A | Fixed | cx_zhang | 0 | 2 | Couch Software | Couch | YES | Closed | WIP25_SegReq | IQon1 | PR# 159883 PR# 915108 | WIP25 | July 20, 2020 8:01:44 PM GMT+08:00
CTUDT00630136 | Refresh | Closed | lin_s | June 22, 2020 3:48:27 PM GMT+08:00 | FOK - Adjusting knob on phantom holder is bent. some phantom damage. And phantom only include system phantom label, no head body phantom diff. | 4.1.10 |  | The workaround is to implement below design improvements on Harmonized Phantom. Issue 1&2: Add 1pcs of Adjustment screw and ANTIWEAR CAP into Body section assembly as replacement. Create a new 12NC (459801842051) including the Adjustment screw (459800441421) and ANTIWEAR CAP (459801585991), set it as FRU. Add requirement on the drawing to define that the adjustment screw need to be against the end for protection before putting in the Package box. Issue 3: Add part description on the label of each Phantom sections. Affected items: 459801578781: PHANTOM ASSEMBLY_BODY-40MM 459800442431: PHANTOM ASSEMBLY_BODY-80MM 459800443051: PHANTOM ASSEMBLY_HEAD 459800443701: PHANTOM ASSEMBLY_PHYSICS 459800444431: PHANTOM ASSEMBLY_INFANT Design change shall be initiated by ECR-074469 and implemented by E-127916. | Subsystem | FOK | N/A | Forwarded to Different Database | wang_ada | 0 | 3 |  | Couch | NO | Closed |  | Ovation | PR152140
CTUDT00630136 | Refresh | Closed | lin_s | June 22, 2020 3:48:27 PM GMT+08:00 | FOK - Adjusting knob on phantom holder is bent. some phantom damage. And phantom only include system phantom label, no head body phantom diff. | 4.1.10 |  | The workaround is to implement below design improvements on Harmonized Phantom. Issue 1&2: Add 1pcs of Adjustment screw and ANTIWEAR CAP into Body section assembly as replacement. Create a new 12NC (459801842051) including the Adjustment screw (459800441421) and ANTIWEAR CAP (459801585991), set it as FRU. Add requirement on the drawing to define that the adjustment screw need to be against the end for protection before putting in the Package box. Issue 3: Add part description on the label of each Phantom sections. Affected items: 459801578781: PHANTOM ASSEMBLY_BODY-40MM 459800442431: PHANTOM ASSEMBLY_BODY-80MM 459800443051: PHANTOM ASSEMBLY_HEAD 459800443701: PHANTOM ASSEMBLY_PHYSICS 459800444431: PHANTOM ASSEMBLY_INFANT Design change shall be initiated by ECR-074469 and implemented by E-127916. | Subsystem | FOK | N/A | Forwarded to Different Database | wang_ada | 0 | 3 |  | Couch | NO | Closed |  | Brilliance 64 | PR152140
CTUDT00630136 | Refresh | Closed | lin_s | June 22, 2020 3:48:27 PM GMT+08:00 | FOK - Adjusting knob on phantom holder is bent. some phantom damage. And phantom only include system phantom label, no head body phantom diff. | 4.1.10 |  | The workaround is to implement below design improvements on Harmonized Phantom. Issue 1&2: Add 1pcs of Adjustment screw and ANTIWEAR CAP into Body section assembly as replacement. Create a new 12NC (459801842051) including the Adjustment screw (459800441421) and ANTIWEAR CAP (459801585991), set it as FRU. Add requirement on the drawing to define that the adjustment screw need to be against the end for protection before putting in the Package box. Issue 3: Add part description on the label of each Phantom sections. Affected items: 459801578781: PHANTOM ASSEMBLY_BODY-40MM 459800442431: PHANTOM ASSEMBLY_BODY-80MM 459800443051: PHANTOM ASSEMBLY_HEAD 459800443701: PHANTOM ASSEMBLY_PHYSICS 459800444431: PHANTOM ASSEMBLY_INFANT Design change shall be initiated by ECR-074469 and implemented by E-127916. | Subsystem | FOK | N/A | Forwarded to Different Database | wang_ada | 0 | 3 |  | Couch | NO | Closed |  | Brilliance Big Bore | PR152140
CTUDT00630136 | Refresh | Closed | lin_s | June 22, 2020 3:48:27 PM GMT+08:00 | FOK - Adjusting knob on phantom holder is bent. some phantom damage. And phantom only include system phantom label, no head body phantom diff. | 4.1.10 |  | The workaround is to implement below design improvements on Harmonized Phantom. Issue 1&2: Add 1pcs of Adjustment screw and ANTIWEAR CAP into Body section assembly as replacement. Create a new 12NC (459801842051) including the Adjustment screw (459800441421) and ANTIWEAR CAP (459801585991), set it as FRU. Add requirement on the drawing to define that the adjustment screw need to be against the end for protection before putting in the Package box. Issue 3: Add part description on the label of each Phantom sections. Affected items: 459801578781: PHANTOM ASSEMBLY_BODY-40MM 459800442431: PHANTOM ASSEMBLY_BODY-80MM 459800443051: PHANTOM ASSEMBLY_HEAD 459800443701: PHANTOM ASSEMBLY_PHYSICS 459800444431: PHANTOM ASSEMBLY_INFANT Design change shall be initiated by ECR-074469 and implemented by E-127916. | Subsystem | FOK | N/A | Forwarded to Different Database | wang_ada | 0 | 3 |  | Couch | NO | Closed |  | Ingenuity 64 | PR152140
CTUDT00630136 | Refresh | Closed | lin_s | June 22, 2020 3:48:27 PM GMT+08:00 | FOK - Adjusting knob on phantom holder is bent. some phantom damage. And phantom only include system phantom label, no head body phantom diff. | 4.1.10 |  | The workaround is to implement below design improvements on Harmonized Phantom. Issue 1&2: Add 1pcs of Adjustment screw and ANTIWEAR CAP into Body section assembly as replacement. Create a new 12NC (459801842051) including the Adjustment screw (459800441421) and ANTIWEAR CAP (459801585991), set it as FRU. Add requirement on the drawing to define that the adjustment screw need to be against the end for protection before putting in the Package box. Issue 3: Add part description on the label of each Phantom sections. Affected items: 459801578781: PHANTOM ASSEMBLY_BODY-40MM 459800442431: PHANTOM ASSEMBLY_BODY-80MM 459800443051: PHANTOM ASSEMBLY_HEAD 459800443701: PHANTOM ASSEMBLY_PHYSICS 459800444431: PHANTOM ASSEMBLY_INFANT Design change shall be initiated by ECR-074469 and implemented by E-127916. | Subsystem | FOK | N/A | Forwarded to Different Database | wang_ada | 0 | 3 |  | Couch | NO | Closed |  | IQon1 | PR152140
CTUDT00630134 | Refresh | Closed | lin_s | June 22, 2020 3:22:44 PM GMT+08:00 | FOK - Cable Harness Rear Cover Close Up been pressed by screw hole | 4.1.10 |  | Design change will be done for rear cover drawings to add the cable tie 453567274081 to fix the cable by E-127321. | Subsystem | FOK | N/A | Forwarded to Different Database | shao_a | 0 | 4 |  | GCX | NO | Closed | DMR new version | Ingenuity 64 | PR152152
CTUDT00630113 | Refresh | Closed | zhong_x | June 18, 2020 1:03:43 PM GMT+08:00 | "Recon box(CIRS) cannot Initialize" message pop up when 3D Axial planned with Tilt | BB_WIP30 |  |  | Subsystem |  | N/A | Does Not Apply | 320057847 | 0 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00630112 | Refresh | Opened | dongping_w | June 18, 2020 11:27:32 AM GMT+08:00 | Refresh TRG Monthly CT Number Linearity and In Plane Spatial Integrity Check specs need to update when using with H-Phantom. | WIP24.2 |  | N/A | System |  | N/A |  | rajalakshmi_j | 0 | 2 |  | Support-System | NO | Opened | SP1-Mar-21 | Ovation
CTUDT00630112 | Refresh | Opened | dongping_w | June 18, 2020 11:27:32 AM GMT+08:00 | Refresh TRG Monthly CT Number Linearity and In Plane Spatial Integrity Check specs need to update when using with H-Phantom. | WIP24.2 |  | N/A | System |  | N/A |  | rajalakshmi_j | 0 | 2 |  | Support-System | NO | Opened | SP1-Mar-21 | IQon1
CTUDT00630111 | Refresh | Closed | dongping_w | June 18, 2020 11:24:36 AM GMT+08:00 | Refresh TRG Monthly CT Number Linearity and In Plane Spatial Integrity Check specs need to update when using with H-Phantom. | WIP22.5 |  |  | System |  | N/A | Works As Specified / Not A Defect | lin_s | 0 |  |  | Support-System | NO | Closed |  | Ingenuity 64
CTUDT00630109 | Refresh | Opened | dongping_w | June 18, 2020 10:49:57 AM GMT+08:00 | Refresh TRG Monthly CT Number Linearity and In Plane Spatial Integrity Check specs need to update when using with H-Phantom. | WIP12.6 |  |  | Subsystem | SP1_PM_Approved | na |  | rajalakshmi_j | 0 | 2 |  | Support-System | NO | Opened | SP1-Mar-21-CRN | Brilliance Big Bore
CTUDT00630107 | Refresh | Closed | zhong_x | June 18, 2020 10:33:16 AM GMT+08:00 | "Index was outside the bounds of the array" error occurred during IQ check with Carbon Top couch. | BB_WIP30 |  |  | Subsystem |  | n/a | No Plan to Fix | frider_o | 0 | 3 |  | DMS | NO | Closed | BB_SubNC | Ovation
CTUDT00630107 | Refresh | Closed | zhong_x | June 18, 2020 10:33:16 AM GMT+08:00 | "Index was outside the bounds of the array" error occurred during IQ check with Carbon Top couch. | BB_WIP30 |  |  | Subsystem |  | n/a | No Plan to Fix | frider_o | 0 | 3 |  | DMS | NO | Closed | BB_SubNC | Brilliance Big Bore
CTUDT00630107 | Refresh | Closed | zhong_x | June 18, 2020 10:33:16 AM GMT+08:00 | "Index was outside the bounds of the array" error occurred during IQ check with Carbon Top couch. | BB_WIP30 |  |  | Subsystem |  | n/a | No Plan to Fix | frider_o | 0 | 3 |  | DMS | NO | Closed | BB_SubNC | Ingenuity 64
CTUDT00630107 | Refresh | Closed | zhong_x | June 18, 2020 10:33:16 AM GMT+08:00 | "Index was outside the bounds of the array" error occurred during IQ check with Carbon Top couch. | BB_WIP30 |  |  | Subsystem |  | n/a | No Plan to Fix | frider_o | 0 | 3 |  | DMS | NO | Closed | BB_SubNC | IQon1
CTUDT00630105 | Refresh | Closed | zhong_x | June 18, 2020 9:14:49 AM GMT+08:00 | "There seems to be a problem in the application, Would you like to create a bug report" issue happen with 3D Axial oblique orientation. | BB_WIP30 |  |  | Subsystem |  | Console SRS.13184 | Does Not Apply | s_srivastava | 0 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00630085 | Refresh | Closed | yao_j | June 16, 2020 2:57:07 PM GMT+08:00 | TC Boost Improvement | IB | WIP33 |  | System | SP1_PM_Approved | DHF281164--34.2.1.2.1.8 | Fixed | zhou_x | 0 | 4 | Gantry Software | GCX | NO | Closed | WIP33SubCom | Brilliance 64 |  | WIP33 | October 26, 2020 1:21:02 PM GMT+08:00
CTUDT00630085 | Refresh | Closed | yao_j | June 16, 2020 2:57:07 PM GMT+08:00 | TC Boost Improvement | IB | WIP33 |  | System | SP1_PM_Approved | DHF281164--34.2.1.2.1.8 | Fixed | zhou_x | 0 | 4 | Gantry Software | GCX | NO | Closed | WIP33SubCom | Brilliance Big Bore |  | WIP33 | October 26, 2020 1:21:02 PM GMT+08:00
CTUDT00630085 | Refresh | Closed | yao_j | June 16, 2020 2:57:07 PM GMT+08:00 | TC Boost Improvement | IB | WIP33 |  | System | SP1_PM_Approved | DHF281164--34.2.1.2.1.8 | Fixed | zhou_x | 0 | 4 | Gantry Software | GCX | NO | Closed | WIP33SubCom | Ingenuity 64 |  | WIP33 | October 26, 2020 1:21:02 PM GMT+08:00
CTUDT00630073 | Refresh | Closed | shaham_t | June 15, 2020 9:39:46 PM GMT+08:00 | RFC: SSDE is required to be optional for the Chinese market | 4.7.7, 4.1.10 | WIP25 |  | Subsystem | SP1 and HotFix | Console-SRS.18283 | Fixed | challa_s | 0 | 2 | Console-Console4 | Console4 | NO | Closed | HotFix | Ovation |  | WIP25 | July 6, 2020 2:31:30 PM GMT+08:00
CTUDT00630073 | Refresh | Closed | shaham_t | June 15, 2020 9:39:46 PM GMT+08:00 | RFC: SSDE is required to be optional for the Chinese market | 4.7.7, 4.1.10 | WIP25 |  | Subsystem | SP1 and HotFix | Console-SRS.18283 | Fixed | challa_s | 0 | 2 | Console-Console4 | Console4 | NO | Closed | HotFix | Brilliance 64 |  | WIP25 | July 6, 2020 2:31:30 PM GMT+08:00
CTUDT00630073 | Refresh | Closed | shaham_t | June 15, 2020 9:39:46 PM GMT+08:00 | RFC: SSDE is required to be optional for the Chinese market | 4.7.7, 4.1.10 | WIP25 |  | Subsystem | SP1 and HotFix | Console-SRS.18283 | Fixed | challa_s | 0 | 2 | Console-Console4 | Console4 | NO | Closed | HotFix | Ingenuity 64 |  | WIP25 | July 6, 2020 2:31:30 PM GMT+08:00
CTUDT00630073 | Refresh | Closed | shaham_t | June 15, 2020 9:39:46 PM GMT+08:00 | RFC: SSDE is required to be optional for the Chinese market | 4.7.7, 4.1.10 | WIP25 |  | Subsystem | SP1 and HotFix | Console-SRS.18283 | Fixed | challa_s | 0 | 2 | Console-Console4 | Console4 | NO | Closed | HotFix | IQon1 |  | WIP25 | July 6, 2020 2:31:30 PM GMT+08:00
CTUDT00630045 | Refresh | Postponed | zhang_l | June 12, 2020 10:25:20 PM GMT+08:00 | Cannot complete HCOR calibration by auto centering while Radiology Flat Top installed | WIP22.5 |  |  | Subsystem | CRN SP1 | n/a |  | frider_o | 0 | 3 |  | DMS | NO | Postponed |  | Ovation
CTUDT00630045 | Refresh | Postponed | zhang_l | June 12, 2020 10:25:20 PM GMT+08:00 | Cannot complete HCOR calibration by auto centering while Radiology Flat Top installed | WIP22.5 |  |  | Subsystem | CRN SP1 | n/a |  | frider_o | 0 | 3 |  | DMS | NO | Postponed |  | Brilliance 64
CTUDT00630045 | Refresh | Postponed | zhang_l | June 12, 2020 10:25:20 PM GMT+08:00 | Cannot complete HCOR calibration by auto centering while Radiology Flat Top installed | WIP22.5 |  |  | Subsystem | CRN SP1 | n/a |  | frider_o | 0 | 3 |  | DMS | NO | Postponed |  | Brilliance Big Bore
CTUDT00630045 | Refresh | Postponed | zhang_l | June 12, 2020 10:25:20 PM GMT+08:00 | Cannot complete HCOR calibration by auto centering while Radiology Flat Top installed | WIP22.5 |  |  | Subsystem | CRN SP1 | n/a |  | frider_o | 0 | 3 |  | DMS | NO | Postponed |  | Ingenuity 64
CTUDT00630045 | Refresh | Postponed | zhang_l | June 12, 2020 10:25:20 PM GMT+08:00 | Cannot complete HCOR calibration by auto centering while Radiology Flat Top installed | WIP22.5 |  |  | Subsystem | CRN SP1 | n/a |  | frider_o | 0 | 3 |  | DMS | NO | Postponed |  | IQon1
CTUDT00629986 | Refresh | Closed | 320085947 | June 7, 2020 9:43:44 PM GMT+08:00 | Refresh WIP30- message of missing image appears in different applications when loading the whole study although there’s no a missing image/s. | WIP30 |  | No | Subsystem |  | NA | No Plan to Fix | aswani_v | 0 | 4 |  | iCAP | NO | Closed |  | Ovation
CTUDT00629986 | Refresh | Closed | 320085947 | June 7, 2020 9:43:44 PM GMT+08:00 | Refresh WIP30- message of missing image appears in different applications when loading the whole study although there’s no a missing image/s. | WIP30 |  | No | Subsystem |  | NA | No Plan to Fix | aswani_v | 0 | 4 |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00629986 | Refresh | Closed | 320085947 | June 7, 2020 9:43:44 PM GMT+08:00 | Refresh WIP30- message of missing image appears in different applications when loading the whole study although there’s no a missing image/s. | WIP30 |  | No | Subsystem |  | NA | No Plan to Fix | aswani_v | 0 | 4 |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00629986 | Refresh | Closed | 320085947 | June 7, 2020 9:43:44 PM GMT+08:00 | Refresh WIP30- message of missing image appears in different applications when loading the whole study although there’s no a missing image/s. | WIP30 |  | No | Subsystem |  | NA | No Plan to Fix | aswani_v | 0 | 4 |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00629986 | Refresh | Closed | 320085947 | June 7, 2020 9:43:44 PM GMT+08:00 | Refresh WIP30- message of missing image appears in different applications when loading the whole study although there’s no a missing image/s. | WIP30 |  | No | Subsystem |  | NA | No Plan to Fix | aswani_v | 0 | 4 |  | iCAP | NO | Closed |  | IQon1
CTUDT00629985 | Refresh | Closed | 320085947 | June 7, 2020 9:38:46 PM GMT+08:00 | Refresh WIP30 - There’s no indication the Result “Radiation dose information” doesn’t suppose to open with any of the viewing applications. | WIP30 |  | No | Subsystem |  | NA | No Plan to Fix | aswani_v | 0 | 4 |  | iCAP | NO | Closed |  | Ovation
CTUDT00629985 | Refresh | Closed | 320085947 | June 7, 2020 9:38:46 PM GMT+08:00 | Refresh WIP30 - There’s no indication the Result “Radiation dose information” doesn’t suppose to open with any of the viewing applications. | WIP30 |  | No | Subsystem |  | NA | No Plan to Fix | aswani_v | 0 | 4 |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00629985 | Refresh | Closed | 320085947 | June 7, 2020 9:38:46 PM GMT+08:00 | Refresh WIP30 - There’s no indication the Result “Radiation dose information” doesn’t suppose to open with any of the viewing applications. | WIP30 |  | No | Subsystem |  | NA | No Plan to Fix | aswani_v | 0 | 4 |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00629985 | Refresh | Closed | 320085947 | June 7, 2020 9:38:46 PM GMT+08:00 | Refresh WIP30 - There’s no indication the Result “Radiation dose information” doesn’t suppose to open with any of the viewing applications. | WIP30 |  | No | Subsystem |  | NA | No Plan to Fix | aswani_v | 0 | 4 |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00629985 | Refresh | Closed | 320085947 | June 7, 2020 9:38:46 PM GMT+08:00 | Refresh WIP30 - There’s no indication the Result “Radiation dose information” doesn’t suppose to open with any of the viewing applications. | WIP30 |  | No | Subsystem |  | NA | No Plan to Fix | aswani_v | 0 | 4 |  | iCAP | NO | Closed |  | IQon1
CTUDT00629930 | Refresh | Assigned | lin_s | May 29, 2020 3:04:23 PM GMT+08:00 | No package can enable Encryption license for Refresh products | 4.1.10, 4.8, 4.7.7 |  |  | System | SP1_PM_Approved | N/A |  | s_namburi | 0 | 2 |  | Support-System | NO | Assigned | 20-Dec-2020 | Ovation
CTUDT00629930 | Refresh | Assigned | lin_s | May 29, 2020 3:04:23 PM GMT+08:00 | No package can enable Encryption license for Refresh products | 4.1.10, 4.8, 4.7.7 |  |  | System | SP1_PM_Approved | N/A |  | s_namburi | 0 | 2 |  | Support-System | NO | Assigned | 20-Dec-2020 | Brilliance 64
CTUDT00629930 | Refresh | Assigned | lin_s | May 29, 2020 3:04:23 PM GMT+08:00 | No package can enable Encryption license for Refresh products | 4.1.10, 4.8, 4.7.7 |  |  | System | SP1_PM_Approved | N/A |  | s_namburi | 0 | 2 |  | Support-System | NO | Assigned | 20-Dec-2020 | Brilliance Big Bore
CTUDT00629930 | Refresh | Assigned | lin_s | May 29, 2020 3:04:23 PM GMT+08:00 | No package can enable Encryption license for Refresh products | 4.1.10, 4.8, 4.7.7 |  |  | System | SP1_PM_Approved | N/A |  | s_namburi | 0 | 2 |  | Support-System | NO | Assigned | 20-Dec-2020 | Ingenuity 64
CTUDT00629930 | Refresh | Assigned | lin_s | May 29, 2020 3:04:23 PM GMT+08:00 | No package can enable Encryption license for Refresh products | 4.1.10, 4.8, 4.7.7 |  |  | System | SP1_PM_Approved | N/A |  | s_namburi | 0 | 2 |  | Support-System | NO | Assigned | 20-Dec-2020 | IQon1
CTUDT00629920 | Refresh | Closed | zhang_l | May 27, 2020 10:42:21 PM GMT+08:00 | Gaps of Instruction, Help Page and Physics Spec for Calibration with H-Phantom | WIP30.1 | WIP33 |  | Subsystem | FOK; SP1_PM_Approved | DHF179202 | Fixed | daibes_s | 0 | 4 | DMS Software | DMS | NO | Closed | WIP33SubCom | Ovation | 170518 | WIP33 | October 11, 2020 4:25:36 PM GMT+08:00
CTUDT00629920 | Refresh | Closed | zhang_l | May 27, 2020 10:42:21 PM GMT+08:00 | Gaps of Instruction, Help Page and Physics Spec for Calibration with H-Phantom | WIP30.1 | WIP33 |  | Subsystem | FOK; SP1_PM_Approved | DHF179202 | Fixed | daibes_s | 0 | 4 | DMS Software | DMS | NO | Closed | WIP33SubCom | Brilliance 64 | 170518 | WIP33 | October 11, 2020 4:25:36 PM GMT+08:00
CTUDT00629920 | Refresh | Closed | zhang_l | May 27, 2020 10:42:21 PM GMT+08:00 | Gaps of Instruction, Help Page and Physics Spec for Calibration with H-Phantom | WIP30.1 | WIP33 |  | Subsystem | FOK; SP1_PM_Approved | DHF179202 | Fixed | daibes_s | 0 | 4 | DMS Software | DMS | NO | Closed | WIP33SubCom | Brilliance Big Bore | 170518 | WIP33 | October 11, 2020 4:25:36 PM GMT+08:00
CTUDT00629920 | Refresh | Closed | zhang_l | May 27, 2020 10:42:21 PM GMT+08:00 | Gaps of Instruction, Help Page and Physics Spec for Calibration with H-Phantom | WIP30.1 | WIP33 |  | Subsystem | FOK; SP1_PM_Approved | DHF179202 | Fixed | daibes_s | 0 | 4 | DMS Software | DMS | NO | Closed | WIP33SubCom | Ingenuity 64 | 170518 | WIP33 | October 11, 2020 4:25:36 PM GMT+08:00
CTUDT00629920 | Refresh | Closed | zhang_l | May 27, 2020 10:42:21 PM GMT+08:00 | Gaps of Instruction, Help Page and Physics Spec for Calibration with H-Phantom | WIP30.1 | WIP33 |  | Subsystem | FOK; SP1_PM_Approved | DHF179202 | Fixed | daibes_s | 0 | 4 | DMS Software | DMS | NO | Closed | WIP33SubCom | IQon1 | 170518 | WIP33 | October 11, 2020 4:25:36 PM GMT+08:00
CTUDT00629917 | Refresh | Closed | cohen_y | May 27, 2020 9:10:34 PM GMT+08:00 | Automatic Acceptance does not comply with 61223-3-5, edition 2.0. | WIP24.1 | WIP31 |  | Subsystem | SP1_PM_Approved | DHF170395 | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP31SubCom | Ovation |  | WIP31 | July 5, 2020 6:44:44 PM GMT+08:00
CTUDT00629917 | Refresh | Closed | cohen_y | May 27, 2020 9:10:34 PM GMT+08:00 | Automatic Acceptance does not comply with 61223-3-5, edition 2.0. | WIP24.1 | WIP31 |  | Subsystem | SP1_PM_Approved | DHF170395 | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP31SubCom | Ingenuity 64 |  | WIP31 | July 5, 2020 6:44:44 PM GMT+08:00
CTUDT00629917 | Refresh | Closed | cohen_y | May 27, 2020 9:10:34 PM GMT+08:00 | Automatic Acceptance does not comply with 61223-3-5, edition 2.0. | WIP24.1 | WIP31 |  | Subsystem | SP1_PM_Approved | DHF170395 | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP31SubCom | Ingenuity BigBore |  | WIP31 | July 5, 2020 6:44:44 PM GMT+08:00
CTUDT00629917 | Refresh | Closed | cohen_y | May 27, 2020 9:10:34 PM GMT+08:00 | Automatic Acceptance does not comply with 61223-3-5, edition 2.0. | WIP24.1 | WIP31 |  | Subsystem | SP1_PM_Approved | DHF170395 | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP31SubCom | IQon1 |  | WIP31 | July 5, 2020 6:44:44 PM GMT+08:00
CTUDT00629915 | Refresh | Closed | cohen_y | May 27, 2020 8:40:28 PM GMT+08:00 | CTDI free air should include Pediatric scans to comply with IEC 61223-3-5 Edition 2.0 | WIP24.1 | WIP31 |  | Subsystem | SP1_PM_Approved | DHF170399 | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP31SubCom | Ovation |  | WIP31 | July 5, 2020 6:51:06 PM GMT+08:00
CTUDT00629915 | Refresh | Closed | cohen_y | May 27, 2020 8:40:28 PM GMT+08:00 | CTDI free air should include Pediatric scans to comply with IEC 61223-3-5 Edition 2.0 | WIP24.1 | WIP31 |  | Subsystem | SP1_PM_Approved | DHF170399 | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP31SubCom | Ingenuity 64 |  | WIP31 | July 5, 2020 6:51:06 PM GMT+08:00
CTUDT00629915 | Refresh | Closed | cohen_y | May 27, 2020 8:40:28 PM GMT+08:00 | CTDI free air should include Pediatric scans to comply with IEC 61223-3-5 Edition 2.0 | WIP24.1 | WIP31 |  | Subsystem | SP1_PM_Approved | DHF170399 | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP31SubCom | Ingenuity BigBore |  | WIP31 | July 5, 2020 6:51:06 PM GMT+08:00
CTUDT00629915 | Refresh | Closed | cohen_y | May 27, 2020 8:40:28 PM GMT+08:00 | CTDI free air should include Pediatric scans to comply with IEC 61223-3-5 Edition 2.0 | WIP24.1 | WIP31 |  | Subsystem | SP1_PM_Approved | DHF170399 | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP31SubCom | IQon1 |  | WIP31 | July 5, 2020 6:51:06 PM GMT+08:00
CTUDT00629903 | Refresh | Closed | lei_yang | May 27, 2020 3:39:02 PM GMT+08:00 | "Unexpected error during workflow creation" pops up when launch workflow. | WIP30.1 | WIP31 | Acknowledge the error message and relaunch the application. | Subsystem | SP1-candidate; SP1_PM_Approved; FoK | N/A | Fixed | srushti_k | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP31SegReq | Ovation | PR 192854 | WIP31 | July 17, 2020 3:52:24 PM GMT+08:00
CTUDT00629903 | Refresh | Closed | lei_yang | May 27, 2020 3:39:02 PM GMT+08:00 | "Unexpected error during workflow creation" pops up when launch workflow. | WIP30.1 | WIP31 | Acknowledge the error message and relaunch the application. | Subsystem | SP1-candidate; SP1_PM_Approved; FoK | N/A | Fixed | srushti_k | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP31SegReq | Brilliance 64 | PR 192854 | WIP31 | July 17, 2020 3:52:24 PM GMT+08:00
CTUDT00629903 | Refresh | Closed | lei_yang | May 27, 2020 3:39:02 PM GMT+08:00 | "Unexpected error during workflow creation" pops up when launch workflow. | WIP30.1 | WIP31 | Acknowledge the error message and relaunch the application. | Subsystem | SP1-candidate; SP1_PM_Approved; FoK | N/A | Fixed | srushti_k | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP31SegReq | Brilliance Big Bore | PR 192854 | WIP31 | July 17, 2020 3:52:24 PM GMT+08:00
CTUDT00629903 | Refresh | Closed | lei_yang | May 27, 2020 3:39:02 PM GMT+08:00 | "Unexpected error during workflow creation" pops up when launch workflow. | WIP30.1 | WIP31 | Acknowledge the error message and relaunch the application. | Subsystem | SP1-candidate; SP1_PM_Approved; FoK | N/A | Fixed | srushti_k | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP31SegReq | Ingenuity 64 | PR 192854 | WIP31 | July 17, 2020 3:52:24 PM GMT+08:00
CTUDT00629903 | Refresh | Closed | lei_yang | May 27, 2020 3:39:02 PM GMT+08:00 | "Unexpected error during workflow creation" pops up when launch workflow. | WIP30.1 | WIP31 | Acknowledge the error message and relaunch the application. | Subsystem | SP1-candidate; SP1_PM_Approved; FoK | N/A | Fixed | srushti_k | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP31SegReq | IQon1 | PR 192854 | WIP31 | July 17, 2020 3:52:24 PM GMT+08:00
CTUDT00629900 | Refresh | Closed | lei_yang | May 27, 2020 3:05:39 PM GMT+08:00 | There are some ECs that show 2 duplicate Surviews in the patient scene. | WIP30.1 |  |  | Subsystem | SP1-candidate | NA | No Plan to Fix | challa_s | 0 | 4 |  | Console4 | NO | Closed |  | Ovation
CTUDT00629900 | Refresh | Closed | lei_yang | May 27, 2020 3:05:39 PM GMT+08:00 | There are some ECs that show 2 duplicate Surviews in the patient scene. | WIP30.1 |  |  | Subsystem | SP1-candidate | NA | No Plan to Fix | challa_s | 0 | 4 |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00629900 | Refresh | Closed | lei_yang | May 27, 2020 3:05:39 PM GMT+08:00 | There are some ECs that show 2 duplicate Surviews in the patient scene. | WIP30.1 |  |  | Subsystem | SP1-candidate | NA | No Plan to Fix | challa_s | 0 | 4 |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00629900 | Refresh | Closed | lei_yang | May 27, 2020 3:05:39 PM GMT+08:00 | There are some ECs that show 2 duplicate Surviews in the patient scene. | WIP30.1 |  |  | Subsystem | SP1-candidate | NA | No Plan to Fix | challa_s | 0 | 4 |  | Console4 | NO | Closed |  | IQon1
CTUDT00629831 | Refresh | Closed | zhang_l | May 20, 2020 1:47:27 PM GMT+08:00 | DMS deviations existed after CPM burnt. | WIP30 | WIP31 | build the SP1 1-click with separate labels according to products: * Philips_MS_DMS_FW_4.9.0.20 - Ingenuity+BR64+BB * Philips_MS_DMS_FW_4.9.0.33 - iCT and IQon | Subsystem | SP1_PM_Approved | N/A | Fixed | venugopal_s | 0 | 2 | DMS Electrical | DI | NO | Closed | WIP31SubCom | Brilliance 64 |  | WIP31 | November 27, 2020 4:19:53 PM GMT+08:00
CTUDT00629831 | Refresh | Closed | zhang_l | May 20, 2020 1:47:27 PM GMT+08:00 | DMS deviations existed after CPM burnt. | WIP30 | WIP31 | build the SP1 1-click with separate labels according to products: * Philips_MS_DMS_FW_4.9.0.20 - Ingenuity+BR64+BB * Philips_MS_DMS_FW_4.9.0.33 - iCT and IQon | Subsystem | SP1_PM_Approved | N/A | Fixed | venugopal_s | 0 | 2 | DMS Electrical | DI | NO | Closed | WIP31SubCom | Ingenuity 64 |  | WIP31 | November 27, 2020 4:19:53 PM GMT+08:00
CTUDT00629816 | Refresh | Closed | guy_sheaffer | May 19, 2020 5:03:09 PM GMT+08:00 | Refresh ORE: Acceptance and performance tests fail on iCT | WIP24.1 | WIP24.1 + patch | N/A | Subsystem | Acceptance; RfD | DHF17039, DHF170396 | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP24.2 | Ovation |  | WIP24.1 + patch | May 24, 2020 5:11:35 PM GMT+08:00
CTUDT00629785 | Refresh | Closed | zhang_l | May 15, 2020 11:28:27 PM GMT+08:00 | Auto centering did not work while Radiology Flat Top installed | WIP22.5 |  |  | Subsystem | Auto Centering; SP1-candidate | n/a | No Plan to Fix | frider_o | 0 | 3 |  | DMS | NO | Closed |  | Ovation
CTUDT00629785 | Refresh | Closed | zhang_l | May 15, 2020 11:28:27 PM GMT+08:00 | Auto centering did not work while Radiology Flat Top installed | WIP22.5 |  |  | Subsystem | Auto Centering; SP1-candidate | n/a | No Plan to Fix | frider_o | 0 | 3 |  | DMS | NO | Closed |  | Brilliance Big Bore
CTUDT00629785 | Refresh | Closed | zhang_l | May 15, 2020 11:28:27 PM GMT+08:00 | Auto centering did not work while Radiology Flat Top installed | WIP22.5 |  |  | Subsystem | Auto Centering; SP1-candidate | n/a | No Plan to Fix | frider_o | 0 | 3 |  | DMS | NO | Closed |  | Ingenuity 64
CTUDT00629785 | Refresh | Closed | zhang_l | May 15, 2020 11:28:27 PM GMT+08:00 | Auto centering did not work while Radiology Flat Top installed | WIP22.5 |  |  | Subsystem | Auto Centering; SP1-candidate | n/a | No Plan to Fix | frider_o | 0 | 3 |  | DMS | NO | Closed |  | IQon1
CTUDT00629755 | Refresh | Closed | zhong_x | May 13, 2020 8:59:00 PM GMT+08:00 | “Therapy Top (Optional)” description in Calibration help page was incorrect. | BB_WIP12.6 | WIP31 | N/A | Subsystem | SP1_PM_Approved | DHF179202 | Fixed | daibes_s | 0 | 3 | DMS Software | DMS | NO | Closed | WIP31SubCom | Ovation |  | WIP31 | July 22, 2020 9:23:12 PM GMT+08:00
CTUDT00629755 | Refresh | Closed | zhong_x | May 13, 2020 8:59:00 PM GMT+08:00 | “Therapy Top (Optional)” description in Calibration help page was incorrect. | BB_WIP12.6 | WIP31 | N/A | Subsystem | SP1_PM_Approved | DHF179202 | Fixed | daibes_s | 0 | 3 | DMS Software | DMS | NO | Closed | WIP31SubCom | Brilliance 64 |  | WIP31 | July 22, 2020 9:23:12 PM GMT+08:00
CTUDT00629755 | Refresh | Closed | zhong_x | May 13, 2020 8:59:00 PM GMT+08:00 | “Therapy Top (Optional)” description in Calibration help page was incorrect. | BB_WIP12.6 | WIP31 | N/A | Subsystem | SP1_PM_Approved | DHF179202 | Fixed | daibes_s | 0 | 3 | DMS Software | DMS | NO | Closed | WIP31SubCom | Brilliance Big Bore |  | WIP31 | July 22, 2020 9:23:12 PM GMT+08:00
CTUDT00629755 | Refresh | Closed | zhong_x | May 13, 2020 8:59:00 PM GMT+08:00 | “Therapy Top (Optional)” description in Calibration help page was incorrect. | BB_WIP12.6 | WIP31 | N/A | Subsystem | SP1_PM_Approved | DHF179202 | Fixed | daibes_s | 0 | 3 | DMS Software | DMS | NO | Closed | WIP31SubCom | Ingenuity 64 |  | WIP31 | July 22, 2020 9:23:12 PM GMT+08:00
CTUDT00629755 | Refresh | Closed | zhong_x | May 13, 2020 8:59:00 PM GMT+08:00 | “Therapy Top (Optional)” description in Calibration help page was incorrect. | BB_WIP12.6 | WIP31 | N/A | Subsystem | SP1_PM_Approved | DHF179202 | Fixed | daibes_s | 0 | 3 | DMS Software | DMS | NO | Closed | WIP31SubCom | IQon1 |  | WIP31 | July 22, 2020 9:23:12 PM GMT+08:00
CTUDT00629739 | Refresh | Resolved | PADMAKUMAR_BS | May 13, 2020 12:34:15 PM GMT+08:00 | Firewall configuration restricts FSE to transfer the IST certificate from FSE Laptop locally. | BB_WIP13.4 | WIP31 | to attach IST dongle to the HOST | Subsystem | SP1_PM_Approved | NA | Fixed | dasgupta_s | 0 | 2 | Console Software | Service Tools | NO | Resolved | WIP31SubCom | Ovation | 106696 | WIP31 | July 22, 2020 3:26:26 PM GMT+08:00
CTUDT00629739 | Refresh | Resolved | PADMAKUMAR_BS | May 13, 2020 12:34:15 PM GMT+08:00 | Firewall configuration restricts FSE to transfer the IST certificate from FSE Laptop locally. | BB_WIP13.4 | WIP31 | to attach IST dongle to the HOST | Subsystem | SP1_PM_Approved | NA | Fixed | dasgupta_s | 0 | 2 | Console Software | Service Tools | NO | Resolved | WIP31SubCom | Brilliance 64 | 106696 | WIP31 | July 22, 2020 3:26:26 PM GMT+08:00
CTUDT00629739 | Refresh | Resolved | PADMAKUMAR_BS | May 13, 2020 12:34:15 PM GMT+08:00 | Firewall configuration restricts FSE to transfer the IST certificate from FSE Laptop locally. | BB_WIP13.4 | WIP31 | to attach IST dongle to the HOST | Subsystem | SP1_PM_Approved | NA | Fixed | dasgupta_s | 0 | 2 | Console Software | Service Tools | NO | Resolved | WIP31SubCom | Brilliance Big Bore | 106696 | WIP31 | July 22, 2020 3:26:26 PM GMT+08:00
CTUDT00629739 | Refresh | Resolved | PADMAKUMAR_BS | May 13, 2020 12:34:15 PM GMT+08:00 | Firewall configuration restricts FSE to transfer the IST certificate from FSE Laptop locally. | BB_WIP13.4 | WIP31 | to attach IST dongle to the HOST | Subsystem | SP1_PM_Approved | NA | Fixed | dasgupta_s | 0 | 2 | Console Software | Service Tools | NO | Resolved | WIP31SubCom | Ingenuity 64 | 106696 | WIP31 | July 22, 2020 3:26:26 PM GMT+08:00
CTUDT00629739 | Refresh | Resolved | PADMAKUMAR_BS | May 13, 2020 12:34:15 PM GMT+08:00 | Firewall configuration restricts FSE to transfer the IST certificate from FSE Laptop locally. | BB_WIP13.4 | WIP31 | to attach IST dongle to the HOST | Subsystem | SP1_PM_Approved | NA | Fixed | dasgupta_s | 0 | 2 | Console Software | Service Tools | NO | Resolved | WIP31SubCom | IQon1 | 106696 | WIP31 | July 22, 2020 3:26:26 PM GMT+08:00
CTUDT00629727 | Refresh | Closed | altshuller_mark | May 12, 2020 6:23:10 PM GMT+08:00 | ExamApplication crash when the user tries to drag in ImageViewer during scan | IQon HB1 |  | na | Subsystem | SP1 | N/A | Works As Specified / Not A Defect | aswani_v | 1 | 3 |  | iCAP | YES | Closed |  | IQon1 | 9710457
CTUDT00629726 | Refresh | Closed | altshuller_mark | May 12, 2020 6:17:01 PM GMT+08:00 | A message asking bugrep creation popped up during dynamic scan sequence. | IQon HP1 |  | All the planned images could be generated by offline recon | Subsystem | SP1 | N/A | Works As Specified / Not A Defect | aswani_v | 1 | 4 |  | iCAP | NO | Closed |  | IQon1 | 10015403
CTUDT00629627 | Refresh | Resolved | molinari_c | May 5, 2020 2:27:23 AM GMT+08:00 | Service Requirement is needed for PSC to work with remote IST entitlements transfer | WIP24.1 | WIP31 |  | Subsystem | SP1_PM_Approved | STPS-SRS.1213 & STPS-SRS.1214 | Fixed | yellapurkar_r | 0 | 2 | Service Tools_Software | Service Tools | NO | Resolved | WIP31SubCom | Ovation |  | WIP31 | July 17, 2020 3:24:20 PM GMT+08:00
CTUDT00629627 | Refresh | Resolved | molinari_c | May 5, 2020 2:27:23 AM GMT+08:00 | Service Requirement is needed for PSC to work with remote IST entitlements transfer | WIP24.1 | WIP31 |  | Subsystem | SP1_PM_Approved | STPS-SRS.1213 & STPS-SRS.1214 | Fixed | yellapurkar_r | 0 | 2 | Service Tools_Software | Service Tools | NO | Resolved | WIP31SubCom | Brilliance 64 |  | WIP31 | July 17, 2020 3:24:20 PM GMT+08:00
CTUDT00629627 | Refresh | Resolved | molinari_c | May 5, 2020 2:27:23 AM GMT+08:00 | Service Requirement is needed for PSC to work with remote IST entitlements transfer | WIP24.1 | WIP31 |  | Subsystem | SP1_PM_Approved | STPS-SRS.1213 & STPS-SRS.1214 | Fixed | yellapurkar_r | 0 | 2 | Service Tools_Software | Service Tools | NO | Resolved | WIP31SubCom | Brilliance Big Bore |  | WIP31 | July 17, 2020 3:24:20 PM GMT+08:00
CTUDT00629627 | Refresh | Resolved | molinari_c | May 5, 2020 2:27:23 AM GMT+08:00 | Service Requirement is needed for PSC to work with remote IST entitlements transfer | WIP24.1 | WIP31 |  | Subsystem | SP1_PM_Approved | STPS-SRS.1213 & STPS-SRS.1214 | Fixed | yellapurkar_r | 0 | 2 | Service Tools_Software | Service Tools | NO | Resolved | WIP31SubCom | Ingenuity 64 |  | WIP31 | July 17, 2020 3:24:20 PM GMT+08:00
CTUDT00629627 | Refresh | Resolved | molinari_c | May 5, 2020 2:27:23 AM GMT+08:00 | Service Requirement is needed for PSC to work with remote IST entitlements transfer | WIP24.1 | WIP31 |  | Subsystem | SP1_PM_Approved | STPS-SRS.1213 & STPS-SRS.1214 | Fixed | yellapurkar_r | 0 | 2 | Service Tools_Software | Service Tools | NO | Resolved | WIP31SubCom | IQon1 |  | WIP31 | July 17, 2020 3:24:20 PM GMT+08:00
CTUDT00629559 | Refresh | Closed | 320053770 | April 26, 2020 5:04:00 PM GMT+08:00 | Refresh; WIP24.1; Reliability regression testing: E-stop was opened during Clinical Scan. | WIP24.1 |  |  | Subsystem | EStop | N/A | No Plan to Fix | Wang_Dexin | 0 | 3 |  | GCX | YES | Closed |  | IQon1
CTUDT00629545 | Refresh | Closed | zhang_l | April 23, 2020 10:16:35 AM GMT+08:00 | Cannot create automatic protocol (P3T) while loading Brain Perfusion Jog/ Non-Jog EC and scan time exceed 40 sec | WIP22.5 |  |  | Subsystem | P3T | Console-SRS.6216 | Works As Specified / Not A Defect | s_srivastava | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00629541 | Refresh | Closed | lei_yang | April 23, 2020 9:52:54 AM GMT+08:00 | No ECG offline wave display for Scan1 when the system automatically performs two scans consecutively (Scan1 uncheck cardiac DoseRight, Scan2 check cardiac DoseRight ). | WIP22.5 |  |  | Subsystem | SP1 | NA | No Plan to Fix | datta_i | 0 | 4 |  | Console4 | NO | Closed |  | Ovation
CTUDT00629541 | Refresh | Closed | lei_yang | April 23, 2020 9:52:54 AM GMT+08:00 | No ECG offline wave display for Scan1 when the system automatically performs two scans consecutively (Scan1 uncheck cardiac DoseRight, Scan2 check cardiac DoseRight ). | WIP22.5 |  |  | Subsystem | SP1 | NA | No Plan to Fix | datta_i | 0 | 4 |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00629541 | Refresh | Closed | lei_yang | April 23, 2020 9:52:54 AM GMT+08:00 | No ECG offline wave display for Scan1 when the system automatically performs two scans consecutively (Scan1 uncheck cardiac DoseRight, Scan2 check cardiac DoseRight ). | WIP22.5 |  |  | Subsystem | SP1 | NA | No Plan to Fix | datta_i | 0 | 4 |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00629541 | Refresh | Closed | lei_yang | April 23, 2020 9:52:54 AM GMT+08:00 | No ECG offline wave display for Scan1 when the system automatically performs two scans consecutively (Scan1 uncheck cardiac DoseRight, Scan2 check cardiac DoseRight ). | WIP22.5 |  |  | Subsystem | SP1 | NA | No Plan to Fix | datta_i | 0 | 4 |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00629541 | Refresh | Closed | lei_yang | April 23, 2020 9:52:54 AM GMT+08:00 | No ECG offline wave display for Scan1 when the system automatically performs two scans consecutively (Scan1 uncheck cardiac DoseRight, Scan2 check cardiac DoseRight ). | WIP22.5 |  |  | Subsystem | SP1 | NA | No Plan to Fix | datta_i | 0 | 4 |  | Console4 | NO | Closed |  | IQon1
CTUDT00629540 | Refresh | Closed | zhong_x | April 23, 2020 9:52:17 AM GMT+08:00 | Bigbore Refresh-Total Image numbers displayed queue manager not consistent | BB_WIP12.5 |  |  | Subsystem |  | NA | No Plan to Fix | s_panigrahi | 0 | 4 |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00629538 | Refresh | Closed | zhong_x | April 23, 2020 9:47:15 AM GMT+08:00 | Import DICOM Images from Import image Tool in desktop | BB_WIP12.5 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | s_panigrahi | 0 |  |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00629535 | Refresh | Closed | lei_yang | April 23, 2020 9:37:13 AM GMT+08:00 | After completing two cycles, the scan ruler shows only one cycle completed. | WIP22.5 |  |  | Subsystem |  | NA | No Plan to Fix | datta_i | 0 | 4 |  | Console4 | NO | Closed |  | Ovation
CTUDT00629535 | Refresh | Closed | lei_yang | April 23, 2020 9:37:13 AM GMT+08:00 | After completing two cycles, the scan ruler shows only one cycle completed. | WIP22.5 |  |  | Subsystem |  | NA | No Plan to Fix | datta_i | 0 | 4 |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00629535 | Refresh | Closed | lei_yang | April 23, 2020 9:37:13 AM GMT+08:00 | After completing two cycles, the scan ruler shows only one cycle completed. | WIP22.5 |  |  | Subsystem |  | NA | No Plan to Fix | datta_i | 0 | 4 |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00629535 | Refresh | Closed | lei_yang | April 23, 2020 9:37:13 AM GMT+08:00 | After completing two cycles, the scan ruler shows only one cycle completed. | WIP22.5 |  |  | Subsystem |  | NA | No Plan to Fix | datta_i | 0 | 4 |  | Console4 | NO | Closed |  | IQon1
CTUDT00629534 | Refresh | Closed | zhong_x | April 23, 2020 9:33:54 AM GMT+08:00 | Error"0 out of 1 series received" when query study from ISITE PACS | WIP12.5 |  |  | Subsystem | SP1 | NA | Invalid/Uncontrolled Configuration | lin_s | 0 |  |  | Support-System | NO | Closed |  | Brilliance Big Bore
CTUDT00629532 | Refresh | Closed | wang_l | April 23, 2020 9:16:32 AM GMT+08:00 | Refresh - Miss Hcor Basic and Hcor Slice physics history path for H phantom calibration in Backup/Restore Tool | WIP22.5 |  |  | Subsystem |  | n/a | Works As Specified / Not A Defect | frider_o | 0 |  |  | DMS | NO | Closed |  | Brilliance 64
CTUDT00629532 | Refresh | Closed | wang_l | April 23, 2020 9:16:32 AM GMT+08:00 | Refresh - Miss Hcor Basic and Hcor Slice physics history path for H phantom calibration in Backup/Restore Tool | WIP22.5 |  |  | Subsystem |  | n/a | Works As Specified / Not A Defect | frider_o | 0 |  |  | DMS | NO | Closed |  | Brilliance Big Bore
CTUDT00629532 | Refresh | Closed | wang_l | April 23, 2020 9:16:32 AM GMT+08:00 | Refresh - Miss Hcor Basic and Hcor Slice physics history path for H phantom calibration in Backup/Restore Tool | WIP22.5 |  |  | Subsystem |  | n/a | Works As Specified / Not A Defect | frider_o | 0 |  |  | DMS | NO | Closed |  | Ingenuity 64
CTUDT00629474 | Refresh | Closed | urim | April 16, 2020 6:27:52 PM GMT+08:00 | Refresh Validation: Wrong Phantom for IQ Check is Presented to CT User in On-Screen Help | WIP24.1 |  | Use the iCT System Phantom specified in iCT TRG (459801063557 Rev 05, page 61) | System |  | N/A | No Plan to Fix | braunstein_d | 0 | 3 |  | Support-Physics | NO | Closed |  | Ovation
CTUDT00629413 | Refresh | Closed | rojin_antony | April 7, 2020 7:35:04 PM GMT+08:00 | IQon Spectral CT Customer Release Note 459800975185_D : Validation Finding | WIP24.1 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | charles_j | 0 |  |  | Support-System | NO | Closed |  | IQon1
CTUDT00629409 | Refresh | Closed | glauberman_o | April 7, 2020 2:57:40 PM GMT+08:00 | Val:Refresh:System Error | WIP24.1 |  | Logout and login or host restart | Subsystem |  | NA | No Plan to Fix | s_sa | 0 | 3 |  | iCAP | NO | Closed |  | IQon1
CTUDT00629396 | Refresh | Closed | glauberman_o | April 6, 2020 6:12:46 PM GMT+08:00 | Val: Refresh:Initialization got stuck | WIP24.1 |  |  | Subsystem | Reliability; SP1 | NA | No Plan to Fix | datta_i | 0 | 3 |  | Console4 | NO | Closed |  | Ovation
CTUDT00629396 | Refresh | Closed | glauberman_o | April 6, 2020 6:12:46 PM GMT+08:00 | Val: Refresh:Initialization got stuck | WIP24.1 |  |  | Subsystem | Reliability; SP1 | NA | No Plan to Fix | datta_i | 0 | 3 |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00629396 | Refresh | Closed | glauberman_o | April 6, 2020 6:12:46 PM GMT+08:00 | Val: Refresh:Initialization got stuck | WIP24.1 |  |  | Subsystem | Reliability; SP1 | NA | No Plan to Fix | datta_i | 0 | 3 |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00629396 | Refresh | Closed | glauberman_o | April 6, 2020 6:12:46 PM GMT+08:00 | Val: Refresh:Initialization got stuck | WIP24.1 |  |  | Subsystem | Reliability; SP1 | NA | No Plan to Fix | datta_i | 0 | 3 |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00629396 | Refresh | Closed | glauberman_o | April 6, 2020 6:12:46 PM GMT+08:00 | Val: Refresh:Initialization got stuck | WIP24.1 |  |  | Subsystem | Reliability; SP1 | NA | No Plan to Fix | datta_i | 0 | 3 |  | Console4 | NO | Closed |  | IQon1
CTUDT00629370 | Refresh | Closed | Revital_Cohen | April 2, 2020 4:17:13 PM GMT+08:00 | Refresh Issue in the front page of the localization iFU of some languages | WIP24 |  |  | Subsystem |  | N/A | Works As Specified / Not A Defect | charles_j | 0 | 2 |  | Support-System | NO | Closed | IFU | IQon1
CTUDT00629350 | Refresh | Closed | barnea_d | April 1, 2020 6:29:08 PM GMT+08:00 | Refresh: Automatic centering phantom function on BB with therapy top not found | BB_WIP12.5 | Thrapy_Top_Update4.8.0.10136 |  | Subsystem |  | N/A | Fixed | frider_o | 0 | 3 | DMS Software | DMS | NO | Closed | BB_WIP12.6_SegReq | Brilliance Big Bore |  | Thrapy_Top_Update4.8.0.10136 | April 8, 2020 10:54:45 AM GMT+08:00
CTUDT00629348 | Refresh | Closed | Revital_Cohen | April 1, 2020 4:24:38 PM GMT+08:00 | Refresh: no IFU translation for Vietnamese - system support.875 | WIP24 | DHF298553 Rev.12 |  | Subsystem | SP1 | N/A | Fixed | shaham_t | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | SRS | Ovation |  | DHF298553 Rev.12 | April 13, 2020 2:35:30 PM GMT+08:00
CTUDT00629336 | Refresh | Closed | Revital_Cohen | March 31, 2020 10:04:25 PM GMT+08:00 | Refresh Issues in the front page of the iCT localization iFUs | WIP24 |  |  | Subsystem |  | NA | No Plan to Fix | charles_j | 0 | 4 |  | Support-System | NO | Closed | IFU | Ovation
CTUDT00629331 | Refresh | Closed | dadon_e | March 31, 2020 6:50:10 PM GMT+08:00 | Val: Refresh iCT: (WIP24): ICT Bay 138, Error on get raw data Validation test workflow scenario - Get Raw Data tool | WIP24 |  | Since the raw data will always be available in workflow, the user can always try copying back to other USB \format the USB to NTFS\copy to other storage devices\copy back to local for further processing. | Subsystem | Raw Data Tool | NA | No Plan to Fix | s_sa | 0 | 3 |  | iCAP | NO | Closed |  | Ovation
CTUDT00629327 | Refresh | Closed | urim | March 31, 2020 5:02:59 PM GMT+08:00 | Refresh Validation: Could not Start 1-Click Installation after Failed Installation with USB SSD Device Connected | Refresh WIP 24.1 |  | As per Service manual,user should disconnect any other external disk during host installation. | Subsystem | 1-Click; | STPS-SSRS-1154 | No Plan to Fix | venugopal_s | 0 | 3 |  | DI | NO | Closed |  | Ovation
CTUDT00629327 | Refresh | Closed | urim | March 31, 2020 5:02:59 PM GMT+08:00 | Refresh Validation: Could not Start 1-Click Installation after Failed Installation with USB SSD Device Connected | Refresh WIP 24.1 |  | As per Service manual,user should disconnect any other external disk during host installation. | Subsystem | 1-Click; | STPS-SSRS-1154 | No Plan to Fix | venugopal_s | 0 | 3 |  | DI | NO | Closed |  | IQon1
CTUDT00629216 | Refresh | Closed | mishlovsky_i | March 22, 2020 3:48:01 PM GMT+08:00 | An error appeared while trying to change the preset of the SBI to Neuro | WIP24 | SP1 WIP 30 | Restore previously backed up Presets (if available). OTHERWISE THE SYSTEM MIGHT NEED TO BE RE-INSTALLED! | Subsystem | SP1,CRN | N/A | Fixed | aviw | 0 | 3 | EBW Applications | Support_Spectral | NO | Closed | SP1 | IQon1 |  | SP1 WIP 30 | May 14, 2020 2:49:06 PM GMT+08:00
CTUDT00629206 | Refresh | Closed | glauberman_o | March 20, 2020 2:34:41 PM GMT+08:00 | Val: Refresh:ECG signal was disconnected | WIP24 | WIP31 |  | System | CRN; SP1_PM_Approved | NA | Fixed | modi_s | 0 | 3 | Console Software | Console4 | NO | Closed | WIP31SubCom | Ovation |  | WIP31 | July 27, 2020 2:02:53 PM GMT+08:00
CTUDT00629206 | Refresh | Closed | glauberman_o | March 20, 2020 2:34:41 PM GMT+08:00 | Val: Refresh:ECG signal was disconnected | WIP24 | WIP31 |  | System | CRN; SP1_PM_Approved | NA | Fixed | modi_s | 0 | 3 | Console Software | Console4 | NO | Closed | WIP31SubCom | Brilliance 64 |  | WIP31 | July 27, 2020 2:02:53 PM GMT+08:00
CTUDT00629206 | Refresh | Closed | glauberman_o | March 20, 2020 2:34:41 PM GMT+08:00 | Val: Refresh:ECG signal was disconnected | WIP24 | WIP31 |  | System | CRN; SP1_PM_Approved | NA | Fixed | modi_s | 0 | 3 | Console Software | Console4 | NO | Closed | WIP31SubCom | Brilliance Big Bore |  | WIP31 | July 27, 2020 2:02:53 PM GMT+08:00
CTUDT00629206 | Refresh | Closed | glauberman_o | March 20, 2020 2:34:41 PM GMT+08:00 | Val: Refresh:ECG signal was disconnected | WIP24 | WIP31 |  | System | CRN; SP1_PM_Approved | NA | Fixed | modi_s | 0 | 3 | Console Software | Console4 | NO | Closed | WIP31SubCom | Ingenuity 64 |  | WIP31 | July 27, 2020 2:02:53 PM GMT+08:00
CTUDT00629206 | Refresh | Closed | glauberman_o | March 20, 2020 2:34:41 PM GMT+08:00 | Val: Refresh:ECG signal was disconnected | WIP24 | WIP31 |  | System | CRN; SP1_PM_Approved | NA | Fixed | modi_s | 0 | 3 | Console Software | Console4 | NO | Closed | WIP31SubCom | IQon1 |  | WIP31 | July 27, 2020 2:02:53 PM GMT+08:00
CTUDT00629205 | Refresh | Closed | lin_s | March 20, 2020 11:10:34 AM GMT+08:00 | Val: Refresh: During Bugrep collection Windows Command Processor has stopped working | WIP22.1 | BB_WIP12.5 WIP22.5 | User can start the bug report collection again. | Subsystem |  | NA | Fixed | srushti_k | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP24.1SegReq | Ovation |  | BB_WIP12.5 WIP22.5 | March 20, 2020 11:27:23 AM GMT+08:00
CTUDT00629205 | Refresh | Closed | lin_s | March 20, 2020 11:10:34 AM GMT+08:00 | Val: Refresh: During Bugrep collection Windows Command Processor has stopped working | WIP22.1 | BB_WIP12.5 WIP22.5 | User can start the bug report collection again. | Subsystem |  | NA | Fixed | srushti_k | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP24.1SegReq | IQon1 |  | BB_WIP12.5 WIP22.5 | March 20, 2020 11:27:23 AM GMT+08:00
CTUDT00629200 | Refresh | Closed | herip_yapoujian | March 19, 2020 10:45:32 PM GMT+08:00 | Val: Refresh: ECG wave in Cardiac viewer is showing wrong annotation of adjusted phase bar. | WIP24 |  |  | Subsystem | Local Phase Editing |  | Duplicate | sanjeeva_a | 0 |  |  |  |  | Closed |  | Ovation
CTUDT00629181 | Refresh | Closed | lin_s | March 18, 2020 5:18:49 PM GMT+08:00 | Auto storage option settings are not retained in Exam cards when exam cards are exported using Exam card manger | WIP22.4 | 459801587361 Rev F 459801587351 Rev F 459801587381 Rev E 459801587391 Rev D | Service user needs to export the exam cards from 4.1.7 (Ingen, Br64, ovation) machine as detailed below: 1. Connect backup USB drive to HOST 2. Take the user exam card backup by copying exam cards to the USB drive from below path: - Copy Orion_64 folder from D:\Pms\Data\ExamCardsRepository for Ingenuity and Brilliance 64 Scanners - Copy Ovation_128 folder from D:\Pms\Data\ExamCardsRepository for iCT Scanner - Copy Ovation_64 folder from D:\Pms\Data\ExamCardsRepository for iCTSP Scanner In the above procedure, all the auto storage settings will be exported to the Exam cards. | Subsystem |  | N/A | Fixed | shi_b | 0 | 2 | Service Tools_Software | Support-Service | NO | Closed | Service Manual next version | Brilliance 64 |  | 459801587361 Rev F 459801587351 Rev F 459801587381 Rev E 459801587391 Rev D | June 30, 2020 3:18:31 PM GMT+08:00
CTUDT00629181 | Refresh | Closed | lin_s | March 18, 2020 5:18:49 PM GMT+08:00 | Auto storage option settings are not retained in Exam cards when exam cards are exported using Exam card manger | WIP22.4 | 459801587361 Rev F 459801587351 Rev F 459801587381 Rev E 459801587391 Rev D | Service user needs to export the exam cards from 4.1.7 (Ingen, Br64, ovation) machine as detailed below: 1. Connect backup USB drive to HOST 2. Take the user exam card backup by copying exam cards to the USB drive from below path: - Copy Orion_64 folder from D:\Pms\Data\ExamCardsRepository for Ingenuity and Brilliance 64 Scanners - Copy Ovation_128 folder from D:\Pms\Data\ExamCardsRepository for iCT Scanner - Copy Ovation_64 folder from D:\Pms\Data\ExamCardsRepository for iCTSP Scanner In the above procedure, all the auto storage settings will be exported to the Exam cards. | Subsystem |  | N/A | Fixed | shi_b | 0 | 2 | Service Tools_Software | Support-Service | NO | Closed | Service Manual next version | Ingenuity 64 |  | 459801587361 Rev F 459801587351 Rev F 459801587381 Rev E 459801587391 Rev D | June 30, 2020 3:18:31 PM GMT+08:00
CTUDT00629180 | Refresh | Closed | lin_s | March 18, 2020 3:31:29 PM GMT+08:00 | Refresh BB - Remote to CIRS failed from Remote Viewer because IP retrieved as 192.168.0.201 | BB_WIP12.4 | 459801587331 Rev G 459801587341 Rev G 459801587371 Rev G 459801587121 Rev F 459801587361 Rev F 459801587351 Rev F 459801587381 Rev E 459801587391 Rev D 459801587191 Rev E | If user cannot remote into CIRS through service tools and CIRS is reporting 201 instead of 200, these steps can be tried to address the issue. 1. Re-starting the CIRS 2. Re-start ServiceNode: a. Use mstsc to remote into CIRS b. Open Services shortcut on the desktop c. Restart the ServiceNodeHost service d. Wait approximately 1 minute e. Retry remote connection through remote viewer 3. ServiceNode Delayed Start Potential Fix a. Use mstsc to remote into CIRS b. Open Services shortcut on the desktop c. Double click ServiceNodeHost service d. Change Startup type to “Automatic (Delayed Start)” e. Click Apply, then OK f. Restart the CIRS For O-level users, they may connect directly to the CIRS using a keyboard, mouse, and monitor if necessary. | Subsystem | FOK | NA | Fixed | shi_b | 0 | 3 | CIRS_SW_Installation | Support-Service | NO | Closed | SubCom-24-Mar-2020 | Brilliance 64 |  | 459801587331 Rev G 459801587341 Rev G 459801587371 Rev G 459801587121 Rev F 459801587361 Rev F 459801587351 Rev F 459801587381 Rev E 459801587391 Rev D 459801587191 Rev E | March 24, 2020 10:49:26 AM GMT+08:00
CTUDT00629180 | Refresh | Closed | lin_s | March 18, 2020 3:31:29 PM GMT+08:00 | Refresh BB - Remote to CIRS failed from Remote Viewer because IP retrieved as 192.168.0.201 | BB_WIP12.4 | 459801587331 Rev G 459801587341 Rev G 459801587371 Rev G 459801587121 Rev F 459801587361 Rev F 459801587351 Rev F 459801587381 Rev E 459801587391 Rev D 459801587191 Rev E | If user cannot remote into CIRS through service tools and CIRS is reporting 201 instead of 200, these steps can be tried to address the issue. 1. Re-starting the CIRS 2. Re-start ServiceNode: a. Use mstsc to remote into CIRS b. Open Services shortcut on the desktop c. Restart the ServiceNodeHost service d. Wait approximately 1 minute e. Retry remote connection through remote viewer 3. ServiceNode Delayed Start Potential Fix a. Use mstsc to remote into CIRS b. Open Services shortcut on the desktop c. Double click ServiceNodeHost service d. Change Startup type to “Automatic (Delayed Start)” e. Click Apply, then OK f. Restart the CIRS For O-level users, they may connect directly to the CIRS using a keyboard, mouse, and monitor if necessary. | Subsystem | FOK | NA | Fixed | shi_b | 0 | 3 | CIRS_SW_Installation | Support-Service | NO | Closed | SubCom-24-Mar-2020 | Brilliance Big Bore |  | 459801587331 Rev G 459801587341 Rev G 459801587371 Rev G 459801587121 Rev F 459801587361 Rev F 459801587351 Rev F 459801587381 Rev E 459801587391 Rev D 459801587191 Rev E | March 24, 2020 10:49:26 AM GMT+08:00
CTUDT00629180 | Refresh | Closed | lin_s | March 18, 2020 3:31:29 PM GMT+08:00 | Refresh BB - Remote to CIRS failed from Remote Viewer because IP retrieved as 192.168.0.201 | BB_WIP12.4 | 459801587331 Rev G 459801587341 Rev G 459801587371 Rev G 459801587121 Rev F 459801587361 Rev F 459801587351 Rev F 459801587381 Rev E 459801587391 Rev D 459801587191 Rev E | If user cannot remote into CIRS through service tools and CIRS is reporting 201 instead of 200, these steps can be tried to address the issue. 1. Re-starting the CIRS 2. Re-start ServiceNode: a. Use mstsc to remote into CIRS b. Open Services shortcut on the desktop c. Restart the ServiceNodeHost service d. Wait approximately 1 minute e. Retry remote connection through remote viewer 3. ServiceNode Delayed Start Potential Fix a. Use mstsc to remote into CIRS b. Open Services shortcut on the desktop c. Double click ServiceNodeHost service d. Change Startup type to “Automatic (Delayed Start)” e. Click Apply, then OK f. Restart the CIRS For O-level users, they may connect directly to the CIRS using a keyboard, mouse, and monitor if necessary. | Subsystem | FOK | NA | Fixed | shi_b | 0 | 3 | CIRS_SW_Installation | Support-Service | NO | Closed | SubCom-24-Mar-2020 | Ingenuity 64 |  | 459801587331 Rev G 459801587341 Rev G 459801587371 Rev G 459801587121 Rev F 459801587361 Rev F 459801587351 Rev F 459801587381 Rev E 459801587391 Rev D 459801587191 Rev E | March 24, 2020 10:49:26 AM GMT+08:00
CTUDT00629179 | Refresh | Closed | lin_s | March 18, 2020 3:22:43 PM GMT+08:00 | System not prompt for a password change for IT user first time login after reset password via Service tool. | BB_WIP12.4 WIP22.4 | 459801587121 Rev F 459801587191 Rev E | There are two ways to reset the password for the Factory Hospital Administrator user. 1. Philips_service user can use IT-Admin tool to change the password. Philips_Service user can launch IT-Admin tool from C:\Pms\Syste\Philips.PmsCT.AdminTool.Hosting.exe 2. If user uses the STPS tool to reset the password, then after resetting the password, user needs perform the below setps. 1. Launch windows user management tool using "Lusrmgr.msc". 2. Go to IT user and open it properties. 3. Set the flag "User must change password at next logon". 4. Uncheck the flag “Account is Locked out” check box. (This is required when the IT user is locked out due to wrong password entries). Ignore if already unchecked. | Subsystem |  | N/A | Fixed | shi_b | 0 | 2 | Service Tools_Software | Support-Service | NO | Closed | Service Manual New Version | Brilliance 64 |  | 459801587121 Rev F 459801587191 Rev E | March 24, 2020 10:49:47 AM GMT+08:00
CTUDT00629179 | Refresh | Closed | lin_s | March 18, 2020 3:22:43 PM GMT+08:00 | System not prompt for a password change for IT user first time login after reset password via Service tool. | BB_WIP12.4 WIP22.4 | 459801587121 Rev F 459801587191 Rev E | There are two ways to reset the password for the Factory Hospital Administrator user. 1. Philips_service user can use IT-Admin tool to change the password. Philips_Service user can launch IT-Admin tool from C:\Pms\Syste\Philips.PmsCT.AdminTool.Hosting.exe 2. If user uses the STPS tool to reset the password, then after resetting the password, user needs perform the below setps. 1. Launch windows user management tool using "Lusrmgr.msc". 2. Go to IT user and open it properties. 3. Set the flag "User must change password at next logon". 4. Uncheck the flag “Account is Locked out” check box. (This is required when the IT user is locked out due to wrong password entries). Ignore if already unchecked. | Subsystem |  | N/A | Fixed | shi_b | 0 | 2 | Service Tools_Software | Support-Service | NO | Closed | Service Manual New Version | Brilliance Big Bore |  | 459801587121 Rev F 459801587191 Rev E | March 24, 2020 10:49:47 AM GMT+08:00
CTUDT00629179 | Refresh | Closed | lin_s | March 18, 2020 3:22:43 PM GMT+08:00 | System not prompt for a password change for IT user first time login after reset password via Service tool. | BB_WIP12.4 WIP22.4 | 459801587121 Rev F 459801587191 Rev E | There are two ways to reset the password for the Factory Hospital Administrator user. 1. Philips_service user can use IT-Admin tool to change the password. Philips_Service user can launch IT-Admin tool from C:\Pms\Syste\Philips.PmsCT.AdminTool.Hosting.exe 2. If user uses the STPS tool to reset the password, then after resetting the password, user needs perform the below setps. 1. Launch windows user management tool using "Lusrmgr.msc". 2. Go to IT user and open it properties. 3. Set the flag "User must change password at next logon". 4. Uncheck the flag “Account is Locked out” check box. (This is required when the IT user is locked out due to wrong password entries). Ignore if already unchecked. | Subsystem |  | N/A | Fixed | shi_b | 0 | 2 | Service Tools_Software | Support-Service | NO | Closed | Service Manual New Version | Ingenuity 64 |  | 459801587121 Rev F 459801587191 Rev E | March 24, 2020 10:49:47 AM GMT+08:00
CTUDT00629174 | Refresh | Closed | lin_s | March 18, 2020 10:26:29 AM GMT+08:00 | [HTML IFU]It displayed MK not Macedonian in the language selection list | WIP22.4, WIP12.4 | ING V4.1.10 RfD BB V4.8 RfD | N/A | Subsystem |  | N/A | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | 05-Mar-2020 | Brilliance Big Bore |  | ING V4.1.10 RfD BB V4.8 RfD | March 18, 2020 11:59:48 AM GMT+08:00
CTUDT00629174 | Refresh | Closed | lin_s | March 18, 2020 10:26:29 AM GMT+08:00 | [HTML IFU]It displayed MK not Macedonian in the language selection list | WIP22.4, WIP12.4 | ING V4.1.10 RfD BB V4.8 RfD | N/A | Subsystem |  | N/A | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | 05-Mar-2020 | Ingenuity 64 |  | ING V4.1.10 RfD BB V4.8 RfD | March 18, 2020 11:59:48 AM GMT+08:00
CTUDT00629155 | Refresh | Closed | bandil_j | March 16, 2020 2:20:09 PM GMT+08:00 | Refresh; IQon: User Reliability issues: WIP24:WF crash after Bolus EC scan. | WIP24 | WIP24.1 |  | Subsystem | Blocker ; PM_Approved | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | YES | Closed | WIP24.1SegReq | Ovation |  | WIP24.1 | March 17, 2020 8:02:19 PM GMT+08:00
CTUDT00629155 | Refresh | Closed | bandil_j | March 16, 2020 2:20:09 PM GMT+08:00 | Refresh; IQon: User Reliability issues: WIP24:WF crash after Bolus EC scan. | WIP24 | WIP24.1 |  | Subsystem | Blocker ; PM_Approved | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | YES | Closed | WIP24.1SegReq | IQon1 |  | WIP24.1 | March 17, 2020 8:02:19 PM GMT+08:00
CTUDT00629146 | Refresh | Closed | l_yiming | March 13, 2020 3:34:50 PM GMT+08:00 | System Crush During Performing Report Approval | WIP22.5 | WIP30 WIP30.1 |  | System | SP1 | NA | Fixed | s_sa | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP30.1SubCom | Brilliance 64 |  | WIP30 WIP30.1 | June 1, 2020 6:51:36 PM GMT+08:00
CTUDT00629146 | Refresh | Closed | l_yiming | March 13, 2020 3:34:50 PM GMT+08:00 | System Crush During Performing Report Approval | WIP22.5 | WIP30 WIP30.1 |  | System | SP1 | NA | Fixed | s_sa | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP30.1SubCom | Brilliance Big Bore |  | WIP30 WIP30.1 | June 1, 2020 6:51:36 PM GMT+08:00
CTUDT00629146 | Refresh | Closed | l_yiming | March 13, 2020 3:34:50 PM GMT+08:00 | System Crush During Performing Report Approval | WIP22.5 | WIP30 WIP30.1 |  | System | SP1 | NA | Fixed | s_sa | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP30.1SubCom | Ingenuity 64 |  | WIP30 WIP30.1 | June 1, 2020 6:51:36 PM GMT+08:00
CTUDT00629082 | Refresh | Closed | s_namburi | March 9, 2020 5:18:16 PM GMT+08:00 | Auto storage option settings are not retained in Exam cards when exam cards are exported using Exam card manger | WIP22.4 | DMR356920 Rev10 | Service user needs to export the exam cards from 4.1.7 (Ingen, Br64, ovation) machine as detailed below: 1. Connect backup USB drive to HOST 2. Take the user exam card backup by copying exam cards to the USB drive from below path: - Copy Orion_64 folder from D:\Pms\Data\ExamCardsRepository for Ingenuity and Brilliance 64 Scanners - Copy Ovation_128 folder from D:\Pms\Data\ExamCardsRepository for iCT Scanner - Copy Ovation_64 folder from D:\Pms\Data\ExamCardsRepository for iCTSP Scanner In the above procedure, all the auto storage settings will be exported to the Exam cards. | Subsystem | SP1 | NA | Fixed | shaham_t | 0 | 2 | Service Documentation | Support-System | NO | Closed | WIP24.1SubCom_26-Mar-2020 | Ovation |  | DMR356920 Rev10 | March 30, 2020 2:22:08 PM GMT+08:00
CTUDT00629065 | Refresh | Closed | cherylli | March 7, 2020 1:09:58 PM GMT+08:00 | BR64 User document: Issues found in 459801608571_D ECB | WIP22.5 | 459801608571_E | N/A | Subsystem |  | N/A | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | ECB new version | Brilliance 64 |  | 459801608571_E | March 13, 2020 12:58:15 PM GMT+08:00
CTUDT00629053 | Refresh | Closed | bingxin_lin | March 6, 2020 3:30:16 PM GMT+08:00 | Move to inner laser button on gantry control panel did not work after using rocker switch on IVC box | BB_WIP12.5 |  | Solution 1: When "move to inner laser" button on gantry control panel did not work, we can press “move to outer laser” button on gantry control panel or “move to inner/outer laser” ,"Go to saved work position" ,"Go to scan position" button on IVC box to active the "move to inner laser" button on gantry control panel Solution 2: when you pressed "move to inner /outer laser" button on gantry control panel , do not use "rocker switch" , "Vertical up/down" ,"tilt backward/forward" on IVC box , this issue will not occur | Subsystem |  | N/A | No Plan to Fix | Mark_Huang | 0 | 4 |  | GCX | NO | Closed |  | Brilliance Big Bore
CTUDT00629044 | Refresh | Closed | 320053769 | March 5, 2020 8:51:23 PM GMT+08:00 | When loading a result or a study to quick review, the written in the left corner is blurred or grayed out. | WIP24 |  | NA | Subsystem |  | NA | Works As Specified / Not A Defect | aswani_v | 0 |  |  | iCAP | NO | Closed |  | Ovation
CTUDT00629044 | Refresh | Closed | 320053769 | March 5, 2020 8:51:23 PM GMT+08:00 | When loading a result or a study to quick review, the written in the left corner is blurred or grayed out. | WIP24 |  | NA | Subsystem |  | NA | Works As Specified / Not A Defect | aswani_v | 0 |  |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00629044 | Refresh | Closed | 320053769 | March 5, 2020 8:51:23 PM GMT+08:00 | When loading a result or a study to quick review, the written in the left corner is blurred or grayed out. | WIP24 |  | NA | Subsystem |  | NA | Works As Specified / Not A Defect | aswani_v | 0 |  |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00629044 | Refresh | Closed | 320053769 | March 5, 2020 8:51:23 PM GMT+08:00 | When loading a result or a study to quick review, the written in the left corner is blurred or grayed out. | WIP24 |  | NA | Subsystem |  | NA | Works As Specified / Not A Defect | aswani_v | 0 |  |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00629044 | Refresh | Closed | 320053769 | March 5, 2020 8:51:23 PM GMT+08:00 | When loading a result or a study to quick review, the written in the left corner is blurred or grayed out. | WIP24 |  | NA | Subsystem |  | NA | Works As Specified / Not A Defect | aswani_v | 0 |  |  | iCAP | NO | Closed |  | IQon1
CTUDT00629041 | Refresh | Closed | shterner_a | March 5, 2020 7:19:48 PM GMT+08:00 | Val: Refresh iCTSP Exam Card Book 459801656451 Rev D requires updates | WIP24 | iCT v4.1.10 RfD | N/A | Subsystem |  | N/A | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | 10-MAR-2020 | Ovation |  | iCT v4.1.10 RfD | March 11, 2020 12:16:25 PM GMT+08:00
CTUDT00629012 | Refresh | Closed | glauberman_o | March 4, 2020 7:20:58 PM GMT+08:00 | Val: Refresh:Copy to SSD Failed | WIP23.1 |  | Clicking on "Retry" transfers are successful. | Subsystem | CRN; | NA | No Plan to Fix | s_panigrahi | 0 | 3 |  | iCAP | NO | Closed |  | Ovation
CTUDT00629012 | Refresh | Closed | glauberman_o | March 4, 2020 7:20:58 PM GMT+08:00 | Val: Refresh:Copy to SSD Failed | WIP23.1 |  | Clicking on "Retry" transfers are successful. | Subsystem | CRN; | NA | No Plan to Fix | s_panigrahi | 0 | 3 |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00629012 | Refresh | Closed | glauberman_o | March 4, 2020 7:20:58 PM GMT+08:00 | Val: Refresh:Copy to SSD Failed | WIP23.1 |  | Clicking on "Retry" transfers are successful. | Subsystem | CRN; | NA | No Plan to Fix | s_panigrahi | 0 | 3 |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00629012 | Refresh | Closed | glauberman_o | March 4, 2020 7:20:58 PM GMT+08:00 | Val: Refresh:Copy to SSD Failed | WIP23.1 |  | Clicking on "Retry" transfers are successful. | Subsystem | CRN; | NA | No Plan to Fix | s_panigrahi | 0 | 3 |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00629012 | Refresh | Closed | glauberman_o | March 4, 2020 7:20:58 PM GMT+08:00 | Val: Refresh:Copy to SSD Failed | WIP23.1 |  | Clicking on "Retry" transfers are successful. | Subsystem | CRN; | NA | No Plan to Fix | s_panigrahi | 0 | 3 |  | iCAP | NO | Closed |  | IQon1
CTUDT00628987 | Refresh | Closed | 320053770 | March 3, 2020 10:20:39 PM GMT+08:00 | Refresh:IQON:WIP24:Acquisition done outside of plan box (plan on surview) | WIP24 |  | User can click on the plan box before pressing go button and plan box will be corrected | Subsystem | Blocker | N/A | No Plan to Fix | navale_s | 0 | 4 |  | Console4 | YES | Closed |  | Ovation
CTUDT00628987 | Refresh | Closed | 320053770 | March 3, 2020 10:20:39 PM GMT+08:00 | Refresh:IQON:WIP24:Acquisition done outside of plan box (plan on surview) | WIP24 |  | User can click on the plan box before pressing go button and plan box will be corrected | Subsystem | Blocker | N/A | No Plan to Fix | navale_s | 0 | 4 |  | Console4 | YES | Closed |  | Brilliance 64
CTUDT00628987 | Refresh | Closed | 320053770 | March 3, 2020 10:20:39 PM GMT+08:00 | Refresh:IQON:WIP24:Acquisition done outside of plan box (plan on surview) | WIP24 |  | User can click on the plan box before pressing go button and plan box will be corrected | Subsystem | Blocker | N/A | No Plan to Fix | navale_s | 0 | 4 |  | Console4 | YES | Closed |  | Brilliance Big Bore
CTUDT00628987 | Refresh | Closed | 320053770 | March 3, 2020 10:20:39 PM GMT+08:00 | Refresh:IQON:WIP24:Acquisition done outside of plan box (plan on surview) | WIP24 |  | User can click on the plan box before pressing go button and plan box will be corrected | Subsystem | Blocker | N/A | No Plan to Fix | navale_s | 0 | 4 |  | Console4 | YES | Closed |  | Ingenuity 64
CTUDT00628987 | Refresh | Closed | 320053770 | March 3, 2020 10:20:39 PM GMT+08:00 | Refresh:IQON:WIP24:Acquisition done outside of plan box (plan on surview) | WIP24 |  | User can click on the plan box before pressing go button and plan box will be corrected | Subsystem | Blocker | N/A | No Plan to Fix | navale_s | 0 | 4 |  | Console4 | YES | Closed |  | IQon1
CTUDT00628957 | Refresh | Closed | Jose_D_Sergio_Almeida | February 28, 2020 7:12:16 PM GMT+08:00 | Val:Refresh: Spectral CT Viewer crashed while loading Chest/Abdomen/Pelvis study | WIP23.1 | WIP24.1 |  | Subsystem | Blocker; PM_Approved | NA | Fixed | aviw | 0 | 2 | SDViewer.exe | Support_Spectral | NO | Closed | WIP24.1SubCom | IQon1 |  | WIP24.1 | March 11, 2020 3:13:43 PM GMT+08:00
CTUDT00628943 | Refresh | Closed | shi_z | February 28, 2020 1:05:17 PM GMT+08:00 | Infant Performance test failed with H-Phantom. | WIP22.2 |  |  | System |  |  | Works As Specified / Not A Defect | shi_z | 0 |  |  |  |  | Closed |  | Brilliance 64
CTUDT00628919 | Refresh | Closed | panchami_pradeep | February 27, 2020 2:54:28 PM GMT+08:00 | Refresh: “Failed to get session from service manager” error was displayed while launching service tools after changing date. | WIP24 | DMR356920 Rev10 | restart the system when the date time is updated in case if this issue is observed, | Subsystem | Users | NA | Fixed | shaham_t | 0 | 4 | Service Documentation | Support-System | NO | Closed | WIP24.1SubCom_26-Mar-2020 | IQon1 |  | DMR356920 Rev10 | March 29, 2020 6:07:25 PM GMT+08:00
CTUDT00628911 | Refresh | Closed | zhang_l | February 27, 2020 10:54:09 AM GMT+08:00 | Lack of Reference Exam Card "Arterial Phase" and "Portal/Venous Phase" based on Reference Exam Card Specification | WIP22.5 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | lin_s | 0 |  |  |  |  | Closed |  | Brilliance 64
CTUDT00628911 | Refresh | Closed | zhang_l | February 27, 2020 10:54:09 AM GMT+08:00 | Lack of Reference Exam Card "Arterial Phase" and "Portal/Venous Phase" based on Reference Exam Card Specification | WIP22.5 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | lin_s | 0 |  |  |  |  | Closed |  | Ingenuity 64
CTUDT00628907 | Refresh | Closed | Ashish_Sharma | February 27, 2020 3:03:03 AM GMT+08:00 | While performing calcium scoring along with coronary CT , cardiac wave disappeared from UI during workflow but showing heart rate | WIP23.1 |  |  | Subsystem |  | NA | Duplicate | benny_w | 0 |  |  | GCX | NO | Closed |  | Ovation
CTUDT00628907 | Refresh | Closed | Ashish_Sharma | February 27, 2020 3:03:03 AM GMT+08:00 | While performing calcium scoring along with coronary CT , cardiac wave disappeared from UI during workflow but showing heart rate | WIP23.1 |  |  | Subsystem |  | NA | Duplicate | benny_w | 0 |  |  | GCX | NO | Closed |  | IQon1
CTUDT00628905 | Refresh | Closed | Ashish_Sharma | February 26, 2020 11:48:38 PM GMT+08:00 | While performing calcium scoring along with coronary CT , Error message popped up and system automatically logged out with blue screen. | WIP23.1 |  | to restart the host prior gantry restart | Subsystem |  | N/A | No Plan to Fix | yang_h | 0 | 3 |  | GCX | YES | Closed |  | Ovation
CTUDT00628905 | Refresh | Closed | Ashish_Sharma | February 26, 2020 11:48:38 PM GMT+08:00 | While performing calcium scoring along with coronary CT , Error message popped up and system automatically logged out with blue screen. | WIP23.1 |  | to restart the host prior gantry restart | Subsystem |  | N/A | No Plan to Fix | yang_h | 0 | 3 |  | GCX | YES | Closed |  | IQon1
CTUDT00628891 | Refresh | Closed | dongping_w | February 26, 2020 11:18:36 AM GMT+08:00 | Val: Refresh: During Bugrep collection Windows Command Processor has stopped working | WIP22.5 |  | User can start the bug report collection again. | Subsystem | FoK; | N/A | Does Not Apply | tiwari_g | 0 | 4 |  | Service Tools | NO | Closed |  | Ovation | PR 192814
CTUDT00628891 | Refresh | Closed | dongping_w | February 26, 2020 11:18:36 AM GMT+08:00 | Val: Refresh: During Bugrep collection Windows Command Processor has stopped working | WIP22.5 |  | User can start the bug report collection again. | Subsystem | FoK; | N/A | Does Not Apply | tiwari_g | 0 | 4 |  | Service Tools | NO | Closed |  | Brilliance 64 | PR 192814
CTUDT00628891 | Refresh | Closed | dongping_w | February 26, 2020 11:18:36 AM GMT+08:00 | Val: Refresh: During Bugrep collection Windows Command Processor has stopped working | WIP22.5 |  | User can start the bug report collection again. | Subsystem | FoK; | N/A | Does Not Apply | tiwari_g | 0 | 4 |  | Service Tools | NO | Closed |  | Brilliance Big Bore | PR 192814
CTUDT00628891 | Refresh | Closed | dongping_w | February 26, 2020 11:18:36 AM GMT+08:00 | Val: Refresh: During Bugrep collection Windows Command Processor has stopped working | WIP22.5 |  | User can start the bug report collection again. | Subsystem | FoK; | N/A | Does Not Apply | tiwari_g | 0 | 4 |  | Service Tools | NO | Closed |  | Ingenuity 64 | PR 192814
CTUDT00628891 | Refresh | Closed | dongping_w | February 26, 2020 11:18:36 AM GMT+08:00 | Val: Refresh: During Bugrep collection Windows Command Processor has stopped working | WIP22.5 |  | User can start the bug report collection again. | Subsystem | FoK; | N/A | Does Not Apply | tiwari_g | 0 | 4 |  | Service Tools | NO | Closed |  | IQon1 | PR 192814
CTUDT00628880 | Refresh | Closed | rojin_antony | February 25, 2020 6:59:33 PM GMT+08:00 | Val: Refresh:Manually trigger the clinical scan : ‘Start Scan’ message was not displayed | WIP23.1 |  |  | Subsystem | UI | N/A | No Plan to Fix | datta_i | 0 | 3 |  | Console4 | YES | Closed |  | IQon1
CTUDT00628877 | Refresh | Closed | l_yiming | February 25, 2020 9:33:52 AM GMT+08:00 | Refresh: View 2 continue to perform offline recon during scan | WIP22.4 |  | Not Required. | Subsystem |  | NA | Works As Specified / Not A Defect | kapadia_v | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00628874 | Refresh | Closed | vinoliya_j | February 24, 2020 11:35:56 PM GMT+08:00 | 459800975185_B Iqon 4.7.7 Customer release notes validation comments | WIP23.1 | IQon v4.7.7 | N/A | Subsystem | Documentation | N/A | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | 05-Mar-2020 | IQon1 |  | IQon v4.7.7 | March 6, 2020 2:12:37 PM GMT+08:00
CTUDT00628872 | Refresh | Closed | vinoliya_j | February 24, 2020 11:25:05 PM GMT+08:00 | Val: Refresh: Manual typing of threshold value in the bolus tracking graph has some refresh issues | WIP23.1 |  |  | Subsystem | Bolus Tracking | N/A | No Plan to Fix | navale_s | 0 | 4 |  | Console4 | NO | Closed |  | IQon1
CTUDT00628870 | Refresh | Closed | vinoliya_j | February 24, 2020 11:07:40 PM GMT+08:00 | Val: Refresh : Voice manager crashed while recording a phrase for the user created language | WIP23.1 |  | Re-launch the Voice manager | Subsystem | Voice Recording | N/A | No Plan to Fix | navale_s | 0 | 4 |  | Console4 | NO | Closed |  | IQon1 | NA
CTUDT00628867 | Refresh | Closed | Ashish_Sharma | February 24, 2020 10:40:28 PM GMT+08:00 | Short Tube conditioning failed | WIP23.1 |  | Verify that all parameters are in accordance with Service manual before starting STC | Subsystem |  | NA | No Plan to Fix | benny_w | 0 | 3 |  | GCX | NO | Closed |  | Ovation
CTUDT00628863 | Refresh | Closed | Ashish_Sharma | February 24, 2020 10:35:01 PM GMT+08:00 | While planning to perform short tube conditioning, System crash | WIP23.1 |  |  | Subsystem | Reliability; Blocker | NA | Duplicate | s_sa | 0 |  |  | iCAP | NO | Closed |  | Ovation
CTUDT00628863 | Refresh | Closed | Ashish_Sharma | February 24, 2020 10:35:01 PM GMT+08:00 | While planning to perform short tube conditioning, System crash | WIP23.1 |  |  | Subsystem | Reliability; Blocker | NA | Duplicate | s_sa | 0 |  |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00628863 | Refresh | Closed | Ashish_Sharma | February 24, 2020 10:35:01 PM GMT+08:00 | While planning to perform short tube conditioning, System crash | WIP23.1 |  |  | Subsystem | Reliability; Blocker | NA | Duplicate | s_sa | 0 |  |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00628863 | Refresh | Closed | Ashish_Sharma | February 24, 2020 10:35:01 PM GMT+08:00 | While planning to perform short tube conditioning, System crash | WIP23.1 |  |  | Subsystem | Reliability; Blocker | NA | Duplicate | s_sa | 0 |  |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00628863 | Refresh | Closed | Ashish_Sharma | February 24, 2020 10:35:01 PM GMT+08:00 | While planning to perform short tube conditioning, System crash | WIP23.1 |  |  | Subsystem | Reliability; Blocker | NA | Duplicate | s_sa | 0 |  |  | iCAP | NO | Closed |  | IQon1
CTUDT00628862 | Refresh | Closed | herip_yapoujian | February 24, 2020 10:24:10 PM GMT+08:00 | Val: Refresh: Unexpected error during Workflow creation | WIP23.1 |  |  | Subsystem |  | NA | Duplicate | s_panigrahi | 0 |  |  | iCAP | NO | Closed |  | Ovation
CTUDT00628862 | Refresh | Closed | herip_yapoujian | February 24, 2020 10:24:10 PM GMT+08:00 | Val: Refresh: Unexpected error during Workflow creation | WIP23.1 |  |  | Subsystem |  | NA | Duplicate | s_panigrahi | 0 |  |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00628862 | Refresh | Closed | herip_yapoujian | February 24, 2020 10:24:10 PM GMT+08:00 | Val: Refresh: Unexpected error during Workflow creation | WIP23.1 |  |  | Subsystem |  | NA | Duplicate | s_panigrahi | 0 |  |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00628862 | Refresh | Closed | herip_yapoujian | February 24, 2020 10:24:10 PM GMT+08:00 | Val: Refresh: Unexpected error during Workflow creation | WIP23.1 |  |  | Subsystem |  | NA | Duplicate | s_panigrahi | 0 |  |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00628862 | Refresh | Closed | herip_yapoujian | February 24, 2020 10:24:10 PM GMT+08:00 | Val: Refresh: Unexpected error during Workflow creation | WIP23.1 |  |  | Subsystem |  | NA | Duplicate | s_panigrahi | 0 |  |  | iCAP | NO | Closed |  | IQon1
CTUDT00628859 | Refresh | Closed | shiraguppi_k | February 24, 2020 9:27:33 PM GMT+08:00 | Val: Refresh: System crash with Bugrep error while Performing Carotid CTA | WIP23.1 |  | System Restart is needed | Subsystem |  | NA | Duplicate | s_sa | 0 |  |  | iCAP | YES | Closed |  | Ovation
CTUDT00628859 | Refresh | Closed | shiraguppi_k | February 24, 2020 9:27:33 PM GMT+08:00 | Val: Refresh: System crash with Bugrep error while Performing Carotid CTA | WIP23.1 |  | System Restart is needed | Subsystem |  | NA | Duplicate | s_sa | 0 |  |  | iCAP | YES | Closed |  | Brilliance 64
CTUDT00628859 | Refresh | Closed | shiraguppi_k | February 24, 2020 9:27:33 PM GMT+08:00 | Val: Refresh: System crash with Bugrep error while Performing Carotid CTA | WIP23.1 |  | System Restart is needed | Subsystem |  | NA | Duplicate | s_sa | 0 |  |  | iCAP | YES | Closed |  | Brilliance Big Bore
CTUDT00628859 | Refresh | Closed | shiraguppi_k | February 24, 2020 9:27:33 PM GMT+08:00 | Val: Refresh: System crash with Bugrep error while Performing Carotid CTA | WIP23.1 |  | System Restart is needed | Subsystem |  | NA | Duplicate | s_sa | 0 |  |  | iCAP | YES | Closed |  | Ingenuity 64
CTUDT00628859 | Refresh | Closed | shiraguppi_k | February 24, 2020 9:27:33 PM GMT+08:00 | Val: Refresh: System crash with Bugrep error while Performing Carotid CTA | WIP23.1 |  | System Restart is needed | Subsystem |  | NA | Duplicate | s_sa | 0 |  |  | iCAP | YES | Closed |  | IQon1
CTUDT00628856 | Refresh | Closed | glauberman_o | February 24, 2020 8:50:37 PM GMT+08:00 | Val:Refresh:Error Patient Directory | WIP23.1 |  | Repeat copying | Subsystem | USP - PD | NA | No Plan to Fix | s_panigrahi | 0 | 3 |  | iCAP | NO | Closed |  | Ovation
CTUDT00628856 | Refresh | Closed | glauberman_o | February 24, 2020 8:50:37 PM GMT+08:00 | Val:Refresh:Error Patient Directory | WIP23.1 |  | Repeat copying | Subsystem | USP - PD | NA | No Plan to Fix | s_panigrahi | 0 | 3 |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00628856 | Refresh | Closed | glauberman_o | February 24, 2020 8:50:37 PM GMT+08:00 | Val:Refresh:Error Patient Directory | WIP23.1 |  | Repeat copying | Subsystem | USP - PD | NA | No Plan to Fix | s_panigrahi | 0 | 3 |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00628856 | Refresh | Closed | glauberman_o | February 24, 2020 8:50:37 PM GMT+08:00 | Val:Refresh:Error Patient Directory | WIP23.1 |  | Repeat copying | Subsystem | USP - PD | NA | No Plan to Fix | s_panigrahi | 0 | 3 |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00628856 | Refresh | Closed | glauberman_o | February 24, 2020 8:50:37 PM GMT+08:00 | Val:Refresh:Error Patient Directory | WIP23.1 |  | Repeat copying | Subsystem | USP - PD | NA | No Plan to Fix | s_panigrahi | 0 | 3 |  | iCAP | NO | Closed |  | IQon1
CTUDT00628827 | Refresh | Assigned | wang_l | February 21, 2020 7:15:32 PM GMT+08:00 | Refresh - The bugreport does not contain ingenuity user EC in \Console4\User ExamCards | WIP22.4 |  | The exam cards can be collected by exporting using ExamCard Manager by Senior Clinical User (With ExamCard Manager access)/FSE. | Subsystem | CRN; SP2_Candidate | NA |  | janakiraman_j | 0 | 2 |  | Service Tools | NO | Assigned | SP2 | Ovation
CTUDT00628827 | Refresh | Assigned | wang_l | February 21, 2020 7:15:32 PM GMT+08:00 | Refresh - The bugreport does not contain ingenuity user EC in \Console4\User ExamCards | WIP22.4 |  | The exam cards can be collected by exporting using ExamCard Manager by Senior Clinical User (With ExamCard Manager access)/FSE. | Subsystem | CRN; SP2_Candidate | NA |  | janakiraman_j | 0 | 2 |  | Service Tools | NO | Assigned | SP2 | Brilliance 64
CTUDT00628827 | Refresh | Assigned | wang_l | February 21, 2020 7:15:32 PM GMT+08:00 | Refresh - The bugreport does not contain ingenuity user EC in \Console4\User ExamCards | WIP22.4 |  | The exam cards can be collected by exporting using ExamCard Manager by Senior Clinical User (With ExamCard Manager access)/FSE. | Subsystem | CRN; SP2_Candidate | NA |  | janakiraman_j | 0 | 2 |  | Service Tools | NO | Assigned | SP2 | Ingenuity 64
CTUDT00628817 | Refresh | Closed | urim | February 20, 2020 10:57:02 PM GMT+08:00 | Refresh Val: Exam Card Book Missing Acronyms and Terms Definition | EC Book Rev D | iCT v4.1.10 RfD BR64 v4.1.10 RfD IQon v4.7.7 RfD | N/A | Subsystem |  | N/A | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | 05-Mar-2020 | Ovation |  | iCT v4.1.10 RfD BR64 v4.1.10 RfD IQon v4.7.7 RfD | March 6, 2020 2:31:20 PM GMT+08:00
CTUDT00628817 | Refresh | Closed | urim | February 20, 2020 10:57:02 PM GMT+08:00 | Refresh Val: Exam Card Book Missing Acronyms and Terms Definition | EC Book Rev D | iCT v4.1.10 RfD BR64 v4.1.10 RfD IQon v4.7.7 RfD | N/A | Subsystem |  | N/A | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | 05-Mar-2020 | Brilliance 64 |  | iCT v4.1.10 RfD BR64 v4.1.10 RfD IQon v4.7.7 RfD | March 6, 2020 2:31:20 PM GMT+08:00
CTUDT00628817 | Refresh | Closed | urim | February 20, 2020 10:57:02 PM GMT+08:00 | Refresh Val: Exam Card Book Missing Acronyms and Terms Definition | EC Book Rev D | iCT v4.1.10 RfD BR64 v4.1.10 RfD IQon v4.7.7 RfD | N/A | Subsystem |  | N/A | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | 05-Mar-2020 | IQon1 |  | iCT v4.1.10 RfD BR64 v4.1.10 RfD IQon v4.7.7 RfD | March 6, 2020 2:31:20 PM GMT+08:00
CTUDT00628814 | Refresh | Closed | Ashish_Sharma | February 20, 2020 9:12:47 PM GMT+08:00 | While planning to perform upper extremity study, gantry error message appeared. | WIP23.1 |  |  | Subsystem | Reliability | N/A | Works As Specified / Not A Defect | ivanf | 0 |  |  | DMS | NO | Closed | SubNC | Ovation
CTUDT00628811 | Refresh | Closed | dannyb | February 20, 2020 4:55:06 PM GMT+08:00 | Refresh/sCTV – Conventional viewport is displayed in RGB when exiting compare mode (two SBIs) | WIP24 | WIP24.1 |  | Subsystem | PM_Approved | NA | Fixed | aviw | 0 | 3 | Console USP_EBW | Support_Spectral | NO | Closed | WIP24.1SegReq | IQon1 |  | WIP24.1 | February 26, 2020 6:35:49 PM GMT+08:00
CTUDT00628808 | Refresh | Closed | dannyb | February 20, 2020 4:52:09 PM GMT+08:00 | Refresh/sCTV – Crash when selecting Tile View in Batch and datatype is spectral. | WIP24 | WIP24.1 |  | Subsystem | PM_Approved | NA | Fixed | aviw | 0 | 2 | Console USP_EBW | Support_Spectral | NO | Closed | WIP24.1SegReq | IQon1 |  | WIP24.1 | February 26, 2020 6:55:25 PM GMT+08:00
CTUDT00628806 | Refresh | Closed | dannyb | February 20, 2020 4:45:40 PM GMT+08:00 | Refresh/sCTV – Z Effective image appears BW when sent to Film. | WIP24 | WIP24.1 | Film display can be used. | Subsystem | Blocker; PM_Approved | NA | Fixed | aswani_v | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP24.1SegReq | IQon1 |  | WIP24.1 | March 12, 2020 5:12:46 PM GMT+08:00
CTUDT00628801 | Refresh | Closed | dannyb | February 20, 2020 4:40:36 PM GMT+08:00 | Refresh/sCTV – Batch is being reset when in Conventional and selecting Tile View | WIP24 |  | do not use Tile View with a batch | Subsystem | sCTV, CRN | NA | No Plan to Fix | aviw | 0 | 3 |  | Support_Spectral | NO | Closed |  | IQon1
CTUDT00628798 | Refresh | Closed | dannyb | February 20, 2020 4:38:53 PM GMT+08:00 | Refresh/sCTV - First viewport in magic glass show Casupp76 instead Casupp with the set index | WIP24 |  |  | Subsystem | sCTV | NA | No Plan to Fix | aviw | 0 | 4 |  | Support_Spectral | NO | Closed |  | IQon1
CTUDT00628797 | Refresh | Closed | cherylli | February 20, 2020 4:31:22 PM GMT+08:00 | BR64 User document: Issues found in 459801096633_d QRG | WIP22.4 | Br64 v4.1.10 RfD | N/A | Subsystem |  | N/A | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | 05-Mar-2020 | Brilliance 64 |  | Br64 v4.1.10 RfD | March 6, 2020 2:26:45 PM GMT+08:00
CTUDT00628794 | Refresh | Closed | zhong_x | February 20, 2020 4:27:58 PM GMT+08:00 | [Refresh] IVC Refresh: Go to inner or outer laser led status of IVC box sometimes not synced with the button on gantry control panel | WIP12.4 |  | Complete the Laser In/Out motion using Gantry panel if it is started from Gantry Panel when this occur. | Subsystem |  | Console-SRS.15655 | No Plan to Fix | 320021835 | 0 | 4 |  | Console4 | YES | Closed |  | Brilliance Big Bore
CTUDT00628793 | Refresh | Closed | cherylli | February 20, 2020 4:22:37 PM GMT+08:00 | BR64 User document: Issues found in 459801608571_c ECB | WIP22.4 | Br64 v4.1.10 RfD | N/A | Subsystem |  | N/A | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | 05-Mar-2020 | Brilliance 64 |  | Br64 v4.1.10 RfD | March 6, 2020 2:23:46 PM GMT+08:00
CTUDT00628791 | Refresh | Closed | cherylli | February 20, 2020 4:16:10 PM GMT+08:00 | BR64 User document: Issues found in 459801063624_d TRG | WIP22.4 | Br64 v4.1.10 RfD | N/A | Subsystem |  | N/A | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | 05-Mar-2020 | Brilliance 64 |  | Br64 v4.1.10 RfD | March 6, 2020 2:20:46 PM GMT+08:00
CTUDT00628788 | Refresh | Closed | wang_l | February 20, 2020 4:01:06 PM GMT+08:00 | Refresh BB - Remote to CIRS failed from Remote Viewer because IP retrieved as 192.168.0.201 | BB_WIP12.4 | DMR356920 Rev10 | If user cannot remote into CIRS through service tools and CIRS is reporting 201 instead of 200, these steps can be tried to address the issue. 1. Re-starting the CIRS 2. Re-start ServiceNode: a. Use mstsc to remote into CIRS b. Open Services shortcut on the desktop c. Restart the ServiceNodeHost service d. Wait approximately 1 minute e. Retry remote connection through remote viewer 3. ServiceNode Delayed Start Potential Fix a. Use mstsc to remote into CIRS b. Open Services shortcut on the desktop c. Double click ServiceNodeHost service d. Change Startup type to “Automatic (Delayed Start)” e. Click Apply, then OK f. Restart the CIRS For O-level users, they may connect directly to the CIRS using a keyboard, mouse, and monitor if necessary. | Subsystem |  | N/A | Fixed | shaham_t | 0 | 3 | Service Documentation | Support-System | NO | Closed | WIP24.1SubCom_26-Mar-2020 | Ovation |  | DMR356920 Rev10 | March 30, 2020 3:41:53 PM GMT+08:00
CTUDT00628784 | Refresh | Closed | cherylli | February 20, 2020 3:47:23 PM GMT+08:00 | BR64 User document: Issues found in 459801063533_d Instruction For Use | WIP22.4 | Br64 v4.1.10 RfD | N/A | Subsystem |  | N/A | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | 05-Mar-2020 | Brilliance 64 |  | Br64 v4.1.10 RfD | March 6, 2020 2:16:38 PM GMT+08:00
CTUDT00628781 | Refresh | Closed | zhong_x | February 20, 2020 3:27:16 PM GMT+08:00 | [Refresh] Breath Rate value and irradiation indicator are not displayed. | WIP12.4 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | zhong_x | 0 |  |  | Support-System | NO | Closed |  | Brilliance Big Bore
CTUDT00628775 | Refresh | Closed | zhong_x | February 20, 2020 9:32:40 AM GMT+08:00 | System not prompt for a password change for IT user first time login after reset password via Service tool. | BB_WIP12.4 | DMR356920 Rev10 | There are two ways to reset the password for the Factory Hospital Administrator user. 1. Philips_service user can use IT-Admin tool to change the password. Philips_Service user can launch IT-Admin tool from C:\Pms\Syste\Philips.PmsCT.AdminTool.Hosting.exe 2. If user uses the STPS tool to reset the password, then after resetting the password, user needs perform the below setps. 1. Launch windows user management tool using "Lusrmgr.msc". 2. Go to IT user and open it properties. 3. Set the flag "User must change password at next logon". 4. Uncheck the flag “Account is Locked out” check box. (This is required when the IT user is locked out due to wrong password entries). Ignore if already unchecked. | Subsystem |  | NA | No Plan to Fix | p_debasmita | 0 | 2 | Service Documentation | Service Tools | NO | Closed | WIP31SegReq | Ovation |  | DMR356920 Rev10 | March 30, 2020 2:51:06 PM GMT+08:00
CTUDT00628775 | Refresh | Closed | zhong_x | February 20, 2020 9:32:40 AM GMT+08:00 | System not prompt for a password change for IT user first time login after reset password via Service tool. | BB_WIP12.4 | DMR356920 Rev10 | There are two ways to reset the password for the Factory Hospital Administrator user. 1. Philips_service user can use IT-Admin tool to change the password. Philips_Service user can launch IT-Admin tool from C:\Pms\Syste\Philips.PmsCT.AdminTool.Hosting.exe 2. If user uses the STPS tool to reset the password, then after resetting the password, user needs perform the below setps. 1. Launch windows user management tool using "Lusrmgr.msc". 2. Go to IT user and open it properties. 3. Set the flag "User must change password at next logon". 4. Uncheck the flag “Account is Locked out” check box. (This is required when the IT user is locked out due to wrong password entries). Ignore if already unchecked. | Subsystem |  | NA | No Plan to Fix | p_debasmita | 0 | 2 | Service Documentation | Service Tools | NO | Closed | WIP31SegReq | Brilliance 64 |  | DMR356920 Rev10 | March 30, 2020 2:51:06 PM GMT+08:00
CTUDT00628775 | Refresh | Closed | zhong_x | February 20, 2020 9:32:40 AM GMT+08:00 | System not prompt for a password change for IT user first time login after reset password via Service tool. | BB_WIP12.4 | DMR356920 Rev10 | There are two ways to reset the password for the Factory Hospital Administrator user. 1. Philips_service user can use IT-Admin tool to change the password. Philips_Service user can launch IT-Admin tool from C:\Pms\Syste\Philips.PmsCT.AdminTool.Hosting.exe 2. If user uses the STPS tool to reset the password, then after resetting the password, user needs perform the below setps. 1. Launch windows user management tool using "Lusrmgr.msc". 2. Go to IT user and open it properties. 3. Set the flag "User must change password at next logon". 4. Uncheck the flag “Account is Locked out” check box. (This is required when the IT user is locked out due to wrong password entries). Ignore if already unchecked. | Subsystem |  | NA | No Plan to Fix | p_debasmita | 0 | 2 | Service Documentation | Service Tools | NO | Closed | WIP31SegReq | Brilliance Big Bore |  | DMR356920 Rev10 | March 30, 2020 2:51:06 PM GMT+08:00
CTUDT00628775 | Refresh | Closed | zhong_x | February 20, 2020 9:32:40 AM GMT+08:00 | System not prompt for a password change for IT user first time login after reset password via Service tool. | BB_WIP12.4 | DMR356920 Rev10 | There are two ways to reset the password for the Factory Hospital Administrator user. 1. Philips_service user can use IT-Admin tool to change the password. Philips_Service user can launch IT-Admin tool from C:\Pms\Syste\Philips.PmsCT.AdminTool.Hosting.exe 2. If user uses the STPS tool to reset the password, then after resetting the password, user needs perform the below setps. 1. Launch windows user management tool using "Lusrmgr.msc". 2. Go to IT user and open it properties. 3. Set the flag "User must change password at next logon". 4. Uncheck the flag “Account is Locked out” check box. (This is required when the IT user is locked out due to wrong password entries). Ignore if already unchecked. | Subsystem |  | NA | No Plan to Fix | p_debasmita | 0 | 2 | Service Documentation | Service Tools | NO | Closed | WIP31SegReq | Ingenuity 64 |  | DMR356920 Rev10 | March 30, 2020 2:51:06 PM GMT+08:00
CTUDT00628775 | Refresh | Closed | zhong_x | February 20, 2020 9:32:40 AM GMT+08:00 | System not prompt for a password change for IT user first time login after reset password via Service tool. | BB_WIP12.4 | DMR356920 Rev10 | There are two ways to reset the password for the Factory Hospital Administrator user. 1. Philips_service user can use IT-Admin tool to change the password. Philips_Service user can launch IT-Admin tool from C:\Pms\Syste\Philips.PmsCT.AdminTool.Hosting.exe 2. If user uses the STPS tool to reset the password, then after resetting the password, user needs perform the below setps. 1. Launch windows user management tool using "Lusrmgr.msc". 2. Go to IT user and open it properties. 3. Set the flag "User must change password at next logon". 4. Uncheck the flag “Account is Locked out” check box. (This is required when the IT user is locked out due to wrong password entries). Ignore if already unchecked. | Subsystem |  | NA | No Plan to Fix | p_debasmita | 0 | 2 | Service Documentation | Service Tools | NO | Closed | WIP31SegReq | IQon1 |  | DMR356920 Rev10 | March 30, 2020 2:51:06 PM GMT+08:00
CTUDT00628772 | Refresh | Closed | shtark_o | February 19, 2020 11:02:53 PM GMT+08:00 | Visual diagnostic CIRS Recon SWBP test is faling | IQON GPU 4.7.5.43533 | DMR356920 Rev10 | Documentation for operations and service would need to be updated to not select this test when testing on a single rack. If the other tests are passing, it is an indication that the recon server is working as expected. This test was just provided to show test coverage for 2D reconstructions. | Subsystem |  | CIRS.SSRS.5362 | Fixed | shaham_t | 0 | 2 | Service Documentation | Support-System | NO | Closed | WIP24.1SubCom_26-Mar-2020 | IQon1 |  | DMR356920 Rev10 | March 30, 2020 3:42:28 PM GMT+08:00
CTUDT00628768 | Refresh | Closed | shterner_a | February 19, 2020 8:35:24 PM GMT+08:00 | Val: Refresh Serviceability: Error "ReconServer.exe has stopped working" popped up when "Create test raw dataset" was launched from PSC-Utilities | WIP23.1 |  |  | Subsystem |  | N/A | No Plan to Fix | bandil_j | 0 | 4 |  | Support-System | NO | Closed |  | Ovation
CTUDT00628768 | Refresh | Closed | shterner_a | February 19, 2020 8:35:24 PM GMT+08:00 | Val: Refresh Serviceability: Error "ReconServer.exe has stopped working" popped up when "Create test raw dataset" was launched from PSC-Utilities | WIP23.1 |  |  | Subsystem |  | N/A | No Plan to Fix | bandil_j | 0 | 4 |  | Support-System | NO | Closed |  | Brilliance 64
CTUDT00628768 | Refresh | Closed | shterner_a | February 19, 2020 8:35:24 PM GMT+08:00 | Val: Refresh Serviceability: Error "ReconServer.exe has stopped working" popped up when "Create test raw dataset" was launched from PSC-Utilities | WIP23.1 |  |  | Subsystem |  | N/A | No Plan to Fix | bandil_j | 0 | 4 |  | Support-System | NO | Closed |  | Brilliance Big Bore
CTUDT00628768 | Refresh | Closed | shterner_a | February 19, 2020 8:35:24 PM GMT+08:00 | Val: Refresh Serviceability: Error "ReconServer.exe has stopped working" popped up when "Create test raw dataset" was launched from PSC-Utilities | WIP23.1 |  |  | Subsystem |  | N/A | No Plan to Fix | bandil_j | 0 | 4 |  | Support-System | NO | Closed |  | Ingenuity 64
CTUDT00628768 | Refresh | Closed | shterner_a | February 19, 2020 8:35:24 PM GMT+08:00 | Val: Refresh Serviceability: Error "ReconServer.exe has stopped working" popped up when "Create test raw dataset" was launched from PSC-Utilities | WIP23.1 |  |  | Subsystem |  | N/A | No Plan to Fix | bandil_j | 0 | 4 |  | Support-System | NO | Closed |  | IQon1
CTUDT00628764 | Refresh | Closed | cai_t | February 19, 2020 8:25:51 PM GMT+08:00 | [HTML IFU]It displayed MK not Macedonian in the language selection list | WIP22.4 | ING v4.1.10 RfD BB v4.8 RfD Br64 v4.1.10 RfD | N/A | Subsystem |  | N/A | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | 05-Mar-2020 | Brilliance 64 |  | ING v4.1.10 RfD BB v4.8 RfD Br64 v4.1.10 RfD | March 28, 2020 6:08:52 PM GMT+08:00
CTUDT00628761 | Refresh | Closed | cai_t | February 19, 2020 8:08:50 PM GMT+08:00 | [HTML IFU]The index content was messy code in Kazakh language | WIP22.4 | v4.8 RfD v4.1.10 RfD |  | Subsystem |  | N/A | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | 05-Mar-2020 | Brilliance Big Bore |  | v4.8 RfD v4.1.10 RfD | March 13, 2020 12:52:18 PM GMT+08:00
CTUDT00628761 | Refresh | Closed | cai_t | February 19, 2020 8:08:50 PM GMT+08:00 | [HTML IFU]The index content was messy code in Kazakh language | WIP22.4 | v4.8 RfD v4.1.10 RfD |  | Subsystem |  | N/A | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | 05-Mar-2020 | Ingenuity 64 |  | v4.8 RfD v4.1.10 RfD | March 13, 2020 12:52:18 PM GMT+08:00
CTUDT00628748 | Refresh | Closed | herip_yapoujian | February 18, 2020 10:19:33 PM GMT+08:00 | Val: Refresh: ExamCardManager has stopped working | WIP23.1 |  | User can load the EC to plan and on performing surview, Min/Max mAs can be edited. In plan, the combobox doesn't allow the user to enter the special characters. Relaunch ECM tool again | Subsystem |  | NA | No Plan to Fix | kr_m | 0 | 3 |  | Console4 | NO | Closed |  | Ovation
CTUDT00628748 | Refresh | Closed | herip_yapoujian | February 18, 2020 10:19:33 PM GMT+08:00 | Val: Refresh: ExamCardManager has stopped working | WIP23.1 |  | User can load the EC to plan and on performing surview, Min/Max mAs can be edited. In plan, the combobox doesn't allow the user to enter the special characters. Relaunch ECM tool again | Subsystem |  | NA | No Plan to Fix | kr_m | 0 | 3 |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00628748 | Refresh | Closed | herip_yapoujian | February 18, 2020 10:19:33 PM GMT+08:00 | Val: Refresh: ExamCardManager has stopped working | WIP23.1 |  | User can load the EC to plan and on performing surview, Min/Max mAs can be edited. In plan, the combobox doesn't allow the user to enter the special characters. Relaunch ECM tool again | Subsystem |  | NA | No Plan to Fix | kr_m | 0 | 3 |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00628748 | Refresh | Closed | herip_yapoujian | February 18, 2020 10:19:33 PM GMT+08:00 | Val: Refresh: ExamCardManager has stopped working | WIP23.1 |  | User can load the EC to plan and on performing surview, Min/Max mAs can be edited. In plan, the combobox doesn't allow the user to enter the special characters. Relaunch ECM tool again | Subsystem |  | NA | No Plan to Fix | kr_m | 0 | 3 |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00628748 | Refresh | Closed | herip_yapoujian | February 18, 2020 10:19:33 PM GMT+08:00 | Val: Refresh: ExamCardManager has stopped working | WIP23.1 |  | User can load the EC to plan and on performing surview, Min/Max mAs can be edited. In plan, the combobox doesn't allow the user to enter the special characters. Relaunch ECM tool again | Subsystem |  | NA | No Plan to Fix | kr_m | 0 | 3 |  | Console4 | NO | Closed |  | IQon1
CTUDT00628746 | Refresh | Closed | shterner_a | February 18, 2020 10:07:24 PM GMT+08:00 | Val: Refresh Serviceability: Philips Support Connect (PSC) crashed when testing Controller utilities. | WIP23.1 |  | Relaunch PSC | Subsystem |  | NA | No Plan to Fix | Sandeep_Kumar | 0 | 3 |  | Service Tools | NO | Closed |  | IQon1
CTUDT00628744 | Refresh | Closed | shterner_a | February 18, 2020 9:38:12 PM GMT+08:00 | Val: Refresh Serviceability: Test Case Modification for Cycle 2 | WIP23.1 | DHF305294 rev 04 |  | Subsystem |  | no | Fixed | dadon_e | 0 | 2 | Validation Documentation | Support-System | NO | Closed | SubCom-17-Mar-2020 | IQon1 |  | DHF305294 rev 04 | March 29, 2020 5:46:08 PM GMT+08:00
CTUDT00628739 | Refresh | Postponed | Bat-El_Asaraf | February 18, 2020 7:33:12 PM GMT+08:00 | Val: Refresh: User password change policy is incorrect | WIP23.1 |  | NA | System | Passwords ; | NA |  | jose_s | 0 | 2 |  | Support-Security | NO | Postponed |  | Ovation
CTUDT00628739 | Refresh | Postponed | Bat-El_Asaraf | February 18, 2020 7:33:12 PM GMT+08:00 | Val: Refresh: User password change policy is incorrect | WIP23.1 |  | NA | System | Passwords ; | NA |  | jose_s | 0 | 2 |  | Support-Security | NO | Postponed |  | Brilliance 64
CTUDT00628739 | Refresh | Postponed | Bat-El_Asaraf | February 18, 2020 7:33:12 PM GMT+08:00 | Val: Refresh: User password change policy is incorrect | WIP23.1 |  | NA | System | Passwords ; | NA |  | jose_s | 0 | 2 |  | Support-Security | NO | Postponed |  | Brilliance Big Bore
CTUDT00628739 | Refresh | Postponed | Bat-El_Asaraf | February 18, 2020 7:33:12 PM GMT+08:00 | Val: Refresh: User password change policy is incorrect | WIP23.1 |  | NA | System | Passwords ; | NA |  | jose_s | 0 | 2 |  | Support-Security | NO | Postponed |  | Ingenuity 64
CTUDT00628739 | Refresh | Postponed | Bat-El_Asaraf | February 18, 2020 7:33:12 PM GMT+08:00 | Val: Refresh: User password change policy is incorrect | WIP23.1 |  | NA | System | Passwords ; | NA |  | jose_s | 0 | 2 |  | Support-Security | NO | Postponed |  | IQon1
CTUDT00628735 | Refresh | Closed | Sandeep_Kumar | February 18, 2020 6:21:27 PM GMT+08:00 | Exam Cards with Unicode characters are not collected in bugrep | WIP24 |  | The missing exam cards (which have Unicode characters in its name) can be collected by exporting using ExamCard Manager by Senior Clinical User (With ExamCard Manager access)/Service User. | Subsystem | CRN | NA | No Plan to Fix | Sandeep_Kumar | 0 | 3 |  | Service Tools | NO | Closed |  | Ovation
CTUDT00628735 | Refresh | Closed | Sandeep_Kumar | February 18, 2020 6:21:27 PM GMT+08:00 | Exam Cards with Unicode characters are not collected in bugrep | WIP24 |  | The missing exam cards (which have Unicode characters in its name) can be collected by exporting using ExamCard Manager by Senior Clinical User (With ExamCard Manager access)/Service User. | Subsystem | CRN | NA | No Plan to Fix | Sandeep_Kumar | 0 | 3 |  | Service Tools | NO | Closed |  | Brilliance 64
CTUDT00628735 | Refresh | Closed | Sandeep_Kumar | February 18, 2020 6:21:27 PM GMT+08:00 | Exam Cards with Unicode characters are not collected in bugrep | WIP24 |  | The missing exam cards (which have Unicode characters in its name) can be collected by exporting using ExamCard Manager by Senior Clinical User (With ExamCard Manager access)/Service User. | Subsystem | CRN | NA | No Plan to Fix | Sandeep_Kumar | 0 | 3 |  | Service Tools | NO | Closed |  | Brilliance Big Bore
CTUDT00628735 | Refresh | Closed | Sandeep_Kumar | February 18, 2020 6:21:27 PM GMT+08:00 | Exam Cards with Unicode characters are not collected in bugrep | WIP24 |  | The missing exam cards (which have Unicode characters in its name) can be collected by exporting using ExamCard Manager by Senior Clinical User (With ExamCard Manager access)/Service User. | Subsystem | CRN | NA | No Plan to Fix | Sandeep_Kumar | 0 | 3 |  | Service Tools | NO | Closed |  | Ingenuity 64
CTUDT00628735 | Refresh | Closed | Sandeep_Kumar | February 18, 2020 6:21:27 PM GMT+08:00 | Exam Cards with Unicode characters are not collected in bugrep | WIP24 |  | The missing exam cards (which have Unicode characters in its name) can be collected by exporting using ExamCard Manager by Senior Clinical User (With ExamCard Manager access)/Service User. | Subsystem | CRN | NA | No Plan to Fix | Sandeep_Kumar | 0 | 3 |  | Service Tools | NO | Closed |  | IQon1
CTUDT00628733 | Refresh | Closed | shiraguppi_k | February 18, 2020 5:29:55 PM GMT+08:00 | Val: Refresh: In Cardiac Viewer ECG Waveform showing wrong Phase point and phase bar after editing the 75% local phase point editing of cycle5, Its showing 75% instead of 60% | WIP23.1 |  | NA | Subsystem | ECG GUI ; | NA | No Plan to Fix | kr_m | 0 | 2 |  | Console4 | NO | Closed |  | Ovation
CTUDT00628733 | Refresh | Closed | shiraguppi_k | February 18, 2020 5:29:55 PM GMT+08:00 | Val: Refresh: In Cardiac Viewer ECG Waveform showing wrong Phase point and phase bar after editing the 75% local phase point editing of cycle5, Its showing 75% instead of 60% | WIP23.1 |  | NA | Subsystem | ECG GUI ; | NA | No Plan to Fix | kr_m | 0 | 2 |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00628733 | Refresh | Closed | shiraguppi_k | February 18, 2020 5:29:55 PM GMT+08:00 | Val: Refresh: In Cardiac Viewer ECG Waveform showing wrong Phase point and phase bar after editing the 75% local phase point editing of cycle5, Its showing 75% instead of 60% | WIP23.1 |  | NA | Subsystem | ECG GUI ; | NA | No Plan to Fix | kr_m | 0 | 2 |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00628733 | Refresh | Closed | shiraguppi_k | February 18, 2020 5:29:55 PM GMT+08:00 | Val: Refresh: In Cardiac Viewer ECG Waveform showing wrong Phase point and phase bar after editing the 75% local phase point editing of cycle5, Its showing 75% instead of 60% | WIP23.1 |  | NA | Subsystem | ECG GUI ; | NA | No Plan to Fix | kr_m | 0 | 2 |  | Console4 | NO | Closed |  | IQon1
CTUDT00628731 | Refresh | Closed | dannyb | February 18, 2020 4:49:46 PM GMT+08:00 | sCTV crashes when pressing Compare (to get out of) two SBIs with spectral ROI | WIP24 | WIP24.1 | Relaunch the Spectral CT Viewer | Subsystem | sCTV; PM_Approved | NA | Fixed | aviw | 0 | 2 | Console USP_EBW | Support_Spectral | NO | Closed | WIP24.1SegReq | IQon1 |  | WIP24.1 | February 26, 2020 6:28:10 PM GMT+08:00
CTUDT00628728 | Refresh | Closed | glauberman_o | February 18, 2020 3:26:06 PM GMT+08:00 | Headline :Val: Refresh:3D modulation message popped up twice | WIP23.1 |  |  | Subsystem | Val; UI | NA | No Plan to Fix | datta_i | 0 | 4 |  | Console4 | NO | Closed |  | IQon1
CTUDT00628721 | Refresh | Closed | wei_f | February 17, 2020 7:41:02 PM GMT+08:00 | Eject the USB mobile hard disk(HDD or SSD) from Patient Directory device list, show error message and failed to eject | WIP22.4 |  | User has to hard eject the HDD or SSD. | Subsystem | CRN | NA | No Plan to Fix | s_panigrahi | 0 | 4 |  | iCAP | NO | Closed |  | Ovation
CTUDT00628721 | Refresh | Closed | wei_f | February 17, 2020 7:41:02 PM GMT+08:00 | Eject the USB mobile hard disk(HDD or SSD) from Patient Directory device list, show error message and failed to eject | WIP22.4 |  | User has to hard eject the HDD or SSD. | Subsystem | CRN | NA | No Plan to Fix | s_panigrahi | 0 | 4 |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00628721 | Refresh | Closed | wei_f | February 17, 2020 7:41:02 PM GMT+08:00 | Eject the USB mobile hard disk(HDD or SSD) from Patient Directory device list, show error message and failed to eject | WIP22.4 |  | User has to hard eject the HDD or SSD. | Subsystem | CRN | NA | No Plan to Fix | s_panigrahi | 0 | 4 |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00628721 | Refresh | Closed | wei_f | February 17, 2020 7:41:02 PM GMT+08:00 | Eject the USB mobile hard disk(HDD or SSD) from Patient Directory device list, show error message and failed to eject | WIP22.4 |  | User has to hard eject the HDD or SSD. | Subsystem | CRN | NA | No Plan to Fix | s_panigrahi | 0 | 4 |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00628721 | Refresh | Closed | wei_f | February 17, 2020 7:41:02 PM GMT+08:00 | Eject the USB mobile hard disk(HDD or SSD) from Patient Directory device list, show error message and failed to eject | WIP22.4 |  | User has to hard eject the HDD or SSD. | Subsystem | CRN | NA | No Plan to Fix | s_panigrahi | 0 | 4 |  | iCAP | NO | Closed |  | IQon1
CTUDT00628705 | Refresh | Closed | shterner_a | February 16, 2020 5:58:46 PM GMT+08:00 | Val: Refresh: Clinical Test Modification for Cycle 2 | WIP23.1 | DHF305293 Rev.05 |  | Subsystem |  | No | Fixed | shterner_a | 0 | 2 | Test Design | Support-System | NO | Closed | SubCom-09-Mar-2020 | IQon1 |  | DHF305293 Rev.05 | March 8, 2020 3:49:37 PM GMT+08:00
CTUDT00628699 | Refresh | Closed | shaham_t | February 13, 2020 10:16:55 PM GMT+08:00 | Service Validation: Restore crashed after reinstallation | WIP24 | DMR356920 Rev10 |  | System |  | NA | Fixed | shaham_t | 0 | 2 | Service Documentation | Support-System | NO | Closed | WIP24.1SubCom_26-Mar-2020 | Ovation |  | DMR356920 Rev10 | March 29, 2020 5:16:16 PM GMT+08:00
CTUDT00628699 | Refresh | Closed | shaham_t | February 13, 2020 10:16:55 PM GMT+08:00 | Service Validation: Restore crashed after reinstallation | WIP24 | DMR356920 Rev10 |  | System |  | NA | Fixed | shaham_t | 0 | 2 | Service Documentation | Support-System | NO | Closed | WIP24.1SubCom_26-Mar-2020 | IQon1 |  | DMR356920 Rev10 | March 29, 2020 5:16:16 PM GMT+08:00
CTUDT00628699 | Refresh | Closed | shaham_t | February 13, 2020 10:16:55 PM GMT+08:00 | Service Validation: Restore crashed after reinstallation | WIP24 | DMR356920 Rev10 |  | System |  | NA | Fixed | shaham_t | 0 | 2 | Service Documentation | Support-System | NO | Closed | WIP24.1SubCom_26-Mar-2020 | Hawk |  | DMR356920 Rev10 | March 29, 2020 5:16:16 PM GMT+08:00
CTUDT00628698 | Refresh | Closed | shiraguppi_k | February 13, 2020 10:14:18 PM GMT+08:00 | Val: Refresh : While performing CCT Single scan, after 3 exposures Gantry Cannot Initialize message popped and scan got Paused | WIP23.1 |  |  | Subsystem |  |  | System Repaired | yu_x | 0 |  |  | GCX | YES | Closed |  | Ovation
CTUDT00628691 | Refresh | Resolved | shiraguppi_k | February 13, 2020 9:04:03 PM GMT+08:00 | Val:Refresh: Patient registration page blank after end exam and Patient Directory stopped working while executing Performance Test | WIP23.1 | WIP30 | For the PD crash issue , clicking on any of the option will re-initialize PD and make the PD active. | System | SP1; FOK | N/A | Fixed | s_sa | 0 | 3 | Console USP_EBW | iCAP | NO | Resolved | WIP30SubCom | Ovation | PR409880 PR293042 PR295524 PR510004 | WIP30 | May 13, 2020 6:41:20 PM GMT+08:00
CTUDT00628691 | Refresh | Resolved | shiraguppi_k | February 13, 2020 9:04:03 PM GMT+08:00 | Val:Refresh: Patient registration page blank after end exam and Patient Directory stopped working while executing Performance Test | WIP23.1 | WIP30 | For the PD crash issue , clicking on any of the option will re-initialize PD and make the PD active. | System | SP1; FOK | N/A | Fixed | s_sa | 0 | 3 | Console USP_EBW | iCAP | NO | Resolved | WIP30SubCom | Brilliance 64 | PR409880 PR293042 PR295524 PR510004 | WIP30 | May 13, 2020 6:41:20 PM GMT+08:00
CTUDT00628691 | Refresh | Resolved | shiraguppi_k | February 13, 2020 9:04:03 PM GMT+08:00 | Val:Refresh: Patient registration page blank after end exam and Patient Directory stopped working while executing Performance Test | WIP23.1 | WIP30 | For the PD crash issue , clicking on any of the option will re-initialize PD and make the PD active. | System | SP1; FOK | N/A | Fixed | s_sa | 0 | 3 | Console USP_EBW | iCAP | NO | Resolved | WIP30SubCom | Brilliance Big Bore | PR409880 PR293042 PR295524 PR510004 | WIP30 | May 13, 2020 6:41:20 PM GMT+08:00
CTUDT00628691 | Refresh | Resolved | shiraguppi_k | February 13, 2020 9:04:03 PM GMT+08:00 | Val:Refresh: Patient registration page blank after end exam and Patient Directory stopped working while executing Performance Test | WIP23.1 | WIP30 | For the PD crash issue , clicking on any of the option will re-initialize PD and make the PD active. | System | SP1; FOK | N/A | Fixed | s_sa | 0 | 3 | Console USP_EBW | iCAP | NO | Resolved | WIP30SubCom | Ingenuity 64 | PR409880 PR293042 PR295524 PR510004 | WIP30 | May 13, 2020 6:41:20 PM GMT+08:00
CTUDT00628691 | Refresh | Resolved | shiraguppi_k | February 13, 2020 9:04:03 PM GMT+08:00 | Val:Refresh: Patient registration page blank after end exam and Patient Directory stopped working while executing Performance Test | WIP23.1 | WIP30 | For the PD crash issue , clicking on any of the option will re-initialize PD and make the PD active. | System | SP1; FOK | N/A | Fixed | s_sa | 0 | 3 | Console USP_EBW | iCAP | NO | Resolved | WIP30SubCom | IQon1 | PR409880 PR293042 PR295524 PR510004 | WIP30 | May 13, 2020 6:41:20 PM GMT+08:00
CTUDT00628688 | Refresh | Closed | vinoliya_j | February 13, 2020 8:43:21 PM GMT+08:00 | 459800917216_C - iQon Technical reference guide validation comments | WIP22.1 | IQon v4.7.7 RfD | N/A | Subsystem |  | N/A | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | 05-Mar-2020 | IQon1 |  | IQon v4.7.7 RfD | March 5, 2020 2:39:12 PM GMT+08:00
CTUDT00628682 | Refresh | Closed | zhong_x | February 13, 2020 10:54:33 AM GMT+08:00 | Varian: Disconnect the cable of file server during Varian scan, ‘ExamAppliation.exe has stopped working’ pop up after scan completed. | WIP12.2 |  |  | Subsystem | Blocker; | N/A | No Plan to Fix | navale_s | 0 | 3 |  | Console4 | YES | Closed |  | Brilliance Big Bore
CTUDT00628680 | Refresh | Closed | zhong_x | February 13, 2020 10:36:24 AM GMT+08:00 | Varian: Disconnect the cable of file server after loading patient to plan scene, wave will not be imported after scan is completed. | WIP12.2 |  |  | Subsystem |  | N/A | No Plan to Fix | navale_s | 0 | 3 |  | Console4 | YES | Closed |  | Brilliance 64
CTUDT00628680 | Refresh | Closed | zhong_x | February 13, 2020 10:36:24 AM GMT+08:00 | Varian: Disconnect the cable of file server after loading patient to plan scene, wave will not be imported after scan is completed. | WIP12.2 |  |  | Subsystem |  | N/A | No Plan to Fix | navale_s | 0 | 3 |  | Console4 | YES | Closed |  | Brilliance Big Bore
CTUDT00628680 | Refresh | Closed | zhong_x | February 13, 2020 10:36:24 AM GMT+08:00 | Varian: Disconnect the cable of file server after loading patient to plan scene, wave will not be imported after scan is completed. | WIP12.2 |  |  | Subsystem |  | N/A | No Plan to Fix | navale_s | 0 | 3 |  | Console4 | YES | Closed |  | Ingenuity 64
CTUDT00628680 | Refresh | Closed | zhong_x | February 13, 2020 10:36:24 AM GMT+08:00 | Varian: Disconnect the cable of file server after loading patient to plan scene, wave will not be imported after scan is completed. | WIP12.2 |  |  | Subsystem |  | N/A | No Plan to Fix | navale_s | 0 | 3 |  | Console4 | YES | Closed |  | Ingenuity TF 64
CTUDT00628680 | Refresh | Closed | zhong_x | February 13, 2020 10:36:24 AM GMT+08:00 | Varian: Disconnect the cable of file server after loading patient to plan scene, wave will not be imported after scan is completed. | WIP12.2 |  |  | Subsystem |  | N/A | No Plan to Fix | navale_s | 0 | 3 |  | Console4 | YES | Closed |  | Veyron
CTUDT00628678 | Refresh | Closed | vinoliya_j | February 13, 2020 1:50:56 AM GMT+08:00 | 459801063557_C - iCT Technical reference guide validation comments | WIP23.1 | iCT v4.1.10 RfD | N/A | Subsystem |  | N/A | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | 05-Mar-2020 | Ovation |  | iCT v4.1.10 RfD | March 5, 2020 2:31:19 PM GMT+08:00
CTUDT00628676 | Refresh | Closed | vinoliya_j | February 13, 2020 1:21:02 AM GMT+08:00 | 459801583751_C - iCT exam card book validation comments | WIP23.1 | iCT v4.1.10 RfD | N/A | Subsystem |  | N/A | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | 05-Mar-2020 | Ovation |  | iCT v4.1.10 RfD | March 5, 2020 2:12:43 PM GMT+08:00
CTUDT00628674 | Refresh | Closed | shiraguppi_k | February 12, 2020 11:33:17 PM GMT+08:00 | 459801063515_C - iCT IFU validation comments | WIP23.1 | v4.1.10 RfD | N/A | Subsystem |  | N/A | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | 05-Mar-2020 | Ovation |  | v4.1.10 RfD | March 5, 2020 2:08:44 PM GMT+08:00
CTUDT00628672 | Refresh | Closed | Noy_Seelig | February 12, 2020 9:34:57 PM GMT+08:00 | Refresh: Air calibration in CT user displays wrong progress after estop opens. | WIP24 |  |  | Subsystem |  | n/a | No Plan to Fix | frider_o | 0 | 4 |  | DMS | NO | Closed | SubNC | Ovation
CTUDT00628668 | Refresh | Closed | ran_y | February 12, 2020 8:48:44 PM GMT+08:00 | Refresh; ICT\ICTSP:WIP24: Body - Ultra high - Filter E - Max. iDose Level = 7 instead of 6. | WIP24 | DHF298553 Rev.10 DHF298554 Rev.12 |  | Subsystem | Requirements; Verified1 | CT-NM.SRS.IQ.3129 | Fixed | shaham_t | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP24SegReq | Ovation |  | DHF298553 Rev.10 DHF298554 Rev.12 | February 25, 2020 6:15:38 PM GMT+08:00
CTUDT00628668 | Refresh | Closed | ran_y | February 12, 2020 8:48:44 PM GMT+08:00 | Refresh; ICT\ICTSP:WIP24: Body - Ultra high - Filter E - Max. iDose Level = 7 instead of 6. | WIP24 | DHF298553 Rev.10 DHF298554 Rev.12 |  | Subsystem | Requirements; Verified1 | CT-NM.SRS.IQ.3129 | Fixed | shaham_t | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP24SegReq | Brilliance 64 |  | DHF298553 Rev.10 DHF298554 Rev.12 | February 25, 2020 6:15:38 PM GMT+08:00
CTUDT00628668 | Refresh | Closed | ran_y | February 12, 2020 8:48:44 PM GMT+08:00 | Refresh; ICT\ICTSP:WIP24: Body - Ultra high - Filter E - Max. iDose Level = 7 instead of 6. | WIP24 | DHF298553 Rev.10 DHF298554 Rev.12 |  | Subsystem | Requirements; Verified1 | CT-NM.SRS.IQ.3129 | Fixed | shaham_t | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP24SegReq | Ingenuity 64 |  | DHF298553 Rev.10 DHF298554 Rev.12 | February 25, 2020 6:15:38 PM GMT+08:00
CTUDT00628666 | Refresh | Closed | shiraguppi_k | February 12, 2020 8:09:21 PM GMT+08:00 | Val: Refresh : Physics Quality Assurance stopped working message popped up when opening Air calibration report from QA Tab. | WIP23.1 | WIP31 | User shall wait for 30 to 40 seconds for report generation. | Subsystem | QA; SP1_PM_Approved | n/a | Fixed | elviral | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP31SubCom | Ovation |  | WIP31 | July 5, 2020 6:41:53 PM GMT+08:00
CTUDT00628666 | Refresh | Closed | shiraguppi_k | February 12, 2020 8:09:21 PM GMT+08:00 | Val: Refresh : Physics Quality Assurance stopped working message popped up when opening Air calibration report from QA Tab. | WIP23.1 | WIP31 | User shall wait for 30 to 40 seconds for report generation. | Subsystem | QA; SP1_PM_Approved | n/a | Fixed | elviral | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP31SubCom | Brilliance 64 |  | WIP31 | July 5, 2020 6:41:53 PM GMT+08:00
CTUDT00628666 | Refresh | Closed | shiraguppi_k | February 12, 2020 8:09:21 PM GMT+08:00 | Val: Refresh : Physics Quality Assurance stopped working message popped up when opening Air calibration report from QA Tab. | WIP23.1 | WIP31 | User shall wait for 30 to 40 seconds for report generation. | Subsystem | QA; SP1_PM_Approved | n/a | Fixed | elviral | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP31SubCom | Brilliance Big Bore |  | WIP31 | July 5, 2020 6:41:53 PM GMT+08:00
CTUDT00628666 | Refresh | Closed | shiraguppi_k | February 12, 2020 8:09:21 PM GMT+08:00 | Val: Refresh : Physics Quality Assurance stopped working message popped up when opening Air calibration report from QA Tab. | WIP23.1 | WIP31 | User shall wait for 30 to 40 seconds for report generation. | Subsystem | QA; SP1_PM_Approved | n/a | Fixed | elviral | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP31SubCom | Ingenuity 64 |  | WIP31 | July 5, 2020 6:41:53 PM GMT+08:00
CTUDT00628666 | Refresh | Closed | shiraguppi_k | February 12, 2020 8:09:21 PM GMT+08:00 | Val: Refresh : Physics Quality Assurance stopped working message popped up when opening Air calibration report from QA Tab. | WIP23.1 | WIP31 | User shall wait for 30 to 40 seconds for report generation. | Subsystem | QA; SP1_PM_Approved | n/a | Fixed | elviral | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP31SubCom | IQon1 |  | WIP31 | July 5, 2020 6:41:53 PM GMT+08:00
CTUDT00628661 | Refresh | Closed | dannyb | February 11, 2020 10:59:07 PM GMT+08:00 | sCTV crashes when swapping viewports and moving the magic glass | WIP24 |  |  | Subsystem |  | NA | No Plan to Fix | aviw | 0 | 3 |  | Support_Spectral | NO | Closed |  | IQon1
CTUDT00628658 | Refresh | Closed | vinoliya_j | February 11, 2020 9:47:54 PM GMT+08:00 | 459800619353_C - iCT User guide validation comments | WIP23.1 | iCT v4.1.10 RfD | n/A | Subsystem |  | N/A | Fixed | charles_j | 0 | 2 | User Documentation | Support-Service | NO | Closed | 05-Mar-2020 | Ovation |  | iCT v4.1.10 RfD | March 5, 2020 2:16:39 PM GMT+08:00
CTUDT00628645 | Refresh | Closed | dannyb | February 10, 2020 10:58:20 PM GMT+08:00 | sCTV on Refresh: Crash when loading to sCTV SBI and reconed MPR and selecting Slab2x2 | WIP24 | WIP24.1 | Load the SBI only, create the MPR in the application. | System | CRN; SP1 | NA | Fixed | aviw | 0 | 3 | SDViewer | Support_Spectral | NO | Closed | WIP31SegReq | IQon1 |  | WIP24.1 | July 5, 2020 3:29:03 PM GMT+08:00
CTUDT00628643 | Refresh | Closed | cai_t | February 10, 2020 8:47:47 PM GMT+08:00 | Big Bore MultiLanguage IFU has some errors | BB_WIP12 | v4.8 RfD | N/A | Subsystem |  | N/A | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | 17-Feb-2020 | Brilliance Big Bore |  | v4.8 RfD | February 14, 2020 7:14:19 PM GMT+08:00
CTUDT00628641 | Refresh | Closed | cai_t | February 10, 2020 8:43:33 PM GMT+08:00 | Ingenuity MultiLanguage IFU has some errors | WIP22.4 | v4.1.10 RfD | N/A | Subsystem |  | N/A | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | 17-Feb-2020 | Ingenuity 64 |  | v4.1.10 RfD | February 14, 2020 7:10:39 PM GMT+08:00
CTUDT00628639 | Refresh | Closed | cai_t | February 10, 2020 8:35:50 PM GMT+08:00 | Br64 MultiLanguage IFU has some errors | WIP22.4 | Br64 v4.1.10 RfD |  | Subsystem | IFU | N/A | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | 31-Mar-2020 | Brilliance 64 |  | Br64 v4.1.10 RfD | March 28, 2020 6:33:16 PM GMT+08:00
CTUDT00628634 | Refresh | Closed | cai_t | February 10, 2020 7:09:55 PM GMT+08:00 | Vietnamese regulatory Label and Albanian, Bosnian and Vietnamese IFUs are not ready | WIP22.4 | DHF298554 Rev11 |  | Subsystem |  | CT-NM.SRS.SysSupport.806, CT-NM.SRS.SysSupport.807, CT-NM.SRS.SysSupport.875, CT-NM.SRS.SysSupport.878 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | SRS next revision | Brilliance 64 |  | DHF298554 Rev11 | February 21, 2020 2:08:32 PM GMT+08:00
CTUDT00628626 | Refresh | Closed | glauberman_o | February 10, 2020 5:34:48 PM GMT+08:00 | Headline:Val: Refresh:ExamApplication Form Error | WIP23.1 |  |  | Subsystem |  |  | System Repaired | yang_h | 0 |  |  | GCX | YES | Closed |  | IQon1
CTUDT00628606 | Refresh | Closed | 320053770 | February 9, 2020 9:01:56 PM GMT+08:00 | Refresh; IQon: User Reliability issues: WIP24:WF crash after Bolus EC scan. | WIP24 | WIP24.1 WIP22.5 WIP12.5 |  | Subsystem |  | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | YES | Closed | WIP24.1SegReq | Brilliance 64 |  | WIP24.1 WIP22.5 WIP12.5 | March 16, 2020 4:54:15 PM GMT+08:00
CTUDT00628606 | Refresh | Closed | 320053770 | February 9, 2020 9:01:56 PM GMT+08:00 | Refresh; IQon: User Reliability issues: WIP24:WF crash after Bolus EC scan. | WIP24 | WIP24.1 WIP22.5 WIP12.5 |  | Subsystem |  | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | YES | Closed | WIP24.1SegReq | Brilliance Big Bore |  | WIP24.1 WIP22.5 WIP12.5 | March 16, 2020 4:54:15 PM GMT+08:00
CTUDT00628606 | Refresh | Closed | 320053770 | February 9, 2020 9:01:56 PM GMT+08:00 | Refresh; IQon: User Reliability issues: WIP24:WF crash after Bolus EC scan. | WIP24 | WIP24.1 WIP22.5 WIP12.5 |  | Subsystem |  | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | YES | Closed | WIP24.1SegReq | Ingenuity 64 |  | WIP24.1 WIP22.5 WIP12.5 | March 16, 2020 4:54:15 PM GMT+08:00
CTUDT00628597 | Refresh | Closed | vinoliya_j | February 7, 2020 8:04:46 PM GMT+08:00 | 459800975174_C - Iqon User guide validation comments | WIP22.1 | IQon v4.7.7 RfD | N/A | Subsystem |  | N/A | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | 05-Mar-2020 | IQon1 |  | IQon v4.7.7 RfD | March 5, 2020 2:48:12 PM GMT+08:00
CTUDT00628594 | Refresh | Closed | vinoliya_j | February 7, 2020 7:42:00 PM GMT+08:00 | 459800917206_C - Iqon Instructions for use validation comments | WIP22.1 | IQon IFU v4.7.7 RfD | N/A | Subsystem |  | N/A | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | 05-Mar-2020 | IQon1 |  | IQon IFU v4.7.7 RfD | March 5, 2020 2:43:33 PM GMT+08:00
CTUDT00628583 | Refresh | Closed | shaham_t | February 6, 2020 10:41:19 PM GMT+08:00 | Refresh; Upgrade tool usage is too complicated, lacking critical information regarding progress status. | WIP23 | 459801587391 Rev C 459801587381 Rev D 459801587371 Rev F |  | Subsystem |  | NA | Fixed | zhang_l | 0 | 2 | Service Tools_Software | Support-Service | NO | Closed | SubCom-21-Feb-2020 | Brilliance 64 |  | 459801587391 Rev C 459801587381 Rev D 459801587371 Rev F | February 23, 2020 3:01:56 PM GMT+08:00
CTUDT00628583 | Refresh | Closed | shaham_t | February 6, 2020 10:41:19 PM GMT+08:00 | Refresh; Upgrade tool usage is too complicated, lacking critical information regarding progress status. | WIP23 | 459801587391 Rev C 459801587381 Rev D 459801587371 Rev F |  | Subsystem |  | NA | Fixed | zhang_l | 0 | 2 | Service Tools_Software | Support-Service | NO | Closed | SubCom-21-Feb-2020 | Brilliance Big Bore |  | 459801587391 Rev C 459801587381 Rev D 459801587371 Rev F | February 23, 2020 3:01:56 PM GMT+08:00
CTUDT00628583 | Refresh | Closed | shaham_t | February 6, 2020 10:41:19 PM GMT+08:00 | Refresh; Upgrade tool usage is too complicated, lacking critical information regarding progress status. | WIP23 | 459801587391 Rev C 459801587381 Rev D 459801587371 Rev F |  | Subsystem |  | NA | Fixed | zhang_l | 0 | 2 | Service Tools_Software | Support-Service | NO | Closed | SubCom-21-Feb-2020 | Ingenuity 64 |  | 459801587391 Rev C 459801587381 Rev D 459801587371 Rev F | February 23, 2020 3:01:56 PM GMT+08:00
CTUDT00628581 | Refresh | Closed | 320053770 | February 6, 2020 9:32:20 PM GMT+08:00 | Refresh; IQon: WIP24: WF crash during WIP installation | WIP24 |  |  | Subsystem |  |  | Duplicate | s_sa | 0 |  |  | iCAP | NO | Closed |  | IQon1
CTUDT00628576 | Refresh | Closed | pB_s | February 6, 2020 7:52:19 PM GMT+08:00 | All results are displayed with same procedure description in View tab and PD even though study got split | WIP24 |  |  | Subsystem |  | NA | No Plan to Fix | datta_i | 0 | 2 |  | Console4 | NO | Closed |  | IQon1
CTUDT00628551 | Refresh | Closed | 320053770 | February 4, 2020 9:30:49 PM GMT+08:00 | Refresh; ICT: WIP23.1: E-stop was opened during a reliability test. | WIP23.1 |  |  | Subsystem |  |  | Invalid/Uncontrolled Configuration | benny_w | 0 |  |  | GCX | NO | Closed |  | Ovation
CTUDT00628548 | Refresh | Closed | Revital_Cohen | February 4, 2020 8:54:15 PM GMT+08:00 | Refresh: The two Requirements:43 and 599 according to the TRF\CSA reports are N\A | WIP23 |  |  | Subsystem |  |  | Does Not Apply | tamari_o | 0 |  |  | Support-System | NO | Closed |  | Ovation
CTUDT00628545 | Refresh | Closed | Revital_Cohen | February 4, 2020 8:13:44 PM GMT+08:00 | Refresh: Missing clauses in TRF/CSA Reports | WIP23 |  |  | Subsystem |  |  | Does Not Apply | tamari_o | 0 |  |  |  |  | Closed |  | Ovation
CTUDT00628545 | Refresh | Closed | Revital_Cohen | February 4, 2020 8:13:44 PM GMT+08:00 | Refresh: Missing clauses in TRF/CSA Reports | WIP23 |  |  | Subsystem |  |  | Does Not Apply | tamari_o | 0 |  |  |  |  | Closed |  | IQon1
CTUDT00628542 | Refresh | Closed | sun_q | February 4, 2020 5:41:57 PM GMT+08:00 | No Device Size, Weight and Power Supply information in Russia System Regulatory Label | 459801771921 RevA | Ingenuity/BB gantry SSRS rev07 |  | Subsystem |  | GCX.SSRS.iCT.20475 | Fixed | wang_k | 0 | 2 | Design Documentation | GCX | NO | Closed | ssrs NEXT VERSION | Brilliance Big Bore |  | Ingenuity/BB gantry SSRS rev07 | February 11, 2020 4:26:59 PM GMT+08:00
CTUDT00628542 | Refresh | Closed | sun_q | February 4, 2020 5:41:57 PM GMT+08:00 | No Device Size, Weight and Power Supply information in Russia System Regulatory Label | 459801771921 RevA | Ingenuity/BB gantry SSRS rev07 |  | Subsystem |  | GCX.SSRS.iCT.20475 | Fixed | wang_k | 0 | 2 | Design Documentation | GCX | NO | Closed | ssrs NEXT VERSION | Ingenuity 64 |  | Ingenuity/BB gantry SSRS rev07 | February 11, 2020 4:26:59 PM GMT+08:00
CTUDT00628539 | Refresh | Closed | vinoliya_j | February 3, 2020 11:45:18 PM GMT+08:00 | 459801656511_C - Iqon Exam card book Validation comments | WIP22.1 | IQon v4.7.7 RfD | N/A | Subsystem |  | N/A | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | 05-Mar-2020 | IQon1 |  | IQon v4.7.7 RfD | March 6, 2020 2:04:47 PM GMT+08:00
CTUDT00628534 | Refresh | Closed | 320053770 | February 3, 2020 10:31:03 PM GMT+08:00 | Refresh; ICT: WIP23.1: ICT 459801496591- Service tools user guide is not covering CT-NM.SRS.SecPri.233. | Wip23 | 459801787151 Rev A |  | System |  | N/A | Fixed | golden_d | 0 | 2 | Service Documentation | Support-Service | NO | Closed | SubCom-05-Mar-2020 | Ovation |  | 459801787151 Rev A | March 3, 2020 3:27:38 PM GMT+08:00
CTUDT00628534 | Refresh | Closed | 320053770 | February 3, 2020 10:31:03 PM GMT+08:00 | Refresh; ICT: WIP23.1: ICT 459801496591- Service tools user guide is not covering CT-NM.SRS.SecPri.233. | Wip23 | 459801787151 Rev A |  | System |  | N/A | Fixed | golden_d | 0 | 2 | Service Documentation | Support-Service | NO | Closed | SubCom-05-Mar-2020 | IQon1 |  | 459801787151 Rev A | March 3, 2020 3:27:38 PM GMT+08:00
CTUDT00628532 | Refresh | Closed | 320053769 | February 3, 2020 10:15:51 PM GMT+08:00 | WIP23.1 – “Critical Error”, “examApplication.exe is not responding” and “system error” warnings appeared | WIP23.1 |  |  | Subsystem |  | NA | Duplicate | datta_i | 0 | 2 |  | Console4 | NO | Closed |  | Ovation
CTUDT00628532 | Refresh | Closed | 320053769 | February 3, 2020 10:15:51 PM GMT+08:00 | WIP23.1 – “Critical Error”, “examApplication.exe is not responding” and “system error” warnings appeared | WIP23.1 |  |  | Subsystem |  | NA | Duplicate | datta_i | 0 | 2 |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00628532 | Refresh | Closed | 320053769 | February 3, 2020 10:15:51 PM GMT+08:00 | WIP23.1 – “Critical Error”, “examApplication.exe is not responding” and “system error” warnings appeared | WIP23.1 |  |  | Subsystem |  | NA | Duplicate | datta_i | 0 | 2 |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00628532 | Refresh | Closed | 320053769 | February 3, 2020 10:15:51 PM GMT+08:00 | WIP23.1 – “Critical Error”, “examApplication.exe is not responding” and “system error” warnings appeared | WIP23.1 |  |  | Subsystem |  | NA | Duplicate | datta_i | 0 | 2 |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00628532 | Refresh | Closed | 320053769 | February 3, 2020 10:15:51 PM GMT+08:00 | WIP23.1 – “Critical Error”, “examApplication.exe is not responding” and “system error” warnings appeared | WIP23.1 |  |  | Subsystem |  | NA | Duplicate | datta_i | 0 | 2 |  | Console4 | NO | Closed |  | IQon1
CTUDT00628527 | Refresh | Closed | 320053769 | February 3, 2020 9:33:21 PM GMT+08:00 | Test Display images at the system monitor TG18-LN8-01 to TG18-LN8-18 labels aren’t equal to test display image at "DIN6868-157 Images Testing" CD (459801424391). | WIP22 | WIP24 |  | Subsystem | Test Design | CT-NM.SRS.RegStand.691 | Fixed | tamari_o | 0 | 2 | Test Design | Support-System | NO | Closed | SubCom-25-Feb-2020 | IQon1 |  | WIP24 | March 2, 2020 1:14:05 AM GMT+08:00
CTUDT00628521 | Refresh | Closed | 320053769 | February 3, 2020 7:40:56 PM GMT+08:00 | Customer Information Sheet (459801424351) isn’t compatible with the newer version of Snagit. | WIP22 | WIP24 |  | Subsystem |  | CT-NM.SRS.RegStand.691 & CT-NM.SRS.RegStand.692 | Fixed | tamari_o | 0 | 3 | Test Design | Support-System | NO | Closed | SubCom-25-Feb-2020 | IQon1 |  | WIP24 | February 26, 2020 7:31:51 PM GMT+08:00
CTUDT00628519 | Refresh | Closed | bandil_j | February 3, 2020 5:23:16 PM GMT+08:00 | WIP23.1 – IDose result of OMAR isn't reconstructed when clicking on the end exam button before clicking on final recon button. | WIP23.1 | BB_WIP12.4 WIP22.4 | User can use offline recon to create Non O-MAR result. User can also use Create Final Recon to create all the results. | Subsystem |  | N/A | Fixed | kapadia_v | 0 | 3 | Console Software | Console4 | NO | Closed | WIP12.4 & WIP22.4 | Brilliance 64 |  | BB_WIP12.4 WIP22.4 | February 7, 2020 2:00:29 PM GMT+08:00
CTUDT00628519 | Refresh | Closed | bandil_j | February 3, 2020 5:23:16 PM GMT+08:00 | WIP23.1 – IDose result of OMAR isn't reconstructed when clicking on the end exam button before clicking on final recon button. | WIP23.1 | BB_WIP12.4 WIP22.4 | User can use offline recon to create Non O-MAR result. User can also use Create Final Recon to create all the results. | Subsystem |  | N/A | Fixed | kapadia_v | 0 | 3 | Console Software | Console4 | NO | Closed | WIP12.4 & WIP22.4 | Brilliance Big Bore |  | BB_WIP12.4 WIP22.4 | February 7, 2020 2:00:29 PM GMT+08:00
CTUDT00628519 | Refresh | Closed | bandil_j | February 3, 2020 5:23:16 PM GMT+08:00 | WIP23.1 – IDose result of OMAR isn't reconstructed when clicking on the end exam button before clicking on final recon button. | WIP23.1 | BB_WIP12.4 WIP22.4 | User can use offline recon to create Non O-MAR result. User can also use Create Final Recon to create all the results. | Subsystem |  | N/A | Fixed | kapadia_v | 0 | 3 | Console Software | Console4 | NO | Closed | WIP12.4 & WIP22.4 | Ingenuity 64 |  | BB_WIP12.4 WIP22.4 | February 7, 2020 2:00:29 PM GMT+08:00
CTUDT00628517 | Refresh | Closed | bandil_j | February 3, 2020 5:10:35 PM GMT+08:00 | Incorrect procedure description assigned to results in a split study. | WIP22.3 | BB_WIP12.4 WIP22.4 |  | Subsystem | Blocker | N/A | Fixed | pandey_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP12.4 & WIP22.4 | Brilliance 64 |  | BB_WIP12.4 WIP22.4 | February 7, 2020 11:20:24 PM GMT+08:00
CTUDT00628517 | Refresh | Closed | bandil_j | February 3, 2020 5:10:35 PM GMT+08:00 | Incorrect procedure description assigned to results in a split study. | WIP22.3 | BB_WIP12.4 WIP22.4 |  | Subsystem | Blocker | N/A | Fixed | pandey_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP12.4 & WIP22.4 | Brilliance Big Bore |  | BB_WIP12.4 WIP22.4 | February 7, 2020 11:20:24 PM GMT+08:00
CTUDT00628517 | Refresh | Closed | bandil_j | February 3, 2020 5:10:35 PM GMT+08:00 | Incorrect procedure description assigned to results in a split study. | WIP22.3 | BB_WIP12.4 WIP22.4 |  | Subsystem | Blocker | N/A | Fixed | pandey_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP12.4 & WIP22.4 | Ingenuity 64 |  | BB_WIP12.4 WIP22.4 | February 7, 2020 11:20:24 PM GMT+08:00
CTUDT00628509 | Refresh | Closed | dadon_e | February 3, 2020 4:30:42 AM GMT+08:00 | Refresh, Val (WIP22.1): IQon Bay 140, error during launch LAN config | WIP22.1 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | s_panigrahi | 1 |  |  | iCAP | NO | Closed |  | IQon1
CTUDT00628505 | Refresh | Closed | 320053769 | February 2, 2020 11:01:49 PM GMT+08:00 | If DIN 6861-2 is enabled in preferences, when choosing result DICOM Info - Din DICOM tags values columns appear one column prior to non-DIN DICOM tags value. | WIP23.1 |  |  | Subsystem |  | NA | No Plan to Fix | srushti_k | 0 | 4 |  | iCAP | NO | Closed |  | Ovation
CTUDT00628505 | Refresh | Closed | 320053769 | February 2, 2020 11:01:49 PM GMT+08:00 | If DIN 6861-2 is enabled in preferences, when choosing result DICOM Info - Din DICOM tags values columns appear one column prior to non-DIN DICOM tags value. | WIP23.1 |  |  | Subsystem |  | NA | No Plan to Fix | srushti_k | 0 | 4 |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00628505 | Refresh | Closed | 320053769 | February 2, 2020 11:01:49 PM GMT+08:00 | If DIN 6861-2 is enabled in preferences, when choosing result DICOM Info - Din DICOM tags values columns appear one column prior to non-DIN DICOM tags value. | WIP23.1 |  |  | Subsystem |  | NA | No Plan to Fix | srushti_k | 0 | 4 |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00628505 | Refresh | Closed | 320053769 | February 2, 2020 11:01:49 PM GMT+08:00 | If DIN 6861-2 is enabled in preferences, when choosing result DICOM Info - Din DICOM tags values columns appear one column prior to non-DIN DICOM tags value. | WIP23.1 |  |  | Subsystem |  | NA | No Plan to Fix | srushti_k | 0 | 4 |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00628505 | Refresh | Closed | 320053769 | February 2, 2020 11:01:49 PM GMT+08:00 | If DIN 6861-2 is enabled in preferences, when choosing result DICOM Info - Din DICOM tags values columns appear one column prior to non-DIN DICOM tags value. | WIP23.1 |  |  | Subsystem |  | NA | No Plan to Fix | srushti_k | 0 | 4 |  | iCAP | NO | Closed |  | IQon1
CTUDT00628497 | Refresh | Closed | dadon_e | February 2, 2020 6:07:20 PM GMT+08:00 | Refresh, Val (WIP22.1): Severe system slowness due to wrong OS setting - enable write cache (calibration took 6 hours) | WIP22.1 | DMR356920 Rev.09 459801587391 Rev C 459801587381 Rev D 459801587371 Rev F 459801587331 Rev F 459801587341 Rev F 459801587361 Rev E 459801587351 Rev E | User can check(enable) the “Enable write caching on the device” once the installation is complete. | Subsystem | Val; Calibrations; Verified1 | NA | Fixed | shaham_t | 0 | 3 | Console Software | Support-Service | NO | Closed | SubCom-21-Feb-2020 | Ovation |  | DMR356920 Rev.09 459801587391 Rev C 459801587381 Rev D 459801587371 Rev F 459801587331 Rev F 459801587341 Rev F 459801587361 Rev E 459801587351 Rev E | February 23, 2020 3:01:40 PM GMT+08:00
CTUDT00628497 | Refresh | Closed | dadon_e | February 2, 2020 6:07:20 PM GMT+08:00 | Refresh, Val (WIP22.1): Severe system slowness due to wrong OS setting - enable write cache (calibration took 6 hours) | WIP22.1 | DMR356920 Rev.09 459801587391 Rev C 459801587381 Rev D 459801587371 Rev F 459801587331 Rev F 459801587341 Rev F 459801587361 Rev E 459801587351 Rev E | User can check(enable) the “Enable write caching on the device” once the installation is complete. | Subsystem | Val; Calibrations; Verified1 | NA | Fixed | shaham_t | 0 | 3 | Console Software | Support-Service | NO | Closed | SubCom-21-Feb-2020 | Brilliance 64 |  | DMR356920 Rev.09 459801587391 Rev C 459801587381 Rev D 459801587371 Rev F 459801587331 Rev F 459801587341 Rev F 459801587361 Rev E 459801587351 Rev E | February 23, 2020 3:01:40 PM GMT+08:00
CTUDT00628497 | Refresh | Closed | dadon_e | February 2, 2020 6:07:20 PM GMT+08:00 | Refresh, Val (WIP22.1): Severe system slowness due to wrong OS setting - enable write cache (calibration took 6 hours) | WIP22.1 | DMR356920 Rev.09 459801587391 Rev C 459801587381 Rev D 459801587371 Rev F 459801587331 Rev F 459801587341 Rev F 459801587361 Rev E 459801587351 Rev E | User can check(enable) the “Enable write caching on the device” once the installation is complete. | Subsystem | Val; Calibrations; Verified1 | NA | Fixed | shaham_t | 0 | 3 | Console Software | Support-Service | NO | Closed | SubCom-21-Feb-2020 | Brilliance Big Bore |  | DMR356920 Rev.09 459801587391 Rev C 459801587381 Rev D 459801587371 Rev F 459801587331 Rev F 459801587341 Rev F 459801587361 Rev E 459801587351 Rev E | February 23, 2020 3:01:40 PM GMT+08:00
CTUDT00628497 | Refresh | Closed | dadon_e | February 2, 2020 6:07:20 PM GMT+08:00 | Refresh, Val (WIP22.1): Severe system slowness due to wrong OS setting - enable write cache (calibration took 6 hours) | WIP22.1 | DMR356920 Rev.09 459801587391 Rev C 459801587381 Rev D 459801587371 Rev F 459801587331 Rev F 459801587341 Rev F 459801587361 Rev E 459801587351 Rev E | User can check(enable) the “Enable write caching on the device” once the installation is complete. | Subsystem | Val; Calibrations; Verified1 | NA | Fixed | shaham_t | 0 | 3 | Console Software | Support-Service | NO | Closed | SubCom-21-Feb-2020 | Ingenuity 64 |  | DMR356920 Rev.09 459801587391 Rev C 459801587381 Rev D 459801587371 Rev F 459801587331 Rev F 459801587341 Rev F 459801587361 Rev E 459801587351 Rev E | February 23, 2020 3:01:40 PM GMT+08:00
CTUDT00628497 | Refresh | Closed | dadon_e | February 2, 2020 6:07:20 PM GMT+08:00 | Refresh, Val (WIP22.1): Severe system slowness due to wrong OS setting - enable write cache (calibration took 6 hours) | WIP22.1 | DMR356920 Rev.09 459801587391 Rev C 459801587381 Rev D 459801587371 Rev F 459801587331 Rev F 459801587341 Rev F 459801587361 Rev E 459801587351 Rev E | User can check(enable) the “Enable write caching on the device” once the installation is complete. | Subsystem | Val; Calibrations; Verified1 | NA | Fixed | shaham_t | 0 | 3 | Console Software | Support-Service | NO | Closed | SubCom-21-Feb-2020 | IQon1 |  | DMR356920 Rev.09 459801587391 Rev C 459801587381 Rev D 459801587371 Rev F 459801587331 Rev F 459801587341 Rev F 459801587361 Rev E 459801587351 Rev E | February 23, 2020 3:01:40 PM GMT+08:00
CTUDT00628484 | Refresh | Closed | richards_b | January 31, 2020 3:09:09 AM GMT+08:00 | 3D Helical additional image upload reli issue | WIP23.1 | WIP24 BB_WIP12.4 WIP22.4 | N/A | Subsystem | GPU; PM_Approved | N/A | Fixed | richards_b | 0 | 4 | CIRS_SW_ImgNodeSvr | CIRS | NO | Closed | WIP24SubCom | Ovation |  | WIP24 BB_WIP12.4 WIP22.4 | February 6, 2020 5:53:51 AM GMT+08:00
CTUDT00628484 | Refresh | Closed | richards_b | January 31, 2020 3:09:09 AM GMT+08:00 | 3D Helical additional image upload reli issue | WIP23.1 | WIP24 BB_WIP12.4 WIP22.4 | N/A | Subsystem | GPU; PM_Approved | N/A | Fixed | richards_b | 0 | 4 | CIRS_SW_ImgNodeSvr | CIRS | NO | Closed | WIP24SubCom | Brilliance 64 |  | WIP24 BB_WIP12.4 WIP22.4 | February 6, 2020 5:53:51 AM GMT+08:00
CTUDT00628484 | Refresh | Closed | richards_b | January 31, 2020 3:09:09 AM GMT+08:00 | 3D Helical additional image upload reli issue | WIP23.1 | WIP24 BB_WIP12.4 WIP22.4 | N/A | Subsystem | GPU; PM_Approved | N/A | Fixed | richards_b | 0 | 4 | CIRS_SW_ImgNodeSvr | CIRS | NO | Closed | WIP24SubCom | Brilliance Big Bore |  | WIP24 BB_WIP12.4 WIP22.4 | February 6, 2020 5:53:51 AM GMT+08:00
CTUDT00628484 | Refresh | Closed | richards_b | January 31, 2020 3:09:09 AM GMT+08:00 | 3D Helical additional image upload reli issue | WIP23.1 | WIP24 BB_WIP12.4 WIP22.4 | N/A | Subsystem | GPU; PM_Approved | N/A | Fixed | richards_b | 0 | 4 | CIRS_SW_ImgNodeSvr | CIRS | NO | Closed | WIP24SubCom | Ingenuity 64 |  | WIP24 BB_WIP12.4 WIP22.4 | February 6, 2020 5:53:51 AM GMT+08:00
CTUDT00628484 | Refresh | Closed | richards_b | January 31, 2020 3:09:09 AM GMT+08:00 | 3D Helical additional image upload reli issue | WIP23.1 | WIP24 BB_WIP12.4 WIP22.4 | N/A | Subsystem | GPU; PM_Approved | N/A | Fixed | richards_b | 0 | 4 | CIRS_SW_ImgNodeSvr | CIRS | NO | Closed | WIP24SubCom | IQon1 |  | WIP24 BB_WIP12.4 WIP22.4 | February 6, 2020 5:53:51 AM GMT+08:00
CTUDT00628479 | Refresh | Closed | herip_yapoujian | January 31, 2020 2:34:07 AM GMT+08:00 | Val: Refresh: ROI shifting In Bolus Tracking Tests | WIP22.1 |  | NA | Subsystem | Auto ROI | NA | Does Not Apply | kr_m | 0 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00628452 | Refresh | Closed | vineetha_s | January 30, 2020 2:53:57 PM GMT+08:00 | While loading DCM file from iWatt, Operator Name field in console is coming blank. | BB_WIP12.3 |  | User can type the operator name. | Subsystem | Connectivity; | N/A | No Plan to Fix | navale_s | 0 | 4 |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00628449 | Refresh | Closed | vineetha_s | January 30, 2020 2:41:34 PM GMT+08:00 | System created multiple studies in PD while selecting the DCM file with English characters | WIP22.3 |  |  | Subsystem | Chinese | N/A | Duplicate | pandey_a | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00628441 | Refresh | Closed | 320053769 | January 29, 2020 10:31:32 PM GMT+08:00 | WIP23.1 – IDose result of OMAR isn't reconstructed when clicking on the end exam button before clicking on final recon button. | WIP23.1 | WIP24 | User can use offline recon to create Non O-MAR result. User can also use Create Final Recon to create all the results. | Subsystem | PM_Approved | NA | Fixed | kapadia_v | 0 | 3 | Console Software | Console4 | NO | Closed | WIP24SubCom | Ovation |  | WIP24 | February 3, 2020 9:07:56 PM GMT+08:00
CTUDT00628441 | Refresh | Closed | 320053769 | January 29, 2020 10:31:32 PM GMT+08:00 | WIP23.1 – IDose result of OMAR isn't reconstructed when clicking on the end exam button before clicking on final recon button. | WIP23.1 | WIP24 | User can use offline recon to create Non O-MAR result. User can also use Create Final Recon to create all the results. | Subsystem | PM_Approved | NA | Fixed | kapadia_v | 0 | 3 | Console Software | Console4 | NO | Closed | WIP24SubCom | IQon1 |  | WIP24 | February 3, 2020 9:07:56 PM GMT+08:00
CTUDT00628402 | Refresh | Closed | 320053768 | January 27, 2020 9:49:09 PM GMT+08:00 | Infant Body Performance fails on ICT with the Harmonized Phantom. | WIP23.1 | WIP24 |  | Subsystem | Blocker; PM_Approved | n/a | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP24SegReq | Ovation |  | WIP24 | January 30, 2020 3:26:26 PM GMT+08:00
CTUDT00628400 | Refresh | Closed | 320053770 | January 27, 2020 9:29:03 PM GMT+08:00 | Refresh; ICT: WIP23.1: Scanner Name was changed from the name which was selected during 1-click installation. | WIP23.1 | DMR356920 Rev09 | Incase issue occours, Service personal can still update the "Computer name" using Service tools in Computer and Network configuration settings. procedure updated in Refresh SW Installation specification | Subsystem |  | N/A | Fixed | shaham_t | 0 | 3 | Service Documentation | Support-System | NO | Closed | SubCom_11-Feb-2020 | Ovation |  | DMR356920 Rev09 | February 17, 2020 8:07:54 PM GMT+08:00
CTUDT00628398 | Refresh | Closed | 320053770 | January 27, 2020 9:18:48 PM GMT+08:00 | Refresh; ICT: WIP23.1: Acceptance test fails. | WIP23.1 | WIP24 |  | Subsystem | Blocker; PM_Approved | n/a | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP24SegReq | Ovation |  | WIP24 | January 30, 2020 3:28:16 PM GMT+08:00
CTUDT00628398 | Refresh | Closed | 320053770 | January 27, 2020 9:18:48 PM GMT+08:00 | Refresh; ICT: WIP23.1: Acceptance test fails. | WIP23.1 | WIP24 |  | Subsystem | Blocker; PM_Approved | n/a | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP24SegReq | IQon1 |  | WIP24 | January 30, 2020 3:28:16 PM GMT+08:00
CTUDT00628394 | Refresh | Closed | yatziv_l | January 27, 2020 6:12:05 PM GMT+08:00 | Refresh: Constancy: Uniformity typical kV (120 kV) limits of Adult Body on Head to be updated | WIP23.1 | WIP31 |  | Subsystem | SP1_PM_Approved | N/A | Fixed | elviral | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP31SubCom | Brilliance Big Bore |  | WIP31 | July 5, 2020 6:43:20 PM GMT+08:00
CTUDT00628394 | Refresh | Closed | yatziv_l | January 27, 2020 6:12:05 PM GMT+08:00 | Refresh: Constancy: Uniformity typical kV (120 kV) limits of Adult Body on Head to be updated | WIP23.1 | WIP31 |  | Subsystem | SP1_PM_Approved | N/A | Fixed | elviral | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP31SubCom | Ingenuity 64 |  | WIP31 | July 5, 2020 6:43:20 PM GMT+08:00
CTUDT00628375 | Refresh | Closed | glauberman_o | January 26, 2020 9:42:40 PM GMT+08:00 | Val: Refresh: Error Exam Application changing the Lateral Surview to Dual | WIP22.1 |  | start exam with current patient and not changing surview | Subsystem | Reliability; Validation; Blocker | NA | No Plan to Fix | datta_i | 0 | 3 |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00628375 | Refresh | Closed | glauberman_o | January 26, 2020 9:42:40 PM GMT+08:00 | Val: Refresh: Error Exam Application changing the Lateral Surview to Dual | WIP22.1 |  | start exam with current patient and not changing surview | Subsystem | Reliability; Validation; Blocker | NA | No Plan to Fix | datta_i | 0 | 3 |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00628375 | Refresh | Closed | glauberman_o | January 26, 2020 9:42:40 PM GMT+08:00 | Val: Refresh: Error Exam Application changing the Lateral Surview to Dual | WIP22.1 |  | start exam with current patient and not changing surview | Subsystem | Reliability; Validation; Blocker | NA | No Plan to Fix | datta_i | 0 | 3 |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00628375 | Refresh | Closed | glauberman_o | January 26, 2020 9:42:40 PM GMT+08:00 | Val: Refresh: Error Exam Application changing the Lateral Surview to Dual | WIP22.1 |  | start exam with current patient and not changing surview | Subsystem | Reliability; Validation; Blocker | NA | No Plan to Fix | datta_i | 0 | 3 |  | Console4 | NO | Closed |  | IQon1
CTUDT00628354 | Refresh | Closed | mN_m | January 24, 2020 5:27:31 PM GMT+08:00 | Bugrep error displayed while performing “Resp Single Phase” scan using Varian device. | BB_WIP12.3 |  |  | Subsystem |  |  | Duplicate | s_sa | 0 |  |  | iCAP | YES | Closed |  | Brilliance Big Bore
CTUDT00628352 | Refresh | Closed | 320053770 | January 24, 2020 4:42:59 AM GMT+08:00 | Refresh; IQon\ICT: Upgrade tool usage is too complicated, lacking critical information regarding progress status. | WIP23 | DMR356920 Rev09 |  | Subsystem | Upgrade utility | NA | Fixed | shaham_t | 0 | 2 | Validation Documentation | Support-System | NO | Closed | SubCom_11-Feb-2020 | Ovation |  | DMR356920 Rev09 | February 17, 2020 7:21:36 PM GMT+08:00
CTUDT00628352 | Refresh | Closed | 320053770 | January 24, 2020 4:42:59 AM GMT+08:00 | Refresh; IQon\ICT: Upgrade tool usage is too complicated, lacking critical information regarding progress status. | WIP23 | DMR356920 Rev09 |  | Subsystem | Upgrade utility | NA | Fixed | shaham_t | 0 | 2 | Validation Documentation | Support-System | NO | Closed | SubCom_11-Feb-2020 | IQon1 |  | DMR356920 Rev09 | February 17, 2020 7:21:36 PM GMT+08:00
CTUDT00628339 | Refresh | Closed | 320053769 | January 23, 2020 7:44:40 PM GMT+08:00 | IQon and iCT IFU - under section "Device Specific Features and Instructions", page number in line Dose Efficiency Warning Message is cut off and missing details. | WIP23.1 | iCT IFU v4.1.10 RfD IQon IFU v4.7.7 RfD | N/A | Subsystem | IFU | CT-NM.SRS.RegStand.728 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | 05-Mar-2020 | Ovation |  | iCT IFU v4.1.10 RfD IQon IFU v4.7.7 RfD | March 5, 2020 2:26:49 PM GMT+08:00
CTUDT00628339 | Refresh | Closed | 320053769 | January 23, 2020 7:44:40 PM GMT+08:00 | IQon and iCT IFU - under section "Device Specific Features and Instructions", page number in line Dose Efficiency Warning Message is cut off and missing details. | WIP23.1 | iCT IFU v4.1.10 RfD IQon IFU v4.7.7 RfD | N/A | Subsystem | IFU | CT-NM.SRS.RegStand.728 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | 05-Mar-2020 | IQon1 |  | iCT IFU v4.1.10 RfD IQon IFU v4.7.7 RfD | March 5, 2020 2:26:49 PM GMT+08:00
CTUDT00628333 | Refresh | Closed | mN_m | January 23, 2020 6:27:00 PM GMT+08:00 | Incorrect procedure description assigned to results in a split study. | WIP22.3 | WIP24 |  | Subsystem | Procedure description; Blocker; PM_Approved | N/A | Fixed | pandey_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP24SubCom | Ovation |  | WIP24 | February 3, 2020 11:43:54 PM GMT+08:00
CTUDT00628333 | Refresh | Closed | mN_m | January 23, 2020 6:27:00 PM GMT+08:00 | Incorrect procedure description assigned to results in a split study. | WIP22.3 | WIP24 |  | Subsystem | Procedure description; Blocker; PM_Approved | N/A | Fixed | pandey_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP24SubCom | IQon1 |  | WIP24 | February 3, 2020 11:43:54 PM GMT+08:00
CTUDT00628329 | Refresh | Closed | kumar_tp | January 23, 2020 5:09:13 PM GMT+08:00 | While keyboard localization in Italian, letter "€" is displayed in Plan viewer as "?" after surveiw scan. | BB_WIP12.3 |  |  | Subsystem | CRN; | None | No Plan to Fix | deshbhratar_m | 0 | 3 |  | Console4 | NO | Closed |  | Ovation
CTUDT00628329 | Refresh | Closed | kumar_tp | January 23, 2020 5:09:13 PM GMT+08:00 | While keyboard localization in Italian, letter "€" is displayed in Plan viewer as "?" after surveiw scan. | BB_WIP12.3 |  |  | Subsystem | CRN; | None | No Plan to Fix | deshbhratar_m | 0 | 3 |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00628329 | Refresh | Closed | kumar_tp | January 23, 2020 5:09:13 PM GMT+08:00 | While keyboard localization in Italian, letter "€" is displayed in Plan viewer as "?" after surveiw scan. | BB_WIP12.3 |  |  | Subsystem | CRN; | None | No Plan to Fix | deshbhratar_m | 0 | 3 |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00628329 | Refresh | Closed | kumar_tp | January 23, 2020 5:09:13 PM GMT+08:00 | While keyboard localization in Italian, letter "€" is displayed in Plan viewer as "?" after surveiw scan. | BB_WIP12.3 |  |  | Subsystem | CRN; | None | No Plan to Fix | deshbhratar_m | 0 | 3 |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00628329 | Refresh | Closed | kumar_tp | January 23, 2020 5:09:13 PM GMT+08:00 | While keyboard localization in Italian, letter "€" is displayed in Plan viewer as "?" after surveiw scan. | BB_WIP12.3 |  |  | Subsystem | CRN; | None | No Plan to Fix | deshbhratar_m | 0 | 3 |  | Console4 | NO | Closed |  | IQon1
CTUDT00628317 | Refresh | Closed | gushansky_m | January 22, 2020 11:04:06 PM GMT+08:00 | Refresh : Remote Console connection over RSN from PRS portal doesn't work | WIP22 | 459800958861 Rev K 459801497481 Rev C | Connect using Remote Console (IST) M2M | System | SRN | NA | Fixed | golden_d | 0 | 3 | Service Documentation | Support-Service | NO | Closed | SubCom-02-Mar-2020 | iCT4 |  | 459800958861 Rev K 459801497481 Rev C | March 3, 2020 8:43:04 PM GMT+08:00
CTUDT00628317 | Refresh | Closed | gushansky_m | January 22, 2020 11:04:06 PM GMT+08:00 | Refresh : Remote Console connection over RSN from PRS portal doesn't work | WIP22 | 459800958861 Rev K 459801497481 Rev C | Connect using Remote Console (IST) M2M | System | SRN | NA | Fixed | golden_d | 0 | 3 | Service Documentation | Support-Service | NO | Closed | SubCom-02-Mar-2020 | IQon1 |  | 459800958861 Rev K 459801497481 Rev C | March 3, 2020 8:43:04 PM GMT+08:00
CTUDT00628310 | Refresh | Closed | glauberman_o | January 22, 2020 7:48:06 PM GMT+08:00 | Val: Refresh: During Bugrep collection Windows Command Processor has stopped working | WIP22.1 | BB_WIP12.5 WIP22.5 | User can start the bug report collection again. | Subsystem | Bugrep; Validation; Blocker; Verified1 | NA | Fixed | srushti_k | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP24.1SegReq | Brilliance 64 |  | BB_WIP12.5 WIP22.5 | February 26, 2020 2:14:18 PM GMT+08:00
CTUDT00628310 | Refresh | Closed | glauberman_o | January 22, 2020 7:48:06 PM GMT+08:00 | Val: Refresh: During Bugrep collection Windows Command Processor has stopped working | WIP22.1 | BB_WIP12.5 WIP22.5 | User can start the bug report collection again. | Subsystem | Bugrep; Validation; Blocker; Verified1 | NA | Fixed | srushti_k | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP24.1SegReq | Brilliance Big Bore |  | BB_WIP12.5 WIP22.5 | February 26, 2020 2:14:18 PM GMT+08:00
CTUDT00628310 | Refresh | Closed | glauberman_o | January 22, 2020 7:48:06 PM GMT+08:00 | Val: Refresh: During Bugrep collection Windows Command Processor has stopped working | WIP22.1 | BB_WIP12.5 WIP22.5 | User can start the bug report collection again. | Subsystem | Bugrep; Validation; Blocker; Verified1 | NA | Fixed | srushti_k | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP24.1SegReq | Ingenuity 64 |  | BB_WIP12.5 WIP22.5 | February 26, 2020 2:14:18 PM GMT+08:00
CTUDT00628307 | Refresh | Closed | zhong_x | January 22, 2020 4:12:14 PM GMT+08:00 | [Refresh] The latest two studies missed after config ISP | BB_WIP12.2 |  | rebuild the local database | Subsystem | CRN | NA | No Plan to Fix | s_panigrahi | 0 | 4 |  | iCAP | YES | Closed |  | Ovation
CTUDT00628307 | Refresh | Closed | zhong_x | January 22, 2020 4:12:14 PM GMT+08:00 | [Refresh] The latest two studies missed after config ISP | BB_WIP12.2 |  | rebuild the local database | Subsystem | CRN | NA | No Plan to Fix | s_panigrahi | 0 | 4 |  | iCAP | YES | Closed |  | Brilliance 64
CTUDT00628307 | Refresh | Closed | zhong_x | January 22, 2020 4:12:14 PM GMT+08:00 | [Refresh] The latest two studies missed after config ISP | BB_WIP12.2 |  | rebuild the local database | Subsystem | CRN | NA | No Plan to Fix | s_panigrahi | 0 | 4 |  | iCAP | YES | Closed |  | Brilliance Big Bore
CTUDT00628307 | Refresh | Closed | zhong_x | January 22, 2020 4:12:14 PM GMT+08:00 | [Refresh] The latest two studies missed after config ISP | BB_WIP12.2 |  | rebuild the local database | Subsystem | CRN | NA | No Plan to Fix | s_panigrahi | 0 | 4 |  | iCAP | YES | Closed |  | Ingenuity 64
CTUDT00628307 | Refresh | Closed | zhong_x | January 22, 2020 4:12:14 PM GMT+08:00 | [Refresh] The latest two studies missed after config ISP | BB_WIP12.2 |  | rebuild the local database | Subsystem | CRN | NA | No Plan to Fix | s_panigrahi | 0 | 4 |  | iCAP | YES | Closed |  | IQon1
CTUDT00628300 | Refresh | Closed | 320053770 | January 21, 2020 7:02:10 PM GMT+08:00 | Refresh; IQon\ICT: GPC installation is not possible to perform by following DMR356920 installation instructions. | WIP23 | DMR356920 Rev09 |  | Subsystem |  | N/A | Fixed | shaham_t | 0 | 2 | Service Documentation | Support-System | NO | Closed | SubCom_09-Feb-2020 | Ovation |  | DMR356920 Rev09 | February 17, 2020 4:26:30 PM GMT+08:00
CTUDT00628300 | Refresh | Closed | 320053770 | January 21, 2020 7:02:10 PM GMT+08:00 | Refresh; IQon\ICT: GPC installation is not possible to perform by following DMR356920 installation instructions. | WIP23 | DMR356920 Rev09 |  | Subsystem |  | N/A | Fixed | shaham_t | 0 | 2 | Service Documentation | Support-System | NO | Closed | SubCom_09-Feb-2020 | IQon1 |  | DMR356920 Rev09 | February 17, 2020 4:26:30 PM GMT+08:00
CTUDT00628298 | Refresh | Closed | 320053770 | January 21, 2020 6:50:42 PM GMT+08:00 | Refresh; IQon; CIRS installation is not possible to perform by following DMR356920 installation instructions. | WIP23 | WIP24 | Keyboard and mouse can be used to complete SW installation and configuration. | Subsystem | Blocker; PM_Approved | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_Installation | CIRS | NO | Closed | WIP24SubCom | IQon1 |  | WIP24 | February 6, 2020 5:54:19 AM GMT+08:00
CTUDT00628297 | Refresh | Closed | bandil_j | January 21, 2020 5:26:03 PM GMT+08:00 | WIP19 - Snagit editor has stopped Working | WIP19 | BB_WIP12.4 WIP22.4 | This issue does not occur for all the users. One can login with different user and use the snagit. If it is occurring for clinical user i.e. CT, one can create new Clinical user (CT1) using user management tool and use CT1 user. With New CT1 user Snagit will work as expected. To resolve the Crash, Snagit tool can be uninstall from the impacted system. There are few points needs to be consider before uninstalling Snagit. 1.FSE needs to make sure that default printer is not set to Snagit. 2.Run Snagit button will be disabled in Preferences Save Image screen. 3.PrintSc key will not be useful, as it is configured to launch with Snagit. | Subsystem |  | N/A | Fixed | kapadia_v | 0 | 4 | Console Software | Console4 | NO | Closed | WIP12.4&WIP22.4 | Brilliance 64 |  | BB_WIP12.4 WIP22.4 | February 6, 2020 7:30:25 PM GMT+08:00
CTUDT00628297 | Refresh | Closed | bandil_j | January 21, 2020 5:26:03 PM GMT+08:00 | WIP19 - Snagit editor has stopped Working | WIP19 | BB_WIP12.4 WIP22.4 | This issue does not occur for all the users. One can login with different user and use the snagit. If it is occurring for clinical user i.e. CT, one can create new Clinical user (CT1) using user management tool and use CT1 user. With New CT1 user Snagit will work as expected. To resolve the Crash, Snagit tool can be uninstall from the impacted system. There are few points needs to be consider before uninstalling Snagit. 1.FSE needs to make sure that default printer is not set to Snagit. 2.Run Snagit button will be disabled in Preferences Save Image screen. 3.PrintSc key will not be useful, as it is configured to launch with Snagit. | Subsystem |  | N/A | Fixed | kapadia_v | 0 | 4 | Console Software | Console4 | NO | Closed | WIP12.4&WIP22.4 | Brilliance Big Bore |  | BB_WIP12.4 WIP22.4 | February 6, 2020 7:30:25 PM GMT+08:00
CTUDT00628297 | Refresh | Closed | bandil_j | January 21, 2020 5:26:03 PM GMT+08:00 | WIP19 - Snagit editor has stopped Working | WIP19 | BB_WIP12.4 WIP22.4 | This issue does not occur for all the users. One can login with different user and use the snagit. If it is occurring for clinical user i.e. CT, one can create new Clinical user (CT1) using user management tool and use CT1 user. With New CT1 user Snagit will work as expected. To resolve the Crash, Snagit tool can be uninstall from the impacted system. There are few points needs to be consider before uninstalling Snagit. 1.FSE needs to make sure that default printer is not set to Snagit. 2.Run Snagit button will be disabled in Preferences Save Image screen. 3.PrintSc key will not be useful, as it is configured to launch with Snagit. | Subsystem |  | N/A | Fixed | kapadia_v | 0 | 4 | Console Software | Console4 | NO | Closed | WIP12.4&WIP22.4 | Ingenuity 64 |  | BB_WIP12.4 WIP22.4 | February 6, 2020 7:30:25 PM GMT+08:00
CTUDT00628277 | Refresh | Closed | 320053768 | January 20, 2020 7:51:00 PM GMT+08:00 | Uniformity-Adult-Body test is failed, uniformity values are out of limits for iQon/iCT. | WIP23 | wip24 DHF298556 Rev.11 DHF298553 Rev.10 DHF170395 Rev.15 | Re-running HCOR and then acceptance again until test passes | Subsystem | IQ; SP1 | CT-NM.SRS.IQ.2208, DHF170395 | Fixed | shaham_t | 0 | 3 | DMS Calibrations | Support-System | NO | Closed | SubCom_28-Feb-2020 | Ovation |  | wip24 DHF298556 Rev.11 DHF298553 Rev.10 DHF170395 Rev.15 | March 3, 2020 5:43:16 PM GMT+08:00
CTUDT00628277 | Refresh | Closed | 320053768 | January 20, 2020 7:51:00 PM GMT+08:00 | Uniformity-Adult-Body test is failed, uniformity values are out of limits for iQon/iCT. | WIP23 | wip24 DHF298556 Rev.11 DHF298553 Rev.10 DHF170395 Rev.15 | Re-running HCOR and then acceptance again until test passes | Subsystem | IQ; SP1 | CT-NM.SRS.IQ.2208, DHF170395 | Fixed | shaham_t | 0 | 3 | DMS Calibrations | Support-System | NO | Closed | SubCom_28-Feb-2020 | IQon1 |  | wip24 DHF298556 Rev.11 DHF298553 Rev.10 DHF170395 Rev.15 | March 3, 2020 5:43:16 PM GMT+08:00
CTUDT00628262 | Refresh | Closed | glauberman_o | January 19, 2020 9:33:09 PM GMT+08:00 | Val: Refresh: Software error during liver3phase with P3T | WIP22.1 |  | Perform offline reconstruction if any results not reconstructed | Subsystem | Validation; P3T; | NA | No Plan to Fix | tiwari_g | 0 | 3 |  | Console4 | YES | Closed | WIP31SegReq | Ovation
CTUDT00628262 | Refresh | Closed | glauberman_o | January 19, 2020 9:33:09 PM GMT+08:00 | Val: Refresh: Software error during liver3phase with P3T | WIP22.1 |  | Perform offline reconstruction if any results not reconstructed | Subsystem | Validation; P3T; | NA | No Plan to Fix | tiwari_g | 0 | 3 |  | Console4 | YES | Closed | WIP31SegReq | Brilliance 64
CTUDT00628262 | Refresh | Closed | glauberman_o | January 19, 2020 9:33:09 PM GMT+08:00 | Val: Refresh: Software error during liver3phase with P3T | WIP22.1 |  | Perform offline reconstruction if any results not reconstructed | Subsystem | Validation; P3T; | NA | No Plan to Fix | tiwari_g | 0 | 3 |  | Console4 | YES | Closed | WIP31SegReq | Brilliance Big Bore
CTUDT00628262 | Refresh | Closed | glauberman_o | January 19, 2020 9:33:09 PM GMT+08:00 | Val: Refresh: Software error during liver3phase with P3T | WIP22.1 |  | Perform offline reconstruction if any results not reconstructed | Subsystem | Validation; P3T; | NA | No Plan to Fix | tiwari_g | 0 | 3 |  | Console4 | YES | Closed | WIP31SegReq | Ingenuity 64
CTUDT00628262 | Refresh | Closed | glauberman_o | January 19, 2020 9:33:09 PM GMT+08:00 | Val: Refresh: Software error during liver3phase with P3T | WIP22.1 |  | Perform offline reconstruction if any results not reconstructed | Subsystem | Validation; P3T; | NA | No Plan to Fix | tiwari_g | 0 | 3 |  | Console4 | YES | Closed | WIP31SegReq | IQon1
CTUDT00628259 | Refresh | Closed | guy_sheaffer | January 19, 2020 4:34:56 PM GMT+08:00 | Val: Refresh: Patient Directory has stopped working | WIP22.1 |  |  | Subsystem | Blocker |  | Duplicate | s_sa | 0 |  |  | iCAP | NO | Closed | WIP24SegReq | Ovation
CTUDT00628259 | Refresh | Closed | guy_sheaffer | January 19, 2020 4:34:56 PM GMT+08:00 | Val: Refresh: Patient Directory has stopped working | WIP22.1 |  |  | Subsystem | Blocker |  | Duplicate | s_sa | 0 |  |  | iCAP | NO | Closed | WIP24SegReq | IQon1
CTUDT00628258 | Refresh | Closed | guy_sheaffer | January 19, 2020 4:32:03 PM GMT+08:00 | Val: Refresh: Critical error occurred when trying to copy images to PACS | WIP22.1 | WIP24 | Relaunch the application. | Subsystem | Validation; Blocker; PM_Approved | NA | Fixed | sagar_v | 0 | 4 | Console Software | Console4 | NO | Closed | WIP24SubCom | Ovation |  | WIP24 | January 28, 2020 2:37:38 PM GMT+08:00
CTUDT00628258 | Refresh | Closed | guy_sheaffer | January 19, 2020 4:32:03 PM GMT+08:00 | Val: Refresh: Critical error occurred when trying to copy images to PACS | WIP22.1 | WIP24 | Relaunch the application. | Subsystem | Validation; Blocker; PM_Approved | NA | Fixed | sagar_v | 0 | 4 | Console Software | Console4 | NO | Closed | WIP24SubCom | IQon1 |  | WIP24 | January 28, 2020 2:37:38 PM GMT+08:00
CTUDT00628245 | Refresh | Closed | s_namburi | January 16, 2020 2:07:53 PM GMT+08:00 | IQon Refresh: backup of approx. 4 GB takes more than 2 hours | WIP22.1 |  |  | Subsystem |  | NA | Does Not Apply | sanjeeva_a | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed |  | Brilliance 64
CTUDT00628245 | Refresh | Closed | s_namburi | January 16, 2020 2:07:53 PM GMT+08:00 | IQon Refresh: backup of approx. 4 GB takes more than 2 hours | WIP22.1 |  |  | Subsystem |  | NA | Does Not Apply | sanjeeva_a | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed |  | Brilliance Big Bore
CTUDT00628245 | Refresh | Closed | s_namburi | January 16, 2020 2:07:53 PM GMT+08:00 | IQon Refresh: backup of approx. 4 GB takes more than 2 hours | WIP22.1 |  |  | Subsystem |  | NA | Does Not Apply | sanjeeva_a | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed |  | Ingenuity 64
CTUDT00628240 | Refresh | Closed | dongping_w | January 16, 2020 12:37:32 PM GMT+08:00 | Chinese study transfer to ISP failed. | BB12.2 | BB_WIP12.3 |  | Subsystem |  | NA | Fixed | s_sa | 0 | 2 | Console Software | iCAP | NO | Closed | BB_wip12.3 | Brilliance Big Bore |  | BB_WIP12.3 | January 20, 2020 8:42:00 PM GMT+08:00
CTUDT00628235 | Refresh | Closed | lin_s | January 15, 2020 6:19:41 PM GMT+08:00 | Val: Refresh: The system can't copy studies to SSD external device | WIP22.1 | BB_WIP12.3 | 1. Use conventional USB device instead of SSD one 2. Define local folder on the SSD device thru service user/Lan Config with certain name. After that this folder is always usable for clinical user | Subsystem | PM_Approved_BB_WIP12.3 | N/A | Fixed | s_sa | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | BB WIP12.3 | Brilliance Big Bore |  | BB_WIP12.3 | January 20, 2020 8:41:19 PM GMT+08:00
CTUDT00628234 | Refresh | Closed | lin_s | January 15, 2020 5:52:00 PM GMT+08:00 | load patient containing Chinese character from iWATT HIS/RIS and try to perform scan, pop up CIRS error and cannot perform the scan. | WIP12.2 | BB_WIP12.3 |  | Subsystem | PM_Approved_BB_WIP12.3 | CT-NM.SRS.RegStand.686 | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | BB_WIP12.3 | Brilliance Big Bore |  | BB_WIP12.3 | January 20, 2020 8:40:19 PM GMT+08:00
CTUDT00628225 | Refresh | Closed | zhang_lawrence | January 15, 2020 1:30:11 PM GMT+08:00 | Sub items in "XRay Tests" group service tool Visual Diagnostics can not pass | WIP23 | WIP23.1 |  | Subsystem | Visual Diagnostics; Blocker | GCX.SSRS.iCT.13024 | Fixed | zhang_lawrence | 0 | 2 | Gantry Software | GCX | NO | Closed | WIP24SubCom | Ovation |  | WIP23.1 | January 16, 2020 10:07:14 AM GMT+08:00
CTUDT00628222 | Refresh | Closed | shterner_a | January 14, 2020 9:13:30 PM GMT+08:00 | Refresh iCT: Requirement CT-NM.SRS.SysPerf.527 is missing tag to iCT | WIP23 |  |  | Subsystem |  |  | Does Not Apply | shterner_a | 0 |  |  |  |  | Closed |  | Ovation
CTUDT00628215 | Refresh | Closed | lin_s | January 13, 2020 8:40:46 PM GMT+08:00 | WIP22 – the popups: “the selected exam cards cannot be loaded” “there seems to be a problem in the application” and “system error” appeared after choosing “coronary CTA” EC. | WIP22.1 | WIP22.1 | Log out and log in to console, EC can be loaded normally. | System |  | N/A | Fixed | sanjeeva_a | 0 | 3 | Console Software | Console4 | NO | Closed | WIP24SegReq | Ovation |  | WIP22.1 | January 15, 2020 6:15:52 PM GMT+08:00
CTUDT00628215 | Refresh | Closed | lin_s | January 13, 2020 8:40:46 PM GMT+08:00 | WIP22 – the popups: “the selected exam cards cannot be loaded” “there seems to be a problem in the application” and “system error” appeared after choosing “coronary CTA” EC. | WIP22.1 | WIP22.1 | Log out and log in to console, EC can be loaded normally. | System |  | N/A | Fixed | sanjeeva_a | 0 | 3 | Console Software | Console4 | NO | Closed | WIP24SegReq | IQon1 |  | WIP22.1 | January 15, 2020 6:15:52 PM GMT+08:00
CTUDT00628202 | Refresh | Resolved | 320053767 | January 13, 2020 5:09:27 PM GMT+08:00 | Missing series after scanning | WIP22.1 | WIP34 | If data exists, User may perform offline recon with the data that is available. | Subsystem | SP1_PM_Approved; Missing Series | N/A | Fixed | richards_b | 0 | 3 | CIRS_SW_Supervisor | CIRS | YES | Resolved | WIP34SubCom | Ovation |  | WIP34 | November 17, 2020 5:23:01 AM GMT+08:00
CTUDT00628202 | Refresh | Resolved | 320053767 | January 13, 2020 5:09:27 PM GMT+08:00 | Missing series after scanning | WIP22.1 | WIP34 | If data exists, User may perform offline recon with the data that is available. | Subsystem | SP1_PM_Approved; Missing Series | N/A | Fixed | richards_b | 0 | 3 | CIRS_SW_Supervisor | CIRS | YES | Resolved | WIP34SubCom | Brilliance 64 |  | WIP34 | November 17, 2020 5:23:01 AM GMT+08:00
CTUDT00628202 | Refresh | Resolved | 320053767 | January 13, 2020 5:09:27 PM GMT+08:00 | Missing series after scanning | WIP22.1 | WIP34 | If data exists, User may perform offline recon with the data that is available. | Subsystem | SP1_PM_Approved; Missing Series | N/A | Fixed | richards_b | 0 | 3 | CIRS_SW_Supervisor | CIRS | YES | Resolved | WIP34SubCom | Brilliance Big Bore |  | WIP34 | November 17, 2020 5:23:01 AM GMT+08:00
CTUDT00628202 | Refresh | Resolved | 320053767 | January 13, 2020 5:09:27 PM GMT+08:00 | Missing series after scanning | WIP22.1 | WIP34 | If data exists, User may perform offline recon with the data that is available. | Subsystem | SP1_PM_Approved; Missing Series | N/A | Fixed | richards_b | 0 | 3 | CIRS_SW_Supervisor | CIRS | YES | Resolved | WIP34SubCom | Ingenuity 64 |  | WIP34 | November 17, 2020 5:23:01 AM GMT+08:00
CTUDT00628202 | Refresh | Resolved | 320053767 | January 13, 2020 5:09:27 PM GMT+08:00 | Missing series after scanning | WIP22.1 | WIP34 | If data exists, User may perform offline recon with the data that is available. | Subsystem | SP1_PM_Approved; Missing Series | N/A | Fixed | richards_b | 0 | 3 | CIRS_SW_Supervisor | CIRS | YES | Resolved | WIP34SubCom | IQon1 |  | WIP34 | November 17, 2020 5:23:01 AM GMT+08:00
CTUDT00628201 | Refresh | Closed | lin_s | January 13, 2020 2:34:52 PM GMT+08:00 | Insecure BIOS. | WIP21 | DMR356920 Rev07 459801587361 Rev D 459801587351 Rev D 459801587331 Rev E 459801587341 Rev E |  | Subsystem |  | N/A | Fixed | shaham_t | 0 | 2 | Service Documentation | Support-System | NO | Closed | WIP22.2SegReq | Ovation |  | DMR356920 Rev07 459801587361 Rev D 459801587351 Rev D 459801587331 Rev E 459801587341 Rev E | January 15, 2020 12:22:51 AM GMT+08:00
CTUDT00628201 | Refresh | Closed | lin_s | January 13, 2020 2:34:52 PM GMT+08:00 | Insecure BIOS. | WIP21 | DMR356920 Rev07 459801587361 Rev D 459801587351 Rev D 459801587331 Rev E 459801587341 Rev E |  | Subsystem |  | N/A | Fixed | shaham_t | 0 | 2 | Service Documentation | Support-System | NO | Closed | WIP22.2SegReq | IQon1 |  | DMR356920 Rev07 459801587361 Rev D 459801587351 Rev D 459801587331 Rev E 459801587341 Rev E | January 15, 2020 12:22:51 AM GMT+08:00
CTUDT00628200 | Refresh | Closed | lin_s | January 13, 2020 2:33:01 PM GMT+08:00 | IFU inconsistency for refresh products in the cardiac IMR section | IFU Version a | v4.7.7 | N/A | Subsystem |  | N/A | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | IFU | IQon1 |  | v4.7.7 | January 17, 2020 4:25:18 PM GMT+08:00
CTUDT00628199 | Refresh | Closed | lin_s | January 13, 2020 2:22:26 PM GMT+08:00 | At 20mGy CTDIvol, while keeping high contrast resolution improved, IMR Noise reduction smaller than 90% compared to FBP. | WIP19 | WIP23 | N/A | Subsystem |  | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_ImgNodeSvr | CIRS | NO | Closed | WIP23SegReq | Ovation |  | WIP23 | January 15, 2020 6:07:44 PM GMT+08:00
CTUDT00628199 | Refresh | Closed | lin_s | January 13, 2020 2:22:26 PM GMT+08:00 | At 20mGy CTDIvol, while keeping high contrast resolution improved, IMR Noise reduction smaller than 90% compared to FBP. | WIP19 | WIP23 | N/A | Subsystem |  | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_ImgNodeSvr | CIRS | NO | Closed | WIP23SegReq | IQon1 |  | WIP23 | January 15, 2020 6:07:44 PM GMT+08:00
CTUDT00628198 | Refresh | Closed | lin_s | January 13, 2020 2:19:50 PM GMT+08:00 | Accept as four words when typing two Chinese Character | BB_WIP9 | WIP21.1 |  | Subsystem |  | N/A | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP24SubCom | Ovation |  | WIP21.1 | January 28, 2020 2:08:12 PM GMT+08:00
CTUDT00628198 | Refresh | Closed | lin_s | January 13, 2020 2:19:50 PM GMT+08:00 | Accept as four words when typing two Chinese Character | BB_WIP9 | WIP21.1 |  | Subsystem |  | N/A | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP24SubCom | IQon1 |  | WIP21.1 | January 28, 2020 2:08:12 PM GMT+08:00
CTUDT00628197 | Refresh | Closed | lin_s | January 13, 2020 2:18:24 PM GMT+08:00 | Refresh: Network and Computer Name update by olevel cannot take effect | BB_WIP9 | WIP22 |  | System |  | N/A | Cannot Reproduce | Sandeep_Kumar | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP24SegReq | Ovation |  | WIP22 | February 20, 2020 5:58:12 PM GMT+08:00
CTUDT00628197 | Refresh | Closed | lin_s | January 13, 2020 2:18:24 PM GMT+08:00 | Refresh: Network and Computer Name update by olevel cannot take effect | BB_WIP9 | WIP22 |  | System |  | N/A | Cannot Reproduce | Sandeep_Kumar | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP24SegReq | IQon1 |  | WIP22 | February 20, 2020 5:58:12 PM GMT+08:00
CTUDT00628196 | Refresh | Closed | lin_s | January 13, 2020 2:16:39 PM GMT+08:00 | Refresh: Missing DICOM tags as per recommendation of DIN6862-2 Standard | WIP19 | WIP21 | N/A | Subsystem |  | Console-SRS.17972,Console-SRS.2133 | Fixed | sanjeeva_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP23SubCom | Ovation |  | WIP21 | January 15, 2020 8:41:06 PM GMT+08:00
CTUDT00628196 | Refresh | Closed | lin_s | January 13, 2020 2:16:39 PM GMT+08:00 | Refresh: Missing DICOM tags as per recommendation of DIN6862-2 Standard | WIP19 | WIP21 | N/A | Subsystem |  | Console-SRS.17972,Console-SRS.2133 | Fixed | sanjeeva_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP23SubCom | IQon1 |  | WIP21 | January 15, 2020 8:41:06 PM GMT+08:00
CTUDT00628194 | Refresh | Closed | lin_s | January 13, 2020 2:14:48 PM GMT+08:00 | The OffFocal Table name isn't passed in the MDU | WIP18 | WIP19 |  | System |  | N/A | Fixed | tamari_o | 0 | 2 | Console Software | Console4 | NO | Closed | WIP23SegReq | Ovation |  | WIP19 | January 15, 2020 6:03:38 PM GMT+08:00
CTUDT00628194 | Refresh | Closed | lin_s | January 13, 2020 2:14:48 PM GMT+08:00 | The OffFocal Table name isn't passed in the MDU | WIP18 | WIP19 |  | System |  | N/A | Fixed | tamari_o | 0 | 2 | Console Software | Console4 | NO | Closed | WIP23SegReq | IQon1 |  | WIP19 | January 15, 2020 6:03:38 PM GMT+08:00
CTUDT00628193 | Refresh | Closed | lin_s | January 13, 2020 2:12:42 PM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.RegStand.741_Cannot find "labeling and labels symbols shall comply with ISO 15223-1:2016 " in TRG file | WIP16 | WIP20 |  | Subsystem |  | N/A | Fixed | shaham_t | 0 | 2 | User Documentation | Support-System | NO | Closed | PostValidation | Ovation |  | WIP20 | January 15, 2020 7:53:03 PM GMT+08:00
CTUDT00628193 | Refresh | Closed | lin_s | January 13, 2020 2:12:42 PM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.RegStand.741_Cannot find "labeling and labels symbols shall comply with ISO 15223-1:2016 " in TRG file | WIP16 | WIP20 |  | Subsystem |  | N/A | Fixed | shaham_t | 0 | 2 | User Documentation | Support-System | NO | Closed | PostValidation | IQon1 |  | WIP20 | January 15, 2020 7:53:03 PM GMT+08:00
CTUDT00628192 | Refresh | Closed | lin_s | January 13, 2020 2:10:32 PM GMT+08:00 | [Refresh][IFU Kits]CT-NM.SRS.SysSupport.781_ instructions for performing system startup/shutdown of the system for TRG is different with IFU/Maintance file | WIP16 | WIP20 |  | Subsystem |  | N/A | Fixed | shaham_t | 0 | 2 | User Documentation | Support-System | NO | Closed | PostValidation | Ovation |  | WIP20 | January 15, 2020 7:23:26 PM GMT+08:00
CTUDT00628192 | Refresh | Closed | lin_s | January 13, 2020 2:10:32 PM GMT+08:00 | [Refresh][IFU Kits]CT-NM.SRS.SysSupport.781_ instructions for performing system startup/shutdown of the system for TRG is different with IFU/Maintance file | WIP16 | WIP20 |  | Subsystem |  | N/A | Fixed | shaham_t | 0 | 2 | User Documentation | Support-System | NO | Closed | PostValidation | IQon1 |  | WIP20 | January 15, 2020 7:23:26 PM GMT+08:00
CTUDT00628191 | Refresh | Closed | lin_s | January 13, 2020 2:07:03 PM GMT+08:00 | IOCC_Refresh: During Modality Worklist query, received response shows invalid characters for patient name in console UI for ISO-IR 13, ISO-IR 87 | WIP12 | DCS |  | Subsystem |  | N/A | Fixed | sanjeeva_a | 0 | 2 | Verification Documentation | iCAP | NO | Closed | WIP23SegReq | Ovation |  | DCS | January 15, 2020 5:57:23 PM GMT+08:00
CTUDT00628191 | Refresh | Closed | lin_s | January 13, 2020 2:07:03 PM GMT+08:00 | IOCC_Refresh: During Modality Worklist query, received response shows invalid characters for patient name in console UI for ISO-IR 13, ISO-IR 87 | WIP12 | DCS |  | Subsystem |  | N/A | Fixed | sanjeeva_a | 0 | 2 | Verification Documentation | iCAP | NO | Closed | WIP23SegReq | IQon1 |  | DCS | January 15, 2020 5:57:23 PM GMT+08:00
CTUDT00628190 | Refresh | Closed | lin_s | January 13, 2020 2:05:23 PM GMT+08:00 | Refresh: Cannot fill Chinese characters in some user controls when IM is Chinese Language | WIP11.1 | See associated AR | 1. Select the Chinese keyboard from preference => regional settings. 2. Perform logoff & login two times. 3. Select the Chinese keyboard by using WIN+SPACE. 4. Press SHIFT once before typing anything in Chinese. This is needed because, Microsoft Pinyin Input profile has two input languages(Chinese and English). To toggle we need use SHIFT key. | Subsystem |  | N/A | Fixed | charles_j | 0 | 2 | Console USP_EBW | Support-System | NO | Closed | WIP23SegReq | Ovation |  | See associated AR | January 15, 2020 5:45:43 PM GMT+08:00
CTUDT00628190 | Refresh | Closed | lin_s | January 13, 2020 2:05:23 PM GMT+08:00 | Refresh: Cannot fill Chinese characters in some user controls when IM is Chinese Language | WIP11.1 | See associated AR | 1. Select the Chinese keyboard from preference => regional settings. 2. Perform logoff & login two times. 3. Select the Chinese keyboard by using WIN+SPACE. 4. Press SHIFT once before typing anything in Chinese. This is needed because, Microsoft Pinyin Input profile has two input languages(Chinese and English). To toggle we need use SHIFT key. | Subsystem |  | N/A | Fixed | charles_j | 0 | 2 | Console USP_EBW | Support-System | NO | Closed | WIP23SegReq | IQon1 |  | See associated AR | January 15, 2020 5:45:43 PM GMT+08:00
CTUDT00628189 | Refresh | Closed | lin_s | January 13, 2020 2:02:46 PM GMT+08:00 | Annotation Values are not shown and application crashes on loading CT RT study to Quick Review/CTV application. | WIP11 | WIP23 |  | Subsystem |  | N/A | Fixed | aswani_v | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP23SubCom | Ovation |  | WIP23 | January 15, 2020 5:40:43 PM GMT+08:00
CTUDT00628189 | Refresh | Closed | lin_s | January 13, 2020 2:02:46 PM GMT+08:00 | Annotation Values are not shown and application crashes on loading CT RT study to Quick Review/CTV application. | WIP11 | WIP23 |  | Subsystem |  | N/A | Fixed | aswani_v | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP23SubCom | IQon1 |  | WIP23 | January 15, 2020 5:40:43 PM GMT+08:00
CTUDT00628188 | Refresh | Closed | lin_s | January 13, 2020 1:59:56 PM GMT+08:00 | FoK: UCL: Exam Application crash during the acquisition | V4.7.7 and V4.1.10 | WIP12 |  | System |  | N/A | Fixed | tamari_o | 0 | 3 | Console Software | Console4 | YES | Closed | WIP24SegReq | Ovation | 9413694 | WIP12 | January 15, 2020 5:33:36 PM GMT+08:00
CTUDT00628188 | Refresh | Closed | lin_s | January 13, 2020 1:59:56 PM GMT+08:00 | FoK: UCL: Exam Application crash during the acquisition | V4.7.7 and V4.1.10 | WIP12 |  | System |  | N/A | Fixed | tamari_o | 0 | 3 | Console Software | Console4 | YES | Closed | WIP24SegReq | IQon1 | 9413694 | WIP12 | January 15, 2020 5:33:36 PM GMT+08:00
CTUDT00628187 | Refresh | Closed | lin_s | January 13, 2020 1:57:02 PM GMT+08:00 | Refresh: When keyboard language is changed to "Chinese" in preferences and Chinese keyboard is used, Data cannot be type in Chinese instead it is taking the characters in English itself. | WIP8 | WIP12 |  | Subsystem |  | N/A | Fixed | shaham_t | 0 | 2 | USP | iCAP | NO | Closed | WIP9SegReq | Ovation |  | WIP12 | January 15, 2020 1:40:08 AM GMT+08:00
CTUDT00628187 | Refresh | Closed | lin_s | January 13, 2020 1:57:02 PM GMT+08:00 | Refresh: When keyboard language is changed to "Chinese" in preferences and Chinese keyboard is used, Data cannot be type in Chinese instead it is taking the characters in English itself. | WIP8 | WIP12 |  | Subsystem |  | N/A | Fixed | shaham_t | 0 | 2 | USP | iCAP | NO | Closed | WIP9SegReq | IQon1 |  | WIP12 | January 15, 2020 1:40:08 AM GMT+08:00
CTUDT00628186 | Refresh | Closed | lin_s | January 13, 2020 1:55:10 PM GMT+08:00 | Crescent artifact in pediatric brain with Ingenuity v4.1.4 FOK | v4.1.4 FOK | WIP17 | Use of High Resolution for brain imaging | System |  | N/A | Fixed | shaham_t | 0 | 2 | CIRS_SW_Installation | Support-Physics | YES | Closed | WIP17SubCom | Ovation | CFR 9367697 PR 4918461 | WIP17 | January 15, 2020 1:33:23 AM GMT+08:00
CTUDT00628185 | Refresh | Closed | lin_s | January 13, 2020 1:53:09 PM GMT+08:00 | Pediatric Head image quality exhibits halo artifact at the bone brain interface | 4.1.4.xx001 | WIP17 | Use of High Resolution for brain imaging | System |  | N/A | Fixed | shaham_t | 0 | 2 | CIRS_SW_Installation | Support-Physics | YES | Closed | WIP17SubCom | Ovation | PR3934861 PR3960567 | WIP17 | January 15, 2020 1:13:25 AM GMT+08:00
CTUDT00628184 | Refresh | Closed | lin_s | January 13, 2020 1:49:14 PM GMT+08:00 | IOCC: CT system proposes following SOP Classes which are not claimed in DCS. | WIP507 | DCS for IQon Refresh |  | Subsystem |  | N/A | Fixed | shaham_t | 0 | 4 | Design Documentation | iCAP | NO | Closed | WIP10SegReq | Ovation |  | DCS for IQon Refresh | January 15, 2020 1:02:04 AM GMT+08:00
CTUDT00628184 | Refresh | Closed | lin_s | January 13, 2020 1:49:14 PM GMT+08:00 | IOCC: CT system proposes following SOP Classes which are not claimed in DCS. | WIP507 | DCS for IQon Refresh |  | Subsystem |  | N/A | Fixed | shaham_t | 0 | 4 | Design Documentation | iCAP | NO | Closed | WIP10SegReq | IQon1 |  | DCS for IQon Refresh | January 15, 2020 1:02:04 AM GMT+08:00
CTUDT00628183 | Refresh | Closed | lin_s | January 13, 2020 1:47:12 PM GMT+08:00 | IOCC: CT system accepts following SOP Classes which are not claimed in DCS. | WIP507 | DCS for IQon Refresh |  | Subsystem |  | N/A | Fixed | shaham_t | 0 | 4 | Design Documentation | iCAP | NO | Closed | WIP10SegReq | Ovation |  | DCS for IQon Refresh | January 15, 2020 12:13:03 AM GMT+08:00
CTUDT00628183 | Refresh | Closed | lin_s | January 13, 2020 1:47:12 PM GMT+08:00 | IOCC: CT system accepts following SOP Classes which are not claimed in DCS. | WIP507 | DCS for IQon Refresh |  | Subsystem |  | N/A | Fixed | shaham_t | 0 | 4 | Design Documentation | iCAP | NO | Closed | WIP10SegReq | IQon1 |  | DCS for IQon Refresh | January 15, 2020 12:13:03 AM GMT+08:00
CTUDT00628182 | Refresh | Closed | lin_s | January 13, 2020 1:43:44 PM GMT+08:00 | Refresh- Sleep mode causing disconnection of Host and Gantry | WIP7.1 | WIP10 | Manually configure Console to never sleep on power plugged-in. | System |  | N/A | Fixed | shaham_t | 0 | 3 | Console Software | Support-Security | NO | Closed | WIP9 | Ovation |  | WIP10 | January 14, 2020 11:52:11 PM GMT+08:00
CTUDT00628182 | Refresh | Closed | lin_s | January 13, 2020 1:43:44 PM GMT+08:00 | Refresh- Sleep mode causing disconnection of Host and Gantry | WIP7.1 | WIP10 | Manually configure Console to never sleep on power plugged-in. | System |  | N/A | Fixed | shaham_t | 0 | 3 | Console Software | Support-Security | NO | Closed | WIP9 | IQon1 |  | WIP10 | January 14, 2020 11:52:11 PM GMT+08:00
CTUDT00628176 | Refresh | Closed | zhong_x | January 11, 2020 1:52:00 PM GMT+08:00 | Varian waveform cannot be exported from the Varian to the CT after one minutes. | WIP12.2 | BB_WIP12.3 | No workaround but if the Export of wave is done within 1 minute of scan end this issue will not happen. | Subsystem | PM_Approved_BB_WIP12.3 | NA | Fixed | ravindra_a | 0 | 2 | Console Software | Console4 | YES | Closed | BB WIP12.3 | Brilliance Big Bore |  | BB_WIP12.3 | January 20, 2020 9:24:22 PM GMT+08:00
CTUDT00628175 | Refresh | Closed | lin_s | January 11, 2020 1:09:08 PM GMT+08:00 | Refresh: IP subnet still display 192.168. but we select 10.168 during one click installation | WIP22 | DMR356920 Rev08 | ComputerAndNetworkConfiguration tool can be used post installation to update the CIRS IP Address value. User can select all the options that can be updated during one click installation with this tool and update it. | Subsystem |  | N/A | Fixed | shaham_t | 0 | 3 | Service Documentation | Support-System | NO | Closed | SubCom 23-Jan-2020 | Ovation |  | DMR356920 Rev08 | January 27, 2020 2:15:05 AM GMT+08:00
CTUDT00628175 | Refresh | Closed | lin_s | January 11, 2020 1:09:08 PM GMT+08:00 | Refresh: IP subnet still display 192.168. but we select 10.168 during one click installation | WIP22 | DMR356920 Rev08 | ComputerAndNetworkConfiguration tool can be used post installation to update the CIRS IP Address value. User can select all the options that can be updated during one click installation with this tool and update it. | Subsystem |  | N/A | Fixed | shaham_t | 0 | 3 | Service Documentation | Support-System | NO | Closed | SubCom 23-Jan-2020 | IQon1 |  | DMR356920 Rev08 | January 27, 2020 2:15:05 AM GMT+08:00
CTUDT00628168 | Refresh | Closed | lei_yang | January 10, 2020 2:50:25 PM GMT+08:00 | Warning message ‘Exam Application has stopped working’ Popped up and Workflow exited when exporting the waveform from the Varian viewer. | WIP22.2 | WIP22.3 | No workaround | Subsystem | PM_Approved_WIP22.3 | NA | Fixed | bhawna_g | 0 | 3 | Console Software | Console4 | YES | Closed | WIP22.3 | Brilliance 64 |  | WIP22.3 | January 16, 2020 9:39:08 PM GMT+08:00
CTUDT00628168 | Refresh | Closed | lei_yang | January 10, 2020 2:50:25 PM GMT+08:00 | Warning message ‘Exam Application has stopped working’ Popped up and Workflow exited when exporting the waveform from the Varian viewer. | WIP22.2 | WIP22.3 | No workaround | Subsystem | PM_Approved_WIP22.3 | NA | Fixed | bhawna_g | 0 | 3 | Console Software | Console4 | YES | Closed | WIP22.3 | Ingenuity 64 |  | WIP22.3 | January 16, 2020 9:39:08 PM GMT+08:00
CTUDT00628166 | Refresh | Closed | glauberman_o | January 10, 2020 2:53:26 AM GMT+08:00 | Val: Refresh: Not all images loaded after editing demographic information of the patient | WIP22.1 |  | Unselect Radiation Dose Information Series from the study and then load it to CTV. | Subsystem | CT Viewer, Blocker | NA | No Plan to Fix | aswani_v | 0 | 4 |  | iCAP | NO | Closed |  | Ovation
CTUDT00628166 | Refresh | Closed | glauberman_o | January 10, 2020 2:53:26 AM GMT+08:00 | Val: Refresh: Not all images loaded after editing demographic information of the patient | WIP22.1 |  | Unselect Radiation Dose Information Series from the study and then load it to CTV. | Subsystem | CT Viewer, Blocker | NA | No Plan to Fix | aswani_v | 0 | 4 |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00628166 | Refresh | Closed | glauberman_o | January 10, 2020 2:53:26 AM GMT+08:00 | Val: Refresh: Not all images loaded after editing demographic information of the patient | WIP22.1 |  | Unselect Radiation Dose Information Series from the study and then load it to CTV. | Subsystem | CT Viewer, Blocker | NA | No Plan to Fix | aswani_v | 0 | 4 |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00628166 | Refresh | Closed | glauberman_o | January 10, 2020 2:53:26 AM GMT+08:00 | Val: Refresh: Not all images loaded after editing demographic information of the patient | WIP22.1 |  | Unselect Radiation Dose Information Series from the study and then load it to CTV. | Subsystem | CT Viewer, Blocker | NA | No Plan to Fix | aswani_v | 0 | 4 |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00628166 | Refresh | Closed | glauberman_o | January 10, 2020 2:53:26 AM GMT+08:00 | Val: Refresh: Not all images loaded after editing demographic information of the patient | WIP22.1 |  | Unselect Radiation Dose Information Series from the study and then load it to CTV. | Subsystem | CT Viewer, Blocker | NA | No Plan to Fix | aswani_v | 0 | 4 |  | iCAP | NO | Closed |  | IQon1
CTUDT00628157 | Refresh | Closed | shterner_a | January 9, 2020 4:57:17 PM GMT+08:00 | Val: Refresh: Test Case Modification for Cycle One | WIP22.1 | DHF305293 Rev.03 |  | Subsystem | Test Design, Blocker | NA | Fixed | shterner_a | 0 | 2 | Validation Documentation | Support-System | NO | Closed | SubCom-30-Jan-2020 | IQon1 |  | DHF305293 Rev.03 | February 3, 2020 5:04:35 PM GMT+08:00
CTUDT00628152 | Refresh | Closed | lin_s | January 9, 2020 1:38:47 PM GMT+08:00 | Refresh - Ingenuity: one File upgrade failed because CalibrationTables_AirDataVarianceTable_CalibrationTables_AirDataVarianceTable_AirDataVarianceTable not found | WIP22.2 |  | Tube output calibration need to be run again. | Subsystem |  | N/A | Works As Specified / Not A Defect | zhang_l | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | Service Manual | Ingenuity 64
CTUDT00628149 | Refresh | Closed | lin_s | January 9, 2020 1:12:53 PM GMT+08:00 | Refresh Upgrade - Gantry calibration records are not upgraded from 4.1.7 to 4.1.10 | WIP21 | WIP22.2 |  | Subsystem |  | N/A | Fixed | lei_yang | 0 | 2 | Service Tools_Software | Console4 | NO | Closed | Service Manual Update | Ingenuity 64 |  | WIP22.2 | January 9, 2020 1:32:11 PM GMT+08:00
CTUDT00628148 | Refresh | Resolved | lin_s | January 9, 2020 12:00:35 PM GMT+08:00 | Refresh - Ingenuity: one File upgrade failed because CalibrationTables_AirDataVarianceTable_CalibrationTables_AirDataVarianceTable_AirDataVarianceTable not found | WIP22.2 | 459801587391 Rev C WIP30 | Tube output calibration need to be run again. | System | Service manual Update; SP1_PM_Approved | NA | Fixed | chakraborty_d | 0 | 2 | Service Tools_Software | Service Tools | NO | Resolved | WIP31SubCom | Brilliance 64 |  | 459801587391 Rev C WIP30 | July 7, 2020 5:00:27 PM GMT+08:00
CTUDT00628148 | Refresh | Resolved | lin_s | January 9, 2020 12:00:35 PM GMT+08:00 | Refresh - Ingenuity: one File upgrade failed because CalibrationTables_AirDataVarianceTable_CalibrationTables_AirDataVarianceTable_AirDataVarianceTable not found | WIP22.2 | 459801587391 Rev C WIP30 | Tube output calibration need to be run again. | System | Service manual Update; SP1_PM_Approved | NA | Fixed | chakraborty_d | 0 | 2 | Service Tools_Software | Service Tools | NO | Resolved | WIP31SubCom | Ingenuity 64 |  | 459801587391 Rev C WIP30 | July 7, 2020 5:00:27 PM GMT+08:00
CTUDT00628143 | Refresh | Closed | glauberman_o | January 8, 2020 9:10:49 PM GMT+08:00 | Val: Refresh: The system can't copy studies to SSD external device | WIP22.1 | WIP23.1 WIP22.3 | 1. Use conventional USB device instead of SDD one 2. Define local folder on the SDD device thru service user/Lan Config with certain name. After that this folder is always usable for clinical user | Subsystem | PM_Approved_WIP22.3 | NA | Fixed | s_panigrahi | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP24SegReq | Ovation |  | WIP23.1 WIP22.3 | January 15, 2020 3:50:42 PM GMT+08:00
CTUDT00628143 | Refresh | Closed | glauberman_o | January 8, 2020 9:10:49 PM GMT+08:00 | Val: Refresh: The system can't copy studies to SSD external device | WIP22.1 | WIP23.1 WIP22.3 | 1. Use conventional USB device instead of SDD one 2. Define local folder on the SDD device thru service user/Lan Config with certain name. After that this folder is always usable for clinical user | Subsystem | PM_Approved_WIP22.3 | NA | Fixed | s_panigrahi | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP24SegReq | Brilliance 64 |  | WIP23.1 WIP22.3 | January 15, 2020 3:50:42 PM GMT+08:00
CTUDT00628143 | Refresh | Closed | glauberman_o | January 8, 2020 9:10:49 PM GMT+08:00 | Val: Refresh: The system can't copy studies to SSD external device | WIP22.1 | WIP23.1 WIP22.3 | 1. Use conventional USB device instead of SDD one 2. Define local folder on the SDD device thru service user/Lan Config with certain name. After that this folder is always usable for clinical user | Subsystem | PM_Approved_WIP22.3 | NA | Fixed | s_panigrahi | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP24SegReq | Ingenuity 64 |  | WIP23.1 WIP22.3 | January 15, 2020 3:50:42 PM GMT+08:00
CTUDT00628143 | Refresh | Closed | glauberman_o | January 8, 2020 9:10:49 PM GMT+08:00 | Val: Refresh: The system can't copy studies to SSD external device | WIP22.1 | WIP23.1 WIP22.3 | 1. Use conventional USB device instead of SDD one 2. Define local folder on the SDD device thru service user/Lan Config with certain name. After that this folder is always usable for clinical user | Subsystem | PM_Approved_WIP22.3 | NA | Fixed | s_panigrahi | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP24SegReq | IQon1 |  | WIP23.1 WIP22.3 | January 15, 2020 3:50:42 PM GMT+08:00
CTUDT00628141 | Refresh | Closed | hilas | January 8, 2020 8:26:28 PM GMT+08:00 | SRS requirements CT-NM.SRS.ProductConfig.177 and CT-NM.SRS.ProductConfig.184 should be removed from IQon product | WIP23 | DHF298556 Rev.11 |  | Subsystem | Requirement Update | ProductConfig.177, ProductConfig.184 | Fixed | shaham_t | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | SubCom-15-Feb-2020 | IQon1 |  | DHF298556 Rev.11 | February 26, 2020 4:12:24 PM GMT+08:00
CTUDT00628139 | Refresh | Closed | glauberman_o | January 8, 2020 7:20:00 PM GMT+08:00 | Val: Refresh: Patient Directory has stopped working | WIP22.1 |  |  | Subsystem | Patient Directory; Blocker | NA | Works As Specified / Not A Defect | srushti_k | 0 |  |  | iCAP | NO | Closed |  | Ovation
CTUDT00628139 | Refresh | Closed | glauberman_o | January 8, 2020 7:20:00 PM GMT+08:00 | Val: Refresh: Patient Directory has stopped working | WIP22.1 |  |  | Subsystem | Patient Directory; Blocker | NA | Works As Specified / Not A Defect | srushti_k | 0 |  |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00628139 | Refresh | Closed | glauberman_o | January 8, 2020 7:20:00 PM GMT+08:00 | Val: Refresh: Patient Directory has stopped working | WIP22.1 |  |  | Subsystem | Patient Directory; Blocker | NA | Works As Specified / Not A Defect | srushti_k | 0 |  |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00628139 | Refresh | Closed | glauberman_o | January 8, 2020 7:20:00 PM GMT+08:00 | Val: Refresh: Patient Directory has stopped working | WIP22.1 |  |  | Subsystem | Patient Directory; Blocker | NA | Works As Specified / Not A Defect | srushti_k | 0 |  |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00628139 | Refresh | Closed | glauberman_o | January 8, 2020 7:20:00 PM GMT+08:00 | Val: Refresh: Patient Directory has stopped working | WIP22.1 |  |  | Subsystem | Patient Directory; Blocker | NA | Works As Specified / Not A Defect | srushti_k | 0 |  |  | iCAP | NO | Closed |  | IQon1
CTUDT00628127 | Refresh | Closed | Bat-El_Asaraf | January 7, 2020 10:16:31 PM GMT+08:00 | Val: Refresh: Liver DRI area missing from images | WIP22.1 |  | User can right click on the surview and select Context menu option "Define Liver Area" to show the lines on surview. If "Use Previous Surview" used then re do Surview and then use context menu option. | Subsystem | Liver DRI; Blocker;CRN | N/A | No Plan to Fix | bhawna_g | 0 | 3 |  | Console4 | YES | Closed |  | Ovation
CTUDT00628127 | Refresh | Closed | Bat-El_Asaraf | January 7, 2020 10:16:31 PM GMT+08:00 | Val: Refresh: Liver DRI area missing from images | WIP22.1 |  | User can right click on the surview and select Context menu option "Define Liver Area" to show the lines on surview. If "Use Previous Surview" used then re do Surview and then use context menu option. | Subsystem | Liver DRI; Blocker;CRN | N/A | No Plan to Fix | bhawna_g | 0 | 3 |  | Console4 | YES | Closed |  | Brilliance 64
CTUDT00628127 | Refresh | Closed | Bat-El_Asaraf | January 7, 2020 10:16:31 PM GMT+08:00 | Val: Refresh: Liver DRI area missing from images | WIP22.1 |  | User can right click on the surview and select Context menu option "Define Liver Area" to show the lines on surview. If "Use Previous Surview" used then re do Surview and then use context menu option. | Subsystem | Liver DRI; Blocker;CRN | N/A | No Plan to Fix | bhawna_g | 0 | 3 |  | Console4 | YES | Closed |  | Brilliance Big Bore
CTUDT00628127 | Refresh | Closed | Bat-El_Asaraf | January 7, 2020 10:16:31 PM GMT+08:00 | Val: Refresh: Liver DRI area missing from images | WIP22.1 |  | User can right click on the surview and select Context menu option "Define Liver Area" to show the lines on surview. If "Use Previous Surview" used then re do Surview and then use context menu option. | Subsystem | Liver DRI; Blocker;CRN | N/A | No Plan to Fix | bhawna_g | 0 | 3 |  | Console4 | YES | Closed |  | Ingenuity 64
CTUDT00628127 | Refresh | Closed | Bat-El_Asaraf | January 7, 2020 10:16:31 PM GMT+08:00 | Val: Refresh: Liver DRI area missing from images | WIP22.1 |  | User can right click on the surview and select Context menu option "Define Liver Area" to show the lines on surview. If "Use Previous Surview" used then re do Surview and then use context menu option. | Subsystem | Liver DRI; Blocker;CRN | N/A | No Plan to Fix | bhawna_g | 0 | 3 |  | Console4 | YES | Closed |  | IQon1
CTUDT00628118 | Refresh | Closed | glauberman_o | January 7, 2020 9:32:58 PM GMT+08:00 | Val: Refresh: Critical error occurred when trying to copy images to PACS | WIP22.1 | BB_WIP12.4 WIP22.4 | Relaunch the application. | Subsystem | Remote Device; Blocker | NA | Fixed | sagar_v | 0 | 4 | Console Software | Console4 | NO | Closed | WIP12.4&WIP22.4 | Brilliance 64 |  | BB_WIP12.4 WIP22.4 | February 7, 2020 2:30:49 PM GMT+08:00
CTUDT00628118 | Refresh | Closed | glauberman_o | January 7, 2020 9:32:58 PM GMT+08:00 | Val: Refresh: Critical error occurred when trying to copy images to PACS | WIP22.1 | BB_WIP12.4 WIP22.4 | Relaunch the application. | Subsystem | Remote Device; Blocker | NA | Fixed | sagar_v | 0 | 4 | Console Software | Console4 | NO | Closed | WIP12.4&WIP22.4 | Brilliance Big Bore |  | BB_WIP12.4 WIP22.4 | February 7, 2020 2:30:49 PM GMT+08:00
CTUDT00628118 | Refresh | Closed | glauberman_o | January 7, 2020 9:32:58 PM GMT+08:00 | Val: Refresh: Critical error occurred when trying to copy images to PACS | WIP22.1 | BB_WIP12.4 WIP22.4 | Relaunch the application. | Subsystem | Remote Device; Blocker | NA | Fixed | sagar_v | 0 | 4 | Console Software | Console4 | NO | Closed | WIP12.4&WIP22.4 | Ingenuity 64 |  | BB_WIP12.4 WIP22.4 | February 7, 2020 2:30:49 PM GMT+08:00
CTUDT00628112 | Refresh | Closed | wei_f | January 7, 2020 10:12:10 AM GMT+08:00 | In plan page, enter Chinese characters in Label, show as "???" | WIP12.2 |  |  | Subsystem |  |  | Duplicate | navale_s | 0 |  |  |  |  | Closed |  | Brilliance Big Bore
CTUDT00628112 | Refresh | Closed | wei_f | January 7, 2020 10:12:10 AM GMT+08:00 | In plan page, enter Chinese characters in Label, show as "???" | WIP12.2 |  |  | Subsystem |  |  | Duplicate | navale_s | 0 |  |  |  |  | Closed |  | Ingenuity 64
CTUDT00628110 | Refresh | Closed | wei_f | January 7, 2020 8:55:38 AM GMT+08:00 | Set Preference setting as default with expection check in“ include_mpps_related_attributes_in_the_IOD” from connectivity, load patient containing Chinese character from iWATT HIS/RIS and try to perform scan, “Recon box error “occurred. | WIP12 |  |  | Subsystem |  |  | Duplicate | s_sa | 0 |  |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00628097 | Refresh | Closed | yatziv_l | January 5, 2020 10:57:32 PM GMT+08:00 | Val: Refresh: Wrong instructions for Phantom section in Constancy test | 4.8 & 4.1.10 | WIP31 |  | Subsystem | Constancy Test; SP1_PM_Approved | N\A | Fixed | frider_o | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP31SubCom | Brilliance Big Bore |  | WIP31 | July 8, 2020 4:58:35 PM GMT+08:00
CTUDT00628097 | Refresh | Closed | yatziv_l | January 5, 2020 10:57:32 PM GMT+08:00 | Val: Refresh: Wrong instructions for Phantom section in Constancy test | 4.8 & 4.1.10 | WIP31 |  | Subsystem | Constancy Test; SP1_PM_Approved | N\A | Fixed | frider_o | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP31SubCom | Ingenuity 64 |  | WIP31 | July 8, 2020 4:58:35 PM GMT+08:00
CTUDT00628093 | Refresh | Closed | yatziv_l | January 5, 2020 7:36:09 PM GMT+08:00 | Refresh: iCT and IQon: DMS Intermittent control link frame errors causing header wrong identification | 4.7.5, 4.1.10 | WIP23.1 |  | Subsystem |  | N/A | Fixed | elkarif_y | 0 | 4 | DMS Electrical | DMS | NO | Closed | WIP23.1SubCom | Ovation |  | WIP23.1 | January 9, 2020 5:21:12 PM GMT+08:00
CTUDT00628093 | Refresh | Closed | yatziv_l | January 5, 2020 7:36:09 PM GMT+08:00 | Refresh: iCT and IQon: DMS Intermittent control link frame errors causing header wrong identification | 4.7.5, 4.1.10 | WIP23.1 |  | Subsystem |  | N/A | Fixed | elkarif_y | 0 | 4 | DMS Electrical | DMS | NO | Closed | WIP23.1SubCom | IQon1 |  | WIP23.1 | January 9, 2020 5:21:12 PM GMT+08:00
CTUDT00628086 | Refresh | Closed | wei_f | January 3, 2020 11:59:22 AM GMT+08:00 | UI display issue of fetched DICOM info from IWATT HIS/RIS for full_width numbers | BB_WIP12.2 |  | Use halfwidth Chinese characters. | Subsystem | CRN | N/A | No Plan to Fix | navale_s | 0 | 4 |  | Console4 | NO | Closed | SubNC | Ovation
CTUDT00628086 | Refresh | Closed | wei_f | January 3, 2020 11:59:22 AM GMT+08:00 | UI display issue of fetched DICOM info from IWATT HIS/RIS for full_width numbers | BB_WIP12.2 |  | Use halfwidth Chinese characters. | Subsystem | CRN | N/A | No Plan to Fix | navale_s | 0 | 4 |  | Console4 | NO | Closed | SubNC | Brilliance 64
CTUDT00628086 | Refresh | Closed | wei_f | January 3, 2020 11:59:22 AM GMT+08:00 | UI display issue of fetched DICOM info from IWATT HIS/RIS for full_width numbers | BB_WIP12.2 |  | Use halfwidth Chinese characters. | Subsystem | CRN | N/A | No Plan to Fix | navale_s | 0 | 4 |  | Console4 | NO | Closed | SubNC | Brilliance Big Bore
CTUDT00628086 | Refresh | Closed | wei_f | January 3, 2020 11:59:22 AM GMT+08:00 | UI display issue of fetched DICOM info from IWATT HIS/RIS for full_width numbers | BB_WIP12.2 |  | Use halfwidth Chinese characters. | Subsystem | CRN | N/A | No Plan to Fix | navale_s | 0 | 4 |  | Console4 | NO | Closed | SubNC | Ingenuity 64
CTUDT00628086 | Refresh | Closed | wei_f | January 3, 2020 11:59:22 AM GMT+08:00 | UI display issue of fetched DICOM info from IWATT HIS/RIS for full_width numbers | BB_WIP12.2 |  | Use halfwidth Chinese characters. | Subsystem | CRN | N/A | No Plan to Fix | navale_s | 0 | 4 |  | Console4 | NO | Closed | SubNC | IQon1
CTUDT00628070 | Refresh | Closed | shterner_a | January 2, 2020 7:28:20 PM GMT+08:00 | IQon Refresh: backup of approx. 4 GB takes more than 2 hours | WIP22.1 | WIP23.1 |  | System | Val Backup; Blocker | NA | Fixed | tamari_o | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP24SegReq | Ovation |  | WIP23.1 | January 16, 2020 12:55:54 PM GMT+08:00
CTUDT00628070 | Refresh | Closed | shterner_a | January 2, 2020 7:28:20 PM GMT+08:00 | IQon Refresh: backup of approx. 4 GB takes more than 2 hours | WIP22.1 | WIP23.1 |  | System | Val Backup; Blocker | NA | Fixed | tamari_o | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP24SegReq | IQon1 |  | WIP23.1 | January 16, 2020 12:55:54 PM GMT+08:00
CTUDT00628055 | Refresh | Closed | 320053769 | January 2, 2020 3:25:08 PM GMT+08:00 | WIP23 - licensing tab in IQon, iCT and iCTSP doesn’t show all licensing options. | WIP23 | WIP23.1 | NA | Subsystem |  | NA | Fixed | s_panigrahi | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP23.1SegReq | Ovation |  | WIP23.1 | January 7, 2020 7:38:03 PM GMT+08:00
CTUDT00628055 | Refresh | Closed | 320053769 | January 2, 2020 3:25:08 PM GMT+08:00 | WIP23 - licensing tab in IQon, iCT and iCTSP doesn’t show all licensing options. | WIP23 | WIP23.1 | NA | Subsystem |  | NA | Fixed | s_panigrahi | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP23.1SegReq | IQon1 |  | WIP23.1 | January 7, 2020 7:38:03 PM GMT+08:00
CTUDT00628035 | Refresh | Closed | Bat-El_Asaraf | December 31, 2019 9:11:31 PM GMT+08:00 | Val: Refresh: Wrong instructions for Phantom section in Constancy test | 4.7.7.43057 | wip24 |  | Subsystem | Constancy Test; Blocker; PM_Approved | N/A | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP24SubCom | Ovation |  | wip24 | January 26, 2020 8:33:52 PM GMT+08:00
CTUDT00628035 | Refresh | Closed | Bat-El_Asaraf | December 31, 2019 9:11:31 PM GMT+08:00 | Val: Refresh: Wrong instructions for Phantom section in Constancy test | 4.7.7.43057 | wip24 |  | Subsystem | Constancy Test; Blocker; PM_Approved | N/A | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP24SubCom | IQon1 |  | wip24 | January 26, 2020 8:33:52 PM GMT+08:00
CTUDT00628000 | Refresh | Closed | gushansky_m | December 30, 2019 5:11:21 PM GMT+08:00 | Refresh : Scan Based Profile failed in Visual Diagnostic | WIP23 | WIP23.1 |  | System | Blocker | NA | Fixed | chakraborty_d | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP23.1SubCom | Ovation |  | WIP23.1 | February 26, 2020 1:15:10 PM GMT+08:00
CTUDT00627976 | Refresh | Closed | gushansky_m | December 24, 2019 7:52:04 PM GMT+08:00 | Refresh : Image Rotation Calibration failed | WIP22.1 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | altshuller_mark | 0 |  |  | Support-System | NO | Closed |  | IQon1
CTUDT00627958 | Refresh | Closed | gushansky_m | December 23, 2019 9:29:51 PM GMT+08:00 | Refresh:Bugrep Silent command doesn't work for user remoteservice vs SSH connection | WIP22.1 |  |  | Subsystem |  | N/A | Works As Specified / Not A Defect | s_shabeer | 0 |  |  | Service Tools | NO | Closed | Not Defect | Ovation
CTUDT00627958 | Refresh | Closed | gushansky_m | December 23, 2019 9:29:51 PM GMT+08:00 | Refresh:Bugrep Silent command doesn't work for user remoteservice vs SSH connection | WIP22.1 |  |  | Subsystem |  | N/A | Works As Specified / Not A Defect | s_shabeer | 0 |  |  | Service Tools | NO | Closed | Not Defect | Brilliance 64
CTUDT00627958 | Refresh | Closed | gushansky_m | December 23, 2019 9:29:51 PM GMT+08:00 | Refresh:Bugrep Silent command doesn't work for user remoteservice vs SSH connection | WIP22.1 |  |  | Subsystem |  | N/A | Works As Specified / Not A Defect | s_shabeer | 0 |  |  | Service Tools | NO | Closed | Not Defect | Ingenuity 64
CTUDT00627958 | Refresh | Closed | gushansky_m | December 23, 2019 9:29:51 PM GMT+08:00 | Refresh:Bugrep Silent command doesn't work for user remoteservice vs SSH connection | WIP22.1 |  |  | Subsystem |  | N/A | Works As Specified / Not A Defect | s_shabeer | 0 |  |  | Service Tools | NO | Closed | Not Defect | IQon1
CTUDT00627952 | Refresh | Closed | wang_l | December 23, 2019 6:43:03 PM GMT+08:00 | Reliability BB - Patient table cannot comply during moving couch for CCT scan | BB_WIP12 |  |  | Subsystem |  |  | Does Not Apply | cx_zhang | 0 |  |  | GCX | NO | Closed |  | Brilliance Big Bore
CTUDT00627950 | Refresh | Closed | wang_l | December 23, 2019 2:23:58 PM GMT+08:00 | Reliability - Ingenuity: system stucked in Initializing.. for more than 10 mins during clinical scan | WIP22.2 |  | User can cancel the scan and plan the same using Previous surview. | System |  | NA | Works As Specified / Not A Defect | kr_m | 0 |  |  | Console4 | NO | Closed | SubNC | Ingenuity 64
CTUDT00627948 | Refresh | Closed | l_yiming | December 23, 2019 10:08:46 AM GMT+08:00 | Refresh Ingenuity User document : Issues found in 459801587101_D Exam cards book | WIP22 | v4.1.10 | N/A | Subsystem |  | N/A | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | 10-JAN-2020 | Ingenuity 64 |  | v4.1.10 | January 9, 2020 1:12:04 PM GMT+08:00
CTUDT00627933 | Refresh | Closed | lin_s | December 20, 2019 12:30:27 PM GMT+08:00 | Refresh - Backup/Restore always show some files not found | WIP22.1 |  |  | Subsystem | SRN | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 |  | Console4 | NO | Closed |  | Ovation
CTUDT00627933 | Refresh | Closed | lin_s | December 20, 2019 12:30:27 PM GMT+08:00 | Refresh - Backup/Restore always show some files not found | WIP22.1 |  |  | Subsystem | SRN | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 |  | Console4 | NO | Closed |  | IQon1
CTUDT00627930 | Refresh | Closed | 320053770 | December 20, 2019 3:02:04 AM GMT+08:00 | Refresh; WIP22.1; IQon\ICT\ICTsp; While running the Reliability test, the ESTOP opens with an error message. | WIP22.1 | WIP23.1 |  | Subsystem |  | N/A | Fixed | tamari_o | 0 | 3 | DMS Electrical | Support-System | NO | Closed | WIP23.1SubCom | Ovation |  | WIP23.1 | January 9, 2020 5:19:18 PM GMT+08:00
CTUDT00627930 | Refresh | Closed | 320053770 | December 20, 2019 3:02:04 AM GMT+08:00 | Refresh; WIP22.1; IQon\ICT\ICTsp; While running the Reliability test, the ESTOP opens with an error message. | WIP22.1 | WIP23.1 |  | Subsystem |  | N/A | Fixed | tamari_o | 0 | 3 | DMS Electrical | Support-System | NO | Closed | WIP23.1SubCom | IQon1 |  | WIP23.1 | January 9, 2020 5:19:18 PM GMT+08:00
CTUDT00627922 | Refresh | Closed | levi_a | December 19, 2019 9:22:31 PM GMT+08:00 | Refresh:Localization: In French Harmonized help page was not fully translated. | WIP 12.2 |  |  | Subsystem | localization | DMS-SRS-1206 | No Plan to Fix | yatziv_l | 0 | 4 |  | DMS | NO | Closed | SubNC | Ovation
CTUDT00627922 | Refresh | Closed | levi_a | December 19, 2019 9:22:31 PM GMT+08:00 | Refresh:Localization: In French Harmonized help page was not fully translated. | WIP 12.2 |  |  | Subsystem | localization | DMS-SRS-1206 | No Plan to Fix | yatziv_l | 0 | 4 |  | DMS | NO | Closed | SubNC | Brilliance Big Bore
CTUDT00627922 | Refresh | Closed | levi_a | December 19, 2019 9:22:31 PM GMT+08:00 | Refresh:Localization: In French Harmonized help page was not fully translated. | WIP 12.2 |  |  | Subsystem | localization | DMS-SRS-1206 | No Plan to Fix | yatziv_l | 0 | 4 |  | DMS | NO | Closed | SubNC | Ingenuity 64
CTUDT00627922 | Refresh | Closed | levi_a | December 19, 2019 9:22:31 PM GMT+08:00 | Refresh:Localization: In French Harmonized help page was not fully translated. | WIP 12.2 |  |  | Subsystem | localization | DMS-SRS-1206 | No Plan to Fix | yatziv_l | 0 | 4 |  | DMS | NO | Closed | SubNC | IQon1
CTUDT00627912 | Refresh | Closed | Noy_Seelig | December 19, 2019 6:54:56 PM GMT+08:00 | Refresh: "Configuration Matrix" sentence in the Harmonized Help Page was not translated in Italian | 22.2 |  |  | Subsystem | localization | DMS-SRS-1206 | No Plan to Fix | yatziv_l | 0 | 4 |  | DMS | NO | Closed | SubNC | Ovation
CTUDT00627912 | Refresh | Closed | Noy_Seelig | December 19, 2019 6:54:56 PM GMT+08:00 | Refresh: "Configuration Matrix" sentence in the Harmonized Help Page was not translated in Italian | 22.2 |  |  | Subsystem | localization | DMS-SRS-1206 | No Plan to Fix | yatziv_l | 0 | 4 |  | DMS | NO | Closed | SubNC | Brilliance Big Bore
CTUDT00627912 | Refresh | Closed | Noy_Seelig | December 19, 2019 6:54:56 PM GMT+08:00 | Refresh: "Configuration Matrix" sentence in the Harmonized Help Page was not translated in Italian | 22.2 |  |  | Subsystem | localization | DMS-SRS-1206 | No Plan to Fix | yatziv_l | 0 | 4 |  | DMS | NO | Closed | SubNC | Ingenuity 64
CTUDT00627912 | Refresh | Closed | Noy_Seelig | December 19, 2019 6:54:56 PM GMT+08:00 | Refresh: "Configuration Matrix" sentence in the Harmonized Help Page was not translated in Italian | 22.2 |  |  | Subsystem | localization | DMS-SRS-1206 | No Plan to Fix | yatziv_l | 0 | 4 |  | DMS | NO | Closed | SubNC | IQon1
CTUDT00627910 | Refresh | Closed | gushansky_m | December 19, 2019 6:12:35 PM GMT+08:00 | Refresh: Load tube default doesn't bring Tube Serial Number (CT-NM.SRS.Service.500) | WIP22.1 | WIP23.1 |  | Subsystem | QA; Blocker | N/A | Fixed | zhou_x | 0 | 2 | Gantry Calibrations | GCX | NO | Closed | WIP23.1SegReq | IQon1 |  | WIP23.1 | January 10, 2020 10:23:36 AM GMT+08:00
CTUDT00627907 | Refresh | Closed | ding_f | December 19, 2019 3:21:55 PM GMT+08:00 | Refresh Bigbore: Performance Test instructions for H phantom has no statement about regulatory standards 21CFR1020.30 and 21CFR1020.33 | BB_WIP12.2 |  |  | Subsystem |  |  | Does Not Apply | yeshan_m | 0 |  |  | Support-System | NO | Closed | SubCom-25-Dec-2019 | Brilliance Big Bore
CTUDT00627902 | Refresh | Closed | wang_l | December 19, 2019 9:40:59 AM GMT+08:00 | Refresh BB - Gantry and IVC communicate error when set all IP to 10.168.0. | WIP12.2 | WIP23.1 |  | System | PM_Approved_BB_WIP12.3 | N/A | Fixed | yang_h | 0 | 3 | Gantry Software | GCX | NO | Closed | BB_WIP12.3 | Brilliance Big Bore |  | WIP23.1 | January 19, 2020 2:46:04 PM GMT+08:00
CTUDT00627898 | Refresh | Closed | wang_l | December 18, 2019 8:09:43 PM GMT+08:00 | Refresh - Ingenuity: one File upgrade failed because CalibrationTables_AirDataVarianceTable_CalibrationTables_AirDataVarianceTable_AirDataVarianceTable not found | WIP22.2 | WIP23.1 | Tube output calibration need to be run again. | System |  | n/a | Fixed | chakraborty_d | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP23.1SubCom | Ovation |  | WIP23.1 | January 6, 2020 5:51:53 PM GMT+08:00
CTUDT00627896 | Refresh | Accepted | zhong_x | December 18, 2019 6:51:33 PM GMT+08:00 | IVC Motion function failure by quickly up and down couch. | WIP12.2 |  | User shall wait for previous motion to complete means till the couch stops completely, User can visually see the couch motion and then use the motion touch buttons. | System | Console, SP2_Candidate | Console-SRS.15696;Console-SRS.15697 |  | reddy_g | 0 | 3 |  | Console4 | NO | Accepted | Works As Designed | Brilliance Big Bore
CTUDT00627894 | Refresh | Closed | zhong_x | December 18, 2019 3:34:24 PM GMT+08:00 | [Refresh]IVC entered BIOS automatically during second reboot after switch boot mode from PXE boot to Local | WIP12.2 |  | If any system with the dual-BIOS encounters the issue, the workaround is to simply hold the power button until the SBC shuts down (About 10 seconds, but do not pull the 48V power plug). Then power it back on normally with the power button. This is all it takes to clear the condition and it should not occur again even after switching between PXE and optimized boot BIOS'. | Subsystem |  | NA | Change Proposal | wang_k | 0 | 3 |  | Support-System | NO | Closed |  | Brilliance Big Bore
CTUDT00627892 | Refresh | Closed | zhong_x | December 18, 2019 3:27:23 PM GMT+08:00 | [Refresh] Refresh IVC:file missed after burning IVC ISO image to USB drive | WIP12.2 | DHF318575 |  | Subsystem |  | N/A | Fixed | pengfei_cai | 0 | 4 | Documentation | DI | NO | Closed | DI Release Note next version | Brilliance Big Bore |  | DHF318575 | December 20, 2019 1:10:25 PM GMT+08:00
CTUDT00627879 | Refresh | Closed | 320053768 | December 17, 2019 7:13:12 PM GMT+08:00 | The " [1][IQ-Dose_Mangement][DoseRightmAsValue]-[2407]” test failed, the displayed mAs and the calculated mAs have different values. | WIP22.1 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | tamari_o | 0 |  |  | Support-System | NO | Closed | SubCom-13-Feb-2020 | IQon1
CTUDT00627877 | Refresh | Closed | 320053768 | December 17, 2019 6:48:30 PM GMT+08:00 | The " [1][IQ-PH][O-MAR Artifacts][IQon]-[3071]” test failed because spectral(1) and spectral (0) images 120 kv, show more artifact in new versus old ones. | WIP21 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | tamari_o | 0 |  |  | Support-System | NO | Closed | SubCom-13-Feb-2020 | IQon1
CTUDT00627873 | Refresh | Closed | gT_s | December 17, 2019 3:52:04 PM GMT+08:00 | Refresh : The following tag 00E1:1063 in recon mdu files in the bugrep is not getting scrubbed. | WIP22.2 | DHF303211 Rev21 DHF316705 Rev21 |  | Subsystem |  | Console-SRS.8551 | Fixed | mistry_h | 0 | 4 | Requirements Documentation | Console4 | NO | Closed | SSRS Update | Ovation |  | DHF303211 Rev21 DHF316705 Rev21 | December 19, 2019 2:34:59 PM GMT+08:00
CTUDT00627873 | Refresh | Closed | gT_s | December 17, 2019 3:52:04 PM GMT+08:00 | Refresh : The following tag 00E1:1063 in recon mdu files in the bugrep is not getting scrubbed. | WIP22.2 | DHF303211 Rev21 DHF316705 Rev21 |  | Subsystem |  | Console-SRS.8551 | Fixed | mistry_h | 0 | 4 | Requirements Documentation | Console4 | NO | Closed | SSRS Update | Brilliance 64 |  | DHF303211 Rev21 DHF316705 Rev21 | December 19, 2019 2:34:59 PM GMT+08:00
CTUDT00627873 | Refresh | Closed | gT_s | December 17, 2019 3:52:04 PM GMT+08:00 | Refresh : The following tag 00E1:1063 in recon mdu files in the bugrep is not getting scrubbed. | WIP22.2 | DHF303211 Rev21 DHF316705 Rev21 |  | Subsystem |  | Console-SRS.8551 | Fixed | mistry_h | 0 | 4 | Requirements Documentation | Console4 | NO | Closed | SSRS Update | Ingenuity 64 |  | DHF303211 Rev21 DHF316705 Rev21 | December 19, 2019 2:34:59 PM GMT+08:00
CTUDT00627873 | Refresh | Closed | gT_s | December 17, 2019 3:52:04 PM GMT+08:00 | Refresh : The following tag 00E1:1063 in recon mdu files in the bugrep is not getting scrubbed. | WIP22.2 | DHF303211 Rev21 DHF316705 Rev21 |  | Subsystem |  | Console-SRS.8551 | Fixed | mistry_h | 0 | 4 | Requirements Documentation | Console4 | NO | Closed | SSRS Update | Ingenuity BigBore |  | DHF303211 Rev21 DHF316705 Rev21 | December 19, 2019 2:34:59 PM GMT+08:00
CTUDT00627873 | Refresh | Closed | gT_s | December 17, 2019 3:52:04 PM GMT+08:00 | Refresh : The following tag 00E1:1063 in recon mdu files in the bugrep is not getting scrubbed. | WIP22.2 | DHF303211 Rev21 DHF316705 Rev21 |  | Subsystem |  | Console-SRS.8551 | Fixed | mistry_h | 0 | 4 | Requirements Documentation | Console4 | NO | Closed | SSRS Update | IQon1 |  | DHF303211 Rev21 DHF316705 Rev21 | December 19, 2019 2:34:59 PM GMT+08:00
CTUDT00627868 | Refresh | Closed | vineetha_s | December 17, 2019 2:48:39 PM GMT+08:00 | Not able to enter 4 numerical characters in DOB field (YYYY) through Chinese Keyboard in Patient Details Page | WIP22.2 |  | Select the Year Text Box and press Alt + Shift. Now type the Year | Subsystem | CRN; | NA | No Plan to Fix | bhawna_g | 0 | 3 |  | Console4 | NO | Closed |  | Ovation
CTUDT00627868 | Refresh | Closed | vineetha_s | December 17, 2019 2:48:39 PM GMT+08:00 | Not able to enter 4 numerical characters in DOB field (YYYY) through Chinese Keyboard in Patient Details Page | WIP22.2 |  | Select the Year Text Box and press Alt + Shift. Now type the Year | Subsystem | CRN; | NA | No Plan to Fix | bhawna_g | 0 | 3 |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00627868 | Refresh | Closed | vineetha_s | December 17, 2019 2:48:39 PM GMT+08:00 | Not able to enter 4 numerical characters in DOB field (YYYY) through Chinese Keyboard in Patient Details Page | WIP22.2 |  | Select the Year Text Box and press Alt + Shift. Now type the Year | Subsystem | CRN; | NA | No Plan to Fix | bhawna_g | 0 | 3 |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00627868 | Refresh | Closed | vineetha_s | December 17, 2019 2:48:39 PM GMT+08:00 | Not able to enter 4 numerical characters in DOB field (YYYY) through Chinese Keyboard in Patient Details Page | WIP22.2 |  | Select the Year Text Box and press Alt + Shift. Now type the Year | Subsystem | CRN; | NA | No Plan to Fix | bhawna_g | 0 | 3 |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00627868 | Refresh | Closed | vineetha_s | December 17, 2019 2:48:39 PM GMT+08:00 | Not able to enter 4 numerical characters in DOB field (YYYY) through Chinese Keyboard in Patient Details Page | WIP22.2 |  | Select the Year Text Box and press Alt + Shift. Now type the Year | Subsystem | CRN; | NA | No Plan to Fix | bhawna_g | 0 | 3 |  | Console4 | NO | Closed |  | IQon1
CTUDT00627866 | Refresh | Closed | wei_f | December 17, 2019 11:51:53 AM GMT+08:00 | load patient containing Chinese character from iWATT HIS/RIS and try to perform scan, pop up CIRS error and cannot perform the scan. | WIP22.2 | WIP23.1 WIP22.3 |  | Subsystem | PM_Approved_WIP22.3 | N/A | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP24SegReq | Ovation |  | WIP23.1 WIP22.3 | January 14, 2020 9:35:14 PM GMT+08:00
CTUDT00627866 | Refresh | Closed | wei_f | December 17, 2019 11:51:53 AM GMT+08:00 | load patient containing Chinese character from iWATT HIS/RIS and try to perform scan, pop up CIRS error and cannot perform the scan. | WIP22.2 | WIP23.1 WIP22.3 |  | Subsystem | PM_Approved_WIP22.3 | N/A | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP24SegReq | Ingenuity 64 |  | WIP23.1 WIP22.3 | January 14, 2020 9:35:14 PM GMT+08:00
CTUDT00627866 | Refresh | Closed | wei_f | December 17, 2019 11:51:53 AM GMT+08:00 | load patient containing Chinese character from iWATT HIS/RIS and try to perform scan, pop up CIRS error and cannot perform the scan. | WIP22.2 | WIP23.1 WIP22.3 |  | Subsystem | PM_Approved_WIP22.3 | N/A | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP24SegReq | IQon1 |  | WIP23.1 WIP22.3 | January 14, 2020 9:35:14 PM GMT+08:00
CTUDT00627864 | Refresh | Closed | wang_l | December 17, 2019 10:38:14 AM GMT+08:00 | Refresh - Backup/Restore always show some files not found | WIP22.2 |  |  | Subsystem |  | N/A | Change Proposal | lin_s | 0 | 4 |  | Support-System | NO | Closed |  | Brilliance Big Bore
CTUDT00627864 | Refresh | Closed | wang_l | December 17, 2019 10:38:14 AM GMT+08:00 | Refresh - Backup/Restore always show some files not found | WIP22.2 |  |  | Subsystem |  | N/A | Change Proposal | lin_s | 0 | 4 |  | Support-System | NO | Closed |  | Ingenuity 64
CTUDT00627862 | Refresh | Closed | zhu_je | December 17, 2019 9:57:38 AM GMT+08:00 | [Refresh] Saline or Contrast volume does not appear on the progress ruler on the bottom of the table if injection time less than 16s | BB_WIP12 |  | User may see the volume information when hover on injection balloon or User may explicitly open injection parameter and see the injected solution volume information. | Subsystem |  | NA | No Plan to Fix | sagar_v | 0 | 4 |  | Console4 | NO | Closed | SubNC | Ovation
CTUDT00627862 | Refresh | Closed | zhu_je | December 17, 2019 9:57:38 AM GMT+08:00 | [Refresh] Saline or Contrast volume does not appear on the progress ruler on the bottom of the table if injection time less than 16s | BB_WIP12 |  | User may see the volume information when hover on injection balloon or User may explicitly open injection parameter and see the injected solution volume information. | Subsystem |  | NA | No Plan to Fix | sagar_v | 0 | 4 |  | Console4 | NO | Closed | SubNC | Brilliance 64
CTUDT00627862 | Refresh | Closed | zhu_je | December 17, 2019 9:57:38 AM GMT+08:00 | [Refresh] Saline or Contrast volume does not appear on the progress ruler on the bottom of the table if injection time less than 16s | BB_WIP12 |  | User may see the volume information when hover on injection balloon or User may explicitly open injection parameter and see the injected solution volume information. | Subsystem |  | NA | No Plan to Fix | sagar_v | 0 | 4 |  | Console4 | NO | Closed | SubNC | Brilliance Big Bore
CTUDT00627862 | Refresh | Closed | zhu_je | December 17, 2019 9:57:38 AM GMT+08:00 | [Refresh] Saline or Contrast volume does not appear on the progress ruler on the bottom of the table if injection time less than 16s | BB_WIP12 |  | User may see the volume information when hover on injection balloon or User may explicitly open injection parameter and see the injected solution volume information. | Subsystem |  | NA | No Plan to Fix | sagar_v | 0 | 4 |  | Console4 | NO | Closed | SubNC | Ingenuity 64
CTUDT00627862 | Refresh | Closed | zhu_je | December 17, 2019 9:57:38 AM GMT+08:00 | [Refresh] Saline or Contrast volume does not appear on the progress ruler on the bottom of the table if injection time less than 16s | BB_WIP12 |  | User may see the volume information when hover on injection balloon or User may explicitly open injection parameter and see the injected solution volume information. | Subsystem |  | NA | No Plan to Fix | sagar_v | 0 | 4 |  | Console4 | NO | Closed | SubNC | IQon1
CTUDT00627854 | Refresh | Closed | cohen_y | December 16, 2019 8:44:00 PM GMT+08:00 | Refresh: Visual Diagnostics: Bad Detectors Calibration is not performed during Modules Test | WIP22.1 | WIP23.1 | Continue without Bad Detectors Calibration Tests. | Subsystem | Bad Detectors | N/A | Fixed | kovalev_a | 0 | 2 | DMS Software | DMS | NO | Closed | WIP23.1SubCom | Ovation |  | WIP23.1 | January 2, 2020 7:40:54 PM GMT+08:00
CTUDT00627847 | Refresh | Closed | yatziv_l | December 16, 2019 3:31:42 PM GMT+08:00 | Refresh--Need automatic centering phantom function with therapy top | 4.1.10, 4.7.7 | WIP31 | Manual centering | Subsystem | SP1_PM_Approved | N/A | Fixed | frider_o | 0 | 3 | DMS Software | DMS | NO | Closed | WIP31SubCom | Ovation |  | WIP31 | July 22, 2020 5:17:36 PM GMT+08:00
CTUDT00627847 | Refresh | Closed | yatziv_l | December 16, 2019 3:31:42 PM GMT+08:00 | Refresh--Need automatic centering phantom function with therapy top | 4.1.10, 4.7.7 | WIP31 | Manual centering | Subsystem | SP1_PM_Approved | N/A | Fixed | frider_o | 0 | 3 | DMS Software | DMS | NO | Closed | WIP31SubCom | IQon1 |  | WIP31 | July 22, 2020 5:17:36 PM GMT+08:00
CTUDT00627845 | Refresh | Closed | 320053768 | December 16, 2019 3:59:01 AM GMT+08:00 | The test failed while measuring the ratio between "MTF with IMR" vs. "MTF, iDose level 0”. | WIP22 |  |  | Subsystem | IQ; Blocker |  | Works As Specified / Not A Defect | richards_b | 0 |  |  | Support-Physics | NO | Closed |  | IQon1
CTUDT00627843 | Refresh | Closed | Ashish_Sharma | December 15, 2019 11:55:27 PM GMT+08:00 | Ingenuity User document : Issues found in 459801587101_C Exam cards book | WIP21 | v4.1.10 | N/A | Subsystem |  | N/A | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | 20-Dec-2019 | Ingenuity 64 |  | v4.1.10 | December 17, 2019 8:17:28 PM GMT+08:00
CTUDT00627840 | Refresh | Closed | 320053767 | December 15, 2019 11:27:01 PM GMT+08:00 | IQ PEP 1.3 : "High limit"is lower than "expected" when running [IQ-PH][Gammex472_SpecialRecons][IQon]-[2454][2456][2457][2460][2818][2819] | WIP21 |  |  | Subsystem | Blocker |  | Works As Specified / Not A Defect | 320053767 | 0 |  |  |  |  | Closed |  | IQon1
CTUDT00627839 | Refresh | Closed | vinoliya_j | December 15, 2019 11:26:54 PM GMT+08:00 | 459801063604_A - iPatient 4.1.10(Ingenuity and Brilliance 64) Customer release notes - Validation comments | WIP21 | v4.1.10 | N/A | Subsystem |  | N/A | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | 20-Dec-2019 | Ingenuity 64 |  | v4.1.10 | December 17, 2019 8:14:08 PM GMT+08:00
CTUDT00627836 | Refresh | Closed | cai_t | December 15, 2019 9:44:57 PM GMT+08:00 | Refresh:Calibration:It popups a warning The phantom is not detected when do Tube Alignment calibration with automatic centering. | BB_WIP12.2 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | yatziv_l | 0 |  |  | DMS | NO | Closed |  | Brilliance 64
CTUDT00627836 | Refresh | Closed | cai_t | December 15, 2019 9:44:57 PM GMT+08:00 | Refresh:Calibration:It popups a warning The phantom is not detected when do Tube Alignment calibration with automatic centering. | BB_WIP12.2 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | yatziv_l | 0 |  |  | DMS | NO | Closed |  | Brilliance Big Bore
CTUDT00627836 | Refresh | Closed | cai_t | December 15, 2019 9:44:57 PM GMT+08:00 | Refresh:Calibration:It popups a warning The phantom is not detected when do Tube Alignment calibration with automatic centering. | BB_WIP12.2 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | yatziv_l | 0 |  |  | DMS | NO | Closed |  | Ingenuity 64
CTUDT00627834 | Refresh | Closed | dannyb | December 15, 2019 9:34:42 PM GMT+08:00 | Modes Refresh test design | WIP21 | WIP23 |  | Subsystem |  | CT-NM.SRS.Modes.1196;CT-NM.SRS.Modes.1811;CT-NM.SRS.Modes.1244 | Fixed | tamari_o | 0 | 2 | Test Design | Support-System | NO | Closed | SubCom-06-Jan-2020 | Ovation |  | WIP23 | January 8, 2020 7:00:43 PM GMT+08:00
CTUDT00627834 | Refresh | Closed | dannyb | December 15, 2019 9:34:42 PM GMT+08:00 | Modes Refresh test design | WIP21 | WIP23 |  | Subsystem |  | CT-NM.SRS.Modes.1196;CT-NM.SRS.Modes.1811;CT-NM.SRS.Modes.1244 | Fixed | tamari_o | 0 | 2 | Test Design | Support-System | NO | Closed | SubCom-06-Jan-2020 | IQon1 |  | WIP23 | January 8, 2020 7:00:43 PM GMT+08:00
CTUDT00627831 | Refresh | Closed | zhong_x | December 14, 2019 9:28:08 AM GMT+08:00 | “Recon box(CIRS) cannot initialize…” message pop up during CCT Scan | WIP12 |  | For Scenario1: Change the display mode to 3 and then change it back to Single. For Scenario2: Change the coverage to some other value and then change it back to 12. | Subsystem |  | NA | No Plan to Fix | kr_m | 0 | 3 |  | Console4 | NO | Closed | SubNC | Brilliance Big Bore
CTUDT00627825 | Refresh | Closed | org | December 13, 2019 5:44:04 PM GMT+08:00 | Small focal spot thickness correction factors wrong in Dose.cfg causing false negative CTDI Free air failure | WIP22 | WIP23 |  | System |  | N/A | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP23SegReq | Ovation |  | WIP23 | December 16, 2019 2:21:52 PM GMT+08:00
CTUDT00627825 | Refresh | Closed | org | December 13, 2019 5:44:04 PM GMT+08:00 | Small focal spot thickness correction factors wrong in Dose.cfg causing false negative CTDI Free air failure | WIP22 | WIP23 |  | System |  | N/A | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP23SegReq | IQon1 |  | WIP23 | December 16, 2019 2:21:52 PM GMT+08:00
CTUDT00627822 | Refresh | Closed | 320053770 | December 13, 2019 2:10:49 AM GMT+08:00 | Refresh – Wip22.1:IQon: Helical IQ test is paused before the Scan is completed. | WIP 22.1 |  |  | Subsystem | IQ Test; | No | No Plan to Fix | benny_w | 0 | 3 |  | GCX | YES | Closed |  | Ovation
CTUDT00627822 | Refresh | Closed | 320053770 | December 13, 2019 2:10:49 AM GMT+08:00 | Refresh – Wip22.1:IQon: Helical IQ test is paused before the Scan is completed. | WIP 22.1 |  |  | Subsystem | IQ Test; | No | No Plan to Fix | benny_w | 0 | 3 |  | GCX | YES | Closed |  | IQon1
CTUDT00627816 | Refresh | Closed | Noy_Seelig | December 12, 2019 9:35:34 PM GMT+08:00 | Refresh: IQ Check in CT user: Start Button In Instructions Window Is Not Translated In German, Norwegian and Danish | WIP22.1 |  |  | Subsystem | IQ Check |  | Works As Specified / Not A Defect | frider_o | 0 |  |  | DMS | NO | Closed |  | Ovation
CTUDT00627814 | Refresh | Closed | 320053770 | December 12, 2019 8:31:39 PM GMT+08:00 | Refresh – Wip22.1:IQon: When Performing a Scan and Recon of many results, The Console ResultsEngine crush and the Reconstruction slows down. | WIP22.1 |  |  | Subsystem |  |  | Duplicate | cohen_r | 0 |  |  | Support_Spectral | NO | Closed |  | IQon1
CTUDT00627812 | Refresh | Scored | wang_l | December 12, 2019 7:46:55 PM GMT+08:00 | Refresh: IP subnet still display 192.168. but we select 10.168 during one click installation | WIP22 | 459801587361 Rev D 459801587331 Rev E | ComputerAndNetworkConfiguration tool can be used post installation to update the CIRS IP Address value. User can select all the options that can be updated during one click installation with this tool and update it. | System | SP2_Candidate | NA |  | mukherjee_s | 0 | 3 | Service Documentation | Console4 | NO | Scored | SP1-TBD | Ovation |  | 459801587361 Rev D 459801587331 Rev E | January 11, 2020 1:14:25 PM GMT+08:00
CTUDT00627812 | Refresh | Scored | wang_l | December 12, 2019 7:46:55 PM GMT+08:00 | Refresh: IP subnet still display 192.168. but we select 10.168 during one click installation | WIP22 | 459801587361 Rev D 459801587331 Rev E | ComputerAndNetworkConfiguration tool can be used post installation to update the CIRS IP Address value. User can select all the options that can be updated during one click installation with this tool and update it. | System | SP2_Candidate | NA |  | mukherjee_s | 0 | 3 | Service Documentation | Console4 | NO | Scored | SP1-TBD | Brilliance 64 |  | 459801587361 Rev D 459801587331 Rev E | January 11, 2020 1:14:25 PM GMT+08:00
CTUDT00627812 | Refresh | Scored | wang_l | December 12, 2019 7:46:55 PM GMT+08:00 | Refresh: IP subnet still display 192.168. but we select 10.168 during one click installation | WIP22 | 459801587361 Rev D 459801587331 Rev E | ComputerAndNetworkConfiguration tool can be used post installation to update the CIRS IP Address value. User can select all the options that can be updated during one click installation with this tool and update it. | System | SP2_Candidate | NA |  | mukherjee_s | 0 | 3 | Service Documentation | Console4 | NO | Scored | SP1-TBD | Brilliance Big Bore |  | 459801587361 Rev D 459801587331 Rev E | January 11, 2020 1:14:25 PM GMT+08:00
CTUDT00627812 | Refresh | Scored | wang_l | December 12, 2019 7:46:55 PM GMT+08:00 | Refresh: IP subnet still display 192.168. but we select 10.168 during one click installation | WIP22 | 459801587361 Rev D 459801587331 Rev E | ComputerAndNetworkConfiguration tool can be used post installation to update the CIRS IP Address value. User can select all the options that can be updated during one click installation with this tool and update it. | System | SP2_Candidate | NA |  | mukherjee_s | 0 | 3 | Service Documentation | Console4 | NO | Scored | SP1-TBD | Ingenuity 64 |  | 459801587361 Rev D 459801587331 Rev E | January 11, 2020 1:14:25 PM GMT+08:00
CTUDT00627812 | Refresh | Scored | wang_l | December 12, 2019 7:46:55 PM GMT+08:00 | Refresh: IP subnet still display 192.168. but we select 10.168 during one click installation | WIP22 | 459801587361 Rev D 459801587331 Rev E | ComputerAndNetworkConfiguration tool can be used post installation to update the CIRS IP Address value. User can select all the options that can be updated during one click installation with this tool and update it. | System | SP2_Candidate | NA |  | mukherjee_s | 0 | 3 | Service Documentation | Console4 | NO | Scored | SP1-TBD | IQon1 |  | 459801587361 Rev D 459801587331 Rev E | January 11, 2020 1:14:25 PM GMT+08:00
CTUDT00627809 | Refresh | Closed | vinoliya_j | December 12, 2019 6:37:51 PM GMT+08:00 | 459801011775_A - Big Bore v4.8 Customer release notes - Validation comments | BB_WIP11 | v4.8 | N/A | Subsystem |  | N/A | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | 10-JAN-2020 | Brilliance Big Bore |  | v4.8 | January 9, 2020 12:14:19 PM GMT+08:00
CTUDT00627805 | Refresh | Closed | shiraguppi_k | December 12, 2019 1:45:14 PM GMT+08:00 | 459801594951_C - Big Bore v4.8 Exam Cards Book - Validation comments | BB_WIP11 | E-121770 | N/A | Subsystem | IFU | N/A | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | 20-Dec-2019 | Brilliance Big Bore |  | E-121770 | December 17, 2019 7:58:03 PM GMT+08:00
CTUDT00627795 | Refresh | Closed | wang_l | December 11, 2019 8:57:58 PM GMT+08:00 | Refresh: The system does not allow multiple remote connections at a time with LOTS capability | BB_WIP12 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | Sandeep_Kumar | 0 |  |  | Service Tools | NO | Closed |  | Brilliance Big Bore
CTUDT00627788 | Refresh | Closed | vineetha_s | December 11, 2019 6:43:55 PM GMT+08:00 | Login screen is not displayed after the first restart. | WIP22.2 |  | Yes, there are two workarounds: 1. Press Alt+TAB keys to get login screen to the front (or) 2. Restart the system to get login screen automatically This workaround is required only once. On subsequent restarts login screen comes automatically. | Subsystem | CRN | NA | No Plan to Fix | ravindra_a | 0 | 4 |  | Console4 | NO | Closed | SubNC | Ovation
CTUDT00627788 | Refresh | Closed | vineetha_s | December 11, 2019 6:43:55 PM GMT+08:00 | Login screen is not displayed after the first restart. | WIP22.2 |  | Yes, there are two workarounds: 1. Press Alt+TAB keys to get login screen to the front (or) 2. Restart the system to get login screen automatically This workaround is required only once. On subsequent restarts login screen comes automatically. | Subsystem | CRN | NA | No Plan to Fix | ravindra_a | 0 | 4 |  | Console4 | NO | Closed | SubNC | Brilliance 64
CTUDT00627788 | Refresh | Closed | vineetha_s | December 11, 2019 6:43:55 PM GMT+08:00 | Login screen is not displayed after the first restart. | WIP22.2 |  | Yes, there are two workarounds: 1. Press Alt+TAB keys to get login screen to the front (or) 2. Restart the system to get login screen automatically This workaround is required only once. On subsequent restarts login screen comes automatically. | Subsystem | CRN | NA | No Plan to Fix | ravindra_a | 0 | 4 |  | Console4 | NO | Closed | SubNC | Brilliance Big Bore
CTUDT00627788 | Refresh | Closed | vineetha_s | December 11, 2019 6:43:55 PM GMT+08:00 | Login screen is not displayed after the first restart. | WIP22.2 |  | Yes, there are two workarounds: 1. Press Alt+TAB keys to get login screen to the front (or) 2. Restart the system to get login screen automatically This workaround is required only once. On subsequent restarts login screen comes automatically. | Subsystem | CRN | NA | No Plan to Fix | ravindra_a | 0 | 4 |  | Console4 | NO | Closed | SubNC | Ingenuity 64
CTUDT00627788 | Refresh | Closed | vineetha_s | December 11, 2019 6:43:55 PM GMT+08:00 | Login screen is not displayed after the first restart. | WIP22.2 |  | Yes, there are two workarounds: 1. Press Alt+TAB keys to get login screen to the front (or) 2. Restart the system to get login screen automatically This workaround is required only once. On subsequent restarts login screen comes automatically. | Subsystem | CRN | NA | No Plan to Fix | ravindra_a | 0 | 4 |  | Console4 | NO | Closed | SubNC | IQon1
CTUDT00627775 | Refresh | Closed | zhong_x | December 11, 2019 4:05:50 PM GMT+08:00 | Failed to update interface box Firmware(Invalid Response) and the error Communication interface error | WIP12 |  |  | Subsystem | IVC; Blocker | na | Works As Specified / Not A Defect | 320021835 | 0 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00627761 | Refresh | Closed | l_yiming | December 10, 2019 4:56:12 PM GMT+08:00 | Refresh: Reference protocols Mapping failed | WIP21 |  | yes, User may explicitly select exam card from the respective group and/or create User ECs and perform mapping to it. | Subsystem | Blocker | NA | No Plan to Fix | sagar_v | 0 | 4 |  | Console4 | NO | Closed |  | Ovation
CTUDT00627761 | Refresh | Closed | l_yiming | December 10, 2019 4:56:12 PM GMT+08:00 | Refresh: Reference protocols Mapping failed | WIP21 |  | yes, User may explicitly select exam card from the respective group and/or create User ECs and perform mapping to it. | Subsystem | Blocker | NA | No Plan to Fix | sagar_v | 0 | 4 |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00627761 | Refresh | Closed | l_yiming | December 10, 2019 4:56:12 PM GMT+08:00 | Refresh: Reference protocols Mapping failed | WIP21 |  | yes, User may explicitly select exam card from the respective group and/or create User ECs and perform mapping to it. | Subsystem | Blocker | NA | No Plan to Fix | sagar_v | 0 | 4 |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00627761 | Refresh | Closed | l_yiming | December 10, 2019 4:56:12 PM GMT+08:00 | Refresh: Reference protocols Mapping failed | WIP21 |  | yes, User may explicitly select exam card from the respective group and/or create User ECs and perform mapping to it. | Subsystem | Blocker | NA | No Plan to Fix | sagar_v | 0 | 4 |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00627761 | Refresh | Closed | l_yiming | December 10, 2019 4:56:12 PM GMT+08:00 | Refresh: Reference protocols Mapping failed | WIP21 |  | yes, User may explicitly select exam card from the respective group and/or create User ECs and perform mapping to it. | Subsystem | Blocker | NA | No Plan to Fix | sagar_v | 0 | 4 |  | Console4 | NO | Closed |  | IQon1
CTUDT00627739 | Refresh | Closed | zhu_je | December 9, 2019 4:05:41 PM GMT+08:00 | [Refresh] Tool Manual Exposure was occurred in document ING Image Quality and Dose Testing for v4.1.10 O-Level | WIP22 | 459801587241 Rev C | N/A | Subsystem |  | CT-NM.SRS.Service.2919 | Fixed | shi_b | 0 | 2 | Service Documentation | Support-Service | NO | Closed | Service Manual Update | Ingenuity 64 |  | 459801587241 Rev C | December 24, 2019 10:07:14 AM GMT+08:00
CTUDT00627737 | Refresh | Closed | cohen_y | December 9, 2019 2:15:31 AM GMT+08:00 | Refresh: Phantom Centering: Tilt of Harmonized Head Phantom is not always recognized | WIP22.1 |  | Manually adjust phantom position | Subsystem | Calibrations; SP1_PM_Approved | N/A | No Plan to Fix | frider_o | 0 | 4 |  | DMS | NO | Closed | SubNC | Ovation
CTUDT00627737 | Refresh | Closed | cohen_y | December 9, 2019 2:15:31 AM GMT+08:00 | Refresh: Phantom Centering: Tilt of Harmonized Head Phantom is not always recognized | WIP22.1 |  | Manually adjust phantom position | Subsystem | Calibrations; SP1_PM_Approved | N/A | No Plan to Fix | frider_o | 0 | 4 |  | DMS | NO | Closed | SubNC | Brilliance 64
CTUDT00627737 | Refresh | Closed | cohen_y | December 9, 2019 2:15:31 AM GMT+08:00 | Refresh: Phantom Centering: Tilt of Harmonized Head Phantom is not always recognized | WIP22.1 |  | Manually adjust phantom position | Subsystem | Calibrations; SP1_PM_Approved | N/A | No Plan to Fix | frider_o | 0 | 4 |  | DMS | NO | Closed | SubNC | Brilliance Big Bore
CTUDT00627737 | Refresh | Closed | cohen_y | December 9, 2019 2:15:31 AM GMT+08:00 | Refresh: Phantom Centering: Tilt of Harmonized Head Phantom is not always recognized | WIP22.1 |  | Manually adjust phantom position | Subsystem | Calibrations; SP1_PM_Approved | N/A | No Plan to Fix | frider_o | 0 | 4 |  | DMS | NO | Closed | SubNC | Ingenuity 64
CTUDT00627737 | Refresh | Closed | cohen_y | December 9, 2019 2:15:31 AM GMT+08:00 | Refresh: Phantom Centering: Tilt of Harmonized Head Phantom is not always recognized | WIP22.1 |  | Manually adjust phantom position | Subsystem | Calibrations; SP1_PM_Approved | N/A | No Plan to Fix | frider_o | 0 | 4 |  | DMS | NO | Closed | SubNC | IQon1
CTUDT00627689 | Refresh | Closed | cai_t | December 8, 2019 12:42:04 PM GMT+08:00 | MultiLanguage IFU has some errors | BB_WIP12 | v4.8 Refresh | N/A | Subsystem |  | N/A | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | 15-Jan-2020 | Brilliance 64 |  | v4.8 Refresh | January 24, 2020 2:17:17 PM GMT+08:00
CTUDT00627689 | Refresh | Closed | cai_t | December 8, 2019 12:42:04 PM GMT+08:00 | MultiLanguage IFU has some errors | BB_WIP12 | v4.8 Refresh | N/A | Subsystem |  | N/A | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | 15-Jan-2020 | Brilliance Big Bore |  | v4.8 Refresh | January 24, 2020 2:17:17 PM GMT+08:00
CTUDT00627689 | Refresh | Closed | cai_t | December 8, 2019 12:42:04 PM GMT+08:00 | MultiLanguage IFU has some errors | BB_WIP12 | v4.8 Refresh | N/A | Subsystem |  | N/A | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | 15-Jan-2020 | Ingenuity 64 |  | v4.8 Refresh | January 24, 2020 2:17:17 PM GMT+08:00
CTUDT00627687 | Refresh | Closed | lakshmikanth_s | December 7, 2019 7:11:58 PM GMT+08:00 | Images were not reconstructed for result (2.3) for cardiac helical scan. | BB_WIP12 | BB_WIP12.2 WIP22.2 |  | Subsystem |  | N/A | Fixed | bhawna_g | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP12.2_SubCom | Ovation |  | BB_WIP12.2 WIP22.2 | December 9, 2019 10:19:09 PM GMT+08:00
CTUDT00627687 | Refresh | Closed | lakshmikanth_s | December 7, 2019 7:11:58 PM GMT+08:00 | Images were not reconstructed for result (2.3) for cardiac helical scan. | BB_WIP12 | BB_WIP12.2 WIP22.2 |  | Subsystem |  | N/A | Fixed | bhawna_g | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP12.2_SubCom | Brilliance 64 |  | BB_WIP12.2 WIP22.2 | December 9, 2019 10:19:09 PM GMT+08:00
CTUDT00627687 | Refresh | Closed | lakshmikanth_s | December 7, 2019 7:11:58 PM GMT+08:00 | Images were not reconstructed for result (2.3) for cardiac helical scan. | BB_WIP12 | BB_WIP12.2 WIP22.2 |  | Subsystem |  | N/A | Fixed | bhawna_g | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP12.2_SubCom | Brilliance Big Bore |  | BB_WIP12.2 WIP22.2 | December 9, 2019 10:19:09 PM GMT+08:00
CTUDT00627687 | Refresh | Closed | lakshmikanth_s | December 7, 2019 7:11:58 PM GMT+08:00 | Images were not reconstructed for result (2.3) for cardiac helical scan. | BB_WIP12 | BB_WIP12.2 WIP22.2 |  | Subsystem |  | N/A | Fixed | bhawna_g | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP12.2_SubCom | Ingenuity 64 |  | BB_WIP12.2 WIP22.2 | December 9, 2019 10:19:09 PM GMT+08:00
CTUDT00627687 | Refresh | Closed | lakshmikanth_s | December 7, 2019 7:11:58 PM GMT+08:00 | Images were not reconstructed for result (2.3) for cardiac helical scan. | BB_WIP12 | BB_WIP12.2 WIP22.2 |  | Subsystem |  | N/A | Fixed | bhawna_g | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP12.2_SubCom | IQon1 |  | BB_WIP12.2 WIP22.2 | December 9, 2019 10:19:09 PM GMT+08:00
CTUDT00627685 | Refresh | Closed | dongping_w | December 7, 2019 10:27:13 AM GMT+08:00 | Refresh--Need automatic centering phantom function with therapy top | BB_WIP12 | WIP31 WIP30.1 | Negative impact with workaround as user can skip the auto centering and do centering manually. | Subsystem | SP1_PM_Approved | N/A | Fixed | frider_o | 0 | 3 | DMS Software | DMS | NO | Closed | WIP31SubCom | Brilliance 64 |  | WIP31 WIP30.1 | July 22, 2020 5:18:40 PM GMT+08:00
CTUDT00627685 | Refresh | Closed | dongping_w | December 7, 2019 10:27:13 AM GMT+08:00 | Refresh--Need automatic centering phantom function with therapy top | BB_WIP12 | WIP31 WIP30.1 | Negative impact with workaround as user can skip the auto centering and do centering manually. | Subsystem | SP1_PM_Approved | N/A | Fixed | frider_o | 0 | 3 | DMS Software | DMS | NO | Closed | WIP31SubCom | Ingenuity 64 |  | WIP31 WIP30.1 | July 22, 2020 5:18:40 PM GMT+08:00
CTUDT00627683 | Refresh | Closed | richards_b | December 7, 2019 4:57:35 AM GMT+08:00 | 3D Axial Back projection on GPU Reli issue | WIP22 | WIP22.2 |  | Subsystem |  | CIRS.SSRS.4587, CIRS.SSRS.4868, CIRS.SSRS.4869, CIRS.SSRS.4870, CIRS.SSRS.5202 | Fixed | richards_b | 0 | 4 | CIRS_SW_ImgNodeSvr | CIRS | YES | Closed | WIP22.2SegReq | Ovation |  | WIP22.2 | December 9, 2019 8:53:56 PM GMT+08:00
CTUDT00627683 | Refresh | Closed | richards_b | December 7, 2019 4:57:35 AM GMT+08:00 | 3D Axial Back projection on GPU Reli issue | WIP22 | WIP22.2 |  | Subsystem |  | CIRS.SSRS.4587, CIRS.SSRS.4868, CIRS.SSRS.4869, CIRS.SSRS.4870, CIRS.SSRS.5202 | Fixed | richards_b | 0 | 4 | CIRS_SW_ImgNodeSvr | CIRS | YES | Closed | WIP22.2SegReq | Brilliance 64 |  | WIP22.2 | December 9, 2019 8:53:56 PM GMT+08:00
CTUDT00627683 | Refresh | Closed | richards_b | December 7, 2019 4:57:35 AM GMT+08:00 | 3D Axial Back projection on GPU Reli issue | WIP22 | WIP22.2 |  | Subsystem |  | CIRS.SSRS.4587, CIRS.SSRS.4868, CIRS.SSRS.4869, CIRS.SSRS.4870, CIRS.SSRS.5202 | Fixed | richards_b | 0 | 4 | CIRS_SW_ImgNodeSvr | CIRS | YES | Closed | WIP22.2SegReq | Brilliance Big Bore |  | WIP22.2 | December 9, 2019 8:53:56 PM GMT+08:00
CTUDT00627683 | Refresh | Closed | richards_b | December 7, 2019 4:57:35 AM GMT+08:00 | 3D Axial Back projection on GPU Reli issue | WIP22 | WIP22.2 |  | Subsystem |  | CIRS.SSRS.4587, CIRS.SSRS.4868, CIRS.SSRS.4869, CIRS.SSRS.4870, CIRS.SSRS.5202 | Fixed | richards_b | 0 | 4 | CIRS_SW_ImgNodeSvr | CIRS | YES | Closed | WIP22.2SegReq | Ingenuity 64 |  | WIP22.2 | December 9, 2019 8:53:56 PM GMT+08:00
CTUDT00627683 | Refresh | Closed | richards_b | December 7, 2019 4:57:35 AM GMT+08:00 | 3D Axial Back projection on GPU Reli issue | WIP22 | WIP22.2 |  | Subsystem |  | CIRS.SSRS.4587, CIRS.SSRS.4868, CIRS.SSRS.4869, CIRS.SSRS.4870, CIRS.SSRS.5202 | Fixed | richards_b | 0 | 4 | CIRS_SW_ImgNodeSvr | CIRS | YES | Closed | WIP22.2SegReq | IQon1 |  | WIP22.2 | December 9, 2019 8:53:56 PM GMT+08:00
CTUDT00627675 | Refresh | Closed | richards_b | December 6, 2019 10:40:52 AM GMT+08:00 | CIRS Serviceability/Installation Updates | WIP22 | WIP22.2 |  | Subsystem |  | CIRS.SSRS.4922, CIRS.SSRS.5084, CIRS.SSRS.5085, CIRS.SSRS.6659 | Fixed | richards_b | 0 | 3 | CIRS_SW_Installation | CIRS | NO | Closed | BB_WIP12.2_SegReq | Ovation |  | WIP22.2 | December 9, 2019 8:53:29 PM GMT+08:00
CTUDT00627675 | Refresh | Closed | richards_b | December 6, 2019 10:40:52 AM GMT+08:00 | CIRS Serviceability/Installation Updates | WIP22 | WIP22.2 |  | Subsystem |  | CIRS.SSRS.4922, CIRS.SSRS.5084, CIRS.SSRS.5085, CIRS.SSRS.6659 | Fixed | richards_b | 0 | 3 | CIRS_SW_Installation | CIRS | NO | Closed | BB_WIP12.2_SegReq | Brilliance 64 |  | WIP22.2 | December 9, 2019 8:53:29 PM GMT+08:00
CTUDT00627675 | Refresh | Closed | richards_b | December 6, 2019 10:40:52 AM GMT+08:00 | CIRS Serviceability/Installation Updates | WIP22 | WIP22.2 |  | Subsystem |  | CIRS.SSRS.4922, CIRS.SSRS.5084, CIRS.SSRS.5085, CIRS.SSRS.6659 | Fixed | richards_b | 0 | 3 | CIRS_SW_Installation | CIRS | NO | Closed | BB_WIP12.2_SegReq | Brilliance Big Bore |  | WIP22.2 | December 9, 2019 8:53:29 PM GMT+08:00
CTUDT00627675 | Refresh | Closed | richards_b | December 6, 2019 10:40:52 AM GMT+08:00 | CIRS Serviceability/Installation Updates | WIP22 | WIP22.2 |  | Subsystem |  | CIRS.SSRS.4922, CIRS.SSRS.5084, CIRS.SSRS.5085, CIRS.SSRS.6659 | Fixed | richards_b | 0 | 3 | CIRS_SW_Installation | CIRS | NO | Closed | BB_WIP12.2_SegReq | Ingenuity 64 |  | WIP22.2 | December 9, 2019 8:53:29 PM GMT+08:00
CTUDT00627675 | Refresh | Closed | richards_b | December 6, 2019 10:40:52 AM GMT+08:00 | CIRS Serviceability/Installation Updates | WIP22 | WIP22.2 |  | Subsystem |  | CIRS.SSRS.4922, CIRS.SSRS.5084, CIRS.SSRS.5085, CIRS.SSRS.6659 | Fixed | richards_b | 0 | 3 | CIRS_SW_Installation | CIRS | NO | Closed | BB_WIP12.2_SegReq | IQon1 |  | WIP22.2 | December 9, 2019 8:53:29 PM GMT+08:00
CTUDT00627673 | Refresh | Closed | manju_kolkar | December 6, 2019 12:17:49 AM GMT+08:00 | Result was disabled after clicking on “Accept Changes” or “Shorten Scan” buttons. | WIP21 |  | Recreate results with acceptable parameters | Subsystem |  | N/A | No Plan to Fix | navale_s | 0 | 4 |  | Console4 | NO | Closed | WAS | IQon1
CTUDT00627651 | Refresh | Closed | zhong_x | December 5, 2019 6:54:23 PM GMT+08:00 | Constancy test instruction is incorrect. | WIP12 | WIP22.2 |  | Subsystem | Verified1 | n/a | Fixed | frider_o | 0 | 3 | DMS Software | DMS | NO | Closed | BB_WIP12.2_SubCom | Ovation |  | WIP22.2 | December 9, 2019 7:06:03 PM GMT+08:00
CTUDT00627651 | Refresh | Closed | zhong_x | December 5, 2019 6:54:23 PM GMT+08:00 | Constancy test instruction is incorrect. | WIP12 | WIP22.2 |  | Subsystem | Verified1 | n/a | Fixed | frider_o | 0 | 3 | DMS Software | DMS | NO | Closed | BB_WIP12.2_SubCom | Brilliance Big Bore |  | WIP22.2 | December 9, 2019 7:06:03 PM GMT+08:00
CTUDT00627651 | Refresh | Closed | zhong_x | December 5, 2019 6:54:23 PM GMT+08:00 | Constancy test instruction is incorrect. | WIP12 | WIP22.2 |  | Subsystem | Verified1 | n/a | Fixed | frider_o | 0 | 3 | DMS Software | DMS | NO | Closed | BB_WIP12.2_SubCom | Ingenuity 64 |  | WIP22.2 | December 9, 2019 7:06:03 PM GMT+08:00
CTUDT00627651 | Refresh | Closed | zhong_x | December 5, 2019 6:54:23 PM GMT+08:00 | Constancy test instruction is incorrect. | WIP12 | WIP22.2 |  | Subsystem | Verified1 | n/a | Fixed | frider_o | 0 | 3 | DMS Software | DMS | NO | Closed | BB_WIP12.2_SubCom | IQon1 |  | WIP22.2 | December 9, 2019 7:06:03 PM GMT+08:00
CTUDT00627647 | Refresh | Closed | manju_kolkar | December 5, 2019 5:12:57 PM GMT+08:00 | Wrong error message was displayed while enabling/disabled P3T when injector is non-idle | WIP21 |  |  | Subsystem |  | N/A | No Plan to Fix | navale_s | 0 | 3 |  | Console4 | NO | Closed | SubNC | Ovation
CTUDT00627647 | Refresh | Closed | manju_kolkar | December 5, 2019 5:12:57 PM GMT+08:00 | Wrong error message was displayed while enabling/disabled P3T when injector is non-idle | WIP21 |  |  | Subsystem |  | N/A | No Plan to Fix | navale_s | 0 | 3 |  | Console4 | NO | Closed | SubNC | Brilliance 64
CTUDT00627647 | Refresh | Closed | manju_kolkar | December 5, 2019 5:12:57 PM GMT+08:00 | Wrong error message was displayed while enabling/disabled P3T when injector is non-idle | WIP21 |  |  | Subsystem |  | N/A | No Plan to Fix | navale_s | 0 | 3 |  | Console4 | NO | Closed | SubNC | Brilliance Big Bore
CTUDT00627647 | Refresh | Closed | manju_kolkar | December 5, 2019 5:12:57 PM GMT+08:00 | Wrong error message was displayed while enabling/disabled P3T when injector is non-idle | WIP21 |  |  | Subsystem |  | N/A | No Plan to Fix | navale_s | 0 | 3 |  | Console4 | NO | Closed | SubNC | Ingenuity 64
CTUDT00627647 | Refresh | Closed | manju_kolkar | December 5, 2019 5:12:57 PM GMT+08:00 | Wrong error message was displayed while enabling/disabled P3T when injector is non-idle | WIP21 |  |  | Subsystem |  | N/A | No Plan to Fix | navale_s | 0 | 3 |  | Console4 | NO | Closed | SubNC | IQon1
CTUDT00627635 | Refresh | Closed | Jose_D_Sergio_Almeida | December 5, 2019 2:02:51 PM GMT+08:00 | The table height Changed Press enable message occurs even after the table height is changed and acquisition completed. | WIP21 |  |  | Subsystem |  | N/A | No Plan to Fix | bhawna_g | 0 | 3 |  | Console4 | NO | Closed | SubNC | Ovation | NA
CTUDT00627635 | Refresh | Closed | Jose_D_Sergio_Almeida | December 5, 2019 2:02:51 PM GMT+08:00 | The table height Changed Press enable message occurs even after the table height is changed and acquisition completed. | WIP21 |  |  | Subsystem |  | N/A | No Plan to Fix | bhawna_g | 0 | 3 |  | Console4 | NO | Closed | SubNC | Brilliance 64 | NA
CTUDT00627635 | Refresh | Closed | Jose_D_Sergio_Almeida | December 5, 2019 2:02:51 PM GMT+08:00 | The table height Changed Press enable message occurs even after the table height is changed and acquisition completed. | WIP21 |  |  | Subsystem |  | N/A | No Plan to Fix | bhawna_g | 0 | 3 |  | Console4 | NO | Closed | SubNC | Brilliance Big Bore | NA
CTUDT00627635 | Refresh | Closed | Jose_D_Sergio_Almeida | December 5, 2019 2:02:51 PM GMT+08:00 | The table height Changed Press enable message occurs even after the table height is changed and acquisition completed. | WIP21 |  |  | Subsystem |  | N/A | No Plan to Fix | bhawna_g | 0 | 3 |  | Console4 | NO | Closed | SubNC | Ingenuity 64 | NA
CTUDT00627635 | Refresh | Closed | Jose_D_Sergio_Almeida | December 5, 2019 2:02:51 PM GMT+08:00 | The table height Changed Press enable message occurs even after the table height is changed and acquisition completed. | WIP21 |  |  | Subsystem |  | N/A | No Plan to Fix | bhawna_g | 0 | 3 |  | Console4 | NO | Closed | SubNC | IQon1 | NA
CTUDT00627631 | Refresh | Closed | 320053769 | December 5, 2019 12:37:07 AM GMT+08:00 | WIP22 – Software error appeared while executing trauma automation test | WIP22.1 |  | Use offline reconstruction | Subsystem | Performance; Blocker | N/A | Does Not Apply | navale_s | 0 | 4 |  | Console4 | NO | Closed | SubNC | Ovation
CTUDT00627631 | Refresh | Closed | 320053769 | December 5, 2019 12:37:07 AM GMT+08:00 | WIP22 – Software error appeared while executing trauma automation test | WIP22.1 |  | Use offline reconstruction | Subsystem | Performance; Blocker | N/A | Does Not Apply | navale_s | 0 | 4 |  | Console4 | NO | Closed | SubNC | Brilliance 64
CTUDT00627631 | Refresh | Closed | 320053769 | December 5, 2019 12:37:07 AM GMT+08:00 | WIP22 – Software error appeared while executing trauma automation test | WIP22.1 |  | Use offline reconstruction | Subsystem | Performance; Blocker | N/A | Does Not Apply | navale_s | 0 | 4 |  | Console4 | NO | Closed | SubNC | Brilliance Big Bore
CTUDT00627631 | Refresh | Closed | 320053769 | December 5, 2019 12:37:07 AM GMT+08:00 | WIP22 – Software error appeared while executing trauma automation test | WIP22.1 |  | Use offline reconstruction | Subsystem | Performance; Blocker | N/A | Does Not Apply | navale_s | 0 | 4 |  | Console4 | NO | Closed | SubNC | Ingenuity 64
CTUDT00627631 | Refresh | Closed | 320053769 | December 5, 2019 12:37:07 AM GMT+08:00 | WIP22 – Software error appeared while executing trauma automation test | WIP22.1 |  | Use offline reconstruction | Subsystem | Performance; Blocker | N/A | Does Not Apply | navale_s | 0 | 4 |  | Console4 | NO | Closed | SubNC | IQon1
CTUDT00627622 | Refresh | Closed | 320053769 | December 4, 2019 10:17:09 PM GMT+08:00 | WIP22 – while launching WF “A critical error has occurred, starting log-out process” warning appeared | WIP22.1 |  | This Use case is observed in Philips_Service user. User can restart the application if it crashes. | Subsystem |  | NA | No Plan to Fix | kapadia_v | 0 | 4 |  | Console4 | NO | Closed | SubNC | Ovation
CTUDT00627612 | Refresh | Closed | vineetha_s | December 4, 2019 7:51:28 PM GMT+08:00 | Patient name, Patient ID & Accession Number are not displaying in Chinese in report PDF if Save type is "DICOM Embedded PDF" | WIP22 |  | When the reporting language is set to Chinese, the user is able to view the report with the respective details. | Subsystem | CRN | NA | No Plan to Fix | s_sa | 0 | 3 |  | iCAP | NO | Closed |  | Ovation
CTUDT00627612 | Refresh | Closed | vineetha_s | December 4, 2019 7:51:28 PM GMT+08:00 | Patient name, Patient ID & Accession Number are not displaying in Chinese in report PDF if Save type is "DICOM Embedded PDF" | WIP22 |  | When the reporting language is set to Chinese, the user is able to view the report with the respective details. | Subsystem | CRN | NA | No Plan to Fix | s_sa | 0 | 3 |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00627612 | Refresh | Closed | vineetha_s | December 4, 2019 7:51:28 PM GMT+08:00 | Patient name, Patient ID & Accession Number are not displaying in Chinese in report PDF if Save type is "DICOM Embedded PDF" | WIP22 |  | When the reporting language is set to Chinese, the user is able to view the report with the respective details. | Subsystem | CRN | NA | No Plan to Fix | s_sa | 0 | 3 |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00627612 | Refresh | Closed | vineetha_s | December 4, 2019 7:51:28 PM GMT+08:00 | Patient name, Patient ID & Accession Number are not displaying in Chinese in report PDF if Save type is "DICOM Embedded PDF" | WIP22 |  | When the reporting language is set to Chinese, the user is able to view the report with the respective details. | Subsystem | CRN | NA | No Plan to Fix | s_sa | 0 | 3 |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00627612 | Refresh | Closed | vineetha_s | December 4, 2019 7:51:28 PM GMT+08:00 | Patient name, Patient ID & Accession Number are not displaying in Chinese in report PDF if Save type is "DICOM Embedded PDF" | WIP22 |  | When the reporting language is set to Chinese, the user is able to view the report with the respective details. | Subsystem | CRN | NA | No Plan to Fix | s_sa | 0 | 3 |  | iCAP | NO | Closed |  | IQon1
CTUDT00627597 | Refresh | Closed | l_yiming | December 4, 2019 9:59:28 AM GMT+08:00 | Refresh: In Local Directory Deleting Patient Data Failed | WIP21 |  |  | Subsystem |  | NA | No Plan to Fix | s_panigrahi | 0 | 4 |  | iCAP | NO | Closed |  | Ovation
CTUDT00627597 | Refresh | Closed | l_yiming | December 4, 2019 9:59:28 AM GMT+08:00 | Refresh: In Local Directory Deleting Patient Data Failed | WIP21 |  |  | Subsystem |  | NA | No Plan to Fix | s_panigrahi | 0 | 4 |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00627597 | Refresh | Closed | l_yiming | December 4, 2019 9:59:28 AM GMT+08:00 | Refresh: In Local Directory Deleting Patient Data Failed | WIP21 |  |  | Subsystem |  | NA | No Plan to Fix | s_panigrahi | 0 | 4 |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00627597 | Refresh | Closed | l_yiming | December 4, 2019 9:59:28 AM GMT+08:00 | Refresh: In Local Directory Deleting Patient Data Failed | WIP21 |  |  | Subsystem |  | NA | No Plan to Fix | s_panigrahi | 0 | 4 |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00627597 | Refresh | Closed | l_yiming | December 4, 2019 9:59:28 AM GMT+08:00 | Refresh: In Local Directory Deleting Patient Data Failed | WIP21 |  |  | Subsystem |  | NA | No Plan to Fix | s_panigrahi | 0 | 4 |  | iCAP | NO | Closed |  | IQon1
CTUDT00627583 | Refresh | Closed | Ashish_Sharma | December 3, 2019 11:50:50 PM GMT+08:00 | When loading CS S&S COR S&S exam card ,arrhythmia handling is not proper for prospective step & Shoot cardiac scan. | WIP21 |  |  | Subsystem | Blocker |  | Works As Specified / Not A Defect | tan_j | 0 |  |  | GCX | NO | Closed |  | Ingenuity 64
CTUDT00627572 | Refresh | Closed | manjunath_mk | December 3, 2019 6:15:21 PM GMT+08:00 | IVC does not log accurate positions for Vertical and tilt motions | WIP12 | DHF330449 Rev17 Big Bore Refresh HOST Verification Test Protocol | None | Subsystem |  | N/A | Fixed | mN_m | 0 | 3 | Test Design | Console4 | NO | Closed | WIP23SegReq | Brilliance Big Bore |  | DHF330449 Rev17 Big Bore Refresh HOST Verification Test Protocol | December 5, 2019 7:19:35 PM GMT+08:00
CTUDT00627569 | Refresh | Closed | lei_yang | December 3, 2019 5:41:42 PM GMT+08:00 | The scan ruler is stuck and popup error message When perform axial non Jog scan with 15 groups,255 Cycles. | WIP22 |  |  | Subsystem | Test Design | NA | Invalid/Uncontrolled Configuration | lei_yang | 0 | 4 |  | Support-System | NO | Closed | SubNC | Ovation
CTUDT00627569 | Refresh | Closed | lei_yang | December 3, 2019 5:41:42 PM GMT+08:00 | The scan ruler is stuck and popup error message When perform axial non Jog scan with 15 groups,255 Cycles. | WIP22 |  |  | Subsystem | Test Design | NA | Invalid/Uncontrolled Configuration | lei_yang | 0 | 4 |  | Support-System | NO | Closed | SubNC | Brilliance 64
CTUDT00627569 | Refresh | Closed | lei_yang | December 3, 2019 5:41:42 PM GMT+08:00 | The scan ruler is stuck and popup error message When perform axial non Jog scan with 15 groups,255 Cycles. | WIP22 |  |  | Subsystem | Test Design | NA | Invalid/Uncontrolled Configuration | lei_yang | 0 | 4 |  | Support-System | NO | Closed | SubNC | Brilliance Big Bore
CTUDT00627569 | Refresh | Closed | lei_yang | December 3, 2019 5:41:42 PM GMT+08:00 | The scan ruler is stuck and popup error message When perform axial non Jog scan with 15 groups,255 Cycles. | WIP22 |  |  | Subsystem | Test Design | NA | Invalid/Uncontrolled Configuration | lei_yang | 0 | 4 |  | Support-System | NO | Closed | SubNC | Ingenuity 64
CTUDT00627569 | Refresh | Closed | lei_yang | December 3, 2019 5:41:42 PM GMT+08:00 | The scan ruler is stuck and popup error message When perform axial non Jog scan with 15 groups,255 Cycles. | WIP22 |  |  | Subsystem | Test Design | NA | Invalid/Uncontrolled Configuration | lei_yang | 0 | 4 |  | Support-System | NO | Closed | SubNC | IQon1
CTUDT00627560 | Refresh | Closed | 320036209 | December 3, 2019 4:09:59 PM GMT+08:00 | Refresh Service Val_BB RT service validation comments of v4.2 to v4.8 | 4.2.1.17001 | 459801587371 Rev C |  | Subsystem |  | NA | Fixed | 320036209 | 0 | 2 | Service Documentation | Support-Service | NO | Closed | SubCom-13-Dec-2019 | Brilliance Big Bore |  | 459801587371 Rev C | December 19, 2019 3:57:26 PM GMT+08:00
CTUDT00627558 | Refresh | Closed | manjunath_mk | December 3, 2019 3:52:45 PM GMT+08:00 | System throws error pop up while performing coronary exam with multiple results with more number of images | WIP12 | DHF330449 Rev17 Big Bore Refresh HOST Verification Test Protocol | Have maximum of 11 results instead of 12 . | Subsystem | Test Design | NA | Fixed | mN_m | 0 | 4 | Test Design | Console4 | NO | Closed | Test Design | Brilliance Big Bore |  | DHF330449 Rev17 Big Bore Refresh HOST Verification Test Protocol | December 11, 2019 2:21:31 AM GMT+08:00
CTUDT00627551 | Refresh | Closed | 320053769 | December 3, 2019 12:45:10 AM GMT+08:00 | WIP22 – the popups: “the selected exam cards cannot be loaded” “there seems to be a problem in the application” and “system error” appeared after choosing “coronary CTA” EC. | WIP22.1 | BB_WIP12.2 WIP22.2 | Log out and log in to console, EC can be loaded normally. | System | EC; Verified 1 | NA | Fixed | tamari_o | 0 | 3 | Console Software | Console4 | NO | Closed | WIP22.2SubCom | Brilliance 64 |  | BB_WIP12.2 WIP22.2 | December 5, 2019 8:08:29 PM GMT+08:00
CTUDT00627551 | Refresh | Closed | 320053769 | December 3, 2019 12:45:10 AM GMT+08:00 | WIP22 – the popups: “the selected exam cards cannot be loaded” “there seems to be a problem in the application” and “system error” appeared after choosing “coronary CTA” EC. | WIP22.1 | BB_WIP12.2 WIP22.2 | Log out and log in to console, EC can be loaded normally. | System | EC; Verified 1 | NA | Fixed | tamari_o | 0 | 3 | Console Software | Console4 | NO | Closed | WIP22.2SubCom | Brilliance Big Bore |  | BB_WIP12.2 WIP22.2 | December 5, 2019 8:08:29 PM GMT+08:00
CTUDT00627551 | Refresh | Closed | 320053769 | December 3, 2019 12:45:10 AM GMT+08:00 | WIP22 – the popups: “the selected exam cards cannot be loaded” “there seems to be a problem in the application” and “system error” appeared after choosing “coronary CTA” EC. | WIP22.1 | BB_WIP12.2 WIP22.2 | Log out and log in to console, EC can be loaded normally. | System | EC; Verified 1 | NA | Fixed | tamari_o | 0 | 3 | Console Software | Console4 | NO | Closed | WIP22.2SubCom | Ingenuity 64 |  | BB_WIP12.2 WIP22.2 | December 5, 2019 8:08:29 PM GMT+08:00
CTUDT00627546 | Refresh | Closed | guy_sheaffer | December 2, 2019 5:26:54 PM GMT+08:00 | FSTracking_PSD_params_Helios.cfg - remove from upgrade profile | WIP22.1 | WIP23 |  | Subsystem | Backup Restore | N/A | Fixed | zhou_x | 0 | 2 | Gantry Calibrations | GCX | YES | Closed | WIP23SegReq | Ovation |  | WIP23 | December 10, 2019 9:51:01 AM GMT+08:00
CTUDT00627546 | Refresh | Closed | guy_sheaffer | December 2, 2019 5:26:54 PM GMT+08:00 | FSTracking_PSD_params_Helios.cfg - remove from upgrade profile | WIP22.1 | WIP23 |  | Subsystem | Backup Restore | N/A | Fixed | zhou_x | 0 | 2 | Gantry Calibrations | GCX | YES | Closed | WIP23SegReq | IQon1 |  | WIP23 | December 10, 2019 9:51:01 AM GMT+08:00
CTUDT00627544 | Refresh | Closed | zhang_t | December 2, 2019 4:03:14 PM GMT+08:00 | Some performance degrade in recent WIP compared with previous released WIP20 | WIP21 |  |  | Subsystem | System Performance |  | Duplicate | roy_r | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00627542 | Refresh | Closed | gushansky_m | December 2, 2019 3:22:26 PM GMT+08:00 | Refresh: Visual Diagnostics . Data-Path diagnostic failed | WIP21 |  |  | Subsystem | VD; Blocker |  | System Repaired | zhang_lawrence | 0 |  |  | GCX | NO | Closed |  | iCT4
CTUDT00627540 | Refresh | Closed | gushansky_m | December 2, 2019 2:57:02 PM GMT+08:00 | Refresh: Filament Diagnostics. No FSU rail voltage results. CT-NM.SRS.Service.2021 | WIP21 | WIP23 |  | Subsystem | VD; Blocker | Service.2021 | Fixed | zhang_lawrence | 0 | 2 | Gantry Software | GCX | NO | Closed | WIP23SegReq | IQon1 |  | WIP23 | December 17, 2019 5:25:38 PM GMT+08:00
CTUDT00627538 | Refresh | Closed | chandran_a | December 2, 2019 1:20:42 PM GMT+08:00 | Bypassing Thick Client Application Restriction | WIP21 |  |  | Subsystem | Security | N/A | No Plan to Fix | dasgupta_s | 0 | 3 |  | Service Tools | NO | Closed | SubNC | Ingenuity BigBore
CTUDT00627532 | Refresh | Closed | 320053770 | December 1, 2019 8:54:09 PM GMT+08:00 | Refresh – Wip21:IQon: when performing Dual surview with high CTDIvol value – the WF plan screen freezes after confirming popup message. | WIP21 | WIP22.2 BB_WIP12.2 |  | Subsystem | UI Bug | N/A | Fixed | datta_i | 0 | 2 | Console Software | Console4 | YES | Closed | WIP22.2SegReq | Ovation | PR9502317 PR9812971 | WIP22.2 BB_WIP12.2 | December 5, 2019 11:38:20 AM GMT+08:00
CTUDT00627532 | Refresh | Closed | 320053770 | December 1, 2019 8:54:09 PM GMT+08:00 | Refresh – Wip21:IQon: when performing Dual surview with high CTDIvol value – the WF plan screen freezes after confirming popup message. | WIP21 | WIP22.2 BB_WIP12.2 |  | Subsystem | UI Bug | N/A | Fixed | datta_i | 0 | 2 | Console Software | Console4 | YES | Closed | WIP22.2SegReq | Brilliance 64 | PR9502317 PR9812971 | WIP22.2 BB_WIP12.2 | December 5, 2019 11:38:20 AM GMT+08:00
CTUDT00627532 | Refresh | Closed | 320053770 | December 1, 2019 8:54:09 PM GMT+08:00 | Refresh – Wip21:IQon: when performing Dual surview with high CTDIvol value – the WF plan screen freezes after confirming popup message. | WIP21 | WIP22.2 BB_WIP12.2 |  | Subsystem | UI Bug | N/A | Fixed | datta_i | 0 | 2 | Console Software | Console4 | YES | Closed | WIP22.2SegReq | Brilliance Big Bore | PR9502317 PR9812971 | WIP22.2 BB_WIP12.2 | December 5, 2019 11:38:20 AM GMT+08:00
CTUDT00627532 | Refresh | Closed | 320053770 | December 1, 2019 8:54:09 PM GMT+08:00 | Refresh – Wip21:IQon: when performing Dual surview with high CTDIvol value – the WF plan screen freezes after confirming popup message. | WIP21 | WIP22.2 BB_WIP12.2 |  | Subsystem | UI Bug | N/A | Fixed | datta_i | 0 | 2 | Console Software | Console4 | YES | Closed | WIP22.2SegReq | Ingenuity 64 | PR9502317 PR9812971 | WIP22.2 BB_WIP12.2 | December 5, 2019 11:38:20 AM GMT+08:00
CTUDT00627532 | Refresh | Closed | 320053770 | December 1, 2019 8:54:09 PM GMT+08:00 | Refresh – Wip21:IQon: when performing Dual surview with high CTDIvol value – the WF plan screen freezes after confirming popup message. | WIP21 | WIP22.2 BB_WIP12.2 |  | Subsystem | UI Bug | N/A | Fixed | datta_i | 0 | 2 | Console Software | Console4 | YES | Closed | WIP22.2SegReq | IQon1 | PR9502317 PR9812971 | WIP22.2 BB_WIP12.2 | December 5, 2019 11:38:20 AM GMT+08:00
CTUDT00627529 | Refresh | Closed | 320053769 | December 1, 2019 5:09:59 PM GMT+08:00 | WIP22 - While saving a remote system in LAN Config, “Lanconfig.exe has stopped working” warning appeared and afterward while launching WF “unexpected error during Workflow creation” warning appeared. | WIP22.1 |  | After entering CT user and then logging back to service user WF opened successfully and the new added “remote system” appeared in remote devices in PD. | Subsystem |  | NA | No Plan to Fix | srushti_k | 0 | 4 |  | iCAP | NO | Closed |  | Ovation
CTUDT00627529 | Refresh | Closed | 320053769 | December 1, 2019 5:09:59 PM GMT+08:00 | WIP22 - While saving a remote system in LAN Config, “Lanconfig.exe has stopped working” warning appeared and afterward while launching WF “unexpected error during Workflow creation” warning appeared. | WIP22.1 |  | After entering CT user and then logging back to service user WF opened successfully and the new added “remote system” appeared in remote devices in PD. | Subsystem |  | NA | No Plan to Fix | srushti_k | 0 | 4 |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00627529 | Refresh | Closed | 320053769 | December 1, 2019 5:09:59 PM GMT+08:00 | WIP22 - While saving a remote system in LAN Config, “Lanconfig.exe has stopped working” warning appeared and afterward while launching WF “unexpected error during Workflow creation” warning appeared. | WIP22.1 |  | After entering CT user and then logging back to service user WF opened successfully and the new added “remote system” appeared in remote devices in PD. | Subsystem |  | NA | No Plan to Fix | srushti_k | 0 | 4 |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00627529 | Refresh | Closed | 320053769 | December 1, 2019 5:09:59 PM GMT+08:00 | WIP22 - While saving a remote system in LAN Config, “Lanconfig.exe has stopped working” warning appeared and afterward while launching WF “unexpected error during Workflow creation” warning appeared. | WIP22.1 |  | After entering CT user and then logging back to service user WF opened successfully and the new added “remote system” appeared in remote devices in PD. | Subsystem |  | NA | No Plan to Fix | srushti_k | 0 | 4 |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00627529 | Refresh | Closed | 320053769 | December 1, 2019 5:09:59 PM GMT+08:00 | WIP22 - While saving a remote system in LAN Config, “Lanconfig.exe has stopped working” warning appeared and afterward while launching WF “unexpected error during Workflow creation” warning appeared. | WIP22.1 |  | After entering CT user and then logging back to service user WF opened successfully and the new added “remote system” appeared in remote devices in PD. | Subsystem |  | NA | No Plan to Fix | srushti_k | 0 | 4 |  | iCAP | NO | Closed |  | IQon1
CTUDT00627525 | Refresh | Closed | braunstein_d | December 1, 2019 3:57:51 PM GMT+08:00 | Refresh: Performance: IMR auto test on Big Bore bands/rings on images | 4.2.2.17xxx-WIP04 | WIP22.2 |  | Subsystem | Performance; Blocker ; Verified1 | DHF170396 Rev 21 | Fixed | elviral | 0 | 3 | DMS Calibrations | DMS | NO | Closed | BB_WIP12.2_SubCom | Brilliance Big Bore |  | WIP22.2 | December 8, 2019 3:49:38 PM GMT+08:00
CTUDT00627523 | Refresh | Closed | wei_f | November 30, 2019 4:09:03 PM GMT+08:00 | Setting the NTP server failed on host | WIP12 |  | NA | Subsystem | STPS | NA | No Plan to Fix | kr_m | 0 | 4 |  | Service Tools | NO | Closed |  | Brilliance Big Bore
CTUDT00627521 | Refresh | Closed | mN_m | November 29, 2019 8:58:47 PM GMT+08:00 | Priority transfer setting is enabled for Refresh Big Bore in preferences. | BB_WIP12 | WIP22.2 |  | Subsystem |  | NA | Fixed | srushti_k | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP22.2SegReq | Brilliance Big Bore |  | WIP22.2 | December 3, 2019 7:25:36 PM GMT+08:00
CTUDT00627519 | Refresh | Closed | manju_kolkar | November 29, 2019 8:19:15 PM GMT+08:00 | Update of breathing waveform took more than 0.2 sec | BB_WIP12 | Host Verification Test Protocol DHF318051 Rev 25 Big Bore Refresh HOST Verification Test Protocol DHF330449 Rev 17 |  | Subsystem |  | NA | Fixed | jindal_k | 0 | 2 | Test Design | Console4 | NO | Closed | SubVerification case update | Brilliance 64 |  | Host Verification Test Protocol DHF318051 Rev 25 Big Bore Refresh HOST Verification Test Protocol DHF330449 Rev 17 | December 5, 2019 7:08:24 PM GMT+08:00
CTUDT00627519 | Refresh | Closed | manju_kolkar | November 29, 2019 8:19:15 PM GMT+08:00 | Update of breathing waveform took more than 0.2 sec | BB_WIP12 | Host Verification Test Protocol DHF318051 Rev 25 Big Bore Refresh HOST Verification Test Protocol DHF330449 Rev 17 |  | Subsystem |  | NA | Fixed | jindal_k | 0 | 2 | Test Design | Console4 | NO | Closed | SubVerification case update | Brilliance Big Bore |  | Host Verification Test Protocol DHF318051 Rev 25 Big Bore Refresh HOST Verification Test Protocol DHF330449 Rev 17 | December 5, 2019 7:08:24 PM GMT+08:00
CTUDT00627519 | Refresh | Closed | manju_kolkar | November 29, 2019 8:19:15 PM GMT+08:00 | Update of breathing waveform took more than 0.2 sec | BB_WIP12 | Host Verification Test Protocol DHF318051 Rev 25 Big Bore Refresh HOST Verification Test Protocol DHF330449 Rev 17 |  | Subsystem |  | NA | Fixed | jindal_k | 0 | 2 | Test Design | Console4 | NO | Closed | SubVerification case update | Ingenuity 64 |  | Host Verification Test Protocol DHF318051 Rev 25 Big Bore Refresh HOST Verification Test Protocol DHF330449 Rev 17 | December 5, 2019 7:08:24 PM GMT+08:00
CTUDT00627511 | Refresh | Closed | chandran_a | November 29, 2019 6:31:26 PM GMT+08:00 | Insecure BIOS. | WIP21 | DMR356920 Rev07 459801587361 Rev D 459801587351 Rev D 459801587331 Rev E 459801587341 Rev E |  | Subsystem |  | N/A | Fixed | shaham_t | 0 | 2 | Service Documentation | Support-System | NO | Closed | WIP22.2SegReq | Brilliance 64 |  | DMR356920 Rev07 459801587361 Rev D 459801587351 Rev D 459801587331 Rev E 459801587341 Rev E | January 5, 2020 3:15:33 PM GMT+08:00
CTUDT00627511 | Refresh | Closed | chandran_a | November 29, 2019 6:31:26 PM GMT+08:00 | Insecure BIOS. | WIP21 | DMR356920 Rev07 459801587361 Rev D 459801587351 Rev D 459801587331 Rev E 459801587341 Rev E |  | Subsystem |  | N/A | Fixed | shaham_t | 0 | 2 | Service Documentation | Support-System | NO | Closed | WIP22.2SegReq | Brilliance Big Bore |  | DMR356920 Rev07 459801587361 Rev D 459801587351 Rev D 459801587331 Rev E 459801587341 Rev E | January 5, 2020 3:15:33 PM GMT+08:00
CTUDT00627511 | Refresh | Closed | chandran_a | November 29, 2019 6:31:26 PM GMT+08:00 | Insecure BIOS. | WIP21 | DMR356920 Rev07 459801587361 Rev D 459801587351 Rev D 459801587331 Rev E 459801587341 Rev E |  | Subsystem |  | N/A | Fixed | shaham_t | 0 | 2 | Service Documentation | Support-System | NO | Closed | WIP22.2SegReq | Ingenuity 64 |  | DMR356920 Rev07 459801587361 Rev D 459801587351 Rev D 459801587331 Rev E 459801587341 Rev E | January 5, 2020 3:15:33 PM GMT+08:00
CTUDT00627503 | Refresh | Closed | chandran_a | November 29, 2019 5:45:19 PM GMT+08:00 | CT user is able to copy files into System drive from external data storage. | WIP21 | BB_WIP12.2 WIP22.2 |  | Subsystem |  | N/A | Fixed | kapadia_v | 0 | 3 | Console Software | Support-Security | NO | Closed | WIP22.2SegReq | Ovation |  | BB_WIP12.2 WIP22.2 | December 3, 2019 9:34:11 PM GMT+08:00
CTUDT00627503 | Refresh | Closed | chandran_a | November 29, 2019 5:45:19 PM GMT+08:00 | CT user is able to copy files into System drive from external data storage. | WIP21 | BB_WIP12.2 WIP22.2 |  | Subsystem |  | N/A | Fixed | kapadia_v | 0 | 3 | Console Software | Support-Security | NO | Closed | WIP22.2SegReq | Brilliance 64 |  | BB_WIP12.2 WIP22.2 | December 3, 2019 9:34:11 PM GMT+08:00
CTUDT00627503 | Refresh | Closed | chandran_a | November 29, 2019 5:45:19 PM GMT+08:00 | CT user is able to copy files into System drive from external data storage. | WIP21 | BB_WIP12.2 WIP22.2 |  | Subsystem |  | N/A | Fixed | kapadia_v | 0 | 3 | Console Software | Support-Security | NO | Closed | WIP22.2SegReq | Brilliance Big Bore |  | BB_WIP12.2 WIP22.2 | December 3, 2019 9:34:11 PM GMT+08:00
CTUDT00627503 | Refresh | Closed | chandran_a | November 29, 2019 5:45:19 PM GMT+08:00 | CT user is able to copy files into System drive from external data storage. | WIP21 | BB_WIP12.2 WIP22.2 |  | Subsystem |  | N/A | Fixed | kapadia_v | 0 | 3 | Console Software | Support-Security | NO | Closed | WIP22.2SegReq | Ingenuity 64 |  | BB_WIP12.2 WIP22.2 | December 3, 2019 9:34:11 PM GMT+08:00
CTUDT00627503 | Refresh | Closed | chandran_a | November 29, 2019 5:45:19 PM GMT+08:00 | CT user is able to copy files into System drive from external data storage. | WIP21 | BB_WIP12.2 WIP22.2 |  | Subsystem |  | N/A | Fixed | kapadia_v | 0 | 3 | Console Software | Support-Security | NO | Closed | WIP22.2SegReq | IQon1 |  | BB_WIP12.2 WIP22.2 | December 3, 2019 9:34:11 PM GMT+08:00
CTUDT00627498 | Refresh | Closed | nanjunda_s | November 29, 2019 5:18:01 PM GMT+08:00 | Tolerance mismatch in the requirement and in the Reference exam card Spec. | wip21.1 | DHF303211 Rev20 DHF316705 Rev20 | N/A | Subsystem |  | Console-SRS.15416,17158,17157,17156,17603 | Fixed | roy_r | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | Doc update | Ovation |  | DHF303211 Rev20 DHF316705 Rev20 | December 4, 2019 12:17:23 PM GMT+08:00
CTUDT00627498 | Refresh | Closed | nanjunda_s | November 29, 2019 5:18:01 PM GMT+08:00 | Tolerance mismatch in the requirement and in the Reference exam card Spec. | wip21.1 | DHF303211 Rev20 DHF316705 Rev20 | N/A | Subsystem |  | Console-SRS.15416,17158,17157,17156,17603 | Fixed | roy_r | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | Doc update | Brilliance 64 |  | DHF303211 Rev20 DHF316705 Rev20 | December 4, 2019 12:17:23 PM GMT+08:00
CTUDT00627498 | Refresh | Closed | nanjunda_s | November 29, 2019 5:18:01 PM GMT+08:00 | Tolerance mismatch in the requirement and in the Reference exam card Spec. | wip21.1 | DHF303211 Rev20 DHF316705 Rev20 | N/A | Subsystem |  | Console-SRS.15416,17158,17157,17156,17603 | Fixed | roy_r | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | Doc update | Brilliance Big Bore |  | DHF303211 Rev20 DHF316705 Rev20 | December 4, 2019 12:17:23 PM GMT+08:00
CTUDT00627498 | Refresh | Closed | nanjunda_s | November 29, 2019 5:18:01 PM GMT+08:00 | Tolerance mismatch in the requirement and in the Reference exam card Spec. | wip21.1 | DHF303211 Rev20 DHF316705 Rev20 | N/A | Subsystem |  | Console-SRS.15416,17158,17157,17156,17603 | Fixed | roy_r | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | Doc update | Ingenuity 64 |  | DHF303211 Rev20 DHF316705 Rev20 | December 4, 2019 12:17:23 PM GMT+08:00
CTUDT00627498 | Refresh | Closed | nanjunda_s | November 29, 2019 5:18:01 PM GMT+08:00 | Tolerance mismatch in the requirement and in the Reference exam card Spec. | wip21.1 | DHF303211 Rev20 DHF316705 Rev20 | N/A | Subsystem |  | Console-SRS.15416,17158,17157,17156,17603 | Fixed | roy_r | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | Doc update | IQon1 |  | DHF303211 Rev20 DHF316705 Rev20 | December 4, 2019 12:17:23 PM GMT+08:00
CTUDT00627490 | Refresh | Closed | 310117562 | November 28, 2019 11:33:49 PM GMT+08:00 | No clear indication regarding side table controller (IVC HW) orientation | WIP11.1 |  | The current table side position is shown in Menu -> Table Position, User can check the current table position by selecting Table Position from Menu. | System |  | RFC | No Plan to Fix | reddy_g | 0 | 2 |  | Console4 | NO | Closed | SubNC | Brilliance Big Bore
CTUDT00627488 | Refresh | Closed | 310117562 | November 28, 2019 11:16:33 PM GMT+08:00 | Joystick mouse is difficult to use | WIP11.1 |  |  | Subsystem |  | NA | No Plan to Fix | wang_k | 0 | 4 |  | GCX | NO | Closed |  | Brilliance Big Bore
CTUDT00627486 | Refresh | Closed | 310117562 | November 28, 2019 9:36:32 PM GMT+08:00 | Plan box doesn't show the correct scan location and confuses the users | WIP11.1 |  | When user moves couch to one of the defined needle positions, the shot is taken at the same position. | Subsystem |  | RFC | No Plan to Fix | reddy_g | 0 | 3 |  | Console4 | NO | Closed | SubNC | Brilliance Big Bore
CTUDT00627483 | Refresh | Closed | wei_f | November 28, 2019 7:47:08 PM GMT+08:00 | The QuickReview.exe has stopped working when load the study to the Quick Review. | WIP12 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | aswani_v | 1 |  |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00627481 | Refresh | Closed | 320021835 | November 28, 2019 7:34:33 PM GMT+08:00 | Tilt does not stop immediately when button in IVC touch panel is released. | BB_WIP12 | BB_WIP12.2 |  | Subsystem |  | NA | Fixed | 320021835 | 0 | 3 | IVC Software | Console4 | NO | Closed | WIP12.2SubCom | Brilliance Big Bore |  | BB_WIP12.2 | November 29, 2019 7:51:40 PM GMT+08:00
CTUDT00627478 | Refresh | Closed | vineetha_s | November 28, 2019 2:45:30 PM GMT+08:00 | Spectral CT Viewer is getting crashed during autolaunch | WIP22 |  | 1. Wait till all the series are loaded in SCTV and then select. 2. Use load series to SCTV from the Patient Directory instead of autolaunch | Subsystem | SCTV | NA | No Plan to Fix | aswani_v | 0 | 4 |  | iCAP | NO | Closed |  | IQon1
CTUDT00627476 | Refresh | Closed | roy_r | November 28, 2019 2:02:13 PM GMT+08:00 | Refresh: Post Process images from EBW apps do not fill DICOM tags mandated as per DIN6862-2 standard | WIP21 |  |  | Subsystem |  | NA | No Plan to Fix | aswani_v | 0 | 2 |  | iCAP | NO | Closed |  | Ovation
CTUDT00627476 | Refresh | Closed | roy_r | November 28, 2019 2:02:13 PM GMT+08:00 | Refresh: Post Process images from EBW apps do not fill DICOM tags mandated as per DIN6862-2 standard | WIP21 |  |  | Subsystem |  | NA | No Plan to Fix | aswani_v | 0 | 2 |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00627476 | Refresh | Closed | roy_r | November 28, 2019 2:02:13 PM GMT+08:00 | Refresh: Post Process images from EBW apps do not fill DICOM tags mandated as per DIN6862-2 standard | WIP21 |  |  | Subsystem |  | NA | No Plan to Fix | aswani_v | 0 | 2 |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00627476 | Refresh | Closed | roy_r | November 28, 2019 2:02:13 PM GMT+08:00 | Refresh: Post Process images from EBW apps do not fill DICOM tags mandated as per DIN6862-2 standard | WIP21 |  |  | Subsystem |  | NA | No Plan to Fix | aswani_v | 0 | 2 |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00627476 | Refresh | Closed | roy_r | November 28, 2019 2:02:13 PM GMT+08:00 | Refresh: Post Process images from EBW apps do not fill DICOM tags mandated as per DIN6862-2 standard | WIP21 |  |  | Subsystem |  | NA | No Plan to Fix | aswani_v | 0 | 2 |  | iCAP | NO | Closed |  | IQon1
CTUDT00627474 | Refresh | Closed | wang_l | November 28, 2019 12:53:41 PM GMT+08:00 | Refresh Reliability - Coronary recon results was not generated | BB WIP11.1 |  | Result can be generated using Offline recon | Subsystem |  | NA | No Plan to Fix | bhawna_g | 0 | 3 |  | Console4 | NO | Closed | SUbNC | Brilliance Big Bore
CTUDT00627472 | Refresh | Closed | wang_l | November 28, 2019 12:37:14 PM GMT+08:00 | Refresh Reliability - BB:Error “There seems to be a problem in the application Would you like to create a Bug Report” once click "End Exam and Continue Working" | BB WIP11.1 |  |  | Subsystem |  | NA | Duplicate | s_panigrahi | 0 |  |  | iCAP | NO | Closed |  | Ovation
CTUDT00627472 | Refresh | Closed | wang_l | November 28, 2019 12:37:14 PM GMT+08:00 | Refresh Reliability - BB:Error “There seems to be a problem in the application Would you like to create a Bug Report” once click "End Exam and Continue Working" | BB WIP11.1 |  |  | Subsystem |  | NA | Duplicate | s_panigrahi | 0 |  |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00627472 | Refresh | Closed | wang_l | November 28, 2019 12:37:14 PM GMT+08:00 | Refresh Reliability - BB:Error “There seems to be a problem in the application Would you like to create a Bug Report” once click "End Exam and Continue Working" | BB WIP11.1 |  |  | Subsystem |  | NA | Duplicate | s_panigrahi | 0 |  |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00627472 | Refresh | Closed | wang_l | November 28, 2019 12:37:14 PM GMT+08:00 | Refresh Reliability - BB:Error “There seems to be a problem in the application Would you like to create a Bug Report” once click "End Exam and Continue Working" | BB WIP11.1 |  |  | Subsystem |  | NA | Duplicate | s_panigrahi | 0 |  |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00627472 | Refresh | Closed | wang_l | November 28, 2019 12:37:14 PM GMT+08:00 | Refresh Reliability - BB:Error “There seems to be a problem in the application Would you like to create a Bug Report” once click "End Exam and Continue Working" | BB WIP11.1 |  |  | Subsystem |  | NA | Duplicate | s_panigrahi | 0 |  |  | iCAP | NO | Closed |  | IQon1
CTUDT00627468 | Refresh | Closed | cai_t | November 27, 2019 9:46:12 PM GMT+08:00 | Microsoft IME has stopped working message popup when login system with CT user. | BB_WIP12 |  | As a workaround , please logout and login again and press Win+Space. This should resolve the issue. | Subsystem | CRN | NA | No Plan to Fix | s_panigrahi | 0 | 3 |  | iCAP | NO | Closed |  | Ovation
CTUDT00627468 | Refresh | Closed | cai_t | November 27, 2019 9:46:12 PM GMT+08:00 | Microsoft IME has stopped working message popup when login system with CT user. | BB_WIP12 |  | As a workaround , please logout and login again and press Win+Space. This should resolve the issue. | Subsystem | CRN | NA | No Plan to Fix | s_panigrahi | 0 | 3 |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00627468 | Refresh | Closed | cai_t | November 27, 2019 9:46:12 PM GMT+08:00 | Microsoft IME has stopped working message popup when login system with CT user. | BB_WIP12 |  | As a workaround , please logout and login again and press Win+Space. This should resolve the issue. | Subsystem | CRN | NA | No Plan to Fix | s_panigrahi | 0 | 3 |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00627468 | Refresh | Closed | cai_t | November 27, 2019 9:46:12 PM GMT+08:00 | Microsoft IME has stopped working message popup when login system with CT user. | BB_WIP12 |  | As a workaround , please logout and login again and press Win+Space. This should resolve the issue. | Subsystem | CRN | NA | No Plan to Fix | s_panigrahi | 0 | 3 |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00627468 | Refresh | Closed | cai_t | November 27, 2019 9:46:12 PM GMT+08:00 | Microsoft IME has stopped working message popup when login system with CT user. | BB_WIP12 |  | As a workaround , please logout and login again and press Win+Space. This should resolve the issue. | Subsystem | CRN | NA | No Plan to Fix | s_panigrahi | 0 | 3 |  | iCAP | NO | Closed |  | IQon1
CTUDT00627454 | Refresh | Closed | manjunath_mk | November 27, 2019 4:36:44 PM GMT+08:00 | IVC touch panels non functional when language is changed from console | WIP12 |  | Manually select the change panel position option using the options in Hamburger menu and select approriate panel position. | Subsystem |  | NA | No Plan to Fix | 320021835 | 0 | 3 |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00627449 | Refresh | Closed | manjunath_mk | November 27, 2019 4:30:23 PM GMT+08:00 | Localization message is wrong in IVC and information truncated in About page | WIP12 | BB_WIP12.2 |  | Subsystem |  | NA | Fixed | Snehashis_Dey | 0 | 4 | IVC Software | Console4 | NO | Closed | BB_WIP12.2_SubCom | Brilliance Big Bore |  | BB_WIP12.2 | November 28, 2019 10:07:55 PM GMT+08:00
CTUDT00627448 | Refresh | Closed | roy_r | November 27, 2019 4:30:00 PM GMT+08:00 | Refresh: DICOM tag SOURCE_IMAGE_SEQUENCE(0008,2112) is incorrectly filled for Spectral Axial images | WIP21 | WIP23.1 | N/A | Subsystem | Blocker | N/A | Fixed | aviw | 0 | 4 | Console Software | Support_Spectral | NO | Closed | SubCom-06-Jan-2020 | IQon1 |  | WIP23.1 | January 14, 2020 12:29:36 AM GMT+08:00
CTUDT00627445 | Refresh | Closed | roy_r | November 27, 2019 4:10:16 PM GMT+08:00 | Preference option desired for enabling/disabling DIN6862-compliance | WIP21 | WIP22.2 BB_WIP12.2 | None | Subsystem | Blocker | Console-SRS.17972 | Fixed | ravindra_a | 0 | 4 | Service Tools_Software | Console4 | NO | Closed | WIP23SegReq | Ovation |  | WIP22.2 BB_WIP12.2 | December 4, 2019 12:39:43 AM GMT+08:00
CTUDT00627445 | Refresh | Closed | roy_r | November 27, 2019 4:10:16 PM GMT+08:00 | Preference option desired for enabling/disabling DIN6862-compliance | WIP21 | WIP22.2 BB_WIP12.2 | None | Subsystem | Blocker | Console-SRS.17972 | Fixed | ravindra_a | 0 | 4 | Service Tools_Software | Console4 | NO | Closed | WIP23SegReq | Brilliance 64 |  | WIP22.2 BB_WIP12.2 | December 4, 2019 12:39:43 AM GMT+08:00
CTUDT00627445 | Refresh | Closed | roy_r | November 27, 2019 4:10:16 PM GMT+08:00 | Preference option desired for enabling/disabling DIN6862-compliance | WIP21 | WIP22.2 BB_WIP12.2 | None | Subsystem | Blocker | Console-SRS.17972 | Fixed | ravindra_a | 0 | 4 | Service Tools_Software | Console4 | NO | Closed | WIP23SegReq | Brilliance Big Bore |  | WIP22.2 BB_WIP12.2 | December 4, 2019 12:39:43 AM GMT+08:00
CTUDT00627445 | Refresh | Closed | roy_r | November 27, 2019 4:10:16 PM GMT+08:00 | Preference option desired for enabling/disabling DIN6862-compliance | WIP21 | WIP22.2 BB_WIP12.2 | None | Subsystem | Blocker | Console-SRS.17972 | Fixed | ravindra_a | 0 | 4 | Service Tools_Software | Console4 | NO | Closed | WIP23SegReq | Ingenuity 64 |  | WIP22.2 BB_WIP12.2 | December 4, 2019 12:39:43 AM GMT+08:00
CTUDT00627445 | Refresh | Closed | roy_r | November 27, 2019 4:10:16 PM GMT+08:00 | Preference option desired for enabling/disabling DIN6862-compliance | WIP21 | WIP22.2 BB_WIP12.2 | None | Subsystem | Blocker | Console-SRS.17972 | Fixed | ravindra_a | 0 | 4 | Service Tools_Software | Console4 | NO | Closed | WIP23SegReq | IQon1 |  | WIP22.2 BB_WIP12.2 | December 4, 2019 12:39:43 AM GMT+08:00
CTUDT00627443 | Refresh | Closed | rojin_antony | November 27, 2019 3:48:58 PM GMT+08:00 | Images reconstructed wrongly with blank space in-between | BB_WIP11 |  | The user would be able to perform offline recon to correct the issue. | Subsystem | Recon | N/A | Documentation Created/Updated | dongping_w | 0 | 3 | CIRS_SW_Installation | Support-System | YES | Closed | BB_WIP12_SegReq | Brilliance Big Bore |  |  | December 5, 2019 3:33:02 PM GMT+08:00
CTUDT00627440 | Refresh | Closed | rojin_antony | November 27, 2019 3:38:19 PM GMT+08:00 | Patient Directory Stopped working | BB_WIP11 |  | Select any of the options from the given options of the pop up message box displayed which will restore PD. | Subsystem | PD | NA | No Plan to Fix | s_panigrahi | 0 | 3 |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00627439 | Refresh | Closed | wang_l | November 27, 2019 3:35:01 PM GMT+08:00 | Refresh Reliability -BB: Gantry cannot initialize for clinical scan because of “S_MASTER_SEQUENCER_ACQ_NOT_OK” | BB_WIP11.1 |  | The user can acknowledge the error and re-submit the scan request. | Subsystem | ACQ_NOT_OK | N/A | Duplicate | dongping_w | 0 |  |  | Support-System | NO | Closed | SubCom-16-Dec-2019 | Brilliance Big Bore
CTUDT00627436 | Refresh | Closed | zhang_t | November 27, 2019 2:02:52 PM GMT+08:00 | View2 tab was not displayed automatically once click Recon button or End Exam and Continue Working for CCT EC. | BB_WIP12 |  | Since the View2 button is already activated, the user can click on View2 button on the right monitor to open View2. | Subsystem |  | N/A | No Plan to Fix | pandey_a | 0 | 4 |  | Console4 | NO | Closed | SubNC | Ovation
CTUDT00627436 | Refresh | Closed | zhang_t | November 27, 2019 2:02:52 PM GMT+08:00 | View2 tab was not displayed automatically once click Recon button or End Exam and Continue Working for CCT EC. | BB_WIP12 |  | Since the View2 button is already activated, the user can click on View2 button on the right monitor to open View2. | Subsystem |  | N/A | No Plan to Fix | pandey_a | 0 | 4 |  | Console4 | NO | Closed | SubNC | Brilliance 64
CTUDT00627436 | Refresh | Closed | zhang_t | November 27, 2019 2:02:52 PM GMT+08:00 | View2 tab was not displayed automatically once click Recon button or End Exam and Continue Working for CCT EC. | BB_WIP12 |  | Since the View2 button is already activated, the user can click on View2 button on the right monitor to open View2. | Subsystem |  | N/A | No Plan to Fix | pandey_a | 0 | 4 |  | Console4 | NO | Closed | SubNC | Brilliance Big Bore
CTUDT00627436 | Refresh | Closed | zhang_t | November 27, 2019 2:02:52 PM GMT+08:00 | View2 tab was not displayed automatically once click Recon button or End Exam and Continue Working for CCT EC. | BB_WIP12 |  | Since the View2 button is already activated, the user can click on View2 button on the right monitor to open View2. | Subsystem |  | N/A | No Plan to Fix | pandey_a | 0 | 4 |  | Console4 | NO | Closed | SubNC | Ingenuity 64
CTUDT00627436 | Refresh | Closed | zhang_t | November 27, 2019 2:02:52 PM GMT+08:00 | View2 tab was not displayed automatically once click Recon button or End Exam and Continue Working for CCT EC. | BB_WIP12 |  | Since the View2 button is already activated, the user can click on View2 button on the right monitor to open View2. | Subsystem |  | N/A | No Plan to Fix | pandey_a | 0 | 4 |  | Console4 | NO | Closed | SubNC | IQon1
CTUDT00627435 | Refresh | Closed | wang_l | November 27, 2019 1:55:17 PM GMT+08:00 | Refresh Reliability - Go button was disabled when End exam and load one new EC | BB_WIP11.1 |  |  | Subsystem | Workflow | N/A | No Plan to Fix | singh_j | 0 | 4 |  | Console4 | NO | Closed | SubNC | Brilliance Big Bore
CTUDT00627433 | Refresh | Closed | yuan_s | November 27, 2019 9:05:46 AM GMT+08:00 | The height of IQon Packaging failed the requirement defined in CIRS.SSRS.6666. | WIP22 | NA | N/A | Subsystem |  | CIRS.SSRS.6666 | Fixed | tysl_k | 0 | 3 | CIRS_HW_Generic | Support-System | NO | Closed | SubCom-13-Dec-2019 | IQon1 |  | NA | December 17, 2019 2:50:38 AM GMT+08:00
CTUDT00627423 | Refresh | Closed | ding_s | November 26, 2019 6:41:49 PM GMT+08:00 | The Rotation Time mentioned in requirements of SRS_Rev 09 is not consistent with the value displayed in Workflow UI. | WIP22 | DHF298554 Rev10 |  | Subsystem | SRS | CT-NM.SRS.SysPerf.582;CT-NM.SRS.SysPerf.804;CT-NM.SRS.SysPerf.771;CT-NM.SRS.SysPerf.774;CT-NM.SRS.SysPerf.769 | Fixed | ding_s | 0 | 3 | Requirements Documentation | Support-System | NO | Closed | SubCom-16-Dec-2019 | Ingenuity 64 |  | DHF298554 Rev10 | December 13, 2019 5:03:53 PM GMT+08:00
CTUDT00627421 | Refresh | Closed | Ashish_Sharma | November 26, 2019 5:25:49 PM GMT+08:00 | When refreshing the Patient directory with multiple clicks , Bugrep message appeared and system respond slowly. | WIP21 |  | Restart the system | Subsystem | Patient Directory | NA | No Plan to Fix | s_panigrahi | 0 | 3 |  | iCAP | NO | Closed |  | Ovation
CTUDT00627421 | Refresh | Closed | Ashish_Sharma | November 26, 2019 5:25:49 PM GMT+08:00 | When refreshing the Patient directory with multiple clicks , Bugrep message appeared and system respond slowly. | WIP21 |  | Restart the system | Subsystem | Patient Directory | NA | No Plan to Fix | s_panigrahi | 0 | 3 |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00627421 | Refresh | Closed | Ashish_Sharma | November 26, 2019 5:25:49 PM GMT+08:00 | When refreshing the Patient directory with multiple clicks , Bugrep message appeared and system respond slowly. | WIP21 |  | Restart the system | Subsystem | Patient Directory | NA | No Plan to Fix | s_panigrahi | 0 | 3 |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00627421 | Refresh | Closed | Ashish_Sharma | November 26, 2019 5:25:49 PM GMT+08:00 | When refreshing the Patient directory with multiple clicks , Bugrep message appeared and system respond slowly. | WIP21 |  | Restart the system | Subsystem | Patient Directory | NA | No Plan to Fix | s_panigrahi | 0 | 3 |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00627421 | Refresh | Closed | Ashish_Sharma | November 26, 2019 5:25:49 PM GMT+08:00 | When refreshing the Patient directory with multiple clicks , Bugrep message appeared and system respond slowly. | WIP21 |  | Restart the system | Subsystem | Patient Directory | NA | No Plan to Fix | s_panigrahi | 0 | 3 |  | iCAP | NO | Closed |  | IQon1
CTUDT00627419 | Refresh | Closed | 320053769 | November 26, 2019 4:54:41 PM GMT+08:00 | There was no option to send images from film tab to report. | WIP21 | DHF298554 Rev10 DHF298553 Rev.09 |  | Subsystem | SRS | Workflow.1798 | Fixed | bandil_j | 0 | 3 | Requirements Documentation | Support-System | NO | Closed | SubCom-19-Dec-2019 | Ovation |  | DHF298554 Rev10 DHF298553 Rev.09 | December 22, 2019 10:13:16 PM GMT+08:00
CTUDT00627419 | Refresh | Closed | 320053769 | November 26, 2019 4:54:41 PM GMT+08:00 | There was no option to send images from film tab to report. | WIP21 | DHF298554 Rev10 DHF298553 Rev.09 |  | Subsystem | SRS | Workflow.1798 | Fixed | bandil_j | 0 | 3 | Requirements Documentation | Support-System | NO | Closed | SubCom-19-Dec-2019 | Brilliance 64 |  | DHF298554 Rev10 DHF298553 Rev.09 | December 22, 2019 10:13:16 PM GMT+08:00
CTUDT00627419 | Refresh | Closed | 320053769 | November 26, 2019 4:54:41 PM GMT+08:00 | There was no option to send images from film tab to report. | WIP21 | DHF298554 Rev10 DHF298553 Rev.09 |  | Subsystem | SRS | Workflow.1798 | Fixed | bandil_j | 0 | 3 | Requirements Documentation | Support-System | NO | Closed | SubCom-19-Dec-2019 | Ingenuity 64 |  | DHF298554 Rev10 DHF298553 Rev.09 | December 22, 2019 10:13:16 PM GMT+08:00
CTUDT00627416 | Refresh | Closed | yosfan_s | November 26, 2019 4:04:11 PM GMT+08:00 | Refresh: Quality Assurance: Reasons Matrix: Selected procedures are not as defined in Spec | WIP17 | WIP20 |  | Subsystem | Verified1 | N/A | Fixed | yatziv_l | 0 | 3 | Design Documentation | DMS | NO | Closed | WIP22SubCom | Ovation |  | WIP20 | November 26, 2019 11:48:44 PM GMT+08:00
CTUDT00627414 | Refresh | Closed | yosfan_s | November 26, 2019 3:56:15 PM GMT+08:00 | The limits of some plots generated during FEE test do not match the spec | Wip8 | WIP10 |  | Subsystem |  | N/A | Fixed | yatziv_l | 0 | 2 | DMS Software | DMS | NO | Closed | WIP22SubCom | Ovation |  | WIP10 | November 26, 2019 11:52:13 PM GMT+08:00
CTUDT00627410 | Refresh | Closed | guy_sheaffer | November 25, 2019 10:25:55 PM GMT+08:00 | Sticky Keys popup allows access to control panel when logged in as clinical user | WIP21 | WIP22.2 BB_WIP12.2 | N/A | Subsystem | Blocker | N/A | Fixed | Prabhakar_Priyank_Muppidi | 0 | 2 | Console Software | Support-Security | NO | Closed | WIP22.2SubCom | Ovation |  | WIP22.2 BB_WIP12.2 | November 29, 2019 1:31:41 PM GMT+08:00
CTUDT00627410 | Refresh | Closed | guy_sheaffer | November 25, 2019 10:25:55 PM GMT+08:00 | Sticky Keys popup allows access to control panel when logged in as clinical user | WIP21 | WIP22.2 BB_WIP12.2 | N/A | Subsystem | Blocker | N/A | Fixed | Prabhakar_Priyank_Muppidi | 0 | 2 | Console Software | Support-Security | NO | Closed | WIP22.2SubCom | Brilliance 64 |  | WIP22.2 BB_WIP12.2 | November 29, 2019 1:31:41 PM GMT+08:00
CTUDT00627410 | Refresh | Closed | guy_sheaffer | November 25, 2019 10:25:55 PM GMT+08:00 | Sticky Keys popup allows access to control panel when logged in as clinical user | WIP21 | WIP22.2 BB_WIP12.2 | N/A | Subsystem | Blocker | N/A | Fixed | Prabhakar_Priyank_Muppidi | 0 | 2 | Console Software | Support-Security | NO | Closed | WIP22.2SubCom | Brilliance Big Bore |  | WIP22.2 BB_WIP12.2 | November 29, 2019 1:31:41 PM GMT+08:00
CTUDT00627410 | Refresh | Closed | guy_sheaffer | November 25, 2019 10:25:55 PM GMT+08:00 | Sticky Keys popup allows access to control panel when logged in as clinical user | WIP21 | WIP22.2 BB_WIP12.2 | N/A | Subsystem | Blocker | N/A | Fixed | Prabhakar_Priyank_Muppidi | 0 | 2 | Console Software | Support-Security | NO | Closed | WIP22.2SubCom | Ingenuity 64 |  | WIP22.2 BB_WIP12.2 | November 29, 2019 1:31:41 PM GMT+08:00
CTUDT00627410 | Refresh | Closed | guy_sheaffer | November 25, 2019 10:25:55 PM GMT+08:00 | Sticky Keys popup allows access to control panel when logged in as clinical user | WIP21 | WIP22.2 BB_WIP12.2 | N/A | Subsystem | Blocker | N/A | Fixed | Prabhakar_Priyank_Muppidi | 0 | 2 | Console Software | Support-Security | NO | Closed | WIP22.2SubCom | IQon1 |  | WIP22.2 BB_WIP12.2 | November 29, 2019 1:31:41 PM GMT+08:00
CTUDT00627405 | Refresh | Closed | swamypreethi | November 25, 2019 2:17:20 PM GMT+08:00 | IOCC: DIN - Presence of Reason for the Requested Procedure (0040,1002), Reason for Requested Procedure Code Sequence (0040,100A) at root level instead of sequence level | WIP22 | WIP22.2 BB_WIP12.2 |  | Subsystem | Blocker | NA | Fixed | ravindra_a | 0 | 3 | Console Software | Console4 | NO | Closed | WIP22.2SubCom | Ovation |  | WIP22.2 BB_WIP12.2 | November 29, 2019 1:36:33 PM GMT+08:00
CTUDT00627405 | Refresh | Closed | swamypreethi | November 25, 2019 2:17:20 PM GMT+08:00 | IOCC: DIN - Presence of Reason for the Requested Procedure (0040,1002), Reason for Requested Procedure Code Sequence (0040,100A) at root level instead of sequence level | WIP22 | WIP22.2 BB_WIP12.2 |  | Subsystem | Blocker | NA | Fixed | ravindra_a | 0 | 3 | Console Software | Console4 | NO | Closed | WIP22.2SubCom | Brilliance 64 |  | WIP22.2 BB_WIP12.2 | November 29, 2019 1:36:33 PM GMT+08:00
CTUDT00627405 | Refresh | Closed | swamypreethi | November 25, 2019 2:17:20 PM GMT+08:00 | IOCC: DIN - Presence of Reason for the Requested Procedure (0040,1002), Reason for Requested Procedure Code Sequence (0040,100A) at root level instead of sequence level | WIP22 | WIP22.2 BB_WIP12.2 |  | Subsystem | Blocker | NA | Fixed | ravindra_a | 0 | 3 | Console Software | Console4 | NO | Closed | WIP22.2SubCom | Ingenuity 64 |  | WIP22.2 BB_WIP12.2 | November 29, 2019 1:36:33 PM GMT+08:00
CTUDT00627405 | Refresh | Closed | swamypreethi | November 25, 2019 2:17:20 PM GMT+08:00 | IOCC: DIN - Presence of Reason for the Requested Procedure (0040,1002), Reason for Requested Procedure Code Sequence (0040,100A) at root level instead of sequence level | WIP22 | WIP22.2 BB_WIP12.2 |  | Subsystem | Blocker | NA | Fixed | ravindra_a | 0 | 3 | Console Software | Console4 | NO | Closed | WIP22.2SubCom | Ingenuity BigBore |  | WIP22.2 BB_WIP12.2 | November 29, 2019 1:36:33 PM GMT+08:00
CTUDT00627405 | Refresh | Closed | swamypreethi | November 25, 2019 2:17:20 PM GMT+08:00 | IOCC: DIN - Presence of Reason for the Requested Procedure (0040,1002), Reason for Requested Procedure Code Sequence (0040,100A) at root level instead of sequence level | WIP22 | WIP22.2 BB_WIP12.2 |  | Subsystem | Blocker | NA | Fixed | ravindra_a | 0 | 3 | Console Software | Console4 | NO | Closed | WIP22.2SubCom | IQon1 |  | WIP22.2 BB_WIP12.2 | November 29, 2019 1:36:33 PM GMT+08:00
CTUDT00627403 | Refresh | Closed | gT_s | November 25, 2019 1:50:02 PM GMT+08:00 | Refresh : Operation could not be completed ( error 0x00000490 ) is displayed after login to the olevel user. | WIP22 | WIP22.2 BB_WIP12.2 | N/A | Subsystem | Service Tools; Blocker | N/A | Fixed | tejaswini_nidamanuri | 0 | 3 | Service Tools_Software | Service Tools | NO | Closed | WIP22.2SegReq | Ovation |  | WIP22.2 BB_WIP12.2 | November 29, 2019 12:23:55 PM GMT+08:00
CTUDT00627403 | Refresh | Closed | gT_s | November 25, 2019 1:50:02 PM GMT+08:00 | Refresh : Operation could not be completed ( error 0x00000490 ) is displayed after login to the olevel user. | WIP22 | WIP22.2 BB_WIP12.2 | N/A | Subsystem | Service Tools; Blocker | N/A | Fixed | tejaswini_nidamanuri | 0 | 3 | Service Tools_Software | Service Tools | NO | Closed | WIP22.2SegReq | Brilliance 64 |  | WIP22.2 BB_WIP12.2 | November 29, 2019 12:23:55 PM GMT+08:00
CTUDT00627403 | Refresh | Closed | gT_s | November 25, 2019 1:50:02 PM GMT+08:00 | Refresh : Operation could not be completed ( error 0x00000490 ) is displayed after login to the olevel user. | WIP22 | WIP22.2 BB_WIP12.2 | N/A | Subsystem | Service Tools; Blocker | N/A | Fixed | tejaswini_nidamanuri | 0 | 3 | Service Tools_Software | Service Tools | NO | Closed | WIP22.2SegReq | Brilliance Big Bore |  | WIP22.2 BB_WIP12.2 | November 29, 2019 12:23:55 PM GMT+08:00
CTUDT00627403 | Refresh | Closed | gT_s | November 25, 2019 1:50:02 PM GMT+08:00 | Refresh : Operation could not be completed ( error 0x00000490 ) is displayed after login to the olevel user. | WIP22 | WIP22.2 BB_WIP12.2 | N/A | Subsystem | Service Tools; Blocker | N/A | Fixed | tejaswini_nidamanuri | 0 | 3 | Service Tools_Software | Service Tools | NO | Closed | WIP22.2SegReq | Ingenuity 64 |  | WIP22.2 BB_WIP12.2 | November 29, 2019 12:23:55 PM GMT+08:00
CTUDT00627403 | Refresh | Closed | gT_s | November 25, 2019 1:50:02 PM GMT+08:00 | Refresh : Operation could not be completed ( error 0x00000490 ) is displayed after login to the olevel user. | WIP22 | WIP22.2 BB_WIP12.2 | N/A | Subsystem | Service Tools; Blocker | N/A | Fixed | tejaswini_nidamanuri | 0 | 3 | Service Tools_Software | Service Tools | NO | Closed | WIP22.2SegReq | IQon1 |  | WIP22.2 BB_WIP12.2 | November 29, 2019 12:23:55 PM GMT+08:00
CTUDT00627396 | Refresh | Closed | Noy_Seelig | November 24, 2019 4:40:11 PM GMT+08:00 | Refresh: "Minutes left" string in Temperature Stabilization countdown bar is not translated | WIP21 |  |  | Subsystem | localization | N/A | No Plan to Fix | yatziv_l | 0 | 4 |  | DMS | NO | Closed | SubNC | Ovation
CTUDT00627396 | Refresh | Closed | Noy_Seelig | November 24, 2019 4:40:11 PM GMT+08:00 | Refresh: "Minutes left" string in Temperature Stabilization countdown bar is not translated | WIP21 |  |  | Subsystem | localization | N/A | No Plan to Fix | yatziv_l | 0 | 4 |  | DMS | NO | Closed | SubNC | IQon1
CTUDT00627393 | Refresh | Closed | shiraguppi_k | November 22, 2019 5:14:32 PM GMT+08:00 | Big Bore: BB Refresh Validation: Pulmo Viewer Crashed with Bug Rep error | BB_WIP11 |  | select single series to load in pulmo viewer. | Subsystem | Blocker | NA | No Plan to Fix | aswani_v | 0 | 3 |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00627391 | Refresh | Closed | manjunath_mk | November 22, 2019 2:44:31 PM GMT+08:00 | IVC does not move back to CT mode when logged out from service mode | WIP11.1 | BB_WIP12.2 | No | Subsystem |  | NA | Fixed | 320021835 | 0 | 2 | IVC Software | Console4 | NO | Closed | BB_WIP12_SubCom | Brilliance Big Bore |  | BB_WIP12.2 | November 25, 2019 1:21:00 PM GMT+08:00
CTUDT00627388 | Refresh | Closed | 320053770 | November 21, 2019 11:11:13 PM GMT+08:00 | Refresh – Wip21:IQon: ‘Iodine Removed’ Spectral result is not reconstructed. | WIP21 | WIP23.1 | Recreate the result (either in the exam application of from the spectral viewer) | Subsystem | Spectral; Blocker | NA | Fixed | aviw | 0 | 3 | Console Software | Support_Spectral | YES | Closed | WIP23SegReq | IQon1 |  | WIP23.1 | December 17, 2019 5:19:27 PM GMT+08:00
CTUDT00627381 | Refresh | Closed | lin_s | November 21, 2019 4:06:58 PM GMT+08:00 | ING:CRS.6694 is not applicable for Refresh anymore after brand naming integrated. | DHF314349 | DHF316435 Rev04 |  | Subsystem |  | CT-NM.PRS.CRS.6694 | Fixed | lin_s | 0 | 4 | Requirements Documentation | Support-System | NO | Closed | SubCom-16-Dec-2019 | Ingenuity 64 |  | DHF316435 Rev04 | December 13, 2019 5:06:06 PM GMT+08:00
CTUDT00627378 | Refresh | Closed | ding_s | November 21, 2019 1:12:54 PM GMT+08:00 | 1-Click Installation Process prompted again after close the installation Wizard, after automatic rebooting, it displayed no boot device found. | WIP21 | 459801587391 Rev C |  | Subsystem | Installation; Service Manual; | Host-Security-SwRS-058 | Fixed | zhang_l | 0 | 3 | Console Software | Support-Service | NO | Closed | SubCom-21-Feb-2020 | Brilliance 64 |  | 459801587391 Rev C | February 23, 2020 3:01:32 PM GMT+08:00
CTUDT00627376 | Refresh | Closed | wang_l | November 21, 2019 10:47:09 AM GMT+08:00 | Refresh Upgrade - Gantry calibration records are not upgraded from 4.1.7 to 4.1.10 | WIP21 | WIP22.2 WIP23 |  | System | Blocker; Upgrade; Verified1 | N/A | Fixed | chakraborty_d | 0 | 2 | Service Tools_Software | Console4 | NO | Closed | WIP22.2SubCom | Ovation |  | WIP22.2 WIP23 | December 12, 2019 7:05:15 PM GMT+08:00
CTUDT00627376 | Refresh | Closed | wang_l | November 21, 2019 10:47:09 AM GMT+08:00 | Refresh Upgrade - Gantry calibration records are not upgraded from 4.1.7 to 4.1.10 | WIP21 | WIP22.2 WIP23 |  | System | Blocker; Upgrade; Verified1 | N/A | Fixed | chakraborty_d | 0 | 2 | Service Tools_Software | Console4 | NO | Closed | WIP22.2SubCom | Brilliance 64 |  | WIP22.2 WIP23 | December 12, 2019 7:05:15 PM GMT+08:00
CTUDT00627376 | Refresh | Closed | wang_l | November 21, 2019 10:47:09 AM GMT+08:00 | Refresh Upgrade - Gantry calibration records are not upgraded from 4.1.7 to 4.1.10 | WIP21 | WIP22.2 WIP23 |  | System | Blocker; Upgrade; Verified1 | N/A | Fixed | chakraborty_d | 0 | 2 | Service Tools_Software | Console4 | NO | Closed | WIP22.2SubCom | IQon1 |  | WIP22.2 WIP23 | December 12, 2019 7:05:15 PM GMT+08:00
CTUDT00627362 | Refresh | Closed | shiraguppi_k | November 20, 2019 9:41:48 PM GMT+08:00 | Big Bore: BB Refresh Validation: Left Monitor Patient Registration Tab was Disabled After End Exam of CCT SCAN | BB_WIP11 |  |  | Subsystem |  | N/A | No Plan to Fix | navale_s | 0 | 3 |  | Console4 | NO | Closed | SubNC | Brilliance Big Bore
CTUDT00627346 | Refresh | Closed | Kevin_ZHAO | November 20, 2019 2:44:25 PM GMT+08:00 | Refresh ING/BR64 Test case modification Cycle 2 | WIP21 | DHF305263 Rev03 |  | Subsystem |  | None | Fixed | Kevin_ZHAO | 0 | 3 | Validation Documentation | Support-System | NO | Closed | 19-Dec-2019-SubCom | Ingenuity 64 |  | DHF305263 Rev03 | December 18, 2019 4:54:12 PM GMT+08:00
CTUDT00627344 | Refresh | Closed | wei_f | November 20, 2019 1:08:26 PM GMT+08:00 | Read the Chinese study from the DVD by running the application "RunDlg.exe" failed | WIP21 |  |  | Subsystem | Chinese Character; | NA | No Plan to Fix | s_sa | 0 | 3 |  | iCAP | NO | Closed |  | Ovation
CTUDT00627344 | Refresh | Closed | wei_f | November 20, 2019 1:08:26 PM GMT+08:00 | Read the Chinese study from the DVD by running the application "RunDlg.exe" failed | WIP21 |  |  | Subsystem | Chinese Character; | NA | No Plan to Fix | s_sa | 0 | 3 |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00627344 | Refresh | Closed | wei_f | November 20, 2019 1:08:26 PM GMT+08:00 | Read the Chinese study from the DVD by running the application "RunDlg.exe" failed | WIP21 |  |  | Subsystem | Chinese Character; | NA | No Plan to Fix | s_sa | 0 | 3 |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00627344 | Refresh | Closed | wei_f | November 20, 2019 1:08:26 PM GMT+08:00 | Read the Chinese study from the DVD by running the application "RunDlg.exe" failed | WIP21 |  |  | Subsystem | Chinese Character; | NA | No Plan to Fix | s_sa | 0 | 3 |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00627344 | Refresh | Closed | wei_f | November 20, 2019 1:08:26 PM GMT+08:00 | Read the Chinese study from the DVD by running the application "RunDlg.exe" failed | WIP21 |  |  | Subsystem | Chinese Character; | NA | No Plan to Fix | s_sa | 0 | 3 |  | iCAP | NO | Closed |  | IQon1
CTUDT00627335 | Refresh | Closed | lakshmikanth_s | November 19, 2019 7:39:56 PM GMT+08:00 | The ECG data files which when converted to .mdu(as per verification test case) contained Patient Information. | WIP11.1 | BB_WIP12.2 WIP22.2 |  | Subsystem | Service Tools, Blocker | N/A | Fixed | pandey_a | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP12_SegReq | Ovation |  | BB_WIP12.2 WIP22.2 | December 9, 2019 10:08:25 PM GMT+08:00
CTUDT00627335 | Refresh | Closed | lakshmikanth_s | November 19, 2019 7:39:56 PM GMT+08:00 | The ECG data files which when converted to .mdu(as per verification test case) contained Patient Information. | WIP11.1 | BB_WIP12.2 WIP22.2 |  | Subsystem | Service Tools, Blocker | N/A | Fixed | pandey_a | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP12_SegReq | Brilliance 64 |  | BB_WIP12.2 WIP22.2 | December 9, 2019 10:08:25 PM GMT+08:00
CTUDT00627335 | Refresh | Closed | lakshmikanth_s | November 19, 2019 7:39:56 PM GMT+08:00 | The ECG data files which when converted to .mdu(as per verification test case) contained Patient Information. | WIP11.1 | BB_WIP12.2 WIP22.2 |  | Subsystem | Service Tools, Blocker | N/A | Fixed | pandey_a | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP12_SegReq | Brilliance Big Bore |  | BB_WIP12.2 WIP22.2 | December 9, 2019 10:08:25 PM GMT+08:00
CTUDT00627335 | Refresh | Closed | lakshmikanth_s | November 19, 2019 7:39:56 PM GMT+08:00 | The ECG data files which when converted to .mdu(as per verification test case) contained Patient Information. | WIP11.1 | BB_WIP12.2 WIP22.2 |  | Subsystem | Service Tools, Blocker | N/A | Fixed | pandey_a | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP12_SegReq | Ingenuity 64 |  | BB_WIP12.2 WIP22.2 | December 9, 2019 10:08:25 PM GMT+08:00
CTUDT00627335 | Refresh | Closed | lakshmikanth_s | November 19, 2019 7:39:56 PM GMT+08:00 | The ECG data files which when converted to .mdu(as per verification test case) contained Patient Information. | WIP11.1 | BB_WIP12.2 WIP22.2 |  | Subsystem | Service Tools, Blocker | N/A | Fixed | pandey_a | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP12_SegReq | IQon1 |  | BB_WIP12.2 WIP22.2 | December 9, 2019 10:08:25 PM GMT+08:00
CTUDT00627333 | Refresh | Closed | shtark_o | November 19, 2019 6:57:41 PM GMT+08:00 | Phantom calibration stuck during the calibration in few of the cases | Ovation 4.1.7 CIRS Version -5.1.0.0204 | WIP22 | N/A | Subsystem | Blocker; Verified1 | CIRS.SSRS.6582 | Fixed | richards_b | 0 | 2 | CIRS_SW_ImgNodeSvr | CIRS | NO | Closed | WIP22SubCom | Ovation |  | WIP22 | November 25, 2019 9:08:15 PM GMT+08:00
CTUDT00627333 | Refresh | Closed | shtark_o | November 19, 2019 6:57:41 PM GMT+08:00 | Phantom calibration stuck during the calibration in few of the cases | Ovation 4.1.7 CIRS Version -5.1.0.0204 | WIP22 | N/A | Subsystem | Blocker; Verified1 | CIRS.SSRS.6582 | Fixed | richards_b | 0 | 2 | CIRS_SW_ImgNodeSvr | CIRS | NO | Closed | WIP22SubCom | Brilliance 64 |  | WIP22 | November 25, 2019 9:08:15 PM GMT+08:00
CTUDT00627333 | Refresh | Closed | shtark_o | November 19, 2019 6:57:41 PM GMT+08:00 | Phantom calibration stuck during the calibration in few of the cases | Ovation 4.1.7 CIRS Version -5.1.0.0204 | WIP22 | N/A | Subsystem | Blocker; Verified1 | CIRS.SSRS.6582 | Fixed | richards_b | 0 | 2 | CIRS_SW_ImgNodeSvr | CIRS | NO | Closed | WIP22SubCom | Brilliance Big Bore |  | WIP22 | November 25, 2019 9:08:15 PM GMT+08:00
CTUDT00627333 | Refresh | Closed | shtark_o | November 19, 2019 6:57:41 PM GMT+08:00 | Phantom calibration stuck during the calibration in few of the cases | Ovation 4.1.7 CIRS Version -5.1.0.0204 | WIP22 | N/A | Subsystem | Blocker; Verified1 | CIRS.SSRS.6582 | Fixed | richards_b | 0 | 2 | CIRS_SW_ImgNodeSvr | CIRS | NO | Closed | WIP22SubCom | Ingenuity 64 |  | WIP22 | November 25, 2019 9:08:15 PM GMT+08:00
CTUDT00627333 | Refresh | Closed | shtark_o | November 19, 2019 6:57:41 PM GMT+08:00 | Phantom calibration stuck during the calibration in few of the cases | Ovation 4.1.7 CIRS Version -5.1.0.0204 | WIP22 | N/A | Subsystem | Blocker; Verified1 | CIRS.SSRS.6582 | Fixed | richards_b | 0 | 2 | CIRS_SW_ImgNodeSvr | CIRS | NO | Closed | WIP22SubCom | IQon1 |  | WIP22 | November 25, 2019 9:08:15 PM GMT+08:00
CTUDT00627331 | Refresh | Closed | vinoliya_j | November 18, 2019 11:45:23 PM GMT+08:00 | Ingenuity User document : Issues found in 459801656031_b ACR guide | WIP19 | v4.1.10 | N/A | Subsystem |  | N/A | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | 6-Dec-2019 | Ingenuity 64 |  | v4.1.10 | December 9, 2019 10:32:53 AM GMT+08:00
CTUDT00627329 | Refresh | Closed | vinoliya_j | November 18, 2019 11:39:56 PM GMT+08:00 | Ingenuity User document : Issues found in 459801096624_b QRG | WIP19 | v4.1.10 | N/A | Subsystem |  | N/A | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | 6-Dec-2019 | Ingenuity 64 |  | v4.1.10 | December 9, 2019 10:32:30 AM GMT+08:00
CTUDT00627327 | Refresh | Closed | vinoliya_j | November 18, 2019 11:33:57 PM GMT+08:00 | Ingenuity User document : Issues found in 459801587101_b Exam cards book | WIP19 | v4.1.10 | N/A | Subsystem |  | N/A | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | 6-Dec-2019 | Ingenuity 64 |  | v4.1.10 | December 9, 2019 10:32:04 AM GMT+08:00
CTUDT00627325 | Refresh | Closed | vinoliya_j | November 18, 2019 11:27:42 PM GMT+08:00 | Ingenuity User document : Issues found in 459801063524_b Instructions for use | WIP19 | v4.1.10 | N/A | Subsystem |  | N/A | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | 6-Dec-2019 | Ingenuity 64 |  | v4.1.10 | December 9, 2019 10:31:42 AM GMT+08:00
CTUDT00627320 | Refresh | Closed | 320053770 | November 18, 2019 7:25:48 PM GMT+08:00 | Refresh – Wip21:IQon\ICT: Report editor tool crash after a few seconds. | WIP21 |  |  | Subsystem | EBW |  | Duplicate | s_sa | 0 |  |  | iCAP | NO | Closed |  | Ovation
CTUDT00627320 | Refresh | Closed | 320053770 | November 18, 2019 7:25:48 PM GMT+08:00 | Refresh – Wip21:IQon\ICT: Report editor tool crash after a few seconds. | WIP21 |  |  | Subsystem | EBW |  | Duplicate | s_sa | 0 |  |  | iCAP | NO | Closed |  | IQon1
CTUDT00627319 | Refresh | Closed | shaham_t | November 18, 2019 5:54:58 PM GMT+08:00 | IB: S_XSC_ERROR_CX_MAIN_CONTROLLER “Invalid data from RIX” error happened | 4.7.5 | WIP22 WIP23.1 |  | System |  | n/a | Fixed | elkarif_y | 0 | 3 | DMS Software | DMS | YES | Closed | WIP23.1SubCom | Ovation | PR#9371227 OneEMS #0114287279 | WIP22 WIP23.1 | January 9, 2020 5:17:02 PM GMT+08:00
CTUDT00627319 | Refresh | Closed | shaham_t | November 18, 2019 5:54:58 PM GMT+08:00 | IB: S_XSC_ERROR_CX_MAIN_CONTROLLER “Invalid data from RIX” error happened | 4.7.5 | WIP22 WIP23.1 |  | System |  | n/a | Fixed | elkarif_y | 0 | 3 | DMS Software | DMS | YES | Closed | WIP23.1SubCom | IQon1 | PR#9371227 OneEMS #0114287279 | WIP22 WIP23.1 | January 9, 2020 5:17:02 PM GMT+08:00
CTUDT00627316 | Refresh | Closed | burke_c | November 18, 2019 5:26:53 PM GMT+08:00 | After 10 acquisitions, the exam summary page places ... instead of series numbers | BB_WIP11 |  | NA | Subsystem |  | NA | No Plan to Fix | kr_m | 0 | 4 |  | Console4 | NO | Closed | SubNC | Ovation
CTUDT00627316 | Refresh | Closed | burke_c | November 18, 2019 5:26:53 PM GMT+08:00 | After 10 acquisitions, the exam summary page places ... instead of series numbers | BB_WIP11 |  | NA | Subsystem |  | NA | No Plan to Fix | kr_m | 0 | 4 |  | Console4 | NO | Closed | SubNC | Brilliance 64
CTUDT00627316 | Refresh | Closed | burke_c | November 18, 2019 5:26:53 PM GMT+08:00 | After 10 acquisitions, the exam summary page places ... instead of series numbers | BB_WIP11 |  | NA | Subsystem |  | NA | No Plan to Fix | kr_m | 0 | 4 |  | Console4 | NO | Closed | SubNC | Brilliance Big Bore
CTUDT00627316 | Refresh | Closed | burke_c | November 18, 2019 5:26:53 PM GMT+08:00 | After 10 acquisitions, the exam summary page places ... instead of series numbers | BB_WIP11 |  | NA | Subsystem |  | NA | No Plan to Fix | kr_m | 0 | 4 |  | Console4 | NO | Closed | SubNC | Ingenuity 64
CTUDT00627314 | Refresh | Closed | qiao_d | November 18, 2019 2:40:56 PM GMT+08:00 | iCT visual diagnostics X-Ray system test can not pass | WIP21 | WIP23.1 |  | Subsystem | Visual Diagnostic; Blocker | GCX.SSRS.iCT.13778 | Fixed | zhang_lawrence | 0 | 2 | Gantry Software | GCX | NO | Closed | WIP23.1SubCom | Ovation |  | WIP23.1 | January 6, 2020 1:52:55 PM GMT+08:00
CTUDT00627310 | Refresh | Closed | 320036209 | November 18, 2019 10:07:56 AM GMT+08:00 | Refresh Service Val_BB RT service validation comments | BB_WIP11 | Service Manual Update |  | Subsystem |  | NA | Fixed | 320036209 | 0 | 4 | Service Documentation | Support-Service | NO | Closed | SubCom-13-Dec-2019 | Brilliance Big Bore |  | Service Manual Update | December 19, 2019 3:57:06 PM GMT+08:00
CTUDT00627308 | Refresh | Closed | qiao_d | November 17, 2019 10:33:29 PM GMT+08:00 | ICT visual diagnostics | WIP22 |  |  | Subsystem |  |  | Duplicate | qiao_d | 0 |  |  |  |  | Closed |  | iCT4
CTUDT00627306 | Refresh | Closed | Valeria_Kashtanov | November 17, 2019 10:03:39 PM GMT+08:00 | Refresh: Acceptance Test: Scan parameters are not according to spec | WIP21 | DHF170395 Rev12 | N/A | Subsystem |  | DHF170395 | Fixed | richards_b | 0 | 2 | Design Documentation | Support-Physics | NO | Closed | SpecUpdate | IQon1 |  | DHF170395 Rev12 | January 16, 2020 8:35:41 PM GMT+08:00
CTUDT00627301 | Refresh | Closed | cohen_y | November 17, 2019 8:25:55 PM GMT+08:00 | Refresh: Calibrations: Constancy Test: Harmonized: Uniformity limit in case 3 Adult Body (on Head) is not according to Spec | WIP21 | WIP22 | N/A | Subsystem | Constancy | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP22SubCom | Ovation |  | WIP22 | November 19, 2019 3:12:53 PM GMT+08:00
CTUDT00627298 | Refresh | Closed | pick_a | November 17, 2019 6:12:30 PM GMT+08:00 | Wrong version of TightVNC is specified in Ingenuity TRG | WIP22 | v4.1.10 | N/A | Subsystem | IFU | N/A | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | 6-Dec-2019 | Ingenuity 64 |  | v4.1.10 | December 9, 2019 10:31:21 AM GMT+08:00
CTUDT00627296 | Refresh | Closed | pick_a | November 17, 2019 5:59:07 PM GMT+08:00 | Wrong version of TightVNC is specified in TRG | WIP22 | v4.8 | N/A | Subsystem |  | N/A | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | 6-Dec-2019 | Brilliance Big Bore |  | v4.8 | December 9, 2019 10:30:59 AM GMT+08:00
CTUDT00627296 | Refresh | Closed | pick_a | November 17, 2019 5:59:07 PM GMT+08:00 | Wrong version of TightVNC is specified in TRG | WIP22 | v4.8 | N/A | Subsystem |  | N/A | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | 6-Dec-2019 | IQon1 |  | v4.8 | December 9, 2019 10:30:59 AM GMT+08:00
CTUDT00627294 | Refresh | Closed | pick_a | November 17, 2019 5:32:40 PM GMT+08:00 | Password is specified in installation manual | WIP22 |  |  | System |  | NA | Does Not Apply | shi_b | 0 |  |  | Support-Service | NO | Closed |  | Brilliance 64
CTUDT00627294 | Refresh | Closed | pick_a | November 17, 2019 5:32:40 PM GMT+08:00 | Password is specified in installation manual | WIP22 |  |  | System |  | NA | Does Not Apply | shi_b | 0 |  |  | Support-Service | NO | Closed |  | Ingenuity 64
CTUDT00627291 | Refresh | Closed | shi_b | November 15, 2019 10:29:59 PM GMT+08:00 | Refresh Service Val_3D Modulation & Z-Modulation test failed | BB_WIP11 |  |  | Subsystem | CIRS, Blocker | NA | Duplicate | lin_s | 0 |  |  | Support-System | NO | Closed |  | Brilliance Big Bore
CTUDT00627287 | Refresh | Closed | vinoliya_j | November 15, 2019 6:26:02 PM GMT+08:00 | 459801656431_B - BigBore ACR guide - Validation Comments | BB_WIP9 | E-121408 | N/A | Subsystem |  | N/A | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | 6-Dec-2019 | Brilliance Big Bore |  | E-121408 | December 9, 2019 10:29:23 AM GMT+08:00
CTUDT00627285 | Refresh | Closed | vinoliya_j | November 15, 2019 6:15:34 PM GMT+08:00 | 459800863014_B - BigBore IFU - Validation Comments | BB_WIP9 | E-121408 | N/A | Subsystem |  | N/A | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | 6-Dec-2019 | Brilliance Big Bore |  | E-121408 | December 9, 2019 10:28:52 AM GMT+08:00
CTUDT00627283 | Refresh | Closed | 320053770 | November 15, 2019 5:51:17 PM GMT+08:00 | Refresh – Wip21:CTV/SCTV viewer cannot open data from remote stations\displays "white snow" upon loading. | WIP21 | WIP23 |  | Subsystem | Test Design | NA | Fixed | tamari_o | 0 | 4 | Test Design | Support-System | NO | Closed | SubCom-09-Jan-2020 | Ovation |  | WIP23 | January 21, 2020 11:05:06 AM GMT+08:00
CTUDT00627283 | Refresh | Closed | 320053770 | November 15, 2019 5:51:17 PM GMT+08:00 | Refresh – Wip21:CTV/SCTV viewer cannot open data from remote stations\displays "white snow" upon loading. | WIP21 | WIP23 |  | Subsystem | Test Design | NA | Fixed | tamari_o | 0 | 4 | Test Design | Support-System | NO | Closed | SubCom-09-Jan-2020 | IQon1 |  | WIP23 | January 21, 2020 11:05:06 AM GMT+08:00
CTUDT00627280 | Refresh | Closed | vinoliya_j | November 15, 2019 4:55:21 PM GMT+08:00 | 459801594951_B - BigBore Exam Cards Book - Validation comments | BB_WIP9 | E-121408 | N/A | Subsystem |  | N/A | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | 6-Dec-2019 | Brilliance Big Bore |  | E-121408 | December 9, 2019 10:27:37 AM GMT+08:00
CTUDT00627273 | Refresh | Closed | cai_t | November 15, 2019 3:49:35 PM GMT+08:00 | Refresh:IVC: Go to inner laser button did not work on HW IVC box after Pressing Go to inner laser button on Gantry control panel | WIP11 | BB_WIP12 |  | Subsystem |  | Console-SRS.17809, Console-SRS.17810 | Fixed | 320021835 | 0 | 2 | IVC Software | Console4 | NO | Closed | BB_WIP12_SubCom | Brilliance Big Bore |  | BB_WIP12 | November 15, 2019 8:47:35 PM GMT+08:00
CTUDT00627271 | Refresh | Closed | zhang_t | November 15, 2019 3:18:07 PM GMT+08:00 | Refresh: CCT Fluoro - Status monitor was not correct, it displayed Scanning while Xray OFF | BB WIP11.1 |  |  | Subsystem | Blocker |  | Duplicate | kr_m | 0 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00627269 | Refresh | Closed | wang_l | November 15, 2019 2:03:57 PM GMT+08:00 | Refresh Reliability - Prompt error message of 'A problem has occurred' during Automatic reliability testing. | BB_WIP11 |  | NA | Subsystem | GCX, Blocker | NA | Duplicate | tan_j | 0 |  |  | GCX | YES | Closed |  | Brilliance Big Bore
CTUDT00627266 | Refresh | Closed | 320036209 | November 15, 2019 11:09:41 AM GMT+08:00 | phantom is not detected | BB_WIP11 |  |  | Subsystem |  |  | Does Not Apply | shen_j | 0 |  |  |  |  | Closed |  | Brilliance Big Bore
CTUDT00627264 | Refresh | Closed | cohen_y | November 14, 2019 11:22:58 PM GMT+08:00 | Refresh: Calibrations: FEE Test: Slice Profile Edge Flatness limit is not as defined in Spec | WIP21 | WIP22 |  | Subsystem | FEE test | n/a | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP22SubCom | Ovation |  | WIP22 | November 19, 2019 3:12:43 PM GMT+08:00
CTUDT00627262 | Refresh | Closed | shi_b | November 14, 2019 10:41:41 PM GMT+08:00 | Refresh Service Val_Remote connect to S2 Server via PRS is failed due to there is no ReconRD_S2 | BB_WIP11 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | s_shabeer | 0 |  |  | Service Tools | NO | Closed | WAS | Brilliance 64
CTUDT00627262 | Refresh | Closed | shi_b | November 14, 2019 10:41:41 PM GMT+08:00 | Refresh Service Val_Remote connect to S2 Server via PRS is failed due to there is no ReconRD_S2 | BB_WIP11 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | s_shabeer | 0 |  |  | Service Tools | NO | Closed | WAS | Brilliance Big Bore
CTUDT00627262 | Refresh | Closed | shi_b | November 14, 2019 10:41:41 PM GMT+08:00 | Refresh Service Val_Remote connect to S2 Server via PRS is failed due to there is no ReconRD_S2 | BB_WIP11 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | s_shabeer | 0 |  |  | Service Tools | NO | Closed | WAS | Ingenuity 64
CTUDT00627259 | Refresh | Closed | kapadia_v | November 14, 2019 7:00:24 PM GMT+08:00 | Result engine crashes, when CIRS sends duplicate images to console host. | WIP20 | WIP22 | NA | Subsystem |  | NA | Fixed | kapadia_v | 0 | 4 | Console Software | Console4 | NO | Closed | WIP22SegReq | IQon1 |  | WIP22 | November 15, 2019 8:59:23 PM GMT+08:00
CTUDT00627256 | Refresh | Closed | dongping_w | November 14, 2019 2:23:47 PM GMT+08:00 | Prompt error message of 'A problem has occurred' during Automatic reliability testing. | BB_WIP11 |  |  | Subsystem |  |  | Duplicate | tan_j | 0 |  |  | GCX | NO | Closed |  | Brilliance Big Bore
CTUDT00627254 | Refresh | Closed | wang_l | November 14, 2019 2:11:36 PM GMT+08:00 | Refresh Reliability - Refresh Reliability - "Gantry cannot initialize.." during clinical scan | WIP21 |  | The user can acknowledge the error and re-submit the scan request. | Subsystem | DMS, Blocker | N/A | No Plan to Fix | richards_b | 0 | 4 |  | CIRS | YES | Closed |  | Ingenuity 64
CTUDT00627251 | Refresh | Closed | wang_l | November 14, 2019 2:00:22 PM GMT+08:00 | Refresh Reliability - "Gantry Acquisition cannot comply" during clinical scan | BB_WIP11 |  |  | Subsystem | CIRS; | NA | Duplicate | yeshan_m | 1 |  |  | Support-System | NO | Closed |  | Brilliance Big Bore
CTUDT00627249 | Refresh | Closed | 320021835 | November 14, 2019 1:28:32 PM GMT+08:00 | The disabled look of increments list is not according to the UI specification | BB_WIP11 | BB_WIP12 | Case 1. None Case 2. None. | Subsystem |  | NA | Fixed | 320021835 | 0 | 3 | IVC Software | Console4 | NO | Closed | BB_WIP12_SubCom | Brilliance Big Bore |  | BB_WIP12 | November 15, 2019 7:43:16 PM GMT+08:00
CTUDT00627245 | Refresh | Closed | tamari_o | November 13, 2019 10:52:21 PM GMT+08:00 | Refresh remove the couch diagnostic tool in Refresh for iCT and IQon | WIP21 | DHF298553 Rev.09 DHF298556 Rev.10 |  | Subsystem |  | CT-NM.SRS.Service.2050 and CT-NM.PRS.CSR.282 | Fixed | bandil_j | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | SubCom-19-Dec-2019 | IQon1 |  | DHF298553 Rev.09 DHF298556 Rev.10 | December 22, 2019 10:11:33 PM GMT+08:00
CTUDT00627243 | Refresh | Closed | manju_kolkar | November 13, 2019 10:35:05 PM GMT+08:00 | Message was not localized in selected Non-English language | WIP11 | BB_WIP12 |  | Subsystem |  | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP12_SubCom | Brilliance Big Bore |  | BB_WIP12 | November 14, 2019 7:25:16 PM GMT+08:00
CTUDT00627239 | Refresh | Closed | Valeria_Kashtanov | November 13, 2019 8:19:47 PM GMT+08:00 | Refresh: Acceptance Test: Slice Thickness limits are not according to spec | WIP21 | WIP22 | N/A | Subsystem | Acceptance test | N/A | Fixed | elviral | 0 | 2 | DMS Software | DMS | NO | Closed | WIP22SubCom | IQon1 |  | WIP22 | November 19, 2019 3:10:01 PM GMT+08:00
CTUDT00627237 | Refresh | Closed | Sharmila_Vasireddy | November 13, 2019 6:24:02 PM GMT+08:00 | Not able to verify below reserved keywords in console | WIP21 | DHF303211 Rev19 | N/A | Subsystem |  | Console-SRS.16125, Console-SRS.15977 | Fixed | roy_r | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | SSRS Update | Ovation |  | DHF303211 Rev19 | November 14, 2019 7:20:52 PM GMT+08:00
CTUDT00627237 | Refresh | Closed | Sharmila_Vasireddy | November 13, 2019 6:24:02 PM GMT+08:00 | Not able to verify below reserved keywords in console | WIP21 | DHF303211 Rev19 | N/A | Subsystem |  | Console-SRS.16125, Console-SRS.15977 | Fixed | roy_r | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | SSRS Update | Brilliance 64 |  | DHF303211 Rev19 | November 14, 2019 7:20:52 PM GMT+08:00
CTUDT00627237 | Refresh | Closed | Sharmila_Vasireddy | November 13, 2019 6:24:02 PM GMT+08:00 | Not able to verify below reserved keywords in console | WIP21 | DHF303211 Rev19 | N/A | Subsystem |  | Console-SRS.16125, Console-SRS.15977 | Fixed | roy_r | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | SSRS Update | Ingenuity 64 |  | DHF303211 Rev19 | November 14, 2019 7:20:52 PM GMT+08:00
CTUDT00627237 | Refresh | Closed | Sharmila_Vasireddy | November 13, 2019 6:24:02 PM GMT+08:00 | Not able to verify below reserved keywords in console | WIP21 | DHF303211 Rev19 | N/A | Subsystem |  | Console-SRS.16125, Console-SRS.15977 | Fixed | roy_r | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | SSRS Update | IQon1 |  | DHF303211 Rev19 | November 14, 2019 7:20:52 PM GMT+08:00
CTUDT00627232 | Refresh | Closed | 320036209 | November 13, 2019 12:10:28 PM GMT+08:00 | Service manual comments/redlines for BB RT service validation | BB_WIP11 |  |  | System |  | NA | Duplicate | shi_b | 0 |  |  | Support-Service | NO | Closed |  | Brilliance Big Bore
CTUDT00627230 | Refresh | Closed | wang_l | November 13, 2019 11:01:03 AM GMT+08:00 | Refresh Reiability - Estop opened during scanning | BB_WIP11 |  |  | Subsystem | Gantry | NA | No Plan to Fix | wang_k | 0 | 3 |  | GCX | YES | Closed |  | Brilliance Big Bore
CTUDT00627227 | Refresh | Closed | dannyb | November 12, 2019 9:45:19 PM GMT+08:00 | Refresh: Beat is being skipped when performing S&S at 35 bpm. | WIP21 | DHF298553 Rev.10 DHF298556 Rev.11 DHF298293 Rev.10 DHF298290 Rev.09 |  | Subsystem | Test Design | CT-NM.PRS.CRS.6430; CT-NM.SRS.Modes.1427 | Fixed | shaham_t | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | SubCom-15-Feb-2020 | Ovation |  | DHF298553 Rev.10 DHF298556 Rev.11 DHF298293 Rev.10 DHF298290 Rev.09 | February 25, 2020 6:25:11 PM GMT+08:00
CTUDT00627221 | Refresh | Closed | golzban_v | November 12, 2019 7:33:03 PM GMT+08:00 | Perfusion: The system allows planning any number of cycles/groups | WIP20 |  |  | Subsystem | Perfusion | N/A | Works As Specified / Not A Defect | guy_sheaffer | 0 |  |  | Support-System | NO | Closed | WAS | iCT4
CTUDT00627221 | Refresh | Closed | golzban_v | November 12, 2019 7:33:03 PM GMT+08:00 | Perfusion: The system allows planning any number of cycles/groups | WIP20 |  |  | Subsystem | Perfusion | N/A | Works As Specified / Not A Defect | guy_sheaffer | 0 |  |  | Support-System | NO | Closed | WAS | IQon1
CTUDT00627219 | Refresh | Closed | mN_m | November 12, 2019 6:13:48 PM GMT+08:00 | Console installation exceeds 90 minutes. | WIP11.1 |  | Disconnect the HOST computer from the hospital LAN before installation | Subsystem | Test Design | NA | Works As Specified / Not A Defect | panchami_pradeep | 0 | 3 |  | Console4 | NO | Closed | WAS | Ovation
CTUDT00627219 | Refresh | Closed | mN_m | November 12, 2019 6:13:48 PM GMT+08:00 | Console installation exceeds 90 minutes. | WIP11.1 |  | Disconnect the HOST computer from the hospital LAN before installation | Subsystem | Test Design | NA | Works As Specified / Not A Defect | panchami_pradeep | 0 | 3 |  | Console4 | NO | Closed | WAS | Brilliance 64
CTUDT00627219 | Refresh | Closed | mN_m | November 12, 2019 6:13:48 PM GMT+08:00 | Console installation exceeds 90 minutes. | WIP11.1 |  | Disconnect the HOST computer from the hospital LAN before installation | Subsystem | Test Design | NA | Works As Specified / Not A Defect | panchami_pradeep | 0 | 3 |  | Console4 | NO | Closed | WAS | Brilliance Big Bore
CTUDT00627219 | Refresh | Closed | mN_m | November 12, 2019 6:13:48 PM GMT+08:00 | Console installation exceeds 90 minutes. | WIP11.1 |  | Disconnect the HOST computer from the hospital LAN before installation | Subsystem | Test Design | NA | Works As Specified / Not A Defect | panchami_pradeep | 0 | 3 |  | Console4 | NO | Closed | WAS | Ingenuity 64
CTUDT00627219 | Refresh | Closed | mN_m | November 12, 2019 6:13:48 PM GMT+08:00 | Console installation exceeds 90 minutes. | WIP11.1 |  | Disconnect the HOST computer from the hospital LAN before installation | Subsystem | Test Design | NA | Works As Specified / Not A Defect | panchami_pradeep | 0 | 3 |  | Console4 | NO | Closed | WAS | IQon1
CTUDT00627217 | Refresh | Closed | Jose_D_Sergio_Almeida | November 12, 2019 6:05:25 PM GMT+08:00 | The waveform is skipped but scan showed as radiated in red color in ECG waveform for Coronary S&S | WIP19 |  | NA | Subsystem |  | NA | Works As Specified / Not A Defect | tan_j | 0 |  |  | GCX | NO | Closed |  | Ingenuity 64 | NA
CTUDT00627214 | Refresh | Closed | Sateesh_S | November 12, 2019 5:38:30 PM GMT+08:00 | Actual SSDE is not calculating correctly for Cardiac scan with DoseRight enabled | WIP11 | BB_WIP12.2 WIP22.2 |  | Subsystem | SSDE | N/A | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | NO | Closed | WIP22SubCom | Ovation |  | BB_WIP12.2 WIP22.2 | December 2, 2019 8:32:39 PM GMT+08:00
CTUDT00627214 | Refresh | Closed | Sateesh_S | November 12, 2019 5:38:30 PM GMT+08:00 | Actual SSDE is not calculating correctly for Cardiac scan with DoseRight enabled | WIP11 | BB_WIP12.2 WIP22.2 |  | Subsystem | SSDE | N/A | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | NO | Closed | WIP22SubCom | Brilliance 64 |  | BB_WIP12.2 WIP22.2 | December 2, 2019 8:32:39 PM GMT+08:00
CTUDT00627214 | Refresh | Closed | Sateesh_S | November 12, 2019 5:38:30 PM GMT+08:00 | Actual SSDE is not calculating correctly for Cardiac scan with DoseRight enabled | WIP11 | BB_WIP12.2 WIP22.2 |  | Subsystem | SSDE | N/A | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | NO | Closed | WIP22SubCom | Brilliance Big Bore |  | BB_WIP12.2 WIP22.2 | December 2, 2019 8:32:39 PM GMT+08:00
CTUDT00627214 | Refresh | Closed | Sateesh_S | November 12, 2019 5:38:30 PM GMT+08:00 | Actual SSDE is not calculating correctly for Cardiac scan with DoseRight enabled | WIP11 | BB_WIP12.2 WIP22.2 |  | Subsystem | SSDE | N/A | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | NO | Closed | WIP22SubCom | Ingenuity 64 |  | BB_WIP12.2 WIP22.2 | December 2, 2019 8:32:39 PM GMT+08:00
CTUDT00627214 | Refresh | Closed | Sateesh_S | November 12, 2019 5:38:30 PM GMT+08:00 | Actual SSDE is not calculating correctly for Cardiac scan with DoseRight enabled | WIP11 | BB_WIP12.2 WIP22.2 |  | Subsystem | SSDE | N/A | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | NO | Closed | WIP22SubCom | IQon1 |  | BB_WIP12.2 WIP22.2 | December 2, 2019 8:32:39 PM GMT+08:00
CTUDT00627212 | Refresh | Closed | panchami_pradeep | November 12, 2019 4:54:45 PM GMT+08:00 | Refresh: System error for S&S scan with largest spectral result when Console does not send phase points to CIRS. | WIP 21 | WIP22 | Set conventional result larger or equal to spectral one | Subsystem |  | Console-SSRS.17861 | Fixed | bhawna_g | 0 | 4 | Console Software | Console4 | NO | Closed | WIP22SubCom | IQon1 |  | WIP22 | November 15, 2019 9:24:09 PM GMT+08:00
CTUDT00627209 | Refresh | Closed | zhu_je | November 12, 2019 4:29:38 PM GMT+08:00 | [Refresh] Recording missing with repeating “Add Silence Before/Remove Silence Before” operation many times. | WIP21 |  |  | Subsystem | Voice Manager | N/A | No Plan to Fix | navale_s | 0 | 3 |  | Console4 | NO | Closed | SubNC | Ovation
CTUDT00627209 | Refresh | Closed | zhu_je | November 12, 2019 4:29:38 PM GMT+08:00 | [Refresh] Recording missing with repeating “Add Silence Before/Remove Silence Before” operation many times. | WIP21 |  |  | Subsystem | Voice Manager | N/A | No Plan to Fix | navale_s | 0 | 3 |  | Console4 | NO | Closed | SubNC | Brilliance Big Bore
CTUDT00627209 | Refresh | Closed | zhu_je | November 12, 2019 4:29:38 PM GMT+08:00 | [Refresh] Recording missing with repeating “Add Silence Before/Remove Silence Before” operation many times. | WIP21 |  |  | Subsystem | Voice Manager | N/A | No Plan to Fix | navale_s | 0 | 3 |  | Console4 | NO | Closed | SubNC | Ingenuity 64
CTUDT00627209 | Refresh | Closed | zhu_je | November 12, 2019 4:29:38 PM GMT+08:00 | [Refresh] Recording missing with repeating “Add Silence Before/Remove Silence Before” operation many times. | WIP21 |  |  | Subsystem | Voice Manager | N/A | No Plan to Fix | navale_s | 0 | 3 |  | Console4 | NO | Closed | SubNC | IQon1
CTUDT00627208 | Refresh | Closed | chakraborty_d | November 12, 2019 4:26:00 PM GMT+08:00 | Strings "Encrypted" and "Decrypted" are not translated when the regional locale of the system is changed to languages other than English. | WIP21.1 |  |  | Subsystem |  | NA | No Plan to Fix | chakraborty_d | 0 | 4 |  | Console4 | NO | Closed | SubNC | Ovation
CTUDT00627208 | Refresh | Closed | chakraborty_d | November 12, 2019 4:26:00 PM GMT+08:00 | Strings "Encrypted" and "Decrypted" are not translated when the regional locale of the system is changed to languages other than English. | WIP21.1 |  |  | Subsystem |  | NA | No Plan to Fix | chakraborty_d | 0 | 4 |  | Console4 | NO | Closed | SubNC | Brilliance 64
CTUDT00627208 | Refresh | Closed | chakraborty_d | November 12, 2019 4:26:00 PM GMT+08:00 | Strings "Encrypted" and "Decrypted" are not translated when the regional locale of the system is changed to languages other than English. | WIP21.1 |  |  | Subsystem |  | NA | No Plan to Fix | chakraborty_d | 0 | 4 |  | Console4 | NO | Closed | SubNC | Brilliance Big Bore
CTUDT00627208 | Refresh | Closed | chakraborty_d | November 12, 2019 4:26:00 PM GMT+08:00 | Strings "Encrypted" and "Decrypted" are not translated when the regional locale of the system is changed to languages other than English. | WIP21.1 |  |  | Subsystem |  | NA | No Plan to Fix | chakraborty_d | 0 | 4 |  | Console4 | NO | Closed | SubNC | Ingenuity 64
CTUDT00627208 | Refresh | Closed | chakraborty_d | November 12, 2019 4:26:00 PM GMT+08:00 | Strings "Encrypted" and "Decrypted" are not translated when the regional locale of the system is changed to languages other than English. | WIP21.1 |  |  | Subsystem |  | NA | No Plan to Fix | chakraborty_d | 0 | 4 |  | Console4 | NO | Closed | SubNC | IQon1
CTUDT00627206 | Refresh | Closed | l_yiming | November 12, 2019 3:38:49 PM GMT+08:00 | Refresh Ingenuity: Previous recon images were not displayed in view2 | WIP19 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | kr_m | 0 |  |  | Console4 | NO | Closed | WAS | Ingenuity 64
CTUDT00627203 | Refresh | Closed | l_yiming | November 12, 2019 1:53:28 PM GMT+08:00 | Refresh Ingenuity Validation: PACS connect issue | WIP19 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | TEST_tl_2 | 0 |  |  |  |  | Closed |  | Ingenuity 64
CTUDT00627200 | Refresh | Closed | Kevin_ZHAO | November 11, 2019 10:02:15 PM GMT+08:00 | Refresh Ingenuity Validation: HisRis connect issue | WIP19 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | Kevin_ZHAO | 0 |  |  | Support-System | NO | Closed |  | Ovation
CTUDT00627200 | Refresh | Closed | Kevin_ZHAO | November 11, 2019 10:02:15 PM GMT+08:00 | Refresh Ingenuity Validation: HisRis connect issue | WIP19 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | Kevin_ZHAO | 0 |  |  | Support-System | NO | Closed |  | Brilliance 64
CTUDT00627200 | Refresh | Closed | Kevin_ZHAO | November 11, 2019 10:02:15 PM GMT+08:00 | Refresh Ingenuity Validation: HisRis connect issue | WIP19 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | Kevin_ZHAO | 0 |  |  | Support-System | NO | Closed |  | Brilliance Big Bore
CTUDT00627200 | Refresh | Closed | Kevin_ZHAO | November 11, 2019 10:02:15 PM GMT+08:00 | Refresh Ingenuity Validation: HisRis connect issue | WIP19 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | Kevin_ZHAO | 0 |  |  | Support-System | NO | Closed |  | Ingenuity 64
CTUDT00627200 | Refresh | Closed | Kevin_ZHAO | November 11, 2019 10:02:15 PM GMT+08:00 | Refresh Ingenuity Validation: HisRis connect issue | WIP19 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | Kevin_ZHAO | 0 |  |  | Support-System | NO | Closed |  | IQon1
CTUDT00627191 | Refresh | Closed | levy_s | November 10, 2019 7:46:28 PM GMT+08:00 | Refresh: Phase Roll back mechanism is not consistent | WIP21 |  |  | Subsystem |  |  | Not Enough Information | levy_s | 0 |  |  |  |  | Closed |  | Ovation
CTUDT00627189 | Refresh | Closed | levy_s | November 10, 2019 7:42:58 PM GMT+08:00 | Refresh: Phase tolerance doesn’t appear on the online ECG | WIP21 |  |  | Subsystem | CRN | N/A | No Plan to Fix | yang_h | 0 | 3 |  | GCX | NO | Closed |  | Ovation
CTUDT00627187 | Refresh | Closed | levy_s | November 10, 2019 7:40:12 PM GMT+08:00 | Refresh: Recon box (CIRS) cannot initialize error message when trying to scan the clinical scan | WIP21 |  | If the user tries to start an acquisition and it fails due to CIRS not being able to initialize, then the user will need to re-try again. If that fails, then the user may need to re-boot the system. | Subsystem | CIRS initialization | N/A | Works As Specified / Not A Defect | richards_b | 0 |  |  | CIRS | NO | Closed |  | IQon1
CTUDT00627185 | Refresh | Closed | levy_s | November 10, 2019 7:37:29 PM GMT+08:00 | Refresh: Initialization time is long | WIP21 |  | NA | Subsystem |  | NA | No Plan to Fix | kr_m | 0 | 4 |  | Console4 | NO | Closed | SubNC | IQon1
CTUDT00627183 | Refresh | Resolved | levy_s | November 10, 2019 7:31:56 PM GMT+08:00 | Refresh: The plan boxes gets stuck when planning around 20 results | WIP21 | WIP35 | Plan only SBI and thin axials during scan planning and all other results produce offline | Subsystem | Performance; SP1_PM_Approved | NA | Fixed | veeraswamy_k | 0 | 2 | Console Software | Console4 | NO | Resolved | WIP35SubCom | Ovation |  | WIP35 | November 27, 2020 1:03:38 AM GMT+08:00
CTUDT00627183 | Refresh | Resolved | levy_s | November 10, 2019 7:31:56 PM GMT+08:00 | Refresh: The plan boxes gets stuck when planning around 20 results | WIP21 | WIP35 | Plan only SBI and thin axials during scan planning and all other results produce offline | Subsystem | Performance; SP1_PM_Approved | NA | Fixed | veeraswamy_k | 0 | 2 | Console Software | Console4 | NO | Resolved | WIP35SubCom | Brilliance 64 |  | WIP35 | November 27, 2020 1:03:38 AM GMT+08:00
CTUDT00627183 | Refresh | Resolved | levy_s | November 10, 2019 7:31:56 PM GMT+08:00 | Refresh: The plan boxes gets stuck when planning around 20 results | WIP21 | WIP35 | Plan only SBI and thin axials during scan planning and all other results produce offline | Subsystem | Performance; SP1_PM_Approved | NA | Fixed | veeraswamy_k | 0 | 2 | Console Software | Console4 | NO | Resolved | WIP35SubCom | Brilliance Big Bore |  | WIP35 | November 27, 2020 1:03:38 AM GMT+08:00
CTUDT00627183 | Refresh | Resolved | levy_s | November 10, 2019 7:31:56 PM GMT+08:00 | Refresh: The plan boxes gets stuck when planning around 20 results | WIP21 | WIP35 | Plan only SBI and thin axials during scan planning and all other results produce offline | Subsystem | Performance; SP1_PM_Approved | NA | Fixed | veeraswamy_k | 0 | 2 | Console Software | Console4 | NO | Resolved | WIP35SubCom | Ingenuity 64 |  | WIP35 | November 27, 2020 1:03:38 AM GMT+08:00
CTUDT00627183 | Refresh | Resolved | levy_s | November 10, 2019 7:31:56 PM GMT+08:00 | Refresh: The plan boxes gets stuck when planning around 20 results | WIP21 | WIP35 | Plan only SBI and thin axials during scan planning and all other results produce offline | Subsystem | Performance; SP1_PM_Approved | NA | Fixed | veeraswamy_k | 0 | 2 | Console Software | Console4 | NO | Resolved | WIP35SubCom | IQon1 |  | WIP35 | November 27, 2020 1:03:38 AM GMT+08:00
CTUDT00627178 | Refresh | Closed | barnea_d | November 10, 2019 1:18:42 PM GMT+08:00 | Refresh: WIP 21: iCT: Calibrations: Performance Test: Bed height is not properly positioned | WIP21 | WIP22 |  | Subsystem | Calibrations, Verified1 | n/a | Fixed | frider_o | 0 | 3 | DMS Software | DMS | NO | Closed | WIP22SubCom | Ovation |  | WIP22 | November 12, 2019 11:55:28 PM GMT+08:00
CTUDT00627175 | Refresh | Closed | cherylli | November 10, 2019 1:06:05 AM GMT+08:00 | Refresh Val: water HU has no accuracy of 0 +6.5/-1.5 in IQ body test with system phantom | BB_WIP9 |  | Run HCOR and re-perform IQ check | Subsystem | IQ Check | N/A | Works As Specified / Not A Defect | cherylli | 1 |  |  | Support-System | NO | Closed |  | Brilliance Big Bore
CTUDT00627173 | Refresh | Closed | Kevin_ZHAO | November 10, 2019 12:04:24 AM GMT+08:00 | BB Refresh validation:Calcium Scoring: it is possible to choose a recon with a phase different than scanned | BB_WIP9 |  |  | Subsystem |  |  | Duplicate | pandey_a | 0 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00627171 | Refresh | Closed | Kevin_ZHAO | November 10, 2019 12:01:45 AM GMT+08:00 | BB Refresh validation: Cardiac Phases Recon error | BB_WIP9 |  |  | Subsystem |  |  | Duplicate | singh_j | 0 |  |  |  |  | Closed |  | Brilliance Big Bore
CTUDT00627169 | Refresh | Closed | Kevin_ZHAO | November 9, 2019 11:35:13 PM GMT+08:00 | BB Refresh validation: Critical error after closing funtional CT | BB_WIP9 |  |  | Subsystem | EBW |  | Works As Specified / Not A Defect | aswani_v | 0 |  |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00627164 | Refresh | Closed | chang_zhiguo | November 9, 2019 1:20:46 PM GMT+08:00 | BB refresh validation: The scan bar minimize button and Extend button interlapped | BB_WIP9 |  |  | Subsystem |  |  | Change Proposal | chang_zhiguo | 0 |  |  |  |  | Closed |  | Brilliance Big Bore
CTUDT00627160 | Refresh | Closed | chang_zhiguo | November 9, 2019 12:44:45 PM GMT+08:00 | BB refresh validation: pop up error message of “Quickreview.exe stop working.” | BB_WIP9 |  |  | Subsystem | EBW |  | Works As Specified / Not A Defect | aswani_v | 1 |  |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00627158 | Refresh | Closed | Ashish_Sharma | November 9, 2019 6:03:21 AM GMT+08:00 | Ingenuity Validation: In axial sinus scan ,slab, volume ,endo tabs were inactive after loading added/reconed result during acquisition to CT viewer. | WIP19 |  | Full volume of images are available if the increment is set as contiguous. Slab, Volume and Endo will be active if volume of images are available. | Subsystem | Test Design | None | Works As Specified / Not A Defect | Ashish_Sharma | 0 | 3 |  | Support-System | NO | Closed | WAS | Ingenuity 64
CTUDT00627156 | Refresh | Closed | Ashish_Sharma | November 9, 2019 5:59:07 AM GMT+08:00 | In axial sinus scan ,on loading data to recon tab for doing reconstruction ,plan box cannot be reduced. | WIP19 |  |  | Subsystem | Scan Planning | NA | Works As Specified / Not A Defect | bhawna_g | 0 |  |  | Console4 | NO | Closed | WAS | Ovation
CTUDT00627156 | Refresh | Closed | Ashish_Sharma | November 9, 2019 5:59:07 AM GMT+08:00 | In axial sinus scan ,on loading data to recon tab for doing reconstruction ,plan box cannot be reduced. | WIP19 |  |  | Subsystem | Scan Planning | NA | Works As Specified / Not A Defect | bhawna_g | 0 |  |  | Console4 | NO | Closed | WAS | Brilliance 64
CTUDT00627156 | Refresh | Closed | Ashish_Sharma | November 9, 2019 5:59:07 AM GMT+08:00 | In axial sinus scan ,on loading data to recon tab for doing reconstruction ,plan box cannot be reduced. | WIP19 |  |  | Subsystem | Scan Planning | NA | Works As Specified / Not A Defect | bhawna_g | 0 |  |  | Console4 | NO | Closed | WAS | Ingenuity 64
CTUDT00627156 | Refresh | Closed | Ashish_Sharma | November 9, 2019 5:59:07 AM GMT+08:00 | In axial sinus scan ,on loading data to recon tab for doing reconstruction ,plan box cannot be reduced. | WIP19 |  |  | Subsystem | Scan Planning | NA | Works As Specified / Not A Defect | bhawna_g | 0 |  |  | Console4 | NO | Closed | WAS | IQon1
CTUDT00627154 | Refresh | Closed | Ashish_Sharma | November 9, 2019 5:51:31 AM GMT+08:00 | Scan time inactive on GUI for pediatric gated helical CTA | WIP19 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | Ashish_Sharma | 0 |  |  |  |  | Closed |  | Ingenuity 64
CTUDT00627152 | Refresh | Closed | Ashish_Sharma | November 9, 2019 5:19:41 AM GMT+08:00 | In non gated helical chest infant study , "Cardiac Dose Right" option was not enabled. | WIP19 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | Ashish_Sharma | 0 |  |  |  |  | Closed |  | Ingenuity 64
CTUDT00627150 | Refresh | Closed | shiraguppi_k | November 8, 2019 6:03:05 PM GMT+08:00 | Big Bore: BB Refresh Validation: After re-login In ECM Non-English Auto voice language of an EC is changing back to English. | BB_WIP9 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | datta_i | 0 |  |  | Console4 | NO | Closed | WAS | Brilliance Big Bore
CTUDT00627148 | Refresh | Accepted | ding_s | November 8, 2019 5:10:57 PM GMT+08:00 | Prompt error message of 'A problem has occurred' and 'Initialization failed' during Automatic reliability testing. S_MPLTRANSPORTREC_DEFRAG_TOUCAN_ERROR | WIP21 |  |  | Subsystem | ;Blocker |  |  | wang_l | 1 |  |  | Support-System | NO | Accepted |  | Ingenuity 64
CTUDT00627144 | Refresh | Closed | Kevin_ZHAO | November 8, 2019 4:52:51 PM GMT+08:00 | BB Refresh Validation: IQ check head failed | BB_WIP9 |  | N/A | Subsystem | Test Design | N/A | Works As Specified / Not A Defect | Kevin_ZHAO | 1 | 4 |  | Support-System | NO | Closed | WIP22SegReq | Brilliance Big Bore
CTUDT00627143 | Refresh | Closed | wei_f | November 8, 2019 4:46:29 PM GMT+08:00 | Color images did not converted to Gray scale. | WIP20 |  |  | Subsystem |  | N/A | Works As Specified / Not A Defect | wei_f | 0 |  | Design Documentation | Support-System | NO | Closed |  | Brilliance 64
CTUDT00627143 | Refresh | Closed | wei_f | November 8, 2019 4:46:29 PM GMT+08:00 | Color images did not converted to Gray scale. | WIP20 |  |  | Subsystem |  | N/A | Works As Specified / Not A Defect | wei_f | 0 |  | Design Documentation | Support-System | NO | Closed |  | Ingenuity 64
CTUDT00627140 | Refresh | Closed | lei_yang | November 8, 2019 4:15:32 PM GMT+08:00 | DHF232540 Calibraiton Main Menu Parts Replacement Spec isn't applicable Refresh BR64. | WIP21 | DHF232540 Rev15 | N/A | Subsystem |  | DHF232540 | Fixed | richards_b | 0 | 2 | Design Documentation | Support-Physics | NO | Closed | SPEC NEW REVISION | Brilliance 64 |  | DHF232540 Rev15 | November 16, 2019 3:57:39 AM GMT+08:00
CTUDT00627136 | Refresh | Closed | sun_q | November 8, 2019 4:07:08 PM GMT+08:00 | CT Value difference is more than +/-1 HU between the images reconstructed by the partial raw data and full raw data | WIP21 | WIP21 | Measure the ROI within the Head and/or Body sections of the phantom. Do not place the ROI on the phantom holder. | Subsystem | ;Blocker | CT-NM.SRS.Service.2911 | Fixed | wang_l | 0 | 4 | Design Documentation | Support-System | NO | Closed | WIP22SegReq | Brilliance 64 |  | WIP21 | November 19, 2019 11:47:21 AM GMT+08:00
CTUDT00627136 | Refresh | Closed | sun_q | November 8, 2019 4:07:08 PM GMT+08:00 | CT Value difference is more than +/-1 HU between the images reconstructed by the partial raw data and full raw data | WIP21 | WIP21 | Measure the ROI within the Head and/or Body sections of the phantom. Do not place the ROI on the phantom holder. | Subsystem | ;Blocker | CT-NM.SRS.Service.2911 | Fixed | wang_l | 0 | 4 | Design Documentation | Support-System | NO | Closed | WIP22SegReq | Brilliance Big Bore |  | WIP21 | November 19, 2019 11:47:21 AM GMT+08:00
CTUDT00627136 | Refresh | Closed | sun_q | November 8, 2019 4:07:08 PM GMT+08:00 | CT Value difference is more than +/-1 HU between the images reconstructed by the partial raw data and full raw data | WIP21 | WIP21 | Measure the ROI within the Head and/or Body sections of the phantom. Do not place the ROI on the phantom holder. | Subsystem | ;Blocker | CT-NM.SRS.Service.2911 | Fixed | wang_l | 0 | 4 | Design Documentation | Support-System | NO | Closed | WIP22SegReq | Ingenuity 64 |  | WIP21 | November 19, 2019 11:47:21 AM GMT+08:00
CTUDT00627134 | Refresh | Closed | zyanjiang | November 8, 2019 3:11:26 PM GMT+08:00 | Ingenuity validation: A critical error has occurred, starting log out process | WIP19 |  |  | Subsystem |  | NA | No Plan to Fix | datta_i | 0 | 4 |  | Console4 | NO | Closed | SubNC | Ingenuity 64
CTUDT00627128 | Refresh | Closed | cherylli | November 8, 2019 1:13:17 AM GMT+08:00 | Refresh Val: While logging in with Service mode, Workfolw flashed back to Windows Page | BB_WIP9 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | kapadia_v | 0 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00627126 | Refresh | Closed | cherylli | November 8, 2019 1:03:12 AM GMT+08:00 | Refresh Val: Overlap display of scan time of CCT Fluoro in time bar | BB_WIP9 |  |  | Subsystem |  |  | Duplicate | pandey_a | 0 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00627124 | Refresh | Closed | cherylli | November 8, 2019 12:30:29 AM GMT+08:00 | Refresh Val: CCT Fluoro error | BB_WIP9 |  | N/A | Subsystem |  | NA | Works As Specified / Not A Defect | kapadia_v | 0 |  |  | Console4 | YES | Closed | WAS/ Duplicate | Brilliance Big Bore
CTUDT00627120 | Refresh | Closed | marino | November 7, 2019 9:57:26 PM GMT+08:00 | A network interface on a CIRS computer is not properly configured after installation. | WIP21 | WIP22 | Service would need to connect a PS keyboard and mouse to correct the NIC when identified. | Subsystem | Blocker; Verified1 | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_Installation | CIRS | NO | Closed | WIP22SubCom | Ovation |  | WIP22 | November 25, 2019 9:08:45 PM GMT+08:00
CTUDT00627120 | Refresh | Closed | marino | November 7, 2019 9:57:26 PM GMT+08:00 | A network interface on a CIRS computer is not properly configured after installation. | WIP21 | WIP22 | Service would need to connect a PS keyboard and mouse to correct the NIC when identified. | Subsystem | Blocker; Verified1 | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_Installation | CIRS | NO | Closed | WIP22SubCom | Brilliance 64 |  | WIP22 | November 25, 2019 9:08:45 PM GMT+08:00
CTUDT00627120 | Refresh | Closed | marino | November 7, 2019 9:57:26 PM GMT+08:00 | A network interface on a CIRS computer is not properly configured after installation. | WIP21 | WIP22 | Service would need to connect a PS keyboard and mouse to correct the NIC when identified. | Subsystem | Blocker; Verified1 | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_Installation | CIRS | NO | Closed | WIP22SubCom | Brilliance Big Bore |  | WIP22 | November 25, 2019 9:08:45 PM GMT+08:00
CTUDT00627120 | Refresh | Closed | marino | November 7, 2019 9:57:26 PM GMT+08:00 | A network interface on a CIRS computer is not properly configured after installation. | WIP21 | WIP22 | Service would need to connect a PS keyboard and mouse to correct the NIC when identified. | Subsystem | Blocker; Verified1 | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_Installation | CIRS | NO | Closed | WIP22SubCom | Ingenuity 64 |  | WIP22 | November 25, 2019 9:08:45 PM GMT+08:00
CTUDT00627120 | Refresh | Closed | marino | November 7, 2019 9:57:26 PM GMT+08:00 | A network interface on a CIRS computer is not properly configured after installation. | WIP21 | WIP22 | Service would need to connect a PS keyboard and mouse to correct the NIC when identified. | Subsystem | Blocker; Verified1 | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_Installation | CIRS | NO | Closed | WIP22SubCom | IQon1 |  | WIP22 | November 25, 2019 9:08:45 PM GMT+08:00
CTUDT00627118 | Refresh | Closed | zhu_je | November 7, 2019 9:17:21 PM GMT+08:00 | [Refresh]Default mAs for some scan type are not same as the values in file Scan Type Classifications for DHF251214 Rev07 | WIP21 | WIP21 | Change the mAs value manually through the Exam Card Manager during the creation of a new Exam Card | Subsystem | Blocker | N/A | Fixed | dunitza_m | 0 | 3 | Design Documentation | Support-System | NO | Closed | WIP22SegReq | Brilliance 64 |  | WIP21 | November 19, 2019 8:43:59 PM GMT+08:00
CTUDT00627116 | Refresh | Closed | gushansky_m | November 7, 2019 8:26:25 PM GMT+08:00 | Refresh: Partial raw data. Offline recon was stalled | WIP21 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | gushansky_m | 0 |  |  | Support-System | NO | Closed |  | IQon1
CTUDT00627114 | Refresh | Closed | rojin_antony | November 7, 2019 7:50:32 PM GMT+08:00 | More than half of the ECG wave R to R interval is marked as radiated (Red in Color ) during Calcium score with phase 75% | BB_WIP9 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | dongping_w | 0 |  |  | Support-System | NO | Closed |  | Brilliance Big Bore
CTUDT00627108 | Refresh | Closed | rojin_antony | November 7, 2019 5:42:25 PM GMT+08:00 | Portion of patient name is hidden on left monitor if the total number of letters are more than 18 | BB_WIP9 |  | yes, User may see the full name just beside in annotation . image attached | Subsystem |  | NA | No Plan to Fix | sagar_v | 0 | 4 |  | Console4 | NO | Closed | WAS | Ovation
CTUDT00627108 | Refresh | Closed | rojin_antony | November 7, 2019 5:42:25 PM GMT+08:00 | Portion of patient name is hidden on left monitor if the total number of letters are more than 18 | BB_WIP9 |  | yes, User may see the full name just beside in annotation . image attached | Subsystem |  | NA | No Plan to Fix | sagar_v | 0 | 4 |  | Console4 | NO | Closed | WAS | Brilliance 64
CTUDT00627108 | Refresh | Closed | rojin_antony | November 7, 2019 5:42:25 PM GMT+08:00 | Portion of patient name is hidden on left monitor if the total number of letters are more than 18 | BB_WIP9 |  | yes, User may see the full name just beside in annotation . image attached | Subsystem |  | NA | No Plan to Fix | sagar_v | 0 | 4 |  | Console4 | NO | Closed | WAS | Brilliance Big Bore
CTUDT00627108 | Refresh | Closed | rojin_antony | November 7, 2019 5:42:25 PM GMT+08:00 | Portion of patient name is hidden on left monitor if the total number of letters are more than 18 | BB_WIP9 |  | yes, User may see the full name just beside in annotation . image attached | Subsystem |  | NA | No Plan to Fix | sagar_v | 0 | 4 |  | Console4 | NO | Closed | WAS | Ingenuity 64
CTUDT00627108 | Refresh | Closed | rojin_antony | November 7, 2019 5:42:25 PM GMT+08:00 | Portion of patient name is hidden on left monitor if the total number of letters are more than 18 | BB_WIP9 |  | yes, User may see the full name just beside in annotation . image attached | Subsystem |  | NA | No Plan to Fix | sagar_v | 0 | 4 |  | Console4 | NO | Closed | WAS | IQon1
CTUDT00627103 | Refresh | Closed | sun_q | November 7, 2019 2:16:18 PM GMT+08:00 | IVC-"Go to scan positions" and "Go to saved work positions" led behavior is abnormal when worked with "Move to selected position"(On Console) | BB_WIP11 |  |  | System |  | Cosnole-SRS.16539 | No Plan to Fix | madamsetti_k | 0 | 2 |  | Console4 | NO | Closed | SubNC | Brilliance Big Bore
CTUDT00627101 | Refresh | Closed | shiraguppi_k | November 7, 2019 2:07:53 PM GMT+08:00 | Big Bore: BB Refresh Validation: After repeating current/same Cardiac scan system Freezed/Crashed | BB_WIP9 |  |  | Subsystem |  | NA | Duplicate | prasad_r | 0 |  |  | Console4 | YES | Closed | Duplicate | Brilliance Big Bore
CTUDT00627099 | Refresh | Closed | shiraguppi_k | November 7, 2019 2:01:41 PM GMT+08:00 | Big Bore: BB Refresh Validation: In Cardiac scan after click on start final recon, Cardiac results are not reconstructing. | BB_WIP9 |  | If a reconstruction failed, the user will have the ability to reconstruct the data offline to get the expected results. | Subsystem |  | N/A | Works As Specified / Not A Defect | chang_zhiguo | 1 |  |  | Support-System | NO | Closed |  | Brilliance Big Bore
CTUDT00627097 | Refresh | Closed | richards_b | November 7, 2019 11:20:42 AM GMT+08:00 | CIRS Performance Improvement moving WARP processing to GPU | WIP21 | WIP21.2 | Perform WARP on CPU | Subsystem | Verified1 | N/A | Fixed | richards_b | 0 | 3 | CIRS_SW_ImgNodeSvr | CIRS | NO | Closed | WIP22SubCom | Ovation |  | WIP21.2 | November 14, 2019 5:35:31 AM GMT+08:00
CTUDT00627097 | Refresh | Closed | richards_b | November 7, 2019 11:20:42 AM GMT+08:00 | CIRS Performance Improvement moving WARP processing to GPU | WIP21 | WIP21.2 | Perform WARP on CPU | Subsystem | Verified1 | N/A | Fixed | richards_b | 0 | 3 | CIRS_SW_ImgNodeSvr | CIRS | NO | Closed | WIP22SubCom | Brilliance 64 |  | WIP21.2 | November 14, 2019 5:35:31 AM GMT+08:00
CTUDT00627097 | Refresh | Closed | richards_b | November 7, 2019 11:20:42 AM GMT+08:00 | CIRS Performance Improvement moving WARP processing to GPU | WIP21 | WIP21.2 | Perform WARP on CPU | Subsystem | Verified1 | N/A | Fixed | richards_b | 0 | 3 | CIRS_SW_ImgNodeSvr | CIRS | NO | Closed | WIP22SubCom | Brilliance Big Bore |  | WIP21.2 | November 14, 2019 5:35:31 AM GMT+08:00
CTUDT00627097 | Refresh | Closed | richards_b | November 7, 2019 11:20:42 AM GMT+08:00 | CIRS Performance Improvement moving WARP processing to GPU | WIP21 | WIP21.2 | Perform WARP on CPU | Subsystem | Verified1 | N/A | Fixed | richards_b | 0 | 3 | CIRS_SW_ImgNodeSvr | CIRS | NO | Closed | WIP22SubCom | Ingenuity 64 |  | WIP21.2 | November 14, 2019 5:35:31 AM GMT+08:00
CTUDT00627097 | Refresh | Closed | richards_b | November 7, 2019 11:20:42 AM GMT+08:00 | CIRS Performance Improvement moving WARP processing to GPU | WIP21 | WIP21.2 | Perform WARP on CPU | Subsystem | Verified1 | N/A | Fixed | richards_b | 0 | 3 | CIRS_SW_ImgNodeSvr | CIRS | NO | Closed | WIP22SubCom | IQon1 |  | WIP21.2 | November 14, 2019 5:35:31 AM GMT+08:00
CTUDT00627095 | Refresh | Closed | Kevin_ZHAO | November 7, 2019 11:08:26 AM GMT+08:00 | BB Refresh Validation: Calibrations-Constancy test failed | BB_WIP9 |  |  | Subsystem |  |  | Duplicate | yatziv_l | 0 |  |  | DMS | YES | Closed |  | Brilliance Big Bore
CTUDT00627090 | Refresh | Closed | dunitza_m | November 7, 2019 12:30:22 AM GMT+08:00 | Increased CTDI values in Exam Cards for iCT Refresh WIP19 | WIP19 |  | User can manually decrease the CTDI level during planning. | Subsystem | Exam Cards | NA | Duplicate | dunitza_m | 0 |  |  | Support-System | NO | Closed | EC Spec Update | Ovation
CTUDT00627087 | Refresh | Closed | panchami_pradeep | November 6, 2019 10:20:32 PM GMT+08:00 | Refresh: Unable to add Breathing Lights for the phrases in Voice manager. | WIP21 | WIP22 |  | Subsystem |  | N/A | Fixed | navale_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP22SubCom | Ovation |  | WIP22 | November 8, 2019 10:08:52 PM GMT+08:00
CTUDT00627087 | Refresh | Closed | panchami_pradeep | November 6, 2019 10:20:32 PM GMT+08:00 | Refresh: Unable to add Breathing Lights for the phrases in Voice manager. | WIP21 | WIP22 |  | Subsystem |  | N/A | Fixed | navale_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP22SubCom | IQon1 |  | WIP22 | November 8, 2019 10:08:52 PM GMT+08:00
CTUDT00627083 | Refresh | Closed | cherylli | November 6, 2019 10:12:52 PM GMT+08:00 | Refresh Val: Patient UI with information was frozen | BB_WIP9 |  |  | Subsystem |  | N/A | Works As Specified / Not A Defect | pandey_a | 0 |  |  | Console4 | NO | Closed | WAS/Not an issue | Brilliance Big Bore
CTUDT00627082 | Refresh | Closed | panchami_pradeep | November 6, 2019 10:09:39 PM GMT+08:00 | Refresh: 0.42 Option is not present in RT drop down in iQon | WIP21 | DHF303211 Rev19 | N/A | Subsystem |  | Console-SRS.17735 | Fixed | roy_r | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | SSRS Update | IQon1 |  | DHF303211 Rev19 | November 14, 2019 7:21:19 PM GMT+08:00
CTUDT00627080 | Refresh | Closed | cherylli | November 6, 2019 9:54:33 PM GMT+08:00 | Refresh Validation: Inconsistent recon parameters in offline recon (edit results in slab viewer) | BB_WIP9 |  |  | Subsystem |  |  | Duplicate | ravindra_a | 0 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00627077 | Refresh | Closed | mN_m | November 6, 2019 9:20:42 PM GMT+08:00 | Console allowed to delete the not saved reconstructed conventional matching result when the conventional result is edited and duplicated. | WIP21 |  | Perform offline reconstruction if any result is desired | Subsystem |  | NA | No Plan to Fix | challa_s | 0 | 4 |  | Console4 | NO | Closed | SubNC | IQon1
CTUDT00627076 | Refresh | Closed | cherylli | November 6, 2019 9:10:03 PM GMT+08:00 | Refresh Val: Snagit Editor stopped working | BB_WIP9 | BB_WIP12 WIP22 | User can close the popup message box and continue working. | Subsystem |  | NA | Fixed | dasgupta_s | 0 | 4 | Console Software | Console4 | NO | Closed | WIP22SubCom | Brilliance Big Bore |  | BB_WIP12 WIP22 | November 12, 2019 10:41:15 PM GMT+08:00
CTUDT00627074 | Refresh | Closed | vineetha_s | November 6, 2019 8:06:13 PM GMT+08:00 | A suggestion was not present in the report for the change of scan type while converting protocols and exam cards from Z440 to Z4 | WIP21 | DHF303211 Rev19 DHF316705 Rev19 | NA | Subsystem |  | Console-SRS.3483 | Fixed | roy_r | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | SSRS Update | Ovation |  | DHF303211 Rev19 DHF316705 Rev19 | November 14, 2019 7:20:33 PM GMT+08:00
CTUDT00627074 | Refresh | Closed | vineetha_s | November 6, 2019 8:06:13 PM GMT+08:00 | A suggestion was not present in the report for the change of scan type while converting protocols and exam cards from Z440 to Z4 | WIP21 | DHF303211 Rev19 DHF316705 Rev19 | NA | Subsystem |  | Console-SRS.3483 | Fixed | roy_r | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | SSRS Update | Brilliance 64 |  | DHF303211 Rev19 DHF316705 Rev19 | November 14, 2019 7:20:33 PM GMT+08:00
CTUDT00627074 | Refresh | Closed | vineetha_s | November 6, 2019 8:06:13 PM GMT+08:00 | A suggestion was not present in the report for the change of scan type while converting protocols and exam cards from Z440 to Z4 | WIP21 | DHF303211 Rev19 DHF316705 Rev19 | NA | Subsystem |  | Console-SRS.3483 | Fixed | roy_r | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | SSRS Update | Brilliance Big Bore |  | DHF303211 Rev19 DHF316705 Rev19 | November 14, 2019 7:20:33 PM GMT+08:00
CTUDT00627074 | Refresh | Closed | vineetha_s | November 6, 2019 8:06:13 PM GMT+08:00 | A suggestion was not present in the report for the change of scan type while converting protocols and exam cards from Z440 to Z4 | WIP21 | DHF303211 Rev19 DHF316705 Rev19 | NA | Subsystem |  | Console-SRS.3483 | Fixed | roy_r | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | SSRS Update | Ingenuity 64 |  | DHF303211 Rev19 DHF316705 Rev19 | November 14, 2019 7:20:33 PM GMT+08:00
CTUDT00627074 | Refresh | Closed | vineetha_s | November 6, 2019 8:06:13 PM GMT+08:00 | A suggestion was not present in the report for the change of scan type while converting protocols and exam cards from Z440 to Z4 | WIP21 | DHF303211 Rev19 DHF316705 Rev19 | NA | Subsystem |  | Console-SRS.3483 | Fixed | roy_r | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | SSRS Update | IQon1 |  | DHF303211 Rev19 DHF316705 Rev19 | November 14, 2019 7:20:33 PM GMT+08:00
CTUDT00627071 | Refresh | Closed | rojin_antony | November 6, 2019 7:12:22 PM GMT+08:00 | The Edited patient study does not support Slab ,Volume and Endo options in CT Viewer | BB_WIP9 |  | CT viewer is offline tool and user can load original series if required | Subsystem |  | na | Works As Specified / Not A Defect | s_sa | 0 |  |  | iCAP | NO | Closed |  | Ovation
CTUDT00627071 | Refresh | Closed | rojin_antony | November 6, 2019 7:12:22 PM GMT+08:00 | The Edited patient study does not support Slab ,Volume and Endo options in CT Viewer | BB_WIP9 |  | CT viewer is offline tool and user can load original series if required | Subsystem |  | na | Works As Specified / Not A Defect | s_sa | 0 |  |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00627071 | Refresh | Closed | rojin_antony | November 6, 2019 7:12:22 PM GMT+08:00 | The Edited patient study does not support Slab ,Volume and Endo options in CT Viewer | BB_WIP9 |  | CT viewer is offline tool and user can load original series if required | Subsystem |  | na | Works As Specified / Not A Defect | s_sa | 0 |  |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00627071 | Refresh | Closed | rojin_antony | November 6, 2019 7:12:22 PM GMT+08:00 | The Edited patient study does not support Slab ,Volume and Endo options in CT Viewer | BB_WIP9 |  | CT viewer is offline tool and user can load original series if required | Subsystem |  | na | Works As Specified / Not A Defect | s_sa | 0 |  |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00627071 | Refresh | Closed | rojin_antony | November 6, 2019 7:12:22 PM GMT+08:00 | The Edited patient study does not support Slab ,Volume and Endo options in CT Viewer | BB_WIP9 |  | CT viewer is offline tool and user can load original series if required | Subsystem |  | na | Works As Specified / Not A Defect | s_sa | 0 |  |  | iCAP | NO | Closed |  | IQon1
CTUDT00627068 | Refresh | Closed | Jose_D_Sergio_Almeida | November 6, 2019 6:40:00 PM GMT+08:00 | Ingenuity validation: Unable to change from gated to non-gated for Step& shoot acquisition | WIP19 |  |  | Subsystem |  | NA | No Plan to Fix | prasad_r | 0 | 3 |  | Console4 | NO | Closed |  | Ovation | NA
CTUDT00627068 | Refresh | Closed | Jose_D_Sergio_Almeida | November 6, 2019 6:40:00 PM GMT+08:00 | Ingenuity validation: Unable to change from gated to non-gated for Step& shoot acquisition | WIP19 |  |  | Subsystem |  | NA | No Plan to Fix | prasad_r | 0 | 3 |  | Console4 | NO | Closed |  | Brilliance 64 | NA
CTUDT00627068 | Refresh | Closed | Jose_D_Sergio_Almeida | November 6, 2019 6:40:00 PM GMT+08:00 | Ingenuity validation: Unable to change from gated to non-gated for Step& shoot acquisition | WIP19 |  |  | Subsystem |  | NA | No Plan to Fix | prasad_r | 0 | 3 |  | Console4 | NO | Closed |  | Brilliance Big Bore | NA
CTUDT00627068 | Refresh | Closed | Jose_D_Sergio_Almeida | November 6, 2019 6:40:00 PM GMT+08:00 | Ingenuity validation: Unable to change from gated to non-gated for Step& shoot acquisition | WIP19 |  |  | Subsystem |  | NA | No Plan to Fix | prasad_r | 0 | 3 |  | Console4 | NO | Closed |  | Ingenuity 64 | NA
CTUDT00627068 | Refresh | Closed | Jose_D_Sergio_Almeida | November 6, 2019 6:40:00 PM GMT+08:00 | Ingenuity validation: Unable to change from gated to non-gated for Step& shoot acquisition | WIP19 |  |  | Subsystem |  | NA | No Plan to Fix | prasad_r | 0 | 3 |  | Console4 | NO | Closed |  | IQon1 | NA
CTUDT00627066 | Refresh | Closed | chang_zhiguo | November 6, 2019 5:51:28 PM GMT+08:00 | BB refresh validation: The patient not moved to history, after finished the scan and ended the study. | BB_WIP9 |  |  | Subsystem |  | N/A | Works As Specified / Not A Defect | pandey_a | 0 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00627064 | Refresh | Closed | zyanjiang | November 6, 2019 4:58:53 PM GMT+08:00 | Ingenuity validation: WC/WW of locator and tracker are 300/500 for preference coronary CTA exam card | WIP19 |  |  | Subsystem |  |  | Duplicate | pandey_a | 0 |  |  |  |  | Closed |  | Ingenuity 64
CTUDT00627062 | Refresh | Closed | wang_l | November 6, 2019 3:33:08 PM GMT+08:00 | Refresh Reliability - CCT Work Position radio button was clicked but does not have any response | BB_WIP11 |  |  | Subsystem |  | Console-SRS.15820 | Works As Specified / Not A Defect | tan_j | 0 |  |  | GCX | NO | Closed | WAS | Brilliance Big Bore
CTUDT00627060 | Refresh | Closed | l_yiming | November 6, 2019 3:08:49 PM GMT+08:00 | ConsoleHowDoI Stop Working | WIP19 |  |  | Subsystem |  |  | Duplicate | pandey_a | 0 |  |  |  |  | Closed |  | Ingenuity 64
CTUDT00627058 | Refresh | Closed | l_yiming | November 6, 2019 2:48:19 PM GMT+08:00 | Ingenuity validation: Recon Thickness Wrongly Indicated | WIP19 |  | Check images which display corrected thickness | Subsystem | CRN | NA | No Plan to Fix | ravindra_a | 0 | 3 |  | Console4 | NO | Closed | SubNC | Ovation
CTUDT00627058 | Refresh | Closed | l_yiming | November 6, 2019 2:48:19 PM GMT+08:00 | Ingenuity validation: Recon Thickness Wrongly Indicated | WIP19 |  | Check images which display corrected thickness | Subsystem | CRN | NA | No Plan to Fix | ravindra_a | 0 | 3 |  | Console4 | NO | Closed | SubNC | Brilliance 64
CTUDT00627058 | Refresh | Closed | l_yiming | November 6, 2019 2:48:19 PM GMT+08:00 | Ingenuity validation: Recon Thickness Wrongly Indicated | WIP19 |  | Check images which display corrected thickness | Subsystem | CRN | NA | No Plan to Fix | ravindra_a | 0 | 3 |  | Console4 | NO | Closed | SubNC | Brilliance Big Bore
CTUDT00627058 | Refresh | Closed | l_yiming | November 6, 2019 2:48:19 PM GMT+08:00 | Ingenuity validation: Recon Thickness Wrongly Indicated | WIP19 |  | Check images which display corrected thickness | Subsystem | CRN | NA | No Plan to Fix | ravindra_a | 0 | 3 |  | Console4 | NO | Closed | SubNC | Ingenuity 64
CTUDT00627058 | Refresh | Closed | l_yiming | November 6, 2019 2:48:19 PM GMT+08:00 | Ingenuity validation: Recon Thickness Wrongly Indicated | WIP19 |  | Check images which display corrected thickness | Subsystem | CRN | NA | No Plan to Fix | ravindra_a | 0 | 3 |  | Console4 | NO | Closed | SubNC | IQon1
CTUDT00627056 | Refresh | Closed | zhu_je | November 6, 2019 1:56:45 PM GMT+08:00 | [BB Refresh] Series in studies repetitive display Studies tab during refresh in reconstruction tab. | BB_WIP11 |  | click on Refresh button on Studies tab or select some other study. | Subsystem | Patient Directory | NA | No Plan to Fix | s_panigrahi | 0 | 3 |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00627054 | Refresh | Closed | wang_l | November 6, 2019 10:26:55 AM GMT+08:00 | Refresh BB - Unexpected error "ConsoleHowDoI has stopped working" prompte during normal operation | BB_WIP11 | WIP22 BB_WIP12 |  | Subsystem |  | N/A | Fixed | pandey_a | 0 | 4 | Console Software | Console4 | NO | Closed | WIP22SubCom | Ovation |  | WIP22 BB_WIP12 | November 8, 2019 8:59:47 PM GMT+08:00
CTUDT00627054 | Refresh | Closed | wang_l | November 6, 2019 10:26:55 AM GMT+08:00 | Refresh BB - Unexpected error "ConsoleHowDoI has stopped working" prompte during normal operation | BB_WIP11 | WIP22 BB_WIP12 |  | Subsystem |  | N/A | Fixed | pandey_a | 0 | 4 | Console Software | Console4 | NO | Closed | WIP22SubCom | Brilliance 64 |  | WIP22 BB_WIP12 | November 8, 2019 8:59:47 PM GMT+08:00
CTUDT00627054 | Refresh | Closed | wang_l | November 6, 2019 10:26:55 AM GMT+08:00 | Refresh BB - Unexpected error "ConsoleHowDoI has stopped working" prompte during normal operation | BB_WIP11 | WIP22 BB_WIP12 |  | Subsystem |  | N/A | Fixed | pandey_a | 0 | 4 | Console Software | Console4 | NO | Closed | WIP22SubCom | Brilliance Big Bore |  | WIP22 BB_WIP12 | November 8, 2019 8:59:47 PM GMT+08:00
CTUDT00627054 | Refresh | Closed | wang_l | November 6, 2019 10:26:55 AM GMT+08:00 | Refresh BB - Unexpected error "ConsoleHowDoI has stopped working" prompte during normal operation | BB_WIP11 | WIP22 BB_WIP12 |  | Subsystem |  | N/A | Fixed | pandey_a | 0 | 4 | Console Software | Console4 | NO | Closed | WIP22SubCom | Ingenuity 64 |  | WIP22 BB_WIP12 | November 8, 2019 8:59:47 PM GMT+08:00
CTUDT00627054 | Refresh | Closed | wang_l | November 6, 2019 10:26:55 AM GMT+08:00 | Refresh BB - Unexpected error "ConsoleHowDoI has stopped working" prompte during normal operation | BB_WIP11 | WIP22 BB_WIP12 |  | Subsystem |  | N/A | Fixed | pandey_a | 0 | 4 | Console Software | Console4 | NO | Closed | WIP22SubCom | IQon1 |  | WIP22 BB_WIP12 | November 8, 2019 8:59:47 PM GMT+08:00
CTUDT00627052 | Refresh | Closed | martinez_c | November 6, 2019 5:22:37 AM GMT+08:00 | One of two GPU driver versions installed during CIRS OS install. | CIRS-5.1.0.0203 | WIP21.2 | N/A | Subsystem | Verified1 | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_Installation | CIRS | NO | Closed | WIP22SegReq | Ovation |  | WIP21.2 | November 14, 2019 5:35:02 AM GMT+08:00
CTUDT00627052 | Refresh | Closed | martinez_c | November 6, 2019 5:22:37 AM GMT+08:00 | One of two GPU driver versions installed during CIRS OS install. | CIRS-5.1.0.0203 | WIP21.2 | N/A | Subsystem | Verified1 | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_Installation | CIRS | NO | Closed | WIP22SegReq | Brilliance 64 |  | WIP21.2 | November 14, 2019 5:35:02 AM GMT+08:00
CTUDT00627052 | Refresh | Closed | martinez_c | November 6, 2019 5:22:37 AM GMT+08:00 | One of two GPU driver versions installed during CIRS OS install. | CIRS-5.1.0.0203 | WIP21.2 | N/A | Subsystem | Verified1 | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_Installation | CIRS | NO | Closed | WIP22SegReq | Brilliance Big Bore |  | WIP21.2 | November 14, 2019 5:35:02 AM GMT+08:00
CTUDT00627052 | Refresh | Closed | martinez_c | November 6, 2019 5:22:37 AM GMT+08:00 | One of two GPU driver versions installed during CIRS OS install. | CIRS-5.1.0.0203 | WIP21.2 | N/A | Subsystem | Verified1 | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_Installation | CIRS | NO | Closed | WIP22SegReq | Ingenuity 64 |  | WIP21.2 | November 14, 2019 5:35:02 AM GMT+08:00
CTUDT00627052 | Refresh | Closed | martinez_c | November 6, 2019 5:22:37 AM GMT+08:00 | One of two GPU driver versions installed during CIRS OS install. | CIRS-5.1.0.0203 | WIP21.2 | N/A | Subsystem | Verified1 | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_Installation | CIRS | NO | Closed | WIP22SegReq | IQon1 |  | WIP21.2 | November 14, 2019 5:35:02 AM GMT+08:00
CTUDT00627051 | Refresh | Closed | altshuller_mark | November 5, 2019 9:50:30 PM GMT+08:00 | Locator does not appear after dual surview, if tracker dose is too high | IQon HB1 | WIP31 | Immediately click on the “Accumulated CTDIvol” popup once it appears on the screen. | System | SP1 | N/A | Fixed | janakiraman_j | 0 | 2 | Console Software | Console4 | YES | Closed | WIP31SegReq | Ovation | PR.9502317 PR. 9812971 | WIP31 | July 24, 2020 7:05:51 PM GMT+08:00
CTUDT00627051 | Refresh | Closed | altshuller_mark | November 5, 2019 9:50:30 PM GMT+08:00 | Locator does not appear after dual surview, if tracker dose is too high | IQon HB1 | WIP31 | Immediately click on the “Accumulated CTDIvol” popup once it appears on the screen. | System | SP1 | N/A | Fixed | janakiraman_j | 0 | 2 | Console Software | Console4 | YES | Closed | WIP31SegReq | IQon1 | PR.9502317 PR. 9812971 | WIP31 | July 24, 2020 7:05:51 PM GMT+08:00
CTUDT00627049 | Refresh | Closed | ding_f | November 5, 2019 8:18:12 PM GMT+08:00 | The Motion speed controlled by rocker switch do not match with Sliding Gantry specification in MotionSpeed.cfg file. | WIP11 | BB_WIP12 | can manually change the freemotionspeeds table value in C:\Pms\Config\ScannerFramework\IVMotionControlService\MotionSpeed.cfg as per the sliding gantry values in sliding gantry configuration. | System | ;Blocker | Console-SRS.16541 | Fixed | li_e | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP12_SubCom | Brilliance Big Bore |  | BB_WIP12 | November 8, 2019 10:49:34 PM GMT+08:00
CTUDT00627047 | Refresh | Closed | vineetha_s | November 5, 2019 6:51:56 PM GMT+08:00 | the deviations found in the windowing presets for spectral results - Iodine density , Calcium Supression & Z Effective | WIP21 | DHF303211 Rev19 | N/A | Subsystem |  | Console-SRS.16960 | Fixed | roy_r | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | SSRS Update | IQon1 |  | DHF303211 Rev19 | November 14, 2019 7:23:12 PM GMT+08:00
CTUDT00627042 | Refresh | Closed | 320053769 | November 5, 2019 5:27:57 PM GMT+08:00 | In recording popup message, the maximum disk space display doesn’t match the storage capacity of the disc, and record button is disabled (grayed out). | WIP21 |  |  | Subsystem | CDR;Blocker | NA | Works As Specified / Not A Defect | srushti_k | 0 |  |  | iCAP | NO | Closed |  | Ovation
CTUDT00627042 | Refresh | Closed | 320053769 | November 5, 2019 5:27:57 PM GMT+08:00 | In recording popup message, the maximum disk space display doesn’t match the storage capacity of the disc, and record button is disabled (grayed out). | WIP21 |  |  | Subsystem | CDR;Blocker | NA | Works As Specified / Not A Defect | srushti_k | 0 |  |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00627042 | Refresh | Closed | 320053769 | November 5, 2019 5:27:57 PM GMT+08:00 | In recording popup message, the maximum disk space display doesn’t match the storage capacity of the disc, and record button is disabled (grayed out). | WIP21 |  |  | Subsystem | CDR;Blocker | NA | Works As Specified / Not A Defect | srushti_k | 0 |  |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00627042 | Refresh | Closed | 320053769 | November 5, 2019 5:27:57 PM GMT+08:00 | In recording popup message, the maximum disk space display doesn’t match the storage capacity of the disc, and record button is disabled (grayed out). | WIP21 |  |  | Subsystem | CDR;Blocker | NA | Works As Specified / Not A Defect | srushti_k | 0 |  |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00627042 | Refresh | Closed | 320053769 | November 5, 2019 5:27:57 PM GMT+08:00 | In recording popup message, the maximum disk space display doesn’t match the storage capacity of the disc, and record button is disabled (grayed out). | WIP21 |  |  | Subsystem | CDR;Blocker | NA | Works As Specified / Not A Defect | srushti_k | 0 |  |  | iCAP | NO | Closed |  | IQon1
CTUDT00627041 | Refresh | Closed | Jose_D_Sergio_Almeida | November 5, 2019 5:22:48 PM GMT+08:00 | Incorrect Window Width and Window level for Pediatric Step&Shoot reference EC | WIP19 |  | The user can change window width level, in ECM or during planning, or on eview Before Save phase or if required can do offline reconstruction with changed window width and window level. | Subsystem | CRN | N/A | No Plan to Fix | dunitza_m | 0 | 2 |  | Support-System | NO | Closed |  | Ingenuity 64 | NA
CTUDT00627039 | Refresh | Closed | 320053769 | November 5, 2019 5:06:59 PM GMT+08:00 | In recording popup message, the letter “E” at “maximum disk space” row is cut/hidden | WIP21 |  |  | Subsystem | GUI | NA | No Plan to Fix | srushti_k | 0 | 3 |  | iCAP | NO | Closed |  | Ovation
CTUDT00627039 | Refresh | Closed | 320053769 | November 5, 2019 5:06:59 PM GMT+08:00 | In recording popup message, the letter “E” at “maximum disk space” row is cut/hidden | WIP21 |  |  | Subsystem | GUI | NA | No Plan to Fix | srushti_k | 0 | 3 |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00627039 | Refresh | Closed | 320053769 | November 5, 2019 5:06:59 PM GMT+08:00 | In recording popup message, the letter “E” at “maximum disk space” row is cut/hidden | WIP21 |  |  | Subsystem | GUI | NA | No Plan to Fix | srushti_k | 0 | 3 |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00627039 | Refresh | Closed | 320053769 | November 5, 2019 5:06:59 PM GMT+08:00 | In recording popup message, the letter “E” at “maximum disk space” row is cut/hidden | WIP21 |  |  | Subsystem | GUI | NA | No Plan to Fix | srushti_k | 0 | 3 |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00627039 | Refresh | Closed | 320053769 | November 5, 2019 5:06:59 PM GMT+08:00 | In recording popup message, the letter “E” at “maximum disk space” row is cut/hidden | WIP21 |  |  | Subsystem | GUI | NA | No Plan to Fix | srushti_k | 0 | 3 |  | iCAP | NO | Closed |  | IQon1
CTUDT00627034 | Refresh | Closed | wang_l | November 5, 2019 4:38:00 PM GMT+08:00 | Refresh Reliability - End Exam does not work | BB_WIP11 |  | Ideal suggestion is to restart the system once and try. | Subsystem | Blocker | NA | Works As Specified / Not A Defect | li_e | 1 |  |  | Support-System | NO | Closed |  | Ovation
CTUDT00627034 | Refresh | Closed | wang_l | November 5, 2019 4:38:00 PM GMT+08:00 | Refresh Reliability - End Exam does not work | BB_WIP11 |  | Ideal suggestion is to restart the system once and try. | Subsystem | Blocker | NA | Works As Specified / Not A Defect | li_e | 1 |  |  | Support-System | NO | Closed |  | Brilliance 64
CTUDT00627034 | Refresh | Closed | wang_l | November 5, 2019 4:38:00 PM GMT+08:00 | Refresh Reliability - End Exam does not work | BB_WIP11 |  | Ideal suggestion is to restart the system once and try. | Subsystem | Blocker | NA | Works As Specified / Not A Defect | li_e | 1 |  |  | Support-System | NO | Closed |  | Brilliance Big Bore
CTUDT00627034 | Refresh | Closed | wang_l | November 5, 2019 4:38:00 PM GMT+08:00 | Refresh Reliability - End Exam does not work | BB_WIP11 |  | Ideal suggestion is to restart the system once and try. | Subsystem | Blocker | NA | Works As Specified / Not A Defect | li_e | 1 |  |  | Support-System | NO | Closed |  | Ingenuity 64
CTUDT00627034 | Refresh | Closed | wang_l | November 5, 2019 4:38:00 PM GMT+08:00 | Refresh Reliability - End Exam does not work | BB_WIP11 |  | Ideal suggestion is to restart the system once and try. | Subsystem | Blocker | NA | Works As Specified / Not A Defect | li_e | 1 |  |  | Support-System | NO | Closed |  | IQon1
CTUDT00627032 | Refresh | Closed | wang_l | November 5, 2019 4:31:01 PM GMT+08:00 | Refresh Reliability - dose alert "mAs value has changed from xxx mAs to xxx mAs. Please ensure mAs is appropriate for current patient" occurred during surview scan | WIP21 |  | Its working as specified | Subsystem | Blocker | None | Works As Specified / Not A Defect | jose_s | 0 |  |  | Console4 | NO | Closed | WAS | Ovation
CTUDT00627032 | Refresh | Closed | wang_l | November 5, 2019 4:31:01 PM GMT+08:00 | Refresh Reliability - dose alert "mAs value has changed from xxx mAs to xxx mAs. Please ensure mAs is appropriate for current patient" occurred during surview scan | WIP21 |  | Its working as specified | Subsystem | Blocker | None | Works As Specified / Not A Defect | jose_s | 0 |  |  | Console4 | NO | Closed | WAS | Brilliance 64
CTUDT00627032 | Refresh | Closed | wang_l | November 5, 2019 4:31:01 PM GMT+08:00 | Refresh Reliability - dose alert "mAs value has changed from xxx mAs to xxx mAs. Please ensure mAs is appropriate for current patient" occurred during surview scan | WIP21 |  | Its working as specified | Subsystem | Blocker | None | Works As Specified / Not A Defect | jose_s | 0 |  |  | Console4 | NO | Closed | WAS | Brilliance Big Bore
CTUDT00627032 | Refresh | Closed | wang_l | November 5, 2019 4:31:01 PM GMT+08:00 | Refresh Reliability - dose alert "mAs value has changed from xxx mAs to xxx mAs. Please ensure mAs is appropriate for current patient" occurred during surview scan | WIP21 |  | Its working as specified | Subsystem | Blocker | None | Works As Specified / Not A Defect | jose_s | 0 |  |  | Console4 | NO | Closed | WAS | Ingenuity 64
CTUDT00627032 | Refresh | Closed | wang_l | November 5, 2019 4:31:01 PM GMT+08:00 | Refresh Reliability - dose alert "mAs value has changed from xxx mAs to xxx mAs. Please ensure mAs is appropriate for current patient" occurred during surview scan | WIP21 |  | Its working as specified | Subsystem | Blocker | None | Works As Specified / Not A Defect | jose_s | 0 |  |  | Console4 | NO | Closed | WAS | IQon1
CTUDT00627031 | Refresh | Closed | l_yiming | November 5, 2019 4:14:29 PM GMT+08:00 | Ingenuity Validation: Reference Images did not arrive in CCT Single and Continous procedure | WIP19 |  |  | Subsystem | CCT |  | Duplicate | prasad_r | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00627028 | Refresh | Closed | navale_s | November 5, 2019 3:09:47 PM GMT+08:00 | Localization strings are getting cropped in the context menu for IVC in CCT Scene or Result viewer, not able to read the options. | V4.8 |  | User can visually see feedback after selection and correct it if required | Subsystem | FOK | NA | No Plan to Fix | aswani_v | 0 | 3 |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00627026 | Refresh | Closed | jose_s | November 5, 2019 2:28:03 PM GMT+08:00 | Software and Hardware Interventional Controls videos are not translated | BB_WIP11 | BB_WIP12 | English translation is available | Subsystem |  | None | Fixed | jose_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP22SubCom | Brilliance Big Bore |  | BB_WIP12 | November 7, 2019 1:26:05 PM GMT+08:00
CTUDT00627021 | Refresh | Closed | wang_l | November 5, 2019 12:37:41 PM GMT+08:00 | Refresh Reliability - Warning message “Warning: Current performance does not enable real time image display” displayed for CCT Fluoro | BB_WIP11 |  |  | Subsystem |  | CT-NM.SRS.Modes.1474 | No Plan to Fix | bandil_j | 0 | 4 |  | Support-System | NO | Closed |  | Brilliance Big Bore
CTUDT00627018 | Refresh | Closed | 320053769 | November 5, 2019 3:38:03 AM GMT+08:00 | Scan is prevented when disk space of the scanner is below 19.4 GB. | WIP21 |  |  | Subsystem | Disk Space; Blocker |  | Duplicate | pandey_a | 0 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00627016 | Refresh | Closed | Ashish_Sharma | November 5, 2019 1:29:49 AM GMT+08:00 | Ingenuity validation: When loading non gated helical chest exam card 0-18m ,added a gated results with 75% ,nongated result changed to gated result. | WIP19 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | Ashish_Sharma | 1 |  |  | Support-System | NO | Closed |  | Ingenuity 64
CTUDT00627014 | Refresh | Closed | gushansky_m | November 4, 2019 11:26:09 PM GMT+08:00 | Refresh : Service: X-Ray doesn't return data according to CT-NM.SRS.Service.2015 | WIP20 |  |  | Subsystem | Visual Diagnostics - Gantry;Blocker |  | Works As Specified / Not A Defect | benny_w | 0 |  |  | GCX | NO | Closed |  | IQon1
CTUDT00627012 | Refresh | Closed | rojin_antony | November 4, 2019 7:33:41 PM GMT+08:00 | Ring Artifact were noticed in pelvic region during CCT Single , Continuous and Fluro | BB_WIP9 |  | N/A | Subsystem |  | N/A | Works As Specified / Not A Defect | richards_b | 0 |  |  | Support-Physics | NO | Closed |  | Brilliance Big Bore
CTUDT00627010 | Refresh | Closed | rojin_antony | November 4, 2019 6:56:58 PM GMT+08:00 | BB Refresh Validation: Exam Summary Dose information does not have Total CTDIvol and Total SSDE | BB_WIP9 |  | NA | Subsystem |  | Console SRS.2166 | Works As Specified / Not A Defect | morton_t | 0 | 2 |  | Support-System | NO | Closed | WAS/Change proposal | Ovation
CTUDT00627010 | Refresh | Closed | rojin_antony | November 4, 2019 6:56:58 PM GMT+08:00 | BB Refresh Validation: Exam Summary Dose information does not have Total CTDIvol and Total SSDE | BB_WIP9 |  | NA | Subsystem |  | Console SRS.2166 | Works As Specified / Not A Defect | morton_t | 0 | 2 |  | Support-System | NO | Closed | WAS/Change proposal | Brilliance 64
CTUDT00627010 | Refresh | Closed | rojin_antony | November 4, 2019 6:56:58 PM GMT+08:00 | BB Refresh Validation: Exam Summary Dose information does not have Total CTDIvol and Total SSDE | BB_WIP9 |  | NA | Subsystem |  | Console SRS.2166 | Works As Specified / Not A Defect | morton_t | 0 | 2 |  | Support-System | NO | Closed | WAS/Change proposal | Brilliance Big Bore
CTUDT00627010 | Refresh | Closed | rojin_antony | November 4, 2019 6:56:58 PM GMT+08:00 | BB Refresh Validation: Exam Summary Dose information does not have Total CTDIvol and Total SSDE | BB_WIP9 |  | NA | Subsystem |  | Console SRS.2166 | Works As Specified / Not A Defect | morton_t | 0 | 2 |  | Support-System | NO | Closed | WAS/Change proposal | Ingenuity 64
CTUDT00627010 | Refresh | Closed | rojin_antony | November 4, 2019 6:56:58 PM GMT+08:00 | BB Refresh Validation: Exam Summary Dose information does not have Total CTDIvol and Total SSDE | BB_WIP9 |  | NA | Subsystem |  | Console SRS.2166 | Works As Specified / Not A Defect | morton_t | 0 | 2 |  | Support-System | NO | Closed | WAS/Change proposal | IQon1
CTUDT00627007 | Refresh | Closed | chang_zhiguo | November 4, 2019 5:25:28 PM GMT+08:00 | BB refresh validation: 3D-modulation and Z-modulation is unavailable for the user mode | BB_WIP9 |  | Login using service credential to enable Z-MOD or 3d MOD | Subsystem |  | N/A | Works As Specified / Not A Defect | roy_r | 0 |  |  | Console4 | NO | Closed | WAS | Brilliance Big Bore
CTUDT00627005 | Refresh | Closed | chang_zhiguo | November 4, 2019 5:20:49 PM GMT+08:00 | BB refresh validation: No pediatric EC available in Orthoped EC group | BB_WIP9 |  | User can create a Pediatric Extremity EC in ECM and perform scan using the same. | Subsystem |  | NA | Works As Specified / Not A Defect | kr_m | 0 |  |  | Console4 | NO | Closed | WAS | Brilliance Big Bore
CTUDT00627002 | Refresh | Closed | chang_zhiguo | November 4, 2019 5:11:05 PM GMT+08:00 | BB refresh validation: Unexpected error messages pop up while planning the scan in the plan UI | BB_WIP9 |  |  | Subsystem |  |  | Duplicate | pandey_a | 0 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00626997 | Refresh | Closed | zhu_je | November 4, 2019 4:38:13 PM GMT+08:00 | [Refresh][IVC] Go to outer laser led status was not right when return back to normal mode | BB_WIP11 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | Snehashis_Dey | 0 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00626995 | Refresh | Closed | chang_zhiguo | November 4, 2019 4:33:19 PM GMT+08:00 | BB refresh validation: YC image quality is not sharp enough, the cortical edge of bone was blur, hardly to see the details clearly. | BB_WIP9 |  | The end user would have the ability to perform offline reconstruction if they did not like the reconstruction with YC. The end user could use YC + Enhancement or YD for a sharper appearance. | Subsystem | CRN | N/A | No Plan to Fix | richards_b | 0 | 3 |  | Support-Physics | NO | Closed |  | Brilliance Big Bore
CTUDT00626993 | Refresh | Closed | zyanjiang | November 4, 2019 4:21:07 PM GMT+08:00 | Ingenuity validation: TTP and PID display in different decimal places | WIP19 |  | NA | Subsystem |  | NA | No Plan to Fix | aswani_v | 0 | 3 |  | iCAP | NO | Closed |  | Ovation
CTUDT00626993 | Refresh | Closed | zyanjiang | November 4, 2019 4:21:07 PM GMT+08:00 | Ingenuity validation: TTP and PID display in different decimal places | WIP19 |  | NA | Subsystem |  | NA | No Plan to Fix | aswani_v | 0 | 3 |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00626993 | Refresh | Closed | zyanjiang | November 4, 2019 4:21:07 PM GMT+08:00 | Ingenuity validation: TTP and PID display in different decimal places | WIP19 |  | NA | Subsystem |  | NA | No Plan to Fix | aswani_v | 0 | 3 |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00626993 | Refresh | Closed | zyanjiang | November 4, 2019 4:21:07 PM GMT+08:00 | Ingenuity validation: TTP and PID display in different decimal places | WIP19 |  | NA | Subsystem |  | NA | No Plan to Fix | aswani_v | 0 | 3 |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00626993 | Refresh | Closed | zyanjiang | November 4, 2019 4:21:07 PM GMT+08:00 | Ingenuity validation: TTP and PID display in different decimal places | WIP19 |  | NA | Subsystem |  | NA | No Plan to Fix | aswani_v | 0 | 3 |  | iCAP | NO | Closed |  | IQon1
CTUDT00626992 | Refresh | Closed | wang_l | November 4, 2019 4:21:04 PM GMT+08:00 | Refresh BB - Right panel display blank and contents are hidden | BB_WIP11 |  |  | Subsystem |  | NA | No Plan to Fix | s_panigrahi | 0 | 4 |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00626989 | Refresh | Closed | wang_l | November 4, 2019 4:17:52 PM GMT+08:00 | Refresh Reliability - Review tab was not opened after recon finished for EC with CT Viewer was selected | BB_WIP11 |  | yes, user may explicitly open the CTviewer. | Subsystem |  | NA | Works As Specified / Not A Defect | wang_l | 1 |  |  | Support-System | NO | Closed |  | Brilliance Big Bore
CTUDT00626987 | Refresh | Closed | zyanjiang | November 4, 2019 4:16:37 PM GMT+08:00 | Ingenuity validation: Overlap display when changing UI from functional CT to Local. | WIP19 |  | User can click on arrow to close menu before moving to other screens. | Subsystem |  | NA | Works As Specified / Not A Defect | aswani_v | 0 |  |  | iCAP | NO | Closed |  | Ovation
CTUDT00626987 | Refresh | Closed | zyanjiang | November 4, 2019 4:16:37 PM GMT+08:00 | Ingenuity validation: Overlap display when changing UI from functional CT to Local. | WIP19 |  | User can click on arrow to close menu before moving to other screens. | Subsystem |  | NA | Works As Specified / Not A Defect | aswani_v | 0 |  |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00626987 | Refresh | Closed | zyanjiang | November 4, 2019 4:16:37 PM GMT+08:00 | Ingenuity validation: Overlap display when changing UI from functional CT to Local. | WIP19 |  | User can click on arrow to close menu before moving to other screens. | Subsystem |  | NA | Works As Specified / Not A Defect | aswani_v | 0 |  |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00626987 | Refresh | Closed | zyanjiang | November 4, 2019 4:16:37 PM GMT+08:00 | Ingenuity validation: Overlap display when changing UI from functional CT to Local. | WIP19 |  | User can click on arrow to close menu before moving to other screens. | Subsystem |  | NA | Works As Specified / Not A Defect | aswani_v | 0 |  |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00626987 | Refresh | Closed | zyanjiang | November 4, 2019 4:16:37 PM GMT+08:00 | Ingenuity validation: Overlap display when changing UI from functional CT to Local. | WIP19 |  | User can click on arrow to close menu before moving to other screens. | Subsystem |  | NA | Works As Specified / Not A Defect | aswani_v | 0 |  |  | iCAP | NO | Closed |  | IQon1
CTUDT00626984 | Refresh | Closed | shiraguppi_k | November 4, 2019 4:08:33 PM GMT+08:00 | Big Bore: BB Refresh Validation: In Preference Spectral Setting options available | BB_WIP9 | WIP22 |  | Subsystem |  | NA | Fixed | s_panigrahi | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP22SubCom | Brilliance Big Bore | NA | WIP22 | November 14, 2019 2:38:24 PM GMT+08:00
CTUDT00626983 | Refresh | Closed | zyanjiang | November 4, 2019 4:08:06 PM GMT+08:00 | Ingenuity validation:no need to record procedure description when using “Current” function | WIP19 |  |  | Subsystem | Procedure description | N/A | Duplicate | pandey_a | 0 |  |  | Console4 | NO | Closed | WAS | Ingenuity 64
CTUDT00626980 | Refresh | Closed | shiraguppi_k | November 4, 2019 4:02:08 PM GMT+08:00 | Big Bore: BB Refresh Validation: In Patient Registration page if you click on New its asking Overwrite current patient? | BB_WIP9 | BB_WIP12 |  | Subsystem | Patient Data | N/A | Fixed | pandey_a | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP11_SubCom | Brilliance Big Bore | NA | BB_WIP12 | November 6, 2019 10:55:33 PM GMT+08:00
CTUDT00626974 | Refresh | Closed | wang_l | November 4, 2019 2:29:48 PM GMT+08:00 | Refresh Reliability - “inializing…” take too much time after click Go button for the clinical series | WIP21 |  |  | Subsystem | Performance | n/a | No Plan to Fix | frider_o | 0 | 4 |  | DMS | NO | Closed | SubNC | Ovation
CTUDT00626974 | Refresh | Closed | wang_l | November 4, 2019 2:29:48 PM GMT+08:00 | Refresh Reliability - “inializing…” take too much time after click Go button for the clinical series | WIP21 |  |  | Subsystem | Performance | n/a | No Plan to Fix | frider_o | 0 | 4 |  | DMS | NO | Closed | SubNC | Brilliance Big Bore
CTUDT00626974 | Refresh | Closed | wang_l | November 4, 2019 2:29:48 PM GMT+08:00 | Refresh Reliability - “inializing…” take too much time after click Go button for the clinical series | WIP21 |  |  | Subsystem | Performance | n/a | No Plan to Fix | frider_o | 0 | 4 |  | DMS | NO | Closed | SubNC | Ingenuity 64
CTUDT00626974 | Refresh | Closed | wang_l | November 4, 2019 2:29:48 PM GMT+08:00 | Refresh Reliability - “inializing…” take too much time after click Go button for the clinical series | WIP21 |  |  | Subsystem | Performance | n/a | No Plan to Fix | frider_o | 0 | 4 |  | DMS | NO | Closed | SubNC | IQon1
CTUDT00626972 | Refresh | Closed | cherylli | November 3, 2019 9:55:23 PM GMT+08:00 | Refresh Val: Wrong display of Procedure name in Directory | BB_WIP9 |  |  | Subsystem |  | N/A | Works As Specified / Not A Defect | cherylli | 0 |  |  | Support-System | NO | Closed | WAS | Brilliance Big Bore
CTUDT00626970 | Refresh | Closed | shaham_t | November 3, 2019 6:52:22 PM GMT+08:00 | RegStand.567 or RegStand.568 untagging from Refresh | SRS | DHF298554 Rev09 DHF298553 Rev.09 |  | Subsystem | Requirements | RegStand.567 | Fixed | bandil_j | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | SubCom-19-Dec-2019 | Ovation |  | DHF298554 Rev09 DHF298553 Rev.09 | December 22, 2019 10:09:43 PM GMT+08:00
CTUDT00626970 | Refresh | Closed | shaham_t | November 3, 2019 6:52:22 PM GMT+08:00 | RegStand.567 or RegStand.568 untagging from Refresh | SRS | DHF298554 Rev09 DHF298553 Rev.09 |  | Subsystem | Requirements | RegStand.567 | Fixed | bandil_j | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | SubCom-19-Dec-2019 | Brilliance Big Bore |  | DHF298554 Rev09 DHF298553 Rev.09 | December 22, 2019 10:09:43 PM GMT+08:00
CTUDT00626970 | Refresh | Closed | shaham_t | November 3, 2019 6:52:22 PM GMT+08:00 | RegStand.567 or RegStand.568 untagging from Refresh | SRS | DHF298554 Rev09 DHF298553 Rev.09 |  | Subsystem | Requirements | RegStand.567 | Fixed | bandil_j | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | SubCom-19-Dec-2019 | Ingenuity 64 |  | DHF298554 Rev09 DHF298553 Rev.09 | December 22, 2019 10:09:43 PM GMT+08:00
CTUDT00626965 | Refresh | Closed | zyanjiang | November 3, 2019 4:34:38 PM GMT+08:00 | Ingenuity validation:inappropriate pop up message for not match between age group and Exam card | WIP19 |  | Not Required. | Subsystem | Rule System | NA | Works As Specified / Not A Defect | kapadia_v | 0 |  |  | Console4 | NO | Closed | WAS | Ingenuity 64
CTUDT00626964 | Refresh | Closed | zyanjiang | November 3, 2019 4:34:11 PM GMT+08:00 | Ingenuity validation:when tilt the gantry the increment value will increase for brain axial(2D/3D) scan | WIP19 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | zyanjiang | 0 |  |  |  |  | Closed |  | Ingenuity 64
CTUDT00626961 | Refresh | Closed | zyanjiang | November 3, 2019 4:24:34 PM GMT+08:00 | Ingenuity validation:No new study function for completed studies | WIP19 |  |  | Subsystem |  |  | Change Proposal | zyanjiang | 0 |  |  |  |  | Closed |  | Ingenuity 64
CTUDT00626960 | Refresh | Closed | zyanjiang | November 3, 2019 4:24:06 PM GMT+08:00 | Ingenuity validation:IMR information can’t be displayed like iDose | WIP19 |  |  | Subsystem |  |  | Change Proposal | zyanjiang | 0 |  |  |  |  | Closed |  | Ingenuity 64
CTUDT00626958 | Refresh | Closed | zyanjiang | November 3, 2019 4:17:57 PM GMT+08:00 | Ingenuity validation:Raw data always display as “0” | WIP19 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | zyanjiang | 0 |  |  |  |  | Closed |  | Ingenuity 64
CTUDT00626956 | Refresh | Closed | zyanjiang | November 3, 2019 4:08:28 PM GMT+08:00 | Ingenuity validation:there is a quick flash when use quick review to load radiation dose information(SR) | WIP19 |  | NA | Subsystem | EBW | NA | Works As Specified / Not A Defect | aswani_v | 0 |  |  | iCAP | NO | Closed |  | Ovation
CTUDT00626956 | Refresh | Closed | zyanjiang | November 3, 2019 4:08:28 PM GMT+08:00 | Ingenuity validation:there is a quick flash when use quick review to load radiation dose information(SR) | WIP19 |  | NA | Subsystem | EBW | NA | Works As Specified / Not A Defect | aswani_v | 0 |  |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00626956 | Refresh | Closed | zyanjiang | November 3, 2019 4:08:28 PM GMT+08:00 | Ingenuity validation:there is a quick flash when use quick review to load radiation dose information(SR) | WIP19 |  | NA | Subsystem | EBW | NA | Works As Specified / Not A Defect | aswani_v | 0 |  |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00626956 | Refresh | Closed | zyanjiang | November 3, 2019 4:08:28 PM GMT+08:00 | Ingenuity validation:there is a quick flash when use quick review to load radiation dose information(SR) | WIP19 |  | NA | Subsystem | EBW | NA | Works As Specified / Not A Defect | aswani_v | 0 |  |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00626956 | Refresh | Closed | zyanjiang | November 3, 2019 4:08:28 PM GMT+08:00 | Ingenuity validation:there is a quick flash when use quick review to load radiation dose information(SR) | WIP19 |  | NA | Subsystem | EBW | NA | Works As Specified / Not A Defect | aswani_v | 0 |  |  | iCAP | NO | Closed |  | IQon1
CTUDT00626954 | Refresh | Closed | zyanjiang | November 3, 2019 3:58:17 PM GMT+08:00 | Ingenuity validation:when age information filled,the age group should be fixed | WIP19 |  | Not Required. | Subsystem |  | NA | Works As Specified / Not A Defect | kapadia_v | 0 |  |  | Console4 | NO | Closed | WAS | Ingenuity 64
CTUDT00626952 | Refresh | Closed | zyanjiang | November 3, 2019 3:40:59 PM GMT+08:00 | Ingenuity validation:constancy test failed | WIP19 |  | Run HCOR and re-perform Acceptance/Constancy | Subsystem |  | N/A | Works As Specified / Not A Defect | zyanjiang | 1 |  |  | Support-System | NO | Closed |  | Ingenuity 64
CTUDT00626950 | Refresh | Closed | chang_zhiguo | November 2, 2019 8:26:47 AM GMT+08:00 | BB refresh validation: Reportable ring artifact for knee scan | BB_WIP9 |  | The end user would have the ability to perform offline reconstruction if they did not like the reconstruction with YD. The end user could use YC + Enhancement in order to avoid rings. | System |  | N/A | Works As Specified / Not A Defect | richards_b | 0 |  |  | Support-Physics | NO | Closed |  | Ovation
CTUDT00626950 | Refresh | Closed | chang_zhiguo | November 2, 2019 8:26:47 AM GMT+08:00 | BB refresh validation: Reportable ring artifact for knee scan | BB_WIP9 |  | The end user would have the ability to perform offline reconstruction if they did not like the reconstruction with YD. The end user could use YC + Enhancement in order to avoid rings. | System |  | N/A | Works As Specified / Not A Defect | richards_b | 0 |  |  | Support-Physics | NO | Closed |  | Brilliance 64
CTUDT00626950 | Refresh | Closed | chang_zhiguo | November 2, 2019 8:26:47 AM GMT+08:00 | BB refresh validation: Reportable ring artifact for knee scan | BB_WIP9 |  | The end user would have the ability to perform offline reconstruction if they did not like the reconstruction with YD. The end user could use YC + Enhancement in order to avoid rings. | System |  | N/A | Works As Specified / Not A Defect | richards_b | 0 |  |  | Support-Physics | NO | Closed |  | Brilliance Big Bore
CTUDT00626950 | Refresh | Closed | chang_zhiguo | November 2, 2019 8:26:47 AM GMT+08:00 | BB refresh validation: Reportable ring artifact for knee scan | BB_WIP9 |  | The end user would have the ability to perform offline reconstruction if they did not like the reconstruction with YD. The end user could use YC + Enhancement in order to avoid rings. | System |  | N/A | Works As Specified / Not A Defect | richards_b | 0 |  |  | Support-Physics | NO | Closed |  | Ingenuity 64
CTUDT00626950 | Refresh | Closed | chang_zhiguo | November 2, 2019 8:26:47 AM GMT+08:00 | BB refresh validation: Reportable ring artifact for knee scan | BB_WIP9 |  | The end user would have the ability to perform offline reconstruction if they did not like the reconstruction with YD. The end user could use YC + Enhancement in order to avoid rings. | System |  | N/A | Works As Specified / Not A Defect | richards_b | 0 |  |  | Support-Physics | NO | Closed |  | IQon1
CTUDT00626948 | Refresh | Closed | chang_zhiguo | November 2, 2019 8:11:57 AM GMT+08:00 | BB refresh validation: The thickness available options not updating accordingly after the rendering mode change from MIP/minIP to Average. | BB_WIP9 | BB_WIP12 WIP22 |  | Subsystem |  | NA | Fixed | datta_i | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP11_SubCom | Ovation |  | BB_WIP12 WIP22 | November 7, 2019 9:18:08 PM GMT+08:00
CTUDT00626948 | Refresh | Closed | chang_zhiguo | November 2, 2019 8:11:57 AM GMT+08:00 | BB refresh validation: The thickness available options not updating accordingly after the rendering mode change from MIP/minIP to Average. | BB_WIP9 | BB_WIP12 WIP22 |  | Subsystem |  | NA | Fixed | datta_i | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP11_SubCom | Brilliance 64 |  | BB_WIP12 WIP22 | November 7, 2019 9:18:08 PM GMT+08:00
CTUDT00626948 | Refresh | Closed | chang_zhiguo | November 2, 2019 8:11:57 AM GMT+08:00 | BB refresh validation: The thickness available options not updating accordingly after the rendering mode change from MIP/minIP to Average. | BB_WIP9 | BB_WIP12 WIP22 |  | Subsystem |  | NA | Fixed | datta_i | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP11_SubCom | Brilliance Big Bore |  | BB_WIP12 WIP22 | November 7, 2019 9:18:08 PM GMT+08:00
CTUDT00626948 | Refresh | Closed | chang_zhiguo | November 2, 2019 8:11:57 AM GMT+08:00 | BB refresh validation: The thickness available options not updating accordingly after the rendering mode change from MIP/minIP to Average. | BB_WIP9 | BB_WIP12 WIP22 |  | Subsystem |  | NA | Fixed | datta_i | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP11_SubCom | Ingenuity 64 |  | BB_WIP12 WIP22 | November 7, 2019 9:18:08 PM GMT+08:00
CTUDT00626948 | Refresh | Closed | chang_zhiguo | November 2, 2019 8:11:57 AM GMT+08:00 | BB refresh validation: The thickness available options not updating accordingly after the rendering mode change from MIP/minIP to Average. | BB_WIP9 | BB_WIP12 WIP22 |  | Subsystem |  | NA | Fixed | datta_i | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP11_SubCom | IQon1 |  | BB_WIP12 WIP22 | November 7, 2019 9:18:08 PM GMT+08:00
CTUDT00626946 | Refresh | Closed | waggoner_j | November 2, 2019 3:39:20 AM GMT+08:00 | IFU inconsistency for refresh products in the cardiac IMR section | ifu version a | v4.8 | N/A | Subsystem | Verified1 | N/A | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | 20-Dec-2019 | Brilliance Big Bore |  | v4.8 | December 17, 2019 8:01:35 PM GMT+08:00
CTUDT00626944 | Refresh | Closed | zhu_je | November 1, 2019 3:04:14 PM GMT+08:00 | [Refresh] The Image annotation was “Ingenuity CT” for brilliance 64 | WIP21 | WIP22 |  | Subsystem |  | NA | Fixed | prasad_r | 0 | 2 | Console Software | Console4 | NO | Closed | WIP22SubCom | Brilliance 64 |  | WIP22 | November 6, 2019 1:43:54 PM GMT+08:00
CTUDT00626940 | Refresh | Closed | yosfan_s | October 31, 2019 9:09:26 PM GMT+08:00 | Refresh: Beam and DMS Check: Mismatch between the slice direction in the plot tooltip and in the tables of failures | WIP21 |  |  | Subsystem |  | N/A | No Plan to Fix | yatziv_l | 0 | 3 |  | DMS | NO | Closed | SubNC | Ovation
CTUDT00626940 | Refresh | Closed | yosfan_s | October 31, 2019 9:09:26 PM GMT+08:00 | Refresh: Beam and DMS Check: Mismatch between the slice direction in the plot tooltip and in the tables of failures | WIP21 |  |  | Subsystem |  | N/A | No Plan to Fix | yatziv_l | 0 | 3 |  | DMS | NO | Closed | SubNC | Brilliance Big Bore
CTUDT00626940 | Refresh | Closed | yosfan_s | October 31, 2019 9:09:26 PM GMT+08:00 | Refresh: Beam and DMS Check: Mismatch between the slice direction in the plot tooltip and in the tables of failures | WIP21 |  |  | Subsystem |  | N/A | No Plan to Fix | yatziv_l | 0 | 3 |  | DMS | NO | Closed | SubNC | Ingenuity 64
CTUDT00626940 | Refresh | Closed | yosfan_s | October 31, 2019 9:09:26 PM GMT+08:00 | Refresh: Beam and DMS Check: Mismatch between the slice direction in the plot tooltip and in the tables of failures | WIP21 |  |  | Subsystem |  | N/A | No Plan to Fix | yatziv_l | 0 | 3 |  | DMS | NO | Closed | SubNC | IQon1
CTUDT00626938 | Refresh | Closed | PremKumar_Rajendran | October 31, 2019 8:34:50 PM GMT+08:00 | BB Refresh : Localization issues in Interventional Control on console for few texts. | BB_WIP11 | BB_WIP12 |  | Subsystem |  | NA | Fixed | Snehashis_Dey | 0 | 2 | Console Software | Console4 | NO | Closed | WIP22_SubCom | Brilliance Big Bore |  | BB_WIP12 | November 7, 2019 9:44:48 PM GMT+08:00
CTUDT00626933 | Refresh | Closed | Valeria_Kashtanov | October 31, 2019 3:53:21 PM GMT+08:00 | Refresh: Performance Test: Tube Data doesn't appear on Header in Html Report | BB_WIP11 |  |  | Subsystem |  | DHF170396 Rev 21 | No Plan to Fix | yatziv_l | 0 | 3 |  | DMS | NO | Closed | SubNC | Ovation
CTUDT00626933 | Refresh | Closed | Valeria_Kashtanov | October 31, 2019 3:53:21 PM GMT+08:00 | Refresh: Performance Test: Tube Data doesn't appear on Header in Html Report | BB_WIP11 |  |  | Subsystem |  | DHF170396 Rev 21 | No Plan to Fix | yatziv_l | 0 | 3 |  | DMS | NO | Closed | SubNC | Brilliance Big Bore
CTUDT00626933 | Refresh | Closed | Valeria_Kashtanov | October 31, 2019 3:53:21 PM GMT+08:00 | Refresh: Performance Test: Tube Data doesn't appear on Header in Html Report | BB_WIP11 |  |  | Subsystem |  | DHF170396 Rev 21 | No Plan to Fix | yatziv_l | 0 | 3 |  | DMS | NO | Closed | SubNC | Ingenuity 64
CTUDT00626933 | Refresh | Closed | Valeria_Kashtanov | October 31, 2019 3:53:21 PM GMT+08:00 | Refresh: Performance Test: Tube Data doesn't appear on Header in Html Report | BB_WIP11 |  |  | Subsystem |  | DHF170396 Rev 21 | No Plan to Fix | yatziv_l | 0 | 3 |  | DMS | NO | Closed | SubNC | IQon1
CTUDT00626931 | Refresh | Closed | zhong_x | October 31, 2019 1:12:42 PM GMT+08:00 | There was an error “The data that was loaded is not valid. Please reselect data” when load the Dual Energy data in ISP10 workstation. | BB_WIP10 |  |  | Subsystem | CRN | NA | Works As Specified / Not A Defect | zhong_x | 0 |  |  | Support-System | NO | Closed | WAS | Brilliance Big Bore
CTUDT00626929 | Refresh | Closed | 320053769 | October 30, 2019 9:32:48 PM GMT+08:00 | WIP21 - Exam record creation time in IQon is more than 2 seconds | WIP21 |  | No workaround. User has to wait out the delay for about 2-3s. | Subsystem | Performance | NA | No Plan to Fix | ravindra_a | 0 | 2 |  | Console4 | NO | Closed | SubNC | Ovation
CTUDT00626929 | Refresh | Closed | 320053769 | October 30, 2019 9:32:48 PM GMT+08:00 | WIP21 - Exam record creation time in IQon is more than 2 seconds | WIP21 |  | No workaround. User has to wait out the delay for about 2-3s. | Subsystem | Performance | NA | No Plan to Fix | ravindra_a | 0 | 2 |  | Console4 | NO | Closed | SubNC | Brilliance 64
CTUDT00626929 | Refresh | Closed | 320053769 | October 30, 2019 9:32:48 PM GMT+08:00 | WIP21 - Exam record creation time in IQon is more than 2 seconds | WIP21 |  | No workaround. User has to wait out the delay for about 2-3s. | Subsystem | Performance | NA | No Plan to Fix | ravindra_a | 0 | 2 |  | Console4 | NO | Closed | SubNC | Brilliance Big Bore
CTUDT00626929 | Refresh | Closed | 320053769 | October 30, 2019 9:32:48 PM GMT+08:00 | WIP21 - Exam record creation time in IQon is more than 2 seconds | WIP21 |  | No workaround. User has to wait out the delay for about 2-3s. | Subsystem | Performance | NA | No Plan to Fix | ravindra_a | 0 | 2 |  | Console4 | NO | Closed | SubNC | Ingenuity 64
CTUDT00626929 | Refresh | Closed | 320053769 | October 30, 2019 9:32:48 PM GMT+08:00 | WIP21 - Exam record creation time in IQon is more than 2 seconds | WIP21 |  | No workaround. User has to wait out the delay for about 2-3s. | Subsystem | Performance | NA | No Plan to Fix | ravindra_a | 0 | 2 |  | Console4 | NO | Closed | SubNC | IQon1
CTUDT00626926 | Refresh | Closed | Noy_Seelig | October 30, 2019 6:40:02 PM GMT+08:00 | Refresh: Common Functions: Help Page for Manual Centering in Harmonized Phantom Appears as Legacy | WIP 21 | WIP22 |  | Subsystem | Verified1 | n/a | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP22SubCom | Ovation |  | WIP22 | November 7, 2019 10:52:42 PM GMT+08:00
CTUDT00626926 | Refresh | Closed | Noy_Seelig | October 30, 2019 6:40:02 PM GMT+08:00 | Refresh: Common Functions: Help Page for Manual Centering in Harmonized Phantom Appears as Legacy | WIP 21 | WIP22 |  | Subsystem | Verified1 | n/a | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP22SubCom | Brilliance Big Bore |  | WIP22 | November 7, 2019 10:52:42 PM GMT+08:00
CTUDT00626926 | Refresh | Closed | Noy_Seelig | October 30, 2019 6:40:02 PM GMT+08:00 | Refresh: Common Functions: Help Page for Manual Centering in Harmonized Phantom Appears as Legacy | WIP 21 | WIP22 |  | Subsystem | Verified1 | n/a | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP22SubCom | Ingenuity 64 |  | WIP22 | November 7, 2019 10:52:42 PM GMT+08:00
CTUDT00626926 | Refresh | Closed | Noy_Seelig | October 30, 2019 6:40:02 PM GMT+08:00 | Refresh: Common Functions: Help Page for Manual Centering in Harmonized Phantom Appears as Legacy | WIP 21 | WIP22 |  | Subsystem | Verified1 | n/a | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP22SubCom | IQon1 |  | WIP22 | November 7, 2019 10:52:42 PM GMT+08:00
CTUDT00626922 | Refresh | Closed | Valeria_Kashtanov | October 29, 2019 9:48:58 PM GMT+08:00 | Refresh: Performance Test: ROI sizes for measurements section in spec should be clarified | WIP20 | DHF170396 Rev 23 | N/A | System |  | DHF170396 rev 21 | System Repaired | dongping_w | 0 | 2 | Design Documentation | Support-Physics | NO | Closed | Spec update | Ovation |  | DHF170396 Rev 23 | January 13, 2020 2:07:36 PM GMT+08:00
CTUDT00626922 | Refresh | Closed | Valeria_Kashtanov | October 29, 2019 9:48:58 PM GMT+08:00 | Refresh: Performance Test: ROI sizes for measurements section in spec should be clarified | WIP20 | DHF170396 Rev 23 | N/A | System |  | DHF170396 rev 21 | System Repaired | dongping_w | 0 | 2 | Design Documentation | Support-Physics | NO | Closed | Spec update | Brilliance 64 |  | DHF170396 Rev 23 | January 13, 2020 2:07:36 PM GMT+08:00
CTUDT00626922 | Refresh | Closed | Valeria_Kashtanov | October 29, 2019 9:48:58 PM GMT+08:00 | Refresh: Performance Test: ROI sizes for measurements section in spec should be clarified | WIP20 | DHF170396 Rev 23 | N/A | System |  | DHF170396 rev 21 | System Repaired | dongping_w | 0 | 2 | Design Documentation | Support-Physics | NO | Closed | Spec update | Brilliance Big Bore |  | DHF170396 Rev 23 | January 13, 2020 2:07:36 PM GMT+08:00
CTUDT00626922 | Refresh | Closed | Valeria_Kashtanov | October 29, 2019 9:48:58 PM GMT+08:00 | Refresh: Performance Test: ROI sizes for measurements section in spec should be clarified | WIP20 | DHF170396 Rev 23 | N/A | System |  | DHF170396 rev 21 | System Repaired | dongping_w | 0 | 2 | Design Documentation | Support-Physics | NO | Closed | Spec update | Ingenuity 64 |  | DHF170396 Rev 23 | January 13, 2020 2:07:36 PM GMT+08:00
CTUDT00626922 | Refresh | Closed | Valeria_Kashtanov | October 29, 2019 9:48:58 PM GMT+08:00 | Refresh: Performance Test: ROI sizes for measurements section in spec should be clarified | WIP20 | DHF170396 Rev 23 | N/A | System |  | DHF170396 rev 21 | System Repaired | dongping_w | 0 | 2 | Design Documentation | Support-Physics | NO | Closed | Spec update | IQon1 |  | DHF170396 Rev 23 | January 13, 2020 2:07:36 PM GMT+08:00
CTUDT00626920 | Refresh | Closed | levy_s | October 29, 2019 7:34:43 PM GMT+08:00 | Refresh: Estop was opened while scanning dual Surview | WIP20 |  |  | Subsystem | Gantry |  | Works As Specified / Not A Defect | tamari_o | 1 |  |  | Support-System | YES | Closed |  | Ovation
CTUDT00626916 | Refresh | Closed | yang_h | October 29, 2019 11:10:45 AM GMT+08:00 | Refresh: In bolus tracking scan PTD changes from 2 second to 1 second after moving the clinical acquisition plan box slightly | WIP19 |  |  | Subsystem | Post Threshold Delay; Blocker | N/A | No Plan to Fix | yang_h | 0 | 4 |  | GCX | NO | Closed |  | Ovation
CTUDT00626916 | Refresh | Closed | yang_h | October 29, 2019 11:10:45 AM GMT+08:00 | Refresh: In bolus tracking scan PTD changes from 2 second to 1 second after moving the clinical acquisition plan box slightly | WIP19 |  |  | Subsystem | Post Threshold Delay; Blocker | N/A | No Plan to Fix | yang_h | 0 | 4 |  | GCX | NO | Closed |  | IQon1
CTUDT00626912 | Refresh | Closed | 320053769 | October 28, 2019 10:15:37 PM GMT+08:00 | While scanning abdomen EC, scan was paused and “A problem has occurred, if the problem persist, contact service” error appeared. | WIP21 |  |  | Subsystem | Gantry |  | Works As Specified / Not A Defect | tamari_o | 1 |  |  | Support-System | YES | Closed |  | IQon1
CTUDT00626885 | Refresh | Closed | barnea_d | October 27, 2019 10:10:07 PM GMT+08:00 | Refresh: WIP 20: iCT:Calibrations: Performance Test: The Center ROI in Head pins1 scan is elliptical. | WIP20 | WIP22 |  | Subsystem | Verified1 | DHF170396 Rev 20 | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP22SubCom | Ovation |  | WIP22 | November 12, 2019 11:24:03 PM GMT+08:00
CTUDT00626875 | Refresh | Closed | 320053769 | October 24, 2019 7:51:14 PM GMT+08:00 | Missing imported EC - WF doesn’t load all exam cards to Patient scene | WIP21 | WIP21.1 BB_WIP11.1 |  | System | Performance;Verified1 | NA | Fixed | tamari_o | 0 | 3 | Console Software | Console4 | NO | Closed | WIP21SegReq | Ovation |  | WIP21.1 BB_WIP11.1 | October 28, 2019 9:03:31 PM GMT+08:00
CTUDT00626875 | Refresh | Closed | 320053769 | October 24, 2019 7:51:14 PM GMT+08:00 | Missing imported EC - WF doesn’t load all exam cards to Patient scene | WIP21 | WIP21.1 BB_WIP11.1 |  | System | Performance;Verified1 | NA | Fixed | tamari_o | 0 | 3 | Console Software | Console4 | NO | Closed | WIP21SegReq | Brilliance 64 |  | WIP21.1 BB_WIP11.1 | October 28, 2019 9:03:31 PM GMT+08:00
CTUDT00626875 | Refresh | Closed | 320053769 | October 24, 2019 7:51:14 PM GMT+08:00 | Missing imported EC - WF doesn’t load all exam cards to Patient scene | WIP21 | WIP21.1 BB_WIP11.1 |  | System | Performance;Verified1 | NA | Fixed | tamari_o | 0 | 3 | Console Software | Console4 | NO | Closed | WIP21SegReq | Brilliance Big Bore |  | WIP21.1 BB_WIP11.1 | October 28, 2019 9:03:31 PM GMT+08:00
CTUDT00626875 | Refresh | Closed | 320053769 | October 24, 2019 7:51:14 PM GMT+08:00 | Missing imported EC - WF doesn’t load all exam cards to Patient scene | WIP21 | WIP21.1 BB_WIP11.1 |  | System | Performance;Verified1 | NA | Fixed | tamari_o | 0 | 3 | Console Software | Console4 | NO | Closed | WIP21SegReq | Ingenuity 64 |  | WIP21.1 BB_WIP11.1 | October 28, 2019 9:03:31 PM GMT+08:00
CTUDT00626875 | Refresh | Closed | 320053769 | October 24, 2019 7:51:14 PM GMT+08:00 | Missing imported EC - WF doesn’t load all exam cards to Patient scene | WIP21 | WIP21.1 BB_WIP11.1 |  | System | Performance;Verified1 | NA | Fixed | tamari_o | 0 | 3 | Console Software | Console4 | NO | Closed | WIP21SegReq | IQon1 |  | WIP21.1 BB_WIP11.1 | October 28, 2019 9:03:31 PM GMT+08:00
CTUDT00626872 | Refresh | Closed | wang_l | October 24, 2019 4:35:36 PM GMT+08:00 | Refresh Ingenuity: "X-ray system cannot comply. Please retry." error during dual surivew | WIP19 | WIP22 |  | Subsystem |  | NA | Fixed | wang_k | 0 | 3 | Gantry Software | Support-System | NO | Closed | WIP23.1SubCom | Ingenuity 64 |  | WIP22 | January 10, 2020 1:08:48 PM GMT+08:00
CTUDT00626869 | Refresh | Closed | Chakradhar_Betha | October 24, 2019 4:14:23 PM GMT+08:00 | [IB]There is no ELSCINTCT_PHASE_POINTS (01F7:108B) available in the DICOM information for Calcium Scoring 3D with Edit before Final Recon Turned off | 4.1.10 | DHF303211 Rev18 | User can enable/check the Edit Before Final Recon option in show all of the Calcium Scoring 3D scan. | System |  | Console-SRS.17856 | Fixed | roy_r | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | SSRS | Ovation |  | DHF303211 Rev18 | October 25, 2019 12:39:01 PM GMT+08:00
CTUDT00626869 | Refresh | Closed | Chakradhar_Betha | October 24, 2019 4:14:23 PM GMT+08:00 | [IB]There is no ELSCINTCT_PHASE_POINTS (01F7:108B) available in the DICOM information for Calcium Scoring 3D with Edit before Final Recon Turned off | 4.1.10 | DHF303211 Rev18 | User can enable/check the Edit Before Final Recon option in show all of the Calcium Scoring 3D scan. | System |  | Console-SRS.17856 | Fixed | roy_r | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | SSRS | Brilliance 64 |  | DHF303211 Rev18 | October 25, 2019 12:39:01 PM GMT+08:00
CTUDT00626869 | Refresh | Closed | Chakradhar_Betha | October 24, 2019 4:14:23 PM GMT+08:00 | [IB]There is no ELSCINTCT_PHASE_POINTS (01F7:108B) available in the DICOM information for Calcium Scoring 3D with Edit before Final Recon Turned off | 4.1.10 | DHF303211 Rev18 | User can enable/check the Edit Before Final Recon option in show all of the Calcium Scoring 3D scan. | System |  | Console-SRS.17856 | Fixed | roy_r | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | SSRS | Ingenuity 64 |  | DHF303211 Rev18 | October 25, 2019 12:39:01 PM GMT+08:00
CTUDT00626869 | Refresh | Closed | Chakradhar_Betha | October 24, 2019 4:14:23 PM GMT+08:00 | [IB]There is no ELSCINTCT_PHASE_POINTS (01F7:108B) available in the DICOM information for Calcium Scoring 3D with Edit before Final Recon Turned off | 4.1.10 | DHF303211 Rev18 | User can enable/check the Edit Before Final Recon option in show all of the Calcium Scoring 3D scan. | System |  | Console-SRS.17856 | Fixed | roy_r | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | SSRS | IQon1 |  | DHF303211 Rev18 | October 25, 2019 12:39:01 PM GMT+08:00
CTUDT00626862 | Refresh | Closed | roy_r | October 24, 2019 12:25:59 PM GMT+08:00 | Console requirement for Data Buffer Overflow needs update | WIP20 | DHF303211 Rev18 | N/A | Subsystem |  | Console-SRS.5685 | Fixed | roy_r | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | SSRS | Ovation |  | DHF303211 Rev18 | October 25, 2019 12:38:37 PM GMT+08:00
CTUDT00626862 | Refresh | Closed | roy_r | October 24, 2019 12:25:59 PM GMT+08:00 | Console requirement for Data Buffer Overflow needs update | WIP20 | DHF303211 Rev18 | N/A | Subsystem |  | Console-SRS.5685 | Fixed | roy_r | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | SSRS | IQon1 |  | DHF303211 Rev18 | October 25, 2019 12:38:37 PM GMT+08:00
CTUDT00626859 | Refresh | Closed | roy_r | October 24, 2019 12:15:10 PM GMT+08:00 | BB Refresh: Result Planning Limit Spec in Big Bore needs to be improved | WIP20 | BB_WIP11.1 | NA | Subsystem |  | Console-SRS.16080 | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP11.1_SubCom | Brilliance Big Bore |  | BB_WIP11.1 | October 24, 2019 9:20:55 PM GMT+08:00
CTUDT00626853 | Refresh | Closed | marino | October 24, 2019 3:16:02 AM GMT+08:00 | The CIRS SW SSRS is being updated but the system level changes to DOORS have not been implemented. | WIP20 | DHF297589 Rev 17 DHF319071 Rev 06 | N/A | Subsystem | Requirements | N/A | Fixed | richards_b | 0 | 2 | Design Documentation | CIRS | NO | Closed | DocumentationUpdate | Brilliance 64 |  | DHF297589 Rev 17 DHF319071 Rev 06 | January 9, 2020 5:03:36 AM GMT+08:00
CTUDT00626853 | Refresh | Closed | marino | October 24, 2019 3:16:02 AM GMT+08:00 | The CIRS SW SSRS is being updated but the system level changes to DOORS have not been implemented. | WIP20 | DHF297589 Rev 17 DHF319071 Rev 06 | N/A | Subsystem | Requirements | N/A | Fixed | richards_b | 0 | 2 | Design Documentation | CIRS | NO | Closed | DocumentationUpdate | Brilliance Big Bore |  | DHF297589 Rev 17 DHF319071 Rev 06 | January 9, 2020 5:03:36 AM GMT+08:00
CTUDT00626853 | Refresh | Closed | marino | October 24, 2019 3:16:02 AM GMT+08:00 | The CIRS SW SSRS is being updated but the system level changes to DOORS have not been implemented. | WIP20 | DHF297589 Rev 17 DHF319071 Rev 06 | N/A | Subsystem | Requirements | N/A | Fixed | richards_b | 0 | 2 | Design Documentation | CIRS | NO | Closed | DocumentationUpdate | iCT4 |  | DHF297589 Rev 17 DHF319071 Rev 06 | January 9, 2020 5:03:36 AM GMT+08:00
CTUDT00626853 | Refresh | Closed | marino | October 24, 2019 3:16:02 AM GMT+08:00 | The CIRS SW SSRS is being updated but the system level changes to DOORS have not been implemented. | WIP20 | DHF297589 Rev 17 DHF319071 Rev 06 | N/A | Subsystem | Requirements | N/A | Fixed | richards_b | 0 | 2 | Design Documentation | CIRS | NO | Closed | DocumentationUpdate | Ingenuity 64 |  | DHF297589 Rev 17 DHF319071 Rev 06 | January 9, 2020 5:03:36 AM GMT+08:00
CTUDT00626853 | Refresh | Closed | marino | October 24, 2019 3:16:02 AM GMT+08:00 | The CIRS SW SSRS is being updated but the system level changes to DOORS have not been implemented. | WIP20 | DHF297589 Rev 17 DHF319071 Rev 06 | N/A | Subsystem | Requirements | N/A | Fixed | richards_b | 0 | 2 | Design Documentation | CIRS | NO | Closed | DocumentationUpdate | IQon1 |  | DHF297589 Rev 17 DHF319071 Rev 06 | January 9, 2020 5:03:36 AM GMT+08:00
CTUDT00626842 | Refresh | Closed | wang_l | October 23, 2019 8:24:22 PM GMT+08:00 | Refresh BB - View2 tab not closed normally and many process of ViewTwo generated | BB_WIP10 | BB_WIP11.1 |  | Subsystem |  | N/A | Fixed | navale_s | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP12_SegReq | Brilliance Big Bore |  | BB_WIP11.1 | October 24, 2019 9:52:53 PM GMT+08:00
CTUDT00626836 | Refresh | Closed | cherylli | October 23, 2019 3:37:52 PM GMT+08:00 | Refresh Pre-Val: Disable tilt doesn’t work | WIP21 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | cherylli | 0 |  |  |  |  | Closed |  | Ingenuity 64
CTUDT00626833 | Refresh | Closed | balasubramaian_m | October 23, 2019 2:37:22 PM GMT+08:00 | Reliability Requirements to be updated for Big Bore and Ingenuity | SRS DHF298554,DHF298555 | DHF298554 Rev10 |  | Subsystem |  | syssupport.920 & syssupport.528 | Fixed | dongping_w | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | SubCom-20-Dec-2019 | Brilliance Big Bore |  | DHF298554 Rev10 | December 23, 2019 8:07:50 PM GMT+08:00
CTUDT00626833 | Refresh | Closed | balasubramaian_m | October 23, 2019 2:37:22 PM GMT+08:00 | Reliability Requirements to be updated for Big Bore and Ingenuity | SRS DHF298554,DHF298555 | DHF298554 Rev10 |  | Subsystem |  | syssupport.920 & syssupport.528 | Fixed | dongping_w | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | SubCom-20-Dec-2019 | Ingenuity 64 |  | DHF298554 Rev10 | December 23, 2019 8:07:50 PM GMT+08:00
CTUDT00626831 | Refresh | Closed | balasubramaian_m | October 23, 2019 2:09:34 PM GMT+08:00 | Updates required to Reliability Requirements for IQon, iCT/iCT-SP -CCT-NM.SRS.SysSupport.527 and CCT-NM.SRS.SysSupport.696 has to be un-tagged from Refresh. | SRS DHF298553,DHF298556 |  |  | Subsystem |  |  | Does Not Apply | guy_sheaffer | 0 |  |  | Support-System | NO | Closed |  | Ovation
CTUDT00626831 | Refresh | Closed | balasubramaian_m | October 23, 2019 2:09:34 PM GMT+08:00 | Updates required to Reliability Requirements for IQon, iCT/iCT-SP -CCT-NM.SRS.SysSupport.527 and CCT-NM.SRS.SysSupport.696 has to be un-tagged from Refresh. | SRS DHF298553,DHF298556 |  |  | Subsystem |  |  | Does Not Apply | guy_sheaffer | 0 |  |  | Support-System | NO | Closed |  | IQon1
CTUDT00626826 | Refresh | Closed | karkour | October 23, 2019 10:42:49 AM GMT+08:00 | Gantry error S_MC_AXIS_JOG_FAILURE when using IVC HW near Couch limits | BB_WIP10 | WIP22 |  | Subsystem |  | Console-SRS.16541 | Fixed | cx_zhang | 0 | 4 | Couch Software | GCX | NO | Closed | BB_WIP12_SubCom | Brilliance Big Bore |  | WIP22 | November 8, 2019 10:07:18 AM GMT+08:00
CTUDT00626824 | Refresh | Closed | finestine_m | October 22, 2019 9:56:03 PM GMT+08:00 | Refresh: User Access: Manual Hcor Utility Is Missing For O-Level User | WIP20 | DHF232540 Rev15 |  | Subsystem | Reason matrix | DHF232540 Rev 14 | Fixed | richards_b | 0 | 2 | Design Documentation | Support-Physics | NO | Closed | DHF232540 Rev15 | Ovation |  | DHF232540 Rev15 | November 16, 2019 3:58:03 AM GMT+08:00
CTUDT00626818 | Refresh | Closed | barnea_d | October 22, 2019 7:48:00 PM GMT+08:00 | Refresh: User Access: Physics Analysis information is missing for iCT in rev. 14 of DHF232540 | WIP20 | DHF232540 Rev15 |  | Subsystem | reason matrix | DHF232540 Rev 14 | Fixed | richards_b | 0 | 2 | Design Documentation | Support-Physics | NO | Closed | DHF232540 Rev15 | Ovation |  | DHF232540 Rev15 | November 16, 2019 3:57:50 AM GMT+08:00
CTUDT00626806 | Refresh | Closed | vineetha_s | October 22, 2019 6:13:01 PM GMT+08:00 | "Not enough space on the system to download" message was not localized to IT user | WIP20 |  | NA | Subsystem |  | NA | No Plan to Fix | sagar_v | 0 | 4 |  | Console4 | NO | Closed | SubNC | iCT4
CTUDT00626806 | Refresh | Closed | vineetha_s | October 22, 2019 6:13:01 PM GMT+08:00 | "Not enough space on the system to download" message was not localized to IT user | WIP20 |  | NA | Subsystem |  | NA | No Plan to Fix | sagar_v | 0 | 4 |  | Console4 | NO | Closed | SubNC | Ingenuity 64
CTUDT00626806 | Refresh | Closed | vineetha_s | October 22, 2019 6:13:01 PM GMT+08:00 | "Not enough space on the system to download" message was not localized to IT user | WIP20 |  | NA | Subsystem |  | NA | No Plan to Fix | sagar_v | 0 | 4 |  | Console4 | NO | Closed | SubNC | IQon1
CTUDT00626803 | Refresh | Closed | finestine_m | October 22, 2019 3:57:25 PM GMT+08:00 | Refresh: Reason Matrix: iCT- New System Integration Build Reason Not According To The Spec | WIP20 | wip22 |  | Subsystem | reason matrix | n/a | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP22SubCom | Ovation |  | wip22 | October 29, 2019 5:56:08 PM GMT+08:00
CTUDT00626800 | Refresh | Closed | cherylli | October 22, 2019 3:51:42 PM GMT+08:00 | Refresh Pre-Val: Recon server stopped working while performing coronary CTA | BB_WIP10 | BB_WIP12 | Try to offline recon each result separately, not two in parallel. Do the recon after logoff/logon. | Subsystem |  | NA | Fixed | datta_i | 0 | 2 | Console Software | Console4 | YES | Closed | BB_WIP12_SegReq | Brilliance Big Bore |  | BB_WIP12 | November 15, 2019 8:49:36 PM GMT+08:00
CTUDT00626795 | Refresh | Closed | bandil_j | October 22, 2019 2:19:40 PM GMT+08:00 | Installation spec need tp be updated for Ingenuity | 4.1.10 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | bandil_j | 0 |  |  | Support-System | NO | Closed |  | Ingenuity 64
CTUDT00626791 | Refresh | Closed | zyanjiang | October 22, 2019 11:14:24 AM GMT+08:00 | Refresh Pre-Val:Time bar value overlap display | BB_WIP10 |  |  | Subsystem | GUI | NA | No Plan to Fix | Prakash_Chand_Yadav | 0 | 4 |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00626786 | Refresh | Closed | richards_b | October 21, 2019 10:29:03 PM GMT+08:00 | NTP Timesync improvement for CIRS | WIP20 | WIP21.2 | N/A | Subsystem | Verified1 | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_Installation | CIRS | NO | Closed | WIP22SegReq | Ovation |  | WIP21.2 | November 9, 2019 5:17:48 AM GMT+08:00
CTUDT00626786 | Refresh | Closed | richards_b | October 21, 2019 10:29:03 PM GMT+08:00 | NTP Timesync improvement for CIRS | WIP20 | WIP21.2 | N/A | Subsystem | Verified1 | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_Installation | CIRS | NO | Closed | WIP22SegReq | Brilliance 64 |  | WIP21.2 | November 9, 2019 5:17:48 AM GMT+08:00
CTUDT00626786 | Refresh | Closed | richards_b | October 21, 2019 10:29:03 PM GMT+08:00 | NTP Timesync improvement for CIRS | WIP20 | WIP21.2 | N/A | Subsystem | Verified1 | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_Installation | CIRS | NO | Closed | WIP22SegReq | Brilliance Big Bore |  | WIP21.2 | November 9, 2019 5:17:48 AM GMT+08:00
CTUDT00626786 | Refresh | Closed | richards_b | October 21, 2019 10:29:03 PM GMT+08:00 | NTP Timesync improvement for CIRS | WIP20 | WIP21.2 | N/A | Subsystem | Verified1 | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_Installation | CIRS | NO | Closed | WIP22SegReq | Ingenuity 64 |  | WIP21.2 | November 9, 2019 5:17:48 AM GMT+08:00
CTUDT00626786 | Refresh | Closed | richards_b | October 21, 2019 10:29:03 PM GMT+08:00 | NTP Timesync improvement for CIRS | WIP20 | WIP21.2 | N/A | Subsystem | Verified1 | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_Installation | CIRS | NO | Closed | WIP22SegReq | IQon1 |  | WIP21.2 | November 9, 2019 5:17:48 AM GMT+08:00
CTUDT00626784 | Refresh | Closed | roy_r | October 21, 2019 6:32:18 PM GMT+08:00 | SRS needs update to include requirement for DIN6862-2 compliance | DHF298555 Rev07 DHF298556 Rev08,DHF298553 Rev07,DHF298554 Rev07 | DHF298554 Rev08 DHF298556 Rev 09 DHF298553 Rev 08 DHF298555 Rev08 |  | Subsystem |  | N/A | Fixed | bandil_j | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP22SegReq | Ovation |  | DHF298554 Rev08 DHF298556 Rev 09 DHF298553 Rev 08 DHF298555 Rev08 | November 5, 2019 7:05:57 PM GMT+08:00
CTUDT00626784 | Refresh | Closed | roy_r | October 21, 2019 6:32:18 PM GMT+08:00 | SRS needs update to include requirement for DIN6862-2 compliance | DHF298555 Rev07 DHF298556 Rev08,DHF298553 Rev07,DHF298554 Rev07 | DHF298554 Rev08 DHF298556 Rev 09 DHF298553 Rev 08 DHF298555 Rev08 |  | Subsystem |  | N/A | Fixed | bandil_j | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP22SegReq | Brilliance 64 |  | DHF298554 Rev08 DHF298556 Rev 09 DHF298553 Rev 08 DHF298555 Rev08 | November 5, 2019 7:05:57 PM GMT+08:00
CTUDT00626784 | Refresh | Closed | roy_r | October 21, 2019 6:32:18 PM GMT+08:00 | SRS needs update to include requirement for DIN6862-2 compliance | DHF298555 Rev07 DHF298556 Rev08,DHF298553 Rev07,DHF298554 Rev07 | DHF298554 Rev08 DHF298556 Rev 09 DHF298553 Rev 08 DHF298555 Rev08 |  | Subsystem |  | N/A | Fixed | bandil_j | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP22SegReq | Brilliance Big Bore |  | DHF298554 Rev08 DHF298556 Rev 09 DHF298553 Rev 08 DHF298555 Rev08 | November 5, 2019 7:05:57 PM GMT+08:00
CTUDT00626784 | Refresh | Closed | roy_r | October 21, 2019 6:32:18 PM GMT+08:00 | SRS needs update to include requirement for DIN6862-2 compliance | DHF298555 Rev07 DHF298556 Rev08,DHF298553 Rev07,DHF298554 Rev07 | DHF298554 Rev08 DHF298556 Rev 09 DHF298553 Rev 08 DHF298555 Rev08 |  | Subsystem |  | N/A | Fixed | bandil_j | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP22SegReq | Ingenuity 64 |  | DHF298554 Rev08 DHF298556 Rev 09 DHF298553 Rev 08 DHF298555 Rev08 | November 5, 2019 7:05:57 PM GMT+08:00
CTUDT00626784 | Refresh | Closed | roy_r | October 21, 2019 6:32:18 PM GMT+08:00 | SRS needs update to include requirement for DIN6862-2 compliance | DHF298555 Rev07 DHF298556 Rev08,DHF298553 Rev07,DHF298554 Rev07 | DHF298554 Rev08 DHF298556 Rev 09 DHF298553 Rev 08 DHF298555 Rev08 |  | Subsystem |  | N/A | Fixed | bandil_j | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP22SegReq | IQon1 |  | DHF298554 Rev08 DHF298556 Rev 09 DHF298553 Rev 08 DHF298555 Rev08 | November 5, 2019 7:05:57 PM GMT+08:00
CTUDT00626781 | Refresh | Closed | zhu_je | October 21, 2019 6:02:56 PM GMT+08:00 | [Refresh] There is error “There seems to be a problem in the application Would you like to create a Bug Report” if pause second surview during scanning | WIP19 | WIP21.1 WIP11.1 | Start new exam, select previous patient and use previous surview | Subsystem |  | N/A | Fixed | pandey_a | 0 | 4 | Console Software | Console4 | YES | Closed | WIP21SubCom | Ovation |  | WIP21.1 WIP11.1 | October 23, 2019 6:19:28 PM GMT+08:00
CTUDT00626781 | Refresh | Closed | zhu_je | October 21, 2019 6:02:56 PM GMT+08:00 | [Refresh] There is error “There seems to be a problem in the application Would you like to create a Bug Report” if pause second surview during scanning | WIP19 | WIP21.1 WIP11.1 | Start new exam, select previous patient and use previous surview | Subsystem |  | N/A | Fixed | pandey_a | 0 | 4 | Console Software | Console4 | YES | Closed | WIP21SubCom | Brilliance 64 |  | WIP21.1 WIP11.1 | October 23, 2019 6:19:28 PM GMT+08:00
CTUDT00626781 | Refresh | Closed | zhu_je | October 21, 2019 6:02:56 PM GMT+08:00 | [Refresh] There is error “There seems to be a problem in the application Would you like to create a Bug Report” if pause second surview during scanning | WIP19 | WIP21.1 WIP11.1 | Start new exam, select previous patient and use previous surview | Subsystem |  | N/A | Fixed | pandey_a | 0 | 4 | Console Software | Console4 | YES | Closed | WIP21SubCom | Brilliance Big Bore |  | WIP21.1 WIP11.1 | October 23, 2019 6:19:28 PM GMT+08:00
CTUDT00626781 | Refresh | Closed | zhu_je | October 21, 2019 6:02:56 PM GMT+08:00 | [Refresh] There is error “There seems to be a problem in the application Would you like to create a Bug Report” if pause second surview during scanning | WIP19 | WIP21.1 WIP11.1 | Start new exam, select previous patient and use previous surview | Subsystem |  | N/A | Fixed | pandey_a | 0 | 4 | Console Software | Console4 | YES | Closed | WIP21SubCom | Ingenuity 64 |  | WIP21.1 WIP11.1 | October 23, 2019 6:19:28 PM GMT+08:00
CTUDT00626781 | Refresh | Closed | zhu_je | October 21, 2019 6:02:56 PM GMT+08:00 | [Refresh] There is error “There seems to be a problem in the application Would you like to create a Bug Report” if pause second surview during scanning | WIP19 | WIP21.1 WIP11.1 | Start new exam, select previous patient and use previous surview | Subsystem |  | N/A | Fixed | pandey_a | 0 | 4 | Console Software | Console4 | YES | Closed | WIP21SubCom | IQon1 |  | WIP21.1 WIP11.1 | October 23, 2019 6:19:28 PM GMT+08:00
CTUDT00626778 | Refresh | Closed | chakraborty_d | October 21, 2019 2:08:45 PM GMT+08:00 | Messages displayed by the Disk Encryption tool to the clinical user are appearing only in English, even when the regional locale of the system is changed. | WIP20 | BB_WIP11.1 WIP21.1 |  | Subsystem |  | NA | Fixed | chakraborty_d | 0 | 2 | Service Tools_Software | Console4 | NO | Closed | WIP22SubCom | Ovation |  | BB_WIP11.1 WIP21.1 | October 23, 2019 7:34:33 PM GMT+08:00
CTUDT00626778 | Refresh | Closed | chakraborty_d | October 21, 2019 2:08:45 PM GMT+08:00 | Messages displayed by the Disk Encryption tool to the clinical user are appearing only in English, even when the regional locale of the system is changed. | WIP20 | BB_WIP11.1 WIP21.1 |  | Subsystem |  | NA | Fixed | chakraborty_d | 0 | 2 | Service Tools_Software | Console4 | NO | Closed | WIP22SubCom | Brilliance 64 |  | BB_WIP11.1 WIP21.1 | October 23, 2019 7:34:33 PM GMT+08:00
CTUDT00626778 | Refresh | Closed | chakraborty_d | October 21, 2019 2:08:45 PM GMT+08:00 | Messages displayed by the Disk Encryption tool to the clinical user are appearing only in English, even when the regional locale of the system is changed. | WIP20 | BB_WIP11.1 WIP21.1 |  | Subsystem |  | NA | Fixed | chakraborty_d | 0 | 2 | Service Tools_Software | Console4 | NO | Closed | WIP22SubCom | Brilliance Big Bore |  | BB_WIP11.1 WIP21.1 | October 23, 2019 7:34:33 PM GMT+08:00
CTUDT00626778 | Refresh | Closed | chakraborty_d | October 21, 2019 2:08:45 PM GMT+08:00 | Messages displayed by the Disk Encryption tool to the clinical user are appearing only in English, even when the regional locale of the system is changed. | WIP20 | BB_WIP11.1 WIP21.1 |  | Subsystem |  | NA | Fixed | chakraborty_d | 0 | 2 | Service Tools_Software | Console4 | NO | Closed | WIP22SubCom | Ingenuity 64 |  | BB_WIP11.1 WIP21.1 | October 23, 2019 7:34:33 PM GMT+08:00
CTUDT00626778 | Refresh | Closed | chakraborty_d | October 21, 2019 2:08:45 PM GMT+08:00 | Messages displayed by the Disk Encryption tool to the clinical user are appearing only in English, even when the regional locale of the system is changed. | WIP20 | BB_WIP11.1 WIP21.1 |  | Subsystem |  | NA | Fixed | chakraborty_d | 0 | 2 | Service Tools_Software | Console4 | NO | Closed | WIP22SubCom | IQon1 |  | BB_WIP11.1 WIP21.1 | October 23, 2019 7:34:33 PM GMT+08:00
CTUDT00626777 | Refresh | Closed | saranna_cao | October 21, 2019 2:07:50 PM GMT+08:00 | Refresh Pre-Val: Gantry tilt of Oncology hasn’t been disabled when doing oncology brain axial scan. | BB_WIP9 | BB_WIP11.1 |  | Subsystem |  | NA | Fixed | bhawna_g | 0 | 3 | Console Software | Console4 | NO | Closed | WIP21.1 | Brilliance Big Bore |  | BB_WIP11.1 | October 25, 2019 5:45:16 PM GMT+08:00
CTUDT00626774 | Refresh | Closed | shiraguppi_k | October 21, 2019 12:40:00 PM GMT+08:00 | Pre-Val: After scanning Cardiac perfusion scan with Phase 45% after recon, Click on Edit Results in Phase drop down all phases showing and available. | WIP19 |  |  | Subsystem | CRN | NA | No Plan to Fix | bhawna_g | 0 | 3 |  | Console4 | NO | Closed | SubNC | Ovation
CTUDT00626774 | Refresh | Closed | shiraguppi_k | October 21, 2019 12:40:00 PM GMT+08:00 | Pre-Val: After scanning Cardiac perfusion scan with Phase 45% after recon, Click on Edit Results in Phase drop down all phases showing and available. | WIP19 |  |  | Subsystem | CRN | NA | No Plan to Fix | bhawna_g | 0 | 3 |  | Console4 | NO | Closed | SubNC | Brilliance 64
CTUDT00626774 | Refresh | Closed | shiraguppi_k | October 21, 2019 12:40:00 PM GMT+08:00 | Pre-Val: After scanning Cardiac perfusion scan with Phase 45% after recon, Click on Edit Results in Phase drop down all phases showing and available. | WIP19 |  |  | Subsystem | CRN | NA | No Plan to Fix | bhawna_g | 0 | 3 |  | Console4 | NO | Closed | SubNC | Ingenuity 64
CTUDT00626774 | Refresh | Closed | shiraguppi_k | October 21, 2019 12:40:00 PM GMT+08:00 | Pre-Val: After scanning Cardiac perfusion scan with Phase 45% after recon, Click on Edit Results in Phase drop down all phases showing and available. | WIP19 |  |  | Subsystem | CRN | NA | No Plan to Fix | bhawna_g | 0 | 3 |  | Console4 | NO | Closed | SubNC | IQon1
CTUDT00626772 | Refresh | Closed | l_yiming | October 21, 2019 11:29:09 AM GMT+08:00 | Login ct User Failed | WIP20 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | l_yiming | 0 |  |  | Support-System | NO | Closed |  | Ingenuity 64
CTUDT00626770 | Refresh | Closed | l_yiming | October 21, 2019 11:13:40 AM GMT+08:00 | ECM stopped working. | WIP20 | WIP21.1 BB_WIP11.1 |  | Subsystem |  | NA | Fixed | bhawna_g | 0 | 3 | Console Software | Console4 | NO | Closed | WIP22SubCom | Ovation |  | WIP21.1 BB_WIP11.1 | October 25, 2019 4:39:02 PM GMT+08:00
CTUDT00626770 | Refresh | Closed | l_yiming | October 21, 2019 11:13:40 AM GMT+08:00 | ECM stopped working. | WIP20 | WIP21.1 BB_WIP11.1 |  | Subsystem |  | NA | Fixed | bhawna_g | 0 | 3 | Console Software | Console4 | NO | Closed | WIP22SubCom | Brilliance 64 |  | WIP21.1 BB_WIP11.1 | October 25, 2019 4:39:02 PM GMT+08:00
CTUDT00626770 | Refresh | Closed | l_yiming | October 21, 2019 11:13:40 AM GMT+08:00 | ECM stopped working. | WIP20 | WIP21.1 BB_WIP11.1 |  | Subsystem |  | NA | Fixed | bhawna_g | 0 | 3 | Console Software | Console4 | NO | Closed | WIP22SubCom | Brilliance Big Bore |  | WIP21.1 BB_WIP11.1 | October 25, 2019 4:39:02 PM GMT+08:00
CTUDT00626770 | Refresh | Closed | l_yiming | October 21, 2019 11:13:40 AM GMT+08:00 | ECM stopped working. | WIP20 | WIP21.1 BB_WIP11.1 |  | Subsystem |  | NA | Fixed | bhawna_g | 0 | 3 | Console Software | Console4 | NO | Closed | WIP22SubCom | Ingenuity 64 |  | WIP21.1 BB_WIP11.1 | October 25, 2019 4:39:02 PM GMT+08:00
CTUDT00626770 | Refresh | Closed | l_yiming | October 21, 2019 11:13:40 AM GMT+08:00 | ECM stopped working. | WIP20 | WIP21.1 BB_WIP11.1 |  | Subsystem |  | NA | Fixed | bhawna_g | 0 | 3 | Console Software | Console4 | NO | Closed | WIP22SubCom | IQon1 |  | WIP21.1 BB_WIP11.1 | October 25, 2019 4:39:02 PM GMT+08:00
CTUDT00626766 | Refresh | Closed | l_yiming | October 21, 2019 10:46:13 AM GMT+08:00 | Refresh: Loading protocol to scan from injector protocol failed | WIP20 |  | Newly added protocol in injector device will be available in Console from next exam (end current exam to start new exam) | Subsystem |  | None | Works As Specified / Not A Defect | jose_s | 0 |  |  | Console4 | NO | Closed | WAS | Ovation
CTUDT00626766 | Refresh | Closed | l_yiming | October 21, 2019 10:46:13 AM GMT+08:00 | Refresh: Loading protocol to scan from injector protocol failed | WIP20 |  | Newly added protocol in injector device will be available in Console from next exam (end current exam to start new exam) | Subsystem |  | None | Works As Specified / Not A Defect | jose_s | 0 |  |  | Console4 | NO | Closed | WAS | Brilliance 64
CTUDT00626766 | Refresh | Closed | l_yiming | October 21, 2019 10:46:13 AM GMT+08:00 | Refresh: Loading protocol to scan from injector protocol failed | WIP20 |  | Newly added protocol in injector device will be available in Console from next exam (end current exam to start new exam) | Subsystem |  | None | Works As Specified / Not A Defect | jose_s | 0 |  |  | Console4 | NO | Closed | WAS | Brilliance Big Bore
CTUDT00626766 | Refresh | Closed | l_yiming | October 21, 2019 10:46:13 AM GMT+08:00 | Refresh: Loading protocol to scan from injector protocol failed | WIP20 |  | Newly added protocol in injector device will be available in Console from next exam (end current exam to start new exam) | Subsystem |  | None | Works As Specified / Not A Defect | jose_s | 0 |  |  | Console4 | NO | Closed | WAS | Ingenuity 64
CTUDT00626766 | Refresh | Closed | l_yiming | October 21, 2019 10:46:13 AM GMT+08:00 | Refresh: Loading protocol to scan from injector protocol failed | WIP20 |  | Newly added protocol in injector device will be available in Console from next exam (end current exam to start new exam) | Subsystem |  | None | Works As Specified / Not A Defect | jose_s | 0 |  |  | Console4 | NO | Closed | WAS | IQon1
CTUDT00626764 | Refresh | Closed | l_yiming | October 21, 2019 9:55:02 AM GMT+08:00 | Refresh: Injector Initializing Failure | WIP20 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | l_yiming | 0 |  |  |  |  | Closed |  | Brilliance 64
CTUDT00626762 | Refresh | Closed | l_yiming | October 21, 2019 9:36:19 AM GMT+08:00 | Refresh: In Injection Parameters, options are missing. | WIP20 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | pandey_a | 0 |  |  |  |  | Closed |  | Brilliance 64
CTUDT00626760 | Refresh | Closed | li_ni | October 20, 2019 9:50:45 AM GMT+08:00 | After PIM/ECG disconnected, the message shows on console exceed 1s | WIP20 |  |  | Subsystem | Test Design | GCX.SSRS.iCT.12514 | Works As Specified / Not A Defect | li_ni | 0 | 3 | Test Design | GCX | NO | Closed | Test Design | Ovation |  |  | November 1, 2019 5:11:25 PM GMT+08:00
CTUDT00626757 | Refresh | Closed | li_ni | October 20, 2019 9:47:10 AM GMT+08:00 | CRTC Data Logger stops working | WIP20 | WIP22 |  | Subsystem | Logging | N/A | Fixed | zhang_lawrence | 0 | 2 | Gantry Software | GCX | NO | Closed | WIP21SegReq | Ovation |  | WIP22 | October 24, 2019 4:42:14 PM GMT+08:00
CTUDT00626755 | Refresh | Closed | li_ni | October 20, 2019 9:35:14 AM GMT+08:00 | Cardiac Coronary Helical scan can`t produce when using PIMI_GUI pattern "60, 60, 20" | WIP20 |  |  | Subsystem | Cardiac | GCX.SSRS.iCT.16783,GCX.SSRS.iCT.16784 | Works As Specified / Not A Defect | li_ni | 0 | 3 | Test Design | GCX | NO | Closed | Test Design | Ovation |  |  | November 1, 2019 5:11:57 PM GMT+08:00
CTUDT00626751 | Refresh | Closed | marino | October 19, 2019 3:21:21 AM GMT+08:00 | The IMR slow option has not been available for years. | WIP20 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | marino | 0 |  |  | CIRS | NO | Closed |  | Ovation
CTUDT00626751 | Refresh | Closed | marino | October 19, 2019 3:21:21 AM GMT+08:00 | The IMR slow option has not been available for years. | WIP20 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | marino | 0 |  |  | CIRS | NO | Closed |  | Brilliance 64
CTUDT00626751 | Refresh | Closed | marino | October 19, 2019 3:21:21 AM GMT+08:00 | The IMR slow option has not been available for years. | WIP20 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | marino | 0 |  |  | CIRS | NO | Closed |  | Brilliance Big Bore
CTUDT00626751 | Refresh | Closed | marino | October 19, 2019 3:21:21 AM GMT+08:00 | The IMR slow option has not been available for years. | WIP20 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | marino | 0 |  |  | CIRS | NO | Closed |  | Ingenuity 64
CTUDT00626747 | Refresh | Closed | cherylli | October 18, 2019 5:00:10 PM GMT+08:00 | Refresh: Wrong instruction of Bolus Tracking help images | WIP20 |  |  | Subsystem |  | None | No Plan to Fix | jose_s | 0 | 3 | Console Software | Console4 | NO | Closed | SubNC | Ovation
CTUDT00626747 | Refresh | Closed | cherylli | October 18, 2019 5:00:10 PM GMT+08:00 | Refresh: Wrong instruction of Bolus Tracking help images | WIP20 |  |  | Subsystem |  | None | No Plan to Fix | jose_s | 0 | 3 | Console Software | Console4 | NO | Closed | SubNC | Brilliance 64
CTUDT00626747 | Refresh | Closed | cherylli | October 18, 2019 5:00:10 PM GMT+08:00 | Refresh: Wrong instruction of Bolus Tracking help images | WIP20 |  |  | Subsystem |  | None | No Plan to Fix | jose_s | 0 | 3 | Console Software | Console4 | NO | Closed | SubNC | Brilliance Big Bore
CTUDT00626747 | Refresh | Closed | cherylli | October 18, 2019 5:00:10 PM GMT+08:00 | Refresh: Wrong instruction of Bolus Tracking help images | WIP20 |  |  | Subsystem |  | None | No Plan to Fix | jose_s | 0 | 3 | Console Software | Console4 | NO | Closed | SubNC | Ingenuity 64
CTUDT00626747 | Refresh | Closed | cherylli | October 18, 2019 5:00:10 PM GMT+08:00 | Refresh: Wrong instruction of Bolus Tracking help images | WIP20 |  |  | Subsystem |  | None | No Plan to Fix | jose_s | 0 | 3 | Console Software | Console4 | NO | Closed | SubNC | IQon1
CTUDT00626745 | Refresh | Closed | cherylli | October 18, 2019 4:54:24 PM GMT+08:00 | Refresh: The icon of Brain DRI in left Scan Parameters is wrong and unchangeable | WIP20 |  |  | Subsystem |  |  | Duplicate | bhawna_g | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00626743 | Refresh | Closed | cherylli | October 18, 2019 4:49:14 PM GMT+08:00 | Refresh: The icon of Liver DRI in left Scan Parameters is wrong and unchangeable | WIP20 |  |  | Subsystem |  |  | Duplicate | bhawna_g | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00626739 | Refresh | Closed | Bat-El_Asaraf | October 18, 2019 12:17:29 AM GMT+08:00 | Pre-Val Refresh: Absence of phase point (bar) on the shot post disconnection | WIP20 |  |  | Subsystem | CRN | N/A | No Plan to Fix | shaham_t | 0 | 3 | Console Software | Support-System | YES | Closed | WAS/Change Request | Ovation
CTUDT00626739 | Refresh | Closed | Bat-El_Asaraf | October 18, 2019 12:17:29 AM GMT+08:00 | Pre-Val Refresh: Absence of phase point (bar) on the shot post disconnection | WIP20 |  |  | Subsystem | CRN | N/A | No Plan to Fix | shaham_t | 0 | 3 | Console Software | Support-System | YES | Closed | WAS/Change Request | Brilliance 64
CTUDT00626739 | Refresh | Closed | Bat-El_Asaraf | October 18, 2019 12:17:29 AM GMT+08:00 | Pre-Val Refresh: Absence of phase point (bar) on the shot post disconnection | WIP20 |  |  | Subsystem | CRN | N/A | No Plan to Fix | shaham_t | 0 | 3 | Console Software | Support-System | YES | Closed | WAS/Change Request | Ingenuity 64
CTUDT00626739 | Refresh | Closed | Bat-El_Asaraf | October 18, 2019 12:17:29 AM GMT+08:00 | Pre-Val Refresh: Absence of phase point (bar) on the shot post disconnection | WIP20 |  |  | Subsystem | CRN | N/A | No Plan to Fix | shaham_t | 0 | 3 | Console Software | Support-System | YES | Closed | WAS/Change Request | IQon1
CTUDT00626737 | Refresh | Closed | burke_c | October 17, 2019 10:36:02 PM GMT+08:00 | Un-commanded Vertical Height Movement while using the hardware Interventional Control | Refresh wip 109 | BB_WIP11.1 | User can use the CT Box for more precise vertical movement. | Subsystem |  | NA | Fixed | 320021835 | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP12_SubCom | Brilliance Big Bore | found internally no CFR provided | BB_WIP11.1 | October 23, 2019 7:22:25 PM GMT+08:00
CTUDT00626735 | Refresh | Closed | waggoner_j | October 17, 2019 9:50:42 PM GMT+08:00 | Big Bore 4.8. The helical exam cards under the Onco Gating EC Group, all have cardiac image definitions when utilizing IMR | BB_WIP9 | BB_WIP12 WIP22 | User can manually change the suitable image definition before performing the recon. | System |  | NA | Fixed | li_e | 0 | 3 | Console Software | Console4 | NO | Closed | WIP22SubCom | Brilliance 64 |  | BB_WIP12 WIP22 | November 8, 2019 5:05:46 PM GMT+08:00
CTUDT00626735 | Refresh | Closed | waggoner_j | October 17, 2019 9:50:42 PM GMT+08:00 | Big Bore 4.8. The helical exam cards under the Onco Gating EC Group, all have cardiac image definitions when utilizing IMR | BB_WIP9 | BB_WIP12 WIP22 | User can manually change the suitable image definition before performing the recon. | System |  | NA | Fixed | li_e | 0 | 3 | Console Software | Console4 | NO | Closed | WIP22SubCom | Brilliance Big Bore |  | BB_WIP12 WIP22 | November 8, 2019 5:05:46 PM GMT+08:00
CTUDT00626735 | Refresh | Closed | waggoner_j | October 17, 2019 9:50:42 PM GMT+08:00 | Big Bore 4.8. The helical exam cards under the Onco Gating EC Group, all have cardiac image definitions when utilizing IMR | BB_WIP9 | BB_WIP12 WIP22 | User can manually change the suitable image definition before performing the recon. | System |  | NA | Fixed | li_e | 0 | 3 | Console Software | Console4 | NO | Closed | WIP22SubCom | Ingenuity 64 |  | BB_WIP12 WIP22 | November 8, 2019 5:05:46 PM GMT+08:00
CTUDT00626728 | Refresh | Closed | shterner_a | October 17, 2019 7:52:48 PM GMT+08:00 | Refresh PRS: CT-NM.PRS.CRS.5357 - IQon Refresh has no reference protocol for trauma | PRS Qon Refresh DHF298293 Rev 07 | IQon PRS DHF298293 Rev.08 |  | Subsystem | Requirements | CT-NM.PRS.CRS.5357 | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP21SegReq | IQon1 |  | IQon PRS DHF298293 Rev.08 | November 6, 2019 2:20:52 PM GMT+08:00
CTUDT00626723 | Refresh | Closed | wei_f | October 17, 2019 7:11:36 PM GMT+08:00 | Modality value displayed CT in Patient worklist after load the patient details from iWATT HISRIS | WIP19 | WIP20 |  | System |  | CT-NM.SRS.Interop.237 | Fixed | wei_f | 0 | 2 | Test Design | Support-System | NO | Closed | Test Case update | Ingenuity 64 |  | WIP20 | October 29, 2019 4:48:17 PM GMT+08:00
CTUDT00626712 | Refresh | Closed | barnea_d | October 17, 2019 4:39:40 PM GMT+08:00 | Refresh: Special Cases: error message of cracked wedge: missing space bar | BB_WIP10 | WIP21 |  | Subsystem | localization;Verified1 | n/a | Fixed | frider_o | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP21SubCom | Ovation |  | WIP21 | October 17, 2019 9:51:03 PM GMT+08:00
CTUDT00626712 | Refresh | Closed | barnea_d | October 17, 2019 4:39:40 PM GMT+08:00 | Refresh: Special Cases: error message of cracked wedge: missing space bar | BB_WIP10 | WIP21 |  | Subsystem | localization;Verified1 | n/a | Fixed | frider_o | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP21SubCom | Brilliance Big Bore |  | WIP21 | October 17, 2019 9:51:03 PM GMT+08:00
CTUDT00626712 | Refresh | Closed | barnea_d | October 17, 2019 4:39:40 PM GMT+08:00 | Refresh: Special Cases: error message of cracked wedge: missing space bar | BB_WIP10 | WIP21 |  | Subsystem | localization;Verified1 | n/a | Fixed | frider_o | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP21SubCom | Ingenuity 64 |  | WIP21 | October 17, 2019 9:51:03 PM GMT+08:00
CTUDT00626712 | Refresh | Closed | barnea_d | October 17, 2019 4:39:40 PM GMT+08:00 | Refresh: Special Cases: error message of cracked wedge: missing space bar | BB_WIP10 | WIP21 |  | Subsystem | localization;Verified1 | n/a | Fixed | frider_o | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP21SubCom | IQon1 |  | WIP21 | October 17, 2019 9:51:03 PM GMT+08:00
CTUDT00626709 | Refresh | Closed | zhu_je | October 17, 2019 2:16:53 PM GMT+08:00 | [Refresh]Failed assign priority to storage devices although set priority in preference "Save Images" | WIP19 |  |  | Subsystem | Connectivity |  | Works As Specified / Not A Defect | bhawna_g | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00626707 | Refresh | Closed | zyanjiang | October 17, 2019 10:43:24 AM GMT+08:00 | Refresh Pre-Val:Surview should be displayed in realtime | BB_WIP9 |  |  | Subsystem |  |  | Change Proposal | zyanjiang | 0 |  |  |  |  | Closed |  | Brilliance Big Bore
CTUDT00626705 | Refresh | Closed | zyanjiang | October 17, 2019 10:34:06 AM GMT+08:00 | Refresh Pre-Val:Can’t save start value after right click duplicate surview | BB_WIP9 |  |  | Subsystem | Usability | NA | Works As Specified / Not A Defect | challa_s | 0 |  |  | Console4 | NO | Closed | WAS | Brilliance Big Bore
CTUDT00626703 | Refresh | Closed | cherylli | October 17, 2019 9:37:56 AM GMT+08:00 | Refresh Pre-Val: Frozen in offline recon | BB_WIP10 |  | No workaround | Subsystem | Offline recon | NA | No Plan to Fix | ravindra_a | 0 | 4 |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00626701 | Refresh | Closed | zyanjiang | October 17, 2019 8:41:00 AM GMT+08:00 | Refresh Pre-Val:There is no isocenter line on Frontal surview | BB_WIP9 |  | Right Click on Frontal surview and in the Context menu select option "Show Gantry Center Line". This will be persisted in the next exam | Subsystem |  | NA | Works As Specified / Not A Defect | bhawna_g | 0 |  |  | Console4 | NO | Closed | WAS | Brilliance Big Bore
CTUDT00626699 | Refresh | Closed | 320053768 | October 16, 2019 10:42:54 PM GMT+08:00 | Focal spot x position calibration on bay 127 (iQon- gen 5.1) with harmonized phantom orders 2.5 hours to be performed (estimated time as was written in the machine), but took more than 4 hours | WIP20 |  |  | Subsystem |  | N/A | Works As Specified / Not A Defect | tamari_o | 1 |  |  | Support-System | NO | Closed |  | IQon1
CTUDT00626685 | Refresh | Closed | zhong_x | October 16, 2019 7:54:34 PM GMT+08:00 | [Refresh] IVC Refresh: System was hang after loading CCT Continuous Scan | BB_WIP10 |  | User shall logout from Workflow and start Workflow again | Subsystem | IVC | N/A | Works As Specified / Not A Defect | reddy_g | 0 |  |  | Console4 | NO | Closed | Not A Defect | Brilliance Big Bore
CTUDT00626684 | Refresh | Closed | s_namburi | October 16, 2019 7:49:49 PM GMT+08:00 | Refresh : O-Level user cannot connect the HOST to an external NTP server. | WIP18 |  |  | Subsystem |  |  | Duplicate | bandil_j | 0 |  |  |  |  | Closed |  | Ovation
CTUDT00626684 | Refresh | Closed | s_namburi | October 16, 2019 7:49:49 PM GMT+08:00 | Refresh : O-Level user cannot connect the HOST to an external NTP server. | WIP18 |  |  | Subsystem |  |  | Duplicate | bandil_j | 0 |  |  |  |  | Closed |  | Brilliance 64
CTUDT00626684 | Refresh | Closed | s_namburi | October 16, 2019 7:49:49 PM GMT+08:00 | Refresh : O-Level user cannot connect the HOST to an external NTP server. | WIP18 |  |  | Subsystem |  |  | Duplicate | bandil_j | 0 |  |  |  |  | Closed |  | Brilliance Big Bore
CTUDT00626684 | Refresh | Closed | s_namburi | October 16, 2019 7:49:49 PM GMT+08:00 | Refresh : O-Level user cannot connect the HOST to an external NTP server. | WIP18 |  |  | Subsystem |  |  | Duplicate | bandil_j | 0 |  |  |  |  | Closed |  | Ingenuity 64
CTUDT00626684 | Refresh | Closed | s_namburi | October 16, 2019 7:49:49 PM GMT+08:00 | Refresh : O-Level user cannot connect the HOST to an external NTP server. | WIP18 |  |  | Subsystem |  |  | Duplicate | bandil_j | 0 |  |  |  |  | Closed |  | IQon1
CTUDT00626680 | Refresh | Closed | zyanjiang | October 16, 2019 3:56:01 PM GMT+08:00 | Refresh Pre-Val:System error and exit scan interface directly , CT images was not be reconned | BB_WIP9 | WIP21.1 BB_WIP11.1 |  | Subsystem |  | NA | Fixed | prasad_r | 0 | 3 | Console Software | Console4 | YES | Closed | WIP22SubCom | Ovation |  | WIP21.1 BB_WIP11.1 | October 23, 2019 1:04:19 PM GMT+08:00
CTUDT00626680 | Refresh | Closed | zyanjiang | October 16, 2019 3:56:01 PM GMT+08:00 | Refresh Pre-Val:System error and exit scan interface directly , CT images was not be reconned | BB_WIP9 | WIP21.1 BB_WIP11.1 |  | Subsystem |  | NA | Fixed | prasad_r | 0 | 3 | Console Software | Console4 | YES | Closed | WIP22SubCom | Brilliance 64 |  | WIP21.1 BB_WIP11.1 | October 23, 2019 1:04:19 PM GMT+08:00
CTUDT00626680 | Refresh | Closed | zyanjiang | October 16, 2019 3:56:01 PM GMT+08:00 | Refresh Pre-Val:System error and exit scan interface directly , CT images was not be reconned | BB_WIP9 | WIP21.1 BB_WIP11.1 |  | Subsystem |  | NA | Fixed | prasad_r | 0 | 3 | Console Software | Console4 | YES | Closed | WIP22SubCom | Brilliance Big Bore |  | WIP21.1 BB_WIP11.1 | October 23, 2019 1:04:19 PM GMT+08:00
CTUDT00626680 | Refresh | Closed | zyanjiang | October 16, 2019 3:56:01 PM GMT+08:00 | Refresh Pre-Val:System error and exit scan interface directly , CT images was not be reconned | BB_WIP9 | WIP21.1 BB_WIP11.1 |  | Subsystem |  | NA | Fixed | prasad_r | 0 | 3 | Console Software | Console4 | YES | Closed | WIP22SubCom | Ingenuity 64 |  | WIP21.1 BB_WIP11.1 | October 23, 2019 1:04:19 PM GMT+08:00
CTUDT00626680 | Refresh | Closed | zyanjiang | October 16, 2019 3:56:01 PM GMT+08:00 | Refresh Pre-Val:System error and exit scan interface directly , CT images was not be reconned | BB_WIP9 | WIP21.1 BB_WIP11.1 |  | Subsystem |  | NA | Fixed | prasad_r | 0 | 3 | Console Software | Console4 | YES | Closed | WIP22SubCom | IQon1 |  | WIP21.1 BB_WIP11.1 | October 23, 2019 1:04:19 PM GMT+08:00
CTUDT00626674 | Refresh | Closed | azagury_d | October 15, 2019 8:55:23 PM GMT+08:00 | Refresh:Calibrations: Bad Detectors: Bias Out of limits failure code does not match spec. | WIP19 |  | the official report produced on the host (by the calibration does include the reason) | Subsystem | Bad Detectoors | n/a | No Plan to Fix | frider_o | 0 | 3 |  | DMS | NO | Closed | SubNC | IQon1
CTUDT00626672 | Refresh | Closed | vineetha_s | October 15, 2019 7:53:48 PM GMT+08:00 | When we perform the Clinical Scan with the User created Language and Phrase, A Gantry error message is displayed, after upgrade in IQON from 4.7.5 to 4.7.7 | WIP20 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | s_shabeer | 0 |  | Console Software | Service Tools | YES | Closed | WAS | IQon1
CTUDT00626664 | Refresh | Closed | cherylli | October 15, 2019 5:24:27 PM GMT+08:00 | Refresh Pre-Val: Cannot adjust scan range in plan box of continued series | BB_WIP10 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | Prakash_Chand_Yadav | 0 |  | Console Software | Console4 | NO | Closed | WAS/Change request | Brilliance Big Bore
CTUDT00626660 | Refresh | Closed | lin_s | October 15, 2019 3:57:19 PM GMT+08:00 | Brilliance 64 is EOP and upgrade does not sale to China marketing. So Regstand.225 & Regstand.4 are not applicable for Brilliance 64. | SRS Rev07 | ING, BR64 SRS DHF298554 Rev08 |  | Subsystem |  | CT-NM.SRS.RegStand.225 & CT-NM.SRS.RegStand.4 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | SRS | Brilliance 64 |  | ING, BR64 SRS DHF298554 Rev08 | November 4, 2019 10:27:46 AM GMT+08:00
CTUDT00626655 | Refresh | Closed | cherylli | October 15, 2019 9:38:40 AM GMT+08:00 | Refresh: crushed while ending study | BB_WIP10 |  | Since the issue is seen once, the potential workaround is a system restart. | Subsystem |  | NA | Works As Specified / Not A Defect | s_sa | 0 |  | Console Software | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00626652 | Refresh | Closed | cherylli | October 15, 2019 9:11:14 AM GMT+08:00 | Refresh Pre-Val: crushed while clicking “Patient” | BB_WIP10 | BB_WIP11 | Clicking on the other controls like patient directory , patient page is displayed correctly | Subsystem |  | N/A | Fixed | navale_s | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP11_SubCom | Brilliance Big Bore |  | BB_WIP11 | October 16, 2019 9:02:55 PM GMT+08:00
CTUDT00626649 | Refresh | Closed | roy_r | October 14, 2019 9:06:42 PM GMT+08:00 | How Do I videos not available for Interventional Controls | BB_WIP10 | BB_WIP11 | Refer to IFU | Subsystem | IVC; How Do I | Console-SRS.16884 | Fixed | vishnu_devi | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP11_SubCom | Brilliance Big Bore |  | BB_WIP11 | October 16, 2019 9:46:54 PM GMT+08:00
CTUDT00626646 | Refresh | Closed | PremKumar_Rajendran | October 14, 2019 7:58:48 PM GMT+08:00 | BB Refresh : The Actual Latency was more than 500 milliseconds for timed scan and tracker scan. | BB_WIP10 | DHF330449 Rev14 |  | Subsystem | Performance | N/A | Fixed | panchami_pradeep | 0 | 4 | Test Design | Console4 | NO | Closed | TD Update | Ovation |  | DHF330449 Rev14 | October 16, 2019 9:24:56 PM GMT+08:00
CTUDT00626646 | Refresh | Closed | PremKumar_Rajendran | October 14, 2019 7:58:48 PM GMT+08:00 | BB Refresh : The Actual Latency was more than 500 milliseconds for timed scan and tracker scan. | BB_WIP10 | DHF330449 Rev14 |  | Subsystem | Performance | N/A | Fixed | panchami_pradeep | 0 | 4 | Test Design | Console4 | NO | Closed | TD Update | Brilliance 64 |  | DHF330449 Rev14 | October 16, 2019 9:24:56 PM GMT+08:00
CTUDT00626646 | Refresh | Closed | PremKumar_Rajendran | October 14, 2019 7:58:48 PM GMT+08:00 | BB Refresh : The Actual Latency was more than 500 milliseconds for timed scan and tracker scan. | BB_WIP10 | DHF330449 Rev14 |  | Subsystem | Performance | N/A | Fixed | panchami_pradeep | 0 | 4 | Test Design | Console4 | NO | Closed | TD Update | Brilliance Big Bore |  | DHF330449 Rev14 | October 16, 2019 9:24:56 PM GMT+08:00
CTUDT00626646 | Refresh | Closed | PremKumar_Rajendran | October 14, 2019 7:58:48 PM GMT+08:00 | BB Refresh : The Actual Latency was more than 500 milliseconds for timed scan and tracker scan. | BB_WIP10 | DHF330449 Rev14 |  | Subsystem | Performance | N/A | Fixed | panchami_pradeep | 0 | 4 | Test Design | Console4 | NO | Closed | TD Update | Ingenuity 64 |  | DHF330449 Rev14 | October 16, 2019 9:24:56 PM GMT+08:00
CTUDT00626646 | Refresh | Closed | PremKumar_Rajendran | October 14, 2019 7:58:48 PM GMT+08:00 | BB Refresh : The Actual Latency was more than 500 milliseconds for timed scan and tracker scan. | BB_WIP10 | DHF330449 Rev14 |  | Subsystem | Performance | N/A | Fixed | panchami_pradeep | 0 | 4 | Test Design | Console4 | NO | Closed | TD Update | IQon1 |  | DHF330449 Rev14 | October 16, 2019 9:24:56 PM GMT+08:00
CTUDT00626644 | Refresh | Closed | manjunath_c | October 14, 2019 7:49:01 PM GMT+08:00 | Deviation in the default parameter values displayed in Exam card manager compared to what is mentioned in the "DHF223377 Rev.22 Bigbore Reference Exam Cards" Specification | WIP20 | WIP12 |  | Subsystem | ECM | N/A | Fixed | dunitza_m | 0 | 2 | Design Documentation | Console4 | NO | Closed | Spec update | Brilliance Big Bore |  | WIP12 | December 2, 2019 8:21:26 PM GMT+08:00
CTUDT00626642 | Refresh | Closed | ding_f | October 14, 2019 2:29:52 PM GMT+08:00 | Air Data Variance calibration table values are incorrect in ReconMdu for scans with all collimations except 64x0.625. | WIP19 | WIP21 |  | Subsystem | IMR, IQ, Blocker21;Verified1 | n/a | Fixed | frider_o | 0 | 2 | DMS Software | DMS | YES | Closed | WIP21SubCom | Ingenuity 64 |  | WIP21 | October 17, 2019 10:02:23 PM GMT+08:00
CTUDT00626636 | Refresh | Closed | ding_f | October 14, 2019 1:12:42 PM GMT+08:00 | Air Data Variance calibration table values are incorrect in ReconMdu for scans with all collimations except 16x1.5. | BB_WIP9 | WIP21 |  | Subsystem | IMR IQ;Blocker21;Verified1 | n/a | Fixed | frider_o | 0 | 2 | DMS Software | DMS | YES | Closed | BB_WIP11_SubCom | Brilliance Big Bore |  | WIP21 | October 17, 2019 10:00:56 PM GMT+08:00
CTUDT00626634 | Refresh | Closed | l_yiming | October 14, 2019 12:02:10 PM GMT+08:00 | Entering in Post Injection Delay value failed | WIP20 |  |  | System |  |  | Not Enough Information | bhawna_g | 0 |  | Console Software | Console4 | NO | Closed |  | Brilliance 64
CTUDT00626634 | Refresh | Closed | l_yiming | October 14, 2019 12:02:10 PM GMT+08:00 | Entering in Post Injection Delay value failed | WIP20 |  |  | System |  |  | Not Enough Information | bhawna_g | 0 |  | Console Software | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00626631 | Refresh | Closed | l_yiming | October 14, 2019 11:49:20 AM GMT+08:00 | Password Protection for Editing Protocol Failed | WIP20 |  |  | System |  | NA | Works As Specified / Not A Defect | Prakash_Chand_Yadav | 0 |  | Console Software | Console4 | NO | Closed | WAS | Brilliance 64
CTUDT00626631 | Refresh | Closed | l_yiming | October 14, 2019 11:49:20 AM GMT+08:00 | Password Protection for Editing Protocol Failed | WIP20 |  |  | System |  | NA | Works As Specified / Not A Defect | Prakash_Chand_Yadav | 0 |  | Console Software | Console4 | NO | Closed | WAS | Ingenuity 64
CTUDT00626629 | Refresh | Closed | l_yiming | October 14, 2019 11:45:33 AM GMT+08:00 | IQ Check-Head Failed | WIP20 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | richards_b | 0 |  |  | Support-Physics | NO | Closed |  | Brilliance 64
CTUDT00626627 | Refresh | Closed | l_yiming | October 14, 2019 11:08:38 AM GMT+08:00 | Refresh Pre-Val: Wrong Render Mode Indications on MPR Images | WIP20 |  | Not Required. | System |  | NA | Works As Specified / Not A Defect | kapadia_v | 0 |  | Console Software | Console4 | NO | Closed | WAS | Brilliance 64
CTUDT00626627 | Refresh | Closed | l_yiming | October 14, 2019 11:08:38 AM GMT+08:00 | Refresh Pre-Val: Wrong Render Mode Indications on MPR Images | WIP20 |  | Not Required. | System |  | NA | Works As Specified / Not A Defect | kapadia_v | 0 |  | Console Software | Console4 | NO | Closed | WAS | Ingenuity 64
CTUDT00626625 | Refresh | Closed | l_yiming | October 14, 2019 10:23:28 AM GMT+08:00 | Refresh Pre-Val: Lack of Reference Dual Energy Exam Card | WIP20 |  | NA | System |  | NA | Works As Specified / Not A Defect | kr_m | 0 |  | Console Software | Console4 | NO | Closed | WAS | Brilliance 64
CTUDT00626625 | Refresh | Closed | l_yiming | October 14, 2019 10:23:28 AM GMT+08:00 | Refresh Pre-Val: Lack of Reference Dual Energy Exam Card | WIP20 |  | NA | System |  | NA | Works As Specified / Not A Defect | kr_m | 0 |  | Console Software | Console4 | NO | Closed | WAS | Ingenuity 64
CTUDT00626623 | Refresh | Closed | l_yiming | October 14, 2019 10:03:23 AM GMT+08:00 | Refresh Pre-Val: Long Initial Time | BB_WIP10 |  | NA | System |  | NA | Change Proposal | tan_j | 0 |  | Console Software | GCX | NO | Closed |  | Brilliance Big Bore
CTUDT00626620 | Refresh | Closed | ravindra_a | October 12, 2019 7:00:17 PM GMT+08:00 | Few parameters are missing and there is mismatch in “Show all” and “Print and text tab” for Surview and results | 4.8 | BB_WIP11 | NA | Subsystem |  | NA | Fixed | ravindra_a | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP11_SubCom | Brilliance Big Bore |  | BB_WIP11 | October 17, 2019 1:50:35 PM GMT+08:00
CTUDT00626618 | Refresh | Closed | cherylli | October 12, 2019 5:54:40 PM GMT+08:00 | Refresh Pre-Val: “Crosscheck Slice Location” needs improvement | BB_WIP10 |  | On Replan user can should update the slice location by changing the start/end on Plan box for Result 3.1. | Subsystem |  | NA | Works As Specified / Not A Defect | bhawna_g | 0 |  | Console Software | Console4 | NO | Closed | WAS | Brilliance Big Bore
CTUDT00626617 | Refresh | Closed | kapadia_v | October 12, 2019 5:44:26 PM GMT+08:00 | Refresh: While running performance scenarios “ExamManager has stopped working” error prevents the user from executing the scan. | WIP10 | BB_WIP11 | N/A | System |  | NA | Fixed | li_e | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP11_SubCom | Brilliance Big Bore |  | BB_WIP11 | October 14, 2019 1:13:53 PM GMT+08:00
CTUDT00626615 | Refresh | Closed | cherylli | October 12, 2019 5:26:44 PM GMT+08:00 | Refresh Pre-Val: Wrong time stamp in Image Directory of “Resp Single Phase” scans | BB_WIP10 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | Prakash_Chand_Yadav | 0 |  | Console Software | Console4 | NO | Closed | WAS/Change request | Brilliance Big Bore
CTUDT00626613 | Refresh | Closed | cherylli | October 12, 2019 5:07:33 PM GMT+08:00 | Refresh Pre-Val: Incorrect message: “irradiation but no reconstructed images” | BB_WIP10 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | Prakash_Chand_Yadav | 0 |  | Console Software | Console4 | NO | Closed | WAS/Change request | Brilliance Big Bore
CTUDT00626610 | Refresh | Closed | manjunath_mk | October 12, 2019 2:28:49 PM GMT+08:00 | Localization issues in IVC messages and some messages are truncated in different languages | BB_WIP10 | BB_WIP11 |  | Subsystem |  | NA | Fixed | Snehashis_Dey | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP11_SubCom | Brilliance Big Bore |  | BB_WIP11 | October 15, 2019 8:47:08 PM GMT+08:00
CTUDT00626608 | Refresh | Closed | manjunath_mk | October 12, 2019 1:57:51 PM GMT+08:00 | Interface box Firmware update failed on performing update all. | BB_WIP10 | BB_WIP11 | Scenario 1: Restart the unit and try to install the failed firmware individually. Scenario 2: Nothing. | Subsystem |  | NA | Fixed | 320021835 | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP11_SubCom | Brilliance Big Bore |  | BB_WIP11 | October 15, 2019 7:11:37 PM GMT+08:00
CTUDT00626606 | Refresh | Closed | navale_s | October 12, 2019 1:19:06 PM GMT+08:00 | Refresh:HPZ4: No sound heard at the CT Box speaker from voice manager. | 4.8 | BB_WIP11 | N/A | Subsystem |  | N/A | Fixed | navale_s | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP11_SubCom | Brilliance Big Bore |  | BB_WIP11 | October 17, 2019 8:01:01 PM GMT+08:00
CTUDT00626602 | Refresh | Closed | kapadia_v | October 11, 2019 8:24:31 PM GMT+08:00 | On selecting the result edited in Slab viewer, ECG series associated with the result is not getting selected. | BB_WIP9 | BB_WIP11 | NA | Subsystem |  | NA | Fixed | sagar_v | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP11_SubCom | Brilliance Big Bore |  | BB_WIP11 | October 12, 2019 2:26:15 PM GMT+08:00
CTUDT00626600 | Refresh | Closed | kapadia_v | October 11, 2019 7:29:04 PM GMT+08:00 | Pre-Val: A Error pop up displayed during online reconstruction "Not all Images were created Please use the recon option to create the missing images". | BB_WIP8 | BB_WIP11 |  | Subsystem |  | NA | Fixed | sagar_v | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP11_SubCom | Brilliance Big Bore |  | BB_WIP11 | October 12, 2019 3:46:22 PM GMT+08:00
CTUDT00626599 | Refresh | Closed | manjunath_c | October 11, 2019 7:26:14 PM GMT+08:00 | Deviation in the parameter values displayed in Exam Card Manager compared to what is mentioned in the Bigbore Scan Type Specification - DHF209557 Rev.12 | WIP10 |  |  | Subsystem |  |  | Duplicate | dunitza_m | 0 |  |  | Support-System | NO | Closed |  | Brilliance Big Bore
CTUDT00626596 | Refresh | Closed | cherylli | October 11, 2019 6:28:01 PM GMT+08:00 | Refresh Pre-Val: No Continue option to move on the CCT scan | BB_WIP10 |  |  | Subsystem |  | N/A | Works As Specified / Not A Defect | roy_r | 0 |  | Console Software | Console4 | NO | Closed | WAS/Change Request | Ovation
CTUDT00626596 | Refresh | Closed | cherylli | October 11, 2019 6:28:01 PM GMT+08:00 | Refresh Pre-Val: No Continue option to move on the CCT scan | BB_WIP10 |  |  | Subsystem |  | N/A | Works As Specified / Not A Defect | roy_r | 0 |  | Console Software | Console4 | NO | Closed | WAS/Change Request | Brilliance 64
CTUDT00626596 | Refresh | Closed | cherylli | October 11, 2019 6:28:01 PM GMT+08:00 | Refresh Pre-Val: No Continue option to move on the CCT scan | BB_WIP10 |  |  | Subsystem |  | N/A | Works As Specified / Not A Defect | roy_r | 0 |  | Console Software | Console4 | NO | Closed | WAS/Change Request | Brilliance Big Bore
CTUDT00626596 | Refresh | Closed | cherylli | October 11, 2019 6:28:01 PM GMT+08:00 | Refresh Pre-Val: No Continue option to move on the CCT scan | BB_WIP10 |  |  | Subsystem |  | N/A | Works As Specified / Not A Defect | roy_r | 0 |  | Console Software | Console4 | NO | Closed | WAS/Change Request | Ingenuity 64
CTUDT00626596 | Refresh | Closed | cherylli | October 11, 2019 6:28:01 PM GMT+08:00 | Refresh Pre-Val: No Continue option to move on the CCT scan | BB_WIP10 |  |  | Subsystem |  | N/A | Works As Specified / Not A Defect | roy_r | 0 |  | Console Software | Console4 | NO | Closed | WAS/Change Request | IQon1
CTUDT00626594 | Refresh | Closed | cherylli | October 11, 2019 4:16:24 PM GMT+08:00 | Refresh Pre-Val: The numbers of scan series are NOT successive to previous | BB_WIP10 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | datta_i | 0 |  | Console Software | Console4 | NO | Closed | WAS | Brilliance Big Bore
CTUDT00626593 | Refresh | Closed | bhawna_g | October 11, 2019 3:10:50 PM GMT+08:00 | Refresh: “Wait for the wave…” will disappear though wave is not arrived. | 4.8 | BB_WIP11 | Do not press "Start Final Recon" till scan ends and/or the wave appears. | Subsystem |  | NA | Fixed | bhawna_g | 0 | 3 | Console Software | Console4 | YES | Closed | BB_WIP11_SubCom | Brilliance Big Bore |  | BB_WIP11 | October 12, 2019 7:08:24 PM GMT+08:00
CTUDT00626588 | Refresh | Closed | cherylli | October 11, 2019 1:20:53 PM GMT+08:00 | Refresh Pre-Val: Cardiac DoseRight was available in Non-Gated helical cardiac scan | BB_WIP10 | BB_WIP11 |  | Subsystem |  | Console-SRS.15128 | Fixed | wilsonraja | 0 | 2 | Console Software | Console4 | NO | Closed | WIP21SubCom | Brilliance Big Bore |  | BB_WIP11 | October 16, 2019 4:49:48 PM GMT+08:00
CTUDT00626586 | Refresh | Closed | kapadia_v | October 10, 2019 7:35:40 PM GMT+08:00 | Incorrect message was appeared if an imported language has the same name as an existing language for French and Spanish. | BB_WIP10 | BB_WIP11 | NA | Subsystem |  | NA | Fixed | kapadia_v | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP11_SubCom | Brilliance Big Bore |  | BB_WIP11 | October 12, 2019 8:18:14 PM GMT+08:00
CTUDT00626584 | Refresh | Closed | 320053767 | October 10, 2019 2:57:40 PM GMT+08:00 | wrong xml file parameter (performance_scenario_ovation128.xml, performance_scenario_ovation64.xml) | WIP20 | WIP21 |  | Subsystem | Performance Test | n/a | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP21SubCom | Ovation |  | WIP21 | October 17, 2019 9:55:47 PM GMT+08:00
CTUDT00626582 | Refresh | Closed | manjunath_c | October 10, 2019 2:17:59 PM GMT+08:00 | Deviation in the parameter values displayed in Exam Card Manager compared to what is mentioned in the iQon QA Protocol Scan Type Specification - DHF179204 Rev06 | WIP19 | WIP21 | User can manually change the values. | Subsystem |  | NA | Fixed | braunstein_d | 0 | 2 | Console Software | Support-Physics | NO | Closed | WIP21SubCom | IQon1 |  | WIP21 | October 27, 2019 8:18:26 PM GMT+08:00
CTUDT00626579 | Refresh | Closed | zhang_t | October 10, 2019 1:45:56 PM GMT+08:00 | WIP10 - System Error appeared when planning CCT scan while operating IVC HW unit | BB_WIP10 |  | Since the issue is seen once, the potential workaround is a system restart. | Subsystem | IVC | NA | No Plan to Fix | shaham_t | 0 | 3 | Console Software | iCAP | YES | Closed |  | Brilliance Big Bore
CTUDT00626574 | Refresh | Closed | Kevin_ZHAO | October 10, 2019 11:35:05 AM GMT+08:00 | Refresh Pre val:Ingenuity and Br64 miss requirement for RGSC | WIP19 | ING, BR64 CRS DHF316435 Rev03 |  | Subsystem |  | CT-NM.PRS.CRS.4878 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP21SegReq | Brilliance 64 |  | ING, BR64 CRS DHF316435 Rev03 | November 4, 2019 10:26:15 AM GMT+08:00
CTUDT00626574 | Refresh | Closed | Kevin_ZHAO | October 10, 2019 11:35:05 AM GMT+08:00 | Refresh Pre val:Ingenuity and Br64 miss requirement for RGSC | WIP19 | ING, BR64 CRS DHF316435 Rev03 |  | Subsystem |  | CT-NM.PRS.CRS.4878 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP21SegReq | Ingenuity 64 |  | ING, BR64 CRS DHF316435 Rev03 | November 4, 2019 10:26:15 AM GMT+08:00
CTUDT00626571 | Refresh | Closed | zhang_t | October 10, 2019 9:33:34 AM GMT+08:00 | S_COUCHMGR_MOTION_START_ERROR, patient table cannot comply, please retry message popup when using IVC HW | BB_WIP10 |  | Since issue is sporadic user can retry the couch movement after clicking GO again | Subsystem | IV Control | console-SRS.16541 | Duplicate | cx_zhang | 0 |  |  | GCX | NO | Closed |  | Brilliance Big Bore
CTUDT00626568 | Refresh | Closed | manjunath_c | October 9, 2019 8:49:05 PM GMT+08:00 | Deviation in the parameter values displayed in Exam Card Manager compared to what is mentioned in the iQon Scan Type Specification - DHF282681 Rev03 | WIP19 | DHF318051 Rev23 Host verification Test Protocol | Change manually the default values to the desirable ones | Subsystem |  | N/A | Fixed | nanjunda_s | 0 | 3 | Test Design | Console4 | NO | Closed | TD Update | IQon1 |  | DHF318051 Rev23 Host verification Test Protocol | November 5, 2019 9:18:35 PM GMT+08:00
CTUDT00626565 | Refresh | Closed | manjunath_c | October 9, 2019 8:43:47 PM GMT+08:00 | Deviation in the parameter values displayed in Exam Card Manager compared to what is mentioned in the iCTSP Scan Type Specification - DHF285448 Rev05 | WIP20 | DHF318051 Rev23 Host verification protocol DHF285448 Rev.06 iCT-SP Clinical Scan Types Specification | N/A | Subsystem |  | N/A | Fixed | dunitza_m | 0 | 3 | Design Documentation | Support-System | NO | Closed | TD Update | Ovation |  | DHF318051 Rev23 Host verification protocol DHF285448 Rev.06 iCT-SP Clinical Scan Types Specification | November 28, 2019 4:20:48 PM GMT+08:00
CTUDT00626563 | Refresh | Closed | Sharmila_Vasireddy | October 9, 2019 8:40:36 PM GMT+08:00 | BB refresh: In Console the message (#509) was displaying improperly for pulmo exam cards | BB_WIP10 | DHF303221 Rev17 DHF316705 Rev17 | N/A | Subsystem |  | Console-SRS.17715 | Fixed | roy_r | 0 | 4 | Requirements Documentation | Console4 | NO | Closed | WIP21SegReq | Brilliance Big Bore |  | DHF303221 Rev17 DHF316705 Rev17 | October 16, 2019 3:01:44 PM GMT+08:00
CTUDT00626562 | Refresh | Closed | manjunath_c | October 9, 2019 8:33:56 PM GMT+08:00 | Deviation in the parameter values displayed in Exam Card Manager compared to what is mentioned in the iCT Scan Type Specification - DHF285446 Rev05 | WIP19 | WIP22 | No workaround as it is a documentation issue | Subsystem |  | N/A | Fixed | dunitza_m | 0 | 2 | Design Documentation | Support-System | NO | Closed | Spec update | Ovation |  | WIP22 | November 28, 2019 4:16:21 PM GMT+08:00
CTUDT00626558 | Refresh | Closed | g_di | October 9, 2019 4:01:38 PM GMT+08:00 | Few parameters are missing and there is mismatch in “Show all” and “Print and text tab” for Surview and results | WIP20 | WIP21 | NA | Subsystem |  | NA | Fixed | ravindra_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP21SubCom | Ovation |  | WIP21 | October 17, 2019 1:41:18 PM GMT+08:00
CTUDT00626558 | Refresh | Closed | g_di | October 9, 2019 4:01:38 PM GMT+08:00 | Few parameters are missing and there is mismatch in “Show all” and “Print and text tab” for Surview and results | WIP20 | WIP21 | NA | Subsystem |  | NA | Fixed | ravindra_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP21SubCom | Brilliance 64 |  | WIP21 | October 17, 2019 1:41:18 PM GMT+08:00
CTUDT00626558 | Refresh | Closed | g_di | October 9, 2019 4:01:38 PM GMT+08:00 | Few parameters are missing and there is mismatch in “Show all” and “Print and text tab” for Surview and results | WIP20 | WIP21 | NA | Subsystem |  | NA | Fixed | ravindra_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP21SubCom | Ingenuity 64 |  | WIP21 | October 17, 2019 1:41:18 PM GMT+08:00
CTUDT00626558 | Refresh | Closed | g_di | October 9, 2019 4:01:38 PM GMT+08:00 | Few parameters are missing and there is mismatch in “Show all” and “Print and text tab” for Surview and results | WIP20 | WIP21 | NA | Subsystem |  | NA | Fixed | ravindra_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP21SubCom | IQon1 |  | WIP21 | October 17, 2019 1:41:18 PM GMT+08:00
CTUDT00626556 | Refresh | Closed | wang_l | October 9, 2019 1:19:09 PM GMT+08:00 | BB: System requirements for Big bore include UHR Assembly but BB does not has UHR | BB_WIP9 | BB SRS DHF298555 Rev08 |  | Subsystem |  | CT-NM.SRS.Service.2424 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | SRS | Brilliance Big Bore |  | BB SRS DHF298555 Rev08 | November 4, 2019 10:22:36 AM GMT+08:00
CTUDT00626552 | Refresh | Closed | levy_s | October 7, 2019 9:05:38 PM GMT+08:00 | Refresh: During selection of results by double click in the Series list, automatic scroll occurs, causing selection of un-intendent result | WIP20 | WIP21 BB_WIP11 | User can select intended result by single mouse click insted mouse double click. | Subsystem | Performance | NA | No Plan to Fix | kapadia_v | 0 | 3 | Console Software | Console4 | NO | Closed | SubNC | Ovation |  | WIP21 BB_WIP11 | October 14, 2019 9:14:50 PM GMT+08:00
CTUDT00626552 | Refresh | Closed | levy_s | October 7, 2019 9:05:38 PM GMT+08:00 | Refresh: During selection of results by double click in the Series list, automatic scroll occurs, causing selection of un-intendent result | WIP20 | WIP21 BB_WIP11 | User can select intended result by single mouse click insted mouse double click. | Subsystem | Performance | NA | No Plan to Fix | kapadia_v | 0 | 3 | Console Software | Console4 | NO | Closed | SubNC | Brilliance Big Bore |  | WIP21 BB_WIP11 | October 14, 2019 9:14:50 PM GMT+08:00
CTUDT00626552 | Refresh | Closed | levy_s | October 7, 2019 9:05:38 PM GMT+08:00 | Refresh: During selection of results by double click in the Series list, automatic scroll occurs, causing selection of un-intendent result | WIP20 | WIP21 BB_WIP11 | User can select intended result by single mouse click insted mouse double click. | Subsystem | Performance | NA | No Plan to Fix | kapadia_v | 0 | 3 | Console Software | Console4 | NO | Closed | SubNC | Ingenuity 64 |  | WIP21 BB_WIP11 | October 14, 2019 9:14:50 PM GMT+08:00
CTUDT00626552 | Refresh | Closed | levy_s | October 7, 2019 9:05:38 PM GMT+08:00 | Refresh: During selection of results by double click in the Series list, automatic scroll occurs, causing selection of un-intendent result | WIP20 | WIP21 BB_WIP11 | User can select intended result by single mouse click insted mouse double click. | Subsystem | Performance | NA | No Plan to Fix | kapadia_v | 0 | 3 | Console Software | Console4 | NO | Closed | SubNC | IQon1 |  | WIP21 BB_WIP11 | October 14, 2019 9:14:50 PM GMT+08:00
CTUDT00626550 | Refresh | Closed | Valeria_Kashtanov | October 7, 2019 6:57:32 PM GMT+08:00 | Refresh: Phantom Centering: Steps Phantom Centering is not performed | WIP19 | WIP21 |  | Subsystem | Steps | n/a | Fixed | frider_o | 0 | 4 | DMS Software | DMS | NO | Closed | WIP21SegReq | Ovation |  | WIP21 | October 15, 2019 9:01:34 PM GMT+08:00
CTUDT00626548 | Refresh | Closed | barnea_d | October 7, 2019 4:27:06 PM GMT+08:00 | Refresh:Calibrations:Performance Test:The Center ROI in Head pins1 scan is elliptical. | WIP20 | WIP21 | N/A | Subsystem | Performance;Verified1 | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP21SubCom | Ingenuity 64 |  | WIP21 | October 15, 2019 4:13:19 PM GMT+08:00
CTUDT00626544 | Refresh | Closed | dunitza_m | October 6, 2019 9:15:27 PM GMT+08:00 | ECG Gated – delay End Exam option prior ECG wave arrival | 4.1.7 |  |  | Subsystem |  | N/A | Change Proposal | shaham_t | 0 |  |  | Support-System | YES | Closed |  | Ovation
CTUDT00626544 | Refresh | Closed | dunitza_m | October 6, 2019 9:15:27 PM GMT+08:00 | ECG Gated – delay End Exam option prior ECG wave arrival | 4.1.7 |  |  | Subsystem |  | N/A | Change Proposal | shaham_t | 0 |  |  | Support-System | YES | Closed |  | Brilliance 64
CTUDT00626544 | Refresh | Closed | dunitza_m | October 6, 2019 9:15:27 PM GMT+08:00 | ECG Gated – delay End Exam option prior ECG wave arrival | 4.1.7 |  |  | Subsystem |  | N/A | Change Proposal | shaham_t | 0 |  |  | Support-System | YES | Closed |  | Ingenuity 64
CTUDT00626544 | Refresh | Closed | dunitza_m | October 6, 2019 9:15:27 PM GMT+08:00 | ECG Gated – delay End Exam option prior ECG wave arrival | 4.1.7 |  |  | Subsystem |  | N/A | Change Proposal | shaham_t | 0 |  |  | Support-System | YES | Closed |  | IQon1
CTUDT00626539 | Refresh | Closed | pB_s | October 4, 2019 6:39:45 PM GMT+08:00 | Resp 4D CT scan type is present in the scan type drop down when the " Retrospective Tagging" option not selected in license key | WIP20 | WIP21 | NA | Subsystem |  | NA | Fixed | ravindra_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP21SubCom | Brilliance 64 |  | WIP21 | October 12, 2019 5:35:43 PM GMT+08:00
CTUDT00626539 | Refresh | Closed | pB_s | October 4, 2019 6:39:45 PM GMT+08:00 | Resp 4D CT scan type is present in the scan type drop down when the " Retrospective Tagging" option not selected in license key | WIP20 | WIP21 | NA | Subsystem |  | NA | Fixed | ravindra_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP21SubCom | Ingenuity 64 |  | WIP21 | October 12, 2019 5:35:43 PM GMT+08:00
CTUDT00626537 | Refresh | Closed | manjunath_c | October 4, 2019 3:50:40 PM GMT+08:00 | Deviation in the default parameter values displayed in Exam card manager compared to what is mentioned in the "DHF251245 - Rev.02 Br64 Reference Exam Cards" Specification | WIP19 | WIP22 | No workaround as it is a documentation issue | Subsystem |  | N/A | Fixed | dunitza_m | 0 | 2 | Design Documentation | Support-System | NO | Closed | EC spec update | Brilliance 64 |  | WIP22 | November 28, 2019 4:52:32 PM GMT+08:00
CTUDT00626535 | Refresh | Closed | manjunath_c | October 4, 2019 3:39:17 PM GMT+08:00 | Deviation in the default parameter values displayed in Exam card manager compared to what is mentioned in the "DHF282684 - Rev.05 iQon Reference Exam Cards" Specification | WIP19 | WIP22.1 | This issue relates to Refrnce Exam Cards, it is a spec update issue therefore no workaround is available. | Subsystem |  | N/A | Fixed | dunitza_m | 0 | 2 | Design Documentation | Support-System | NO | Closed | EC Spec update; SubCom-05-Dec-2019 | IQon1 |  | WIP22.1 | December 11, 2019 11:19:35 PM GMT+08:00
CTUDT00626533 | Refresh | Closed | panchami_pradeep | October 4, 2019 11:26:57 AM GMT+08:00 | Refresh: “Wait for the wave…” will disappear though wave is not arrived. | WIP20 | WIP21 | Do not press "Start Final Recon" till scan ends and/or the wave appears. | Subsystem |  | N/A | Fixed | bhawna_g | 0 | 3 | Console Software | Console4 | YES | Closed | WIP21SegReq | Ovation |  | WIP21 | October 12, 2019 1:16:50 AM GMT+08:00
CTUDT00626533 | Refresh | Closed | panchami_pradeep | October 4, 2019 11:26:57 AM GMT+08:00 | Refresh: “Wait for the wave…” will disappear though wave is not arrived. | WIP20 | WIP21 | Do not press "Start Final Recon" till scan ends and/or the wave appears. | Subsystem |  | N/A | Fixed | bhawna_g | 0 | 3 | Console Software | Console4 | YES | Closed | WIP21SegReq | Brilliance 64 |  | WIP21 | October 12, 2019 1:16:50 AM GMT+08:00
CTUDT00626533 | Refresh | Closed | panchami_pradeep | October 4, 2019 11:26:57 AM GMT+08:00 | Refresh: “Wait for the wave…” will disappear though wave is not arrived. | WIP20 | WIP21 | Do not press "Start Final Recon" till scan ends and/or the wave appears. | Subsystem |  | N/A | Fixed | bhawna_g | 0 | 3 | Console Software | Console4 | YES | Closed | WIP21SegReq | Ingenuity 64 |  | WIP21 | October 12, 2019 1:16:50 AM GMT+08:00
CTUDT00626533 | Refresh | Closed | panchami_pradeep | October 4, 2019 11:26:57 AM GMT+08:00 | Refresh: “Wait for the wave…” will disappear though wave is not arrived. | WIP20 | WIP21 | Do not press "Start Final Recon" till scan ends and/or the wave appears. | Subsystem |  | N/A | Fixed | bhawna_g | 0 | 3 | Console Software | Console4 | YES | Closed | WIP21SegReq | IQon1 |  | WIP21 | October 12, 2019 1:16:50 AM GMT+08:00
CTUDT00626531 | Refresh | Closed | Bat-El_Asaraf | October 3, 2019 11:47:17 PM GMT+08:00 | Pre-Val Refresh: Bugrep collect crashes | WIP20 |  |  | Subsystem |  | NA | Invalid/Uncontrolled Configuration | Prakash_Chand_Yadav | 0 |  | Console Software | Console4 | NO | Closed | Invalid Configuration/Not A defect | Brilliance Big Bore
CTUDT00626531 | Refresh | Closed | Bat-El_Asaraf | October 3, 2019 11:47:17 PM GMT+08:00 | Pre-Val Refresh: Bugrep collect crashes | WIP20 |  |  | Subsystem |  | NA | Invalid/Uncontrolled Configuration | Prakash_Chand_Yadav | 0 |  | Console Software | Console4 | NO | Closed | Invalid Configuration/Not A defect | IQon1
CTUDT00626529 | Refresh | Closed | Bat-El_Asaraf | October 3, 2019 11:36:05 PM GMT+08:00 | Pre-Val Refresh: Restore created user accounts fails | WIP20 | WIP22 BB_WIP12 | User could try this operation using IT Admin tool. | Subsystem |  | NA | Fixed | kapadia_v | 0 | 3 | Console Software | Console4 | NO | Closed | WIP21.1 | Ovation |  | WIP22 BB_WIP12 | November 8, 2019 10:14:38 PM GMT+08:00
CTUDT00626529 | Refresh | Closed | Bat-El_Asaraf | October 3, 2019 11:36:05 PM GMT+08:00 | Pre-Val Refresh: Restore created user accounts fails | WIP20 | WIP22 BB_WIP12 | User could try this operation using IT Admin tool. | Subsystem |  | NA | Fixed | kapadia_v | 0 | 3 | Console Software | Console4 | NO | Closed | WIP21.1 | Brilliance 64 |  | WIP22 BB_WIP12 | November 8, 2019 10:14:38 PM GMT+08:00
CTUDT00626529 | Refresh | Closed | Bat-El_Asaraf | October 3, 2019 11:36:05 PM GMT+08:00 | Pre-Val Refresh: Restore created user accounts fails | WIP20 | WIP22 BB_WIP12 | User could try this operation using IT Admin tool. | Subsystem |  | NA | Fixed | kapadia_v | 0 | 3 | Console Software | Console4 | NO | Closed | WIP21.1 | Brilliance Big Bore |  | WIP22 BB_WIP12 | November 8, 2019 10:14:38 PM GMT+08:00
CTUDT00626529 | Refresh | Closed | Bat-El_Asaraf | October 3, 2019 11:36:05 PM GMT+08:00 | Pre-Val Refresh: Restore created user accounts fails | WIP20 | WIP22 BB_WIP12 | User could try this operation using IT Admin tool. | Subsystem |  | NA | Fixed | kapadia_v | 0 | 3 | Console Software | Console4 | NO | Closed | WIP21.1 | Ingenuity 64 |  | WIP22 BB_WIP12 | November 8, 2019 10:14:38 PM GMT+08:00
CTUDT00626529 | Refresh | Closed | Bat-El_Asaraf | October 3, 2019 11:36:05 PM GMT+08:00 | Pre-Val Refresh: Restore created user accounts fails | WIP20 | WIP22 BB_WIP12 | User could try this operation using IT Admin tool. | Subsystem |  | NA | Fixed | kapadia_v | 0 | 3 | Console Software | Console4 | NO | Closed | WIP21.1 | IQon1 |  | WIP22 BB_WIP12 | November 8, 2019 10:14:38 PM GMT+08:00
CTUDT00626526 | Refresh | Closed | madankumar_s | October 3, 2019 10:24:56 PM GMT+08:00 | On selecting the result edited in Slab viewer, ECG series associated with the result is not getting selected. | WIP20 | WIP21 | NA | Subsystem |  | NA | Fixed | sagar_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP21SubCom | Ovation |  | WIP21 | October 12, 2019 2:35:59 PM GMT+08:00
CTUDT00626526 | Refresh | Closed | madankumar_s | October 3, 2019 10:24:56 PM GMT+08:00 | On selecting the result edited in Slab viewer, ECG series associated with the result is not getting selected. | WIP20 | WIP21 | NA | Subsystem |  | NA | Fixed | sagar_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP21SubCom | Brilliance 64 |  | WIP21 | October 12, 2019 2:35:59 PM GMT+08:00
CTUDT00626526 | Refresh | Closed | madankumar_s | October 3, 2019 10:24:56 PM GMT+08:00 | On selecting the result edited in Slab viewer, ECG series associated with the result is not getting selected. | WIP20 | WIP21 | NA | Subsystem |  | NA | Fixed | sagar_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP21SubCom | Ingenuity 64 |  | WIP21 | October 12, 2019 2:35:59 PM GMT+08:00
CTUDT00626526 | Refresh | Closed | madankumar_s | October 3, 2019 10:24:56 PM GMT+08:00 | On selecting the result edited in Slab viewer, ECG series associated with the result is not getting selected. | WIP20 | WIP21 | NA | Subsystem |  | NA | Fixed | sagar_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP21SubCom | IQon1 |  | WIP21 | October 12, 2019 2:35:59 PM GMT+08:00
CTUDT00626524 | Refresh | Closed | barnea_d | October 3, 2019 8:46:07 PM GMT+08:00 | Refresh:Localization: In French, Italian, Norwegian and Swedish Harmonized help page was not fully translated. | WIP20 | WIP21 WIP22.2 |  | Subsystem | localization;Verified1 | N/A | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP22.2SubCom | Ovation |  | WIP21 WIP22.2 | December 9, 2019 3:58:56 PM GMT+08:00
CTUDT00626524 | Refresh | Closed | barnea_d | October 3, 2019 8:46:07 PM GMT+08:00 | Refresh:Localization: In French, Italian, Norwegian and Swedish Harmonized help page was not fully translated. | WIP20 | WIP21 WIP22.2 |  | Subsystem | localization;Verified1 | N/A | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP22.2SubCom | Brilliance 64 |  | WIP21 WIP22.2 | December 9, 2019 3:58:56 PM GMT+08:00
CTUDT00626524 | Refresh | Closed | barnea_d | October 3, 2019 8:46:07 PM GMT+08:00 | Refresh:Localization: In French, Italian, Norwegian and Swedish Harmonized help page was not fully translated. | WIP20 | WIP21 WIP22.2 |  | Subsystem | localization;Verified1 | N/A | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP22.2SubCom | Brilliance Big Bore |  | WIP21 WIP22.2 | December 9, 2019 3:58:56 PM GMT+08:00
CTUDT00626524 | Refresh | Closed | barnea_d | October 3, 2019 8:46:07 PM GMT+08:00 | Refresh:Localization: In French, Italian, Norwegian and Swedish Harmonized help page was not fully translated. | WIP20 | WIP21 WIP22.2 |  | Subsystem | localization;Verified1 | N/A | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP22.2SubCom | Ingenuity 64 |  | WIP21 WIP22.2 | December 9, 2019 3:58:56 PM GMT+08:00
CTUDT00626524 | Refresh | Closed | barnea_d | October 3, 2019 8:46:07 PM GMT+08:00 | Refresh:Localization: In French, Italian, Norwegian and Swedish Harmonized help page was not fully translated. | WIP20 | WIP21 WIP22.2 |  | Subsystem | localization;Verified1 | N/A | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP22.2SubCom | IQon1 |  | WIP21 WIP22.2 | December 9, 2019 3:58:56 PM GMT+08:00
CTUDT00626522 | Refresh | Closed | roy_r | October 3, 2019 8:28:43 PM GMT+08:00 | CT SRS needs update | N/A | ING, BR64 SRS DHF298554 Rev08 BB SRS DHF298555 Rev08 iCT, SRS DHF298553 Rev.08 IQon SRS DHF298556 Rev.09 |  | Subsystem | Verified1 | several SRS | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP21SegReq | Ovation |  | ING, BR64 SRS DHF298554 Rev08 BB SRS DHF298555 Rev08 iCT, SRS DHF298553 Rev.08 IQon SRS DHF298556 Rev.09 | November 6, 2019 2:18:20 PM GMT+08:00
CTUDT00626522 | Refresh | Closed | roy_r | October 3, 2019 8:28:43 PM GMT+08:00 | CT SRS needs update | N/A | ING, BR64 SRS DHF298554 Rev08 BB SRS DHF298555 Rev08 iCT, SRS DHF298553 Rev.08 IQon SRS DHF298556 Rev.09 |  | Subsystem | Verified1 | several SRS | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP21SegReq | Brilliance 64 |  | ING, BR64 SRS DHF298554 Rev08 BB SRS DHF298555 Rev08 iCT, SRS DHF298553 Rev.08 IQon SRS DHF298556 Rev.09 | November 6, 2019 2:18:20 PM GMT+08:00
CTUDT00626522 | Refresh | Closed | roy_r | October 3, 2019 8:28:43 PM GMT+08:00 | CT SRS needs update | N/A | ING, BR64 SRS DHF298554 Rev08 BB SRS DHF298555 Rev08 iCT, SRS DHF298553 Rev.08 IQon SRS DHF298556 Rev.09 |  | Subsystem | Verified1 | several SRS | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP21SegReq | Brilliance Big Bore |  | ING, BR64 SRS DHF298554 Rev08 BB SRS DHF298555 Rev08 iCT, SRS DHF298553 Rev.08 IQon SRS DHF298556 Rev.09 | November 6, 2019 2:18:20 PM GMT+08:00
CTUDT00626522 | Refresh | Closed | roy_r | October 3, 2019 8:28:43 PM GMT+08:00 | CT SRS needs update | N/A | ING, BR64 SRS DHF298554 Rev08 BB SRS DHF298555 Rev08 iCT, SRS DHF298553 Rev.08 IQon SRS DHF298556 Rev.09 |  | Subsystem | Verified1 | several SRS | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP21SegReq | Ingenuity 64 |  | ING, BR64 SRS DHF298554 Rev08 BB SRS DHF298555 Rev08 iCT, SRS DHF298553 Rev.08 IQon SRS DHF298556 Rev.09 | November 6, 2019 2:18:20 PM GMT+08:00
CTUDT00626522 | Refresh | Closed | roy_r | October 3, 2019 8:28:43 PM GMT+08:00 | CT SRS needs update | N/A | ING, BR64 SRS DHF298554 Rev08 BB SRS DHF298555 Rev08 iCT, SRS DHF298553 Rev.08 IQon SRS DHF298556 Rev.09 |  | Subsystem | Verified1 | several SRS | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP21SegReq | IQon1 |  | ING, BR64 SRS DHF298554 Rev08 BB SRS DHF298555 Rev08 iCT, SRS DHF298553 Rev.08 IQon SRS DHF298556 Rev.09 | November 6, 2019 2:18:20 PM GMT+08:00
CTUDT00626519 | Refresh | Closed | manjunath_c | October 3, 2019 6:21:58 PM GMT+08:00 | Deviation in the default parameter values displayed in Exam card manager compared to what is mentioned in the "DHF316206 - Rev.02 iCTSP Reference Exam Cards" Specification | WIP19 | WIP22 | No workaround as it is a spec issue | Subsystem | ECM | N/A | Fixed | dunitza_m | 0 | 3 | Design Documentation | Support-System | NO | Closed | Doc update | Ovation |  | WIP22 | December 3, 2019 1:37:22 AM GMT+08:00
CTUDT00626504 | Refresh | Closed | finestine_m | October 3, 2019 3:57:40 PM GMT+08:00 | Refresh: User Access: BDC Appears For O-Level User | WIP20 | WIP21 |  | Subsystem | user access;Verified1 | na | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP21SubCom | Ovation |  | WIP21 | October 7, 2019 10:23:19 PM GMT+08:00
CTUDT00626504 | Refresh | Closed | finestine_m | October 3, 2019 3:57:40 PM GMT+08:00 | Refresh: User Access: BDC Appears For O-Level User | WIP20 | WIP21 |  | Subsystem | user access;Verified1 | na | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP21SubCom | Brilliance Big Bore |  | WIP21 | October 7, 2019 10:23:19 PM GMT+08:00
CTUDT00626504 | Refresh | Closed | finestine_m | October 3, 2019 3:57:40 PM GMT+08:00 | Refresh: User Access: BDC Appears For O-Level User | WIP20 | WIP21 |  | Subsystem | user access;Verified1 | na | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP21SubCom | Ingenuity 64 |  | WIP21 | October 7, 2019 10:23:19 PM GMT+08:00
CTUDT00626504 | Refresh | Closed | finestine_m | October 3, 2019 3:57:40 PM GMT+08:00 | Refresh: User Access: BDC Appears For O-Level User | WIP20 | WIP21 |  | Subsystem | user access;Verified1 | na | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP21SubCom | IQon1 |  | WIP21 | October 7, 2019 10:23:19 PM GMT+08:00
CTUDT00626500 | Refresh | Closed | manjunath_c | October 3, 2019 3:38:22 PM GMT+08:00 | Deviation in the default parameter values displayed in Exam card manager compared to what is mentioned in the "DHF316204 - Rev.02 iCT Reference Exam Cards" Specification | WIP19 | WIP22.2 | No workaround | Subsystem | ECM | N/A | Fixed | dunitza_m | 0 | 3 | Design Documentation | Support-System | NO | Closed | SubCom-05-Dec-2019 | Ovation |  | WIP22.2 | December 11, 2019 11:18:55 PM GMT+08:00
CTUDT00626498 | Refresh | Closed | shterner_a | October 3, 2019 3:25:22 PM GMT+08:00 | Pre Val Refresh: PRS requirement CT-NM.PRS.CRS.6936 is missing for some Refresh products | DHF298290 Rev 06 | iCT SRS DHF298553 Rev 08 iCT PRS DHF298290 Rev 07 ING, BR64 CRS DHF316435 Rev 03 ING, BR64 SRS DHF298554 Rev 08 |  | Subsystem | Requirements | CT-NM.PRS.CRS.6936, CT-NM.SRS.Modes.984 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP21SegReq | Ovation |  | iCT SRS DHF298553 Rev 08 iCT PRS DHF298290 Rev 07 ING, BR64 CRS DHF316435 Rev 03 ING, BR64 SRS DHF298554 Rev 08 | November 6, 2019 11:15:06 PM GMT+08:00
CTUDT00626498 | Refresh | Closed | shterner_a | October 3, 2019 3:25:22 PM GMT+08:00 | Pre Val Refresh: PRS requirement CT-NM.PRS.CRS.6936 is missing for some Refresh products | DHF298290 Rev 06 | iCT SRS DHF298553 Rev 08 iCT PRS DHF298290 Rev 07 ING, BR64 CRS DHF316435 Rev 03 ING, BR64 SRS DHF298554 Rev 08 |  | Subsystem | Requirements | CT-NM.PRS.CRS.6936, CT-NM.SRS.Modes.984 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP21SegReq | Brilliance 64 |  | iCT SRS DHF298553 Rev 08 iCT PRS DHF298290 Rev 07 ING, BR64 CRS DHF316435 Rev 03 ING, BR64 SRS DHF298554 Rev 08 | November 6, 2019 11:15:06 PM GMT+08:00
CTUDT00626498 | Refresh | Closed | shterner_a | October 3, 2019 3:25:22 PM GMT+08:00 | Pre Val Refresh: PRS requirement CT-NM.PRS.CRS.6936 is missing for some Refresh products | DHF298290 Rev 06 | iCT SRS DHF298553 Rev 08 iCT PRS DHF298290 Rev 07 ING, BR64 CRS DHF316435 Rev 03 ING, BR64 SRS DHF298554 Rev 08 |  | Subsystem | Requirements | CT-NM.PRS.CRS.6936, CT-NM.SRS.Modes.984 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP21SegReq | Ingenuity 64 |  | iCT SRS DHF298553 Rev 08 iCT PRS DHF298290 Rev 07 ING, BR64 CRS DHF316435 Rev 03 ING, BR64 SRS DHF298554 Rev 08 | November 6, 2019 11:15:06 PM GMT+08:00
CTUDT00626497 | Refresh | Closed | bhawna_g | October 3, 2019 2:42:51 PM GMT+08:00 | Refresh: Cardiac: Improper message displayed when ECG signal was OFF and press ‘Go’ for Cardiac scan | 4.8 | BB_WIP11 |  | Subsystem |  | N/A | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | NO | Closed | WIP21SegReq | Brilliance Big Bore |  | BB_WIP11 | October 9, 2019 8:03:45 PM GMT+08:00
CTUDT00626488 | Refresh | Closed | ding_f | October 3, 2019 10:04:54 AM GMT+08:00 | At 20mGy CTDIvol, while keeping high contrast resolution improved, IMR Noise reduction smaller than 90% compared to FBP. | WIP19; WIP9 | WIP21.2 | N/A | System | ;Blocker; Verified1 | CT-NM.SRS.IQ.2540, 3008 | Fixed | richards_b | 0 | 2 | CIRS_SW_ImgNodeSvr | CIRS | NO | Closed | WIP22SegReq | Brilliance 64 |  | WIP21.2 | November 9, 2019 5:30:41 AM GMT+08:00
CTUDT00626488 | Refresh | Closed | ding_f | October 3, 2019 10:04:54 AM GMT+08:00 | At 20mGy CTDIvol, while keeping high contrast resolution improved, IMR Noise reduction smaller than 90% compared to FBP. | WIP19; WIP9 | WIP21.2 | N/A | System | ;Blocker; Verified1 | CT-NM.SRS.IQ.2540, 3008 | Fixed | richards_b | 0 | 2 | CIRS_SW_ImgNodeSvr | CIRS | NO | Closed | WIP22SegReq | Brilliance Big Bore |  | WIP21.2 | November 9, 2019 5:30:41 AM GMT+08:00
CTUDT00626488 | Refresh | Closed | ding_f | October 3, 2019 10:04:54 AM GMT+08:00 | At 20mGy CTDIvol, while keeping high contrast resolution improved, IMR Noise reduction smaller than 90% compared to FBP. | WIP19; WIP9 | WIP21.2 | N/A | System | ;Blocker; Verified1 | CT-NM.SRS.IQ.2540, 3008 | Fixed | richards_b | 0 | 2 | CIRS_SW_ImgNodeSvr | CIRS | NO | Closed | WIP22SegReq | Ingenuity 64 |  | WIP21.2 | November 9, 2019 5:30:41 AM GMT+08:00
CTUDT00626483 | Refresh | Closed | nanjunda_s | October 2, 2019 3:11:02 PM GMT+08:00 | Warning message not displayed for duplicated/added exam cards when Injector summary is not closed/reviewed in the injector. | WIP20 |  | Once the first scan is done, you need to configure the injection for the next step before the scan. This needs to be followed every time if multiple steps with injection are added by duplicating/adding a new step thereafter. | Subsystem |  | NA | No Plan to Fix | datta_i | 0 | 4 | Console Software | Console4 | NO | Closed |  | Ovation
CTUDT00626483 | Refresh | Closed | nanjunda_s | October 2, 2019 3:11:02 PM GMT+08:00 | Warning message not displayed for duplicated/added exam cards when Injector summary is not closed/reviewed in the injector. | WIP20 |  | Once the first scan is done, you need to configure the injection for the next step before the scan. This needs to be followed every time if multiple steps with injection are added by duplicating/adding a new step thereafter. | Subsystem |  | NA | No Plan to Fix | datta_i | 0 | 4 | Console Software | Console4 | NO | Closed |  | Brilliance 64
CTUDT00626483 | Refresh | Closed | nanjunda_s | October 2, 2019 3:11:02 PM GMT+08:00 | Warning message not displayed for duplicated/added exam cards when Injector summary is not closed/reviewed in the injector. | WIP20 |  | Once the first scan is done, you need to configure the injection for the next step before the scan. This needs to be followed every time if multiple steps with injection are added by duplicating/adding a new step thereafter. | Subsystem |  | NA | No Plan to Fix | datta_i | 0 | 4 | Console Software | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00626483 | Refresh | Closed | nanjunda_s | October 2, 2019 3:11:02 PM GMT+08:00 | Warning message not displayed for duplicated/added exam cards when Injector summary is not closed/reviewed in the injector. | WIP20 |  | Once the first scan is done, you need to configure the injection for the next step before the scan. This needs to be followed every time if multiple steps with injection are added by duplicating/adding a new step thereafter. | Subsystem |  | NA | No Plan to Fix | datta_i | 0 | 4 | Console Software | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00626483 | Refresh | Closed | nanjunda_s | October 2, 2019 3:11:02 PM GMT+08:00 | Warning message not displayed for duplicated/added exam cards when Injector summary is not closed/reviewed in the injector. | WIP20 |  | Once the first scan is done, you need to configure the injection for the next step before the scan. This needs to be followed every time if multiple steps with injection are added by duplicating/adding a new step thereafter. | Subsystem |  | NA | No Plan to Fix | datta_i | 0 | 4 | Console Software | Console4 | NO | Closed |  | IQon1
CTUDT00626480 | Refresh | Closed | madankumar_s | October 1, 2019 7:22:59 PM GMT+08:00 | Incorrect message was appeared if an imported language has the same name as an existing language for French and Spanish. | WIP20 | WIP21 |  | Subsystem |  | None | Fixed | kapadia_v | 0 | 4 | Console Software | Console4 | NO | Closed | WIP21SegReq | Ovation |  | WIP21 | October 10, 2019 6:36:15 PM GMT+08:00
CTUDT00626480 | Refresh | Closed | madankumar_s | October 1, 2019 7:22:59 PM GMT+08:00 | Incorrect message was appeared if an imported language has the same name as an existing language for French and Spanish. | WIP20 | WIP21 |  | Subsystem |  | None | Fixed | kapadia_v | 0 | 4 | Console Software | Console4 | NO | Closed | WIP21SegReq | Brilliance 64 |  | WIP21 | October 10, 2019 6:36:15 PM GMT+08:00
CTUDT00626480 | Refresh | Closed | madankumar_s | October 1, 2019 7:22:59 PM GMT+08:00 | Incorrect message was appeared if an imported language has the same name as an existing language for French and Spanish. | WIP20 | WIP21 |  | Subsystem |  | None | Fixed | kapadia_v | 0 | 4 | Console Software | Console4 | NO | Closed | WIP21SegReq | Brilliance Big Bore |  | WIP21 | October 10, 2019 6:36:15 PM GMT+08:00
CTUDT00626480 | Refresh | Closed | madankumar_s | October 1, 2019 7:22:59 PM GMT+08:00 | Incorrect message was appeared if an imported language has the same name as an existing language for French and Spanish. | WIP20 | WIP21 |  | Subsystem |  | None | Fixed | kapadia_v | 0 | 4 | Console Software | Console4 | NO | Closed | WIP21SegReq | Ingenuity 64 |  | WIP21 | October 10, 2019 6:36:15 PM GMT+08:00
CTUDT00626480 | Refresh | Closed | madankumar_s | October 1, 2019 7:22:59 PM GMT+08:00 | Incorrect message was appeared if an imported language has the same name as an existing language for French and Spanish. | WIP20 | WIP21 |  | Subsystem |  | None | Fixed | kapadia_v | 0 | 4 | Console Software | Console4 | NO | Closed | WIP21SegReq | IQon1 |  | WIP21 | October 10, 2019 6:36:15 PM GMT+08:00
CTUDT00626474 | Refresh | Closed | bommiReddy_r | September 30, 2019 9:52:11 PM GMT+08:00 | Refresh: Cardiac: Improper message displayed when ECG signal was OFF and press ‘Go’ for Cardiac scan | WIP20 | WIP21 |  | Subsystem |  | N/A | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | NO | Closed | WIP21SubCom | Ovation |  | WIP21 | October 4, 2019 9:23:55 PM GMT+08:00
CTUDT00626474 | Refresh | Closed | bommiReddy_r | September 30, 2019 9:52:11 PM GMT+08:00 | Refresh: Cardiac: Improper message displayed when ECG signal was OFF and press ‘Go’ for Cardiac scan | WIP20 | WIP21 |  | Subsystem |  | N/A | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | NO | Closed | WIP21SubCom | Brilliance 64 |  | WIP21 | October 4, 2019 9:23:55 PM GMT+08:00
CTUDT00626474 | Refresh | Closed | bommiReddy_r | September 30, 2019 9:52:11 PM GMT+08:00 | Refresh: Cardiac: Improper message displayed when ECG signal was OFF and press ‘Go’ for Cardiac scan | WIP20 | WIP21 |  | Subsystem |  | N/A | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | NO | Closed | WIP21SubCom | Ingenuity 64 |  | WIP21 | October 4, 2019 9:23:55 PM GMT+08:00
CTUDT00626474 | Refresh | Closed | bommiReddy_r | September 30, 2019 9:52:11 PM GMT+08:00 | Refresh: Cardiac: Improper message displayed when ECG signal was OFF and press ‘Go’ for Cardiac scan | WIP20 | WIP21 |  | Subsystem |  | N/A | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | NO | Closed | WIP21SubCom | IQon1 |  | WIP21 | October 4, 2019 9:23:55 PM GMT+08:00
CTUDT00626473 | Refresh | Closed | jotawar_d | September 30, 2019 8:30:05 PM GMT+08:00 | While performing CCT scans, the reference image was not showing the result image. | WIP18 | WIP21 |  | Subsystem |  | Console-SRS.17727 | Fixed | vishnu_devi | 0 | 2 | Console Software | Console4 | NO | Closed | WIP21SubCom | Ovation |  | WIP21 | October 4, 2019 6:58:14 PM GMT+08:00
CTUDT00626473 | Refresh | Closed | jotawar_d | September 30, 2019 8:30:05 PM GMT+08:00 | While performing CCT scans, the reference image was not showing the result image. | WIP18 | WIP21 |  | Subsystem |  | Console-SRS.17727 | Fixed | vishnu_devi | 0 | 2 | Console Software | Console4 | NO | Closed | WIP21SubCom | Brilliance 64 |  | WIP21 | October 4, 2019 6:58:14 PM GMT+08:00
CTUDT00626473 | Refresh | Closed | jotawar_d | September 30, 2019 8:30:05 PM GMT+08:00 | While performing CCT scans, the reference image was not showing the result image. | WIP18 | WIP21 |  | Subsystem |  | Console-SRS.17727 | Fixed | vishnu_devi | 0 | 2 | Console Software | Console4 | NO | Closed | WIP21SubCom | Ingenuity 64 |  | WIP21 | October 4, 2019 6:58:14 PM GMT+08:00
CTUDT00626473 | Refresh | Closed | jotawar_d | September 30, 2019 8:30:05 PM GMT+08:00 | While performing CCT scans, the reference image was not showing the result image. | WIP18 | WIP21 |  | Subsystem |  | Console-SRS.17727 | Fixed | vishnu_devi | 0 | 2 | Console Software | Console4 | NO | Closed | WIP21SubCom | IQon1 |  | WIP21 | October 4, 2019 6:58:14 PM GMT+08:00
CTUDT00626467 | Refresh | Closed | zhang_l | September 30, 2019 4:53:03 PM GMT+08:00 | Chinese study transfer to ISP failed. | WIP19 | WIP23.1 WIP22.3 |  | Subsystem | Blocker | NA | Fixed | s_sa | 0 | 2 | Console Software | iCAP | NO | Closed | WIP24SegReq | Ovation |  | WIP23.1 WIP22.3 | January 14, 2020 9:41:30 PM GMT+08:00
CTUDT00626467 | Refresh | Closed | zhang_l | September 30, 2019 4:53:03 PM GMT+08:00 | Chinese study transfer to ISP failed. | WIP19 | WIP23.1 WIP22.3 |  | Subsystem | Blocker | NA | Fixed | s_sa | 0 | 2 | Console Software | iCAP | NO | Closed | WIP24SegReq | Brilliance 64 |  | WIP23.1 WIP22.3 | January 14, 2020 9:41:30 PM GMT+08:00
CTUDT00626467 | Refresh | Closed | zhang_l | September 30, 2019 4:53:03 PM GMT+08:00 | Chinese study transfer to ISP failed. | WIP19 | WIP23.1 WIP22.3 |  | Subsystem | Blocker | NA | Fixed | s_sa | 0 | 2 | Console Software | iCAP | NO | Closed | WIP24SegReq | Ingenuity 64 |  | WIP23.1 WIP22.3 | January 14, 2020 9:41:30 PM GMT+08:00
CTUDT00626467 | Refresh | Closed | zhang_l | September 30, 2019 4:53:03 PM GMT+08:00 | Chinese study transfer to ISP failed. | WIP19 | WIP23.1 WIP22.3 |  | Subsystem | Blocker | NA | Fixed | s_sa | 0 | 2 | Console Software | iCAP | NO | Closed | WIP24SegReq | IQon1 |  | WIP23.1 WIP22.3 | January 14, 2020 9:41:30 PM GMT+08:00
CTUDT00626466 | Refresh | Closed | roy_r | September 30, 2019 4:32:00 PM GMT+08:00 | Console requirements need update | WIP19 | DHF316705 Rev16 DHF303211 Rev16 |  | Subsystem | Requirements | Console-SRS.16125,15977,2856,17457 | Fixed | roy_r | 0 | 3 | Requirements Documentation | Console4 | NO | Closed | SubCom | Ovation |  | DHF316705 Rev16 DHF303211 Rev16 | October 3, 2019 4:53:58 PM GMT+08:00
CTUDT00626466 | Refresh | Closed | roy_r | September 30, 2019 4:32:00 PM GMT+08:00 | Console requirements need update | WIP19 | DHF316705 Rev16 DHF303211 Rev16 |  | Subsystem | Requirements | Console-SRS.16125,15977,2856,17457 | Fixed | roy_r | 0 | 3 | Requirements Documentation | Console4 | NO | Closed | SubCom | Brilliance 64 |  | DHF316705 Rev16 DHF303211 Rev16 | October 3, 2019 4:53:58 PM GMT+08:00
CTUDT00626466 | Refresh | Closed | roy_r | September 30, 2019 4:32:00 PM GMT+08:00 | Console requirements need update | WIP19 | DHF316705 Rev16 DHF303211 Rev16 |  | Subsystem | Requirements | Console-SRS.16125,15977,2856,17457 | Fixed | roy_r | 0 | 3 | Requirements Documentation | Console4 | NO | Closed | SubCom | Brilliance Big Bore |  | DHF316705 Rev16 DHF303211 Rev16 | October 3, 2019 4:53:58 PM GMT+08:00
CTUDT00626466 | Refresh | Closed | roy_r | September 30, 2019 4:32:00 PM GMT+08:00 | Console requirements need update | WIP19 | DHF316705 Rev16 DHF303211 Rev16 |  | Subsystem | Requirements | Console-SRS.16125,15977,2856,17457 | Fixed | roy_r | 0 | 3 | Requirements Documentation | Console4 | NO | Closed | SubCom | Ingenuity 64 |  | DHF316705 Rev16 DHF303211 Rev16 | October 3, 2019 4:53:58 PM GMT+08:00
CTUDT00626466 | Refresh | Closed | roy_r | September 30, 2019 4:32:00 PM GMT+08:00 | Console requirements need update | WIP19 | DHF316705 Rev16 DHF303211 Rev16 |  | Subsystem | Requirements | Console-SRS.16125,15977,2856,17457 | Fixed | roy_r | 0 | 3 | Requirements Documentation | Console4 | NO | Closed | SubCom | IQon1 |  | DHF316705 Rev16 DHF303211 Rev16 | October 3, 2019 4:53:58 PM GMT+08:00
CTUDT00626464 | Refresh | Closed | roy_r | September 30, 2019 4:24:51 PM GMT+08:00 | Refresh: DMS Diagnostics Tool is not available for Service User. | WIP19 | WIP21 |  | Subsystem |  | STPS-SSRS-1213 | Fixed | Sandeep_Kumar | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP21SubCom | Ovation |  | WIP21 | October 10, 2019 9:19:44 PM GMT+08:00
CTUDT00626464 | Refresh | Closed | roy_r | September 30, 2019 4:24:51 PM GMT+08:00 | Refresh: DMS Diagnostics Tool is not available for Service User. | WIP19 | WIP21 |  | Subsystem |  | STPS-SSRS-1213 | Fixed | Sandeep_Kumar | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP21SubCom | Brilliance 64 |  | WIP21 | October 10, 2019 9:19:44 PM GMT+08:00
CTUDT00626464 | Refresh | Closed | roy_r | September 30, 2019 4:24:51 PM GMT+08:00 | Refresh: DMS Diagnostics Tool is not available for Service User. | WIP19 | WIP21 |  | Subsystem |  | STPS-SSRS-1213 | Fixed | Sandeep_Kumar | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP21SubCom | Ingenuity 64 |  | WIP21 | October 10, 2019 9:19:44 PM GMT+08:00
CTUDT00626464 | Refresh | Closed | roy_r | September 30, 2019 4:24:51 PM GMT+08:00 | Refresh: DMS Diagnostics Tool is not available for Service User. | WIP19 | WIP21 |  | Subsystem |  | STPS-SSRS-1213 | Fixed | Sandeep_Kumar | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP21SubCom | IQon1 |  | WIP21 | October 10, 2019 9:19:44 PM GMT+08:00
CTUDT00626461 | Refresh | Closed | roy_r | September 30, 2019 3:59:01 PM GMT+08:00 | BB Refresh: IVC does not boot up within 30 secs. | BB_WIP9 | DHF316705 Rev16 |  | Subsystem | Requirement | Console-SRS.16545 | Fixed | roy_r | 0 | 3 | HOST SSRS | Console4 | NO | Closed | SubCom | Brilliance Big Bore |  | DHF316705 Rev16 | October 3, 2019 4:53:20 PM GMT+08:00
CTUDT00626460 | Refresh | Closed | jotawar_d | September 30, 2019 3:16:27 PM GMT+08:00 | Refresh - Different thickness for IMR result when performing extension | BB_WIP10 | BB_WIP11 | User can perform offline recon with any thickness. | Subsystem |  | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP11_SubCom | Brilliance Big Bore |  | BB_WIP11 | October 9, 2019 9:29:02 PM GMT+08:00
CTUDT00626459 | Refresh | Closed | jotawar_d | September 30, 2019 3:15:01 PM GMT+08:00 | Layout was not changed to Best Fit after performing Clinical scan and Recon. | BB_WIP10 | BB_WIP11 | Yes, User can manually change the layout to "Best fit" | Subsystem |  | NA | Fixed | prasad_r | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP11_SubCom | Brilliance Big Bore |  | BB_WIP11 | October 9, 2019 1:48:48 PM GMT+08:00
CTUDT00626457 | Refresh | Closed | ye_h | September 30, 2019 11:23:05 AM GMT+08:00 | BB Biocompatibility TRR DHF350963 Rev00 missing IVC biocompatibility statement | BB_WIP9 |  |  | Subsystem | IVC | NA | Does Not Apply | wang_k | 0 |  |  | Support-System | NO | Closed |  | Brilliance Big Bore
CTUDT00626453 | Refresh | Closed | roy_r | September 30, 2019 1:53:20 AM GMT+08:00 | Refresh: New Branding requirements for Ingenuity scanners need to be implemented in Refresh | N/A | WIP21 |  | Subsystem | Branding | NA | Fixed | prasad_r | 0 | 2 | Console Software | Console4 | NO | Closed | WIP21SubCom | Ingenuity 64 |  | WIP21 | October 12, 2019 6:41:06 PM GMT+08:00
CTUDT00626451 | Refresh | Closed | mN_m | September 29, 2019 5:26:38 PM GMT+08:00 | Refresh: Images were taken at some different position which is not the current couch position when CCT scan is repeated. | WIP20 |  |  | Subsystem |  |  | Invalid/Uncontrolled Configuration | bandil_j | 0 |  | Console Software | Support-System | NO | Closed |  | Brilliance Big Bore
CTUDT00626451 | Refresh | Closed | mN_m | September 29, 2019 5:26:38 PM GMT+08:00 | Refresh: Images were taken at some different position which is not the current couch position when CCT scan is repeated. | WIP20 |  |  | Subsystem |  |  | Invalid/Uncontrolled Configuration | bandil_j | 0 |  | Console Software | Support-System | NO | Closed |  | Ingenuity 64
CTUDT00626449 | Refresh | Closed | zhang_t | September 29, 2019 3:59:11 PM GMT+08:00 | ResultEngine stopped working popup when click logout button | WIP20 | WIP22 |  | System |  | NA | Fixed | navale_s | 0 | 4 | Console Software | Console4 | NO | Closed | WIP21SubCom | Ovation |  | WIP22 | November 7, 2019 1:39:47 PM GMT+08:00
CTUDT00626449 | Refresh | Closed | zhang_t | September 29, 2019 3:59:11 PM GMT+08:00 | ResultEngine stopped working popup when click logout button | WIP20 | WIP22 |  | System |  | NA | Fixed | navale_s | 0 | 4 | Console Software | Console4 | NO | Closed | WIP21SubCom | Brilliance 64 |  | WIP22 | November 7, 2019 1:39:47 PM GMT+08:00
CTUDT00626449 | Refresh | Closed | zhang_t | September 29, 2019 3:59:11 PM GMT+08:00 | ResultEngine stopped working popup when click logout button | WIP20 | WIP22 |  | System |  | NA | Fixed | navale_s | 0 | 4 | Console Software | Console4 | NO | Closed | WIP21SubCom | Ingenuity 64 |  | WIP22 | November 7, 2019 1:39:47 PM GMT+08:00
CTUDT00626449 | Refresh | Closed | zhang_t | September 29, 2019 3:59:11 PM GMT+08:00 | ResultEngine stopped working popup when click logout button | WIP20 | WIP22 |  | System |  | NA | Fixed | navale_s | 0 | 4 | Console Software | Console4 | NO | Closed | WIP21SubCom | IQon1 |  | WIP22 | November 7, 2019 1:39:47 PM GMT+08:00
CTUDT00626439 | Refresh | Closed | zhu_je | September 27, 2019 9:01:21 AM GMT+08:00 | [Refresh][IVC]There was no rocker switch angle on IVC box during motion using rocker switch | BB_WIP9 |  |  | Subsystem | IVC | None | Works As Specified / Not A Defect | zhu_je | 1 | 2 |  | Support-System | NO | Closed | Test Design Update | Brilliance Big Bore
CTUDT00626437 | Refresh | Closed | rakic_k | September 27, 2019 3:32:31 AM GMT+08:00 | Missing files in the CIRS BugRep | WIP20 | WIP21 |  | Subsystem | Bugrep; Verified1 | CIRS.SSRS.4919 | Fixed | richards_b | 0 | 2 | CIRS_SW_ServiceNode | CIRS | NO | Closed | WIP21SubCom | Ovation |  | WIP21 | October 21, 2019 9:59:24 PM GMT+08:00
CTUDT00626437 | Refresh | Closed | rakic_k | September 27, 2019 3:32:31 AM GMT+08:00 | Missing files in the CIRS BugRep | WIP20 | WIP21 |  | Subsystem | Bugrep; Verified1 | CIRS.SSRS.4919 | Fixed | richards_b | 0 | 2 | CIRS_SW_ServiceNode | CIRS | NO | Closed | WIP21SubCom | Brilliance 64 |  | WIP21 | October 21, 2019 9:59:24 PM GMT+08:00
CTUDT00626437 | Refresh | Closed | rakic_k | September 27, 2019 3:32:31 AM GMT+08:00 | Missing files in the CIRS BugRep | WIP20 | WIP21 |  | Subsystem | Bugrep; Verified1 | CIRS.SSRS.4919 | Fixed | richards_b | 0 | 2 | CIRS_SW_ServiceNode | CIRS | NO | Closed | WIP21SubCom | Brilliance Big Bore |  | WIP21 | October 21, 2019 9:59:24 PM GMT+08:00
CTUDT00626437 | Refresh | Closed | rakic_k | September 27, 2019 3:32:31 AM GMT+08:00 | Missing files in the CIRS BugRep | WIP20 | WIP21 |  | Subsystem | Bugrep; Verified1 | CIRS.SSRS.4919 | Fixed | richards_b | 0 | 2 | CIRS_SW_ServiceNode | CIRS | NO | Closed | WIP21SubCom | Ingenuity 64 |  | WIP21 | October 21, 2019 9:59:24 PM GMT+08:00
CTUDT00626437 | Refresh | Closed | rakic_k | September 27, 2019 3:32:31 AM GMT+08:00 | Missing files in the CIRS BugRep | WIP20 | WIP21 |  | Subsystem | Bugrep; Verified1 | CIRS.SSRS.4919 | Fixed | richards_b | 0 | 2 | CIRS_SW_ServiceNode | CIRS | NO | Closed | WIP21SubCom | IQon1 |  | WIP21 | October 21, 2019 9:59:24 PM GMT+08:00
CTUDT00626434 | Refresh | Closed | jose_s | September 26, 2019 5:57:13 PM GMT+08:00 | Refresh: Left panel freeze infrequently when the user choose to End Exam. | BB_WIP9 | BB_WIP12 |  | System | Left panel Frozen | NA | Fixed | tamari_o | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP11_SubCom | Brilliance Big Bore |  | BB_WIP12 | November 15, 2019 4:27:51 PM GMT+08:00
CTUDT00626432 | Refresh | Closed | kr_m | September 26, 2019 4:17:42 PM GMT+08:00 | [Refresh]: Default scan type parameters for Body Perfusion-Helical cannot be tested | WIP19 |  | the User can manually change the values. | Subsystem |  | NA | System Repaired | zhu_je | 0 | 2 | Verification Documentation | Support-System | NO | Closed | WIP21SegReq | Ingenuity 64 |  |  | October 16, 2019 8:35:03 AM GMT+08:00
CTUDT00626431 | Refresh | Closed | bhawna_g | September 26, 2019 1:29:51 PM GMT+08:00 | Pre-validation-if I drag plan box out of surview at superior side, we are getting SSDE value to max possible drag ,contrary to this if we are dragging plan box at inferior side ,we are not getting planned SSDE values. It showing NA | 4.8 | BB_WIP11 |  | Subsystem |  | NA | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP11_SubCom | Brilliance Big Bore |  | BB_WIP11 | September 30, 2019 7:44:07 PM GMT+08:00
CTUDT00626429 | Refresh | Closed | marino | September 26, 2019 4:08:14 AM GMT+08:00 | List(Partial)MDU errors cannot be caused, & unexpected reading types apply to all streams. | WIP20 | DHF297589 Rev.15 | N/A | Subsystem | MDU error messages; Verified1 | CIRS.SSRS.5147 & 5176 & 5049 | Fixed | richards_b | 0 | 2 | Requirements Documentation | CIRS | NO | Closed | WIP21SegReq | Brilliance 64 |  | DHF297589 Rev.15 | October 29, 2019 5:53:52 AM GMT+08:00
CTUDT00626429 | Refresh | Closed | marino | September 26, 2019 4:08:14 AM GMT+08:00 | List(Partial)MDU errors cannot be caused, & unexpected reading types apply to all streams. | WIP20 | DHF297589 Rev.15 | N/A | Subsystem | MDU error messages; Verified1 | CIRS.SSRS.5147 & 5176 & 5049 | Fixed | richards_b | 0 | 2 | Requirements Documentation | CIRS | NO | Closed | WIP21SegReq | Brilliance Big Bore |  | DHF297589 Rev.15 | October 29, 2019 5:53:52 AM GMT+08:00
CTUDT00626429 | Refresh | Closed | marino | September 26, 2019 4:08:14 AM GMT+08:00 | List(Partial)MDU errors cannot be caused, & unexpected reading types apply to all streams. | WIP20 | DHF297589 Rev.15 | N/A | Subsystem | MDU error messages; Verified1 | CIRS.SSRS.5147 & 5176 & 5049 | Fixed | richards_b | 0 | 2 | Requirements Documentation | CIRS | NO | Closed | WIP21SegReq | iCT4 |  | DHF297589 Rev.15 | October 29, 2019 5:53:52 AM GMT+08:00
CTUDT00626429 | Refresh | Closed | marino | September 26, 2019 4:08:14 AM GMT+08:00 | List(Partial)MDU errors cannot be caused, & unexpected reading types apply to all streams. | WIP20 | DHF297589 Rev.15 | N/A | Subsystem | MDU error messages; Verified1 | CIRS.SSRS.5147 & 5176 & 5049 | Fixed | richards_b | 0 | 2 | Requirements Documentation | CIRS | NO | Closed | WIP21SegReq | Ingenuity 64 |  | DHF297589 Rev.15 | October 29, 2019 5:53:52 AM GMT+08:00
CTUDT00626429 | Refresh | Closed | marino | September 26, 2019 4:08:14 AM GMT+08:00 | List(Partial)MDU errors cannot be caused, & unexpected reading types apply to all streams. | WIP20 | DHF297589 Rev.15 | N/A | Subsystem | MDU error messages; Verified1 | CIRS.SSRS.5147 & 5176 & 5049 | Fixed | richards_b | 0 | 2 | Requirements Documentation | CIRS | NO | Closed | WIP21SegReq | IQon1 |  | DHF297589 Rev.15 | October 29, 2019 5:53:52 AM GMT+08:00
CTUDT00626423 | Refresh | Closed | ding_f | September 25, 2019 9:37:25 PM GMT+08:00 | Refresh BigBore: Performance Test instructions for legacy system phantom not ready. | BB_WIP9 | DMR350183 Rev04 DMR365968 Rev03 | N/A | Subsystem | Legacy Phantom, Performance Instruction | CT-NM.SRS.IQ.3128 | Does Not Apply | yeshan_m | 0 | 2 | Physics Inspection | Support-System | YES | Closed | WIP21SegReq | Brilliance Big Bore |  | DMR350183 Rev04 DMR365968 Rev03 | November 25, 2019 4:52:30 PM GMT+08:00
CTUDT00626421 | Refresh | Closed | lakshmikanth_s | September 25, 2019 9:32:23 PM GMT+08:00 | Recon is not happening for edit result on slab viewer | BB_WIP9 |  |  | Subsystem |  | Console-SRS.15999 | Works As Specified / Not A Defect | nanjunda_s | 0 |  | Console Software | Console4 | NO | Closed | Not A Defect | Brilliance Big Bore
CTUDT00626421 | Refresh | Closed | lakshmikanth_s | September 25, 2019 9:32:23 PM GMT+08:00 | Recon is not happening for edit result on slab viewer | BB_WIP9 |  |  | Subsystem |  | Console-SRS.15999 | Works As Specified / Not A Defect | nanjunda_s | 0 |  | Console Software | Console4 | NO | Closed | Not A Defect | Ingenuity 64
CTUDT00626418 | Refresh | Closed | mN_m | September 25, 2019 9:12:08 PM GMT+08:00 | BBRefresh: By default none of the option is selected (highlighted) in the network boot settings after the fresh installation. | BB_WIP9 | BB_WIP11 | Select the desired option. | Subsystem | Network boot setting | NA | Fixed | Snehashis_Dey | 0 | 3 | IVC Software | Console4 | NO | Closed | BB_WIP11_SubCom | Brilliance Big Bore |  | BB_WIP11 | October 11, 2019 1:15:11 PM GMT+08:00
CTUDT00626416 | Refresh | Closed | cai_t | September 25, 2019 9:08:51 PM GMT+08:00 | Refresh: The maximum noise is 75.5dB which is more than 70dB in scan room. | BB_WIP9 | DHF326015_Rev06 DHF330999 Rev06 |  | Subsystem |  | CTUDT00626416 | Fixed | cai_t | 0 | 3 | Gantry Mechanical | Support-System | NO | Closed | WIP21SegReq | Brilliance 64 |  | DHF326015_Rev06 DHF330999 Rev06 | October 31, 2019 6:09:05 PM GMT+08:00
CTUDT00626416 | Refresh | Closed | cai_t | September 25, 2019 9:08:51 PM GMT+08:00 | Refresh: The maximum noise is 75.5dB which is more than 70dB in scan room. | BB_WIP9 | DHF326015_Rev06 DHF330999 Rev06 |  | Subsystem |  | CTUDT00626416 | Fixed | cai_t | 0 | 3 | Gantry Mechanical | Support-System | NO | Closed | WIP21SegReq | Brilliance Big Bore |  | DHF326015_Rev06 DHF330999 Rev06 | October 31, 2019 6:09:05 PM GMT+08:00
CTUDT00626416 | Refresh | Closed | cai_t | September 25, 2019 9:08:51 PM GMT+08:00 | Refresh: The maximum noise is 75.5dB which is more than 70dB in scan room. | BB_WIP9 | DHF326015_Rev06 DHF330999 Rev06 |  | Subsystem |  | CTUDT00626416 | Fixed | cai_t | 0 | 3 | Gantry Mechanical | Support-System | NO | Closed | WIP21SegReq | Ingenuity 64 |  | DHF326015_Rev06 DHF330999 Rev06 | October 31, 2019 6:09:05 PM GMT+08:00
CTUDT00626415 | Refresh | Closed | ding_s | September 25, 2019 9:03:34 PM GMT+08:00 | There was only 0.4 sec Rotation option, but no 0.42 sec Rotation option in KGS website. | WIP19 | DHF215793-728232 - Brilliance 64 Upgrade PCM -Rev17 |  | Subsystem |  | NA | Works As Specified / Not A Defect | hong_x | 0 | 2 | PCM | Support-System | NO | Closed | WIP21SegReq | Brilliance 64 |  | DHF215793-728232 - Brilliance 64 Upgrade PCM -Rev17 | October 31, 2019 9:43:08 AM GMT+08:00
CTUDT00626415 | Refresh | Closed | ding_s | September 25, 2019 9:03:34 PM GMT+08:00 | There was only 0.4 sec Rotation option, but no 0.42 sec Rotation option in KGS website. | WIP19 | DHF215793-728232 - Brilliance 64 Upgrade PCM -Rev17 |  | Subsystem |  | NA | Works As Specified / Not A Defect | hong_x | 0 | 2 | PCM | Support-System | NO | Closed | WIP21SegReq | Ingenuity 64 |  | DHF215793-728232 - Brilliance 64 Upgrade PCM -Rev17 | October 31, 2019 9:43:08 AM GMT+08:00
CTUDT00626413 | Refresh | Closed | ding_f | September 25, 2019 8:39:35 PM GMT+08:00 | Refresh Ingenuity: Performance Test instructions for legacy system phantom not ready. | WIP19 | DMR350176 DMR366373 DMR366374 | N/A | Subsystem | Legacy Phantom, Performance Instruction | CT-NM.SRS.IQ.2541 | Fixed | ding_f | 0 | 2 | Physics Instruction | Support-System | YES | Closed | WIP21SegReq | Ingenuity 64 |  | DMR350176 DMR366373 DMR366374 | December 13, 2019 4:38:00 PM GMT+08:00
CTUDT00626412 | Refresh | Closed | liu_bo | September 25, 2019 8:35:50 PM GMT+08:00 | TubeCollimatorFrontBladePos is not correct. | WIP19 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | Wang_Zhen | 0 |  |  | GCX | NO | Closed |  | Ovation
CTUDT00626407 | Refresh | Closed | gT_s | September 25, 2019 6:30:47 PM GMT+08:00 | Refresh : The connection to CIRS is not happening from the PRS portal for Brilliance64, it displays the authentication error. | WIP19 |  |  | Subsystem |  | NA | Invalid/Uncontrolled Configuration | Sandeep_Kumar | 0 |  |  | Service Tools | NO | Closed | Invalid Configuration | Brilliance 64
CTUDT00626405 | Refresh | Closed | datta_i | September 25, 2019 5:56:28 PM GMT+08:00 | Image jumped to the first picture after clicking any key on the keyboard. | 4.8 | BB_WIP11 |  | Subsystem |  | NA | Fixed | challa_s | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP11_SubCom | Brilliance Big Bore |  | BB_WIP11 | September 27, 2019 5:48:58 PM GMT+08:00
CTUDT00626403 | Refresh | Closed | wei_f | September 25, 2019 4:28:13 PM GMT+08:00 | some of The 3rd party software list on the host do not exist in the Refresh BB Configuration Management Plan. | BB_WIP9 | DHF316550 Rev 04 |  | Subsystem | 3rd-party SW | NA | Fixed | venugopal_s | 0 | 2 | Requirements Documentation | DI | NO | Closed | WIP22SegReq | Brilliance 64 |  | DHF316550 Rev 04 | November 19, 2019 1:18:40 PM GMT+08:00
CTUDT00626403 | Refresh | Closed | wei_f | September 25, 2019 4:28:13 PM GMT+08:00 | some of The 3rd party software list on the host do not exist in the Refresh BB Configuration Management Plan. | BB_WIP9 | DHF316550 Rev 04 |  | Subsystem | 3rd-party SW | NA | Fixed | venugopal_s | 0 | 2 | Requirements Documentation | DI | NO | Closed | WIP22SegReq | Brilliance Big Bore |  | DHF316550 Rev 04 | November 19, 2019 1:18:40 PM GMT+08:00
CTUDT00626403 | Refresh | Closed | wei_f | September 25, 2019 4:28:13 PM GMT+08:00 | some of The 3rd party software list on the host do not exist in the Refresh BB Configuration Management Plan. | BB_WIP9 | DHF316550 Rev 04 |  | Subsystem | 3rd-party SW | NA | Fixed | venugopal_s | 0 | 2 | Requirements Documentation | DI | NO | Closed | WIP22SegReq | Ingenuity 64 |  | DHF316550 Rev 04 | November 19, 2019 1:18:40 PM GMT+08:00
CTUDT00626402 | Refresh | Closed | navale_s | September 25, 2019 1:28:55 PM GMT+08:00 | BB Refresh : For Resp 4D CT EC, Phase results are not removed from the offline viewer and in series list | 4.1.10 | WIP21 |  | Subsystem |  | N/A | Fixed | Prakash_Chand_Yadav | 0 | 3 | Console Software | Console4 | NO | Closed | WIP21SubCom | Brilliance 64 |  | WIP21 | October 1, 2019 7:40:42 PM GMT+08:00
CTUDT00626402 | Refresh | Closed | navale_s | September 25, 2019 1:28:55 PM GMT+08:00 | BB Refresh : For Resp 4D CT EC, Phase results are not removed from the offline viewer and in series list | 4.1.10 | WIP21 |  | Subsystem |  | N/A | Fixed | Prakash_Chand_Yadav | 0 | 3 | Console Software | Console4 | NO | Closed | WIP21SubCom | Ingenuity 64 |  | WIP21 | October 1, 2019 7:40:42 PM GMT+08:00
CTUDT00626398 | Refresh | Closed | Sateesh_S | September 25, 2019 11:25:31 AM GMT+08:00 | Actual SSDE is not displaying correctly for the second time pedal press and release for CCT Fluoro | BB_WIP9 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | jindal_k | 0 | 2 | Console Software | Console4 | NO | Closed | WAS | Brilliance Big Bore
CTUDT00626396 | Refresh | Closed | wang_l | September 25, 2019 10:10:16 AM GMT+08:00 | Refresh BB: IVC upgrade failed | BB_WIP9 | BB_WIP11 |  | Subsystem | IVC; | STPS-SSRS-1352, STPS-SSRS-1376 | Fixed | 320021835 | 0 | 2 | IVC Software | Console4 | NO | Closed | BB_WIP11_SubCom | Brilliance Big Bore |  | BB_WIP11 | October 3, 2019 7:26:27 PM GMT+08:00
CTUDT00626394 | Refresh | Closed | wei_f | September 25, 2019 9:23:57 AM GMT+08:00 | Accept as four words when typing two Chinese Character | BB_WIP9 | WIP21.1 |  | Subsystem | Verified1 | NA | Fixed | srushti_k | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP22SegReq | Brilliance 64 |  | WIP21.1 | October 23, 2019 8:13:43 PM GMT+08:00
CTUDT00626394 | Refresh | Closed | wei_f | September 25, 2019 9:23:57 AM GMT+08:00 | Accept as four words when typing two Chinese Character | BB_WIP9 | WIP21.1 |  | Subsystem | Verified1 | NA | Fixed | srushti_k | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP22SegReq | Brilliance Big Bore |  | WIP21.1 | October 23, 2019 8:13:43 PM GMT+08:00
CTUDT00626394 | Refresh | Closed | wei_f | September 25, 2019 9:23:57 AM GMT+08:00 | Accept as four words when typing two Chinese Character | BB_WIP9 | WIP21.1 |  | Subsystem | Verified1 | NA | Fixed | srushti_k | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP22SegReq | Ingenuity 64 |  | WIP21.1 | October 23, 2019 8:13:43 PM GMT+08:00
CTUDT00626392 | Refresh | Closed | hilas | September 24, 2019 10:16:58 PM GMT+08:00 | Refresh: Focal Spot X position and Focal Spot Z position calibrations generate wrong parameters in XRTParams_HeliosData.cfg | WIP19, WIP20 | WIP22 |  | Subsystem |  | N/A | Fixed | zhou_x | 0 | 2 | DMS Calibrations | GCX | YES | Closed | WIP21SegReq | Ovation |  | WIP22 | October 18, 2019 1:05:54 PM GMT+08:00
CTUDT00626392 | Refresh | Closed | hilas | September 24, 2019 10:16:58 PM GMT+08:00 | Refresh: Focal Spot X position and Focal Spot Z position calibrations generate wrong parameters in XRTParams_HeliosData.cfg | WIP19, WIP20 | WIP22 |  | Subsystem |  | N/A | Fixed | zhou_x | 0 | 2 | DMS Calibrations | GCX | YES | Closed | WIP21SegReq | IQon1 |  | WIP22 | October 18, 2019 1:05:54 PM GMT+08:00
CTUDT00626390 | Refresh | Closed | sunilkumar_j | September 24, 2019 9:40:37 PM GMT+08:00 | 'Use spectral recon of:' checkbox was checked for Child spectral result after changing the location. | WIP19 |  | Select any other result from the steplist. Again click on the child result. | System | CRN | N/A | No Plan to Fix | pandey_a | 0 | 3 | Console Software | Console4 | NO | Closed | WIP21SegReq | IQon1
CTUDT00626387 | Refresh | Closed | attali_s | September 24, 2019 8:46:55 PM GMT+08:00 | Refresh: Beam and DMS Check: Mismatch between the slice direction on the plots and in the tables of failures | WIP19 | wip21 |  | Subsystem | Beam and DMS check | na | Fixed | yatziv_l | 0 | 2 | DMS Software | DMS | NO | Closed | WIP21SubCom | Ovation |  | wip21 | October 31, 2019 9:51:09 PM GMT+08:00
CTUDT00626387 | Refresh | Closed | attali_s | September 24, 2019 8:46:55 PM GMT+08:00 | Refresh: Beam and DMS Check: Mismatch between the slice direction on the plots and in the tables of failures | WIP19 | wip21 |  | Subsystem | Beam and DMS check | na | Fixed | yatziv_l | 0 | 2 | DMS Software | DMS | NO | Closed | WIP21SubCom | Brilliance Big Bore |  | wip21 | October 31, 2019 9:51:09 PM GMT+08:00
CTUDT00626387 | Refresh | Closed | attali_s | September 24, 2019 8:46:55 PM GMT+08:00 | Refresh: Beam and DMS Check: Mismatch between the slice direction on the plots and in the tables of failures | WIP19 | wip21 |  | Subsystem | Beam and DMS check | na | Fixed | yatziv_l | 0 | 2 | DMS Software | DMS | NO | Closed | WIP21SubCom | Ingenuity 64 |  | wip21 | October 31, 2019 9:51:09 PM GMT+08:00
CTUDT00626387 | Refresh | Closed | attali_s | September 24, 2019 8:46:55 PM GMT+08:00 | Refresh: Beam and DMS Check: Mismatch between the slice direction on the plots and in the tables of failures | WIP19 | wip21 |  | Subsystem | Beam and DMS check | na | Fixed | yatziv_l | 0 | 2 | DMS Software | DMS | NO | Closed | WIP21SubCom | IQon1 |  | wip21 | October 31, 2019 9:51:09 PM GMT+08:00
CTUDT00626382 | Refresh | Closed | lin_s | September 24, 2019 9:49:18 AM GMT+08:00 | No SSRS link to SystemSupport.913 for BR64 | 4.1.10 | WIP20 |  | Subsystem |  | GCX.SSRS.iCT.11549 | Fixed | wang_k | 0 | 2 | Requirements Documentation | GCX | NO | Closed | WIP21SegReq | Brilliance 64 |  | WIP20 | October 22, 2019 5:02:17 PM GMT+08:00
CTUDT00626380 | Refresh | Closed | guy_sheaffer | September 24, 2019 2:22:09 AM GMT+08:00 | SRS requirement Syssupport.655 is not tagged for iCT-SP | WIP20 | iCT, SRS DHF298553 Rev.08 |  | Subsystem | SRS | Syssupport.655 | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP21SegReq | Ovation |  | iCT, SRS DHF298553 Rev.08 | November 6, 2019 1:48:33 PM GMT+08:00
CTUDT00626378 | Refresh | Closed | zhu_je | September 23, 2019 8:48:31 PM GMT+08:00 | [Refresh]The naming of Logout screen/Open workflow Splash screen/Help page/Login CT Splash screen was "Ingenuity CT powered by iPatient" althought the license key installed was Ingenuity Core or Ingenuity Core128 | WIP19 | DHF298554 Rev08 |  | Subsystem |  | N/A | Fixed | bandil_j | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP21SegReq | Ingenuity 64 |  | DHF298554 Rev08 | November 5, 2019 9:31:15 PM GMT+08:00
CTUDT00626376 | Refresh | Closed | Noy_Seelig | September 23, 2019 8:33:22 PM GMT+08:00 | Refresh: Temperature Notification: In Norwegian the strings are not translated and appear incorrect | WIP19 | WIP21 |  | Subsystem |  | DMS-SRS-1206 | Fixed | franco_a | 0 | 3 | DMS Software | DMS | NO | Closed | WIP21SegReq | Ovation |  | WIP21 | October 6, 2019 3:10:49 PM GMT+08:00
CTUDT00626376 | Refresh | Closed | Noy_Seelig | September 23, 2019 8:33:22 PM GMT+08:00 | Refresh: Temperature Notification: In Norwegian the strings are not translated and appear incorrect | WIP19 | WIP21 |  | Subsystem |  | DMS-SRS-1206 | Fixed | franco_a | 0 | 3 | DMS Software | DMS | NO | Closed | WIP21SegReq | IQon1 |  | WIP21 | October 6, 2019 3:10:49 PM GMT+08:00
CTUDT00626373 | Refresh | Closed | hilas | September 23, 2019 7:53:17 PM GMT+08:00 | Refresh iCT/iCT SP: Reliability Test in Quality Assurance failed | WIP19 | N/A |  | Subsystem |  | n/a | Fixed | hilas | 0 | 2 | Manufacturing Documentation | Support-System | NO | Closed | SubCom-31-Mar-2020 | Ovation |  | N/A | April 1, 2020 2:40:45 AM GMT+08:00
CTUDT00626372 | Refresh | Closed | jose_s | September 23, 2019 6:06:57 PM GMT+08:00 | Refresh: Missing DICOM tags as per recommendation of DIN6862-2 Standard | BB_WIP10 | BB_WIP11 | N/A | Subsystem |  | New ID | Fixed | deshbhratar_m | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP11_SubCom | Brilliance Big Bore |  | BB_WIP11 | October 15, 2019 12:11:01 PM GMT+08:00
CTUDT00626370 | Refresh | Closed | kumar_tp | September 23, 2019 5:55:59 PM GMT+08:00 | [Refresh] Not able to do upgrade from 4.1.7 to 4.1.10 1-click Installation in Z440 Machine | WIP19 | WIP21 BB_WIP11 |  | Subsystem |  | NA | Fixed | venkatachalam_p | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP21SubCom | Ovation |  | WIP21 BB_WIP11 | October 16, 2019 8:15:01 PM GMT+08:00
CTUDT00626370 | Refresh | Closed | kumar_tp | September 23, 2019 5:55:59 PM GMT+08:00 | [Refresh] Not able to do upgrade from 4.1.7 to 4.1.10 1-click Installation in Z440 Machine | WIP19 | WIP21 BB_WIP11 |  | Subsystem |  | NA | Fixed | venkatachalam_p | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP21SubCom | Brilliance 64 |  | WIP21 BB_WIP11 | October 16, 2019 8:15:01 PM GMT+08:00
CTUDT00626370 | Refresh | Closed | kumar_tp | September 23, 2019 5:55:59 PM GMT+08:00 | [Refresh] Not able to do upgrade from 4.1.7 to 4.1.10 1-click Installation in Z440 Machine | WIP19 | WIP21 BB_WIP11 |  | Subsystem |  | NA | Fixed | venkatachalam_p | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP21SubCom | Brilliance Big Bore |  | WIP21 BB_WIP11 | October 16, 2019 8:15:01 PM GMT+08:00
CTUDT00626370 | Refresh | Closed | kumar_tp | September 23, 2019 5:55:59 PM GMT+08:00 | [Refresh] Not able to do upgrade from 4.1.7 to 4.1.10 1-click Installation in Z440 Machine | WIP19 | WIP21 BB_WIP11 |  | Subsystem |  | NA | Fixed | venkatachalam_p | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP21SubCom | Ingenuity 64 |  | WIP21 BB_WIP11 | October 16, 2019 8:15:01 PM GMT+08:00
CTUDT00626370 | Refresh | Closed | kumar_tp | September 23, 2019 5:55:59 PM GMT+08:00 | [Refresh] Not able to do upgrade from 4.1.7 to 4.1.10 1-click Installation in Z440 Machine | WIP19 | WIP21 BB_WIP11 |  | Subsystem |  | NA | Fixed | venkatachalam_p | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP21SubCom | IQon1 |  | WIP21 BB_WIP11 | October 16, 2019 8:15:01 PM GMT+08:00
CTUDT00626366 | Refresh | Closed | wang_l | September 23, 2019 2:46:11 PM GMT+08:00 | Refresh: patient page was not displayed when wait for more than 80 seconds once click End Exam | WIP19 |  |  | Subsystem |  | CT-NM.SRS.Workflow.776 | Works As Specified / Not A Defect | wang_l | 1 |  | Console Software | Support-System | NO | Closed |  | Ingenuity 64
CTUDT00626364 | Refresh | Closed | mN_m | September 23, 2019 1:32:21 PM GMT+08:00 | BB Refresh: Move to inner laser button LED pulsating even after completing the laser motion. | BB_WIP10 | BB_WIP11 |  | Subsystem | IVC | Console-SRS.15655, Console-SRS.15656 | Fixed | 320021835 | 0 | 2 | IVC Software | Console4 | NO | Closed | BB_WIP11_SubCom | Brilliance Big Bore |  | BB_WIP11 | September 27, 2019 7:53:24 PM GMT+08:00
CTUDT00626361 | Refresh | Closed | sun_q | September 23, 2019 1:24:32 PM GMT+08:00 | BB-Error message popped up for Axial Jog with 127 cycles after scan finished | BB_WIP9 | DHF342221 Rev04 |  | Subsystem |  | NA | Fixed | sun_q | 0 | 4 | Gantry Software | GCX | NO | Closed | WIP21SegReq | Brilliance Big Bore |  | DHF342221 Rev04 | October 21, 2019 10:08:38 AM GMT+08:00
CTUDT00626360 | Refresh | Closed | barnea_d | September 23, 2019 1:10:20 PM GMT+08:00 | refresh: acceptance and constancy test: constancy: image parameters are not as defined in the spec | BB_WIP9 | WIP21 |  | Subsystem | Verified1 | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP21SubCom | Brilliance Big Bore |  | WIP21 | October 15, 2019 9:50:29 PM GMT+08:00
CTUDT00626357 | Refresh | Closed | zhu_je | September 22, 2019 8:01:00 PM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.RegStand.711_Cannot find evidence of HU conversion in Br64 TRG | WIP19 | WIP20 | No | Subsystem |  | CT-NM.SRS.RegStand.711 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | Post Validation | Brilliance 64 |  | WIP20 | November 6, 2019 2:47:02 PM GMT+08:00
CTUDT00626355 | Refresh | Closed | swack_g | September 22, 2019 1:49:41 PM GMT+08:00 | Refresh: Physics Analysis: Missing brief summary in MainStatForm window | wip20 |  |  | Subsystem | Physics analysis | n/a | Invalid/Uncontrolled Configuration | frider_o | 0 | 2 |  | DMS | NO | Closed | WIP21SubCom | IQon1
CTUDT00626353 | Refresh | Closed | cai_t | September 21, 2019 7:57:40 PM GMT+08:00 | Refresh: ACR: The phantom is not detected while doing ACR calibration in Harmonized system. | BB_WIP9 |  |  | Subsystem | ACR calibration; H-Phantom | n/a | Works As Specified / Not A Defect | frider_o | 0 |  |  | DMS | NO | Closed |  | Brilliance 64
CTUDT00626353 | Refresh | Closed | cai_t | September 21, 2019 7:57:40 PM GMT+08:00 | Refresh: ACR: The phantom is not detected while doing ACR calibration in Harmonized system. | BB_WIP9 |  |  | Subsystem | ACR calibration; H-Phantom | n/a | Works As Specified / Not A Defect | frider_o | 0 |  |  | DMS | NO | Closed |  | Brilliance Big Bore
CTUDT00626353 | Refresh | Closed | cai_t | September 21, 2019 7:57:40 PM GMT+08:00 | Refresh: ACR: The phantom is not detected while doing ACR calibration in Harmonized system. | BB_WIP9 |  |  | Subsystem | ACR calibration; H-Phantom | n/a | Works As Specified / Not A Defect | frider_o | 0 |  |  | DMS | NO | Closed |  | Ingenuity 64
CTUDT00626351 | Refresh | Closed | wei_f | September 21, 2019 3:35:38 PM GMT+08:00 | The list of ports open on the system did not match the DHF314398 Appendix Active Connections file | BB_WIP9 | WIP21 BB_WIP11 | N/A | Subsystem |  | CT-NM.SRS.SecPri.342 | Fixed | dasgupta_s | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP21SubCom | Ovation |  | WIP21 BB_WIP11 | November 8, 2019 7:34:39 PM GMT+08:00
CTUDT00626351 | Refresh | Closed | wei_f | September 21, 2019 3:35:38 PM GMT+08:00 | The list of ports open on the system did not match the DHF314398 Appendix Active Connections file | BB_WIP9 | WIP21 BB_WIP11 | N/A | Subsystem |  | CT-NM.SRS.SecPri.342 | Fixed | dasgupta_s | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP21SubCom | Brilliance 64 |  | WIP21 BB_WIP11 | November 8, 2019 7:34:39 PM GMT+08:00
CTUDT00626351 | Refresh | Closed | wei_f | September 21, 2019 3:35:38 PM GMT+08:00 | The list of ports open on the system did not match the DHF314398 Appendix Active Connections file | BB_WIP9 | WIP21 BB_WIP11 | N/A | Subsystem |  | CT-NM.SRS.SecPri.342 | Fixed | dasgupta_s | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP21SubCom | Brilliance Big Bore |  | WIP21 BB_WIP11 | November 8, 2019 7:34:39 PM GMT+08:00
CTUDT00626351 | Refresh | Closed | wei_f | September 21, 2019 3:35:38 PM GMT+08:00 | The list of ports open on the system did not match the DHF314398 Appendix Active Connections file | BB_WIP9 | WIP21 BB_WIP11 | N/A | Subsystem |  | CT-NM.SRS.SecPri.342 | Fixed | dasgupta_s | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP21SubCom | Ingenuity 64 |  | WIP21 BB_WIP11 | November 8, 2019 7:34:39 PM GMT+08:00
CTUDT00626351 | Refresh | Closed | wei_f | September 21, 2019 3:35:38 PM GMT+08:00 | The list of ports open on the system did not match the DHF314398 Appendix Active Connections file | BB_WIP9 | WIP21 BB_WIP11 | N/A | Subsystem |  | CT-NM.SRS.SecPri.342 | Fixed | dasgupta_s | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP21SubCom | IQon1 |  | WIP21 BB_WIP11 | November 8, 2019 7:34:39 PM GMT+08:00
CTUDT00626349 | Refresh | Closed | zhu_je | September 21, 2019 9:15:16 AM GMT+08:00 | [Refresh]Default mAs/Default Pitch/Default Rotation Time for some scan type are not same the values in file Scan Type Classifications for DHF251214 Rev06 | WIP19 | WIP21 |  | Subsystem |  | NA | Fixed | kr_m | 0 | 3 | Console Software | Console4 | NO | Closed | WIP21_SubCom | Ingenuity 64 |  | WIP21 | October 12, 2019 5:46:13 PM GMT+08:00
CTUDT00626346 | Refresh | Closed | Sharmila_Vasireddy | September 20, 2019 5:34:20 PM GMT+08:00 | BB refresh: In PD, after recon and End Exam; there are 2 (301 – Exam Summary) Series were generated | BB_WIP9 | BB_WIP11 |  | Subsystem |  | NA | Fixed | vishnu_devi | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP11_SubCom | Brilliance Big Bore |  | BB_WIP11 | September 26, 2019 5:08:35 PM GMT+08:00
CTUDT00626343 | Refresh | Closed | bagdady_b | September 19, 2019 9:09:29 PM GMT+08:00 | Refresh - sCTV: Magic glass is not presenting the correct data from the marked area. | WIP19 | WIP20 |  | System | Magic Glass; Blocker | CT-NM.SRS.TA.435 | Fixed | aviw | 0 | 2 | BaseSpectral | Support_Spectral | NO | Closed | WIP21SegReq | IQon1 |  | WIP20 | February 16, 2020 2:43:48 PM GMT+08:00
CTUDT00626342 | Refresh | Closed | swack_g | September 19, 2019 9:08:23 PM GMT+08:00 | Refresh: Physics Analysis: In Raw Viewer tab, cannot pan left and right. | WIP20 |  |  | Subsystem | Physics Analysis | n/a | No Plan to Fix | frider_o | 0 | 3 |  | DMS | NO | Closed |  | IQon1
CTUDT00626339 | Refresh | Closed | swack_g | September 19, 2019 8:53:55 PM GMT+08:00 | Refresh: Physics Analysis: In Calibration Tables tab, Missing column header for the Slice column | WIP20 |  |  | Subsystem | Physics Analysis | n/a | No Plan to Fix | frider_o | 0 | 3 |  | DMS | NO | Closed |  | IQon1
CTUDT00626333 | Refresh | Closed | swack_g | September 19, 2019 8:39:57 PM GMT+08:00 | Refresh: Physics Analysis: In Calibration Tables tab, STD and Avg rows disappear after showing Bad Detectors vector | WIP20 |  |  | Subsystem | Physics Analysis | n/a | No Plan to Fix | frider_o | 0 | 4 |  | DMS | NO | Closed | SubNC | IQon1
CTUDT00626330 | Refresh | Closed | zhang_l | September 19, 2019 8:00:52 PM GMT+08:00 | Four buttons in the lower right corner overlaped with the data. | WIP19 |  |  | Subsystem |  | NA | No Plan to Fix | aswani_v | 0 | 3 | Console USP_EBW | iCAP | NO | Closed |  | Ovation
CTUDT00626330 | Refresh | Closed | zhang_l | September 19, 2019 8:00:52 PM GMT+08:00 | Four buttons in the lower right corner overlaped with the data. | WIP19 |  |  | Subsystem |  | NA | No Plan to Fix | aswani_v | 0 | 3 | Console USP_EBW | iCAP | NO | Closed |  | Brilliance 64
CTUDT00626330 | Refresh | Closed | zhang_l | September 19, 2019 8:00:52 PM GMT+08:00 | Four buttons in the lower right corner overlaped with the data. | WIP19 |  |  | Subsystem |  | NA | No Plan to Fix | aswani_v | 0 | 3 | Console USP_EBW | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00626330 | Refresh | Closed | zhang_l | September 19, 2019 8:00:52 PM GMT+08:00 | Four buttons in the lower right corner overlaped with the data. | WIP19 |  |  | Subsystem |  | NA | No Plan to Fix | aswani_v | 0 | 3 | Console USP_EBW | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00626330 | Refresh | Closed | zhang_l | September 19, 2019 8:00:52 PM GMT+08:00 | Four buttons in the lower right corner overlaped with the data. | WIP19 |  |  | Subsystem |  | NA | No Plan to Fix | aswani_v | 0 | 3 | Console USP_EBW | iCAP | NO | Closed |  | IQon1
CTUDT00626324 | Refresh | Closed | liu_bo | September 19, 2019 6:10:54 PM GMT+08:00 | Headline The follow reason matrix doesn't match with iQon Gantry SSRS. New Site Installation and Tube Replacement | WIP19 | Gantry SSRS 06 |  | Subsystem |  | GCX.SSRS.iCT.19246 | Fixed | bruce_li | 0 | 3 | Requirements Documentation | GCX | NO | Closed | WIP21SegReq | IQon1 |  | Gantry SSRS 06 | October 22, 2019 9:13:54 AM GMT+08:00
CTUDT00626321 | Refresh | Closed | zhang_l | September 19, 2019 4:37:15 PM GMT+08:00 | Image jumped to the first picture after clicking any key on the keyboard. | WIP19 | WIP21 |  | Subsystem |  | NA | Fixed | challa_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP21SubCom | Ovation |  | WIP21 | September 25, 2019 5:43:54 PM GMT+08:00
CTUDT00626321 | Refresh | Closed | zhang_l | September 19, 2019 4:37:15 PM GMT+08:00 | Image jumped to the first picture after clicking any key on the keyboard. | WIP19 | WIP21 |  | Subsystem |  | NA | Fixed | challa_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP21SubCom | Brilliance 64 |  | WIP21 | September 25, 2019 5:43:54 PM GMT+08:00
CTUDT00626321 | Refresh | Closed | zhang_l | September 19, 2019 4:37:15 PM GMT+08:00 | Image jumped to the first picture after clicking any key on the keyboard. | WIP19 | WIP21 |  | Subsystem |  | NA | Fixed | challa_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP21SubCom | Ingenuity 64 |  | WIP21 | September 25, 2019 5:43:54 PM GMT+08:00
CTUDT00626321 | Refresh | Closed | zhang_l | September 19, 2019 4:37:15 PM GMT+08:00 | Image jumped to the first picture after clicking any key on the keyboard. | WIP19 | WIP21 |  | Subsystem |  | NA | Fixed | challa_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP21SubCom | IQon1 |  | WIP21 | September 25, 2019 5:43:54 PM GMT+08:00
CTUDT00626317 | Refresh | Closed | zhu_je | September 19, 2019 9:26:30 AM GMT+08:00 | [Refresh]Maximum pitch and pitch menu for scan type Resp. Free Breathing and Infant Body are not same the values in file Scan Type Classifications for DHF251214 Rev06 | WIP19 |  |  | Subsystem | Test Design | CT-NM.SRS.Workflow.1843 | Documentation Created/Updated | zhu_je | 0 | 2 | Test Design | Support-System | NO | Closed | test case update | Ingenuity 64 |  |  | October 15, 2019 7:04:36 PM GMT+08:00
CTUDT00626312 | Refresh | Closed | Sharmila_Vasireddy | September 18, 2019 9:35:36 PM GMT+08:00 | BB refresh: “PIPE” configuration was not uninstalled after US-DoD config tool start | BB_WIP9 | BB_WIP11 | Philips_service user can manually uninstall the PIPE MSI. | Subsystem | Service Tools | NA | Fixed | prasad_r | 0 | 3 | Console Software | Service Tools | NO | Closed | BB_WIP11_SubCom | Brilliance Big Bore |  | BB_WIP11 | September 26, 2019 6:00:28 PM GMT+08:00
CTUDT00626309 | Refresh | Closed | vineetha_s | September 18, 2019 8:32:44 PM GMT+08:00 | Refresh: Chinese Keyboard Localization Issues | WIP19 |  |  | Subsystem | Localization;CRN | NA | No Plan to Fix | srushti_k | 0 | 4 |  | iCAP | NO | Closed |  | Ovation
CTUDT00626309 | Refresh | Closed | vineetha_s | September 18, 2019 8:32:44 PM GMT+08:00 | Refresh: Chinese Keyboard Localization Issues | WIP19 |  |  | Subsystem | Localization;CRN | NA | No Plan to Fix | srushti_k | 0 | 4 |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00626309 | Refresh | Closed | vineetha_s | September 18, 2019 8:32:44 PM GMT+08:00 | Refresh: Chinese Keyboard Localization Issues | WIP19 |  |  | Subsystem | Localization;CRN | NA | No Plan to Fix | srushti_k | 0 | 4 |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00626309 | Refresh | Closed | vineetha_s | September 18, 2019 8:32:44 PM GMT+08:00 | Refresh: Chinese Keyboard Localization Issues | WIP19 |  |  | Subsystem | Localization;CRN | NA | No Plan to Fix | srushti_k | 0 | 4 |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00626309 | Refresh | Closed | vineetha_s | September 18, 2019 8:32:44 PM GMT+08:00 | Refresh: Chinese Keyboard Localization Issues | WIP19 |  |  | Subsystem | Localization;CRN | NA | No Plan to Fix | srushti_k | 0 | 4 |  | iCAP | NO | Closed |  | IQon1
CTUDT00626307 | Refresh | Closed | ran_y | September 18, 2019 8:29:01 PM GMT+08:00 | Workflow Refresh test design | WIP19 | WIP22 |  | Subsystem |  | CT-NM.workflow.2279;CT-NM.workflow.2176 | Fixed | tamari_o | 0 | 2 | Verification Documentation | Support-System | NO | Closed | WIP22SegReq | Ovation |  | WIP22 | November 19, 2019 1:39:13 PM GMT+08:00
CTUDT00626307 | Refresh | Closed | ran_y | September 18, 2019 8:29:01 PM GMT+08:00 | Workflow Refresh test design | WIP19 | WIP22 |  | Subsystem |  | CT-NM.workflow.2279;CT-NM.workflow.2176 | Fixed | tamari_o | 0 | 2 | Verification Documentation | Support-System | NO | Closed | WIP22SegReq | IQon1 |  | WIP22 | November 19, 2019 1:39:13 PM GMT+08:00
CTUDT00626305 | Refresh | Closed | nanjunda_s | September 18, 2019 8:05:13 PM GMT+08:00 | Deviation in the parameter values displayed in Exam Card Manager compared to what is mentioned in the QA exam cards in the Specification | WIP19 | WIP21 | User can manually select the values from the dropdown or type for few parameters. | Subsystem | ECM | NA | Fixed | braunstein_d | 0 | 2 | Console Software | Support-Physics | YES | Closed | WIP21SubCom | Brilliance 64 |  | WIP21 | October 27, 2019 8:20:23 PM GMT+08:00
CTUDT00626305 | Refresh | Closed | nanjunda_s | September 18, 2019 8:05:13 PM GMT+08:00 | Deviation in the parameter values displayed in Exam Card Manager compared to what is mentioned in the QA exam cards in the Specification | WIP19 | WIP21 | User can manually select the values from the dropdown or type for few parameters. | Subsystem | ECM | NA | Fixed | braunstein_d | 0 | 2 | Console Software | Support-Physics | YES | Closed | WIP21SubCom | Ingenuity 64 |  | WIP21 | October 27, 2019 8:20:23 PM GMT+08:00
CTUDT00626303 | Refresh | Closed | manjunath_c | September 18, 2019 7:53:20 PM GMT+08:00 | Deviation in the parameter values displayed in Exam Card Manager compared to what is mentioned in the Scan Type Specification - DHF251214 Rev06 | WIP19 | WIP20 | No workaround for the clinical user | Subsystem | ECM | NA | Fixed | dunitza_m | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | SPEC UPDATE | Brilliance 64 |  | WIP20 | November 7, 2019 8:46:27 PM GMT+08:00
CTUDT00626303 | Refresh | Closed | manjunath_c | September 18, 2019 7:53:20 PM GMT+08:00 | Deviation in the parameter values displayed in Exam Card Manager compared to what is mentioned in the Scan Type Specification - DHF251214 Rev06 | WIP19 | WIP20 | No workaround for the clinical user | Subsystem | ECM | NA | Fixed | dunitza_m | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | SPEC UPDATE | Ingenuity 64 |  | WIP20 | November 7, 2019 8:46:27 PM GMT+08:00
CTUDT00626300 | Refresh | Closed | manjunath_c | September 18, 2019 6:56:58 PM GMT+08:00 | Deviation in the default parameter values displayed in Exam card manager compared to what is mentioned in the "DHF251235 - Rev.03 Ingenuity Reference Exam Cards" Specification | WIP19 | WIP22 | In Clinical use it is possible to change parameters through the ECM. No workaround for documentation issue | Subsystem | ECM | N/A | Fixed | dunitza_m | 0 | 2 | Console Software | Support-System | NO | Closed | DHF251235 - Rev.04 | Ingenuity 64 |  | WIP22 | November 28, 2019 4:16:06 PM GMT+08:00
CTUDT00626298 | Refresh | Closed | mN_m | September 18, 2019 5:43:37 PM GMT+08:00 | BBRefresh: Vertical motion is not continuous, the couch stops at different location in between. | BB_WIP9 |  |  | Subsystem | Couch; IVC |  | Works As Specified / Not A Defect | mN_m | 1 |  |  | Console4 | NO | Closed | TBR | Brilliance Big Bore
CTUDT00626296 | Refresh | Closed | mN_m | September 18, 2019 5:23:54 PM GMT+08:00 | BB Refresh: Couch moves to a previously saved position when the position is updated after the radio button is selected for all the three controlled motion from IV console. | BB_WIP9 | BB_WIP12 |  | Subsystem | IVC;CCT; Blocker21 | N/A | Fixed | navale_s | 0 | 2 | Console Software | Console4 | YES | Closed | BB_WIP10_SubCom | Brilliance Big Bore |  | BB_WIP12 | November 15, 2019 4:22:39 PM GMT+08:00
CTUDT00626293 | Refresh | Closed | zhu_je | September 18, 2019 4:58:20 PM GMT+08:00 | [Refresh] Default rotation time and Rotation Time Menu for some scan type are not match with values in file Scan Type Classifications for DHF251214 Rev06 | WIP19 |  |  | Subsystem | ECM |  | Duplicate | dunitza_m | 0 |  |  | Support-System | NO | Closed |  | Ingenuity 64
CTUDT00626290 | Refresh | Closed | Ashish_Sharma | September 18, 2019 3:22:11 PM GMT+08:00 | Pre-validation-if I drag plan box out of surview at superior side, we are getting SSDE value to max possible drag ,contrary to this if we are dragging plan box at inferior side ,we are not getting planned SSDE values. It showing NA | WIP19 | WIP21 |  | Subsystem |  | NA | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | NO | Closed | WIP21SubCom | Ovation |  | WIP21 | September 27, 2019 1:39:30 PM GMT+08:00
CTUDT00626290 | Refresh | Closed | Ashish_Sharma | September 18, 2019 3:22:11 PM GMT+08:00 | Pre-validation-if I drag plan box out of surview at superior side, we are getting SSDE value to max possible drag ,contrary to this if we are dragging plan box at inferior side ,we are not getting planned SSDE values. It showing NA | WIP19 | WIP21 |  | Subsystem |  | NA | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | NO | Closed | WIP21SubCom | Brilliance 64 |  | WIP21 | September 27, 2019 1:39:30 PM GMT+08:00
CTUDT00626290 | Refresh | Closed | Ashish_Sharma | September 18, 2019 3:22:11 PM GMT+08:00 | Pre-validation-if I drag plan box out of surview at superior side, we are getting SSDE value to max possible drag ,contrary to this if we are dragging plan box at inferior side ,we are not getting planned SSDE values. It showing NA | WIP19 | WIP21 |  | Subsystem |  | NA | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | NO | Closed | WIP21SubCom | Ingenuity 64 |  | WIP21 | September 27, 2019 1:39:30 PM GMT+08:00
CTUDT00626290 | Refresh | Closed | Ashish_Sharma | September 18, 2019 3:22:11 PM GMT+08:00 | Pre-validation-if I drag plan box out of surview at superior side, we are getting SSDE value to max possible drag ,contrary to this if we are dragging plan box at inferior side ,we are not getting planned SSDE values. It showing NA | WIP19 | WIP21 |  | Subsystem |  | NA | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | NO | Closed | WIP21SubCom | IQon1 |  | WIP21 | September 27, 2019 1:39:30 PM GMT+08:00
CTUDT00626288 | Refresh | Closed | levy_s | September 18, 2019 2:14:52 PM GMT+08:00 | Refresh: While scanning EC with IMR and spectral results the images are not reconstructed | WIP19 |  | Offline recon | Subsystem | CRN; | Console-SRS.9804 | No Plan to Fix | singh_j | 0 | 3 |  | Console4 | YES | Closed | SubNC | IQon1
CTUDT00626285 | Refresh | Closed | zhu_je | September 18, 2019 11:01:14 AM GMT+08:00 | [Refresh] Report templates page was disappeared after loading for a few seconds | BB_WIP9 | WIP21.1 |  | Subsystem | Report template page;Verified1 | CT-NM.SRS.Workflow.1861 | Fixed | s_panigrahi | 0 | 2 | Console USP_EBW | iCAP | NO | Closed |  | Ovation |  | WIP21.1 | October 23, 2019 12:44:11 PM GMT+08:00
CTUDT00626285 | Refresh | Closed | zhu_je | September 18, 2019 11:01:14 AM GMT+08:00 | [Refresh] Report templates page was disappeared after loading for a few seconds | BB_WIP9 | WIP21.1 |  | Subsystem | Report template page;Verified1 | CT-NM.SRS.Workflow.1861 | Fixed | s_panigrahi | 0 | 2 | Console USP_EBW | iCAP | NO | Closed |  | Brilliance 64 |  | WIP21.1 | October 23, 2019 12:44:11 PM GMT+08:00
CTUDT00626285 | Refresh | Closed | zhu_je | September 18, 2019 11:01:14 AM GMT+08:00 | [Refresh] Report templates page was disappeared after loading for a few seconds | BB_WIP9 | WIP21.1 |  | Subsystem | Report template page;Verified1 | CT-NM.SRS.Workflow.1861 | Fixed | s_panigrahi | 0 | 2 | Console USP_EBW | iCAP | NO | Closed |  | Brilliance Big Bore |  | WIP21.1 | October 23, 2019 12:44:11 PM GMT+08:00
CTUDT00626285 | Refresh | Closed | zhu_je | September 18, 2019 11:01:14 AM GMT+08:00 | [Refresh] Report templates page was disappeared after loading for a few seconds | BB_WIP9 | WIP21.1 |  | Subsystem | Report template page;Verified1 | CT-NM.SRS.Workflow.1861 | Fixed | s_panigrahi | 0 | 2 | Console USP_EBW | iCAP | NO | Closed |  | Ingenuity 64 |  | WIP21.1 | October 23, 2019 12:44:11 PM GMT+08:00
CTUDT00626285 | Refresh | Closed | zhu_je | September 18, 2019 11:01:14 AM GMT+08:00 | [Refresh] Report templates page was disappeared after loading for a few seconds | BB_WIP9 | WIP21.1 |  | Subsystem | Report template page;Verified1 | CT-NM.SRS.Workflow.1861 | Fixed | s_panigrahi | 0 | 2 | Console USP_EBW | iCAP | NO | Closed |  | IQon1 |  | WIP21.1 | October 23, 2019 12:44:11 PM GMT+08:00
CTUDT00626283 | Refresh | Closed | zhu_je | September 18, 2019 10:52:00 AM GMT+08:00 | [Refresh] Default Pitch , Maximum Pitch , Pitch Menu ,thickness menu and Default mAs of some ECs did not match with the values specified in DHF209557 Rev12 scan type document Big Bore v4.8 Scan Type Classification.xlsx | BB_WIP9 | DHF209557 Rev15 |  | Subsystem | Blocker | NA | Fixed | dunitza_m | 0 | 2 | Design Documentation | Support-System | NO | Closed | Spec update | Brilliance Big Bore |  | DHF209557 Rev15 | December 12, 2019 6:23:08 PM GMT+08:00
CTUDT00626281 | Refresh | Closed | madankumar_s | September 17, 2019 8:45:33 PM GMT+08:00 | Layout was not changed to Best Fit after performing Clinical scan and Recon. | WIP19 | WIP21 | Yes, User can manually change the layout to "Best fit" | Subsystem | Acquisition Viewer | NA | Fixed | kr_m | 0 | 3 | Console Software | Console4 | NO | Closed | WIP21SubCom | Ovation |  | WIP21 | September 26, 2019 5:47:34 PM GMT+08:00
CTUDT00626281 | Refresh | Closed | madankumar_s | September 17, 2019 8:45:33 PM GMT+08:00 | Layout was not changed to Best Fit after performing Clinical scan and Recon. | WIP19 | WIP21 | Yes, User can manually change the layout to "Best fit" | Subsystem | Acquisition Viewer | NA | Fixed | kr_m | 0 | 3 | Console Software | Console4 | NO | Closed | WIP21SubCom | Brilliance 64 |  | WIP21 | September 26, 2019 5:47:34 PM GMT+08:00
CTUDT00626281 | Refresh | Closed | madankumar_s | September 17, 2019 8:45:33 PM GMT+08:00 | Layout was not changed to Best Fit after performing Clinical scan and Recon. | WIP19 | WIP21 | Yes, User can manually change the layout to "Best fit" | Subsystem | Acquisition Viewer | NA | Fixed | kr_m | 0 | 3 | Console Software | Console4 | NO | Closed | WIP21SubCom | Ingenuity 64 |  | WIP21 | September 26, 2019 5:47:34 PM GMT+08:00
CTUDT00626281 | Refresh | Closed | madankumar_s | September 17, 2019 8:45:33 PM GMT+08:00 | Layout was not changed to Best Fit after performing Clinical scan and Recon. | WIP19 | WIP21 | Yes, User can manually change the layout to "Best fit" | Subsystem | Acquisition Viewer | NA | Fixed | kr_m | 0 | 3 | Console Software | Console4 | NO | Closed | WIP21SubCom | IQon1 |  | WIP21 | September 26, 2019 5:47:34 PM GMT+08:00
CTUDT00626280 | Refresh | Closed | deshbhratar_m | September 17, 2019 7:44:36 PM GMT+08:00 | Liver DRI value is not updating in Scan parameter Tab | 4.8 | BB_WIP11 | Yes.To see the selected Liver DRI value on the Scan Parameters Toolbox, user has to perform a scan .After the scan the updated Liver DRI value shall be reflected under Parameters Panel region | Subsystem | Liver DRI | NA | Fixed | deshbhratar_m | 0 | 2 | Console Software | Console4 | YES | Closed | BB_WIP11_SubCom | Brilliance Big Bore |  | BB_WIP11 | September 23, 2019 5:53:07 PM GMT+08:00
CTUDT00626279 | Refresh | Closed | deshbhratar_m | September 17, 2019 7:08:07 PM GMT+08:00 | Bug rep error is prompts and after click on "Ok" on the pop up message, "System error" was displayed | 4.8 | BB_WIP11 |  | Subsystem | System Error displayed | None | Fixed | deshbhratar_m | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP11_SubCom | Brilliance Big Bore |  | BB_WIP11 | September 23, 2019 6:25:37 PM GMT+08:00
CTUDT00626277 | Refresh | Closed | bommiReddy_r | September 17, 2019 7:00:50 PM GMT+08:00 | Refresh:HW Scoping: Images from CDR to DVD burning process was not successful | WIP19 |  |  | Subsystem | CD-DVD operation | NA | Invalid/Uncontrolled Configuration | s_sa | 0 |  |  | iCAP | NO | Closed |  | Ovation
CTUDT00626277 | Refresh | Closed | bommiReddy_r | September 17, 2019 7:00:50 PM GMT+08:00 | Refresh:HW Scoping: Images from CDR to DVD burning process was not successful | WIP19 |  |  | Subsystem | CD-DVD operation | NA | Invalid/Uncontrolled Configuration | s_sa | 0 |  |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00626277 | Refresh | Closed | bommiReddy_r | September 17, 2019 7:00:50 PM GMT+08:00 | Refresh:HW Scoping: Images from CDR to DVD burning process was not successful | WIP19 |  |  | Subsystem | CD-DVD operation | NA | Invalid/Uncontrolled Configuration | s_sa | 0 |  |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00626277 | Refresh | Closed | bommiReddy_r | September 17, 2019 7:00:50 PM GMT+08:00 | Refresh:HW Scoping: Images from CDR to DVD burning process was not successful | WIP19 |  |  | Subsystem | CD-DVD operation | NA | Invalid/Uncontrolled Configuration | s_sa | 0 |  |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00626277 | Refresh | Closed | bommiReddy_r | September 17, 2019 7:00:50 PM GMT+08:00 | Refresh:HW Scoping: Images from CDR to DVD burning process was not successful | WIP19 |  |  | Subsystem | CD-DVD operation | NA | Invalid/Uncontrolled Configuration | s_sa | 0 |  |  | iCAP | NO | Closed |  | IQon1
CTUDT00626275 | Refresh | Closed | navale_s | September 17, 2019 4:35:39 PM GMT+08:00 | [IVC] Localization is not present for the newly added requirements for Interventional controls | v4.8 | BB_WIP11 |  | Subsystem | Localization | Console-SRS.17821, Console-SRS.17796 | Fixed | wilsonraja | 0 | 2 | Console Software | Console4 | YES | Closed | BB_WIP11_SubCom | Brilliance Big Bore |  | BB_WIP11 | October 10, 2019 11:44:32 PM GMT+08:00
CTUDT00626272 | Refresh | Closed | pB_s | September 17, 2019 3:51:22 PM GMT+08:00 | "Backup Date & Time" label is not translated properly to Dutch language under Restore User Accounts section in IT tool | WIP19 |  |  | Subsystem | Backup/Restore |  | Works As Specified / Not A Defect | prabhu_m | 0 |  |  | Service Tools | NO | Closed |  | Ingenuity 64
CTUDT00626268 | Refresh | Closed | cohen_y | September 17, 2019 3:08:28 AM GMT+08:00 | Refresh: Performance Test: One Part Phantom: Recon Matrix of Head CT13, 14, 15 and Body CT11, 12 cases are not according to Spec | WIP19 | N/A |  | Subsystem | Documentation | N/A | Fixed | braunstein_d | 0 | 2 | Design Documentation | Support-Physics | NO | Closed | WIP21SegReq | Ingenuity 64 |  | N/A | October 27, 2019 8:33:14 PM GMT+08:00
CTUDT00626266 | Refresh | Closed | marino | September 16, 2019 10:26:26 PM GMT+08:00 | The average TRT across the helical protocols of interest is longer across two CIRS computers than it is across one CIRS computer. | WIP20 |  |  | Subsystem | Performance | na | Works As Specified / Not A Defect | marino | 0 | 2 |  | CIRS | NO | Closed | SRS | Ovation
CTUDT00626266 | Refresh | Closed | marino | September 16, 2019 10:26:26 PM GMT+08:00 | The average TRT across the helical protocols of interest is longer across two CIRS computers than it is across one CIRS computer. | WIP20 |  |  | Subsystem | Performance | na | Works As Specified / Not A Defect | marino | 0 | 2 |  | CIRS | NO | Closed | SRS | Brilliance 64
CTUDT00626266 | Refresh | Closed | marino | September 16, 2019 10:26:26 PM GMT+08:00 | The average TRT across the helical protocols of interest is longer across two CIRS computers than it is across one CIRS computer. | WIP20 |  |  | Subsystem | Performance | na | Works As Specified / Not A Defect | marino | 0 | 2 |  | CIRS | NO | Closed | SRS | Ingenuity 64
CTUDT00626266 | Refresh | Closed | marino | September 16, 2019 10:26:26 PM GMT+08:00 | The average TRT across the helical protocols of interest is longer across two CIRS computers than it is across one CIRS computer. | WIP20 |  |  | Subsystem | Performance | na | Works As Specified / Not A Defect | marino | 0 | 2 |  | CIRS | NO | Closed | SRS | IQon1
CTUDT00626264 | Refresh | Closed | Sharmila_Vasireddy | September 16, 2019 10:05:33 PM GMT+08:00 | BB Refresh : For Resp 4D CT EC, Phase results are not removed from the offline viewer and in series list | BB_WIP9 | BB_WIP11 |  | Subsystem | Respiratory | N/A | Fixed | navale_s | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP11_SubCom | Brilliance Big Bore |  | BB_WIP11 | September 23, 2019 8:14:41 PM GMT+08:00
CTUDT00626261 | Refresh | Closed | wang_l | September 16, 2019 9:22:15 PM GMT+08:00 | Refresh BB: Inherit issue for system requirement SysSupport.645 and Workflow.2481,2805 | BB_WIP9 | D-394190 |  | Subsystem | Requirments | CT-NM.SRS.SysSupport.645 | Fixed | dongping_w | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | SubCom-25-Nov-2019 | Brilliance Big Bore |  | D-394190 | December 17, 2019 11:10:05 AM GMT+08:00
CTUDT00626256 | Refresh | Closed | Jose_D_Sergio_Almeida | September 16, 2019 5:10:36 PM GMT+08:00 | Pre-Validation: Mismatch in 3D modulation mAs profile for IN and OUT scan direction. | WIP19 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | morton_t | 0 |  |  | CIRS | NO | Closed |  | Ingenuity 64
CTUDT00626254 | Refresh | Closed | lei_yang | September 16, 2019 3:34:36 PM GMT+08:00 | The waveform was not imported to Console When the Varian device and the console patient had the same ID. | WIP19 | BB_WIP11.1 WIP21.1 |  | Subsystem |  | NA | Fixed | navale_s | 0 | 2 | Console Software | Console4 | YES | Closed | WIP22SubCom | Brilliance 64 |  | BB_WIP11.1 WIP21.1 | October 23, 2019 7:00:29 PM GMT+08:00
CTUDT00626254 | Refresh | Closed | lei_yang | September 16, 2019 3:34:36 PM GMT+08:00 | The waveform was not imported to Console When the Varian device and the console patient had the same ID. | WIP19 | BB_WIP11.1 WIP21.1 |  | Subsystem |  | NA | Fixed | navale_s | 0 | 2 | Console Software | Console4 | YES | Closed | WIP22SubCom | Brilliance Big Bore |  | BB_WIP11.1 WIP21.1 | October 23, 2019 7:00:29 PM GMT+08:00
CTUDT00626254 | Refresh | Closed | lei_yang | September 16, 2019 3:34:36 PM GMT+08:00 | The waveform was not imported to Console When the Varian device and the console patient had the same ID. | WIP19 | BB_WIP11.1 WIP21.1 |  | Subsystem |  | NA | Fixed | navale_s | 0 | 2 | Console Software | Console4 | YES | Closed | WIP22SubCom | Ingenuity 64 |  | BB_WIP11.1 WIP21.1 | October 23, 2019 7:00:29 PM GMT+08:00
CTUDT00626252 | Refresh | Closed | lei_yang | September 16, 2019 3:14:46 PM GMT+08:00 | While performing CCT scans, the reference image was not showing the result image. | BB_WIP9 | BB_WIP11 |  | Subsystem | CCT | Console-SRS.17727 | Fixed | vishnu_devi | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP11_SubCom | Brilliance Big Bore |  | BB_WIP11 | October 3, 2019 7:33:04 PM GMT+08:00
CTUDT00626247 | Refresh | Closed | wang_l | September 16, 2019 9:27:13 AM GMT+08:00 | Refresh: CT-NM.SRS.Service.3000 and service.2999 are not applicable any more since roll back was removed | WIP19 | DHF298554 Rev08 DHF298555 Rev08 iCT, SRS DHF298553 Rev.08 IQon SRS DHF298556 Rev.09 |  | Subsystem | Patch Management;Verified1 | Service.2999 and Service.3000 | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP21SegReq | Ovation |  | DHF298554 Rev08 DHF298555 Rev08 iCT, SRS DHF298553 Rev.08 IQon SRS DHF298556 Rev.09 | November 6, 2019 2:08:12 PM GMT+08:00
CTUDT00626247 | Refresh | Closed | wang_l | September 16, 2019 9:27:13 AM GMT+08:00 | Refresh: CT-NM.SRS.Service.3000 and service.2999 are not applicable any more since roll back was removed | WIP19 | DHF298554 Rev08 DHF298555 Rev08 iCT, SRS DHF298553 Rev.08 IQon SRS DHF298556 Rev.09 |  | Subsystem | Patch Management;Verified1 | Service.2999 and Service.3000 | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP21SegReq | Brilliance 64 |  | DHF298554 Rev08 DHF298555 Rev08 iCT, SRS DHF298553 Rev.08 IQon SRS DHF298556 Rev.09 | November 6, 2019 2:08:12 PM GMT+08:00
CTUDT00626247 | Refresh | Closed | wang_l | September 16, 2019 9:27:13 AM GMT+08:00 | Refresh: CT-NM.SRS.Service.3000 and service.2999 are not applicable any more since roll back was removed | WIP19 | DHF298554 Rev08 DHF298555 Rev08 iCT, SRS DHF298553 Rev.08 IQon SRS DHF298556 Rev.09 |  | Subsystem | Patch Management;Verified1 | Service.2999 and Service.3000 | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP21SegReq | Brilliance Big Bore |  | DHF298554 Rev08 DHF298555 Rev08 iCT, SRS DHF298553 Rev.08 IQon SRS DHF298556 Rev.09 | November 6, 2019 2:08:12 PM GMT+08:00
CTUDT00626247 | Refresh | Closed | wang_l | September 16, 2019 9:27:13 AM GMT+08:00 | Refresh: CT-NM.SRS.Service.3000 and service.2999 are not applicable any more since roll back was removed | WIP19 | DHF298554 Rev08 DHF298555 Rev08 iCT, SRS DHF298553 Rev.08 IQon SRS DHF298556 Rev.09 |  | Subsystem | Patch Management;Verified1 | Service.2999 and Service.3000 | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP21SegReq | Ingenuity 64 |  | DHF298554 Rev08 DHF298555 Rev08 iCT, SRS DHF298553 Rev.08 IQon SRS DHF298556 Rev.09 | November 6, 2019 2:08:12 PM GMT+08:00
CTUDT00626247 | Refresh | Closed | wang_l | September 16, 2019 9:27:13 AM GMT+08:00 | Refresh: CT-NM.SRS.Service.3000 and service.2999 are not applicable any more since roll back was removed | WIP19 | DHF298554 Rev08 DHF298555 Rev08 iCT, SRS DHF298553 Rev.08 IQon SRS DHF298556 Rev.09 |  | Subsystem | Patch Management;Verified1 | Service.2999 and Service.3000 | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP21SegReq | IQon1 |  | DHF298554 Rev08 DHF298555 Rev08 iCT, SRS DHF298553 Rev.08 IQon SRS DHF298556 Rev.09 | November 6, 2019 2:08:12 PM GMT+08:00
CTUDT00626245 | Refresh | Closed | 320053769 | September 15, 2019 10:06:14 PM GMT+08:00 | WIP19 - Snagit editor has stopped Working | WIP19 | WIP24 | This issue does not occur for all the users. One can login with different user and use the snagit. If it is occurring for clinical user i.e. CT, one can create new Clinical user (CT1) using user management tool and use CT1 user. With New CT1 user Snagit will work as expected. To resolve the Crash, Snagit tool can be uninstall from the impacted system. There are few points needs to be consider before uninstalling Snagit. 1.FSE needs to make sure that default printer is not set to Snagit. 2.Run Snagit button will be disabled in Preferences Save Image screen. 3.PrintSc key will not be useful, as it is configured to launch with Snagit. | Subsystem | SnagIt, PM_Approved, Blocker | NA | Fixed | kapadia_v | 0 | 4 | Console Software | Console4 | NO | Closed | WIP24SubCom | Ovation |  | WIP24 | January 29, 2020 8:17:10 PM GMT+08:00
CTUDT00626245 | Refresh | Closed | 320053769 | September 15, 2019 10:06:14 PM GMT+08:00 | WIP19 - Snagit editor has stopped Working | WIP19 | WIP24 | This issue does not occur for all the users. One can login with different user and use the snagit. If it is occurring for clinical user i.e. CT, one can create new Clinical user (CT1) using user management tool and use CT1 user. With New CT1 user Snagit will work as expected. To resolve the Crash, Snagit tool can be uninstall from the impacted system. There are few points needs to be consider before uninstalling Snagit. 1.FSE needs to make sure that default printer is not set to Snagit. 2.Run Snagit button will be disabled in Preferences Save Image screen. 3.PrintSc key will not be useful, as it is configured to launch with Snagit. | Subsystem | SnagIt, PM_Approved, Blocker | NA | Fixed | kapadia_v | 0 | 4 | Console Software | Console4 | NO | Closed | WIP24SubCom | IQon1 |  | WIP24 | January 29, 2020 8:17:10 PM GMT+08:00
CTUDT00626242 | Refresh | Closed | bommiReddy_r | September 13, 2019 9:41:55 PM GMT+08:00 | Refresh: Localization: "User accounts restore was successful.<OK>" was not localized as per the selected language | WIP19 | WIP21 BB_WIP11 |  | Subsystem |  | NA | Fixed | vishnu_devi | 0 | 4 | Console Software | Console4 | NO | Closed | WIP21SubCom | Ovation |  | WIP21 BB_WIP11 | September 19, 2019 5:36:59 PM GMT+08:00
CTUDT00626242 | Refresh | Closed | bommiReddy_r | September 13, 2019 9:41:55 PM GMT+08:00 | Refresh: Localization: "User accounts restore was successful.<OK>" was not localized as per the selected language | WIP19 | WIP21 BB_WIP11 |  | Subsystem |  | NA | Fixed | vishnu_devi | 0 | 4 | Console Software | Console4 | NO | Closed | WIP21SubCom | Brilliance 64 |  | WIP21 BB_WIP11 | September 19, 2019 5:36:59 PM GMT+08:00
CTUDT00626242 | Refresh | Closed | bommiReddy_r | September 13, 2019 9:41:55 PM GMT+08:00 | Refresh: Localization: "User accounts restore was successful.<OK>" was not localized as per the selected language | WIP19 | WIP21 BB_WIP11 |  | Subsystem |  | NA | Fixed | vishnu_devi | 0 | 4 | Console Software | Console4 | NO | Closed | WIP21SubCom | Brilliance Big Bore |  | WIP21 BB_WIP11 | September 19, 2019 5:36:59 PM GMT+08:00
CTUDT00626242 | Refresh | Closed | bommiReddy_r | September 13, 2019 9:41:55 PM GMT+08:00 | Refresh: Localization: "User accounts restore was successful.<OK>" was not localized as per the selected language | WIP19 | WIP21 BB_WIP11 |  | Subsystem |  | NA | Fixed | vishnu_devi | 0 | 4 | Console Software | Console4 | NO | Closed | WIP21SubCom | Ingenuity 64 |  | WIP21 BB_WIP11 | September 19, 2019 5:36:59 PM GMT+08:00
CTUDT00626242 | Refresh | Closed | bommiReddy_r | September 13, 2019 9:41:55 PM GMT+08:00 | Refresh: Localization: "User accounts restore was successful.<OK>" was not localized as per the selected language | WIP19 | WIP21 BB_WIP11 |  | Subsystem |  | NA | Fixed | vishnu_devi | 0 | 4 | Console Software | Console4 | NO | Closed | WIP21SubCom | IQon1 |  | WIP21 BB_WIP11 | September 19, 2019 5:36:59 PM GMT+08:00
CTUDT00626240 | Refresh | Closed | rojin_antony | September 13, 2019 5:37:44 PM GMT+08:00 | Pre-Val: Windows Tab with running applications icons are visible on the left monitor | BB_WIP8 | BB_WIP11 |  | Subsystem | Accessibility | N/A | Fixed | vishnu_devi | 0 | 2 | Console Software | Support-Security | NO | Closed | BB_WIP11_SubCom | Brilliance Big Bore |  | BB_WIP11 | October 12, 2019 7:48:37 PM GMT+08:00
CTUDT00626238 | Refresh | Closed | rojin_antony | September 13, 2019 5:28:04 PM GMT+08:00 | Pre-Val: System crashed with BSOD while performing a SnagIt Video recording | BB_WIP9 |  | NA | Subsystem |  | NA | Works As Specified / Not A Defect | rojin_antony | 1 |  | Console Software | Support-System | NO | Closed | TBR | Brilliance Big Bore
CTUDT00626236 | Refresh | Closed | rojin_antony | September 13, 2019 5:19:25 PM GMT+08:00 | Pre-Val: IVC hardware Joystick mouse sensitivity is low and difficult to control the mouse movements . Right click context menu is displayed only intermittent with joystick right button. | BB_WIP9 |  |  | Subsystem | IVC;CRN | N/A | No Plan to Fix | pandey_a | 0 | 4 |  | Console4 | NO | Closed | Not a Defect | Brilliance Big Bore
CTUDT00626234 | Refresh | Closed | rojin_antony | September 13, 2019 5:01:11 PM GMT+08:00 | Pre-Val: A Error pop up displayed during online reconstruction "Not all Images were created Please use the recon option to create the missing images". | BB_WIP8 | WIP21 |  | Subsystem |  | NA | Fixed | sagar_v | 0 | 4 | Console Software | Console4 | NO | Closed | WIP21SubCom | Ovation |  | WIP21 | October 10, 2019 1:38:44 PM GMT+08:00
CTUDT00626234 | Refresh | Closed | rojin_antony | September 13, 2019 5:01:11 PM GMT+08:00 | Pre-Val: A Error pop up displayed during online reconstruction "Not all Images were created Please use the recon option to create the missing images". | BB_WIP8 | WIP21 |  | Subsystem |  | NA | Fixed | sagar_v | 0 | 4 | Console Software | Console4 | NO | Closed | WIP21SubCom | Brilliance 64 |  | WIP21 | October 10, 2019 1:38:44 PM GMT+08:00
CTUDT00626234 | Refresh | Closed | rojin_antony | September 13, 2019 5:01:11 PM GMT+08:00 | Pre-Val: A Error pop up displayed during online reconstruction "Not all Images were created Please use the recon option to create the missing images". | BB_WIP8 | WIP21 |  | Subsystem |  | NA | Fixed | sagar_v | 0 | 4 | Console Software | Console4 | NO | Closed | WIP21SubCom | Ingenuity 64 |  | WIP21 | October 10, 2019 1:38:44 PM GMT+08:00
CTUDT00626234 | Refresh | Closed | rojin_antony | September 13, 2019 5:01:11 PM GMT+08:00 | Pre-Val: A Error pop up displayed during online reconstruction "Not all Images were created Please use the recon option to create the missing images". | BB_WIP8 | WIP21 |  | Subsystem |  | NA | Fixed | sagar_v | 0 | 4 | Console Software | Console4 | NO | Closed | WIP21SubCom | Access 16 |  | WIP21 | October 10, 2019 1:38:44 PM GMT+08:00
CTUDT00626234 | Refresh | Closed | rojin_antony | September 13, 2019 5:01:11 PM GMT+08:00 | Pre-Val: A Error pop up displayed during online reconstruction "Not all Images were created Please use the recon option to create the missing images". | BB_WIP8 | WIP21 |  | Subsystem |  | NA | Fixed | sagar_v | 0 | 4 | Console Software | Console4 | NO | Closed | WIP21SubCom | IQon1 |  | WIP21 | October 10, 2019 1:38:44 PM GMT+08:00
CTUDT00626231 | Refresh | Closed | rojin_antony | September 13, 2019 4:52:04 PM GMT+08:00 | Pre-Val: Emergency Login is having Olevel access in Quality Assurance Tab | BB_WIP8 |  |  | Subsystem | O-level access;Blocker | N/A | Change Proposal | bandil_j | 0 |  |  | Support-System | NO | Closed |  | Ovation
CTUDT00626231 | Refresh | Closed | rojin_antony | September 13, 2019 4:52:04 PM GMT+08:00 | Pre-Val: Emergency Login is having Olevel access in Quality Assurance Tab | BB_WIP8 |  |  | Subsystem | O-level access;Blocker | N/A | Change Proposal | bandil_j | 0 |  |  | Support-System | NO | Closed |  | Brilliance 64
CTUDT00626231 | Refresh | Closed | rojin_antony | September 13, 2019 4:52:04 PM GMT+08:00 | Pre-Val: Emergency Login is having Olevel access in Quality Assurance Tab | BB_WIP8 |  |  | Subsystem | O-level access;Blocker | N/A | Change Proposal | bandil_j | 0 |  |  | Support-System | NO | Closed |  | Brilliance Big Bore
CTUDT00626231 | Refresh | Closed | rojin_antony | September 13, 2019 4:52:04 PM GMT+08:00 | Pre-Val: Emergency Login is having Olevel access in Quality Assurance Tab | BB_WIP8 |  |  | Subsystem | O-level access;Blocker | N/A | Change Proposal | bandil_j | 0 |  |  | Support-System | NO | Closed |  | Ingenuity 64
CTUDT00626231 | Refresh | Closed | rojin_antony | September 13, 2019 4:52:04 PM GMT+08:00 | Pre-Val: Emergency Login is having Olevel access in Quality Assurance Tab | BB_WIP8 |  |  | Subsystem | O-level access;Blocker | N/A | Change Proposal | bandil_j | 0 |  |  | Support-System | NO | Closed |  | IQon1
CTUDT00626230 | Refresh | Closed | Jose_D_Sergio_Almeida | September 13, 2019 4:51:06 PM GMT+08:00 | Pre-Val: Message not shown when planned beyond surview. | WIP19 |  |  | Subsystem |  | N/A | Works As Specified / Not A Defect | pandey_a | 0 |  |  | Console4 | NO | Closed | WAS | Ingenuity 64
CTUDT00626228 | Refresh | Closed | bhawna_g | September 12, 2019 9:40:00 PM GMT+08:00 | BigBore Refresh: Planned Average Scan size is getting updated after the clinical scan for Perfusion Jog exam cards. | 4.1.10 and 4.7.7 | WIP21 |  | Subsystem |  | N/A | Fixed | pandey_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP21SubCom | Ovation |  | WIP21 | September 18, 2019 8:23:02 PM GMT+08:00
CTUDT00626228 | Refresh | Closed | bhawna_g | September 12, 2019 9:40:00 PM GMT+08:00 | BigBore Refresh: Planned Average Scan size is getting updated after the clinical scan for Perfusion Jog exam cards. | 4.1.10 and 4.7.7 | WIP21 |  | Subsystem |  | N/A | Fixed | pandey_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP21SubCom | Brilliance 64 |  | WIP21 | September 18, 2019 8:23:02 PM GMT+08:00
CTUDT00626228 | Refresh | Closed | bhawna_g | September 12, 2019 9:40:00 PM GMT+08:00 | BigBore Refresh: Planned Average Scan size is getting updated after the clinical scan for Perfusion Jog exam cards. | 4.1.10 and 4.7.7 | WIP21 |  | Subsystem |  | N/A | Fixed | pandey_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP21SubCom | Ingenuity 64 |  | WIP21 | September 18, 2019 8:23:02 PM GMT+08:00
CTUDT00626228 | Refresh | Closed | bhawna_g | September 12, 2019 9:40:00 PM GMT+08:00 | BigBore Refresh: Planned Average Scan size is getting updated after the clinical scan for Perfusion Jog exam cards. | 4.1.10 and 4.7.7 | WIP21 |  | Subsystem |  | N/A | Fixed | pandey_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP21SubCom | IQon1 |  | WIP21 | September 18, 2019 8:23:02 PM GMT+08:00
CTUDT00626227 | Refresh | Closed | altshuller_mark | September 12, 2019 6:08:39 PM GMT+08:00 | Pre-Val Refresh: Emergency user access is disabled by default | WIP19 | 459801587331 Rev C 459801587341 Rev C 459801587361 Rev C 459801587351 Rev C | NA. | System |  | N/A | Fixed | li_e | 0 | 2 | Console Software | Support-Service | NO | Closed | WIP21SegReq | Brilliance 64 |  | 459801587331 Rev C 459801587341 Rev C 459801587361 Rev C 459801587351 Rev C | November 19, 2019 5:53:29 PM GMT+08:00
CTUDT00626227 | Refresh | Closed | altshuller_mark | September 12, 2019 6:08:39 PM GMT+08:00 | Pre-Val Refresh: Emergency user access is disabled by default | WIP19 | 459801587331 Rev C 459801587341 Rev C 459801587361 Rev C 459801587351 Rev C | NA. | System |  | N/A | Fixed | li_e | 0 | 2 | Console Software | Support-Service | NO | Closed | WIP21SegReq | Brilliance Big Bore |  | 459801587331 Rev C 459801587341 Rev C 459801587361 Rev C 459801587351 Rev C | November 19, 2019 5:53:29 PM GMT+08:00
CTUDT00626227 | Refresh | Closed | altshuller_mark | September 12, 2019 6:08:39 PM GMT+08:00 | Pre-Val Refresh: Emergency user access is disabled by default | WIP19 | 459801587331 Rev C 459801587341 Rev C 459801587361 Rev C 459801587351 Rev C | NA. | System |  | N/A | Fixed | li_e | 0 | 2 | Console Software | Support-Service | NO | Closed | WIP21SegReq | Ingenuity 64 |  | 459801587331 Rev C 459801587341 Rev C 459801587361 Rev C 459801587351 Rev C | November 19, 2019 5:53:29 PM GMT+08:00
CTUDT00626224 | Refresh | Closed | 320053769 | September 11, 2019 10:51:19 PM GMT+08:00 | WIP19 - System Error warning appeared after performing CCT scan and returning to patient scene . | WIP19 |  | Please restart the system | Subsystem |  | NA | Works As Specified / Not A Defect | tamari_o | 1 |  | Console Software | Support-System | NO | Closed |  | Ovation
CTUDT00626224 | Refresh | Closed | 320053769 | September 11, 2019 10:51:19 PM GMT+08:00 | WIP19 - System Error warning appeared after performing CCT scan and returning to patient scene . | WIP19 |  | Please restart the system | Subsystem |  | NA | Works As Specified / Not A Defect | tamari_o | 1 |  | Console Software | Support-System | NO | Closed |  | Brilliance 64
CTUDT00626224 | Refresh | Closed | 320053769 | September 11, 2019 10:51:19 PM GMT+08:00 | WIP19 - System Error warning appeared after performing CCT scan and returning to patient scene . | WIP19 |  | Please restart the system | Subsystem |  | NA | Works As Specified / Not A Defect | tamari_o | 1 |  | Console Software | Support-System | NO | Closed |  | Brilliance Big Bore
CTUDT00626224 | Refresh | Closed | 320053769 | September 11, 2019 10:51:19 PM GMT+08:00 | WIP19 - System Error warning appeared after performing CCT scan and returning to patient scene . | WIP19 |  | Please restart the system | Subsystem |  | NA | Works As Specified / Not A Defect | tamari_o | 1 |  | Console Software | Support-System | NO | Closed |  | Ingenuity 64
CTUDT00626224 | Refresh | Closed | 320053769 | September 11, 2019 10:51:19 PM GMT+08:00 | WIP19 - System Error warning appeared after performing CCT scan and returning to patient scene . | WIP19 |  | Please restart the system | Subsystem |  | NA | Works As Specified / Not A Defect | tamari_o | 1 |  | Console Software | Support-System | NO | Closed |  | IQon1
CTUDT00626221 | Refresh | Closed | Bat-El_Asaraf | September 11, 2019 9:26:15 PM GMT+08:00 | Pre-Val Refresh: Help page is incorrect for IQ check directions (Harmonized Phantom) | WIP19 |  |  | Subsystem |  | n/a | Works As Specified / Not A Defect | yatziv_l | 0 | 3 |  | DMS | NO | Closed | WIP20SegReq | Ovation
CTUDT00626221 | Refresh | Closed | Bat-El_Asaraf | September 11, 2019 9:26:15 PM GMT+08:00 | Pre-Val Refresh: Help page is incorrect for IQ check directions (Harmonized Phantom) | WIP19 |  |  | Subsystem |  | n/a | Works As Specified / Not A Defect | yatziv_l | 0 | 3 |  | DMS | NO | Closed | WIP20SegReq | Brilliance 64
CTUDT00626221 | Refresh | Closed | Bat-El_Asaraf | September 11, 2019 9:26:15 PM GMT+08:00 | Pre-Val Refresh: Help page is incorrect for IQ check directions (Harmonized Phantom) | WIP19 |  |  | Subsystem |  | n/a | Works As Specified / Not A Defect | yatziv_l | 0 | 3 |  | DMS | NO | Closed | WIP20SegReq | IQon1
CTUDT00626220 | Refresh | Closed | yatziv_l | September 11, 2019 9:19:19 PM GMT+08:00 | Refresh: CT-NM.SRS.SysSupport.420 should define the used Norwegian tongue which is Bokmal | N/A |  |  | Subsystem |  | CT-NM.SRS.SysSupport.420 | No Plan to Fix | bandil_j | 0 | 3 |  | Support-System | NO | Closed |  | Ovation
CTUDT00626220 | Refresh | Closed | yatziv_l | September 11, 2019 9:19:19 PM GMT+08:00 | Refresh: CT-NM.SRS.SysSupport.420 should define the used Norwegian tongue which is Bokmal | N/A |  |  | Subsystem |  | CT-NM.SRS.SysSupport.420 | No Plan to Fix | bandil_j | 0 | 3 |  | Support-System | NO | Closed |  | Brilliance 64
CTUDT00626220 | Refresh | Closed | yatziv_l | September 11, 2019 9:19:19 PM GMT+08:00 | Refresh: CT-NM.SRS.SysSupport.420 should define the used Norwegian tongue which is Bokmal | N/A |  |  | Subsystem |  | CT-NM.SRS.SysSupport.420 | No Plan to Fix | bandil_j | 0 | 3 |  | Support-System | NO | Closed |  | Brilliance Big Bore
CTUDT00626220 | Refresh | Closed | yatziv_l | September 11, 2019 9:19:19 PM GMT+08:00 | Refresh: CT-NM.SRS.SysSupport.420 should define the used Norwegian tongue which is Bokmal | N/A |  |  | Subsystem |  | CT-NM.SRS.SysSupport.420 | No Plan to Fix | bandil_j | 0 | 3 |  | Support-System | NO | Closed |  | Ingenuity 64
CTUDT00626220 | Refresh | Closed | yatziv_l | September 11, 2019 9:19:19 PM GMT+08:00 | Refresh: CT-NM.SRS.SysSupport.420 should define the used Norwegian tongue which is Bokmal | N/A |  |  | Subsystem |  | CT-NM.SRS.SysSupport.420 | No Plan to Fix | bandil_j | 0 | 3 |  | Support-System | NO | Closed |  | IQon1
CTUDT00626218 | Refresh | Closed | Bat-El_Asaraf | September 11, 2019 8:32:03 PM GMT+08:00 | Pre-Val Refresh: Not all clinical images created | WIP19 |  |  | Subsystem |  | N/A | Does Not Apply | tamari_o | 1 |  |  | Support-System | NO | Closed | WAS | Ovation
CTUDT00626218 | Refresh | Closed | Bat-El_Asaraf | September 11, 2019 8:32:03 PM GMT+08:00 | Pre-Val Refresh: Not all clinical images created | WIP19 |  |  | Subsystem |  | N/A | Does Not Apply | tamari_o | 1 |  |  | Support-System | NO | Closed | WAS | Brilliance 64
CTUDT00626218 | Refresh | Closed | Bat-El_Asaraf | September 11, 2019 8:32:03 PM GMT+08:00 | Pre-Val Refresh: Not all clinical images created | WIP19 |  |  | Subsystem |  | N/A | Does Not Apply | tamari_o | 1 |  |  | Support-System | NO | Closed | WAS | Brilliance Big Bore
CTUDT00626218 | Refresh | Closed | Bat-El_Asaraf | September 11, 2019 8:32:03 PM GMT+08:00 | Pre-Val Refresh: Not all clinical images created | WIP19 |  |  | Subsystem |  | N/A | Does Not Apply | tamari_o | 1 |  |  | Support-System | NO | Closed | WAS | Ingenuity 64
CTUDT00626218 | Refresh | Closed | Bat-El_Asaraf | September 11, 2019 8:32:03 PM GMT+08:00 | Pre-Val Refresh: Not all clinical images created | WIP19 |  |  | Subsystem |  | N/A | Does Not Apply | tamari_o | 1 |  |  | Support-System | NO | Closed | WAS | IQon1
CTUDT00626213 | Refresh | Closed | Noy_Seelig | September 11, 2019 7:08:54 PM GMT+08:00 | Refresh: Pre WIP 10 calibrations: A Plane Opening fails due to wrong mapping of colliomation | BB_WIP9 | WIP20 |  | Subsystem | Blocker20;Calibration | NA | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP20SubCom | Brilliance Big Bore |  | WIP20 | September 18, 2019 2:44:36 PM GMT+08:00
CTUDT00626210 | Refresh | Closed | wang_l | September 11, 2019 4:41:05 PM GMT+08:00 | Refresh Ingenuity: Voice Restore stopped working during Upgrade from 4.1.7 to 4.1.10 | WIP19 | WIP21 |  | Subsystem |  | NA | Fixed | chakraborty_d | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP21SubCom | Ovation |  | WIP21 | September 18, 2019 5:53:56 PM GMT+08:00
CTUDT00626210 | Refresh | Closed | wang_l | September 11, 2019 4:41:05 PM GMT+08:00 | Refresh Ingenuity: Voice Restore stopped working during Upgrade from 4.1.7 to 4.1.10 | WIP19 | WIP21 |  | Subsystem |  | NA | Fixed | chakraborty_d | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP21SubCom | Brilliance 64 |  | WIP21 | September 18, 2019 5:53:56 PM GMT+08:00
CTUDT00626210 | Refresh | Closed | wang_l | September 11, 2019 4:41:05 PM GMT+08:00 | Refresh Ingenuity: Voice Restore stopped working during Upgrade from 4.1.7 to 4.1.10 | WIP19 | WIP21 |  | Subsystem |  | NA | Fixed | chakraborty_d | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP21SubCom | Ingenuity 64 |  | WIP21 | September 18, 2019 5:53:56 PM GMT+08:00
CTUDT00626210 | Refresh | Closed | wang_l | September 11, 2019 4:41:05 PM GMT+08:00 | Refresh Ingenuity: Voice Restore stopped working during Upgrade from 4.1.7 to 4.1.10 | WIP19 | WIP21 |  | Subsystem |  | NA | Fixed | chakraborty_d | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP21SubCom | IQon1 |  | WIP21 | September 18, 2019 5:53:56 PM GMT+08:00
CTUDT00626208 | Refresh | Closed | ding_s | September 11, 2019 4:32:01 PM GMT+08:00 | Scanning is disabled when free space of D drive is less than 19.5Gb but more than 15Gb on Host. | WIP19 | ING, BR64 SRS DHF298554 Rev08 iCT, SRS DHF298553 Rev.08 |  | Subsystem | Verified1 | N/A | Fixed | guy_sheaffer | 0 | 2 | Console Software | Support-System | NO | Closed | WIP21SegReq | Ovation |  | ING, BR64 SRS DHF298554 Rev08 iCT, SRS DHF298553 Rev.08 | November 6, 2019 2:07:01 PM GMT+08:00
CTUDT00626208 | Refresh | Closed | ding_s | September 11, 2019 4:32:01 PM GMT+08:00 | Scanning is disabled when free space of D drive is less than 19.5Gb but more than 15Gb on Host. | WIP19 | ING, BR64 SRS DHF298554 Rev08 iCT, SRS DHF298553 Rev.08 |  | Subsystem | Verified1 | N/A | Fixed | guy_sheaffer | 0 | 2 | Console Software | Support-System | NO | Closed | WIP21SegReq | Brilliance 64 |  | ING, BR64 SRS DHF298554 Rev08 iCT, SRS DHF298553 Rev.08 | November 6, 2019 2:07:01 PM GMT+08:00
CTUDT00626208 | Refresh | Closed | ding_s | September 11, 2019 4:32:01 PM GMT+08:00 | Scanning is disabled when free space of D drive is less than 19.5Gb but more than 15Gb on Host. | WIP19 | ING, BR64 SRS DHF298554 Rev08 iCT, SRS DHF298553 Rev.08 |  | Subsystem | Verified1 | N/A | Fixed | guy_sheaffer | 0 | 2 | Console Software | Support-System | NO | Closed | WIP21SegReq | Ingenuity 64 |  | ING, BR64 SRS DHF298554 Rev08 iCT, SRS DHF298553 Rev.08 | November 6, 2019 2:07:01 PM GMT+08:00
CTUDT00626207 | Refresh | Closed | lei_yang | September 11, 2019 4:05:04 PM GMT+08:00 | Refresh:Pulmo wave displays abnormal while starting IBoxInjector and inputting pulmonary waveform by bellow. | BB_WIP9 |  | NA | Subsystem | Pulmo wave display | NA | Invalid/Uncontrolled Configuration | lei_yang | 0 |  |  | Support-System | NO | Closed | WAS | Brilliance Big Bore
CTUDT00626201 | Refresh | Closed | zhu_je | September 11, 2019 11:41:58 AM GMT+08:00 | [Refresh] [IVC] When chose Blue form System Diagnostics and then click return button , the light state of Power on /off led was still abnormal(Green) | BB_WIP9 | BB_WIP11.1 | When this issue happen, do the following 1. Go to diagnostic page again using the option in Hamburger menu 2. Click the 'White' button and the power button will become 'white' now. 3. Exit the diagnostics page using the '<' button on top left | Subsystem | IVC | NA | Fixed | Snehashis_Dey | 0 | 2 | IVC Software | Console4 | NO | Closed | BB_WIP11_SubCom | Brilliance Big Bore |  | BB_WIP11.1 | October 31, 2019 8:53:10 PM GMT+08:00
CTUDT00626199 | Refresh | Closed | zhu_je | September 11, 2019 11:21:34 AM GMT+08:00 | [Refresh] [IVC]HW Control Box did not display the 'Connecting with Console' message when the timer was counting down after 30 sec if end workflow during clean | BB_WIP9 | BB_WIP10 |  | Subsystem | IVC | NA | Fixed | Snehashis_Dey | 0 | 2 | IVC Software | Console4 | NO | Closed | BB_WIP11_SubCom | Brilliance Big Bore |  | BB_WIP10 | September 18, 2019 5:21:26 PM GMT+08:00
CTUDT00626197 | Refresh | Closed | ding_s | September 11, 2019 10:22:08 AM GMT+08:00 | There was no instruction to enable upgrade to the latest released configuration for installed base customers with released versions older than 4.1.7 | WIP19 | DHF326014_Rev05 |  | Subsystem | Upgrade | NA | Fixed | ding_f | 0 | 2 | Test Design | Support-System | NO | Closed | WIP21SegReq | Brilliance 64 |  | DHF326014_Rev05 | November 4, 2019 4:30:55 PM GMT+08:00
CTUDT00626197 | Refresh | Closed | ding_s | September 11, 2019 10:22:08 AM GMT+08:00 | There was no instruction to enable upgrade to the latest released configuration for installed base customers with released versions older than 4.1.7 | WIP19 | DHF326014_Rev05 |  | Subsystem | Upgrade | NA | Fixed | ding_f | 0 | 2 | Test Design | Support-System | NO | Closed | WIP21SegReq | Brilliance Big Bore |  | DHF326014_Rev05 | November 4, 2019 4:30:55 PM GMT+08:00
CTUDT00626197 | Refresh | Closed | ding_s | September 11, 2019 10:22:08 AM GMT+08:00 | There was no instruction to enable upgrade to the latest released configuration for installed base customers with released versions older than 4.1.7 | WIP19 | DHF326014_Rev05 |  | Subsystem | Upgrade | NA | Fixed | ding_f | 0 | 2 | Test Design | Support-System | NO | Closed | WIP21SegReq | Ingenuity 64 |  | DHF326014_Rev05 | November 4, 2019 4:30:55 PM GMT+08:00
CTUDT00626193 | Refresh | Closed | madankumar_s | September 10, 2019 8:48:01 PM GMT+08:00 | In PSC Tool Color Scheme Selection Tool is wrongly displayed under System in Configurations tab. | WIP19 | WIP20 |  | Subsystem | System Configuration | NA | Fixed | Sandeep_Kumar | 0 | 3 | Service Tools_Software | Service Tools | NO | Closed | WIP21SubCom | Ovation |  | WIP20 | September 13, 2019 7:27:04 PM GMT+08:00
CTUDT00626191 | Refresh | Closed | 320053769 | September 10, 2019 8:39:23 PM GMT+08:00 | Report approval fails when a report has more than 230 images. | WIP19 |  | NA | Subsystem | Patient Directory | NA | No Plan to Fix | s_panigrahi | 0 | 3 |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00626191 | Refresh | Closed | 320053769 | September 10, 2019 8:39:23 PM GMT+08:00 | Report approval fails when a report has more than 230 images. | WIP19 |  | NA | Subsystem | Patient Directory | NA | No Plan to Fix | s_panigrahi | 0 | 3 |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00626191 | Refresh | Closed | 320053769 | September 10, 2019 8:39:23 PM GMT+08:00 | Report approval fails when a report has more than 230 images. | WIP19 |  | NA | Subsystem | Patient Directory | NA | No Plan to Fix | s_panigrahi | 0 | 3 |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00626191 | Refresh | Closed | 320053769 | September 10, 2019 8:39:23 PM GMT+08:00 | Report approval fails when a report has more than 230 images. | WIP19 |  | NA | Subsystem | Patient Directory | NA | No Plan to Fix | s_panigrahi | 0 | 3 |  | iCAP | NO | Closed |  | IQon1
CTUDT00626185 | Refresh | Closed | shiraguppi_k | September 10, 2019 5:49:19 PM GMT+08:00 | Pre-Val: Bolus Tracker Image WW & WC is different than default selected WW & WC. | WIP19 | WIP21 | Change WW & WC values using ECM or during planning. | Subsystem | Bolus Tracking | N/A | Fixed | pandey_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP21SegReq | Brilliance 64 |  | WIP21 | October 14, 2019 9:54:45 PM GMT+08:00
CTUDT00626185 | Refresh | Closed | shiraguppi_k | September 10, 2019 5:49:19 PM GMT+08:00 | Pre-Val: Bolus Tracker Image WW & WC is different than default selected WW & WC. | WIP19 | WIP21 | Change WW & WC values using ECM or during planning. | Subsystem | Bolus Tracking | N/A | Fixed | pandey_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP21SegReq | Ingenuity 64 |  | WIP21 | October 14, 2019 9:54:45 PM GMT+08:00
CTUDT00626183 | Refresh | Closed | shiraguppi_k | September 10, 2019 5:40:34 PM GMT+08:00 | Pre-Val: While saving EC in ECM, Exam Card name is showing previously selected EC Name. | WIP19 |  |  | Subsystem | ECM | NA | Works As Specified / Not A Defect | challa_s | 0 |  |  | Console4 | NO | Closed | WAS | Ingenuity 64
CTUDT00626181 | Refresh | Closed | g_di | September 10, 2019 5:02:00 PM GMT+08:00 | "Script Error” message pops-up when there is a change in the “From date” in Time Window section of Universal Log Viewer | WIP19 |  |  | Subsystem | Log Viewer | NA | No Plan to Fix | s_shabeer | 0 | 4 |  | Service Tools | NO | Closed | SubNC | Ovation
CTUDT00626181 | Refresh | Closed | g_di | September 10, 2019 5:02:00 PM GMT+08:00 | "Script Error” message pops-up when there is a change in the “From date” in Time Window section of Universal Log Viewer | WIP19 |  |  | Subsystem | Log Viewer | NA | No Plan to Fix | s_shabeer | 0 | 4 |  | Service Tools | NO | Closed | SubNC | Brilliance Big Bore
CTUDT00626181 | Refresh | Closed | g_di | September 10, 2019 5:02:00 PM GMT+08:00 | "Script Error” message pops-up when there is a change in the “From date” in Time Window section of Universal Log Viewer | WIP19 |  |  | Subsystem | Log Viewer | NA | No Plan to Fix | s_shabeer | 0 | 4 |  | Service Tools | NO | Closed | SubNC | Ingenuity 64
CTUDT00626181 | Refresh | Closed | g_di | September 10, 2019 5:02:00 PM GMT+08:00 | "Script Error” message pops-up when there is a change in the “From date” in Time Window section of Universal Log Viewer | WIP19 |  |  | Subsystem | Log Viewer | NA | No Plan to Fix | s_shabeer | 0 | 4 |  | Service Tools | NO | Closed | SubNC | IQon1
CTUDT00626177 | Refresh | Closed | cai_t | September 10, 2019 2:03:27 PM GMT+08:00 | Refresh: Tube current tolerance is 20.6% which is higher than required value 20% for 80kV | BB_WIP9 | wip21 |  | Subsystem | Tube Current;Verified1 | CT-NM.SRS.Operations.251 | Fixed | cai_t | 0 | 2 | DMS Calibrations | Support-System | NO | Closed | WIP21SegReq | Brilliance Big Bore |  | wip21 | November 21, 2019 9:53:42 AM GMT+08:00
CTUDT00626176 | Refresh | Closed | datta_i | September 10, 2019 1:45:04 PM GMT+08:00 | The system and software error is displayed on performing cardiac and non-cardiac scan | 4.8 | BB_WIP10 |  | Subsystem | Scan abort | NA | Fixed | challa_s | 0 | 2 | Console Software | Console4 | YES | Closed | BB_WIP10_SubCom | Brilliance Big Bore |  | BB_WIP10 | September 13, 2019 1:01:07 PM GMT+08:00
CTUDT00626174 | Refresh | Closed | lei_yang | September 10, 2019 8:34:45 AM GMT+08:00 | The Plan scene does not disappear after clicking "End the exam" on the varian no signal to establish pop-up window. | BB_WIP8 |  |  | Subsystem | Pop up message | N/A | Works As Specified / Not A Defect | pandey_a | 0 |  |  | Console4 | NO | Closed | WAS | Brilliance 64
CTUDT00626174 | Refresh | Closed | lei_yang | September 10, 2019 8:34:45 AM GMT+08:00 | The Plan scene does not disappear after clicking "End the exam" on the varian no signal to establish pop-up window. | BB_WIP8 |  |  | Subsystem | Pop up message | N/A | Works As Specified / Not A Defect | pandey_a | 0 |  |  | Console4 | NO | Closed | WAS | Brilliance Big Bore
CTUDT00626174 | Refresh | Closed | lei_yang | September 10, 2019 8:34:45 AM GMT+08:00 | The Plan scene does not disappear after clicking "End the exam" on the varian no signal to establish pop-up window. | BB_WIP8 |  |  | Subsystem | Pop up message | N/A | Works As Specified / Not A Defect | pandey_a | 0 |  |  | Console4 | NO | Closed | WAS | Ingenuity 64
CTUDT00626171 | Refresh | Closed | wei_f | September 9, 2019 7:49:05 PM GMT+08:00 | The system did not notifie the user that password has expired and prompted to change the password. | WIP19 | WIP19 |  | Subsystem | Login |  | Works As Specified / Not A Defect | wei_f | 0 |  | Test Design | Support-System | NO | Closed |  | Brilliance 64 |  | WIP19
CTUDT00626171 | Refresh | Closed | wei_f | September 9, 2019 7:49:05 PM GMT+08:00 | The system did not notifie the user that password has expired and prompted to change the password. | WIP19 | WIP19 |  | Subsystem | Login |  | Works As Specified / Not A Defect | wei_f | 0 |  | Test Design | Support-System | NO | Closed |  | Ingenuity 64 |  | WIP19
CTUDT00626167 | Refresh | Closed | Dana_Kamyshev | September 9, 2019 6:50:45 PM GMT+08:00 | Refresh - Localization issue in SCTV | WIP19 | WIP21 |  | Subsystem | Spectral CT Viewer | NA | Fixed | aswani_v | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP21SegReq | IQon1 |  | WIP21 | September 30, 2019 2:02:21 PM GMT+08:00
CTUDT00626159 | Refresh | Closed | shiraguppi_k | September 9, 2019 5:27:19 PM GMT+08:00 | Pre-Val: All Phase Point Editing option is enabled in ECG Viewer after reconstruction (Both places View & Cardiac Viewer) | WIP19 |  | NA | Subsystem |  | Console-SRS.8387 | No Plan to Fix | sagar_v | 0 | 4 |  | Console4 | NO | Closed | SubNC | Ovation
CTUDT00626159 | Refresh | Closed | shiraguppi_k | September 9, 2019 5:27:19 PM GMT+08:00 | Pre-Val: All Phase Point Editing option is enabled in ECG Viewer after reconstruction (Both places View & Cardiac Viewer) | WIP19 |  | NA | Subsystem |  | Console-SRS.8387 | No Plan to Fix | sagar_v | 0 | 4 |  | Console4 | NO | Closed | SubNC | Brilliance 64
CTUDT00626159 | Refresh | Closed | shiraguppi_k | September 9, 2019 5:27:19 PM GMT+08:00 | Pre-Val: All Phase Point Editing option is enabled in ECG Viewer after reconstruction (Both places View & Cardiac Viewer) | WIP19 |  | NA | Subsystem |  | Console-SRS.8387 | No Plan to Fix | sagar_v | 0 | 4 |  | Console4 | NO | Closed | SubNC | Ingenuity 64
CTUDT00626159 | Refresh | Closed | shiraguppi_k | September 9, 2019 5:27:19 PM GMT+08:00 | Pre-Val: All Phase Point Editing option is enabled in ECG Viewer after reconstruction (Both places View & Cardiac Viewer) | WIP19 |  | NA | Subsystem |  | Console-SRS.8387 | No Plan to Fix | sagar_v | 0 | 4 |  | Console4 | NO | Closed | SubNC | IQon1
CTUDT00626156 | Refresh | Closed | shiraguppi_k | September 9, 2019 5:16:40 PM GMT+08:00 | Pre-Val: System crashed with Bug rep error with system error occurred while performing First time Pulmo Resp Single Phase scan. | WIP19 | WIP21 | Restart the host | Subsystem | Respiratory | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | YES | Closed | BBWIP21_SubCom | Brilliance Big Bore |  | WIP21 | October 9, 2019 3:05:04 PM GMT+08:00
CTUDT00626153 | Refresh | Closed | wei_f | September 9, 2019 4:11:59 PM GMT+08:00 | In patient list switch to DVD-RAM folder, select a study and right click, the delete button is grayed out. | WIP19 | WIP19 |  | Subsystem | CD-DVD operation | CT-NM.SRS.Interop.344 | Fixed | wei_f | 0 | 2 | Test Design | Support-System | NO | Closed | WIP20SubCom | Ovation |  | WIP19 | November 4, 2019 1:50:36 PM GMT+08:00
CTUDT00626153 | Refresh | Closed | wei_f | September 9, 2019 4:11:59 PM GMT+08:00 | In patient list switch to DVD-RAM folder, select a study and right click, the delete button is grayed out. | WIP19 | WIP19 |  | Subsystem | CD-DVD operation | CT-NM.SRS.Interop.344 | Fixed | wei_f | 0 | 2 | Test Design | Support-System | NO | Closed | WIP20SubCom | Brilliance 64 |  | WIP19 | November 4, 2019 1:50:36 PM GMT+08:00
CTUDT00626153 | Refresh | Closed | wei_f | September 9, 2019 4:11:59 PM GMT+08:00 | In patient list switch to DVD-RAM folder, select a study and right click, the delete button is grayed out. | WIP19 | WIP19 |  | Subsystem | CD-DVD operation | CT-NM.SRS.Interop.344 | Fixed | wei_f | 0 | 2 | Test Design | Support-System | NO | Closed | WIP20SubCom | Brilliance Big Bore |  | WIP19 | November 4, 2019 1:50:36 PM GMT+08:00
CTUDT00626153 | Refresh | Closed | wei_f | September 9, 2019 4:11:59 PM GMT+08:00 | In patient list switch to DVD-RAM folder, select a study and right click, the delete button is grayed out. | WIP19 | WIP19 |  | Subsystem | CD-DVD operation | CT-NM.SRS.Interop.344 | Fixed | wei_f | 0 | 2 | Test Design | Support-System | NO | Closed | WIP20SubCom | Ingenuity 64 |  | WIP19 | November 4, 2019 1:50:36 PM GMT+08:00
CTUDT00626153 | Refresh | Closed | wei_f | September 9, 2019 4:11:59 PM GMT+08:00 | In patient list switch to DVD-RAM folder, select a study and right click, the delete button is grayed out. | WIP19 | WIP19 |  | Subsystem | CD-DVD operation | CT-NM.SRS.Interop.344 | Fixed | wei_f | 0 | 2 | Test Design | Support-System | NO | Closed | WIP20SubCom | IQon1 |  | WIP19 | November 4, 2019 1:50:36 PM GMT+08:00
CTUDT00626141 | Refresh | Closed | cai_t | September 7, 2019 12:25:09 PM GMT+08:00 | The height of BB gantry crate excluding the tray cannot fit the ISO type container inside height. | BB_WIP9 | DHF298554 Rev08 DHF298555 Rev08 |  | Subsystem | Gantry dimension | na | Fixed | wang_k | 0 | 2 | SRS | Support-System | NO | Closed | SRS Rev08 | Brilliance 64 |  | DHF298554 Rev08 DHF298555 Rev08 | November 4, 2019 4:44:07 PM GMT+08:00
CTUDT00626141 | Refresh | Closed | cai_t | September 7, 2019 12:25:09 PM GMT+08:00 | The height of BB gantry crate excluding the tray cannot fit the ISO type container inside height. | BB_WIP9 | DHF298554 Rev08 DHF298555 Rev08 |  | Subsystem | Gantry dimension | na | Fixed | wang_k | 0 | 2 | SRS | Support-System | NO | Closed | SRS Rev08 | Brilliance Big Bore |  | DHF298554 Rev08 DHF298555 Rev08 | November 4, 2019 4:44:07 PM GMT+08:00
CTUDT00626141 | Refresh | Closed | cai_t | September 7, 2019 12:25:09 PM GMT+08:00 | The height of BB gantry crate excluding the tray cannot fit the ISO type container inside height. | BB_WIP9 | DHF298554 Rev08 DHF298555 Rev08 |  | Subsystem | Gantry dimension | na | Fixed | wang_k | 0 | 2 | SRS | Support-System | NO | Closed | SRS Rev08 | Ingenuity 64 |  | DHF298554 Rev08 DHF298555 Rev08 | November 4, 2019 4:44:07 PM GMT+08:00
CTUDT00626136 | Refresh | Closed | zhong_x | September 7, 2019 10:57:50 AM GMT+08:00 | Recon Speed was less than 20 IPS with Gen5.1 Console HW FBP recon type | BB_WIP9 | WIP21 | N/A | Subsystem | Recon performance | CT-NM.SRS.SysPerf.698 | Fixed | richards_b | 0 | 2 | CIRS_SW_ImgNodeSvr | CIRS | NO | Closed | WIP21SubCom | Brilliance Big Bore |  | WIP21 | October 21, 2019 10:02:20 PM GMT+08:00
CTUDT00626134 | Refresh | Closed | dang_h | September 6, 2019 11:49:52 PM GMT+08:00 | Bone beam hardening artifacts are not reduced as expected after applying ABC | WIP19 | WIP20 | N/A | Subsystem | IQ, ABC; Blocker20; Verified1 | CT-NM.SRS.IQ.1408 | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP20SubCom | Ovation |  | WIP20 | September 19, 2019 5:38:53 PM GMT+08:00
CTUDT00626134 | Refresh | Closed | dang_h | September 6, 2019 11:49:52 PM GMT+08:00 | Bone beam hardening artifacts are not reduced as expected after applying ABC | WIP19 | WIP20 | N/A | Subsystem | IQ, ABC; Blocker20; Verified1 | CT-NM.SRS.IQ.1408 | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP20SubCom | Ingenuity 64 |  | WIP20 | September 19, 2019 5:38:53 PM GMT+08:00
CTUDT00626133 | Refresh | Closed | datta_i | September 6, 2019 7:31:45 PM GMT+08:00 | Refresh: Localization: Import language pop-up box header if the name is same is getting cropped for French and Spanish | 4.8 | BB_WIP10 |  | Subsystem |  | NA | Fixed | challa_s | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP10_SubCom | Brilliance Big Bore |  | BB_WIP10 | September 11, 2019 3:14:16 PM GMT+08:00
CTUDT00626131 | Refresh | Closed | lakshmikanth_s | September 6, 2019 6:41:08 PM GMT+08:00 | Not able to connect remotely to HOST through ISSL and RSN | WIP19 | WIP21.1 BB_WIP11.1 | N/A | Subsystem | Remote Access Tool | N/A | Fixed | dasgupta_s | 0 | 2 | Service Tools_Software | Support-Security | NO | Closed | WIP21SubCom | Ovation |  | WIP21.1 BB_WIP11.1 | October 1, 2019 3:11:01 PM GMT+08:00
CTUDT00626131 | Refresh | Closed | lakshmikanth_s | September 6, 2019 6:41:08 PM GMT+08:00 | Not able to connect remotely to HOST through ISSL and RSN | WIP19 | WIP21.1 BB_WIP11.1 | N/A | Subsystem | Remote Access Tool | N/A | Fixed | dasgupta_s | 0 | 2 | Service Tools_Software | Support-Security | NO | Closed | WIP21SubCom | Brilliance 64 |  | WIP21.1 BB_WIP11.1 | October 1, 2019 3:11:01 PM GMT+08:00
CTUDT00626131 | Refresh | Closed | lakshmikanth_s | September 6, 2019 6:41:08 PM GMT+08:00 | Not able to connect remotely to HOST through ISSL and RSN | WIP19 | WIP21.1 BB_WIP11.1 | N/A | Subsystem | Remote Access Tool | N/A | Fixed | dasgupta_s | 0 | 2 | Service Tools_Software | Support-Security | NO | Closed | WIP21SubCom | Brilliance Big Bore |  | WIP21.1 BB_WIP11.1 | October 1, 2019 3:11:01 PM GMT+08:00
CTUDT00626131 | Refresh | Closed | lakshmikanth_s | September 6, 2019 6:41:08 PM GMT+08:00 | Not able to connect remotely to HOST through ISSL and RSN | WIP19 | WIP21.1 BB_WIP11.1 | N/A | Subsystem | Remote Access Tool | N/A | Fixed | dasgupta_s | 0 | 2 | Service Tools_Software | Support-Security | NO | Closed | WIP21SubCom | Ingenuity 64 |  | WIP21.1 BB_WIP11.1 | October 1, 2019 3:11:01 PM GMT+08:00
CTUDT00626131 | Refresh | Closed | lakshmikanth_s | September 6, 2019 6:41:08 PM GMT+08:00 | Not able to connect remotely to HOST through ISSL and RSN | WIP19 | WIP21.1 BB_WIP11.1 | N/A | Subsystem | Remote Access Tool | N/A | Fixed | dasgupta_s | 0 | 2 | Service Tools_Software | Support-Security | NO | Closed | WIP21SubCom | IQon1 |  | WIP21.1 BB_WIP11.1 | October 1, 2019 3:11:01 PM GMT+08:00
CTUDT00626128 | Refresh | Closed | madankumar_s | September 6, 2019 5:34:10 PM GMT+08:00 | The clock on the bottom of the left monitor, was not synced with the clock on the right monitor on performing clinical scan. | WIP19 |  | No impact on any Test Design, it was an observation noted during the sanity and doesn't need any test design updation | Subsystem | Test Design | None | No Plan to Fix | babu_s | 0 | 4 | Console Software | Console4 | NO | Closed |  | IQon1
CTUDT00626126 | Refresh | Closed | bommiReddy_r | September 6, 2019 3:16:44 PM GMT+08:00 | Refresh: Localization: <OK> was not localized in the console messages as per Spanish language | WIP19 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | vishnu_devi | 0 |  |  | Console4 | NO | Closed | WAS | Ingenuity 64
CTUDT00626124 | Refresh | Closed | azagury_d | September 5, 2019 9:26:51 PM GMT+08:00 | Refresh: Localization: Cracked wedge and Contrast warning messages have buttons that are not transated ("Continue" and "Cancel") | WIP19 | WIP21 |  | Subsystem | Verified1 | n/a | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP20SubCom | Ovation |  | WIP21 | September 10, 2019 10:17:13 PM GMT+08:00
CTUDT00626124 | Refresh | Closed | azagury_d | September 5, 2019 9:26:51 PM GMT+08:00 | Refresh: Localization: Cracked wedge and Contrast warning messages have buttons that are not transated ("Continue" and "Cancel") | WIP19 | WIP21 |  | Subsystem | Verified1 | n/a | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP20SubCom | Brilliance Big Bore |  | WIP21 | September 10, 2019 10:17:13 PM GMT+08:00
CTUDT00626124 | Refresh | Closed | azagury_d | September 5, 2019 9:26:51 PM GMT+08:00 | Refresh: Localization: Cracked wedge and Contrast warning messages have buttons that are not transated ("Continue" and "Cancel") | WIP19 | WIP21 |  | Subsystem | Verified1 | n/a | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP20SubCom | Ingenuity 64 |  | WIP21 | September 10, 2019 10:17:13 PM GMT+08:00
CTUDT00626124 | Refresh | Closed | azagury_d | September 5, 2019 9:26:51 PM GMT+08:00 | Refresh: Localization: Cracked wedge and Contrast warning messages have buttons that are not transated ("Continue" and "Cancel") | WIP19 | WIP21 |  | Subsystem | Verified1 | n/a | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP20SubCom | IQon1 |  | WIP21 | September 10, 2019 10:17:13 PM GMT+08:00
CTUDT00626122 | Refresh | Closed | datta_i | September 5, 2019 8:44:39 PM GMT+08:00 | EC was disappeared in group of Reference Set | 4.8 | BB_WIP10 |  | Subsystem |  | NA | Fixed | rath_s | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP10_SubCom | Brilliance Big Bore |  | BB_WIP10 | September 9, 2019 7:43:05 PM GMT+08:00
CTUDT00626119 | Refresh | Closed | jose_s | September 5, 2019 4:46:39 PM GMT+08:00 | X-ray on indicator "white arrow" was missing in offline Viewer of single phase scan. See attached screen Capture. | BB_WIP9 | BB_WIP11 |  | Subsystem |  | NA | Fixed | prasad_r | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP10_SubCom | Brilliance Big Bore |  | BB_WIP11 | October 10, 2019 9:43:26 PM GMT+08:00
CTUDT00626116 | Refresh | Closed | 320053767 | September 5, 2019 4:21:33 PM GMT+08:00 | Refresh: Harmonized Phantom: Wrong automatic phantom centering during Phantom calibration. | WIP19 | WIP21 |  | Subsystem | Centering | NaN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP21SubCom | Ovation |  | WIP21 | October 15, 2019 9:03:45 PM GMT+08:00
CTUDT00626115 | Refresh | Closed | finestine_m | September 5, 2019 4:20:42 PM GMT+08:00 | Refresh: Localization: "IQ Check" Calibration's Name Is Translated In Several Languages | WIP19 |  |  | Subsystem |  |  | Duplicate | yatziv_l | 0 |  |  | DMS | NO | Closed |  | Ingenuity 64
CTUDT00626109 | Refresh | Closed | nanjunda_s | September 5, 2019 3:18:36 PM GMT+08:00 | Bug rep error is prompts and after click on "Ok" on the pop up message, "System error" was displayed | WIP19 | WIP21 |  | Subsystem |  | NA | Fixed | deshbhratar_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP21SubCom | Ovation |  | WIP21 | September 17, 2019 8:21:02 PM GMT+08:00
CTUDT00626109 | Refresh | Closed | nanjunda_s | September 5, 2019 3:18:36 PM GMT+08:00 | Bug rep error is prompts and after click on "Ok" on the pop up message, "System error" was displayed | WIP19 | WIP21 |  | Subsystem |  | NA | Fixed | deshbhratar_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP21SubCom | Brilliance 64 |  | WIP21 | September 17, 2019 8:21:02 PM GMT+08:00
CTUDT00626109 | Refresh | Closed | nanjunda_s | September 5, 2019 3:18:36 PM GMT+08:00 | Bug rep error is prompts and after click on "Ok" on the pop up message, "System error" was displayed | WIP19 | WIP21 |  | Subsystem |  | NA | Fixed | deshbhratar_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP21SubCom | Ingenuity 64 |  | WIP21 | September 17, 2019 8:21:02 PM GMT+08:00
CTUDT00626109 | Refresh | Closed | nanjunda_s | September 5, 2019 3:18:36 PM GMT+08:00 | Bug rep error is prompts and after click on "Ok" on the pop up message, "System error" was displayed | WIP19 | WIP21 |  | Subsystem |  | NA | Fixed | deshbhratar_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP21SubCom | IQon1 |  | WIP21 | September 17, 2019 8:21:02 PM GMT+08:00
CTUDT00626105 | Refresh | Closed | madankumar_s | September 5, 2019 1:41:28 PM GMT+08:00 | NULL Values are getting saved in password expiry field and password expiry warning field. | WIP19 |  | No workaround but there is clear information to the user on the tool regarding the valid range of values that has to be given. | Subsystem |  | NA | No Plan to Fix | ravindra_a | 0 | 4 | Console Software | Console4 | NO | Closed | SubNC | Ovation
CTUDT00626105 | Refresh | Closed | madankumar_s | September 5, 2019 1:41:28 PM GMT+08:00 | NULL Values are getting saved in password expiry field and password expiry warning field. | WIP19 |  | No workaround but there is clear information to the user on the tool regarding the valid range of values that has to be given. | Subsystem |  | NA | No Plan to Fix | ravindra_a | 0 | 4 | Console Software | Console4 | NO | Closed | SubNC | Brilliance 64
CTUDT00626105 | Refresh | Closed | madankumar_s | September 5, 2019 1:41:28 PM GMT+08:00 | NULL Values are getting saved in password expiry field and password expiry warning field. | WIP19 |  | No workaround but there is clear information to the user on the tool regarding the valid range of values that has to be given. | Subsystem |  | NA | No Plan to Fix | ravindra_a | 0 | 4 | Console Software | Console4 | NO | Closed | SubNC | Brilliance Big Bore
CTUDT00626105 | Refresh | Closed | madankumar_s | September 5, 2019 1:41:28 PM GMT+08:00 | NULL Values are getting saved in password expiry field and password expiry warning field. | WIP19 |  | No workaround but there is clear information to the user on the tool regarding the valid range of values that has to be given. | Subsystem |  | NA | No Plan to Fix | ravindra_a | 0 | 4 | Console Software | Console4 | NO | Closed | SubNC | Ingenuity 64
CTUDT00626105 | Refresh | Closed | madankumar_s | September 5, 2019 1:41:28 PM GMT+08:00 | NULL Values are getting saved in password expiry field and password expiry warning field. | WIP19 |  | No workaround but there is clear information to the user on the tool regarding the valid range of values that has to be given. | Subsystem |  | NA | No Plan to Fix | ravindra_a | 0 | 4 | Console Software | Console4 | NO | Closed | SubNC | IQon1
CTUDT00626103 | Refresh | Closed | gT_s | September 5, 2019 10:49:32 AM GMT+08:00 | Refresh : The installation failed is getting displayed while installing the valid software package | WIP19 | WIP21 BB_WIP11 | Change the windows installer service as automatic and start the service for MSI packages. Enable windows update service and start the service for security patches. | Subsystem |  | None | Fixed | jose_s | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP21SubCom | Ovation |  | WIP21 BB_WIP11 | September 20, 2019 1:12:15 PM GMT+08:00
CTUDT00626103 | Refresh | Closed | gT_s | September 5, 2019 10:49:32 AM GMT+08:00 | Refresh : The installation failed is getting displayed while installing the valid software package | WIP19 | WIP21 BB_WIP11 | Change the windows installer service as automatic and start the service for MSI packages. Enable windows update service and start the service for security patches. | Subsystem |  | None | Fixed | jose_s | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP21SubCom | Brilliance 64 |  | WIP21 BB_WIP11 | September 20, 2019 1:12:15 PM GMT+08:00
CTUDT00626103 | Refresh | Closed | gT_s | September 5, 2019 10:49:32 AM GMT+08:00 | Refresh : The installation failed is getting displayed while installing the valid software package | WIP19 | WIP21 BB_WIP11 | Change the windows installer service as automatic and start the service for MSI packages. Enable windows update service and start the service for security patches. | Subsystem |  | None | Fixed | jose_s | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP21SubCom | Brilliance Big Bore |  | WIP21 BB_WIP11 | September 20, 2019 1:12:15 PM GMT+08:00
CTUDT00626103 | Refresh | Closed | gT_s | September 5, 2019 10:49:32 AM GMT+08:00 | Refresh : The installation failed is getting displayed while installing the valid software package | WIP19 | WIP21 BB_WIP11 | Change the windows installer service as automatic and start the service for MSI packages. Enable windows update service and start the service for security patches. | Subsystem |  | None | Fixed | jose_s | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP21SubCom | Ingenuity 64 |  | WIP21 BB_WIP11 | September 20, 2019 1:12:15 PM GMT+08:00
CTUDT00626103 | Refresh | Closed | gT_s | September 5, 2019 10:49:32 AM GMT+08:00 | Refresh : The installation failed is getting displayed while installing the valid software package | WIP19 | WIP21 BB_WIP11 | Change the windows installer service as automatic and start the service for MSI packages. Enable windows update service and start the service for security patches. | Subsystem |  | None | Fixed | jose_s | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP21SubCom | IQon1 |  | WIP21 BB_WIP11 | September 20, 2019 1:12:15 PM GMT+08:00
CTUDT00626101 | Refresh | Closed | zhang_t | September 5, 2019 9:42:08 AM GMT+08:00 | [BB Refresh][IVC] IVC not work when Host changed to 10.168.0. network environment and IVC install failed when Host is in 10.168.0. network config | BB_WIP9 | BB_WIP11 | Install using sub net 192.168.x.x and after install perform above step | Subsystem | IVC | NA | Fixed | 320021835 | 0 | 3 | IVC Software | Console4 | NO | Closed | BB_WIP11_SubCom | Brilliance Big Bore |  | BB_WIP11 | October 14, 2019 7:41:41 PM GMT+08:00
CTUDT00626099 | Refresh | Closed | Valeria_Kashtanov | September 4, 2019 10:39:42 PM GMT+08:00 | Refresh: Calibrations: ACR Test: Pediatric Head scan Focal Spot size is not according to spec | WIP19 | WIP20 |  | Subsystem | ACR Test; Blocker20 | NaN | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP20SubCom | Ingenuity 64 |  | WIP20 | September 19, 2019 5:25:00 PM GMT+08:00
CTUDT00626096 | Refresh | Closed | Prakash_Chand_Yadav | September 4, 2019 10:08:39 PM GMT+08:00 | Refresh: Localization: Import language pop-up box header if the name is same is getting cropped for French and Spanish | WIP19 | WIP20 |  | Subsystem |  | None | Fixed | challa_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP20SubCom | Ovation |  | WIP20 | September 6, 2019 8:02:40 PM GMT+08:00
CTUDT00626096 | Refresh | Closed | Prakash_Chand_Yadav | September 4, 2019 10:08:39 PM GMT+08:00 | Refresh: Localization: Import language pop-up box header if the name is same is getting cropped for French and Spanish | WIP19 | WIP20 |  | Subsystem |  | None | Fixed | challa_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP20SubCom | Brilliance 64 |  | WIP20 | September 6, 2019 8:02:40 PM GMT+08:00
CTUDT00626096 | Refresh | Closed | Prakash_Chand_Yadav | September 4, 2019 10:08:39 PM GMT+08:00 | Refresh: Localization: Import language pop-up box header if the name is same is getting cropped for French and Spanish | WIP19 | WIP20 |  | Subsystem |  | None | Fixed | challa_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP20SubCom | Ingenuity 64 |  | WIP20 | September 6, 2019 8:02:40 PM GMT+08:00
CTUDT00626096 | Refresh | Closed | Prakash_Chand_Yadav | September 4, 2019 10:08:39 PM GMT+08:00 | Refresh: Localization: Import language pop-up box header if the name is same is getting cropped for French and Spanish | WIP19 | WIP20 |  | Subsystem |  | None | Fixed | challa_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP20SubCom | IQon1 |  | WIP20 | September 6, 2019 8:02:40 PM GMT+08:00
CTUDT00626094 | Refresh | Closed | bhawna_g | September 4, 2019 7:56:19 PM GMT+08:00 | BigBore Refresh: Planned and Actual SSDE and Average scan size values are not showing for extended scan in acquisition viewer | 4.1.10 and 4.7.7 | WIP21 |  | Subsystem |  | N/A | Fixed | pandey_a | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP10_SubCom | Ovation |  | WIP21 | September 16, 2019 9:34:53 PM GMT+08:00
CTUDT00626094 | Refresh | Closed | bhawna_g | September 4, 2019 7:56:19 PM GMT+08:00 | BigBore Refresh: Planned and Actual SSDE and Average scan size values are not showing for extended scan in acquisition viewer | 4.1.10 and 4.7.7 | WIP21 |  | Subsystem |  | N/A | Fixed | pandey_a | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP10_SubCom | Brilliance 64 |  | WIP21 | September 16, 2019 9:34:53 PM GMT+08:00
CTUDT00626094 | Refresh | Closed | bhawna_g | September 4, 2019 7:56:19 PM GMT+08:00 | BigBore Refresh: Planned and Actual SSDE and Average scan size values are not showing for extended scan in acquisition viewer | 4.1.10 and 4.7.7 | WIP21 |  | Subsystem |  | N/A | Fixed | pandey_a | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP10_SubCom | Ingenuity 64 |  | WIP21 | September 16, 2019 9:34:53 PM GMT+08:00
CTUDT00626094 | Refresh | Closed | bhawna_g | September 4, 2019 7:56:19 PM GMT+08:00 | BigBore Refresh: Planned and Actual SSDE and Average scan size values are not showing for extended scan in acquisition viewer | 4.1.10 and 4.7.7 | WIP21 |  | Subsystem |  | N/A | Fixed | pandey_a | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP10_SubCom | IQon1 |  | WIP21 | September 16, 2019 9:34:53 PM GMT+08:00
CTUDT00626091 | Refresh | Closed | Valeria_Kashtanov | September 4, 2019 7:14:46 PM GMT+08:00 | Refresh: Calibrations: Performance Infant Test: Central Uniformity ROI diameter of Spectral Head cases is not according to spec | WIP18 | WIP20 | N/A | Subsystem | Performance; Blocker20 | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP20SubCom | IQon1 |  | WIP20 | September 9, 2019 8:32:46 PM GMT+08:00
CTUDT00626089 | Refresh | Closed | shazeeda_d | September 4, 2019 6:50:37 PM GMT+08:00 | Translation UI items:Few items are not localized in respective languages | 4.7.7.43048 |  | - "Windowing" translation works as specified in showall and it's translated strings in context menu option is taken from EBW localization list. - "Close" translation in showall is taken from EBW localization list . - "Close" in "ExamCardManager->EcSharing-> Print&text" is taken from Console localization list. Please find the attachment. | Subsystem |  | NA | Works As Specified / Not A Defect | vishnu_devi | 0 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00626089 | Refresh | Closed | shazeeda_d | September 4, 2019 6:50:37 PM GMT+08:00 | Translation UI items:Few items are not localized in respective languages | 4.7.7.43048 |  | - "Windowing" translation works as specified in showall and it's translated strings in context menu option is taken from EBW localization list. - "Close" translation in showall is taken from EBW localization list . - "Close" in "ExamCardManager->EcSharing-> Print&text" is taken from Console localization list. Please find the attachment. | Subsystem |  | NA | Works As Specified / Not A Defect | vishnu_devi | 0 |  |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00626089 | Refresh | Closed | shazeeda_d | September 4, 2019 6:50:37 PM GMT+08:00 | Translation UI items:Few items are not localized in respective languages | 4.7.7.43048 |  | - "Windowing" translation works as specified in showall and it's translated strings in context menu option is taken from EBW localization list. - "Close" translation in showall is taken from EBW localization list . - "Close" in "ExamCardManager->EcSharing-> Print&text" is taken from Console localization list. Please find the attachment. | Subsystem |  | NA | Works As Specified / Not A Defect | vishnu_devi | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00626089 | Refresh | Closed | shazeeda_d | September 4, 2019 6:50:37 PM GMT+08:00 | Translation UI items:Few items are not localized in respective languages | 4.7.7.43048 |  | - "Windowing" translation works as specified in showall and it's translated strings in context menu option is taken from EBW localization list. - "Close" translation in showall is taken from EBW localization list . - "Close" in "ExamCardManager->EcSharing-> Print&text" is taken from Console localization list. Please find the attachment. | Subsystem |  | NA | Works As Specified / Not A Defect | vishnu_devi | 0 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00626086 | Refresh | Closed | yatziv_l | September 4, 2019 5:00:29 PM GMT+08:00 | RFC: Air Calibration: Adding 1sec and 1.5sec modes for UHR cases in Ingenuity | N.A | WIP20 |  | Subsystem | Air Calibration | n/a | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP20SubCom | Ingenuity 64 |  | WIP20 | September 11, 2019 9:06:48 PM GMT+08:00
CTUDT00626085 | Refresh | Closed | wang_l | September 4, 2019 4:56:28 PM GMT+08:00 | Refresh: Inherit issue for CT-NM.SRS.Service.232 | WIP19 | ING, BR64 SRS DHF298554 Rev08 BB SRS DHF298555 Rev08 |  | Subsystem |  | Service.232 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP20SubCom | Brilliance 64 |  | ING, BR64 SRS DHF298554 Rev08 BB SRS DHF298555 Rev08 | November 4, 2019 10:20:04 AM GMT+08:00
CTUDT00626085 | Refresh | Closed | wang_l | September 4, 2019 4:56:28 PM GMT+08:00 | Refresh: Inherit issue for CT-NM.SRS.Service.232 | WIP19 | ING, BR64 SRS DHF298554 Rev08 BB SRS DHF298555 Rev08 |  | Subsystem |  | Service.232 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP20SubCom | Brilliance Big Bore |  | ING, BR64 SRS DHF298554 Rev08 BB SRS DHF298555 Rev08 | November 4, 2019 10:20:04 AM GMT+08:00
CTUDT00626085 | Refresh | Closed | wang_l | September 4, 2019 4:56:28 PM GMT+08:00 | Refresh: Inherit issue for CT-NM.SRS.Service.232 | WIP19 | ING, BR64 SRS DHF298554 Rev08 BB SRS DHF298555 Rev08 |  | Subsystem |  | Service.232 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP20SubCom | Ingenuity 64 |  | ING, BR64 SRS DHF298554 Rev08 BB SRS DHF298555 Rev08 | November 4, 2019 10:20:04 AM GMT+08:00
CTUDT00626083 | Refresh | Closed | zhu_je | September 4, 2019 4:45:54 PM GMT+08:00 | [Refresh][PCM] Cannot find label for language “Serbian” Br64 PCM | WIP19 | DHF215793 - 728232 - Brilliance 64 upgrades PCM- Rev17 | [Refresh][PCM] Cannot find label for language “Serbian” Br64 PCM | Subsystem |  | NA | Fixed | hong_x | 0 | 2 | PCM | Support-System | NO | Closed | pcm | Brilliance 64 |  | DHF215793 - 728232 - Brilliance 64 upgrades PCM- Rev17 | November 4, 2019 10:21:50 AM GMT+08:00
CTUDT00626081 | Refresh | Closed | yatziv_l | September 4, 2019 4:26:00 PM GMT+08:00 | Refresh: Manual Bad Detectors: Manual Bad Detectors for O-level should appear in Utilities tab | WIP19 | WIP20 |  | Subsystem | Utilities tab;Verified1 | n/a | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP20SubCom | Ovation |  | WIP20 | September 9, 2019 7:54:24 PM GMT+08:00
CTUDT00626081 | Refresh | Closed | yatziv_l | September 4, 2019 4:26:00 PM GMT+08:00 | Refresh: Manual Bad Detectors: Manual Bad Detectors for O-level should appear in Utilities tab | WIP19 | WIP20 |  | Subsystem | Utilities tab;Verified1 | n/a | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP20SubCom | Brilliance Big Bore |  | WIP20 | September 9, 2019 7:54:24 PM GMT+08:00
CTUDT00626081 | Refresh | Closed | yatziv_l | September 4, 2019 4:26:00 PM GMT+08:00 | Refresh: Manual Bad Detectors: Manual Bad Detectors for O-level should appear in Utilities tab | WIP19 | WIP20 |  | Subsystem | Utilities tab;Verified1 | n/a | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP20SubCom | Ingenuity 64 |  | WIP20 | September 9, 2019 7:54:24 PM GMT+08:00
CTUDT00626081 | Refresh | Closed | yatziv_l | September 4, 2019 4:26:00 PM GMT+08:00 | Refresh: Manual Bad Detectors: Manual Bad Detectors for O-level should appear in Utilities tab | WIP19 | WIP20 |  | Subsystem | Utilities tab;Verified1 | n/a | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP20SubCom | IQon1 |  | WIP20 | September 9, 2019 7:54:24 PM GMT+08:00
CTUDT00626078 | Refresh | Closed | barnea_d | September 4, 2019 3:30:28 PM GMT+08:00 | iCT,WIP19, steps test,CSOF phantom-steps Z failed due to couch alignment value | WIP19 | WIP20 |  | Subsystem | Steps test | NaN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP20SubCom | Ovation |  | WIP20 | September 10, 2019 8:13:17 PM GMT+08:00
CTUDT00626075 | Refresh | Closed | zhang_t | September 4, 2019 3:03:36 PM GMT+08:00 | [BB Refresh][IVC] When couch reach the extreme Out position via Rocker Switch, the increment movement In Button was enabled but not function | BB_WIP9 |  | User can move the couch to a position > 73.0 mm, then the motion command to a position between 23.0 and 73.0 is possible. | Subsystem | IVC | NA | Works As Specified / Not A Defect | zhong_x | 0 | 4 |  | GCX | NO | Closed |  | Brilliance Big Bore
CTUDT00626073 | Refresh | Closed | zhang_t | September 4, 2019 2:55:02 PM GMT+08:00 | [BB Refresh][IVC] The increment value was changed to 1 slice from other selection automatically after reaching extreme position or after pressing IVC buttons when couch in extreme position. | BB_WIP9 | BB_WIP11 | After the selection gets changed, the user can manually again select the previous selected value. | Subsystem | IVC | NA | Fixed | Snehashis_Dey | 0 | 3 | IVC Software | Console4 | NO | Closed | BB_WIP10_SubCom | Brilliance Big Bore |  | BB_WIP11 | October 1, 2019 8:08:01 PM GMT+08:00
CTUDT00626070 | Refresh | Closed | zhang_t | September 4, 2019 2:37:31 PM GMT+08:00 | [BB Refresh][IVC] The Horizontal table Speed has no selection every time restart IVC | BB_WIP9 | BB_WIP10 |  | Subsystem | IVC;Blocker20 | NA | Fixed | Snehashis_Dey | 0 | 3 | IVC Software | Console4 | NO | Closed | BB_WIP10_SubCom | Brilliance Big Bore |  | BB_WIP10 | September 16, 2019 8:23:21 PM GMT+08:00
CTUDT00626068 | Refresh | Closed | zhang_t | September 4, 2019 2:22:12 PM GMT+08:00 | [BB Refresh][IVC] The couch position in CT box and IVC SW didn't match sometimes when reach the extreme horizontal position | BB_WIP9 | BB_WIP10 |  | Subsystem | IVC; couch position | Console-SRS.15851 | Fixed | Snehashis_Dey | 0 | 2 | Console Software | Console4 | YES | Closed | BB_WIP10_SubCom | Brilliance Big Bore |  | BB_WIP10 | September 17, 2019 5:53:10 PM GMT+08:00
CTUDT00626064 | Refresh | Closed | gT_s | September 4, 2019 11:21:16 AM GMT+08:00 | Refresh : The localization is not working for the software packages as per selected languages. | WIP19 |  |  | Subsystem | Software Packages | None | No Plan to Fix | sanjeeva_a | 0 | 2 |  | Console4 | NO | Closed | SubNC | Ovation
CTUDT00626064 | Refresh | Closed | gT_s | September 4, 2019 11:21:16 AM GMT+08:00 | Refresh : The localization is not working for the software packages as per selected languages. | WIP19 |  |  | Subsystem | Software Packages | None | No Plan to Fix | sanjeeva_a | 0 | 2 |  | Console4 | NO | Closed | SubNC | Brilliance 64
CTUDT00626064 | Refresh | Closed | gT_s | September 4, 2019 11:21:16 AM GMT+08:00 | Refresh : The localization is not working for the software packages as per selected languages. | WIP19 |  |  | Subsystem | Software Packages | None | No Plan to Fix | sanjeeva_a | 0 | 2 |  | Console4 | NO | Closed | SubNC | Brilliance Big Bore
CTUDT00626064 | Refresh | Closed | gT_s | September 4, 2019 11:21:16 AM GMT+08:00 | Refresh : The localization is not working for the software packages as per selected languages. | WIP19 |  |  | Subsystem | Software Packages | None | No Plan to Fix | sanjeeva_a | 0 | 2 |  | Console4 | NO | Closed | SubNC | Ingenuity 64
CTUDT00626064 | Refresh | Closed | gT_s | September 4, 2019 11:21:16 AM GMT+08:00 | Refresh : The localization is not working for the software packages as per selected languages. | WIP19 |  |  | Subsystem | Software Packages | None | No Plan to Fix | sanjeeva_a | 0 | 2 |  | Console4 | NO | Closed | SubNC | IQon1
CTUDT00626062 | Refresh | Closed | gT_s | September 3, 2019 8:40:49 PM GMT+08:00 | Refresh : The cancelling the software package is not happening when user clicks on cancel button. | WIP19 |  | In this scenario, GUI will be updated with cancelled state if user restarts the system. | Subsystem |  | None | No Plan to Fix | jose_s | 0 | 4 | Service Tools_Software | Service Tools | NO | Closed | SubNC | Ovation
CTUDT00626062 | Refresh | Closed | gT_s | September 3, 2019 8:40:49 PM GMT+08:00 | Refresh : The cancelling the software package is not happening when user clicks on cancel button. | WIP19 |  | In this scenario, GUI will be updated with cancelled state if user restarts the system. | Subsystem |  | None | No Plan to Fix | jose_s | 0 | 4 | Service Tools_Software | Service Tools | NO | Closed | SubNC | Brilliance 64
CTUDT00626062 | Refresh | Closed | gT_s | September 3, 2019 8:40:49 PM GMT+08:00 | Refresh : The cancelling the software package is not happening when user clicks on cancel button. | WIP19 |  | In this scenario, GUI will be updated with cancelled state if user restarts the system. | Subsystem |  | None | No Plan to Fix | jose_s | 0 | 4 | Service Tools_Software | Service Tools | NO | Closed | SubNC | Brilliance Big Bore
CTUDT00626062 | Refresh | Closed | gT_s | September 3, 2019 8:40:49 PM GMT+08:00 | Refresh : The cancelling the software package is not happening when user clicks on cancel button. | WIP19 |  | In this scenario, GUI will be updated with cancelled state if user restarts the system. | Subsystem |  | None | No Plan to Fix | jose_s | 0 | 4 | Service Tools_Software | Service Tools | NO | Closed | SubNC | Ingenuity 64
CTUDT00626062 | Refresh | Closed | gT_s | September 3, 2019 8:40:49 PM GMT+08:00 | Refresh : The cancelling the software package is not happening when user clicks on cancel button. | WIP19 |  | In this scenario, GUI will be updated with cancelled state if user restarts the system. | Subsystem |  | None | No Plan to Fix | jose_s | 0 | 4 | Service Tools_Software | Service Tools | NO | Closed | SubNC | IQon1
CTUDT00626059 | Refresh | Closed | beck_d | September 3, 2019 8:33:23 PM GMT+08:00 | Refresh: Calibrations: Image Ranking Tool appears in Reasons Matrix Menu Spec although it is not implemented | WIP19 | DHF232540 Rev14 |  | Subsystem | Reason Matrix | N/A | Fixed | richards_b | 0 | 2 | Design Documentation | Support-Physics | NO | Closed | WIP20SubCom | Ingenuity 64 |  | DHF232540 Rev14 | September 19, 2019 10:13:48 AM GMT+08:00
CTUDT00626057 | Refresh | Closed | Bat-El_Asaraf | September 3, 2019 7:59:00 PM GMT+08:00 | Pre-Val Refresh: Emergency user access is disabled by default | WIP19 | WIP20 | NA. | Subsystem | Install spec | N/A | Fixed | hilas | 0 | 2 | Console Software | Support-System | NO | Closed | SegRec | Ovation |  | WIP20 | September 26, 2019 3:27:15 PM GMT+08:00
CTUDT00626057 | Refresh | Closed | Bat-El_Asaraf | September 3, 2019 7:59:00 PM GMT+08:00 | Pre-Val Refresh: Emergency user access is disabled by default | WIP19 | WIP20 | NA. | Subsystem | Install spec | N/A | Fixed | hilas | 0 | 2 | Console Software | Support-System | NO | Closed | SegRec | IQon1 |  | WIP20 | September 26, 2019 3:27:15 PM GMT+08:00
CTUDT00626056 | Refresh | Closed | swack_g | September 3, 2019 7:48:47 PM GMT+08:00 | iCT,WIP19,SHADOW test, channel shadow xz- wrong limits | WIP19 | WIP20 |  | Subsystem | Shadow; Blocker20 | NaN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP20SubCom | Ovation |  | WIP20 | September 9, 2019 7:27:40 PM GMT+08:00
CTUDT00626053 | Refresh | Closed | 320053769 | September 3, 2019 7:02:35 PM GMT+08:00 | WIP19 - After recording to CDR failed “system error”, “system is still initializing. Please wait”, “ExamApplication.exe has stopped working” pop-ups appeared | WIP19 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | veerabhadraiah | 1 |  | Console USP_EBW | iCAP | NO | Closed |  | Ovation
CTUDT00626053 | Refresh | Closed | 320053769 | September 3, 2019 7:02:35 PM GMT+08:00 | WIP19 - After recording to CDR failed “system error”, “system is still initializing. Please wait”, “ExamApplication.exe has stopped working” pop-ups appeared | WIP19 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | veerabhadraiah | 1 |  | Console USP_EBW | iCAP | NO | Closed |  | Brilliance 64
CTUDT00626053 | Refresh | Closed | 320053769 | September 3, 2019 7:02:35 PM GMT+08:00 | WIP19 - After recording to CDR failed “system error”, “system is still initializing. Please wait”, “ExamApplication.exe has stopped working” pop-ups appeared | WIP19 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | veerabhadraiah | 1 |  | Console USP_EBW | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00626053 | Refresh | Closed | 320053769 | September 3, 2019 7:02:35 PM GMT+08:00 | WIP19 - After recording to CDR failed “system error”, “system is still initializing. Please wait”, “ExamApplication.exe has stopped working” pop-ups appeared | WIP19 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | veerabhadraiah | 1 |  | Console USP_EBW | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00626053 | Refresh | Closed | 320053769 | September 3, 2019 7:02:35 PM GMT+08:00 | WIP19 - After recording to CDR failed “system error”, “system is still initializing. Please wait”, “ExamApplication.exe has stopped working” pop-ups appeared | WIP19 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | veerabhadraiah | 1 |  | Console USP_EBW | iCAP | NO | Closed |  | IQon1
CTUDT00626050 | Refresh | Closed | Dana_Kamyshev | September 3, 2019 6:55:46 PM GMT+08:00 | Refresh - System error before clinical scan | WIP19 |  | User can press "OK" on system error message and continue working. | Subsystem |  | None | Duplicate | prasad_r | 0 |  | Console Software | Console4 | NO | Closed | Duplicate | Ovation
CTUDT00626050 | Refresh | Closed | Dana_Kamyshev | September 3, 2019 6:55:46 PM GMT+08:00 | Refresh - System error before clinical scan | WIP19 |  | User can press "OK" on system error message and continue working. | Subsystem |  | None | Duplicate | prasad_r | 0 |  | Console Software | Console4 | NO | Closed | Duplicate | Brilliance 64
CTUDT00626050 | Refresh | Closed | Dana_Kamyshev | September 3, 2019 6:55:46 PM GMT+08:00 | Refresh - System error before clinical scan | WIP19 |  | User can press "OK" on system error message and continue working. | Subsystem |  | None | Duplicate | prasad_r | 0 |  | Console Software | Console4 | NO | Closed | Duplicate | Brilliance Big Bore
CTUDT00626050 | Refresh | Closed | Dana_Kamyshev | September 3, 2019 6:55:46 PM GMT+08:00 | Refresh - System error before clinical scan | WIP19 |  | User can press "OK" on system error message and continue working. | Subsystem |  | None | Duplicate | prasad_r | 0 |  | Console Software | Console4 | NO | Closed | Duplicate | Ingenuity 64
CTUDT00626050 | Refresh | Closed | Dana_Kamyshev | September 3, 2019 6:55:46 PM GMT+08:00 | Refresh - System error before clinical scan | WIP19 |  | User can press "OK" on system error message and continue working. | Subsystem |  | None | Duplicate | prasad_r | 0 |  | Console Software | Console4 | NO | Closed | Duplicate | IQon1
CTUDT00626048 | Refresh | Closed | bagdady_b | September 3, 2019 6:40:24 PM GMT+08:00 | Refresh: sCTV - Subtraction of batch images from original scanned result has residuals | WIP19 | WIP20 |  | Subsystem | Test Design | CT-NM.SRS.TA.395], CT-NM.SRS.TA.389 | Fixed | tamari_o | 0 | 4 | Test Design | Support-System | NO | Closed | SubCom-09-Jan-2020 | IQon1 |  | WIP20 | January 31, 2020 1:11:21 AM GMT+08:00
CTUDT00626045 | Refresh | Closed | 320053769 | September 3, 2019 5:03:48 PM GMT+08:00 | WIP19 - “System Error” warning pooped while scanning. | WIP19 |  | User can press "OK" on system error message and continue working. | Subsystem |  | None | Duplicate | jose_s | 0 |  | Console Software | Console4 | NO | Closed | Duplicate | Ovation
CTUDT00626045 | Refresh | Closed | 320053769 | September 3, 2019 5:03:48 PM GMT+08:00 | WIP19 - “System Error” warning pooped while scanning. | WIP19 |  | User can press "OK" on system error message and continue working. | Subsystem |  | None | Duplicate | jose_s | 0 |  | Console Software | Console4 | NO | Closed | Duplicate | Brilliance 64
CTUDT00626045 | Refresh | Closed | 320053769 | September 3, 2019 5:03:48 PM GMT+08:00 | WIP19 - “System Error” warning pooped while scanning. | WIP19 |  | User can press "OK" on system error message and continue working. | Subsystem |  | None | Duplicate | jose_s | 0 |  | Console Software | Console4 | NO | Closed | Duplicate | Brilliance Big Bore
CTUDT00626045 | Refresh | Closed | 320053769 | September 3, 2019 5:03:48 PM GMT+08:00 | WIP19 - “System Error” warning pooped while scanning. | WIP19 |  | User can press "OK" on system error message and continue working. | Subsystem |  | None | Duplicate | jose_s | 0 |  | Console Software | Console4 | NO | Closed | Duplicate | Ingenuity 64
CTUDT00626045 | Refresh | Closed | 320053769 | September 3, 2019 5:03:48 PM GMT+08:00 | WIP19 - “System Error” warning pooped while scanning. | WIP19 |  | User can press "OK" on system error message and continue working. | Subsystem |  | None | Duplicate | jose_s | 0 |  | Console Software | Console4 | NO | Closed | Duplicate | IQon1
CTUDT00626043 | Refresh | Closed | wang_l | September 3, 2019 4:04:53 PM GMT+08:00 | Refresh: Script error when set Automatic IP to No if hospital Lan was not connected | BB_WIP9 |  |  | Subsystem | Computer and Network configuration | NA | Works As Specified / Not A Defect | lin_s | 0 |  |  | Support-System | NO | Closed | WAS | Brilliance Big Bore
CTUDT00626043 | Refresh | Closed | wang_l | September 3, 2019 4:04:53 PM GMT+08:00 | Refresh: Script error when set Automatic IP to No if hospital Lan was not connected | BB_WIP9 |  |  | Subsystem | Computer and Network configuration | NA | Works As Specified / Not A Defect | lin_s | 0 |  |  | Support-System | NO | Closed | WAS | Ingenuity 64
CTUDT00626040 | Refresh | Closed | wang_l | September 3, 2019 3:43:28 PM GMT+08:00 | Refresh: Network and Computer Name update by olevel cannot take effect | BB_WIP9 | WIP22 BB_WIP12 |  | System | Computer configuration;Verified1 | NA | Fixed | li_e | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP21SubCom | Brilliance 64 |  | WIP22 BB_WIP12 | November 14, 2019 1:29:53 PM GMT+08:00
CTUDT00626040 | Refresh | Closed | wang_l | September 3, 2019 3:43:28 PM GMT+08:00 | Refresh: Network and Computer Name update by olevel cannot take effect | BB_WIP9 | WIP22 BB_WIP12 |  | System | Computer configuration;Verified1 | NA | Fixed | li_e | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP21SubCom | Brilliance Big Bore |  | WIP22 BB_WIP12 | November 14, 2019 1:29:53 PM GMT+08:00
CTUDT00626040 | Refresh | Closed | wang_l | September 3, 2019 3:43:28 PM GMT+08:00 | Refresh: Network and Computer Name update by olevel cannot take effect | BB_WIP9 | WIP22 BB_WIP12 |  | System | Computer configuration;Verified1 | NA | Fixed | li_e | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP21SubCom | Ingenuity 64 |  | WIP22 BB_WIP12 | November 14, 2019 1:29:53 PM GMT+08:00
CTUDT00626038 | Refresh | Closed | Bat-El_Asaraf | September 3, 2019 3:41:08 PM GMT+08:00 | Pre-Val Refresh: Missing display of number of trials before locking the account | WIP19 |  | NA | Subsystem |  | NA | Works As Specified / Not A Defect | kapadia_v | 0 |  | Console Software | Console4 | NO | Closed | WAS | Ovation
CTUDT00626038 | Refresh | Closed | Bat-El_Asaraf | September 3, 2019 3:41:08 PM GMT+08:00 | Pre-Val Refresh: Missing display of number of trials before locking the account | WIP19 |  | NA | Subsystem |  | NA | Works As Specified / Not A Defect | kapadia_v | 0 |  | Console Software | Console4 | NO | Closed | WAS | Brilliance 64
CTUDT00626038 | Refresh | Closed | Bat-El_Asaraf | September 3, 2019 3:41:08 PM GMT+08:00 | Pre-Val Refresh: Missing display of number of trials before locking the account | WIP19 |  | NA | Subsystem |  | NA | Works As Specified / Not A Defect | kapadia_v | 0 |  | Console Software | Console4 | NO | Closed | WAS | Brilliance Big Bore
CTUDT00626038 | Refresh | Closed | Bat-El_Asaraf | September 3, 2019 3:41:08 PM GMT+08:00 | Pre-Val Refresh: Missing display of number of trials before locking the account | WIP19 |  | NA | Subsystem |  | NA | Works As Specified / Not A Defect | kapadia_v | 0 |  | Console Software | Console4 | NO | Closed | WAS | Ingenuity 64
CTUDT00626038 | Refresh | Closed | Bat-El_Asaraf | September 3, 2019 3:41:08 PM GMT+08:00 | Pre-Val Refresh: Missing display of number of trials before locking the account | WIP19 |  | NA | Subsystem |  | NA | Works As Specified / Not A Defect | kapadia_v | 0 |  | Console Software | Console4 | NO | Closed | WAS | IQon1
CTUDT00626037 | Refresh | Closed | or_shmuel | September 3, 2019 3:18:43 PM GMT+08:00 | Refresh:freeze screen on initializing status for approx. 7 min | WIP19 |  |  | Subsystem | Workflow |  | Works As Specified / Not A Defect | tamari_o | 1 |  |  | Support-System | NO | Closed |  | IQon1
CTUDT00626035 | Refresh | Closed | Bat-El_Asaraf | September 3, 2019 2:52:51 PM GMT+08:00 | Pre-Val Refresh: Log in User name is not displayed on right monitor | BB_WIP9 | WIP21 |  | System | Right Monitor | NA | Fixed | srushti_k | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | BB_WIP11_SubCom | Brilliance Big Bore |  | WIP21 | October 9, 2019 4:58:53 PM GMT+08:00
CTUDT00626032 | Refresh | Closed | zhu_je | September 3, 2019 12:23:53 PM GMT+08:00 | [Refresh] “System Error” message pop up when opened workflow after importing an existed language to Voice Manager | BB_WIP9 |  |  | Subsystem | Import to Voice Manager | NA | Duplicate | prasad_r | 0 |  |  | Console4 | NO | Closed | Duplicate | Brilliance Big Bore
CTUDT00626028 | Refresh | Closed | levy_s | September 2, 2019 7:24:35 PM GMT+08:00 | Refresh: Slow performance is evidence in the compression phase when collecting a BugRep. | WIP19 | BB_WIP11.1 WIP21.1 |  | Subsystem | Performance, Blocker | NA | Fixed | Prakash_Chand_Yadav | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP21SubCom | Ovation |  | BB_WIP11.1 WIP21.1 | October 24, 2019 2:24:10 PM GMT+08:00
CTUDT00626028 | Refresh | Closed | levy_s | September 2, 2019 7:24:35 PM GMT+08:00 | Refresh: Slow performance is evidence in the compression phase when collecting a BugRep. | WIP19 | BB_WIP11.1 WIP21.1 |  | Subsystem | Performance, Blocker | NA | Fixed | Prakash_Chand_Yadav | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP21SubCom | Brilliance 64 |  | BB_WIP11.1 WIP21.1 | October 24, 2019 2:24:10 PM GMT+08:00
CTUDT00626028 | Refresh | Closed | levy_s | September 2, 2019 7:24:35 PM GMT+08:00 | Refresh: Slow performance is evidence in the compression phase when collecting a BugRep. | WIP19 | BB_WIP11.1 WIP21.1 |  | Subsystem | Performance, Blocker | NA | Fixed | Prakash_Chand_Yadav | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP21SubCom | Brilliance Big Bore |  | BB_WIP11.1 WIP21.1 | October 24, 2019 2:24:10 PM GMT+08:00
CTUDT00626028 | Refresh | Closed | levy_s | September 2, 2019 7:24:35 PM GMT+08:00 | Refresh: Slow performance is evidence in the compression phase when collecting a BugRep. | WIP19 | BB_WIP11.1 WIP21.1 |  | Subsystem | Performance, Blocker | NA | Fixed | Prakash_Chand_Yadav | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP21SubCom | Ingenuity 64 |  | BB_WIP11.1 WIP21.1 | October 24, 2019 2:24:10 PM GMT+08:00
CTUDT00626028 | Refresh | Closed | levy_s | September 2, 2019 7:24:35 PM GMT+08:00 | Refresh: Slow performance is evidence in the compression phase when collecting a BugRep. | WIP19 | BB_WIP11.1 WIP21.1 |  | Subsystem | Performance, Blocker | NA | Fixed | Prakash_Chand_Yadav | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP21SubCom | IQon1 |  | BB_WIP11.1 WIP21.1 | October 24, 2019 2:24:10 PM GMT+08:00
CTUDT00626026 | Refresh | Closed | finestine_m | September 2, 2019 7:10:58 PM GMT+08:00 | Refresh: Constancy test: Harmonized phantom- some limits and scan parameters are not according to spec | WIP18 | WIP20 |  | Subsystem | Constancy; Blocker20 | NaN | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP20SubCom | IQon1 |  | WIP20 | September 9, 2019 7:15:21 PM GMT+08:00
CTUDT00626024 | Refresh | Closed | bagdady_b | September 2, 2019 6:56:46 PM GMT+08:00 | Refresh - Different thickness for IMR result when performing extension | WIP19 | WIP21 |  | Subsystem |  | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP20SubCom | Ovation |  | WIP21 | September 27, 2019 2:17:14 PM GMT+08:00
CTUDT00626024 | Refresh | Closed | bagdady_b | September 2, 2019 6:56:46 PM GMT+08:00 | Refresh - Different thickness for IMR result when performing extension | WIP19 | WIP21 |  | Subsystem |  | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP20SubCom | Brilliance 64 |  | WIP21 | September 27, 2019 2:17:14 PM GMT+08:00
CTUDT00626024 | Refresh | Closed | bagdady_b | September 2, 2019 6:56:46 PM GMT+08:00 | Refresh - Different thickness for IMR result when performing extension | WIP19 | WIP21 |  | Subsystem |  | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP20SubCom | Ingenuity 64 |  | WIP21 | September 27, 2019 2:17:14 PM GMT+08:00
CTUDT00626024 | Refresh | Closed | bagdady_b | September 2, 2019 6:56:46 PM GMT+08:00 | Refresh - Different thickness for IMR result when performing extension | WIP19 | WIP21 |  | Subsystem |  | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP20SubCom | IQon1 |  | WIP21 | September 27, 2019 2:17:14 PM GMT+08:00
CTUDT00626021 | Refresh | Closed | alex_nihaev | September 2, 2019 3:45:10 PM GMT+08:00 | Refresh:WIP19: Calibrations: IQ Check: Legacy Phantom: Noise failed above limit for Head and Body | WIP19 | WIP20 |  | Subsystem | Blocker20 | N.A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP20SubCom | Ovation |  | WIP20 | September 18, 2019 2:49:33 PM GMT+08:00
CTUDT00626019 | Refresh | Closed | cai_t | September 2, 2019 1:28:39 PM GMT+08:00 | Multi language IFUs(Instructions for Use) are not ready | WIP19 | WIP20 | No | Subsystem |  | CT-NM.SRS.SysSupport.496 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | Post Validation | Brilliance 64 |  | WIP20 | November 6, 2019 2:59:49 PM GMT+08:00
CTUDT00626019 | Refresh | Closed | cai_t | September 2, 2019 1:28:39 PM GMT+08:00 | Multi language IFUs(Instructions for Use) are not ready | WIP19 | WIP20 | No | Subsystem |  | CT-NM.SRS.SysSupport.496 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | Post Validation | Ingenuity 64 |  | WIP20 | November 6, 2019 2:59:49 PM GMT+08:00
CTUDT00626017 | Refresh | Closed | liu_bo | September 2, 2019 12:40:32 PM GMT+08:00 | There is not FFSX Position & FFSZ Position in Dipole Current Report Traces. | WIP18 | Gantry SSRS V04 |  | Subsystem | Visual Diagnostics | GCX-SSRS-iCT-20735 | Fixed | bruce_li | 0 | 2 | Requirements Documentation | GCX | NO | Closed | WAS | Ovation |  | Gantry SSRS V04 | September 26, 2019 8:26:15 AM GMT+08:00
CTUDT00626015 | Refresh | Closed | cai_t | September 2, 2019 12:26:34 PM GMT+08:00 | Multi language IFUs(Instructions for Use) are not ready | BB_WIP8 | WIP20 | No | Subsystem |  | CT-NM.SRS.SysSupport.496 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | Post Validation | Brilliance Big Bore |  | WIP20 | November 6, 2019 2:58:07 PM GMT+08:00
CTUDT00626012 | Refresh | Closed | finestine_m | September 1, 2019 9:13:22 PM GMT+08:00 | Refresh: Performance Infant: Wrong FOV For Body Scans | WIP18 | WIP22 |  | Subsystem | Performance; Blocker20 | NaN | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP20SubCom | IQon1 |  | WIP22 | November 18, 2019 9:10:26 PM GMT+08:00
CTUDT00626010 | Refresh | Closed | 320053769 | September 1, 2019 5:18:31 PM GMT+08:00 | Queue manager fails to transfer studies with ~3000 (or more) images to CDR | WIP19 |  | Retry transfer | Subsystem |  | NA | No Plan to Fix | s_panigrahi | 0 | 3 | Console USP_EBW | Support-System | NO | Closed |  | Ovation
CTUDT00626010 | Refresh | Closed | 320053769 | September 1, 2019 5:18:31 PM GMT+08:00 | Queue manager fails to transfer studies with ~3000 (or more) images to CDR | WIP19 |  | Retry transfer | Subsystem |  | NA | No Plan to Fix | s_panigrahi | 0 | 3 | Console USP_EBW | Support-System | NO | Closed |  | Brilliance 64
CTUDT00626010 | Refresh | Closed | 320053769 | September 1, 2019 5:18:31 PM GMT+08:00 | Queue manager fails to transfer studies with ~3000 (or more) images to CDR | WIP19 |  | Retry transfer | Subsystem |  | NA | No Plan to Fix | s_panigrahi | 0 | 3 | Console USP_EBW | Support-System | NO | Closed |  | Brilliance Big Bore
CTUDT00626010 | Refresh | Closed | 320053769 | September 1, 2019 5:18:31 PM GMT+08:00 | Queue manager fails to transfer studies with ~3000 (or more) images to CDR | WIP19 |  | Retry transfer | Subsystem |  | NA | No Plan to Fix | s_panigrahi | 0 | 3 | Console USP_EBW | Support-System | NO | Closed |  | Ingenuity 64
CTUDT00626010 | Refresh | Closed | 320053769 | September 1, 2019 5:18:31 PM GMT+08:00 | Queue manager fails to transfer studies with ~3000 (or more) images to CDR | WIP19 |  | Retry transfer | Subsystem |  | NA | No Plan to Fix | s_panigrahi | 0 | 3 | Console USP_EBW | Support-System | NO | Closed |  | IQon1
CTUDT00626008 | Refresh | Closed | liu_bo | September 1, 2019 4:42:49 PM GMT+08:00 | X-Ray Performance Tests missing test item. | WIP18 | WIP22 DHF298554 Rev09 DHF298553 Rev.09 DHF298556 Rev.10 |  | System | Gantry Calibration; | CT-NM.SRS.Service.1872 , CT-NM.SRS.Service.1873 | Fixed | bandil_j | 0 | 2 | Service Tools_Software | Support-System | NO | Closed | SubCom-19-Dec-2019 | Ovation |  | WIP22 DHF298554 Rev09 DHF298553 Rev.09 DHF298556 Rev.10 | December 20, 2019 3:51:17 AM GMT+08:00
CTUDT00626008 | Refresh | Closed | liu_bo | September 1, 2019 4:42:49 PM GMT+08:00 | X-Ray Performance Tests missing test item. | WIP18 | WIP22 DHF298554 Rev09 DHF298553 Rev.09 DHF298556 Rev.10 |  | System | Gantry Calibration; | CT-NM.SRS.Service.1872 , CT-NM.SRS.Service.1873 | Fixed | bandil_j | 0 | 2 | Service Tools_Software | Support-System | NO | Closed | SubCom-19-Dec-2019 | Brilliance 64 |  | WIP22 DHF298554 Rev09 DHF298553 Rev.09 DHF298556 Rev.10 | December 20, 2019 3:51:17 AM GMT+08:00
CTUDT00626008 | Refresh | Closed | liu_bo | September 1, 2019 4:42:49 PM GMT+08:00 | X-Ray Performance Tests missing test item. | WIP18 | WIP22 DHF298554 Rev09 DHF298553 Rev.09 DHF298556 Rev.10 |  | System | Gantry Calibration; | CT-NM.SRS.Service.1872 , CT-NM.SRS.Service.1873 | Fixed | bandil_j | 0 | 2 | Service Tools_Software | Support-System | NO | Closed | SubCom-19-Dec-2019 | Brilliance Big Bore |  | WIP22 DHF298554 Rev09 DHF298553 Rev.09 DHF298556 Rev.10 | December 20, 2019 3:51:17 AM GMT+08:00
CTUDT00626008 | Refresh | Closed | liu_bo | September 1, 2019 4:42:49 PM GMT+08:00 | X-Ray Performance Tests missing test item. | WIP18 | WIP22 DHF298554 Rev09 DHF298553 Rev.09 DHF298556 Rev.10 |  | System | Gantry Calibration; | CT-NM.SRS.Service.1872 , CT-NM.SRS.Service.1873 | Fixed | bandil_j | 0 | 2 | Service Tools_Software | Support-System | NO | Closed | SubCom-19-Dec-2019 | Ingenuity 64 |  | WIP22 DHF298554 Rev09 DHF298553 Rev.09 DHF298556 Rev.10 | December 20, 2019 3:51:17 AM GMT+08:00
CTUDT00626008 | Refresh | Closed | liu_bo | September 1, 2019 4:42:49 PM GMT+08:00 | X-Ray Performance Tests missing test item. | WIP18 | WIP22 DHF298554 Rev09 DHF298553 Rev.09 DHF298556 Rev.10 |  | System | Gantry Calibration; | CT-NM.SRS.Service.1872 , CT-NM.SRS.Service.1873 | Fixed | bandil_j | 0 | 2 | Service Tools_Software | Support-System | NO | Closed | SubCom-19-Dec-2019 | IQon1 |  | WIP22 DHF298554 Rev09 DHF298553 Rev.09 DHF298556 Rev.10 | December 20, 2019 3:51:17 AM GMT+08:00
CTUDT00626007 | Refresh | Closed | borise | September 1, 2019 4:24:55 PM GMT+08:00 | Spectral algorithm failed to produce Spectral images | 4.7.5 | WIP20 | Request the result again (from the Exam Application or VieweTwo of sCTV). | System | Spectral Algorithm | NA | Fixed | aviw | 0 | 3 | SpectralAlgorithms | Support_Spectral | YES | Closed | WIP20SubCom | IQon1 | PR 9430129 | WIP20 | September 2, 2019 5:11:21 PM GMT+08:00
CTUDT00626004 | Refresh | Closed | roy_r | August 30, 2019 5:57:04 PM GMT+08:00 | Refresh: Missing DICOM tags as per recommendation of DIN6862-2 Standard | WIP19 | WIP21 | N/A | Subsystem | Verified1 | Console-SRS.17972,Console-SRS.2133 | Fixed | deshbhratar_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP21SubCom | Brilliance 64 |  | WIP21 | October 14, 2019 6:58:20 PM GMT+08:00
CTUDT00626004 | Refresh | Closed | roy_r | August 30, 2019 5:57:04 PM GMT+08:00 | Refresh: Missing DICOM tags as per recommendation of DIN6862-2 Standard | WIP19 | WIP21 | N/A | Subsystem | Verified1 | Console-SRS.17972,Console-SRS.2133 | Fixed | deshbhratar_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP21SubCom | Ingenuity 64 |  | WIP21 | October 14, 2019 6:58:20 PM GMT+08:00
CTUDT00626002 | Refresh | Closed | ravindra_a | August 30, 2019 5:30:15 PM GMT+08:00 | Refresh: Installation time in scanner exceeds 90 minutes. | 4.8 | BB_WIP10 | No workaround | Subsystem |  | NA | Fixed | ravindra_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP20SubCom | Brilliance Big Bore |  | BB_WIP10 | September 9, 2019 2:26:31 PM GMT+08:00
CTUDT00626001 | Refresh | Closed | mN_m | August 30, 2019 5:30:10 PM GMT+08:00 | Refresh: Few Console Requirements need update. | WIP20 | DHF303211 Rev15 DHF316705 Rev14 | N/A | Subsystem |  | Console-SRS.8792,Console-SRS.12904,Console-SRS.15102 | Fixed | roy_r | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | SubCom | Ovation |  | DHF303211 Rev15 DHF316705 Rev14 | September 9, 2019 12:52:15 PM GMT+08:00
CTUDT00626001 | Refresh | Closed | mN_m | August 30, 2019 5:30:10 PM GMT+08:00 | Refresh: Few Console Requirements need update. | WIP20 | DHF303211 Rev15 DHF316705 Rev14 | N/A | Subsystem |  | Console-SRS.8792,Console-SRS.12904,Console-SRS.15102 | Fixed | roy_r | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | SubCom | Brilliance 64 |  | DHF303211 Rev15 DHF316705 Rev14 | September 9, 2019 12:52:15 PM GMT+08:00
CTUDT00626001 | Refresh | Closed | mN_m | August 30, 2019 5:30:10 PM GMT+08:00 | Refresh: Few Console Requirements need update. | WIP20 | DHF303211 Rev15 DHF316705 Rev14 | N/A | Subsystem |  | Console-SRS.8792,Console-SRS.12904,Console-SRS.15102 | Fixed | roy_r | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | SubCom | Brilliance Big Bore |  | DHF303211 Rev15 DHF316705 Rev14 | September 9, 2019 12:52:15 PM GMT+08:00
CTUDT00626001 | Refresh | Closed | mN_m | August 30, 2019 5:30:10 PM GMT+08:00 | Refresh: Few Console Requirements need update. | WIP20 | DHF303211 Rev15 DHF316705 Rev14 | N/A | Subsystem |  | Console-SRS.8792,Console-SRS.12904,Console-SRS.15102 | Fixed | roy_r | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | SubCom | Ingenuity 64 |  | DHF303211 Rev15 DHF316705 Rev14 | September 9, 2019 12:52:15 PM GMT+08:00
CTUDT00626001 | Refresh | Closed | mN_m | August 30, 2019 5:30:10 PM GMT+08:00 | Refresh: Few Console Requirements need update. | WIP20 | DHF303211 Rev15 DHF316705 Rev14 | N/A | Subsystem |  | Console-SRS.8792,Console-SRS.12904,Console-SRS.15102 | Fixed | roy_r | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | SubCom | IQon1 |  | DHF303211 Rev15 DHF316705 Rev14 | September 9, 2019 12:52:15 PM GMT+08:00
CTUDT00625997 | Refresh | Closed | PremKumar_Rajendran | August 30, 2019 5:17:44 PM GMT+08:00 | BB Refresh : The context menu was displayed with "Edit Auto Voice, Remove Injection & Remove Timing" options when scan was Duplicated/Pasted | BB_WIP8 |  |  | Subsystem | Context menu display |  | Duplicate | kr_m | 0 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00625995 | Refresh | Closed | mN_m | August 30, 2019 5:10:21 PM GMT+08:00 | BB Refresh: Few IVC requirements needs update. | BB_WIP9 | BB_WIP10 | N/A | Subsystem |  | Console-SRS.17755,STPS-SSRS-1390, STPS-SSRS-1352,STPS-SSRS-1376 | Fixed | 320021835 | 0 | 2 | Console Software | Console4 | NO | Closed | SubCom | Brilliance Big Bore |  | BB_WIP10 | September 10, 2019 2:16:06 PM GMT+08:00
CTUDT00625987 | Refresh | Closed | madankumar_s | August 30, 2019 1:38:04 PM GMT+08:00 | Few string was not localized as per selected language in CCT exam card for "Print and Text" in EC Sharing Tab. | WIP18 |  |  | Subsystem |  | NA | No Plan to Fix | vishnu_devi | 0 | 4 | Console Software | Console4 | NO | Closed | SubNC | Ovation
CTUDT00625987 | Refresh | Closed | madankumar_s | August 30, 2019 1:38:04 PM GMT+08:00 | Few string was not localized as per selected language in CCT exam card for "Print and Text" in EC Sharing Tab. | WIP18 |  |  | Subsystem |  | NA | No Plan to Fix | vishnu_devi | 0 | 4 | Console Software | Console4 | NO | Closed | SubNC | Brilliance 64
CTUDT00625987 | Refresh | Closed | madankumar_s | August 30, 2019 1:38:04 PM GMT+08:00 | Few string was not localized as per selected language in CCT exam card for "Print and Text" in EC Sharing Tab. | WIP18 |  |  | Subsystem |  | NA | No Plan to Fix | vishnu_devi | 0 | 4 | Console Software | Console4 | NO | Closed | SubNC | Brilliance Big Bore
CTUDT00625987 | Refresh | Closed | madankumar_s | August 30, 2019 1:38:04 PM GMT+08:00 | Few string was not localized as per selected language in CCT exam card for "Print and Text" in EC Sharing Tab. | WIP18 |  |  | Subsystem |  | NA | No Plan to Fix | vishnu_devi | 0 | 4 | Console Software | Console4 | NO | Closed | SubNC | Ingenuity 64
CTUDT00625987 | Refresh | Closed | madankumar_s | August 30, 2019 1:38:04 PM GMT+08:00 | Few string was not localized as per selected language in CCT exam card for "Print and Text" in EC Sharing Tab. | WIP18 |  |  | Subsystem |  | NA | No Plan to Fix | vishnu_devi | 0 | 4 | Console Software | Console4 | NO | Closed | SubNC | IQon1
CTUDT00625984 | Refresh | Closed | liu_bo | August 30, 2019 1:18:01 PM GMT+08:00 | ‘Sending GetReadyAns’ event with ‘Status’ set to Wrong. | WIP18 |  |  | Subsystem |  | 6698 | Works As Specified / Not A Defect | liu_bo | 0 | 2 | Gantry Software | GCX | NO | Closed | WIP20SegReq | Ovation
CTUDT00625979 | Refresh | Closed | bagdady_b | August 29, 2019 9:12:35 PM GMT+08:00 | Refresh - Unable to produce STL model from Calcium Scoring application | WIP19 |  |  | Subsystem | STL Model Saving;Blocker | CT-NM.SRS.TA.515 | Does Not Apply | tamari_o | 0 | 3 | Console USP_EBW | Support-System | NO | Closed | WIP22SegReq | Ovation
CTUDT00625979 | Refresh | Closed | bagdady_b | August 29, 2019 9:12:35 PM GMT+08:00 | Refresh - Unable to produce STL model from Calcium Scoring application | WIP19 |  |  | Subsystem | STL Model Saving;Blocker | CT-NM.SRS.TA.515 | Does Not Apply | tamari_o | 0 | 3 | Console USP_EBW | Support-System | NO | Closed | WIP22SegReq | Brilliance 64
CTUDT00625979 | Refresh | Closed | bagdady_b | August 29, 2019 9:12:35 PM GMT+08:00 | Refresh - Unable to produce STL model from Calcium Scoring application | WIP19 |  |  | Subsystem | STL Model Saving;Blocker | CT-NM.SRS.TA.515 | Does Not Apply | tamari_o | 0 | 3 | Console USP_EBW | Support-System | NO | Closed | WIP22SegReq | Brilliance Big Bore
CTUDT00625979 | Refresh | Closed | bagdady_b | August 29, 2019 9:12:35 PM GMT+08:00 | Refresh - Unable to produce STL model from Calcium Scoring application | WIP19 |  |  | Subsystem | STL Model Saving;Blocker | CT-NM.SRS.TA.515 | Does Not Apply | tamari_o | 0 | 3 | Console USP_EBW | Support-System | NO | Closed | WIP22SegReq | Ingenuity 64
CTUDT00625979 | Refresh | Closed | bagdady_b | August 29, 2019 9:12:35 PM GMT+08:00 | Refresh - Unable to produce STL model from Calcium Scoring application | WIP19 |  |  | Subsystem | STL Model Saving;Blocker | CT-NM.SRS.TA.515 | Does Not Apply | tamari_o | 0 | 3 | Console USP_EBW | Support-System | NO | Closed | WIP22SegReq | IQon1
CTUDT00625977 | Refresh | Closed | levy_s | August 29, 2019 9:00:41 PM GMT+08:00 | Refresh: Infrequently it is impossible to reach utilities when pressing the “Tools” icon. | WIP19 |  | Logout and login | Subsystem | Quality Assurance | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | Ovation
CTUDT00625977 | Refresh | Closed | levy_s | August 29, 2019 9:00:41 PM GMT+08:00 | Refresh: Infrequently it is impossible to reach utilities when pressing the “Tools” icon. | WIP19 |  | Logout and login | Subsystem | Quality Assurance | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00625977 | Refresh | Closed | levy_s | August 29, 2019 9:00:41 PM GMT+08:00 | Refresh: Infrequently it is impossible to reach utilities when pressing the “Tools” icon. | WIP19 |  | Logout and login | Subsystem | Quality Assurance | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00625977 | Refresh | Closed | levy_s | August 29, 2019 9:00:41 PM GMT+08:00 | Refresh: Infrequently it is impossible to reach utilities when pressing the “Tools” icon. | WIP19 |  | Logout and login | Subsystem | Quality Assurance | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00625977 | Refresh | Closed | levy_s | August 29, 2019 9:00:41 PM GMT+08:00 | Refresh: Infrequently it is impossible to reach utilities when pressing the “Tools” icon. | WIP19 |  | Logout and login | Subsystem | Quality Assurance | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | IQon1
CTUDT00625975 | Refresh | Closed | levy_s | August 29, 2019 8:14:33 PM GMT+08:00 | Refresh: Importing user Exam Cards got stuck at 70% . | WIP19 | WIP20 | in case of multiple exam cards, discarding an exam card in between discarding the rest of the exam cards for further process. | Subsystem | ECM | None | Fixed | datta_i | 0 | 2 | Console Software | Console4 | NO | Closed | WIP20SubCom | Ovation |  | WIP20 | September 6, 2019 7:28:07 PM GMT+08:00
CTUDT00625975 | Refresh | Closed | levy_s | August 29, 2019 8:14:33 PM GMT+08:00 | Refresh: Importing user Exam Cards got stuck at 70% . | WIP19 | WIP20 | in case of multiple exam cards, discarding an exam card in between discarding the rest of the exam cards for further process. | Subsystem | ECM | None | Fixed | datta_i | 0 | 2 | Console Software | Console4 | NO | Closed | WIP20SubCom | IQon1 |  | WIP20 | September 6, 2019 7:28:07 PM GMT+08:00
CTUDT00625973 | Refresh | Closed | 320053769 | August 29, 2019 7:38:53 PM GMT+08:00 | WIP19 – While collection bugrep “Window Command Processor” popup appears. | WIP19 |  |  | Subsystem | BugRep | NA | Works As Specified / Not A Defect | s_sa | 1 |  | Service Tools_Software | iCAP | NO | Closed |  | Ovation
CTUDT00625973 | Refresh | Closed | 320053769 | August 29, 2019 7:38:53 PM GMT+08:00 | WIP19 – While collection bugrep “Window Command Processor” popup appears. | WIP19 |  |  | Subsystem | BugRep | NA | Works As Specified / Not A Defect | s_sa | 1 |  | Service Tools_Software | iCAP | NO | Closed |  | Brilliance 64
CTUDT00625973 | Refresh | Closed | 320053769 | August 29, 2019 7:38:53 PM GMT+08:00 | WIP19 – While collection bugrep “Window Command Processor” popup appears. | WIP19 |  |  | Subsystem | BugRep | NA | Works As Specified / Not A Defect | s_sa | 1 |  | Service Tools_Software | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00625973 | Refresh | Closed | 320053769 | August 29, 2019 7:38:53 PM GMT+08:00 | WIP19 – While collection bugrep “Window Command Processor” popup appears. | WIP19 |  |  | Subsystem | BugRep | NA | Works As Specified / Not A Defect | s_sa | 1 |  | Service Tools_Software | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00625973 | Refresh | Closed | 320053769 | August 29, 2019 7:38:53 PM GMT+08:00 | WIP19 – While collection bugrep “Window Command Processor” popup appears. | WIP19 |  |  | Subsystem | BugRep | NA | Works As Specified / Not A Defect | s_sa | 1 |  | Service Tools_Software | iCAP | NO | Closed |  | IQon1
CTUDT00625967 | Refresh | Closed | Sateesh_S | August 29, 2019 6:07:18 PM GMT+08:00 | BigBore Refresh: Planned Average Scan size is getting updated after the clinical scan for Perfusion Jog exam cards. | BB_WIP8 | BB_WIP10 |  | Subsystem | average scan size; Blocker | N/A | Fixed | pandey_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP21SegReq | Brilliance Big Bore |  | BB_WIP10 | September 12, 2019 10:01:02 PM GMT+08:00
CTUDT00625964 | Refresh | Closed | bhawna_g | August 29, 2019 6:04:15 PM GMT+08:00 | Refresh: Bolus Tracking: "Please wait for the waves to be displayed” message was displayed after pressing on the “Repeat Tracker” option | 4.8 | BB_WIP11 |  | Subsystem | Bolus Tracking; TAVI | NA | Fixed | wilsonraja | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP10_SubCom; TAVI | Brilliance Big Bore |  | BB_WIP11 | September 25, 2019 6:21:22 PM GMT+08:00
CTUDT00625962 | Refresh | Closed | bhawna_g | August 29, 2019 6:02:43 PM GMT+08:00 | Cardiac/Pulmo helical recon phase tagging doesn’t match actual recon | 4.8 | BB_WIP10 |  | Subsystem | Phase tagging; TAVI | Console-SRS.3216 | Fixed | wilsonraja | 0 | 2 | Console Software | Console4 | YES | Closed | BB_WIP10_SubCom; TAVI | Brilliance Big Bore | PR9391150 | BB_WIP10 | September 17, 2019 4:58:57 PM GMT+08:00
CTUDT00625957 | Refresh | Closed | panchami_pradeep | August 29, 2019 5:08:49 PM GMT+08:00 | Refresh: C & W is not displayed in Film tab images and CTV. | WIP19 | WIP21 |  | Subsystem | Verified1 | NA | Fixed | aswani_v | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP21SegReq | Ovation |  | WIP21 | October 9, 2019 7:12:46 PM GMT+08:00
CTUDT00625957 | Refresh | Closed | panchami_pradeep | August 29, 2019 5:08:49 PM GMT+08:00 | Refresh: C & W is not displayed in Film tab images and CTV. | WIP19 | WIP21 |  | Subsystem | Verified1 | NA | Fixed | aswani_v | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP21SegReq | Brilliance 64 |  | WIP21 | October 9, 2019 7:12:46 PM GMT+08:00
CTUDT00625957 | Refresh | Closed | panchami_pradeep | August 29, 2019 5:08:49 PM GMT+08:00 | Refresh: C & W is not displayed in Film tab images and CTV. | WIP19 | WIP21 |  | Subsystem | Verified1 | NA | Fixed | aswani_v | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP21SegReq | Brilliance Big Bore |  | WIP21 | October 9, 2019 7:12:46 PM GMT+08:00
CTUDT00625957 | Refresh | Closed | panchami_pradeep | August 29, 2019 5:08:49 PM GMT+08:00 | Refresh: C & W is not displayed in Film tab images and CTV. | WIP19 | WIP21 |  | Subsystem | Verified1 | NA | Fixed | aswani_v | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP21SegReq | Ingenuity 64 |  | WIP21 | October 9, 2019 7:12:46 PM GMT+08:00
CTUDT00625957 | Refresh | Closed | panchami_pradeep | August 29, 2019 5:08:49 PM GMT+08:00 | Refresh: C & W is not displayed in Film tab images and CTV. | WIP19 | WIP21 |  | Subsystem | Verified1 | NA | Fixed | aswani_v | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP21SegReq | IQon1 |  | WIP21 | October 9, 2019 7:12:46 PM GMT+08:00
CTUDT00625955 | Refresh | Closed | ding_s | August 29, 2019 4:09:28 PM GMT+08:00 | 19 inches LCD monitor is marked as 'Not Assigned' in latest BigBore RT PCM. | BB_WIP8 | Rev 05 |  | Subsystem | PCM | CT-NM.SRS.ProductConfig.267 | Fixed | dongping_w | 0 | 3 | Design Documentation | Support-System | NO | Closed | PCM | Brilliance Big Bore |  | Rev 05 | November 7, 2019 11:22:20 AM GMT+08:00
CTUDT00625950 | Refresh | Closed | sowmya_k | August 29, 2019 3:11:15 PM GMT+08:00 | Refresh: Localization: The part of the message “The selected name already exists. Please select a new name.” was cropped | BB_WIP9 | BB_WIP10 |  | Subsystem | warning message | NA | Fixed | Prakash_Chand_Yadav | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP9_SubCom | Brilliance Big Bore |  | BB_WIP10 | September 4, 2019 9:51:15 PM GMT+08:00
CTUDT00625948 | Refresh | Closed | ding_f | August 29, 2019 2:40:04 PM GMT+08:00 | “Phantom 2 sl. 57 det.288 is out of EP limit” error occurs during Phantom Calibration. | WIP19 |  |  | Subsystem | Phantom Calibration; Blocker |  | Works As Specified / Not A Defect | braunstein_d | 0 |  |  | Support-Physics | NO | Closed |  | Ingenuity 64
CTUDT00625943 | Refresh | Closed | zhang_t | August 29, 2019 12:08:12 PM GMT+08:00 | Refresh - Wip19:: Philips_Serive user cannot save external NTP server IP via service tool. | BB_WIP9, WIP19 | BB_WIP10 WIP20 |  | Subsystem | NTP | NA | Fixed | Sandeep_Kumar | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP20SubCom | Ovation |  | BB_WIP10 WIP20 | September 13, 2019 7:19:15 PM GMT+08:00
CTUDT00625943 | Refresh | Closed | zhang_t | August 29, 2019 12:08:12 PM GMT+08:00 | Refresh - Wip19:: Philips_Serive user cannot save external NTP server IP via service tool. | BB_WIP9, WIP19 | BB_WIP10 WIP20 |  | Subsystem | NTP | NA | Fixed | Sandeep_Kumar | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP20SubCom | Brilliance 64 |  | BB_WIP10 WIP20 | September 13, 2019 7:19:15 PM GMT+08:00
CTUDT00625943 | Refresh | Closed | zhang_t | August 29, 2019 12:08:12 PM GMT+08:00 | Refresh - Wip19:: Philips_Serive user cannot save external NTP server IP via service tool. | BB_WIP9, WIP19 | BB_WIP10 WIP20 |  | Subsystem | NTP | NA | Fixed | Sandeep_Kumar | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP20SubCom | Brilliance Big Bore |  | BB_WIP10 WIP20 | September 13, 2019 7:19:15 PM GMT+08:00
CTUDT00625943 | Refresh | Closed | zhang_t | August 29, 2019 12:08:12 PM GMT+08:00 | Refresh - Wip19:: Philips_Serive user cannot save external NTP server IP via service tool. | BB_WIP9, WIP19 | BB_WIP10 WIP20 |  | Subsystem | NTP | NA | Fixed | Sandeep_Kumar | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP20SubCom | Ingenuity 64 |  | BB_WIP10 WIP20 | September 13, 2019 7:19:15 PM GMT+08:00
CTUDT00625941 | Refresh | Closed | liu_bo | August 29, 2019 11:10:19 AM GMT+08:00 | Use Gate test tool, execute script failed. | WIP18 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | bruce_li | 0 |  | Gantry Software | GCX | NO | Closed |  | Ovation
CTUDT00625939 | Refresh | Closed | cohen_y | August 28, 2019 8:31:02 PM GMT+08:00 | Refresh: Calibrations: Crosstalk Test: Crosstalk Z fails due to Slice 1 values | BB_WIP9 | WIP20 | N/A | Subsystem | crosstalk test | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | BB_WIP10_SubCom | Brilliance Big Bore |  | WIP20 | September 9, 2019 6:49:25 PM GMT+08:00
CTUDT00625935 | Refresh | Closed | golzban_v | August 28, 2019 6:20:27 PM GMT+08:00 | Refresh: The system allows planning tracker sequence with cycle time smaller than 2*rotation time | WIP19 | WIP21 BB_WIP11 | Set rotation time nanually | Subsystem | General Bug; Blocker | NA | Fixed | datta_i | 0 | 3 | Gantry Software | Console4 | NO | Closed | WIP21SubCom | Ovation |  | WIP21 BB_WIP11 | October 16, 2019 8:27:25 PM GMT+08:00
CTUDT00625935 | Refresh | Closed | golzban_v | August 28, 2019 6:20:27 PM GMT+08:00 | Refresh: The system allows planning tracker sequence with cycle time smaller than 2*rotation time | WIP19 | WIP21 BB_WIP11 | Set rotation time nanually | Subsystem | General Bug; Blocker | NA | Fixed | datta_i | 0 | 3 | Gantry Software | Console4 | NO | Closed | WIP21SubCom | Brilliance 64 |  | WIP21 BB_WIP11 | October 16, 2019 8:27:25 PM GMT+08:00
CTUDT00625935 | Refresh | Closed | golzban_v | August 28, 2019 6:20:27 PM GMT+08:00 | Refresh: The system allows planning tracker sequence with cycle time smaller than 2*rotation time | WIP19 | WIP21 BB_WIP11 | Set rotation time nanually | Subsystem | General Bug; Blocker | NA | Fixed | datta_i | 0 | 3 | Gantry Software | Console4 | NO | Closed | WIP21SubCom | Brilliance Big Bore |  | WIP21 BB_WIP11 | October 16, 2019 8:27:25 PM GMT+08:00
CTUDT00625935 | Refresh | Closed | golzban_v | August 28, 2019 6:20:27 PM GMT+08:00 | Refresh: The system allows planning tracker sequence with cycle time smaller than 2*rotation time | WIP19 | WIP21 BB_WIP11 | Set rotation time nanually | Subsystem | General Bug; Blocker | NA | Fixed | datta_i | 0 | 3 | Gantry Software | Console4 | NO | Closed | WIP21SubCom | Ingenuity 64 |  | WIP21 BB_WIP11 | October 16, 2019 8:27:25 PM GMT+08:00
CTUDT00625935 | Refresh | Closed | golzban_v | August 28, 2019 6:20:27 PM GMT+08:00 | Refresh: The system allows planning tracker sequence with cycle time smaller than 2*rotation time | WIP19 | WIP21 BB_WIP11 | Set rotation time nanually | Subsystem | General Bug; Blocker | NA | Fixed | datta_i | 0 | 3 | Gantry Software | Console4 | NO | Closed | WIP21SubCom | IQon1 |  | WIP21 BB_WIP11 | October 16, 2019 8:27:25 PM GMT+08:00
CTUDT00625932 | Refresh | Closed | golzban_v | August 28, 2019 6:09:55 PM GMT+08:00 | Refresh: The system allows planning a bolus tracking sequence with minimum post threshold delay smaller than 2 seconds. | WIP19 | DHF298554 Rev09 DHF298553 Rev.09 DHF298556 Rev.10 |  | Subsystem | Bolus Tracking | CT-NM.SRS.Modes.1810 | Fixed | bandil_j | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | SubCom-19-Dec-2019 | Ovation |  | DHF298554 Rev09 DHF298553 Rev.09 DHF298556 Rev.10 | December 20, 2019 3:48:42 AM GMT+08:00
CTUDT00625932 | Refresh | Closed | golzban_v | August 28, 2019 6:09:55 PM GMT+08:00 | Refresh: The system allows planning a bolus tracking sequence with minimum post threshold delay smaller than 2 seconds. | WIP19 | DHF298554 Rev09 DHF298553 Rev.09 DHF298556 Rev.10 |  | Subsystem | Bolus Tracking | CT-NM.SRS.Modes.1810 | Fixed | bandil_j | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | SubCom-19-Dec-2019 | IQon1 |  | DHF298554 Rev09 DHF298553 Rev.09 DHF298556 Rev.10 | December 20, 2019 3:48:42 AM GMT+08:00
CTUDT00625928 | Refresh | Closed | shterner_a | August 28, 2019 4:00:59 PM GMT+08:00 | Pre Val: Refresh PRS: CRS.6933 and CRS.6934 missing tag to iCT iCT SP | WIP19 | DHF298293 IQon PRS Rev.08 and DHF298290 iCT PRS Rev.07 ING, BR64 CRS DHF316435 Rev03 |  | Subsystem | Verified 1 | CT-NM.PRS.CRS.6933, 6934 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP20SegReq | Ovation |  | DHF298293 IQon PRS Rev.08 and DHF298290 iCT PRS Rev.07 ING, BR64 CRS DHF316435 Rev03 | November 6, 2019 8:35:10 AM GMT+08:00
CTUDT00625928 | Refresh | Closed | shterner_a | August 28, 2019 4:00:59 PM GMT+08:00 | Pre Val: Refresh PRS: CRS.6933 and CRS.6934 missing tag to iCT iCT SP | WIP19 | DHF298293 IQon PRS Rev.08 and DHF298290 iCT PRS Rev.07 ING, BR64 CRS DHF316435 Rev03 |  | Subsystem | Verified 1 | CT-NM.PRS.CRS.6933, 6934 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP20SegReq | Brilliance 64 |  | DHF298293 IQon PRS Rev.08 and DHF298290 iCT PRS Rev.07 ING, BR64 CRS DHF316435 Rev03 | November 6, 2019 8:35:10 AM GMT+08:00
CTUDT00625928 | Refresh | Closed | shterner_a | August 28, 2019 4:00:59 PM GMT+08:00 | Pre Val: Refresh PRS: CRS.6933 and CRS.6934 missing tag to iCT iCT SP | WIP19 | DHF298293 IQon PRS Rev.08 and DHF298290 iCT PRS Rev.07 ING, BR64 CRS DHF316435 Rev03 |  | Subsystem | Verified 1 | CT-NM.PRS.CRS.6933, 6934 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP20SegReq | Ingenuity 64 |  | DHF298293 IQon PRS Rev.08 and DHF298290 iCT PRS Rev.07 ING, BR64 CRS DHF316435 Rev03 | November 6, 2019 8:35:10 AM GMT+08:00
CTUDT00625920 | Refresh | Closed | jose_s | August 28, 2019 12:31:04 PM GMT+08:00 | Refresh: some exam card import failed if import too much by one time | BB_WIP9 | BB_WIP10 |  | Subsystem | Exam card | None | Fixed | deshbhratar_m | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP9_SubCom | Brilliance Big Bore |  | BB_WIP10 | September 6, 2019 7:20:46 PM GMT+08:00
CTUDT00625918 | Refresh | Closed | ding_f | August 28, 2019 10:38:17 AM GMT+08:00 | Refresh: Image Rotation images with obvious Ring artifacts. | WIP19/WIP9 | WIP20 |  | Subsystem | calibration;Blocker20 | N/A | Fixed | zhou_x | 0 | 2 | Gantry Calibrations | GCX | NO | Closed | WIP20SubCom | Brilliance 64 |  | WIP20 | September 3, 2019 4:27:39 PM GMT+08:00
CTUDT00625918 | Refresh | Closed | ding_f | August 28, 2019 10:38:17 AM GMT+08:00 | Refresh: Image Rotation images with obvious Ring artifacts. | WIP19/WIP9 | WIP20 |  | Subsystem | calibration;Blocker20 | N/A | Fixed | zhou_x | 0 | 2 | Gantry Calibrations | GCX | NO | Closed | WIP20SubCom | Brilliance Big Bore |  | WIP20 | September 3, 2019 4:27:39 PM GMT+08:00
CTUDT00625918 | Refresh | Closed | ding_f | August 28, 2019 10:38:17 AM GMT+08:00 | Refresh: Image Rotation images with obvious Ring artifacts. | WIP19/WIP9 | WIP20 |  | Subsystem | calibration;Blocker20 | N/A | Fixed | zhou_x | 0 | 2 | Gantry Calibrations | GCX | NO | Closed | WIP20SubCom | Ingenuity 64 |  | WIP20 | September 3, 2019 4:27:39 PM GMT+08:00
CTUDT00625916 | Refresh | Closed | morton_t | August 28, 2019 3:51:53 AM GMT+08:00 | Refresh: Acceptance: Incorrect Uniformity limits for Ingenuity scanner and Legacy Phantom | WIP19 | WIP20 | N/A | Subsystem | Calibration; Blocker20 | DHF170395 | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP20SubCom | Brilliance 64 |  | WIP20 | September 12, 2019 7:42:41 PM GMT+08:00
CTUDT00625916 | Refresh | Closed | morton_t | August 28, 2019 3:51:53 AM GMT+08:00 | Refresh: Acceptance: Incorrect Uniformity limits for Ingenuity scanner and Legacy Phantom | WIP19 | WIP20 | N/A | Subsystem | Calibration; Blocker20 | DHF170395 | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP20SubCom | Ingenuity 64 |  | WIP20 | September 12, 2019 7:42:41 PM GMT+08:00
CTUDT00625912 | Refresh | Closed | wang_l | August 27, 2019 9:17:35 PM GMT+08:00 | Refresh: "A Problem has occurred, please retry. " error displayed during load patient | BB_WIP9 |  |  | Subsystem | Pateient loading | N/A | Works As Specified / Not A Defect | wang_l | 1 |  |  | Support-System | NO | Closed |  | Brilliance Big Bore
CTUDT00625910 | Refresh | Closed | bagdady_b | August 27, 2019 8:28:11 PM GMT+08:00 | Refresh - Change Patient Details breaks the perfusion volume separation | WIP19 | WIP22.2 | Use original images | Subsystem |  | NA | Fixed | srushti_k | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP22.2SubCom | Ovation |  | WIP22.2 | December 2, 2019 1:29:36 PM GMT+08:00
CTUDT00625910 | Refresh | Closed | bagdady_b | August 27, 2019 8:28:11 PM GMT+08:00 | Refresh - Change Patient Details breaks the perfusion volume separation | WIP19 | WIP22.2 | Use original images | Subsystem |  | NA | Fixed | srushti_k | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP22.2SubCom | Brilliance 64 |  | WIP22.2 | December 2, 2019 1:29:36 PM GMT+08:00
CTUDT00625910 | Refresh | Closed | bagdady_b | August 27, 2019 8:28:11 PM GMT+08:00 | Refresh - Change Patient Details breaks the perfusion volume separation | WIP19 | WIP22.2 | Use original images | Subsystem |  | NA | Fixed | srushti_k | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP22.2SubCom | Brilliance Big Bore |  | WIP22.2 | December 2, 2019 1:29:36 PM GMT+08:00
CTUDT00625910 | Refresh | Closed | bagdady_b | August 27, 2019 8:28:11 PM GMT+08:00 | Refresh - Change Patient Details breaks the perfusion volume separation | WIP19 | WIP22.2 | Use original images | Subsystem |  | NA | Fixed | srushti_k | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP22.2SubCom | Ingenuity 64 |  | WIP22.2 | December 2, 2019 1:29:36 PM GMT+08:00
CTUDT00625910 | Refresh | Closed | bagdady_b | August 27, 2019 8:28:11 PM GMT+08:00 | Refresh - Change Patient Details breaks the perfusion volume separation | WIP19 | WIP22.2 | Use original images | Subsystem |  | NA | Fixed | srushti_k | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP22.2SubCom | IQon1 |  | WIP22.2 | December 2, 2019 1:29:36 PM GMT+08:00
CTUDT00625909 | Refresh | Closed | datta_i | August 27, 2019 6:25:34 PM GMT+08:00 | Refresh - Thickness field is editable during CCT Continuous scan with display set to 3 | 4.8 | BB_WIP10 |  | Subsystem |  | NA | Fixed | datta_i | 0 | 2 | Console Software | Console4 | YES | Closed | BB_WIP10_SubCom | Brilliance Big Bore |  | BB_WIP10 | September 5, 2019 12:41:32 PM GMT+08:00
CTUDT00625905 | Refresh | Closed | wang_l | August 27, 2019 4:43:29 PM GMT+08:00 | Refresh Ingenuity: DFS Position hold on about 3~5 minutes during calibration and then recover itself | WIP19 | WIP20 |  | Subsystem | Calibration;Performance | N/A | Fixed | frider_o | 0 | 3 | DMS Software | DMS | NO | Closed | WIP20SubCom | Ingenuity 64 |  | WIP20 | September 8, 2019 7:42:37 PM GMT+08:00
CTUDT00625903 | Refresh | Closed | Sateesh_S | August 27, 2019 3:19:08 PM GMT+08:00 | BigBore Refresh: Planned and Actual SSDE and Average scan size values are not showing for extended scan in acquisition viewer | BB_WIP8 | BB_WIP10 |  | Subsystem |  | N/A | Fixed | pandey_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP20SegReq | Brilliance Big Bore |  | BB_WIP10 | September 4, 2019 8:27:59 PM GMT+08:00
CTUDT00625900 | Refresh | Closed | 320053768 | August 26, 2019 10:30:05 PM GMT+08:00 | Constancy test failure after “object is currently in use elsewhere” error. | WIP19 |  | Run the calibration without opening the image viewer | Subsystem | Release Notes | n/a | No Plan to Fix | aswani_v | 0 | 3 |  | iCAP | NO | Closed |  | Ovation
CTUDT00625900 | Refresh | Closed | 320053768 | August 26, 2019 10:30:05 PM GMT+08:00 | Constancy test failure after “object is currently in use elsewhere” error. | WIP19 |  | Run the calibration without opening the image viewer | Subsystem | Release Notes | n/a | No Plan to Fix | aswani_v | 0 | 3 |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00625900 | Refresh | Closed | 320053768 | August 26, 2019 10:30:05 PM GMT+08:00 | Constancy test failure after “object is currently in use elsewhere” error. | WIP19 |  | Run the calibration without opening the image viewer | Subsystem | Release Notes | n/a | No Plan to Fix | aswani_v | 0 | 3 |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00625900 | Refresh | Closed | 320053768 | August 26, 2019 10:30:05 PM GMT+08:00 | Constancy test failure after “object is currently in use elsewhere” error. | WIP19 |  | Run the calibration without opening the image viewer | Subsystem | Release Notes | n/a | No Plan to Fix | aswani_v | 0 | 3 |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00625900 | Refresh | Closed | 320053768 | August 26, 2019 10:30:05 PM GMT+08:00 | Constancy test failure after “object is currently in use elsewhere” error. | WIP19 |  | Run the calibration without opening the image viewer | Subsystem | Release Notes | n/a | No Plan to Fix | aswani_v | 0 | 3 |  | iCAP | NO | Closed |  | IQon1
CTUDT00625898 | Refresh | Closed | 320053768 | August 26, 2019 10:26:10 PM GMT+08:00 | Quality assurance crashed while performing “performance” test on iCTSP. | WIP19 |  |  | Subsystem | IQ system performance test;Blocker20 |  | Duplicate | altshuller_mark | 0 |  |  | DMS | NO | Closed |  | Ovation
CTUDT00625894 | Refresh | Closed | 320053768 | August 26, 2019 9:31:16 PM GMT+08:00 | “UHR focal spot x position” and “focal spot x position” calibrations failure (incorrect scan combination) | WIP19 | WIP20 |  | Subsystem | Gantry Calib;Blocker20 | N/A | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP20SegReq | Ovation |  | WIP20 | September 10, 2019 10:19:38 PM GMT+08:00
CTUDT00625892 | Refresh | Closed | 320053768 | August 26, 2019 9:05:28 PM GMT+08:00 | “UHR focal spot x position” calibration failure (incorrect parameter in XML file for Harmonized Phantom) | WIP19 | WIP20 |  | Subsystem | Gantry Calib;Blocker20 | N/A | Fixed | zhou_x | 0 | 2 | Gantry Calibrations | GCX | NO | Closed | WIP20SubCom | Ovation |  | WIP20 | September 10, 2019 7:09:06 PM GMT+08:00
CTUDT00625889 | Refresh | Closed | Noy_Seelig | August 26, 2019 8:24:42 PM GMT+08:00 | Refresh: Pre WIP 19 Testing: Tube Data appeared twice in HTML Report in Performance Test | BB_WIP8 | WIP21 |  | Subsystem | Verified1 | n/a | Fixed | frider_o | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP21SubCom | Ovation |  | WIP21 | September 24, 2019 8:17:09 PM GMT+08:00
CTUDT00625889 | Refresh | Closed | Noy_Seelig | August 26, 2019 8:24:42 PM GMT+08:00 | Refresh: Pre WIP 19 Testing: Tube Data appeared twice in HTML Report in Performance Test | BB_WIP8 | WIP21 |  | Subsystem | Verified1 | n/a | Fixed | frider_o | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP21SubCom | Brilliance 64 |  | WIP21 | September 24, 2019 8:17:09 PM GMT+08:00
CTUDT00625889 | Refresh | Closed | Noy_Seelig | August 26, 2019 8:24:42 PM GMT+08:00 | Refresh: Pre WIP 19 Testing: Tube Data appeared twice in HTML Report in Performance Test | BB_WIP8 | WIP21 |  | Subsystem | Verified1 | n/a | Fixed | frider_o | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP21SubCom | Brilliance Big Bore |  | WIP21 | September 24, 2019 8:17:09 PM GMT+08:00
CTUDT00625889 | Refresh | Closed | Noy_Seelig | August 26, 2019 8:24:42 PM GMT+08:00 | Refresh: Pre WIP 19 Testing: Tube Data appeared twice in HTML Report in Performance Test | BB_WIP8 | WIP21 |  | Subsystem | Verified1 | n/a | Fixed | frider_o | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP21SubCom | Ingenuity 64 |  | WIP21 | September 24, 2019 8:17:09 PM GMT+08:00
CTUDT00625889 | Refresh | Closed | Noy_Seelig | August 26, 2019 8:24:42 PM GMT+08:00 | Refresh: Pre WIP 19 Testing: Tube Data appeared twice in HTML Report in Performance Test | BB_WIP8 | WIP21 |  | Subsystem | Verified1 | n/a | Fixed | frider_o | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP21SubCom | IQon1 |  | WIP21 | September 24, 2019 8:17:09 PM GMT+08:00
CTUDT00625887 | Refresh | Closed | finestine_m | August 26, 2019 7:42:04 PM GMT+08:00 | Refresh: Beam and DMS Check: Wrong Limits in "Normalized scan to air ratio per sector k”. | WIP18 | WIP20 |  | Subsystem | DMS Calib; Verified1 | N/A | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP20SubCom | Ovation |  | WIP20 | September 9, 2019 8:20:52 PM GMT+08:00
CTUDT00625887 | Refresh | Closed | finestine_m | August 26, 2019 7:42:04 PM GMT+08:00 | Refresh: Beam and DMS Check: Wrong Limits in "Normalized scan to air ratio per sector k”. | WIP18 | WIP20 |  | Subsystem | DMS Calib; Verified1 | N/A | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP20SubCom | Brilliance Big Bore |  | WIP20 | September 9, 2019 8:20:52 PM GMT+08:00
CTUDT00625887 | Refresh | Closed | finestine_m | August 26, 2019 7:42:04 PM GMT+08:00 | Refresh: Beam and DMS Check: Wrong Limits in "Normalized scan to air ratio per sector k”. | WIP18 | WIP20 |  | Subsystem | DMS Calib; Verified1 | N/A | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP20SubCom | Ingenuity 64 |  | WIP20 | September 9, 2019 8:20:52 PM GMT+08:00
CTUDT00625887 | Refresh | Closed | finestine_m | August 26, 2019 7:42:04 PM GMT+08:00 | Refresh: Beam and DMS Check: Wrong Limits in "Normalized scan to air ratio per sector k”. | WIP18 | WIP20 |  | Subsystem | DMS Calib; Verified1 | N/A | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP20SubCom | IQon1 |  | WIP20 | September 9, 2019 8:20:52 PM GMT+08:00
CTUDT00625885 | Refresh | Closed | mN_m | August 26, 2019 6:21:36 PM GMT+08:00 | BB Refresh: Just a single press and release of “Go to Saved position” button takes the couch to the saved position and it doesn’t stops in the middle. | BB_WIP9 |  |  | Subsystem | Couch movement | NA | Works As Specified / Not A Defect | cx_zhang | 1 |  |  | GCX | NO | Closed |  | Brilliance Big Bore
CTUDT00625883 | Refresh | Closed | mN_m | August 26, 2019 6:13:24 PM GMT+08:00 | BBRefresh: "Go to Saved position" button is not functional and keeps blinking and couch doesn't move when the couch is at extreme position. | BB_WIP9 | BB_WIP11 |  | Subsystem |  | NA | Fixed | Snehashis_Dey | 0 | 2 | IVC Software | Console4 | YES | Closed | BB_WIP11_SubCom | Brilliance Big Bore |  | BB_WIP11 | October 10, 2019 8:50:47 PM GMT+08:00
CTUDT00625880 | Refresh | Closed | zhang_t | August 26, 2019 2:27:26 PM GMT+08:00 | Refresh IVC: cannot use service tool -> connect to controller to download GHOST CPLD | BB_WIP9 | BB_WIP10 WIP20 | N/A | Subsystem |  | N/A | Fixed | Mark_Huang | 0 | 2 | Gantry Software | GCX | NO | Closed | WIP20SegReq | Ovation |  | BB_WIP10 WIP20 | September 4, 2019 1:08:38 PM GMT+08:00
CTUDT00625880 | Refresh | Closed | zhang_t | August 26, 2019 2:27:26 PM GMT+08:00 | Refresh IVC: cannot use service tool -> connect to controller to download GHOST CPLD | BB_WIP9 | BB_WIP10 WIP20 | N/A | Subsystem |  | N/A | Fixed | Mark_Huang | 0 | 2 | Gantry Software | GCX | NO | Closed | WIP20SegReq | Brilliance 64 |  | BB_WIP10 WIP20 | September 4, 2019 1:08:38 PM GMT+08:00
CTUDT00625880 | Refresh | Closed | zhang_t | August 26, 2019 2:27:26 PM GMT+08:00 | Refresh IVC: cannot use service tool -> connect to controller to download GHOST CPLD | BB_WIP9 | BB_WIP10 WIP20 | N/A | Subsystem |  | N/A | Fixed | Mark_Huang | 0 | 2 | Gantry Software | GCX | NO | Closed | WIP20SegReq | Brilliance Big Bore |  | BB_WIP10 WIP20 | September 4, 2019 1:08:38 PM GMT+08:00
CTUDT00625880 | Refresh | Closed | zhang_t | August 26, 2019 2:27:26 PM GMT+08:00 | Refresh IVC: cannot use service tool -> connect to controller to download GHOST CPLD | BB_WIP9 | BB_WIP10 WIP20 | N/A | Subsystem |  | N/A | Fixed | Mark_Huang | 0 | 2 | Gantry Software | GCX | NO | Closed | WIP20SegReq | Ingenuity 64 |  | BB_WIP10 WIP20 | September 4, 2019 1:08:38 PM GMT+08:00
CTUDT00625880 | Refresh | Closed | zhang_t | August 26, 2019 2:27:26 PM GMT+08:00 | Refresh IVC: cannot use service tool -> connect to controller to download GHOST CPLD | BB_WIP9 | BB_WIP10 WIP20 | N/A | Subsystem |  | N/A | Fixed | Mark_Huang | 0 | 2 | Gantry Software | GCX | NO | Closed | WIP20SegReq | IQon1 |  | BB_WIP10 WIP20 | September 4, 2019 1:08:38 PM GMT+08:00
CTUDT00625879 | Refresh | Closed | azagury_d | August 26, 2019 2:25:04 PM GMT+08:00 | Refresh: Visual Diagnostics: After Module test fails, FRU tab remains empty. | WIP18 | WIP20 |  | Subsystem | Visual Diagnostic | DMS-SRS-475 | Fixed | kovalev_a | 0 | 2 | DMS Software | DMS | NO | Closed | WIP21SegReq | Ovation |  | WIP20 | September 19, 2019 2:24:23 PM GMT+08:00
CTUDT00625875 | Refresh | Closed | cohen_y | August 26, 2019 12:59:39 PM GMT+08:00 | Refresh: Localization: ‘IQ Check’ is translated in Beam and DMS Cracked Wedge warning message | Pre WIP19 | WIP20 |  | Subsystem | Localization; Verified1 | n/a | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP20SubCom | Ovation |  | WIP20 | September 12, 2019 7:41:08 PM GMT+08:00
CTUDT00625875 | Refresh | Closed | cohen_y | August 26, 2019 12:59:39 PM GMT+08:00 | Refresh: Localization: ‘IQ Check’ is translated in Beam and DMS Cracked Wedge warning message | Pre WIP19 | WIP20 |  | Subsystem | Localization; Verified1 | n/a | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP20SubCom | Brilliance 64 |  | WIP20 | September 12, 2019 7:41:08 PM GMT+08:00
CTUDT00625875 | Refresh | Closed | cohen_y | August 26, 2019 12:59:39 PM GMT+08:00 | Refresh: Localization: ‘IQ Check’ is translated in Beam and DMS Cracked Wedge warning message | Pre WIP19 | WIP20 |  | Subsystem | Localization; Verified1 | n/a | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP20SubCom | Brilliance Big Bore |  | WIP20 | September 12, 2019 7:41:08 PM GMT+08:00
CTUDT00625875 | Refresh | Closed | cohen_y | August 26, 2019 12:59:39 PM GMT+08:00 | Refresh: Localization: ‘IQ Check’ is translated in Beam and DMS Cracked Wedge warning message | Pre WIP19 | WIP20 |  | Subsystem | Localization; Verified1 | n/a | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP20SubCom | Ingenuity 64 |  | WIP20 | September 12, 2019 7:41:08 PM GMT+08:00
CTUDT00625875 | Refresh | Closed | cohen_y | August 26, 2019 12:59:39 PM GMT+08:00 | Refresh: Localization: ‘IQ Check’ is translated in Beam and DMS Cracked Wedge warning message | Pre WIP19 | WIP20 |  | Subsystem | Localization; Verified1 | n/a | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP20SubCom | IQon1 |  | WIP20 | September 12, 2019 7:41:08 PM GMT+08:00
CTUDT00625873 | Refresh | Closed | cohen_y | August 26, 2019 12:55:43 PM GMT+08:00 | Refresh: Localization: Therapy top instructions are cut | Pre WIP19 | wip22 |  | Subsystem | Verified1 | DMS-SRS-1206 | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP22SubCom | Ovation |  | wip22 | November 5, 2019 4:44:23 PM GMT+08:00
CTUDT00625873 | Refresh | Closed | cohen_y | August 26, 2019 12:55:43 PM GMT+08:00 | Refresh: Localization: Therapy top instructions are cut | Pre WIP19 | wip22 |  | Subsystem | Verified1 | DMS-SRS-1206 | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP22SubCom | Brilliance 64 |  | wip22 | November 5, 2019 4:44:23 PM GMT+08:00
CTUDT00625873 | Refresh | Closed | cohen_y | August 26, 2019 12:55:43 PM GMT+08:00 | Refresh: Localization: Therapy top instructions are cut | Pre WIP19 | wip22 |  | Subsystem | Verified1 | DMS-SRS-1206 | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP22SubCom | Brilliance Big Bore |  | wip22 | November 5, 2019 4:44:23 PM GMT+08:00
CTUDT00625873 | Refresh | Closed | cohen_y | August 26, 2019 12:55:43 PM GMT+08:00 | Refresh: Localization: Therapy top instructions are cut | Pre WIP19 | wip22 |  | Subsystem | Verified1 | DMS-SRS-1206 | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP22SubCom | Ingenuity 64 |  | wip22 | November 5, 2019 4:44:23 PM GMT+08:00
CTUDT00625873 | Refresh | Closed | cohen_y | August 26, 2019 12:55:43 PM GMT+08:00 | Refresh: Localization: Therapy top instructions are cut | Pre WIP19 | wip22 |  | Subsystem | Verified1 | DMS-SRS-1206 | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP22SubCom | IQon1 |  | wip22 | November 5, 2019 4:44:23 PM GMT+08:00
CTUDT00625872 | Refresh | Closed | wang_l | August 26, 2019 12:55:03 PM GMT+08:00 | Refresh: CT-NM.SRS.Service.1130 should be removed since URL Restriction tool was no longer supported | WIP19 | ING, BR64 SRS DHF298554 Rev08 iCT, SRS DHF298553 Rev.08 |  | Subsystem | Verified1 | SRS.Service.1130 | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | SRS next version | Ovation |  | ING, BR64 SRS DHF298554 Rev08 iCT, SRS DHF298553 Rev.08 | November 6, 2019 2:06:03 PM GMT+08:00
CTUDT00625872 | Refresh | Closed | wang_l | August 26, 2019 12:55:03 PM GMT+08:00 | Refresh: CT-NM.SRS.Service.1130 should be removed since URL Restriction tool was no longer supported | WIP19 | ING, BR64 SRS DHF298554 Rev08 iCT, SRS DHF298553 Rev.08 |  | Subsystem | Verified1 | SRS.Service.1130 | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | SRS next version | Brilliance 64 |  | ING, BR64 SRS DHF298554 Rev08 iCT, SRS DHF298553 Rev.08 | November 6, 2019 2:06:03 PM GMT+08:00
CTUDT00625872 | Refresh | Closed | wang_l | August 26, 2019 12:55:03 PM GMT+08:00 | Refresh: CT-NM.SRS.Service.1130 should be removed since URL Restriction tool was no longer supported | WIP19 | ING, BR64 SRS DHF298554 Rev08 iCT, SRS DHF298553 Rev.08 |  | Subsystem | Verified1 | SRS.Service.1130 | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | SRS next version | Ingenuity 64 |  | ING, BR64 SRS DHF298554 Rev08 iCT, SRS DHF298553 Rev.08 | November 6, 2019 2:06:03 PM GMT+08:00
CTUDT00625870 | Refresh | Closed | cohen_y | August 26, 2019 12:50:42 PM GMT+08:00 | Refresh: Localization: Beam and DMS Check messages are translated in Chinese when the interference language is Norwegian | Pre WIP19 | WIP20 |  | Subsystem | Verified1 | n/a | Fixed | frider_o | 0 | 2 | Couch Software | DMS | NO | Closed | WIP20SubCom | Ovation |  | WIP20 | September 4, 2019 5:50:16 PM GMT+08:00
CTUDT00625870 | Refresh | Closed | cohen_y | August 26, 2019 12:50:42 PM GMT+08:00 | Refresh: Localization: Beam and DMS Check messages are translated in Chinese when the interference language is Norwegian | Pre WIP19 | WIP20 |  | Subsystem | Verified1 | n/a | Fixed | frider_o | 0 | 2 | Couch Software | DMS | NO | Closed | WIP20SubCom | Brilliance 64 |  | WIP20 | September 4, 2019 5:50:16 PM GMT+08:00
CTUDT00625870 | Refresh | Closed | cohen_y | August 26, 2019 12:50:42 PM GMT+08:00 | Refresh: Localization: Beam and DMS Check messages are translated in Chinese when the interference language is Norwegian | Pre WIP19 | WIP20 |  | Subsystem | Verified1 | n/a | Fixed | frider_o | 0 | 2 | Couch Software | DMS | NO | Closed | WIP20SubCom | Brilliance Big Bore |  | WIP20 | September 4, 2019 5:50:16 PM GMT+08:00
CTUDT00625870 | Refresh | Closed | cohen_y | August 26, 2019 12:50:42 PM GMT+08:00 | Refresh: Localization: Beam and DMS Check messages are translated in Chinese when the interference language is Norwegian | Pre WIP19 | WIP20 |  | Subsystem | Verified1 | n/a | Fixed | frider_o | 0 | 2 | Couch Software | DMS | NO | Closed | WIP20SubCom | Ingenuity 64 |  | WIP20 | September 4, 2019 5:50:16 PM GMT+08:00
CTUDT00625870 | Refresh | Closed | cohen_y | August 26, 2019 12:50:42 PM GMT+08:00 | Refresh: Localization: Beam and DMS Check messages are translated in Chinese when the interference language is Norwegian | Pre WIP19 | WIP20 |  | Subsystem | Verified1 | n/a | Fixed | frider_o | 0 | 2 | Couch Software | DMS | NO | Closed | WIP20SubCom | IQon1 |  | WIP20 | September 4, 2019 5:50:16 PM GMT+08:00
CTUDT00625868 | Refresh | Closed | cohen_y | August 26, 2019 12:44:09 PM GMT+08:00 | Refresh: Localization: Some fields are not translated in the Report Summary of Constancy Test and IQ Check | Pre WIP19 |  |  | Subsystem | Localization | n/a | No Plan to Fix | frider_o | 0 | 4 |  | DMS | NO | Closed | SubNC | Ovation
CTUDT00625868 | Refresh | Closed | cohen_y | August 26, 2019 12:44:09 PM GMT+08:00 | Refresh: Localization: Some fields are not translated in the Report Summary of Constancy Test and IQ Check | Pre WIP19 |  |  | Subsystem | Localization | n/a | No Plan to Fix | frider_o | 0 | 4 |  | DMS | NO | Closed | SubNC | Brilliance Big Bore
CTUDT00625868 | Refresh | Closed | cohen_y | August 26, 2019 12:44:09 PM GMT+08:00 | Refresh: Localization: Some fields are not translated in the Report Summary of Constancy Test and IQ Check | Pre WIP19 |  |  | Subsystem | Localization | n/a | No Plan to Fix | frider_o | 0 | 4 |  | DMS | NO | Closed | SubNC | Ingenuity 64
CTUDT00625868 | Refresh | Closed | cohen_y | August 26, 2019 12:44:09 PM GMT+08:00 | Refresh: Localization: Some fields are not translated in the Report Summary of Constancy Test and IQ Check | Pre WIP19 |  |  | Subsystem | Localization | n/a | No Plan to Fix | frider_o | 0 | 4 |  | DMS | NO | Closed | SubNC | IQon1
CTUDT00625866 | Refresh | Closed | zhang_t | August 26, 2019 10:41:39 AM GMT+08:00 | Refresh: cannot open iDMS Utility and DMC Utility via service tool | BB_WIP9 |  |  | Subsystem |  |  | Duplicate | Sandeep_Kumar | 0 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00625864 | Refresh | Closed | Noy_Seelig | August 25, 2019 9:19:06 PM GMT+08:00 | Refresh: Performance Test: ROI appears as ellipse shape instead round, wrong measurement. | WIP18 | WIP20 | N/A | Subsystem | Performance Test;Blocker20 | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP20SubCom | IQon1 |  | WIP20 | August 28, 2019 4:51:38 PM GMT+08:00
CTUDT00625863 | Refresh | Closed | altshuller_mark | August 25, 2019 8:10:01 PM GMT+08:00 | “PhysicsQualityAssurance has stopped working” Pop Up during Perform DFS Calibration | WIP19 |  | N/A | Subsystem |  |  | Duplicate | yatziv_l | 0 | 2 |  | DMS | NO | Closed |  | Ovation
CTUDT00625863 | Refresh | Closed | altshuller_mark | August 25, 2019 8:10:01 PM GMT+08:00 | “PhysicsQualityAssurance has stopped working” Pop Up during Perform DFS Calibration | WIP19 |  | N/A | Subsystem |  |  | Duplicate | yatziv_l | 0 | 2 |  | DMS | NO | Closed |  | Brilliance 64
CTUDT00625863 | Refresh | Closed | altshuller_mark | August 25, 2019 8:10:01 PM GMT+08:00 | “PhysicsQualityAssurance has stopped working” Pop Up during Perform DFS Calibration | WIP19 |  | N/A | Subsystem |  |  | Duplicate | yatziv_l | 0 | 2 |  | DMS | NO | Closed |  | Ingenuity 64
CTUDT00625863 | Refresh | Closed | altshuller_mark | August 25, 2019 8:10:01 PM GMT+08:00 | “PhysicsQualityAssurance has stopped working” Pop Up during Perform DFS Calibration | WIP19 |  | N/A | Subsystem |  |  | Duplicate | yatziv_l | 0 | 2 |  | DMS | NO | Closed |  | IQon1
CTUDT00625859 | Refresh | Closed | attali_s | August 25, 2019 6:44:10 PM GMT+08:00 | Refresh: Performance Test: Wrong upper Acrylic limit in Head lin2b case with Two-Part Phantom | WIP18 | WIP20 | N/A | Subsystem | Performance Test;Blocker20 | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP20SubCom | IQon1 |  | WIP20 | September 3, 2019 10:09:06 PM GMT+08:00
CTUDT00625857 | Refresh | Closed | brutzkus_u | August 25, 2019 5:56:33 PM GMT+08:00 | Rotor Balancing upon UHR Replacement | WIP19 | WIP20 |  | Subsystem |  | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP20SubCom | Ovation |  | WIP20 | September 18, 2019 2:59:01 PM GMT+08:00
CTUDT00625853 | Refresh | Closed | wang_l | August 25, 2019 11:29:41 AM GMT+08:00 | Refresh: some exam card import failed if import too much by one time | WIP18 | WIP20 |  | Subsystem | ECM;Performance | None | Fixed | deshbhratar_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP20SubCom | Ovation |  | WIP20 | September 6, 2019 6:40:38 PM GMT+08:00
CTUDT00625853 | Refresh | Closed | wang_l | August 25, 2019 11:29:41 AM GMT+08:00 | Refresh: some exam card import failed if import too much by one time | WIP18 | WIP20 |  | Subsystem | ECM;Performance | None | Fixed | deshbhratar_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP20SubCom | Brilliance 64 |  | WIP20 | September 6, 2019 6:40:38 PM GMT+08:00
CTUDT00625853 | Refresh | Closed | wang_l | August 25, 2019 11:29:41 AM GMT+08:00 | Refresh: some exam card import failed if import too much by one time | WIP18 | WIP20 |  | Subsystem | ECM;Performance | None | Fixed | deshbhratar_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP20SubCom | Ingenuity 64 |  | WIP20 | September 6, 2019 6:40:38 PM GMT+08:00
CTUDT00625853 | Refresh | Closed | wang_l | August 25, 2019 11:29:41 AM GMT+08:00 | Refresh: some exam card import failed if import too much by one time | WIP18 | WIP20 |  | Subsystem | ECM;Performance | None | Fixed | deshbhratar_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP20SubCom | IQon1 |  | WIP20 | September 6, 2019 6:40:38 PM GMT+08:00
CTUDT00625847 | Refresh | Closed | zhu_je | August 24, 2019 10:34:49 AM GMT+08:00 | [Refresh][PCM] Cannot find label for language “Montenegran” Br64 PCM | WIP18 | DHF215793 - 728232 - Brilliance 64 Upgrades PCM - Rev17 | [Refresh][PCM] Cannot find label for language “Montenegran” Br64 PCM | Subsystem | PCM | NA | Fixed | hong_x | 0 | 2 | PCM | Support-System | NO | Closed | Next PCM version | Ovation |  | DHF215793 - 728232 - Brilliance 64 Upgrades PCM - Rev17 | October 31, 2019 9:43:50 AM GMT+08:00
CTUDT00625847 | Refresh | Closed | zhu_je | August 24, 2019 10:34:49 AM GMT+08:00 | [Refresh][PCM] Cannot find label for language “Montenegran” Br64 PCM | WIP18 | DHF215793 - 728232 - Brilliance 64 Upgrades PCM - Rev17 | [Refresh][PCM] Cannot find label for language “Montenegran” Br64 PCM | Subsystem | PCM | NA | Fixed | hong_x | 0 | 2 | PCM | Support-System | NO | Closed | Next PCM version | Brilliance 64 |  | DHF215793 - 728232 - Brilliance 64 Upgrades PCM - Rev17 | October 31, 2019 9:43:50 AM GMT+08:00
CTUDT00625847 | Refresh | Closed | zhu_je | August 24, 2019 10:34:49 AM GMT+08:00 | [Refresh][PCM] Cannot find label for language “Montenegran” Br64 PCM | WIP18 | DHF215793 - 728232 - Brilliance 64 Upgrades PCM - Rev17 | [Refresh][PCM] Cannot find label for language “Montenegran” Br64 PCM | Subsystem | PCM | NA | Fixed | hong_x | 0 | 2 | PCM | Support-System | NO | Closed | Next PCM version | Ingenuity 64 |  | DHF215793 - 728232 - Brilliance 64 Upgrades PCM - Rev17 | October 31, 2019 9:43:50 AM GMT+08:00
CTUDT00625845 | Refresh | Closed | Sharmila_Vasireddy | August 23, 2019 3:42:59 PM GMT+08:00 | BB Refresh: In IQ Test Exam group, few exam cards are displaying for CT user | BB_WIP8 | BB_WIP8 |  | Subsystem | Patient Details | NA | Fixed | richards_b | 0 | 2 | Design Documentation | Support-Physics | NO | Closed | BB_WIP10_SubCom | Brilliance Big Bore |  | BB_WIP8 | October 4, 2019 9:49:32 PM GMT+08:00
CTUDT00625843 | Refresh | Closed | vineetha_s | August 23, 2019 3:31:38 PM GMT+08:00 | The system and software error is displayed on performing cardiac and non-cardiac scan | WIP18 | WIP20 |  | Subsystem |  | NA | Fixed | challa_s | 0 | 2 | Console Software | Console4 | YES | Closed | WIP20SubCom | Ovation |  | WIP20 | September 6, 2019 1:04:10 PM GMT+08:00
CTUDT00625843 | Refresh | Closed | vineetha_s | August 23, 2019 3:31:38 PM GMT+08:00 | The system and software error is displayed on performing cardiac and non-cardiac scan | WIP18 | WIP20 |  | Subsystem |  | NA | Fixed | challa_s | 0 | 2 | Console Software | Console4 | YES | Closed | WIP20SubCom | Brilliance 64 |  | WIP20 | September 6, 2019 1:04:10 PM GMT+08:00
CTUDT00625843 | Refresh | Closed | vineetha_s | August 23, 2019 3:31:38 PM GMT+08:00 | The system and software error is displayed on performing cardiac and non-cardiac scan | WIP18 | WIP20 |  | Subsystem |  | NA | Fixed | challa_s | 0 | 2 | Console Software | Console4 | YES | Closed | WIP20SubCom | Ingenuity 64 |  | WIP20 | September 6, 2019 1:04:10 PM GMT+08:00
CTUDT00625843 | Refresh | Closed | vineetha_s | August 23, 2019 3:31:38 PM GMT+08:00 | The system and software error is displayed on performing cardiac and non-cardiac scan | WIP18 | WIP20 |  | Subsystem |  | NA | Fixed | challa_s | 0 | 2 | Console Software | Console4 | YES | Closed | WIP20SubCom | IQon1 |  | WIP20 | September 6, 2019 1:04:10 PM GMT+08:00
CTUDT00625839 | Refresh | Closed | 320053767 | August 22, 2019 11:10:56 PM GMT+08:00 | Performance Infant system test failure (incorrect factor in cfg file for Harmonized Phantom ) | WIP19 |  |  | Subsystem |  |  | Not Enough Information | 320053767 | 0 |  |  | DMS |  | Closed |  | IQon1
CTUDT00625837 | Refresh | Closed | tamari_o | August 22, 2019 10:56:23 PM GMT+08:00 | Refresh WIP19 - Tube Alignment failure (incorrect bed position parameter for Harmonized Phantom ) | WIP19 | WIP20 |  | Subsystem | Calibrations;Blocker20; Verified1 | NaN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP20SubCom | Ovation |  | WIP20 | August 29, 2019 4:05:22 PM GMT+08:00
CTUDT00625837 | Refresh | Closed | tamari_o | August 22, 2019 10:56:23 PM GMT+08:00 | Refresh WIP19 - Tube Alignment failure (incorrect bed position parameter for Harmonized Phantom ) | WIP19 | WIP20 |  | Subsystem | Calibrations;Blocker20; Verified1 | NaN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP20SubCom | Brilliance Big Bore |  | WIP20 | August 29, 2019 4:05:22 PM GMT+08:00
CTUDT00625837 | Refresh | Closed | tamari_o | August 22, 2019 10:56:23 PM GMT+08:00 | Refresh WIP19 - Tube Alignment failure (incorrect bed position parameter for Harmonized Phantom ) | WIP19 | WIP20 |  | Subsystem | Calibrations;Blocker20; Verified1 | NaN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP20SubCom | Ingenuity 64 |  | WIP20 | August 29, 2019 4:05:22 PM GMT+08:00
CTUDT00625837 | Refresh | Closed | tamari_o | August 22, 2019 10:56:23 PM GMT+08:00 | Refresh WIP19 - Tube Alignment failure (incorrect bed position parameter for Harmonized Phantom ) | WIP19 | WIP20 |  | Subsystem | Calibrations;Blocker20; Verified1 | NaN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP20SubCom | IQon1 |  | WIP20 | August 29, 2019 4:05:22 PM GMT+08:00
CTUDT00625831 | Refresh | Closed | levy_s | August 22, 2019 10:19:16 PM GMT+08:00 | Refresh: Slow GUI performance after scanning the locator | WIP18 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | tamari_o | 1 |  |  | Support-System | NO | Closed |  | IQon1
CTUDT00625830 | Refresh | Closed | 320053767 | August 22, 2019 10:18:49 PM GMT+08:00 | DDC calibration failure (incorrect factor in XML file for Harmonized Phantom ) | WIP19 | WIP20 |  | Subsystem | Calibrations;Blocker20 | NaN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP20SubCom | IQon1 |  | WIP20 | August 29, 2019 4:40:08 PM GMT+08:00
CTUDT00625828 | Refresh | Closed | 320053767 | August 22, 2019 9:58:24 PM GMT+08:00 | Focal spot X position values out of range GLC | WIP19 | WIP20 |  | Subsystem | Gantry Calibrations;Blocker20 | N/A | Fixed | zhou_x | 0 | 2 | Gantry Calibrations | GCX | NO | Closed | WIP20SegReq | Ovation |  | WIP20 | September 12, 2019 6:27:02 PM GMT+08:00
CTUDT00625828 | Refresh | Closed | 320053767 | August 22, 2019 9:58:24 PM GMT+08:00 | Focal spot X position values out of range GLC | WIP19 | WIP20 |  | Subsystem | Gantry Calibrations;Blocker20 | N/A | Fixed | zhou_x | 0 | 2 | Gantry Calibrations | GCX | NO | Closed | WIP20SegReq | IQon1 |  | WIP20 | September 12, 2019 6:27:02 PM GMT+08:00
CTUDT00625826 | Refresh | Closed | zhang_l | August 22, 2019 9:23:35 PM GMT+08:00 | “Diisconnected from CIRS Server 2. Local disconnection” Typo Error while close CIRS2 remote from remote viewer | WIP19 |  |  | Subsystem |  | NA | Duplicate | chakraborty_d | 0 |  |  | Service Tools | NO | Closed | WIP20SubCom | Ovation
CTUDT00625826 | Refresh | Closed | zhang_l | August 22, 2019 9:23:35 PM GMT+08:00 | “Diisconnected from CIRS Server 2. Local disconnection” Typo Error while close CIRS2 remote from remote viewer | WIP19 |  |  | Subsystem |  | NA | Duplicate | chakraborty_d | 0 |  |  | Service Tools | NO | Closed | WIP20SubCom | Brilliance 64
CTUDT00625826 | Refresh | Closed | zhang_l | August 22, 2019 9:23:35 PM GMT+08:00 | “Diisconnected from CIRS Server 2. Local disconnection” Typo Error while close CIRS2 remote from remote viewer | WIP19 |  |  | Subsystem |  | NA | Duplicate | chakraborty_d | 0 |  |  | Service Tools | NO | Closed | WIP20SubCom | Brilliance Big Bore
CTUDT00625826 | Refresh | Closed | zhang_l | August 22, 2019 9:23:35 PM GMT+08:00 | “Diisconnected from CIRS Server 2. Local disconnection” Typo Error while close CIRS2 remote from remote viewer | WIP19 |  |  | Subsystem |  | NA | Duplicate | chakraborty_d | 0 |  |  | Service Tools | NO | Closed | WIP20SubCom | Ingenuity 64
CTUDT00625826 | Refresh | Closed | zhang_l | August 22, 2019 9:23:35 PM GMT+08:00 | “Diisconnected from CIRS Server 2. Local disconnection” Typo Error while close CIRS2 remote from remote viewer | WIP19 |  |  | Subsystem |  | NA | Duplicate | chakraborty_d | 0 |  |  | Service Tools | NO | Closed | WIP20SubCom | IQon1
CTUDT00625821 | Refresh | Closed | liu_bo | August 22, 2019 7:03:53 PM GMT+08:00 | There is not Speed Profile test item on the left column. | WIP18 |  | By communicating with testers, it was found that the problem was incorrect in describing the navigation path of virtual diagnosis in service tools in test cases. This is not a problem, gantry software does not need to make any changes. Change the Path form : "Service Tools – Diagnostics – Gantry - Visual Diagnostics" to "Service Tools – Diagnostics – System - Visual Diagnostics" in the test case. | Subsystem | Visual Diagnostics | NA | Works As Specified / Not A Defect | zhang_lawrence | 0 |  |  | GCX | NO | Closed |  | Ovation
CTUDT00625816 | Refresh | Closed | gushansky_m | August 22, 2019 6:51:34 PM GMT+08:00 | Refresh: Spelling error in software distribution tool | WIP18 | BB_WIP10 WIP20 | None | Subsystem |  | None | Fixed | tejaswini_nidamanuri | 0 | 3 | Service Tools_Software | Service Tools | NO | Closed | WIP20SubCom | Ovation |  | BB_WIP10 WIP20 | August 28, 2019 1:04:03 PM GMT+08:00
CTUDT00625816 | Refresh | Closed | gushansky_m | August 22, 2019 6:51:34 PM GMT+08:00 | Refresh: Spelling error in software distribution tool | WIP18 | BB_WIP10 WIP20 | None | Subsystem |  | None | Fixed | tejaswini_nidamanuri | 0 | 3 | Service Tools_Software | Service Tools | NO | Closed | WIP20SubCom | Brilliance 64 |  | BB_WIP10 WIP20 | August 28, 2019 1:04:03 PM GMT+08:00
CTUDT00625816 | Refresh | Closed | gushansky_m | August 22, 2019 6:51:34 PM GMT+08:00 | Refresh: Spelling error in software distribution tool | WIP18 | BB_WIP10 WIP20 | None | Subsystem |  | None | Fixed | tejaswini_nidamanuri | 0 | 3 | Service Tools_Software | Service Tools | NO | Closed | WIP20SubCom | Brilliance Big Bore |  | BB_WIP10 WIP20 | August 28, 2019 1:04:03 PM GMT+08:00
CTUDT00625816 | Refresh | Closed | gushansky_m | August 22, 2019 6:51:34 PM GMT+08:00 | Refresh: Spelling error in software distribution tool | WIP18 | BB_WIP10 WIP20 | None | Subsystem |  | None | Fixed | tejaswini_nidamanuri | 0 | 3 | Service Tools_Software | Service Tools | NO | Closed | WIP20SubCom | Ingenuity 64 |  | BB_WIP10 WIP20 | August 28, 2019 1:04:03 PM GMT+08:00
CTUDT00625816 | Refresh | Closed | gushansky_m | August 22, 2019 6:51:34 PM GMT+08:00 | Refresh: Spelling error in software distribution tool | WIP18 | BB_WIP10 WIP20 | None | Subsystem |  | None | Fixed | tejaswini_nidamanuri | 0 | 3 | Service Tools_Software | Service Tools | NO | Closed | WIP20SubCom | IQon1 |  | BB_WIP10 WIP20 | August 28, 2019 1:04:03 PM GMT+08:00
CTUDT00625815 | Refresh | Closed | liu_bo | August 22, 2019 6:50:24 PM GMT+08:00 | There is not GLC_FDOM_PeakAve folder in Gantry directory of BugReport, but exist extra ChostDataFiles folder | WIP18 |  | Communicate with testers and check the "GantryBugRepMap.xml" in the "C:\PmsCT\Config\Service\Bugrep" folder 1: The "GLC_FDOM_PeakAve" folder Collection config item in the config file("GantryBugRepMap.xml") <Folder Id="GLC_FDOM_PeakAve" SourcePath="D:\Data_GantryTextLogs\GantryLimitsCalc\FDOM_PeakAve" DestinationSubFolder="GLC_FDOM_PeakAve" IncludeInMiniBugrep="true"></Folder> If System do not implement FreeDOM dose modulation scan ,there is nothing in the fold "D:\Data_GantryTextLogs\GantryLimitsCalc\FDOM_PeakAve", so bugrep will not do the copy and it will not create "GLC_FDOM_PeakAve" folder in the Gantry bugrep folder. This is normal. 2: The "ChostDataFiles" folder is the config files of Chost Service Tool, They are designed to support the newly added Chost service tools. "ChostDataFiles" folder will in the Gantry bugrep folder. This is normal. The conclusion is that this result is not a problem and no change is required in the Gantry SW. | Subsystem |  | NA | Works As Specified / Not A Defect | zhang_lawrence | 0 |  |  | GCX | NO | Closed |  | Ovation
CTUDT00625808 | Refresh | Closed | wei_f | August 22, 2019 6:27:56 PM GMT+08:00 | Copy studies which contain SR to ISP local, the SR series was not displayed in ISP local folder | BB_WIP8 |  |  | Subsystem | CD-DVD operation | NA | Does Not Apply | srushti_k | 0 |  |  | iCAP | NO | Closed |  | Ovation
CTUDT00625808 | Refresh | Closed | wei_f | August 22, 2019 6:27:56 PM GMT+08:00 | Copy studies which contain SR to ISP local, the SR series was not displayed in ISP local folder | BB_WIP8 |  |  | Subsystem | CD-DVD operation | NA | Does Not Apply | srushti_k | 0 |  |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00625808 | Refresh | Closed | wei_f | August 22, 2019 6:27:56 PM GMT+08:00 | Copy studies which contain SR to ISP local, the SR series was not displayed in ISP local folder | BB_WIP8 |  |  | Subsystem | CD-DVD operation | NA | Does Not Apply | srushti_k | 0 |  |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00625808 | Refresh | Closed | wei_f | August 22, 2019 6:27:56 PM GMT+08:00 | Copy studies which contain SR to ISP local, the SR series was not displayed in ISP local folder | BB_WIP8 |  |  | Subsystem | CD-DVD operation | NA | Does Not Apply | srushti_k | 0 |  |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00625808 | Refresh | Closed | wei_f | August 22, 2019 6:27:56 PM GMT+08:00 | Copy studies which contain SR to ISP local, the SR series was not displayed in ISP local folder | BB_WIP8 |  |  | Subsystem | CD-DVD operation | NA | Does Not Apply | srushti_k | 0 |  |  | iCAP | NO | Closed |  | IQon1
CTUDT00625806 | Refresh | Closed | ding_s | August 22, 2019 6:01:23 PM GMT+08:00 | 'Sampled Collimation' did not display in result series. | WIP18 | WIP19 |  | Subsystem |  | N/A | Fixed | ding_s | 0 | 2 | Test Design | Support-System | NO | Closed | WAS | Ingenuity 64 |  | WIP19 | November 4, 2019 1:45:12 PM GMT+08:00
CTUDT00625794 | Refresh | Closed | zhang_l | August 22, 2019 4:22:30 PM GMT+08:00 | Air calibration paused about 2~3 mins | WIP19 |  | N/A | Subsystem |  | N/A | Duplicate | frider_o | 0 |  |  | DMS | NO | Closed |  | Ingenuity 64
CTUDT00625789 | Refresh | Closed | zhong_x | August 22, 2019 4:08:49 PM GMT+08:00 | “Diisconnected from IVC.Local disconnection” Typo Error while close IVC remote from remote viewer | BB_WIP9 | BB_WIP10 |  | Subsystem |  | NA | Fixed | chakraborty_d | 0 | 4 | Service Tools_Software | Service Tools | NO | Closed | BB_WIP10_SubCom | Brilliance Big Bore |  | BB_WIP10 | August 29, 2019 7:49:18 PM GMT+08:00
CTUDT00625786 | Refresh | Closed | zhong_x | August 22, 2019 3:59:54 PM GMT+08:00 | Lock Screen didn’t show entirely in IVC remote view, Need to drag the bar in bottom. | BB_WIP9 |  |  | Subsystem |  | NA | No plan to Fix | chakraborty_d | 0 | 4 |  | Service Tools | NO | Closed | BB_SubNC | Brilliance Big Bore
CTUDT00625781 | Refresh | Closed | zhong_x | August 22, 2019 3:48:32 PM GMT+08:00 | Move Couch up and down are not working in IVC remote viewer while the displayed couch up and down value keep syncing with CT Box. | BB_WIP9 |  |  | Subsystem | IVC | NA | Works As Specified / Not A Defect | ye_h | 0 | 2 |  | Support-System | NO | Closed | BB_SubNC | Brilliance Big Bore
CTUDT00625776 | Refresh | Closed | zhong_x | August 22, 2019 3:36:46 PM GMT+08:00 | In remote viewer, Need to click restore down” before drag the remote window | BB_WIP9 |  |  | Subsystem |  | NA | No Plan to Fix | chakraborty_d | 0 | 4 |  | Service Tools | NO | Closed | BB_SubNC | Brilliance Big Bore
CTUDT00625768 | Refresh | Closed | wei_f | August 22, 2019 3:20:39 PM GMT+08:00 | Load Chinese Bookmarks to CT Viewer failed. | BB_WIP8 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | wei_f | 1 |  |  | Support-System | NO | Closed |  | Ovation
CTUDT00625768 | Refresh | Closed | wei_f | August 22, 2019 3:20:39 PM GMT+08:00 | Load Chinese Bookmarks to CT Viewer failed. | BB_WIP8 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | wei_f | 1 |  |  | Support-System | NO | Closed |  | Brilliance 64
CTUDT00625768 | Refresh | Closed | wei_f | August 22, 2019 3:20:39 PM GMT+08:00 | Load Chinese Bookmarks to CT Viewer failed. | BB_WIP8 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | wei_f | 1 |  |  | Support-System | NO | Closed |  | Brilliance Big Bore
CTUDT00625768 | Refresh | Closed | wei_f | August 22, 2019 3:20:39 PM GMT+08:00 | Load Chinese Bookmarks to CT Viewer failed. | BB_WIP8 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | wei_f | 1 |  |  | Support-System | NO | Closed |  | Ingenuity 64
CTUDT00625768 | Refresh | Closed | wei_f | August 22, 2019 3:20:39 PM GMT+08:00 | Load Chinese Bookmarks to CT Viewer failed. | BB_WIP8 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | wei_f | 1 |  |  | Support-System | NO | Closed |  | IQon1
CTUDT00625766 | Refresh | Closed | ding_s | August 22, 2019 3:11:41 PM GMT+08:00 | Loaded Resp 4D CT exam card failed. | WIP18 | WIP20 |  | Subsystem |  | NA | Fixed | m_nanda | 0 | 3 | Console Software | Console4 | NO | Closed | WIP20SubCom | Brilliance 64 |  | WIP20 | September 6, 2019 3:09:21 PM GMT+08:00
CTUDT00625766 | Refresh | Closed | ding_s | August 22, 2019 3:11:41 PM GMT+08:00 | Loaded Resp 4D CT exam card failed. | WIP18 | WIP20 |  | Subsystem |  | NA | Fixed | m_nanda | 0 | 3 | Console Software | Console4 | NO | Closed | WIP20SubCom | Ingenuity 64 |  | WIP20 | September 6, 2019 3:09:21 PM GMT+08:00
CTUDT00625762 | Refresh | Closed | ding_s | August 22, 2019 2:29:06 PM GMT+08:00 | Did not find instruction of altitude range of 10000m or equivalent pressure range of 30kPa in service documents. | WIP18 | 453567548104 Rev B |  | System |  | CT-NM.SRS.SysSupport.529 | Fixed | li_e | 0 | 2 | Service Documentation | Support-Service | NO | Closed | WIP20SubCom | Brilliance 64 |  | 453567548104 Rev B | November 19, 2019 5:52:59 PM GMT+08:00
CTUDT00625760 | Refresh | Closed | ding_s | August 22, 2019 2:21:19 PM GMT+08:00 | Did not find instruction of altitude range and pressure range for system in service documents. | WIP18 | 453567548104 Rev B |  | System |  | CT-NM.SRS.SysSupport.239 | Fixed | li_e | 0 | 2 | Service Documentation | Support-Service | NO | Closed | WIP20SubCom | Brilliance 64 |  | 453567548104 Rev B | November 19, 2019 5:52:22 PM GMT+08:00
CTUDT00625758 | Refresh | Closed | ding_s | August 22, 2019 1:49:38 PM GMT+08:00 | Did not find instruction of total heat output generated by equipment in service documents. | WIP18 | 453567548104 Rev B |  | System |  | CT-NM.SRS.SysSupport.199 | Fixed | li_e | 0 | 2 | Service Documentation | Support-Service | NO | Closed | WIP20SubCom | Brilliance 64 |  | 453567548104 Rev B | November 19, 2019 5:51:54 PM GMT+08:00
CTUDT00625754 | Refresh | Closed | ding_s | August 22, 2019 1:43:18 PM GMT+08:00 | Did not find instruction of operate and scan up to an altitude 9842 feet (3,000 meter) [70 kilopascals] in service documents. | WIP18 | 453567548104 Rev B |  | Subsystem |  | CT-NM.SRS.SysSupport.197 | Fixed | li_e | 0 | 2 | Service Documentation | Support-System | NO | Closed | WIP20SubCom | Brilliance 64 |  | 453567548104 Rev B | October 15, 2019 2:00:49 PM GMT+08:00
CTUDT00625753 | Refresh | Closed | liu_bo | August 22, 2019 1:42:34 PM GMT+08:00 | After restart control panel manually, panel display error dialog. | WIP18 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | Mark_Huang | 0 |  |  | GCX | NO | Closed |  | Ovation
CTUDT00625751 | Refresh | Closed | zhong_x | August 22, 2019 1:37:17 PM GMT+08:00 | IVC Touch screen is not sensitive, sometimes need to press twice. | BB_WIP9 |  | If touch sound not played on increment selection change, it mean windows did not detect touch as click. Redo touch. | Subsystem |  | None | No Plan to Fix | lin_s | 0 | 4 |  | Support-System | NO | Closed | WAS | Brilliance Big Bore
CTUDT00625749 | Refresh | Closed | ding_s | August 22, 2019 1:34:11 PM GMT+08:00 | Did not find instruction about ‘align the lasers’ in service documents. | WIP18 |  |  | System |  | CT-NM.SRS.Service.2431 | Works As Specified / Not A Defect | lin_s | 0 |  |  | Support-Service | NO | Closed |  | Brilliance 64
CTUDT00625748 | Refresh | Closed | zhong_x | August 22, 2019 1:29:12 PM GMT+08:00 | System was hang after click End Exam in CCT continuous Scan | BB_WIP9 |  |  | Subsystem | IVC |  | Works As Specified / Not A Defect | zhong_x | 1 |  |  | Support-System | NO | Closed |  | Brilliance Big Bore
CTUDT00625746 | Refresh | Closed | zhong_x | August 22, 2019 1:19:10 PM GMT+08:00 | EC was disappeared in group of Reference Set | BB_WIP9 | WIP20 |  | Subsystem |  | NA | Fixed | rath_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP20SubCom | Ovation |  | WIP20 | September 5, 2019 8:36:14 PM GMT+08:00
CTUDT00625746 | Refresh | Closed | zhong_x | August 22, 2019 1:19:10 PM GMT+08:00 | EC was disappeared in group of Reference Set | BB_WIP9 | WIP20 |  | Subsystem |  | NA | Fixed | rath_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP20SubCom | Brilliance 64 |  | WIP20 | September 5, 2019 8:36:14 PM GMT+08:00
CTUDT00625746 | Refresh | Closed | zhong_x | August 22, 2019 1:19:10 PM GMT+08:00 | EC was disappeared in group of Reference Set | BB_WIP9 | WIP20 |  | Subsystem |  | NA | Fixed | rath_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP20SubCom | Ingenuity 64 |  | WIP20 | September 5, 2019 8:36:14 PM GMT+08:00
CTUDT00625746 | Refresh | Closed | zhong_x | August 22, 2019 1:19:10 PM GMT+08:00 | EC was disappeared in group of Reference Set | BB_WIP9 | WIP20 |  | Subsystem |  | NA | Fixed | rath_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP20SubCom | IQon1 |  | WIP20 | September 5, 2019 8:36:14 PM GMT+08:00
CTUDT00625743 | Refresh | Closed | Sandeep_Kumar | August 22, 2019 1:04:09 PM GMT+08:00 | Console Status not displayed in System Status tool. | BB_WIP9 | BB_WIP10 |  | Subsystem |  | NA | Does Not Apply | s_shabeer | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | Does Not Apply | Brilliance Big Bore |  | BB_WIP10 | August 23, 2019 7:53:39 PM GMT+08:00
CTUDT00625741 | Refresh | Closed | kapadia_v | August 22, 2019 12:54:17 PM GMT+08:00 | In Emergency login user is not able to proceed the scan when scan exceeds dose limits in dose management | BB_WIP9 | BB_WIP11 | Decrease dose below dose limits | Subsystem |  | NA | Fixed | kr_m | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP10_SubCom | Brilliance Big Bore |  | BB_WIP11 | October 10, 2019 7:38:22 PM GMT+08:00
CTUDT00625740 | Refresh | Closed | liu_bo | August 22, 2019 12:42:27 PM GMT+08:00 | After remote GPC successful, user cannot enter into GPC windows explore , only display panel. | WIP18 |  |  | Subsystem | Service Tools |  | Duplicate | kline_b | 0 |  |  | GCX | NO | Closed |  | Ovation
CTUDT00625738 | Refresh | Closed | ding_s | August 22, 2019 11:15:41 AM GMT+08:00 | Did not find instruction of Image Quality and Dose test procedures in service documents. | WIP18 |  |  | System |  | CT-NM.SRS.RegStand.740 | Works As Specified / Not A Defect | lin_s | 0 |  |  | Support-Service | NO | Closed |  | Brilliance 64
CTUDT00625738 | Refresh | Closed | ding_s | August 22, 2019 11:15:41 AM GMT+08:00 | Did not find instruction of Image Quality and Dose test procedures in service documents. | WIP18 |  |  | System |  | CT-NM.SRS.RegStand.740 | Works As Specified / Not A Defect | lin_s | 0 |  |  | Support-Service | NO | Closed |  | Brilliance Big Bore
CTUDT00625738 | Refresh | Closed | ding_s | August 22, 2019 11:15:41 AM GMT+08:00 | Did not find instruction of Image Quality and Dose test procedures in service documents. | WIP18 |  |  | System |  | CT-NM.SRS.RegStand.740 | Works As Specified / Not A Defect | lin_s | 0 |  |  | Support-Service | NO | Closed |  | Ingenuity 64
CTUDT00625736 | Refresh | Closed | ding_s | August 22, 2019 11:05:36 AM GMT+08:00 | Did not find instruction of how to load and display the SMPTE in service documents. | WIP18 | 459801587191 Rev B 459801587121 Rev B |  | System |  | CT-NM.SRS.Service.3056 | Fixed | li_e | 0 | 2 | Service Documentation | Support-Service | NO | Closed | WIP19SubCom | Brilliance 64 |  | 459801587191 Rev B 459801587121 Rev B | November 19, 2019 5:50:56 PM GMT+08:00
CTUDT00625736 | Refresh | Closed | ding_s | August 22, 2019 11:05:36 AM GMT+08:00 | Did not find instruction of how to load and display the SMPTE in service documents. | WIP18 | 459801587191 Rev B 459801587121 Rev B |  | System |  | CT-NM.SRS.Service.3056 | Fixed | li_e | 0 | 2 | Service Documentation | Support-Service | NO | Closed | WIP19SubCom | Brilliance Big Bore |  | 459801587191 Rev B 459801587121 Rev B | November 19, 2019 5:50:56 PM GMT+08:00
CTUDT00625736 | Refresh | Closed | ding_s | August 22, 2019 11:05:36 AM GMT+08:00 | Did not find instruction of how to load and display the SMPTE in service documents. | WIP18 | 459801587191 Rev B 459801587121 Rev B |  | System |  | CT-NM.SRS.Service.3056 | Fixed | li_e | 0 | 2 | Service Documentation | Support-Service | NO | Closed | WIP19SubCom | Ingenuity 64 |  | 459801587191 Rev B 459801587121 Rev B | November 19, 2019 5:50:56 PM GMT+08:00
CTUDT00625734 | Refresh | Closed | ding_s | August 22, 2019 10:55:31 AM GMT+08:00 | Did not find instruction of how to activate CopSSH in software installation manual. | WIP18 |  |  | System |  | CT-NM.SRS.Service.3064 | Works As Specified / Not A Defect | lin_s | 0 |  |  | Support-Service | NO | Closed |  | Brilliance 64
CTUDT00625734 | Refresh | Closed | ding_s | August 22, 2019 10:55:31 AM GMT+08:00 | Did not find instruction of how to activate CopSSH in software installation manual. | WIP18 |  |  | System |  | CT-NM.SRS.Service.3064 | Works As Specified / Not A Defect | lin_s | 0 |  |  | Support-Service | NO | Closed |  | Brilliance Big Bore
CTUDT00625734 | Refresh | Closed | ding_s | August 22, 2019 10:55:31 AM GMT+08:00 | Did not find instruction of how to activate CopSSH in software installation manual. | WIP18 |  |  | System |  | CT-NM.SRS.Service.3064 | Works As Specified / Not A Defect | lin_s | 0 |  |  | Support-Service | NO | Closed |  | Ingenuity 64
CTUDT00625732 | Refresh | Closed | ding_s | August 22, 2019 10:51:59 AM GMT+08:00 | Did not find instruction of unique administrator and password of the system to customer in installation service documents. | WIP18 | 459801587331 Rev C 459801587341 Rev C 459801587361 Rev C 459801587351 Rev C |  | System |  | CT-NM.SRS.SecPri.373 | Fixed | li_e | 0 | 2 | Service Documentation | Support-Service | NO | Closed | WIP21SegReq | Brilliance 64 |  | 459801587331 Rev C 459801587341 Rev C 459801587361 Rev C 459801587351 Rev C | November 19, 2019 5:51:22 PM GMT+08:00
CTUDT00625732 | Refresh | Closed | ding_s | August 22, 2019 10:51:59 AM GMT+08:00 | Did not find instruction of unique administrator and password of the system to customer in installation service documents. | WIP18 | 459801587331 Rev C 459801587341 Rev C 459801587361 Rev C 459801587351 Rev C |  | System |  | CT-NM.SRS.SecPri.373 | Fixed | li_e | 0 | 2 | Service Documentation | Support-Service | NO | Closed | WIP21SegReq | Brilliance Big Bore |  | 459801587331 Rev C 459801587341 Rev C 459801587361 Rev C 459801587351 Rev C | November 19, 2019 5:51:22 PM GMT+08:00
CTUDT00625732 | Refresh | Closed | ding_s | August 22, 2019 10:51:59 AM GMT+08:00 | Did not find instruction of unique administrator and password of the system to customer in installation service documents. | WIP18 | 459801587331 Rev C 459801587341 Rev C 459801587361 Rev C 459801587351 Rev C |  | System |  | CT-NM.SRS.SecPri.373 | Fixed | li_e | 0 | 2 | Service Documentation | Support-Service | NO | Closed | WIP21SegReq | Ingenuity 64 |  | 459801587331 Rev C 459801587341 Rev C 459801587361 Rev C 459801587351 Rev C | November 19, 2019 5:51:22 PM GMT+08:00
CTUDT00625730 | Refresh | Closed | ding_s | August 22, 2019 10:42:08 AM GMT+08:00 | Did not find guidance of HIPAA Security Standards - Technical Safeguards in service documents. | WIP18 |  |  | System |  |  | Duplicate | lin_s | 0 |  |  | Support-Service | NO | Closed |  | Brilliance 64
CTUDT00625730 | Refresh | Closed | ding_s | August 22, 2019 10:42:08 AM GMT+08:00 | Did not find guidance of HIPAA Security Standards - Technical Safeguards in service documents. | WIP18 |  |  | System |  |  | Duplicate | lin_s | 0 |  |  | Support-Service | NO | Closed |  | Ingenuity 64
CTUDT00625728 | Refresh | Closed | ding_s | August 22, 2019 10:37:33 AM GMT+08:00 | Did not find instruction of Ingenuity and Br64 in seismic service documents. | WIP18 | 459801587391 Rev C |  | System |  | CT-NM.SRS.SysSupport.863 | Fixed | shi_b | 0 | 2 | Service Documentation | Support-Service | NO | Closed | SubCom-21-Feb-2020 | Brilliance 64 |  | 459801587391 Rev C | February 23, 2020 3:00:58 PM GMT+08:00
CTUDT00625726 | Refresh | Closed | wei_f | August 22, 2019 9:48:05 AM GMT+08:00 | The description of CT-NM.SRS.Interop.328 is wrong. | BB_WIP8 | DHF298554 Rev08 DHF298555 Rev08 iCT, SRS DHF298553 Rev.08 IQon SRS DHF298556 Rev.09 |  | Subsystem | Requirements;Verified1 | CT-NM.SRS.Interop.328 | Fixed | guy_sheaffer | 0 | 2 | SRS | Support-System | NO | Closed | BB SRS Rev08 | Ovation |  | DHF298554 Rev08 DHF298555 Rev08 iCT, SRS DHF298553 Rev.08 IQon SRS DHF298556 Rev.09 | November 6, 2019 2:09:41 PM GMT+08:00
CTUDT00625726 | Refresh | Closed | wei_f | August 22, 2019 9:48:05 AM GMT+08:00 | The description of CT-NM.SRS.Interop.328 is wrong. | BB_WIP8 | DHF298554 Rev08 DHF298555 Rev08 iCT, SRS DHF298553 Rev.08 IQon SRS DHF298556 Rev.09 |  | Subsystem | Requirements;Verified1 | CT-NM.SRS.Interop.328 | Fixed | guy_sheaffer | 0 | 2 | SRS | Support-System | NO | Closed | BB SRS Rev08 | Brilliance 64 |  | DHF298554 Rev08 DHF298555 Rev08 iCT, SRS DHF298553 Rev.08 IQon SRS DHF298556 Rev.09 | November 6, 2019 2:09:41 PM GMT+08:00
CTUDT00625726 | Refresh | Closed | wei_f | August 22, 2019 9:48:05 AM GMT+08:00 | The description of CT-NM.SRS.Interop.328 is wrong. | BB_WIP8 | DHF298554 Rev08 DHF298555 Rev08 iCT, SRS DHF298553 Rev.08 IQon SRS DHF298556 Rev.09 |  | Subsystem | Requirements;Verified1 | CT-NM.SRS.Interop.328 | Fixed | guy_sheaffer | 0 | 2 | SRS | Support-System | NO | Closed | BB SRS Rev08 | Brilliance Big Bore |  | DHF298554 Rev08 DHF298555 Rev08 iCT, SRS DHF298553 Rev.08 IQon SRS DHF298556 Rev.09 | November 6, 2019 2:09:41 PM GMT+08:00
CTUDT00625726 | Refresh | Closed | wei_f | August 22, 2019 9:48:05 AM GMT+08:00 | The description of CT-NM.SRS.Interop.328 is wrong. | BB_WIP8 | DHF298554 Rev08 DHF298555 Rev08 iCT, SRS DHF298553 Rev.08 IQon SRS DHF298556 Rev.09 |  | Subsystem | Requirements;Verified1 | CT-NM.SRS.Interop.328 | Fixed | guy_sheaffer | 0 | 2 | SRS | Support-System | NO | Closed | BB SRS Rev08 | Ingenuity 64 |  | DHF298554 Rev08 DHF298555 Rev08 iCT, SRS DHF298553 Rev.08 IQon SRS DHF298556 Rev.09 | November 6, 2019 2:09:41 PM GMT+08:00
CTUDT00625726 | Refresh | Closed | wei_f | August 22, 2019 9:48:05 AM GMT+08:00 | The description of CT-NM.SRS.Interop.328 is wrong. | BB_WIP8 | DHF298554 Rev08 DHF298555 Rev08 iCT, SRS DHF298553 Rev.08 IQon SRS DHF298556 Rev.09 |  | Subsystem | Requirements;Verified1 | CT-NM.SRS.Interop.328 | Fixed | guy_sheaffer | 0 | 2 | SRS | Support-System | NO | Closed | BB SRS Rev08 | IQon1 |  | DHF298554 Rev08 DHF298555 Rev08 iCT, SRS DHF298553 Rev.08 IQon SRS DHF298556 Rev.09 | November 6, 2019 2:09:41 PM GMT+08:00
CTUDT00625724 | Refresh | Closed | zhong_x | August 22, 2019 9:34:11 AM GMT+08:00 | User can access device manager and user account management screen before enter the IVC SW | BB_WIP9 | BB_WIP11.1 | N/A | Subsystem | IVC | N/A | Fixed | reddy_g | 0 | 2 | IVC Software | Console4 | YES | Closed | BB_WIP11_SubCom | Brilliance Big Bore |  | BB_WIP11.1 | November 7, 2019 2:06:11 PM GMT+08:00
CTUDT00625722 | Refresh | Closed | ding_s | August 22, 2019 9:30:11 AM GMT+08:00 | In safety service manual of Ingenuity and Br64, the Caution instruction does not meet CT-NM.SRS.Service.2950 | WIP18 | 459800093391 Rev K 459801263801 Rev B 459801263811 Rev B | N/A | Subsystem |  | CT-NM.SRS.Service.2950 | Fixed | li_e | 0 | 2 | Service Documentation | Support-System | NO | Closed | Next version of safety service manual | Brilliance 64 |  | 459800093391 Rev K 459801263801 Rev B 459801263811 Rev B | October 15, 2019 1:58:15 PM GMT+08:00
CTUDT00625722 | Refresh | Closed | ding_s | August 22, 2019 9:30:11 AM GMT+08:00 | In safety service manual of Ingenuity and Br64, the Caution instruction does not meet CT-NM.SRS.Service.2950 | WIP18 | 459800093391 Rev K 459801263801 Rev B 459801263811 Rev B | N/A | Subsystem |  | CT-NM.SRS.Service.2950 | Fixed | li_e | 0 | 2 | Service Documentation | Support-System | NO | Closed | Next version of safety service manual | Brilliance Big Bore |  | 459800093391 Rev K 459801263801 Rev B 459801263811 Rev B | October 15, 2019 1:58:15 PM GMT+08:00
CTUDT00625722 | Refresh | Closed | ding_s | August 22, 2019 9:30:11 AM GMT+08:00 | In safety service manual of Ingenuity and Br64, the Caution instruction does not meet CT-NM.SRS.Service.2950 | WIP18 | 459800093391 Rev K 459801263801 Rev B 459801263811 Rev B | N/A | Subsystem |  | CT-NM.SRS.Service.2950 | Fixed | li_e | 0 | 2 | Service Documentation | Support-System | NO | Closed | Next version of safety service manual | Ingenuity 64 |  | 459800093391 Rev K 459801263801 Rev B 459801263811 Rev B | October 15, 2019 1:58:15 PM GMT+08:00
CTUDT00625720 | Refresh | Closed | zhong_x | August 22, 2019 9:12:09 AM GMT+08:00 | “PhysicsQualityAssurance has stopped working” Pop Up during Perform DFS Calibration | BB_WIP9 | WIP20 | N/A | Subsystem | Blocker20 | N/A | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP20SubCom | Ovation |  | WIP20 | September 5, 2019 7:05:11 PM GMT+08:00
CTUDT00625720 | Refresh | Closed | zhong_x | August 22, 2019 9:12:09 AM GMT+08:00 | “PhysicsQualityAssurance has stopped working” Pop Up during Perform DFS Calibration | BB_WIP9 | WIP20 | N/A | Subsystem | Blocker20 | N/A | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP20SubCom | Brilliance Big Bore |  | WIP20 | September 5, 2019 7:05:11 PM GMT+08:00
CTUDT00625720 | Refresh | Closed | zhong_x | August 22, 2019 9:12:09 AM GMT+08:00 | “PhysicsQualityAssurance has stopped working” Pop Up during Perform DFS Calibration | BB_WIP9 | WIP20 | N/A | Subsystem | Blocker20 | N/A | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP20SubCom | Ingenuity 64 |  | WIP20 | September 5, 2019 7:05:11 PM GMT+08:00
CTUDT00625720 | Refresh | Closed | zhong_x | August 22, 2019 9:12:09 AM GMT+08:00 | “PhysicsQualityAssurance has stopped working” Pop Up during Perform DFS Calibration | BB_WIP9 | WIP20 | N/A | Subsystem | Blocker20 | N/A | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP20SubCom | IQon1 |  | WIP20 | September 5, 2019 7:05:11 PM GMT+08:00
CTUDT00625718 | Refresh | Closed | ding_s | August 22, 2019 9:07:47 AM GMT+08:00 | Did not find evidence of CT-NM.SRS.Service.2919 in service documents. | WIP18 |  |  | Subsystem |  | CT-NM.SRS.Service.2919 | Duplicate | lin_s | 0 |  |  | Support-Service | NO | Closed |  | Brilliance 64
CTUDT00625718 | Refresh | Closed | ding_s | August 22, 2019 9:07:47 AM GMT+08:00 | Did not find evidence of CT-NM.SRS.Service.2919 in service documents. | WIP18 |  |  | Subsystem |  | CT-NM.SRS.Service.2919 | Duplicate | lin_s | 0 |  |  | Support-Service | NO | Closed |  | Ingenuity 64
CTUDT00625716 | Refresh | Closed | ding_s | August 22, 2019 8:58:22 AM GMT+08:00 | Did not find evidence of CT-NM.SRS.SecPri.351 in service documents. | WIP18 |  |  | Subsystem |  | CT-NM.SRS.SecPri.351 | Works As Specified / Not A Defect | ding_s | 0 |  |  | Support-Service | NO | Closed |  | Brilliance 64
CTUDT00625716 | Refresh | Closed | ding_s | August 22, 2019 8:58:22 AM GMT+08:00 | Did not find evidence of CT-NM.SRS.SecPri.351 in service documents. | WIP18 |  |  | Subsystem |  | CT-NM.SRS.SecPri.351 | Works As Specified / Not A Defect | ding_s | 0 |  |  | Support-Service | NO | Closed |  | Ingenuity 64
CTUDT00625714 | Refresh | Closed | ding_f | August 17, 2019 2:03:58 PM GMT+08:00 | Refresh: Harmonized Phantom Performance Test- IMR body 1a Noise value not within the range. | WIP19 | WIP20 WIP21 | N/A | Subsystem | Blocker21:Performance Test with H Phantom;Verified | DHF170396 | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP20SubCom | Ingenuity 64 |  | WIP20 WIP21 | October 7, 2019 10:51:27 PM GMT+08:00
CTUDT00625712 | Refresh | Closed | wang_l | August 17, 2019 1:37:17 PM GMT+08:00 | Refresh: The Perfusion group was missed during importing perfusion Exam Card | WIP19 | WIP20 | Add perfusion group to Exam Card manually | Subsystem |  | NA | Fixed | prasad_r | 0 | 3 | Console Software | Console4 | NO | Closed | WIP21SubCom | Ovation |  | WIP20 | September 13, 2019 5:24:09 PM GMT+08:00
CTUDT00625712 | Refresh | Closed | wang_l | August 17, 2019 1:37:17 PM GMT+08:00 | Refresh: The Perfusion group was missed during importing perfusion Exam Card | WIP19 | WIP20 | Add perfusion group to Exam Card manually | Subsystem |  | NA | Fixed | prasad_r | 0 | 3 | Console Software | Console4 | NO | Closed | WIP21SubCom | Brilliance 64 |  | WIP20 | September 13, 2019 5:24:09 PM GMT+08:00
CTUDT00625712 | Refresh | Closed | wang_l | August 17, 2019 1:37:17 PM GMT+08:00 | Refresh: The Perfusion group was missed during importing perfusion Exam Card | WIP19 | WIP20 | Add perfusion group to Exam Card manually | Subsystem |  | NA | Fixed | prasad_r | 0 | 3 | Console Software | Console4 | NO | Closed | WIP21SubCom | Ingenuity 64 |  | WIP20 | September 13, 2019 5:24:09 PM GMT+08:00
CTUDT00625712 | Refresh | Closed | wang_l | August 17, 2019 1:37:17 PM GMT+08:00 | Refresh: The Perfusion group was missed during importing perfusion Exam Card | WIP19 | WIP20 | Add perfusion group to Exam Card manually | Subsystem |  | NA | Fixed | prasad_r | 0 | 3 | Console Software | Console4 | NO | Closed | WIP21SubCom | IQon1 |  | WIP20 | September 13, 2019 5:24:09 PM GMT+08:00
CTUDT00625710 | Refresh | Closed | zhong_x | August 17, 2019 12:19:40 PM GMT+08:00 | “The system cannot find the file specified” error occurred during A-plane Encoder calibration. | BB_WIP9 | WIP20 |  | Subsystem | Calibration;Blocker20 | n/a | Fixed | zhou_x | 0 | 2 | Gantry Calibrations | GCX | NO | Closed | WIP20SubCom | Brilliance 64 |  | WIP20 | September 3, 2019 4:30:55 PM GMT+08:00
CTUDT00625710 | Refresh | Closed | zhong_x | August 17, 2019 12:19:40 PM GMT+08:00 | “The system cannot find the file specified” error occurred during A-plane Encoder calibration. | BB_WIP9 | WIP20 |  | Subsystem | Calibration;Blocker20 | n/a | Fixed | zhou_x | 0 | 2 | Gantry Calibrations | GCX | NO | Closed | WIP20SubCom | Brilliance Big Bore |  | WIP20 | September 3, 2019 4:30:55 PM GMT+08:00
CTUDT00625710 | Refresh | Closed | zhong_x | August 17, 2019 12:19:40 PM GMT+08:00 | “The system cannot find the file specified” error occurred during A-plane Encoder calibration. | BB_WIP9 | WIP20 |  | Subsystem | Calibration;Blocker20 | n/a | Fixed | zhou_x | 0 | 2 | Gantry Calibrations | GCX | NO | Closed | WIP20SubCom | Ingenuity 64 |  | WIP20 | September 3, 2019 4:30:55 PM GMT+08:00
CTUDT00625708 | Refresh | Closed | zhong_x | August 17, 2019 10:42:36 AM GMT+08:00 | Tilt button is not working in interventional control Box | BB_WIP9 | Rev 15 | Use gantry panel for tilt change | Subsystem | IVC | Unknown | Fixed | dongping_w | 0 | 2 | CIRS_SW_Installation | Support-System | NO | Closed | SegRec | Brilliance Big Bore |  | Rev 15 | October 31, 2019 10:00:37 AM GMT+08:00
CTUDT00625706 | Refresh | Closed | zhong_x | August 17, 2019 9:46:02 AM GMT+08:00 | CANMSEgine Module has stopped working Pop Up randomly. | BB_WIP9 | WIP20 BB_WIP10 | No any function impact. Close pop up will be OK. | Subsystem |  | N/A | Fixed | Mark_Huang | 0 | 4 | Gantry Software | GCX | NO | Closed | WIP20SegReq | Ovation |  | WIP20 BB_WIP10 | September 4, 2019 1:04:35 PM GMT+08:00
CTUDT00625706 | Refresh | Closed | zhong_x | August 17, 2019 9:46:02 AM GMT+08:00 | CANMSEgine Module has stopped working Pop Up randomly. | BB_WIP9 | WIP20 BB_WIP10 | No any function impact. Close pop up will be OK. | Subsystem |  | N/A | Fixed | Mark_Huang | 0 | 4 | Gantry Software | GCX | NO | Closed | WIP20SegReq | Brilliance 64 |  | WIP20 BB_WIP10 | September 4, 2019 1:04:35 PM GMT+08:00
CTUDT00625706 | Refresh | Closed | zhong_x | August 17, 2019 9:46:02 AM GMT+08:00 | CANMSEgine Module has stopped working Pop Up randomly. | BB_WIP9 | WIP20 BB_WIP10 | No any function impact. Close pop up will be OK. | Subsystem |  | N/A | Fixed | Mark_Huang | 0 | 4 | Gantry Software | GCX | NO | Closed | WIP20SegReq | Brilliance Big Bore |  | WIP20 BB_WIP10 | September 4, 2019 1:04:35 PM GMT+08:00
CTUDT00625706 | Refresh | Closed | zhong_x | August 17, 2019 9:46:02 AM GMT+08:00 | CANMSEgine Module has stopped working Pop Up randomly. | BB_WIP9 | WIP20 BB_WIP10 | No any function impact. Close pop up will be OK. | Subsystem |  | N/A | Fixed | Mark_Huang | 0 | 4 | Gantry Software | GCX | NO | Closed | WIP20SegReq | Ingenuity 64 |  | WIP20 BB_WIP10 | September 4, 2019 1:04:35 PM GMT+08:00
CTUDT00625706 | Refresh | Closed | zhong_x | August 17, 2019 9:46:02 AM GMT+08:00 | CANMSEgine Module has stopped working Pop Up randomly. | BB_WIP9 | WIP20 BB_WIP10 | No any function impact. Close pop up will be OK. | Subsystem |  | N/A | Fixed | Mark_Huang | 0 | 4 | Gantry Software | GCX | NO | Closed | WIP20SegReq | IQon1 |  | WIP20 BB_WIP10 | September 4, 2019 1:04:35 PM GMT+08:00
CTUDT00625703 | Refresh | Closed | wang_l | August 16, 2019 4:31:23 PM GMT+08:00 | Refresh BB: Error message ""Gantry cannot initializing.."occurred during load patient | BB_WIP8 |  |  | Subsystem |  | NA | Duplicate | wang_k | 0 |  |  | GCX | NO | Closed |  | Brilliance Big Bore
CTUDT00625701 | Refresh | Closed | wang_l | August 16, 2019 2:25:38 PM GMT+08:00 | Refresh: Cannot find service manual which can cover requirement CT-NM.SRS.Service.3056 | BB_WIP8 |  |  | Subsystem |  | CT-NM.SRS.Service.3056 | Duplicate | ye_h | 0 |  |  | Support-Service | NO | Closed |  | Brilliance Big Bore
CTUDT00625699 | Refresh | Closed | wang_l | August 16, 2019 2:15:55 PM GMT+08:00 | Refresh: Cannot find service manual which can cover requirement CT-NM.SRS.SecPri.373 | BB_WIP8 |  |  | Subsystem |  |  | Duplicate | shi_b | 0 |  |  | Support-Service | NO | Closed |  | Brilliance Big Bore
CTUDT00625697 | Refresh | Closed | wang_l | August 16, 2019 1:56:13 PM GMT+08:00 | Refresh: Cannot find service manual which can cover requirement CT-NM.SRS.SecPri.374 | BB_WIP8 | 459800093391 Rev L 459801263801 Rev C 459801263811 Rev C |  | System |  | CT-NM.SRS.SecPri.374 | Fixed | li_e | 0 | 2 | Service Documentation | Support-Service | NO | Closed | NA | Brilliance 64 |  | 459800093391 Rev L 459801263801 Rev C 459801263811 Rev C | November 19, 2019 5:49:48 PM GMT+08:00
CTUDT00625697 | Refresh | Closed | wang_l | August 16, 2019 1:56:13 PM GMT+08:00 | Refresh: Cannot find service manual which can cover requirement CT-NM.SRS.SecPri.374 | BB_WIP8 | 459800093391 Rev L 459801263801 Rev C 459801263811 Rev C |  | System |  | CT-NM.SRS.SecPri.374 | Fixed | li_e | 0 | 2 | Service Documentation | Support-Service | NO | Closed | NA | Brilliance Big Bore |  | 459800093391 Rev L 459801263801 Rev C 459801263811 Rev C | November 19, 2019 5:49:48 PM GMT+08:00
CTUDT00625697 | Refresh | Closed | wang_l | August 16, 2019 1:56:13 PM GMT+08:00 | Refresh: Cannot find service manual which can cover requirement CT-NM.SRS.SecPri.374 | BB_WIP8 | 459800093391 Rev L 459801263801 Rev C 459801263811 Rev C |  | System |  | CT-NM.SRS.SecPri.374 | Fixed | li_e | 0 | 2 | Service Documentation | Support-Service | NO | Closed | NA | Ingenuity 64 |  | 459800093391 Rev L 459801263801 Rev C 459801263811 Rev C | November 19, 2019 5:49:48 PM GMT+08:00
CTUDT00625695 | Refresh | Closed | wang_l | August 16, 2019 1:20:22 PM GMT+08:00 | Refresh: Service document Remote Service Manual: The screenshot and contents did not match with Refresh system current behavior | BB_WIP8 | 459801587281 Rev B |  | System |  | NA | Fixed | li_e | 0 | 2 | Service Documentation | Support-Service | NO | Closed | BB_WIP9_SubCom | Brilliance Big Bore |  | 459801587281 Rev B | November 19, 2019 5:49:07 PM GMT+08:00
CTUDT00625693 | Refresh | Closed | wang_l | August 16, 2019 12:09:14 PM GMT+08:00 | Refresh: Service document BR Big Bore Repair & Replacement Manual: the screenshot related to UI did not match with Refresh system current behavior | BB_WIP8 | 453567142131 Rev AC | Yes, follow the manual description and software UI to perform action, even the UI is slightly different but still can be performed. | System |  | NA | Fixed | li_e | 0 | 2 | Service Documentation | Support-Service | NO | Closed | BB_WIP9_SubCom | Brilliance Big Bore |  | 453567142131 Rev AC | November 19, 2019 5:42:36 PM GMT+08:00
CTUDT00625691 | Refresh | Closed | wang_l | August 16, 2019 11:43:59 AM GMT+08:00 | Refresh: Cannot find service manual which can cover requirement CT-NM.SRS.Service.2418 | BB_WIP8 |  |  | Subsystem |  | CT-NM.SRS.Service.2418 | Works As Specified / Not A Defect | lin_s | 0 |  |  | Support-Service | NO | Closed |  | Brilliance Big Bore
CTUDT00625689 | Refresh | Closed | wang_l | August 16, 2019 11:33:08 AM GMT+08:00 | Refresh: Cannot find service manual which can cover requirement CT-NM.SRS.Service.2919 | BB_WIP8 |  |  | Subsystem |  | CT-NM.SRS.Service.2919 | Documentation Created/Updated | wang_l | 0 | 3 | Test Design | Support-System | NO | Closed | BB_WIP13_SegReq | Brilliance 64 |  |  | November 4, 2019 2:57:18 PM GMT+08:00
CTUDT00625689 | Refresh | Closed | wang_l | August 16, 2019 11:33:08 AM GMT+08:00 | Refresh: Cannot find service manual which can cover requirement CT-NM.SRS.Service.2919 | BB_WIP8 |  |  | Subsystem |  | CT-NM.SRS.Service.2919 | Documentation Created/Updated | wang_l | 0 | 3 | Test Design | Support-System | NO | Closed | BB_WIP13_SegReq | Brilliance Big Bore |  |  | November 4, 2019 2:57:18 PM GMT+08:00
CTUDT00625689 | Refresh | Closed | wang_l | August 16, 2019 11:33:08 AM GMT+08:00 | Refresh: Cannot find service manual which can cover requirement CT-NM.SRS.Service.2919 | BB_WIP8 |  |  | Subsystem |  | CT-NM.SRS.Service.2919 | Documentation Created/Updated | wang_l | 0 | 3 | Test Design | Support-System | NO | Closed | BB_WIP13_SegReq | Ingenuity 64 |  |  | November 4, 2019 2:57:18 PM GMT+08:00
CTUDT00625687 | Refresh | Closed | wang_l | August 16, 2019 11:06:23 AM GMT+08:00 | Refresh: Cannot find service manual which can cover requirement CT-NM.SRS.SecPri.351 | BB_WIP8 |  |  | Subsystem |  | CT-NM.SRS.SecPri.351 | Works As Specified / Not A Defect | lin_s | 0 |  |  | Support-Service | NO | Closed |  | Brilliance Big Bore
CTUDT00625685 | Refresh | Closed | ding_s | August 16, 2019 9:45:00 AM GMT+08:00 | There is no available Gated Axial protocol with IMR recon mode. | BB_WIP9 | DHF298555 rev09 |  | System | requirement correction | CT-NM.SRS.Workflow.1873;CT-NM.SRS.ProductConfig.145 | Fixed | zhong_x | 0 | 2 | Design Documentation | Support-System | NO | Closed | SubCom-25-Nov-2019 | Brilliance Big Bore |  | DHF298555 rev09 | December 3, 2019 4:38:26 PM GMT+08:00
CTUDT00625683 | Refresh | Closed | zhong_x | August 16, 2019 9:32:23 AM GMT+08:00 | Once installation got stuck, “Windows Script error” pop up in next attempt and cannot be bypassed | BB_WIP8.1 | WIP20 |  | Subsystem |  | STPS-SSRS-1154 | Fixed | zhang_t | 0 | 3 | Service Documentation | DI | NO | Closed | WIP20SubCom | Ovation |  | WIP20 | October 1, 2019 1:38:10 PM GMT+08:00
CTUDT00625683 | Refresh | Closed | zhong_x | August 16, 2019 9:32:23 AM GMT+08:00 | Once installation got stuck, “Windows Script error” pop up in next attempt and cannot be bypassed | BB_WIP8.1 | WIP20 |  | Subsystem |  | STPS-SSRS-1154 | Fixed | zhang_t | 0 | 3 | Service Documentation | DI | NO | Closed | WIP20SubCom | Brilliance 64 |  | WIP20 | October 1, 2019 1:38:10 PM GMT+08:00
CTUDT00625683 | Refresh | Closed | zhong_x | August 16, 2019 9:32:23 AM GMT+08:00 | Once installation got stuck, “Windows Script error” pop up in next attempt and cannot be bypassed | BB_WIP8.1 | WIP20 |  | Subsystem |  | STPS-SSRS-1154 | Fixed | zhang_t | 0 | 3 | Service Documentation | DI | NO | Closed | WIP20SubCom | Brilliance Big Bore |  | WIP20 | October 1, 2019 1:38:10 PM GMT+08:00
CTUDT00625683 | Refresh | Closed | zhong_x | August 16, 2019 9:32:23 AM GMT+08:00 | Once installation got stuck, “Windows Script error” pop up in next attempt and cannot be bypassed | BB_WIP8.1 | WIP20 |  | Subsystem |  | STPS-SSRS-1154 | Fixed | zhang_t | 0 | 3 | Service Documentation | DI | NO | Closed | WIP20SubCom | Ingenuity 64 |  | WIP20 | October 1, 2019 1:38:10 PM GMT+08:00
CTUDT00625683 | Refresh | Closed | zhong_x | August 16, 2019 9:32:23 AM GMT+08:00 | Once installation got stuck, “Windows Script error” pop up in next attempt and cannot be bypassed | BB_WIP8.1 | WIP20 |  | Subsystem |  | STPS-SSRS-1154 | Fixed | zhang_t | 0 | 3 | Service Documentation | DI | NO | Closed | WIP20SubCom | IQon1 |  | WIP20 | October 1, 2019 1:38:10 PM GMT+08:00
CTUDT00625680 | Refresh | Closed | golzban_v | August 15, 2019 8:59:02 PM GMT+08:00 | CT-NM.SRS.Modes.296: Incorrect phrasing and contradiction with CT-NM.SRS.Modes.295 | N\A | iCT, IQon SRS DHF298556 Rev.09 and DHF298290 Rev.07 ING, BR64 SRS DHF298554 Rev08 |  | Subsystem | SRS | CT-NM.SRS.Modes.296 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP20SubCom | Ovation |  | iCT, IQon SRS DHF298556 Rev.09 and DHF298290 Rev.07 ING, BR64 SRS DHF298554 Rev08 | November 5, 2019 9:17:29 PM GMT+08:00
CTUDT00625680 | Refresh | Closed | golzban_v | August 15, 2019 8:59:02 PM GMT+08:00 | CT-NM.SRS.Modes.296: Incorrect phrasing and contradiction with CT-NM.SRS.Modes.295 | N\A | iCT, IQon SRS DHF298556 Rev.09 and DHF298290 Rev.07 ING, BR64 SRS DHF298554 Rev08 |  | Subsystem | SRS | CT-NM.SRS.Modes.296 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP20SubCom | Brilliance 64 |  | iCT, IQon SRS DHF298556 Rev.09 and DHF298290 Rev.07 ING, BR64 SRS DHF298554 Rev08 | November 5, 2019 9:17:29 PM GMT+08:00
CTUDT00625680 | Refresh | Closed | golzban_v | August 15, 2019 8:59:02 PM GMT+08:00 | CT-NM.SRS.Modes.296: Incorrect phrasing and contradiction with CT-NM.SRS.Modes.295 | N\A | iCT, IQon SRS DHF298556 Rev.09 and DHF298290 Rev.07 ING, BR64 SRS DHF298554 Rev08 |  | Subsystem | SRS | CT-NM.SRS.Modes.296 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP20SubCom | Ingenuity 64 |  | iCT, IQon SRS DHF298556 Rev.09 and DHF298290 Rev.07 ING, BR64 SRS DHF298554 Rev08 | November 5, 2019 9:17:29 PM GMT+08:00
CTUDT00625680 | Refresh | Closed | golzban_v | August 15, 2019 8:59:02 PM GMT+08:00 | CT-NM.SRS.Modes.296: Incorrect phrasing and contradiction with CT-NM.SRS.Modes.295 | N\A | iCT, IQon SRS DHF298556 Rev.09 and DHF298290 Rev.07 ING, BR64 SRS DHF298554 Rev08 |  | Subsystem | SRS | CT-NM.SRS.Modes.296 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP20SubCom | IQon1 |  | iCT, IQon SRS DHF298556 Rev.09 and DHF298290 Rev.07 ING, BR64 SRS DHF298554 Rev08 | November 5, 2019 9:17:29 PM GMT+08:00
CTUDT00625679 | Refresh | Closed | krisher_i | August 15, 2019 8:38:05 PM GMT+08:00 | Refresh- End Exam takes too long | WIP18 |  |  | Subsystem | Performance |  | Works As Specified / Not A Defect | tamari_o | 1 |  |  | Support-System | NO | Closed | TBR-2-WIP24 | IQon1
CTUDT00625677 | Refresh | Closed | golzban_v | August 15, 2019 7:54:32 PM GMT+08:00 | CT-NM.SRS.Modes.1483 contradicts CT-NM.SRS.Modes.1484 | N/A | DHF298554 Rev08 iCT, SRS DHF298553 Rev.08 IQon SRS DHF298556 Rev.09 |  | Subsystem | Verified1 | Modes.1483, Modes.1808 | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | NA | Ovation |  | DHF298554 Rev08 iCT, SRS DHF298553 Rev.08 IQon SRS DHF298556 Rev.09 | November 6, 2019 1:52:48 PM GMT+08:00
CTUDT00625677 | Refresh | Closed | golzban_v | August 15, 2019 7:54:32 PM GMT+08:00 | CT-NM.SRS.Modes.1483 contradicts CT-NM.SRS.Modes.1484 | N/A | DHF298554 Rev08 iCT, SRS DHF298553 Rev.08 IQon SRS DHF298556 Rev.09 |  | Subsystem | Verified1 | Modes.1483, Modes.1808 | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | NA | Brilliance 64 |  | DHF298554 Rev08 iCT, SRS DHF298553 Rev.08 IQon SRS DHF298556 Rev.09 | November 6, 2019 1:52:48 PM GMT+08:00
CTUDT00625677 | Refresh | Closed | golzban_v | August 15, 2019 7:54:32 PM GMT+08:00 | CT-NM.SRS.Modes.1483 contradicts CT-NM.SRS.Modes.1484 | N/A | DHF298554 Rev08 iCT, SRS DHF298553 Rev.08 IQon SRS DHF298556 Rev.09 |  | Subsystem | Verified1 | Modes.1483, Modes.1808 | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | NA | Ingenuity 64 |  | DHF298554 Rev08 iCT, SRS DHF298553 Rev.08 IQon SRS DHF298556 Rev.09 | November 6, 2019 1:52:48 PM GMT+08:00
CTUDT00625677 | Refresh | Closed | golzban_v | August 15, 2019 7:54:32 PM GMT+08:00 | CT-NM.SRS.Modes.1483 contradicts CT-NM.SRS.Modes.1484 | N/A | DHF298554 Rev08 iCT, SRS DHF298553 Rev.08 IQon SRS DHF298556 Rev.09 |  | Subsystem | Verified1 | Modes.1483, Modes.1808 | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | NA | IQon1 |  | DHF298554 Rev08 iCT, SRS DHF298553 Rev.08 IQon SRS DHF298556 Rev.09 | November 6, 2019 1:52:48 PM GMT+08:00
CTUDT00625675 | Refresh | Closed | 320050709 | August 15, 2019 1:38:54 PM GMT+08:00 | Miss the information and warning in the user manual | IFU 459801559891 | WIP20 | NA | Subsystem |  | RMM.503 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | YES | Closed | Before SVAL | Ovation |  | WIP20 | August 27, 2019 1:12:51 PM GMT+08:00
CTUDT00625675 | Refresh | Closed | 320050709 | August 15, 2019 1:38:54 PM GMT+08:00 | Miss the information and warning in the user manual | IFU 459801559891 | WIP20 | NA | Subsystem |  | RMM.503 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | YES | Closed | Before SVAL | Brilliance 64 |  | WIP20 | August 27, 2019 1:12:51 PM GMT+08:00
CTUDT00625675 | Refresh | Closed | 320050709 | August 15, 2019 1:38:54 PM GMT+08:00 | Miss the information and warning in the user manual | IFU 459801559891 | WIP20 | NA | Subsystem |  | RMM.503 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | YES | Closed | Before SVAL | Brilliance Big Bore |  | WIP20 | August 27, 2019 1:12:51 PM GMT+08:00
CTUDT00625675 | Refresh | Closed | 320050709 | August 15, 2019 1:38:54 PM GMT+08:00 | Miss the information and warning in the user manual | IFU 459801559891 | WIP20 | NA | Subsystem |  | RMM.503 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | YES | Closed | Before SVAL | Ingenuity 64 |  | WIP20 | August 27, 2019 1:12:51 PM GMT+08:00
CTUDT00625675 | Refresh | Closed | 320050709 | August 15, 2019 1:38:54 PM GMT+08:00 | Miss the information and warning in the user manual | IFU 459801559891 | WIP20 | NA | Subsystem |  | RMM.503 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | YES | Closed | Before SVAL | IQon1 |  | WIP20 | August 27, 2019 1:12:51 PM GMT+08:00
CTUDT00625672 | Refresh | Closed | cai_t | August 14, 2019 10:04:42 PM GMT+08:00 | The footswitch color is not right. | BB_WIP8 | DHF317593 Rev05 |  | Subsystem |  | DHF173447-2573, DHF173447-2770, | Fixed | zhang_je | 0 | 2 | Requirements Documentation | Couch | NO | Closed | Couch SSRS next version | Brilliance 64 |  | DHF317593 Rev05 | October 22, 2019 2:37:20 PM GMT+08:00
CTUDT00625672 | Refresh | Closed | cai_t | August 14, 2019 10:04:42 PM GMT+08:00 | The footswitch color is not right. | BB_WIP8 | DHF317593 Rev05 |  | Subsystem |  | DHF173447-2573, DHF173447-2770, | Fixed | zhang_je | 0 | 2 | Requirements Documentation | Couch | NO | Closed | Couch SSRS next version | Brilliance Big Bore |  | DHF317593 Rev05 | October 22, 2019 2:37:20 PM GMT+08:00
CTUDT00625672 | Refresh | Closed | cai_t | August 14, 2019 10:04:42 PM GMT+08:00 | The footswitch color is not right. | BB_WIP8 | DHF317593 Rev05 |  | Subsystem |  | DHF173447-2573, DHF173447-2770, | Fixed | zhang_je | 0 | 2 | Requirements Documentation | Couch | NO | Closed | Couch SSRS next version | Ingenuity 64 |  | DHF317593 Rev05 | October 22, 2019 2:37:20 PM GMT+08:00
CTUDT00625669 | Refresh | Closed | panchami_pradeep | August 14, 2019 7:21:57 PM GMT+08:00 | Refresh: Installation time in scanner exceeds 90 minutes. | WIP19 Sanity | WIP20 | No workaround | Subsystem |  | NA | Fixed | ravindra_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP20SubCom | Ovation |  | WIP20 | August 30, 2019 5:43:35 PM GMT+08:00
CTUDT00625669 | Refresh | Closed | panchami_pradeep | August 14, 2019 7:21:57 PM GMT+08:00 | Refresh: Installation time in scanner exceeds 90 minutes. | WIP19 Sanity | WIP20 | No workaround | Subsystem |  | NA | Fixed | ravindra_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP20SubCom | Brilliance 64 |  | WIP20 | August 30, 2019 5:43:35 PM GMT+08:00
CTUDT00625669 | Refresh | Closed | panchami_pradeep | August 14, 2019 7:21:57 PM GMT+08:00 | Refresh: Installation time in scanner exceeds 90 minutes. | WIP19 Sanity | WIP20 | No workaround | Subsystem |  | NA | Fixed | ravindra_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP20SubCom | Ingenuity 64 |  | WIP20 | August 30, 2019 5:43:35 PM GMT+08:00
CTUDT00625669 | Refresh | Closed | panchami_pradeep | August 14, 2019 7:21:57 PM GMT+08:00 | Refresh: Installation time in scanner exceeds 90 minutes. | WIP19 Sanity | WIP20 | No workaround | Subsystem |  | NA | Fixed | ravindra_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP20SubCom | IQon1 |  | WIP20 | August 30, 2019 5:43:35 PM GMT+08:00
CTUDT00625662 | Refresh | Closed | g_di | August 14, 2019 6:34:56 PM GMT+08:00 | User is able to login with the Special characters which are not allowed to in the “Emergency User field” | WIP18 | WIP20 BB_WIP10 | No workaround but user can delete the special character entered. | Subsystem |  | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP20SubCom | Ovation |  | WIP20 BB_WIP10 | September 16, 2019 3:35:53 PM GMT+08:00
CTUDT00625662 | Refresh | Closed | g_di | August 14, 2019 6:34:56 PM GMT+08:00 | User is able to login with the Special characters which are not allowed to in the “Emergency User field” | WIP18 | WIP20 BB_WIP10 | No workaround but user can delete the special character entered. | Subsystem |  | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP20SubCom | Brilliance 64 |  | WIP20 BB_WIP10 | September 16, 2019 3:35:53 PM GMT+08:00
CTUDT00625662 | Refresh | Closed | g_di | August 14, 2019 6:34:56 PM GMT+08:00 | User is able to login with the Special characters which are not allowed to in the “Emergency User field” | WIP18 | WIP20 BB_WIP10 | No workaround but user can delete the special character entered. | Subsystem |  | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP20SubCom | Brilliance Big Bore |  | WIP20 BB_WIP10 | September 16, 2019 3:35:53 PM GMT+08:00
CTUDT00625662 | Refresh | Closed | g_di | August 14, 2019 6:34:56 PM GMT+08:00 | User is able to login with the Special characters which are not allowed to in the “Emergency User field” | WIP18 | WIP20 BB_WIP10 | No workaround but user can delete the special character entered. | Subsystem |  | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP20SubCom | Ingenuity 64 |  | WIP20 BB_WIP10 | September 16, 2019 3:35:53 PM GMT+08:00
CTUDT00625662 | Refresh | Closed | g_di | August 14, 2019 6:34:56 PM GMT+08:00 | User is able to login with the Special characters which are not allowed to in the “Emergency User field” | WIP18 | WIP20 BB_WIP10 | No workaround but user can delete the special character entered. | Subsystem |  | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP20SubCom | IQon1 |  | WIP20 BB_WIP10 | September 16, 2019 3:35:53 PM GMT+08:00
CTUDT00625658 | Refresh | Closed | jotawar_d | August 14, 2019 6:30:38 PM GMT+08:00 | BB Refresh : The message was truncated in German Language on field “mAs for Average Adult:” | WIP19 | WIP20 |  | Subsystem |  | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP20SubCom | Ovation |  | WIP20 | August 23, 2019 3:29:35 PM GMT+08:00
CTUDT00625658 | Refresh | Closed | jotawar_d | August 14, 2019 6:30:38 PM GMT+08:00 | BB Refresh : The message was truncated in German Language on field “mAs for Average Adult:” | WIP19 | WIP20 |  | Subsystem |  | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP20SubCom | Brilliance 64 |  | WIP20 | August 23, 2019 3:29:35 PM GMT+08:00
CTUDT00625658 | Refresh | Closed | jotawar_d | August 14, 2019 6:30:38 PM GMT+08:00 | BB Refresh : The message was truncated in German Language on field “mAs for Average Adult:” | WIP19 | WIP20 |  | Subsystem |  | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP20SubCom | Ingenuity 64 |  | WIP20 | August 23, 2019 3:29:35 PM GMT+08:00
CTUDT00625658 | Refresh | Closed | jotawar_d | August 14, 2019 6:30:38 PM GMT+08:00 | BB Refresh : The message was truncated in German Language on field “mAs for Average Adult:” | WIP19 | WIP20 |  | Subsystem |  | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP20SubCom | IQon1 |  | WIP20 | August 23, 2019 3:29:35 PM GMT+08:00
CTUDT00625654 | Refresh | Closed | 320053770 | August 14, 2019 5:56:38 PM GMT+08:00 | Refresh - Wip18: IQon&ICt: OS services information for the system are not consistent with “DHF314402_Appendix A - Services Configured v02” | WIP18 |  |  | Subsystem |  |  | Duplicate | dasgupta_s | 0 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00625654 | Refresh | Closed | 320053770 | August 14, 2019 5:56:38 PM GMT+08:00 | Refresh - Wip18: IQon&ICt: OS services information for the system are not consistent with “DHF314402_Appendix A - Services Configured v02” | WIP18 |  |  | Subsystem |  |  | Duplicate | dasgupta_s | 0 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00625648 | Refresh | Closed | sowmya_k | August 14, 2019 4:55:41 PM GMT+08:00 | “ Phase “ field is missing in show-all under ECG gating section | WIP8 | BB_WIP10 | As issue happen in ECM user can use any cardiac exam card to changed to desired cardiac scan rather then using Abdomen exam card . | Subsystem |  | N/A | Fixed | pandey_a | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP10_SubCom | Brilliance Big Bore |  | BB_WIP10 | August 23, 2019 8:34:52 PM GMT+08:00
CTUDT00625647 | Refresh | Closed | wang_l | August 14, 2019 4:55:03 PM GMT+08:00 | Refresh Ingenuity: Constancy Test: Stage Resolution 3 didn’t pass | WIP18 |  | Re-run acceptance test | Subsystem | Blocker: Constancy Test with Legacy Phantom | N/A | Works As Specified / Not A Defect | zhang_l | 0 |  |  | Support-System | NO | Closed |  | Ingenuity 64
CTUDT00625646 | Refresh | Closed | sowmya_k | August 14, 2019 4:53:47 PM GMT+08:00 | BB Refresh : Pitch, Rotation Time, Scan Time fields were not Localized in UI on Dose right warning message | WIP19 | WIP20 |  | Subsystem |  | NA | Fixed | challa_s | 0 | 3 | Console Software | Console4 | YES | Closed | WIP20SubCom | Ovation |  | WIP20 | August 23, 2019 2:46:33 PM GMT+08:00
CTUDT00625646 | Refresh | Closed | sowmya_k | August 14, 2019 4:53:47 PM GMT+08:00 | BB Refresh : Pitch, Rotation Time, Scan Time fields were not Localized in UI on Dose right warning message | WIP19 | WIP20 |  | Subsystem |  | NA | Fixed | challa_s | 0 | 3 | Console Software | Console4 | YES | Closed | WIP20SubCom | Brilliance 64 |  | WIP20 | August 23, 2019 2:46:33 PM GMT+08:00
CTUDT00625646 | Refresh | Closed | sowmya_k | August 14, 2019 4:53:47 PM GMT+08:00 | BB Refresh : Pitch, Rotation Time, Scan Time fields were not Localized in UI on Dose right warning message | WIP19 | WIP20 |  | Subsystem |  | NA | Fixed | challa_s | 0 | 3 | Console Software | Console4 | YES | Closed | WIP20SubCom | Ingenuity 64 |  | WIP20 | August 23, 2019 2:46:33 PM GMT+08:00
CTUDT00625646 | Refresh | Closed | sowmya_k | August 14, 2019 4:53:47 PM GMT+08:00 | BB Refresh : Pitch, Rotation Time, Scan Time fields were not Localized in UI on Dose right warning message | WIP19 | WIP20 |  | Subsystem |  | NA | Fixed | challa_s | 0 | 3 | Console Software | Console4 | YES | Closed | WIP20SubCom | IQon1 |  | WIP20 | August 23, 2019 2:46:33 PM GMT+08:00
CTUDT00625644 | Refresh | Closed | 320053770 | August 14, 2019 4:34:41 PM GMT+08:00 | Refresh : O-Level user cannot connect the HOST to an external NTP server. | WIP18 | DHF298555 Rev08 DHF298553 Rev 08 DHF298556 Rev 09 DHF298554 Rev08 | N/A | Subsystem | NTP | STPS-SSRS-1214 , STPS-SSRS-1229, STPS-SSRS-1213 | Fixed | bandil_j | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | SRS Update | Ovation |  | DHF298555 Rev08 DHF298553 Rev 08 DHF298556 Rev 09 DHF298554 Rev08 | November 6, 2019 2:35:59 PM GMT+08:00
CTUDT00625644 | Refresh | Closed | 320053770 | August 14, 2019 4:34:41 PM GMT+08:00 | Refresh : O-Level user cannot connect the HOST to an external NTP server. | WIP18 | DHF298555 Rev08 DHF298553 Rev 08 DHF298556 Rev 09 DHF298554 Rev08 | N/A | Subsystem | NTP | STPS-SSRS-1214 , STPS-SSRS-1229, STPS-SSRS-1213 | Fixed | bandil_j | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | SRS Update | Brilliance 64 |  | DHF298555 Rev08 DHF298553 Rev 08 DHF298556 Rev 09 DHF298554 Rev08 | November 6, 2019 2:35:59 PM GMT+08:00
CTUDT00625644 | Refresh | Closed | 320053770 | August 14, 2019 4:34:41 PM GMT+08:00 | Refresh : O-Level user cannot connect the HOST to an external NTP server. | WIP18 | DHF298555 Rev08 DHF298553 Rev 08 DHF298556 Rev 09 DHF298554 Rev08 | N/A | Subsystem | NTP | STPS-SSRS-1214 , STPS-SSRS-1229, STPS-SSRS-1213 | Fixed | bandil_j | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | SRS Update | Brilliance Big Bore |  | DHF298555 Rev08 DHF298553 Rev 08 DHF298556 Rev 09 DHF298554 Rev08 | November 6, 2019 2:35:59 PM GMT+08:00
CTUDT00625644 | Refresh | Closed | 320053770 | August 14, 2019 4:34:41 PM GMT+08:00 | Refresh : O-Level user cannot connect the HOST to an external NTP server. | WIP18 | DHF298555 Rev08 DHF298553 Rev 08 DHF298556 Rev 09 DHF298554 Rev08 | N/A | Subsystem | NTP | STPS-SSRS-1214 , STPS-SSRS-1229, STPS-SSRS-1213 | Fixed | bandil_j | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | SRS Update | Ingenuity 64 |  | DHF298555 Rev08 DHF298553 Rev 08 DHF298556 Rev 09 DHF298554 Rev08 | November 6, 2019 2:35:59 PM GMT+08:00
CTUDT00625644 | Refresh | Closed | 320053770 | August 14, 2019 4:34:41 PM GMT+08:00 | Refresh : O-Level user cannot connect the HOST to an external NTP server. | WIP18 | DHF298555 Rev08 DHF298553 Rev 08 DHF298556 Rev 09 DHF298554 Rev08 | N/A | Subsystem | NTP | STPS-SSRS-1214 , STPS-SSRS-1229, STPS-SSRS-1213 | Fixed | bandil_j | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | SRS Update | IQon1 |  | DHF298555 Rev08 DHF298553 Rev 08 DHF298556 Rev 09 DHF298554 Rev08 | November 6, 2019 2:35:59 PM GMT+08:00
CTUDT00625638 | Refresh | Closed | wang_l | August 14, 2019 10:29:25 AM GMT+08:00 | Refresh: Filters displayed in View Logs as Markers tool does not match with requirement | BB_WIP8 | DHF298554 Rev08 DHF298556 Rev 09 DHF298553 Rev 08 DHF298555 Rev08 |  | Subsystem |  | CT-NM.SRS.Service.2749;CT-NM.SRS.Service.1895 | Fixed | bandil_j | 0 | 3 | Requirements Documentation | Support-System | NO | Closed | Requirements | Ovation |  | DHF298554 Rev08 DHF298556 Rev 09 DHF298553 Rev 08 DHF298555 Rev08 | November 6, 2019 2:25:04 PM GMT+08:00
CTUDT00625638 | Refresh | Closed | wang_l | August 14, 2019 10:29:25 AM GMT+08:00 | Refresh: Filters displayed in View Logs as Markers tool does not match with requirement | BB_WIP8 | DHF298554 Rev08 DHF298556 Rev 09 DHF298553 Rev 08 DHF298555 Rev08 |  | Subsystem |  | CT-NM.SRS.Service.2749;CT-NM.SRS.Service.1895 | Fixed | bandil_j | 0 | 3 | Requirements Documentation | Support-System | NO | Closed | Requirements | Brilliance 64 |  | DHF298554 Rev08 DHF298556 Rev 09 DHF298553 Rev 08 DHF298555 Rev08 | November 6, 2019 2:25:04 PM GMT+08:00
CTUDT00625638 | Refresh | Closed | wang_l | August 14, 2019 10:29:25 AM GMT+08:00 | Refresh: Filters displayed in View Logs as Markers tool does not match with requirement | BB_WIP8 | DHF298554 Rev08 DHF298556 Rev 09 DHF298553 Rev 08 DHF298555 Rev08 |  | Subsystem |  | CT-NM.SRS.Service.2749;CT-NM.SRS.Service.1895 | Fixed | bandil_j | 0 | 3 | Requirements Documentation | Support-System | NO | Closed | Requirements | Brilliance Big Bore |  | DHF298554 Rev08 DHF298556 Rev 09 DHF298553 Rev 08 DHF298555 Rev08 | November 6, 2019 2:25:04 PM GMT+08:00
CTUDT00625638 | Refresh | Closed | wang_l | August 14, 2019 10:29:25 AM GMT+08:00 | Refresh: Filters displayed in View Logs as Markers tool does not match with requirement | BB_WIP8 | DHF298554 Rev08 DHF298556 Rev 09 DHF298553 Rev 08 DHF298555 Rev08 |  | Subsystem |  | CT-NM.SRS.Service.2749;CT-NM.SRS.Service.1895 | Fixed | bandil_j | 0 | 3 | Requirements Documentation | Support-System | NO | Closed | Requirements | Ingenuity 64 |  | DHF298554 Rev08 DHF298556 Rev 09 DHF298553 Rev 08 DHF298555 Rev08 | November 6, 2019 2:25:04 PM GMT+08:00
CTUDT00625638 | Refresh | Closed | wang_l | August 14, 2019 10:29:25 AM GMT+08:00 | Refresh: Filters displayed in View Logs as Markers tool does not match with requirement | BB_WIP8 | DHF298554 Rev08 DHF298556 Rev 09 DHF298553 Rev 08 DHF298555 Rev08 |  | Subsystem |  | CT-NM.SRS.Service.2749;CT-NM.SRS.Service.1895 | Fixed | bandil_j | 0 | 3 | Requirements Documentation | Support-System | NO | Closed | Requirements | IQon1 |  | DHF298554 Rev08 DHF298556 Rev 09 DHF298553 Rev 08 DHF298555 Rev08 | November 6, 2019 2:25:04 PM GMT+08:00
CTUDT00625635 | Refresh | Closed | golzban_v | August 13, 2019 9:35:59 PM GMT+08:00 | Wrong target release value in CT-NM.SRS.Modes.263 | N/A | IQon SRS DHF298556 Rev.09 | N/A | Subsystem |  | CT-NM.SRS.Modes.263 | Fixed | guy_sheaffer | 0 | 3 | Requirements Documentation | Support-System | NO | Closed | SRS | IQon1 |  | IQon SRS DHF298556 Rev.09 | November 5, 2019 2:53:42 PM GMT+08:00
CTUDT00625634 | Refresh | Closed | Valeria_Kashtanov | August 13, 2019 9:26:28 PM GMT+08:00 | Refresh: Calibrations: ACR Test: Wrong scan parameters: FOV and Window W | WIP18 | WIP20 | N/A | Subsystem | ACR Test;Blocker20 | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP20SegReq | Ovation |  | WIP20 | August 25, 2019 7:32:05 PM GMT+08:00
CTUDT00625633 | Refresh | Closed | jose_s | August 13, 2019 8:35:50 PM GMT+08:00 | Refresh Pre-Val:after type"lung" in the "label" of parameter interface of Exam card ,only click enter button can make it work | BB_WIP8 | BB_WIP10 | After editing the label, user has to press Enter key or press any where else in UI other than other result series for applying the changes. | Subsystem |  | None | Fixed | tejaswini_nidamanuri | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP10_SubCom | Brilliance Big Bore |  | BB_WIP10 | August 22, 2019 6:40:59 PM GMT+08:00
CTUDT00625626 | Refresh | Closed | lin_s | August 13, 2019 3:27:58 PM GMT+08:00 | Vietnamese label requirement need be defined in SRS and SSRS | WIP18 | ECR-070903 |  | Subsystem |  | SRS Update | Does Not Apply | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | SRS/SSRS | Ovation |  | ECR-070903 | November 7, 2019 2:52:30 PM GMT+08:00
CTUDT00625626 | Refresh | Closed | lin_s | August 13, 2019 3:27:58 PM GMT+08:00 | Vietnamese label requirement need be defined in SRS and SSRS | WIP18 | ECR-070903 |  | Subsystem |  | SRS Update | Does Not Apply | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | SRS/SSRS | Brilliance 64 |  | ECR-070903 | November 7, 2019 2:52:30 PM GMT+08:00
CTUDT00625626 | Refresh | Closed | lin_s | August 13, 2019 3:27:58 PM GMT+08:00 | Vietnamese label requirement need be defined in SRS and SSRS | WIP18 | ECR-070903 |  | Subsystem |  | SRS Update | Does Not Apply | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | SRS/SSRS | Ingenuity 64 |  | ECR-070903 | November 7, 2019 2:52:30 PM GMT+08:00
CTUDT00625623 | Refresh | Closed | jose_s | August 13, 2019 1:31:03 PM GMT+08:00 | Refresh: System error displayed once click End Exam button | BB_WIP9 | BB_WIP10 | User can press "OK" on system error message and continue working. | Subsystem |  | None | Fixed | tejaswini_nidamanuri | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP10_SubCom | Brilliance Big Bore |  | BB_WIP10 | August 22, 2019 6:23:11 PM GMT+08:00
CTUDT00625619 | Refresh | Closed | richards_b | August 13, 2019 6:10:39 AM GMT+08:00 | Ing: The OffFocal Table name for Child is set to Adult instead of Child | WIP19 | WIP20 | N/A | Subsystem | Blocker20 | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP20SubCom | Ingenuity 64 |  | WIP20 | September 8, 2019 4:34:04 PM GMT+08:00
CTUDT00625617 | Refresh | Closed | richards_b | August 13, 2019 5:37:37 AM GMT+08:00 | Spatial Resolution Acceptance Tolerance Updates with Harmonized Phantom | WIP19 | WIP20 |  | Subsystem | Blocker20; Verified1 | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP20SubCom | Ovation |  | WIP20 | September 19, 2019 3:55:03 PM GMT+08:00
CTUDT00625617 | Refresh | Closed | richards_b | August 13, 2019 5:37:37 AM GMT+08:00 | Spatial Resolution Acceptance Tolerance Updates with Harmonized Phantom | WIP19 | WIP20 |  | Subsystem | Blocker20; Verified1 | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP20SubCom | Brilliance 64 |  | WIP20 | September 19, 2019 3:55:03 PM GMT+08:00
CTUDT00625617 | Refresh | Closed | richards_b | August 13, 2019 5:37:37 AM GMT+08:00 | Spatial Resolution Acceptance Tolerance Updates with Harmonized Phantom | WIP19 | WIP20 |  | Subsystem | Blocker20; Verified1 | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP20SubCom | Brilliance Big Bore |  | WIP20 | September 19, 2019 3:55:03 PM GMT+08:00
CTUDT00625617 | Refresh | Closed | richards_b | August 13, 2019 5:37:37 AM GMT+08:00 | Spatial Resolution Acceptance Tolerance Updates with Harmonized Phantom | WIP19 | WIP20 |  | Subsystem | Blocker20; Verified1 | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP20SubCom | Ingenuity 64 |  | WIP20 | September 19, 2019 3:55:03 PM GMT+08:00
CTUDT00625617 | Refresh | Closed | richards_b | August 13, 2019 5:37:37 AM GMT+08:00 | Spatial Resolution Acceptance Tolerance Updates with Harmonized Phantom | WIP19 | WIP20 |  | Subsystem | Blocker20; Verified1 | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP20SubCom | IQon1 |  | WIP20 | September 19, 2019 3:55:03 PM GMT+08:00
CTUDT00625615 | Refresh | Closed | levy_s | August 12, 2019 9:28:22 PM GMT+08:00 | Refresh: Scanning fifteen patients takes more than one hour | WIP18 |  |  | Subsystem | Performance |  | Works As Specified / Not A Defect | tamari_o | 1 |  |  | Support-System | NO | Closed | TBR-2-WIP19 | IQon1
CTUDT00625613 | Refresh | Closed | lin_s | August 12, 2019 9:21:25 PM GMT+08:00 | No STPS SSRS link to Serivce.3061 for Refresh ING, BR64 | WIP19 | WIP19 |  | Subsystem |  | SRS.Serivce.3061 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | SRS DHF298554 Rev07 | Brilliance 64 |  | WIP19 | August 27, 2019 5:58:02 PM GMT+08:00
CTUDT00625613 | Refresh | Closed | lin_s | August 12, 2019 9:21:25 PM GMT+08:00 | No STPS SSRS link to Serivce.3061 for Refresh ING, BR64 | WIP19 | WIP19 |  | Subsystem |  | SRS.Serivce.3061 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | SRS DHF298554 Rev07 | Ingenuity 64 |  | WIP19 | August 27, 2019 5:58:02 PM GMT+08:00
CTUDT00625611 | Refresh | Closed | levy_s | August 12, 2019 9:07:26 PM GMT+08:00 | Refresh: First preview image and first reconstructed image are presented with delay | WIP18 |  |  | Subsystem | Performance |  | Works As Specified / Not A Defect | tamari_o | 1 |  |  | Support-System | NO | Closed |  | Ovation
CTUDT00625608 | Refresh | Closed | gT_s | August 12, 2019 7:39:16 PM GMT+08:00 | Refresh : The localization is not working for the different languages while performing the remote connection from PRS. | WIP18 | DHF318051 Rev 21 |  | Subsystem |  | NA | Works As Specified / Not A Defect | jindal_k | 0 | 2 | Test Design | Service Tools | NO | Closed | WIP20SubCom | Ovation |  | DHF318051 Rev 21 | October 16, 2019 1:53:34 PM GMT+08:00
CTUDT00625608 | Refresh | Closed | gT_s | August 12, 2019 7:39:16 PM GMT+08:00 | Refresh : The localization is not working for the different languages while performing the remote connection from PRS. | WIP18 | DHF318051 Rev 21 |  | Subsystem |  | NA | Works As Specified / Not A Defect | jindal_k | 0 | 2 | Test Design | Service Tools | NO | Closed | WIP20SubCom | Brilliance 64 |  | DHF318051 Rev 21 | October 16, 2019 1:53:34 PM GMT+08:00
CTUDT00625608 | Refresh | Closed | gT_s | August 12, 2019 7:39:16 PM GMT+08:00 | Refresh : The localization is not working for the different languages while performing the remote connection from PRS. | WIP18 | DHF318051 Rev 21 |  | Subsystem |  | NA | Works As Specified / Not A Defect | jindal_k | 0 | 2 | Test Design | Service Tools | NO | Closed | WIP20SubCom | Brilliance Big Bore |  | DHF318051 Rev 21 | October 16, 2019 1:53:34 PM GMT+08:00
CTUDT00625608 | Refresh | Closed | gT_s | August 12, 2019 7:39:16 PM GMT+08:00 | Refresh : The localization is not working for the different languages while performing the remote connection from PRS. | WIP18 | DHF318051 Rev 21 |  | Subsystem |  | NA | Works As Specified / Not A Defect | jindal_k | 0 | 2 | Test Design | Service Tools | NO | Closed | WIP20SubCom | Ingenuity 64 |  | DHF318051 Rev 21 | October 16, 2019 1:53:34 PM GMT+08:00
CTUDT00625608 | Refresh | Closed | gT_s | August 12, 2019 7:39:16 PM GMT+08:00 | Refresh : The localization is not working for the different languages while performing the remote connection from PRS. | WIP18 | DHF318051 Rev 21 |  | Subsystem |  | NA | Works As Specified / Not A Defect | jindal_k | 0 | 2 | Test Design | Service Tools | NO | Closed | WIP20SubCom | IQon1 |  | DHF318051 Rev 21 | October 16, 2019 1:53:34 PM GMT+08:00
CTUDT00625607 | Refresh | Closed | bommiReddy_r | August 12, 2019 7:27:09 PM GMT+08:00 | Refresh: Localization: The part of the message “The selected name already exists. Please select a new name.” was cropped | WIP18 | WIP20 |  | Subsystem |  | NA | Fixed | Prakash_Chand_Yadav | 0 | 2 | Console Software | Console4 | NO | Closed | WIP20SubCom | Ovation |  | WIP20 | September 3, 2019 3:51:12 PM GMT+08:00
CTUDT00625607 | Refresh | Closed | bommiReddy_r | August 12, 2019 7:27:09 PM GMT+08:00 | Refresh: Localization: The part of the message “The selected name already exists. Please select a new name.” was cropped | WIP18 | WIP20 |  | Subsystem |  | NA | Fixed | Prakash_Chand_Yadav | 0 | 2 | Console Software | Console4 | NO | Closed | WIP20SubCom | Brilliance 64 |  | WIP20 | September 3, 2019 3:51:12 PM GMT+08:00
CTUDT00625607 | Refresh | Closed | bommiReddy_r | August 12, 2019 7:27:09 PM GMT+08:00 | Refresh: Localization: The part of the message “The selected name already exists. Please select a new name.” was cropped | WIP18 | WIP20 |  | Subsystem |  | NA | Fixed | Prakash_Chand_Yadav | 0 | 2 | Console Software | Console4 | NO | Closed | WIP20SubCom | Ingenuity 64 |  | WIP20 | September 3, 2019 3:51:12 PM GMT+08:00
CTUDT00625607 | Refresh | Closed | bommiReddy_r | August 12, 2019 7:27:09 PM GMT+08:00 | Refresh: Localization: The part of the message “The selected name already exists. Please select a new name.” was cropped | WIP18 | WIP20 |  | Subsystem |  | NA | Fixed | Prakash_Chand_Yadav | 0 | 2 | Console Software | Console4 | NO | Closed | WIP20SubCom | IQon1 |  | WIP20 | September 3, 2019 3:51:12 PM GMT+08:00
CTUDT00625605 | Refresh | Closed | atin_s | August 12, 2019 5:59:37 PM GMT+08:00 | Cardiac/Pulmo helical recon phase tagging doesn’t match actual recon | 4.8 |  |  | Subsystem | TAVI | N/A | Duplicate | atin_s | 0 | 2 |  | Console4 | YES | Closed | BB_WIP11_SubCom | Brilliance Big Bore | PR9391150
CTUDT00625597 | Refresh | Closed | PremKumar_Rajendran | August 12, 2019 5:02:29 PM GMT+08:00 | BB Refresh - Console was crashed and unable to continue further due to Continuous Bug rep errors and System error. | BB_WIP8 | BB_WIP10 |  | Subsystem |  | NA | Fixed | challa_s | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP10_SubCom | Brilliance Big Bore |  | BB_WIP10 | August 27, 2019 8:53:20 PM GMT+08:00
CTUDT00625596 | Refresh | Closed | ding_f | August 12, 2019 5:01:40 PM GMT+08:00 | IMRSettings.cfg file for BB is not same as Ingenuity one | WIP8.1 | BB_WIP9 | N/A | System | Blocker19 | All IMR IQ related Reqts | Fixed | li_e | 0 | 2 | Console Software | Console4 | YES | Closed | BB_WIP9_SubCom | Brilliance Big Bore |  | BB_WIP9 | August 14, 2019 1:25:14 PM GMT+08:00
CTUDT00625592 | Refresh | Closed | cai_t | August 12, 2019 3:31:55 PM GMT+08:00 | The Warning Symbol ISO7010 - W001 was missed on several translation of Table Power Disconnect Warning Label | BB_WIP8 | WIP20 | The workaround is to add the missing warning symbol on all affected labels. Design change shall be initiated by ECR-070686 and implemented by E-117360. | Subsystem |  | N/A | Fixed | cai_t | 0 | 2 | Couch Mechanical | Couch | NO | Closed | DMR Drawing | Ovation |  | WIP20 | December 13, 2019 4:55:57 PM GMT+08:00
CTUDT00625592 | Refresh | Closed | cai_t | August 12, 2019 3:31:55 PM GMT+08:00 | The Warning Symbol ISO7010 - W001 was missed on several translation of Table Power Disconnect Warning Label | BB_WIP8 | WIP20 | The workaround is to add the missing warning symbol on all affected labels. Design change shall be initiated by ECR-070686 and implemented by E-117360. | Subsystem |  | N/A | Fixed | cai_t | 0 | 2 | Couch Mechanical | Couch | NO | Closed | DMR Drawing | Brilliance 64 |  | WIP20 | December 13, 2019 4:55:57 PM GMT+08:00
CTUDT00625592 | Refresh | Closed | cai_t | August 12, 2019 3:31:55 PM GMT+08:00 | The Warning Symbol ISO7010 - W001 was missed on several translation of Table Power Disconnect Warning Label | BB_WIP8 | WIP20 | The workaround is to add the missing warning symbol on all affected labels. Design change shall be initiated by ECR-070686 and implemented by E-117360. | Subsystem |  | N/A | Fixed | cai_t | 0 | 2 | Couch Mechanical | Couch | NO | Closed | DMR Drawing | Brilliance Big Bore |  | WIP20 | December 13, 2019 4:55:57 PM GMT+08:00
CTUDT00625592 | Refresh | Closed | cai_t | August 12, 2019 3:31:55 PM GMT+08:00 | The Warning Symbol ISO7010 - W001 was missed on several translation of Table Power Disconnect Warning Label | BB_WIP8 | WIP20 | The workaround is to add the missing warning symbol on all affected labels. Design change shall be initiated by ECR-070686 and implemented by E-117360. | Subsystem |  | N/A | Fixed | cai_t | 0 | 2 | Couch Mechanical | Couch | NO | Closed | DMR Drawing | Ingenuity 64 |  | WIP20 | December 13, 2019 4:55:57 PM GMT+08:00
CTUDT00625592 | Refresh | Closed | cai_t | August 12, 2019 3:31:55 PM GMT+08:00 | The Warning Symbol ISO7010 - W001 was missed on several translation of Table Power Disconnect Warning Label | BB_WIP8 | WIP20 | The workaround is to add the missing warning symbol on all affected labels. Design change shall be initiated by ECR-070686 and implemented by E-117360. | Subsystem |  | N/A | Fixed | cai_t | 0 | 2 | Couch Mechanical | Couch | NO | Closed | DMR Drawing | Ingenuity 16 |  | WIP20 | December 13, 2019 4:55:57 PM GMT+08:00
CTUDT00625592 | Refresh | Closed | cai_t | August 12, 2019 3:31:55 PM GMT+08:00 | The Warning Symbol ISO7010 - W001 was missed on several translation of Table Power Disconnect Warning Label | BB_WIP8 | WIP20 | The workaround is to add the missing warning symbol on all affected labels. Design change shall be initiated by ECR-070686 and implemented by E-117360. | Subsystem |  | N/A | Fixed | cai_t | 0 | 2 | Couch Mechanical | Couch | NO | Closed | DMR Drawing | IQon1 |  | WIP20 | December 13, 2019 4:55:57 PM GMT+08:00
CTUDT00625591 | Refresh | Closed | wang_l | August 12, 2019 3:28:24 PM GMT+08:00 | Refresh: "X-ray system cannot comply. Please retry." error during dual surivew | WIP18 |  |  | Subsystem | Performance | NA | Duplicate | wang_k | 0 |  |  | GCX | YES | Closed |  | Brilliance Big Bore
CTUDT00625588 | Refresh | Closed | 320053769 | August 12, 2019 2:23:15 PM GMT+08:00 | Refresh - When copying a study from portal 9/10/11 to Local the copy fails and reverse. | WIP17 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | s_sa | 0 |  |  | iCAP | NO | Closed |  | Ovation
CTUDT00625588 | Refresh | Closed | 320053769 | August 12, 2019 2:23:15 PM GMT+08:00 | Refresh - When copying a study from portal 9/10/11 to Local the copy fails and reverse. | WIP17 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | s_sa | 0 |  |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00625588 | Refresh | Closed | 320053769 | August 12, 2019 2:23:15 PM GMT+08:00 | Refresh - When copying a study from portal 9/10/11 to Local the copy fails and reverse. | WIP17 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | s_sa | 0 |  |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00625588 | Refresh | Closed | 320053769 | August 12, 2019 2:23:15 PM GMT+08:00 | Refresh - When copying a study from portal 9/10/11 to Local the copy fails and reverse. | WIP17 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | s_sa | 0 |  |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00625588 | Refresh | Closed | 320053769 | August 12, 2019 2:23:15 PM GMT+08:00 | Refresh - When copying a study from portal 9/10/11 to Local the copy fails and reverse. | WIP17 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | s_sa | 0 |  |  | iCAP | NO | Closed |  | IQon1
CTUDT00625581 | Refresh | Closed | PremKumar_Rajendran | August 9, 2019 1:21:52 PM GMT+08:00 | BB Refresh - Unhandled exception error has occurred when trying to add Language using Import button in Voice Manager | BB_WIP8 | BB_WIP10 |  | Subsystem |  | NA | Fixed | challa_s | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP10_SubCom | Brilliance Big Bore |  | BB_WIP10 | August 27, 2019 8:53:04 PM GMT+08:00
CTUDT00625580 | Refresh | Closed | manju_kolkar | August 9, 2019 1:05:33 PM GMT+08:00 | Big Bore Refresh - EBW Issue - Localization | BB_WIP8 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | manju_kolkar | 0 | 4 |  | Console4 | NO | Closed | WAS | Ovation
CTUDT00625580 | Refresh | Closed | manju_kolkar | August 9, 2019 1:05:33 PM GMT+08:00 | Big Bore Refresh - EBW Issue - Localization | BB_WIP8 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | manju_kolkar | 0 | 4 |  | Console4 | NO | Closed | WAS | Brilliance 64
CTUDT00625580 | Refresh | Closed | manju_kolkar | August 9, 2019 1:05:33 PM GMT+08:00 | Big Bore Refresh - EBW Issue - Localization | BB_WIP8 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | manju_kolkar | 0 | 4 |  | Console4 | NO | Closed | WAS | Brilliance Big Bore
CTUDT00625580 | Refresh | Closed | manju_kolkar | August 9, 2019 1:05:33 PM GMT+08:00 | Big Bore Refresh - EBW Issue - Localization | BB_WIP8 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | manju_kolkar | 0 | 4 |  | Console4 | NO | Closed | WAS | Ingenuity 64
CTUDT00625580 | Refresh | Closed | manju_kolkar | August 9, 2019 1:05:33 PM GMT+08:00 | Big Bore Refresh - EBW Issue - Localization | BB_WIP8 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | manju_kolkar | 0 | 4 |  | Console4 | NO | Closed | WAS | IQon1
CTUDT00625578 | Refresh | Closed | shekhar_p | August 9, 2019 12:11:38 AM GMT+08:00 | 140 KV is missing in KV drop down for Biopsy EC. | WIP18 | WIP20 |  | Subsystem |  | NA | Fixed | m_nanda | 0 | 2 | Console Software | Console4 | NO | Closed | WIP20SegReq | Ovation |  | WIP20 | August 14, 2019 5:09:33 PM GMT+08:00
CTUDT00625575 | Refresh | Closed | shekhar_p | August 8, 2019 9:03:11 PM GMT+08:00 | In Emergency login user is not able to proceed the scan when scan exceeds dose limits in dose management | WIP18 | WIP21 | Decrease dose below dose limits | Subsystem |  | N/A | Fixed | kr_m | 0 | 3 | Console Software | Console4 | NO | Closed | WIP20SubCom | Ovation |  | WIP21 | October 10, 2019 4:50:57 PM GMT+08:00
CTUDT00625575 | Refresh | Closed | shekhar_p | August 8, 2019 9:03:11 PM GMT+08:00 | In Emergency login user is not able to proceed the scan when scan exceeds dose limits in dose management | WIP18 | WIP21 | Decrease dose below dose limits | Subsystem |  | N/A | Fixed | kr_m | 0 | 3 | Console Software | Console4 | NO | Closed | WIP20SubCom | Brilliance 64 |  | WIP21 | October 10, 2019 4:50:57 PM GMT+08:00
CTUDT00625575 | Refresh | Closed | shekhar_p | August 8, 2019 9:03:11 PM GMT+08:00 | In Emergency login user is not able to proceed the scan when scan exceeds dose limits in dose management | WIP18 | WIP21 | Decrease dose below dose limits | Subsystem |  | N/A | Fixed | kr_m | 0 | 3 | Console Software | Console4 | NO | Closed | WIP20SubCom | Ingenuity 64 |  | WIP21 | October 10, 2019 4:50:57 PM GMT+08:00
CTUDT00625575 | Refresh | Closed | shekhar_p | August 8, 2019 9:03:11 PM GMT+08:00 | In Emergency login user is not able to proceed the scan when scan exceeds dose limits in dose management | WIP18 | WIP21 | Decrease dose below dose limits | Subsystem |  | N/A | Fixed | kr_m | 0 | 3 | Console Software | Console4 | NO | Closed | WIP20SubCom | IQon1 |  | WIP21 | October 10, 2019 4:50:57 PM GMT+08:00
CTUDT00625570 | Refresh | Closed | shiraguppi_k | August 8, 2019 8:05:44 PM GMT+08:00 | Pre Val: All Service related Calibrations are available for Senior Tech login | WIP18 | WIP20 |  | Subsystem | Calibration | n/a | Fixed | frider_o | 0 | 3 | DMS Software | DMS | NO | Closed | WIP20SubCom | Ingenuity 64 |  | WIP20 | September 5, 2019 10:46:33 PM GMT+08:00
CTUDT00625565 | Refresh | Closed | nanjunda_s | August 8, 2019 6:50:11 PM GMT+08:00 | In the view2,Images are reconstructed even though partial raw data is not available for the position(-900.mm) | WIP18 |  | NA | Subsystem |  | NA | No Plan to Fix | ravindra_a | 0 | 4 |  | Console4 | NO | Closed |  | Ovation
CTUDT00625564 | Refresh | Closed | shterner_a | August 8, 2019 6:45:28 PM GMT+08:00 | Refresh IQon/iCT/Ingenuity PRS discrepancies | WIP18 | WIP20 | None Required. Not a hazard. | Subsystem |  | CT-NM.PRS.CRS.7189, CT-NM.PRS.CRS.7190 and CT-NM.PRS.CRS.4782 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP20SegReq | Ovation |  | WIP20 | September 10, 2019 9:49:39 AM GMT+08:00
CTUDT00625564 | Refresh | Closed | shterner_a | August 8, 2019 6:45:28 PM GMT+08:00 | Refresh IQon/iCT/Ingenuity PRS discrepancies | WIP18 | WIP20 | None Required. Not a hazard. | Subsystem |  | CT-NM.PRS.CRS.7189, CT-NM.PRS.CRS.7190 and CT-NM.PRS.CRS.4782 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP20SegReq | Brilliance 64 |  | WIP20 | September 10, 2019 9:49:39 AM GMT+08:00
CTUDT00625564 | Refresh | Closed | shterner_a | August 8, 2019 6:45:28 PM GMT+08:00 | Refresh IQon/iCT/Ingenuity PRS discrepancies | WIP18 | WIP20 | None Required. Not a hazard. | Subsystem |  | CT-NM.PRS.CRS.7189, CT-NM.PRS.CRS.7190 and CT-NM.PRS.CRS.4782 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP20SegReq | Ingenuity 64 |  | WIP20 | September 10, 2019 9:49:39 AM GMT+08:00
CTUDT00625564 | Refresh | Closed | shterner_a | August 8, 2019 6:45:28 PM GMT+08:00 | Refresh IQon/iCT/Ingenuity PRS discrepancies | WIP18 | WIP20 | None Required. Not a hazard. | Subsystem |  | CT-NM.PRS.CRS.7189, CT-NM.PRS.CRS.7190 and CT-NM.PRS.CRS.4782 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP20SegReq | IQon1 |  | WIP20 | September 10, 2019 9:49:39 AM GMT+08:00
CTUDT00625562 | Refresh | Closed | deshbhratar_m | August 8, 2019 5:58:56 PM GMT+08:00 | Refresh : The application displayed system error and ECG/Pulmo signals stopped getting displayed in system. | 4.8 |  |  | Subsystem |  |  | Does Not Apply | deshbhratar_m | 0 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00625556 | Refresh | Closed | deshbhratar_m | August 8, 2019 1:56:03 PM GMT+08:00 | Dose modulation is turned off for the remainder of the scan after the arrhythmic cycle acquisition proceeded. | 4.8 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | sagar_v | 0 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00625553 | Refresh | Closed | zhang_t | August 8, 2019 1:23:10 PM GMT+08:00 | Refresh IVC: Patient table cannot comply. Please retry message popup when moving couch using rocker switch in IVC . | BB_WIP8 | WIP20 | No workaround, besides using the swicth at a moderate speed. | Subsystem | IVC;Blocker20 | NA | Duplicate | Wang_Dexin | 0 | 2 | Gantry Software | GCX | YES | Closed | BB_WIP9_SubCom | Brilliance Big Bore |  | WIP20 | November 12, 2019 9:16:22 AM GMT+08:00
CTUDT00625551 | Refresh | Closed | zhong_x | August 8, 2019 11:03:11 AM GMT+08:00 | X-ray on indicator "white arrow" was missing in offline Viewer of single phase scan. See attached screen Capture. | BB_WIP8 | WIP21 |  | Subsystem | respiratory wave | NA | Fixed | prasad_r | 0 | 2 | Console Software | Console4 | NO | Closed | WIP20SubCom | Brilliance 64 |  | WIP21 | October 10, 2019 1:37:43 PM GMT+08:00
CTUDT00625551 | Refresh | Closed | zhong_x | August 8, 2019 11:03:11 AM GMT+08:00 | X-ray on indicator "white arrow" was missing in offline Viewer of single phase scan. See attached screen Capture. | BB_WIP8 | WIP21 |  | Subsystem | respiratory wave | NA | Fixed | prasad_r | 0 | 2 | Console Software | Console4 | NO | Closed | WIP20SubCom | Ingenuity 64 |  | WIP21 | October 10, 2019 1:37:43 PM GMT+08:00
CTUDT00625541 | Refresh | Closed | 320053767 | August 7, 2019 8:42:58 PM GMT+08:00 | Wrong pitch definition in help file. | WIP18 | WIP20 |  | Subsystem |  | NaN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP20SubCom | Ovation |  | WIP20 | August 29, 2019 7:21:14 PM GMT+08:00
CTUDT00625541 | Refresh | Closed | 320053767 | August 7, 2019 8:42:58 PM GMT+08:00 | Wrong pitch definition in help file. | WIP18 | WIP20 |  | Subsystem |  | NaN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP20SubCom | IQon1 |  | WIP20 | August 29, 2019 7:21:14 PM GMT+08:00
CTUDT00625536 | Refresh | Closed | PremKumar_Rajendran | August 7, 2019 5:25:22 PM GMT+08:00 | Refresh - Message "Please fill in the user name" was displayed and not localized as expected in UI on selected language and Message "Please fill in the user name" was displayed when performed fourth scan on German language. | BB_WIP8 |  | No Workaround | Subsystem |  | Console-SRS.15526 | Works As Specified / Not A Defect | nayak_r | 0 | 2 |  | Console4 | NO | Closed | WAS | Ovation
CTUDT00625536 | Refresh | Closed | PremKumar_Rajendran | August 7, 2019 5:25:22 PM GMT+08:00 | Refresh - Message "Please fill in the user name" was displayed and not localized as expected in UI on selected language and Message "Please fill in the user name" was displayed when performed fourth scan on German language. | BB_WIP8 |  | No Workaround | Subsystem |  | Console-SRS.15526 | Works As Specified / Not A Defect | nayak_r | 0 | 2 |  | Console4 | NO | Closed | WAS | Brilliance 64
CTUDT00625536 | Refresh | Closed | PremKumar_Rajendran | August 7, 2019 5:25:22 PM GMT+08:00 | Refresh - Message "Please fill in the user name" was displayed and not localized as expected in UI on selected language and Message "Please fill in the user name" was displayed when performed fourth scan on German language. | BB_WIP8 |  | No Workaround | Subsystem |  | Console-SRS.15526 | Works As Specified / Not A Defect | nayak_r | 0 | 2 |  | Console4 | NO | Closed | WAS | Brilliance Big Bore
CTUDT00625536 | Refresh | Closed | PremKumar_Rajendran | August 7, 2019 5:25:22 PM GMT+08:00 | Refresh - Message "Please fill in the user name" was displayed and not localized as expected in UI on selected language and Message "Please fill in the user name" was displayed when performed fourth scan on German language. | BB_WIP8 |  | No Workaround | Subsystem |  | Console-SRS.15526 | Works As Specified / Not A Defect | nayak_r | 0 | 2 |  | Console4 | NO | Closed | WAS | Ingenuity 64
CTUDT00625536 | Refresh | Closed | PremKumar_Rajendran | August 7, 2019 5:25:22 PM GMT+08:00 | Refresh - Message "Please fill in the user name" was displayed and not localized as expected in UI on selected language and Message "Please fill in the user name" was displayed when performed fourth scan on German language. | BB_WIP8 |  | No Workaround | Subsystem |  | Console-SRS.15526 | Works As Specified / Not A Defect | nayak_r | 0 | 2 |  | Console4 | NO | Closed | WAS | IQon1
CTUDT00625533 | Refresh | Closed | PremKumar_Rajendran | August 7, 2019 3:03:09 PM GMT+08:00 | BB Refresh - The password for the account has expired. Please change your password <OK> and Message The password for the account has expired was displayed and not localized as expected in UI in selected language. | BB_WIP8 |  | NA | Subsystem |  | NA | Works As Specified / Not A Defect | kr_m | 0 |  |  | Console4 | NO | Closed | WAS | Brilliance Big Bore
CTUDT00625524 | Refresh | Closed | shekhar_p | August 6, 2019 10:23:29 PM GMT+08:00 | System crashes with Bugrep error | WIP18 | WIP20 |  | Subsystem |  | N/A | Fixed | pandey_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP20SubCom | Brilliance 64 |  | WIP20 | August 26, 2019 9:00:03 PM GMT+08:00
CTUDT00625524 | Refresh | Closed | shekhar_p | August 6, 2019 10:23:29 PM GMT+08:00 | System crashes with Bugrep error | WIP18 | WIP20 |  | Subsystem |  | N/A | Fixed | pandey_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP20SubCom | Ingenuity 64 |  | WIP20 | August 26, 2019 9:00:03 PM GMT+08:00
CTUDT00625522 | Refresh | Closed | shekhar_p | August 6, 2019 9:38:29 PM GMT+08:00 | “ Phase “ field is missing in show-all under ECG gating section | WIP18 | WIP20 | As issue happen in ECM user can use any cardiac exam card to changed to desired cardiac scan rather then using Abdomen exam card . | Subsystem | ECM | N/A | Fixed | pandey_a | 0 | 3 | Console Software | Console4 | NO | Closed | WIP20SubCom | Ovation |  | WIP20 | August 22, 2019 5:22:41 PM GMT+08:00
CTUDT00625522 | Refresh | Closed | shekhar_p | August 6, 2019 9:38:29 PM GMT+08:00 | “ Phase “ field is missing in show-all under ECG gating section | WIP18 | WIP20 | As issue happen in ECM user can use any cardiac exam card to changed to desired cardiac scan rather then using Abdomen exam card . | Subsystem | ECM | N/A | Fixed | pandey_a | 0 | 3 | Console Software | Console4 | NO | Closed | WIP20SubCom | Brilliance 64 |  | WIP20 | August 22, 2019 5:22:41 PM GMT+08:00
CTUDT00625522 | Refresh | Closed | shekhar_p | August 6, 2019 9:38:29 PM GMT+08:00 | “ Phase “ field is missing in show-all under ECG gating section | WIP18 | WIP20 | As issue happen in ECM user can use any cardiac exam card to changed to desired cardiac scan rather then using Abdomen exam card . | Subsystem | ECM | N/A | Fixed | pandey_a | 0 | 3 | Console Software | Console4 | NO | Closed | WIP20SubCom | Ingenuity 64 |  | WIP20 | August 22, 2019 5:22:41 PM GMT+08:00
CTUDT00625522 | Refresh | Closed | shekhar_p | August 6, 2019 9:38:29 PM GMT+08:00 | “ Phase “ field is missing in show-all under ECG gating section | WIP18 | WIP20 | As issue happen in ECM user can use any cardiac exam card to changed to desired cardiac scan rather then using Abdomen exam card . | Subsystem | ECM | N/A | Fixed | pandey_a | 0 | 3 | Console Software | Console4 | NO | Closed | WIP20SubCom | IQon1 |  | WIP20 | August 22, 2019 5:22:41 PM GMT+08:00
CTUDT00625520 | Refresh | Closed | v_sumanth | August 6, 2019 9:19:08 PM GMT+08:00 | The OffFocal Table name isn't passed in the MDU | 4.8 |  | N/A | Subsystem |  | N/A | Works As Specified / Not A Defect | roy_r | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP10_SubCom | Brilliance Big Bore
CTUDT00625519 | Refresh | Closed | PremKumar_Rajendran | August 6, 2019 9:13:54 PM GMT+08:00 | BB Refresh : Pitch, Rotation Time, Scan Time fields were not Localized in UI on Dose right warning message | BB_WIP8 | BB_WIP10 |  | Subsystem |  | NA | Fixed | challa_s | 0 | 3 | Console Software | Console4 | YES | Closed | BB_WIP10_SubCom | Brilliance Big Bore |  | BB_WIP10 | August 16, 2019 9:26:13 PM GMT+08:00
CTUDT00625517 | Refresh | Closed | yatziv_l | August 6, 2019 9:02:36 PM GMT+08:00 | Refresh: Calibrations: Localization: Phantom centering Help HTML page containing English in the pictures | WIP15 | WIP20 |  | Subsystem | Localization; Verified1 | DMS-SRS-1206 | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP20SubCom | Ovation |  | WIP20 | September 4, 2019 7:39:35 PM GMT+08:00
CTUDT00625517 | Refresh | Closed | yatziv_l | August 6, 2019 9:02:36 PM GMT+08:00 | Refresh: Calibrations: Localization: Phantom centering Help HTML page containing English in the pictures | WIP15 | WIP20 |  | Subsystem | Localization; Verified1 | DMS-SRS-1206 | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP20SubCom | Brilliance Big Bore |  | WIP20 | September 4, 2019 7:39:35 PM GMT+08:00
CTUDT00625517 | Refresh | Closed | yatziv_l | August 6, 2019 9:02:36 PM GMT+08:00 | Refresh: Calibrations: Localization: Phantom centering Help HTML page containing English in the pictures | WIP15 | WIP20 |  | Subsystem | Localization; Verified1 | DMS-SRS-1206 | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP20SubCom | Ingenuity 64 |  | WIP20 | September 4, 2019 7:39:35 PM GMT+08:00
CTUDT00625517 | Refresh | Closed | yatziv_l | August 6, 2019 9:02:36 PM GMT+08:00 | Refresh: Calibrations: Localization: Phantom centering Help HTML page containing English in the pictures | WIP15 | WIP20 |  | Subsystem | Localization; Verified1 | DMS-SRS-1206 | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP20SubCom | IQon1 |  | WIP20 | September 4, 2019 7:39:35 PM GMT+08:00
CTUDT00625515 | Refresh | Closed | PremKumar_Rajendran | August 6, 2019 8:35:32 PM GMT+08:00 | BB Refresh : The message was truncated in German Language on field “mAs for Average Adult:” | BB_WIP8 | BB_WIP10 |  | Subsystem | Localization | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP10_SubCom | Brilliance Big Bore |  | BB_WIP10 | August 14, 2019 9:48:31 PM GMT+08:00
CTUDT00625510 | Refresh | Closed | shekhar_p | August 6, 2019 6:53:00 PM GMT+08:00 | How do I videos are not playing | WIP18 | WIP19 | None | Subsystem |  | None | Fixed | jose_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP19SubCom | Ovation |  | WIP19 | August 9, 2019 4:18:59 PM GMT+08:00
CTUDT00625510 | Refresh | Closed | shekhar_p | August 6, 2019 6:53:00 PM GMT+08:00 | How do I videos are not playing | WIP18 | WIP19 | None | Subsystem |  | None | Fixed | jose_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP19SubCom | Brilliance 64 |  | WIP19 | August 9, 2019 4:18:59 PM GMT+08:00
CTUDT00625510 | Refresh | Closed | shekhar_p | August 6, 2019 6:53:00 PM GMT+08:00 | How do I videos are not playing | WIP18 | WIP19 | None | Subsystem |  | None | Fixed | jose_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP19SubCom | Ingenuity 64 |  | WIP19 | August 9, 2019 4:18:59 PM GMT+08:00
CTUDT00625510 | Refresh | Closed | shekhar_p | August 6, 2019 6:53:00 PM GMT+08:00 | How do I videos are not playing | WIP18 | WIP19 | None | Subsystem |  | None | Fixed | jose_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP19SubCom | IQon1 |  | WIP19 | August 9, 2019 4:18:59 PM GMT+08:00
CTUDT00625508 | Refresh | Closed | ravindra_a | August 6, 2019 5:03:01 PM GMT+08:00 | Refresh : Actual Latency for closing of Edit Results was more than 1000 milliseconds | 4.8 | BB_WIP9 | no workaround for this scenario | Subsystem | Performance | NA | Fixed | ravindra_a | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP9_SubCom | Brilliance Big Bore |  | BB_WIP9 | August 12, 2019 2:13:24 PM GMT+08:00
CTUDT00625505 | Refresh | Closed | s_sa | August 6, 2019 3:04:54 PM GMT+08:00 | Refresh : The configuration test failed while performing the remote connection between system and Iwatt simulator when we add device through Olevel or IT user | WIP18 | WIP20 BB_WIP10 | NO. | Subsystem | Blocker | NA | Fixed | kapadia_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP20SubCom | Ovation |  | WIP20 BB_WIP10 | September 9, 2019 9:12:41 PM GMT+08:00
CTUDT00625505 | Refresh | Closed | s_sa | August 6, 2019 3:04:54 PM GMT+08:00 | Refresh : The configuration test failed while performing the remote connection between system and Iwatt simulator when we add device through Olevel or IT user | WIP18 | WIP20 BB_WIP10 | NO. | Subsystem | Blocker | NA | Fixed | kapadia_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP20SubCom | Brilliance 64 |  | WIP20 BB_WIP10 | September 9, 2019 9:12:41 PM GMT+08:00
CTUDT00625505 | Refresh | Closed | s_sa | August 6, 2019 3:04:54 PM GMT+08:00 | Refresh : The configuration test failed while performing the remote connection between system and Iwatt simulator when we add device through Olevel or IT user | WIP18 | WIP20 BB_WIP10 | NO. | Subsystem | Blocker | NA | Fixed | kapadia_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP20SubCom | Brilliance Big Bore |  | WIP20 BB_WIP10 | September 9, 2019 9:12:41 PM GMT+08:00
CTUDT00625505 | Refresh | Closed | s_sa | August 6, 2019 3:04:54 PM GMT+08:00 | Refresh : The configuration test failed while performing the remote connection between system and Iwatt simulator when we add device through Olevel or IT user | WIP18 | WIP20 BB_WIP10 | NO. | Subsystem | Blocker | NA | Fixed | kapadia_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP20SubCom | Ingenuity 64 |  | WIP20 BB_WIP10 | September 9, 2019 9:12:41 PM GMT+08:00
CTUDT00625505 | Refresh | Closed | s_sa | August 6, 2019 3:04:54 PM GMT+08:00 | Refresh : The configuration test failed while performing the remote connection between system and Iwatt simulator when we add device through Olevel or IT user | WIP18 | WIP20 BB_WIP10 | NO. | Subsystem | Blocker | NA | Fixed | kapadia_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP20SubCom | IQon1 |  | WIP20 BB_WIP10 | September 9, 2019 9:12:41 PM GMT+08:00
CTUDT00625504 | Refresh | Closed | cai_t | August 6, 2019 2:59:29 PM GMT+08:00 | Refresh Calibrations:FEE test:Signal Test failed. | BB_WIP8 | WIP19 |  | Subsystem | Calibration;Blocker19 | n/a | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP19SubCom | Brilliance Big Bore |  | WIP19 | August 12, 2019 10:46:55 PM GMT+08:00
CTUDT00625501 | Refresh | Closed | sun_q | August 6, 2019 10:41:08 AM GMT+08:00 | The Warning Symbol ISO7010 - W001 was missed on several translation of UPS Disconnect Warning label | WIP18 | E-117360 | not applicable | Subsystem |  | N/A | Fixed | zhu_l | 0 | 4 | Design Documentation | Support-System | NO | Closed | ECR | Brilliance Big Bore |  | E-117360 | November 8, 2019 5:22:55 PM GMT+08:00
CTUDT00625501 | Refresh | Closed | sun_q | August 6, 2019 10:41:08 AM GMT+08:00 | The Warning Symbol ISO7010 - W001 was missed on several translation of UPS Disconnect Warning label | WIP18 | E-117360 | not applicable | Subsystem |  | N/A | Fixed | zhu_l | 0 | 4 | Design Documentation | Support-System | NO | Closed | ECR | Ingenuity 64 |  | E-117360 | November 8, 2019 5:22:55 PM GMT+08:00
CTUDT00625499 | Refresh | Closed | ding_s | August 6, 2019 10:18:17 AM GMT+08:00 | Clinical User created with IT has O-Level QA tools | BB_WIP8 |  |  | Subsystem | QA Tools |  | Duplicate | altshuller_mark | 0 |  |  | DMS | NO | Closed |  | Brilliance Big Bore
CTUDT00625497 | Refresh | Closed | zhang_t | August 6, 2019 9:43:56 AM GMT+08:00 | Refresh: IQ - Filter tables in recon MDU under legacy phantom config are incorrect. | WIP18 | WIP19 |  | Subsystem | MDU;blocker19 | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP19SubCom | Ovation |  | WIP19 | August 12, 2019 4:06:01 PM GMT+08:00
CTUDT00625497 | Refresh | Closed | zhang_t | August 6, 2019 9:43:56 AM GMT+08:00 | Refresh: IQ - Filter tables in recon MDU under legacy phantom config are incorrect. | WIP18 | WIP19 |  | Subsystem | MDU;blocker19 | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP19SubCom | Ingenuity 64 |  | WIP19 | August 12, 2019 4:06:01 PM GMT+08:00
CTUDT00625494 | Refresh | Closed | ding_f | August 6, 2019 8:51:59 AM GMT+08:00 | Harmonized Infant Body phantom: IQ check(CT, Noise, Unif, LC) tool fails to measure data on some water phantom slices. | WIP18 |  | N/A | Subsystem | Harmonize Phantom;Blocker | N/A | Works As Specified / Not A Defect | ding_f | 0 | 2 | Verification Documentation | Support-System | NO | Closed | BB_WIP9_SubCom | Brilliance 64 |  |  | October 15, 2019 6:37:30 PM GMT+08:00
CTUDT00625494 | Refresh | Closed | ding_f | August 6, 2019 8:51:59 AM GMT+08:00 | Harmonized Infant Body phantom: IQ check(CT, Noise, Unif, LC) tool fails to measure data on some water phantom slices. | WIP18 |  | N/A | Subsystem | Harmonize Phantom;Blocker | N/A | Works As Specified / Not A Defect | ding_f | 0 | 2 | Verification Documentation | Support-System | NO | Closed | BB_WIP9_SubCom | Ingenuity 64 |  |  | October 15, 2019 6:37:30 PM GMT+08:00
CTUDT00625491 | Refresh | Closed | 320053767 | August 5, 2019 9:35:40 PM GMT+08:00 | iCt iCt SP H-phantom wrong limits - IQ test failed because of wrong limit parameter . | WIP18 | WIP20 |  | Subsystem | IQ Test | N/A | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP19SubCom | Ovation |  | WIP20 | August 11, 2019 8:16:39 PM GMT+08:00
CTUDT00625484 | Refresh | Closed | lin_s | August 5, 2019 6:21:35 PM GMT+08:00 | Gantry subsystem evidence for TEAL PDU are not included in DHF322920 | WIP18 | DHF322920 Rev09 |  | Subsystem |  | GCX.SSRS.iCT.20404 | Fixed | sun_q | 0 | 3 | Verification Documentation | GCX | NO | Closed | WIP19SubCom | Brilliance 64 |  | DHF322920 Rev09 | September 17, 2019 3:35:20 PM GMT+08:00
CTUDT00625484 | Refresh | Closed | lin_s | August 5, 2019 6:21:35 PM GMT+08:00 | Gantry subsystem evidence for TEAL PDU are not included in DHF322920 | WIP18 | DHF322920 Rev09 |  | Subsystem |  | GCX.SSRS.iCT.20404 | Fixed | sun_q | 0 | 3 | Verification Documentation | GCX | NO | Closed | WIP19SubCom | Brilliance Big Bore |  | DHF322920 Rev09 | September 17, 2019 3:35:20 PM GMT+08:00
CTUDT00625484 | Refresh | Closed | lin_s | August 5, 2019 6:21:35 PM GMT+08:00 | Gantry subsystem evidence for TEAL PDU are not included in DHF322920 | WIP18 | DHF322920 Rev09 |  | Subsystem |  | GCX.SSRS.iCT.20404 | Fixed | sun_q | 0 | 3 | Verification Documentation | GCX | NO | Closed | WIP19SubCom | Ingenuity 64 |  | DHF322920 Rev09 | September 17, 2019 3:35:20 PM GMT+08:00
CTUDT00625484 | Refresh | Closed | lin_s | August 5, 2019 6:21:35 PM GMT+08:00 | Gantry subsystem evidence for TEAL PDU are not included in DHF322920 | WIP18 | DHF322920 Rev09 |  | Subsystem |  | GCX.SSRS.iCT.20404 | Fixed | sun_q | 0 | 3 | Verification Documentation | GCX | NO | Closed | WIP19SubCom | Ingenuity BigBore |  | DHF322920 Rev09 | September 17, 2019 3:35:20 PM GMT+08:00
CTUDT00625481 | Refresh | Closed | lin_s | August 5, 2019 6:09:18 PM GMT+08:00 | GCX.SSRS.iCT.20405 need untag from Briliance 64 for it is not created for TEAL PDU, not for Teal PDU. | WIP18 | WIP19 |  | Subsystem |  | GCX.SSRS.iCT.20405 | Fixed | wang_k | 0 | 2 | Design Documentation | GCX | NO | Closed | Gantry SSRS DHF317363 Rev04 | Brilliance 64 |  | WIP19 | September 17, 2019 3:33:16 PM GMT+08:00
CTUDT00625480 | Refresh | Closed | levy_s | August 5, 2019 6:07:04 PM GMT+08:00 | Refresh: Left panel freeze infrequently when the user choose to End Exam. | WIP18 | WIP20 |  | System | Performance;Verified1 | NA | Fixed | tamari_o | 0 | 2 | Console Software | Console4 | NO | Closed | WIP20SubCom | Ovation |  | WIP20 | September 13, 2019 2:32:15 PM GMT+08:00
CTUDT00625480 | Refresh | Closed | levy_s | August 5, 2019 6:07:04 PM GMT+08:00 | Refresh: Left panel freeze infrequently when the user choose to End Exam. | WIP18 | WIP20 |  | System | Performance;Verified1 | NA | Fixed | tamari_o | 0 | 2 | Console Software | Console4 | NO | Closed | WIP20SubCom | Brilliance 64 |  | WIP20 | September 13, 2019 2:32:15 PM GMT+08:00
CTUDT00625480 | Refresh | Closed | levy_s | August 5, 2019 6:07:04 PM GMT+08:00 | Refresh: Left panel freeze infrequently when the user choose to End Exam. | WIP18 | WIP20 |  | System | Performance;Verified1 | NA | Fixed | tamari_o | 0 | 2 | Console Software | Console4 | NO | Closed | WIP20SubCom | Ingenuity 64 |  | WIP20 | September 13, 2019 2:32:15 PM GMT+08:00
CTUDT00625480 | Refresh | Closed | levy_s | August 5, 2019 6:07:04 PM GMT+08:00 | Refresh: Left panel freeze infrequently when the user choose to End Exam. | WIP18 | WIP20 |  | System | Performance;Verified1 | NA | Fixed | tamari_o | 0 | 2 | Console Software | Console4 | NO | Closed | WIP20SubCom | IQon1 |  | WIP20 | September 13, 2019 2:32:15 PM GMT+08:00
CTUDT00625478 | Refresh | Closed | rojin_antony | August 5, 2019 5:01:08 PM GMT+08:00 | Refresh Pre-Val : While doing an Air Cal in Service login the newly opened window is showing OLevel | WIP18 |  |  | Subsystem | Access Control | n/a | Works As Specified / Not A Defect | yatziv_l | 0 |  |  | DMS | NO | Closed |  | Ingenuity 64
CTUDT00625476 | Refresh | Closed | levy_s | August 5, 2019 4:43:57 PM GMT+08:00 | Refresh: “Gantry cannot initialize” error appear while trying to execute Surview scan and loading a patient | WIP18 |  |  | Subsystem | Gantry |  | Works As Specified / Not A Defect | tamari_o | 1 |  |  | Support-System | NO | Closed |  | Ovation
CTUDT00625476 | Refresh | Closed | levy_s | August 5, 2019 4:43:57 PM GMT+08:00 | Refresh: “Gantry cannot initialize” error appear while trying to execute Surview scan and loading a patient | WIP18 |  |  | Subsystem | Gantry |  | Works As Specified / Not A Defect | tamari_o | 1 |  |  | Support-System | NO | Closed |  | IQon1
CTUDT00625469 | Refresh | Closed | 320053767 | August 4, 2019 9:48:46 PM GMT+08:00 | Low limit wrong parameter (Harmonized phantom) | WIP18 | WIP19 |  | Subsystem |  | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP20SegReq | Ovation |  | WIP19 | August 12, 2019 3:07:11 PM GMT+08:00
CTUDT00625467 | Refresh | Closed | 320053767 | August 4, 2019 4:12:41 PM GMT+08:00 | Noise limits incorrect in acceptance test. | WIP18 | WIP19 |  | Subsystem |  | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP20SegReq | Ovation |  | WIP19 | August 12, 2019 3:15:35 PM GMT+08:00
CTUDT00625464 | Refresh | Closed | lei_yang | August 3, 2019 11:50:31 AM GMT+08:00 | The actual recon phases were inconsistent in Log viewer and CIRS Event Viewer and plan recon phase. | WIP18 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | lin_s | 0 | 2 |  | Support-System | NO | Closed | WIP19SubCom | Brilliance Big Bore
CTUDT00625464 | Refresh | Closed | lei_yang | August 3, 2019 11:50:31 AM GMT+08:00 | The actual recon phases were inconsistent in Log viewer and CIRS Event Viewer and plan recon phase. | WIP18 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | lin_s | 0 | 2 |  | Support-System | NO | Closed | WIP19SubCom | Ingenuity 64
CTUDT00625461 | Refresh | Closed | rojin_antony | August 2, 2019 5:05:17 PM GMT+08:00 | Refresh Pre-Val : Liver Area DRI is Editable during Offline Reconstruction | WIP18 |  |  | Subsystem |  | NA | No Plan to Fix | sagar_v | 0 | 4 |  | Console4 | NO | Closed |  | Ovation
CTUDT00625461 | Refresh | Closed | rojin_antony | August 2, 2019 5:05:17 PM GMT+08:00 | Refresh Pre-Val : Liver Area DRI is Editable during Offline Reconstruction | WIP18 |  |  | Subsystem |  | NA | No Plan to Fix | sagar_v | 0 | 4 |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00625461 | Refresh | Closed | rojin_antony | August 2, 2019 5:05:17 PM GMT+08:00 | Refresh Pre-Val : Liver Area DRI is Editable during Offline Reconstruction | WIP18 |  |  | Subsystem |  | NA | No Plan to Fix | sagar_v | 0 | 4 |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00625461 | Refresh | Closed | rojin_antony | August 2, 2019 5:05:17 PM GMT+08:00 | Refresh Pre-Val : Liver Area DRI is Editable during Offline Reconstruction | WIP18 |  |  | Subsystem |  | NA | No Plan to Fix | sagar_v | 0 | 4 |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00625461 | Refresh | Closed | rojin_antony | August 2, 2019 5:05:17 PM GMT+08:00 | Refresh Pre-Val : Liver Area DRI is Editable during Offline Reconstruction | WIP18 |  |  | Subsystem |  | NA | No Plan to Fix | sagar_v | 0 | 4 |  | Console4 | NO | Closed |  | IQon1
CTUDT00625459 | Refresh | Closed | gT_s | August 2, 2019 4:53:54 PM GMT+08:00 | Refresh : The application displayed system error and ECG/Pulmo signals stopped getting displayed in system. | WIP18 | WIP21 BB_WIP11 |  | Subsystem |  | None | Fixed | prasad_r | 0 | 3 | Service Tools_Software | Console4 | NO | Closed | WIP21SubCom | Brilliance 64 |  | WIP21 BB_WIP11 | September 17, 2019 3:10:24 PM GMT+08:00
CTUDT00625459 | Refresh | Closed | gT_s | August 2, 2019 4:53:54 PM GMT+08:00 | Refresh : The application displayed system error and ECG/Pulmo signals stopped getting displayed in system. | WIP18 | WIP21 BB_WIP11 |  | Subsystem |  | None | Fixed | prasad_r | 0 | 3 | Service Tools_Software | Console4 | NO | Closed | WIP21SubCom | Brilliance Big Bore |  | WIP21 BB_WIP11 | September 17, 2019 3:10:24 PM GMT+08:00
CTUDT00625459 | Refresh | Closed | gT_s | August 2, 2019 4:53:54 PM GMT+08:00 | Refresh : The application displayed system error and ECG/Pulmo signals stopped getting displayed in system. | WIP18 | WIP21 BB_WIP11 |  | Subsystem |  | None | Fixed | prasad_r | 0 | 3 | Service Tools_Software | Console4 | NO | Closed | WIP21SubCom | Ingenuity 64 |  | WIP21 BB_WIP11 | September 17, 2019 3:10:24 PM GMT+08:00
CTUDT00625455 | Refresh | Closed | rojin_antony | August 2, 2019 2:57:50 PM GMT+08:00 | Refresh Pre-Val : Acquired Brain soft tissue images are totally white in color even if the protocol preferences are set at correct window width and center | WIP18 |  |  | Subsystem | Exam Cards |  | Duplicate | richards_b | 0 |  |  | CIRS | NO | Closed |  | Ingenuity 64
CTUDT00625454 | Refresh | Closed | gT_s | August 2, 2019 2:53:32 PM GMT+08:00 | Refresh :The warning message is not getting displayed when the breath rate is below 5 bpm. | WIP18 | WIP20 |  | Subsystem |  | NA | Fixed | m_nanda | 0 | 2 | Console Software | Console4 | YES | Closed | WIP20SubCom | Brilliance 64 |  | WIP20 | August 29, 2019 7:27:38 PM GMT+08:00
CTUDT00625454 | Refresh | Closed | gT_s | August 2, 2019 2:53:32 PM GMT+08:00 | Refresh :The warning message is not getting displayed when the breath rate is below 5 bpm. | WIP18 | WIP20 |  | Subsystem |  | NA | Fixed | m_nanda | 0 | 2 | Console Software | Console4 | YES | Closed | WIP20SubCom | Ingenuity 64 |  | WIP20 | August 29, 2019 7:27:38 PM GMT+08:00
CTUDT00625453 | Refresh | Closed | ravindra_a | August 2, 2019 2:31:03 PM GMT+08:00 | Refresh: ING, Resp.4D CT EC is displayed in EC group but the selected exam cards cannot be used when loading it to plan while Basic license option Key is used. | 4.8 | BB_WIP9 | No | Subsystem |  | NA | Fixed | prasad_r | 0 | 4 | Console Software | Console4 | NO | Closed | WIP19SubCom | Brilliance Big Bore |  | BB_WIP9 | August 7, 2019 4:37:34 PM GMT+08:00
CTUDT00625451 | Refresh | Closed | sun_q | August 2, 2019 11:05:37 AM GMT+08:00 | No warning instructions regarding the limitations of the rear interlock and voltages that remain live. | WIP18 | 459800093391 Rev K 459801263801 Rev B 459801263811 Rev B | No, since this is documentation issue. | System |  | GCX.SSRS.iCT.16913 | Fixed | shi_b | 0 | 2 | Service Documentation | Support-Service | YES | Closed | BB_WIP10_SubCom | Brilliance 64 |  | 459800093391 Rev K 459801263801 Rev B 459801263811 Rev B | October 15, 2019 1:56:39 PM GMT+08:00
CTUDT00625451 | Refresh | Closed | sun_q | August 2, 2019 11:05:37 AM GMT+08:00 | No warning instructions regarding the limitations of the rear interlock and voltages that remain live. | WIP18 | 459800093391 Rev K 459801263801 Rev B 459801263811 Rev B | No, since this is documentation issue. | System |  | GCX.SSRS.iCT.16913 | Fixed | shi_b | 0 | 2 | Service Documentation | Support-Service | YES | Closed | BB_WIP10_SubCom | Brilliance Big Bore |  | 459800093391 Rev K 459801263801 Rev B 459801263811 Rev B | October 15, 2019 1:56:39 PM GMT+08:00
CTUDT00625451 | Refresh | Closed | sun_q | August 2, 2019 11:05:37 AM GMT+08:00 | No warning instructions regarding the limitations of the rear interlock and voltages that remain live. | WIP18 | 459800093391 Rev K 459801263801 Rev B 459801263811 Rev B | No, since this is documentation issue. | System |  | GCX.SSRS.iCT.16913 | Fixed | shi_b | 0 | 2 | Service Documentation | Support-Service | YES | Closed | BB_WIP10_SubCom | Ingenuity 64 |  | 459800093391 Rev K 459801263801 Rev B 459801263811 Rev B | October 15, 2019 1:56:39 PM GMT+08:00
CTUDT00625449 | Refresh | Closed | sun_q | August 2, 2019 10:34:26 AM GMT+08:00 | IVC:CANMM_MOVE_MARKER_STATUS_BKST is not displayed in CanTrace | BB_WIP8 | BB_WIP9 |  | Subsystem | IVC | GCX.SSRS.iCT.20993 | Fixed | wang_k | 0 | 2 | Design Documentation | GCX | NO | Closed | BB_WIP10_SubCom | Brilliance Big Bore |  | BB_WIP9 | September 19, 2019 9:15:58 AM GMT+08:00
CTUDT00625433 | Refresh | Closed | 320053770 | August 1, 2019 7:46:56 PM GMT+08:00 | Refresh - Wip18: IQon&ICt – after trying to connect to the GPC via Service tools Remote Viewer – an error pops up. | WIP18 | WIP19 |  | Subsystem |  | NA | Fixed | prasad_r | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP19SubCom | Ovation |  | WIP19 | August 6, 2019 7:38:02 PM GMT+08:00
CTUDT00625433 | Refresh | Closed | 320053770 | August 1, 2019 7:46:56 PM GMT+08:00 | Refresh - Wip18: IQon&ICt – after trying to connect to the GPC via Service tools Remote Viewer – an error pops up. | WIP18 | WIP19 |  | Subsystem |  | NA | Fixed | prasad_r | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP19SubCom | IQon1 |  | WIP19 | August 6, 2019 7:38:02 PM GMT+08:00
CTUDT00625428 | Refresh | Closed | 320053767 | August 1, 2019 6:10:32 PM GMT+08:00 | Refresh wip18 Bay 131 Performance test failure. | WIP18 | WIP19 |  | Subsystem |  | n/a | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP20SegReq | Ovation |  | WIP19 | August 6, 2019 11:32:12 PM GMT+08:00
CTUDT00625427 | Refresh | Closed | v_sumanth | August 1, 2019 5:47:15 PM GMT+08:00 | Refresh : The YES & No button is not localized as per selected Norwegian language. | 4.8 | BB_WIP9 |  | Subsystem |  | NA | Fixed | rath_s | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP9_SubCom | Brilliance Big Bore |  | BB_WIP9 | August 8, 2019 5:58:09 PM GMT+08:00
CTUDT00625417 | Refresh | Closed | wang_l | August 1, 2019 4:48:46 PM GMT+08:00 | Refresh: Unexpected error lead to "There is not enough disk space to perform the requested reconstruction | WIP18 |  |  | Subsystem | Test Design;Performance | NA | Works As Specified / Not A Defect | wang_l | 0 |  |  | Support-System | NO | Closed | WAS | Ovation
CTUDT00625417 | Refresh | Closed | wang_l | August 1, 2019 4:48:46 PM GMT+08:00 | Refresh: Unexpected error lead to "There is not enough disk space to perform the requested reconstruction | WIP18 |  |  | Subsystem | Test Design;Performance | NA | Works As Specified / Not A Defect | wang_l | 0 |  |  | Support-System | NO | Closed | WAS | Brilliance 64
CTUDT00625417 | Refresh | Closed | wang_l | August 1, 2019 4:48:46 PM GMT+08:00 | Refresh: Unexpected error lead to "There is not enough disk space to perform the requested reconstruction | WIP18 |  |  | Subsystem | Test Design;Performance | NA | Works As Specified / Not A Defect | wang_l | 0 |  |  | Support-System | NO | Closed | WAS | Brilliance Big Bore
CTUDT00625417 | Refresh | Closed | wang_l | August 1, 2019 4:48:46 PM GMT+08:00 | Refresh: Unexpected error lead to "There is not enough disk space to perform the requested reconstruction | WIP18 |  |  | Subsystem | Test Design;Performance | NA | Works As Specified / Not A Defect | wang_l | 0 |  |  | Support-System | NO | Closed | WAS | Ingenuity 64
CTUDT00625417 | Refresh | Closed | wang_l | August 1, 2019 4:48:46 PM GMT+08:00 | Refresh: Unexpected error lead to "There is not enough disk space to perform the requested reconstruction | WIP18 |  |  | Subsystem | Test Design;Performance | NA | Works As Specified / Not A Defect | wang_l | 0 |  |  | Support-System | NO | Closed | WAS | IQon1
CTUDT00625412 | Refresh | Closed | v_sumanth | August 1, 2019 4:17:54 PM GMT+08:00 | Getting an error message "ViewTwo has stopped working" when performing reconstruction in View2 with more number of results | 4.8 | BB_WIP9 | User can reload the series in View 2 again taking each time less results | Subsystem |  | NA | Fixed | v_sumanth | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP9_SubCom | Brilliance Big Bore |  | BB_WIP9 | August 6, 2019 7:42:59 PM GMT+08:00
CTUDT00625406 | Refresh | Closed | wang_l | August 1, 2019 3:07:46 PM GMT+08:00 | Refresh: System error displayed once click End Exam button | WIP18 | WIP20 | User can press "OK" on system error message and continue working. | System | Verified1 | Console.SRS-2047 | Fixed | li_e | 0 | 4 | Console Software | Console4 | NO | Closed | WIP20SubCom | Ovation |  | WIP20 | August 13, 2019 2:26:01 PM GMT+08:00
CTUDT00625406 | Refresh | Closed | wang_l | August 1, 2019 3:07:46 PM GMT+08:00 | Refresh: System error displayed once click End Exam button | WIP18 | WIP20 | User can press "OK" on system error message and continue working. | System | Verified1 | Console.SRS-2047 | Fixed | li_e | 0 | 4 | Console Software | Console4 | NO | Closed | WIP20SubCom | Brilliance 64 |  | WIP20 | August 13, 2019 2:26:01 PM GMT+08:00
CTUDT00625406 | Refresh | Closed | wang_l | August 1, 2019 3:07:46 PM GMT+08:00 | Refresh: System error displayed once click End Exam button | WIP18 | WIP20 | User can press "OK" on system error message and continue working. | System | Verified1 | Console.SRS-2047 | Fixed | li_e | 0 | 4 | Console Software | Console4 | NO | Closed | WIP20SubCom | Ingenuity 64 |  | WIP20 | August 13, 2019 2:26:01 PM GMT+08:00
CTUDT00625406 | Refresh | Closed | wang_l | August 1, 2019 3:07:46 PM GMT+08:00 | Refresh: System error displayed once click End Exam button | WIP18 | WIP20 | User can press "OK" on system error message and continue working. | System | Verified1 | Console.SRS-2047 | Fixed | li_e | 0 | 4 | Console Software | Console4 | NO | Closed | WIP20SubCom | IQon1 |  | WIP20 | August 13, 2019 2:26:01 PM GMT+08:00
CTUDT00625405 | Refresh | Closed | atin_s | August 1, 2019 2:25:36 PM GMT+08:00 | Cardiac/Pulmo helical recon phase tagging doesn’t match actual recon | 4.1.10 | WIP20 |  | Subsystem | TAVI | Console-SRS.3216 | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | YES | Closed | WIP21SegReq | Ovation | PR9391150 | WIP20 | September 5, 2019 7:22:06 PM GMT+08:00
CTUDT00625405 | Refresh | Closed | atin_s | August 1, 2019 2:25:36 PM GMT+08:00 | Cardiac/Pulmo helical recon phase tagging doesn’t match actual recon | 4.1.10 | WIP20 |  | Subsystem | TAVI | Console-SRS.3216 | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | YES | Closed | WIP21SegReq | Brilliance 64 | PR9391150 | WIP20 | September 5, 2019 7:22:06 PM GMT+08:00
CTUDT00625405 | Refresh | Closed | atin_s | August 1, 2019 2:25:36 PM GMT+08:00 | Cardiac/Pulmo helical recon phase tagging doesn’t match actual recon | 4.1.10 | WIP20 |  | Subsystem | TAVI | Console-SRS.3216 | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | YES | Closed | WIP21SegReq | Ingenuity 64 | PR9391150 | WIP20 | September 5, 2019 7:22:06 PM GMT+08:00
CTUDT00625405 | Refresh | Closed | atin_s | August 1, 2019 2:25:36 PM GMT+08:00 | Cardiac/Pulmo helical recon phase tagging doesn’t match actual recon | 4.1.10 | WIP20 |  | Subsystem | TAVI | Console-SRS.3216 | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | YES | Closed | WIP21SegReq | IQon1 | PR9391150 | WIP20 | September 5, 2019 7:22:06 PM GMT+08:00
CTUDT00625404 | Refresh | Closed | atin_s | August 1, 2019 2:19:34 PM GMT+08:00 | Wrong ECG phase tagging after signal disconnection during S&S scan | 4.1.10 | WIP20 |  | Subsystem | TAVI | Console-SRS.3220 and Console-SRS.3224 | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | YES | Closed | WIP21SubCom | Ovation | N/A 8844848 8837662 8844846 | WIP20 | September 5, 2019 7:21:18 PM GMT+08:00
CTUDT00625404 | Refresh | Closed | atin_s | August 1, 2019 2:19:34 PM GMT+08:00 | Wrong ECG phase tagging after signal disconnection during S&S scan | 4.1.10 | WIP20 |  | Subsystem | TAVI | Console-SRS.3220 and Console-SRS.3224 | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | YES | Closed | WIP21SubCom | Brilliance 64 | N/A 8844848 8837662 8844846 | WIP20 | September 5, 2019 7:21:18 PM GMT+08:00
CTUDT00625404 | Refresh | Closed | atin_s | August 1, 2019 2:19:34 PM GMT+08:00 | Wrong ECG phase tagging after signal disconnection during S&S scan | 4.1.10 | WIP20 |  | Subsystem | TAVI | Console-SRS.3220 and Console-SRS.3224 | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | YES | Closed | WIP21SubCom | Ingenuity 64 | N/A 8844848 8837662 8844846 | WIP20 | September 5, 2019 7:21:18 PM GMT+08:00
CTUDT00625404 | Refresh | Closed | atin_s | August 1, 2019 2:19:34 PM GMT+08:00 | Wrong ECG phase tagging after signal disconnection during S&S scan | 4.1.10 | WIP20 |  | Subsystem | TAVI | Console-SRS.3220 and Console-SRS.3224 | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | YES | Closed | WIP21SubCom | IQon1 | N/A 8844848 8837662 8844846 | WIP20 | September 5, 2019 7:21:18 PM GMT+08:00
CTUDT00625401 | Refresh | Closed | charles_j | August 1, 2019 1:17:07 PM GMT+08:00 | Remove Tech Pubs from Inheritor list for the attached requirements | WIP18 | WIP19 BB_WIP9 |  | Subsystem |  | SRS in attachments | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP20SegReq | Ovation |  | WIP19 BB_WIP9 | September 5, 2019 4:31:19 PM GMT+08:00
CTUDT00625401 | Refresh | Closed | charles_j | August 1, 2019 1:17:07 PM GMT+08:00 | Remove Tech Pubs from Inheritor list for the attached requirements | WIP18 | WIP19 BB_WIP9 |  | Subsystem |  | SRS in attachments | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP20SegReq | Brilliance 64 |  | WIP19 BB_WIP9 | September 5, 2019 4:31:19 PM GMT+08:00
CTUDT00625401 | Refresh | Closed | charles_j | August 1, 2019 1:17:07 PM GMT+08:00 | Remove Tech Pubs from Inheritor list for the attached requirements | WIP18 | WIP19 BB_WIP9 |  | Subsystem |  | SRS in attachments | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP20SegReq | Brilliance Big Bore |  | WIP19 BB_WIP9 | September 5, 2019 4:31:19 PM GMT+08:00
CTUDT00625401 | Refresh | Closed | charles_j | August 1, 2019 1:17:07 PM GMT+08:00 | Remove Tech Pubs from Inheritor list for the attached requirements | WIP18 | WIP19 BB_WIP9 |  | Subsystem |  | SRS in attachments | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP20SegReq | Ingenuity 64 |  | WIP19 BB_WIP9 | September 5, 2019 4:31:19 PM GMT+08:00
CTUDT00625400 | Refresh | Closed | yuan_s | August 1, 2019 1:05:22 PM GMT+08:00 | 1 Click Non-Standard installlation: TightVNC is not installed in Host computer. | WIP18 | WIP19 BB_WIP9 |  | Subsystem | Installation;Blocker19 | NA | Fixed | Sandeep_Kumar | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP19SubCom | Ovation |  | WIP19 BB_WIP9 | August 6, 2019 2:52:08 PM GMT+08:00
CTUDT00625400 | Refresh | Closed | yuan_s | August 1, 2019 1:05:22 PM GMT+08:00 | 1 Click Non-Standard installlation: TightVNC is not installed in Host computer. | WIP18 | WIP19 BB_WIP9 |  | Subsystem | Installation;Blocker19 | NA | Fixed | Sandeep_Kumar | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP19SubCom | Brilliance 64 |  | WIP19 BB_WIP9 | August 6, 2019 2:52:08 PM GMT+08:00
CTUDT00625400 | Refresh | Closed | yuan_s | August 1, 2019 1:05:22 PM GMT+08:00 | 1 Click Non-Standard installlation: TightVNC is not installed in Host computer. | WIP18 | WIP19 BB_WIP9 |  | Subsystem | Installation;Blocker19 | NA | Fixed | Sandeep_Kumar | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP19SubCom | Brilliance Big Bore |  | WIP19 BB_WIP9 | August 6, 2019 2:52:08 PM GMT+08:00
CTUDT00625400 | Refresh | Closed | yuan_s | August 1, 2019 1:05:22 PM GMT+08:00 | 1 Click Non-Standard installlation: TightVNC is not installed in Host computer. | WIP18 | WIP19 BB_WIP9 |  | Subsystem | Installation;Blocker19 | NA | Fixed | Sandeep_Kumar | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP19SubCom | Ingenuity 64 |  | WIP19 BB_WIP9 | August 6, 2019 2:52:08 PM GMT+08:00
CTUDT00625400 | Refresh | Closed | yuan_s | August 1, 2019 1:05:22 PM GMT+08:00 | 1 Click Non-Standard installlation: TightVNC is not installed in Host computer. | WIP18 | WIP19 BB_WIP9 |  | Subsystem | Installation;Blocker19 | NA | Fixed | Sandeep_Kumar | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP19SubCom | IQon1 |  | WIP19 BB_WIP9 | August 6, 2019 2:52:08 PM GMT+08:00
CTUDT00625391 | Refresh | Closed | wang_l | August 1, 2019 9:46:50 AM GMT+08:00 | Refresh: Fields displayed in HIS/RIS basic mode did not match with requirement | WIP18 | WIP19 BB_WIP9 |  | Subsystem |  | NA | Fixed | guy_sheaffer | 0 | 3 | Requirements Documentation | iCAP | NO | Closed | WIP20SubCom | Ovation |  | WIP19 BB_WIP9 | September 5, 2019 4:30:36 PM GMT+08:00
CTUDT00625391 | Refresh | Closed | wang_l | August 1, 2019 9:46:50 AM GMT+08:00 | Refresh: Fields displayed in HIS/RIS basic mode did not match with requirement | WIP18 | WIP19 BB_WIP9 |  | Subsystem |  | NA | Fixed | guy_sheaffer | 0 | 3 | Requirements Documentation | iCAP | NO | Closed | WIP20SubCom | Brilliance 64 |  | WIP19 BB_WIP9 | September 5, 2019 4:30:36 PM GMT+08:00
CTUDT00625391 | Refresh | Closed | wang_l | August 1, 2019 9:46:50 AM GMT+08:00 | Refresh: Fields displayed in HIS/RIS basic mode did not match with requirement | WIP18 | WIP19 BB_WIP9 |  | Subsystem |  | NA | Fixed | guy_sheaffer | 0 | 3 | Requirements Documentation | iCAP | NO | Closed | WIP20SubCom | Ingenuity 64 |  | WIP19 BB_WIP9 | September 5, 2019 4:30:36 PM GMT+08:00
CTUDT00625391 | Refresh | Closed | wang_l | August 1, 2019 9:46:50 AM GMT+08:00 | Refresh: Fields displayed in HIS/RIS basic mode did not match with requirement | WIP18 | WIP19 BB_WIP9 |  | Subsystem |  | NA | Fixed | guy_sheaffer | 0 | 3 | Requirements Documentation | iCAP | NO | Closed | WIP20SubCom | IQon1 |  | WIP19 BB_WIP9 | September 5, 2019 4:30:36 PM GMT+08:00
CTUDT00625389 | Refresh | Closed | altshuller_mark | July 31, 2019 3:11:41 PM GMT+08:00 | IQon HB: Waikato: TAVi scan failure | IQon HB | WIP19 | Change the protocol either by reducing the scan length, or rotation time ot increasing the collimation to 64x0.625. avoid using rotation speed of 0.27-0.33 sec\rot while collimation is 32*0.625 and scan time is above ~35 sec | Subsystem |  | N/A | Fixed | rodavsky_l | 0 | 3 | DMS Electrical | DMS | YES | Closed | WIP19SubCom | Ovation | PR 8389836 | WIP19 | September 2, 2019 8:40:38 PM GMT+08:00
CTUDT00625389 | Refresh | Closed | altshuller_mark | July 31, 2019 3:11:41 PM GMT+08:00 | IQon HB: Waikato: TAVi scan failure | IQon HB | WIP19 | Change the protocol either by reducing the scan length, or rotation time ot increasing the collimation to 64x0.625. avoid using rotation speed of 0.27-0.33 sec\rot while collimation is 32*0.625 and scan time is above ~35 sec | Subsystem |  | N/A | Fixed | rodavsky_l | 0 | 3 | DMS Electrical | DMS | YES | Closed | WIP19SubCom | IQon1 | PR 8389836 | WIP19 | September 2, 2019 8:40:38 PM GMT+08:00
CTUDT00625374 | Refresh | Closed | zhu_je | July 31, 2019 10:39:31 AM GMT+08:00 | [Refresh]IV controls command was not ignored,and the couch kept moving in the diection of switch. | BB_WIP8 |  |  | Subsystem | Couch | NA | Works As Specified / Not A Defect | Snehashis_Dey | 0 |  |  | Console4 | NO | Closed | WAS | Brilliance Big Bore
CTUDT00625371 | Refresh | Closed | zhu_je | July 31, 2019 10:33:44 AM GMT+08:00 | [Refresh] IV Console HW is displayed in KGS and Preference Licensing page. | BB_WIP8 | WIP19 |  | Subsystem | Preferences | NA | Fixed | srushti_k | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP19SubCom | Brilliance Big Bore |  | WIP19 | August 16, 2019 3:16:43 PM GMT+08:00
CTUDT00625370 | Refresh | Closed | sun_q | July 31, 2019 10:26:59 AM GMT+08:00 | Refresh-BB [IVC] Move Inner/Outer laser button led is abnormal. | BB_WIP8 | BB_WIP10 | Hardware button LEDs represents the right state of operation, but not in Gantry Panel | Subsystem | GCP - LED;Blocker20 | Console-SRS.17809,Console-SRS.17810 | Fixed | reddy_g | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP9_SubCom | Brilliance Big Bore |  | BB_WIP10 | September 18, 2019 7:27:54 PM GMT+08:00
CTUDT00625368 | Refresh | Closed | nanjunda_s | July 31, 2019 2:41:47 AM GMT+08:00 | Message is not displayed as per the requirement | WIP18 | DHF303211 Rev14 |  | Subsystem | System | Console-SRS.5685 | Fixed | roy_r | 0 | 3 | HOST SSRS | Console4 | NO | Closed | SubCom | Ovation |  | DHF303211 Rev14 | August 5, 2019 2:37:36 PM GMT+08:00
CTUDT00625368 | Refresh | Closed | nanjunda_s | July 31, 2019 2:41:47 AM GMT+08:00 | Message is not displayed as per the requirement | WIP18 | DHF303211 Rev14 |  | Subsystem | System | Console-SRS.5685 | Fixed | roy_r | 0 | 3 | HOST SSRS | Console4 | NO | Closed | SubCom | IQon1 |  | DHF303211 Rev14 | August 5, 2019 2:37:36 PM GMT+08:00
CTUDT00625366 | Refresh | Closed | Valeria_Kashtanov | July 30, 2019 7:35:25 PM GMT+08:00 | Refresh: Calibrations: Performance Test: Additional ROI coordinates of Adult Off-Center CT15a case are incorrect | BB_WIP8 | WIP19 | N/A | Subsystem | Performance Test; Verified1 | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP19SubCom | Brilliance Big Bore |  | WIP19 | August 4, 2019 2:41:00 PM GMT+08:00
CTUDT00625363 | Refresh | Closed | 320053767 | July 30, 2019 6:37:53 PM GMT+08:00 | Missing Phantom tables. | WIP18 |  |  | Subsystem | Blocker | n/a | Duplicate | yatziv_l | 0 |  |  | DMS | YES | Closed |  | Ovation
CTUDT00625358 | Refresh | Closed | sun_q | July 30, 2019 3:59:17 PM GMT+08:00 | China YY0310-2015 Regulatory report is not ready | WIP18 | WIP20 |  | System | Requirements | GCX.SSRS.iCT.21070; GCX.SSRS.iCT.21071 | Fixed | zhu_l | 0 | 2 | Test Design | Support-System | NO | Closed | SubCom-Feb-2020 | Brilliance Big Bore |  | WIP20 | January 17, 2020 1:17:40 PM GMT+08:00
CTUDT00625356 | Refresh | Closed | sun_q | July 30, 2019 3:39:39 PM GMT+08:00 | Seismic report and label are not ready | WIP18 | WIP20 |  | System | Labels | GCX.SSRS.iCT.19379, GCX.SSRS.iCT.20292, GCX.SSRS.iCT.20294, GCX.SSRS.iCT.20295, GCX.SSRS.iCT.20296 | Fixed | zhu_l | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | SubCom-Jul-2020 | Brilliance Big Bore |  | WIP20 | January 7, 2020 5:08:42 PM GMT+08:00
CTUDT00625356 | Refresh | Closed | sun_q | July 30, 2019 3:39:39 PM GMT+08:00 | Seismic report and label are not ready | WIP18 | WIP20 |  | System | Labels | GCX.SSRS.iCT.19379, GCX.SSRS.iCT.20292, GCX.SSRS.iCT.20294, GCX.SSRS.iCT.20295, GCX.SSRS.iCT.20296 | Fixed | zhu_l | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | SubCom-Jul-2020 | Ingenuity 64 |  | WIP20 | January 7, 2020 5:08:42 PM GMT+08:00
CTUDT00625354 | Refresh | Closed | sun_q | July 30, 2019 2:47:20 PM GMT+08:00 | Refresh ING/BB Regulatory label are not ready | WIP18 | WIP20 |  | System | Labels | GCX.SSRS.iCT.20548, GCX.SSRS.iCT.20383, GCX.SSRS.iCT.20384, GCX.SSRS.iCT.20114, GCX.SSRS.iCT.20278, GCX.SSRS.iCT.20280, GCX.SSRS.iCT.20279, GCX.SSRS.iCT.20281, GCX.SSRS.iCT.20395, GCX.SSRS.iCT.20396, GCX.SSRS.iCT.20476, GCX.SSRS.iCT.20115, GCX.SSRS.iCT. | Fixed | zhu_l | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | SubCom-12-Dec-2019 | Brilliance Big Bore |  | WIP20 | January 17, 2020 1:38:46 PM GMT+08:00
CTUDT00625354 | Refresh | Closed | sun_q | July 30, 2019 2:47:20 PM GMT+08:00 | Refresh ING/BB Regulatory label are not ready | WIP18 | WIP20 |  | System | Labels | GCX.SSRS.iCT.20548, GCX.SSRS.iCT.20383, GCX.SSRS.iCT.20384, GCX.SSRS.iCT.20114, GCX.SSRS.iCT.20278, GCX.SSRS.iCT.20280, GCX.SSRS.iCT.20279, GCX.SSRS.iCT.20281, GCX.SSRS.iCT.20395, GCX.SSRS.iCT.20396, GCX.SSRS.iCT.20476, GCX.SSRS.iCT.20115, GCX.SSRS.iCT. | Fixed | zhu_l | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | SubCom-12-Dec-2019 | Ingenuity 64 |  | WIP20 | January 17, 2020 1:38:46 PM GMT+08:00
CTUDT00625352 | Refresh | Closed | wang_l | July 30, 2019 2:00:51 PM GMT+08:00 | Refresh: Device Management tool still displayed in Service Tool | WIP18 | WIP19 BB_WIP9 |  | Subsystem | Device Management | NA | Fixed | s_shabeer | 0 | 4 | Service Tools_Software | Service Tools | NO | Closed | WIP19SubCom | Ovation |  | WIP19 BB_WIP9 | August 5, 2019 9:09:37 PM GMT+08:00
CTUDT00625352 | Refresh | Closed | wang_l | July 30, 2019 2:00:51 PM GMT+08:00 | Refresh: Device Management tool still displayed in Service Tool | WIP18 | WIP19 BB_WIP9 |  | Subsystem | Device Management | NA | Fixed | s_shabeer | 0 | 4 | Service Tools_Software | Service Tools | NO | Closed | WIP19SubCom | Brilliance 64 |  | WIP19 BB_WIP9 | August 5, 2019 9:09:37 PM GMT+08:00
CTUDT00625352 | Refresh | Closed | wang_l | July 30, 2019 2:00:51 PM GMT+08:00 | Refresh: Device Management tool still displayed in Service Tool | WIP18 | WIP19 BB_WIP9 |  | Subsystem | Device Management | NA | Fixed | s_shabeer | 0 | 4 | Service Tools_Software | Service Tools | NO | Closed | WIP19SubCom | Brilliance Big Bore |  | WIP19 BB_WIP9 | August 5, 2019 9:09:37 PM GMT+08:00
CTUDT00625352 | Refresh | Closed | wang_l | July 30, 2019 2:00:51 PM GMT+08:00 | Refresh: Device Management tool still displayed in Service Tool | WIP18 | WIP19 BB_WIP9 |  | Subsystem | Device Management | NA | Fixed | s_shabeer | 0 | 4 | Service Tools_Software | Service Tools | NO | Closed | WIP19SubCom | Ingenuity 64 |  | WIP19 BB_WIP9 | August 5, 2019 9:09:37 PM GMT+08:00
CTUDT00625352 | Refresh | Closed | wang_l | July 30, 2019 2:00:51 PM GMT+08:00 | Refresh: Device Management tool still displayed in Service Tool | WIP18 | WIP19 BB_WIP9 |  | Subsystem | Device Management | NA | Fixed | s_shabeer | 0 | 4 | Service Tools_Software | Service Tools | NO | Closed | WIP19SubCom | IQon1 |  | WIP19 BB_WIP9 | August 5, 2019 9:09:37 PM GMT+08:00
CTUDT00625351 | Refresh | Closed | kapadia_v | July 30, 2019 12:00:36 PM GMT+08:00 | [Refresh] “Microsoft. NET Framework” error pop up when try to delete a renamed language set | 4.8 | BB_WIP9 | None | Subsystem | Voice Manager | NA | Fixed | jose_s | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP9_SubCom | Brilliance Big Bore |  | BB_WIP9 | August 2, 2019 8:35:22 PM GMT+08:00
CTUDT00625349 | Refresh | Closed | wang_l | July 30, 2019 11:32:09 AM GMT+08:00 | Refresh: Not all images generated for reconstruction with MPR | WIP18 |  | Perform Offline Recon | Subsystem | CIRS | N/A | No Plan to Fix | richards_b | 0 | 4 |  | CIRS | NO | Closed | WIP20SubCom | Ovation
CTUDT00625349 | Refresh | Closed | wang_l | July 30, 2019 11:32:09 AM GMT+08:00 | Refresh: Not all images generated for reconstruction with MPR | WIP18 |  | Perform Offline Recon | Subsystem | CIRS | N/A | No Plan to Fix | richards_b | 0 | 4 |  | CIRS | NO | Closed | WIP20SubCom | Brilliance 64
CTUDT00625349 | Refresh | Closed | wang_l | July 30, 2019 11:32:09 AM GMT+08:00 | Refresh: Not all images generated for reconstruction with MPR | WIP18 |  | Perform Offline Recon | Subsystem | CIRS | N/A | No Plan to Fix | richards_b | 0 | 4 |  | CIRS | NO | Closed | WIP20SubCom | Brilliance Big Bore
CTUDT00625349 | Refresh | Closed | wang_l | July 30, 2019 11:32:09 AM GMT+08:00 | Refresh: Not all images generated for reconstruction with MPR | WIP18 |  | Perform Offline Recon | Subsystem | CIRS | N/A | No Plan to Fix | richards_b | 0 | 4 |  | CIRS | NO | Closed | WIP20SubCom | Ingenuity 64
CTUDT00625349 | Refresh | Closed | wang_l | July 30, 2019 11:32:09 AM GMT+08:00 | Refresh: Not all images generated for reconstruction with MPR | WIP18 |  | Perform Offline Recon | Subsystem | CIRS | N/A | No Plan to Fix | richards_b | 0 | 4 |  | CIRS | NO | Closed | WIP20SubCom | IQon1
CTUDT00625347 | Refresh | Closed | cohen_y | July 29, 2019 10:11:55 PM GMT+08:00 | Refresh: Calibrations: Running Copy Tables script can cause an overwrite of OffFocal_120kV_FSSmall_adult.tbl in CIRS and therefore a Recon Error | WIP18 | WIP19 | Perform off-focal calibrations | Subsystem | CIRS | n/a | Fixed | frider_o | 0 | 3 | DMS Calibrations | DMS | YES | Closed | WIP19SubCom | IQon1 |  | WIP19 | August 6, 2019 4:10:24 PM GMT+08:00
CTUDT00625345 | Refresh | Closed | cohen_y | July 29, 2019 9:53:02 PM GMT+08:00 | Refresh: Calibrations: Steps Test: Steps Z Scan fails due to Surview Length error | BB_WIP8 | WIP19 |  | Subsystem | Calibrations | NaN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP19SubCom | Brilliance Big Bore |  | WIP19 | August 4, 2019 3:34:12 PM GMT+08:00
CTUDT00625343 | Refresh | Closed | frider_o | July 29, 2019 9:15:04 PM GMT+08:00 | The OffFocal Table name isn't passed in the MDU | WIP18 | WIP19 |  | System | MDU;blocker19;Verified1 | NA | Fixed | tamari_o | 0 | 2 | Console Software | Console4 | NO | Closed | WIP19SubCom | Brilliance 64 |  | WIP19 | August 6, 2019 11:30:54 PM GMT+08:00
CTUDT00625343 | Refresh | Closed | frider_o | July 29, 2019 9:15:04 PM GMT+08:00 | The OffFocal Table name isn't passed in the MDU | WIP18 | WIP19 |  | System | MDU;blocker19;Verified1 | NA | Fixed | tamari_o | 0 | 2 | Console Software | Console4 | NO | Closed | WIP19SubCom | Ingenuity 64 |  | WIP19 | August 6, 2019 11:30:54 PM GMT+08:00
CTUDT00625340 | Refresh | Closed | azagury_d | July 29, 2019 8:21:38 PM GMT+08:00 | Refresh: Calibrations: Acceptance and Constancy: Constancy in CT Summary does not appear in the report. | BB_WIP8 | WIP19 | N/A | Subsystem | Acceptance/Constancy | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP19SubCom | Ovation |  | WIP19 | August 7, 2019 4:00:30 PM GMT+08:00
CTUDT00625340 | Refresh | Closed | azagury_d | July 29, 2019 8:21:38 PM GMT+08:00 | Refresh: Calibrations: Acceptance and Constancy: Constancy in CT Summary does not appear in the report. | BB_WIP8 | WIP19 | N/A | Subsystem | Acceptance/Constancy | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP19SubCom | Brilliance Big Bore |  | WIP19 | August 7, 2019 4:00:30 PM GMT+08:00
CTUDT00625340 | Refresh | Closed | azagury_d | July 29, 2019 8:21:38 PM GMT+08:00 | Refresh: Calibrations: Acceptance and Constancy: Constancy in CT Summary does not appear in the report. | BB_WIP8 | WIP19 | N/A | Subsystem | Acceptance/Constancy | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP19SubCom | Ingenuity 64 |  | WIP19 | August 7, 2019 4:00:30 PM GMT+08:00
CTUDT00625340 | Refresh | Closed | azagury_d | July 29, 2019 8:21:38 PM GMT+08:00 | Refresh: Calibrations: Acceptance and Constancy: Constancy in CT Summary does not appear in the report. | BB_WIP8 | WIP19 | N/A | Subsystem | Acceptance/Constancy | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP19SubCom | IQon1 |  | WIP19 | August 7, 2019 4:00:30 PM GMT+08:00
CTUDT00625338 | Refresh | Closed | lakshmikanth_s | July 29, 2019 8:06:14 PM GMT+08:00 | Big Bore: All saved needles are not displayed in context menu | BB_WIP7 |  |  | Subsystem | CCT | Console-SRS.17821, Console-SRS.17804, Console-SRS.17807 | Does Not Apply | navale_s | 0 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00625336 | Refresh | Closed | Noy_Seelig | July 29, 2019 7:47:40 PM GMT+08:00 | Refresh: Calibrations: Beam and DMS Check: couch moves in RnD user | WIP18 | N/A |  | Subsystem | Calibrations | n/a | Fixed | yosfan_s | 0 | 3 | Test Design | DMS | NO | Closed | WIP19SubCom | Ovation |  | N/A | August 7, 2019 3:27:36 PM GMT+08:00
CTUDT00625336 | Refresh | Closed | Noy_Seelig | July 29, 2019 7:47:40 PM GMT+08:00 | Refresh: Calibrations: Beam and DMS Check: couch moves in RnD user | WIP18 | N/A |  | Subsystem | Calibrations | n/a | Fixed | yosfan_s | 0 | 3 | Test Design | DMS | NO | Closed | WIP19SubCom | Brilliance Big Bore |  | N/A | August 7, 2019 3:27:36 PM GMT+08:00
CTUDT00625336 | Refresh | Closed | Noy_Seelig | July 29, 2019 7:47:40 PM GMT+08:00 | Refresh: Calibrations: Beam and DMS Check: couch moves in RnD user | WIP18 | N/A |  | Subsystem | Calibrations | n/a | Fixed | yosfan_s | 0 | 3 | Test Design | DMS | NO | Closed | WIP19SubCom | Ingenuity 64 |  | N/A | August 7, 2019 3:27:36 PM GMT+08:00
CTUDT00625336 | Refresh | Closed | Noy_Seelig | July 29, 2019 7:47:40 PM GMT+08:00 | Refresh: Calibrations: Beam and DMS Check: couch moves in RnD user | WIP18 | N/A |  | Subsystem | Calibrations | n/a | Fixed | yosfan_s | 0 | 3 | Test Design | DMS | NO | Closed | WIP19SubCom | IQon1 |  | N/A | August 7, 2019 3:27:36 PM GMT+08:00
CTUDT00625335 | Refresh | Closed | wang_l | July 29, 2019 7:46:23 PM GMT+08:00 | Refresh: The configured default printer does not take effect for other users | BB_WIP8 | WIP20 BB_WIP10 | NA | System | Printer Config Tool;Blocker;Verified1 | NA | Fixed | tejaswini_nidamanuri | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP19SegReq | Ovation |  | WIP20 BB_WIP10 | September 20, 2019 1:11:29 PM GMT+08:00
CTUDT00625335 | Refresh | Closed | wang_l | July 29, 2019 7:46:23 PM GMT+08:00 | Refresh: The configured default printer does not take effect for other users | BB_WIP8 | WIP20 BB_WIP10 | NA | System | Printer Config Tool;Blocker;Verified1 | NA | Fixed | tejaswini_nidamanuri | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP19SegReq | Brilliance 64 |  | WIP20 BB_WIP10 | September 20, 2019 1:11:29 PM GMT+08:00
CTUDT00625335 | Refresh | Closed | wang_l | July 29, 2019 7:46:23 PM GMT+08:00 | Refresh: The configured default printer does not take effect for other users | BB_WIP8 | WIP20 BB_WIP10 | NA | System | Printer Config Tool;Blocker;Verified1 | NA | Fixed | tejaswini_nidamanuri | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP19SegReq | Brilliance Big Bore |  | WIP20 BB_WIP10 | September 20, 2019 1:11:29 PM GMT+08:00
CTUDT00625335 | Refresh | Closed | wang_l | July 29, 2019 7:46:23 PM GMT+08:00 | Refresh: The configured default printer does not take effect for other users | BB_WIP8 | WIP20 BB_WIP10 | NA | System | Printer Config Tool;Blocker;Verified1 | NA | Fixed | tejaswini_nidamanuri | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP19SegReq | Ingenuity 64 |  | WIP20 BB_WIP10 | September 20, 2019 1:11:29 PM GMT+08:00
CTUDT00625335 | Refresh | Closed | wang_l | July 29, 2019 7:46:23 PM GMT+08:00 | Refresh: The configured default printer does not take effect for other users | BB_WIP8 | WIP20 BB_WIP10 | NA | System | Printer Config Tool;Blocker;Verified1 | NA | Fixed | tejaswini_nidamanuri | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP19SegReq | IQon1 |  | WIP20 BB_WIP10 | September 20, 2019 1:11:29 PM GMT+08:00
CTUDT00625331 | Refresh | Closed | 320053767 | July 29, 2019 7:05:13 PM GMT+08:00 | Incorrect pin position. | WIP18 |  |  | Subsystem |  |  | Not Enough Information | 320053767 | 0 |  |  |  |  | Closed |  | Ovation
CTUDT00625331 | Refresh | Closed | 320053767 | July 29, 2019 7:05:13 PM GMT+08:00 | Incorrect pin position. | WIP18 |  |  | Subsystem |  |  | Not Enough Information | 320053767 | 0 |  |  |  |  | Closed |  | IQon1
CTUDT00625330 | Refresh | Closed | PremKumar_Rajendran | July 29, 2019 6:56:13 PM GMT+08:00 | Refresh - DLP values mismatched between Show-all parameters vs Dose Log Viewer | WIP18 |  |  | Subsystem | Dose Log Viewer | NA | Works As Specified / Not A Defect | m_nanda | 0 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00625330 | Refresh | Closed | PremKumar_Rajendran | July 29, 2019 6:56:13 PM GMT+08:00 | Refresh - DLP values mismatched between Show-all parameters vs Dose Log Viewer | WIP18 |  |  | Subsystem | Dose Log Viewer | NA | Works As Specified / Not A Defect | m_nanda | 0 |  |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00625330 | Refresh | Closed | PremKumar_Rajendran | July 29, 2019 6:56:13 PM GMT+08:00 | Refresh - DLP values mismatched between Show-all parameters vs Dose Log Viewer | WIP18 |  |  | Subsystem | Dose Log Viewer | NA | Works As Specified / Not A Defect | m_nanda | 0 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00625330 | Refresh | Closed | PremKumar_Rajendran | July 29, 2019 6:56:13 PM GMT+08:00 | Refresh - DLP values mismatched between Show-all parameters vs Dose Log Viewer | WIP18 |  |  | Subsystem | Dose Log Viewer | NA | Works As Specified / Not A Defect | m_nanda | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00625330 | Refresh | Closed | PremKumar_Rajendran | July 29, 2019 6:56:13 PM GMT+08:00 | Refresh - DLP values mismatched between Show-all parameters vs Dose Log Viewer | WIP18 |  |  | Subsystem | Dose Log Viewer | NA | Works As Specified / Not A Defect | m_nanda | 0 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00625324 | Refresh | Closed | 320053767 | July 29, 2019 4:36:08 PM GMT+08:00 | Wrong Images in calibration help application | WIP18 |  |  | Subsystem |  | N/A | No Plan to Fix | braunstein_d | 0 | 3 | DMS Calibrations | Support-Physics | NO | Closed | WIP21SegReq | Ovation
CTUDT00625324 | Refresh | Closed | 320053767 | July 29, 2019 4:36:08 PM GMT+08:00 | Wrong Images in calibration help application | WIP18 |  |  | Subsystem |  | N/A | No Plan to Fix | braunstein_d | 0 | 3 | DMS Calibrations | Support-Physics | NO | Closed | WIP21SegReq | IQon1
CTUDT00625321 | Refresh | Closed | PremKumar_Rajendran | July 29, 2019 3:28:27 PM GMT+08:00 | BB Refresh - When launching from O-Level User iDMS Utility tool throwing an unhandled exception error | BB_WIP7 | BB_WIP10 |  | Subsystem | Olevel | DMS-SRS-656 | Fixed | s_shabeer | 0 | 2 | DMS Software | Service Tools | NO | Closed | WIP19SubCom | Brilliance Big Bore |  | BB_WIP10 | August 16, 2019 1:55:24 PM GMT+08:00
CTUDT00625318 | Refresh | Closed | 320053769 | July 29, 2019 2:35:49 PM GMT+08:00 | Burning a study with 200 images (of matrix 512X512) to CD takes longer than 3 minutes. | WIP18 | ING, BR64 SRS DHF298554 Rev08 BB SRS DHF298555 Rev08 Updated with D-389368 for iCT/iCT-SP and IQon |  | Subsystem | Requirement;Verified1 | CT-NM.SRS.Interop.176 | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP20SegReq | Ovation |  | ING, BR64 SRS DHF298554 Rev08 BB SRS DHF298555 Rev08 Updated with D-389368 for iCT/iCT-SP and IQon | November 6, 2019 1:51:30 PM GMT+08:00
CTUDT00625318 | Refresh | Closed | 320053769 | July 29, 2019 2:35:49 PM GMT+08:00 | Burning a study with 200 images (of matrix 512X512) to CD takes longer than 3 minutes. | WIP18 | ING, BR64 SRS DHF298554 Rev08 BB SRS DHF298555 Rev08 Updated with D-389368 for iCT/iCT-SP and IQon |  | Subsystem | Requirement;Verified1 | CT-NM.SRS.Interop.176 | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP20SegReq | Brilliance 64 |  | ING, BR64 SRS DHF298554 Rev08 BB SRS DHF298555 Rev08 Updated with D-389368 for iCT/iCT-SP and IQon | November 6, 2019 1:51:30 PM GMT+08:00
CTUDT00625318 | Refresh | Closed | 320053769 | July 29, 2019 2:35:49 PM GMT+08:00 | Burning a study with 200 images (of matrix 512X512) to CD takes longer than 3 minutes. | WIP18 | ING, BR64 SRS DHF298554 Rev08 BB SRS DHF298555 Rev08 Updated with D-389368 for iCT/iCT-SP and IQon |  | Subsystem | Requirement;Verified1 | CT-NM.SRS.Interop.176 | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP20SegReq | Brilliance Big Bore |  | ING, BR64 SRS DHF298554 Rev08 BB SRS DHF298555 Rev08 Updated with D-389368 for iCT/iCT-SP and IQon | November 6, 2019 1:51:30 PM GMT+08:00
CTUDT00625318 | Refresh | Closed | 320053769 | July 29, 2019 2:35:49 PM GMT+08:00 | Burning a study with 200 images (of matrix 512X512) to CD takes longer than 3 minutes. | WIP18 | ING, BR64 SRS DHF298554 Rev08 BB SRS DHF298555 Rev08 Updated with D-389368 for iCT/iCT-SP and IQon |  | Subsystem | Requirement;Verified1 | CT-NM.SRS.Interop.176 | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP20SegReq | Ingenuity 64 |  | ING, BR64 SRS DHF298554 Rev08 BB SRS DHF298555 Rev08 Updated with D-389368 for iCT/iCT-SP and IQon | November 6, 2019 1:51:30 PM GMT+08:00
CTUDT00625318 | Refresh | Closed | 320053769 | July 29, 2019 2:35:49 PM GMT+08:00 | Burning a study with 200 images (of matrix 512X512) to CD takes longer than 3 minutes. | WIP18 | ING, BR64 SRS DHF298554 Rev08 BB SRS DHF298555 Rev08 Updated with D-389368 for iCT/iCT-SP and IQon |  | Subsystem | Requirement;Verified1 | CT-NM.SRS.Interop.176 | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP20SegReq | IQon1 |  | ING, BR64 SRS DHF298554 Rev08 BB SRS DHF298555 Rev08 Updated with D-389368 for iCT/iCT-SP and IQon | November 6, 2019 1:51:30 PM GMT+08:00
CTUDT00625315 | Refresh | Closed | ding_s | July 29, 2019 12:11:28 PM GMT+08:00 | IMR recon mode for Gated Axial exam card is not available when already enabled IMR function. | BB_WIP8 |  | N/A | Subsystem | License | N/A | Works As Specified / Not A Defect | kr_m | 0 |  |  | Console4 | NO | Closed | WAS | Brilliance Big Bore
CTUDT00625313 | Refresh | Closed | ding_s | July 29, 2019 9:39:35 AM GMT+08:00 | There is no 'Sampled Collimation' displayed in QA Head Axial 3D and QA Body Axial 3D scan type, cannot set 32 images. | BB_WIP6 | BB_WIP9 |  | Subsystem | Plan;Blocker19 | N/A | Fixed | pandey_a | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP9_SegReq | Brilliance Big Bore |  | BB_WIP9 | August 6, 2019 8:19:12 PM GMT+08:00
CTUDT00625311 | Refresh | Closed | Noy_Seelig | July 28, 2019 11:53:22 PM GMT+08:00 | Refresh: Calibrations: Tube Current, Scan Time, mAs Test: Wrong status in HTML report for all scans in RnD user | BB_WIP8 | WIP19 |  | Subsystem | Calibrations | n/a | Fixed | bikman | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP19SubCom | Brilliance Big Bore |  | WIP19 | August 7, 2019 7:09:00 PM GMT+08:00
CTUDT00625309 | Refresh | Closed | Noy_Seelig | July 28, 2019 11:39:04 PM GMT+08:00 | Refresh: Calibrations: Tube Current, Scan Time, mAs Test: Wrong limits in HTML Report in Production user | BB_WIP8 | WIP19 |  | Subsystem | Calibrations | n/a | Fixed | bikman | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP19SubCom | Brilliance Big Bore |  | WIP19 | August 7, 2019 7:07:15 PM GMT+08:00
CTUDT00625306 | Refresh | Closed | ding_s | July 28, 2019 4:40:12 PM GMT+08:00 | Refresh BigBore doesn't support concurrent reconstruction, CT-NM.SRS.ProductConfig.511 should be removed from SRS. | BB_WIP6 | BB_WIP9 |  | Subsystem | Requirements | CT-NM.SRS.ProductConfig.511 | Fixed | ye_h | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | BB_WIP9_SubCom | Brilliance Big Bore |  | BB_WIP9 | September 4, 2019 10:15:23 AM GMT+08:00
CTUDT00625302 | Refresh | Closed | zyanjiang | July 26, 2019 3:07:14 PM GMT+08:00 | Refresh Pre-Val: the volume of contrast agent bar can change in real time | BB_WIP6 |  | None | Subsystem | Injector | None | Works As Specified / Not A Defect | jose_s | 0 |  |  | Console4 | NO | Closed | WAS | Brilliance Big Bore
CTUDT00625300 | Refresh | Closed | zyanjiang | July 26, 2019 2:58:03 PM GMT+08:00 | Refresh Pre-Val: accession number to split studies have same DLP and requested procedure | BB_WIP6 | BB_WIP9 | NA | Subsystem | Exam Summary | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP9_SubCom | Brilliance Big Bore |  | BB_WIP9 | August 9, 2019 3:45:32 PM GMT+08:00
CTUDT00625298 | Refresh | Closed | zyanjiang | July 26, 2019 2:41:12 PM GMT+08:00 | Refresh Pre-Val:Stuck in worklist | BB_WIP6 |  |  | Subsystem | Patient Scene |  | Duplicate | kr_m | 0 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00625296 | Refresh | Closed | manju_kolkar | July 26, 2019 1:35:56 PM GMT+08:00 | Big Bore Refresh - “Review Before Save” and “Edit Result” options were disabled | BB_WIP7 |  |  | Subsystem | Crosscheck | NA | Works As Specified / Not A Defect | prasad_r | 0 |  |  | Console4 | NO | Closed | WAS | Brilliance Big Bore
CTUDT00625293 | Refresh | Closed | wang_l | July 26, 2019 1:03:22 PM GMT+08:00 | Refresh - The infant phantom and eye phantom are N/A in the help page, but actually exist in the Phantom Centering customized list. | BB_WIP8 |  |  | Subsystem | Calibrations | N/A | Works As Specified / Not A Defect | yatziv_l | 0 | 2 |  | DMS | NO | Closed | WIP20SubCom | Ovation
CTUDT00625293 | Refresh | Closed | wang_l | July 26, 2019 1:03:22 PM GMT+08:00 | Refresh - The infant phantom and eye phantom are N/A in the help page, but actually exist in the Phantom Centering customized list. | BB_WIP8 |  |  | Subsystem | Calibrations | N/A | Works As Specified / Not A Defect | yatziv_l | 0 | 2 |  | DMS | NO | Closed | WIP20SubCom | Brilliance 64
CTUDT00625293 | Refresh | Closed | wang_l | July 26, 2019 1:03:22 PM GMT+08:00 | Refresh - The infant phantom and eye phantom are N/A in the help page, but actually exist in the Phantom Centering customized list. | BB_WIP8 |  |  | Subsystem | Calibrations | N/A | Works As Specified / Not A Defect | yatziv_l | 0 | 2 |  | DMS | NO | Closed | WIP20SubCom | Brilliance Big Bore
CTUDT00625293 | Refresh | Closed | wang_l | July 26, 2019 1:03:22 PM GMT+08:00 | Refresh - The infant phantom and eye phantom are N/A in the help page, but actually exist in the Phantom Centering customized list. | BB_WIP8 |  |  | Subsystem | Calibrations | N/A | Works As Specified / Not A Defect | yatziv_l | 0 | 2 |  | DMS | NO | Closed | WIP20SubCom | Ingenuity 64
CTUDT00625293 | Refresh | Closed | wang_l | July 26, 2019 1:03:22 PM GMT+08:00 | Refresh - The infant phantom and eye phantom are N/A in the help page, but actually exist in the Phantom Centering customized list. | BB_WIP8 |  |  | Subsystem | Calibrations | N/A | Works As Specified / Not A Defect | yatziv_l | 0 | 2 |  | DMS | NO | Closed | WIP20SubCom | IQon1
CTUDT00625289 | Refresh | Closed | attali_s | July 25, 2019 10:33:43 PM GMT+08:00 | Refresh: Confusing final status after re-performing failed cases | WIP18 |  |  | Subsystem | Calibrations | n/a | No Plan to Fix | yatziv_l | 0 | 4 |  | DMS | NO | Closed |  | Ovation
CTUDT00625289 | Refresh | Closed | attali_s | July 25, 2019 10:33:43 PM GMT+08:00 | Refresh: Confusing final status after re-performing failed cases | WIP18 |  |  | Subsystem | Calibrations | n/a | No Plan to Fix | yatziv_l | 0 | 4 |  | DMS | NO | Closed |  | Brilliance Big Bore
CTUDT00625289 | Refresh | Closed | attali_s | July 25, 2019 10:33:43 PM GMT+08:00 | Refresh: Confusing final status after re-performing failed cases | WIP18 |  |  | Subsystem | Calibrations | n/a | No Plan to Fix | yatziv_l | 0 | 4 |  | DMS | NO | Closed |  | Ingenuity 64
CTUDT00625289 | Refresh | Closed | attali_s | July 25, 2019 10:33:43 PM GMT+08:00 | Refresh: Confusing final status after re-performing failed cases | WIP18 |  |  | Subsystem | Calibrations | n/a | No Plan to Fix | yatziv_l | 0 | 4 |  | DMS | NO | Closed |  | IQon1
CTUDT00625286 | Refresh | Closed | attali_s | July 25, 2019 10:20:24 PM GMT+08:00 | Refresh: ACR Test: Scan Parameters (Resolution and mAs) are not as defined in the spec | WIP18 | WIP19 |  | Subsystem | Calibrations | NA | Fixed | Kostya_Baklan | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP19SubCom | Ingenuity 64 |  | WIP19 | July 31, 2019 6:07:41 PM GMT+08:00
CTUDT00625285 | Refresh | Closed | Valeria_Kashtanov | July 25, 2019 10:10:50 PM GMT+08:00 | Refresh: Calibrations: Hcor Slice Infant, RT of head and body cases is not according to spec | WIP18 | WIP19 | N/A | Subsystem | Calibrations | N/A | Fixed | elviral | 0 | 4 | DMS Calibrations | DMS | YES | Closed | WIP19SubCom | Ovation |  | WIP19 | August 12, 2019 8:20:21 PM GMT+08:00
CTUDT00625285 | Refresh | Closed | Valeria_Kashtanov | July 25, 2019 10:10:50 PM GMT+08:00 | Refresh: Calibrations: Hcor Slice Infant, RT of head and body cases is not according to spec | WIP18 | WIP19 | N/A | Subsystem | Calibrations | N/A | Fixed | elviral | 0 | 4 | DMS Calibrations | DMS | YES | Closed | WIP19SubCom | Ingenuity 64 |  | WIP19 | August 12, 2019 8:20:21 PM GMT+08:00
CTUDT00625283 | Refresh | Closed | Valeria_Kashtanov | July 25, 2019 9:46:20 PM GMT+08:00 | Refresh: Calibrations: Performance Test: Noise ROI radius of IMR 1a case is not according to spec | WIP18 |  |  | Subsystem | Performance Test | N/A | Duplicate | yatziv_l | 0 |  |  | DMS | NO | Closed |  | Ingenuity 64
CTUDT00625281 | Refresh | Closed | datta_i | July 25, 2019 8:31:55 PM GMT+08:00 | The console allowed to add duplicate exam cards in the suggested EC group | 4.8 | BB_WIP9 |  | Subsystem | ECM | NA | Fixed | datta_i | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP9_SubCom | Brilliance Big Bore |  | BB_WIP9 | July 30, 2019 7:21:27 PM GMT+08:00
CTUDT00625274 | Refresh | Closed | gT_s | July 25, 2019 6:39:48 PM GMT+08:00 | Refresh : The YES & No button is not localized as per selected Norwegian language. | WIP18 | WIP19 |  | Subsystem | Localization | NA | Fixed | rath_s | 0 | 4 | Console Software | Console4 | NO | Closed | WIP19SubCom | Ovation |  | WIP19 | August 2, 2019 1:41:48 PM GMT+08:00
CTUDT00625274 | Refresh | Closed | gT_s | July 25, 2019 6:39:48 PM GMT+08:00 | Refresh : The YES & No button is not localized as per selected Norwegian language. | WIP18 | WIP19 |  | Subsystem | Localization | NA | Fixed | rath_s | 0 | 4 | Console Software | Console4 | NO | Closed | WIP19SubCom | Brilliance 64 |  | WIP19 | August 2, 2019 1:41:48 PM GMT+08:00
CTUDT00625274 | Refresh | Closed | gT_s | July 25, 2019 6:39:48 PM GMT+08:00 | Refresh : The YES & No button is not localized as per selected Norwegian language. | WIP18 | WIP19 |  | Subsystem | Localization | NA | Fixed | rath_s | 0 | 4 | Console Software | Console4 | NO | Closed | WIP19SubCom | Ingenuity 64 |  | WIP19 | August 2, 2019 1:41:48 PM GMT+08:00
CTUDT00625274 | Refresh | Closed | gT_s | July 25, 2019 6:39:48 PM GMT+08:00 | Refresh : The YES & No button is not localized as per selected Norwegian language. | WIP18 | WIP19 |  | Subsystem | Localization | NA | Fixed | rath_s | 0 | 4 | Console Software | Console4 | NO | Closed | WIP19SubCom | IQon1 |  | WIP19 | August 2, 2019 1:41:48 PM GMT+08:00
CTUDT00625273 | Refresh | Closed | 320053770 | July 25, 2019 6:38:53 PM GMT+08:00 | Refresh WIP 18: ICT & IQon: At the Reconstruction tab in the Patient Directory - the Time&date are stuck together with no space between them (Regression from HB).” | WIP18 | WIP19 | None | Subsystem | Patient Directory | Console.SRS-3323 | Fixed | srushti_k | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP19SegReq | Ovation |  | WIP19 | August 16, 2019 5:20:03 PM GMT+08:00
CTUDT00625273 | Refresh | Closed | 320053770 | July 25, 2019 6:38:53 PM GMT+08:00 | Refresh WIP 18: ICT & IQon: At the Reconstruction tab in the Patient Directory - the Time&date are stuck together with no space between them (Regression from HB).” | WIP18 | WIP19 | None | Subsystem | Patient Directory | Console.SRS-3323 | Fixed | srushti_k | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP19SegReq | Brilliance 64 |  | WIP19 | August 16, 2019 5:20:03 PM GMT+08:00
CTUDT00625273 | Refresh | Closed | 320053770 | July 25, 2019 6:38:53 PM GMT+08:00 | Refresh WIP 18: ICT & IQon: At the Reconstruction tab in the Patient Directory - the Time&date are stuck together with no space between them (Regression from HB).” | WIP18 | WIP19 | None | Subsystem | Patient Directory | Console.SRS-3323 | Fixed | srushti_k | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP19SegReq | Brilliance Big Bore |  | WIP19 | August 16, 2019 5:20:03 PM GMT+08:00
CTUDT00625273 | Refresh | Closed | 320053770 | July 25, 2019 6:38:53 PM GMT+08:00 | Refresh WIP 18: ICT & IQon: At the Reconstruction tab in the Patient Directory - the Time&date are stuck together with no space between them (Regression from HB).” | WIP18 | WIP19 | None | Subsystem | Patient Directory | Console.SRS-3323 | Fixed | srushti_k | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP19SegReq | Ingenuity 64 |  | WIP19 | August 16, 2019 5:20:03 PM GMT+08:00
CTUDT00625273 | Refresh | Closed | 320053770 | July 25, 2019 6:38:53 PM GMT+08:00 | Refresh WIP 18: ICT & IQon: At the Reconstruction tab in the Patient Directory - the Time&date are stuck together with no space between them (Regression from HB).” | WIP18 | WIP19 | None | Subsystem | Patient Directory | Console.SRS-3323 | Fixed | srushti_k | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP19SegReq | IQon1 |  | WIP19 | August 16, 2019 5:20:03 PM GMT+08:00
CTUDT00625270 | Refresh | Closed | 320053770 | July 25, 2019 6:26:35 PM GMT+08:00 | WIP 18: ICT & IQon: in service tools - RDP remote viewer, when connecting to any system(CIRS\GPC), the boundaries of the opened window doesn’t match the monitor.” | WIP18 | WIP20 BB_WIP10 | User can maximize the remote viewer and continue using it. | Subsystem | RDP;Blocker20 | N/A | Fixed | chakraborty_d | 0 | 3 | Service Tools_Software | Service Tools | NO | Closed | WIP20SubCom | Ovation |  | WIP20 BB_WIP10 | September 4, 2019 7:04:33 PM GMT+08:00
CTUDT00625270 | Refresh | Closed | 320053770 | July 25, 2019 6:26:35 PM GMT+08:00 | WIP 18: ICT & IQon: in service tools - RDP remote viewer, when connecting to any system(CIRS\GPC), the boundaries of the opened window doesn’t match the monitor.” | WIP18 | WIP20 BB_WIP10 | User can maximize the remote viewer and continue using it. | Subsystem | RDP;Blocker20 | N/A | Fixed | chakraborty_d | 0 | 3 | Service Tools_Software | Service Tools | NO | Closed | WIP20SubCom | Brilliance 64 |  | WIP20 BB_WIP10 | September 4, 2019 7:04:33 PM GMT+08:00
CTUDT00625270 | Refresh | Closed | 320053770 | July 25, 2019 6:26:35 PM GMT+08:00 | WIP 18: ICT & IQon: in service tools - RDP remote viewer, when connecting to any system(CIRS\GPC), the boundaries of the opened window doesn’t match the monitor.” | WIP18 | WIP20 BB_WIP10 | User can maximize the remote viewer and continue using it. | Subsystem | RDP;Blocker20 | N/A | Fixed | chakraborty_d | 0 | 3 | Service Tools_Software | Service Tools | NO | Closed | WIP20SubCom | Brilliance Big Bore |  | WIP20 BB_WIP10 | September 4, 2019 7:04:33 PM GMT+08:00
CTUDT00625270 | Refresh | Closed | 320053770 | July 25, 2019 6:26:35 PM GMT+08:00 | WIP 18: ICT & IQon: in service tools - RDP remote viewer, when connecting to any system(CIRS\GPC), the boundaries of the opened window doesn’t match the monitor.” | WIP18 | WIP20 BB_WIP10 | User can maximize the remote viewer and continue using it. | Subsystem | RDP;Blocker20 | N/A | Fixed | chakraborty_d | 0 | 3 | Service Tools_Software | Service Tools | NO | Closed | WIP20SubCom | Ingenuity 64 |  | WIP20 BB_WIP10 | September 4, 2019 7:04:33 PM GMT+08:00
CTUDT00625270 | Refresh | Closed | 320053770 | July 25, 2019 6:26:35 PM GMT+08:00 | WIP 18: ICT & IQon: in service tools - RDP remote viewer, when connecting to any system(CIRS\GPC), the boundaries of the opened window doesn’t match the monitor.” | WIP18 | WIP20 BB_WIP10 | User can maximize the remote viewer and continue using it. | Subsystem | RDP;Blocker20 | N/A | Fixed | chakraborty_d | 0 | 3 | Service Tools_Software | Service Tools | NO | Closed | WIP20SubCom | IQon1 |  | WIP20 BB_WIP10 | September 4, 2019 7:04:33 PM GMT+08:00
CTUDT00625265 | Refresh | Closed | Valeria_Kashtanov | July 25, 2019 6:02:09 PM GMT+08:00 | Refresh: Calibrations: Performance Infant Test: Window value of Body Infant scans is not according to spec | WIP18 | WIP19 | N/A | Subsystem | Performance Test; Verified1 | N/A | Fixed | elviral | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP19SubCom | Ovation |  | WIP19 | August 4, 2019 3:51:21 PM GMT+08:00
CTUDT00625265 | Refresh | Closed | Valeria_Kashtanov | July 25, 2019 6:02:09 PM GMT+08:00 | Refresh: Calibrations: Performance Infant Test: Window value of Body Infant scans is not according to spec | WIP18 | WIP19 | N/A | Subsystem | Performance Test; Verified1 | N/A | Fixed | elviral | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP19SubCom | Brilliance Big Bore |  | WIP19 | August 4, 2019 3:51:21 PM GMT+08:00
CTUDT00625265 | Refresh | Closed | Valeria_Kashtanov | July 25, 2019 6:02:09 PM GMT+08:00 | Refresh: Calibrations: Performance Infant Test: Window value of Body Infant scans is not according to spec | WIP18 | WIP19 | N/A | Subsystem | Performance Test; Verified1 | N/A | Fixed | elviral | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP19SubCom | Ingenuity 64 |  | WIP19 | August 4, 2019 3:51:21 PM GMT+08:00
CTUDT00625265 | Refresh | Closed | Valeria_Kashtanov | July 25, 2019 6:02:09 PM GMT+08:00 | Refresh: Calibrations: Performance Infant Test: Window value of Body Infant scans is not according to spec | WIP18 | WIP19 | N/A | Subsystem | Performance Test; Verified1 | N/A | Fixed | elviral | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP19SubCom | IQon1 |  | WIP19 | August 4, 2019 3:51:21 PM GMT+08:00
CTUDT00625263 | Refresh | Closed | zhang_t | July 25, 2019 4:56:53 PM GMT+08:00 | Refresh: NTP time sync- the registry of HOST W32tm server is NoSync after 1 click installation, it shall be NTP. | WIP18 | WIP19 BB_WIP9 | Manually reconfigure after installation | System | Installation;Blocker19 | N/A | Fixed | dasgupta_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP19SubCom | Ovation |  | WIP19 BB_WIP9 | July 31, 2019 2:09:31 PM GMT+08:00
CTUDT00625263 | Refresh | Closed | zhang_t | July 25, 2019 4:56:53 PM GMT+08:00 | Refresh: NTP time sync- the registry of HOST W32tm server is NoSync after 1 click installation, it shall be NTP. | WIP18 | WIP19 BB_WIP9 | Manually reconfigure after installation | System | Installation;Blocker19 | N/A | Fixed | dasgupta_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP19SubCom | Brilliance 64 |  | WIP19 BB_WIP9 | July 31, 2019 2:09:31 PM GMT+08:00
CTUDT00625263 | Refresh | Closed | zhang_t | July 25, 2019 4:56:53 PM GMT+08:00 | Refresh: NTP time sync- the registry of HOST W32tm server is NoSync after 1 click installation, it shall be NTP. | WIP18 | WIP19 BB_WIP9 | Manually reconfigure after installation | System | Installation;Blocker19 | N/A | Fixed | dasgupta_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP19SubCom | Brilliance Big Bore |  | WIP19 BB_WIP9 | July 31, 2019 2:09:31 PM GMT+08:00
CTUDT00625263 | Refresh | Closed | zhang_t | July 25, 2019 4:56:53 PM GMT+08:00 | Refresh: NTP time sync- the registry of HOST W32tm server is NoSync after 1 click installation, it shall be NTP. | WIP18 | WIP19 BB_WIP9 | Manually reconfigure after installation | System | Installation;Blocker19 | N/A | Fixed | dasgupta_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP19SubCom | Ingenuity 64 |  | WIP19 BB_WIP9 | July 31, 2019 2:09:31 PM GMT+08:00
CTUDT00625263 | Refresh | Closed | zhang_t | July 25, 2019 4:56:53 PM GMT+08:00 | Refresh: NTP time sync- the registry of HOST W32tm server is NoSync after 1 click installation, it shall be NTP. | WIP18 | WIP19 BB_WIP9 | Manually reconfigure after installation | System | Installation;Blocker19 | N/A | Fixed | dasgupta_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP19SubCom | IQon1 |  | WIP19 BB_WIP9 | July 31, 2019 2:09:31 PM GMT+08:00
CTUDT00625261 | Refresh | Closed | Valeria_Kashtanov | July 25, 2019 4:07:07 PM GMT+08:00 | Refresh: Calibrations: Performance Test: X-Ray Tube Data is duplicated in Html Report | WIP18 | WIP19 |  | Subsystem | Performance Test;Verified1 | NA | Fixed | Kostya_Baklan | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP19SubCom | Ovation |  | WIP19 | August 4, 2019 4:00:47 PM GMT+08:00
CTUDT00625261 | Refresh | Closed | Valeria_Kashtanov | July 25, 2019 4:07:07 PM GMT+08:00 | Refresh: Calibrations: Performance Test: X-Ray Tube Data is duplicated in Html Report | WIP18 | WIP19 |  | Subsystem | Performance Test;Verified1 | NA | Fixed | Kostya_Baklan | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP19SubCom | Brilliance Big Bore |  | WIP19 | August 4, 2019 4:00:47 PM GMT+08:00
CTUDT00625261 | Refresh | Closed | Valeria_Kashtanov | July 25, 2019 4:07:07 PM GMT+08:00 | Refresh: Calibrations: Performance Test: X-Ray Tube Data is duplicated in Html Report | WIP18 | WIP19 |  | Subsystem | Performance Test;Verified1 | NA | Fixed | Kostya_Baklan | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP19SubCom | Ingenuity 64 |  | WIP19 | August 4, 2019 4:00:47 PM GMT+08:00
CTUDT00625261 | Refresh | Closed | Valeria_Kashtanov | July 25, 2019 4:07:07 PM GMT+08:00 | Refresh: Calibrations: Performance Test: X-Ray Tube Data is duplicated in Html Report | WIP18 | WIP19 |  | Subsystem | Performance Test;Verified1 | NA | Fixed | Kostya_Baklan | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP19SubCom | IQon1 |  | WIP19 | August 4, 2019 4:00:47 PM GMT+08:00
CTUDT00625259 | Refresh | Closed | nanjunda_s | July 25, 2019 3:14:23 AM GMT+08:00 | Getting an error message when performed Couch tests in the Visual diagnostics tool | WIP18 | WIP19 |  | Subsystem | Visual Diagostics | NA | Fixed | Sandeep_Kumar | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP19SubCom | Ovation |  | WIP19 | August 5, 2019 9:11:01 PM GMT+08:00
CTUDT00625259 | Refresh | Closed | nanjunda_s | July 25, 2019 3:14:23 AM GMT+08:00 | Getting an error message when performed Couch tests in the Visual diagnostics tool | WIP18 | WIP19 |  | Subsystem | Visual Diagostics | NA | Fixed | Sandeep_Kumar | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP19SubCom | Brilliance 64 |  | WIP19 | August 5, 2019 9:11:01 PM GMT+08:00
CTUDT00625259 | Refresh | Closed | nanjunda_s | July 25, 2019 3:14:23 AM GMT+08:00 | Getting an error message when performed Couch tests in the Visual diagnostics tool | WIP18 | WIP19 |  | Subsystem | Visual Diagostics | NA | Fixed | Sandeep_Kumar | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP19SubCom | IQon1 |  | WIP19 | August 5, 2019 9:11:01 PM GMT+08:00
CTUDT00625256 | Refresh | Closed | nanjunda_s | July 25, 2019 2:46:09 AM GMT+08:00 | Olevel user can able to access US-DOD config tool | WIP18 | WIP19 BB_WIP9 |  | Subsystem | Olevel | NA | Fixed | tejaswini_nidamanuri | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP19SubCom | Ovation |  | WIP19 BB_WIP9 | July 30, 2019 8:34:52 PM GMT+08:00
CTUDT00625256 | Refresh | Closed | nanjunda_s | July 25, 2019 2:46:09 AM GMT+08:00 | Olevel user can able to access US-DOD config tool | WIP18 | WIP19 BB_WIP9 |  | Subsystem | Olevel | NA | Fixed | tejaswini_nidamanuri | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP19SubCom | Brilliance 64 |  | WIP19 BB_WIP9 | July 30, 2019 8:34:52 PM GMT+08:00
CTUDT00625256 | Refresh | Closed | nanjunda_s | July 25, 2019 2:46:09 AM GMT+08:00 | Olevel user can able to access US-DOD config tool | WIP18 | WIP19 BB_WIP9 |  | Subsystem | Olevel | NA | Fixed | tejaswini_nidamanuri | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP19SubCom | Brilliance Big Bore |  | WIP19 BB_WIP9 | July 30, 2019 8:34:52 PM GMT+08:00
CTUDT00625256 | Refresh | Closed | nanjunda_s | July 25, 2019 2:46:09 AM GMT+08:00 | Olevel user can able to access US-DOD config tool | WIP18 | WIP19 BB_WIP9 |  | Subsystem | Olevel | NA | Fixed | tejaswini_nidamanuri | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP19SubCom | Ingenuity 64 |  | WIP19 BB_WIP9 | July 30, 2019 8:34:52 PM GMT+08:00
CTUDT00625256 | Refresh | Closed | nanjunda_s | July 25, 2019 2:46:09 AM GMT+08:00 | Olevel user can able to access US-DOD config tool | WIP18 | WIP19 BB_WIP9 |  | Subsystem | Olevel | NA | Fixed | tejaswini_nidamanuri | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP19SubCom | IQon1 |  | WIP19 BB_WIP9 | July 30, 2019 8:34:52 PM GMT+08:00
CTUDT00625251 | Refresh | Closed | bommiReddy_r | July 24, 2019 8:14:58 PM GMT+08:00 | Refresh: Bolus Tracking: "Please wait for the waves to be displayed” message was displayed after pressing on the “Repeat Tracker” option | WIP18 | WIP20 |  | Subsystem | Bolus Tracking;TAVI | NA | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | NO | Closed | WIP21SubCom | Ovation |  | WIP20 | September 5, 2019 7:57:44 PM GMT+08:00
CTUDT00625251 | Refresh | Closed | bommiReddy_r | July 24, 2019 8:14:58 PM GMT+08:00 | Refresh: Bolus Tracking: "Please wait for the waves to be displayed” message was displayed after pressing on the “Repeat Tracker” option | WIP18 | WIP20 |  | Subsystem | Bolus Tracking;TAVI | NA | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | NO | Closed | WIP21SubCom | Brilliance 64 |  | WIP20 | September 5, 2019 7:57:44 PM GMT+08:00
CTUDT00625251 | Refresh | Closed | bommiReddy_r | July 24, 2019 8:14:58 PM GMT+08:00 | Refresh: Bolus Tracking: "Please wait for the waves to be displayed” message was displayed after pressing on the “Repeat Tracker” option | WIP18 | WIP20 |  | Subsystem | Bolus Tracking;TAVI | NA | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | NO | Closed | WIP21SubCom | Ingenuity 64 |  | WIP20 | September 5, 2019 7:57:44 PM GMT+08:00
CTUDT00625251 | Refresh | Closed | bommiReddy_r | July 24, 2019 8:14:58 PM GMT+08:00 | Refresh: Bolus Tracking: "Please wait for the waves to be displayed” message was displayed after pressing on the “Repeat Tracker” option | WIP18 | WIP20 |  | Subsystem | Bolus Tracking;TAVI | NA | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | NO | Closed | WIP21SubCom | IQon1 |  | WIP20 | September 5, 2019 7:57:44 PM GMT+08:00
CTUDT00625249 | Refresh | Closed | hatti_c | July 24, 2019 6:22:21 PM GMT+08:00 | CT-NM.SRS.Service.1707 and NM.SRS.Service.2747 has to be untagged from Refresh as URL tool is not applicable to WIN10-Refresh | WIP17 | WIP19 BB_WIP9 |  | Subsystem | SRS | CT-NM.SRS.Service.1707 and NM.SRS.Service.2747 | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP19SegReq | Ovation |  | WIP19 BB_WIP9 | September 5, 2019 4:29:25 PM GMT+08:00
CTUDT00625249 | Refresh | Closed | hatti_c | July 24, 2019 6:22:21 PM GMT+08:00 | CT-NM.SRS.Service.1707 and NM.SRS.Service.2747 has to be untagged from Refresh as URL tool is not applicable to WIN10-Refresh | WIP17 | WIP19 BB_WIP9 |  | Subsystem | SRS | CT-NM.SRS.Service.1707 and NM.SRS.Service.2747 | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP19SegReq | Brilliance 64 |  | WIP19 BB_WIP9 | September 5, 2019 4:29:25 PM GMT+08:00
CTUDT00625249 | Refresh | Closed | hatti_c | July 24, 2019 6:22:21 PM GMT+08:00 | CT-NM.SRS.Service.1707 and NM.SRS.Service.2747 has to be untagged from Refresh as URL tool is not applicable to WIN10-Refresh | WIP17 | WIP19 BB_WIP9 |  | Subsystem | SRS | CT-NM.SRS.Service.1707 and NM.SRS.Service.2747 | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP19SegReq | Brilliance Big Bore |  | WIP19 BB_WIP9 | September 5, 2019 4:29:25 PM GMT+08:00
CTUDT00625249 | Refresh | Closed | hatti_c | July 24, 2019 6:22:21 PM GMT+08:00 | CT-NM.SRS.Service.1707 and NM.SRS.Service.2747 has to be untagged from Refresh as URL tool is not applicable to WIN10-Refresh | WIP17 | WIP19 BB_WIP9 |  | Subsystem | SRS | CT-NM.SRS.Service.1707 and NM.SRS.Service.2747 | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP19SegReq | Ingenuity 64 |  | WIP19 BB_WIP9 | September 5, 2019 4:29:25 PM GMT+08:00
CTUDT00625249 | Refresh | Closed | hatti_c | July 24, 2019 6:22:21 PM GMT+08:00 | CT-NM.SRS.Service.1707 and NM.SRS.Service.2747 has to be untagged from Refresh as URL tool is not applicable to WIN10-Refresh | WIP17 | WIP19 BB_WIP9 |  | Subsystem | SRS | CT-NM.SRS.Service.1707 and NM.SRS.Service.2747 | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP19SegReq | IQon1 |  | WIP19 BB_WIP9 | September 5, 2019 4:29:25 PM GMT+08:00
CTUDT00625247 | Refresh | Closed | zhu_je | July 24, 2019 6:18:04 PM GMT+08:00 | [Refresh]When press ? in the help page of ACR Test by harmonized phantom, the help page was system phantom. | BB_WIP8 |  |  | Subsystem | Calibrations | NaN | Works As Specified / Not A Defect | yatziv_l | 0 | 2 |  | DMS | NO | Closed | WIP19SubCom | Brilliance Big Bore
CTUDT00625242 | Refresh | Closed | finestine_m | July 23, 2019 9:48:05 PM GMT+08:00 | Refresh: CT User Missing Translation in Instructions Window for Constancy Test | WIP18 | WIP19 |  | Subsystem | Acceptance/Constancy; Verified1 | NaN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP19SubCom | Ovation |  | WIP19 | August 8, 2019 5:41:33 PM GMT+08:00
CTUDT00625242 | Refresh | Closed | finestine_m | July 23, 2019 9:48:05 PM GMT+08:00 | Refresh: CT User Missing Translation in Instructions Window for Constancy Test | WIP18 | WIP19 |  | Subsystem | Acceptance/Constancy; Verified1 | NaN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP19SubCom | Brilliance Big Bore |  | WIP19 | August 8, 2019 5:41:33 PM GMT+08:00
CTUDT00625242 | Refresh | Closed | finestine_m | July 23, 2019 9:48:05 PM GMT+08:00 | Refresh: CT User Missing Translation in Instructions Window for Constancy Test | WIP18 | WIP19 |  | Subsystem | Acceptance/Constancy; Verified1 | NaN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP19SubCom | Ingenuity 64 |  | WIP19 | August 8, 2019 5:41:33 PM GMT+08:00
CTUDT00625242 | Refresh | Closed | finestine_m | July 23, 2019 9:48:05 PM GMT+08:00 | Refresh: CT User Missing Translation in Instructions Window for Constancy Test | WIP18 | WIP19 |  | Subsystem | Acceptance/Constancy; Verified1 | NaN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP19SubCom | IQon1 |  | WIP19 | August 8, 2019 5:41:33 PM GMT+08:00
CTUDT00625240 | Refresh | Closed | cohen_y | July 23, 2019 9:22:58 PM GMT+08:00 | Refresh: IQ Check: Window Center of Body scan with Harmonized Phantom is not as defined in Spec | BB_WIP7 | WIP19 | N/A | Subsystem | IQ Check | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | BB_WIP9_SegReq | Brilliance Big Bore |  | WIP19 | July 31, 2019 10:05:20 PM GMT+08:00
CTUDT00625237 | Refresh | Closed | alex_nihaev | July 23, 2019 9:01:20 PM GMT+08:00 | Refresh: Calibrations: HCOR: HCOR Slice: H-Phantom: Wrong ROI radius for Body and Head cases | WIP18 | WIP20 |  | Subsystem | Specifications;Blocker20 | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP19SubCom | Ingenuity 64 |  | WIP20 | August 25, 2019 7:18:52 PM GMT+08:00
CTUDT00625235 | Refresh | Closed | Noy_Seelig | July 23, 2019 8:24:19 PM GMT+08:00 | Refresh: User Access: White Spot test appears in Spec but not in SW | WIP18 | N/A |  | Subsystem | White Spot | N/A | Fixed | braunstein_d | 0 | 2 | Design Documentation | Support-Physics | NO | Closed | BB_WIP9_SubCom | Ingenuity 64 |  | N/A | September 19, 2019 2:49:26 PM GMT+08:00
CTUDT00625233 | Refresh | Closed | cohen_y | July 23, 2019 7:55:32 PM GMT+08:00 | Refresh: Bad Detectors Calibration: Bias limits in Production user are incorrect in DHF302480 Rev 03 | BB_WIP7 | WIP19 |  | Subsystem | Bad Detectors | DHF3020480 (r4) | Fixed | bikman | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP19SubCom | Brilliance Big Bore |  | WIP19 | July 31, 2019 5:33:03 PM GMT+08:00
CTUDT00625231 | Refresh | Closed | ravindra_a | July 23, 2019 7:29:55 PM GMT+08:00 | Big Bore Refresh -Average Scan size is not update for CCT scan after each shots | WIP18 | WIP19 | No | Subsystem | SSDE | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | YES | Closed | WIP19SubCom | Ovation |  | WIP19 | July 30, 2019 1:21:56 PM GMT+08:00
CTUDT00625231 | Refresh | Closed | ravindra_a | July 23, 2019 7:29:55 PM GMT+08:00 | Big Bore Refresh -Average Scan size is not update for CCT scan after each shots | WIP18 | WIP19 | No | Subsystem | SSDE | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | YES | Closed | WIP19SubCom | Brilliance 64 |  | WIP19 | July 30, 2019 1:21:56 PM GMT+08:00
CTUDT00625231 | Refresh | Closed | ravindra_a | July 23, 2019 7:29:55 PM GMT+08:00 | Big Bore Refresh -Average Scan size is not update for CCT scan after each shots | WIP18 | WIP19 | No | Subsystem | SSDE | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | YES | Closed | WIP19SubCom | Ingenuity 64 |  | WIP19 | July 30, 2019 1:21:56 PM GMT+08:00
CTUDT00625231 | Refresh | Closed | ravindra_a | July 23, 2019 7:29:55 PM GMT+08:00 | Big Bore Refresh -Average Scan size is not update for CCT scan after each shots | WIP18 | WIP19 | No | Subsystem | SSDE | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | YES | Closed | WIP19SubCom | IQon1 |  | WIP19 | July 30, 2019 1:21:56 PM GMT+08:00
CTUDT00625228 | Refresh | Closed | wang_l | July 23, 2019 6:31:58 PM GMT+08:00 | Refresh: System stuck during Air calibration and phantom calibration | WIP18 | WIP19 |  | Subsystem | Calibrations | n/a | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP19SegReq | Ovation |  | WIP19 | August 6, 2019 11:31:44 PM GMT+08:00
CTUDT00625228 | Refresh | Closed | wang_l | July 23, 2019 6:31:58 PM GMT+08:00 | Refresh: System stuck during Air calibration and phantom calibration | WIP18 | WIP19 |  | Subsystem | Calibrations | n/a | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP19SegReq | Brilliance 64 |  | WIP19 | August 6, 2019 11:31:44 PM GMT+08:00
CTUDT00625228 | Refresh | Closed | wang_l | July 23, 2019 6:31:58 PM GMT+08:00 | Refresh: System stuck during Air calibration and phantom calibration | WIP18 | WIP19 |  | Subsystem | Calibrations | n/a | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP19SegReq | Ingenuity 64 |  | WIP19 | August 6, 2019 11:31:44 PM GMT+08:00
CTUDT00625228 | Refresh | Closed | wang_l | July 23, 2019 6:31:58 PM GMT+08:00 | Refresh: System stuck during Air calibration and phantom calibration | WIP18 | WIP19 |  | Subsystem | Calibrations | n/a | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP19SegReq | Ingenuity BigBore |  | WIP19 | August 6, 2019 11:31:44 PM GMT+08:00
CTUDT00625228 | Refresh | Closed | wang_l | July 23, 2019 6:31:58 PM GMT+08:00 | Refresh: System stuck during Air calibration and phantom calibration | WIP18 | WIP19 |  | Subsystem | Calibrations | n/a | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP19SegReq | IQon1 |  | WIP19 | August 6, 2019 11:31:44 PM GMT+08:00
CTUDT00625226 | Refresh | Closed | Noy_Seelig | July 23, 2019 6:18:46 PM GMT+08:00 | Refresh: Calibrations: In Instructions of IQ check, Therapy Top Instructions were not translated to other languages but English | WIP18 |  |  | Subsystem | IQ Check |  | Duplicate | frider_o | 0 |  |  | DMS | NO | Closed |  | Ingenuity 64
CTUDT00625223 | Refresh | Closed | nanjunda_s | July 23, 2019 5:36:47 PM GMT+08:00 | Patient Birth time is not available to query in worklist. | WIP17 | DHF303211_Rev14_Host_SSRS_for_Refresh DHF316705 Rev.13-Host_SSRS_for_BB-Refresh |  | Subsystem | DICOM | Console-SRS.15999 | Fixed | hatti_c | 0 | 3 | Requirements Documentation | Console4 | NO | Closed | SubCom | Ovation |  | DHF303211_Rev14_Host_SSRS_for_Refresh DHF316705 Rev.13-Host_SSRS_for_BB-Refresh | August 5, 2019 2:30:11 PM GMT+08:00
CTUDT00625223 | Refresh | Closed | nanjunda_s | July 23, 2019 5:36:47 PM GMT+08:00 | Patient Birth time is not available to query in worklist. | WIP17 | DHF303211_Rev14_Host_SSRS_for_Refresh DHF316705 Rev.13-Host_SSRS_for_BB-Refresh |  | Subsystem | DICOM | Console-SRS.15999 | Fixed | hatti_c | 0 | 3 | Requirements Documentation | Console4 | NO | Closed | SubCom | Brilliance 64 |  | DHF303211_Rev14_Host_SSRS_for_Refresh DHF316705 Rev.13-Host_SSRS_for_BB-Refresh | August 5, 2019 2:30:11 PM GMT+08:00
CTUDT00625223 | Refresh | Closed | nanjunda_s | July 23, 2019 5:36:47 PM GMT+08:00 | Patient Birth time is not available to query in worklist. | WIP17 | DHF303211_Rev14_Host_SSRS_for_Refresh DHF316705 Rev.13-Host_SSRS_for_BB-Refresh |  | Subsystem | DICOM | Console-SRS.15999 | Fixed | hatti_c | 0 | 3 | Requirements Documentation | Console4 | NO | Closed | SubCom | Brilliance Big Bore |  | DHF303211_Rev14_Host_SSRS_for_Refresh DHF316705 Rev.13-Host_SSRS_for_BB-Refresh | August 5, 2019 2:30:11 PM GMT+08:00
CTUDT00625223 | Refresh | Closed | nanjunda_s | July 23, 2019 5:36:47 PM GMT+08:00 | Patient Birth time is not available to query in worklist. | WIP17 | DHF303211_Rev14_Host_SSRS_for_Refresh DHF316705 Rev.13-Host_SSRS_for_BB-Refresh |  | Subsystem | DICOM | Console-SRS.15999 | Fixed | hatti_c | 0 | 3 | Requirements Documentation | Console4 | NO | Closed | SubCom | Ingenuity 64 |  | DHF303211_Rev14_Host_SSRS_for_Refresh DHF316705 Rev.13-Host_SSRS_for_BB-Refresh | August 5, 2019 2:30:11 PM GMT+08:00
CTUDT00625223 | Refresh | Closed | nanjunda_s | July 23, 2019 5:36:47 PM GMT+08:00 | Patient Birth time is not available to query in worklist. | WIP17 | DHF303211_Rev14_Host_SSRS_for_Refresh DHF316705 Rev.13-Host_SSRS_for_BB-Refresh |  | Subsystem | DICOM | Console-SRS.15999 | Fixed | hatti_c | 0 | 3 | Requirements Documentation | Console4 | NO | Closed | SubCom | IQon1 |  | DHF303211_Rev14_Host_SSRS_for_Refresh DHF316705 Rev.13-Host_SSRS_for_BB-Refresh | August 5, 2019 2:30:11 PM GMT+08:00
CTUDT00625208 | Refresh | Closed | bhawna_g | July 23, 2019 4:41:41 PM GMT+08:00 | [Refresh] “Export Failed” error Pop up when tried to export a renamed language set | 4.8 | WIP19 |  | Subsystem |  | NA | Fixed | 320057847 | 0 | 2 | Console Software | Console4 | NO | Closed | WIP19SegReq | Ovation |  | WIP19 | July 30, 2019 7:13:26 PM GMT+08:00
CTUDT00625208 | Refresh | Closed | bhawna_g | July 23, 2019 4:41:41 PM GMT+08:00 | [Refresh] “Export Failed” error Pop up when tried to export a renamed language set | 4.8 | WIP19 |  | Subsystem |  | NA | Fixed | 320057847 | 0 | 2 | Console Software | Console4 | NO | Closed | WIP19SegReq | Brilliance 64 |  | WIP19 | July 30, 2019 7:13:26 PM GMT+08:00
CTUDT00625208 | Refresh | Closed | bhawna_g | July 23, 2019 4:41:41 PM GMT+08:00 | [Refresh] “Export Failed” error Pop up when tried to export a renamed language set | 4.8 | WIP19 |  | Subsystem |  | NA | Fixed | 320057847 | 0 | 2 | Console Software | Console4 | NO | Closed | WIP19SegReq | Ingenuity 64 |  | WIP19 | July 30, 2019 7:13:26 PM GMT+08:00
CTUDT00625208 | Refresh | Closed | bhawna_g | July 23, 2019 4:41:41 PM GMT+08:00 | [Refresh] “Export Failed” error Pop up when tried to export a renamed language set | 4.8 | WIP19 |  | Subsystem |  | NA | Fixed | 320057847 | 0 | 2 | Console Software | Console4 | NO | Closed | WIP19SegReq | IQon1 |  | WIP19 | July 30, 2019 7:13:26 PM GMT+08:00
CTUDT00625207 | Refresh | Closed | bhawna_g | July 23, 2019 4:37:28 PM GMT+08:00 | [Refresh] Offline ECG Wave is not shown in acquisition scene (but shown in view scene) after click final recon for Coronary CTA EC | 4.8 | BB_WIP9 |  | System | OfflineWave | NA | Fixed | datta_i | 0 | 3 | Console Software | Console4 | YES | Closed | BB_WIP9_SegReq | Brilliance Big Bore |  | BB_WIP9 | August 1, 2019 6:30:20 PM GMT+08:00
CTUDT00625197 | Refresh | Closed | lin_s | July 23, 2019 10:12:48 AM GMT+08:00 | Brilliance 64 is EOP and upgrade does not sale to China marketing. So Regstand.751 is not applicable for Brilliance 64. | WIP18 | WIP19 |  | Subsystem | Requirements | RegStand.751 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP19SubCom | Brilliance 64 |  | WIP19 | August 27, 2019 6:03:26 PM GMT+08:00
CTUDT00625195 | Refresh | Closed | lei_yang | July 22, 2019 9:45:58 PM GMT+08:00 | Some values in the "Image Parameters" for Locator are not the same as the actual value. | BB_WIP6 |  |  | Subsystem | Bolus | NA | Works As Specified / Not A Defect | chokkalingam_r | 0 |  |  | Console4 | NO | Closed | WAS | Brilliance Big Bore
CTUDT00625193 | Refresh | Closed | pandey_a | July 22, 2019 8:17:28 PM GMT+08:00 | BBRefresh: Wait for Wave Dialog pops up when doing non-gated 4DCT pulmo scan | BB_WIP8 | BB_WIP9 |  | Subsystem | Respiratory | N/A | Fixed | pandey_a | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP9_SubCom | Brilliance Big Bore |  | BB_WIP9 | July 30, 2019 4:35:35 PM GMT+08:00
CTUDT00625184 | Refresh | Closed | cohen_y | July 22, 2019 5:32:33 PM GMT+08:00 | Refresh: Acceptance/Contancy: Angular Sampling in Resolution 2 case is Normal and not High | BB_WIP7 | WIP19 |  | Subsystem | Acceptance/Constancy; Verified1 | n/a | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP19SubCom | Brilliance Big Bore |  | WIP19 | August 11, 2019 8:14:13 PM GMT+08:00
CTUDT00625184 | Refresh | Closed | cohen_y | July 22, 2019 5:32:33 PM GMT+08:00 | Refresh: Acceptance/Contancy: Angular Sampling in Resolution 2 case is Normal and not High | BB_WIP7 | WIP19 |  | Subsystem | Acceptance/Constancy; Verified1 | n/a | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP19SubCom | Ingenuity 64 |  | WIP19 | August 11, 2019 8:14:13 PM GMT+08:00
CTUDT00625184 | Refresh | Closed | cohen_y | July 22, 2019 5:32:33 PM GMT+08:00 | Refresh: Acceptance/Contancy: Angular Sampling in Resolution 2 case is Normal and not High | BB_WIP7 | WIP19 |  | Subsystem | Acceptance/Constancy; Verified1 | n/a | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP19SubCom | IQon1 |  | WIP19 | August 11, 2019 8:14:13 PM GMT+08:00
CTUDT00625182 | Refresh | Closed | Noy_Seelig | July 22, 2019 4:23:38 PM GMT+08:00 | Refresh: Calibration: Hcor Basic Infant Test: Wrong position of Body section taken from Adult instead of Child | WIP17 | N/A | N/A | Subsystem | Test Design | IQon | Fixed | yosfan_s | 0 | 2 | Test Design | DMS | NO | Closed | WIP19SubCom | IQon1 |  | N/A | August 7, 2019 2:23:13 PM GMT+08:00
CTUDT00625180 | Refresh | Closed | ding_f | July 22, 2019 3:04:17 PM GMT+08:00 | "Tube Current, Scan Time, mAs Test" failed with half the nominal Shot Time and mAs values for RT=2.0sec stage. | WIP8 | WIP19 |  | Subsystem | Calibrations;blocker19 | NaN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP19SubCom | Brilliance Big Bore |  | WIP19 | August 7, 2019 9:05:24 PM GMT+08:00
CTUDT00625178 | Refresh | Closed | ding_f | July 22, 2019 1:54:16 PM GMT+08:00 | “IMR Noise Reduction Check” shall be included in system performance test requirement | WIP8 | BB_WIP9 |  | Subsystem | Requirements | CT-NM.SRS.IQ.2791 | Fixed | ye_h | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | BB_WIP9_SubCom | Brilliance Big Bore |  | BB_WIP9 | September 3, 2019 2:29:01 PM GMT+08:00
CTUDT00625175 | Refresh | Closed | lei_yang | July 22, 2019 9:25:14 AM GMT+08:00 | "Display" option cannot be set to "3" with CCT Biopsy EC. | WIP16 |  |  | Subsystem | CCT;Blocker | NA | Works As Specified / Not A Defect | lin_s | 0 | 2 |  | Support-System | NO | Closed | BB_WIP9_SubCom | Ingenuity 64
CTUDT00625171 | Refresh | Closed | Prakash_Chand_Yadav | July 19, 2019 5:17:55 PM GMT+08:00 | [Refresh] Change couch type tool is allowing to select more then one couch type. | WIP18 | WIP19 |  | Subsystem | Service Tools | None | Fixed | Prakash_Chand_Yadav | 0 | 4 | Console Software | Service Tools | NO | Closed | WIP19SubCom | IQon1 |  | WIP19 | July 23, 2019 2:00:41 PM GMT+08:00
CTUDT00625170 | Refresh | Closed | bhawna_g | July 19, 2019 2:40:26 PM GMT+08:00 | Refresh Pre-Val: Wrong representation of the scan on the resp waveform (Resp Single Phase; User Defined) | 4.1.10 | WIP19 |  | System | Respiratory | N/A | Fixed | pandey_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP19SubCom | Brilliance 64 |  | WIP19 | July 25, 2019 8:06:11 PM GMT+08:00
CTUDT00625170 | Refresh | Closed | bhawna_g | July 19, 2019 2:40:26 PM GMT+08:00 | Refresh Pre-Val: Wrong representation of the scan on the resp waveform (Resp Single Phase; User Defined) | 4.1.10 | WIP19 |  | System | Respiratory | N/A | Fixed | pandey_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP19SubCom | Ingenuity 64 |  | WIP19 | July 25, 2019 8:06:11 PM GMT+08:00
CTUDT00625169 | Refresh | Closed | deshbhratar_m | July 19, 2019 2:02:46 PM GMT+08:00 | [Refresh] There is system error if click repeat or next after finish perfusion scan | 4.8 | BB_WIP9 |  | Subsystem | System | NA | Fixed | deshbhratar_m | 0 | 3 | Console Software | Console4 | YES | Closed | BB_WIP9_SegReq | Brilliance Big Bore |  | BB_WIP9 | July 30, 2019 7:45:35 PM GMT+08:00
CTUDT00625167 | Refresh | Closed | v_sumanth | July 19, 2019 1:47:21 PM GMT+08:00 | Service tools show different couch type than the right ones via one-click installation. | 4.8 | BB_WIP9 |  | Subsystem | Service Tools | None | Fixed | Prakash_Chand_Yadav | 0 | 3 | Console Software | Service Tools | NO | Closed | BB_WIP9_SubCom | Brilliance Big Bore |  | BB_WIP9 | July 23, 2019 5:21:56 PM GMT+08:00
CTUDT00625165 | Refresh | Closed | lei_yang | July 19, 2019 9:14:35 AM GMT+08:00 | Dose modulation is turned off for the remainder of the scan after the arrhythmic cycle acquisition proceeded. | WIP16 | WIP20 DHF303211 Rev15 DHF298554 Rev08 |  | Subsystem | DoseRight;Blocker | N/A | Fixed | bandil_j | 0 | 2 | Console Software | Support-System | YES | Closed | WIP20SubCom | Brilliance 64 |  | WIP20 DHF303211 Rev15 DHF298554 Rev08 | November 5, 2019 6:59:21 PM GMT+08:00
CTUDT00625165 | Refresh | Closed | lei_yang | July 19, 2019 9:14:35 AM GMT+08:00 | Dose modulation is turned off for the remainder of the scan after the arrhythmic cycle acquisition proceeded. | WIP16 | WIP20 DHF303211 Rev15 DHF298554 Rev08 |  | Subsystem | DoseRight;Blocker | N/A | Fixed | bandil_j | 0 | 2 | Console Software | Support-System | YES | Closed | WIP20SubCom | Ingenuity 64 |  | WIP20 DHF303211 Rev15 DHF298554 Rev08 | November 5, 2019 6:59:21 PM GMT+08:00
CTUDT00625163 | Refresh | Closed | Sateesh_S | July 18, 2019 11:54:04 PM GMT+08:00 | Big Bore Refresh -Average Scan size is not update for CCT scan after each shots | BB_WIP7 | BB_WIP9 | No | Subsystem | CCT | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | YES | Closed | BB_WIP9_SubCom | Brilliance Big Bore |  | BB_WIP9 | July 25, 2019 6:10:17 PM GMT+08:00
CTUDT00625161 | Refresh | Closed | wang_l | July 18, 2019 8:47:56 PM GMT+08:00 | Refresh: SFTP connection failed by philips_service user which not match with CT-NM.SRS.Service.179 | WIP18 | WIP19 BB_WIP9 |  | Subsystem | Service Tools | Service.179 | Fixed | ye_h | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP19SegReq | Brilliance 64 |  | WIP19 BB_WIP9 | September 3, 2019 3:01:04 PM GMT+08:00
CTUDT00625161 | Refresh | Closed | wang_l | July 18, 2019 8:47:56 PM GMT+08:00 | Refresh: SFTP connection failed by philips_service user which not match with CT-NM.SRS.Service.179 | WIP18 | WIP19 BB_WIP9 |  | Subsystem | Service Tools | Service.179 | Fixed | ye_h | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP19SegReq | Brilliance Big Bore |  | WIP19 BB_WIP9 | September 3, 2019 3:01:04 PM GMT+08:00
CTUDT00625161 | Refresh | Closed | wang_l | July 18, 2019 8:47:56 PM GMT+08:00 | Refresh: SFTP connection failed by philips_service user which not match with CT-NM.SRS.Service.179 | WIP18 | WIP19 BB_WIP9 |  | Subsystem | Service Tools | Service.179 | Fixed | ye_h | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP19SegReq | Ingenuity 64 |  | WIP19 BB_WIP9 | September 3, 2019 3:01:04 PM GMT+08:00
CTUDT00625158 | Refresh | Closed | wang_l | July 18, 2019 7:35:05 PM GMT+08:00 | Refresh BB: Aplane Reader is displayed even if without IST certificate authorization | BB_WIP8 | BB_WIP9 |  | Subsystem | Service Tools;Blocker | CT-NM.SRS.Service.2992 | Fixed | ye_h | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | BB_WIP9_SubCom | Brilliance Big Bore |  | BB_WIP9 | September 3, 2019 2:30:58 PM GMT+08:00
CTUDT00625156 | Refresh | Closed | wang_l | July 18, 2019 7:04:06 PM GMT+08:00 | Refresh BB: Requirement CT-NM.SRS.Service.3045 does not inherit DMS but contents mentioned DMS | BB_WIP8 | BB_WIP9 |  | Subsystem | Requirements;Blocker | CT-NM.SRS.Service.3045 | Fixed | ye_h | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | BB_WIP9_SubCom | Brilliance Big Bore |  | BB_WIP9 | September 3, 2019 2:32:53 PM GMT+08:00
CTUDT00625154 | Refresh | Closed | wang_l | July 18, 2019 6:40:35 PM GMT+08:00 | Refresh: Voltage of the Slice Width, Z, and filter of the Aplane Collimator also displayed but CT-NM.SRS.Service.3063 not covered | WIP18 | WIP19 |  | Subsystem | Requirements | Service.3063 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | SRS DHF298554 Rev07 | Brilliance 64 |  | WIP19 | August 27, 2019 5:37:13 PM GMT+08:00
CTUDT00625154 | Refresh | Closed | wang_l | July 18, 2019 6:40:35 PM GMT+08:00 | Refresh: Voltage of the Slice Width, Z, and filter of the Aplane Collimator also displayed but CT-NM.SRS.Service.3063 not covered | WIP18 | WIP19 |  | Subsystem | Requirements | Service.3063 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | SRS DHF298554 Rev07 | Ingenuity 64 |  | WIP19 | August 27, 2019 5:37:13 PM GMT+08:00
CTUDT00625152 | Refresh | Closed | wang_l | July 18, 2019 6:22:02 PM GMT+08:00 | Refresh: Requirement CT-NM.SRS.Service.1113 was not fulfilled because olevel user cannot access to Device Management tool | WIP18 | WIP19 |  | Subsystem | Requirement | service.1113 | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP20SegReq | Ovation |  | WIP19 | September 5, 2019 4:28:06 PM GMT+08:00
CTUDT00625152 | Refresh | Closed | wang_l | July 18, 2019 6:22:02 PM GMT+08:00 | Refresh: Requirement CT-NM.SRS.Service.1113 was not fulfilled because olevel user cannot access to Device Management tool | WIP18 | WIP19 |  | Subsystem | Requirement | service.1113 | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP20SegReq | Brilliance 64 |  | WIP19 | September 5, 2019 4:28:06 PM GMT+08:00
CTUDT00625152 | Refresh | Closed | wang_l | July 18, 2019 6:22:02 PM GMT+08:00 | Refresh: Requirement CT-NM.SRS.Service.1113 was not fulfilled because olevel user cannot access to Device Management tool | WIP18 | WIP19 |  | Subsystem | Requirement | service.1113 | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP20SegReq | Ingenuity 64 |  | WIP19 | September 5, 2019 4:28:06 PM GMT+08:00
CTUDT00625152 | Refresh | Closed | wang_l | July 18, 2019 6:22:02 PM GMT+08:00 | Refresh: Requirement CT-NM.SRS.Service.1113 was not fulfilled because olevel user cannot access to Device Management tool | WIP18 | WIP19 |  | Subsystem | Requirement | service.1113 | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP20SegReq | IQon1 |  | WIP19 | September 5, 2019 4:28:06 PM GMT+08:00
CTUDT00625150 | Refresh | Closed | 320053770 | July 18, 2019 5:43:56 PM GMT+08:00 | WIP 18: ICT IMR: in service tools - RDP remote viewer, when connecting to CIRS S1 – an error pops “Failed to instantiate…”. | WIP18 |  |  | Subsystem | CIRS |  | Duplicate | richards_b | 0 |  |  | CIRS | NO | Closed |  | Ovation
CTUDT00625148 | Refresh | Closed | 320053770 | July 18, 2019 5:29:34 PM GMT+08:00 | Refresh WIP 18: while inside an RDP remote control window, ALT–Tab scrolling doesn’t work. | WIP18 | WIP19 BB_WIP9 |  | Subsystem | RDP;Blocker19 | N/A | Fixed | chakraborty_d | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP19SubCom | Ovation |  | WIP19 BB_WIP9 | August 2, 2019 2:16:28 PM GMT+08:00
CTUDT00625148 | Refresh | Closed | 320053770 | July 18, 2019 5:29:34 PM GMT+08:00 | Refresh WIP 18: while inside an RDP remote control window, ALT–Tab scrolling doesn’t work. | WIP18 | WIP19 BB_WIP9 |  | Subsystem | RDP;Blocker19 | N/A | Fixed | chakraborty_d | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP19SubCom | Brilliance 64 |  | WIP19 BB_WIP9 | August 2, 2019 2:16:28 PM GMT+08:00
CTUDT00625148 | Refresh | Closed | 320053770 | July 18, 2019 5:29:34 PM GMT+08:00 | Refresh WIP 18: while inside an RDP remote control window, ALT–Tab scrolling doesn’t work. | WIP18 | WIP19 BB_WIP9 |  | Subsystem | RDP;Blocker19 | N/A | Fixed | chakraborty_d | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP19SubCom | Brilliance Big Bore |  | WIP19 BB_WIP9 | August 2, 2019 2:16:28 PM GMT+08:00
CTUDT00625148 | Refresh | Closed | 320053770 | July 18, 2019 5:29:34 PM GMT+08:00 | Refresh WIP 18: while inside an RDP remote control window, ALT–Tab scrolling doesn’t work. | WIP18 | WIP19 BB_WIP9 |  | Subsystem | RDP;Blocker19 | N/A | Fixed | chakraborty_d | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP19SubCom | Ingenuity 64 |  | WIP19 BB_WIP9 | August 2, 2019 2:16:28 PM GMT+08:00
CTUDT00625148 | Refresh | Closed | 320053770 | July 18, 2019 5:29:34 PM GMT+08:00 | Refresh WIP 18: while inside an RDP remote control window, ALT–Tab scrolling doesn’t work. | WIP18 | WIP19 BB_WIP9 |  | Subsystem | RDP;Blocker19 | N/A | Fixed | chakraborty_d | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP19SubCom | IQon1 |  | WIP19 BB_WIP9 | August 2, 2019 2:16:28 PM GMT+08:00
CTUDT00625145 | Refresh | Closed | wang_l | July 18, 2019 3:42:56 PM GMT+08:00 | Refresh:Calibration Tube current,Scan time,mAs test failed for BR64 MRC600 | WIP18 |  |  | Subsystem | Calibrations - Tube Current | n/a | Works As Specified / Not A Defect | braunstein_d | 0 |  |  | Support-Physics | NO | Closed |  | Brilliance 64
CTUDT00625137 | Refresh | Closed | manjunath_mk | July 18, 2019 12:16:16 PM GMT+08:00 | 'Edit auto voice' option is available for second acquisition in scan ruler for Dual energy Exam cards. | BB_WIP7 | BB_WIP9 |  | Subsystem | Plan | NA | Fixed | jose_s | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP9_SubCom | Ovation |  | BB_WIP9 | August 1, 2019 6:34:33 PM GMT+08:00
CTUDT00625137 | Refresh | Closed | manjunath_mk | July 18, 2019 12:16:16 PM GMT+08:00 | 'Edit auto voice' option is available for second acquisition in scan ruler for Dual energy Exam cards. | BB_WIP7 | BB_WIP9 |  | Subsystem | Plan | NA | Fixed | jose_s | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP9_SubCom | Brilliance 64 |  | BB_WIP9 | August 1, 2019 6:34:33 PM GMT+08:00
CTUDT00625137 | Refresh | Closed | manjunath_mk | July 18, 2019 12:16:16 PM GMT+08:00 | 'Edit auto voice' option is available for second acquisition in scan ruler for Dual energy Exam cards. | BB_WIP7 | BB_WIP9 |  | Subsystem | Plan | NA | Fixed | jose_s | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP9_SubCom | Brilliance Big Bore |  | BB_WIP9 | August 1, 2019 6:34:33 PM GMT+08:00
CTUDT00625137 | Refresh | Closed | manjunath_mk | July 18, 2019 12:16:16 PM GMT+08:00 | 'Edit auto voice' option is available for second acquisition in scan ruler for Dual energy Exam cards. | BB_WIP7 | BB_WIP9 |  | Subsystem | Plan | NA | Fixed | jose_s | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP9_SubCom | Ingenuity 64 |  | BB_WIP9 | August 1, 2019 6:34:33 PM GMT+08:00
CTUDT00625135 | Refresh | Closed | 320053767 | July 17, 2019 7:22:46 PM GMT+08:00 | No colors in Z effective secondary series. | WIP18 |  |  | Subsystem | Spectral CT Viewer | NA | Works As Specified / Not A Defect | aviw | 0 | 2 |  | Support_Spectral | NO | Closed | WIP19SegReq | IQon1
CTUDT00625132 | Refresh | Closed | shekhar_p | July 17, 2019 6:06:02 PM GMT+08:00 | For Ultra High resolution for all age groups iDose level is not showing 1 to 7 | WIP17 | WIP19 |  | Subsystem | iDose | NA | Fixed | chokkalingam_r | 0 | 2 | Console Software | Console4 | NO | Closed | WIP19SubCom | Ovation |  | WIP19 | July 25, 2019 5:31:14 PM GMT+08:00
CTUDT00625132 | Refresh | Closed | shekhar_p | July 17, 2019 6:06:02 PM GMT+08:00 | For Ultra High resolution for all age groups iDose level is not showing 1 to 7 | WIP17 | WIP19 |  | Subsystem | iDose | NA | Fixed | chokkalingam_r | 0 | 2 | Console Software | Console4 | NO | Closed | WIP19SubCom | Brilliance 64 |  | WIP19 | July 25, 2019 5:31:14 PM GMT+08:00
CTUDT00625132 | Refresh | Closed | shekhar_p | July 17, 2019 6:06:02 PM GMT+08:00 | For Ultra High resolution for all age groups iDose level is not showing 1 to 7 | WIP17 | WIP19 |  | Subsystem | iDose | NA | Fixed | chokkalingam_r | 0 | 2 | Console Software | Console4 | NO | Closed | WIP19SubCom | Ingenuity 64 |  | WIP19 | July 25, 2019 5:31:14 PM GMT+08:00
CTUDT00625130 | Refresh | Closed | charles_j | July 17, 2019 5:37:13 PM GMT+08:00 | In CT-NM.SRS.IQ.1434, remove Tech Pubs as Inheritor and add Service Pubs as Inheritor. | WIP18 | DHF298555 Rev 06 - BB SRS DHF298554 Rev 06 - Ing/BR64 SRS DHF298553 Rev 06 - iCT/iCT-SP SRS DHF298556 Rev 07 - IQon SRS |  | System | Requirements | CT-NM.SRS.IQ.1434 | Fixed | mistry_h | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | BB_WIP9_SubCom | Ovation |  | DHF298555 Rev 06 - BB SRS DHF298554 Rev 06 - Ing/BR64 SRS DHF298553 Rev 06 - iCT/iCT-SP SRS DHF298556 Rev 07 - IQon SRS | July 26, 2019 8:19:26 PM GMT+08:00
CTUDT00625130 | Refresh | Closed | charles_j | July 17, 2019 5:37:13 PM GMT+08:00 | In CT-NM.SRS.IQ.1434, remove Tech Pubs as Inheritor and add Service Pubs as Inheritor. | WIP18 | DHF298555 Rev 06 - BB SRS DHF298554 Rev 06 - Ing/BR64 SRS DHF298553 Rev 06 - iCT/iCT-SP SRS DHF298556 Rev 07 - IQon SRS |  | System | Requirements | CT-NM.SRS.IQ.1434 | Fixed | mistry_h | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | BB_WIP9_SubCom | Brilliance 64 |  | DHF298555 Rev 06 - BB SRS DHF298554 Rev 06 - Ing/BR64 SRS DHF298553 Rev 06 - iCT/iCT-SP SRS DHF298556 Rev 07 - IQon SRS | July 26, 2019 8:19:26 PM GMT+08:00
CTUDT00625130 | Refresh | Closed | charles_j | July 17, 2019 5:37:13 PM GMT+08:00 | In CT-NM.SRS.IQ.1434, remove Tech Pubs as Inheritor and add Service Pubs as Inheritor. | WIP18 | DHF298555 Rev 06 - BB SRS DHF298554 Rev 06 - Ing/BR64 SRS DHF298553 Rev 06 - iCT/iCT-SP SRS DHF298556 Rev 07 - IQon SRS |  | System | Requirements | CT-NM.SRS.IQ.1434 | Fixed | mistry_h | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | BB_WIP9_SubCom | Brilliance Big Bore |  | DHF298555 Rev 06 - BB SRS DHF298554 Rev 06 - Ing/BR64 SRS DHF298553 Rev 06 - iCT/iCT-SP SRS DHF298556 Rev 07 - IQon SRS | July 26, 2019 8:19:26 PM GMT+08:00
CTUDT00625130 | Refresh | Closed | charles_j | July 17, 2019 5:37:13 PM GMT+08:00 | In CT-NM.SRS.IQ.1434, remove Tech Pubs as Inheritor and add Service Pubs as Inheritor. | WIP18 | DHF298555 Rev 06 - BB SRS DHF298554 Rev 06 - Ing/BR64 SRS DHF298553 Rev 06 - iCT/iCT-SP SRS DHF298556 Rev 07 - IQon SRS |  | System | Requirements | CT-NM.SRS.IQ.1434 | Fixed | mistry_h | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | BB_WIP9_SubCom | Ingenuity 64 |  | DHF298555 Rev 06 - BB SRS DHF298554 Rev 06 - Ing/BR64 SRS DHF298553 Rev 06 - iCT/iCT-SP SRS DHF298556 Rev 07 - IQon SRS | July 26, 2019 8:19:26 PM GMT+08:00
CTUDT00625130 | Refresh | Closed | charles_j | July 17, 2019 5:37:13 PM GMT+08:00 | In CT-NM.SRS.IQ.1434, remove Tech Pubs as Inheritor and add Service Pubs as Inheritor. | WIP18 | DHF298555 Rev 06 - BB SRS DHF298554 Rev 06 - Ing/BR64 SRS DHF298553 Rev 06 - iCT/iCT-SP SRS DHF298556 Rev 07 - IQon SRS |  | System | Requirements | CT-NM.SRS.IQ.1434 | Fixed | mistry_h | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | BB_WIP9_SubCom | IQon1 |  | DHF298555 Rev 06 - BB SRS DHF298554 Rev 06 - Ing/BR64 SRS DHF298553 Rev 06 - iCT/iCT-SP SRS DHF298556 Rev 07 - IQon SRS | July 26, 2019 8:19:26 PM GMT+08:00
CTUDT00625125 | Refresh | Closed | attali_s | July 17, 2019 2:59:57 PM GMT+08:00 | Refresh: Acceptance/Constancy: Wrong limit for Pediatric Body case with HPhantom | WIP17 | WIP19 |  | Subsystem | Acceptance/Constancy | NA | Fixed | Kostya_Baklan | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP19SegReq | Ovation |  | WIP19 | July 31, 2019 6:08:47 PM GMT+08:00
CTUDT00625121 | Refresh | Closed | wang_k | July 17, 2019 10:14:14 AM GMT+08:00 | SRS input and verification evidence are not sufficient for Translate languages of laser label | WIP17 | WIP19 BB_WIP9 |  | Subsystem | Requirements | GCX.SSRS.iCT.21066 | Fixed | ye_h | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP19SegReq | Brilliance 64 |  | WIP19 BB_WIP9 | September 3, 2019 2:55:40 PM GMT+08:00
CTUDT00625121 | Refresh | Closed | wang_k | July 17, 2019 10:14:14 AM GMT+08:00 | SRS input and verification evidence are not sufficient for Translate languages of laser label | WIP17 | WIP19 BB_WIP9 |  | Subsystem | Requirements | GCX.SSRS.iCT.21066 | Fixed | ye_h | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP19SegReq | Brilliance Big Bore |  | WIP19 BB_WIP9 | September 3, 2019 2:55:40 PM GMT+08:00
CTUDT00625121 | Refresh | Closed | wang_k | July 17, 2019 10:14:14 AM GMT+08:00 | SRS input and verification evidence are not sufficient for Translate languages of laser label | WIP17 | WIP19 BB_WIP9 |  | Subsystem | Requirements | GCX.SSRS.iCT.21066 | Fixed | ye_h | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP19SegReq | Ingenuity 64 |  | WIP19 BB_WIP9 | September 3, 2019 2:55:40 PM GMT+08:00
CTUDT00625113 | Refresh | Closed | finestine_m | July 16, 2019 7:52:21 PM GMT+08:00 | Refresh: Beam and DMS: Cracked Wedge Not Detected | BB_WIP7 | N/A |  | Subsystem | Calibrations | NA | Fixed | yosfan_s | 0 | 2 | Test Design | DMS | NO | Closed | WIP19SubCom | Brilliance Big Bore |  | N/A | August 5, 2019 6:58:19 PM GMT+08:00
CTUDT00625111 | Refresh | Closed | zhu_je | July 16, 2019 7:30:13 PM GMT+08:00 | [Refresh] There is no DVD in Copy to list, although DVD has configured in "Lan config->Local Device" | WIP17 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | zhu_je | 0 |  |  |  |  | Closed |  | Brilliance 64
CTUDT00625109 | Refresh | Closed | zhu_je | July 16, 2019 7:27:50 PM GMT+08:00 | [Refresh]Radiation Dose Information image has not sent to iSite Pacs and EBW | WIP17 |  | NA | Subsystem | Test | NA | Works As Specified / Not A Defect | zhu_je | 1 | 2 |  | Support-System | NO | Closed | WIP19SegReq | Ovation
CTUDT00625109 | Refresh | Closed | zhu_je | July 16, 2019 7:27:50 PM GMT+08:00 | [Refresh]Radiation Dose Information image has not sent to iSite Pacs and EBW | WIP17 |  | NA | Subsystem | Test | NA | Works As Specified / Not A Defect | zhu_je | 1 | 2 |  | Support-System | NO | Closed | WIP19SegReq | Brilliance 64
CTUDT00625109 | Refresh | Closed | zhu_je | July 16, 2019 7:27:50 PM GMT+08:00 | [Refresh]Radiation Dose Information image has not sent to iSite Pacs and EBW | WIP17 |  | NA | Subsystem | Test | NA | Works As Specified / Not A Defect | zhu_je | 1 | 2 |  | Support-System | NO | Closed | WIP19SegReq | Brilliance Big Bore
CTUDT00625109 | Refresh | Closed | zhu_je | July 16, 2019 7:27:50 PM GMT+08:00 | [Refresh]Radiation Dose Information image has not sent to iSite Pacs and EBW | WIP17 |  | NA | Subsystem | Test | NA | Works As Specified / Not A Defect | zhu_je | 1 | 2 |  | Support-System | NO | Closed | WIP19SegReq | Ingenuity 64
CTUDT00625109 | Refresh | Closed | zhu_je | July 16, 2019 7:27:50 PM GMT+08:00 | [Refresh]Radiation Dose Information image has not sent to iSite Pacs and EBW | WIP17 |  | NA | Subsystem | Test | NA | Works As Specified / Not A Defect | zhu_je | 1 | 2 |  | Support-System | NO | Closed | WIP19SegReq | IQon1
CTUDT00625107 | Refresh | Closed | 320053768 | July 16, 2019 7:21:24 PM GMT+08:00 | system error issue while performing a perfusion scan with 80kv and 32x1.25 collimation in iCTSP configuration | WIP17 | WIP19 | Perform scan with different parameters | Subsystem | System;Blocker19 | NaN | Fixed | daibes_s | 0 | 3 | DMS Calibrations | DMS | YES | Closed | WIP19SubCom | Ovation |  | WIP19 | August 4, 2019 3:47:59 PM GMT+08:00
CTUDT00625103 | Refresh | Closed | zhu_je | July 16, 2019 6:34:37 PM GMT+08:00 | [Refresh] In local folder, cannot load Radiation Dose Information series into CT viewer or Quick Review application | WIP17 |  |  | Subsystem | CT Viewer | NA | Works As Specified / Not A Defect | aswani_v | 0 |  |  | iCAP | NO | Closed |  | Ovation
CTUDT00625103 | Refresh | Closed | zhu_je | July 16, 2019 6:34:37 PM GMT+08:00 | [Refresh] In local folder, cannot load Radiation Dose Information series into CT viewer or Quick Review application | WIP17 |  |  | Subsystem | CT Viewer | NA | Works As Specified / Not A Defect | aswani_v | 0 |  |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00625103 | Refresh | Closed | zhu_je | July 16, 2019 6:34:37 PM GMT+08:00 | [Refresh] In local folder, cannot load Radiation Dose Information series into CT viewer or Quick Review application | WIP17 |  |  | Subsystem | CT Viewer | NA | Works As Specified / Not A Defect | aswani_v | 0 |  |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00625103 | Refresh | Closed | zhu_je | July 16, 2019 6:34:37 PM GMT+08:00 | [Refresh] In local folder, cannot load Radiation Dose Information series into CT viewer or Quick Review application | WIP17 |  |  | Subsystem | CT Viewer | NA | Works As Specified / Not A Defect | aswani_v | 0 |  |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00625103 | Refresh | Closed | zhu_je | July 16, 2019 6:34:37 PM GMT+08:00 | [Refresh] In local folder, cannot load Radiation Dose Information series into CT viewer or Quick Review application | WIP17 |  |  | Subsystem | CT Viewer | NA | Works As Specified / Not A Defect | aswani_v | 0 |  |  | iCAP | NO | Closed |  | IQon1
CTUDT00625101 | Refresh | Closed | zhu_je | July 16, 2019 6:24:15 PM GMT+08:00 | [Refresh] Offline ECG Wave is not shown in acquisition scene (but shown in view scene) after click final recon for Coronary CTA EC | WIP17 | WIP19 |  | System | OfflineWave | NA | Fixed | datta_i | 0 | 3 | Console Software | Console4 | YES | Closed | wip9 and wip19 | Ovation |  | WIP19 | August 1, 2019 6:09:46 PM GMT+08:00
CTUDT00625101 | Refresh | Closed | zhu_je | July 16, 2019 6:24:15 PM GMT+08:00 | [Refresh] Offline ECG Wave is not shown in acquisition scene (but shown in view scene) after click final recon for Coronary CTA EC | WIP17 | WIP19 |  | System | OfflineWave | NA | Fixed | datta_i | 0 | 3 | Console Software | Console4 | YES | Closed | wip9 and wip19 | Brilliance 64 |  | WIP19 | August 1, 2019 6:09:46 PM GMT+08:00
CTUDT00625101 | Refresh | Closed | zhu_je | July 16, 2019 6:24:15 PM GMT+08:00 | [Refresh] Offline ECG Wave is not shown in acquisition scene (but shown in view scene) after click final recon for Coronary CTA EC | WIP17 | WIP19 |  | System | OfflineWave | NA | Fixed | datta_i | 0 | 3 | Console Software | Console4 | YES | Closed | wip9 and wip19 | Ingenuity 64 |  | WIP19 | August 1, 2019 6:09:46 PM GMT+08:00
CTUDT00625101 | Refresh | Closed | zhu_je | July 16, 2019 6:24:15 PM GMT+08:00 | [Refresh] Offline ECG Wave is not shown in acquisition scene (but shown in view scene) after click final recon for Coronary CTA EC | WIP17 | WIP19 |  | System | OfflineWave | NA | Fixed | datta_i | 0 | 3 | Console Software | Console4 | YES | Closed | wip9 and wip19 | IQon1 |  | WIP19 | August 1, 2019 6:09:46 PM GMT+08:00
CTUDT00625098 | Refresh | Closed | 320053770 | July 16, 2019 6:11:05 PM GMT+08:00 | WIP 18:ICT IMR: when pressing on "start final recon" an error pops up – “CIRS cannot initialize.” | WIP18 | WIP18.1 |  | Subsystem | Blocker; Verified1 | CIRS.SSRS.6627, CIRS.SSRS.5369, CIRS.SSRS.5372, CIRS.SSRS.6650, CIRS.SSRS.6649 | Fixed | richards_b | 0 | 2 | CIRS_SW_ImgNodeSvr | CIRS | YES | Closed | WIP18.1SubCom | Ovation |  | WIP18.1 | July 30, 2019 9:39:50 PM GMT+08:00
CTUDT00625098 | Refresh | Closed | 320053770 | July 16, 2019 6:11:05 PM GMT+08:00 | WIP 18:ICT IMR: when pressing on "start final recon" an error pops up – “CIRS cannot initialize.” | WIP18 | WIP18.1 |  | Subsystem | Blocker; Verified1 | CIRS.SSRS.6627, CIRS.SSRS.5369, CIRS.SSRS.5372, CIRS.SSRS.6650, CIRS.SSRS.6649 | Fixed | richards_b | 0 | 2 | CIRS_SW_ImgNodeSvr | CIRS | YES | Closed | WIP18.1SubCom | Brilliance 64 |  | WIP18.1 | July 30, 2019 9:39:50 PM GMT+08:00
CTUDT00625098 | Refresh | Closed | 320053770 | July 16, 2019 6:11:05 PM GMT+08:00 | WIP 18:ICT IMR: when pressing on "start final recon" an error pops up – “CIRS cannot initialize.” | WIP18 | WIP18.1 |  | Subsystem | Blocker; Verified1 | CIRS.SSRS.6627, CIRS.SSRS.5369, CIRS.SSRS.5372, CIRS.SSRS.6650, CIRS.SSRS.6649 | Fixed | richards_b | 0 | 2 | CIRS_SW_ImgNodeSvr | CIRS | YES | Closed | WIP18.1SubCom | Brilliance Big Bore |  | WIP18.1 | July 30, 2019 9:39:50 PM GMT+08:00
CTUDT00625098 | Refresh | Closed | 320053770 | July 16, 2019 6:11:05 PM GMT+08:00 | WIP 18:ICT IMR: when pressing on "start final recon" an error pops up – “CIRS cannot initialize.” | WIP18 | WIP18.1 |  | Subsystem | Blocker; Verified1 | CIRS.SSRS.6627, CIRS.SSRS.5369, CIRS.SSRS.5372, CIRS.SSRS.6650, CIRS.SSRS.6649 | Fixed | richards_b | 0 | 2 | CIRS_SW_ImgNodeSvr | CIRS | YES | Closed | WIP18.1SubCom | Ingenuity 64 |  | WIP18.1 | July 30, 2019 9:39:50 PM GMT+08:00
CTUDT00625095 | Refresh | Closed | sshahar | July 16, 2019 5:57:43 PM GMT+08:00 | Refresh/Calibrations/Legacy phantom/UHR Comb Align Eye Phantom pin is not detected in 4th scan | WIP18 | WIP19 |  | Subsystem | Blocker19 | n/a | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP19SubCom | Ovation |  | WIP19 | August 7, 2019 8:43:53 PM GMT+08:00
CTUDT00625091 | Refresh | Closed | PremKumar_Rajendran | July 16, 2019 5:10:21 PM GMT+08:00 | Refresh : Mismatch in warning message between the requirement and UI | WIP17 | DHF303211_Rev14_Host_SSRS_for_Refresh DHF316705 Rev.13-Host_SSRS_for_BB-Refresh | None | Subsystem | Login | Console.SRS-16315 | Fixed | hatti_c | 0 | 4 | Requirements Documentation | Console4 | NO | Closed | WIP19SubCom | Ovation |  | DHF303211_Rev14_Host_SSRS_for_Refresh DHF316705 Rev.13-Host_SSRS_for_BB-Refresh | August 5, 2019 2:27:56 PM GMT+08:00
CTUDT00625091 | Refresh | Closed | PremKumar_Rajendran | July 16, 2019 5:10:21 PM GMT+08:00 | Refresh : Mismatch in warning message between the requirement and UI | WIP17 | DHF303211_Rev14_Host_SSRS_for_Refresh DHF316705 Rev.13-Host_SSRS_for_BB-Refresh | None | Subsystem | Login | Console.SRS-16315 | Fixed | hatti_c | 0 | 4 | Requirements Documentation | Console4 | NO | Closed | WIP19SubCom | Brilliance 64 |  | DHF303211_Rev14_Host_SSRS_for_Refresh DHF316705 Rev.13-Host_SSRS_for_BB-Refresh | August 5, 2019 2:27:56 PM GMT+08:00
CTUDT00625091 | Refresh | Closed | PremKumar_Rajendran | July 16, 2019 5:10:21 PM GMT+08:00 | Refresh : Mismatch in warning message between the requirement and UI | WIP17 | DHF303211_Rev14_Host_SSRS_for_Refresh DHF316705 Rev.13-Host_SSRS_for_BB-Refresh | None | Subsystem | Login | Console.SRS-16315 | Fixed | hatti_c | 0 | 4 | Requirements Documentation | Console4 | NO | Closed | WIP19SubCom | Brilliance Big Bore |  | DHF303211_Rev14_Host_SSRS_for_Refresh DHF316705 Rev.13-Host_SSRS_for_BB-Refresh | August 5, 2019 2:27:56 PM GMT+08:00
CTUDT00625091 | Refresh | Closed | PremKumar_Rajendran | July 16, 2019 5:10:21 PM GMT+08:00 | Refresh : Mismatch in warning message between the requirement and UI | WIP17 | DHF303211_Rev14_Host_SSRS_for_Refresh DHF316705 Rev.13-Host_SSRS_for_BB-Refresh | None | Subsystem | Login | Console.SRS-16315 | Fixed | hatti_c | 0 | 4 | Requirements Documentation | Console4 | NO | Closed | WIP19SubCom | Ingenuity 64 |  | DHF303211_Rev14_Host_SSRS_for_Refresh DHF316705 Rev.13-Host_SSRS_for_BB-Refresh | August 5, 2019 2:27:56 PM GMT+08:00
CTUDT00625091 | Refresh | Closed | PremKumar_Rajendran | July 16, 2019 5:10:21 PM GMT+08:00 | Refresh : Mismatch in warning message between the requirement and UI | WIP17 | DHF303211_Rev14_Host_SSRS_for_Refresh DHF316705 Rev.13-Host_SSRS_for_BB-Refresh | None | Subsystem | Login | Console.SRS-16315 | Fixed | hatti_c | 0 | 4 | Requirements Documentation | Console4 | NO | Closed | WIP19SubCom | IQon1 |  | DHF303211_Rev14_Host_SSRS_for_Refresh DHF316705 Rev.13-Host_SSRS_for_BB-Refresh | August 5, 2019 2:27:56 PM GMT+08:00
CTUDT00625086 | Refresh | Closed | zhu_je | July 16, 2019 4:39:27 PM GMT+08:00 | [Refresh] Cannot import language set from USB drive and DVD in CT user mode | BB_WIP6 | BB_WIP9 |  | Subsystem | Voice Manager | NA | Fixed | prasad_r | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP9_SubCom | Brilliance Big Bore |  | BB_WIP9 | July 23, 2019 7:36:57 PM GMT+08:00
CTUDT00625083 | Refresh | Closed | 320053770 | July 16, 2019 4:37:25 PM GMT+08:00 | WIP 18 IQon: When pressing "start final recon" an error pops saying – “CIRS cannot initialize.” | WIP18 |  |  | Subsystem | Blocker | n/a | Invalid/Uncontrolled Configuration | frider_o | 0 |  |  | DMS | NO | Closed |  | IQon1
CTUDT00625082 | Refresh | Closed | zhu_je | July 16, 2019 4:37:09 PM GMT+08:00 | [Refresh] “Export Failed” error Pop up when tried to export a renamed language set | BB_WIP6 | BB_WIP9 |  | Subsystem | Voice Manager | NA | Fixed | 320057847 | 0 | 2 | Console Software | Console4 | NO | Closed | WIP19SubCom | Brilliance Big Bore |  | BB_WIP9 | July 26, 2019 4:35:23 PM GMT+08:00
CTUDT00625080 | Refresh | Closed | zhu_je | July 16, 2019 4:33:46 PM GMT+08:00 | [Refresh] “Microsoft. NET Framework” error pop up when try to delete a renamed language set | BB_WIP6 | WIP19 | None | Subsystem | Voice Manager;Blocker | CT-NM.SRS.Workflow.549 | Fixed | jose_s | 0 | 4 | Console Software | Console4 | NO | Closed | WIP19SubCom | Ovation |  | WIP19 | July 30, 2019 7:16:04 PM GMT+08:00
CTUDT00625080 | Refresh | Closed | zhu_je | July 16, 2019 4:33:46 PM GMT+08:00 | [Refresh] “Microsoft. NET Framework” error pop up when try to delete a renamed language set | BB_WIP6 | WIP19 | None | Subsystem | Voice Manager;Blocker | CT-NM.SRS.Workflow.549 | Fixed | jose_s | 0 | 4 | Console Software | Console4 | NO | Closed | WIP19SubCom | Brilliance 64 |  | WIP19 | July 30, 2019 7:16:04 PM GMT+08:00
CTUDT00625080 | Refresh | Closed | zhu_je | July 16, 2019 4:33:46 PM GMT+08:00 | [Refresh] “Microsoft. NET Framework” error pop up when try to delete a renamed language set | BB_WIP6 | WIP19 | None | Subsystem | Voice Manager;Blocker | CT-NM.SRS.Workflow.549 | Fixed | jose_s | 0 | 4 | Console Software | Console4 | NO | Closed | WIP19SubCom | Ingenuity 64 |  | WIP19 | July 30, 2019 7:16:04 PM GMT+08:00
CTUDT00625080 | Refresh | Closed | zhu_je | July 16, 2019 4:33:46 PM GMT+08:00 | [Refresh] “Microsoft. NET Framework” error pop up when try to delete a renamed language set | BB_WIP6 | WIP19 | None | Subsystem | Voice Manager;Blocker | CT-NM.SRS.Workflow.549 | Fixed | jose_s | 0 | 4 | Console Software | Console4 | NO | Closed | WIP19SubCom | IQon1 |  | WIP19 | July 30, 2019 7:16:04 PM GMT+08:00
CTUDT00625076 | Refresh | Closed | 320053767 | July 16, 2019 2:09:56 PM GMT+08:00 | Exam Card parameters are changed when loaded to plan. | WIP18 |  |  | Subsystem | ECM | NA | Invalid/Uncontrolled Configuration | rath_s | 1 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00625074 | Refresh | Closed | charles_j | July 16, 2019 1:36:52 PM GMT+08:00 | The description of CT-NM.SRS.RegStand.734 should be changed | WIP18 | WIP19 |  | Subsystem | Requirement | RegStand.734 | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP19 & WIP9 | Ovation |  | WIP19 | September 5, 2019 4:42:13 PM GMT+08:00
CTUDT00625074 | Refresh | Closed | charles_j | July 16, 2019 1:36:52 PM GMT+08:00 | The description of CT-NM.SRS.RegStand.734 should be changed | WIP18 | WIP19 |  | Subsystem | Requirement | RegStand.734 | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP19 & WIP9 | Brilliance 64 |  | WIP19 | September 5, 2019 4:42:13 PM GMT+08:00
CTUDT00625074 | Refresh | Closed | charles_j | July 16, 2019 1:36:52 PM GMT+08:00 | The description of CT-NM.SRS.RegStand.734 should be changed | WIP18 | WIP19 |  | Subsystem | Requirement | RegStand.734 | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP19 & WIP9 | Brilliance Big Bore |  | WIP19 | September 5, 2019 4:42:13 PM GMT+08:00
CTUDT00625074 | Refresh | Closed | charles_j | July 16, 2019 1:36:52 PM GMT+08:00 | The description of CT-NM.SRS.RegStand.734 should be changed | WIP18 | WIP19 |  | Subsystem | Requirement | RegStand.734 | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP19 & WIP9 | Ingenuity 64 |  | WIP19 | September 5, 2019 4:42:13 PM GMT+08:00
CTUDT00625074 | Refresh | Closed | charles_j | July 16, 2019 1:36:52 PM GMT+08:00 | The description of CT-NM.SRS.RegStand.734 should be changed | WIP18 | WIP19 |  | Subsystem | Requirement | RegStand.734 | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP19 & WIP9 | IQon1 |  | WIP19 | September 5, 2019 4:42:13 PM GMT+08:00
CTUDT00625072 | Refresh | Closed | charles_j | July 16, 2019 1:33:53 PM GMT+08:00 | The description of CT-NM.SRS.RegStand.733 should be changed | WIP18 | WIP19 |  | Subsystem | Requirement | RegStand.733 | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP19 & WIP9 | Ovation |  | WIP19 | September 5, 2019 4:26:01 PM GMT+08:00
CTUDT00625072 | Refresh | Closed | charles_j | July 16, 2019 1:33:53 PM GMT+08:00 | The description of CT-NM.SRS.RegStand.733 should be changed | WIP18 | WIP19 |  | Subsystem | Requirement | RegStand.733 | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP19 & WIP9 | Brilliance 64 |  | WIP19 | September 5, 2019 4:26:01 PM GMT+08:00
CTUDT00625072 | Refresh | Closed | charles_j | July 16, 2019 1:33:53 PM GMT+08:00 | The description of CT-NM.SRS.RegStand.733 should be changed | WIP18 | WIP19 |  | Subsystem | Requirement | RegStand.733 | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP19 & WIP9 | Brilliance Big Bore |  | WIP19 | September 5, 2019 4:26:01 PM GMT+08:00
CTUDT00625072 | Refresh | Closed | charles_j | July 16, 2019 1:33:53 PM GMT+08:00 | The description of CT-NM.SRS.RegStand.733 should be changed | WIP18 | WIP19 |  | Subsystem | Requirement | RegStand.733 | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP19 & WIP9 | Ingenuity 64 |  | WIP19 | September 5, 2019 4:26:01 PM GMT+08:00
CTUDT00625072 | Refresh | Closed | charles_j | July 16, 2019 1:33:53 PM GMT+08:00 | The description of CT-NM.SRS.RegStand.733 should be changed | WIP18 | WIP19 |  | Subsystem | Requirement | RegStand.733 | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP19 & WIP9 | IQon1 |  | WIP19 | September 5, 2019 4:26:01 PM GMT+08:00
CTUDT00625070 | Refresh | Closed | charles_j | July 16, 2019 1:25:19 PM GMT+08:00 | The description of CT-NM.SRS.RegStand.732 should be changed | WIP18 | WIP19 |  | Subsystem | Requirement | RegStand.732 | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP19 & WIP9 | Ovation |  | WIP19 | September 5, 2019 4:24:37 PM GMT+08:00
CTUDT00625070 | Refresh | Closed | charles_j | July 16, 2019 1:25:19 PM GMT+08:00 | The description of CT-NM.SRS.RegStand.732 should be changed | WIP18 | WIP19 |  | Subsystem | Requirement | RegStand.732 | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP19 & WIP9 | Brilliance 64 |  | WIP19 | September 5, 2019 4:24:37 PM GMT+08:00
CTUDT00625070 | Refresh | Closed | charles_j | July 16, 2019 1:25:19 PM GMT+08:00 | The description of CT-NM.SRS.RegStand.732 should be changed | WIP18 | WIP19 |  | Subsystem | Requirement | RegStand.732 | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP19 & WIP9 | Brilliance Big Bore |  | WIP19 | September 5, 2019 4:24:37 PM GMT+08:00
CTUDT00625070 | Refresh | Closed | charles_j | July 16, 2019 1:25:19 PM GMT+08:00 | The description of CT-NM.SRS.RegStand.732 should be changed | WIP18 | WIP19 |  | Subsystem | Requirement | RegStand.732 | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP19 & WIP9 | Ingenuity 64 |  | WIP19 | September 5, 2019 4:24:37 PM GMT+08:00
CTUDT00625070 | Refresh | Closed | charles_j | July 16, 2019 1:25:19 PM GMT+08:00 | The description of CT-NM.SRS.RegStand.732 should be changed | WIP18 | WIP19 |  | Subsystem | Requirement | RegStand.732 | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP19 & WIP9 | IQon1 |  | WIP19 | September 5, 2019 4:24:37 PM GMT+08:00
CTUDT00625066 | Refresh | Closed | azagury_d | July 15, 2019 10:20:30 PM GMT+08:00 | Refresh: Calibrations: HCOR Basic: Body case (Case 5) fails due to phantom misalignment | BB_WIP7 | WIP19 |  | Subsystem |  | NaN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP19SubCom | Brilliance Big Bore |  | WIP19 | August 4, 2019 2:04:34 PM GMT+08:00
CTUDT00625065 | Refresh | Closed | finestine_m | July 15, 2019 10:09:47 PM GMT+08:00 | Refresh: Common Functions: Wrong Eye Phantom Distance From Center In Phantom Centering Instructions | BB_WIP7 |  |  | Subsystem | Calibrations | n/a | Works As Specified / Not A Defect | yatziv_l | 0 |  |  | DMS | NO | Closed |  | Brilliance Big Bore
CTUDT00625062 | Refresh | Closed | azagury_d | July 15, 2019 8:31:35 PM GMT+08:00 | Refresh: Calibration: Harmonized and c-sof tilt measurement is incorrect. | WIP17 | WIP19 | Ensure that no phantom tilt manually | Subsystem | Calibrations | NaN | Fixed | daibes_s | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP19SubCom | Ovation |  | WIP19 | August 5, 2019 10:45:18 PM GMT+08:00
CTUDT00625062 | Refresh | Closed | azagury_d | July 15, 2019 8:31:35 PM GMT+08:00 | Refresh: Calibration: Harmonized and c-sof tilt measurement is incorrect. | WIP17 | WIP19 | Ensure that no phantom tilt manually | Subsystem | Calibrations | NaN | Fixed | daibes_s | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP19SubCom | Brilliance Big Bore |  | WIP19 | August 5, 2019 10:45:18 PM GMT+08:00
CTUDT00625062 | Refresh | Closed | azagury_d | July 15, 2019 8:31:35 PM GMT+08:00 | Refresh: Calibration: Harmonized and c-sof tilt measurement is incorrect. | WIP17 | WIP19 | Ensure that no phantom tilt manually | Subsystem | Calibrations | NaN | Fixed | daibes_s | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP19SubCom | Ingenuity 64 |  | WIP19 | August 5, 2019 10:45:18 PM GMT+08:00
CTUDT00625062 | Refresh | Closed | azagury_d | July 15, 2019 8:31:35 PM GMT+08:00 | Refresh: Calibration: Harmonized and c-sof tilt measurement is incorrect. | WIP17 | WIP19 | Ensure that no phantom tilt manually | Subsystem | Calibrations | NaN | Fixed | daibes_s | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP19SubCom | IQon1 |  | WIP19 | August 5, 2019 10:45:18 PM GMT+08:00
CTUDT00625060 | Refresh | Closed | cai_t | July 15, 2019 8:22:00 PM GMT+08:00 | When the interface language is set as Dutch,Error occurred after choosing Reference Exam Cards in patient screen | WIP16 | WIP19 |  | Subsystem | Localization | NA | Fixed | chokkalingam_r | 0 | 2 | Console Software | Console4 | NO | Closed | WIP19SubCom | Ovation |  | WIP19 | July 18, 2019 6:55:00 PM GMT+08:00
CTUDT00625060 | Refresh | Closed | cai_t | July 15, 2019 8:22:00 PM GMT+08:00 | When the interface language is set as Dutch,Error occurred after choosing Reference Exam Cards in patient screen | WIP16 | WIP19 |  | Subsystem | Localization | NA | Fixed | chokkalingam_r | 0 | 2 | Console Software | Console4 | NO | Closed | WIP19SubCom | Brilliance 64 |  | WIP19 | July 18, 2019 6:55:00 PM GMT+08:00
CTUDT00625060 | Refresh | Closed | cai_t | July 15, 2019 8:22:00 PM GMT+08:00 | When the interface language is set as Dutch,Error occurred after choosing Reference Exam Cards in patient screen | WIP16 | WIP19 |  | Subsystem | Localization | NA | Fixed | chokkalingam_r | 0 | 2 | Console Software | Console4 | NO | Closed | WIP19SubCom | Brilliance Big Bore |  | WIP19 | July 18, 2019 6:55:00 PM GMT+08:00
CTUDT00625060 | Refresh | Closed | cai_t | July 15, 2019 8:22:00 PM GMT+08:00 | When the interface language is set as Dutch,Error occurred after choosing Reference Exam Cards in patient screen | WIP16 | WIP19 |  | Subsystem | Localization | NA | Fixed | chokkalingam_r | 0 | 2 | Console Software | Console4 | NO | Closed | WIP19SubCom | Ingenuity 64 |  | WIP19 | July 18, 2019 6:55:00 PM GMT+08:00
CTUDT00625060 | Refresh | Closed | cai_t | July 15, 2019 8:22:00 PM GMT+08:00 | When the interface language is set as Dutch,Error occurred after choosing Reference Exam Cards in patient screen | WIP16 | WIP19 |  | Subsystem | Localization | NA | Fixed | chokkalingam_r | 0 | 2 | Console Software | Console4 | NO | Closed | WIP19SubCom | IQon1 |  | WIP19 | July 18, 2019 6:55:00 PM GMT+08:00
CTUDT00625059 | Refresh | Closed | ding_s | July 15, 2019 8:13:00 PM GMT+08:00 | There were more than one reconstruction paralleled reconstructing when disable Concurrent Reconstruction. | WIP16 | DHF303211 Rev15 ING, BR64 SRS DHF298554 Rev08 iCT, SRS DHF298553 Rev.08 IQon SRS DHF298556 Rev.09 |  | Subsystem | License;Verified1 | Console-SRS.17505, ProductConfig.511 | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WAS | Ovation |  | DHF303211 Rev15 ING, BR64 SRS DHF298554 Rev08 iCT, SRS DHF298553 Rev.08 IQon SRS DHF298556 Rev.09 | November 6, 2019 1:50:02 PM GMT+08:00
CTUDT00625059 | Refresh | Closed | ding_s | July 15, 2019 8:13:00 PM GMT+08:00 | There were more than one reconstruction paralleled reconstructing when disable Concurrent Reconstruction. | WIP16 | DHF303211 Rev15 ING, BR64 SRS DHF298554 Rev08 iCT, SRS DHF298553 Rev.08 IQon SRS DHF298556 Rev.09 |  | Subsystem | License;Verified1 | Console-SRS.17505, ProductConfig.511 | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WAS | Brilliance 64 |  | DHF303211 Rev15 ING, BR64 SRS DHF298554 Rev08 iCT, SRS DHF298553 Rev.08 IQon SRS DHF298556 Rev.09 | November 6, 2019 1:50:02 PM GMT+08:00
CTUDT00625059 | Refresh | Closed | ding_s | July 15, 2019 8:13:00 PM GMT+08:00 | There were more than one reconstruction paralleled reconstructing when disable Concurrent Reconstruction. | WIP16 | DHF303211 Rev15 ING, BR64 SRS DHF298554 Rev08 iCT, SRS DHF298553 Rev.08 IQon SRS DHF298556 Rev.09 |  | Subsystem | License;Verified1 | Console-SRS.17505, ProductConfig.511 | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WAS | Ingenuity 64 |  | DHF303211 Rev15 ING, BR64 SRS DHF298554 Rev08 iCT, SRS DHF298553 Rev.08 IQon SRS DHF298556 Rev.09 | November 6, 2019 1:50:02 PM GMT+08:00
CTUDT00625056 | Refresh | Closed | lin_s | July 15, 2019 7:40:00 PM GMT+08:00 | [Refresh] : No SRS link to CT-NM.PRS.CRS.5398 for Refresh Brilliance64 | WIP18 | WIP19 |  | Subsystem | Requirements | Syssupport.203 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP19SegReq | Brilliance 64 |  | WIP19 | August 27, 2019 5:16:54 PM GMT+08:00
CTUDT00625055 | Refresh | Closed | kapadia_v | July 15, 2019 7:19:03 PM GMT+08:00 | Cut option was not present in edited results context menu | BB_WIP8 | BB_WIP9 | Copy and paste and then delete. | Subsystem | System | Console-SRS.9081 | Fixed | prasad_r | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP9_SubCom | Brilliance Big Bore |  | BB_WIP9 | July 19, 2019 8:40:54 PM GMT+08:00
CTUDT00625051 | Refresh | Closed | azagury_d | July 15, 2019 6:39:52 PM GMT+08:00 | Refresh: Calibrations: Performance: Mismatch in iDose limits between spec and machine | BB_WIP7 | N/A |  | Subsystem |  | N/A | Fixed | braunstein_d | 0 | 2 | Design Documentation | Support-Physics | NO | Closed | BB_WIP9_SegReq | Brilliance Big Bore |  | N/A | September 16, 2019 3:59:02 PM GMT+08:00
CTUDT00625050 | Refresh | Closed | kapadia_v | July 15, 2019 5:49:28 PM GMT+08:00 | Val: System Error during Acquisition Duplication in Active Scan | WIP7 | BB_WIP9 | instead of cutting and pasting the clinical step, user can cut and paste the uncompleted surview. | System | Planning | Console-SRS.4997 | Fixed | jose_s | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP9_SubCom | Brilliance Big Bore |  | BB_WIP9 | July 19, 2019 7:37:51 PM GMT+08:00
CTUDT00625048 | Refresh | Closed | yatziv_l | July 15, 2019 4:13:15 PM GMT+08:00 | Refresh: BB DMS should be removed from CT-NM.SRS.RegStand.741 | N/A | BB_WIP9 |  | Subsystem | Requirements | CT-NM.SRS.RegStand.741 | Fixed | ye_h | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | BB_WIP9_SubCom | Brilliance Big Bore |  | BB_WIP9 | September 10, 2019 4:39:42 PM GMT+08:00
CTUDT00625045 | Refresh | Closed | zyanjiang | July 15, 2019 3:43:55 PM GMT+08:00 | Refresh Pre-Val:after type"lung" in the "label" of parameter interface of Exam card ,only click enter button can make it work | BB_WIP6 | WIP20 | After editing the label, one pressing enter the data will be updated. | Subsystem | ExamCardManager | NA | Fixed | jose_s | 0 | 3 | Console Software | Console4 | NO | Closed | WIP19SegReq | Ovation |  | WIP20 | August 14, 2019 5:13:34 PM GMT+08:00
CTUDT00625045 | Refresh | Closed | zyanjiang | July 15, 2019 3:43:55 PM GMT+08:00 | Refresh Pre-Val:after type"lung" in the "label" of parameter interface of Exam card ,only click enter button can make it work | BB_WIP6 | WIP20 | After editing the label, one pressing enter the data will be updated. | Subsystem | ExamCardManager | NA | Fixed | jose_s | 0 | 3 | Console Software | Console4 | NO | Closed | WIP19SegReq | Brilliance 64 |  | WIP20 | August 14, 2019 5:13:34 PM GMT+08:00
CTUDT00625045 | Refresh | Closed | zyanjiang | July 15, 2019 3:43:55 PM GMT+08:00 | Refresh Pre-Val:after type"lung" in the "label" of parameter interface of Exam card ,only click enter button can make it work | BB_WIP6 | WIP20 | After editing the label, one pressing enter the data will be updated. | Subsystem | ExamCardManager | NA | Fixed | jose_s | 0 | 3 | Console Software | Console4 | NO | Closed | WIP19SegReq | Ingenuity 64 |  | WIP20 | August 14, 2019 5:13:34 PM GMT+08:00
CTUDT00625045 | Refresh | Closed | zyanjiang | July 15, 2019 3:43:55 PM GMT+08:00 | Refresh Pre-Val:after type"lung" in the "label" of parameter interface of Exam card ,only click enter button can make it work | BB_WIP6 | WIP20 | After editing the label, one pressing enter the data will be updated. | Subsystem | ExamCardManager | NA | Fixed | jose_s | 0 | 3 | Console Software | Console4 | NO | Closed | WIP19SegReq | IQon1 |  | WIP20 | August 14, 2019 5:13:34 PM GMT+08:00
CTUDT00625044 | Refresh | Closed | finestine_m | July 15, 2019 3:33:08 PM GMT+08:00 | Refresh: Performance: Wrong Noise ROI Radius For Case Body 1a | WIP17 | WIP19 |  | Subsystem | Calibrations; Verified1 | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP19SubCom | Ovation |  | WIP19 | August 4, 2019 2:52:32 PM GMT+08:00
CTUDT00625044 | Refresh | Closed | finestine_m | July 15, 2019 3:33:08 PM GMT+08:00 | Refresh: Performance: Wrong Noise ROI Radius For Case Body 1a | WIP17 | WIP19 |  | Subsystem | Calibrations; Verified1 | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP19SubCom | Ingenuity 64 |  | WIP19 | August 4, 2019 2:52:32 PM GMT+08:00
CTUDT00625044 | Refresh | Closed | finestine_m | July 15, 2019 3:33:08 PM GMT+08:00 | Refresh: Performance: Wrong Noise ROI Radius For Case Body 1a | WIP17 | WIP19 |  | Subsystem | Calibrations; Verified1 | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP19SubCom | IQon1 |  | WIP19 | August 4, 2019 2:52:32 PM GMT+08:00
CTUDT00625042 | Refresh | Closed | sun_q | July 15, 2019 1:56:27 PM GMT+08:00 | Gantry UPS Disconnect Warning Label is not existed in the back of the Gantry. | WIP16 | WIP18.1 |  | Subsystem | Labels | GCX.SSRS.iCT.19777 | Fixed | zhu_l | 0 | 2 | Design Documentation | GCX | YES | Closed | WIP19SubCom | Ingenuity 64 |  | WIP18.1 | August 14, 2019 1:24:34 PM GMT+08:00
CTUDT00625038 | Refresh | Closed | wang_l | July 14, 2019 9:44:34 PM GMT+08:00 | Refresh: The Last Name of the patient Selected from the HIS/RIS DB or Schedule DB display incorrect | WIP16 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | wang_l | 0 |  |  |  |  | Closed |  | Ingenuity 64
CTUDT00625026 | Refresh | Closed | ding_s | July 12, 2019 1:33:12 PM GMT+08:00 | Elevated head holder should be removed from CT-NM.SRS.SysPerf.245 | BB_WIP6 |  |  | Subsystem | Requirements;Blocker | N/A | Works As Specified / Not A Defect | wang_ada | 0 |  |  | Couch | NO | Closed |  | Brilliance 64
CTUDT00625026 | Refresh | Closed | ding_s | July 12, 2019 1:33:12 PM GMT+08:00 | Elevated head holder should be removed from CT-NM.SRS.SysPerf.245 | BB_WIP6 |  |  | Subsystem | Requirements;Blocker | N/A | Works As Specified / Not A Defect | wang_ada | 0 |  |  | Couch | NO | Closed |  | Brilliance Big Bore
CTUDT00625026 | Refresh | Closed | ding_s | July 12, 2019 1:33:12 PM GMT+08:00 | Elevated head holder should be removed from CT-NM.SRS.SysPerf.245 | BB_WIP6 |  |  | Subsystem | Requirements;Blocker | N/A | Works As Specified / Not A Defect | wang_ada | 0 |  |  | Couch | NO | Closed |  | Ingenuity 64
CTUDT00625024 | Refresh | Closed | zyanjiang | July 12, 2019 11:21:53 AM GMT+08:00 | Refresh Pre-Val:Bolus track:after locator finished,move the table height,there will be a popup"scan cant be performed…." | BB_WIP6 |  |  | Subsystem |  |  | Not Enough Information | zyanjiang | 0 |  |  | Console4 |  | Closed |  | Brilliance Big Bore
CTUDT00625022 | Refresh | Closed | zyanjiang | July 12, 2019 11:16:38 AM GMT+08:00 | Refresh Pre-Val:In the left corner ,name of disk space display incomplete (truncation issue) | BB_WIP6 |  |  | Subsystem |  |  | Duplicate | s_sa | 0 |  |  |  |  | Closed |  | Brilliance Big Bore
CTUDT00625020 | Refresh | Closed | zyanjiang | July 12, 2019 11:11:47 AM GMT+08:00 | Refresh Pre-Val:after click “bug report”,red alert behind mandatory item will flashing several times | BB_WIP6 |  |  | Subsystem | Bugrep | NA | Works As Specified / Not A Defect | Prakash_Chand_Yadav | 0 |  |  | Console4 | NO | Closed | WAS | Ovation
CTUDT00625020 | Refresh | Closed | zyanjiang | July 12, 2019 11:11:47 AM GMT+08:00 | Refresh Pre-Val:after click “bug report”,red alert behind mandatory item will flashing several times | BB_WIP6 |  |  | Subsystem | Bugrep | NA | Works As Specified / Not A Defect | Prakash_Chand_Yadav | 0 |  |  | Console4 | NO | Closed | WAS | Brilliance Big Bore
CTUDT00625020 | Refresh | Closed | zyanjiang | July 12, 2019 11:11:47 AM GMT+08:00 | Refresh Pre-Val:after click “bug report”,red alert behind mandatory item will flashing several times | BB_WIP6 |  |  | Subsystem | Bugrep | NA | Works As Specified / Not A Defect | Prakash_Chand_Yadav | 0 |  |  | Console4 | NO | Closed | WAS | Ingenuity 64
CTUDT00625020 | Refresh | Closed | zyanjiang | July 12, 2019 11:11:47 AM GMT+08:00 | Refresh Pre-Val:after click “bug report”,red alert behind mandatory item will flashing several times | BB_WIP6 |  |  | Subsystem | Bugrep | NA | Works As Specified / Not A Defect | Prakash_Chand_Yadav | 0 |  |  | Console4 | NO | Closed | WAS | IQon1
CTUDT00625018 | Refresh | Closed | zyanjiang | July 12, 2019 11:07:01 AM GMT+08:00 | Refresh Pre-Val:when choose“dual surview”during scanning,can’t change the direction,it turns grey | BB_WIP6 | DHF303211_Rev14_Host_SSRS_for_Refresh DHF316705 Rev.13-Host_SSRS_for_BB-Refresh | N/A | Subsystem | DualSurview | NA | Fixed | hatti_c | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | WAS | Brilliance Big Bore |  | DHF303211_Rev14_Host_SSRS_for_Refresh DHF316705 Rev.13-Host_SSRS_for_BB-Refresh | August 5, 2019 2:28:59 PM GMT+08:00
CTUDT00625016 | Refresh | Closed | zyanjiang | July 12, 2019 11:01:50 AM GMT+08:00 | Refresh Pre-Val:Extend scan rules | BB_WIP6 |  | Not Required. | Subsystem | Extend | NA | Works As Specified / Not A Defect | kapadia_v | 0 |  |  | Console4 | NO | Closed | WAS | Brilliance Big Bore
CTUDT00625014 | Refresh | Closed | zyanjiang | July 12, 2019 10:53:55 AM GMT+08:00 | Refresh Pre-Val:truncation in "preferences-reconstruction-oncology-CT simulation Exam card presets" | BB_WIP6 |  |  | Subsystem |  |  | Duplicate | s_sa | 0 |  |  |  |  | Closed |  | Brilliance Big Bore
CTUDT00625012 | Refresh | Closed | zyanjiang | July 12, 2019 10:44:16 AM GMT+08:00 | Refresh Pre-Val:Hardware failure | BB_WIP6 |  |  | Subsystem | GantryPanel | NA | Invalid/Uncontrolled Configuration | wang_k | 1 |  |  | GCX | NO | Closed |  | Brilliance Big Bore
CTUDT00625010 | Refresh | Closed | zyanjiang | July 12, 2019 10:33:00 AM GMT+08:00 | Refresh Pre-Val:The highest table position value is variable | BB_WIP6 |  |  | Subsystem | Couch | DHF173447-2106. | Works As Specified / Not A Defect | zhang_je | 0 |  |  | Couch | NO | Closed |  | Ovation
CTUDT00625010 | Refresh | Closed | zyanjiang | July 12, 2019 10:33:00 AM GMT+08:00 | Refresh Pre-Val:The highest table position value is variable | BB_WIP6 |  |  | Subsystem | Couch | DHF173447-2106. | Works As Specified / Not A Defect | zhang_je | 0 |  |  | Couch | NO | Closed |  | Brilliance 64
CTUDT00625010 | Refresh | Closed | zyanjiang | July 12, 2019 10:33:00 AM GMT+08:00 | Refresh Pre-Val:The highest table position value is variable | BB_WIP6 |  |  | Subsystem | Couch | DHF173447-2106. | Works As Specified / Not A Defect | zhang_je | 0 |  |  | Couch | NO | Closed |  | Brilliance Big Bore
CTUDT00625010 | Refresh | Closed | zyanjiang | July 12, 2019 10:33:00 AM GMT+08:00 | Refresh Pre-Val:The highest table position value is variable | BB_WIP6 |  |  | Subsystem | Couch | DHF173447-2106. | Works As Specified / Not A Defect | zhang_je | 0 |  |  | Couch | NO | Closed |  | Ingenuity 64
CTUDT00625010 | Refresh | Closed | zyanjiang | July 12, 2019 10:33:00 AM GMT+08:00 | Refresh Pre-Val:The highest table position value is variable | BB_WIP6 |  |  | Subsystem | Couch | DHF173447-2106. | Works As Specified / Not A Defect | zhang_je | 0 |  |  | Couch | NO | Closed |  | Ingenuity 16
CTUDT00625010 | Refresh | Closed | zyanjiang | July 12, 2019 10:33:00 AM GMT+08:00 | Refresh Pre-Val:The highest table position value is variable | BB_WIP6 |  |  | Subsystem | Couch | DHF173447-2106. | Works As Specified / Not A Defect | zhang_je | 0 |  |  | Couch | NO | Closed |  | IQon1
CTUDT00625007 | Refresh | Closed | zyanjiang | July 12, 2019 9:57:54 AM GMT+08:00 | Refresh Pre-Val:Emergency login failed | BB_WIP6 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | mistry_h | 0 |  |  |  |  | Closed |  | Brilliance Big Bore
CTUDT00625003 | Refresh | Closed | bommiReddy_r | July 11, 2019 10:34:26 PM GMT+08:00 | Refresh: Recon: "IBC structure under ELSCINTCT_IBC_PARAM_TAG” is missing in Recon MDU file if "Harmonized Phantom" option was unchecked | WIP17 | DHF318051Rev19 |  | Subsystem | MDU | Console-SRS.8171 | Fixed | shekhar_p | 0 | 2 | Test Design | Console4 | NO | Closed | WIP19SegReq | Ovation |  | DHF318051Rev19 | August 5, 2019 5:01:44 PM GMT+08:00
CTUDT00625003 | Refresh | Closed | bommiReddy_r | July 11, 2019 10:34:26 PM GMT+08:00 | Refresh: Recon: "IBC structure under ELSCINTCT_IBC_PARAM_TAG” is missing in Recon MDU file if "Harmonized Phantom" option was unchecked | WIP17 | DHF318051Rev19 |  | Subsystem | MDU | Console-SRS.8171 | Fixed | shekhar_p | 0 | 2 | Test Design | Console4 | NO | Closed | WIP19SegReq | Brilliance 64 |  | DHF318051Rev19 | August 5, 2019 5:01:44 PM GMT+08:00
CTUDT00625003 | Refresh | Closed | bommiReddy_r | July 11, 2019 10:34:26 PM GMT+08:00 | Refresh: Recon: "IBC structure under ELSCINTCT_IBC_PARAM_TAG” is missing in Recon MDU file if "Harmonized Phantom" option was unchecked | WIP17 | DHF318051Rev19 |  | Subsystem | MDU | Console-SRS.8171 | Fixed | shekhar_p | 0 | 2 | Test Design | Console4 | NO | Closed | WIP19SegReq | Ingenuity 64 |  | DHF318051Rev19 | August 5, 2019 5:01:44 PM GMT+08:00
CTUDT00625001 | Refresh | Closed | Sharmila_Vasireddy | July 11, 2019 8:12:52 PM GMT+08:00 | The console allowed to add duplicate exam cards in the suggested EC group | BB_WIP7 | WIP19 |  | Subsystem | ECM | NA | Fixed | datta_i | 0 | 4 | Console Software | Console4 | NO | Closed | WIP19SegReq | Ovation |  | WIP19 | July 26, 2019 8:32:49 PM GMT+08:00
CTUDT00625001 | Refresh | Closed | Sharmila_Vasireddy | July 11, 2019 8:12:52 PM GMT+08:00 | The console allowed to add duplicate exam cards in the suggested EC group | BB_WIP7 | WIP19 |  | Subsystem | ECM | NA | Fixed | datta_i | 0 | 4 | Console Software | Console4 | NO | Closed | WIP19SegReq | Brilliance 64 |  | WIP19 | July 26, 2019 8:32:49 PM GMT+08:00
CTUDT00625001 | Refresh | Closed | Sharmila_Vasireddy | July 11, 2019 8:12:52 PM GMT+08:00 | The console allowed to add duplicate exam cards in the suggested EC group | BB_WIP7 | WIP19 |  | Subsystem | ECM | NA | Fixed | datta_i | 0 | 4 | Console Software | Console4 | NO | Closed | WIP19SegReq | Ingenuity 64 |  | WIP19 | July 26, 2019 8:32:49 PM GMT+08:00
CTUDT00625001 | Refresh | Closed | Sharmila_Vasireddy | July 11, 2019 8:12:52 PM GMT+08:00 | The console allowed to add duplicate exam cards in the suggested EC group | BB_WIP7 | WIP19 |  | Subsystem | ECM | NA | Fixed | datta_i | 0 | 4 | Console Software | Console4 | NO | Closed | WIP19SegReq | IQon1 |  | WIP19 | July 26, 2019 8:32:49 PM GMT+08:00
CTUDT00624999 | Refresh | Closed | Sharmila_Vasireddy | July 11, 2019 8:07:46 PM GMT+08:00 | “Average planned mA” was displaying in the acquisition viewer | BB_WIP7 | DHF303211_Rev14_Host_SSRS_for_Refresh DHF316705 Rev.13-Host_SSRS_for_BB-Refresh |  | Subsystem | System | Console-SRS.16127 | Fixed | hatti_c | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | WIP19SubCom | Ovation |  | DHF303211_Rev14_Host_SSRS_for_Refresh DHF316705 Rev.13-Host_SSRS_for_BB-Refresh | August 5, 2019 2:23:31 PM GMT+08:00
CTUDT00624999 | Refresh | Closed | Sharmila_Vasireddy | July 11, 2019 8:07:46 PM GMT+08:00 | “Average planned mA” was displaying in the acquisition viewer | BB_WIP7 | DHF303211_Rev14_Host_SSRS_for_Refresh DHF316705 Rev.13-Host_SSRS_for_BB-Refresh |  | Subsystem | System | Console-SRS.16127 | Fixed | hatti_c | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | WIP19SubCom | Brilliance 64 |  | DHF303211_Rev14_Host_SSRS_for_Refresh DHF316705 Rev.13-Host_SSRS_for_BB-Refresh | August 5, 2019 2:23:31 PM GMT+08:00
CTUDT00624999 | Refresh | Closed | Sharmila_Vasireddy | July 11, 2019 8:07:46 PM GMT+08:00 | “Average planned mA” was displaying in the acquisition viewer | BB_WIP7 | DHF303211_Rev14_Host_SSRS_for_Refresh DHF316705 Rev.13-Host_SSRS_for_BB-Refresh |  | Subsystem | System | Console-SRS.16127 | Fixed | hatti_c | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | WIP19SubCom | Brilliance Big Bore |  | DHF303211_Rev14_Host_SSRS_for_Refresh DHF316705 Rev.13-Host_SSRS_for_BB-Refresh | August 5, 2019 2:23:31 PM GMT+08:00
CTUDT00624999 | Refresh | Closed | Sharmila_Vasireddy | July 11, 2019 8:07:46 PM GMT+08:00 | “Average planned mA” was displaying in the acquisition viewer | BB_WIP7 | DHF303211_Rev14_Host_SSRS_for_Refresh DHF316705 Rev.13-Host_SSRS_for_BB-Refresh |  | Subsystem | System | Console-SRS.16127 | Fixed | hatti_c | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | WIP19SubCom | Ingenuity 64 |  | DHF303211_Rev14_Host_SSRS_for_Refresh DHF316705 Rev.13-Host_SSRS_for_BB-Refresh | August 5, 2019 2:23:31 PM GMT+08:00
CTUDT00624999 | Refresh | Closed | Sharmila_Vasireddy | July 11, 2019 8:07:46 PM GMT+08:00 | “Average planned mA” was displaying in the acquisition viewer | BB_WIP7 | DHF303211_Rev14_Host_SSRS_for_Refresh DHF316705 Rev.13-Host_SSRS_for_BB-Refresh |  | Subsystem | System | Console-SRS.16127 | Fixed | hatti_c | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | WIP19SubCom | IQon1 |  | DHF303211_Rev14_Host_SSRS_for_Refresh DHF316705 Rev.13-Host_SSRS_for_BB-Refresh | August 5, 2019 2:23:31 PM GMT+08:00
CTUDT00624996 | Refresh | Closed | bommiReddy_r | July 11, 2019 6:23:07 PM GMT+08:00 | Refresh:Pulmo: Not able to change the binning mode from linear to Amplitude in View2 | WIP17 |  |  | Subsystem | View2 | NA | Works As Specified / Not A Defect | prasad_r | 0 |  |  | Console4 | NO | Closed | WAS | Ingenuity 64
CTUDT00624984 | Refresh | Closed | cherylli | July 11, 2019 4:11:05 PM GMT+08:00 | Refresh Pre-Val: No IMR dicom tag in Ing | WIP16 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | mistry_h | 0 |  |  |  |  | Closed |  | Ingenuity 64
CTUDT00624980 | Refresh | Closed | cherylli | July 11, 2019 4:00:31 PM GMT+08:00 | Refresh Pre-Val: No IMR dicom tag in BB | BB_WIP7 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | mistry_h | 0 |  |  |  |  | Closed |  | Brilliance Big Bore
CTUDT00624973 | Refresh | Closed | hilas | July 10, 2019 9:57:11 PM GMT+08:00 | Refresh: (WIP16 - iCT) HCOR calibration with H-phantom is failing on case 84 on ICT scanner | WIP16 | WIP19 | Run after restart. | Subsystem | Calibrations | N/A | Fixed | elviral | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP19SubCom | Ovation |  | WIP19 | August 4, 2019 3:43:40 PM GMT+08:00
CTUDT00624971 | Refresh | Closed | alex_nihaev | July 10, 2019 8:59:41 PM GMT+08:00 | Refresh: BB: Constancy test: H Phantom: Wrong mAs scan parameter for Axial 2D body,100KV case | BB_WIP7 | WIP19 |  | Subsystem | Acceptance/Constancy;Verified1 | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP19SubCom | Brilliance Big Bore |  | WIP19 | July 31, 2019 10:13:34 PM GMT+08:00
CTUDT00624969 | Refresh | Closed | panchami_pradeep | July 10, 2019 7:55:31 PM GMT+08:00 | Refresh: Remote Viewer tool is giving error while trying to connect. | WIP18 |  |  | Subsystem | Remote Configuration |  | Duplicate | richards_b | 0 |  |  | CIRS | NO | Closed |  | Ingenuity 64
CTUDT00624964 | Refresh | Closed | manju_kolkar | July 10, 2019 5:23:48 PM GMT+08:00 | Big Bore Refresh: Clinical scan uses previously scanned Surview image even though it is planned after unperformed Surview | BB_WIP7 | BB_WIP9 |  | Subsystem | Plan | Console-SRS.5557 | Fixed | sagar_v | 0 | 4 | Console Software | Console4 | NO | Closed | WIP19SubCom | Brilliance Big Bore |  | BB_WIP9 | July 29, 2019 2:17:54 PM GMT+08:00
CTUDT00624962 | Refresh | Closed | wei_f | July 10, 2019 5:05:00 PM GMT+08:00 | Most of The Third Party Software that listed in the DHF316550 were not installed on the system. | BB_WIP6 | BB_WIP9 |  | System | Security;Blocker | NA | Fixed | wei_f | 0 | 2 | Test Design | Support-System | NO | Closed | BB_WIP9_SegReq | Brilliance 64 |  | BB_WIP9 | August 22, 2019 5:02:39 PM GMT+08:00
CTUDT00624962 | Refresh | Closed | wei_f | July 10, 2019 5:05:00 PM GMT+08:00 | Most of The Third Party Software that listed in the DHF316550 were not installed on the system. | BB_WIP6 | BB_WIP9 |  | System | Security;Blocker | NA | Fixed | wei_f | 0 | 2 | Test Design | Support-System | NO | Closed | BB_WIP9_SegReq | Brilliance Big Bore |  | BB_WIP9 | August 22, 2019 5:02:39 PM GMT+08:00
CTUDT00624962 | Refresh | Closed | wei_f | July 10, 2019 5:05:00 PM GMT+08:00 | Most of The Third Party Software that listed in the DHF316550 were not installed on the system. | BB_WIP6 | BB_WIP9 |  | System | Security;Blocker | NA | Fixed | wei_f | 0 | 2 | Test Design | Support-System | NO | Closed | BB_WIP9_SegReq | Ingenuity 64 |  | BB_WIP9 | August 22, 2019 5:02:39 PM GMT+08:00
CTUDT00624956 | Refresh | Closed | wei_f | July 10, 2019 1:07:44 PM GMT+08:00 | when other users login to the system, the system time did not display the time setting by philips service user | WIP16 |  | N/A | Subsystem | System | NA | Invalid/Uncontrolled Configuration | dasgupta_s | 0 | 3 |  | Console4 | YES | Closed | WIP19SegReq | Ovation
CTUDT00624956 | Refresh | Closed | wei_f | July 10, 2019 1:07:44 PM GMT+08:00 | when other users login to the system, the system time did not display the time setting by philips service user | WIP16 |  | N/A | Subsystem | System | NA | Invalid/Uncontrolled Configuration | dasgupta_s | 0 | 3 |  | Console4 | YES | Closed | WIP19SegReq | Brilliance 64
CTUDT00624956 | Refresh | Closed | wei_f | July 10, 2019 1:07:44 PM GMT+08:00 | when other users login to the system, the system time did not display the time setting by philips service user | WIP16 |  | N/A | Subsystem | System | NA | Invalid/Uncontrolled Configuration | dasgupta_s | 0 | 3 |  | Console4 | YES | Closed | WIP19SegReq | Brilliance Big Bore
CTUDT00624956 | Refresh | Closed | wei_f | July 10, 2019 1:07:44 PM GMT+08:00 | when other users login to the system, the system time did not display the time setting by philips service user | WIP16 |  | N/A | Subsystem | System | NA | Invalid/Uncontrolled Configuration | dasgupta_s | 0 | 3 |  | Console4 | YES | Closed | WIP19SegReq | Ingenuity 64
CTUDT00624956 | Refresh | Closed | wei_f | July 10, 2019 1:07:44 PM GMT+08:00 | when other users login to the system, the system time did not display the time setting by philips service user | WIP16 |  | N/A | Subsystem | System | NA | Invalid/Uncontrolled Configuration | dasgupta_s | 0 | 3 |  | Console4 | YES | Closed | WIP19SegReq | IQon1
CTUDT00624953 | Refresh | Closed | wang_l | July 9, 2019 9:32:30 PM GMT+08:00 | Refresh: ECM open with error by CT user when changed UI language | WIP16 |  |  | Subsystem | ECM | NA | Duplicate | prasad_r | 0 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00624953 | Refresh | Closed | wang_l | July 9, 2019 9:32:30 PM GMT+08:00 | Refresh: ECM open with error by CT user when changed UI language | WIP16 |  |  | Subsystem | ECM | NA | Duplicate | prasad_r | 0 |  |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00624953 | Refresh | Closed | wang_l | July 9, 2019 9:32:30 PM GMT+08:00 | Refresh: ECM open with error by CT user when changed UI language | WIP16 |  |  | Subsystem | ECM | NA | Duplicate | prasad_r | 0 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00624953 | Refresh | Closed | wang_l | July 9, 2019 9:32:30 PM GMT+08:00 | Refresh: ECM open with error by CT user when changed UI language | WIP16 |  |  | Subsystem | ECM | NA | Duplicate | prasad_r | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00624953 | Refresh | Closed | wang_l | July 9, 2019 9:32:30 PM GMT+08:00 | Refresh: ECM open with error by CT user when changed UI language | WIP16 |  |  | Subsystem | ECM | NA | Duplicate | prasad_r | 0 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00624950 | Refresh | Closed | wei_f | July 9, 2019 9:17:53 PM GMT+08:00 | Failed to copy studies from DVD-RAM to local. | WIP16 | WIP19 |  | Subsystem | CD-DVD operation | NA | Fixed | s_sa | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP19SubCom | Ovation |  | WIP19 | July 25, 2019 4:44:34 PM GMT+08:00
CTUDT00624950 | Refresh | Closed | wei_f | July 9, 2019 9:17:53 PM GMT+08:00 | Failed to copy studies from DVD-RAM to local. | WIP16 | WIP19 |  | Subsystem | CD-DVD operation | NA | Fixed | s_sa | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP19SubCom | Brilliance 64 |  | WIP19 | July 25, 2019 4:44:34 PM GMT+08:00
CTUDT00624950 | Refresh | Closed | wei_f | July 9, 2019 9:17:53 PM GMT+08:00 | Failed to copy studies from DVD-RAM to local. | WIP16 | WIP19 |  | Subsystem | CD-DVD operation | NA | Fixed | s_sa | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP19SubCom | Brilliance Big Bore |  | WIP19 | July 25, 2019 4:44:34 PM GMT+08:00
CTUDT00624950 | Refresh | Closed | wei_f | July 9, 2019 9:17:53 PM GMT+08:00 | Failed to copy studies from DVD-RAM to local. | WIP16 | WIP19 |  | Subsystem | CD-DVD operation | NA | Fixed | s_sa | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP19SubCom | Ingenuity 64 |  | WIP19 | July 25, 2019 4:44:34 PM GMT+08:00
CTUDT00624950 | Refresh | Closed | wei_f | July 9, 2019 9:17:53 PM GMT+08:00 | Failed to copy studies from DVD-RAM to local. | WIP16 | WIP19 |  | Subsystem | CD-DVD operation | NA | Fixed | s_sa | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP19SubCom | IQon1 |  | WIP19 | July 25, 2019 4:44:34 PM GMT+08:00
CTUDT00624948 | Refresh | Closed | wang_l | July 9, 2019 8:51:35 PM GMT+08:00 | Refresh: Script error displayed once configure network by olevel user | WIP16 | BB_WIP9 WIP19 |  | Subsystem | Service Tools | NA | Fixed | Sandeep_Kumar | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP19SubCom | Brilliance 64 |  | BB_WIP9 WIP19 | July 19, 2019 9:27:49 PM GMT+08:00
CTUDT00624948 | Refresh | Closed | wang_l | July 9, 2019 8:51:35 PM GMT+08:00 | Refresh: Script error displayed once configure network by olevel user | WIP16 | BB_WIP9 WIP19 |  | Subsystem | Service Tools | NA | Fixed | Sandeep_Kumar | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP19SubCom | Brilliance Big Bore |  | BB_WIP9 WIP19 | July 19, 2019 9:27:49 PM GMT+08:00
CTUDT00624948 | Refresh | Closed | wang_l | July 9, 2019 8:51:35 PM GMT+08:00 | Refresh: Script error displayed once configure network by olevel user | WIP16 | BB_WIP9 WIP19 |  | Subsystem | Service Tools | NA | Fixed | Sandeep_Kumar | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP19SubCom | Ingenuity 64 |  | BB_WIP9 WIP19 | July 19, 2019 9:27:49 PM GMT+08:00
CTUDT00624948 | Refresh | Closed | wang_l | July 9, 2019 8:51:35 PM GMT+08:00 | Refresh: Script error displayed once configure network by olevel user | WIP16 | BB_WIP9 WIP19 |  | Subsystem | Service Tools | NA | Fixed | Sandeep_Kumar | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP19SubCom | IQon1 |  | BB_WIP9 WIP19 | July 19, 2019 9:27:49 PM GMT+08:00
CTUDT00624945 | Refresh | Closed | wei_f | July 9, 2019 8:11:57 PM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.Service.156 was not inherited by Tech_Pub | BB_WIP6 | DHF298554 Rev07 DHF298553 Rev.07 DHF298556 Rev.08 |  | Subsystem | Requirements;Blocker | CT-NM.SRS.Service.156 | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP19 & WIP9 | Ovation |  | DHF298554 Rev07 DHF298553 Rev.07 DHF298556 Rev.08 | September 5, 2019 4:11:50 PM GMT+08:00
CTUDT00624945 | Refresh | Closed | wei_f | July 9, 2019 8:11:57 PM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.Service.156 was not inherited by Tech_Pub | BB_WIP6 | DHF298554 Rev07 DHF298553 Rev.07 DHF298556 Rev.08 |  | Subsystem | Requirements;Blocker | CT-NM.SRS.Service.156 | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP19 & WIP9 | Brilliance 64 |  | DHF298554 Rev07 DHF298553 Rev.07 DHF298556 Rev.08 | September 5, 2019 4:11:50 PM GMT+08:00
CTUDT00624945 | Refresh | Closed | wei_f | July 9, 2019 8:11:57 PM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.Service.156 was not inherited by Tech_Pub | BB_WIP6 | DHF298554 Rev07 DHF298553 Rev.07 DHF298556 Rev.08 |  | Subsystem | Requirements;Blocker | CT-NM.SRS.Service.156 | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP19 & WIP9 | Brilliance Big Bore |  | DHF298554 Rev07 DHF298553 Rev.07 DHF298556 Rev.08 | September 5, 2019 4:11:50 PM GMT+08:00
CTUDT00624945 | Refresh | Closed | wei_f | July 9, 2019 8:11:57 PM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.Service.156 was not inherited by Tech_Pub | BB_WIP6 | DHF298554 Rev07 DHF298553 Rev.07 DHF298556 Rev.08 |  | Subsystem | Requirements;Blocker | CT-NM.SRS.Service.156 | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP19 & WIP9 | Ingenuity 64 |  | DHF298554 Rev07 DHF298553 Rev.07 DHF298556 Rev.08 | September 5, 2019 4:11:50 PM GMT+08:00
CTUDT00624945 | Refresh | Closed | wei_f | July 9, 2019 8:11:57 PM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.Service.156 was not inherited by Tech_Pub | BB_WIP6 | DHF298554 Rev07 DHF298553 Rev.07 DHF298556 Rev.08 |  | Subsystem | Requirements;Blocker | CT-NM.SRS.Service.156 | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP19 & WIP9 | IQon1 |  | DHF298554 Rev07 DHF298553 Rev.07 DHF298556 Rev.08 | September 5, 2019 4:11:50 PM GMT+08:00
CTUDT00624943 | Refresh | Closed | wei_f | July 9, 2019 8:03:39 PM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.Service.161 was not inherited by Tech_Pub | BB_WIP6 | BB_WIP9 |  | Subsystem | Requirements;Blocker | CT-NM.SRS.Service.161 | Fixed | ye_h | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | PostValidation | Ovation |  | BB_WIP9 | September 10, 2019 4:37:24 PM GMT+08:00
CTUDT00624943 | Refresh | Closed | wei_f | July 9, 2019 8:03:39 PM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.Service.161 was not inherited by Tech_Pub | BB_WIP6 | BB_WIP9 |  | Subsystem | Requirements;Blocker | CT-NM.SRS.Service.161 | Fixed | ye_h | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | PostValidation | Brilliance 64 |  | BB_WIP9 | September 10, 2019 4:37:24 PM GMT+08:00
CTUDT00624943 | Refresh | Closed | wei_f | July 9, 2019 8:03:39 PM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.Service.161 was not inherited by Tech_Pub | BB_WIP6 | BB_WIP9 |  | Subsystem | Requirements;Blocker | CT-NM.SRS.Service.161 | Fixed | ye_h | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | PostValidation | Brilliance Big Bore |  | BB_WIP9 | September 10, 2019 4:37:24 PM GMT+08:00
CTUDT00624943 | Refresh | Closed | wei_f | July 9, 2019 8:03:39 PM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.Service.161 was not inherited by Tech_Pub | BB_WIP6 | BB_WIP9 |  | Subsystem | Requirements;Blocker | CT-NM.SRS.Service.161 | Fixed | ye_h | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | PostValidation | Ingenuity 64 |  | BB_WIP9 | September 10, 2019 4:37:24 PM GMT+08:00
CTUDT00624943 | Refresh | Closed | wei_f | July 9, 2019 8:03:39 PM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.Service.161 was not inherited by Tech_Pub | BB_WIP6 | BB_WIP9 |  | Subsystem | Requirements;Blocker | CT-NM.SRS.Service.161 | Fixed | ye_h | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | PostValidation | IQon1 |  | BB_WIP9 | September 10, 2019 4:37:24 PM GMT+08:00
CTUDT00624941 | Refresh | Closed | cohen_y | July 9, 2019 7:03:00 PM GMT+08:00 | Refresh: Performance Test: Harmonized: Filter value of Head lin 2, 3 and 4 are not as defined in Spec | BB_WIP7 | N/A |  | Subsystem | Performance Test | N/A | Fixed | richards_b | 0 | 2 | Design Documentation | Support-Physics | NO | Closed | BB_WIP9_SubCom | Brilliance Big Bore |  | N/A | September 13, 2019 10:34:06 PM GMT+08:00
CTUDT00624937 | Refresh | Closed | panchami_pradeep | July 9, 2019 4:59:06 PM GMT+08:00 | CT password field is grayed out, after checking "use Remote IST certificate" and enter ct user name. | WIP17 | WIP19 |  | Subsystem |  | NA | Fixed | rath_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP19SubCom | Ovation |  | WIP19 | July 16, 2019 4:57:30 PM GMT+08:00
CTUDT00624937 | Refresh | Closed | panchami_pradeep | July 9, 2019 4:59:06 PM GMT+08:00 | CT password field is grayed out, after checking "use Remote IST certificate" and enter ct user name. | WIP17 | WIP19 |  | Subsystem |  | NA | Fixed | rath_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP19SubCom | Ingenuity 64 |  | WIP19 | July 16, 2019 4:57:30 PM GMT+08:00
CTUDT00624937 | Refresh | Closed | panchami_pradeep | July 9, 2019 4:59:06 PM GMT+08:00 | CT password field is grayed out, after checking "use Remote IST certificate" and enter ct user name. | WIP17 | WIP19 |  | Subsystem |  | NA | Fixed | rath_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP19SubCom | IQon1 |  | WIP19 | July 16, 2019 4:57:30 PM GMT+08:00
CTUDT00624933 | Refresh | Closed | nanjunda_s | July 9, 2019 4:20:09 PM GMT+08:00 | Getting an error message "ViewTwo has stopped working" when performing reconstruction in View2 with more number of results | wip17 | WIP19 | User can reload the series in View 2 again taking each time less results | Subsystem | View Two | NA | Fixed | v_sumanth | 0 | 4 | Console Software | Console4 | NO | Closed | WIP19SubCom | Ovation |  | WIP19 | August 1, 2019 7:39:47 PM GMT+08:00
CTUDT00624933 | Refresh | Closed | nanjunda_s | July 9, 2019 4:20:09 PM GMT+08:00 | Getting an error message "ViewTwo has stopped working" when performing reconstruction in View2 with more number of results | wip17 | WIP19 | User can reload the series in View 2 again taking each time less results | Subsystem | View Two | NA | Fixed | v_sumanth | 0 | 4 | Console Software | Console4 | NO | Closed | WIP19SubCom | Brilliance 64 |  | WIP19 | August 1, 2019 7:39:47 PM GMT+08:00
CTUDT00624933 | Refresh | Closed | nanjunda_s | July 9, 2019 4:20:09 PM GMT+08:00 | Getting an error message "ViewTwo has stopped working" when performing reconstruction in View2 with more number of results | wip17 | WIP19 | User can reload the series in View 2 again taking each time less results | Subsystem | View Two | NA | Fixed | v_sumanth | 0 | 4 | Console Software | Console4 | NO | Closed | WIP19SubCom | Ingenuity 64 |  | WIP19 | August 1, 2019 7:39:47 PM GMT+08:00
CTUDT00624933 | Refresh | Closed | nanjunda_s | July 9, 2019 4:20:09 PM GMT+08:00 | Getting an error message "ViewTwo has stopped working" when performing reconstruction in View2 with more number of results | wip17 | WIP19 | User can reload the series in View 2 again taking each time less results | Subsystem | View Two | NA | Fixed | v_sumanth | 0 | 4 | Console Software | Console4 | NO | Closed | WIP19SubCom | IQon1 |  | WIP19 | August 1, 2019 7:39:47 PM GMT+08:00
CTUDT00624932 | Refresh | Closed | Valeria_Kashtanov | July 9, 2019 4:10:05 PM GMT+08:00 | Refresh: Calibrations: Acceptance and Constancy: ROIs in Slice Thickness cases with 2 scans are not as defined in spec | WIP17 | N/A | N/A | Subsystem | Acceptance/Constancy | N/A | Fixed | yosfan_s | 0 | 2 | Test Design | DMS | NO | Closed | WIP19SubCom | Ovation |  | N/A | August 5, 2019 6:58:50 PM GMT+08:00
CTUDT00624932 | Refresh | Closed | Valeria_Kashtanov | July 9, 2019 4:10:05 PM GMT+08:00 | Refresh: Calibrations: Acceptance and Constancy: ROIs in Slice Thickness cases with 2 scans are not as defined in spec | WIP17 | N/A | N/A | Subsystem | Acceptance/Constancy | N/A | Fixed | yosfan_s | 0 | 2 | Test Design | DMS | NO | Closed | WIP19SubCom | Brilliance Big Bore |  | N/A | August 5, 2019 6:58:50 PM GMT+08:00
CTUDT00624932 | Refresh | Closed | Valeria_Kashtanov | July 9, 2019 4:10:05 PM GMT+08:00 | Refresh: Calibrations: Acceptance and Constancy: ROIs in Slice Thickness cases with 2 scans are not as defined in spec | WIP17 | N/A | N/A | Subsystem | Acceptance/Constancy | N/A | Fixed | yosfan_s | 0 | 2 | Test Design | DMS | NO | Closed | WIP19SubCom | Ingenuity 64 |  | N/A | August 5, 2019 6:58:50 PM GMT+08:00
CTUDT00624932 | Refresh | Closed | Valeria_Kashtanov | July 9, 2019 4:10:05 PM GMT+08:00 | Refresh: Calibrations: Acceptance and Constancy: ROIs in Slice Thickness cases with 2 scans are not as defined in spec | WIP17 | N/A | N/A | Subsystem | Acceptance/Constancy | N/A | Fixed | yosfan_s | 0 | 2 | Test Design | DMS | NO | Closed | WIP19SubCom | IQon1 |  | N/A | August 5, 2019 6:58:50 PM GMT+08:00
CTUDT00624930 | Refresh | Closed | nanjunda_s | July 9, 2019 4:05:37 PM GMT+08:00 | The value of calculated "mAs for Average Patient" and show all "mAs for Average Patient" value is not matching and more than tolerance | WIP17 | DHF318051 Rev 17 |  | Subsystem | Test Design | Console.SRS.13561 | Fixed | nanjunda_s | 0 | 4 | Test Design | Console4 | NO | Closed | WAS | IQon1 | NA | DHF318051 Rev 17 | July 22, 2019 8:15:38 PM GMT+08:00
CTUDT00624928 | Refresh | Closed | Sateesh_S | July 9, 2019 3:09:47 PM GMT+08:00 | BB Refresh: In log viewer the BIG message “BIG: SUITE_DOOR_OPEN” was not displaying when we open the door during scanning | BB_WIP7 | BB_WIP9 |  | Subsystem | Log Viewer | NA | Fixed | rath_s | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP9_SubCom | Brilliance Big Bore |  | BB_WIP9 | July 22, 2019 4:54:59 PM GMT+08:00
CTUDT00624926 | Refresh | Closed | sunilkumar_j | July 9, 2019 2:27:16 PM GMT+08:00 | Ingenuity [Refresh WIP 17] : User was not able to login to service mode successfully. | WIP17 | WIP21.1 BB_WIP11.1 | • Register the device from PSC: o In Remote configuration, provide connection Type, PRS Server and Device Unique ID. o Click Connect. • Login to the PRS Portal • Search the Device Unique ID in the search bar. • Connection Status should be Online. • Ask Clinical person in the hospital to Sign-out from CT user. • Click Remote Desktop(IST) from PRS portal. • Once the disclaimer from PRS pops up to agree, don’t click Agree button. • Ask the clinical person to login using Philips_service using IST Certificate checkbox. • Then from PRS portal, Agree the disclaimer. • Ask clinical person to accept the disclaimer within Philips user in Hospital. | Subsystem | Login | NA | Fixed | venkatachalam_p | 0 | 2 | Service Tools_Software | Console4 | NO | Closed | WIP20SegReq | Ovation |  | WIP21.1 BB_WIP11.1 | October 9, 2019 7:48:57 PM GMT+08:00
CTUDT00624926 | Refresh | Closed | sunilkumar_j | July 9, 2019 2:27:16 PM GMT+08:00 | Ingenuity [Refresh WIP 17] : User was not able to login to service mode successfully. | WIP17 | WIP21.1 BB_WIP11.1 | • Register the device from PSC: o In Remote configuration, provide connection Type, PRS Server and Device Unique ID. o Click Connect. • Login to the PRS Portal • Search the Device Unique ID in the search bar. • Connection Status should be Online. • Ask Clinical person in the hospital to Sign-out from CT user. • Click Remote Desktop(IST) from PRS portal. • Once the disclaimer from PRS pops up to agree, don’t click Agree button. • Ask the clinical person to login using Philips_service using IST Certificate checkbox. • Then from PRS portal, Agree the disclaimer. • Ask clinical person to accept the disclaimer within Philips user in Hospital. | Subsystem | Login | NA | Fixed | venkatachalam_p | 0 | 2 | Service Tools_Software | Console4 | NO | Closed | WIP20SegReq | Brilliance 64 |  | WIP21.1 BB_WIP11.1 | October 9, 2019 7:48:57 PM GMT+08:00
CTUDT00624926 | Refresh | Closed | sunilkumar_j | July 9, 2019 2:27:16 PM GMT+08:00 | Ingenuity [Refresh WIP 17] : User was not able to login to service mode successfully. | WIP17 | WIP21.1 BB_WIP11.1 | • Register the device from PSC: o In Remote configuration, provide connection Type, PRS Server and Device Unique ID. o Click Connect. • Login to the PRS Portal • Search the Device Unique ID in the search bar. • Connection Status should be Online. • Ask Clinical person in the hospital to Sign-out from CT user. • Click Remote Desktop(IST) from PRS portal. • Once the disclaimer from PRS pops up to agree, don’t click Agree button. • Ask the clinical person to login using Philips_service using IST Certificate checkbox. • Then from PRS portal, Agree the disclaimer. • Ask clinical person to accept the disclaimer within Philips user in Hospital. | Subsystem | Login | NA | Fixed | venkatachalam_p | 0 | 2 | Service Tools_Software | Console4 | NO | Closed | WIP20SegReq | Brilliance Big Bore |  | WIP21.1 BB_WIP11.1 | October 9, 2019 7:48:57 PM GMT+08:00
CTUDT00624926 | Refresh | Closed | sunilkumar_j | July 9, 2019 2:27:16 PM GMT+08:00 | Ingenuity [Refresh WIP 17] : User was not able to login to service mode successfully. | WIP17 | WIP21.1 BB_WIP11.1 | • Register the device from PSC: o In Remote configuration, provide connection Type, PRS Server and Device Unique ID. o Click Connect. • Login to the PRS Portal • Search the Device Unique ID in the search bar. • Connection Status should be Online. • Ask Clinical person in the hospital to Sign-out from CT user. • Click Remote Desktop(IST) from PRS portal. • Once the disclaimer from PRS pops up to agree, don’t click Agree button. • Ask the clinical person to login using Philips_service using IST Certificate checkbox. • Then from PRS portal, Agree the disclaimer. • Ask clinical person to accept the disclaimer within Philips user in Hospital. | Subsystem | Login | NA | Fixed | venkatachalam_p | 0 | 2 | Service Tools_Software | Console4 | NO | Closed | WIP20SegReq | Ingenuity 64 |  | WIP21.1 BB_WIP11.1 | October 9, 2019 7:48:57 PM GMT+08:00
CTUDT00624926 | Refresh | Closed | sunilkumar_j | July 9, 2019 2:27:16 PM GMT+08:00 | Ingenuity [Refresh WIP 17] : User was not able to login to service mode successfully. | WIP17 | WIP21.1 BB_WIP11.1 | • Register the device from PSC: o In Remote configuration, provide connection Type, PRS Server and Device Unique ID. o Click Connect. • Login to the PRS Portal • Search the Device Unique ID in the search bar. • Connection Status should be Online. • Ask Clinical person in the hospital to Sign-out from CT user. • Click Remote Desktop(IST) from PRS portal. • Once the disclaimer from PRS pops up to agree, don’t click Agree button. • Ask the clinical person to login using Philips_service using IST Certificate checkbox. • Then from PRS portal, Agree the disclaimer. • Ask clinical person to accept the disclaimer within Philips user in Hospital. | Subsystem | Login | NA | Fixed | venkatachalam_p | 0 | 2 | Service Tools_Software | Console4 | NO | Closed | WIP20SegReq | IQon1 |  | WIP21.1 BB_WIP11.1 | October 9, 2019 7:48:57 PM GMT+08:00
CTUDT00624923 | Refresh | Closed | zhu_je | July 9, 2019 12:17:12 PM GMT+08:00 | [Refresh][PCM] Cannot find label for language “Russian” and “Vietnam” in PCM | WIP16 | BB_WIP10 BB_WIP11 BB_WIP12 BB_WIP13 BB_WIP2 BB_WIP3 BB_WIP4 BB_WIP5 BB_WIP5.1 BB_WIP6 BB_WIP7 BB_WIP8 BB_WIP8.1 BB_WIP9 WIP11.1 WIP12 WIP13 WIP14 WIP15 WIP16 WIP17 WIP18 WIP18.1 WIP19 WIP20 | CT-NM.SRS.SysSupport.819 The system shall provide labels in Russian According to the SRS, the system shall provide labels in Russian, the labels mean the "safety and warning labels". we have the Russian warning label kit in the Agile, but due to Russian accept the ENG warning labels as well, the Russian languages warning kit never go to the PCM. Due to this DEFECT, The Product manager XieHong is working on the add the Russian language warning kit into PCM. CT-NM.SRS.SysSupport.878 The system shall provide regulatory product and regulatory packaging labels for Vietnam. The system shall provide regulatory product and regulatory packing labels for Vietnam, this is not a requirement to check the Vietnam languages warning kit in PCM, this should not go to the DEFECT. | Subsystem | PCM;Blocker | CT-NM.SRS.SysSupport.819, CT-NM.SRS.SysSupport.878 | Fixed | zhu_l | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | BB PCM next version | Brilliance 64 |  | BB_WIP10 BB_WIP11 BB_WIP12 BB_WIP13 BB_WIP2 BB_WIP3 BB_WIP4 BB_WIP5 BB_WIP5.1 BB_WIP6 BB_WIP7 BB_WIP8 BB_WIP8.1 BB_WIP9 WIP11.1 WIP12 WIP13 WIP14 WIP15 WIP16 WIP17 WIP18 WIP18.1 WIP19 WIP20 | November 7, 2019 2:08:11 PM GMT+08:00
CTUDT00624923 | Refresh | Closed | zhu_je | July 9, 2019 12:17:12 PM GMT+08:00 | [Refresh][PCM] Cannot find label for language “Russian” and “Vietnam” in PCM | WIP16 | BB_WIP10 BB_WIP11 BB_WIP12 BB_WIP13 BB_WIP2 BB_WIP3 BB_WIP4 BB_WIP5 BB_WIP5.1 BB_WIP6 BB_WIP7 BB_WIP8 BB_WIP8.1 BB_WIP9 WIP11.1 WIP12 WIP13 WIP14 WIP15 WIP16 WIP17 WIP18 WIP18.1 WIP19 WIP20 | CT-NM.SRS.SysSupport.819 The system shall provide labels in Russian According to the SRS, the system shall provide labels in Russian, the labels mean the "safety and warning labels". we have the Russian warning label kit in the Agile, but due to Russian accept the ENG warning labels as well, the Russian languages warning kit never go to the PCM. Due to this DEFECT, The Product manager XieHong is working on the add the Russian language warning kit into PCM. CT-NM.SRS.SysSupport.878 The system shall provide regulatory product and regulatory packaging labels for Vietnam. The system shall provide regulatory product and regulatory packing labels for Vietnam, this is not a requirement to check the Vietnam languages warning kit in PCM, this should not go to the DEFECT. | Subsystem | PCM;Blocker | CT-NM.SRS.SysSupport.819, CT-NM.SRS.SysSupport.878 | Fixed | zhu_l | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | BB PCM next version | Ingenuity 64 |  | BB_WIP10 BB_WIP11 BB_WIP12 BB_WIP13 BB_WIP2 BB_WIP3 BB_WIP4 BB_WIP5 BB_WIP5.1 BB_WIP6 BB_WIP7 BB_WIP8 BB_WIP8.1 BB_WIP9 WIP11.1 WIP12 WIP13 WIP14 WIP15 WIP16 WIP17 WIP18 WIP18.1 WIP19 WIP20 | November 7, 2019 2:08:11 PM GMT+08:00
CTUDT00624921 | Refresh | Closed | zhu_je | July 9, 2019 10:04:37 AM GMT+08:00 | [Refresh] Rename language name through double clicking is not working for part of language set in voice manger | BB_WIP6 | BB_WIP9 | Language can be renamed during context menu. | System | Voice Manager | N/A | Fixed | chokkalingam_r | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP9_SubCom | Brilliance Big Bore |  | BB_WIP9 | July 19, 2019 6:02:50 PM GMT+08:00
CTUDT00624919 | Refresh | Closed | zhu_je | July 9, 2019 9:54:55 AM GMT+08:00 | [Refresh] “Rename language” selection is gray out after right click for default factory language in voice manager | BB_WIP6 |  |  | Subsystem | Voice Manager | N/A | Works As Specified / Not A Defect | pandey_a | 0 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00624917 | Refresh | Closed | zhu_je | July 9, 2019 9:27:48 AM GMT+08:00 | [Refresh] Cannot export language set to USB drive in CT user mode | BB_WIP6 | BB_WIP9 |  | Subsystem |  | NA | Fixed | prasad_r | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP9_SubCom | Brilliance Big Bore |  | BB_WIP9 | July 23, 2019 7:35:38 PM GMT+08:00
CTUDT00624915 | Refresh | Closed | zhu_je | July 8, 2019 9:12:33 PM GMT+08:00 | [Refresh] The systemwas not popped up with message #412. "mAs value has changed from xxx mAs to…”' | BB_WIP6 | DHF316705 Rev12 |  | Subsystem | ECM;Blocker | Console-SRS.16005 | Fixed | roy_r | 0 | 2 | HOST SSRS | Console4 | NO | Closed | BB_WIP8_SubCom | Brilliance Big Bore |  | DHF316705 Rev12 | August 1, 2019 6:18:10 PM GMT+08:00
CTUDT00624912 | Refresh | Closed | wang_l | July 8, 2019 4:48:24 PM GMT+08:00 | Refresh: ACR calibration failed for item - Audit Head | WIP17 |  |  | Subsystem | Blocker |  | Duplicate | frider_o | 0 |  |  | DMS | NO | Closed | WIP18SegReq | Ingenuity 64
CTUDT00624910 | Refresh | Closed | zhu_je | July 8, 2019 4:35:51 PM GMT+08:00 | [Refresh] There is system error if click repeat or next after finish perfusion scan | WIP16 | WIP19 |  | Subsystem | System | None | Fixed | deshbhratar_m | 0 | 3 | Console Software | Console4 | YES | Closed | WIP19SubCom | Ovation |  | WIP19 | July 19, 2019 2:27:34 PM GMT+08:00
CTUDT00624910 | Refresh | Closed | zhu_je | July 8, 2019 4:35:51 PM GMT+08:00 | [Refresh] There is system error if click repeat or next after finish perfusion scan | WIP16 | WIP19 |  | Subsystem | System | None | Fixed | deshbhratar_m | 0 | 3 | Console Software | Console4 | YES | Closed | WIP19SubCom | Brilliance 64 |  | WIP19 | July 19, 2019 2:27:34 PM GMT+08:00
CTUDT00624910 | Refresh | Closed | zhu_je | July 8, 2019 4:35:51 PM GMT+08:00 | [Refresh] There is system error if click repeat or next after finish perfusion scan | WIP16 | WIP19 |  | Subsystem | System | None | Fixed | deshbhratar_m | 0 | 3 | Console Software | Console4 | YES | Closed | WIP19SubCom | Ingenuity 64 |  | WIP19 | July 19, 2019 2:27:34 PM GMT+08:00
CTUDT00624910 | Refresh | Closed | zhu_je | July 8, 2019 4:35:51 PM GMT+08:00 | [Refresh] There is system error if click repeat or next after finish perfusion scan | WIP16 | WIP19 |  | Subsystem | System | None | Fixed | deshbhratar_m | 0 | 3 | Console Software | Console4 | YES | Closed | WIP19SubCom | IQon1 |  | WIP19 | July 19, 2019 2:27:34 PM GMT+08:00
CTUDT00624908 | Refresh | Closed | lei_yang | July 8, 2019 4:20:13 PM GMT+08:00 | No host event log about “CardiacPulmoDevice : CRTC…” to certify different arrhythmia handling policies for coronary CTA and Gated CTA. | WIP16 | WIP18 |  | Subsystem | Log Viewer;Blocker | CT-NM.SRS.Modes.1806 | Works As Specified / Not A Defect | lei_yang | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP19SubCom | Ingenuity 64 |  | WIP18 | August 5, 2019 8:40:15 AM GMT+08:00
CTUDT00624906 | Refresh | Closed | lei_yang | July 8, 2019 2:18:33 PM GMT+08:00 | The time-stamp of one of the last images reconstructed was earlier than the time-stamp of the last stopped X-Ray shot. | WIP16 | Refer to Attachments |  | Subsystem | Step&Shoot;Blocker | NA | Fixed | lei_yang | 0 | 2 | Test Design | Support-System | NO | Closed | WIP19SegReq | Ingenuity 64 |  | Refer to Attachments | August 7, 2019 5:07:57 PM GMT+08:00
CTUDT00624904 | Refresh | Closed | ding_s | July 8, 2019 12:32:08 PM GMT+08:00 | Specific functionalities are in KGS basic list, cannot disable. | WIP16 | WIP19 |  | System | License;Blocker | CT-NM.SRS.ProductConfig.168,CT-NM.SRS.ProductConfig.182,CT-NM.SRS.ProductConfig.189 and CT-NM.SRS.ProductConfig.174 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP19SegReq | Ingenuity 64 |  | WIP19 | August 27, 2019 6:25:31 PM GMT+08:00
CTUDT00624899 | Refresh | Closed | alex_nihaev | July 7, 2019 8:36:06 PM GMT+08:00 | Refresh: Performance infant Test: H-Phantom: Infant Head 100kv case is graphically wrong zoomed | WIP17 | WIP19 | N/A | Subsystem | Performance Test | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP19SubCom | IQon1 |  | WIP19 | August 4, 2019 3:18:24 PM GMT+08:00
CTUDT00624895 | Refresh | Closed | g_di | July 5, 2019 8:05:32 PM GMT+08:00 | In Matrix field values are showing wrong when Automatic Matrix is unchecked. | WIP17 | WIP21 | N/A. This is documentation update | Subsystem | Requirement | Console-SRS.15385 | Fixed | datta_i | 0 | 3 | Console Software | Console4 | NO | Closed | SubCom | Ovation |  | WIP21 | September 18, 2019 8:14:57 PM GMT+08:00
CTUDT00624895 | Refresh | Closed | g_di | July 5, 2019 8:05:32 PM GMT+08:00 | In Matrix field values are showing wrong when Automatic Matrix is unchecked. | WIP17 | WIP21 | N/A. This is documentation update | Subsystem | Requirement | Console-SRS.15385 | Fixed | datta_i | 0 | 3 | Console Software | Console4 | NO | Closed | SubCom | Brilliance 64 |  | WIP21 | September 18, 2019 8:14:57 PM GMT+08:00
CTUDT00624895 | Refresh | Closed | g_di | July 5, 2019 8:05:32 PM GMT+08:00 | In Matrix field values are showing wrong when Automatic Matrix is unchecked. | WIP17 | WIP21 | N/A. This is documentation update | Subsystem | Requirement | Console-SRS.15385 | Fixed | datta_i | 0 | 3 | Console Software | Console4 | NO | Closed | SubCom | Ingenuity 64 |  | WIP21 | September 18, 2019 8:14:57 PM GMT+08:00
CTUDT00624893 | Refresh | Closed | g_di | July 5, 2019 7:14:12 PM GMT+08:00 | Increment value is changing in the slab viewer for MPR result | WIP17 | Host Verification Test Protocol DHF318051 Rev 19 |  | Subsystem | Test Design | NA | Fixed | jindal_k | 0 | 4 | Test Design | Console4 | NO | Closed | WIP20SegReq | Ovation |  | Host Verification Test Protocol DHF318051 Rev 19 | August 22, 2019 4:22:40 PM GMT+08:00
CTUDT00624893 | Refresh | Closed | g_di | July 5, 2019 7:14:12 PM GMT+08:00 | Increment value is changing in the slab viewer for MPR result | WIP17 | Host Verification Test Protocol DHF318051 Rev 19 |  | Subsystem | Test Design | NA | Fixed | jindal_k | 0 | 4 | Test Design | Console4 | NO | Closed | WIP20SegReq | Brilliance 64 |  | Host Verification Test Protocol DHF318051 Rev 19 | August 22, 2019 4:22:40 PM GMT+08:00
CTUDT00624893 | Refresh | Closed | g_di | July 5, 2019 7:14:12 PM GMT+08:00 | Increment value is changing in the slab viewer for MPR result | WIP17 | Host Verification Test Protocol DHF318051 Rev 19 |  | Subsystem | Test Design | NA | Fixed | jindal_k | 0 | 4 | Test Design | Console4 | NO | Closed | WIP20SegReq | Brilliance Big Bore |  | Host Verification Test Protocol DHF318051 Rev 19 | August 22, 2019 4:22:40 PM GMT+08:00
CTUDT00624893 | Refresh | Closed | g_di | July 5, 2019 7:14:12 PM GMT+08:00 | Increment value is changing in the slab viewer for MPR result | WIP17 | Host Verification Test Protocol DHF318051 Rev 19 |  | Subsystem | Test Design | NA | Fixed | jindal_k | 0 | 4 | Test Design | Console4 | NO | Closed | WIP20SegReq | Ingenuity 64 |  | Host Verification Test Protocol DHF318051 Rev 19 | August 22, 2019 4:22:40 PM GMT+08:00
CTUDT00624893 | Refresh | Closed | g_di | July 5, 2019 7:14:12 PM GMT+08:00 | Increment value is changing in the slab viewer for MPR result | WIP17 | Host Verification Test Protocol DHF318051 Rev 19 |  | Subsystem | Test Design | NA | Fixed | jindal_k | 0 | 4 | Test Design | Console4 | NO | Closed | WIP20SegReq | IQon1 |  | Host Verification Test Protocol DHF318051 Rev 19 | August 22, 2019 4:22:40 PM GMT+08:00
CTUDT00624890 | Refresh | Closed | lei_yang | July 5, 2019 3:23:13 PM GMT+08:00 | Breathing waveform actual latency was more than 200 milliseconds. | BB_WIP6 | Refer to Attachments |  | Subsystem | Respiratory;Blocker | NA | Fixed | lei_yang | 0 | 3 | Verification Documentation | Support-System | NO | Closed | BB_WIP8_SegReq | Brilliance Big Bore |  | Refer to Attachments | August 22, 2019 5:00:27 PM GMT+08:00
CTUDT00624888 | Refresh | Closed | lei_yang | July 5, 2019 3:16:48 PM GMT+08:00 | Lost "jog" word in the message when the number of groups exceeds the limit. | WIP16 |  |  | Subsystem | Logging | Console-SRS.8792 | Works As Specified / Not A Defect | sagar_v | 0 |  |  | Console4 | NO | Closed | WAS | Ovation
CTUDT00624888 | Refresh | Closed | lei_yang | July 5, 2019 3:16:48 PM GMT+08:00 | Lost "jog" word in the message when the number of groups exceeds the limit. | WIP16 |  |  | Subsystem | Logging | Console-SRS.8792 | Works As Specified / Not A Defect | sagar_v | 0 |  |  | Console4 | NO | Closed | WAS | Brilliance 64
CTUDT00624888 | Refresh | Closed | lei_yang | July 5, 2019 3:16:48 PM GMT+08:00 | Lost "jog" word in the message when the number of groups exceeds the limit. | WIP16 |  |  | Subsystem | Logging | Console-SRS.8792 | Works As Specified / Not A Defect | sagar_v | 0 |  |  | Console4 | NO | Closed | WAS | Brilliance Big Bore
CTUDT00624888 | Refresh | Closed | lei_yang | July 5, 2019 3:16:48 PM GMT+08:00 | Lost "jog" word in the message when the number of groups exceeds the limit. | WIP16 |  |  | Subsystem | Logging | Console-SRS.8792 | Works As Specified / Not A Defect | sagar_v | 0 |  |  | Console4 | NO | Closed | WAS | Ingenuity 64
CTUDT00624888 | Refresh | Closed | lei_yang | July 5, 2019 3:16:48 PM GMT+08:00 | Lost "jog" word in the message when the number of groups exceeds the limit. | WIP16 |  |  | Subsystem | Logging | Console-SRS.8792 | Works As Specified / Not A Defect | sagar_v | 0 |  |  | Console4 | NO | Closed | WAS | IQon1
CTUDT00624887 | Refresh | Closed | bhawna_g | July 5, 2019 12:41:32 PM GMT+08:00 | Respiratory Level of 'Show All' was not identical with the Respiratory Level of 'Scan Parameters' | 4.1.10 | WIP19 |  | Subsystem | Plan | None | Fixed | wilsonraja | 0 | 2 | Console Software | Console4 | YES | Closed | WIP18SubCom | Brilliance 64 |  | WIP19 | July 12, 2019 4:24:36 PM GMT+08:00
CTUDT00624887 | Refresh | Closed | bhawna_g | July 5, 2019 12:41:32 PM GMT+08:00 | Respiratory Level of 'Show All' was not identical with the Respiratory Level of 'Scan Parameters' | 4.1.10 | WIP19 |  | Subsystem | Plan | None | Fixed | wilsonraja | 0 | 2 | Console Software | Console4 | YES | Closed | WIP18SubCom | Ingenuity 64 |  | WIP19 | July 12, 2019 4:24:36 PM GMT+08:00
CTUDT00624884 | Refresh | Closed | 320053769 | July 4, 2019 10:11:53 PM GMT+08:00 | Refresh - Performance issue, login to CT user takes more than 30 seconds | WIP17 |  |  | Subsystem | Login | N/A | Works As Specified / Not A Defect | tamari_o | 1 |  |  | Support-System | NO | Closed |  | IQon1
CTUDT00624881 | Refresh | Closed | wei_f | July 4, 2019 8:28:49 PM GMT+08:00 | cannot access to Device Manager from Olevel Service Tools | BB_WIP7 | DHF298554 Rev06 |  | Subsystem |  | Service.3003 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP18SubCom | Ovation |  | DHF298554 Rev06 | July 25, 2019 4:41:12 PM GMT+08:00
CTUDT00624881 | Refresh | Closed | wei_f | July 4, 2019 8:28:49 PM GMT+08:00 | cannot access to Device Manager from Olevel Service Tools | BB_WIP7 | DHF298554 Rev06 |  | Subsystem |  | Service.3003 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP18SubCom | Brilliance 64 |  | DHF298554 Rev06 | July 25, 2019 4:41:12 PM GMT+08:00
CTUDT00624881 | Refresh | Closed | wei_f | July 4, 2019 8:28:49 PM GMT+08:00 | cannot access to Device Manager from Olevel Service Tools | BB_WIP7 | DHF298554 Rev06 |  | Subsystem |  | Service.3003 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP18SubCom | Brilliance Big Bore |  | DHF298554 Rev06 | July 25, 2019 4:41:12 PM GMT+08:00
CTUDT00624881 | Refresh | Closed | wei_f | July 4, 2019 8:28:49 PM GMT+08:00 | cannot access to Device Manager from Olevel Service Tools | BB_WIP7 | DHF298554 Rev06 |  | Subsystem |  | Service.3003 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP18SubCom | Ingenuity 64 |  | DHF298554 Rev06 | July 25, 2019 4:41:12 PM GMT+08:00
CTUDT00624881 | Refresh | Closed | wei_f | July 4, 2019 8:28:49 PM GMT+08:00 | cannot access to Device Manager from Olevel Service Tools | BB_WIP7 | DHF298554 Rev06 |  | Subsystem |  | Service.3003 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP18SubCom | IQon1 |  | DHF298554 Rev06 | July 25, 2019 4:41:12 PM GMT+08:00
CTUDT00624878 | Refresh | Closed | wei_f | July 4, 2019 7:45:20 PM GMT+08:00 | The console did not display blanks after idle for 20 minutes. | WIP16 |  |  | Subsystem | Screen Blanking |  | Works As Specified / Not A Defect | srushti_k | 1 |  |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00624875 | Refresh | Closed | wei_f | July 4, 2019 6:50:45 PM GMT+08:00 | Value for the some tags did not remain same after de-identification. | WIP16 |  |  | Subsystem | Test Design | NA | Works As Specified / Not A Defect | wei_f | 0 | 4 | Test Design | Support-System | NO | Closed | WIP19SegReq | Ovation |  |  | August 12, 2019 4:23:50 PM GMT+08:00
CTUDT00624875 | Refresh | Closed | wei_f | July 4, 2019 6:50:45 PM GMT+08:00 | Value for the some tags did not remain same after de-identification. | WIP16 |  |  | Subsystem | Test Design | NA | Works As Specified / Not A Defect | wei_f | 0 | 4 | Test Design | Support-System | NO | Closed | WIP19SegReq | Brilliance 64 |  |  | August 12, 2019 4:23:50 PM GMT+08:00
CTUDT00624875 | Refresh | Closed | wei_f | July 4, 2019 6:50:45 PM GMT+08:00 | Value for the some tags did not remain same after de-identification. | WIP16 |  |  | Subsystem | Test Design | NA | Works As Specified / Not A Defect | wei_f | 0 | 4 | Test Design | Support-System | NO | Closed | WIP19SegReq | Brilliance Big Bore |  |  | August 12, 2019 4:23:50 PM GMT+08:00
CTUDT00624875 | Refresh | Closed | wei_f | July 4, 2019 6:50:45 PM GMT+08:00 | Value for the some tags did not remain same after de-identification. | WIP16 |  |  | Subsystem | Test Design | NA | Works As Specified / Not A Defect | wei_f | 0 | 4 | Test Design | Support-System | NO | Closed | WIP19SegReq | Ingenuity 64 |  |  | August 12, 2019 4:23:50 PM GMT+08:00
CTUDT00624875 | Refresh | Closed | wei_f | July 4, 2019 6:50:45 PM GMT+08:00 | Value for the some tags did not remain same after de-identification. | WIP16 |  |  | Subsystem | Test Design | NA | Works As Specified / Not A Defect | wei_f | 0 | 4 | Test Design | Support-System | NO | Closed | WIP19SegReq | IQon1 |  |  | August 12, 2019 4:23:50 PM GMT+08:00
CTUDT00624874 | Refresh | Closed | wei_f | July 4, 2019 6:40:22 PM GMT+08:00 | "0010,1010 PATIENTS_AGE" tag value in Anonymized study was [ ] years after anonymizing the study. | WIP16 |  |  | Subsystem | Test Design | NA | Works As Specified / Not A Defect | wei_f | 0 | 4 | Test Design | Support-System | NO | Closed | WIP19SegReq | Ovation |  |  | August 12, 2019 4:22:33 PM GMT+08:00
CTUDT00624874 | Refresh | Closed | wei_f | July 4, 2019 6:40:22 PM GMT+08:00 | "0010,1010 PATIENTS_AGE" tag value in Anonymized study was [ ] years after anonymizing the study. | WIP16 |  |  | Subsystem | Test Design | NA | Works As Specified / Not A Defect | wei_f | 0 | 4 | Test Design | Support-System | NO | Closed | WIP19SegReq | Brilliance 64 |  |  | August 12, 2019 4:22:33 PM GMT+08:00
CTUDT00624874 | Refresh | Closed | wei_f | July 4, 2019 6:40:22 PM GMT+08:00 | "0010,1010 PATIENTS_AGE" tag value in Anonymized study was [ ] years after anonymizing the study. | WIP16 |  |  | Subsystem | Test Design | NA | Works As Specified / Not A Defect | wei_f | 0 | 4 | Test Design | Support-System | NO | Closed | WIP19SegReq | Brilliance Big Bore |  |  | August 12, 2019 4:22:33 PM GMT+08:00
CTUDT00624874 | Refresh | Closed | wei_f | July 4, 2019 6:40:22 PM GMT+08:00 | "0010,1010 PATIENTS_AGE" tag value in Anonymized study was [ ] years after anonymizing the study. | WIP16 |  |  | Subsystem | Test Design | NA | Works As Specified / Not A Defect | wei_f | 0 | 4 | Test Design | Support-System | NO | Closed | WIP19SegReq | Ingenuity 64 |  |  | August 12, 2019 4:22:33 PM GMT+08:00
CTUDT00624874 | Refresh | Closed | wei_f | July 4, 2019 6:40:22 PM GMT+08:00 | "0010,1010 PATIENTS_AGE" tag value in Anonymized study was [ ] years after anonymizing the study. | WIP16 |  |  | Subsystem | Test Design | NA | Works As Specified / Not A Defect | wei_f | 0 | 4 | Test Design | Support-System | NO | Closed | WIP19SegReq | IQon1 |  |  | August 12, 2019 4:22:33 PM GMT+08:00
CTUDT00624869 | Refresh | Closed | sun_q | July 4, 2019 11:13:45 AM GMT+08:00 | The Warning Symbol ISO7010 - W001 was missed on several translation of Hazardous Voltage Warning label | WIP16 | BB_WIP9 |  | Subsystem | Labels | GCX.SSRS.iCT.19391, GCX.SSRS.iCT.19392 | Fixed | zhu_l | 0 | 4 | Design Documentation | GCX | NO | Closed | WIP19SubCom | Brilliance 64 |  | BB_WIP9 | August 14, 2019 10:35:38 AM GMT+08:00
CTUDT00624869 | Refresh | Closed | sun_q | July 4, 2019 11:13:45 AM GMT+08:00 | The Warning Symbol ISO7010 - W001 was missed on several translation of Hazardous Voltage Warning label | WIP16 | BB_WIP9 |  | Subsystem | Labels | GCX.SSRS.iCT.19391, GCX.SSRS.iCT.19392 | Fixed | zhu_l | 0 | 4 | Design Documentation | GCX | NO | Closed | WIP19SubCom | Ingenuity 64 |  | BB_WIP9 | August 14, 2019 10:35:38 AM GMT+08:00
CTUDT00624867 | Refresh | Closed | sshahar | July 3, 2019 11:13:35 PM GMT+08:00 | Refresh/Wip17/Calibration/Tube alignment calibration failed | WIP17 |  |  | Subsystem | Blocker | n/a | Works As Specified / Not A Defect | zhou_x | 0 |  |  | GCX | NO | Closed |  | Ovation
CTUDT00624864 | Refresh | Closed | manjunath_c | July 3, 2019 5:28:11 PM GMT+08:00 | Cut option was not present in edited results context menu | WIP17 | WIP19 | Copy and paste and then delete. | Subsystem | System | NA | Fixed | prasad_r | 0 | 3 | Console Software | Console4 | NO | Closed | WIP19SubCom | Ovation |  | WIP19 | July 15, 2019 8:14:14 PM GMT+08:00
CTUDT00624864 | Refresh | Closed | manjunath_c | July 3, 2019 5:28:11 PM GMT+08:00 | Cut option was not present in edited results context menu | WIP17 | WIP19 | Copy and paste and then delete. | Subsystem | System | NA | Fixed | prasad_r | 0 | 3 | Console Software | Console4 | NO | Closed | WIP19SubCom | Brilliance 64 |  | WIP19 | July 15, 2019 8:14:14 PM GMT+08:00
CTUDT00624864 | Refresh | Closed | manjunath_c | July 3, 2019 5:28:11 PM GMT+08:00 | Cut option was not present in edited results context menu | WIP17 | WIP19 | Copy and paste and then delete. | Subsystem | System | NA | Fixed | prasad_r | 0 | 3 | Console Software | Console4 | NO | Closed | WIP19SubCom | Ingenuity 64 |  | WIP19 | July 15, 2019 8:14:14 PM GMT+08:00
CTUDT00624864 | Refresh | Closed | manjunath_c | July 3, 2019 5:28:11 PM GMT+08:00 | Cut option was not present in edited results context menu | WIP17 | WIP19 | Copy and paste and then delete. | Subsystem | System | NA | Fixed | prasad_r | 0 | 3 | Console Software | Console4 | NO | Closed | WIP19SubCom | IQon1 |  | WIP19 | July 15, 2019 8:14:14 PM GMT+08:00
CTUDT00624861 | Refresh | Closed | sun_q | July 3, 2019 3:17:29 PM GMT+08:00 | For E-stop button of White Cover Front/Rear Control Panel, the color is not match with Red 20000 | WIP16 | DHF317363 rev03 DHF317353 rev03 |  | Subsystem | Covers | GCX.SSRS.iCT.20318;GCX.SSRS.iCT.20727 | Fixed | wang_k | 0 | 2 | Design Documentation | GCX | NO | Closed | WIP18SubCom | Brilliance Big Bore |  | DHF317363 rev03 DHF317353 rev03 | September 5, 2019 5:00:51 PM GMT+08:00
CTUDT00624861 | Refresh | Closed | sun_q | July 3, 2019 3:17:29 PM GMT+08:00 | For E-stop button of White Cover Front/Rear Control Panel, the color is not match with Red 20000 | WIP16 | DHF317363 rev03 DHF317353 rev03 |  | Subsystem | Covers | GCX.SSRS.iCT.20318;GCX.SSRS.iCT.20727 | Fixed | wang_k | 0 | 2 | Design Documentation | GCX | NO | Closed | WIP18SubCom | Ingenuity 64 |  | DHF317363 rev03 DHF317353 rev03 | September 5, 2019 5:00:51 PM GMT+08:00
CTUDT00624860 | Refresh | Closed | bhawna_g | July 3, 2019 1:07:29 PM GMT+08:00 | An error message appeared “system error. Please end the exam and retry” while performing offline recon when the last name is null | 4.8 | BB_WIP8 | Manually filling Last name in service mode will not cause this issue. | Subsystem |  | NA | Fixed | bhawna_g | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP8_SubCom | Brilliance Big Bore |  | BB_WIP8 | July 4, 2019 4:33:24 PM GMT+08:00
CTUDT00624855 | Refresh | Closed | v_sumanth | July 3, 2019 12:35:52 PM GMT+08:00 | [Refresh] "Rotor_Position_Req" is not appeared in Cantrace after load an EC to scan if setting "Enable Fast Get Ready" | 4.8 | BB_WIP9 |  | System |  | None | Fixed | li_e | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP8_SubCom | Brilliance Big Bore |  | BB_WIP9 | July 23, 2019 3:13:38 PM GMT+08:00
CTUDT00624844 | Refresh | Closed | Sharmila_Vasireddy | July 2, 2019 8:56:33 PM GMT+08:00 | BB refresh: Not able to login with user created account | BB_WIP7 |  | Reset user name and password only with supported symbols. | Subsystem | Login | NA | No Plan to Fix | kapadia_v | 0 | 4 |  | Console4 | NO | Closed | WAS | Brilliance Big Bore
CTUDT00624842 | Refresh | Closed | Sharmila_Vasireddy | July 2, 2019 8:39:00 PM GMT+08:00 | Displaying Wrong user name in the pop up message while creating user with special characters (c%&) | BB_WIP7 |  | NA. | Subsystem | Login | NA | No Plan to Fix | kapadia_v | 0 | 4 |  | Console4 | NO | Closed |  | Ovation
CTUDT00624842 | Refresh | Closed | Sharmila_Vasireddy | July 2, 2019 8:39:00 PM GMT+08:00 | Displaying Wrong user name in the pop up message while creating user with special characters (c%&) | BB_WIP7 |  | NA. | Subsystem | Login | NA | No Plan to Fix | kapadia_v | 0 | 4 |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00624842 | Refresh | Closed | Sharmila_Vasireddy | July 2, 2019 8:39:00 PM GMT+08:00 | Displaying Wrong user name in the pop up message while creating user with special characters (c%&) | BB_WIP7 |  | NA. | Subsystem | Login | NA | No Plan to Fix | kapadia_v | 0 | 4 |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00624842 | Refresh | Closed | Sharmila_Vasireddy | July 2, 2019 8:39:00 PM GMT+08:00 | Displaying Wrong user name in the pop up message while creating user with special characters (c%&) | BB_WIP7 |  | NA. | Subsystem | Login | NA | No Plan to Fix | kapadia_v | 0 | 4 |  | Console4 | NO | Closed |  | IQon1
CTUDT00624831 | Refresh | Closed | zhu_je | July 2, 2019 11:08:54 AM GMT+08:00 | [Refresh] An error ErrorExamCardManager popped up after clicking ‘Print&Text’ button in sharing tab in ECM | WIP16 | WIP20 | Adding license of 0.4 solves the issue | System | ECM | NA | Fixed | deshbhratar_m | 0 | 3 | Console Software | Console4 | NO | Closed | WIP19SubCom | Brilliance 64 |  | WIP20 | August 22, 2019 7:52:38 PM GMT+08:00
CTUDT00624831 | Refresh | Closed | zhu_je | July 2, 2019 11:08:54 AM GMT+08:00 | [Refresh] An error ErrorExamCardManager popped up after clicking ‘Print&Text’ button in sharing tab in ECM | WIP16 | WIP20 | Adding license of 0.4 solves the issue | System | ECM | NA | Fixed | deshbhratar_m | 0 | 3 | Console Software | Console4 | NO | Closed | WIP19SubCom | Ingenuity 64 |  | WIP20 | August 22, 2019 7:52:38 PM GMT+08:00
CTUDT00624829 | Refresh | Closed | zhu_je | July 2, 2019 10:58:50 AM GMT+08:00 | [Refresh] In ECM->EC sharing tab. The selected ECs (with blue color) will be not selected (without blue color) if fold up and then spread out the group | WIP16 |  | Re-select the required exam card. | Subsystem | ECM | NA | No Plan to Fix | kapadia_v | 0 | 4 |  | Console4 | NO | Closed |  | Ovation
CTUDT00624829 | Refresh | Closed | zhu_je | July 2, 2019 10:58:50 AM GMT+08:00 | [Refresh] In ECM->EC sharing tab. The selected ECs (with blue color) will be not selected (without blue color) if fold up and then spread out the group | WIP16 |  | Re-select the required exam card. | Subsystem | ECM | NA | No Plan to Fix | kapadia_v | 0 | 4 |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00624829 | Refresh | Closed | zhu_je | July 2, 2019 10:58:50 AM GMT+08:00 | [Refresh] In ECM->EC sharing tab. The selected ECs (with blue color) will be not selected (without blue color) if fold up and then spread out the group | WIP16 |  | Re-select the required exam card. | Subsystem | ECM | NA | No Plan to Fix | kapadia_v | 0 | 4 |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00624829 | Refresh | Closed | zhu_je | July 2, 2019 10:58:50 AM GMT+08:00 | [Refresh] In ECM->EC sharing tab. The selected ECs (with blue color) will be not selected (without blue color) if fold up and then spread out the group | WIP16 |  | Re-select the required exam card. | Subsystem | ECM | NA | No Plan to Fix | kapadia_v | 0 | 4 |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00624829 | Refresh | Closed | zhu_je | July 2, 2019 10:58:50 AM GMT+08:00 | [Refresh] In ECM->EC sharing tab. The selected ECs (with blue color) will be not selected (without blue color) if fold up and then spread out the group | WIP16 |  | Re-select the required exam card. | Subsystem | ECM | NA | No Plan to Fix | kapadia_v | 0 | 4 |  | Console4 | NO | Closed |  | IQon1
CTUDT00624827 | Refresh | Closed | zhu_je | July 2, 2019 10:12:39 AM GMT+08:00 | [Refresh] if right click mouse on any EC in ECM, it cannot be choose any EC then | WIP16 |  |  | Subsystem | ECM | CT-NM.SRS.Workflow.258 | No Plan to Fix | prasad_r | 0 | 4 |  | Console4 | NO | Closed | WIP19SubCom | Ovation
CTUDT00624827 | Refresh | Closed | zhu_je | July 2, 2019 10:12:39 AM GMT+08:00 | [Refresh] if right click mouse on any EC in ECM, it cannot be choose any EC then | WIP16 |  |  | Subsystem | ECM | CT-NM.SRS.Workflow.258 | No Plan to Fix | prasad_r | 0 | 4 |  | Console4 | NO | Closed | WIP19SubCom | Brilliance 64
CTUDT00624827 | Refresh | Closed | zhu_je | July 2, 2019 10:12:39 AM GMT+08:00 | [Refresh] if right click mouse on any EC in ECM, it cannot be choose any EC then | WIP16 |  |  | Subsystem | ECM | CT-NM.SRS.Workflow.258 | No Plan to Fix | prasad_r | 0 | 4 |  | Console4 | NO | Closed | WIP19SubCom | Brilliance Big Bore
CTUDT00624827 | Refresh | Closed | zhu_je | July 2, 2019 10:12:39 AM GMT+08:00 | [Refresh] if right click mouse on any EC in ECM, it cannot be choose any EC then | WIP16 |  |  | Subsystem | ECM | CT-NM.SRS.Workflow.258 | No Plan to Fix | prasad_r | 0 | 4 |  | Console4 | NO | Closed | WIP19SubCom | Ingenuity 64
CTUDT00624827 | Refresh | Closed | zhu_je | July 2, 2019 10:12:39 AM GMT+08:00 | [Refresh] if right click mouse on any EC in ECM, it cannot be choose any EC then | WIP16 |  |  | Subsystem | ECM | CT-NM.SRS.Workflow.258 | No Plan to Fix | prasad_r | 0 | 4 |  | Console4 | NO | Closed | WIP19SubCom | IQon1
CTUDT00624825 | Refresh | Closed | 320053768 | July 1, 2019 8:21:36 PM GMT+08:00 | An error message appeared “system error. Please end the exam and retry” while performing offline recon when the last name is null | WIP17 | WIP18 | Manually filling Last name in service mode will not cause this issue. | Subsystem | Reconstruction | NA | Fixed | bhawna_g | 0 | 3 | Console Software | Console4 | NO | Closed | WIP19SubCom | Ovation |  | WIP18 | July 4, 2019 1:56:46 PM GMT+08:00
CTUDT00624825 | Refresh | Closed | 320053768 | July 1, 2019 8:21:36 PM GMT+08:00 | An error message appeared “system error. Please end the exam and retry” while performing offline recon when the last name is null | WIP17 | WIP18 | Manually filling Last name in service mode will not cause this issue. | Subsystem | Reconstruction | NA | Fixed | bhawna_g | 0 | 3 | Console Software | Console4 | NO | Closed | WIP19SubCom | Brilliance 64 |  | WIP18 | July 4, 2019 1:56:46 PM GMT+08:00
CTUDT00624825 | Refresh | Closed | 320053768 | July 1, 2019 8:21:36 PM GMT+08:00 | An error message appeared “system error. Please end the exam and retry” while performing offline recon when the last name is null | WIP17 | WIP18 | Manually filling Last name in service mode will not cause this issue. | Subsystem | Reconstruction | NA | Fixed | bhawna_g | 0 | 3 | Console Software | Console4 | NO | Closed | WIP19SubCom | Ingenuity 64 |  | WIP18 | July 4, 2019 1:56:46 PM GMT+08:00
CTUDT00624825 | Refresh | Closed | 320053768 | July 1, 2019 8:21:36 PM GMT+08:00 | An error message appeared “system error. Please end the exam and retry” while performing offline recon when the last name is null | WIP17 | WIP18 | Manually filling Last name in service mode will not cause this issue. | Subsystem | Reconstruction | NA | Fixed | bhawna_g | 0 | 3 | Console Software | Console4 | NO | Closed | WIP19SubCom | IQon1 |  | WIP18 | July 4, 2019 1:56:46 PM GMT+08:00
CTUDT00624823 | Refresh | Closed | 320053769 | July 1, 2019 6:15:06 PM GMT+08:00 | Service tools show different couch type than the right ones via one-click installation. | WIP17 | WIP19 |  | Subsystem | Service Tools | None | Fixed | Prakash_Chand_Yadav | 0 | 3 | Console Software | Service Tools | NO | Closed | WIP19SubCom | Ovation |  | WIP19 | July 19, 2019 5:02:12 PM GMT+08:00
CTUDT00624823 | Refresh | Closed | 320053769 | July 1, 2019 6:15:06 PM GMT+08:00 | Service tools show different couch type than the right ones via one-click installation. | WIP17 | WIP19 |  | Subsystem | Service Tools | None | Fixed | Prakash_Chand_Yadav | 0 | 3 | Console Software | Service Tools | NO | Closed | WIP19SubCom | Brilliance 64 |  | WIP19 | July 19, 2019 5:02:12 PM GMT+08:00
CTUDT00624823 | Refresh | Closed | 320053769 | July 1, 2019 6:15:06 PM GMT+08:00 | Service tools show different couch type than the right ones via one-click installation. | WIP17 | WIP19 |  | Subsystem | Service Tools | None | Fixed | Prakash_Chand_Yadav | 0 | 3 | Console Software | Service Tools | NO | Closed | WIP19SubCom | Ingenuity 64 |  | WIP19 | July 19, 2019 5:02:12 PM GMT+08:00
CTUDT00624823 | Refresh | Closed | 320053769 | July 1, 2019 6:15:06 PM GMT+08:00 | Service tools show different couch type than the right ones via one-click installation. | WIP17 | WIP19 |  | Subsystem | Service Tools | None | Fixed | Prakash_Chand_Yadav | 0 | 3 | Console Software | Service Tools | NO | Closed | WIP19SubCom | IQon1 |  | WIP19 | July 19, 2019 5:02:12 PM GMT+08:00
CTUDT00624817 | Refresh | Closed | alex_nihaev | June 30, 2019 9:15:26 PM GMT+08:00 | Refresh: Radiation Drift Test: Harmonized Holder: Test can not be completed because of wrong Bed position | WIP17 | WIP18 |  | Subsystem | DMS Calibrations | NaN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP18SegReq | IQon1 |  | WIP18 | July 7, 2019 10:11:17 PM GMT+08:00
CTUDT00624814 | Refresh | Closed | wang_l | June 30, 2019 5:51:49 PM GMT+08:00 | Refresh: Cannot register remote device after software package was downloaded from PRS server | WIP17 |  | NA | Subsystem | Blocker | NA | Invalid/Uncontrolled Configuration | wang_l | 0 |  |  | Support-System | NO | Closed |  | Brilliance Big Bore
CTUDT00624814 | Refresh | Closed | wang_l | June 30, 2019 5:51:49 PM GMT+08:00 | Refresh: Cannot register remote device after software package was downloaded from PRS server | WIP17 |  | NA | Subsystem | Blocker | NA | Invalid/Uncontrolled Configuration | wang_l | 0 |  |  | Support-System | NO | Closed |  | Ingenuity 64
CTUDT00624812 | Refresh | Closed | sun_q | June 30, 2019 5:01:26 PM GMT+08:00 | In Quality Assurance, all calibration successfully message is not displayed in History->Message | WIP17 |  |  | Subsystem | Calibrations | N/a | Works As Specified / Not A Defect | frider_o | 0 |  |  | DMS | NO | Closed |  | Brilliance 64
CTUDT00624812 | Refresh | Closed | sun_q | June 30, 2019 5:01:26 PM GMT+08:00 | In Quality Assurance, all calibration successfully message is not displayed in History->Message | WIP17 |  |  | Subsystem | Calibrations | N/a | Works As Specified / Not A Defect | frider_o | 0 |  |  | DMS | NO | Closed |  | Brilliance Big Bore
CTUDT00624812 | Refresh | Closed | sun_q | June 30, 2019 5:01:26 PM GMT+08:00 | In Quality Assurance, all calibration successfully message is not displayed in History->Message | WIP17 |  |  | Subsystem | Calibrations | N/a | Works As Specified / Not A Defect | frider_o | 0 |  |  | DMS | NO | Closed |  | Ingenuity 64
CTUDT00624810 | Refresh | Closed | azagury_d | June 30, 2019 4:15:59 PM GMT+08:00 | Refresh:Calibrations:ACR Test:In ACR test pitch for Adult abdomen scan was out of tolerance. | WIP17 | WIP18 |  | Subsystem | ACR Test | n/a | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP18SubCom | Ingenuity 64 |  | WIP18 | July 8, 2019 9:59:24 PM GMT+08:00
CTUDT00624808 | Refresh | Closed | azagury_d | June 30, 2019 1:29:10 PM GMT+08:00 | Refresh: Calibrations: Beam and DMS Check: Issues with Plot limits | BB_WIP7 | WIP18 |  | Subsystem | Beam and DMS Check; Verified1 | na | Fixed | Kostya_Baklan | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP18SubCom | Ovation |  | WIP18 | July 7, 2019 6:14:20 PM GMT+08:00
CTUDT00624808 | Refresh | Closed | azagury_d | June 30, 2019 1:29:10 PM GMT+08:00 | Refresh: Calibrations: Beam and DMS Check: Issues with Plot limits | BB_WIP7 | WIP18 |  | Subsystem | Beam and DMS Check; Verified1 | na | Fixed | Kostya_Baklan | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP18SubCom | Brilliance Big Bore |  | WIP18 | July 7, 2019 6:14:20 PM GMT+08:00
CTUDT00624808 | Refresh | Closed | azagury_d | June 30, 2019 1:29:10 PM GMT+08:00 | Refresh: Calibrations: Beam and DMS Check: Issues with Plot limits | BB_WIP7 | WIP18 |  | Subsystem | Beam and DMS Check; Verified1 | na | Fixed | Kostya_Baklan | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP18SubCom | Ingenuity 64 |  | WIP18 | July 7, 2019 6:14:20 PM GMT+08:00
CTUDT00624808 | Refresh | Closed | azagury_d | June 30, 2019 1:29:10 PM GMT+08:00 | Refresh: Calibrations: Beam and DMS Check: Issues with Plot limits | BB_WIP7 | WIP18 |  | Subsystem | Beam and DMS Check; Verified1 | na | Fixed | Kostya_Baklan | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP18SubCom | IQon1 |  | WIP18 | July 7, 2019 6:14:20 PM GMT+08:00
CTUDT00624806 | Refresh | Closed | azagury_d | June 30, 2019 1:15:45 PM GMT+08:00 | Refresh: Calibrations: ACR Test: Mismatch in tube data between HTML report and Service tools | BB_WIP7 | N/A |  | Subsystem | ACR Test | n/a | Fixed | yosfan_s | 0 | 3 | Test Design | DMS | NO | Closed | WIP18SubCom | Brilliance Big Bore |  | N/A | July 14, 2019 2:29:35 PM GMT+08:00
CTUDT00624796 | Refresh | Closed | cai_t | June 27, 2019 9:26:00 PM GMT+08:00 | Refresh:Monitor Calibration:Clinical Refrence image should not display "Thickness 3.00 mm Philips, Brilliance Big Bore" | BB_WIP6 | BB_WIP9 |  | Subsystem | Monitor Calibration;Blocker | NA | Fixed | pandey_a | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP9_SubCom | Brilliance Big Bore |  | BB_WIP9 | August 12, 2019 4:19:31 PM GMT+08:00
CTUDT00624795 | Refresh | Closed | wei_f | June 27, 2019 9:10:56 PM GMT+08:00 | The DVD is ejected while the CCT and Bolus Tracking scans is in progress. | BB_WIP6 | WIP19 |  | Subsystem | Test Design | NA | Fixed | s_sa | 0 | 3 | Console USP_EBW | iCAP | YES | Closed | WIP19SegReq | Ovation |  | WIP19 | October 9, 2019 12:35:26 PM GMT+08:00
CTUDT00624795 | Refresh | Closed | wei_f | June 27, 2019 9:10:56 PM GMT+08:00 | The DVD is ejected while the CCT and Bolus Tracking scans is in progress. | BB_WIP6 | WIP19 |  | Subsystem | Test Design | NA | Fixed | s_sa | 0 | 3 | Console USP_EBW | iCAP | YES | Closed | WIP19SegReq | Brilliance 64 |  | WIP19 | October 9, 2019 12:35:26 PM GMT+08:00
CTUDT00624795 | Refresh | Closed | wei_f | June 27, 2019 9:10:56 PM GMT+08:00 | The DVD is ejected while the CCT and Bolus Tracking scans is in progress. | BB_WIP6 | WIP19 |  | Subsystem | Test Design | NA | Fixed | s_sa | 0 | 3 | Console USP_EBW | iCAP | YES | Closed | WIP19SegReq | Brilliance Big Bore |  | WIP19 | October 9, 2019 12:35:26 PM GMT+08:00
CTUDT00624795 | Refresh | Closed | wei_f | June 27, 2019 9:10:56 PM GMT+08:00 | The DVD is ejected while the CCT and Bolus Tracking scans is in progress. | BB_WIP6 | WIP19 |  | Subsystem | Test Design | NA | Fixed | s_sa | 0 | 3 | Console USP_EBW | iCAP | YES | Closed | WIP19SegReq | Ingenuity 64 |  | WIP19 | October 9, 2019 12:35:26 PM GMT+08:00
CTUDT00624795 | Refresh | Closed | wei_f | June 27, 2019 9:10:56 PM GMT+08:00 | The DVD is ejected while the CCT and Bolus Tracking scans is in progress. | BB_WIP6 | WIP19 |  | Subsystem | Test Design | NA | Fixed | s_sa | 0 | 3 | Console USP_EBW | iCAP | YES | Closed | WIP19SegReq | IQon1 |  | WIP19 | October 9, 2019 12:35:26 PM GMT+08:00
CTUDT00624793 | Refresh | Closed | alex_nihaev | June 27, 2019 6:54:08 PM GMT+08:00 | Refresh: Performance Test: Two part Phantom: Body CT4d ROI position is not as defined in Spec | WIP17 | WIP19 | N/A | Subsystem | Performance Test | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP19SubCom | IQon1 |  | WIP19 | August 7, 2019 5:38:14 PM GMT+08:00
CTUDT00624790 | Refresh | Closed | cai_t | June 27, 2019 4:35:21 PM GMT+08:00 | Refresh:HPZ4: No sound heard at the CT Box speaker from voice manager. | V4.8.0.10119 | WIP 21 | N/A | Subsystem | Voice Manager | N/A | Fixed | navale_s | 0 | 2 | Console Software | Console4 | NO | Closed |  | Ovation |  | WIP 21 | October 17, 2019 6:36:40 PM GMT+08:00
CTUDT00624790 | Refresh | Closed | cai_t | June 27, 2019 4:35:21 PM GMT+08:00 | Refresh:HPZ4: No sound heard at the CT Box speaker from voice manager. | V4.8.0.10119 | WIP 21 | N/A | Subsystem | Voice Manager | N/A | Fixed | navale_s | 0 | 2 | Console Software | Console4 | NO | Closed |  | Brilliance 64 |  | WIP 21 | October 17, 2019 6:36:40 PM GMT+08:00
CTUDT00624790 | Refresh | Closed | cai_t | June 27, 2019 4:35:21 PM GMT+08:00 | Refresh:HPZ4: No sound heard at the CT Box speaker from voice manager. | V4.8.0.10119 | WIP 21 | N/A | Subsystem | Voice Manager | N/A | Fixed | navale_s | 0 | 2 | Console Software | Console4 | NO | Closed |  | Ingenuity 64 |  | WIP 21 | October 17, 2019 6:36:40 PM GMT+08:00
CTUDT00624790 | Refresh | Closed | cai_t | June 27, 2019 4:35:21 PM GMT+08:00 | Refresh:HPZ4: No sound heard at the CT Box speaker from voice manager. | V4.8.0.10119 | WIP 21 | N/A | Subsystem | Voice Manager | N/A | Fixed | navale_s | 0 | 2 | Console Software | Console4 | NO | Closed |  | IQon1 |  | WIP 21 | October 17, 2019 6:36:40 PM GMT+08:00
CTUDT00624789 | Refresh | Closed | alex_nihaev | June 27, 2019 4:33:08 PM GMT+08:00 | Refresh: Constancy test: No indication in the HTML report of Constancy test regarding baseline that was used from Acceptance test | WIP17 | WIP18 | N/A | Subsystem | Acceptance/Constancy | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP18SegReq | Ovation |  | WIP18 | July 7, 2019 4:55:41 PM GMT+08:00
CTUDT00624789 | Refresh | Closed | alex_nihaev | June 27, 2019 4:33:08 PM GMT+08:00 | Refresh: Constancy test: No indication in the HTML report of Constancy test regarding baseline that was used from Acceptance test | WIP17 | WIP18 | N/A | Subsystem | Acceptance/Constancy | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP18SegReq | Brilliance Big Bore |  | WIP18 | July 7, 2019 4:55:41 PM GMT+08:00
CTUDT00624789 | Refresh | Closed | alex_nihaev | June 27, 2019 4:33:08 PM GMT+08:00 | Refresh: Constancy test: No indication in the HTML report of Constancy test regarding baseline that was used from Acceptance test | WIP17 | WIP18 | N/A | Subsystem | Acceptance/Constancy | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP18SegReq | Ingenuity 64 |  | WIP18 | July 7, 2019 4:55:41 PM GMT+08:00
CTUDT00624789 | Refresh | Closed | alex_nihaev | June 27, 2019 4:33:08 PM GMT+08:00 | Refresh: Constancy test: No indication in the HTML report of Constancy test regarding baseline that was used from Acceptance test | WIP17 | WIP18 | N/A | Subsystem | Acceptance/Constancy | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP18SegReq | IQon1 |  | WIP18 | July 7, 2019 4:55:41 PM GMT+08:00
CTUDT00624785 | Refresh | Closed | pengfei_cai | June 27, 2019 1:44:08 PM GMT+08:00 | The CCT Fluro continuously display "Scanning..." in the scan indication area after quick pedal press and release . | 4.2 |  | 1. Not to be quick in pressing/releasing of pedal. 2. if issue happened press pedal again. | Subsystem | CCT | Console-SRS.10391 | No Plan to Fix | veeraswamy_k | 0 | 4 |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00624782 | Refresh | Closed | pengfei_cai | June 27, 2019 11:11:17 AM GMT+08:00 | Operator tripped by patient strap | 4.1.7 | Refer to Resolution Description section | Remove the patient strap or put on the table is not used. | System |  | NM.SRS.Workflow.1619 | Fixed | zhou_r | 0 | 3 | Couch Mechanical | Couch | YES | Closed | SegReq | Ovation | PR#8300210 PR#8300684 PR#8300723 | Refer to Resolution Description section | December 16, 2019 6:10:59 PM GMT+08:00
CTUDT00624782 | Refresh | Closed | pengfei_cai | June 27, 2019 11:11:17 AM GMT+08:00 | Operator tripped by patient strap | 4.1.7 | Refer to Resolution Description section | Remove the patient strap or put on the table is not used. | System |  | NM.SRS.Workflow.1619 | Fixed | zhou_r | 0 | 3 | Couch Mechanical | Couch | YES | Closed | SegReq | Brilliance 64 | PR#8300210 PR#8300684 PR#8300723 | Refer to Resolution Description section | December 16, 2019 6:10:59 PM GMT+08:00
CTUDT00624782 | Refresh | Closed | pengfei_cai | June 27, 2019 11:11:17 AM GMT+08:00 | Operator tripped by patient strap | 4.1.7 | Refer to Resolution Description section | Remove the patient strap or put on the table is not used. | System |  | NM.SRS.Workflow.1619 | Fixed | zhou_r | 0 | 3 | Couch Mechanical | Couch | YES | Closed | SegReq | Brilliance Big Bore | PR#8300210 PR#8300684 PR#8300723 | Refer to Resolution Description section | December 16, 2019 6:10:59 PM GMT+08:00
CTUDT00624782 | Refresh | Closed | pengfei_cai | June 27, 2019 11:11:17 AM GMT+08:00 | Operator tripped by patient strap | 4.1.7 | Refer to Resolution Description section | Remove the patient strap or put on the table is not used. | System |  | NM.SRS.Workflow.1619 | Fixed | zhou_r | 0 | 3 | Couch Mechanical | Couch | YES | Closed | SegReq | Ingenuity 64 | PR#8300210 PR#8300684 PR#8300723 | Refer to Resolution Description section | December 16, 2019 6:10:59 PM GMT+08:00
CTUDT00624782 | Refresh | Closed | pengfei_cai | June 27, 2019 11:11:17 AM GMT+08:00 | Operator tripped by patient strap | 4.1.7 | Refer to Resolution Description section | Remove the patient strap or put on the table is not used. | System |  | NM.SRS.Workflow.1619 | Fixed | zhou_r | 0 | 3 | Couch Mechanical | Couch | YES | Closed | SegReq | IQon1 | PR#8300210 PR#8300684 PR#8300723 | Refer to Resolution Description section | December 16, 2019 6:10:59 PM GMT+08:00
CTUDT00624779 | Refresh | Closed | krisher_i | June 26, 2019 10:36:18 PM GMT+08:00 | Inconsistency between Patient Scene and ECM | WIP17 |  | NA | Subsystem | ECM | NA | Works As Specified / Not A Defect | kr_m | 0 |  |  | Console4 | NO | Closed | WAS | Ovation
CTUDT00624775 | Refresh | Closed | Noy_Seelig | June 26, 2019 9:00:11 PM GMT+08:00 | Refresh: Performance Infant Test, Uniformity shown as Passed but doesn't present results | BB_WIP7 | WIP18 |  | Subsystem | Performance Test | NaN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | BB_WIP8_SegReq | Brilliance Big Bore |  | WIP18 | July 7, 2019 10:11:02 PM GMT+08:00
CTUDT00624771 | Refresh | Closed | cai_t | June 26, 2019 8:42:48 PM GMT+08:00 | Refresh:The "X-ray system cannot comply" popups while doing the second surview in dual surview | V4.8.0.10119 |  | User will have to press the button asking to perform the second surview | Subsystem | System | NA | Invalid/Uncontrolled Configuration | ye_h | 0 |  |  | Support-System | YES | Closed |  | Brilliance Big Bore
CTUDT00624770 | Refresh | Closed | Noy_Seelig | June 26, 2019 8:34:44 PM GMT+08:00 | Refresh: Wrong Couch Height in Performance Infant Test | BB_WIP7 |  |  | Subsystem | Performance Test | NA | Invalid/Uncontrolled Configuration | yatziv_l | 0 |  |  | DMS | NO | Closed |  | Brilliance Big Bore
CTUDT00624767 | Refresh | Closed | Noy_Seelig | June 26, 2019 8:30:43 PM GMT+08:00 | Refresh: Acceptance test, wrong limits in Spatial Resolution Pediatric Head Impulse Response. | BB_WIP7 | WIP18 | N/A | Subsystem | Calilbrations | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | BB_WIP9_SubCom | Brilliance Big Bore |  | WIP18 | July 3, 2019 3:52:04 PM GMT+08:00
CTUDT00624765 | Refresh | Closed | datta_i | June 26, 2019 7:44:27 PM GMT+08:00 | Unable to set the start position value in clinical scan. | 4.8 | BB_WIP8 | Use plan on surview or move the coach to the desired position | Subsystem | System | NA | Fixed | datta_i | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP8_SubCom | Brilliance Big Bore |  | BB_WIP8 | June 27, 2019 7:12:10 PM GMT+08:00
CTUDT00624761 | Refresh | Closed | wang_l | June 26, 2019 3:28:52 PM GMT+08:00 | Refresh: Results - Arc Information does not displayed completely for short tube conditioning and long tube conditioning calibration | BB_WIP6 | WIP18 |  | Subsystem | Calibrations | N/A | Fixed | zhou_x | 0 | 4 | Gantry Calibrations | GCX | NO | Closed | WIP19SegReq | Ovation |  | WIP18 | July 11, 2019 3:39:27 PM GMT+08:00
CTUDT00624761 | Refresh | Closed | wang_l | June 26, 2019 3:28:52 PM GMT+08:00 | Refresh: Results - Arc Information does not displayed completely for short tube conditioning and long tube conditioning calibration | BB_WIP6 | WIP18 |  | Subsystem | Calibrations | N/A | Fixed | zhou_x | 0 | 4 | Gantry Calibrations | GCX | NO | Closed | WIP19SegReq | Brilliance Big Bore |  | WIP18 | July 11, 2019 3:39:27 PM GMT+08:00
CTUDT00624761 | Refresh | Closed | wang_l | June 26, 2019 3:28:52 PM GMT+08:00 | Refresh: Results - Arc Information does not displayed completely for short tube conditioning and long tube conditioning calibration | BB_WIP6 | WIP18 |  | Subsystem | Calibrations | N/A | Fixed | zhou_x | 0 | 4 | Gantry Calibrations | GCX | NO | Closed | WIP19SegReq | Ingenuity 64 |  | WIP18 | July 11, 2019 3:39:27 PM GMT+08:00
CTUDT00624761 | Refresh | Closed | wang_l | June 26, 2019 3:28:52 PM GMT+08:00 | Refresh: Results - Arc Information does not displayed completely for short tube conditioning and long tube conditioning calibration | BB_WIP6 | WIP18 |  | Subsystem | Calibrations | N/A | Fixed | zhou_x | 0 | 4 | Gantry Calibrations | GCX | NO | Closed | WIP19SegReq | IQon1 |  | WIP18 | July 11, 2019 3:39:27 PM GMT+08:00
CTUDT00624759 | Refresh | Closed | lei_yang | June 26, 2019 3:15:18 PM GMT+08:00 | The pitch cannot be set to 0.04 when perform 4D CT scan with some specific parameters. | WIP16 | WIP19 |  | System | Plan; Blocker | Modes.1513 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP18SubCom | Brilliance 64 |  | WIP19 | July 18, 2019 5:19:03 PM GMT+08:00
CTUDT00624759 | Refresh | Closed | lei_yang | June 26, 2019 3:15:18 PM GMT+08:00 | The pitch cannot be set to 0.04 when perform 4D CT scan with some specific parameters. | WIP16 | WIP19 |  | System | Plan; Blocker | Modes.1513 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP18SubCom | Ingenuity 64 |  | WIP19 | July 18, 2019 5:19:03 PM GMT+08:00
CTUDT00624753 | Refresh | Closed | PremKumar_Rajendran | June 26, 2019 1:08:05 PM GMT+08:00 | Refresh - ‘Report’ page crashed when click on ‘Print Screen’ using Keyboard While selecting ‘Reporting’ drop downs | WIP17 | WIP19 | load report before using drop down | Subsystem | Localization | NA | Fixed | s_sa | 0 | 4 | Console USP_EBW | iCAP | NO | Closed | WIP19SubCom | Ovation |  | WIP19 | July 18, 2019 2:20:42 PM GMT+08:00
CTUDT00624753 | Refresh | Closed | PremKumar_Rajendran | June 26, 2019 1:08:05 PM GMT+08:00 | Refresh - ‘Report’ page crashed when click on ‘Print Screen’ using Keyboard While selecting ‘Reporting’ drop downs | WIP17 | WIP19 | load report before using drop down | Subsystem | Localization | NA | Fixed | s_sa | 0 | 4 | Console USP_EBW | iCAP | NO | Closed | WIP19SubCom | Brilliance 64 |  | WIP19 | July 18, 2019 2:20:42 PM GMT+08:00
CTUDT00624753 | Refresh | Closed | PremKumar_Rajendran | June 26, 2019 1:08:05 PM GMT+08:00 | Refresh - ‘Report’ page crashed when click on ‘Print Screen’ using Keyboard While selecting ‘Reporting’ drop downs | WIP17 | WIP19 | load report before using drop down | Subsystem | Localization | NA | Fixed | s_sa | 0 | 4 | Console USP_EBW | iCAP | NO | Closed | WIP19SubCom | Brilliance Big Bore |  | WIP19 | July 18, 2019 2:20:42 PM GMT+08:00
CTUDT00624753 | Refresh | Closed | PremKumar_Rajendran | June 26, 2019 1:08:05 PM GMT+08:00 | Refresh - ‘Report’ page crashed when click on ‘Print Screen’ using Keyboard While selecting ‘Reporting’ drop downs | WIP17 | WIP19 | load report before using drop down | Subsystem | Localization | NA | Fixed | s_sa | 0 | 4 | Console USP_EBW | iCAP | NO | Closed | WIP19SubCom | Ingenuity 64 |  | WIP19 | July 18, 2019 2:20:42 PM GMT+08:00
CTUDT00624753 | Refresh | Closed | PremKumar_Rajendran | June 26, 2019 1:08:05 PM GMT+08:00 | Refresh - ‘Report’ page crashed when click on ‘Print Screen’ using Keyboard While selecting ‘Reporting’ drop downs | WIP17 | WIP19 | load report before using drop down | Subsystem | Localization | NA | Fixed | s_sa | 0 | 4 | Console USP_EBW | iCAP | NO | Closed | WIP19SubCom | IQon1 |  | WIP19 | July 18, 2019 2:20:42 PM GMT+08:00
CTUDT00624744 | Refresh | Closed | zhu_je | June 26, 2019 10:52:45 AM GMT+08:00 | [Refresh] There was message “There seems to be a problem in the application would you like to create a Bug Report” during performing two Calcium Scoring 3D scans | WIP16 | WIP19 | NA | Subsystem | System | CT-NM.SRS.Modes.1510 | Fixed | prasad_r | 0 | 2 | Console Software | Console4 | YES | Closed | WIP18SubCom | Ovation |  | WIP19 | July 4, 2019 6:12:04 PM GMT+08:00
CTUDT00624744 | Refresh | Closed | zhu_je | June 26, 2019 10:52:45 AM GMT+08:00 | [Refresh] There was message “There seems to be a problem in the application would you like to create a Bug Report” during performing two Calcium Scoring 3D scans | WIP16 | WIP19 | NA | Subsystem | System | CT-NM.SRS.Modes.1510 | Fixed | prasad_r | 0 | 2 | Console Software | Console4 | YES | Closed | WIP18SubCom | Brilliance 64 |  | WIP19 | July 4, 2019 6:12:04 PM GMT+08:00
CTUDT00624744 | Refresh | Closed | zhu_je | June 26, 2019 10:52:45 AM GMT+08:00 | [Refresh] There was message “There seems to be a problem in the application would you like to create a Bug Report” during performing two Calcium Scoring 3D scans | WIP16 | WIP19 | NA | Subsystem | System | CT-NM.SRS.Modes.1510 | Fixed | prasad_r | 0 | 2 | Console Software | Console4 | YES | Closed | WIP18SubCom | Ingenuity 64 |  | WIP19 | July 4, 2019 6:12:04 PM GMT+08:00
CTUDT00624744 | Refresh | Closed | zhu_je | June 26, 2019 10:52:45 AM GMT+08:00 | [Refresh] There was message “There seems to be a problem in the application would you like to create a Bug Report” during performing two Calcium Scoring 3D scans | WIP16 | WIP19 | NA | Subsystem | System | CT-NM.SRS.Modes.1510 | Fixed | prasad_r | 0 | 2 | Console Software | Console4 | YES | Closed | WIP18SubCom | IQon1 |  | WIP19 | July 4, 2019 6:12:04 PM GMT+08:00
CTUDT00624742 | Refresh | Closed | cai_t | June 26, 2019 9:12:34 AM GMT+08:00 | Refresh:HPZ440: No sound was not heard at the CT Box speaker and Gantry speaker before scanning. | BB_WIP6 |  |  | Subsystem | Voice |  | Duplicate | naveen_k | 0 |  |  | Console4 | YES | Closed |  | Brilliance Big Bore
CTUDT00624740 | Refresh | Closed | PremKumar_Rajendran | June 25, 2019 7:47:13 PM GMT+08:00 | BB Refresh : Additional results created in ‘View’ scene and ‘Patient Directory’ when O-MAR options enabled. | BB_WIP7 |  | NA | Subsystem | OMAR | NA | Works As Specified / Not A Defect | kr_m | 0 |  |  | Console4 | NO | Closed | WAS | Brilliance Big Bore
CTUDT00624733 | Refresh | Closed | naveen_k | June 25, 2019 4:24:45 PM GMT+08:00 | IOCC_Refresh: IHE_Verification :- C_FIND and C_STORE attribute values not same | WIP7 | BB_WIP8 | There is no workaround. | Subsystem | DICOM | NA | Fixed | jayakumar_p | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP8_SubCom | Brilliance Big Bore |  | BB_WIP8 | July 4, 2019 1:17:08 PM GMT+08:00
CTUDT00624726 | Refresh | Closed | zhang_t | June 25, 2019 2:58:00 PM GMT+08:00 | Refresh: The warning message popup every time open QA even after restart Gantry. | WIP17 | WIP18 |  | Subsystem | Calibrations | na | Fixed | Kostya_Baklan | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP18SegReq | Ovation |  | WIP18 | June 30, 2019 9:31:26 PM GMT+08:00
CTUDT00624726 | Refresh | Closed | zhang_t | June 25, 2019 2:58:00 PM GMT+08:00 | Refresh: The warning message popup every time open QA even after restart Gantry. | WIP17 | WIP18 |  | Subsystem | Calibrations | na | Fixed | Kostya_Baklan | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP18SegReq | Ingenuity 64 |  | WIP18 | June 30, 2019 9:31:26 PM GMT+08:00
CTUDT00624726 | Refresh | Closed | zhang_t | June 25, 2019 2:58:00 PM GMT+08:00 | Refresh: The warning message popup every time open QA even after restart Gantry. | WIP17 | WIP18 |  | Subsystem | Calibrations | na | Fixed | Kostya_Baklan | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP18SegReq | IQon1 |  | WIP18 | June 30, 2019 9:31:26 PM GMT+08:00
CTUDT00624723 | Refresh | Closed | lei_yang | June 25, 2019 1:24:47 PM GMT+08:00 | Respiratory Level of 'Show All' was not identical with the Respiratory Level of 'Scan Parameters' | BB_WIP6 | BB_WIP8 |  | Subsystem | Plan | None | Fixed | wilsonraja | 0 | 2 | Console Software | Console4 | YES | Closed | WIP8Sub and WIP18Sub | Brilliance Big Bore |  | BB_WIP8 | July 5, 2019 5:55:09 PM GMT+08:00
CTUDT00624721 | Refresh | Closed | lin_s | June 25, 2019 9:50:04 AM GMT+08:00 | Target markets requirement is not defined in the Customer Requirement Specification (PRS / CRS), not meet Design Input SU-03002 7.2.1.8 Markets in which the product will be sold (e.g. China, Europe, US etc) | PRS/CRS | DHF342944 Rev 01 - BB CRS |  | System | Requirements | CT-NM.PRS.CRS.7294 | Fixed | mistry_h | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | BB_WIP9_SegReq | Brilliance Big Bore |  | DHF342944 Rev 01 - BB CRS | August 7, 2019 9:26:26 AM GMT+08:00
CTUDT00624718 | Refresh | Closed | azagury_d | June 25, 2019 3:33:24 AM GMT+08:00 | Refresh: Calibrations: Beam and DMS Check: Couch leaves scanning area in RnD mode | WIP17 | WIP18 |  | Subsystem | Calibrations | n/a | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP8Sub + WIP18Sub | Ovation |  | WIP18 | June 30, 2019 5:56:11 PM GMT+08:00
CTUDT00624718 | Refresh | Closed | azagury_d | June 25, 2019 3:33:24 AM GMT+08:00 | Refresh: Calibrations: Beam and DMS Check: Couch leaves scanning area in RnD mode | WIP17 | WIP18 |  | Subsystem | Calibrations | n/a | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP8Sub + WIP18Sub | Brilliance Big Bore |  | WIP18 | June 30, 2019 5:56:11 PM GMT+08:00
CTUDT00624718 | Refresh | Closed | azagury_d | June 25, 2019 3:33:24 AM GMT+08:00 | Refresh: Calibrations: Beam and DMS Check: Couch leaves scanning area in RnD mode | WIP17 | WIP18 |  | Subsystem | Calibrations | n/a | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP8Sub + WIP18Sub | Ingenuity 64 |  | WIP18 | June 30, 2019 5:56:11 PM GMT+08:00
CTUDT00624718 | Refresh | Closed | azagury_d | June 25, 2019 3:33:24 AM GMT+08:00 | Refresh: Calibrations: Beam and DMS Check: Couch leaves scanning area in RnD mode | WIP17 | WIP18 |  | Subsystem | Calibrations | n/a | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP8Sub + WIP18Sub | IQon1 |  | WIP18 | June 30, 2019 5:56:11 PM GMT+08:00
CTUDT00624714 | Refresh | Closed | cherylli | June 24, 2019 11:10:31 PM GMT+08:00 | Refresh Pre-Val: Wrong representation of the scan on the resp waveform (Resp Single Phase; User Defined) | BB_WIP6 | BB_WIP9 |  | Subsystem | Respiratory | N/A | Fixed | pandey_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP19 & WIP9 | Brilliance Big Bore |  | BB_WIP9 | July 25, 2019 8:20:22 PM GMT+08:00
CTUDT00624712 | Refresh | Closed | cherylli | June 24, 2019 10:51:36 PM GMT+08:00 | Refresh Pre-Val: Incorrect labelling of the R-R intervals (Resp Single Phase User Defined) | BB_WIP6 |  |  | Subsystem | Respiratory | Console-SRS.11314 | Works As Specified / Not A Defect | kanagaraj_n | 0 |  |  | Console4 | NO | Closed | WAS | Brilliance Big Bore
CTUDT00624710 | Refresh | Closed | cherylli | June 24, 2019 10:26:14 PM GMT+08:00 | Refresh Pre-Val: Skipped R-R intervals of Resp Single Phase although respiratory waveform is consistent and stable | BB_WIP6 |  |  | Subsystem | Respiratory | None | Works As Specified / Not A Defect | tan_j | 0 | 2 |  | GCX | YES | Closed |  | Brilliance 64
CTUDT00624710 | Refresh | Closed | cherylli | June 24, 2019 10:26:14 PM GMT+08:00 | Refresh Pre-Val: Skipped R-R intervals of Resp Single Phase although respiratory waveform is consistent and stable | BB_WIP6 |  |  | Subsystem | Respiratory | None | Works As Specified / Not A Defect | tan_j | 0 | 2 |  | GCX | YES | Closed |  | Brilliance Big Bore
CTUDT00624710 | Refresh | Closed | cherylli | June 24, 2019 10:26:14 PM GMT+08:00 | Refresh Pre-Val: Skipped R-R intervals of Resp Single Phase although respiratory waveform is consistent and stable | BB_WIP6 |  |  | Subsystem | Respiratory | None | Works As Specified / Not A Defect | tan_j | 0 | 2 |  | GCX | YES | Closed |  | Ingenuity 64
CTUDT00624710 | Refresh | Closed | cherylli | June 24, 2019 10:26:14 PM GMT+08:00 | Refresh Pre-Val: Skipped R-R intervals of Resp Single Phase although respiratory waveform is consistent and stable | BB_WIP6 |  |  | Subsystem | Respiratory | None | Works As Specified / Not A Defect | tan_j | 0 | 2 |  | GCX | YES | Closed |  | Ingenuity BigBore
CTUDT00624708 | Refresh | Closed | cherylli | June 24, 2019 9:50:46 PM GMT+08:00 | Refresh Pre-Val: Inconsistent use of terminology: “metal reduction” instead of “OMAR” | BB_WIP6 |  | NA | Subsystem | Terminology | Console.SRS.15388 | Works As Specified / Not A Defect | saha_d | 0 | 3 |  | Console4 | NO | Closed | SubNC | Brilliance Big Bore
CTUDT00624706 | Refresh | Closed | beck_d | June 24, 2019 9:00:14 PM GMT+08:00 | Refresh: IQ Check Test: Harmonized Phantom: Head Noise lower limit is not as defined in spec | BB_WIP7 | WIP18 | N/A | Subsystem | IQ Check | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | YES | Closed | BB_WIP8_SubCom | Brilliance Big Bore |  | WIP18 | June 30, 2019 4:07:46 PM GMT+08:00
CTUDT00624699 | Refresh | Closed | Sharmila_Vasireddy | June 24, 2019 7:56:28 PM GMT+08:00 | BB refresh: The warning message was not localized as per the language set in Preferences | BB_WIP7 | BB_WIP9 |  | Subsystem | Localization | NA | Fixed | chokkalingam_r | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP9_SubCom | Brilliance Big Bore |  | BB_WIP9 | July 19, 2019 4:40:37 PM GMT+08:00
CTUDT00624696 | Refresh | Closed | kapadia_v | June 24, 2019 5:11:07 PM GMT+08:00 | Refresh: Security: Incorrect message displayed when user enters un supported character in emergency user name | 4.8 | BB_WIP8 |  | Subsystem | ECM | NA | Fixed | kapadia_v | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP8_SubCom | Brilliance Big Bore |  | BB_WIP8 | June 28, 2019 5:41:50 PM GMT+08:00
CTUDT00624682 | Refresh | Closed | zhu_je | June 24, 2019 4:34:08 PM GMT+08:00 | [Refresh][IFU Kits][Ing&Br64] There are many requirements linked to Tech_Pub but missing in the traceability DHF303288 Rev 01_User Doc and Localization Plan | WIP16 | WIP19 |  | Subsystem | IFU | Refer to Attachment | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Ovation |  | WIP19 | August 1, 2019 1:43:01 PM GMT+08:00
CTUDT00624682 | Refresh | Closed | zhu_je | June 24, 2019 4:34:08 PM GMT+08:00 | [Refresh][IFU Kits][Ing&Br64] There are many requirements linked to Tech_Pub but missing in the traceability DHF303288 Rev 01_User Doc and Localization Plan | WIP16 | WIP19 |  | Subsystem | IFU | Refer to Attachment | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Brilliance 64 |  | WIP19 | August 1, 2019 1:43:01 PM GMT+08:00
CTUDT00624682 | Refresh | Closed | zhu_je | June 24, 2019 4:34:08 PM GMT+08:00 | [Refresh][IFU Kits][Ing&Br64] There are many requirements linked to Tech_Pub but missing in the traceability DHF303288 Rev 01_User Doc and Localization Plan | WIP16 | WIP19 |  | Subsystem | IFU | Refer to Attachment | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Brilliance Big Bore |  | WIP19 | August 1, 2019 1:43:01 PM GMT+08:00
CTUDT00624682 | Refresh | Closed | zhu_je | June 24, 2019 4:34:08 PM GMT+08:00 | [Refresh][IFU Kits][Ing&Br64] There are many requirements linked to Tech_Pub but missing in the traceability DHF303288 Rev 01_User Doc and Localization Plan | WIP16 | WIP19 |  | Subsystem | IFU | Refer to Attachment | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Ingenuity 64 |  | WIP19 | August 1, 2019 1:43:01 PM GMT+08:00
CTUDT00624682 | Refresh | Closed | zhu_je | June 24, 2019 4:34:08 PM GMT+08:00 | [Refresh][IFU Kits][Ing&Br64] There are many requirements linked to Tech_Pub but missing in the traceability DHF303288 Rev 01_User Doc and Localization Plan | WIP16 | WIP19 |  | Subsystem | IFU | Refer to Attachment | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | IQon1 |  | WIP19 | August 1, 2019 1:43:01 PM GMT+08:00
CTUDT00624680 | Refresh | Closed | cherylli | June 24, 2019 4:05:40 PM GMT+08:00 | Refresh Pre-Val:“LOCALIZER” should be “Surview” | BB_WIP6 |  |  | Subsystem | Terminology | NA | No Plan to Fix | kr_m | 0 | 4 |  | Console4 | NO | Closed | SubNC | Ovation
CTUDT00624680 | Refresh | Closed | cherylli | June 24, 2019 4:05:40 PM GMT+08:00 | Refresh Pre-Val:“LOCALIZER” should be “Surview” | BB_WIP6 |  |  | Subsystem | Terminology | NA | No Plan to Fix | kr_m | 0 | 4 |  | Console4 | NO | Closed | SubNC | Brilliance 64
CTUDT00624680 | Refresh | Closed | cherylli | June 24, 2019 4:05:40 PM GMT+08:00 | Refresh Pre-Val:“LOCALIZER” should be “Surview” | BB_WIP6 |  |  | Subsystem | Terminology | NA | No Plan to Fix | kr_m | 0 | 4 |  | Console4 | NO | Closed | SubNC | Brilliance Big Bore
CTUDT00624680 | Refresh | Closed | cherylli | June 24, 2019 4:05:40 PM GMT+08:00 | Refresh Pre-Val:“LOCALIZER” should be “Surview” | BB_WIP6 |  |  | Subsystem | Terminology | NA | No Plan to Fix | kr_m | 0 | 4 |  | Console4 | NO | Closed | SubNC | Ingenuity 64
CTUDT00624680 | Refresh | Closed | cherylli | June 24, 2019 4:05:40 PM GMT+08:00 | Refresh Pre-Val:“LOCALIZER” should be “Surview” | BB_WIP6 |  |  | Subsystem | Terminology | NA | No Plan to Fix | kr_m | 0 | 4 |  | Console4 | NO | Closed | SubNC | Ingenuity TF 64
CTUDT00624680 | Refresh | Closed | cherylli | June 24, 2019 4:05:40 PM GMT+08:00 | Refresh Pre-Val:“LOCALIZER” should be “Surview” | BB_WIP6 |  |  | Subsystem | Terminology | NA | No Plan to Fix | kr_m | 0 | 4 |  | Console4 | NO | Closed | SubNC | Veyron
CTUDT00624680 | Refresh | Closed | cherylli | June 24, 2019 4:05:40 PM GMT+08:00 | Refresh Pre-Val:“LOCALIZER” should be “Surview” | BB_WIP6 |  |  | Subsystem | Terminology | NA | No Plan to Fix | kr_m | 0 | 4 |  | Console4 | NO | Closed | SubNC | IQon1
CTUDT00624678 | Refresh | Closed | zhang_t | June 24, 2019 3:28:01 PM GMT+08:00 | Refresh: There was no ACQ diagnostic tool in Desktop via remote viewer to CIRS S1 under olevel service. | WIP17 |  |  | Subsystem | Remote Configuration | NA | Works As Specified / Not A Defect | rakic_k | 0 |  |  | CIRS | NO | Closed |  | Ovation
CTUDT00624678 | Refresh | Closed | zhang_t | June 24, 2019 3:28:01 PM GMT+08:00 | Refresh: There was no ACQ diagnostic tool in Desktop via remote viewer to CIRS S1 under olevel service. | WIP17 |  |  | Subsystem | Remote Configuration | NA | Works As Specified / Not A Defect | rakic_k | 0 |  |  | CIRS | NO | Closed |  | Brilliance 64
CTUDT00624678 | Refresh | Closed | zhang_t | June 24, 2019 3:28:01 PM GMT+08:00 | Refresh: There was no ACQ diagnostic tool in Desktop via remote viewer to CIRS S1 under olevel service. | WIP17 |  |  | Subsystem | Remote Configuration | NA | Works As Specified / Not A Defect | rakic_k | 0 |  |  | CIRS | NO | Closed |  | Brilliance Big Bore
CTUDT00624678 | Refresh | Closed | zhang_t | June 24, 2019 3:28:01 PM GMT+08:00 | Refresh: There was no ACQ diagnostic tool in Desktop via remote viewer to CIRS S1 under olevel service. | WIP17 |  |  | Subsystem | Remote Configuration | NA | Works As Specified / Not A Defect | rakic_k | 0 |  |  | CIRS | NO | Closed |  | Ingenuity 64
CTUDT00624678 | Refresh | Closed | zhang_t | June 24, 2019 3:28:01 PM GMT+08:00 | Refresh: There was no ACQ diagnostic tool in Desktop via remote viewer to CIRS S1 under olevel service. | WIP17 |  |  | Subsystem | Remote Configuration | NA | Works As Specified / Not A Defect | rakic_k | 0 |  |  | CIRS | NO | Closed |  | IQon1
CTUDT00624675 | Refresh | Closed | teng_h | June 24, 2019 2:55:23 PM GMT+08:00 | button force and actuation distance of IVC joystick mouse is out of requirement | 4.1.7 | WIP19 |  | Subsystem |  | GCX.SSRS.iCT.20608;GCX.SSRS.iCT.20609 | Fixed | Wang_Zhen | 0 | 2 | Gantry Electrical | GCX | NO | Closed | BB_WIP9_SubCom | Brilliance Big Bore |  | WIP19 | August 6, 2019 7:53:49 PM GMT+08:00
CTUDT00624673 | Refresh | Closed | zhang_t | June 24, 2019 2:50:44 PM GMT+08:00 | Refresh: Sometimes, cannot remote viewer to CIRS after CIRS password changes | WIP17 | 459801587331 Rev C 459801587341 Rev C 459801587361 Rev C 459801587351 Rev C DMR356920 Rev.06 | Service would have to re-set the computer/password and re-install. | Subsystem | Remote Configuration; | N/A | Fixed | shaham_t | 0 | 2 | CIRS_SW_Installation | Support-System | NO | Closed | SubCom-03-Dec-2019 | Ovation |  | 459801587331 Rev C 459801587341 Rev C 459801587361 Rev C 459801587351 Rev C DMR356920 Rev.06 | December 3, 2019 8:27:25 PM GMT+08:00
CTUDT00624673 | Refresh | Closed | zhang_t | June 24, 2019 2:50:44 PM GMT+08:00 | Refresh: Sometimes, cannot remote viewer to CIRS after CIRS password changes | WIP17 | 459801587331 Rev C 459801587341 Rev C 459801587361 Rev C 459801587351 Rev C DMR356920 Rev.06 | Service would have to re-set the computer/password and re-install. | Subsystem | Remote Configuration; | N/A | Fixed | shaham_t | 0 | 2 | CIRS_SW_Installation | Support-System | NO | Closed | SubCom-03-Dec-2019 | Brilliance 64 |  | 459801587331 Rev C 459801587341 Rev C 459801587361 Rev C 459801587351 Rev C DMR356920 Rev.06 | December 3, 2019 8:27:25 PM GMT+08:00
CTUDT00624673 | Refresh | Closed | zhang_t | June 24, 2019 2:50:44 PM GMT+08:00 | Refresh: Sometimes, cannot remote viewer to CIRS after CIRS password changes | WIP17 | 459801587331 Rev C 459801587341 Rev C 459801587361 Rev C 459801587351 Rev C DMR356920 Rev.06 | Service would have to re-set the computer/password and re-install. | Subsystem | Remote Configuration; | N/A | Fixed | shaham_t | 0 | 2 | CIRS_SW_Installation | Support-System | NO | Closed | SubCom-03-Dec-2019 | Brilliance Big Bore |  | 459801587331 Rev C 459801587341 Rev C 459801587361 Rev C 459801587351 Rev C DMR356920 Rev.06 | December 3, 2019 8:27:25 PM GMT+08:00
CTUDT00624673 | Refresh | Closed | zhang_t | June 24, 2019 2:50:44 PM GMT+08:00 | Refresh: Sometimes, cannot remote viewer to CIRS after CIRS password changes | WIP17 | 459801587331 Rev C 459801587341 Rev C 459801587361 Rev C 459801587351 Rev C DMR356920 Rev.06 | Service would have to re-set the computer/password and re-install. | Subsystem | Remote Configuration; | N/A | Fixed | shaham_t | 0 | 2 | CIRS_SW_Installation | Support-System | NO | Closed | SubCom-03-Dec-2019 | Ingenuity 64 |  | 459801587331 Rev C 459801587341 Rev C 459801587361 Rev C 459801587351 Rev C DMR356920 Rev.06 | December 3, 2019 8:27:25 PM GMT+08:00
CTUDT00624673 | Refresh | Closed | zhang_t | June 24, 2019 2:50:44 PM GMT+08:00 | Refresh: Sometimes, cannot remote viewer to CIRS after CIRS password changes | WIP17 | 459801587331 Rev C 459801587341 Rev C 459801587361 Rev C 459801587351 Rev C DMR356920 Rev.06 | Service would have to re-set the computer/password and re-install. | Subsystem | Remote Configuration; | N/A | Fixed | shaham_t | 0 | 2 | CIRS_SW_Installation | Support-System | NO | Closed | SubCom-03-Dec-2019 | IQon1 |  | 459801587331 Rev C 459801587341 Rev C 459801587361 Rev C 459801587351 Rev C DMR356920 Rev.06 | December 3, 2019 8:27:25 PM GMT+08:00
CTUDT00624671 | Refresh | Closed | zhu_je | June 24, 2019 2:09:41 PM GMT+08:00 | [Refresh] The Rotation Time and Pitch values did not automatically update due to scan length changed | WIP16 |  |  | Subsystem | Plan | CT-NM.SRS.Workflow.1760 | Works As Specified / Not A Defect | rath_s | 0 |  |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00624671 | Refresh | Closed | zhu_je | June 24, 2019 2:09:41 PM GMT+08:00 | [Refresh] The Rotation Time and Pitch values did not automatically update due to scan length changed | WIP16 |  |  | Subsystem | Plan | CT-NM.SRS.Workflow.1760 | Works As Specified / Not A Defect | rath_s | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00624668 | Refresh | Closed | zhang_t | June 24, 2019 1:20:15 PM GMT+08:00 | Performance Infant Test failed under Harmonized phantom, the couch auto move to 35 from 122 during testing. | WIP17 |  |  | Subsystem | Performance Test; Blocker |  | Invalid/Uncontrolled Configuration | yatziv_l | 0 |  |  | DMS | NO | Closed | WIP18SegReq | Ingenuity 64
CTUDT00624667 | Refresh | Closed | datta_i | June 24, 2019 12:45:35 PM GMT+08:00 | Mismatch in mAs clipped message | 4.8 | BB_WIP9 |  | Subsystem |  | Console-SRS.4738 | Fixed | datta_i | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP8_SubCom | Brilliance Big Bore |  | BB_WIP9 | July 17, 2019 2:35:00 PM GMT+08:00
CTUDT00624665 | Refresh | Closed | zhang_t | June 24, 2019 11:51:01 AM GMT+08:00 | Refresh: During the calibration FS width, DFS position, UHR DFS position get stuck and cannot continue | WIP17 | WIP19 | Disable the firewall to run the calibrations. | Subsystem | Calibrations; Blocker19 | n/a | Fixed | richards_b | 0 | 2 | Console Software | Support-Physics | NO | Closed | WIP18SegReq | Brilliance 64 |  | WIP19 | August 17, 2019 1:45:55 AM GMT+08:00
CTUDT00624665 | Refresh | Closed | zhang_t | June 24, 2019 11:51:01 AM GMT+08:00 | Refresh: During the calibration FS width, DFS position, UHR DFS position get stuck and cannot continue | WIP17 | WIP19 | Disable the firewall to run the calibrations. | Subsystem | Calibrations; Blocker19 | n/a | Fixed | richards_b | 0 | 2 | Console Software | Support-Physics | NO | Closed | WIP18SegReq | Brilliance Big Bore |  | WIP19 | August 17, 2019 1:45:55 AM GMT+08:00
CTUDT00624665 | Refresh | Closed | zhang_t | June 24, 2019 11:51:01 AM GMT+08:00 | Refresh: During the calibration FS width, DFS position, UHR DFS position get stuck and cannot continue | WIP17 | WIP19 | Disable the firewall to run the calibrations. | Subsystem | Calibrations; Blocker19 | n/a | Fixed | richards_b | 0 | 2 | Console Software | Support-Physics | NO | Closed | WIP18SegReq | Ingenuity 64 |  | WIP19 | August 17, 2019 1:45:55 AM GMT+08:00
CTUDT00624663 | Refresh | Closed | wang_l | June 24, 2019 10:42:01 AM GMT+08:00 | Refresh: Performance issue - Collect bugreport stucked at 49% progress | WIP9.1 |  |  | Subsystem | Bugreport |  | Works As Specified / Not A Defect | wang_l | 1 |  |  | Support-System | NO | Closed |  | Ingenuity 64
CTUDT00624660 | Refresh | Closed | ding_s | June 24, 2019 9:34:05 AM GMT+08:00 | Exam cards in Orthoped group loaded failed. | WIP16 |  | There is no workaround | Subsystem | Plan | NA | Works As Specified / Not A Defect | ding_s | 1 |  |  | Support-System | NO | Closed |  | Ingenuity 64
CTUDT00624655 | Refresh | Closed | liu_j | June 21, 2019 10:47:12 PM GMT+08:00 | Refresh Console SSRS update needed for Console-gantry interface related to move to lasers LEDs | BB_WIP6 | DHF316705 Rev.12 Host SSRS for Refresh BB | N/A | Subsystem | Requirements | N/A | Fixed | hatti_c | 0 | 2 | Requirements Documentation | Console4 | YES | Closed | BB_WIP8_SubCom | Brilliance Big Bore |  | DHF316705 Rev.12 Host SSRS for Refresh BB | July 19, 2019 5:32:57 PM GMT+08:00
CTUDT00624653 | Refresh | Closed | liu_j | June 21, 2019 10:33:42 PM GMT+08:00 | Refresh SRS stopping distance requirement for gantry move to inner/outer lasers is not met | BB_WIP6 | DHF298555 Rev 06 - BB SRS |  | System | Requirements | CT-NM.SRS.Modes.1789 | Fixed | mistry_h | 0 | 2 | Requirements Documentation | Support-System | YES | Closed | BB_WIP8_SubCom | Brilliance Big Bore |  | DHF298555 Rev 06 - BB SRS | July 26, 2019 8:18:02 PM GMT+08:00
CTUDT00624651 | Refresh | Closed | cai_t | June 21, 2019 8:36:24 PM GMT+08:00 | system error popups after save an exam card in Exam Card Manager. | BB_WIP6 |  |  | Subsystem | ECM |  | Works As Specified / Not A Defect | s_sa | 1 |  |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00624649 | Refresh | Closed | zhu_je | June 21, 2019 6:10:57 PM GMT+08:00 | [Refresh] "Rotor_Position_Req" is not appeared in Cantrace after load an EC to scan if setting "Enable Fast Get Ready" | WIP16 | WIP19 |  | System | System;Verified1 | None | Fixed | li_e | 0 | 4 | Console Software | Console4 | NO | Closed | WIP19SubCom | Ovation |  | WIP19 | July 19, 2019 1:04:35 PM GMT+08:00
CTUDT00624649 | Refresh | Closed | zhu_je | June 21, 2019 6:10:57 PM GMT+08:00 | [Refresh] "Rotor_Position_Req" is not appeared in Cantrace after load an EC to scan if setting "Enable Fast Get Ready" | WIP16 | WIP19 |  | System | System;Verified1 | None | Fixed | li_e | 0 | 4 | Console Software | Console4 | NO | Closed | WIP19SubCom | Brilliance 64 |  | WIP19 | July 19, 2019 1:04:35 PM GMT+08:00
CTUDT00624649 | Refresh | Closed | zhu_je | June 21, 2019 6:10:57 PM GMT+08:00 | [Refresh] "Rotor_Position_Req" is not appeared in Cantrace after load an EC to scan if setting "Enable Fast Get Ready" | WIP16 | WIP19 |  | System | System;Verified1 | None | Fixed | li_e | 0 | 4 | Console Software | Console4 | NO | Closed | WIP19SubCom | Ingenuity 64 |  | WIP19 | July 19, 2019 1:04:35 PM GMT+08:00
CTUDT00624649 | Refresh | Closed | zhu_je | June 21, 2019 6:10:57 PM GMT+08:00 | [Refresh] "Rotor_Position_Req" is not appeared in Cantrace after load an EC to scan if setting "Enable Fast Get Ready" | WIP16 | WIP19 |  | System | System;Verified1 | None | Fixed | li_e | 0 | 4 | Console Software | Console4 | NO | Closed | WIP19SubCom | IQon1 |  | WIP19 | July 19, 2019 1:04:35 PM GMT+08:00
CTUDT00624647 | Refresh | Closed | lei_yang | June 21, 2019 5:16:11 PM GMT+08:00 | Overlap is shown when the Peak-Peak interval are close to each other. | WIP16 |  |  | Subsystem | Respiratory | NA | Works As Specified / Not A Defect | chokkalingam_r | 0 |  |  | Console4 | NO | Closed | SubNC | Brilliance 64
CTUDT00624647 | Refresh | Closed | lei_yang | June 21, 2019 5:16:11 PM GMT+08:00 | Overlap is shown when the Peak-Peak interval are close to each other. | WIP16 |  |  | Subsystem | Respiratory | NA | Works As Specified / Not A Defect | chokkalingam_r | 0 |  |  | Console4 | NO | Closed | SubNC | Brilliance Big Bore
CTUDT00624647 | Refresh | Closed | lei_yang | June 21, 2019 5:16:11 PM GMT+08:00 | Overlap is shown when the Peak-Peak interval are close to each other. | WIP16 |  |  | Subsystem | Respiratory | NA | Works As Specified / Not A Defect | chokkalingam_r | 0 |  |  | Console4 | NO | Closed | SubNC | Ingenuity 64
CTUDT00624644 | Refresh | Closed | zhu_je | June 21, 2019 4:37:09 PM GMT+08:00 | [Refresh] When the number of printed pages was more using Snagit9, error occurred for Snagit9 and system error | WIP16 |  |  | Subsystem | SnagIt |  | Works As Specified / Not A Defect | lin_s | 0 |  |  | Support-System | NO | Closed |  | Brilliance 64
CTUDT00624644 | Refresh | Closed | zhu_je | June 21, 2019 4:37:09 PM GMT+08:00 | [Refresh] When the number of printed pages was more using Snagit9, error occurred for Snagit9 and system error | WIP16 |  |  | Subsystem | SnagIt |  | Works As Specified / Not A Defect | lin_s | 0 |  |  | Support-System | NO | Closed |  | Ingenuity 64
CTUDT00624642 | Refresh | Closed | zhu_je | June 21, 2019 3:25:47 PM GMT+08:00 | [Refresh] Sometimes cannot lock the Records in reconstruction tab on the right monitor | WIP16 |  | Do not run refresh until lock is finished and lock flag appears | Subsystem | Patient Directory | NA | No Plan to Fix | shaham_t | 0 | 3 | Console USP_EBW | iCAP | YES | Closed | SubNC | Ovation
CTUDT00624642 | Refresh | Closed | zhu_je | June 21, 2019 3:25:47 PM GMT+08:00 | [Refresh] Sometimes cannot lock the Records in reconstruction tab on the right monitor | WIP16 |  | Do not run refresh until lock is finished and lock flag appears | Subsystem | Patient Directory | NA | No Plan to Fix | shaham_t | 0 | 3 | Console USP_EBW | iCAP | YES | Closed | SubNC | Brilliance 64
CTUDT00624642 | Refresh | Closed | zhu_je | June 21, 2019 3:25:47 PM GMT+08:00 | [Refresh] Sometimes cannot lock the Records in reconstruction tab on the right monitor | WIP16 |  | Do not run refresh until lock is finished and lock flag appears | Subsystem | Patient Directory | NA | No Plan to Fix | shaham_t | 0 | 3 | Console USP_EBW | iCAP | YES | Closed | SubNC | Brilliance Big Bore
CTUDT00624642 | Refresh | Closed | zhu_je | June 21, 2019 3:25:47 PM GMT+08:00 | [Refresh] Sometimes cannot lock the Records in reconstruction tab on the right monitor | WIP16 |  | Do not run refresh until lock is finished and lock flag appears | Subsystem | Patient Directory | NA | No Plan to Fix | shaham_t | 0 | 3 | Console USP_EBW | iCAP | YES | Closed | SubNC | Ingenuity 64
CTUDT00624642 | Refresh | Closed | zhu_je | June 21, 2019 3:25:47 PM GMT+08:00 | [Refresh] Sometimes cannot lock the Records in reconstruction tab on the right monitor | WIP16 |  | Do not run refresh until lock is finished and lock flag appears | Subsystem | Patient Directory | NA | No Plan to Fix | shaham_t | 0 | 3 | Console USP_EBW | iCAP | YES | Closed | SubNC | IQon1
CTUDT00624640 | Refresh | Closed | zhu_je | June 21, 2019 3:03:14 PM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.Service.161_ Cannot find the evidence in TRG and IFU files | WIP16 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.Service.161 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Ovation |  | WIP19 | July 31, 2019 4:43:28 PM GMT+08:00
CTUDT00624640 | Refresh | Closed | zhu_je | June 21, 2019 3:03:14 PM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.Service.161_ Cannot find the evidence in TRG and IFU files | WIP16 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.Service.161 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Brilliance 64 |  | WIP19 | July 31, 2019 4:43:28 PM GMT+08:00
CTUDT00624640 | Refresh | Closed | zhu_je | June 21, 2019 3:03:14 PM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.Service.161_ Cannot find the evidence in TRG and IFU files | WIP16 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.Service.161 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Brilliance Big Bore |  | WIP19 | July 31, 2019 4:43:28 PM GMT+08:00
CTUDT00624640 | Refresh | Closed | zhu_je | June 21, 2019 3:03:14 PM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.Service.161_ Cannot find the evidence in TRG and IFU files | WIP16 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.Service.161 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Ingenuity 64 |  | WIP19 | July 31, 2019 4:43:28 PM GMT+08:00
CTUDT00624640 | Refresh | Closed | zhu_je | June 21, 2019 3:03:14 PM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.Service.161_ Cannot find the evidence in TRG and IFU files | WIP16 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.Service.161 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | IQon1 |  | WIP19 | July 31, 2019 4:43:28 PM GMT+08:00
CTUDT00624638 | Refresh | Closed | zhu_je | June 21, 2019 2:58:05 PM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.Service.156_ Cannot find the evidence in TRG and IFU files | WIP16 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.Service.156 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Ovation |  | WIP19 | July 31, 2019 4:42:28 PM GMT+08:00
CTUDT00624638 | Refresh | Closed | zhu_je | June 21, 2019 2:58:05 PM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.Service.156_ Cannot find the evidence in TRG and IFU files | WIP16 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.Service.156 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Brilliance 64 |  | WIP19 | July 31, 2019 4:42:28 PM GMT+08:00
CTUDT00624638 | Refresh | Closed | zhu_je | June 21, 2019 2:58:05 PM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.Service.156_ Cannot find the evidence in TRG and IFU files | WIP16 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.Service.156 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Brilliance Big Bore |  | WIP19 | July 31, 2019 4:42:28 PM GMT+08:00
CTUDT00624638 | Refresh | Closed | zhu_je | June 21, 2019 2:58:05 PM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.Service.156_ Cannot find the evidence in TRG and IFU files | WIP16 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.Service.156 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Ingenuity 64 |  | WIP19 | July 31, 2019 4:42:28 PM GMT+08:00
CTUDT00624638 | Refresh | Closed | zhu_je | June 21, 2019 2:58:05 PM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.Service.156_ Cannot find the evidence in TRG and IFU files | WIP16 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.Service.156 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | IQon1 |  | WIP19 | July 31, 2019 4:42:28 PM GMT+08:00
CTUDT00624636 | Refresh | Closed | Sharmila_Vasireddy | June 21, 2019 2:19:10 PM GMT+08:00 | BB refresh : Import files window not displayed upon clicking “import” button in voice manager tool | BB_WIP7 | BB_WIP8 | N/A | Subsystem | Voice Manager | N/A | Fixed | sheik_c | 0 | 2 | Console Software | Console4 | NO | Closed | WIP18SubCom | Brilliance Big Bore |  | BB_WIP8 | July 2, 2019 9:38:20 PM GMT+08:00
CTUDT00624634 | Refresh | Closed | lakshmikanth_s | June 21, 2019 2:03:54 PM GMT+08:00 | Refresh: Emergency user was locked after proceeding login attempts allowed value. | WIP17 |  |  | Subsystem | Login |  | Duplicate | sanjeeva_a | 0 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00624634 | Refresh | Closed | lakshmikanth_s | June 21, 2019 2:03:54 PM GMT+08:00 | Refresh: Emergency user was locked after proceeding login attempts allowed value. | WIP17 |  |  | Subsystem | Login |  | Duplicate | sanjeeva_a | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00624630 | Refresh | Closed | cherylli | June 21, 2019 10:46:06 AM GMT+08:00 | Refresh Pre-Val:Not possible to merge 10 phases of 1st 4D scan with 2nd 4D scan | BB_WIP6 |  |  | Subsystem | Merge | Console-SRS.5031 | No Plan to Fix | mistry_h | 0 | 4 |  | Support-System | NO | Closed | WAS | Brilliance Big Bore
CTUDT00624628 | Refresh | Closed | zhu_je | June 21, 2019 9:43:47 AM GMT+08:00 | [Refresh] Exam summary was missed when logout workflow without clicking end exam card | BB_WIP6 | BB_WIP8 |  | System | Exam Summary | NA | Fixed | Prakash_Chand_Yadav | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP8_SubCom | Brilliance Big Bore |  | BB_WIP8 | July 1, 2019 7:52:16 PM GMT+08:00
CTUDT00624625 | Refresh | Closed | Sateesh_S | June 20, 2019 9:16:55 PM GMT+08:00 | Incorrect Audit log information displayed for user account enable or disable. | WIP17 | WIP18 BB_WIP8 | N/A | Subsystem | Audit Trail | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP18SubCom | Ovation |  | WIP18 BB_WIP8 | June 26, 2019 7:34:06 PM GMT+08:00
CTUDT00624625 | Refresh | Closed | Sateesh_S | June 20, 2019 9:16:55 PM GMT+08:00 | Incorrect Audit log information displayed for user account enable or disable. | WIP17 | WIP18 BB_WIP8 | N/A | Subsystem | Audit Trail | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP18SubCom | Brilliance 64 |  | WIP18 BB_WIP8 | June 26, 2019 7:34:06 PM GMT+08:00
CTUDT00624625 | Refresh | Closed | Sateesh_S | June 20, 2019 9:16:55 PM GMT+08:00 | Incorrect Audit log information displayed for user account enable or disable. | WIP17 | WIP18 BB_WIP8 | N/A | Subsystem | Audit Trail | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP18SubCom | Brilliance Big Bore |  | WIP18 BB_WIP8 | June 26, 2019 7:34:06 PM GMT+08:00
CTUDT00624625 | Refresh | Closed | Sateesh_S | June 20, 2019 9:16:55 PM GMT+08:00 | Incorrect Audit log information displayed for user account enable or disable. | WIP17 | WIP18 BB_WIP8 | N/A | Subsystem | Audit Trail | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP18SubCom | Ingenuity 64 |  | WIP18 BB_WIP8 | June 26, 2019 7:34:06 PM GMT+08:00
CTUDT00624623 | Refresh | Closed | Sharmila_Vasireddy | June 20, 2019 8:16:51 PM GMT+08:00 | First log entry was not deleted and contained all 31 days logs in service log book | BB_WIP7 | WIP18 BB_WIP8 |  | Subsystem | Service Tools | NA | Fixed | shreekanth_manniyath | 0 | 4 | Console Software | Service Tools | NO | Closed | WIP18SubCom | Ovation |  | WIP18 BB_WIP8 | June 28, 2019 2:34:57 PM GMT+08:00
CTUDT00624623 | Refresh | Closed | Sharmila_Vasireddy | June 20, 2019 8:16:51 PM GMT+08:00 | First log entry was not deleted and contained all 31 days logs in service log book | BB_WIP7 | WIP18 BB_WIP8 |  | Subsystem | Service Tools | NA | Fixed | shreekanth_manniyath | 0 | 4 | Console Software | Service Tools | NO | Closed | WIP18SubCom | Brilliance 64 |  | WIP18 BB_WIP8 | June 28, 2019 2:34:57 PM GMT+08:00
CTUDT00624623 | Refresh | Closed | Sharmila_Vasireddy | June 20, 2019 8:16:51 PM GMT+08:00 | First log entry was not deleted and contained all 31 days logs in service log book | BB_WIP7 | WIP18 BB_WIP8 |  | Subsystem | Service Tools | NA | Fixed | shreekanth_manniyath | 0 | 4 | Console Software | Service Tools | NO | Closed | WIP18SubCom | Brilliance Big Bore |  | WIP18 BB_WIP8 | June 28, 2019 2:34:57 PM GMT+08:00
CTUDT00624623 | Refresh | Closed | Sharmila_Vasireddy | June 20, 2019 8:16:51 PM GMT+08:00 | First log entry was not deleted and contained all 31 days logs in service log book | BB_WIP7 | WIP18 BB_WIP8 |  | Subsystem | Service Tools | NA | Fixed | shreekanth_manniyath | 0 | 4 | Console Software | Service Tools | NO | Closed | WIP18SubCom | Ingenuity 64 |  | WIP18 BB_WIP8 | June 28, 2019 2:34:57 PM GMT+08:00
CTUDT00624623 | Refresh | Closed | Sharmila_Vasireddy | June 20, 2019 8:16:51 PM GMT+08:00 | First log entry was not deleted and contained all 31 days logs in service log book | BB_WIP7 | WIP18 BB_WIP8 |  | Subsystem | Service Tools | NA | Fixed | shreekanth_manniyath | 0 | 4 | Console Software | Service Tools | NO | Closed | WIP18SubCom | IQon1 |  | WIP18 BB_WIP8 | June 28, 2019 2:34:57 PM GMT+08:00
CTUDT00624620 | Refresh | Closed | gT_s | June 20, 2019 6:49:20 PM GMT+08:00 | Refresh : The rotation time value 1.5 is not present for exam card QA Axial Head 2D/3D | WIP17 | WIP18 |  | Subsystem | ECM | NA | Fixed | kumar-san | 0 | 2 | Console Software | Console4 | NO | Closed | WIP18SubCom | Brilliance 64 |  | WIP18 | June 28, 2019 2:42:51 PM GMT+08:00
CTUDT00624620 | Refresh | Closed | gT_s | June 20, 2019 6:49:20 PM GMT+08:00 | Refresh : The rotation time value 1.5 is not present for exam card QA Axial Head 2D/3D | WIP17 | WIP18 |  | Subsystem | ECM | NA | Fixed | kumar-san | 0 | 2 | Console Software | Console4 | NO | Closed | WIP18SubCom | Ingenuity 64 |  | WIP18 | June 28, 2019 2:42:51 PM GMT+08:00
CTUDT00624613 | Refresh | Closed | lei_yang | June 20, 2019 5:54:02 PM GMT+08:00 | The CCT scans volume display without Center, Head and Feet labels on the viewports. | BB_WIP6 | DHF298555 Rev 06 - BB SRS DHF298554 Rev 06 - Ing/BR64 SRS DHF298553 Rev 06 - iCT/iCT-SP SRS DHF298556 Rev 07 - IQon SRS |  | System | CCT | CT-NM.SRS.Modes.1574 | Fixed | mistry_h | 0 | 3 | Requirements Documentation | Support-System | NO | Closed | WIP18SegReq | Ovation |  | DHF298555 Rev 06 - BB SRS DHF298554 Rev 06 - Ing/BR64 SRS DHF298553 Rev 06 - iCT/iCT-SP SRS DHF298556 Rev 07 - IQon SRS | July 26, 2019 8:18:34 PM GMT+08:00
CTUDT00624613 | Refresh | Closed | lei_yang | June 20, 2019 5:54:02 PM GMT+08:00 | The CCT scans volume display without Center, Head and Feet labels on the viewports. | BB_WIP6 | DHF298555 Rev 06 - BB SRS DHF298554 Rev 06 - Ing/BR64 SRS DHF298553 Rev 06 - iCT/iCT-SP SRS DHF298556 Rev 07 - IQon SRS |  | System | CCT | CT-NM.SRS.Modes.1574 | Fixed | mistry_h | 0 | 3 | Requirements Documentation | Support-System | NO | Closed | WIP18SegReq | Brilliance 64 |  | DHF298555 Rev 06 - BB SRS DHF298554 Rev 06 - Ing/BR64 SRS DHF298553 Rev 06 - iCT/iCT-SP SRS DHF298556 Rev 07 - IQon SRS | July 26, 2019 8:18:34 PM GMT+08:00
CTUDT00624613 | Refresh | Closed | lei_yang | June 20, 2019 5:54:02 PM GMT+08:00 | The CCT scans volume display without Center, Head and Feet labels on the viewports. | BB_WIP6 | DHF298555 Rev 06 - BB SRS DHF298554 Rev 06 - Ing/BR64 SRS DHF298553 Rev 06 - iCT/iCT-SP SRS DHF298556 Rev 07 - IQon SRS |  | System | CCT | CT-NM.SRS.Modes.1574 | Fixed | mistry_h | 0 | 3 | Requirements Documentation | Support-System | NO | Closed | WIP18SegReq | Brilliance Big Bore |  | DHF298555 Rev 06 - BB SRS DHF298554 Rev 06 - Ing/BR64 SRS DHF298553 Rev 06 - iCT/iCT-SP SRS DHF298556 Rev 07 - IQon SRS | July 26, 2019 8:18:34 PM GMT+08:00
CTUDT00624613 | Refresh | Closed | lei_yang | June 20, 2019 5:54:02 PM GMT+08:00 | The CCT scans volume display without Center, Head and Feet labels on the viewports. | BB_WIP6 | DHF298555 Rev 06 - BB SRS DHF298554 Rev 06 - Ing/BR64 SRS DHF298553 Rev 06 - iCT/iCT-SP SRS DHF298556 Rev 07 - IQon SRS |  | System | CCT | CT-NM.SRS.Modes.1574 | Fixed | mistry_h | 0 | 3 | Requirements Documentation | Support-System | NO | Closed | WIP18SegReq | Ingenuity 64 |  | DHF298555 Rev 06 - BB SRS DHF298554 Rev 06 - Ing/BR64 SRS DHF298553 Rev 06 - iCT/iCT-SP SRS DHF298556 Rev 07 - IQon SRS | July 26, 2019 8:18:34 PM GMT+08:00
CTUDT00624613 | Refresh | Closed | lei_yang | June 20, 2019 5:54:02 PM GMT+08:00 | The CCT scans volume display without Center, Head and Feet labels on the viewports. | BB_WIP6 | DHF298555 Rev 06 - BB SRS DHF298554 Rev 06 - Ing/BR64 SRS DHF298553 Rev 06 - iCT/iCT-SP SRS DHF298556 Rev 07 - IQon SRS |  | System | CCT | CT-NM.SRS.Modes.1574 | Fixed | mistry_h | 0 | 3 | Requirements Documentation | Support-System | NO | Closed | WIP18SegReq | IQon1 |  | DHF298555 Rev 06 - BB SRS DHF298554 Rev 06 - Ing/BR64 SRS DHF298553 Rev 06 - iCT/iCT-SP SRS DHF298556 Rev 07 - IQon SRS | July 26, 2019 8:18:34 PM GMT+08:00
CTUDT00624607 | Refresh | Closed | PremKumar_Rajendran | June 20, 2019 3:35:15 PM GMT+08:00 | Refresh - In ‘Report’ page, Localized Languages are not displayed in ‘Reporting, Language, Template’ drop down fields | WIP17 |  |  | Subsystem | Localization | NA | Works As Specified / Not A Defect | s_sa | 0 |  |  | iCAP | NO | Closed |  | Ovation
CTUDT00624607 | Refresh | Closed | PremKumar_Rajendran | June 20, 2019 3:35:15 PM GMT+08:00 | Refresh - In ‘Report’ page, Localized Languages are not displayed in ‘Reporting, Language, Template’ drop down fields | WIP17 |  |  | Subsystem | Localization | NA | Works As Specified / Not A Defect | s_sa | 0 |  |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00624607 | Refresh | Closed | PremKumar_Rajendran | June 20, 2019 3:35:15 PM GMT+08:00 | Refresh - In ‘Report’ page, Localized Languages are not displayed in ‘Reporting, Language, Template’ drop down fields | WIP17 |  |  | Subsystem | Localization | NA | Works As Specified / Not A Defect | s_sa | 0 |  |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00624607 | Refresh | Closed | PremKumar_Rajendran | June 20, 2019 3:35:15 PM GMT+08:00 | Refresh - In ‘Report’ page, Localized Languages are not displayed in ‘Reporting, Language, Template’ drop down fields | WIP17 |  |  | Subsystem | Localization | NA | Works As Specified / Not A Defect | s_sa | 0 |  |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00624607 | Refresh | Closed | PremKumar_Rajendran | June 20, 2019 3:35:15 PM GMT+08:00 | Refresh - In ‘Report’ page, Localized Languages are not displayed in ‘Reporting, Language, Template’ drop down fields | WIP17 |  |  | Subsystem | Localization | NA | Works As Specified / Not A Defect | s_sa | 0 |  |  | iCAP | NO | Closed |  | IQon1
CTUDT00624605 | Refresh | Closed | wei_f | June 20, 2019 3:26:51 PM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.Workflow.2460_corresponding evidence chapter marked error in traceability DHF303288 Rev 01_User Doc and Localization Plan. | BB_WIP6 | WIP19 |  | Subsystem | IFU | CT-NM.SRS.Workflow.2460 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Brilliance Big Bore |  | WIP19 | July 31, 2019 4:41:14 PM GMT+08:00
CTUDT00624603 | Refresh | Closed | wei_f | June 20, 2019 3:13:03 PM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.SecPri.346_corresponding evidence marked error in traceability DHF303288 Rev 01_User Doc and Localization Plan. | BB_WIP6 | WIP19 |  | Subsystem | IFU | CT-NM.SRS.SecPri.346 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Ovation |  | WIP19 | July 31, 2019 4:40:04 PM GMT+08:00
CTUDT00624603 | Refresh | Closed | wei_f | June 20, 2019 3:13:03 PM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.SecPri.346_corresponding evidence marked error in traceability DHF303288 Rev 01_User Doc and Localization Plan. | BB_WIP6 | WIP19 |  | Subsystem | IFU | CT-NM.SRS.SecPri.346 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Brilliance 64 |  | WIP19 | July 31, 2019 4:40:04 PM GMT+08:00
CTUDT00624603 | Refresh | Closed | wei_f | June 20, 2019 3:13:03 PM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.SecPri.346_corresponding evidence marked error in traceability DHF303288 Rev 01_User Doc and Localization Plan. | BB_WIP6 | WIP19 |  | Subsystem | IFU | CT-NM.SRS.SecPri.346 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Brilliance Big Bore |  | WIP19 | July 31, 2019 4:40:04 PM GMT+08:00
CTUDT00624603 | Refresh | Closed | wei_f | June 20, 2019 3:13:03 PM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.SecPri.346_corresponding evidence marked error in traceability DHF303288 Rev 01_User Doc and Localization Plan. | BB_WIP6 | WIP19 |  | Subsystem | IFU | CT-NM.SRS.SecPri.346 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Ingenuity 64 |  | WIP19 | July 31, 2019 4:40:04 PM GMT+08:00
CTUDT00624603 | Refresh | Closed | wei_f | June 20, 2019 3:13:03 PM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.SecPri.346_corresponding evidence marked error in traceability DHF303288 Rev 01_User Doc and Localization Plan. | BB_WIP6 | WIP19 |  | Subsystem | IFU | CT-NM.SRS.SecPri.346 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | IQon1 |  | WIP19 | July 31, 2019 4:40:04 PM GMT+08:00
CTUDT00624598 | Refresh | Closed | 320053767 | June 20, 2019 3:03:46 PM GMT+08:00 | Refresh IQon IFU – could not find information in the IFU documentation for [CT-NM.SRS.IQ.1435] | WIP16 | WIP19 |  | Subsystem | IFU | CT-NM.SRS.IQ.1435 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Ovation |  | WIP19 | July 31, 2019 4:49:33 PM GMT+08:00
CTUDT00624598 | Refresh | Closed | 320053767 | June 20, 2019 3:03:46 PM GMT+08:00 | Refresh IQon IFU – could not find information in the IFU documentation for [CT-NM.SRS.IQ.1435] | WIP16 | WIP19 |  | Subsystem | IFU | CT-NM.SRS.IQ.1435 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Brilliance 64 |  | WIP19 | July 31, 2019 4:49:33 PM GMT+08:00
CTUDT00624598 | Refresh | Closed | 320053767 | June 20, 2019 3:03:46 PM GMT+08:00 | Refresh IQon IFU – could not find information in the IFU documentation for [CT-NM.SRS.IQ.1435] | WIP16 | WIP19 |  | Subsystem | IFU | CT-NM.SRS.IQ.1435 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Brilliance Big Bore |  | WIP19 | July 31, 2019 4:49:33 PM GMT+08:00
CTUDT00624598 | Refresh | Closed | 320053767 | June 20, 2019 3:03:46 PM GMT+08:00 | Refresh IQon IFU – could not find information in the IFU documentation for [CT-NM.SRS.IQ.1435] | WIP16 | WIP19 |  | Subsystem | IFU | CT-NM.SRS.IQ.1435 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Ingenuity 64 |  | WIP19 | July 31, 2019 4:49:33 PM GMT+08:00
CTUDT00624598 | Refresh | Closed | 320053767 | June 20, 2019 3:03:46 PM GMT+08:00 | Refresh IQon IFU – could not find information in the IFU documentation for [CT-NM.SRS.IQ.1435] | WIP16 | WIP19 |  | Subsystem | IFU | CT-NM.SRS.IQ.1435 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | IQon1 |  | WIP19 | July 31, 2019 4:49:33 PM GMT+08:00
CTUDT00624597 | Refresh | Closed | sunilkumar_j | June 20, 2019 2:54:10 PM GMT+08:00 | Refresh: Security: Correct message was not displaying for Unauthorized IST dongle | WIP17 |  | Not Required. | Subsystem | Login | NA | Works As Specified / Not A Defect | kapadia_v | 0 |  |  | Console4 | NO | Closed | WAS | Ingenuity 64
CTUDT00624595 | Refresh | Closed | wei_f | June 20, 2019 2:44:02 PM GMT+08:00 | [Refresh][IFU Kits] There are many requirements linked to Tech_Pub but missing in the traceability DHF303288 Rev 01_User Doc and Localization Plan. | BB_WIP6 |  | None | Subsystem | IFU | Refer Attachment | Duplicate | charles_j | 0 | 2 |  | Support-System | NO | Closed | PostVerification | Ovation
CTUDT00624595 | Refresh | Closed | wei_f | June 20, 2019 2:44:02 PM GMT+08:00 | [Refresh][IFU Kits] There are many requirements linked to Tech_Pub but missing in the traceability DHF303288 Rev 01_User Doc and Localization Plan. | BB_WIP6 |  | None | Subsystem | IFU | Refer Attachment | Duplicate | charles_j | 0 | 2 |  | Support-System | NO | Closed | PostVerification | Brilliance 64
CTUDT00624595 | Refresh | Closed | wei_f | June 20, 2019 2:44:02 PM GMT+08:00 | [Refresh][IFU Kits] There are many requirements linked to Tech_Pub but missing in the traceability DHF303288 Rev 01_User Doc and Localization Plan. | BB_WIP6 |  | None | Subsystem | IFU | Refer Attachment | Duplicate | charles_j | 0 | 2 |  | Support-System | NO | Closed | PostVerification | Brilliance Big Bore
CTUDT00624595 | Refresh | Closed | wei_f | June 20, 2019 2:44:02 PM GMT+08:00 | [Refresh][IFU Kits] There are many requirements linked to Tech_Pub but missing in the traceability DHF303288 Rev 01_User Doc and Localization Plan. | BB_WIP6 |  | None | Subsystem | IFU | Refer Attachment | Duplicate | charles_j | 0 | 2 |  | Support-System | NO | Closed | PostVerification | Ingenuity 64
CTUDT00624595 | Refresh | Closed | wei_f | June 20, 2019 2:44:02 PM GMT+08:00 | [Refresh][IFU Kits] There are many requirements linked to Tech_Pub but missing in the traceability DHF303288 Rev 01_User Doc and Localization Plan. | BB_WIP6 |  | None | Subsystem | IFU | Refer Attachment | Duplicate | charles_j | 0 | 2 |  | Support-System | NO | Closed | PostVerification | IQon1
CTUDT00624593 | Refresh | Closed | wei_f | June 20, 2019 2:17:38 PM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.IQ.1434_There is no section "perform long tube conditioning when tube arcing artifacts are observed in the scanned images" in the IFU kits. | BB_WIP6 |  |  | Subsystem | IFU | CT-NM.SRS.IQ.1434 | Duplicate | charles_j | 0 | 2 |  | Support-System | NO | Closed | PostVerification | Brilliance 64
CTUDT00624593 | Refresh | Closed | wei_f | June 20, 2019 2:17:38 PM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.IQ.1434_There is no section "perform long tube conditioning when tube arcing artifacts are observed in the scanned images" in the IFU kits. | BB_WIP6 |  |  | Subsystem | IFU | CT-NM.SRS.IQ.1434 | Duplicate | charles_j | 0 | 2 |  | Support-System | NO | Closed | PostVerification | Brilliance Big Bore
CTUDT00624593 | Refresh | Closed | wei_f | June 20, 2019 2:17:38 PM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.IQ.1434_There is no section "perform long tube conditioning when tube arcing artifacts are observed in the scanned images" in the IFU kits. | BB_WIP6 |  |  | Subsystem | IFU | CT-NM.SRS.IQ.1434 | Duplicate | charles_j | 0 | 2 |  | Support-System | NO | Closed | PostVerification | iCT4
CTUDT00624593 | Refresh | Closed | wei_f | June 20, 2019 2:17:38 PM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.IQ.1434_There is no section "perform long tube conditioning when tube arcing artifacts are observed in the scanned images" in the IFU kits. | BB_WIP6 |  |  | Subsystem | IFU | CT-NM.SRS.IQ.1434 | Duplicate | charles_j | 0 | 2 |  | Support-System | NO | Closed | PostVerification | Ingenuity 64
CTUDT00624593 | Refresh | Closed | wei_f | June 20, 2019 2:17:38 PM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.IQ.1434_There is no section "perform long tube conditioning when tube arcing artifacts are observed in the scanned images" in the IFU kits. | BB_WIP6 |  |  | Subsystem | IFU | CT-NM.SRS.IQ.1434 | Duplicate | charles_j | 0 | 2 |  | Support-System | NO | Closed | PostVerification | IQon1
CTUDT00624591 | Refresh | Closed | wei_f | June 20, 2019 1:57:01 PM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.Service.162_can not find the section "terminate the remote access connection if the connection between the local user and the remote user over phone call is lost" in the IFU Kits | BB_WIP6 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.Service.162 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | YES | Closed | PostVerification | Ovation |  | WIP19 | July 31, 2019 4:34:11 PM GMT+08:00
CTUDT00624591 | Refresh | Closed | wei_f | June 20, 2019 1:57:01 PM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.Service.162_can not find the section "terminate the remote access connection if the connection between the local user and the remote user over phone call is lost" in the IFU Kits | BB_WIP6 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.Service.162 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | YES | Closed | PostVerification | Brilliance 64 |  | WIP19 | July 31, 2019 4:34:11 PM GMT+08:00
CTUDT00624591 | Refresh | Closed | wei_f | June 20, 2019 1:57:01 PM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.Service.162_can not find the section "terminate the remote access connection if the connection between the local user and the remote user over phone call is lost" in the IFU Kits | BB_WIP6 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.Service.162 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | YES | Closed | PostVerification | Brilliance Big Bore |  | WIP19 | July 31, 2019 4:34:11 PM GMT+08:00
CTUDT00624591 | Refresh | Closed | wei_f | June 20, 2019 1:57:01 PM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.Service.162_can not find the section "terminate the remote access connection if the connection between the local user and the remote user over phone call is lost" in the IFU Kits | BB_WIP6 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.Service.162 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | YES | Closed | PostVerification | Ingenuity 64 |  | WIP19 | July 31, 2019 4:34:11 PM GMT+08:00
CTUDT00624591 | Refresh | Closed | wei_f | June 20, 2019 1:57:01 PM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.Service.162_can not find the section "terminate the remote access connection if the connection between the local user and the remote user over phone call is lost" in the IFU Kits | BB_WIP6 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.Service.162 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | YES | Closed | PostVerification | IQon1 |  | WIP19 | July 31, 2019 4:34:11 PM GMT+08:00
CTUDT00624589 | Refresh | Closed | wei_f | June 20, 2019 1:41:27 PM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.Service.2447_There is no Safety guidelines in Disclaimer and in User manual for the Local user. | BB_WIP6 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.Service.2447 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | YES | Closed | PostVerification | Ovation |  | WIP19 | July 31, 2019 4:33:01 PM GMT+08:00
CTUDT00624589 | Refresh | Closed | wei_f | June 20, 2019 1:41:27 PM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.Service.2447_There is no Safety guidelines in Disclaimer and in User manual for the Local user. | BB_WIP6 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.Service.2447 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | YES | Closed | PostVerification | Brilliance 64 |  | WIP19 | July 31, 2019 4:33:01 PM GMT+08:00
CTUDT00624589 | Refresh | Closed | wei_f | June 20, 2019 1:41:27 PM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.Service.2447_There is no Safety guidelines in Disclaimer and in User manual for the Local user. | BB_WIP6 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.Service.2447 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | YES | Closed | PostVerification | Brilliance Big Bore |  | WIP19 | July 31, 2019 4:33:01 PM GMT+08:00
CTUDT00624589 | Refresh | Closed | wei_f | June 20, 2019 1:41:27 PM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.Service.2447_There is no Safety guidelines in Disclaimer and in User manual for the Local user. | BB_WIP6 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.Service.2447 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | YES | Closed | PostVerification | Ingenuity 64 |  | WIP19 | July 31, 2019 4:33:01 PM GMT+08:00
CTUDT00624589 | Refresh | Closed | wei_f | June 20, 2019 1:41:27 PM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.Service.2447_There is no Safety guidelines in Disclaimer and in User manual for the Local user. | BB_WIP6 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.Service.2447 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | YES | Closed | PostVerification | IQon1 |  | WIP19 | July 31, 2019 4:33:01 PM GMT+08:00
CTUDT00624587 | Refresh | Closed | wei_f | June 20, 2019 1:24:03 PM GMT+08:00 | [Refresh][IFU Kits]CT-NM.SRS.RegStand.741_the symbol is defined as ISO 15223-1:2012 in the TRG, not the ISO15223-1:2016. | BB_WIP6 | WIP19 |  | Subsystem | IFU | CT-NM.SRS.RegStand.741 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Ovation |  | WIP19 | July 31, 2019 4:30:26 PM GMT+08:00
CTUDT00624587 | Refresh | Closed | wei_f | June 20, 2019 1:24:03 PM GMT+08:00 | [Refresh][IFU Kits]CT-NM.SRS.RegStand.741_the symbol is defined as ISO 15223-1:2012 in the TRG, not the ISO15223-1:2016. | BB_WIP6 | WIP19 |  | Subsystem | IFU | CT-NM.SRS.RegStand.741 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Brilliance 64 |  | WIP19 | July 31, 2019 4:30:26 PM GMT+08:00
CTUDT00624587 | Refresh | Closed | wei_f | June 20, 2019 1:24:03 PM GMT+08:00 | [Refresh][IFU Kits]CT-NM.SRS.RegStand.741_the symbol is defined as ISO 15223-1:2012 in the TRG, not the ISO15223-1:2016. | BB_WIP6 | WIP19 |  | Subsystem | IFU | CT-NM.SRS.RegStand.741 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Brilliance Big Bore |  | WIP19 | July 31, 2019 4:30:26 PM GMT+08:00
CTUDT00624587 | Refresh | Closed | wei_f | June 20, 2019 1:24:03 PM GMT+08:00 | [Refresh][IFU Kits]CT-NM.SRS.RegStand.741_the symbol is defined as ISO 15223-1:2012 in the TRG, not the ISO15223-1:2016. | BB_WIP6 | WIP19 |  | Subsystem | IFU | CT-NM.SRS.RegStand.741 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Ingenuity 64 |  | WIP19 | July 31, 2019 4:30:26 PM GMT+08:00
CTUDT00624587 | Refresh | Closed | wei_f | June 20, 2019 1:24:03 PM GMT+08:00 | [Refresh][IFU Kits]CT-NM.SRS.RegStand.741_the symbol is defined as ISO 15223-1:2012 in the TRG, not the ISO15223-1:2016. | BB_WIP6 | WIP19 |  | Subsystem | IFU | CT-NM.SRS.RegStand.741 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | IQon1 |  | WIP19 | July 31, 2019 4:30:26 PM GMT+08:00
CTUDT00624583 | Refresh | Closed | zhu_je | June 20, 2019 11:16:14 AM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.RegStand.741_Cannot find "labeling and labels symbols shall comply with ISO 15223-1:2016 " in TRG file | WIP16 | WIP20 |  | Subsystem | IFU;Verified1 | CT-NM.SRS.RegStand.741 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostValidation | Brilliance 64 |  | WIP20 | November 6, 2019 2:46:13 PM GMT+08:00
CTUDT00624583 | Refresh | Closed | zhu_je | June 20, 2019 11:16:14 AM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.RegStand.741_Cannot find "labeling and labels symbols shall comply with ISO 15223-1:2016 " in TRG file | WIP16 | WIP20 |  | Subsystem | IFU;Verified1 | CT-NM.SRS.RegStand.741 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostValidation | Brilliance Big Bore |  | WIP20 | November 6, 2019 2:46:13 PM GMT+08:00
CTUDT00624583 | Refresh | Closed | zhu_je | June 20, 2019 11:16:14 AM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.RegStand.741_Cannot find "labeling and labels symbols shall comply with ISO 15223-1:2016 " in TRG file | WIP16 | WIP20 |  | Subsystem | IFU;Verified1 | CT-NM.SRS.RegStand.741 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostValidation | Ingenuity 64 |  | WIP20 | November 6, 2019 2:46:13 PM GMT+08:00
CTUDT00624581 | Refresh | Closed | zhu_je | June 20, 2019 11:00:24 AM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.Modes.514_the description in section “Respiratory Mode” in IFU file is not same as requirement | WIP16 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.Modes.514 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Ovation |  | WIP19 | July 31, 2019 4:26:35 PM GMT+08:00
CTUDT00624581 | Refresh | Closed | zhu_je | June 20, 2019 11:00:24 AM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.Modes.514_the description in section “Respiratory Mode” in IFU file is not same as requirement | WIP16 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.Modes.514 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Brilliance 64 |  | WIP19 | July 31, 2019 4:26:35 PM GMT+08:00
CTUDT00624581 | Refresh | Closed | zhu_je | June 20, 2019 11:00:24 AM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.Modes.514_the description in section “Respiratory Mode” in IFU file is not same as requirement | WIP16 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.Modes.514 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Brilliance Big Bore |  | WIP19 | July 31, 2019 4:26:35 PM GMT+08:00
CTUDT00624581 | Refresh | Closed | zhu_je | June 20, 2019 11:00:24 AM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.Modes.514_the description in section “Respiratory Mode” in IFU file is not same as requirement | WIP16 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.Modes.514 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Ingenuity 64 |  | WIP19 | July 31, 2019 4:26:35 PM GMT+08:00
CTUDT00624581 | Refresh | Closed | zhu_je | June 20, 2019 11:00:24 AM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.Modes.514_the description in section “Respiratory Mode” in IFU file is not same as requirement | WIP16 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.Modes.514 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | IQon1 |  | WIP19 | July 31, 2019 4:26:35 PM GMT+08:00
CTUDT00624579 | Refresh | Closed | wang_l | June 20, 2019 10:17:44 AM GMT+08:00 | Refresh BB: Service tools displayed does not match with requirement | BB_WIP6 | DHF298555 Rev 06 - BB SRS DHF298554 Rev 06 - Ing/BR64 SRS DHF298553 Rev 06 - iCT/iCT-SP SRS |  | System | Service Tools | CT-NM.SRS.Service.2992 | Fixed | mistry_h | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | BB_WIP8_SubCom | Brilliance Big Bore |  | DHF298555 Rev 06 - BB SRS DHF298554 Rev 06 - Ing/BR64 SRS DHF298553 Rev 06 - iCT/iCT-SP SRS | July 26, 2019 8:17:46 PM GMT+08:00
CTUDT00624577 | Refresh | Closed | cohen_y | June 20, 2019 1:07:00 AM GMT+08:00 | Refresh: Performance Test: One Part Phantom: iDose CT value limits are not as defined in spec. | BB_WIP7 | N/A |  | Subsystem | Physics | N/A | Fixed | braunstein_d | 0 | 2 | Design Documentation | Support-Physics | NO | Closed | BB_WIP8_SubCom | Brilliance Big Bore |  | N/A | July 15, 2019 6:39:45 PM GMT+08:00
CTUDT00624575 | Refresh | Closed | beck_d | June 20, 2019 12:33:30 AM GMT+08:00 | Refresh: Performance Test: Harmonized Phantom: Head CT3 collimation is not according to spec | BB_WIP7 | WIP18 | N/A | Subsystem | Performance Test | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | YES | Closed | BB_WIP8_SubCom | Brilliance Big Bore |  | WIP18 | June 30, 2019 4:39:41 PM GMT+08:00
CTUDT00624573 | Refresh | Closed | cohen_y | June 20, 2019 12:22:25 AM GMT+08:00 | Refresh: Performance Test: One Part Phantom: Recon Matrix of Surview scans is not according to Spec | BB_WIP7 | N/A | N/A | Subsystem | Documentation | N/A | Fixed | braunstein_d | 0 | 2 | Design Documentation | Support-Physics | NO | Closed | BB_WIP8_SegReq | Brilliance Big Bore |  | N/A | July 4, 2019 8:05:26 PM GMT+08:00
CTUDT00624567 | Refresh | Closed | 320053768 | June 19, 2019 8:08:32 PM GMT+08:00 | Refresh IQon FU – could not find information in the IFU documentation for <CT-NM.SRS.Workflow.2161> | WIP17 | WIP19 |  | Subsystem | IFU | CT-NM.SRS.Workflow.2161 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | YES | Closed | PostVerification | IQon1 |  | WIP19 | July 31, 2019 4:48:44 PM GMT+08:00
CTUDT00624565 | Refresh | Closed | 320053768 | June 19, 2019 8:06:59 PM GMT+08:00 | Refresh IQon IFU – could not find information in all IFU documentation for "CT-NM.SRS.Workflow.2162" | WIP17 | WIP19 |  | Subsystem | IFU | CT-NM.SRS.Workflow.2162 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | IQon1 |  | WIP19 | July 31, 2019 4:44:38 PM GMT+08:00
CTUDT00624556 | Refresh | Closed | lei_yang | June 19, 2019 4:40:49 PM GMT+08:00 | The time-stamps on the reconstructed images did not match the time-stamps of the HOST Log Viewer. | WIP16 |  | N/A | Subsystem |  | N/A | Works As Specified / Not A Defect | Shai_Goldberg | 0 |  |  | Console4 | NO | Closed | WAS | Ovation
CTUDT00624556 | Refresh | Closed | lei_yang | June 19, 2019 4:40:49 PM GMT+08:00 | The time-stamps on the reconstructed images did not match the time-stamps of the HOST Log Viewer. | WIP16 |  | N/A | Subsystem |  | N/A | Works As Specified / Not A Defect | Shai_Goldberg | 0 |  |  | Console4 | NO | Closed | WAS | Brilliance 64
CTUDT00624556 | Refresh | Closed | lei_yang | June 19, 2019 4:40:49 PM GMT+08:00 | The time-stamps on the reconstructed images did not match the time-stamps of the HOST Log Viewer. | WIP16 |  | N/A | Subsystem |  | N/A | Works As Specified / Not A Defect | Shai_Goldberg | 0 |  |  | Console4 | NO | Closed | WAS | Ingenuity 64
CTUDT00624556 | Refresh | Closed | lei_yang | June 19, 2019 4:40:49 PM GMT+08:00 | The time-stamps on the reconstructed images did not match the time-stamps of the HOST Log Viewer. | WIP16 |  | N/A | Subsystem |  | N/A | Works As Specified / Not A Defect | Shai_Goldberg | 0 |  |  | Console4 | NO | Closed | WAS | IQon1
CTUDT00624552 | Refresh | Closed | ding_s | June 19, 2019 3:57:21 PM GMT+08:00 | Licensing in Preferences show wrong functionality which were not included in corresponding product. | WIP16 | WIP22.2 | NA | Subsystem |  | NA | Fixed | s_panigrahi | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | SubCom-25-Nov-2019 | Brilliance 64 |  | WIP22.2 | November 29, 2019 7:22:32 PM GMT+08:00
CTUDT00624552 | Refresh | Closed | ding_s | June 19, 2019 3:57:21 PM GMT+08:00 | Licensing in Preferences show wrong functionality which were not included in corresponding product. | WIP16 | WIP22.2 | NA | Subsystem |  | NA | Fixed | s_panigrahi | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | SubCom-25-Nov-2019 | Brilliance Big Bore |  | WIP22.2 | November 29, 2019 7:22:32 PM GMT+08:00
CTUDT00624552 | Refresh | Closed | ding_s | June 19, 2019 3:57:21 PM GMT+08:00 | Licensing in Preferences show wrong functionality which were not included in corresponding product. | WIP16 | WIP22.2 | NA | Subsystem |  | NA | Fixed | s_panigrahi | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | SubCom-25-Nov-2019 | Ingenuity 64 |  | WIP22.2 | November 29, 2019 7:22:32 PM GMT+08:00
CTUDT00624549 | Refresh | Closed | zhu_je | June 19, 2019 2:18:03 PM GMT+08:00 | [Refresh][IFU Kits]CT-NM.SRS.SysSupport.781_ instructions for performing system startup/shutdown of the system for TRG is different with IFU/Maintance file | WIP16 | WIP20 |  | Subsystem | IFU;Verified1 | CT-NM.SRS.SysSupport.781 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostValidation | Brilliance 64 |  | WIP20 | November 6, 2019 2:46:39 PM GMT+08:00
CTUDT00624549 | Refresh | Closed | zhu_je | June 19, 2019 2:18:03 PM GMT+08:00 | [Refresh][IFU Kits]CT-NM.SRS.SysSupport.781_ instructions for performing system startup/shutdown of the system for TRG is different with IFU/Maintance file | WIP16 | WIP20 |  | Subsystem | IFU;Verified1 | CT-NM.SRS.SysSupport.781 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostValidation | Brilliance Big Bore |  | WIP20 | November 6, 2019 2:46:39 PM GMT+08:00
CTUDT00624549 | Refresh | Closed | zhu_je | June 19, 2019 2:18:03 PM GMT+08:00 | [Refresh][IFU Kits]CT-NM.SRS.SysSupport.781_ instructions for performing system startup/shutdown of the system for TRG is different with IFU/Maintance file | WIP16 | WIP20 |  | Subsystem | IFU;Verified1 | CT-NM.SRS.SysSupport.781 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostValidation | Ingenuity 64 |  | WIP20 | November 6, 2019 2:46:39 PM GMT+08:00
CTUDT00624547 | Refresh | Closed | zhu_je | June 19, 2019 2:07:32 PM GMT+08:00 | [Refresh][IFU Kits]CT-NM.SRS.RegStand.742_ Cannot find the packing integrity in TRG file | WIP16 |  |  | Subsystem | IFU;Blocker | CT-NM.SRS.RegStand.742 | Works As Specified / Not A Defect | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Ovation |  |  | July 31, 2019 4:35:40 PM GMT+08:00
CTUDT00624547 | Refresh | Closed | zhu_je | June 19, 2019 2:07:32 PM GMT+08:00 | [Refresh][IFU Kits]CT-NM.SRS.RegStand.742_ Cannot find the packing integrity in TRG file | WIP16 |  |  | Subsystem | IFU;Blocker | CT-NM.SRS.RegStand.742 | Works As Specified / Not A Defect | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Brilliance 64 |  |  | July 31, 2019 4:35:40 PM GMT+08:00
CTUDT00624547 | Refresh | Closed | zhu_je | June 19, 2019 2:07:32 PM GMT+08:00 | [Refresh][IFU Kits]CT-NM.SRS.RegStand.742_ Cannot find the packing integrity in TRG file | WIP16 |  |  | Subsystem | IFU;Blocker | CT-NM.SRS.RegStand.742 | Works As Specified / Not A Defect | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Ingenuity 64 |  |  | July 31, 2019 4:35:40 PM GMT+08:00
CTUDT00624547 | Refresh | Closed | zhu_je | June 19, 2019 2:07:32 PM GMT+08:00 | [Refresh][IFU Kits]CT-NM.SRS.RegStand.742_ Cannot find the packing integrity in TRG file | WIP16 |  |  | Subsystem | IFU;Blocker | CT-NM.SRS.RegStand.742 | Works As Specified / Not A Defect | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Ingenuity BigBore |  |  | July 31, 2019 4:35:40 PM GMT+08:00
CTUDT00624547 | Refresh | Closed | zhu_je | June 19, 2019 2:07:32 PM GMT+08:00 | [Refresh][IFU Kits]CT-NM.SRS.RegStand.742_ Cannot find the packing integrity in TRG file | WIP16 |  |  | Subsystem | IFU;Blocker | CT-NM.SRS.RegStand.742 | Works As Specified / Not A Defect | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | IQon1 |  |  | July 31, 2019 4:35:40 PM GMT+08:00
CTUDT00624545 | Refresh | Closed | kr_m | June 19, 2019 2:03:21 PM GMT+08:00 | Refresh: The options in the context menu of "Reference Series" displayed the results before the reconstruction is completed. | WIP17 | WIP18 | NA | Subsystem | CCT | NA | Fixed | kr_m | 0 | 4 | Console Software | Console4 | NO | Closed | WIP18SubCom | Ovation |  | WIP18 | June 28, 2019 5:19:48 PM GMT+08:00
CTUDT00624545 | Refresh | Closed | kr_m | June 19, 2019 2:03:21 PM GMT+08:00 | Refresh: The options in the context menu of "Reference Series" displayed the results before the reconstruction is completed. | WIP17 | WIP18 | NA | Subsystem | CCT | NA | Fixed | kr_m | 0 | 4 | Console Software | Console4 | NO | Closed | WIP18SubCom | Brilliance 64 |  | WIP18 | June 28, 2019 5:19:48 PM GMT+08:00
CTUDT00624545 | Refresh | Closed | kr_m | June 19, 2019 2:03:21 PM GMT+08:00 | Refresh: The options in the context menu of "Reference Series" displayed the results before the reconstruction is completed. | WIP17 | WIP18 | NA | Subsystem | CCT | NA | Fixed | kr_m | 0 | 4 | Console Software | Console4 | NO | Closed | WIP18SubCom | Ingenuity 64 |  | WIP18 | June 28, 2019 5:19:48 PM GMT+08:00
CTUDT00624545 | Refresh | Closed | kr_m | June 19, 2019 2:03:21 PM GMT+08:00 | Refresh: The options in the context menu of "Reference Series" displayed the results before the reconstruction is completed. | WIP17 | WIP18 | NA | Subsystem | CCT | NA | Fixed | kr_m | 0 | 4 | Console Software | Console4 | NO | Closed | WIP18SubCom | IQon1 |  | WIP18 | June 28, 2019 5:19:48 PM GMT+08:00
CTUDT00624542 | Refresh | Closed | zhu_je | June 19, 2019 1:52:44 PM GMT+08:00 | [Refresh][IFU Kits]CT-NM.SRS.RegStand.737_ Cannot find subsection entitled "Quality Control Testing Recommendations" in a Pediatric section in IFU file | WIP16 | WIP19 | N/A | Subsystem | Blocker | CT-NM.SRS.RegStand.737 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Ovation |  | WIP19 | August 9, 2019 3:59:25 PM GMT+08:00
CTUDT00624542 | Refresh | Closed | zhu_je | June 19, 2019 1:52:44 PM GMT+08:00 | [Refresh][IFU Kits]CT-NM.SRS.RegStand.737_ Cannot find subsection entitled "Quality Control Testing Recommendations" in a Pediatric section in IFU file | WIP16 | WIP19 | N/A | Subsystem | Blocker | CT-NM.SRS.RegStand.737 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Brilliance 64 |  | WIP19 | August 9, 2019 3:59:25 PM GMT+08:00
CTUDT00624542 | Refresh | Closed | zhu_je | June 19, 2019 1:52:44 PM GMT+08:00 | [Refresh][IFU Kits]CT-NM.SRS.RegStand.737_ Cannot find subsection entitled "Quality Control Testing Recommendations" in a Pediatric section in IFU file | WIP16 | WIP19 | N/A | Subsystem | Blocker | CT-NM.SRS.RegStand.737 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Brilliance Big Bore |  | WIP19 | August 9, 2019 3:59:25 PM GMT+08:00
CTUDT00624542 | Refresh | Closed | zhu_je | June 19, 2019 1:52:44 PM GMT+08:00 | [Refresh][IFU Kits]CT-NM.SRS.RegStand.737_ Cannot find subsection entitled "Quality Control Testing Recommendations" in a Pediatric section in IFU file | WIP16 | WIP19 | N/A | Subsystem | Blocker | CT-NM.SRS.RegStand.737 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Ingenuity 64 |  | WIP19 | August 9, 2019 3:59:25 PM GMT+08:00
CTUDT00624542 | Refresh | Closed | zhu_je | June 19, 2019 1:52:44 PM GMT+08:00 | [Refresh][IFU Kits]CT-NM.SRS.RegStand.737_ Cannot find subsection entitled "Quality Control Testing Recommendations" in a Pediatric section in IFU file | WIP16 | WIP19 | N/A | Subsystem | Blocker | CT-NM.SRS.RegStand.737 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | IQon1 |  | WIP19 | August 9, 2019 3:59:25 PM GMT+08:00
CTUDT00624540 | Refresh | Closed | zhu_je | June 19, 2019 1:47:57 PM GMT+08:00 | [Refresh][IFU Kits]CT-NM.SRS.RegStand.736_ Cannot find subsection entitled "Quality Control Testing Recommendations" in a Pediatric section in IFU file | WIP16 | WIP19 | N/A | Subsystem | IFU;Blocker | CT-NM.SRS.RegStand.736 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Ovation |  | WIP19 | August 9, 2019 3:59:46 PM GMT+08:00
CTUDT00624540 | Refresh | Closed | zhu_je | June 19, 2019 1:47:57 PM GMT+08:00 | [Refresh][IFU Kits]CT-NM.SRS.RegStand.736_ Cannot find subsection entitled "Quality Control Testing Recommendations" in a Pediatric section in IFU file | WIP16 | WIP19 | N/A | Subsystem | IFU;Blocker | CT-NM.SRS.RegStand.736 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Brilliance 64 |  | WIP19 | August 9, 2019 3:59:46 PM GMT+08:00
CTUDT00624540 | Refresh | Closed | zhu_je | June 19, 2019 1:47:57 PM GMT+08:00 | [Refresh][IFU Kits]CT-NM.SRS.RegStand.736_ Cannot find subsection entitled "Quality Control Testing Recommendations" in a Pediatric section in IFU file | WIP16 | WIP19 | N/A | Subsystem | IFU;Blocker | CT-NM.SRS.RegStand.736 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Brilliance Big Bore |  | WIP19 | August 9, 2019 3:59:46 PM GMT+08:00
CTUDT00624540 | Refresh | Closed | zhu_je | June 19, 2019 1:47:57 PM GMT+08:00 | [Refresh][IFU Kits]CT-NM.SRS.RegStand.736_ Cannot find subsection entitled "Quality Control Testing Recommendations" in a Pediatric section in IFU file | WIP16 | WIP19 | N/A | Subsystem | IFU;Blocker | CT-NM.SRS.RegStand.736 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Ingenuity 64 |  | WIP19 | August 9, 2019 3:59:46 PM GMT+08:00
CTUDT00624540 | Refresh | Closed | zhu_je | June 19, 2019 1:47:57 PM GMT+08:00 | [Refresh][IFU Kits]CT-NM.SRS.RegStand.736_ Cannot find subsection entitled "Quality Control Testing Recommendations" in a Pediatric section in IFU file | WIP16 | WIP19 | N/A | Subsystem | IFU;Blocker | CT-NM.SRS.RegStand.736 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | IQon1 |  | WIP19 | August 9, 2019 3:59:46 PM GMT+08:00
CTUDT00624539 | Refresh | Closed | kr_m | June 19, 2019 1:45:34 PM GMT+08:00 | Refresh: Change request to change the message that is notified to the user for the requirement Console-SRS.15087 | WIP17 | WIP18 | NA | Subsystem | System | NA | Fixed | kr_m | 0 | 4 | Console Software | Console4 | NO | Closed | WIP18SubCom | Brilliance 64 |  | WIP18 | June 26, 2019 7:42:57 PM GMT+08:00
CTUDT00624539 | Refresh | Closed | kr_m | June 19, 2019 1:45:34 PM GMT+08:00 | Refresh: Change request to change the message that is notified to the user for the requirement Console-SRS.15087 | WIP17 | WIP18 | NA | Subsystem | System | NA | Fixed | kr_m | 0 | 4 | Console Software | Console4 | NO | Closed | WIP18SubCom | Ingenuity 64 |  | WIP18 | June 26, 2019 7:42:57 PM GMT+08:00
CTUDT00624537 | Refresh | Closed | zhu_je | June 19, 2019 1:43:32 PM GMT+08:00 | [Refresh][IFU Kits]CT-NM.SRS.RegStand.735_ Cannot find Pediatric Safety Checklist that describes any special pediatric issues during pre-acquisition, acquisition, and post-acquisition steps in IFU file | WIP16 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.RegStand.735 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Ovation |  | WIP19 | July 31, 2019 4:59:05 PM GMT+08:00
CTUDT00624537 | Refresh | Closed | zhu_je | June 19, 2019 1:43:32 PM GMT+08:00 | [Refresh][IFU Kits]CT-NM.SRS.RegStand.735_ Cannot find Pediatric Safety Checklist that describes any special pediatric issues during pre-acquisition, acquisition, and post-acquisition steps in IFU file | WIP16 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.RegStand.735 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Brilliance 64 |  | WIP19 | July 31, 2019 4:59:05 PM GMT+08:00
CTUDT00624537 | Refresh | Closed | zhu_je | June 19, 2019 1:43:32 PM GMT+08:00 | [Refresh][IFU Kits]CT-NM.SRS.RegStand.735_ Cannot find Pediatric Safety Checklist that describes any special pediatric issues during pre-acquisition, acquisition, and post-acquisition steps in IFU file | WIP16 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.RegStand.735 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Brilliance Big Bore |  | WIP19 | July 31, 2019 4:59:05 PM GMT+08:00
CTUDT00624537 | Refresh | Closed | zhu_je | June 19, 2019 1:43:32 PM GMT+08:00 | [Refresh][IFU Kits]CT-NM.SRS.RegStand.735_ Cannot find Pediatric Safety Checklist that describes any special pediatric issues during pre-acquisition, acquisition, and post-acquisition steps in IFU file | WIP16 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.RegStand.735 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Ingenuity 64 |  | WIP19 | July 31, 2019 4:59:05 PM GMT+08:00
CTUDT00624537 | Refresh | Closed | zhu_je | June 19, 2019 1:43:32 PM GMT+08:00 | [Refresh][IFU Kits]CT-NM.SRS.RegStand.735_ Cannot find Pediatric Safety Checklist that describes any special pediatric issues during pre-acquisition, acquisition, and post-acquisition steps in IFU file | WIP16 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.RegStand.735 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | IQon1 |  | WIP19 | July 31, 2019 4:59:05 PM GMT+08:00
CTUDT00624535 | Refresh | Closed | zhu_je | June 19, 2019 12:46:20 PM GMT+08:00 | [Refresh][IFU Kits]CT-NM.SRS.RegStand.734_ Cannot find subsection entitled "Instructions for Use" in the section "Pediatric & Small patients" that explains user-configurable imaging parameters that impact dose and image quality in IFU file | WIP16 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.RegStand.734 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Ovation |  | WIP19 | July 31, 2019 4:58:13 PM GMT+08:00
CTUDT00624535 | Refresh | Closed | zhu_je | June 19, 2019 12:46:20 PM GMT+08:00 | [Refresh][IFU Kits]CT-NM.SRS.RegStand.734_ Cannot find subsection entitled "Instructions for Use" in the section "Pediatric & Small patients" that explains user-configurable imaging parameters that impact dose and image quality in IFU file | WIP16 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.RegStand.734 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Brilliance 64 |  | WIP19 | July 31, 2019 4:58:13 PM GMT+08:00
CTUDT00624535 | Refresh | Closed | zhu_je | June 19, 2019 12:46:20 PM GMT+08:00 | [Refresh][IFU Kits]CT-NM.SRS.RegStand.734_ Cannot find subsection entitled "Instructions for Use" in the section "Pediatric & Small patients" that explains user-configurable imaging parameters that impact dose and image quality in IFU file | WIP16 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.RegStand.734 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Brilliance Big Bore |  | WIP19 | July 31, 2019 4:58:13 PM GMT+08:00
CTUDT00624535 | Refresh | Closed | zhu_je | June 19, 2019 12:46:20 PM GMT+08:00 | [Refresh][IFU Kits]CT-NM.SRS.RegStand.734_ Cannot find subsection entitled "Instructions for Use" in the section "Pediatric & Small patients" that explains user-configurable imaging parameters that impact dose and image quality in IFU file | WIP16 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.RegStand.734 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Ingenuity 64 |  | WIP19 | July 31, 2019 4:58:13 PM GMT+08:00
CTUDT00624535 | Refresh | Closed | zhu_je | June 19, 2019 12:46:20 PM GMT+08:00 | [Refresh][IFU Kits]CT-NM.SRS.RegStand.734_ Cannot find subsection entitled "Instructions for Use" in the section "Pediatric & Small patients" that explains user-configurable imaging parameters that impact dose and image quality in IFU file | WIP16 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.RegStand.734 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | IQon1 |  | WIP19 | July 31, 2019 4:58:13 PM GMT+08:00
CTUDT00624533 | Refresh | Closed | zhu_je | June 19, 2019 12:41:59 PM GMT+08:00 | [Refresh][IFU Kits]CT-NM.SRS.RegStand.733_ Cannot find subsection “Instruction for use” in the section "Pediatric & Small patients" that provides radiation dose reduction strategies for pediatric patients in IFU file | WIP16 | WIP19 |  | Subsystem | Blocker | CT-NM.SRS.RegStand.733 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Ovation |  | WIP19 | July 31, 2019 4:57:23 PM GMT+08:00
CTUDT00624533 | Refresh | Closed | zhu_je | June 19, 2019 12:41:59 PM GMT+08:00 | [Refresh][IFU Kits]CT-NM.SRS.RegStand.733_ Cannot find subsection “Instruction for use” in the section "Pediatric & Small patients" that provides radiation dose reduction strategies for pediatric patients in IFU file | WIP16 | WIP19 |  | Subsystem | Blocker | CT-NM.SRS.RegStand.733 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Brilliance 64 |  | WIP19 | July 31, 2019 4:57:23 PM GMT+08:00
CTUDT00624533 | Refresh | Closed | zhu_je | June 19, 2019 12:41:59 PM GMT+08:00 | [Refresh][IFU Kits]CT-NM.SRS.RegStand.733_ Cannot find subsection “Instruction for use” in the section "Pediatric & Small patients" that provides radiation dose reduction strategies for pediatric patients in IFU file | WIP16 | WIP19 |  | Subsystem | Blocker | CT-NM.SRS.RegStand.733 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Brilliance Big Bore |  | WIP19 | July 31, 2019 4:57:23 PM GMT+08:00
CTUDT00624533 | Refresh | Closed | zhu_je | June 19, 2019 12:41:59 PM GMT+08:00 | [Refresh][IFU Kits]CT-NM.SRS.RegStand.733_ Cannot find subsection “Instruction for use” in the section "Pediatric & Small patients" that provides radiation dose reduction strategies for pediatric patients in IFU file | WIP16 | WIP19 |  | Subsystem | Blocker | CT-NM.SRS.RegStand.733 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Ingenuity 64 |  | WIP19 | July 31, 2019 4:57:23 PM GMT+08:00
CTUDT00624533 | Refresh | Closed | zhu_je | June 19, 2019 12:41:59 PM GMT+08:00 | [Refresh][IFU Kits]CT-NM.SRS.RegStand.733_ Cannot find subsection “Instruction for use” in the section "Pediatric & Small patients" that provides radiation dose reduction strategies for pediatric patients in IFU file | WIP16 | WIP19 |  | Subsystem | Blocker | CT-NM.SRS.RegStand.733 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | IQon1 |  | WIP19 | July 31, 2019 4:57:23 PM GMT+08:00
CTUDT00624531 | Refresh | Closed | zhu_je | June 19, 2019 12:32:16 PM GMT+08:00 | [Refresh][IFU Kits]CT-NM.SRS.RegStand.732_ Cannot find subsection “Instruction for use” in the section "Pediatric & Small patients" in IFU doc | WIP16 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.RegStand.732 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Ovation |  | WIP19 | July 31, 2019 4:56:35 PM GMT+08:00
CTUDT00624531 | Refresh | Closed | zhu_je | June 19, 2019 12:32:16 PM GMT+08:00 | [Refresh][IFU Kits]CT-NM.SRS.RegStand.732_ Cannot find subsection “Instruction for use” in the section "Pediatric & Small patients" in IFU doc | WIP16 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.RegStand.732 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Brilliance 64 |  | WIP19 | July 31, 2019 4:56:35 PM GMT+08:00
CTUDT00624531 | Refresh | Closed | zhu_je | June 19, 2019 12:32:16 PM GMT+08:00 | [Refresh][IFU Kits]CT-NM.SRS.RegStand.732_ Cannot find subsection “Instruction for use” in the section "Pediatric & Small patients" in IFU doc | WIP16 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.RegStand.732 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Brilliance Big Bore |  | WIP19 | July 31, 2019 4:56:35 PM GMT+08:00
CTUDT00624531 | Refresh | Closed | zhu_je | June 19, 2019 12:32:16 PM GMT+08:00 | [Refresh][IFU Kits]CT-NM.SRS.RegStand.732_ Cannot find subsection “Instruction for use” in the section "Pediatric & Small patients" in IFU doc | WIP16 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.RegStand.732 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Ingenuity 64 |  | WIP19 | July 31, 2019 4:56:35 PM GMT+08:00
CTUDT00624531 | Refresh | Closed | zhu_je | June 19, 2019 12:32:16 PM GMT+08:00 | [Refresh][IFU Kits]CT-NM.SRS.RegStand.732_ Cannot find subsection “Instruction for use” in the section "Pediatric & Small patients" in IFU doc | WIP16 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.RegStand.732 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | IQon1 |  | WIP19 | July 31, 2019 4:56:35 PM GMT+08:00
CTUDT00624529 | Refresh | Closed | zhu_je | June 19, 2019 12:27:37 PM GMT+08:00 | [Refresh][IFU Kits]CT-NM.SRS.RegStand.731_Cannot find subsection entitled "Device Description" in IFU file | WIP16 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.RegStand.731 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Ovation |  | WIP19 | July 31, 2019 5:01:16 PM GMT+08:00
CTUDT00624529 | Refresh | Closed | zhu_je | June 19, 2019 12:27:37 PM GMT+08:00 | [Refresh][IFU Kits]CT-NM.SRS.RegStand.731_Cannot find subsection entitled "Device Description" in IFU file | WIP16 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.RegStand.731 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Brilliance 64 |  | WIP19 | July 31, 2019 5:01:16 PM GMT+08:00
CTUDT00624529 | Refresh | Closed | zhu_je | June 19, 2019 12:27:37 PM GMT+08:00 | [Refresh][IFU Kits]CT-NM.SRS.RegStand.731_Cannot find subsection entitled "Device Description" in IFU file | WIP16 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.RegStand.731 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Brilliance Big Bore |  | WIP19 | July 31, 2019 5:01:16 PM GMT+08:00
CTUDT00624529 | Refresh | Closed | zhu_je | June 19, 2019 12:27:37 PM GMT+08:00 | [Refresh][IFU Kits]CT-NM.SRS.RegStand.731_Cannot find subsection entitled "Device Description" in IFU file | WIP16 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.RegStand.731 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Ingenuity 64 |  | WIP19 | July 31, 2019 5:01:16 PM GMT+08:00
CTUDT00624529 | Refresh | Closed | zhu_je | June 19, 2019 12:27:37 PM GMT+08:00 | [Refresh][IFU Kits]CT-NM.SRS.RegStand.731_Cannot find subsection entitled "Device Description" in IFU file | WIP16 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.RegStand.731 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | IQon1 |  | WIP19 | July 31, 2019 5:01:16 PM GMT+08:00
CTUDT00624527 | Refresh | Closed | zhu_je | June 18, 2019 9:00:39 PM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.RegStand.726_the content is existing at section “Pediatric and Small Patients”, but it is not in subsection "Introduction” | WIP16 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.RegStand.725, CT-NM.SRS.RegStand.726 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Ovation |  | WIP19 | July 31, 2019 4:55:36 PM GMT+08:00
CTUDT00624527 | Refresh | Closed | zhu_je | June 18, 2019 9:00:39 PM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.RegStand.726_the content is existing at section “Pediatric and Small Patients”, but it is not in subsection "Introduction” | WIP16 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.RegStand.725, CT-NM.SRS.RegStand.726 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Brilliance 64 |  | WIP19 | July 31, 2019 4:55:36 PM GMT+08:00
CTUDT00624527 | Refresh | Closed | zhu_je | June 18, 2019 9:00:39 PM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.RegStand.726_the content is existing at section “Pediatric and Small Patients”, but it is not in subsection "Introduction” | WIP16 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.RegStand.725, CT-NM.SRS.RegStand.726 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Brilliance Big Bore |  | WIP19 | July 31, 2019 4:55:36 PM GMT+08:00
CTUDT00624527 | Refresh | Closed | zhu_je | June 18, 2019 9:00:39 PM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.RegStand.726_the content is existing at section “Pediatric and Small Patients”, but it is not in subsection "Introduction” | WIP16 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.RegStand.725, CT-NM.SRS.RegStand.726 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Ingenuity 64 |  | WIP19 | July 31, 2019 4:55:36 PM GMT+08:00
CTUDT00624527 | Refresh | Closed | zhu_je | June 18, 2019 9:00:39 PM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.RegStand.726_the content is existing at section “Pediatric and Small Patients”, but it is not in subsection "Introduction” | WIP16 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.RegStand.725, CT-NM.SRS.RegStand.726 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | IQon1 |  | WIP19 | July 31, 2019 4:55:36 PM GMT+08:00
CTUDT00624525 | Refresh | Closed | zhu_je | June 18, 2019 8:51:22 PM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.RegStand.724_ There is no 'brief description of exam purpose" in Exam cards file | WIP16 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.RegStand.724 | Works As Specified / Not A Defect | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Ovation |  | WIP19 | August 1, 2019 4:27:20 PM GMT+08:00
CTUDT00624525 | Refresh | Closed | zhu_je | June 18, 2019 8:51:22 PM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.RegStand.724_ There is no 'brief description of exam purpose" in Exam cards file | WIP16 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.RegStand.724 | Works As Specified / Not A Defect | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Brilliance 64 |  | WIP19 | August 1, 2019 4:27:20 PM GMT+08:00
CTUDT00624525 | Refresh | Closed | zhu_je | June 18, 2019 8:51:22 PM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.RegStand.724_ There is no 'brief description of exam purpose" in Exam cards file | WIP16 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.RegStand.724 | Works As Specified / Not A Defect | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Brilliance Big Bore |  | WIP19 | August 1, 2019 4:27:20 PM GMT+08:00
CTUDT00624525 | Refresh | Closed | zhu_je | June 18, 2019 8:51:22 PM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.RegStand.724_ There is no 'brief description of exam purpose" in Exam cards file | WIP16 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.RegStand.724 | Works As Specified / Not A Defect | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Ingenuity 64 |  | WIP19 | August 1, 2019 4:27:20 PM GMT+08:00
CTUDT00624525 | Refresh | Closed | zhu_je | June 18, 2019 8:51:22 PM GMT+08:00 | [Refresh][IFU Kits] CT-NM.SRS.RegStand.724_ There is no 'brief description of exam purpose" in Exam cards file | WIP16 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.RegStand.724 | Works As Specified / Not A Defect | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | IQon1 |  | WIP19 | August 1, 2019 4:27:20 PM GMT+08:00
CTUDT00624523 | Refresh | Closed | sshahar | June 18, 2019 7:50:58 PM GMT+08:00 | Refresh\IFU\guidance on creating a local data disaster recovery plan that includes following information | WIP16 | WIP19 |  | Subsystem | IFU | CT-NM.SRS.SecPri.226 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Ovation |  | WIP19 | July 31, 2019 4:53:01 PM GMT+08:00
CTUDT00624523 | Refresh | Closed | sshahar | June 18, 2019 7:50:58 PM GMT+08:00 | Refresh\IFU\guidance on creating a local data disaster recovery plan that includes following information | WIP16 | WIP19 |  | Subsystem | IFU | CT-NM.SRS.SecPri.226 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Brilliance 64 |  | WIP19 | July 31, 2019 4:53:01 PM GMT+08:00
CTUDT00624523 | Refresh | Closed | sshahar | June 18, 2019 7:50:58 PM GMT+08:00 | Refresh\IFU\guidance on creating a local data disaster recovery plan that includes following information | WIP16 | WIP19 |  | Subsystem | IFU | CT-NM.SRS.SecPri.226 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Brilliance Big Bore |  | WIP19 | July 31, 2019 4:53:01 PM GMT+08:00
CTUDT00624523 | Refresh | Closed | sshahar | June 18, 2019 7:50:58 PM GMT+08:00 | Refresh\IFU\guidance on creating a local data disaster recovery plan that includes following information | WIP16 | WIP19 |  | Subsystem | IFU | CT-NM.SRS.SecPri.226 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Ingenuity 64 |  | WIP19 | July 31, 2019 4:53:01 PM GMT+08:00
CTUDT00624523 | Refresh | Closed | sshahar | June 18, 2019 7:50:58 PM GMT+08:00 | Refresh\IFU\guidance on creating a local data disaster recovery plan that includes following information | WIP16 | WIP19 |  | Subsystem | IFU | CT-NM.SRS.SecPri.226 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | IQon1 |  | WIP19 | July 31, 2019 4:53:01 PM GMT+08:00
CTUDT00624521 | Refresh | Closed | sshahar | June 18, 2019 7:33:10 PM GMT+08:00 | IFU/Missing information for operator and local administrator | WIP16 | WIP19 |  | Subsystem | IFU | CT-NM.SRS.SecPri.315 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Ovation |  | WIP19 | July 31, 2019 4:52:13 PM GMT+08:00
CTUDT00624521 | Refresh | Closed | sshahar | June 18, 2019 7:33:10 PM GMT+08:00 | IFU/Missing information for operator and local administrator | WIP16 | WIP19 |  | Subsystem | IFU | CT-NM.SRS.SecPri.315 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Brilliance 64 |  | WIP19 | July 31, 2019 4:52:13 PM GMT+08:00
CTUDT00624521 | Refresh | Closed | sshahar | June 18, 2019 7:33:10 PM GMT+08:00 | IFU/Missing information for operator and local administrator | WIP16 | WIP19 |  | Subsystem | IFU | CT-NM.SRS.SecPri.315 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Brilliance Big Bore |  | WIP19 | July 31, 2019 4:52:13 PM GMT+08:00
CTUDT00624521 | Refresh | Closed | sshahar | June 18, 2019 7:33:10 PM GMT+08:00 | IFU/Missing information for operator and local administrator | WIP16 | WIP19 |  | Subsystem | IFU | CT-NM.SRS.SecPri.315 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Ingenuity 64 |  | WIP19 | July 31, 2019 4:52:13 PM GMT+08:00
CTUDT00624521 | Refresh | Closed | sshahar | June 18, 2019 7:33:10 PM GMT+08:00 | IFU/Missing information for operator and local administrator | WIP16 | WIP19 |  | Subsystem | IFU | CT-NM.SRS.SecPri.315 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | IQon1 |  | WIP19 | July 31, 2019 4:52:13 PM GMT+08:00
CTUDT00624519 | Refresh | Closed | zhu_je | June 18, 2019 5:06:58 PM GMT+08:00 | [Refresh][IFU] CT-NM.SRS.RegStand.725_There is no subsection entitled "Introduction" in the section "Pediatric & Small Patients" | WIP16 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.RegStand.725 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Ovation |  | WIP19 | July 31, 2019 4:53:46 PM GMT+08:00
CTUDT00624519 | Refresh | Closed | zhu_je | June 18, 2019 5:06:58 PM GMT+08:00 | [Refresh][IFU] CT-NM.SRS.RegStand.725_There is no subsection entitled "Introduction" in the section "Pediatric & Small Patients" | WIP16 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.RegStand.725 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Brilliance 64 |  | WIP19 | July 31, 2019 4:53:46 PM GMT+08:00
CTUDT00624519 | Refresh | Closed | zhu_je | June 18, 2019 5:06:58 PM GMT+08:00 | [Refresh][IFU] CT-NM.SRS.RegStand.725_There is no subsection entitled "Introduction" in the section "Pediatric & Small Patients" | WIP16 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.RegStand.725 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Brilliance Big Bore |  | WIP19 | July 31, 2019 4:53:46 PM GMT+08:00
CTUDT00624519 | Refresh | Closed | zhu_je | June 18, 2019 5:06:58 PM GMT+08:00 | [Refresh][IFU] CT-NM.SRS.RegStand.725_There is no subsection entitled "Introduction" in the section "Pediatric & Small Patients" | WIP16 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.RegStand.725 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | Ingenuity 64 |  | WIP19 | July 31, 2019 4:53:46 PM GMT+08:00
CTUDT00624519 | Refresh | Closed | zhu_je | June 18, 2019 5:06:58 PM GMT+08:00 | [Refresh][IFU] CT-NM.SRS.RegStand.725_There is no subsection entitled "Introduction" in the section "Pediatric & Small Patients" | WIP16 | WIP19 |  | Subsystem | IFU;Blocker | CT-NM.SRS.RegStand.725 | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | NO | Closed | PostVerification | IQon1 |  | WIP19 | July 31, 2019 4:53:46 PM GMT+08:00
CTUDT00624518 | Refresh | Closed | guy_sheaffer | June 18, 2019 4:36:52 PM GMT+08:00 | Cardiac/Gated reconstruction not happened properly | 4.1.7 | WIP21 | 1. User will have to wait till the wave appears in the Offline ECG Viewer and then press Start Final Recon 2. User can perform offline reconstructions from Edit Results or load the result in View2 | System | TAVI | N/A | Fixed | bhawna_g | 0 | 3 | Console Software | Console4 | YES | Closed | WIP21SubCom | Ovation | 8844848 8837662 8844846 | WIP21 | September 24, 2019 9:21:20 PM GMT+08:00
CTUDT00624518 | Refresh | Closed | guy_sheaffer | June 18, 2019 4:36:52 PM GMT+08:00 | Cardiac/Gated reconstruction not happened properly | 4.1.7 | WIP21 | 1. User will have to wait till the wave appears in the Offline ECG Viewer and then press Start Final Recon 2. User can perform offline reconstructions from Edit Results or load the result in View2 | System | TAVI | N/A | Fixed | bhawna_g | 0 | 3 | Console Software | Console4 | YES | Closed | WIP21SubCom | Brilliance 64 | 8844848 8837662 8844846 | WIP21 | September 24, 2019 9:21:20 PM GMT+08:00
CTUDT00624518 | Refresh | Closed | guy_sheaffer | June 18, 2019 4:36:52 PM GMT+08:00 | Cardiac/Gated reconstruction not happened properly | 4.1.7 | WIP21 | 1. User will have to wait till the wave appears in the Offline ECG Viewer and then press Start Final Recon 2. User can perform offline reconstructions from Edit Results or load the result in View2 | System | TAVI | N/A | Fixed | bhawna_g | 0 | 3 | Console Software | Console4 | YES | Closed | WIP21SubCom | Ingenuity 64 | 8844848 8837662 8844846 | WIP21 | September 24, 2019 9:21:20 PM GMT+08:00
CTUDT00624518 | Refresh | Closed | guy_sheaffer | June 18, 2019 4:36:52 PM GMT+08:00 | Cardiac/Gated reconstruction not happened properly | 4.1.7 | WIP21 | 1. User will have to wait till the wave appears in the Offline ECG Viewer and then press Start Final Recon 2. User can perform offline reconstructions from Edit Results or load the result in View2 | System | TAVI | N/A | Fixed | bhawna_g | 0 | 3 | Console Software | Console4 | YES | Closed | WIP21SubCom | IQon1 | 8844848 8837662 8844846 | WIP21 | September 24, 2019 9:21:20 PM GMT+08:00
CTUDT00624516 | Refresh | Closed | swamypreethi | June 18, 2019 4:10:35 PM GMT+08:00 | IOCC_Refresh: Issue in Assisted acquisition protocol setting order change | WIP15 |  | N/A | Subsystem |  | N/A | No Plan to Fix | naveen_k | 0 | 4 |  | Console4 | NO | Closed | SubNC | Ovation
CTUDT00624516 | Refresh | Closed | swamypreethi | June 18, 2019 4:10:35 PM GMT+08:00 | IOCC_Refresh: Issue in Assisted acquisition protocol setting order change | WIP15 |  | N/A | Subsystem |  | N/A | No Plan to Fix | naveen_k | 0 | 4 |  | Console4 | NO | Closed | SubNC | Brilliance 64
CTUDT00624516 | Refresh | Closed | swamypreethi | June 18, 2019 4:10:35 PM GMT+08:00 | IOCC_Refresh: Issue in Assisted acquisition protocol setting order change | WIP15 |  | N/A | Subsystem |  | N/A | No Plan to Fix | naveen_k | 0 | 4 |  | Console4 | NO | Closed | SubNC | Ingenuity 64
CTUDT00624516 | Refresh | Closed | swamypreethi | June 18, 2019 4:10:35 PM GMT+08:00 | IOCC_Refresh: Issue in Assisted acquisition protocol setting order change | WIP15 |  | N/A | Subsystem |  | N/A | No Plan to Fix | naveen_k | 0 | 4 |  | Console4 | NO | Closed | SubNC | Ingenuity BigBore
CTUDT00624516 | Refresh | Closed | swamypreethi | June 18, 2019 4:10:35 PM GMT+08:00 | IOCC_Refresh: Issue in Assisted acquisition protocol setting order change | WIP15 |  | N/A | Subsystem |  | N/A | No Plan to Fix | naveen_k | 0 | 4 |  | Console4 | NO | Closed | SubNC | IQon1
CTUDT00624515 | Refresh | Closed | datta_i | June 18, 2019 3:02:35 PM GMT+08:00 | Max mAs value is shown red in color even after increasing the DRI (to avoid Clipping) | 4.8 | BB_WIP9 | the clipping indication is also available as * sign on the plan on surview | Subsystem |  | NA | Fixed | datta_i | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP8_SubCom | Brilliance Big Bore |  | BB_WIP9 | July 19, 2019 6:45:35 PM GMT+08:00
CTUDT00624511 | Refresh | Closed | cai_t | June 18, 2019 2:31:33 PM GMT+08:00 | Refresh: Quality Assurance: Reliability Test failed one step. | BB_WIP6 | DHF330990 Rev01 |  | Subsystem | Calibrations | Reliability test | Fixed | cai_t | 0 | 4 | Test Design | Support-System | NO | Closed | BB_WIP8_SubCom | Brilliance Big Bore |  | DHF330990 Rev01 | August 4, 2019 9:24:33 AM GMT+08:00
CTUDT00624510 | Refresh | Closed | Valeria_Kashtanov | June 18, 2019 2:20:55 PM GMT+08:00 | Refresh: Calibrations: Hcor Basic Infant: Rotation Time of Head and Body scans is not according to spec | WIP17 | N/A | N/A | Subsystem | Calibrations | N/A | Fixed | braunstein_d | 0 | 2 | Design Documentation | Support-Physics | NO | Closed | WIP18SubCom | Ingenuity 64 |  | N/A | July 4, 2019 7:52:40 PM GMT+08:00
CTUDT00624509 | Refresh | Closed | bhawna_g | June 18, 2019 1:45:27 PM GMT+08:00 | Pre- Val Br BB Refresh: SSDE units does not appear for basic interventional scan | 4.1.10 | WIP18 |  | Subsystem |  | NA | Fixed | bhawna_g | 0 | 4 | Console Software | Console4 | NO | Closed | WIP19SegReq | Ovation |  | WIP18 | June 26, 2019 4:47:27 PM GMT+08:00
CTUDT00624509 | Refresh | Closed | bhawna_g | June 18, 2019 1:45:27 PM GMT+08:00 | Pre- Val Br BB Refresh: SSDE units does not appear for basic interventional scan | 4.1.10 | WIP18 |  | Subsystem |  | NA | Fixed | bhawna_g | 0 | 4 | Console Software | Console4 | NO | Closed | WIP19SegReq | Brilliance 64 |  | WIP18 | June 26, 2019 4:47:27 PM GMT+08:00
CTUDT00624509 | Refresh | Closed | bhawna_g | June 18, 2019 1:45:27 PM GMT+08:00 | Pre- Val Br BB Refresh: SSDE units does not appear for basic interventional scan | 4.1.10 | WIP18 |  | Subsystem |  | NA | Fixed | bhawna_g | 0 | 4 | Console Software | Console4 | NO | Closed | WIP19SegReq | Ingenuity 64 |  | WIP18 | June 26, 2019 4:47:27 PM GMT+08:00
CTUDT00624509 | Refresh | Closed | bhawna_g | June 18, 2019 1:45:27 PM GMT+08:00 | Pre- Val Br BB Refresh: SSDE units does not appear for basic interventional scan | 4.1.10 | WIP18 |  | Subsystem |  | NA | Fixed | bhawna_g | 0 | 4 | Console Software | Console4 | NO | Closed | WIP19SegReq | IQon1 |  | WIP18 | June 26, 2019 4:47:27 PM GMT+08:00
CTUDT00624500 | Refresh | Closed | azagury_d | June 17, 2019 8:57:16 PM GMT+08:00 | Refresh: Backup-Restore: Missing file PHCT_Physics_BackupRestore.xml | WIP17 | N/A - Test Design |  | Subsystem | Test Design | N/A | Fixed | gehtin_e | 0 | 3 | Test Design | DMS | NO | Closed | WIP19SegReq | Ingenuity 64 |  | N/A - Test Design | July 8, 2019 6:08:43 PM GMT+08:00
CTUDT00624498 | Refresh | Closed | 320053769 | June 17, 2019 8:17:40 PM GMT+08:00 | Burning studies to DVD+RW with host Gen 5.3 fails. | WIP17 |  |  | Subsystem | CD-DVD operation | NA | Invalid/Uncontrolled Configuration | s_sa | 0 |  |  | iCAP | NO | Closed |  | Ovation
CTUDT00624498 | Refresh | Closed | 320053769 | June 17, 2019 8:17:40 PM GMT+08:00 | Burning studies to DVD+RW with host Gen 5.3 fails. | WIP17 |  |  | Subsystem | CD-DVD operation | NA | Invalid/Uncontrolled Configuration | s_sa | 0 |  |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00624498 | Refresh | Closed | 320053769 | June 17, 2019 8:17:40 PM GMT+08:00 | Burning studies to DVD+RW with host Gen 5.3 fails. | WIP17 |  |  | Subsystem | CD-DVD operation | NA | Invalid/Uncontrolled Configuration | s_sa | 0 |  |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00624498 | Refresh | Closed | 320053769 | June 17, 2019 8:17:40 PM GMT+08:00 | Burning studies to DVD+RW with host Gen 5.3 fails. | WIP17 |  |  | Subsystem | CD-DVD operation | NA | Invalid/Uncontrolled Configuration | s_sa | 0 |  |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00624498 | Refresh | Closed | 320053769 | June 17, 2019 8:17:40 PM GMT+08:00 | Burning studies to DVD+RW with host Gen 5.3 fails. | WIP17 |  |  | Subsystem | CD-DVD operation | NA | Invalid/Uncontrolled Configuration | s_sa | 0 |  |  | iCAP | NO | Closed |  | IQon1
CTUDT00624480 | Refresh | Closed | finestine_m | June 17, 2019 1:18:22 AM GMT+08:00 | Refresh: Performance Test: Missing Scans In Customization Window | WIP16 | WIP18 |  | Subsystem | Performance Test | n/a | Fixed | Kostya_Baklan | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP18SubCom | Ovation |  | WIP18 | June 27, 2019 3:21:35 PM GMT+08:00
CTUDT00624476 | Refresh | Closed | attali_s | June 16, 2019 7:45:12 PM GMT+08:00 | Refresh: Acceptance/Constancy: Wrong number of scans for some cases | WIP16 | WIP18 | N/A | Subsystem | Acceptance/Constancy | N/A | Fixed | elviral | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP18SegReq | IQon1 |  | WIP18 | June 30, 2019 4:22:07 PM GMT+08:00
CTUDT00624474 | Refresh | Closed | attali_s | June 16, 2019 7:16:24 PM GMT+08:00 | Refresh: Air Calibration: Different values in Logged/Unlogged RefDet FS0 and FS1 between TubeMonitor and Physics Analysis | WIP17 | WIP18 |  | Subsystem | Calibrations | n/a | Fixed | elviral | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP18SubCom | Ovation |  | WIP18 | July 3, 2019 10:40:03 PM GMT+08:00
CTUDT00624474 | Refresh | Closed | attali_s | June 16, 2019 7:16:24 PM GMT+08:00 | Refresh: Air Calibration: Different values in Logged/Unlogged RefDet FS0 and FS1 between TubeMonitor and Physics Analysis | WIP17 | WIP18 |  | Subsystem | Calibrations | n/a | Fixed | elviral | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP18SubCom | Brilliance Big Bore |  | WIP18 | July 3, 2019 10:40:03 PM GMT+08:00
CTUDT00624474 | Refresh | Closed | attali_s | June 16, 2019 7:16:24 PM GMT+08:00 | Refresh: Air Calibration: Different values in Logged/Unlogged RefDet FS0 and FS1 between TubeMonitor and Physics Analysis | WIP17 | WIP18 |  | Subsystem | Calibrations | n/a | Fixed | elviral | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP18SubCom | Ingenuity 64 |  | WIP18 | July 3, 2019 10:40:03 PM GMT+08:00
CTUDT00624474 | Refresh | Closed | attali_s | June 16, 2019 7:16:24 PM GMT+08:00 | Refresh: Air Calibration: Different values in Logged/Unlogged RefDet FS0 and FS1 between TubeMonitor and Physics Analysis | WIP17 | WIP18 |  | Subsystem | Calibrations | n/a | Fixed | elviral | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP18SubCom | IQon1 |  | WIP18 | July 3, 2019 10:40:03 PM GMT+08:00
CTUDT00624472 | Refresh | Closed | alex_nihaev | June 16, 2019 6:49:00 PM GMT+08:00 | Refresh: Calibrations:Phantom Centering Test:Eye Phantom,H Holder, Tilt of the Phantom is not recognized | WIP17 | WIP18 |  | Subsystem | Calibrations | NaN | Fixed | daibes_s | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP18SubCom | Brilliance Big Bore |  | WIP18 | July 7, 2019 10:11:45 PM GMT+08:00
CTUDT00624472 | Refresh | Closed | alex_nihaev | June 16, 2019 6:49:00 PM GMT+08:00 | Refresh: Calibrations:Phantom Centering Test:Eye Phantom,H Holder, Tilt of the Phantom is not recognized | WIP17 | WIP18 |  | Subsystem | Calibrations | NaN | Fixed | daibes_s | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP18SubCom | Ingenuity 64 |  | WIP18 | July 7, 2019 10:11:45 PM GMT+08:00
CTUDT00624466 | Refresh | Closed | waggoner_j | June 15, 2019 12:19:57 AM GMT+08:00 | Update needed to IFU and Training checklists for Refresh product release | BB_WIP6 | WIP19 | Nil | Subsystem | IFU | N/A | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | YES | Closed | PostVerification | Ovation |  | WIP19 | July 31, 2019 4:50:43 PM GMT+08:00
CTUDT00624466 | Refresh | Closed | waggoner_j | June 15, 2019 12:19:57 AM GMT+08:00 | Update needed to IFU and Training checklists for Refresh product release | BB_WIP6 | WIP19 | Nil | Subsystem | IFU | N/A | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | YES | Closed | PostVerification | Brilliance 64 |  | WIP19 | July 31, 2019 4:50:43 PM GMT+08:00
CTUDT00624466 | Refresh | Closed | waggoner_j | June 15, 2019 12:19:57 AM GMT+08:00 | Update needed to IFU and Training checklists for Refresh product release | BB_WIP6 | WIP19 | Nil | Subsystem | IFU | N/A | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | YES | Closed | PostVerification | Brilliance Big Bore |  | WIP19 | July 31, 2019 4:50:43 PM GMT+08:00
CTUDT00624466 | Refresh | Closed | waggoner_j | June 15, 2019 12:19:57 AM GMT+08:00 | Update needed to IFU and Training checklists for Refresh product release | BB_WIP6 | WIP19 | Nil | Subsystem | IFU | N/A | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | YES | Closed | PostVerification | Ingenuity 64 |  | WIP19 | July 31, 2019 4:50:43 PM GMT+08:00
CTUDT00624466 | Refresh | Closed | waggoner_j | June 15, 2019 12:19:57 AM GMT+08:00 | Update needed to IFU and Training checklists for Refresh product release | BB_WIP6 | WIP19 | Nil | Subsystem | IFU | N/A | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | YES | Closed | PostVerification | IQon1 |  | WIP19 | July 31, 2019 4:50:43 PM GMT+08:00
CTUDT00624464 | Refresh | Closed | Prakash_Chand_Yadav | June 14, 2019 7:15:21 PM GMT+08:00 | Status message for spectral result are not getting displayed correctly | WIP17 | WIP18 |  | Subsystem | Plan | None | Fixed | Prakash_Chand_Yadav | 0 | 4 | Console Software | Console4 | NO | Closed | WIP18SubCom | IQon1 |  | WIP18 | June 19, 2019 6:48:10 PM GMT+08:00
CTUDT00624462 | Refresh | Closed | Prakash_Chand_Yadav | June 14, 2019 7:10:14 PM GMT+08:00 | Combobox in front of ‘Use Spectral Recon of:’ does not clears its value | WIP17 | WIP18 | Navigating to the 1st spectral result clears the combobox value in the 2nd spectral result. | Subsystem | Plan | Console-SRS.11598 | Fixed | mY_h | 0 | 3 | Console Software | Console4 | NO | Closed | WIP18SubCom | IQon1 |  | WIP18 | June 26, 2019 2:57:02 PM GMT+08:00
CTUDT00624460 | Refresh | Closed | hatti_c | June 14, 2019 5:39:15 PM GMT+08:00 | SCOE Pentest: Unsigned executable can be executed using run in compatible mode using explorer | WIP14 |  | NA | Subsystem | Security | NA | No Plan to Fix | shaham_t | 0 | 4 |  | Console4 | YES | Closed | SubNC | Ovation
CTUDT00624460 | Refresh | Closed | hatti_c | June 14, 2019 5:39:15 PM GMT+08:00 | SCOE Pentest: Unsigned executable can be executed using run in compatible mode using explorer | WIP14 |  | NA | Subsystem | Security | NA | No Plan to Fix | shaham_t | 0 | 4 |  | Console4 | YES | Closed | SubNC | Brilliance 64
CTUDT00624460 | Refresh | Closed | hatti_c | June 14, 2019 5:39:15 PM GMT+08:00 | SCOE Pentest: Unsigned executable can be executed using run in compatible mode using explorer | WIP14 |  | NA | Subsystem | Security | NA | No Plan to Fix | shaham_t | 0 | 4 |  | Console4 | YES | Closed | SubNC | Brilliance Big Bore
CTUDT00624460 | Refresh | Closed | hatti_c | June 14, 2019 5:39:15 PM GMT+08:00 | SCOE Pentest: Unsigned executable can be executed using run in compatible mode using explorer | WIP14 |  | NA | Subsystem | Security | NA | No Plan to Fix | shaham_t | 0 | 4 |  | Console4 | YES | Closed | SubNC | Ingenuity 64
CTUDT00624460 | Refresh | Closed | hatti_c | June 14, 2019 5:39:15 PM GMT+08:00 | SCOE Pentest: Unsigned executable can be executed using run in compatible mode using explorer | WIP14 |  | NA | Subsystem | Security | NA | No Plan to Fix | shaham_t | 0 | 4 |  | Console4 | YES | Closed | SubNC | IQon1
CTUDT00624455 | Refresh | Closed | zhu_je | June 14, 2019 2:00:10 PM GMT+08:00 | [BB]The system doesn't allow selecting a static injection protocol from the list received from the injector if Automatic P3T is unchecked | BB_WIP6 |  |  | Subsystem | Plan | Console-SRS.916 | Does Not Apply | zhu_je | 0 |  |  | Support-System | NO | Closed | WAS | Brilliance Big Bore
CTUDT00624452 | Refresh | Closed | g_di | June 14, 2019 1:23:33 PM GMT+08:00 | Few parameters are missing and there is mismatch in “Show all” and “Print and text tab”. | WIP17 | WIP18 |  | Subsystem | ECM | NA | Fixed | kumar-san | 0 | 4 | Console Software | Console4 | NO | Closed | WIP18SubCom | IQon1 |  | WIP18 | June 21, 2019 8:12:26 PM GMT+08:00
CTUDT00624450 | Refresh | Closed | Valeria_Kashtanov | June 13, 2019 9:56:12 PM GMT+08:00 | Refresh: HCOR Calibration: Wrong Customization Window for HCOR Slice | WIP17 | N/A | N/A | Subsystem | Calibrations | N/A | Fixed | braunstein_d | 0 | 2 | Design Documentation | Support-Physics | NO | Closed | WIP18SubCom | Ovation |  | N/A | July 4, 2019 7:58:42 PM GMT+08:00
CTUDT00624450 | Refresh | Closed | Valeria_Kashtanov | June 13, 2019 9:56:12 PM GMT+08:00 | Refresh: HCOR Calibration: Wrong Customization Window for HCOR Slice | WIP17 | N/A | N/A | Subsystem | Calibrations | N/A | Fixed | braunstein_d | 0 | 2 | Design Documentation | Support-Physics | NO | Closed | WIP18SubCom | Brilliance Big Bore |  | N/A | July 4, 2019 7:58:42 PM GMT+08:00
CTUDT00624450 | Refresh | Closed | Valeria_Kashtanov | June 13, 2019 9:56:12 PM GMT+08:00 | Refresh: HCOR Calibration: Wrong Customization Window for HCOR Slice | WIP17 | N/A | N/A | Subsystem | Calibrations | N/A | Fixed | braunstein_d | 0 | 2 | Design Documentation | Support-Physics | NO | Closed | WIP18SubCom | Ingenuity 64 |  | N/A | July 4, 2019 7:58:42 PM GMT+08:00
CTUDT00624450 | Refresh | Closed | Valeria_Kashtanov | June 13, 2019 9:56:12 PM GMT+08:00 | Refresh: HCOR Calibration: Wrong Customization Window for HCOR Slice | WIP17 | N/A | N/A | Subsystem | Calibrations | N/A | Fixed | braunstein_d | 0 | 2 | Design Documentation | Support-Physics | NO | Closed | WIP18SubCom | IQon1 |  | N/A | July 4, 2019 7:58:42 PM GMT+08:00
CTUDT00624448 | Refresh | Closed | Valeria_Kashtanov | June 13, 2019 9:33:42 PM GMT+08:00 | Refresh: Calibrations: Hcor Basic: Window values in body and head cases are not according to spec with Harmonized Phantom | WIP17 | WIP18 |  | Subsystem | Calibrations | n/a | Fixed | Kostya_Baklan | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP18SubCom | Ingenuity 64 |  | WIP18 | June 20, 2019 3:01:15 PM GMT+08:00
CTUDT00624434 | Refresh | Closed | cai_t | June 13, 2019 4:25:20 PM GMT+08:00 | Refresh: Localization:Spanish:Tooltip "Add/Replace Exam card" displayed wrong in Spanish. | BB_WIP6 |  | NA | Subsystem | Localization | NA | Works As Specified / Not A Defect | cai_t | 0 | 2 | Test Design | Support-System | NO | Closed | WAS | Brilliance Big Bore |  |  | July 12, 2019 3:35:14 PM GMT+08:00
CTUDT00624427 | Refresh | Closed | zhu_je | June 13, 2019 12:56:46 PM GMT+08:00 | [BB] : Load patient to plan screen was very slowly after changing the host and CIRS’s ip from “192” to “10” | BB_WIP6 | BB_WIP9 |  | Subsystem | Performance;Blocker | NA | Fixed | madamsetti_k | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP9_SubCom | Brilliance Big Bore |  | BB_WIP9 | August 9, 2019 1:27:43 PM GMT+08:00
CTUDT00624423 | Refresh | Closed | cohen_y | June 13, 2019 12:01:25 AM GMT+08:00 | Refresh: Quality Assurance: Reasons Matrix: Harmonized: Hcor Basic and Slice are never selected in relevant reasons | WIP17 | WIP18 |  | Subsystem | Calibrations | na | Fixed | frider_o | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP18SubCom | Ingenuity 64 |  | WIP18 | June 26, 2019 9:25:33 PM GMT+08:00
CTUDT00624421 | Refresh | Closed | zhang_t | June 12, 2019 11:38:30 PM GMT+08:00 | Refresh: While 1 click installation, "Security settings Setup has stopped working” popup | WIP17 | WIP19 |  | Subsystem | 1-click, Blocker19 | Host-Security-SwRS-060 | Fixed | venugopal_s | 0 | 3 | Service Tools_Software | DI | NO | Closed | WIP19SegReq | Brilliance 64 |  | WIP19 | August 8, 2019 4:54:44 PM GMT+08:00
CTUDT00624419 | Refresh | Closed | cohen_y | June 12, 2019 11:01:53 PM GMT+08:00 | Refresh: Quality Assurance: Reasons Matrix: Selected procedures are not as defined in Spec | WIP17 | WIP20 | NA | Subsystem | ;Blocker20; Verified1 | NA | Fixed | elviral | 0 | 3 | Design Documentation | DMS | NO | Closed | WIP18SegReq | Brilliance Big Bore |  | WIP20 | September 4, 2019 12:12:41 AM GMT+08:00
CTUDT00624419 | Refresh | Closed | cohen_y | June 12, 2019 11:01:53 PM GMT+08:00 | Refresh: Quality Assurance: Reasons Matrix: Selected procedures are not as defined in Spec | WIP17 | WIP20 | NA | Subsystem | ;Blocker20; Verified1 | NA | Fixed | elviral | 0 | 3 | Design Documentation | DMS | NO | Closed | WIP18SegReq | Ingenuity 64 |  | WIP20 | September 4, 2019 12:12:41 AM GMT+08:00
CTUDT00624419 | Refresh | Closed | cohen_y | June 12, 2019 11:01:53 PM GMT+08:00 | Refresh: Quality Assurance: Reasons Matrix: Selected procedures are not as defined in Spec | WIP17 | WIP20 | NA | Subsystem | ;Blocker20; Verified1 | NA | Fixed | elviral | 0 | 3 | Design Documentation | DMS | NO | Closed | WIP18SegReq | IQon1 |  | WIP20 | September 4, 2019 12:12:41 AM GMT+08:00
CTUDT00624417 | Refresh | Closed | shterner_a | June 12, 2019 9:28:40 PM GMT+08:00 | Pre- Val Br BB Refresh: SSDE units does not appear for basic interventional scan | BB_WIP6 | BB_WIP8 |  | Subsystem | SSDE | NA | Fixed | bhawna_g | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP8_SubCom | Brilliance Big Bore |  | BB_WIP8 | June 26, 2019 3:12:48 PM GMT+08:00
CTUDT00624414 | Refresh | Closed | naveen_k | June 12, 2019 6:12:23 PM GMT+08:00 | The CCT Single and CCT Continuous display "Scanning..." in the scan indication area after pressing the pedal. | WIP17 | WIP18 |  | Subsystem |  | NA | Fixed | veeraswamy_k | 0 | 4 | Console Software | Console4 | NO | Closed | WIP18SubCom | Ovation |  | WIP18 | June 17, 2019 8:56:48 PM GMT+08:00
CTUDT00624414 | Refresh | Closed | naveen_k | June 12, 2019 6:12:23 PM GMT+08:00 | The CCT Single and CCT Continuous display "Scanning..." in the scan indication area after pressing the pedal. | WIP17 | WIP18 |  | Subsystem |  | NA | Fixed | veeraswamy_k | 0 | 4 | Console Software | Console4 | NO | Closed | WIP18SubCom | Brilliance 64 |  | WIP18 | June 17, 2019 8:56:48 PM GMT+08:00
CTUDT00624414 | Refresh | Closed | naveen_k | June 12, 2019 6:12:23 PM GMT+08:00 | The CCT Single and CCT Continuous display "Scanning..." in the scan indication area after pressing the pedal. | WIP17 | WIP18 |  | Subsystem |  | NA | Fixed | veeraswamy_k | 0 | 4 | Console Software | Console4 | NO | Closed | WIP18SubCom | Ingenuity 64 |  | WIP18 | June 17, 2019 8:56:48 PM GMT+08:00
CTUDT00624414 | Refresh | Closed | naveen_k | June 12, 2019 6:12:23 PM GMT+08:00 | The CCT Single and CCT Continuous display "Scanning..." in the scan indication area after pressing the pedal. | WIP17 | WIP18 |  | Subsystem |  | NA | Fixed | veeraswamy_k | 0 | 4 | Console Software | Console4 | NO | Closed | WIP18SubCom | IQon1 |  | WIP18 | June 17, 2019 8:56:48 PM GMT+08:00
CTUDT00624412 | Refresh | Closed | finestine_m | June 12, 2019 2:29:22 PM GMT+08:00 | Refresh: Performance Test: Wrong Head Scans Limits | WIP16 | WIP18 |  | Subsystem | Performance Test | N/A | Fixed | Kostya_Baklan | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP19SegReq | IQon1 |  | WIP18 | June 20, 2019 2:59:17 PM GMT+08:00
CTUDT00624409 | Refresh | Closed | finestine_m | June 12, 2019 12:33:27 PM GMT+08:00 | Refresh: Performance Test: Missing "Cardiac Image Quality" Table In The Spec For iCT SP | WIP16 | N/A - Test Design | N/A | Subsystem | Test Design | N/A | Fixed | gehtin_e | 0 | 3 | Test Design | DMS | NO | Closed | WIP18SegReq | Ovation |  | N/A - Test Design | July 9, 2019 3:00:59 PM GMT+08:00
CTUDT00624404 | Refresh | Closed | Valeria_Kashtanov | June 12, 2019 2:32:00 AM GMT+08:00 | Refresh: Calibrations: Acceptance-Constancy: Window values of body cases and slice thickness 3 with Harmonized Phantom are not according to Spec | WIP16 | N/A | N/A | Subsystem | Acceptance/Constancy | N/A | Fixed | braunstein_d | 0 | 3 | Design Documentation | Support-Physics | NO | Closed | WIP18SegReq | IQon1 |  | N/A | July 8, 2019 4:31:57 PM GMT+08:00
CTUDT00624402 | Refresh | Closed | Dana_Kamyshev | June 11, 2019 10:33:47 PM GMT+08:00 | Refresh - Patient screen loading time after 'End Exam' and clock freeze | WIP17 | WIP18 | N/A | Subsystem | Performance | NA | Fixed | kapadia_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP18SubCom | Ovation |  | WIP18 | June 26, 2019 4:02:22 PM GMT+08:00
CTUDT00624402 | Refresh | Closed | Dana_Kamyshev | June 11, 2019 10:33:47 PM GMT+08:00 | Refresh - Patient screen loading time after 'End Exam' and clock freeze | WIP17 | WIP18 | N/A | Subsystem | Performance | NA | Fixed | kapadia_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP18SubCom | Brilliance 64 |  | WIP18 | June 26, 2019 4:02:22 PM GMT+08:00
CTUDT00624402 | Refresh | Closed | Dana_Kamyshev | June 11, 2019 10:33:47 PM GMT+08:00 | Refresh - Patient screen loading time after 'End Exam' and clock freeze | WIP17 | WIP18 | N/A | Subsystem | Performance | NA | Fixed | kapadia_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP18SubCom | Ingenuity 16 |  | WIP18 | June 26, 2019 4:02:22 PM GMT+08:00
CTUDT00624402 | Refresh | Closed | Dana_Kamyshev | June 11, 2019 10:33:47 PM GMT+08:00 | Refresh - Patient screen loading time after 'End Exam' and clock freeze | WIP17 | WIP18 | N/A | Subsystem | Performance | NA | Fixed | kapadia_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP18SubCom | IQon1 |  | WIP18 | June 26, 2019 4:02:22 PM GMT+08:00
CTUDT00624400 | Refresh | Closed | sshahar | June 11, 2019 10:11:18 PM GMT+08:00 | Refresh: Security\actual OS services state different from DHF314402 | WIP16 |  |  | Subsystem | System |  | Duplicate | Prabhakar_Priyank_Muppidi | 0 |  |  | Support-Security | NO | Closed | WIP18SubCom | IQon1
CTUDT00624398 | Refresh | Closed | rojin_antony | June 11, 2019 9:13:43 PM GMT+08:00 | CCT center Location is hidden in dual surview display with default zoom ( Zoom 1) | WIP17 |  | 1. User can use Zoom out option to see and edit the value of Center for both the surviews 2. User can click on the Center annotation which displays the all digits of the value. 3. For a Dual surview, User can change the dual viewport mode to single viewport mode by double clicking on the viewport which displays all the digits of Center value. 4. User can open the Show All of the CCT step to view and change the Center value. | Subsystem | Plan | Console-SRS.10192 | No Plan to Fix | sheik_c | 0 | 3 |  | Console4 | NO | Closed |  | Ovation
CTUDT00624398 | Refresh | Closed | rojin_antony | June 11, 2019 9:13:43 PM GMT+08:00 | CCT center Location is hidden in dual surview display with default zoom ( Zoom 1) | WIP17 |  | 1. User can use Zoom out option to see and edit the value of Center for both the surviews 2. User can click on the Center annotation which displays the all digits of the value. 3. For a Dual surview, User can change the dual viewport mode to single viewport mode by double clicking on the viewport which displays all the digits of Center value. 4. User can open the Show All of the CCT step to view and change the Center value. | Subsystem | Plan | Console-SRS.10192 | No Plan to Fix | sheik_c | 0 | 3 |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00624398 | Refresh | Closed | rojin_antony | June 11, 2019 9:13:43 PM GMT+08:00 | CCT center Location is hidden in dual surview display with default zoom ( Zoom 1) | WIP17 |  | 1. User can use Zoom out option to see and edit the value of Center for both the surviews 2. User can click on the Center annotation which displays the all digits of the value. 3. For a Dual surview, User can change the dual viewport mode to single viewport mode by double clicking on the viewport which displays all the digits of Center value. 4. User can open the Show All of the CCT step to view and change the Center value. | Subsystem | Plan | Console-SRS.10192 | No Plan to Fix | sheik_c | 0 | 3 |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00624398 | Refresh | Closed | rojin_antony | June 11, 2019 9:13:43 PM GMT+08:00 | CCT center Location is hidden in dual surview display with default zoom ( Zoom 1) | WIP17 |  | 1. User can use Zoom out option to see and edit the value of Center for both the surviews 2. User can click on the Center annotation which displays the all digits of the value. 3. For a Dual surview, User can change the dual viewport mode to single viewport mode by double clicking on the viewport which displays all the digits of Center value. 4. User can open the Show All of the CCT step to view and change the Center value. | Subsystem | Plan | Console-SRS.10192 | No Plan to Fix | sheik_c | 0 | 3 |  | Console4 | NO | Closed |  | IQon1
CTUDT00624395 | Refresh | Closed | rojin_antony | June 11, 2019 8:41:15 PM GMT+08:00 | In CCT Single & continues "Volume" display mode is not giving the exact pre selected center location in any of the images | WIP17 |  | User can change the display mode to "1"or "3" to find the center image slice | Subsystem | CCT | Console SRS:10192 | No Plan to Fix | kumar-san | 0 | 3 |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00624394 | Refresh | Closed | alex_nihaev | June 11, 2019 8:25:00 PM GMT+08:00 | Refresh: Calibrations: IQ Check, H-Phantom, Body ROIs positions is not as defined in spec | WIP16 | WIP18 |  | Subsystem | Calibrations | n/a | Fixed | Kostya_Baklan | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP18SubCom | IQon1 |  | WIP18 | June 23, 2019 10:28:54 PM GMT+08:00
CTUDT00624391 | Refresh | Closed | rojin_antony | June 11, 2019 7:32:23 PM GMT+08:00 | Liver DRI value is not updating in Scan parameter Tab | WIP16 | WIP21 | Yes.To see the selected Liver DRI value on the Scan Parameters Toolbox, user has to perform a scan .After the scan the updated Liver DRI value shall be reflected under Parameters Panel region | Subsystem |  | None | Fixed | deshbhratar_m | 0 | 2 | Console Software | Console4 | YES | Closed | WIP18SubCom | Ovation |  | WIP21 | September 17, 2019 8:13:10 PM GMT+08:00
CTUDT00624391 | Refresh | Closed | rojin_antony | June 11, 2019 7:32:23 PM GMT+08:00 | Liver DRI value is not updating in Scan parameter Tab | WIP16 | WIP21 | Yes.To see the selected Liver DRI value on the Scan Parameters Toolbox, user has to perform a scan .After the scan the updated Liver DRI value shall be reflected under Parameters Panel region | Subsystem |  | None | Fixed | deshbhratar_m | 0 | 2 | Console Software | Console4 | YES | Closed | WIP18SubCom | Brilliance 64 |  | WIP21 | September 17, 2019 8:13:10 PM GMT+08:00
CTUDT00624391 | Refresh | Closed | rojin_antony | June 11, 2019 7:32:23 PM GMT+08:00 | Liver DRI value is not updating in Scan parameter Tab | WIP16 | WIP21 | Yes.To see the selected Liver DRI value on the Scan Parameters Toolbox, user has to perform a scan .After the scan the updated Liver DRI value shall be reflected under Parameters Panel region | Subsystem |  | None | Fixed | deshbhratar_m | 0 | 2 | Console Software | Console4 | YES | Closed | WIP18SubCom | Ingenuity 64 |  | WIP21 | September 17, 2019 8:13:10 PM GMT+08:00
CTUDT00624391 | Refresh | Closed | rojin_antony | June 11, 2019 7:32:23 PM GMT+08:00 | Liver DRI value is not updating in Scan parameter Tab | WIP16 | WIP21 | Yes.To see the selected Liver DRI value on the Scan Parameters Toolbox, user has to perform a scan .After the scan the updated Liver DRI value shall be reflected under Parameters Panel region | Subsystem |  | None | Fixed | deshbhratar_m | 0 | 2 | Console Software | Console4 | YES | Closed | WIP18SubCom | IQon1 |  | WIP21 | September 17, 2019 8:13:10 PM GMT+08:00
CTUDT00624390 | Refresh | Closed | Dana_Kamyshev | June 11, 2019 7:28:08 PM GMT+08:00 | Refresh - Bolus tracker graph freezes | WIP17 | WIP18 | N/A | Subsystem | Bolus Tracking | N/A | Fixed | naveen_k | 0 | 2 | Console USP_EBW | Console4 | YES | Closed | WIP18SubCom | IQon1 |  | WIP18 | July 12, 2019 7:41:07 PM GMT+08:00
CTUDT00624385 | Refresh | Closed | 320053769 | June 11, 2019 6:06:35 PM GMT+08:00 | WIP17 - CCT fluro does not enable real time image displaying. | WIP17 |  |  | Subsystem | CCT |  | Duplicate | richards_b | 0 |  |  | CIRS | YES | Closed |  | Ovation
CTUDT00624385 | Refresh | Closed | 320053769 | June 11, 2019 6:06:35 PM GMT+08:00 | WIP17 - CCT fluro does not enable real time image displaying. | WIP17 |  |  | Subsystem | CCT |  | Duplicate | richards_b | 0 |  |  | CIRS | YES | Closed |  | IQon1
CTUDT00624383 | Refresh | Closed | 320053769 | June 11, 2019 5:54:28 PM GMT+08:00 | WIP17 - Initializing takes more than 40 seconds. | WIP17 | WIP18 |  | Subsystem | Performance | NA | Works As Specified / Not A Defect | tamari_o | 1 | 3 | Console USP_EBW | Support-System | NO | Closed | WIP18SubCom | Ovation |  | WIP18
CTUDT00624383 | Refresh | Closed | 320053769 | June 11, 2019 5:54:28 PM GMT+08:00 | WIP17 - Initializing takes more than 40 seconds. | WIP17 | WIP18 |  | Subsystem | Performance | NA | Works As Specified / Not A Defect | tamari_o | 1 | 3 | Console USP_EBW | Support-System | NO | Closed | WIP18SubCom | Brilliance 64 |  | WIP18
CTUDT00624383 | Refresh | Closed | 320053769 | June 11, 2019 5:54:28 PM GMT+08:00 | WIP17 - Initializing takes more than 40 seconds. | WIP17 | WIP18 |  | Subsystem | Performance | NA | Works As Specified / Not A Defect | tamari_o | 1 | 3 | Console USP_EBW | Support-System | NO | Closed | WIP18SubCom | Brilliance Big Bore |  | WIP18
CTUDT00624383 | Refresh | Closed | 320053769 | June 11, 2019 5:54:28 PM GMT+08:00 | WIP17 - Initializing takes more than 40 seconds. | WIP17 | WIP18 |  | Subsystem | Performance | NA | Works As Specified / Not A Defect | tamari_o | 1 | 3 | Console USP_EBW | Support-System | NO | Closed | WIP18SubCom | Ingenuity 64 |  | WIP18
CTUDT00624383 | Refresh | Closed | 320053769 | June 11, 2019 5:54:28 PM GMT+08:00 | WIP17 - Initializing takes more than 40 seconds. | WIP17 | WIP18 |  | Subsystem | Performance | NA | Works As Specified / Not A Defect | tamari_o | 1 | 3 | Console USP_EBW | Support-System | NO | Closed | WIP18SubCom | IQon1 |  | WIP18
CTUDT00624381 | Refresh | Closed | rojin_antony | June 11, 2019 5:23:01 PM GMT+08:00 | Max mAs value is shown red in color even after increasing the DRI (to avoid Clipping) | WIP16 | WIP18 | the clipping indication is also available as * sign on the plan on surview | Subsystem | Plan | NA | Fixed | datta_i | 0 | 3 | Console Software | Console4 | NO | Closed | WIP18SubCom | Ovation |  | WIP18 | June 18, 2019 5:20:57 PM GMT+08:00
CTUDT00624381 | Refresh | Closed | rojin_antony | June 11, 2019 5:23:01 PM GMT+08:00 | Max mAs value is shown red in color even after increasing the DRI (to avoid Clipping) | WIP16 | WIP18 | the clipping indication is also available as * sign on the plan on surview | Subsystem | Plan | NA | Fixed | datta_i | 0 | 3 | Console Software | Console4 | NO | Closed | WIP18SubCom | Brilliance 64 |  | WIP18 | June 18, 2019 5:20:57 PM GMT+08:00
CTUDT00624381 | Refresh | Closed | rojin_antony | June 11, 2019 5:23:01 PM GMT+08:00 | Max mAs value is shown red in color even after increasing the DRI (to avoid Clipping) | WIP16 | WIP18 | the clipping indication is also available as * sign on the plan on surview | Subsystem | Plan | NA | Fixed | datta_i | 0 | 3 | Console Software | Console4 | NO | Closed | WIP18SubCom | Ingenuity 64 |  | WIP18 | June 18, 2019 5:20:57 PM GMT+08:00
CTUDT00624381 | Refresh | Closed | rojin_antony | June 11, 2019 5:23:01 PM GMT+08:00 | Max mAs value is shown red in color even after increasing the DRI (to avoid Clipping) | WIP16 | WIP18 | the clipping indication is also available as * sign on the plan on surview | Subsystem | Plan | NA | Fixed | datta_i | 0 | 3 | Console Software | Console4 | NO | Closed | WIP18SubCom | IQon1 |  | WIP18 | June 18, 2019 5:20:57 PM GMT+08:00
CTUDT00624378 | Refresh | Closed | bhawna_g | June 11, 2019 4:58:24 PM GMT+08:00 | IST -ICE Client need to be upgraded to V3.3 to address security Vulnerability of ICEv3.2 | 4.8 |  |  | Subsystem |  |  | Duplicate | sowmya_k | 0 | 2 |  | Service Tools | YES | Closed |  | Brilliance Big Bore
CTUDT00624374 | Refresh | Closed | rojin_antony | June 11, 2019 4:53:56 PM GMT+08:00 | Mismatch in mAs clipped message | WIP16 | WIP19 |  | Subsystem | Plan | NA | Fixed | datta_i | 0 | 4 | Console Software | Console4 | NO | Closed | WIP19SegReq | Ovation |  | WIP19 | July 17, 2019 2:24:50 PM GMT+08:00
CTUDT00624374 | Refresh | Closed | rojin_antony | June 11, 2019 4:53:56 PM GMT+08:00 | Mismatch in mAs clipped message | WIP16 | WIP19 |  | Subsystem | Plan | NA | Fixed | datta_i | 0 | 4 | Console Software | Console4 | NO | Closed | WIP19SegReq | Brilliance 64 |  | WIP19 | July 17, 2019 2:24:50 PM GMT+08:00
CTUDT00624374 | Refresh | Closed | rojin_antony | June 11, 2019 4:53:56 PM GMT+08:00 | Mismatch in mAs clipped message | WIP16 | WIP19 |  | Subsystem | Plan | NA | Fixed | datta_i | 0 | 4 | Console Software | Console4 | NO | Closed | WIP19SegReq | Ingenuity 64 |  | WIP19 | July 17, 2019 2:24:50 PM GMT+08:00
CTUDT00624374 | Refresh | Closed | rojin_antony | June 11, 2019 4:53:56 PM GMT+08:00 | Mismatch in mAs clipped message | WIP16 | WIP19 |  | Subsystem | Plan | NA | Fixed | datta_i | 0 | 4 | Console Software | Console4 | NO | Closed | WIP19SegReq | IQon1 |  | WIP19 | July 17, 2019 2:24:50 PM GMT+08:00
CTUDT00624370 | Refresh | Closed | sshahar | June 11, 2019 4:47:36 PM GMT+08:00 | Refresh: Security\3rd party software settings differences | WIP16 | DHF312898 rev. 04 |  | Subsystem | System,Blocker | N/A | Fixed | borise | 0 | 3 | Design Documentation | Support-System | NO | Closed | WIP20SegReq | Ovation |  | DHF312898 rev. 04 | September 18, 2019 1:36:54 PM GMT+08:00
CTUDT00624370 | Refresh | Closed | sshahar | June 11, 2019 4:47:36 PM GMT+08:00 | Refresh: Security\3rd party software settings differences | WIP16 | DHF312898 rev. 04 |  | Subsystem | System,Blocker | N/A | Fixed | borise | 0 | 3 | Design Documentation | Support-System | NO | Closed | WIP20SegReq | IQon1 |  | DHF312898 rev. 04 | September 18, 2019 1:36:54 PM GMT+08:00
CTUDT00624367 | Refresh | Closed | guy_sheaffer | June 11, 2019 4:06:39 PM GMT+08:00 | Refresh: Beam and DMS Check: issues with reporting of signal time stability. | 4.7.x | WIP18 |  | Subsystem | Beam and DMS Check | n/a | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP18SegReq | IQon1 |  | WIP18 | June 24, 2019 10:46:08 PM GMT+08:00
CTUDT00624366 | Refresh | Closed | 320053770 | June 11, 2019 3:39:48 PM GMT+08:00 | During WIP 17 Sanity tests, in an Air calibration, the Air calibration was paused displaying an error. | wip17 | WIP19 |  | Subsystem | USP | n/a | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP20SegReq | IQon1 |  | WIP19 | August 6, 2019 11:31:58 PM GMT+08:00
CTUDT00624364 | Refresh | Closed | 320053769 | June 11, 2019 2:52:01 PM GMT+08:00 | Portal 6.5 doesn’t configure in refresh machines {iCT, iCT(SP), IQon} | WIP17 |  |  | Subsystem | Remote Configuration |  | Does Not Apply | borise | 0 |  |  | Support-System | NO | Closed |  | Ovation
CTUDT00624364 | Refresh | Closed | 320053769 | June 11, 2019 2:52:01 PM GMT+08:00 | Portal 6.5 doesn’t configure in refresh machines {iCT, iCT(SP), IQon} | WIP17 |  |  | Subsystem | Remote Configuration |  | Does Not Apply | borise | 0 |  |  | Support-System | NO | Closed |  | IQon1
CTUDT00624362 | Refresh | Closed | 320053769 | June 11, 2019 2:32:13 PM GMT+08:00 | When copying a study from portals 9/10/11 to Local the copy fails and reverse. | WIP17 |  |  | Subsystem | Remote Configuration |  | Does Not Apply | borise | 0 |  |  | Support-System | NO | Closed |  | Ovation
CTUDT00624362 | Refresh | Closed | 320053769 | June 11, 2019 2:32:13 PM GMT+08:00 | When copying a study from portals 9/10/11 to Local the copy fails and reverse. | WIP17 |  |  | Subsystem | Remote Configuration |  | Does Not Apply | borise | 0 |  |  | Support-System | NO | Closed |  | IQon1
CTUDT00624360 | Refresh | Closed | Dana_Kamyshev | June 11, 2019 2:15:25 PM GMT+08:00 | Refresh - Chinese letters doesn't display in 'Patient' page | WIP17 |  |  | Subsystem | Localization | N/A | Works As Specified / Not A Defect | wilsonraja | 0 |  |  | Console4 | NO | Closed | WAS | Ovation
CTUDT00624360 | Refresh | Closed | Dana_Kamyshev | June 11, 2019 2:15:25 PM GMT+08:00 | Refresh - Chinese letters doesn't display in 'Patient' page | WIP17 |  |  | Subsystem | Localization | N/A | Works As Specified / Not A Defect | wilsonraja | 0 |  |  | Console4 | NO | Closed | WAS | Brilliance 64
CTUDT00624360 | Refresh | Closed | Dana_Kamyshev | June 11, 2019 2:15:25 PM GMT+08:00 | Refresh - Chinese letters doesn't display in 'Patient' page | WIP17 |  |  | Subsystem | Localization | N/A | Works As Specified / Not A Defect | wilsonraja | 0 |  |  | Console4 | NO | Closed | WAS | Brilliance Big Bore
CTUDT00624360 | Refresh | Closed | Dana_Kamyshev | June 11, 2019 2:15:25 PM GMT+08:00 | Refresh - Chinese letters doesn't display in 'Patient' page | WIP17 |  |  | Subsystem | Localization | N/A | Works As Specified / Not A Defect | wilsonraja | 0 |  |  | Console4 | NO | Closed | WAS | Ingenuity 64
CTUDT00624360 | Refresh | Closed | Dana_Kamyshev | June 11, 2019 2:15:25 PM GMT+08:00 | Refresh - Chinese letters doesn't display in 'Patient' page | WIP17 |  |  | Subsystem | Localization | N/A | Works As Specified / Not A Defect | wilsonraja | 0 |  |  | Console4 | NO | Closed | WAS | IQon1
CTUDT00624356 | Refresh | Closed | Dana_Kamyshev | June 11, 2019 2:10:12 PM GMT+08:00 | Refresh - Double digits display in Chinese keyboard | WIP17 |  |  | Subsystem |  |  | Duplicate | Dana_Kamyshev | 0 |  |  |  |  | Closed |  | IQon1
CTUDT00624355 | Refresh | Closed | zhu_je | June 11, 2019 2:04:47 PM GMT+08:00 | Unable to set the start position value in clinical scan. | WIP16 | WIP18 | Use plan on surview or move the coach to the desired position | Subsystem | System | NA | Fixed | datta_i | 0 | 3 | Console Software | Console4 | NO | Closed | WIP19SubCom | Ovation |  | WIP18 | June 26, 2019 8:30:27 PM GMT+08:00
CTUDT00624355 | Refresh | Closed | zhu_je | June 11, 2019 2:04:47 PM GMT+08:00 | Unable to set the start position value in clinical scan. | WIP16 | WIP18 | Use plan on surview or move the coach to the desired position | Subsystem | System | NA | Fixed | datta_i | 0 | 3 | Console Software | Console4 | NO | Closed | WIP19SubCom | Brilliance 64 |  | WIP18 | June 26, 2019 8:30:27 PM GMT+08:00
CTUDT00624355 | Refresh | Closed | zhu_je | June 11, 2019 2:04:47 PM GMT+08:00 | Unable to set the start position value in clinical scan. | WIP16 | WIP18 | Use plan on surview or move the coach to the desired position | Subsystem | System | NA | Fixed | datta_i | 0 | 3 | Console Software | Console4 | NO | Closed | WIP19SubCom | Ingenuity 64 |  | WIP18 | June 26, 2019 8:30:27 PM GMT+08:00
CTUDT00624355 | Refresh | Closed | zhu_je | June 11, 2019 2:04:47 PM GMT+08:00 | Unable to set the start position value in clinical scan. | WIP16 | WIP18 | Use plan on surview or move the coach to the desired position | Subsystem | System | NA | Fixed | datta_i | 0 | 3 | Console Software | Console4 | NO | Closed | WIP19SubCom | IQon1 |  | WIP18 | June 26, 2019 8:30:27 PM GMT+08:00
CTUDT00624345 | Refresh | Closed | cai_t | June 10, 2019 9:09:40 PM GMT+08:00 | [Refresh]Repeat Tracker button is not visible after pressing Cancel button on Control box after last tracker | WIP16 |  | NA | Subsystem | Tracker | NA | Works As Specified / Not A Defect | cai_t | 0 | 2 | Test Design | Support-System | NO | Closed | WIP18SubCom | Ingenuity 64 |  |  | July 12, 2019 4:40:27 PM GMT+08:00
CTUDT00624344 | Refresh | Closed | Noy_Seelig | June 10, 2019 8:54:13 PM GMT+08:00 | Refresh: IQ Check: No message regarding tilt or phantom not in center. | WIP16 |  |  | Subsystem | IQ Check | n/a | Works As Specified / Not A Defect | yosfan_s | 0 |  |  | DMS | NO | Closed |  | Ovation
CTUDT00624343 | Refresh | Closed | jotawar_d | June 10, 2019 7:42:47 PM GMT+08:00 | [BB Refresh]"MRP/MINP" is displayed on scan result images if render mode=MINIP and orientation=axial | WIP18 |  | NA | Subsystem | DICOM | Console.SRS.12800 | No Plan to Fix | saha_d | 0 | 4 |  | Console4 | NO | Closed | SubNC | Ovation
CTUDT00624343 | Refresh | Closed | jotawar_d | June 10, 2019 7:42:47 PM GMT+08:00 | [BB Refresh]"MRP/MINP" is displayed on scan result images if render mode=MINIP and orientation=axial | WIP18 |  | NA | Subsystem | DICOM | Console.SRS.12800 | No Plan to Fix | saha_d | 0 | 4 |  | Console4 | NO | Closed | SubNC | Brilliance 64
CTUDT00624343 | Refresh | Closed | jotawar_d | June 10, 2019 7:42:47 PM GMT+08:00 | [BB Refresh]"MRP/MINP" is displayed on scan result images if render mode=MINIP and orientation=axial | WIP18 |  | NA | Subsystem | DICOM | Console.SRS.12800 | No Plan to Fix | saha_d | 0 | 4 |  | Console4 | NO | Closed | SubNC | Ingenuity 64
CTUDT00624343 | Refresh | Closed | jotawar_d | June 10, 2019 7:42:47 PM GMT+08:00 | [BB Refresh]"MRP/MINP" is displayed on scan result images if render mode=MINIP and orientation=axial | WIP18 |  | NA | Subsystem | DICOM | Console.SRS.12800 | No Plan to Fix | saha_d | 0 | 4 |  | Console4 | NO | Closed | SubNC | IQon1
CTUDT00624336 | Refresh | Closed | azagury_d | June 10, 2019 3:52:41 PM GMT+08:00 | Refresh: Calibrations: Performance Test Infant : In Body Infant scans in Performance Infant Window W=100, instead of 200. | WIP16 | N/A |  | Subsystem | Performance Test | n/a | Fixed | braunstein_d | 0 | 4 | Design Documentation | Support-Physics | NO | Closed | WIP18SubCom | IQon1 |  | N/A | July 8, 2019 4:32:13 PM GMT+08:00
CTUDT00624329 | Refresh | Closed | lei_yang | June 10, 2019 2:30:49 PM GMT+08:00 | No 'Duplicate' option while right click the Body Perfusion scan in the plan scene. | WIP16 | DHF303211 Rev13 Host SSRS for Refresh DHF316705 Rev.12 Host SSRS for Refresh BB |  | Subsystem | Plan | NA | Fixed | hatti_c | 0 | 3 | Requirements Documentation | Console4 | NO | Closed | WAS | Ingenuity 64 |  | DHF303211 Rev13 Host SSRS for Refresh DHF316705 Rev.12 Host SSRS for Refresh BB | July 19, 2019 5:36:36 PM GMT+08:00
CTUDT00624327 | Refresh | Closed | jayakumar_p | June 10, 2019 2:06:36 PM GMT+08:00 | Refresh: Injection icon showed on duplicated step which has injection set to false. | WIP17 |  | None | Subsystem | Plan | NA | No Plan to Fix | veeraswamy_k | 0 | 4 |  | Console4 | NO | Closed | SubNC | Ovation
CTUDT00624327 | Refresh | Closed | jayakumar_p | June 10, 2019 2:06:36 PM GMT+08:00 | Refresh: Injection icon showed on duplicated step which has injection set to false. | WIP17 |  | None | Subsystem | Plan | NA | No Plan to Fix | veeraswamy_k | 0 | 4 |  | Console4 | NO | Closed | SubNC | Brilliance 64
CTUDT00624327 | Refresh | Closed | jayakumar_p | June 10, 2019 2:06:36 PM GMT+08:00 | Refresh: Injection icon showed on duplicated step which has injection set to false. | WIP17 |  | None | Subsystem | Plan | NA | No Plan to Fix | veeraswamy_k | 0 | 4 |  | Console4 | NO | Closed | SubNC | Brilliance Big Bore
CTUDT00624327 | Refresh | Closed | jayakumar_p | June 10, 2019 2:06:36 PM GMT+08:00 | Refresh: Injection icon showed on duplicated step which has injection set to false. | WIP17 |  | None | Subsystem | Plan | NA | No Plan to Fix | veeraswamy_k | 0 | 4 |  | Console4 | NO | Closed | SubNC | Ingenuity 64
CTUDT00624327 | Refresh | Closed | jayakumar_p | June 10, 2019 2:06:36 PM GMT+08:00 | Refresh: Injection icon showed on duplicated step which has injection set to false. | WIP17 |  | None | Subsystem | Plan | NA | No Plan to Fix | veeraswamy_k | 0 | 4 |  | Console4 | NO | Closed | SubNC | IQon1
CTUDT00624319 | Refresh | Closed | Sandeep_Kumar | June 7, 2019 5:23:25 PM GMT+08:00 | The 'Pause" icon in system does not match with white cover design | 4.8 | BB_WIP8 |  | Subsystem | Status Bar | Console-SRS.16167, STPS-SSRS-1301 | Fixed | Sandeep_Kumar | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP8_SubCom | Brilliance Big Bore |  | BB_WIP8 | June 13, 2019 8:15:58 PM GMT+08:00
CTUDT00624317 | Refresh | Closed | manjunath_c | June 7, 2019 3:44:56 PM GMT+08:00 | [Refresh]: During upgrading from 4.1.7 to 4.1.10 in host installation path an error message was displayed | WIP16 | WIP19 | Tester and Service personal can remove the IPV6 Adress format in the IP adress field and proceed with backup | Subsystem | 1-Click | STPS-SSRS-1157 | Fixed | venugopal_s | 0 | 3 | Service Tools_Software | DI | NO | Closed | WIP19SegReq | Ingenuity 64 |  | WIP19 | July 16, 2019 6:04:42 PM GMT+08:00
CTUDT00624315 | Refresh | Closed | Noy_Seelig | June 7, 2019 1:51:53 AM GMT+08:00 | Refresh: Calibrations: Performance Test Infant wrong limits | WIP16 | WIP18 |  | Subsystem | Performance Test | n/a | Fixed | elviral | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP18SubCom | IQon1 |  | WIP18 | June 13, 2019 5:27:38 PM GMT+08:00
CTUDT00624309 | Refresh | Closed | bhawna_g | June 6, 2019 9:19:43 PM GMT+08:00 | Planned SSDE is getting updated in the Acquisition viewer, tool box and show all after the scan | 4.8 | BB_WIP8 |  | Subsystem | SSDE | NA | Fixed | bhawna_g | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP8_SegReq | Brilliance Big Bore |  | BB_WIP8 | June 13, 2019 2:48:36 PM GMT+08:00
CTUDT00624308 | Refresh | Closed | bhawna_g | June 6, 2019 9:15:11 PM GMT+08:00 | In CCT Fluro SSDE value is displaying as zero in acquisition viewer after releasing the pedal first time | 4.8 | BB_WIP8 |  | Subsystem | SSDE | NA | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | YES | Closed | BB_WIP8_SegReq | Brilliance Big Bore |  | BB_WIP8 | June 11, 2019 6:53:17 PM GMT+08:00
CTUDT00624305 | Refresh | Closed | naveen_k | June 6, 2019 3:21:05 PM GMT+08:00 | [BB Refresh]The "Add injection" and "Remove Injection" are displayed incorrect for scan ruler bubble right click | WIP17 |  |  | Subsystem | Plan | NA | Works As Specified / Not A Defect | jayakumar_p | 0 | 3 |  | Console4 | NO | Closed | WAS | Ovation
CTUDT00624305 | Refresh | Closed | naveen_k | June 6, 2019 3:21:05 PM GMT+08:00 | [BB Refresh]The "Add injection" and "Remove Injection" are displayed incorrect for scan ruler bubble right click | WIP17 |  |  | Subsystem | Plan | NA | Works As Specified / Not A Defect | jayakumar_p | 0 | 3 |  | Console4 | NO | Closed | WAS | Brilliance 64
CTUDT00624305 | Refresh | Closed | naveen_k | June 6, 2019 3:21:05 PM GMT+08:00 | [BB Refresh]The "Add injection" and "Remove Injection" are displayed incorrect for scan ruler bubble right click | WIP17 |  |  | Subsystem | Plan | NA | Works As Specified / Not A Defect | jayakumar_p | 0 | 3 |  | Console4 | NO | Closed | WAS | Ingenuity 64
CTUDT00624305 | Refresh | Closed | naveen_k | June 6, 2019 3:21:05 PM GMT+08:00 | [BB Refresh]The "Add injection" and "Remove Injection" are displayed incorrect for scan ruler bubble right click | WIP17 |  |  | Subsystem | Plan | NA | Works As Specified / Not A Defect | jayakumar_p | 0 | 3 |  | Console4 | NO | Closed | WAS | IQon1
CTUDT00624302 | Refresh | Closed | shekhar_p | June 6, 2019 2:29:58 PM GMT+08:00 | Pop up appears "System error, please restart the system" | WIP16 |  | Workaround for the issue is to restart the machine once the issue comes. | Subsystem | Pause Scan | NA | Works As Specified / Not A Defect | s_sa | 1 |  |  | iCAP | NO | Closed |  | Ovation
CTUDT00624302 | Refresh | Closed | shekhar_p | June 6, 2019 2:29:58 PM GMT+08:00 | Pop up appears "System error, please restart the system" | WIP16 |  | Workaround for the issue is to restart the machine once the issue comes. | Subsystem | Pause Scan | NA | Works As Specified / Not A Defect | s_sa | 1 |  |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00624302 | Refresh | Closed | shekhar_p | June 6, 2019 2:29:58 PM GMT+08:00 | Pop up appears "System error, please restart the system" | WIP16 |  | Workaround for the issue is to restart the machine once the issue comes. | Subsystem | Pause Scan | NA | Works As Specified / Not A Defect | s_sa | 1 |  |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00624302 | Refresh | Closed | shekhar_p | June 6, 2019 2:29:58 PM GMT+08:00 | Pop up appears "System error, please restart the system" | WIP16 |  | Workaround for the issue is to restart the machine once the issue comes. | Subsystem | Pause Scan | NA | Works As Specified / Not A Defect | s_sa | 1 |  |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00624302 | Refresh | Closed | shekhar_p | June 6, 2019 2:29:58 PM GMT+08:00 | Pop up appears "System error, please restart the system" | WIP16 |  | Workaround for the issue is to restart the machine once the issue comes. | Subsystem | Pause Scan | NA | Works As Specified / Not A Defect | s_sa | 1 |  |  | iCAP | NO | Closed |  | IQon1
CTUDT00624298 | Refresh | Closed | nayak_r | June 6, 2019 1:01:40 PM GMT+08:00 | BB Refresh:The VD tool does not display all the sub-systems as referred in SubsystemDiagnosticsBase.cfg file | BB_WIP7 | DHF330449 Rev 08 Host verification Test protocol |  | Subsystem |  | NA | Fixed | nayak_r | 0 | 2 | Test Design | Service Tools | NO | Closed | WIP18SubCom | Brilliance 64 |  | DHF330449 Rev 08 Host verification Test protocol | June 27, 2019 9:18:37 PM GMT+08:00
CTUDT00624298 | Refresh | Closed | nayak_r | June 6, 2019 1:01:40 PM GMT+08:00 | BB Refresh:The VD tool does not display all the sub-systems as referred in SubsystemDiagnosticsBase.cfg file | BB_WIP7 | DHF330449 Rev 08 Host verification Test protocol |  | Subsystem |  | NA | Fixed | nayak_r | 0 | 2 | Test Design | Service Tools | NO | Closed | WIP18SubCom | Brilliance Big Bore |  | DHF330449 Rev 08 Host verification Test protocol | June 27, 2019 9:18:37 PM GMT+08:00
CTUDT00624298 | Refresh | Closed | nayak_r | June 6, 2019 1:01:40 PM GMT+08:00 | BB Refresh:The VD tool does not display all the sub-systems as referred in SubsystemDiagnosticsBase.cfg file | BB_WIP7 | DHF330449 Rev 08 Host verification Test protocol |  | Subsystem |  | NA | Fixed | nayak_r | 0 | 2 | Test Design | Service Tools | NO | Closed | WIP18SubCom | Ingenuity 64 |  | DHF330449 Rev 08 Host verification Test protocol | June 27, 2019 9:18:37 PM GMT+08:00
CTUDT00624296 | Refresh | Closed | finestine_m | June 6, 2019 2:55:58 AM GMT+08:00 | Refresh: Performance Test: Wrong Scan Number For Head Cases With Collimation Of 40 mm | WIP16 | WIP18 |  | Subsystem | Performance Test | n/a | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP18SubCom | Ovation |  | WIP18 | June 16, 2019 9:14:38 PM GMT+08:00
CTUDT00624296 | Refresh | Closed | finestine_m | June 6, 2019 2:55:58 AM GMT+08:00 | Refresh: Performance Test: Wrong Scan Number For Head Cases With Collimation Of 40 mm | WIP16 | WIP18 |  | Subsystem | Performance Test | n/a | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP18SubCom | IQon1 |  | WIP18 | June 16, 2019 9:14:38 PM GMT+08:00
CTUDT00624293 | Refresh | Closed | finestine_m | June 5, 2019 10:55:20 PM GMT+08:00 | Refresh: Performance Test: Wrong Phantom Position in Head Scans | WIP16 | WIP18 WIP23 |  | Subsystem | Performance Test | NaN | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP23SubCom | Ovation |  | WIP18 WIP23 | December 16, 2019 8:37:41 PM GMT+08:00
CTUDT00624291 | Refresh | Closed | zhong_x | June 5, 2019 8:29:43 PM GMT+08:00 | Scanner application was not accessible for emergency user | BB_WIP6 |  | Username should be in block letters ("EMERGENCY"). | Subsystem | Login | NA | Duplicate | kr_m | 0 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00624291 | Refresh | Closed | zhong_x | June 5, 2019 8:29:43 PM GMT+08:00 | Scanner application was not accessible for emergency user | BB_WIP6 |  | Username should be in block letters ("EMERGENCY"). | Subsystem | Login | NA | Duplicate | kr_m | 0 |  |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00624291 | Refresh | Closed | zhong_x | June 5, 2019 8:29:43 PM GMT+08:00 | Scanner application was not accessible for emergency user | BB_WIP6 |  | Username should be in block letters ("EMERGENCY"). | Subsystem | Login | NA | Duplicate | kr_m | 0 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00624291 | Refresh | Closed | zhong_x | June 5, 2019 8:29:43 PM GMT+08:00 | Scanner application was not accessible for emergency user | BB_WIP6 |  | Username should be in block letters ("EMERGENCY"). | Subsystem | Login | NA | Duplicate | kr_m | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00624291 | Refresh | Closed | zhong_x | June 5, 2019 8:29:43 PM GMT+08:00 | Scanner application was not accessible for emergency user | BB_WIP6 |  | Username should be in block letters ("EMERGENCY"). | Subsystem | Login | NA | Duplicate | kr_m | 0 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00624289 | Refresh | Closed | gT_s | June 5, 2019 6:56:57 PM GMT+08:00 | Refresh : The few commands not getting executed in SSH, it displays an error as “Command not found”. | WIP16 | BB_WIP9 WIP19 |  | Subsystem | Remote Connection | NA | Fixed | Sandeep_Kumar | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP18SubCom | Ovation |  | BB_WIP9 WIP19 | September 20, 2019 4:16:34 PM GMT+08:00
CTUDT00624289 | Refresh | Closed | gT_s | June 5, 2019 6:56:57 PM GMT+08:00 | Refresh : The few commands not getting executed in SSH, it displays an error as “Command not found”. | WIP16 | BB_WIP9 WIP19 |  | Subsystem | Remote Connection | NA | Fixed | Sandeep_Kumar | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP18SubCom | Brilliance 64 |  | BB_WIP9 WIP19 | September 20, 2019 4:16:34 PM GMT+08:00
CTUDT00624289 | Refresh | Closed | gT_s | June 5, 2019 6:56:57 PM GMT+08:00 | Refresh : The few commands not getting executed in SSH, it displays an error as “Command not found”. | WIP16 | BB_WIP9 WIP19 |  | Subsystem | Remote Connection | NA | Fixed | Sandeep_Kumar | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP18SubCom | Brilliance Big Bore |  | BB_WIP9 WIP19 | September 20, 2019 4:16:34 PM GMT+08:00
CTUDT00624289 | Refresh | Closed | gT_s | June 5, 2019 6:56:57 PM GMT+08:00 | Refresh : The few commands not getting executed in SSH, it displays an error as “Command not found”. | WIP16 | BB_WIP9 WIP19 |  | Subsystem | Remote Connection | NA | Fixed | Sandeep_Kumar | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP18SubCom | Ingenuity 64 |  | BB_WIP9 WIP19 | September 20, 2019 4:16:34 PM GMT+08:00
CTUDT00624289 | Refresh | Closed | gT_s | June 5, 2019 6:56:57 PM GMT+08:00 | Refresh : The few commands not getting executed in SSH, it displays an error as “Command not found”. | WIP16 | BB_WIP9 WIP19 |  | Subsystem | Remote Connection | NA | Fixed | Sandeep_Kumar | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP18SubCom | IQon1 |  | BB_WIP9 WIP19 | September 20, 2019 4:16:34 PM GMT+08:00
CTUDT00624288 | Refresh | Closed | sheik_c | June 5, 2019 5:52:19 PM GMT+08:00 | [BB Refresh]The voice is not played before scan if the releated voice was added slience and then remove | WIP17 | WIP18 |  | Subsystem |  | Console-SRS.2882,Console-SRS.2884 | Fixed | sheik_c | 0 | 2 | Console Software | Console4 | YES | Closed | WIP18SubCom | Ovation |  | WIP18 | June 7, 2019 1:27:13 PM GMT+08:00
CTUDT00624288 | Refresh | Closed | sheik_c | June 5, 2019 5:52:19 PM GMT+08:00 | [BB Refresh]The voice is not played before scan if the releated voice was added slience and then remove | WIP17 | WIP18 |  | Subsystem |  | Console-SRS.2882,Console-SRS.2884 | Fixed | sheik_c | 0 | 2 | Console Software | Console4 | YES | Closed | WIP18SubCom | Brilliance 64 |  | WIP18 | June 7, 2019 1:27:13 PM GMT+08:00
CTUDT00624288 | Refresh | Closed | sheik_c | June 5, 2019 5:52:19 PM GMT+08:00 | [BB Refresh]The voice is not played before scan if the releated voice was added slience and then remove | WIP17 | WIP18 |  | Subsystem |  | Console-SRS.2882,Console-SRS.2884 | Fixed | sheik_c | 0 | 2 | Console Software | Console4 | YES | Closed | WIP18SubCom | Ingenuity 64 |  | WIP18 | June 7, 2019 1:27:13 PM GMT+08:00
CTUDT00624288 | Refresh | Closed | sheik_c | June 5, 2019 5:52:19 PM GMT+08:00 | [BB Refresh]The voice is not played before scan if the releated voice was added slience and then remove | WIP17 | WIP18 |  | Subsystem |  | Console-SRS.2882,Console-SRS.2884 | Fixed | sheik_c | 0 | 2 | Console Software | Console4 | YES | Closed | WIP18SubCom | IQon1 |  | WIP18 | June 7, 2019 1:27:13 PM GMT+08:00
CTUDT00624284 | Refresh | Closed | lei_yang | June 5, 2019 5:10:49 PM GMT+08:00 | The actual Cycle Time in the Log Viewer isn't equal to the plan Cycle Time value after performing Brain Perfusion Jog exam card. | WIP16 |  |  | Subsystem | Brain Perfusion Jog | NA | Works As Specified / Not A Defect | lei_yang | 0 | 2 | Design Documentation | Support-System | NO | Closed | WAS | Ingenuity 64 |  |  | June 26, 2019 6:36:30 PM GMT+08:00
CTUDT00624271 | Refresh | Closed | gT_s | June 5, 2019 11:05:20 AM GMT+08:00 | Refresh : The application gets crashed while doing the offline recon for multiple results of different spectral type. | WIP17 | WIP18 |  | Subsystem |  | None | Fixed | Prakash_Chand_Yadav | 0 | 2 | Console Software | Console4 | NO | Closed | WIP18SubCom | IQon1 |  | WIP18 | June 7, 2019 1:18:20 PM GMT+08:00
CTUDT00624268 | Refresh | Closed | zhu_je | June 5, 2019 9:45:08 AM GMT+08:00 | [BB]Error occurred during scan when "scan with xray" check box was uncheck and diagnostic level was set to 1 | BB_WIP6 |  |  | Subsystem | System | Console SRS-15466 | Works As Specified / Not A Defect | prasad_r | 0 | 4 |  | Console4 | NO | Closed | WAS | Ovation
CTUDT00624268 | Refresh | Closed | zhu_je | June 5, 2019 9:45:08 AM GMT+08:00 | [BB]Error occurred during scan when "scan with xray" check box was uncheck and diagnostic level was set to 1 | BB_WIP6 |  |  | Subsystem | System | Console SRS-15466 | Works As Specified / Not A Defect | prasad_r | 0 | 4 |  | Console4 | NO | Closed | WAS | Brilliance 64
CTUDT00624268 | Refresh | Closed | zhu_je | June 5, 2019 9:45:08 AM GMT+08:00 | [BB]Error occurred during scan when "scan with xray" check box was uncheck and diagnostic level was set to 1 | BB_WIP6 |  |  | Subsystem | System | Console SRS-15466 | Works As Specified / Not A Defect | prasad_r | 0 | 4 |  | Console4 | NO | Closed | WAS | Brilliance Big Bore
CTUDT00624268 | Refresh | Closed | zhu_je | June 5, 2019 9:45:08 AM GMT+08:00 | [BB]Error occurred during scan when "scan with xray" check box was uncheck and diagnostic level was set to 1 | BB_WIP6 |  |  | Subsystem | System | Console SRS-15466 | Works As Specified / Not A Defect | prasad_r | 0 | 4 |  | Console4 | NO | Closed | WAS | Ingenuity 64
CTUDT00624268 | Refresh | Closed | zhu_je | June 5, 2019 9:45:08 AM GMT+08:00 | [BB]Error occurred during scan when "scan with xray" check box was uncheck and diagnostic level was set to 1 | BB_WIP6 |  |  | Subsystem | System | Console SRS-15466 | Works As Specified / Not A Defect | prasad_r | 0 | 4 |  | Console4 | NO | Closed | WAS | IQon1
CTUDT00624265 | Refresh | Closed | ding_s | June 4, 2019 6:03:12 PM GMT+08:00 | Copy study to DVD-R failed | WIP16 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | ding_s | 0 |  |  |  |  | Closed |  | Ingenuity 64
CTUDT00624263 | Refresh | Closed | zhu_je | June 4, 2019 4:40:10 PM GMT+08:00 | [BB Refresh][IVC]There is error message popping up and scan was paused if moving couch using IVC after clicking go button | BB_WIP6 |  | Since issue is sporadic user can retry the couch movement after clicking GO again | Subsystem | CCT;IVC;Blocker | Console-SRS.16541 | Works As Specified / Not A Defect | reddy_g | 0 | 3 |  | Console4 | NO | Closed | Not Valid scenario | Brilliance Big Bore
CTUDT00624260 | Refresh | Closed | braunstein_d | June 4, 2019 3:48:50 PM GMT+08:00 | H-phantom centering do not work if only Adult Head section is installed. | WIP17 | WIP18 |  | Subsystem | Calbrations | NaN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP18SegReq | Ovation |  | WIP18 | July 7, 2019 10:11:31 PM GMT+08:00
CTUDT00624260 | Refresh | Closed | braunstein_d | June 4, 2019 3:48:50 PM GMT+08:00 | H-phantom centering do not work if only Adult Head section is installed. | WIP17 | WIP18 |  | Subsystem | Calbrations | NaN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP18SegReq | Brilliance Big Bore |  | WIP18 | July 7, 2019 10:11:31 PM GMT+08:00
CTUDT00624260 | Refresh | Closed | braunstein_d | June 4, 2019 3:48:50 PM GMT+08:00 | H-phantom centering do not work if only Adult Head section is installed. | WIP17 | WIP18 |  | Subsystem | Calbrations | NaN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP18SegReq | Ingenuity 64 |  | WIP18 | July 7, 2019 10:11:31 PM GMT+08:00
CTUDT00624258 | Refresh | Closed | ding_s | June 4, 2019 3:15:53 PM GMT+08:00 | When copy images to configured PACS, the progress showed 100%,but the status remained Awaiting Commitment, and no transferred images found in PACS server. | WIP16 |  |  | Subsystem | Connectivity | NA | Works As Specified / Not A Defect | s_sa | 0 | 2 |  | iCAP | NO | Closed | WIP19SegReq | Ovation
CTUDT00624258 | Refresh | Closed | ding_s | June 4, 2019 3:15:53 PM GMT+08:00 | When copy images to configured PACS, the progress showed 100%,but the status remained Awaiting Commitment, and no transferred images found in PACS server. | WIP16 |  |  | Subsystem | Connectivity | NA | Works As Specified / Not A Defect | s_sa | 0 | 2 |  | iCAP | NO | Closed | WIP19SegReq | Brilliance 64
CTUDT00624258 | Refresh | Closed | ding_s | June 4, 2019 3:15:53 PM GMT+08:00 | When copy images to configured PACS, the progress showed 100%,but the status remained Awaiting Commitment, and no transferred images found in PACS server. | WIP16 |  |  | Subsystem | Connectivity | NA | Works As Specified / Not A Defect | s_sa | 0 | 2 |  | iCAP | NO | Closed | WIP19SegReq | Brilliance Big Bore
CTUDT00624258 | Refresh | Closed | ding_s | June 4, 2019 3:15:53 PM GMT+08:00 | When copy images to configured PACS, the progress showed 100%,but the status remained Awaiting Commitment, and no transferred images found in PACS server. | WIP16 |  |  | Subsystem | Connectivity | NA | Works As Specified / Not A Defect | s_sa | 0 | 2 |  | iCAP | NO | Closed | WIP19SegReq | Ingenuity 64
CTUDT00624258 | Refresh | Closed | ding_s | June 4, 2019 3:15:53 PM GMT+08:00 | When copy images to configured PACS, the progress showed 100%,but the status remained Awaiting Commitment, and no transferred images found in PACS server. | WIP16 |  |  | Subsystem | Connectivity | NA | Works As Specified / Not A Defect | s_sa | 0 | 2 |  | iCAP | NO | Closed | WIP19SegReq | IQon1
CTUDT00624255 | Refresh | Closed | manju_kolkar | June 4, 2019 2:33:23 PM GMT+08:00 | Big Bore Refresh - The Emergency user name entered on pop up window was not displayed in User Interface | BB_WIP5.1 | BB_WIP9 | User needs to re-login to console. | Subsystem | Login | Console-SRS.15955 | Fixed | kr_m | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP9_SubCom | Brilliance Big Bore |  | BB_WIP9 | July 22, 2019 4:59:53 PM GMT+08:00
CTUDT00624254 | Refresh | Closed | nayak_r | June 4, 2019 2:28:31 PM GMT+08:00 | BB Refresh:On clicking End exam and Continue, Exam summary was not created with “Patient information and Exam information” pages | BB_WIP7 | BB_WIP8 |  | Subsystem |  | Console.SRS.2155 | Fixed | sagar_v | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP8_SubCom | Brilliance Big Bore |  | BB_WIP8 | June 21, 2019 5:04:26 PM GMT+08:00
CTUDT00624240 | Refresh | Closed | 320053770 | June 4, 2019 2:13:07 AM GMT+08:00 | After WIP 15 – 17 Installation, The GPC didn't sync with the HOST NTP Server. | WIP15 - WIP 17 | WIP18 | The GPC could be set up manually to synchronize time with Host PC. | Subsystem | Installation | NA | Fixed | kline_w | 0 | 2 | Gantry Software | GCX | NO | Closed | WIP19SegReq | Ovation |  | WIP18 | July 9, 2019 11:14:36 PM GMT+08:00
CTUDT00624240 | Refresh | Closed | 320053770 | June 4, 2019 2:13:07 AM GMT+08:00 | After WIP 15 – 17 Installation, The GPC didn't sync with the HOST NTP Server. | WIP15 - WIP 17 | WIP18 | The GPC could be set up manually to synchronize time with Host PC. | Subsystem | Installation | NA | Fixed | kline_w | 0 | 2 | Gantry Software | GCX | NO | Closed | WIP19SegReq | IQon1 |  | WIP18 | July 9, 2019 11:14:36 PM GMT+08:00
CTUDT00624234 | Refresh | Closed | azagury_d | June 3, 2019 9:41:40 PM GMT+08:00 | Refresh: Calibrations: Calibration Main Menu: Mismatches between spec and QA menu in Service user, | WIP16 | N/A |  | Subsystem | Blocker | N/A | Fixed | braunstein_d | 0 | 2 | Design Documentation | Support-Physics | NO | Closed | WIP18SegReq | IQon1 |  | N/A | June 16, 2019 3:00:19 PM GMT+08:00
CTUDT00624232 | Refresh | Closed | azagury_d | June 3, 2019 9:23:20 PM GMT+08:00 | Refresh: Calibrations: Calibration Main Menu: Mismatches between spec and QA menu | WIP16 | WIP18 |  | Subsystem | Calibrations; Verified1 | n/a | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP18SubCom | Ingenuity 64 |  | WIP18 | June 26, 2019 3:01:07 PM GMT+08:00
CTUDT00624232 | Refresh | Closed | azagury_d | June 3, 2019 9:23:20 PM GMT+08:00 | Refresh: Calibrations: Calibration Main Menu: Mismatches between spec and QA menu | WIP16 | WIP18 |  | Subsystem | Calibrations; Verified1 | n/a | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP18SubCom | IQon1 |  | WIP18 | June 26, 2019 3:01:07 PM GMT+08:00
CTUDT00624223 | Refresh | Closed | swamypreethi | June 3, 2019 5:18:19 PM GMT+08:00 | IOCC_Refresh: IHE Verification :- SOP class attributes issues during Performed Procedure Step exception management | WIP15 |  | N/A | Subsystem | DICOM | N/A | Works As Specified / Not A Defect | naveen_k | 0 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00624223 | Refresh | Closed | swamypreethi | June 3, 2019 5:18:19 PM GMT+08:00 | IOCC_Refresh: IHE Verification :- SOP class attributes issues during Performed Procedure Step exception management | WIP15 |  | N/A | Subsystem | DICOM | N/A | Works As Specified / Not A Defect | naveen_k | 0 |  |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00624223 | Refresh | Closed | swamypreethi | June 3, 2019 5:18:19 PM GMT+08:00 | IOCC_Refresh: IHE Verification :- SOP class attributes issues during Performed Procedure Step exception management | WIP15 |  | N/A | Subsystem | DICOM | N/A | Works As Specified / Not A Defect | naveen_k | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00624223 | Refresh | Closed | swamypreethi | June 3, 2019 5:18:19 PM GMT+08:00 | IOCC_Refresh: IHE Verification :- SOP class attributes issues during Performed Procedure Step exception management | WIP15 |  | N/A | Subsystem | DICOM | N/A | Works As Specified / Not A Defect | naveen_k | 0 |  |  | Console4 | NO | Closed |  | Ingenuity BigBore
CTUDT00624223 | Refresh | Closed | swamypreethi | June 3, 2019 5:18:19 PM GMT+08:00 | IOCC_Refresh: IHE Verification :- SOP class attributes issues during Performed Procedure Step exception management | WIP15 |  | N/A | Subsystem | DICOM | N/A | Works As Specified / Not A Defect | naveen_k | 0 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00624218 | Refresh | Closed | swamypreethi | June 3, 2019 5:11:45 PM GMT+08:00 | IOCC_Refresh:IHE Verification:- Missing query options in Broad Worklist Query | WIP15 |  | N/A | Subsystem | DICOM | N/A | No Plan to Fix | naveen_k | 0 | 4 |  | Console4 | NO | Closed |  | Ovation
CTUDT00624218 | Refresh | Closed | swamypreethi | June 3, 2019 5:11:45 PM GMT+08:00 | IOCC_Refresh:IHE Verification:- Missing query options in Broad Worklist Query | WIP15 |  | N/A | Subsystem | DICOM | N/A | No Plan to Fix | naveen_k | 0 | 4 |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00624218 | Refresh | Closed | swamypreethi | June 3, 2019 5:11:45 PM GMT+08:00 | IOCC_Refresh:IHE Verification:- Missing query options in Broad Worklist Query | WIP15 |  | N/A | Subsystem | DICOM | N/A | No Plan to Fix | naveen_k | 0 | 4 |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00624218 | Refresh | Closed | swamypreethi | June 3, 2019 5:11:45 PM GMT+08:00 | IOCC_Refresh:IHE Verification:- Missing query options in Broad Worklist Query | WIP15 |  | N/A | Subsystem | DICOM | N/A | No Plan to Fix | naveen_k | 0 | 4 |  | Console4 | NO | Closed |  | Ingenuity BigBore
CTUDT00624218 | Refresh | Closed | swamypreethi | June 3, 2019 5:11:45 PM GMT+08:00 | IOCC_Refresh:IHE Verification:- Missing query options in Broad Worklist Query | WIP15 |  | N/A | Subsystem | DICOM | N/A | No Plan to Fix | naveen_k | 0 | 4 |  | Console4 | NO | Closed |  | IQon1
CTUDT00624216 | Refresh | Closed | swamypreethi | June 3, 2019 4:58:00 PM GMT+08:00 | IOCC_Refresh: IHE_Verification :- C_FIND and C_STORE attribute values not same | WIP15 | WIP18 | There is no workaround. | Subsystem | DICOM | Console-SRS.12924 | Fixed | jayakumar_p | 0 | 4 | Console Software | Console4 | NO | Closed | WIP19SegReq | Ovation |  | WIP18 | June 21, 2019 7:16:32 PM GMT+08:00
CTUDT00624216 | Refresh | Closed | swamypreethi | June 3, 2019 4:58:00 PM GMT+08:00 | IOCC_Refresh: IHE_Verification :- C_FIND and C_STORE attribute values not same | WIP15 | WIP18 | There is no workaround. | Subsystem | DICOM | Console-SRS.12924 | Fixed | jayakumar_p | 0 | 4 | Console Software | Console4 | NO | Closed | WIP19SegReq | Brilliance 64 |  | WIP18 | June 21, 2019 7:16:32 PM GMT+08:00
CTUDT00624216 | Refresh | Closed | swamypreethi | June 3, 2019 4:58:00 PM GMT+08:00 | IOCC_Refresh: IHE_Verification :- C_FIND and C_STORE attribute values not same | WIP15 | WIP18 | There is no workaround. | Subsystem | DICOM | Console-SRS.12924 | Fixed | jayakumar_p | 0 | 4 | Console Software | Console4 | NO | Closed | WIP19SegReq | Ingenuity 64 |  | WIP18 | June 21, 2019 7:16:32 PM GMT+08:00
CTUDT00624216 | Refresh | Closed | swamypreethi | June 3, 2019 4:58:00 PM GMT+08:00 | IOCC_Refresh: IHE_Verification :- C_FIND and C_STORE attribute values not same | WIP15 | WIP18 | There is no workaround. | Subsystem | DICOM | Console-SRS.12924 | Fixed | jayakumar_p | 0 | 4 | Console Software | Console4 | NO | Closed | WIP19SegReq | IQon1 |  | WIP18 | June 21, 2019 7:16:32 PM GMT+08:00
CTUDT00624211 | Refresh | Closed | Valeria_Kashtanov | June 3, 2019 4:16:17 PM GMT+08:00 | Refresh: Calibrations: Beam and DMS Check: Message about detected contrast is not logged in Service User | WIP16 | WIP18 |  | Subsystem | Calibrations | N/A | Fixed | Kostya_Baklan | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP18SegReq | Ovation |  | WIP18 | June 10, 2019 3:05:19 PM GMT+08:00
CTUDT00624211 | Refresh | Closed | Valeria_Kashtanov | June 3, 2019 4:16:17 PM GMT+08:00 | Refresh: Calibrations: Beam and DMS Check: Message about detected contrast is not logged in Service User | WIP16 | WIP18 |  | Subsystem | Calibrations | N/A | Fixed | Kostya_Baklan | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP18SegReq | Brilliance Big Bore |  | WIP18 | June 10, 2019 3:05:19 PM GMT+08:00
CTUDT00624211 | Refresh | Closed | Valeria_Kashtanov | June 3, 2019 4:16:17 PM GMT+08:00 | Refresh: Calibrations: Beam and DMS Check: Message about detected contrast is not logged in Service User | WIP16 | WIP18 |  | Subsystem | Calibrations | N/A | Fixed | Kostya_Baklan | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP18SegReq | Ingenuity 64 |  | WIP18 | June 10, 2019 3:05:19 PM GMT+08:00
CTUDT00624211 | Refresh | Closed | Valeria_Kashtanov | June 3, 2019 4:16:17 PM GMT+08:00 | Refresh: Calibrations: Beam and DMS Check: Message about detected contrast is not logged in Service User | WIP16 | WIP18 |  | Subsystem | Calibrations | N/A | Fixed | Kostya_Baklan | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP18SegReq | IQon1 |  | WIP18 | June 10, 2019 3:05:19 PM GMT+08:00
CTUDT00624209 | Refresh | Closed | wei_f | June 3, 2019 3:47:01 PM GMT+08:00 | Insert the burnt CD/DVD disc on Win10 system, Open explorer, navigate to the drive and run the application "RunDlg.exe",select the "View Images" option pop up error | BB_WIP6 |  | Please open the CD\DVD in non-scanner machines, and to open in Scanner, launch workflow, and navigate to CD\DVD to view the images. | Subsystem | CD-DVD operation | NA | Works As Specified / Not A Defect | s_sa | 0 |  |  | iCAP | NO | Closed |  | Ovation
CTUDT00624209 | Refresh | Closed | wei_f | June 3, 2019 3:47:01 PM GMT+08:00 | Insert the burnt CD/DVD disc on Win10 system, Open explorer, navigate to the drive and run the application "RunDlg.exe",select the "View Images" option pop up error | BB_WIP6 |  | Please open the CD\DVD in non-scanner machines, and to open in Scanner, launch workflow, and navigate to CD\DVD to view the images. | Subsystem | CD-DVD operation | NA | Works As Specified / Not A Defect | s_sa | 0 |  |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00624209 | Refresh | Closed | wei_f | June 3, 2019 3:47:01 PM GMT+08:00 | Insert the burnt CD/DVD disc on Win10 system, Open explorer, navigate to the drive and run the application "RunDlg.exe",select the "View Images" option pop up error | BB_WIP6 |  | Please open the CD\DVD in non-scanner machines, and to open in Scanner, launch workflow, and navigate to CD\DVD to view the images. | Subsystem | CD-DVD operation | NA | Works As Specified / Not A Defect | s_sa | 0 |  |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00624209 | Refresh | Closed | wei_f | June 3, 2019 3:47:01 PM GMT+08:00 | Insert the burnt CD/DVD disc on Win10 system, Open explorer, navigate to the drive and run the application "RunDlg.exe",select the "View Images" option pop up error | BB_WIP6 |  | Please open the CD\DVD in non-scanner machines, and to open in Scanner, launch workflow, and navigate to CD\DVD to view the images. | Subsystem | CD-DVD operation | NA | Works As Specified / Not A Defect | s_sa | 0 |  |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00624209 | Refresh | Closed | wei_f | June 3, 2019 3:47:01 PM GMT+08:00 | Insert the burnt CD/DVD disc on Win10 system, Open explorer, navigate to the drive and run the application "RunDlg.exe",select the "View Images" option pop up error | BB_WIP6 |  | Please open the CD\DVD in non-scanner machines, and to open in Scanner, launch workflow, and navigate to CD\DVD to view the images. | Subsystem | CD-DVD operation | NA | Works As Specified / Not A Defect | s_sa | 0 |  |  | iCAP | NO | Closed |  | IQon1
CTUDT00624206 | Refresh | Closed | wei_f | June 3, 2019 12:03:16 PM GMT+08:00 | In patient page switch to CDR folder,Record the study to DVD-RAM successful. | WIP16 |  |  | Subsystem | Test Design | NA | Works As Specified / Not A Defect | wei_f | 0 | 4 | Console Interoperability | Support-System | NO | Closed | WIP18SegReq | Ovation |  |  | July 3, 2019 10:40:22 AM GMT+08:00
CTUDT00624206 | Refresh | Closed | wei_f | June 3, 2019 12:03:16 PM GMT+08:00 | In patient page switch to CDR folder,Record the study to DVD-RAM successful. | WIP16 |  |  | Subsystem | Test Design | NA | Works As Specified / Not A Defect | wei_f | 0 | 4 | Console Interoperability | Support-System | NO | Closed | WIP18SegReq | Brilliance 64 |  |  | July 3, 2019 10:40:22 AM GMT+08:00
CTUDT00624206 | Refresh | Closed | wei_f | June 3, 2019 12:03:16 PM GMT+08:00 | In patient page switch to CDR folder,Record the study to DVD-RAM successful. | WIP16 |  |  | Subsystem | Test Design | NA | Works As Specified / Not A Defect | wei_f | 0 | 4 | Console Interoperability | Support-System | NO | Closed | WIP18SegReq | Brilliance Big Bore |  |  | July 3, 2019 10:40:22 AM GMT+08:00
CTUDT00624206 | Refresh | Closed | wei_f | June 3, 2019 12:03:16 PM GMT+08:00 | In patient page switch to CDR folder,Record the study to DVD-RAM successful. | WIP16 |  |  | Subsystem | Test Design | NA | Works As Specified / Not A Defect | wei_f | 0 | 4 | Console Interoperability | Support-System | NO | Closed | WIP18SegReq | Ingenuity 64 |  |  | July 3, 2019 10:40:22 AM GMT+08:00
CTUDT00624206 | Refresh | Closed | wei_f | June 3, 2019 12:03:16 PM GMT+08:00 | In patient page switch to CDR folder,Record the study to DVD-RAM successful. | WIP16 |  |  | Subsystem | Test Design | NA | Works As Specified / Not A Defect | wei_f | 0 | 4 | Console Interoperability | Support-System | NO | Closed | WIP18SegReq | IQon1 |  |  | July 3, 2019 10:40:22 AM GMT+08:00
CTUDT00624204 | Refresh | Closed | gT_s | June 3, 2019 11:39:55 AM GMT+08:00 | Refresh : The configuration test failed while performing the remote connection between system and Iwatt simulator | WIP17 | WIP19 | Please refer to the attachment section for Workaround. Also for testing with iWatt, please ensure that the machine having iWatt is having Firewall disabled. | Subsystem | Remote Configuration;Blocker19 | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP19SubCom | Ovation |  | WIP19 | August 7, 2019 5:33:41 PM GMT+08:00
CTUDT00624204 | Refresh | Closed | gT_s | June 3, 2019 11:39:55 AM GMT+08:00 | Refresh : The configuration test failed while performing the remote connection between system and Iwatt simulator | WIP17 | WIP19 | Please refer to the attachment section for Workaround. Also for testing with iWatt, please ensure that the machine having iWatt is having Firewall disabled. | Subsystem | Remote Configuration;Blocker19 | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP19SubCom | Brilliance 64 |  | WIP19 | August 7, 2019 5:33:41 PM GMT+08:00
CTUDT00624204 | Refresh | Closed | gT_s | June 3, 2019 11:39:55 AM GMT+08:00 | Refresh : The configuration test failed while performing the remote connection between system and Iwatt simulator | WIP17 | WIP19 | Please refer to the attachment section for Workaround. Also for testing with iWatt, please ensure that the machine having iWatt is having Firewall disabled. | Subsystem | Remote Configuration;Blocker19 | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP19SubCom | Brilliance Big Bore |  | WIP19 | August 7, 2019 5:33:41 PM GMT+08:00
CTUDT00624204 | Refresh | Closed | gT_s | June 3, 2019 11:39:55 AM GMT+08:00 | Refresh : The configuration test failed while performing the remote connection between system and Iwatt simulator | WIP17 | WIP19 | Please refer to the attachment section for Workaround. Also for testing with iWatt, please ensure that the machine having iWatt is having Firewall disabled. | Subsystem | Remote Configuration;Blocker19 | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP19SubCom | Ingenuity 64 |  | WIP19 | August 7, 2019 5:33:41 PM GMT+08:00
CTUDT00624204 | Refresh | Closed | gT_s | June 3, 2019 11:39:55 AM GMT+08:00 | Refresh : The configuration test failed while performing the remote connection between system and Iwatt simulator | WIP17 | WIP19 | Please refer to the attachment section for Workaround. Also for testing with iWatt, please ensure that the machine having iWatt is having Firewall disabled. | Subsystem | Remote Configuration;Blocker19 | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP19SubCom | IQon1 |  | WIP19 | August 7, 2019 5:33:41 PM GMT+08:00
CTUDT00624199 | Refresh | Closed | Noy_Seelig | June 2, 2019 11:40:02 PM GMT+08:00 | Refresh: ACR: Pediatric Abdomen cases fail due to calculation errors and no ROIs drawn on images | WIP16 |  | N/A | Subsystem | Calibrations | N/A | System Repaired | yosfan_s | 1 |  |  | DMS | NO | Closed |  | IQon1
CTUDT00624196 | Refresh | Closed | Valeria_Kashtanov | June 2, 2019 9:52:46 PM GMT+08:00 | Refresh: Calibrations: Performance Test: Recon Shift Y in Trauma case is not according to spec | WIP16 | WIP19 |  | Subsystem | Performance Test | N/A | Fixed | braunstein_d | 0 | 2 | Design Documentation | Support-Physics | YES | Closed | WIP21SubCom | Ovation |  | WIP19 | October 27, 2019 8:33:23 PM GMT+08:00
CTUDT00624193 | Refresh | Closed | beck_d | June 2, 2019 8:43:30 PM GMT+08:00 | Refresh: Performance Test: Harmonized: Table height in Body CT12b case is not according to spec | WIP16 | WIP18 | N/A | Subsystem | Performance Test; Verified1 | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP18SubCom | Ovation |  | WIP18 | June 12, 2019 3:06:27 PM GMT+08:00
CTUDT00624193 | Refresh | Closed | beck_d | June 2, 2019 8:43:30 PM GMT+08:00 | Refresh: Performance Test: Harmonized: Table height in Body CT12b case is not according to spec | WIP16 | WIP18 | N/A | Subsystem | Performance Test; Verified1 | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP18SubCom | Brilliance Big Bore |  | WIP18 | June 12, 2019 3:06:27 PM GMT+08:00
CTUDT00624187 | Refresh | Closed | beck_d | June 2, 2019 3:29:03 PM GMT+08:00 | Refresh: Hcor Slice: Harmonized: Hcor Slice Child tables for iCT SP are missing | WIP16 |  |  | Subsystem | Calibrations |  | Duplicate | yatziv_l | 0 |  |  | DMS | NO | Closed |  | Ovation
CTUDT00624184 | Refresh | Closed | zhu_je | May 31, 2019 9:40:00 PM GMT+08:00 | [BB Refresh]The "Add injection" and "Remove Injection" are displayed incorrect for scan ruler bubble right click | BB_WIP6 | BB_WIP8 | NA | Subsystem | Plan | NA | Works As Specified / Not A Defect | jayakumar_p | 0 | 3 | Console Software | Console4 | NO | Closed | WAS | Brilliance Big Bore |  | BB_WIP8
CTUDT00624182 | Refresh | Closed | zhu_je | May 31, 2019 9:20:48 PM GMT+08:00 | [BB Refresh]The max Increment value was not incorrect | BB_WIP6 |  |  | Subsystem | Plan | NA | Works As Specified / Not A Defect | v_sumanth | 0 |  |  | Console4 | NO | Closed | WAS | Brilliance Big Bore
CTUDT00624180 | Refresh | Closed | zhu_je | May 31, 2019 9:13:22 PM GMT+08:00 | [BB Refresh]Scan error if Diagnostic Level other than 0 in preference setting | BB_WIP6 |  | If the the diagnostic level kept as 0 and the Scan with X-Ray check-box is checked then this issue can be avoided . | Subsystem | System | N/A | Works As Specified / Not A Defect | mistry_h | 0 |  |  | Support-System | NO | Closed | BB_WIP8_SubCom | Ovation
CTUDT00624180 | Refresh | Closed | zhu_je | May 31, 2019 9:13:22 PM GMT+08:00 | [BB Refresh]Scan error if Diagnostic Level other than 0 in preference setting | BB_WIP6 |  | If the the diagnostic level kept as 0 and the Scan with X-Ray check-box is checked then this issue can be avoided . | Subsystem | System | N/A | Works As Specified / Not A Defect | mistry_h | 0 |  |  | Support-System | NO | Closed | BB_WIP8_SubCom | Brilliance 64
CTUDT00624180 | Refresh | Closed | zhu_je | May 31, 2019 9:13:22 PM GMT+08:00 | [BB Refresh]Scan error if Diagnostic Level other than 0 in preference setting | BB_WIP6 |  | If the the diagnostic level kept as 0 and the Scan with X-Ray check-box is checked then this issue can be avoided . | Subsystem | System | N/A | Works As Specified / Not A Defect | mistry_h | 0 |  |  | Support-System | NO | Closed | BB_WIP8_SubCom | Brilliance Big Bore
CTUDT00624180 | Refresh | Closed | zhu_je | May 31, 2019 9:13:22 PM GMT+08:00 | [BB Refresh]Scan error if Diagnostic Level other than 0 in preference setting | BB_WIP6 |  | If the the diagnostic level kept as 0 and the Scan with X-Ray check-box is checked then this issue can be avoided . | Subsystem | System | N/A | Works As Specified / Not A Defect | mistry_h | 0 |  |  | Support-System | NO | Closed | BB_WIP8_SubCom | Ingenuity 64
CTUDT00624180 | Refresh | Closed | zhu_je | May 31, 2019 9:13:22 PM GMT+08:00 | [BB Refresh]Scan error if Diagnostic Level other than 0 in preference setting | BB_WIP6 |  | If the the diagnostic level kept as 0 and the Scan with X-Ray check-box is checked then this issue can be avoided . | Subsystem | System | N/A | Works As Specified / Not A Defect | mistry_h | 0 |  |  | Support-System | NO | Closed | BB_WIP8_SubCom | IQon1
CTUDT00624176 | Refresh | Closed | Sharmila_Vasireddy | May 31, 2019 3:39:17 PM GMT+08:00 | Blank desktop was displayed once login Emergency with valid credentials | BB_WIP5.1 | WIP18 BB_WIP8 |  | Subsystem |  | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP18SubCom | Ovation |  | WIP18 BB_WIP8 | June 14, 2019 8:52:15 PM GMT+08:00
CTUDT00624176 | Refresh | Closed | Sharmila_Vasireddy | May 31, 2019 3:39:17 PM GMT+08:00 | Blank desktop was displayed once login Emergency with valid credentials | BB_WIP5.1 | WIP18 BB_WIP8 |  | Subsystem |  | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP18SubCom | Brilliance 64 |  | WIP18 BB_WIP8 | June 14, 2019 8:52:15 PM GMT+08:00
CTUDT00624176 | Refresh | Closed | Sharmila_Vasireddy | May 31, 2019 3:39:17 PM GMT+08:00 | Blank desktop was displayed once login Emergency with valid credentials | BB_WIP5.1 | WIP18 BB_WIP8 |  | Subsystem |  | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP18SubCom | Brilliance Big Bore |  | WIP18 BB_WIP8 | June 14, 2019 8:52:15 PM GMT+08:00
CTUDT00624176 | Refresh | Closed | Sharmila_Vasireddy | May 31, 2019 3:39:17 PM GMT+08:00 | Blank desktop was displayed once login Emergency with valid credentials | BB_WIP5.1 | WIP18 BB_WIP8 |  | Subsystem |  | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP18SubCom | Ingenuity 64 |  | WIP18 BB_WIP8 | June 14, 2019 8:52:15 PM GMT+08:00
CTUDT00624176 | Refresh | Closed | Sharmila_Vasireddy | May 31, 2019 3:39:17 PM GMT+08:00 | Blank desktop was displayed once login Emergency with valid credentials | BB_WIP5.1 | WIP18 BB_WIP8 |  | Subsystem |  | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP18SubCom | IQon1 |  | WIP18 BB_WIP8 | June 14, 2019 8:52:15 PM GMT+08:00
CTUDT00624175 | Refresh | Closed | wang_l | May 31, 2019 3:32:24 PM GMT+08:00 | Refresh: "File is corrupted" is displayed once open imported logs from Log Viewer with format .zip and .log | WIP16 |  |  | Subsystem | Log Viewer | NA | Works As Specified / Not A Defect | s_sa | 0 |  |  | iCAP | NO | Closed |  | Ovation
CTUDT00624175 | Refresh | Closed | wang_l | May 31, 2019 3:32:24 PM GMT+08:00 | Refresh: "File is corrupted" is displayed once open imported logs from Log Viewer with format .zip and .log | WIP16 |  |  | Subsystem | Log Viewer | NA | Works As Specified / Not A Defect | s_sa | 0 |  |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00624175 | Refresh | Closed | wang_l | May 31, 2019 3:32:24 PM GMT+08:00 | Refresh: "File is corrupted" is displayed once open imported logs from Log Viewer with format .zip and .log | WIP16 |  |  | Subsystem | Log Viewer | NA | Works As Specified / Not A Defect | s_sa | 0 |  |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00624175 | Refresh | Closed | wang_l | May 31, 2019 3:32:24 PM GMT+08:00 | Refresh: "File is corrupted" is displayed once open imported logs from Log Viewer with format .zip and .log | WIP16 |  |  | Subsystem | Log Viewer | NA | Works As Specified / Not A Defect | s_sa | 0 |  |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00624175 | Refresh | Closed | wang_l | May 31, 2019 3:32:24 PM GMT+08:00 | Refresh: "File is corrupted" is displayed once open imported logs from Log Viewer with format .zip and .log | WIP16 |  |  | Subsystem | Log Viewer | NA | Works As Specified / Not A Defect | s_sa | 0 |  |  | iCAP | NO | Closed |  | IQon1
CTUDT00624173 | Refresh | Closed | wang_l | May 31, 2019 3:10:52 PM GMT+08:00 | Refresh: Access is denied to collect bug report by command line for BR64 | WIP16 |  |  | Subsystem | Bugreport |  | Invalid/Uncontrolled Configuration | wang_l | 0 |  |  | Support-System | NO | Closed |  | Brilliance 64
CTUDT00624171 | Refresh | Closed | zhu_je | May 31, 2019 1:51:01 PM GMT+08:00 | [BB Refresh]The voice is not played before scan if the releated voice was added slience and then remove | BB_WIP6 | BB_WIP8 |  | Subsystem | Voice Manager | Console-SRS.2882,Console-SRS.2884 | Fixed | sheik_c | 0 | 2 | Console Software | Console4 | YES | Closed | BB_WIP8_SubCom | Brilliance Big Bore |  | BB_WIP8 | June 7, 2019 3:30:08 PM GMT+08:00
CTUDT00624167 | Refresh | Closed | bommiReddy_r | May 30, 2019 8:29:53 PM GMT+08:00 | Refresh: UpgradeUtility: Installation of IQON Refresh 4.7.7 on IQON HB 4.7.5 host computer is failing | WIP16 | WIP20 |  | Subsystem | Installation | N/A | Fixed | hilas | 0 | 2 | Service Documentation | Support-System | NO | Closed | WIP19SubCom | IQon1 |  | WIP20 | September 26, 2019 3:20:20 PM GMT+08:00
CTUDT00624158 | Refresh | Closed | hatti_c | May 30, 2019 6:43:11 PM GMT+08:00 | IST -ICE Client need to be upgraded to V3.3 to address security Vulnerability of ICEv3.2 | WIP16 | WIP18 BB_WIP8 |  | Subsystem | Blocker | N/A | Fixed | s_shabeer | 0 | 2 | Service Tools_Software | Service Tools | YES | Closed | WIP18SubCom | Ovation |  | WIP18 BB_WIP8 | June 26, 2019 8:48:58 PM GMT+08:00
CTUDT00624158 | Refresh | Closed | hatti_c | May 30, 2019 6:43:11 PM GMT+08:00 | IST -ICE Client need to be upgraded to V3.3 to address security Vulnerability of ICEv3.2 | WIP16 | WIP18 BB_WIP8 |  | Subsystem | Blocker | N/A | Fixed | s_shabeer | 0 | 2 | Service Tools_Software | Service Tools | YES | Closed | WIP18SubCom | Brilliance 64 |  | WIP18 BB_WIP8 | June 26, 2019 8:48:58 PM GMT+08:00
CTUDT00624158 | Refresh | Closed | hatti_c | May 30, 2019 6:43:11 PM GMT+08:00 | IST -ICE Client need to be upgraded to V3.3 to address security Vulnerability of ICEv3.2 | WIP16 | WIP18 BB_WIP8 |  | Subsystem | Blocker | N/A | Fixed | s_shabeer | 0 | 2 | Service Tools_Software | Service Tools | YES | Closed | WIP18SubCom | Brilliance Big Bore |  | WIP18 BB_WIP8 | June 26, 2019 8:48:58 PM GMT+08:00
CTUDT00624158 | Refresh | Closed | hatti_c | May 30, 2019 6:43:11 PM GMT+08:00 | IST -ICE Client need to be upgraded to V3.3 to address security Vulnerability of ICEv3.2 | WIP16 | WIP18 BB_WIP8 |  | Subsystem | Blocker | N/A | Fixed | s_shabeer | 0 | 2 | Service Tools_Software | Service Tools | YES | Closed | WIP18SubCom | Ingenuity 64 |  | WIP18 BB_WIP8 | June 26, 2019 8:48:58 PM GMT+08:00
CTUDT00624158 | Refresh | Closed | hatti_c | May 30, 2019 6:43:11 PM GMT+08:00 | IST -ICE Client need to be upgraded to V3.3 to address security Vulnerability of ICEv3.2 | WIP16 | WIP18 BB_WIP8 |  | Subsystem | Blocker | N/A | Fixed | s_shabeer | 0 | 2 | Service Tools_Software | Service Tools | YES | Closed | WIP18SubCom | IQon1 |  | WIP18 BB_WIP8 | June 26, 2019 8:48:58 PM GMT+08:00
CTUDT00624156 | Refresh | Closed | gT_s | May 30, 2019 5:05:34 PM GMT+08:00 | Refresh : The application displays the system error while changing the parameters of the different spectral results. | WIP17 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | babu_s | 1 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00624152 | Refresh | Closed | shekhar_p | May 30, 2019 2:47:30 PM GMT+08:00 | Refresh: System crashes after the surview scan | WIP16 |  |  | Subsystem | System | NA | Duplicate | v_sumanth | 0 |  |  | Console4 | YES | Closed |  | Ingenuity 64
CTUDT00624150 | Refresh | Closed | lei_yang | May 30, 2019 1:41:09 PM GMT+08:00 | The distance from R-tag to the middle of 40% phase trigger window is not correct. | BB_WIP6 |  | No | Subsystem | Cardiac | NO | Works As Specified / Not A Defect | lei_yang | 0 | 2 | Design Documentation | Support-System | NO | Closed | BB_WIP8_SubCom | Brilliance Big Bore |  |  | July 4, 2019 4:30:22 PM GMT+08:00
CTUDT00624150 | Refresh | Closed | lei_yang | May 30, 2019 1:41:09 PM GMT+08:00 | The distance from R-tag to the middle of 40% phase trigger window is not correct. | BB_WIP6 |  | No | Subsystem | Cardiac | NO | Works As Specified / Not A Defect | lei_yang | 0 | 2 | Design Documentation | Support-System | NO | Closed | BB_WIP8_SubCom | Ingenuity 64 |  |  | July 4, 2019 4:30:22 PM GMT+08:00
CTUDT00624150 | Refresh | Closed | lei_yang | May 30, 2019 1:41:09 PM GMT+08:00 | The distance from R-tag to the middle of 40% phase trigger window is not correct. | BB_WIP6 |  | No | Subsystem | Cardiac | NO | Works As Specified / Not A Defect | lei_yang | 0 | 2 | Design Documentation | Support-System | NO | Closed | BB_WIP8_SubCom | Ingenuity 16 |  |  | July 4, 2019 4:30:22 PM GMT+08:00
CTUDT00624150 | Refresh | Closed | lei_yang | May 30, 2019 1:41:09 PM GMT+08:00 | The distance from R-tag to the middle of 40% phase trigger window is not correct. | BB_WIP6 |  | No | Subsystem | Cardiac | NO | Works As Specified / Not A Defect | lei_yang | 0 | 2 | Design Documentation | Support-System | NO | Closed | BB_WIP8_SubCom | Ingenuity BigBore |  |  | July 4, 2019 4:30:22 PM GMT+08:00
CTUDT00624146 | Refresh | Closed | g_di | May 30, 2019 12:56:01 PM GMT+08:00 | Few parameters are not localized for any of the languages in Print and text tab. | WIP16 | WIP18 |  | System | Localization | Console-SRS.784 | Fixed | sheik_c | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP8_SubCom & WIP18_SubCom | Ovation |  | WIP18 | June 10, 2019 4:47:58 PM GMT+08:00
CTUDT00624146 | Refresh | Closed | g_di | May 30, 2019 12:56:01 PM GMT+08:00 | Few parameters are not localized for any of the languages in Print and text tab. | WIP16 | WIP18 |  | System | Localization | Console-SRS.784 | Fixed | sheik_c | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP8_SubCom & WIP18_SubCom | Brilliance 64 |  | WIP18 | June 10, 2019 4:47:58 PM GMT+08:00
CTUDT00624146 | Refresh | Closed | g_di | May 30, 2019 12:56:01 PM GMT+08:00 | Few parameters are not localized for any of the languages in Print and text tab. | WIP16 | WIP18 |  | System | Localization | Console-SRS.784 | Fixed | sheik_c | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP8_SubCom & WIP18_SubCom | Ingenuity 64 |  | WIP18 | June 10, 2019 4:47:58 PM GMT+08:00
CTUDT00624146 | Refresh | Closed | g_di | May 30, 2019 12:56:01 PM GMT+08:00 | Few parameters are not localized for any of the languages in Print and text tab. | WIP16 | WIP18 |  | System | Localization | Console-SRS.784 | Fixed | sheik_c | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP8_SubCom & WIP18_SubCom | IQon1 |  | WIP18 | June 10, 2019 4:47:58 PM GMT+08:00
CTUDT00624144 | Refresh | Closed | wang_l | May 30, 2019 10:20:12 AM GMT+08:00 | Refresh BR64: The Asset Description for Brilliance 64 display as CT_iCT_4.1 | WIP16 | WIP18 |  | Subsystem | Remote Configuration | N/A | Fixed | ranjithkumar_m | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP18SubCom | Brilliance 64 |  | WIP18 | June 12, 2019 6:56:47 PM GMT+08:00
CTUDT00624142 | Refresh | Closed | cohen_y | May 30, 2019 2:33:22 AM GMT+08:00 | Refresh: FEE Test: mAs Linearity: Half Tile limits are not as defined in spec | WIP16 | WIP18 |  | Subsystem | FEE Test | N/A | Fixed | Kostya_Baklan | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP18SubCom | Ovation |  | WIP18 | June 13, 2019 3:45:15 PM GMT+08:00
CTUDT00624140 | Refresh | Closed | cohen_y | May 30, 2019 12:24:01 AM GMT+08:00 | Refresh: FEE Test: iCT SP: Collimation of Reference detector test is 64x0.625 and not as defined in Spec which is 128x0.625 | WIP16 | N/A |  | Subsystem | Calibrations | N/A | Fixed | berent_i | 0 | 2 | Design Documentation | DMS | NO | Closed | WIP18SubCom | Ovation |  | N/A | June 19, 2019 8:12:20 PM GMT+08:00
CTUDT00624137 | Refresh | Closed | gT_s | May 29, 2019 9:00:27 PM GMT+08:00 | Refresh : The spectral label is getting overlapped with image in spectral CT viewer | WIP17 |  |  | Subsystem | Spectral CT Viewer | NA | Works As Specified / Not A Defect | yanovich_z | 0 |  |  | Support_Spectral | NO | Closed |  | IQon1
CTUDT00624135 | Refresh | Closed | hiremath_b | May 29, 2019 8:56:53 PM GMT+08:00 | Refresh-IQon: User is able to select spectral result type when kV is set to 80 OR 100. | WIP16 | WIP18 | NA | Subsystem | System | Console-SRS.11666 | Fixed | Prakash_Chand_Yadav | 0 | 2 | Console Software | Console4 | YES | Closed | WIP20SegReq | IQon1 |  | WIP18 | June 4, 2019 2:52:38 PM GMT+08:00
CTUDT00624134 | Refresh | Closed | Valeria_Kashtanov | May 29, 2019 8:55:58 PM GMT+08:00 | Refresh: Calibrations: UHR Comb Align: Eye Phantom pin is not detected in 4th scan | WIP16 | WIP18 |  | Subsystem | Calibrations | N/A | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP18SubCom | Ovation |  | WIP18 | June 23, 2019 4:35:19 PM GMT+08:00
CTUDT00624130 | Refresh | Closed | Valeria_Kashtanov | May 29, 2019 6:19:44 PM GMT+08:00 | Refresh: Calibrations: ACR Test: Pediatric Head scan fails due to error | WIP16 | WIP18 | resume and run the 4th scan as stand alone scan. | Subsystem |  | n/a | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP19SegReq | Ovation |  | WIP18 | June 24, 2019 11:08:19 PM GMT+08:00
CTUDT00624125 | Refresh | Closed | sowmya_k | May 29, 2019 5:50:41 PM GMT+08:00 | In Big Bore refresh : As per config file “MassCalcuimScoringApplicationMapping.cfg”, the values are not displaying in Dicom tag [0018, 9352] | BB_WIP6 | WIP18 | N/A | Subsystem |  | N/A | Fixed | naveen_k | 0 | 2 | Console Software | Console4 | NO | Closed | WIP18SegReq | Ovation |  | WIP18 | May 30, 2019 6:19:42 PM GMT+08:00
CTUDT00624125 | Refresh | Closed | sowmya_k | May 29, 2019 5:50:41 PM GMT+08:00 | In Big Bore refresh : As per config file “MassCalcuimScoringApplicationMapping.cfg”, the values are not displaying in Dicom tag [0018, 9352] | BB_WIP6 | WIP18 | N/A | Subsystem |  | N/A | Fixed | naveen_k | 0 | 2 | Console Software | Console4 | NO | Closed | WIP18SegReq | Brilliance 64 |  | WIP18 | May 30, 2019 6:19:42 PM GMT+08:00
CTUDT00624125 | Refresh | Closed | sowmya_k | May 29, 2019 5:50:41 PM GMT+08:00 | In Big Bore refresh : As per config file “MassCalcuimScoringApplicationMapping.cfg”, the values are not displaying in Dicom tag [0018, 9352] | BB_WIP6 | WIP18 | N/A | Subsystem |  | N/A | Fixed | naveen_k | 0 | 2 | Console Software | Console4 | NO | Closed | WIP18SegReq | Ingenuity 64 |  | WIP18 | May 30, 2019 6:19:42 PM GMT+08:00
CTUDT00624125 | Refresh | Closed | sowmya_k | May 29, 2019 5:50:41 PM GMT+08:00 | In Big Bore refresh : As per config file “MassCalcuimScoringApplicationMapping.cfg”, the values are not displaying in Dicom tag [0018, 9352] | BB_WIP6 | WIP18 | N/A | Subsystem |  | N/A | Fixed | naveen_k | 0 | 2 | Console Software | Console4 | NO | Closed | WIP18SegReq | IQon1 |  | WIP18 | May 30, 2019 6:19:42 PM GMT+08:00
CTUDT00624123 | Refresh | Closed | zhu_je | May 29, 2019 5:00:41 PM GMT+08:00 | [BB Refresh]"MRP/MINP" is displayed on scan result images if render mode=MINIP and orientation=axial | BB_WIP6 |  | NA | Subsystem |  | Console.SRS.12800 | No Plan to Fix | saha_d | 0 | 4 |  | Console4 | NO | Closed | SubNC | Brilliance Big Bore
CTUDT00624121 | Refresh | Closed | zhu_je | May 29, 2019 4:33:03 PM GMT+08:00 | [BB Refresh]Cannot set width vaule to 0 on preview images before do final recon | BB_WIP6 |  |  | Subsystem | Reconstruction | NA | Does Not Apply | prasad_r | 0 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00624119 | Refresh | Closed | zhu_je | May 29, 2019 4:24:05 PM GMT+08:00 | [BB Refresh][Voice Language] Voice language cannot be selected if clicking dropdown list button | BB_WIP6 | BB_WIP8 |  | Subsystem | Patient Scene | Console-SRS.15988 & Console-SRS.15386 | Fixed | Snehashis_Dey | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP8_SubCom | Brilliance Big Bore |  | BB_WIP8 | June 5, 2019 3:49:13 PM GMT+08:00
CTUDT00624116 | Refresh | Closed | finestine_m | May 29, 2019 3:08:15 PM GMT+08:00 | Refresh: Common Functions: Missing Asics Temperatures In The Results For IQ Tests | WIP16 | WIP18 |  | Subsystem | Calibrations | NA | Fixed | Kostya_Baklan | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP18SegReq | Ovation |  | WIP18 | June 6, 2019 3:46:09 PM GMT+08:00
CTUDT00624116 | Refresh | Closed | finestine_m | May 29, 2019 3:08:15 PM GMT+08:00 | Refresh: Common Functions: Missing Asics Temperatures In The Results For IQ Tests | WIP16 | WIP18 |  | Subsystem | Calibrations | NA | Fixed | Kostya_Baklan | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP18SegReq | IQon1 |  | WIP18 | June 6, 2019 3:46:09 PM GMT+08:00
CTUDT00624114 | Refresh | Closed | lei_yang | May 29, 2019 11:09:20 AM GMT+08:00 | The actual Cycle Time in the Log Viewer rounded to a tenth of an integer isn't equal to the plan Cycle Time value. | WIP16 |  |  | System | Log Viewer | NA | Works As Specified / Not A Defect | lei_yang | 0 | 2 | Design Documentation | GCX | NO | Closed | WIP18SubCom | Ingenuity 64 |  |  | June 26, 2019 6:33:43 PM GMT+08:00
CTUDT00624111 | Refresh | Closed | Valeria_Kashtanov | May 29, 2019 1:55:30 AM GMT+08:00 | Refresh: Calibrations: Performance Infant: iCT SP: Performance Infant body cases fail due to error | WIP16 |  |  | Subsystem | Calibrations |  | Duplicate | yatziv_l | 0 |  |  | DMS | NO | Closed |  | Ovation
CTUDT00624109 | Refresh | Closed | Valeria_Kashtanov | May 28, 2019 10:02:50 PM GMT+08:00 | Refresh: Calibrations: Acceptance and Constancy: iCT SP: No Acceptance Calibration in Quality Assurance menu when Harmonized Phantom is checked | WIP16 |  |  | Subsystem | Acceptance Test |  | Duplicate | yatziv_l | 0 |  |  | DMS | NO | Closed |  | Ovation
CTUDT00624107 | Refresh | Closed | Noy_Seelig | May 28, 2019 9:10:01 PM GMT+08:00 | Refresh: IQ Check fails with error at 25% in all languages but English | WIP16 |  | NA | Subsystem | IQ Check | Console.SRS.16004 | Duplicate | saha_d | 0 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00624107 | Refresh | Closed | Noy_Seelig | May 28, 2019 9:10:01 PM GMT+08:00 | Refresh: IQ Check fails with error at 25% in all languages but English | WIP16 |  | NA | Subsystem | IQ Check | Console.SRS.16004 | Duplicate | saha_d | 0 |  |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00624107 | Refresh | Closed | Noy_Seelig | May 28, 2019 9:10:01 PM GMT+08:00 | Refresh: IQ Check fails with error at 25% in all languages but English | WIP16 |  | NA | Subsystem | IQ Check | Console.SRS.16004 | Duplicate | saha_d | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00624107 | Refresh | Closed | Noy_Seelig | May 28, 2019 9:10:01 PM GMT+08:00 | Refresh: IQ Check fails with error at 25% in all languages but English | WIP16 |  | NA | Subsystem | IQ Check | Console.SRS.16004 | Duplicate | saha_d | 0 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00624105 | Refresh | Closed | gT_s | May 28, 2019 6:58:51 PM GMT+08:00 | Refresh : Clinical user was able to login after the restart the system when installation failure has happened for invalid automatic software package | WIP 16 | WIP17 BB_WIP7 |  | Subsystem | Software Distribution | STPS-SSRS-1263 | Fixed | sagar_v | 0 | 2 | Console Software | Service Tools | YES | Closed | WIP17SubCom | Ovation |  | WIP17 BB_WIP7 | June 6, 2019 1:58:30 PM GMT+08:00
CTUDT00624105 | Refresh | Closed | gT_s | May 28, 2019 6:58:51 PM GMT+08:00 | Refresh : Clinical user was able to login after the restart the system when installation failure has happened for invalid automatic software package | WIP 16 | WIP17 BB_WIP7 |  | Subsystem | Software Distribution | STPS-SSRS-1263 | Fixed | sagar_v | 0 | 2 | Console Software | Service Tools | YES | Closed | WIP17SubCom | Brilliance 64 |  | WIP17 BB_WIP7 | June 6, 2019 1:58:30 PM GMT+08:00
CTUDT00624105 | Refresh | Closed | gT_s | May 28, 2019 6:58:51 PM GMT+08:00 | Refresh : Clinical user was able to login after the restart the system when installation failure has happened for invalid automatic software package | WIP 16 | WIP17 BB_WIP7 |  | Subsystem | Software Distribution | STPS-SSRS-1263 | Fixed | sagar_v | 0 | 2 | Console Software | Service Tools | YES | Closed | WIP17SubCom | Brilliance Big Bore |  | WIP17 BB_WIP7 | June 6, 2019 1:58:30 PM GMT+08:00
CTUDT00624105 | Refresh | Closed | gT_s | May 28, 2019 6:58:51 PM GMT+08:00 | Refresh : Clinical user was able to login after the restart the system when installation failure has happened for invalid automatic software package | WIP 16 | WIP17 BB_WIP7 |  | Subsystem | Software Distribution | STPS-SSRS-1263 | Fixed | sagar_v | 0 | 2 | Console Software | Service Tools | YES | Closed | WIP17SubCom | Ingenuity 64 |  | WIP17 BB_WIP7 | June 6, 2019 1:58:30 PM GMT+08:00
CTUDT00624105 | Refresh | Closed | gT_s | May 28, 2019 6:58:51 PM GMT+08:00 | Refresh : Clinical user was able to login after the restart the system when installation failure has happened for invalid automatic software package | WIP 16 | WIP17 BB_WIP7 |  | Subsystem | Software Distribution | STPS-SSRS-1263 | Fixed | sagar_v | 0 | 2 | Console Software | Service Tools | YES | Closed | WIP17SubCom | IQon1 |  | WIP17 BB_WIP7 | June 6, 2019 1:58:30 PM GMT+08:00
CTUDT00624097 | Refresh | Closed | kapadia_v | May 28, 2019 3:07:51 PM GMT+08:00 | Clear All and few fields were not localized as per selected language. | 4.8 | BB_WIP7 | NA | System |  | Console-SRS.784 | Fixed | saha_d | 0 | 2 | Console Software | Console4 | YES | Closed | BB_WIP7_SubCom | Brilliance Big Bore |  | BB_WIP7 | June 6, 2019 6:02:08 PM GMT+08:00
CTUDT00624096 | Refresh | Closed | bandil_j | May 28, 2019 2:52:41 PM GMT+08:00 | Bottom part of Worklist window in patient page was greyed out | WIP6 | BB_WIP7 | Hover mouse over the controls to reveal UI controls (background is still blocky as shown in the attachment) | Subsystem |  | N/A | Fixed | naveen_k | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP7_SubCom | Brilliance Big Bore |  | BB_WIP7 | May 30, 2019 4:59:15 PM GMT+08:00
CTUDT00624093 | Refresh | Closed | azagury_d | May 27, 2019 9:44:25 PM GMT+08:00 | Refresh: Acceptance/Constancy Tests: Slice Thickness case 4 does not scan | WIP16 | WIP17 | N/A | Subsystem |  | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP17SubCom | Ovation |  | WIP17 | May 30, 2019 2:33:59 PM GMT+08:00
CTUDT00624088 | Refresh | Closed | zhu_je | May 27, 2019 7:53:43 PM GMT+08:00 | [BB Refresh][IVC]The couch was not motion using IV button during CCT fluoro scan | BB_WIP6 |  |  | Subsystem | IVC | NA | Works As Specified / Not A Defect | b_bibin | 0 |  |  | Console4 | NO | Closed | BB_WIP8_SubCom | Brilliance Big Bore
CTUDT00624080 | Refresh | Closed | kokate_k | May 27, 2019 6:31:26 PM GMT+08:00 | system is still initializing. please wait? Pop up on selecting the Examcard | BB_WIP6 |  |  | Subsystem |  |  | Does Not Apply | prabhu_m | 0 | 4 |  | Console4 | NO | Closed |  | Ovation | PR 9064066
CTUDT00624080 | Refresh | Closed | kokate_k | May 27, 2019 6:31:26 PM GMT+08:00 | system is still initializing. please wait? Pop up on selecting the Examcard | BB_WIP6 |  |  | Subsystem |  |  | Does Not Apply | prabhu_m | 0 | 4 |  | Console4 | NO | Closed |  | Brilliance 64 | PR 9064066
CTUDT00624080 | Refresh | Closed | kokate_k | May 27, 2019 6:31:26 PM GMT+08:00 | system is still initializing. please wait? Pop up on selecting the Examcard | BB_WIP6 |  |  | Subsystem |  |  | Does Not Apply | prabhu_m | 0 | 4 |  | Console4 | NO | Closed |  | Brilliance Big Bore | PR 9064066
CTUDT00624080 | Refresh | Closed | kokate_k | May 27, 2019 6:31:26 PM GMT+08:00 | system is still initializing. please wait? Pop up on selecting the Examcard | BB_WIP6 |  |  | Subsystem |  |  | Does Not Apply | prabhu_m | 0 | 4 |  | Console4 | NO | Closed |  | Ingenuity 64 | PR 9064066
CTUDT00624080 | Refresh | Closed | kokate_k | May 27, 2019 6:31:26 PM GMT+08:00 | system is still initializing. please wait? Pop up on selecting the Examcard | BB_WIP6 |  |  | Subsystem |  |  | Does Not Apply | prabhu_m | 0 | 4 |  | Console4 | NO | Closed |  | Ingenuity TF 64 | PR 9064066
CTUDT00624080 | Refresh | Closed | kokate_k | May 27, 2019 6:31:26 PM GMT+08:00 | system is still initializing. please wait? Pop up on selecting the Examcard | BB_WIP6 |  |  | Subsystem |  |  | Does Not Apply | prabhu_m | 0 | 4 |  | Console4 | NO | Closed |  | Veyron | PR 9064066
CTUDT00624080 | Refresh | Closed | kokate_k | May 27, 2019 6:31:26 PM GMT+08:00 | system is still initializing. please wait? Pop up on selecting the Examcard | BB_WIP6 |  |  | Subsystem |  |  | Does Not Apply | prabhu_m | 0 | 4 |  | Console4 | NO | Closed |  | IQon1 | PR 9064066
CTUDT00624079 | Refresh | Closed | kokate_k | May 27, 2019 6:30:54 PM GMT+08:00 | system is still initializing. please wait? Pop up on selecting the Examcard | WIP17 |  |  | Subsystem |  |  | Does Not Apply | prabhu_m | 0 | 4 |  | Console4 | NO | Closed |  | Ovation | PR 9064066
CTUDT00624079 | Refresh | Closed | kokate_k | May 27, 2019 6:30:54 PM GMT+08:00 | system is still initializing. please wait? Pop up on selecting the Examcard | WIP17 |  |  | Subsystem |  |  | Does Not Apply | prabhu_m | 0 | 4 |  | Console4 | NO | Closed |  | Brilliance 64 | PR 9064066
CTUDT00624079 | Refresh | Closed | kokate_k | May 27, 2019 6:30:54 PM GMT+08:00 | system is still initializing. please wait? Pop up on selecting the Examcard | WIP17 |  |  | Subsystem |  |  | Does Not Apply | prabhu_m | 0 | 4 |  | Console4 | NO | Closed |  | Brilliance Big Bore | PR 9064066
CTUDT00624079 | Refresh | Closed | kokate_k | May 27, 2019 6:30:54 PM GMT+08:00 | system is still initializing. please wait? Pop up on selecting the Examcard | WIP17 |  |  | Subsystem |  |  | Does Not Apply | prabhu_m | 0 | 4 |  | Console4 | NO | Closed |  | Ingenuity 64 | PR 9064066
CTUDT00624079 | Refresh | Closed | kokate_k | May 27, 2019 6:30:54 PM GMT+08:00 | system is still initializing. please wait? Pop up on selecting the Examcard | WIP17 |  |  | Subsystem |  |  | Does Not Apply | prabhu_m | 0 | 4 |  | Console4 | NO | Closed |  | Ingenuity TF 64 | PR 9064066
CTUDT00624079 | Refresh | Closed | kokate_k | May 27, 2019 6:30:54 PM GMT+08:00 | system is still initializing. please wait? Pop up on selecting the Examcard | WIP17 |  |  | Subsystem |  |  | Does Not Apply | prabhu_m | 0 | 4 |  | Console4 | NO | Closed |  | Veyron | PR 9064066
CTUDT00624079 | Refresh | Closed | kokate_k | May 27, 2019 6:30:54 PM GMT+08:00 | system is still initializing. please wait? Pop up on selecting the Examcard | WIP17 |  |  | Subsystem |  |  | Does Not Apply | prabhu_m | 0 | 4 |  | Console4 | NO | Closed |  | IQon1 | PR 9064066
CTUDT00624074 | Refresh | Closed | kokate_k | May 27, 2019 6:22:02 PM GMT+08:00 | Unhandled Exception after Dual Surview | BB_WIP6 |  |  | Subsystem |  | Console-SRS.3226,Console-SRS.2374 | Duplicate | Snehashis_Dey | 0 | 3 |  | Console4 | YES | Closed |  | IQon1 | PR# 8613417
CTUDT00624072 | Refresh | Closed | kokate_k | May 27, 2019 6:12:28 PM GMT+08:00 | Unhandled Exception after Dual Surview | WIP17 |  |  | Subsystem |  |  | Does Not Apply | kokate_k | 0 | 3 |  | Console4 | YES | Closed |  | IQon1 | PR# 8613417
CTUDT00624063 | Refresh | Closed | zhu_je | May 27, 2019 5:03:56 PM GMT+08:00 | [BB Refresh][IVC] Move couch using CT-Box, "Go to scan position"/"Go to saved position" turned to twinking if pressed meantime | BB_WIP6 |  |  | Subsystem | IVC | NA | No Plan to Fix | b_bibin | 0 | 4 |  | Console4 | NO | Closed | WAS | Brilliance Big Bore
CTUDT00624057 | Refresh | Closed | zhu_je | May 27, 2019 4:39:42 PM GMT+08:00 | [BB Refresh][IVC] The couch motion was not changed if pressing button move to inner/outer laser button on IVC if couch position was changed | BB_WIP6 | BB_WIP8 |  | Subsystem | IVC | N/A | Fixed | 320021835 | 0 | 4 | IVC Software | Console4 | NO | Closed | BB_WIP8_SubCom | Brilliance Big Bore |  | BB_WIP8 | June 26, 2019 8:39:32 PM GMT+08:00
CTUDT00624055 | Refresh | Closed | zhu_je | May 27, 2019 4:26:29 PM GMT+08:00 | [BB Refresh][IVC] The Console IV UI was not sync if moving couch using IVC HW | BB_WIP6 | BB_WIP8 |  | Subsystem | IVC | Console.SRS-15822 | Fixed | Snehashis_Dey | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP8_SubCom | Brilliance Big Bore |  | BB_WIP8 | June 25, 2019 8:58:04 PM GMT+08:00
CTUDT00624053 | Refresh | Closed | zhu_je | May 27, 2019 4:16:01 PM GMT+08:00 | [BB Refresh][IVC] IVC box display integer numbers numbers rather than to 1 decimal place | BB_WIP6 | BB_WIP7 | Usability confusion only | Subsystem | IVC | Console-SRS.15703 | Fixed | b_bibin | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP7_SubCom | Brilliance Big Bore |  | BB_WIP7 | May 31, 2019 5:27:30 PM GMT+08:00
CTUDT00624051 | Refresh | Closed | zhu_je | May 27, 2019 4:03:11 PM GMT+08:00 | [BB Refresh][IVC]There is error message popping up and scan was paused if moving couch using IVC after clicking go button | BB_WIP6 |  |  | Subsystem | IVC | NA | Works As Specified / Not A Defect | zhu_je | 1 | 4 |  | GCX | NO | Closed | BB_WIP8_SubCom | Brilliance Big Bore
CTUDT00624049 | Refresh | Closed | zhu_je | May 27, 2019 3:56:41 PM GMT+08:00 | [BB Refresh][IVC]sometimes couch up and down didnot work if click icon up/down icon on screen | BB_WIP6 | BB_WIP8 | User can use CTbox or Gantry panel to move couch. | Subsystem | IVC | Console-SRS.15693 | Fixed | b_bibin | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP8_SegReq | Brilliance Big Bore |  | BB_WIP8 | July 16, 2019 12:42:14 PM GMT+08:00
CTUDT00624047 | Refresh | Closed | zhu_je | May 27, 2019 3:50:41 PM GMT+08:00 | [BB Refresh][IVC]The value of innerest and outmost value are different using IVC rocker switch command with using CT-box command | BB_WIP6 |  |  | Subsystem | IVC | NA | Works As Specified / Not A Defect | karkour | 0 |  |  | GCX | NO | Closed |  | Brilliance Big Bore
CTUDT00624045 | Refresh | Closed | zhu_je | May 27, 2019 3:39:58 PM GMT+08:00 | [BB Refresh][IVC]"Go to saved position" was not fine if zero couch on GCP | BB_WIP6 | BB_WIP8 |  | Subsystem | IVC | NA | Fixed | b_bibin | 0 | 2 | Console Software | Console4 | YES | Closed | BB_WIP8_SubCom | Brilliance Big Bore |  | BB_WIP8 | June 18, 2019 8:09:03 PM GMT+08:00
CTUDT00624042 | Refresh | Closed | zhu_je | May 27, 2019 3:03:25 PM GMT+08:00 | [BB Refresh][I2P] I2P window was closed with an error message popping up | BB_WIP6 | BB_WIP7 |  | Subsystem | ImageToPlan | NA | Fixed | g_prasad | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP7_SubCom | Brilliance Big Bore |  | BB_WIP7 | May 30, 2019 4:55:14 PM GMT+08:00
CTUDT00624040 | Refresh | Closed | zhu_je | May 27, 2019 2:31:34 PM GMT+08:00 | [BB Refresh][I2P] I2P window was opened again and still need to input username and password to login | BB_WIP6 | BB_WIP8 |  | System |  | NA | Works As Specified / Not A Defect | zhu_je | 0 | 2 | Test Design | Support-System | NO | Closed | BB_WIP8_SegReq | Brilliance Big Bore |  | BB_WIP8 | August 5, 2019 8:48:27 AM GMT+08:00
CTUDT00624036 | Refresh | Closed | cohen_y | May 27, 2019 3:05:04 AM GMT+08:00 | Refresh: FEE Test: Ops User: Number of CFG tables and MDU pairs is inconsistent Error appears during kV Linearity calculations | WIP16 | WIP17 |  | Subsystem |  | n/a | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP17SubCom | IQon1 |  | WIP17 | May 30, 2019 2:37:38 PM GMT+08:00
CTUDT00624033 | Refresh | Closed | finestine_m | May 27, 2019 12:57:31 AM GMT+08:00 | Refresh: Physics Analysis: Wrong Scan Time in Physics Analysis | WIP16 | WIP18 |  | Subsystem | Physics Analysis | n/a | Fixed | bikman | 0 | 2 | DMS Software | DMS | NO | Closed | WIP18SubCom | Ovation |  | WIP18 | July 3, 2019 3:10:14 PM GMT+08:00
CTUDT00624031 | Refresh | Closed | Noy_Seelig | May 26, 2019 11:39:29 PM GMT+08:00 | Refresh: Wrong Couch position in Performance Infant | WIP16 |  | N/A | Subsystem | Calibrations | N/A | Invalid/Uncontrolled Configuration | elviral | 0 |  |  | DMS | NO | Closed |  | IQon1
CTUDT00624027 | Refresh | Closed | beck_d | May 26, 2019 8:39:06 PM GMT+08:00 | Refresh: Acceptance/Constancy Tests: Harmonized: Slice Thickness case 3 ROIs are not placed in the correct slice according to Spec. | WIP16 | N/A |  | Subsystem |  | N/A | Fixed | braunstein_d | 0 | 2 | Design Documentation | Support-Physics | NO | Closed | WIP19SegReq | Ovation |  | N/A | June 13, 2019 2:56:27 PM GMT+08:00
CTUDT00624024 | Refresh | Closed | urim | May 26, 2019 6:23:14 PM GMT+08:00 | Refresh Pre-Validation: Poor Industrial Design of Harmonized Phantom Holder | WIP14 | WIP20 | Issue 1: The poor ID design is appearance issue, it’s not impact on function. Issue 2 was caused by improper operation, will add WI into the BOM of Phantom holder to guide the operator. Issue 3 is supplier quality issue. Add additional chamfer to avoid it. | Subsystem | Harmonized Phantom | Req ID :DHF173447-2069 | Fixed | wang_ada | 0 | 2 | Couch Mechanical | Couch | YES | Closed | 459800412001 RevC; 459801554641 RevC; 459801566261 RevB | Ovation |  | WIP20 | October 30, 2019 2:48:24 PM GMT+08:00
CTUDT00624024 | Refresh | Closed | urim | May 26, 2019 6:23:14 PM GMT+08:00 | Refresh Pre-Validation: Poor Industrial Design of Harmonized Phantom Holder | WIP14 | WIP20 | Issue 1: The poor ID design is appearance issue, it’s not impact on function. Issue 2 was caused by improper operation, will add WI into the BOM of Phantom holder to guide the operator. Issue 3 is supplier quality issue. Add additional chamfer to avoid it. | Subsystem | Harmonized Phantom | Req ID :DHF173447-2069 | Fixed | wang_ada | 0 | 2 | Couch Mechanical | Couch | YES | Closed | 459800412001 RevC; 459801554641 RevC; 459801566261 RevB | Brilliance 64 |  | WIP20 | October 30, 2019 2:48:24 PM GMT+08:00
CTUDT00624024 | Refresh | Closed | urim | May 26, 2019 6:23:14 PM GMT+08:00 | Refresh Pre-Validation: Poor Industrial Design of Harmonized Phantom Holder | WIP14 | WIP20 | Issue 1: The poor ID design is appearance issue, it’s not impact on function. Issue 2 was caused by improper operation, will add WI into the BOM of Phantom holder to guide the operator. Issue 3 is supplier quality issue. Add additional chamfer to avoid it. | Subsystem | Harmonized Phantom | Req ID :DHF173447-2069 | Fixed | wang_ada | 0 | 2 | Couch Mechanical | Couch | YES | Closed | 459800412001 RevC; 459801554641 RevC; 459801566261 RevB | Brilliance Big Bore |  | WIP20 | October 30, 2019 2:48:24 PM GMT+08:00
CTUDT00624024 | Refresh | Closed | urim | May 26, 2019 6:23:14 PM GMT+08:00 | Refresh Pre-Validation: Poor Industrial Design of Harmonized Phantom Holder | WIP14 | WIP20 | Issue 1: The poor ID design is appearance issue, it’s not impact on function. Issue 2 was caused by improper operation, will add WI into the BOM of Phantom holder to guide the operator. Issue 3 is supplier quality issue. Add additional chamfer to avoid it. | Subsystem | Harmonized Phantom | Req ID :DHF173447-2069 | Fixed | wang_ada | 0 | 2 | Couch Mechanical | Couch | YES | Closed | 459800412001 RevC; 459801554641 RevC; 459801566261 RevB | Ingenuity 64 |  | WIP20 | October 30, 2019 2:48:24 PM GMT+08:00
CTUDT00624024 | Refresh | Closed | urim | May 26, 2019 6:23:14 PM GMT+08:00 | Refresh Pre-Validation: Poor Industrial Design of Harmonized Phantom Holder | WIP14 | WIP20 | Issue 1: The poor ID design is appearance issue, it’s not impact on function. Issue 2 was caused by improper operation, will add WI into the BOM of Phantom holder to guide the operator. Issue 3 is supplier quality issue. Add additional chamfer to avoid it. | Subsystem | Harmonized Phantom | Req ID :DHF173447-2069 | Fixed | wang_ada | 0 | 2 | Couch Mechanical | Couch | YES | Closed | 459800412001 RevC; 459801554641 RevC; 459801566261 RevB | IQon1 |  | WIP20 | October 30, 2019 2:48:24 PM GMT+08:00
CTUDT00624023 | Refresh | Closed | attali_s | May 26, 2019 6:21:56 PM GMT+08:00 | Refresh: Manual Bad Detectors: Bad Detectors are not marked when pressing on the BadDet checkbox | WIP16 | N/A - Test DEsign | n/a | Subsystem | Test Design | DHF302487 ID140 | Fixed | gehtin_e | 0 | 3 | Test Design | DMS | NO | Closed | WIP18SegReq | IQon1 |  | N/A - Test DEsign | July 8, 2019 6:09:24 PM GMT+08:00
CTUDT00624021 | Refresh | Closed | 320053770 | May 26, 2019 3:34:20 PM GMT+08:00 | RunDig.exe app failed on Windows 10. | WIP 16 | WIP18 | Please use PmsDView.exe to view images instead of RunDlg.exe from the same path. | Subsystem | CD-DVD operation | NA | Works As Specified / Not A Defect | s_sa | 1 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP18SegReq | Ovation |  | WIP18 | September 16, 2019 1:16:34 PM GMT+08:00
CTUDT00624021 | Refresh | Closed | 320053770 | May 26, 2019 3:34:20 PM GMT+08:00 | RunDig.exe app failed on Windows 10. | WIP 16 | WIP18 | Please use PmsDView.exe to view images instead of RunDlg.exe from the same path. | Subsystem | CD-DVD operation | NA | Works As Specified / Not A Defect | s_sa | 1 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP18SegReq | Brilliance 64 |  | WIP18 | September 16, 2019 1:16:34 PM GMT+08:00
CTUDT00624021 | Refresh | Closed | 320053770 | May 26, 2019 3:34:20 PM GMT+08:00 | RunDig.exe app failed on Windows 10. | WIP 16 | WIP18 | Please use PmsDView.exe to view images instead of RunDlg.exe from the same path. | Subsystem | CD-DVD operation | NA | Works As Specified / Not A Defect | s_sa | 1 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP18SegReq | Brilliance Big Bore |  | WIP18 | September 16, 2019 1:16:34 PM GMT+08:00
CTUDT00624021 | Refresh | Closed | 320053770 | May 26, 2019 3:34:20 PM GMT+08:00 | RunDig.exe app failed on Windows 10. | WIP 16 | WIP18 | Please use PmsDView.exe to view images instead of RunDlg.exe from the same path. | Subsystem | CD-DVD operation | NA | Works As Specified / Not A Defect | s_sa | 1 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP18SegReq | Ingenuity 64 |  | WIP18 | September 16, 2019 1:16:34 PM GMT+08:00
CTUDT00624021 | Refresh | Closed | 320053770 | May 26, 2019 3:34:20 PM GMT+08:00 | RunDig.exe app failed on Windows 10. | WIP 16 | WIP18 | Please use PmsDView.exe to view images instead of RunDlg.exe from the same path. | Subsystem | CD-DVD operation | NA | Works As Specified / Not A Defect | s_sa | 1 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP18SegReq | IQon1 |  | WIP18 | September 16, 2019 1:16:34 PM GMT+08:00
CTUDT00624019 | Refresh | Closed | 320053770 | May 26, 2019 3:17:48 PM GMT+08:00 | During integration for wip 16 in an ePhi test, while checking Dicom tags exportation, some of the patient data tags are missing. | wip 15 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | s_sa | 0 |  |  | iCAP | NO | Closed |  | Ovation
CTUDT00624019 | Refresh | Closed | 320053770 | May 26, 2019 3:17:48 PM GMT+08:00 | During integration for wip 16 in an ePhi test, while checking Dicom tags exportation, some of the patient data tags are missing. | wip 15 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | s_sa | 0 |  |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00624019 | Refresh | Closed | 320053770 | May 26, 2019 3:17:48 PM GMT+08:00 | During integration for wip 16 in an ePhi test, while checking Dicom tags exportation, some of the patient data tags are missing. | wip 15 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | s_sa | 0 |  |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00624019 | Refresh | Closed | 320053770 | May 26, 2019 3:17:48 PM GMT+08:00 | During integration for wip 16 in an ePhi test, while checking Dicom tags exportation, some of the patient data tags are missing. | wip 15 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | s_sa | 0 |  |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00624019 | Refresh | Closed | 320053770 | May 26, 2019 3:17:48 PM GMT+08:00 | During integration for wip 16 in an ePhi test, while checking Dicom tags exportation, some of the patient data tags are missing. | wip 15 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | s_sa | 0 |  |  | iCAP | NO | Closed |  | IQon1
CTUDT00624017 | Refresh | Closed | bhawna_g | May 24, 2019 7:59:55 PM GMT+08:00 | Refresh: Spectral: KeV field is accepting decimal point for the Spectral result type ‘MonoE’ | 4.8 | BB_WIP7 |  | Subsystem |  | N/A | Fixed | chokkalingam_r | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP7_SegReq | Brilliance Big Bore |  | BB_WIP7 | May 27, 2019 5:05:56 PM GMT+08:00
CTUDT00624011 | Refresh | Closed | bhawna_g | May 24, 2019 6:18:03 PM GMT+08:00 | For Surview Dual, Auto film check box can be enabled, “Auto Film Parameters” are enabled and on selecting “*“in the center field it throws system error. | 4.8 | BB_WIP7 |  | Subsystem |  | NA | Fixed | agrawal_ay | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP7_SubCom | Brilliance Big Bore |  | BB_WIP7 | May 29, 2019 6:57:24 PM GMT+08:00
CTUDT00624010 | Refresh | Closed | ravindra_a | May 24, 2019 6:06:13 PM GMT+08:00 | Refresh: console user messages localization mismatches | 4.8 | BB_WIP7 | NA | Subsystem | Localization | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | YES | Closed | BB_WIP7_SubCom | Brilliance Big Bore |  | BB_WIP7 | May 30, 2019 1:26:58 PM GMT+08:00
CTUDT00624003 | Refresh | Closed | v_sumanth | May 24, 2019 4:54:44 PM GMT+08:00 | BB Refresh :Selecting any configured devices in dose structured report in Preferences->Dose management tab, then local was not selected by default | WIP 16 | WIP18 | None | Subsystem |  | Console-SRS.3405, Console-SRS.3406, Console-SRS.10681 | Fixed | mY_h | 0 | 4 | Console Software | Console4 | NO | Closed | WIP18SubCom | Ovation |  | WIP18 | May 31, 2019 3:08:30 PM GMT+08:00
CTUDT00624003 | Refresh | Closed | v_sumanth | May 24, 2019 4:54:44 PM GMT+08:00 | BB Refresh :Selecting any configured devices in dose structured report in Preferences->Dose management tab, then local was not selected by default | WIP 16 | WIP18 | None | Subsystem |  | Console-SRS.3405, Console-SRS.3406, Console-SRS.10681 | Fixed | mY_h | 0 | 4 | Console Software | Console4 | NO | Closed | WIP18SubCom | Brilliance 64 |  | WIP18 | May 31, 2019 3:08:30 PM GMT+08:00
CTUDT00624003 | Refresh | Closed | v_sumanth | May 24, 2019 4:54:44 PM GMT+08:00 | BB Refresh :Selecting any configured devices in dose structured report in Preferences->Dose management tab, then local was not selected by default | WIP 16 | WIP18 | None | Subsystem |  | Console-SRS.3405, Console-SRS.3406, Console-SRS.10681 | Fixed | mY_h | 0 | 4 | Console Software | Console4 | NO | Closed | WIP18SubCom | Ingenuity 64 |  | WIP18 | May 31, 2019 3:08:30 PM GMT+08:00
CTUDT00624003 | Refresh | Closed | v_sumanth | May 24, 2019 4:54:44 PM GMT+08:00 | BB Refresh :Selecting any configured devices in dose structured report in Preferences->Dose management tab, then local was not selected by default | WIP 16 | WIP18 | None | Subsystem |  | Console-SRS.3405, Console-SRS.3406, Console-SRS.10681 | Fixed | mY_h | 0 | 4 | Console Software | Console4 | NO | Closed | WIP18SubCom | IQon1 |  | WIP18 | May 31, 2019 3:08:30 PM GMT+08:00
CTUDT00624002 | Refresh | Closed | navale_s | May 24, 2019 4:17:00 PM GMT+08:00 | Refresh BR64: The backup Voice language is not upgraded | 4.8 |  | Exam Card can be stored with the correct value | Subsystem |  | NA | Works As Specified / Not A Defect | wang_l | 0 | 3 |  | Support-System | NO | Closed | WAS | Brilliance Big Bore
CTUDT00623999 | Refresh | Closed | wei_f | May 24, 2019 3:41:10 PM GMT+08:00 | Some of the OS services information for the system was not consistent with the DHF314402 Services Configured excel list | WIP16 | BB_WIP12 WIP22 |  | Subsystem | OS hardening | STPS-SSRS-1226 | Fixed | dasgupta_s | 0 | 2 | Console Software | Support-Security | NO | Closed | BB_WIP8_SubCom & WIP18_SubCom | Brilliance 64 |  | BB_WIP12 WIP22 | November 13, 2019 8:14:30 PM GMT+08:00
CTUDT00623999 | Refresh | Closed | wei_f | May 24, 2019 3:41:10 PM GMT+08:00 | Some of the OS services information for the system was not consistent with the DHF314402 Services Configured excel list | WIP16 | BB_WIP12 WIP22 |  | Subsystem | OS hardening | STPS-SSRS-1226 | Fixed | dasgupta_s | 0 | 2 | Console Software | Support-Security | NO | Closed | BB_WIP8_SubCom & WIP18_SubCom | Brilliance Big Bore |  | BB_WIP12 WIP22 | November 13, 2019 8:14:30 PM GMT+08:00
CTUDT00623999 | Refresh | Closed | wei_f | May 24, 2019 3:41:10 PM GMT+08:00 | Some of the OS services information for the system was not consistent with the DHF314402 Services Configured excel list | WIP16 | BB_WIP12 WIP22 |  | Subsystem | OS hardening | STPS-SSRS-1226 | Fixed | dasgupta_s | 0 | 2 | Console Software | Support-Security | NO | Closed | BB_WIP8_SubCom & WIP18_SubCom | Ingenuity 64 |  | BB_WIP12 WIP22 | November 13, 2019 8:14:30 PM GMT+08:00
CTUDT00623999 | Refresh | Closed | wei_f | May 24, 2019 3:41:10 PM GMT+08:00 | Some of the OS services information for the system was not consistent with the DHF314402 Services Configured excel list | WIP16 | BB_WIP12 WIP22 |  | Subsystem | OS hardening | STPS-SSRS-1226 | Fixed | dasgupta_s | 0 | 2 | Console Software | Support-Security | NO | Closed | BB_WIP8_SubCom & WIP18_SubCom | IQon1 |  | BB_WIP12 WIP22 | November 13, 2019 8:14:30 PM GMT+08:00
CTUDT00623997 | Refresh | Closed | wang_l | May 24, 2019 3:13:50 PM GMT+08:00 | Refresh BB: Software Distribution tool open with error | BB_WIP6 | BB_WIP7 | NO | Subsystem | Software Distribution Tool | Console-STPS.274 | Fixed | sagar_v | 0 | 2 | Console Software | Service Tools | NO | Closed | BB_WIP7_SubCom | Brilliance Big Bore |  | BB_WIP7 | May 28, 2019 6:22:23 PM GMT+08:00
CTUDT00623992 | Refresh | Closed | swamypreethi | May 24, 2019 12:05:59 PM GMT+08:00 | IOCC_Refresh: DICOM printer : Max Density attribute is not present | WIP13 |  |  | Subsystem | IOCC - USP | NA | Change Proposal | s_sa | 0 |  |  | iCAP | NO | Closed |  | Ovation
CTUDT00623992 | Refresh | Closed | swamypreethi | May 24, 2019 12:05:59 PM GMT+08:00 | IOCC_Refresh: DICOM printer : Max Density attribute is not present | WIP13 |  |  | Subsystem | IOCC - USP | NA | Change Proposal | s_sa | 0 |  |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00623992 | Refresh | Closed | swamypreethi | May 24, 2019 12:05:59 PM GMT+08:00 | IOCC_Refresh: DICOM printer : Max Density attribute is not present | WIP13 |  |  | Subsystem | IOCC - USP | NA | Change Proposal | s_sa | 0 |  |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00623992 | Refresh | Closed | swamypreethi | May 24, 2019 12:05:59 PM GMT+08:00 | IOCC_Refresh: DICOM printer : Max Density attribute is not present | WIP13 |  |  | Subsystem | IOCC - USP | NA | Change Proposal | s_sa | 0 |  |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00623992 | Refresh | Closed | swamypreethi | May 24, 2019 12:05:59 PM GMT+08:00 | IOCC_Refresh: DICOM printer : Max Density attribute is not present | WIP13 |  |  | Subsystem | IOCC - USP | NA | Change Proposal | s_sa | 0 |  |  | iCAP | NO | Closed |  | IQon1
CTUDT00623989 | Refresh | Closed | wei_f | May 24, 2019 9:44:38 AM GMT+08:00 | after enter Chinese character,Institute Information show as ??? | BB_WIP6 | BB_WIP7 |  | Subsystem |  | NA | Fixed | v_sumanth | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP7_SubCom | Brilliance Big Bore |  | BB_WIP7 | May 28, 2019 7:30:59 PM GMT+08:00
CTUDT00623984 | Refresh | Closed | Sharmila_Vasireddy | May 23, 2019 9:07:39 PM GMT+08:00 | Wrong message was displaying when we are entering wrong password for Philips_Service user | BB_WIP5.1 |  |  | Subsystem |  |  | Duplicate | goswami_s | 0 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00623984 | Refresh | Closed | Sharmila_Vasireddy | May 23, 2019 9:07:39 PM GMT+08:00 | Wrong message was displaying when we are entering wrong password for Philips_Service user | BB_WIP5.1 |  |  | Subsystem |  |  | Duplicate | goswami_s | 0 |  |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00623984 | Refresh | Closed | Sharmila_Vasireddy | May 23, 2019 9:07:39 PM GMT+08:00 | Wrong message was displaying when we are entering wrong password for Philips_Service user | BB_WIP5.1 |  |  | Subsystem |  |  | Duplicate | goswami_s | 0 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00623984 | Refresh | Closed | Sharmila_Vasireddy | May 23, 2019 9:07:39 PM GMT+08:00 | Wrong message was displaying when we are entering wrong password for Philips_Service user | BB_WIP5.1 |  |  | Subsystem |  |  | Duplicate | goswami_s | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00623984 | Refresh | Closed | Sharmila_Vasireddy | May 23, 2019 9:07:39 PM GMT+08:00 | Wrong message was displaying when we are entering wrong password for Philips_Service user | BB_WIP5.1 |  |  | Subsystem |  |  | Duplicate | goswami_s | 0 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00623983 | Refresh | Closed | wang_l | May 23, 2019 9:06:07 PM GMT+08:00 | Refresh BB: Stage Infant Head 100 ,80 and 120 of Performance Test didn't pass for harmonized phantom | BB_WIP6 | N/A |  | Subsystem | Calibrations;BLocker | N/A | Fixed | braunstein_d | 0 | 2 | Design Documentation | Support-Physics | NO | Closed | WIP18SegReq | Brilliance Big Bore |  | N/A | July 8, 2019 4:32:25 PM GMT+08:00
CTUDT00623980 | Refresh | Closed | wang_l | May 23, 2019 8:32:36 PM GMT+08:00 | Refresh BB: Stage IMR noise Body 1a,1b of Performance Test didn't pass for harmonized phantom | BB_WIP6 | DHF170396 Rev20 WIP21 |  | Subsystem | Calibrations; Verified1 | n/a | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP21SubCom | Brilliance Big Bore |  | DHF170396 Rev20 WIP21 | October 6, 2019 11:01:54 PM GMT+08:00
CTUDT00623978 | Refresh | Closed | wang_l | May 23, 2019 8:23:06 PM GMT+08:00 | Refresh BB: Some stage of Constancy Test didn't pass | BB_WIP6 |  |  | Subsystem | Constancy Test;Blocker |  | Duplicate | morton_t | 0 |  |  | Support-Physics | NO | Closed | BB_WIP8_SegReq | Brilliance Big Bore
CTUDT00623976 | Refresh | Closed | wang_l | May 23, 2019 8:17:39 PM GMT+08:00 | Refresh: Some stage of Acceptance/Constancy Test did not pass | BB_WIP6 | WIP20 | No Workaround | Subsystem | Acceptance Test;Blocker19; Verified1 | CT-NM.SRS.IQ.2779 | Fixed | richards_b | 0 | 2 | CIRS_SW_ImgNodeSvr | Support-Physics | NO | Closed | WIP19SubCom | Brilliance 64 |  | WIP20 | August 16, 2019 10:27:32 PM GMT+08:00
CTUDT00623976 | Refresh | Closed | wang_l | May 23, 2019 8:17:39 PM GMT+08:00 | Refresh: Some stage of Acceptance/Constancy Test did not pass | BB_WIP6 | WIP20 | No Workaround | Subsystem | Acceptance Test;Blocker19; Verified1 | CT-NM.SRS.IQ.2779 | Fixed | richards_b | 0 | 2 | CIRS_SW_ImgNodeSvr | Support-Physics | NO | Closed | WIP19SubCom | Brilliance Big Bore |  | WIP20 | August 16, 2019 10:27:32 PM GMT+08:00
CTUDT00623976 | Refresh | Closed | wang_l | May 23, 2019 8:17:39 PM GMT+08:00 | Refresh: Some stage of Acceptance/Constancy Test did not pass | BB_WIP6 | WIP20 | No Workaround | Subsystem | Acceptance Test;Blocker19; Verified1 | CT-NM.SRS.IQ.2779 | Fixed | richards_b | 0 | 2 | CIRS_SW_ImgNodeSvr | Support-Physics | NO | Closed | WIP19SubCom | Ingenuity 64 |  | WIP20 | August 16, 2019 10:27:32 PM GMT+08:00
CTUDT00623974 | Refresh | Closed | wang_l | May 23, 2019 8:13:44 PM GMT+08:00 | Refresh BB: stage body CT8, CT12, idose lev2 and lev7 of Performance Test didn't pass | BB_WIP6 | WIP18 |  | Subsystem | Calibrations; Blocker WIP18 | N/A | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP18SubCom | Brilliance Big Bore |  | WIP18 | July 11, 2019 4:50:22 PM GMT+08:00
CTUDT00623968 | Refresh | Closed | roy_r | May 23, 2019 5:45:19 PM GMT+08:00 | BB Refresh: SRS for Big Bore V4.8 require update to cover Console-SRS.16537, 15462 and STPS-SSRS-1341 | BB_WIP6 | DHF298555 Rev 06 - BB SRS |  | System |  | CT-NM.SRS.Interop.327 | Fixed | mistry_h | 0 | 2 | Requirements Documentation | Support-System | YES | Closed | BB_WIP7_SubCom | Brilliance Big Bore |  | DHF298555 Rev 06 - BB SRS | July 26, 2019 8:17:25 PM GMT+08:00
CTUDT00623963 | Refresh | Closed | bommiReddy_r | May 23, 2019 3:26:59 PM GMT+08:00 | Refresh: Serviceability: After upgrade from IQON HB 4.7.5 to 4.7.7 user was not able to launch ECM | WIP16 |  |  | Subsystem | Upgrade Utility | NA | Duplicate | chokkalingam_r | 0 |  |  | Service Tools | NO | Closed | WIP18SubCom | IQon1
CTUDT00623961 | Refresh | Closed | wang_l | May 23, 2019 3:19:38 PM GMT+08:00 | Refresh: Tube current scan time Calibration failed | WIP16 |  |  | Subsystem | Calibrations |  | Duplicate | altshuller_mark | 0 |  |  | DMS | NO | Closed |  | Brilliance Big Bore
CTUDT00623956 | Refresh | Closed | wang_l | May 23, 2019 10:54:41 AM GMT+08:00 | Refresh: Image Rotation calibration failed | WIP16 | WIP18 |  | Subsystem | Calibrations | n/a | Fixed | zhou_x | 0 | 2 | Gantry Calibrations | GCX | YES | Closed | WIP18SubCom | Ingenuity 64 |  | WIP18 | June 17, 2019 4:23:01 PM GMT+08:00
CTUDT00623954 | Refresh | Closed | wang_l | May 23, 2019 10:17:05 AM GMT+08:00 | Refresh: Software Distribution- Incorrect behavior for Cancel behavior of software package download | WIP16 |  | Yes cleanup the system and manual create configured folder "D:\RemoteSoftwareDistribution\ ", Restart and retry. | Subsystem |  | Console-STPS.1276 | Works As Specified / Not A Defect | sagar_v | 0 |  |  | Service Tools | NO | Closed | WAS | Ovation
CTUDT00623954 | Refresh | Closed | wang_l | May 23, 2019 10:17:05 AM GMT+08:00 | Refresh: Software Distribution- Incorrect behavior for Cancel behavior of software package download | WIP16 |  | Yes cleanup the system and manual create configured folder "D:\RemoteSoftwareDistribution\ ", Restart and retry. | Subsystem |  | Console-STPS.1276 | Works As Specified / Not A Defect | sagar_v | 0 |  |  | Service Tools | NO | Closed | WAS | Brilliance Big Bore
CTUDT00623954 | Refresh | Closed | wang_l | May 23, 2019 10:17:05 AM GMT+08:00 | Refresh: Software Distribution- Incorrect behavior for Cancel behavior of software package download | WIP16 |  | Yes cleanup the system and manual create configured folder "D:\RemoteSoftwareDistribution\ ", Restart and retry. | Subsystem |  | Console-STPS.1276 | Works As Specified / Not A Defect | sagar_v | 0 |  |  | Service Tools | NO | Closed | WAS | Ingenuity 64
CTUDT00623954 | Refresh | Closed | wang_l | May 23, 2019 10:17:05 AM GMT+08:00 | Refresh: Software Distribution- Incorrect behavior for Cancel behavior of software package download | WIP16 |  | Yes cleanup the system and manual create configured folder "D:\RemoteSoftwareDistribution\ ", Restart and retry. | Subsystem |  | Console-STPS.1276 | Works As Specified / Not A Defect | sagar_v | 0 |  |  | Service Tools | NO | Closed | WAS | IQon1
CTUDT00623949 | Refresh | Closed | shekhar_p | May 22, 2019 11:51:42 PM GMT+08:00 | Planned SSDE is getting updated in the Acquisition viewer, tool box and show all after the scan | WIP16 | WIP18 |  | Subsystem |  | N/A | Fixed | bhawna_g | 0 | 4 | Console Software | Console4 | NO | Closed | WIP18SubCom | Ovation |  | WIP18 | June 11, 2019 1:40:51 PM GMT+08:00
CTUDT00623949 | Refresh | Closed | shekhar_p | May 22, 2019 11:51:42 PM GMT+08:00 | Planned SSDE is getting updated in the Acquisition viewer, tool box and show all after the scan | WIP16 | WIP18 |  | Subsystem |  | N/A | Fixed | bhawna_g | 0 | 4 | Console Software | Console4 | NO | Closed | WIP18SubCom | Brilliance 64 |  | WIP18 | June 11, 2019 1:40:51 PM GMT+08:00
CTUDT00623949 | Refresh | Closed | shekhar_p | May 22, 2019 11:51:42 PM GMT+08:00 | Planned SSDE is getting updated in the Acquisition viewer, tool box and show all after the scan | WIP16 | WIP18 |  | Subsystem |  | N/A | Fixed | bhawna_g | 0 | 4 | Console Software | Console4 | NO | Closed | WIP18SubCom | Ingenuity 64 |  | WIP18 | June 11, 2019 1:40:51 PM GMT+08:00
CTUDT00623949 | Refresh | Closed | shekhar_p | May 22, 2019 11:51:42 PM GMT+08:00 | Planned SSDE is getting updated in the Acquisition viewer, tool box and show all after the scan | WIP16 | WIP18 |  | Subsystem |  | N/A | Fixed | bhawna_g | 0 | 4 | Console Software | Console4 | NO | Closed | WIP18SubCom | IQon1 |  | WIP18 | June 11, 2019 1:40:51 PM GMT+08:00
CTUDT00623946 | Refresh | Closed | shekhar_p | May 22, 2019 9:02:06 PM GMT+08:00 | In CCT Fluro SSDE value is displaying as zero in acquisition viewer after releasing the pedal first time | WIP16 | WIP18 |  | Subsystem |  | N/A | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | YES | Closed | WIP18SubCom | Ovation |  | WIP18 | June 6, 2019 1:46:14 PM GMT+08:00
CTUDT00623946 | Refresh | Closed | shekhar_p | May 22, 2019 9:02:06 PM GMT+08:00 | In CCT Fluro SSDE value is displaying as zero in acquisition viewer after releasing the pedal first time | WIP16 | WIP18 |  | Subsystem |  | N/A | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | YES | Closed | WIP18SubCom | Brilliance 64 |  | WIP18 | June 6, 2019 1:46:14 PM GMT+08:00
CTUDT00623946 | Refresh | Closed | shekhar_p | May 22, 2019 9:02:06 PM GMT+08:00 | In CCT Fluro SSDE value is displaying as zero in acquisition viewer after releasing the pedal first time | WIP16 | WIP18 |  | Subsystem |  | N/A | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | YES | Closed | WIP18SubCom | Ingenuity 64 |  | WIP18 | June 6, 2019 1:46:14 PM GMT+08:00
CTUDT00623946 | Refresh | Closed | shekhar_p | May 22, 2019 9:02:06 PM GMT+08:00 | In CCT Fluro SSDE value is displaying as zero in acquisition viewer after releasing the pedal first time | WIP16 | WIP18 |  | Subsystem |  | N/A | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | YES | Closed | WIP18SubCom | IQon1 |  | WIP18 | June 6, 2019 1:46:14 PM GMT+08:00
CTUDT00623931 | Refresh | Closed | alex_nihaev | May 22, 2019 7:19:11 PM GMT+08:00 | Refresh: Voice Manager: Record Phrase functionality is not working | WIP16 |  |  | Subsystem | Voice Manager | NA | Works As Specified / Not A Defect | nanjunda_s | 0 |  |  | Console4 | NO | Closed | WAS | Ingenuity 64
CTUDT00623922 | Refresh | Closed | finestine_m | May 22, 2019 3:51:57 PM GMT+08:00 | Refresh: Beam Quality, Tube Potential Test: No Message Prompted When Some Fields Are Blank In "Dose Meter Details" Window | BB_WIP5.1 |  |  | Subsystem | Beam Quality Tube Potential Test | NaN | Duplicate | daibes_s | 0 |  |  | DMS | NO | Closed |  | Brilliance Big Bore
CTUDT00623920 | Refresh | Closed | swamy_s | May 22, 2019 3:38:14 PM GMT+08:00 | Crash Message:ImageToPlanClient has stopped working, when clicked on close button on Intellispace radiation oncology Page. | BB_WIP6 |  |  | Subsystem |  | NA | Duplicate | g_prasad | 0 | 3 |  | Console4 | NO | Closed | BB_WIP7_SubCom | Brilliance Big Bore
CTUDT00623918 | Refresh | Closed | v_sumanth | May 22, 2019 2:59:20 PM GMT+08:00 | Refresh: Cosmix filter result is stored as 16 bit HU image but it has to be stored as 12 bit HU image. | 4.8 | BB_WIP7 |  | Subsystem |  | Console.SRS.15462 | Fixed | Prakash_Chand_Yadav | 0 | 4 | Console Software | Console4 | NO | Closed | WIP17SubCom | Brilliance Big Bore |  | BB_WIP7 | May 28, 2019 11:35:03 AM GMT+08:00
CTUDT00623914 | Refresh | Closed | wang_l | May 22, 2019 2:54:58 PM GMT+08:00 | Refresh: Harmonized phantom is checked, the Instruction still display system infant phantom for Performance Infant test | WIP16 | WIP19 |  | Subsystem | Performance Infant test | N/A | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP20SubCom | Ovation |  | WIP19 | August 7, 2019 3:56:57 PM GMT+08:00
CTUDT00623914 | Refresh | Closed | wang_l | May 22, 2019 2:54:58 PM GMT+08:00 | Refresh: Harmonized phantom is checked, the Instruction still display system infant phantom for Performance Infant test | WIP16 | WIP19 |  | Subsystem | Performance Infant test | N/A | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP20SubCom | Brilliance Big Bore |  | WIP19 | August 7, 2019 3:56:57 PM GMT+08:00
CTUDT00623914 | Refresh | Closed | wang_l | May 22, 2019 2:54:58 PM GMT+08:00 | Refresh: Harmonized phantom is checked, the Instruction still display system infant phantom for Performance Infant test | WIP16 | WIP19 |  | Subsystem | Performance Infant test | N/A | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP20SubCom | Ingenuity 64 |  | WIP19 | August 7, 2019 3:56:57 PM GMT+08:00
CTUDT00623914 | Refresh | Closed | wang_l | May 22, 2019 2:54:58 PM GMT+08:00 | Refresh: Harmonized phantom is checked, the Instruction still display system infant phantom for Performance Infant test | WIP16 | WIP19 |  | Subsystem | Performance Infant test | N/A | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP20SubCom | IQon1 |  | WIP19 | August 7, 2019 3:56:57 PM GMT+08:00
CTUDT00623908 | Refresh | Closed | swamy_s | May 22, 2019 2:44:44 PM GMT+08:00 | In PSC, The system configuration name still shows " ImageToPlan Configuration Settings". | BB_WIP6 | BB_WIP7 |  | Subsystem |  | NA | Fixed | g_prasad | 0 | 4 | Service Tools_Software | Service Tools | NO | Closed | BB_WIP7_SubCom | Brilliance Big Bore |  | BB_WIP7 | May 30, 2019 1:36:58 AM GMT+08:00
CTUDT00623907 | Refresh | Closed | finestine_m | May 22, 2019 2:43:10 PM GMT+08:00 | Refresh: Acceptance Test: Multiple Stages Fail | BB_WIP5.1 |  |  | Subsystem | Calibrations | N/A | Works As Specified / Not A Defect | finestine_m | 1 |  |  | Support-System | NO | Closed |  | Brilliance Big Bore
CTUDT00623904 | Refresh | Closed | wang_l | May 22, 2019 2:19:03 PM GMT+08:00 | Refresh: Quality Assurance: Reliability Test error | WIP16 | DMR360022 |  | Subsystem | Calibrations; Blocker | n/a | Fixed | lin_s | 0 | 2 | Test Design | Support-System | NO | Closed | WIP18SegReq | Brilliance 64 |  | DMR360022 | July 30, 2019 2:38:32 PM GMT+08:00
CTUDT00623904 | Refresh | Closed | wang_l | May 22, 2019 2:19:03 PM GMT+08:00 | Refresh: Quality Assurance: Reliability Test error | WIP16 | DMR360022 |  | Subsystem | Calibrations; Blocker | n/a | Fixed | lin_s | 0 | 2 | Test Design | Support-System | NO | Closed | WIP18SegReq | Ingenuity 64 |  | DMR360022 | July 30, 2019 2:38:32 PM GMT+08:00
CTUDT00623902 | Refresh | Closed | wang_l | May 22, 2019 2:05:00 PM GMT+08:00 | Refresh: DFS position calibration failed because of out of memory error | WIP16 | WIP22 | Exit and restart the QA application. | Subsystem | Calibrations;Blocker21 | N/A | Fixed | zhou_x | 0 | 3 | Gantry Calibrations | GCX | NO | Closed | WIP21SegReq | Brilliance 64 |  | WIP22 | November 4, 2019 4:57:25 PM GMT+08:00
CTUDT00623902 | Refresh | Closed | wang_l | May 22, 2019 2:05:00 PM GMT+08:00 | Refresh: DFS position calibration failed because of out of memory error | WIP16 | WIP22 | Exit and restart the QA application. | Subsystem | Calibrations;Blocker21 | N/A | Fixed | zhou_x | 0 | 3 | Gantry Calibrations | GCX | NO | Closed | WIP21SegReq | Brilliance Big Bore |  | WIP22 | November 4, 2019 4:57:25 PM GMT+08:00
CTUDT00623902 | Refresh | Closed | wang_l | May 22, 2019 2:05:00 PM GMT+08:00 | Refresh: DFS position calibration failed because of out of memory error | WIP16 | WIP22 | Exit and restart the QA application. | Subsystem | Calibrations;Blocker21 | N/A | Fixed | zhou_x | 0 | 3 | Gantry Calibrations | GCX | NO | Closed | WIP21SegReq | Ingenuity 64 |  | WIP22 | November 4, 2019 4:57:25 PM GMT+08:00
CTUDT00623898 | Refresh | Closed | swamypreethi | May 22, 2019 1:20:58 PM GMT+08:00 | IOCC_Refresh: Acquisition with Chinese character set results in application crash | WIP15 |  |  | Subsystem | Chinese character | NA | Forwarded to Different Database | prabhu_m | 0 | 4 |  | Support-Simulator | NO | Closed | WIP17SegReq | Ovation
CTUDT00623898 | Refresh | Closed | swamypreethi | May 22, 2019 1:20:58 PM GMT+08:00 | IOCC_Refresh: Acquisition with Chinese character set results in application crash | WIP15 |  |  | Subsystem | Chinese character | NA | Forwarded to Different Database | prabhu_m | 0 | 4 |  | Support-Simulator | NO | Closed | WIP17SegReq | Brilliance 64
CTUDT00623898 | Refresh | Closed | swamypreethi | May 22, 2019 1:20:58 PM GMT+08:00 | IOCC_Refresh: Acquisition with Chinese character set results in application crash | WIP15 |  |  | Subsystem | Chinese character | NA | Forwarded to Different Database | prabhu_m | 0 | 4 |  | Support-Simulator | NO | Closed | WIP17SegReq | Brilliance Big Bore
CTUDT00623898 | Refresh | Closed | swamypreethi | May 22, 2019 1:20:58 PM GMT+08:00 | IOCC_Refresh: Acquisition with Chinese character set results in application crash | WIP15 |  |  | Subsystem | Chinese character | NA | Forwarded to Different Database | prabhu_m | 0 | 4 |  | Support-Simulator | NO | Closed | WIP17SegReq | Ingenuity 64
CTUDT00623898 | Refresh | Closed | swamypreethi | May 22, 2019 1:20:58 PM GMT+08:00 | IOCC_Refresh: Acquisition with Chinese character set results in application crash | WIP15 |  |  | Subsystem | Chinese character | NA | Forwarded to Different Database | prabhu_m | 0 | 4 |  | Support-Simulator | NO | Closed | WIP17SegReq | IQon1
CTUDT00623897 | Refresh | Closed | wang_l | May 22, 2019 1:19:41 PM GMT+08:00 | Refresh - SFTP and SSH connectin failed because COPSSH Control Panel open with error | WIP16 | WIP18 BB_WIP8 |  | Subsystem | Remote Connection | N/A | Fixed | dasgupta_s | 0 | 2 | Console Software | Support-Security | NO | Closed | WIP18SubCom | Ovation |  | WIP18 BB_WIP8 | June 21, 2019 1:56:54 PM GMT+08:00
CTUDT00623897 | Refresh | Closed | wang_l | May 22, 2019 1:19:41 PM GMT+08:00 | Refresh - SFTP and SSH connectin failed because COPSSH Control Panel open with error | WIP16 | WIP18 BB_WIP8 |  | Subsystem | Remote Connection | N/A | Fixed | dasgupta_s | 0 | 2 | Console Software | Support-Security | NO | Closed | WIP18SubCom | Brilliance 64 |  | WIP18 BB_WIP8 | June 21, 2019 1:56:54 PM GMT+08:00
CTUDT00623897 | Refresh | Closed | wang_l | May 22, 2019 1:19:41 PM GMT+08:00 | Refresh - SFTP and SSH connectin failed because COPSSH Control Panel open with error | WIP16 | WIP18 BB_WIP8 |  | Subsystem | Remote Connection | N/A | Fixed | dasgupta_s | 0 | 2 | Console Software | Support-Security | NO | Closed | WIP18SubCom | Brilliance Big Bore |  | WIP18 BB_WIP8 | June 21, 2019 1:56:54 PM GMT+08:00
CTUDT00623897 | Refresh | Closed | wang_l | May 22, 2019 1:19:41 PM GMT+08:00 | Refresh - SFTP and SSH connectin failed because COPSSH Control Panel open with error | WIP16 | WIP18 BB_WIP8 |  | Subsystem | Remote Connection | N/A | Fixed | dasgupta_s | 0 | 2 | Console Software | Support-Security | NO | Closed | WIP18SubCom | Ingenuity 64 |  | WIP18 BB_WIP8 | June 21, 2019 1:56:54 PM GMT+08:00
CTUDT00623897 | Refresh | Closed | wang_l | May 22, 2019 1:19:41 PM GMT+08:00 | Refresh - SFTP and SSH connectin failed because COPSSH Control Panel open with error | WIP16 | WIP18 BB_WIP8 |  | Subsystem | Remote Connection | N/A | Fixed | dasgupta_s | 0 | 2 | Console Software | Support-Security | NO | Closed | WIP18SubCom | IQon1 |  | WIP18 BB_WIP8 | June 21, 2019 1:56:54 PM GMT+08:00
CTUDT00623896 | Refresh | Closed | sowmya_k | May 22, 2019 1:04:40 PM GMT+08:00 | BB Refresh:Units for Planned CTDIVol and Planned DLP are incorrectly displayed in Acquisition viewer | WIP16 | WIP18 | No workaround. | Subsystem |  | NA | Fixed | v_su | 0 | 2 | Console Software | Console4 | YES | Closed | WIP17SubCom | Ovation |  | WIP18 | May 31, 2019 7:51:55 PM GMT+08:00
CTUDT00623896 | Refresh | Closed | sowmya_k | May 22, 2019 1:04:40 PM GMT+08:00 | BB Refresh:Units for Planned CTDIVol and Planned DLP are incorrectly displayed in Acquisition viewer | WIP16 | WIP18 | No workaround. | Subsystem |  | NA | Fixed | v_su | 0 | 2 | Console Software | Console4 | YES | Closed | WIP17SubCom | Brilliance 64 |  | WIP18 | May 31, 2019 7:51:55 PM GMT+08:00
CTUDT00623896 | Refresh | Closed | sowmya_k | May 22, 2019 1:04:40 PM GMT+08:00 | BB Refresh:Units for Planned CTDIVol and Planned DLP are incorrectly displayed in Acquisition viewer | WIP16 | WIP18 | No workaround. | Subsystem |  | NA | Fixed | v_su | 0 | 2 | Console Software | Console4 | YES | Closed | WIP17SubCom | Ingenuity 64 |  | WIP18 | May 31, 2019 7:51:55 PM GMT+08:00
CTUDT00623896 | Refresh | Closed | sowmya_k | May 22, 2019 1:04:40 PM GMT+08:00 | BB Refresh:Units for Planned CTDIVol and Planned DLP are incorrectly displayed in Acquisition viewer | WIP16 | WIP18 | No workaround. | Subsystem |  | NA | Fixed | v_su | 0 | 2 | Console Software | Console4 | YES | Closed | WIP17SubCom | IQon1 |  | WIP18 | May 31, 2019 7:51:55 PM GMT+08:00
CTUDT00623894 | Refresh | Closed | panchami_pradeep | May 22, 2019 11:26:05 AM GMT+08:00 | Refresh: PSC: TSSS and TSSC values are not getting updated after upgrade from V4.1.7 to V4.1.10 | WIP16 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | panchami_pradeep | 1 | 2 |  | Service Tools | YES | Closed | WIP18SubCom | Ovation
CTUDT00623894 | Refresh | Closed | panchami_pradeep | May 22, 2019 11:26:05 AM GMT+08:00 | Refresh: PSC: TSSS and TSSC values are not getting updated after upgrade from V4.1.7 to V4.1.10 | WIP16 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | panchami_pradeep | 1 | 2 |  | Service Tools | YES | Closed | WIP18SubCom | Brilliance 64
CTUDT00623894 | Refresh | Closed | panchami_pradeep | May 22, 2019 11:26:05 AM GMT+08:00 | Refresh: PSC: TSSS and TSSC values are not getting updated after upgrade from V4.1.7 to V4.1.10 | WIP16 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | panchami_pradeep | 1 | 2 |  | Service Tools | YES | Closed | WIP18SubCom | Brilliance Big Bore
CTUDT00623894 | Refresh | Closed | panchami_pradeep | May 22, 2019 11:26:05 AM GMT+08:00 | Refresh: PSC: TSSS and TSSC values are not getting updated after upgrade from V4.1.7 to V4.1.10 | WIP16 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | panchami_pradeep | 1 | 2 |  | Service Tools | YES | Closed | WIP18SubCom | Ingenuity 64
CTUDT00623894 | Refresh | Closed | panchami_pradeep | May 22, 2019 11:26:05 AM GMT+08:00 | Refresh: PSC: TSSS and TSSC values are not getting updated after upgrade from V4.1.7 to V4.1.10 | WIP16 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | panchami_pradeep | 1 | 2 |  | Service Tools | YES | Closed | WIP18SubCom | Ingenuity TF 64
CTUDT00623894 | Refresh | Closed | panchami_pradeep | May 22, 2019 11:26:05 AM GMT+08:00 | Refresh: PSC: TSSS and TSSC values are not getting updated after upgrade from V4.1.7 to V4.1.10 | WIP16 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | panchami_pradeep | 1 | 2 |  | Service Tools | YES | Closed | WIP18SubCom | IQon1
CTUDT00623892 | Refresh | Closed | sun_q | May 22, 2019 8:27:13 AM GMT+08:00 | Refresh-BB Move Inner/Outer Laser button function failed | BB_WIP6 | BB_WIP11 |  | Subsystem | Covers | NA | Fixed | kline_w | 0 | 2 | Console Software | GCX | NO | Closed | BB_WIP8_SubCom | Brilliance Big Bore |  | BB_WIP11 | October 17, 2019 10:20:22 AM GMT+08:00
CTUDT00623888 | Refresh | Closed | beck_d | May 22, 2019 1:19:14 AM GMT+08:00 | Refresh: Calibrations: Hcor Slice: Harmonized Phantom: Adult Body mA parameter is not according to spec | WIP15 | WIP18 |  | Subsystem | Calibrations | N/A | Fixed | Kostya_Baklan | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP17SubCom | Ovation |  | WIP18 | May 28, 2019 4:27:30 PM GMT+08:00
CTUDT00623885 | Refresh | Closed | beck_d | May 21, 2019 9:33:40 PM GMT+08:00 | Refresh: Calibrations: Constancy Test: Harmonized Phantom: Adult Body on Head Uniformity limits are not as defined in spec | WIP15 | WIP17 |  | Subsystem | Acceptance/Constancy | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP17SubCom | Ovation |  | WIP17 | May 28, 2019 4:01:10 PM GMT+08:00
CTUDT00623875 | Refresh | Closed | bhawna_g | May 21, 2019 6:19:58 PM GMT+08:00 | Refresh: Console: Short tube conditioning message is displayed even before 12 hours of calibration | 4.8 | BB_WIP7 |  | Subsystem | Calibrations | NA | Fixed | killada_s | 0 | 4 | Couch Software | Console4 | NO | Closed | BB_WIP7_SubCom | Brilliance Big Bore |  | BB_WIP7 | May 23, 2019 5:19:22 PM GMT+08:00
CTUDT00623872 | Refresh | Closed | nanjunda_s | May 21, 2019 6:14:16 PM GMT+08:00 | Getting bug rep error | WIP16 | WIP18 |  | Subsystem |  | None | Fixed | Prakash_Chand_Yadav | 0 | 3 | Console Software | Console4 | NO | Closed | WIP18SegReq | IQon1 |  | WIP18 | May 31, 2019 6:32:44 PM GMT+08:00
CTUDT00623860 | Refresh | Closed | 320053769 | May 21, 2019 1:55:37 PM GMT+08:00 | Non-DICOM images (.jpg, .bpm, .tiff, .avi) are not transferred to different remote device(s) | WIP16 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | 320053769 | 0 |  |  |  |  | Closed |  | Ovation
CTUDT00623860 | Refresh | Closed | 320053769 | May 21, 2019 1:55:37 PM GMT+08:00 | Non-DICOM images (.jpg, .bpm, .tiff, .avi) are not transferred to different remote device(s) | WIP16 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | 320053769 | 0 |  |  |  |  | Closed |  | IQon1
CTUDT00623854 | Refresh | Closed | oreg_o | May 20, 2019 9:32:03 PM GMT+08:00 | Refresh: WIP 15: Performance: Tube data mismatch between Summary and X-ray tube data. | WIP15 | WIP18 |  | Subsystem |  | N/A | No Plan to Fix | yatziv_l | 0 | 3 | DMS Calibrations | DMS | NO | Closed | SubNC | Ovation |  | WIP18 | June 13, 2019 3:37:26 PM GMT+08:00
CTUDT00623854 | Refresh | Closed | oreg_o | May 20, 2019 9:32:03 PM GMT+08:00 | Refresh: WIP 15: Performance: Tube data mismatch between Summary and X-ray tube data. | WIP15 | WIP18 |  | Subsystem |  | N/A | No Plan to Fix | yatziv_l | 0 | 3 | DMS Calibrations | DMS | NO | Closed | SubNC | Ingenuity 64 |  | WIP18 | June 13, 2019 3:37:26 PM GMT+08:00
CTUDT00623854 | Refresh | Closed | oreg_o | May 20, 2019 9:32:03 PM GMT+08:00 | Refresh: WIP 15: Performance: Tube data mismatch between Summary and X-ray tube data. | WIP15 | WIP18 |  | Subsystem |  | N/A | No Plan to Fix | yatziv_l | 0 | 3 | DMS Calibrations | DMS | NO | Closed | SubNC | Ingenuity BigBore |  | WIP18 | June 13, 2019 3:37:26 PM GMT+08:00
CTUDT00623854 | Refresh | Closed | oreg_o | May 20, 2019 9:32:03 PM GMT+08:00 | Refresh: WIP 15: Performance: Tube data mismatch between Summary and X-ray tube data. | WIP15 | WIP18 |  | Subsystem |  | N/A | No Plan to Fix | yatziv_l | 0 | 3 | DMS Calibrations | DMS | NO | Closed | SubNC | IQon1 |  | WIP18 | June 13, 2019 3:37:26 PM GMT+08:00
CTUDT00623852 | Refresh | Closed | g_di | May 20, 2019 8:22:21 PM GMT+08:00 | In Preference, Under ‘Limit Plan geometry’ options are not completely displayed only half of the sentence is displayed and the remaining words are clipped | WIP15 | WIP17 | NA | Subsystem |  | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP18SubCom | Ovation |  | WIP17 | May 30, 2019 5:19:39 PM GMT+08:00
CTUDT00623852 | Refresh | Closed | g_di | May 20, 2019 8:22:21 PM GMT+08:00 | In Preference, Under ‘Limit Plan geometry’ options are not completely displayed only half of the sentence is displayed and the remaining words are clipped | WIP15 | WIP17 | NA | Subsystem |  | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP18SubCom | Brilliance 64 |  | WIP17 | May 30, 2019 5:19:39 PM GMT+08:00
CTUDT00623852 | Refresh | Closed | g_di | May 20, 2019 8:22:21 PM GMT+08:00 | In Preference, Under ‘Limit Plan geometry’ options are not completely displayed only half of the sentence is displayed and the remaining words are clipped | WIP15 | WIP17 | NA | Subsystem |  | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP18SubCom | Ingenuity 64 |  | WIP17 | May 30, 2019 5:19:39 PM GMT+08:00
CTUDT00623852 | Refresh | Closed | g_di | May 20, 2019 8:22:21 PM GMT+08:00 | In Preference, Under ‘Limit Plan geometry’ options are not completely displayed only half of the sentence is displayed and the remaining words are clipped | WIP15 | WIP17 | NA | Subsystem |  | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP18SubCom | IQon1 |  | WIP17 | May 30, 2019 5:19:39 PM GMT+08:00
CTUDT00623850 | Refresh | Closed | Sharmila_Vasireddy | May 20, 2019 8:05:39 PM GMT+08:00 | For Surview Dual, Auto film check box can be enabled, “Auto Film Parameters” are enabled and on selecting “*“in the center field it throws system error. | BB_WIP5.1 | WIP18 |  | Subsystem |  | NA | Fixed | killada_s | 0 | 3 | Console Software | Console4 | NO | Closed | WIP18SubCom | Ovation |  | WIP18 | May 31, 2019 3:08:09 PM GMT+08:00
CTUDT00623850 | Refresh | Closed | Sharmila_Vasireddy | May 20, 2019 8:05:39 PM GMT+08:00 | For Surview Dual, Auto film check box can be enabled, “Auto Film Parameters” are enabled and on selecting “*“in the center field it throws system error. | BB_WIP5.1 | WIP18 |  | Subsystem |  | NA | Fixed | killada_s | 0 | 3 | Console Software | Console4 | NO | Closed | WIP18SubCom | Brilliance 64 |  | WIP18 | May 31, 2019 3:08:09 PM GMT+08:00
CTUDT00623850 | Refresh | Closed | Sharmila_Vasireddy | May 20, 2019 8:05:39 PM GMT+08:00 | For Surview Dual, Auto film check box can be enabled, “Auto Film Parameters” are enabled and on selecting “*“in the center field it throws system error. | BB_WIP5.1 | WIP18 |  | Subsystem |  | NA | Fixed | killada_s | 0 | 3 | Console Software | Console4 | NO | Closed | WIP18SubCom | Ingenuity 64 |  | WIP18 | May 31, 2019 3:08:09 PM GMT+08:00
CTUDT00623850 | Refresh | Closed | Sharmila_Vasireddy | May 20, 2019 8:05:39 PM GMT+08:00 | For Surview Dual, Auto film check box can be enabled, “Auto Film Parameters” are enabled and on selecting “*“in the center field it throws system error. | BB_WIP5.1 | WIP18 |  | Subsystem |  | NA | Fixed | killada_s | 0 | 3 | Console Software | Console4 | NO | Closed | WIP18SubCom | IQon1 |  | WIP18 | May 31, 2019 3:08:09 PM GMT+08:00
CTUDT00623848 | Refresh | Closed | sheik_c | May 20, 2019 7:51:56 PM GMT+08:00 | CIRS subsystem node is not available in System Configuration tool after Encryption | BB_WIP5.1 | BB_WIP9 |  | Subsystem | Encryption | CT-NM.SRS.SecPri.325 | Works As Specified / Not A Defect | wei_f | 1 |  | Verification Documentation | Support-System | NO | Closed |  | Brilliance Big Bore |  | BB_WIP9
CTUDT00623846 | Refresh | Closed | navale_s | May 20, 2019 6:19:18 PM GMT+08:00 | Original EC in Suggested Exam cards group didn't exist after editing | 4.8 | BB_WIP8 | Delete the examcard used for editing from suggested examcard group and create the new examcard. Once the new examcard is created, again add the parent examcard to suggested examcard group. | Subsystem | ECM | N/A | Fixed | Prakash_Chand_Yadav | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP8_SubCom | Brilliance Big Bore |  | BB_WIP8 | July 1, 2019 11:25:04 AM GMT+08:00
CTUDT00623844 | Refresh | Closed | ramesh_t | May 20, 2019 5:55:26 PM GMT+08:00 | Refresh: IQON- The information of the spectral types Electron-Density [%EDW] and Calcium Suppression XXX Index [HU*] are missing in the SDD document. The document has to be updated. | WIP16 | DHF325609 Rev 03 |  | Subsystem |  | NA | Fixed | goswami_s | 0 | 3 | Design Documentation | Console4 | NO | Closed | SegReq | IQon1 |  | DHF325609 Rev 03 | June 4, 2019 8:23:08 PM GMT+08:00
CTUDT00623841 | Refresh | Closed | panchami_pradeep | May 20, 2019 4:13:22 PM GMT+08:00 | Refresh: While extending oblique result, thickness of the extended scan is not as per the scan type document. | WIP16 | DHF318051 Rev 16 |  | Subsystem |  | N/A | Fixed | jindal_k | 0 | 2 | Test Design | Console4 | NO | Closed | SubCom | Ingenuity 64 |  | DHF318051 Rev 16 | June 4, 2019 2:05:11 PM GMT+08:00
CTUDT00623840 | Refresh | Closed | navale_s | May 20, 2019 3:20:43 PM GMT+08:00 | During localization test system crashes. | 4.8 | BB_WIP7 |  | Subsystem | Localization | N/A | Fixed | navale_s | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP7_SubCom | Brilliance Big Bore |  | BB_WIP7 | May 29, 2019 3:38:06 PM GMT+08:00
CTUDT00623837 | Refresh | Closed | Valeria_Kashtanov | May 20, 2019 2:02:49 AM GMT+08:00 | Refresh: Calibrations: Hcor Basic cannot be resumed after cancelling | WIP15 | WIP18 | Restart the calibration | Subsystem | Calibrations | N/A | Fixed | Kostya_Baklan | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP18SegReq | Ovation |  | WIP18 | June 10, 2019 3:06:48 PM GMT+08:00
CTUDT00623834 | Refresh | Closed | finestine_m | May 20, 2019 1:20:04 AM GMT+08:00 | Refresh: Performance: iDose Noise Reduction Check CT Values Are Not The Same In Spec And SW | BB_WIP5.1 | N/A |  | Subsystem | Blocker | N/A | Fixed | braunstein_d | 0 | 2 | Design Documentation | Support-Physics | NO | Closed | WIP18SubCom | Brilliance Big Bore |  | N/A | June 23, 2019 7:11:44 PM GMT+08:00
CTUDT00623832 | Refresh | Closed | beck_d | May 19, 2019 8:45:43 PM GMT+08:00 | Refresh: Calibrations: Performance Infant Test: Harmonized Phantom: Body windowing is not according to spec | WIP15 | WIP17 |  | Subsystem | Calibrations | N/A | Fixed | braunstein_d | 0 | 3 | DMS Calibrations | Support-Physics | NO | Closed | WIP17SubCom | Ovation |  | WIP17 | July 4, 2019 8:12:05 PM GMT+08:00
CTUDT00623832 | Refresh | Closed | beck_d | May 19, 2019 8:45:43 PM GMT+08:00 | Refresh: Calibrations: Performance Infant Test: Harmonized Phantom: Body windowing is not according to spec | WIP15 | WIP17 |  | Subsystem | Calibrations | N/A | Fixed | braunstein_d | 0 | 3 | DMS Calibrations | Support-Physics | NO | Closed | WIP17SubCom | Brilliance Big Bore |  | WIP17 | July 4, 2019 8:12:05 PM GMT+08:00
CTUDT00623830 | Refresh | Closed | beck_d | May 19, 2019 7:41:22 PM GMT+08:00 | Refresh: Calibrations: Performance Test: Body CT1a mAs parameter is not according to spec | WIP15 |  |  | Subsystem |  |  | Duplicate | beck_d | 0 |  |  |  |  | Closed |  | Ovation
CTUDT00623828 | Refresh | Closed | attali_s | May 19, 2019 6:51:16 PM GMT+08:00 | Refresh: Performance Test: The case Body CT2 is missing in the customization window and is not performed | BB_WIP5.1 |  |  | Subsystem | Calibrations | n/A | Works As Specified / Not A Defect | yatziv_l | 0 |  |  | DMS | NO | Closed |  | Brilliance Big Bore
CTUDT00623825 | Refresh | Closed | shterner_a | May 19, 2019 3:33:31 PM GMT+08:00 | RFC : Pre-Val Refresh: In CCT Single mode with Display as Volume length measurement tools are not available after exposure in Acquisition Tab | WIP15 |  | Select Display mode as 1 or 3 the current viewer changes to “CCT” and presentation state to “Acquisition”, according to the change in current viewer i.e. CCT the measurement tools are added. | Subsystem |  | NA | Change Proposal | jayakumar_p | 0 |  |  | Console4 | NO | Closed | SubNC | Ovation
CTUDT00623825 | Refresh | Closed | shterner_a | May 19, 2019 3:33:31 PM GMT+08:00 | RFC : Pre-Val Refresh: In CCT Single mode with Display as Volume length measurement tools are not available after exposure in Acquisition Tab | WIP15 |  | Select Display mode as 1 or 3 the current viewer changes to “CCT” and presentation state to “Acquisition”, according to the change in current viewer i.e. CCT the measurement tools are added. | Subsystem |  | NA | Change Proposal | jayakumar_p | 0 |  |  | Console4 | NO | Closed | SubNC | Brilliance 64
CTUDT00623825 | Refresh | Closed | shterner_a | May 19, 2019 3:33:31 PM GMT+08:00 | RFC : Pre-Val Refresh: In CCT Single mode with Display as Volume length measurement tools are not available after exposure in Acquisition Tab | WIP15 |  | Select Display mode as 1 or 3 the current viewer changes to “CCT” and presentation state to “Acquisition”, according to the change in current viewer i.e. CCT the measurement tools are added. | Subsystem |  | NA | Change Proposal | jayakumar_p | 0 |  |  | Console4 | NO | Closed | SubNC | Brilliance Big Bore
CTUDT00623825 | Refresh | Closed | shterner_a | May 19, 2019 3:33:31 PM GMT+08:00 | RFC : Pre-Val Refresh: In CCT Single mode with Display as Volume length measurement tools are not available after exposure in Acquisition Tab | WIP15 |  | Select Display mode as 1 or 3 the current viewer changes to “CCT” and presentation state to “Acquisition”, according to the change in current viewer i.e. CCT the measurement tools are added. | Subsystem |  | NA | Change Proposal | jayakumar_p | 0 |  |  | Console4 | NO | Closed | SubNC | Ingenuity 64
CTUDT00623825 | Refresh | Closed | shterner_a | May 19, 2019 3:33:31 PM GMT+08:00 | RFC : Pre-Val Refresh: In CCT Single mode with Display as Volume length measurement tools are not available after exposure in Acquisition Tab | WIP15 |  | Select Display mode as 1 or 3 the current viewer changes to “CCT” and presentation state to “Acquisition”, according to the change in current viewer i.e. CCT the measurement tools are added. | Subsystem |  | NA | Change Proposal | jayakumar_p | 0 |  |  | Console4 | NO | Closed | SubNC | IQon1
CTUDT00623823 | Refresh | Closed | shterner_a | May 19, 2019 3:18:02 PM GMT+08:00 | Pre Val Refresh: The Recon Directory stopped responding | WIP16 |  | As a workaround, please click on Refresh when the issue is seen. | Subsystem | Recon Queue Manager | NA | Works As Specified / Not A Defect | s_sa | 1 |  |  | iCAP | NO | Closed |  | Ovation
CTUDT00623823 | Refresh | Closed | shterner_a | May 19, 2019 3:18:02 PM GMT+08:00 | Pre Val Refresh: The Recon Directory stopped responding | WIP16 |  | As a workaround, please click on Refresh when the issue is seen. | Subsystem | Recon Queue Manager | NA | Works As Specified / Not A Defect | s_sa | 1 |  |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00623823 | Refresh | Closed | shterner_a | May 19, 2019 3:18:02 PM GMT+08:00 | Pre Val Refresh: The Recon Directory stopped responding | WIP16 |  | As a workaround, please click on Refresh when the issue is seen. | Subsystem | Recon Queue Manager | NA | Works As Specified / Not A Defect | s_sa | 1 |  |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00623823 | Refresh | Closed | shterner_a | May 19, 2019 3:18:02 PM GMT+08:00 | Pre Val Refresh: The Recon Directory stopped responding | WIP16 |  | As a workaround, please click on Refresh when the issue is seen. | Subsystem | Recon Queue Manager | NA | Works As Specified / Not A Defect | s_sa | 1 |  |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00623823 | Refresh | Closed | shterner_a | May 19, 2019 3:18:02 PM GMT+08:00 | Pre Val Refresh: The Recon Directory stopped responding | WIP16 |  | As a workaround, please click on Refresh when the issue is seen. | Subsystem | Recon Queue Manager | NA | Works As Specified / Not A Defect | s_sa | 1 |  |  | iCAP | NO | Closed |  | IQon1
CTUDT00623821 | Refresh | Closed | shterner_a | May 19, 2019 3:01:22 PM GMT+08:00 | Pre Val Refresh: System Error popped up after initializing the scan | WIP15 |  |  | Subsystem | System |  | Invalid/Uncontrolled Configuration | yatziv_l | 0 |  |  | DMS | NO | Closed |  | IQon1
CTUDT00623819 | Refresh | Closed | shterner_a | May 19, 2019 2:54:03 PM GMT+08:00 | Pre Val Refresh: System Error ,Please end the exam message popped up during Calcium Scoring | WIP16 |  |  | Subsystem | System |  | Invalid/Uncontrolled Configuration | yatziv_l | 0 |  |  | DMS | NO | Closed |  | IQon1
CTUDT00623817 | Refresh | Closed | beck_d | May 19, 2019 2:07:39 PM GMT+08:00 | Refresh: Calibrations: Localization: Phantom centering Help HTML page is not translated correctly | WIP15 | WIP19 |  | Subsystem | Verified1 | N/A | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP19SubCom | Ovation |  | WIP19 | August 6, 2019 10:23:41 PM GMT+08:00
CTUDT00623817 | Refresh | Closed | beck_d | May 19, 2019 2:07:39 PM GMT+08:00 | Refresh: Calibrations: Localization: Phantom centering Help HTML page is not translated correctly | WIP15 | WIP19 |  | Subsystem | Verified1 | N/A | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP19SubCom | Brilliance Big Bore |  | WIP19 | August 6, 2019 10:23:41 PM GMT+08:00
CTUDT00623817 | Refresh | Closed | beck_d | May 19, 2019 2:07:39 PM GMT+08:00 | Refresh: Calibrations: Localization: Phantom centering Help HTML page is not translated correctly | WIP15 | WIP19 |  | Subsystem | Verified1 | N/A | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP19SubCom | Ingenuity 64 |  | WIP19 | August 6, 2019 10:23:41 PM GMT+08:00
CTUDT00623817 | Refresh | Closed | beck_d | May 19, 2019 2:07:39 PM GMT+08:00 | Refresh: Calibrations: Localization: Phantom centering Help HTML page is not translated correctly | WIP15 | WIP19 |  | Subsystem | Verified1 | N/A | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP19SubCom | IQon1 |  | WIP19 | August 6, 2019 10:23:41 PM GMT+08:00
CTUDT00623815 | Refresh | Closed | panchami_pradeep | May 18, 2019 5:29:06 PM GMT+08:00 | Refresh: PSC: While the system is in CT idle mode, packages are not getting downloaded. | WIP16 | BB_WIP8 WIP18 |  | Subsystem | Software Distribution Tool | STPS-SSRS-1187 | Fixed | jose_s | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP18SubCom | Ovation |  | BB_WIP8 WIP18 | June 5, 2019 2:52:44 PM GMT+08:00
CTUDT00623815 | Refresh | Closed | panchami_pradeep | May 18, 2019 5:29:06 PM GMT+08:00 | Refresh: PSC: While the system is in CT idle mode, packages are not getting downloaded. | WIP16 | BB_WIP8 WIP18 |  | Subsystem | Software Distribution Tool | STPS-SSRS-1187 | Fixed | jose_s | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP18SubCom | Brilliance Big Bore |  | BB_WIP8 WIP18 | June 5, 2019 2:52:44 PM GMT+08:00
CTUDT00623815 | Refresh | Closed | panchami_pradeep | May 18, 2019 5:29:06 PM GMT+08:00 | Refresh: PSC: While the system is in CT idle mode, packages are not getting downloaded. | WIP16 | BB_WIP8 WIP18 |  | Subsystem | Software Distribution Tool | STPS-SSRS-1187 | Fixed | jose_s | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP18SubCom | Ingenuity 64 |  | BB_WIP8 WIP18 | June 5, 2019 2:52:44 PM GMT+08:00
CTUDT00623815 | Refresh | Closed | panchami_pradeep | May 18, 2019 5:29:06 PM GMT+08:00 | Refresh: PSC: While the system is in CT idle mode, packages are not getting downloaded. | WIP16 | BB_WIP8 WIP18 |  | Subsystem | Software Distribution Tool | STPS-SSRS-1187 | Fixed | jose_s | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP18SubCom | IQon1 |  | BB_WIP8 WIP18 | June 5, 2019 2:52:44 PM GMT+08:00
CTUDT00623813 | Refresh | Closed | panchami_pradeep | May 18, 2019 4:51:50 PM GMT+08:00 | Refresh: PSC: When the package is browsed from the External Media, while disc space is <11 GB wrong message is displayed. | WIP16 | WIP18 | NA | Subsystem | Service Tool | STPS-SSRS-1292 | Fixed | sagar_v | 0 | 3 | Service Tools_Software | Service Tools | NO | Closed | WIP17SubCom | Ovation |  | WIP18 | August 5, 2019 2:22:02 PM GMT+08:00
CTUDT00623813 | Refresh | Closed | panchami_pradeep | May 18, 2019 4:51:50 PM GMT+08:00 | Refresh: PSC: When the package is browsed from the External Media, while disc space is <11 GB wrong message is displayed. | WIP16 | WIP18 | NA | Subsystem | Service Tool | STPS-SSRS-1292 | Fixed | sagar_v | 0 | 3 | Service Tools_Software | Service Tools | NO | Closed | WIP17SubCom | Brilliance 64 |  | WIP18 | August 5, 2019 2:22:02 PM GMT+08:00
CTUDT00623813 | Refresh | Closed | panchami_pradeep | May 18, 2019 4:51:50 PM GMT+08:00 | Refresh: PSC: When the package is browsed from the External Media, while disc space is <11 GB wrong message is displayed. | WIP16 | WIP18 | NA | Subsystem | Service Tool | STPS-SSRS-1292 | Fixed | sagar_v | 0 | 3 | Service Tools_Software | Service Tools | NO | Closed | WIP17SubCom | Brilliance Big Bore |  | WIP18 | August 5, 2019 2:22:02 PM GMT+08:00
CTUDT00623813 | Refresh | Closed | panchami_pradeep | May 18, 2019 4:51:50 PM GMT+08:00 | Refresh: PSC: When the package is browsed from the External Media, while disc space is <11 GB wrong message is displayed. | WIP16 | WIP18 | NA | Subsystem | Service Tool | STPS-SSRS-1292 | Fixed | sagar_v | 0 | 3 | Service Tools_Software | Service Tools | NO | Closed | WIP17SubCom | Ingenuity 64 |  | WIP18 | August 5, 2019 2:22:02 PM GMT+08:00
CTUDT00623813 | Refresh | Closed | panchami_pradeep | May 18, 2019 4:51:50 PM GMT+08:00 | Refresh: PSC: When the package is browsed from the External Media, while disc space is <11 GB wrong message is displayed. | WIP16 | WIP18 | NA | Subsystem | Service Tool | STPS-SSRS-1292 | Fixed | sagar_v | 0 | 3 | Service Tools_Software | Service Tools | NO | Closed | WIP17SubCom | IQon1 |  | WIP18 | August 5, 2019 2:22:02 PM GMT+08:00
CTUDT00623812 | Refresh | Closed | deshbhratar_m | May 17, 2019 7:18:24 PM GMT+08:00 | Refresh: console user messages localization mismatch | 4.8 | BB_WIP7 | NA | Subsystem | Localization | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP7_SubCom | Brilliance Big Bore |  | BB_WIP7 | May 28, 2019 3:01:07 PM GMT+08:00
CTUDT00623810 | Refresh | Closed | nanjunda_s | May 17, 2019 6:34:24 PM GMT+08:00 | Requirement needs to be update as applications(mentioned in the observations) are available but not license specific. | WIP16 | Host SSRS DHF303211 Rev 12 |  | Subsystem | License | Console-SRS.15142 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | SubCom | Ovation |  | Host SSRS DHF303211 Rev 12 | June 7, 2019 8:02:20 PM GMT+08:00
CTUDT00623810 | Refresh | Closed | nanjunda_s | May 17, 2019 6:34:24 PM GMT+08:00 | Requirement needs to be update as applications(mentioned in the observations) are available but not license specific. | WIP16 | Host SSRS DHF303211 Rev 12 |  | Subsystem | License | Console-SRS.15142 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | SubCom | Brilliance 64 |  | Host SSRS DHF303211 Rev 12 | June 7, 2019 8:02:20 PM GMT+08:00
CTUDT00623810 | Refresh | Closed | nanjunda_s | May 17, 2019 6:34:24 PM GMT+08:00 | Requirement needs to be update as applications(mentioned in the observations) are available but not license specific. | WIP16 | Host SSRS DHF303211 Rev 12 |  | Subsystem | License | Console-SRS.15142 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | SubCom | Ingenuity 64 |  | Host SSRS DHF303211 Rev 12 | June 7, 2019 8:02:20 PM GMT+08:00
CTUDT00623810 | Refresh | Closed | nanjunda_s | May 17, 2019 6:34:24 PM GMT+08:00 | Requirement needs to be update as applications(mentioned in the observations) are available but not license specific. | WIP16 | Host SSRS DHF303211 Rev 12 |  | Subsystem | License | Console-SRS.15142 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | SubCom | IQon1 |  | Host SSRS DHF303211 Rev 12 | June 7, 2019 8:02:20 PM GMT+08:00
CTUDT00623806 | Refresh | Closed | zhang_t | May 17, 2019 3:28:52 PM GMT+08:00 | Refresh ING: Not all images were created using copied partial raw data with specific start and end position from other scanners (same system type) | WIP16 |  |  | Subsystem | Partial raw data copy |  | Duplicate | rakic_k | 0 |  |  | CIRS | NO | Closed |  | Brilliance Big Bore
CTUDT00623806 | Refresh | Closed | zhang_t | May 17, 2019 3:28:52 PM GMT+08:00 | Refresh ING: Not all images were created using copied partial raw data with specific start and end position from other scanners (same system type) | WIP16 |  |  | Subsystem | Partial raw data copy |  | Duplicate | rakic_k | 0 |  |  | CIRS | NO | Closed |  | Ingenuity 64
CTUDT00623802 | Refresh | Closed | Valeria_Kashtanov | May 16, 2019 11:49:15 PM GMT+08:00 | Refresh: Calibrations: ACR Test: ACR Adult Head case fails due to scan initialization failure | WIP15 | WIP18 |  | Subsystem | Calibrations | NA | Fixed | Kostya_Baklan | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP18SubCom | Ingenuity 64 |  | WIP18 | June 6, 2019 3:37:44 PM GMT+08:00
CTUDT00623792 | Refresh | Closed | Valeria_Kashtanov | May 16, 2019 9:16:58 PM GMT+08:00 | Refresh: Calibrations: Hcor Calibration: Error of Phantom misalignment when performing only body cases | BB_WIP5.1 |  |  | Subsystem | Calibrations | NaN | Works As Specified / Not A Defect | yatziv_l | 0 |  |  | DMS | NO | Closed |  | Ovation
CTUDT00623792 | Refresh | Closed | Valeria_Kashtanov | May 16, 2019 9:16:58 PM GMT+08:00 | Refresh: Calibrations: Hcor Calibration: Error of Phantom misalignment when performing only body cases | BB_WIP5.1 |  |  | Subsystem | Calibrations | NaN | Works As Specified / Not A Defect | yatziv_l | 0 |  |  | DMS | NO | Closed |  | Brilliance Big Bore
CTUDT00623792 | Refresh | Closed | Valeria_Kashtanov | May 16, 2019 9:16:58 PM GMT+08:00 | Refresh: Calibrations: Hcor Calibration: Error of Phantom misalignment when performing only body cases | BB_WIP5.1 |  |  | Subsystem | Calibrations | NaN | Works As Specified / Not A Defect | yatziv_l | 0 |  |  | DMS | NO | Closed |  | Ingenuity 64
CTUDT00623792 | Refresh | Closed | Valeria_Kashtanov | May 16, 2019 9:16:58 PM GMT+08:00 | Refresh: Calibrations: Hcor Calibration: Error of Phantom misalignment when performing only body cases | BB_WIP5.1 |  |  | Subsystem | Calibrations | NaN | Works As Specified / Not A Defect | yatziv_l | 0 |  |  | DMS | NO | Closed |  | iCT2
CTUDT00623790 | Refresh | Closed | PremKumar_Rajendran | May 16, 2019 8:29:25 PM GMT+08:00 | BB Refresh : Unnecessary tab ‘Software Distribution’ tool appears in IT login credential | BB_WIP5.1 | DHF316705 Rev10 | N/A | Subsystem | Login | Console-SRS.13133 | Fixed | roy_r | 0 | 2 | HOST SSRS | Console4 | NO | Closed | BB_WIP7_SubCom | Brilliance Big Bore |  | DHF316705 Rev10 | May 27, 2019 4:27:54 PM GMT+08:00
CTUDT00623781 | Refresh | Closed | richards_b | May 16, 2019 10:46:01 AM GMT+08:00 | Bolus scan failed due to watchdog error from CIRS | 3.x and 4.x | WIP19 | N/A | Subsystem | CIRS; Verified1 | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_SeriesProcess | CIRS | YES | Closed | WIP19SubCom | Ovation | 8849486 8084016 | WIP19 | August 16, 2019 9:00:08 PM GMT+08:00
CTUDT00623781 | Refresh | Closed | richards_b | May 16, 2019 10:46:01 AM GMT+08:00 | Bolus scan failed due to watchdog error from CIRS | 3.x and 4.x | WIP19 | N/A | Subsystem | CIRS; Verified1 | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_SeriesProcess | CIRS | YES | Closed | WIP19SubCom | Brilliance 64 | 8849486 8084016 | WIP19 | August 16, 2019 9:00:08 PM GMT+08:00
CTUDT00623781 | Refresh | Closed | richards_b | May 16, 2019 10:46:01 AM GMT+08:00 | Bolus scan failed due to watchdog error from CIRS | 3.x and 4.x | WIP19 | N/A | Subsystem | CIRS; Verified1 | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_SeriesProcess | CIRS | YES | Closed | WIP19SubCom | Brilliance Big Bore | 8849486 8084016 | WIP19 | August 16, 2019 9:00:08 PM GMT+08:00
CTUDT00623781 | Refresh | Closed | richards_b | May 16, 2019 10:46:01 AM GMT+08:00 | Bolus scan failed due to watchdog error from CIRS | 3.x and 4.x | WIP19 | N/A | Subsystem | CIRS; Verified1 | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_SeriesProcess | CIRS | YES | Closed | WIP19SubCom | Ingenuity 64 | 8849486 8084016 | WIP19 | August 16, 2019 9:00:08 PM GMT+08:00
CTUDT00623781 | Refresh | Closed | richards_b | May 16, 2019 10:46:01 AM GMT+08:00 | Bolus scan failed due to watchdog error from CIRS | 3.x and 4.x | WIP19 | N/A | Subsystem | CIRS; Verified1 | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_SeriesProcess | CIRS | YES | Closed | WIP19SubCom | IQon1 | 8849486 8084016 | WIP19 | August 16, 2019 9:00:08 PM GMT+08:00
CTUDT00623775 | Refresh | Closed | cohen_y | May 16, 2019 4:26:24 AM GMT+08:00 | Refresh: CTDI Free Air: Windowing of measurements images are not as defined in spec | WIP15 | WIP19 |  | Subsystem | CTDI free air measurement tool;Verified1 | NaN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP17SubCom | Ovation |  | WIP19 | July 31, 2019 8:33:04 PM GMT+08:00
CTUDT00623775 | Refresh | Closed | cohen_y | May 16, 2019 4:26:24 AM GMT+08:00 | Refresh: CTDI Free Air: Windowing of measurements images are not as defined in spec | WIP15 | WIP19 |  | Subsystem | CTDI free air measurement tool;Verified1 | NaN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP17SubCom | Ingenuity 64 |  | WIP19 | July 31, 2019 8:33:04 PM GMT+08:00
CTUDT00623773 | Refresh | Closed | cohen_y | May 16, 2019 2:05:25 AM GMT+08:00 | Refresh: Performance Test: IMR Cases does not appear in Customization window even though IMR is enabled in license key | WIP15 | WIP17 |  | Subsystem | IMR | N/A | Fixed | Kostya_Baklan | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP17SubCom | Ingenuity 64 |  | WIP17 | May 26, 2019 4:50:32 PM GMT+08:00
CTUDT00623770 | Refresh | Closed | cohen_y | May 16, 2019 12:03:43 AM GMT+08:00 | Refresh: DMS/DataPath Diagnostics: Acquiring CIRS Event log error appears during Data-Path test execution | WIP15 | WIP17 |  | Subsystem | Service Tools | DMS-SRS-658 | Fixed | franco_a | 0 | 2 | DMS Software | DMS | NO | Closed | WIP17SubCom | Ingenuity 64 |  | WIP17 | May 22, 2019 8:33:09 PM GMT+08:00
CTUDT00623769 | Refresh | Closed | liu_bo | May 15, 2019 11:47:27 PM GMT+08:00 | [Refresh-iCT\iQon Gantry]System cannot prompt error message "Pin Offset from isocenter is too large or pin is not in the field of view." during Tube alignment calibration. | WIP16 |  |  | Subsystem |  | n/a | Works As Specified / Not A Defect | zhou_x | 0 | 2 |  | GCX | NO | Closed | WIP17SegReq | Ovation
CTUDT00623769 | Refresh | Closed | liu_bo | May 15, 2019 11:47:27 PM GMT+08:00 | [Refresh-iCT\iQon Gantry]System cannot prompt error message "Pin Offset from isocenter is too large or pin is not in the field of view." during Tube alignment calibration. | WIP16 |  |  | Subsystem |  | n/a | Works As Specified / Not A Defect | zhou_x | 0 | 2 |  | GCX | NO | Closed | WIP17SegReq | IQon1
CTUDT00623767 | Refresh | Closed | g_di | May 15, 2019 10:04:15 PM GMT+08:00 | An error message is displayed upon loading result from patient directory to View2 for all the localization language except language | WIP15 |  |  | Subsystem | Patient Directory | NA | Duplicate | goswami_s | 0 |  |  | Console4 | NO | Closed | WIP17SubCom | Ingenuity 64
CTUDT00623763 | Refresh | Closed | Valeria_Kashtanov | May 15, 2019 8:44:41 PM GMT+08:00 | Refresh: Calibrations: Steps Test: Error appears during Steps Test | WIP15 | WIP18 |  | Subsystem |  | NaN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP17SegReq | Ovation |  | WIP18 | June 17, 2019 10:05:10 PM GMT+08:00
CTUDT00623756 | Refresh | Closed | Valeria_Kashtanov | May 15, 2019 6:48:28 PM GMT+08:00 | Refresh: Calibrations: Performance Test: Wrong mAs scan parameter for Body CT1a case | WIP15 | WIP17 |  | Subsystem | Calibrations | NA | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP17SubCom | Ovation |  | WIP17 | May 28, 2019 2:44:07 PM GMT+08:00
CTUDT00623755 | Refresh | Closed | attali_s | May 15, 2019 6:40:30 PM GMT+08:00 | Refresh: Performance Test: Wrong Couch Height in Performance Infant | WIP15 | WIP18 | N/A | Subsystem |  | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP18SubCom | Ovation |  | WIP18 | June 13, 2019 3:10:35 PM GMT+08:00
CTUDT00623754 | Refresh | Closed | prabhu_m | May 15, 2019 5:27:48 PM GMT+08:00 | As per requirement Use previous surview ID expected latency shows incorrect value | V4.8 | BB_WIP7 |  | Subsystem | Plan | Console-SRS.10103 | Fixed | kumar-san | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP7_SubCom | Brilliance Big Bore |  | BB_WIP7 | May 23, 2019 5:43:05 PM GMT+08:00
CTUDT00623747 | Refresh | Closed | bhawna_g | May 15, 2019 5:10:31 PM GMT+08:00 | Refresh: “System.cfg” file is modified with no apparent reason and prevent the machine functionality | 4.8 | BB_WIP7 |  | Subsystem | Installation | NA | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP7_SubCom | Brilliance Big Bore |  | BB_WIP7 | May 22, 2019 7:56:05 PM GMT+08:00
CTUDT00623743 | Refresh | Closed | finestine_m | May 15, 2019 4:04:15 PM GMT+08:00 | Refresh: Common Functions: Issues with Offline Reconstruction | WIP15 |  |  | Subsystem | Blocker |  | Invalid/Uncontrolled Configuration | rakic_k | 0 |  |  | CIRS | NO | Closed |  | Brilliance Big Bore
CTUDT00623743 | Refresh | Closed | finestine_m | May 15, 2019 4:04:15 PM GMT+08:00 | Refresh: Common Functions: Issues with Offline Reconstruction | WIP15 |  |  | Subsystem | Blocker |  | Invalid/Uncontrolled Configuration | rakic_k | 0 |  |  | CIRS | NO | Closed |  | Ingenuity 16
CTUDT00623743 | Refresh | Closed | finestine_m | May 15, 2019 4:04:15 PM GMT+08:00 | Refresh: Common Functions: Issues with Offline Reconstruction | WIP15 |  |  | Subsystem | Blocker |  | Invalid/Uncontrolled Configuration | rakic_k | 0 |  |  | CIRS | NO | Closed |  | IQon1
CTUDT00623738 | Refresh | Closed | divya_n | May 15, 2019 3:13:28 PM GMT+08:00 | Patient details are displayed with invalid characters when a study with char set ISO-IR 126 and ISO-IR 144 is imported to study grid and launched into sub selection. | WIP16 |  |  | Subsystem | Patient Directory | NA | Does Not Apply | s_sa | 0 |  |  | iCAP | NO | Closed |  | Ovation
CTUDT00623738 | Refresh | Closed | divya_n | May 15, 2019 3:13:28 PM GMT+08:00 | Patient details are displayed with invalid characters when a study with char set ISO-IR 126 and ISO-IR 144 is imported to study grid and launched into sub selection. | WIP16 |  |  | Subsystem | Patient Directory | NA | Does Not Apply | s_sa | 0 |  |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00623738 | Refresh | Closed | divya_n | May 15, 2019 3:13:28 PM GMT+08:00 | Patient details are displayed with invalid characters when a study with char set ISO-IR 126 and ISO-IR 144 is imported to study grid and launched into sub selection. | WIP16 |  |  | Subsystem | Patient Directory | NA | Does Not Apply | s_sa | 0 |  |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00623738 | Refresh | Closed | divya_n | May 15, 2019 3:13:28 PM GMT+08:00 | Patient details are displayed with invalid characters when a study with char set ISO-IR 126 and ISO-IR 144 is imported to study grid and launched into sub selection. | WIP16 |  |  | Subsystem | Patient Directory | NA | Does Not Apply | s_sa | 0 |  |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00623738 | Refresh | Closed | divya_n | May 15, 2019 3:13:28 PM GMT+08:00 | Patient details are displayed with invalid characters when a study with char set ISO-IR 126 and ISO-IR 144 is imported to study grid and launched into sub selection. | WIP16 |  |  | Subsystem | Patient Directory | NA | Does Not Apply | s_sa | 0 |  |  | iCAP | NO | Closed |  | IQon1
CTUDT00623734 | Refresh | Closed | sowmya_k | May 15, 2019 1:25:10 PM GMT+08:00 | Val: "ErrorExamCardManager" Application Error While Creating User Exam Card | BB_WIP6 | BB_WIP7 |  | Subsystem | ECM | NA | Fixed | chokkalingam_r | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP7_SubCom | Brilliance Big Bore |  | BB_WIP7 | May 27, 2019 6:08:09 PM GMT+08:00
CTUDT00623730 | Refresh | Closed | gT_s | May 15, 2019 10:51:10 AM GMT+08:00 | Refresh : A new version of software is available for installation notification is displayed even after the package has been installed | WIP16 | WIP17 BB_WIP7 |  | Subsystem | Software Distribution | STPS-SSRS-1263 | Fixed | sagar_v | 0 | 4 | Console Software | Service Tools | NO | Closed | WIP17SubCom | Ovation |  | WIP17 BB_WIP7 | May 20, 2019 8:29:38 PM GMT+08:00
CTUDT00623730 | Refresh | Closed | gT_s | May 15, 2019 10:51:10 AM GMT+08:00 | Refresh : A new version of software is available for installation notification is displayed even after the package has been installed | WIP16 | WIP17 BB_WIP7 |  | Subsystem | Software Distribution | STPS-SSRS-1263 | Fixed | sagar_v | 0 | 4 | Console Software | Service Tools | NO | Closed | WIP17SubCom | Brilliance 64 |  | WIP17 BB_WIP7 | May 20, 2019 8:29:38 PM GMT+08:00
CTUDT00623730 | Refresh | Closed | gT_s | May 15, 2019 10:51:10 AM GMT+08:00 | Refresh : A new version of software is available for installation notification is displayed even after the package has been installed | WIP16 | WIP17 BB_WIP7 |  | Subsystem | Software Distribution | STPS-SSRS-1263 | Fixed | sagar_v | 0 | 4 | Console Software | Service Tools | NO | Closed | WIP17SubCom | Brilliance Big Bore |  | WIP17 BB_WIP7 | May 20, 2019 8:29:38 PM GMT+08:00
CTUDT00623730 | Refresh | Closed | gT_s | May 15, 2019 10:51:10 AM GMT+08:00 | Refresh : A new version of software is available for installation notification is displayed even after the package has been installed | WIP16 | WIP17 BB_WIP7 |  | Subsystem | Software Distribution | STPS-SSRS-1263 | Fixed | sagar_v | 0 | 4 | Console Software | Service Tools | NO | Closed | WIP17SubCom | Ingenuity 64 |  | WIP17 BB_WIP7 | May 20, 2019 8:29:38 PM GMT+08:00
CTUDT00623730 | Refresh | Closed | gT_s | May 15, 2019 10:51:10 AM GMT+08:00 | Refresh : A new version of software is available for installation notification is displayed even after the package has been installed | WIP16 | WIP17 BB_WIP7 |  | Subsystem | Software Distribution | STPS-SSRS-1263 | Fixed | sagar_v | 0 | 4 | Console Software | Service Tools | NO | Closed | WIP17SubCom | IQon1 |  | WIP17 BB_WIP7 | May 20, 2019 8:29:38 PM GMT+08:00
CTUDT00623727 | Refresh | Closed | oreg_o | May 14, 2019 8:52:24 PM GMT+08:00 | Refresh: "Tube Current, Scan Time, mAs Test", the RT parameter in the requirement is not match to the parameters of ingenuity machine. | WIP15 | N/A |  | Subsystem | Blocker | N/A | Fixed | braunstein_d | 0 | 2 | Design Documentation | Support-Physics | NO | Closed | WIP17SubCom | Ingenuity 64 |  | N/A | June 13, 2019 2:59:46 PM GMT+08:00
CTUDT00623724 | Refresh | Closed | oreg_o | May 14, 2019 8:18:33 PM GMT+08:00 | Refresh: Temperature_Stabilization : the temperature stabilization run for full time and not as defined | WIP15 | N/A |  | Subsystem |  | N/A | Fixed | yosfan_s | 0 | 2 | Test Design | DMS | NO | Closed | WIP18SubCom | Ingenuity 64 |  | N/A | June 16, 2019 3:52:21 PM GMT+08:00
CTUDT00623723 | Refresh | Closed | dadon_e | May 14, 2019 8:16:53 PM GMT+08:00 | Refresh, Pre-Val (WIP16): IQON, Beam and DMS Check calibration failed – unstable channels error | WIP16 |  |  | Subsystem | Calibrations | N/A | Works As Specified / Not A Defect | berent_i | 0 |  |  | DMS | NO | Closed |  | IQon1
CTUDT00623721 | Refresh | Closed | bhawna_g | May 14, 2019 8:12:02 PM GMT+08:00 | Support for Import of Varian wave file with version “1.7” and string change in UI from “RPM” to “Varian” | WIP16 | WIP16 |  | Subsystem | Pulmo | Console-SRS.1530 | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | YES | Closed | WIP17SubCom | Brilliance 64 |  | WIP16 | May 22, 2019 8:15:17 PM GMT+08:00
CTUDT00623721 | Refresh | Closed | bhawna_g | May 14, 2019 8:12:02 PM GMT+08:00 | Support for Import of Varian wave file with version “1.7” and string change in UI from “RPM” to “Varian” | WIP16 | WIP16 |  | Subsystem | Pulmo | Console-SRS.1530 | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | YES | Closed | WIP17SubCom | Ingenuity 64 |  | WIP16 | May 22, 2019 8:15:17 PM GMT+08:00
CTUDT00623719 | Refresh | Closed | dadon_e | May 14, 2019 8:02:11 PM GMT+08:00 | Refresh, Pre-Val (WIP16): IQON, performance test calibration failed - The performance test completed but failed | WIP16 |  |  | Subsystem | Calibrations |  | Works As Specified / Not A Defect | tamari_o | 1 |  |  | Support-System | NO | Closed |  | IQon1
CTUDT00623717 | Refresh | Closed | dadon_e | May 14, 2019 7:41:36 PM GMT+08:00 | Refresh Pre-Val:(WIP16 - IQON) HCOR slice infant failed - error during scan | WIP16 | WIP18 |  | Subsystem | Calibrations | n/a | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP18SegReq | Ovation |  | WIP18 | June 5, 2019 8:10:43 PM GMT+08:00
CTUDT00623717 | Refresh | Closed | dadon_e | May 14, 2019 7:41:36 PM GMT+08:00 | Refresh Pre-Val:(WIP16 - IQON) HCOR slice infant failed - error during scan | WIP16 | WIP18 |  | Subsystem | Calibrations | n/a | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP18SegReq | IQon1 |  | WIP18 | June 5, 2019 8:10:43 PM GMT+08:00
CTUDT00623715 | Refresh | Closed | liu_lc | May 14, 2019 7:22:19 PM GMT+08:00 | [Refresh-iCT\iQon Gantry]The Control panel display is not match the requirement | WIP16 | Gantry SSRS V04 |  | Subsystem | Requirement | GCX.SSRS.iCT.20819 | Fixed | bruce_li | 0 | 3 | Requirements Documentation | GCX | NO | Closed | WIP19SegReq | Ovation |  | Gantry SSRS V04 | July 3, 2019 9:43:52 AM GMT+08:00
CTUDT00623713 | Refresh | Closed | liu_lc | May 14, 2019 7:19:02 PM GMT+08:00 | [Refresh-iCT\iQon Gantry]The color in iCT BOM is not match the CT Box requirement | WIP16 |  |  | Subsystem | System |  | Does Not Apply | dudid | 0 |  |  | Support-System | NO | Closed |  | Ovation
CTUDT00623711 | Refresh | Closed | liu_lc | May 14, 2019 7:15:33 PM GMT+08:00 | [Refresh-iCT\iQon Gantry]Labelling on the CT Control Box doesn't identify the location of the Symbols Glossary and the translation language type is not enough | WIP16 | WIP18 |  | Subsystem | System | GCX.SSRS.iCT.20462 | Fixed | 320025721 | 0 | 2 | Gantry Electrical | GCX | NO | Closed | WIP17SegReq | Ovation |  | WIP18 | July 10, 2019 5:33:57 PM GMT+08:00
CTUDT00623708 | Refresh | Closed | liu_lc | May 14, 2019 7:12:21 PM GMT+08:00 | [Refresh-iCT\iQon Gantry]The parameter on the Compressor Seismic label is not consisted with the preapprove test report on the label | WIP16 | WIP20 |  | Subsystem | System | GCX.SSRS.iCT.20286;GCX.SSRS.iCT.20293 | Fixed | april_z | 0 | 2 | Gantry Mechanical | GCX | NO | Closed | WIP19SegReq | Ovation |  | WIP20 | July 30, 2019 12:54:11 PM GMT+08:00
CTUDT00623706 | Refresh | Closed | liu_lc | May 14, 2019 7:08:00 PM GMT+08:00 | [Refresh-iCT\iQon Gantry]Not all calibration table could be backup and resotred | WIP16 |  |  | Subsystem | Backup/Restore | n/a | Works As Specified / Not A Defect | oren_d | 1 |  |  | GCX | NO | Closed |  | Ovation
CTUDT00623704 | Refresh | Closed | liu_lc | May 14, 2019 7:00:41 PM GMT+08:00 | [Refresh-iCT\iQon Gantry]The parameter on the Cabinet Compressor Seismic label is not consisted with the preapprove test report on the label | WIP16 | WIP20 |  | Subsystem | System | GCX.SSRS.iCT.20299, GCX.SSRS.iCT.20294 | Fixed | april_z | 0 | 2 | Gantry Mechanical | GCX | NO | Closed | WIP19SegReq | Ovation |  | WIP20 | July 30, 2019 12:51:01 PM GMT+08:00
CTUDT00623702 | Refresh | Closed | liu_lc | May 14, 2019 6:57:09 PM GMT+08:00 | [Refresh-iCT\iQon Gantry]FS X Position Calibration parameter is different with requirement | WIP16 | WIP18 |  | Subsystem | Calibrations | n/a | Fixed | zhou_x | 0 | 2 | Gantry Software | GCX | YES | Closed | WIP18SegReq | Ovation |  | WIP18 | June 20, 2019 4:42:09 PM GMT+08:00
CTUDT00623697 | Refresh | Closed | liu_lc | May 14, 2019 6:37:44 PM GMT+08:00 | [Refresh-iCT\iQon Gantry]FS X Position Calibration and UHR FS X Position Calibration mechanism is only FWHM bisector | WIP16 | WIP18 |  | Subsystem | Calibrations | N/A | Fixed | zhou_x | 0 | 2 | Gantry Calibrations | GCX | NO | Closed | WIP18SegReq | Ovation |  | WIP18 | June 25, 2019 10:16:05 AM GMT+08:00
CTUDT00623695 | Refresh | Closed | liu_bo | May 14, 2019 6:33:35 PM GMT+08:00 | [Refresh-iCT\iQon Gantry] Perform A-Plane Opening calibration failed on iCTSP and iQon | WIP16 | WIP17 |  | Subsystem | Calibrations | N/A | Fixed | zhou_x | 0 | 2 | Gantry Calibrations | GCX | NO | Closed | WIP17SegReq | Ovation |  | WIP17 | May 31, 2019 9:36:38 AM GMT+08:00
CTUDT00623695 | Refresh | Closed | liu_bo | May 14, 2019 6:33:35 PM GMT+08:00 | [Refresh-iCT\iQon Gantry] Perform A-Plane Opening calibration failed on iCTSP and iQon | WIP16 | WIP17 |  | Subsystem | Calibrations | N/A | Fixed | zhou_x | 0 | 2 | Gantry Calibrations | GCX | NO | Closed | WIP17SegReq | IQon1 |  | WIP17 | May 31, 2019 9:36:38 AM GMT+08:00
CTUDT00623693 | Refresh | Closed | liu_bo | May 14, 2019 6:29:47 PM GMT+08:00 | [Refresh-iCT\iQon Gantry] The below reason matrix is different with requirement. DMS Replacement, New Site Installation,New System Build, Resolver Replacement, Routine Scheduled Calibration | WIP16 | Gantry SSRS V04 |  | Subsystem | Calibrations | GCX.SSRS.iCT.11994 | Fixed | bruce_li | 0 | 2 | Requirements Documentation | GCX | NO | Closed | WIP18SegReq | Ovation |  | Gantry SSRS V04 | July 4, 2019 7:52:16 AM GMT+08:00
CTUDT00623691 | Refresh | Closed | liu_bo | May 14, 2019 6:26:40 PM GMT+08:00 | [Refresh-iCT\iQon Gantry] The below reason matrix is different with requirement. Generator Replacement, New site Installation,Routine Schedule Calibration | WIP16 | Gantry SSRS V04 |  | Subsystem | Calibrations | GCX.SSRS.iCT.19246 | Fixed | bruce_li | 0 | 2 | Requirements Documentation | GCX | NO | Closed | WIP18SegReq | IQon1 |  | Gantry SSRS V04 | July 4, 2019 7:51:48 AM GMT+08:00
CTUDT00623689 | Refresh | Closed | dadon_e | May 14, 2019 6:18:43 PM GMT+08:00 | Refresh, Pre-Val (WIP16): IQON, A-plane calibration failed - calculation error | WIP16 | WIP17 | N/A | Subsystem | Calibrations | N/A | Fixed | zhou_x | 0 | 2 | Gantry Calibrations | GCX | NO | Closed | WIP17SegReq | IQon1 |  | WIP17 | May 31, 2019 9:34:28 AM GMT+08:00
CTUDT00623688 | Refresh | Closed | liu_bo | May 14, 2019 6:09:17 PM GMT+08:00 | [Refresh-iCT\iQon Gantry] philips_service user and dc_user doesn't have Reliability test in Quality Assurance. Philips_opt and dc_user doesn’t have Advanced Mode in Quality Assurance. | WIP16 | Gantry SSRS V04 |  | Subsystem | Calibrations | n/a | Fixed | bruce_li | 0 | 2 | Requirements Documentation | GCX | NO | Closed | WIP17SegReq | Ovation |  | Gantry SSRS V04 | July 4, 2019 7:51:20 AM GMT+08:00
CTUDT00623685 | Refresh | Closed | liu_bo | May 14, 2019 5:39:48 PM GMT+08:00 | [Refresh-iCT\iQon Gantry] philips_opt and dc_user doesn’t have Advanced Mode in Quality Assurance. | WIP16 | Gantry SSRS V04 |  | Subsystem | Calibrations | GCX.SSRS.iCT.19247 | Fixed | bruce_li | 0 | 2 | Requirements Documentation | GCX | NO | Closed | WIP18SegReq | IQon1 |  | Gantry SSRS V04 | July 4, 2019 7:50:58 AM GMT+08:00
CTUDT00623683 | Refresh | Closed | Sharmila_Vasireddy | May 14, 2019 5:11:16 PM GMT+08:00 | BB refresh: Occurred system error when performing the CCT scan with 255 cycles | BB_WIP5.1 |  |  | Subsystem | CCT | NA | Works As Specified / Not A Defect | rath_s | 0 |  |  | Console4 | NO | Closed | WAS | Brilliance Big Bore
CTUDT00623679 | Refresh | Closed | wang_l | May 14, 2019 10:31:22 AM GMT+08:00 | Refresh BR64: The backup Voice language is not upgraded | WIP15 |  |  | Subsystem | Upgrade | NA | Works As Specified / Not A Defect | wang_l | 1 | 3 | Console Software | Support-System | NO | Closed |  | Ovation
CTUDT00623679 | Refresh | Closed | wang_l | May 14, 2019 10:31:22 AM GMT+08:00 | Refresh BR64: The backup Voice language is not upgraded | WIP15 |  |  | Subsystem | Upgrade | NA | Works As Specified / Not A Defect | wang_l | 1 | 3 | Console Software | Support-System | NO | Closed |  | Brilliance 64
CTUDT00623679 | Refresh | Closed | wang_l | May 14, 2019 10:31:22 AM GMT+08:00 | Refresh BR64: The backup Voice language is not upgraded | WIP15 |  |  | Subsystem | Upgrade | NA | Works As Specified / Not A Defect | wang_l | 1 | 3 | Console Software | Support-System | NO | Closed |  | Ingenuity 64
CTUDT00623679 | Refresh | Closed | wang_l | May 14, 2019 10:31:22 AM GMT+08:00 | Refresh BR64: The backup Voice language is not upgraded | WIP15 |  |  | Subsystem | Upgrade | NA | Works As Specified / Not A Defect | wang_l | 1 | 3 | Console Software | Support-System | NO | Closed |  | IQon1
CTUDT00623675 | Refresh | Closed | dadon_e | May 13, 2019 10:56:06 PM GMT+08:00 | Refresh4.8 Usability – IV SW: Needle Position radio Button is not selected when needle position is selected from the selection list | BB_WIP4 |  |  | Subsystem | IV Controls |  | Change Proposal | mistry_h | 0 |  |  | Support-System | NO | Closed |  | Brilliance Big Bore
CTUDT00623673 | Refresh | Closed | dadon_e | May 13, 2019 10:51:57 PM GMT+08:00 | Refresh4.8 Usability – IV HW: mouse responsiveness is not good | BB_WIP4 | WIP19 |  | Subsystem | IV Controls | DHF317353, Req ID :GCX.SSRS.iCT.20607 | Duplicate | teng_h | 0 | 3 | Gantry Electrical | GCX | NO | Closed | BB_WIP8_SubCom | Brilliance Big Bore |  | WIP19 | August 2, 2019 1:19:58 PM GMT+08:00
CTUDT00623671 | Refresh | Closed | dadon_e | May 13, 2019 10:50:23 PM GMT+08:00 | Refresh4.8 Usability – IV HW: Touch Panel responsiveness is not good | BB_WIP4 | WIP19 |  | Subsystem | IVC Touch Panel | DHF317353, Req ID :GCX.SSRS.iCT.20213 | Fixed | teng_h | 0 | 3 | Gantry Electrical | GCX | NO | Closed | BB_WIP8_SubCom | Brilliance Big Bore |  | WIP19 | August 7, 2019 5:35:19 PM GMT+08:00
CTUDT00623669 | Refresh | Closed | dadon_e | May 13, 2019 10:48:29 PM GMT+08:00 | Refresh4.8 Usability – IV SW: multiple needles - first shot scan position is always needle 1 position, and is inconsistent with repeating shots which relate to bed position. | BB_WIP4 | BB_WIP9 | User can manually update the plan box to match the requried needle position for the first shot. | Subsystem | IV Controls | Console-SRS.15926 | Fixed | sowmya_k | 0 | 3 | Console Software | Console4 | YES | Closed | BB_WIP9_SubCom | Brilliance Big Bore |  | BB_WIP9 | August 8, 2019 5:58:19 PM GMT+08:00
CTUDT00623667 | Refresh | Closed | dadon_e | May 13, 2019 10:46:36 PM GMT+08:00 | Refresh4.8 Usability – IV SW: needle number when opening “New Needle Position” dialog does not support new needle | BB_WIP4 | BB_WIP9 | User can chose to edit the needles before starting interventional scans | Subsystem | IV Controls | N/A | Fixed | sowmya_k | 0 | 3 | Console Software | Console4 | YES | Closed | BB_WIP9_SubCom | Brilliance Big Bore |  | BB_WIP9 | August 2, 2019 8:48:54 PM GMT+08:00
CTUDT00623665 | Refresh | Closed | dadon_e | May 13, 2019 10:44:40 PM GMT+08:00 | Refresh4.8 Usability – IV SW: “Save” on “New Needle Position” dialog does not save the location which is in view upon closing the dialog | BB_WIP4 | BB_WIP9 | User can choose to edit the wrongly saved needle | Subsystem | IV Controls | N/A | Fixed | navale_s | 0 | 3 | Console Software | Console4 | YES | Closed | BB_WIP9_SubCom | Brilliance Big Bore |  | BB_WIP9 | August 8, 2019 9:19:46 PM GMT+08:00
CTUDT00623663 | Refresh | Closed | dadon_e | May 13, 2019 10:41:57 PM GMT+08:00 | Refresh4.8 Usability – IV SW: New Needles dialog : “Save” and “Cancel” location on screen might cause to clicking not the intended button | BB_WIP4 | BB_WIP7 | N/A | Subsystem | IV Controls | Console-SRS.15815 | Fixed | b_bibin | 0 | 5 | Console Software | Console4 | NO | Closed | BB_WIP7_SubCom | Brilliance Big Bore |  | BB_WIP7 | May 30, 2019 5:05:45 PM GMT+08:00
CTUDT00623661 | Refresh | Closed | dadon_e | May 13, 2019 10:39:33 PM GMT+08:00 | Refresh4.8 Usability – IV SW: Edit Needles Position dialog does not have explicit “Save” button | BB_WIP4 | BB_WIP8 | N/A | Subsystem | IV Controls | Console-SRS.15814 | Fixed | Snehashis_Dey | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP7_SubCom | Brilliance Big Bore |  | BB_WIP8 | July 5, 2019 7:14:34 PM GMT+08:00
CTUDT00623659 | Refresh | Closed | dadon_e | May 13, 2019 10:37:22 PM GMT+08:00 | Refresh4.8 Usability – IV HW & SW: symbols are not consistent between HW & SW | BB_WIP4 | BB_WIP8 | N/A | Subsystem | Interventional Control | Console-SRS.15708 | Fixed | Snehashis_Dey | 0 | 5 | Console Software | Console4 | NO | Closed | BB_WIP8_SubCom | Brilliance Big Bore |  | BB_WIP8 | June 25, 2019 8:20:15 PM GMT+08:00
CTUDT00623657 | Refresh | Closed | dadon_e | May 13, 2019 10:34:29 PM GMT+08:00 | Refresh4.8 Usability – IV HW: Position symbols are not intuitive | BB_WIP4 |  | None | Subsystem | Gantry | N/A | Works As Specified / Not A Defect | ye_h | 0 |  |  | Support-System | NO | Closed | BB_SubNC | Brilliance Big Bore
CTUDT00623655 | Refresh | Closed | dadon_e | May 13, 2019 10:30:46 PM GMT+08:00 | Refresh4.8 Usability – IV HW: Using “Annotation” tool for measurement using the Joystick mouse is very hard to do | BB_WIP4 |  |  | Subsystem | IV Controls | FOK | Duplicate | wang_k | 0 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00623652 | Refresh | Closed | dadon_e | May 13, 2019 10:27:32 PM GMT+08:00 | Refresh4.8 Usability – IV SW: button of “move to selected position” is not intuitive as continuous press | BB_WIP4 |  |  | Subsystem | IV Controls |  | Change Proposal | mistry_h | 0 |  |  | Support-System | NO | Closed |  | Brilliance Big Bore
CTUDT00623642 | Refresh | Closed | wei_f | May 13, 2019 7:48:26 PM GMT+08:00 | on service tool Remote viewer, double click the server name, pop up an error message. | WIP16 |  |  | Subsystem | Remote Viewer | N/A | Works As Specified / Not A Defect | prasad_r | 0 |  |  | Service Tools | NO | Closed | WIP18SubCom | Ovation
CTUDT00623642 | Refresh | Closed | wei_f | May 13, 2019 7:48:26 PM GMT+08:00 | on service tool Remote viewer, double click the server name, pop up an error message. | WIP16 |  |  | Subsystem | Remote Viewer | N/A | Works As Specified / Not A Defect | prasad_r | 0 |  |  | Service Tools | NO | Closed | WIP18SubCom | Brilliance 64
CTUDT00623642 | Refresh | Closed | wei_f | May 13, 2019 7:48:26 PM GMT+08:00 | on service tool Remote viewer, double click the server name, pop up an error message. | WIP16 |  |  | Subsystem | Remote Viewer | N/A | Works As Specified / Not A Defect | prasad_r | 0 |  |  | Service Tools | NO | Closed | WIP18SubCom | Brilliance Big Bore
CTUDT00623642 | Refresh | Closed | wei_f | May 13, 2019 7:48:26 PM GMT+08:00 | on service tool Remote viewer, double click the server name, pop up an error message. | WIP16 |  |  | Subsystem | Remote Viewer | N/A | Works As Specified / Not A Defect | prasad_r | 0 |  |  | Service Tools | NO | Closed | WIP18SubCom | Ingenuity 64
CTUDT00623642 | Refresh | Closed | wei_f | May 13, 2019 7:48:26 PM GMT+08:00 | on service tool Remote viewer, double click the server name, pop up an error message. | WIP16 |  |  | Subsystem | Remote Viewer | N/A | Works As Specified / Not A Defect | prasad_r | 0 |  |  | Service Tools | NO | Closed | WIP18SubCom | IQon1
CTUDT00623639 | Refresh | Closed | navale_s | May 13, 2019 6:33:27 PM GMT+08:00 | Console Application closed and r2hImageReady could not send | 4.8 | BB_WIP7 |  | Subsystem | System | N/A | Fixed | navale_s | 0 | 2 | Console Software | Console4 | YES | Closed | BB_WIP7_SubCom | Brilliance Big Bore |  | BB_WIP7 | May 21, 2019 4:26:54 PM GMT+08:00
CTUDT00623637 | Refresh | Closed | zhang_t | May 13, 2019 6:13:10 PM GMT+08:00 | Refresh: Calibrations: Performance Infant Test - Uniformity Head failed | WIP14 | WIP18 |  | Subsystem | Calibrations | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP17SubCom | Ingenuity 64 |  | WIP18 | June 25, 2019 9:24:36 PM GMT+08:00
CTUDT00623621 | Refresh | Closed | ravindra_a | May 10, 2019 5:54:59 PM GMT+08:00 | Loading exam card takes more time | 4.8 | BB_WIP12.2 | No workaround | Subsystem | Exam Summary | NA | Fixed | kapadia_v | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP8_SubCom | Brilliance Big Bore |  | BB_WIP12.2 | November 29, 2019 4:33:46 PM GMT+08:00
CTUDT00623618 | Refresh | Closed | panchami_pradeep | May 10, 2019 2:06:32 PM GMT+08:00 | Refresh: Localization: System crash while trying to Add or Replace EC after surview in Languages other than English. | WIP16 |  |  | Subsystem | Localization | NA | Duplicate | navale_s | 0 | 2 |  | Console4 | YES | Closed | WIP17SubCom | Ingenuity 64
CTUDT00623617 | Refresh | Closed | vinoliya_j | May 10, 2019 2:01:23 PM GMT+08:00 | Refresh: When the ECG wave is edited for a result, but recon is not done, the wave was not reverting to its original wave. | WIP15 |  |  | Subsystem | ECG | NA | Works As Specified / Not A Defect | sowmya_k | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00623613 | Refresh | Closed | panchami_pradeep | May 10, 2019 12:22:21 PM GMT+08:00 | Refresh: Localization: When the console is set up for Norwegian language, right monitor is displayed in English. | WIP16 | WIP17 |  | Subsystem | Localization - USP;Verified1 | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP17SubCom | Ovation |  | WIP17 | May 27, 2019 1:53:05 PM GMT+08:00
CTUDT00623613 | Refresh | Closed | panchami_pradeep | May 10, 2019 12:22:21 PM GMT+08:00 | Refresh: Localization: When the console is set up for Norwegian language, right monitor is displayed in English. | WIP16 | WIP17 |  | Subsystem | Localization - USP;Verified1 | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP17SubCom | Brilliance 64 |  | WIP17 | May 27, 2019 1:53:05 PM GMT+08:00
CTUDT00623613 | Refresh | Closed | panchami_pradeep | May 10, 2019 12:22:21 PM GMT+08:00 | Refresh: Localization: When the console is set up for Norwegian language, right monitor is displayed in English. | WIP16 | WIP17 |  | Subsystem | Localization - USP;Verified1 | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP17SubCom | Brilliance Big Bore |  | WIP17 | May 27, 2019 1:53:05 PM GMT+08:00
CTUDT00623613 | Refresh | Closed | panchami_pradeep | May 10, 2019 12:22:21 PM GMT+08:00 | Refresh: Localization: When the console is set up for Norwegian language, right monitor is displayed in English. | WIP16 | WIP17 |  | Subsystem | Localization - USP;Verified1 | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP17SubCom | Ingenuity 64 |  | WIP17 | May 27, 2019 1:53:05 PM GMT+08:00
CTUDT00623613 | Refresh | Closed | panchami_pradeep | May 10, 2019 12:22:21 PM GMT+08:00 | Refresh: Localization: When the console is set up for Norwegian language, right monitor is displayed in English. | WIP16 | WIP17 |  | Subsystem | Localization - USP;Verified1 | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP17SubCom | IQon1 |  | WIP17 | May 27, 2019 1:53:05 PM GMT+08:00
CTUDT00623612 | Refresh | Closed | lei_yang | May 10, 2019 12:21:17 PM GMT+08:00 | No message prompt when the respiratory analogue signal is not in the range of learning parameters with Resp=User Defined. | WIP16 | WIP19 |  | Subsystem | Respiratory | Console-SRS.15291 | Fixed | Prakash_Chand_Yadav | 0 | 2 | Console Software | Console4 | YES | Closed | WIP18SubCom | Brilliance 64 |  | WIP19 | July 25, 2019 1:55:37 PM GMT+08:00
CTUDT00623612 | Refresh | Closed | lei_yang | May 10, 2019 12:21:17 PM GMT+08:00 | No message prompt when the respiratory analogue signal is not in the range of learning parameters with Resp=User Defined. | WIP16 | WIP19 |  | Subsystem | Respiratory | Console-SRS.15291 | Fixed | Prakash_Chand_Yadav | 0 | 2 | Console Software | Console4 | YES | Closed | WIP18SubCom | Ingenuity 64 |  | WIP19 | July 25, 2019 1:55:37 PM GMT+08:00
CTUDT00623610 | Refresh | Closed | lei_yang | May 10, 2019 11:57:03 AM GMT+08:00 | Trigger window measured on the plot isn't equal to the measured trigger window on the wave -+10%. | WIP16 | Refer to Attachments | NA | System | Cardiac Workflow; Blocker | n/a | Fixed | lei_yang | 0 | 2 | Test Design | GCX | NO | Closed | WIP18.1SubCom | Ingenuity 64 |  | Refer to Attachments | August 7, 2019 5:07:24 PM GMT+08:00
CTUDT00623605 | Refresh | Closed | nayak_r | May 9, 2019 11:19:39 PM GMT+08:00 | BB Refresh:"Raw file does not exist" error popup occurs when we try to load the Partial raw data copied to offline recon | BB_WIP5.1 | DHF330449 Rev 07 Big Bore HOST Verification test protocol |  | Subsystem | Patient Directory | NA | Fixed | vinoliya_j | 0 | 2 | Test Design | Console4 | NO | Closed | WIP17SubCom | Brilliance Big Bore |  | DHF330449 Rev 07 Big Bore HOST Verification test protocol | May 24, 2019 8:59:35 PM GMT+08:00
CTUDT00623600 | Refresh | Closed | shekhar_p | May 9, 2019 7:20:00 PM GMT+08:00 | STPS:US-DoD Config option is not available | WIP16 | WIP17 | None | Subsystem | Service Tools | STPS-SSRS-1339 | Fixed | Snehashis_Dey | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP17SubCom | IQon1 |  | WIP17 | May 17, 2019 2:44:37 PM GMT+08:00
CTUDT00623599 | Refresh | Closed | navale_s | May 9, 2019 7:09:59 PM GMT+08:00 | Refresh: User added voice language cannot be upgraded | 4.8 | BB_WIP9 |  | System | Installation | NA | Fixed | chakraborty_d | 0 | 3 | Service Tools_Software | Service Tools | NO | Closed | BB_WIP7_SubCom | Brilliance Big Bore |  | BB_WIP9 | August 8, 2019 1:05:03 PM GMT+08:00
CTUDT00623593 | Refresh | Closed | jotawar_d | May 9, 2019 6:37:08 PM GMT+08:00 | Refresh: In CCT Fluro Aquisition view, actual CTDIvol and DLPvol values are displayed with ~ 14 decimal places. | 4.8 | BB_WIP8 |  | Subsystem | CCT | Console-SRS.15972 | Fixed | navale_s | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP7_SubCom | Brilliance Big Bore |  | BB_WIP8 | June 11, 2019 4:41:45 PM GMT+08:00
CTUDT00623591 | Refresh | Closed | chen_ho | May 9, 2019 5:06:31 PM GMT+08:00 | Fail to logon as philips_service. "The referenced account is currently locked out and may not be logged on to". | WIP16 |  |  | System |  |  | Invalid/Uncontrolled Configuration | chen_ho | 0 |  |  |  |  | Closed |  | Brilliance 64
CTUDT00623591 | Refresh | Closed | chen_ho | May 9, 2019 5:06:31 PM GMT+08:00 | Fail to logon as philips_service. "The referenced account is currently locked out and may not be logged on to". | WIP16 |  |  | System |  |  | Invalid/Uncontrolled Configuration | chen_ho | 0 |  |  |  |  | Closed |  | iCT2
CTUDT00623588 | Refresh | Closed | PremKumar_Rajendran | May 9, 2019 2:27:34 PM GMT+08:00 | BB Refresh : In Preferences ‘Reconstruction’ tab, under ‘CT Simulation Exam Card Presets’ checkbox texts are truncated | BB_WIP5.1 | WIP17 | NA | Subsystem | Preference | NA | Fixed | s_sa | 0 | 4 | Console USP_EBW | iCAP | NO | Closed | BB_WIP7_SubCom | Brilliance Big Bore |  | WIP17 | May 30, 2019 5:18:55 PM GMT+08:00
CTUDT00623586 | Refresh | Closed | sunilkumar_j | May 9, 2019 2:02:38 PM GMT+08:00 | Ingenuity [Refresh WIP 15]: Few parameters were not displaying properly | WIP15 | DHF318051 Rev 17 |  | Subsystem | Plan | NA | Fixed | jindal_k | 0 | 2 | Test Design | Console4 | NO | Closed | WIP17SubCom | Ovation |  | DHF318051 Rev 17 | June 21, 2019 2:14:53 PM GMT+08:00
CTUDT00623586 | Refresh | Closed | sunilkumar_j | May 9, 2019 2:02:38 PM GMT+08:00 | Ingenuity [Refresh WIP 15]: Few parameters were not displaying properly | WIP15 | DHF318051 Rev 17 |  | Subsystem | Plan | NA | Fixed | jindal_k | 0 | 2 | Test Design | Console4 | NO | Closed | WIP17SubCom | Brilliance 64 |  | DHF318051 Rev 17 | June 21, 2019 2:14:53 PM GMT+08:00
CTUDT00623586 | Refresh | Closed | sunilkumar_j | May 9, 2019 2:02:38 PM GMT+08:00 | Ingenuity [Refresh WIP 15]: Few parameters were not displaying properly | WIP15 | DHF318051 Rev 17 |  | Subsystem | Plan | NA | Fixed | jindal_k | 0 | 2 | Test Design | Console4 | NO | Closed | WIP17SubCom | Ingenuity 64 |  | DHF318051 Rev 17 | June 21, 2019 2:14:53 PM GMT+08:00
CTUDT00623586 | Refresh | Closed | sunilkumar_j | May 9, 2019 2:02:38 PM GMT+08:00 | Ingenuity [Refresh WIP 15]: Few parameters were not displaying properly | WIP15 | DHF318051 Rev 17 |  | Subsystem | Plan | NA | Fixed | jindal_k | 0 | 2 | Test Design | Console4 | NO | Closed | WIP17SubCom | IQon1 |  | DHF318051 Rev 17 | June 21, 2019 2:14:53 PM GMT+08:00
CTUDT00623583 | Refresh | Closed | nayak_r | May 9, 2019 12:47:08 AM GMT+08:00 | Refresh:"CopyToCIRS" operation fails to copy partial raw to CIRS. | BB_WIP5.1/WIP16 & WIP17 | WIP20 | Engineering would still be able to use the data after a clean re-boot or using the UTS commands to upload the partial raw to the CIRS directly. | Subsystem | CopytoCIRS;Blocker20; Verified1 | N/A | Fixed | richards_b | 0 | 4 | CIRS_SW_Supervisor | CIRS | NO | Closed | WIP20SubCom | Ovation |  | WIP20 | September 25, 2019 8:23:25 PM GMT+08:00
CTUDT00623583 | Refresh | Closed | nayak_r | May 9, 2019 12:47:08 AM GMT+08:00 | Refresh:"CopyToCIRS" operation fails to copy partial raw to CIRS. | BB_WIP5.1/WIP16 & WIP17 | WIP20 | Engineering would still be able to use the data after a clean re-boot or using the UTS commands to upload the partial raw to the CIRS directly. | Subsystem | CopytoCIRS;Blocker20; Verified1 | N/A | Fixed | richards_b | 0 | 4 | CIRS_SW_Supervisor | CIRS | NO | Closed | WIP20SubCom | Brilliance 64 |  | WIP20 | September 25, 2019 8:23:25 PM GMT+08:00
CTUDT00623583 | Refresh | Closed | nayak_r | May 9, 2019 12:47:08 AM GMT+08:00 | Refresh:"CopyToCIRS" operation fails to copy partial raw to CIRS. | BB_WIP5.1/WIP16 & WIP17 | WIP20 | Engineering would still be able to use the data after a clean re-boot or using the UTS commands to upload the partial raw to the CIRS directly. | Subsystem | CopytoCIRS;Blocker20; Verified1 | N/A | Fixed | richards_b | 0 | 4 | CIRS_SW_Supervisor | CIRS | NO | Closed | WIP20SubCom | Brilliance Big Bore |  | WIP20 | September 25, 2019 8:23:25 PM GMT+08:00
CTUDT00623583 | Refresh | Closed | nayak_r | May 9, 2019 12:47:08 AM GMT+08:00 | Refresh:"CopyToCIRS" operation fails to copy partial raw to CIRS. | BB_WIP5.1/WIP16 & WIP17 | WIP20 | Engineering would still be able to use the data after a clean re-boot or using the UTS commands to upload the partial raw to the CIRS directly. | Subsystem | CopytoCIRS;Blocker20; Verified1 | N/A | Fixed | richards_b | 0 | 4 | CIRS_SW_Supervisor | CIRS | NO | Closed | WIP20SubCom | Ingenuity 64 |  | WIP20 | September 25, 2019 8:23:25 PM GMT+08:00
CTUDT00623583 | Refresh | Closed | nayak_r | May 9, 2019 12:47:08 AM GMT+08:00 | Refresh:"CopyToCIRS" operation fails to copy partial raw to CIRS. | BB_WIP5.1/WIP16 & WIP17 | WIP20 | Engineering would still be able to use the data after a clean re-boot or using the UTS commands to upload the partial raw to the CIRS directly. | Subsystem | CopytoCIRS;Blocker20; Verified1 | N/A | Fixed | richards_b | 0 | 4 | CIRS_SW_Supervisor | CIRS | NO | Closed | WIP20SubCom | IQon1 |  | WIP20 | September 25, 2019 8:23:25 PM GMT+08:00
CTUDT00623576 | Refresh | Closed | wang_l | May 8, 2019 10:24:13 AM GMT+08:00 | Refresh BB: The product id is still 728244 in system but expected value should be 728242 | BB_WIP5.1 | BB_WIP7 |  | Subsystem | SIDS | CT-NM.SRS.Service.2236 | Fixed | ranjithkumar_m | 0 | 2 | Service Tools_Software | Console4 | NO | Closed | BB_WIP8_SubCom | Brilliance Big Bore |  | BB_WIP7 | May 22, 2019 8:40:22 PM GMT+08:00
CTUDT00623574 | Refresh | Closed | nayak_r | May 8, 2019 4:18:46 AM GMT+08:00 | BB Refresh:Units for Planned CTDIVol and Planned DLP are incorrectly displayed in Acquisition viewer | BB_WIP5.1 | BB_WIP7 | No workaround. | Subsystem | Acquisition Viewer | None | Fixed | killada_s | 0 | 2 | Console Software | Console4 | YES | Closed | BB_WIP7_SubCom | Brilliance Big Bore |  | BB_WIP7 | May 28, 2019 6:26:03 PM GMT+08:00
CTUDT00623572 | Refresh | Closed | 320053769 | May 8, 2019 2:26:36 AM GMT+08:00 | Burning a study with 200 images (of matrix 512X512) to CD takes longer than 3 minutes. | WIP15 |  |  | Subsystem |  |  | Duplicate | veerabhadraiah | 0 |  |  |  |  | Closed |  | IQon1
CTUDT00623570 | Refresh | Closed | Jose_D_Sergio_Almeida | May 8, 2019 2:03:41 AM GMT+08:00 | Incorrect measurements units for Spectral results. | WIP16 | WIP18 |  | Subsystem | sCTV; Blocker | NA | Fixed | aviw | 0 | 2 | Console USP_EBW | Support_Spectral | YES | Closed | WIP18SegReq | IQon1 |  | WIP18 | June 11, 2019 3:25:29 PM GMT+08:00
CTUDT00623568 | Refresh | Closed | nayak_r | May 8, 2019 1:29:34 AM GMT+08:00 | BB Refresh: Gantry cannot initialize issue was observed on trying to extend the completed Abdomen scan | BB_WIP5.1 |  |  | Subsystem | Gantry |  | Invalid/Uncontrolled Configuration | karkour | 0 |  |  | GCX | NO | Closed |  | Brilliance Big Bore
CTUDT00623566 | Refresh | Closed | nayak_r | May 8, 2019 12:51:54 AM GMT+08:00 | BB Refresh: Exam information page in Exam summary displays mismatch for DLP ,SSDE and mAs values with those planned during a scan | BB_WIP5.1 | DHF303211 Rev13 Host SSRS for Refresh DHF316705 Rev.12 Host SSRS for Refresh BB |  | Subsystem | Exam Summary | Console-SRS.15421 | Fixed | hatti_c | 0 | 4 | Host SSRS | Console4 | NO | Closed | WIP17SubCom | Ovation |  | DHF303211 Rev13 Host SSRS for Refresh DHF316705 Rev.12 Host SSRS for Refresh BB | July 19, 2019 2:01:47 PM GMT+08:00
CTUDT00623566 | Refresh | Closed | nayak_r | May 8, 2019 12:51:54 AM GMT+08:00 | BB Refresh: Exam information page in Exam summary displays mismatch for DLP ,SSDE and mAs values with those planned during a scan | BB_WIP5.1 | DHF303211 Rev13 Host SSRS for Refresh DHF316705 Rev.12 Host SSRS for Refresh BB |  | Subsystem | Exam Summary | Console-SRS.15421 | Fixed | hatti_c | 0 | 4 | Host SSRS | Console4 | NO | Closed | WIP17SubCom | Brilliance 64 |  | DHF303211 Rev13 Host SSRS for Refresh DHF316705 Rev.12 Host SSRS for Refresh BB | July 19, 2019 2:01:47 PM GMT+08:00
CTUDT00623566 | Refresh | Closed | nayak_r | May 8, 2019 12:51:54 AM GMT+08:00 | BB Refresh: Exam information page in Exam summary displays mismatch for DLP ,SSDE and mAs values with those planned during a scan | BB_WIP5.1 | DHF303211 Rev13 Host SSRS for Refresh DHF316705 Rev.12 Host SSRS for Refresh BB |  | Subsystem | Exam Summary | Console-SRS.15421 | Fixed | hatti_c | 0 | 4 | Host SSRS | Console4 | NO | Closed | WIP17SubCom | Brilliance Big Bore |  | DHF303211 Rev13 Host SSRS for Refresh DHF316705 Rev.12 Host SSRS for Refresh BB | July 19, 2019 2:01:47 PM GMT+08:00
CTUDT00623566 | Refresh | Closed | nayak_r | May 8, 2019 12:51:54 AM GMT+08:00 | BB Refresh: Exam information page in Exam summary displays mismatch for DLP ,SSDE and mAs values with those planned during a scan | BB_WIP5.1 | DHF303211 Rev13 Host SSRS for Refresh DHF316705 Rev.12 Host SSRS for Refresh BB |  | Subsystem | Exam Summary | Console-SRS.15421 | Fixed | hatti_c | 0 | 4 | Host SSRS | Console4 | NO | Closed | WIP17SubCom | Ingenuity 64 |  | DHF303211 Rev13 Host SSRS for Refresh DHF316705 Rev.12 Host SSRS for Refresh BB | July 19, 2019 2:01:47 PM GMT+08:00
CTUDT00623566 | Refresh | Closed | nayak_r | May 8, 2019 12:51:54 AM GMT+08:00 | BB Refresh: Exam information page in Exam summary displays mismatch for DLP ,SSDE and mAs values with those planned during a scan | BB_WIP5.1 | DHF303211 Rev13 Host SSRS for Refresh DHF316705 Rev.12 Host SSRS for Refresh BB |  | Subsystem | Exam Summary | Console-SRS.15421 | Fixed | hatti_c | 0 | 4 | Host SSRS | Console4 | NO | Closed | WIP17SubCom | IQon1 |  | DHF303211 Rev13 Host SSRS for Refresh DHF316705 Rev.12 Host SSRS for Refresh BB | July 19, 2019 2:01:47 PM GMT+08:00
CTUDT00623563 | Refresh | Closed | levy_s | May 7, 2019 11:48:30 PM GMT+08:00 | Refresh: “RealTimeFSM MFC Application has stopped working” error appear on the GPC | WIP16 | WIP17 | Restart the application | Subsystem | GPC | NA | Fixed | kline_w | 0 | 3 | Gantry Software | GCX | NO | Closed | WIP18SegReq | Ovation |  | WIP17 | August 14, 2019 2:46:53 AM GMT+08:00
CTUDT00623563 | Refresh | Closed | levy_s | May 7, 2019 11:48:30 PM GMT+08:00 | Refresh: “RealTimeFSM MFC Application has stopped working” error appear on the GPC | WIP16 | WIP17 | Restart the application | Subsystem | GPC | NA | Fixed | kline_w | 0 | 3 | Gantry Software | GCX | NO | Closed | WIP18SegReq | IQon1 |  | WIP17 | August 14, 2019 2:46:53 AM GMT+08:00
CTUDT00623561 | Refresh | Closed | 320053770 | May 7, 2019 9:54:33 PM GMT+08:00 | After a System configuration Verification, there is a gantry error until a reset to the TDMC. | WIP 15 |  | reset the TDMC via “DiagSender” interface > Close E-Stop > Perform the scan | Subsystem | System Configuration - TDMC component; Blocker | N/A | No Plan to Fix | ivanf | 0 | 4 |  | DMS | NO | Closed |  | Ovation
CTUDT00623561 | Refresh | Closed | 320053770 | May 7, 2019 9:54:33 PM GMT+08:00 | After a System configuration Verification, there is a gantry error until a reset to the TDMC. | WIP 15 |  | reset the TDMC via “DiagSender” interface > Close E-Stop > Perform the scan | Subsystem | System Configuration - TDMC component; Blocker | N/A | No Plan to Fix | ivanf | 0 | 4 |  | DMS | NO | Closed |  | IQon1
CTUDT00623559 | Refresh | Closed | 320053770 | May 7, 2019 9:41:45 PM GMT+08:00 | During integration for WIP 15. In an ePHI test, Raw data Copy function failed. | WIP 15 |  |  | Subsystem | Patient Directory | None | Works As Specified / Not A Defect | Prakash_Chand_Yadav | 0 |  |  | Console4 | NO | Closed | WAS | Ovation
CTUDT00623559 | Refresh | Closed | 320053770 | May 7, 2019 9:41:45 PM GMT+08:00 | During integration for WIP 15. In an ePHI test, Raw data Copy function failed. | WIP 15 |  |  | Subsystem | Patient Directory | None | Works As Specified / Not A Defect | Prakash_Chand_Yadav | 0 |  |  | Console4 | NO | Closed | WAS | Brilliance 64
CTUDT00623559 | Refresh | Closed | 320053770 | May 7, 2019 9:41:45 PM GMT+08:00 | During integration for WIP 15. In an ePHI test, Raw data Copy function failed. | WIP 15 |  |  | Subsystem | Patient Directory | None | Works As Specified / Not A Defect | Prakash_Chand_Yadav | 0 |  |  | Console4 | NO | Closed | WAS | Brilliance Big Bore
CTUDT00623559 | Refresh | Closed | 320053770 | May 7, 2019 9:41:45 PM GMT+08:00 | During integration for WIP 15. In an ePHI test, Raw data Copy function failed. | WIP 15 |  |  | Subsystem | Patient Directory | None | Works As Specified / Not A Defect | Prakash_Chand_Yadav | 0 |  |  | Console4 | NO | Closed | WAS | Ingenuity 64
CTUDT00623559 | Refresh | Closed | 320053770 | May 7, 2019 9:41:45 PM GMT+08:00 | During integration for WIP 15. In an ePHI test, Raw data Copy function failed. | WIP 15 |  |  | Subsystem | Patient Directory | None | Works As Specified / Not A Defect | Prakash_Chand_Yadav | 0 |  |  | Console4 | NO | Closed | WAS | IQon1
CTUDT00623556 | Refresh | Closed | yosfan_s | May 7, 2019 9:03:46 PM GMT+08:00 | Refresh: Acceptance/Constancy Tests: Harmonized: Resolution case 2 ROIs are not placed in the correct slice according to Spec. | WIP15 | WIP18 WIP18.1 |  | Subsystem | Calibrations;Blocker | N/A | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | YES | Closed | BB_WIP7_SubCom | Ingenuity 64 |  | WIP18 WIP18.1 | July 22, 2019 3:49:05 PM GMT+08:00
CTUDT00623549 | Refresh | Closed | 320053770 | May 7, 2019 8:13:33 PM GMT+08:00 | During integration for WIP 15, in visual diagnostics, there was an error in the DataPath test. | WIP 15 |  |  | Subsystem | Visual Diagnostics |  | Duplicate | venkatachalam_p | 0 |  |  | Service Tools | NO | Closed |  | Ovation
CTUDT00623549 | Refresh | Closed | 320053770 | May 7, 2019 8:13:33 PM GMT+08:00 | During integration for WIP 15, in visual diagnostics, there was an error in the DataPath test. | WIP 15 |  |  | Subsystem | Visual Diagnostics |  | Duplicate | venkatachalam_p | 0 |  |  | Service Tools | NO | Closed |  | IQon1
CTUDT00623546 | Refresh | Closed | yosfan_s | May 7, 2019 8:08:49 PM GMT+08:00 | Refresh: Acceptance/Constancy: Window Center (Level) value of body cases (and slice thickness 3) with Harmonized Phantom is not according to Spec. | WIP15 | WIP18 |  | Subsystem | Calibrations | NA | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP17SubCom | Brilliance Big Bore |  | WIP18 | May 26, 2019 7:35:09 PM GMT+08:00
CTUDT00623546 | Refresh | Closed | yosfan_s | May 7, 2019 8:08:49 PM GMT+08:00 | Refresh: Acceptance/Constancy: Window Center (Level) value of body cases (and slice thickness 3) with Harmonized Phantom is not according to Spec. | WIP15 | WIP18 |  | Subsystem | Calibrations | NA | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP17SubCom | Ingenuity 64 |  | WIP18 | May 26, 2019 7:35:09 PM GMT+08:00
CTUDT00623545 | Refresh | Closed | Jose_D_Sergio_Almeida | May 7, 2019 7:54:31 PM GMT+08:00 | Refresh - Spectral CT viewer viewports are not linked when comparing series | WIP15 | WIP18 | The user will to set the attribute in each viewport | Subsystem | Spectral CT Viewer - requirements | NA | Fixed | aviw | 0 | 3 | SDViewer.exe | Support_Spectral | NO | Closed | WIP18SegReq | IQon1 |  | WIP18 | June 25, 2019 4:36:58 PM GMT+08:00
CTUDT00623544 | Refresh | Closed | Jose_D_Sergio_Almeida | May 7, 2019 7:54:17 PM GMT+08:00 | Refresh - Spectral viewer crashed when creating a temporary series. | WIP16 | WIP18 |  | Subsystem | Spectral CT Viewer ; Blocker | NA | Fixed | aviw | 0 | 2 | SDViewer | Support_Spectral | NO | Closed | WIP19SegReq | IQon1 |  | WIP18 | June 18, 2019 1:39:59 PM GMT+08:00
CTUDT00623541 | Refresh | Closed | jotawar_d | May 7, 2019 7:20:58 PM GMT+08:00 | Exam crashes after setting threshold below -100HU in Bolus sequence | V4.8 | BB_WIP7 | no workaround for this scenario. | Subsystem | System | Console-SRS.15414 | Fixed | Snehashis_Dey | 0 | 3 | Console Software | Console4 | YES | Closed | BB_WIP7_SubCom | Brilliance Big Bore |  | BB_WIP7 | May 16, 2019 2:38:08 PM GMT+08:00
CTUDT00623537 | Refresh | Closed | shterner_a | May 7, 2019 6:01:04 PM GMT+08:00 | Pre-Val Refresh: SSDE, DLP CTDIvol planned and actual parameters should be rearranged when presented on acquisition viewer. | WIP15 |  |  | Subsystem | Acquisition Viewer | N/A | Change Proposal | bhawna_g | 0 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00623537 | Refresh | Closed | shterner_a | May 7, 2019 6:01:04 PM GMT+08:00 | Pre-Val Refresh: SSDE, DLP CTDIvol planned and actual parameters should be rearranged when presented on acquisition viewer. | WIP15 |  |  | Subsystem | Acquisition Viewer | N/A | Change Proposal | bhawna_g | 0 |  |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00623537 | Refresh | Closed | shterner_a | May 7, 2019 6:01:04 PM GMT+08:00 | Pre-Val Refresh: SSDE, DLP CTDIvol planned and actual parameters should be rearranged when presented on acquisition viewer. | WIP15 |  |  | Subsystem | Acquisition Viewer | N/A | Change Proposal | bhawna_g | 0 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00623537 | Refresh | Closed | shterner_a | May 7, 2019 6:01:04 PM GMT+08:00 | Pre-Val Refresh: SSDE, DLP CTDIvol planned and actual parameters should be rearranged when presented on acquisition viewer. | WIP15 |  |  | Subsystem | Acquisition Viewer | N/A | Change Proposal | bhawna_g | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00623537 | Refresh | Closed | shterner_a | May 7, 2019 6:01:04 PM GMT+08:00 | Pre-Val Refresh: SSDE, DLP CTDIvol planned and actual parameters should be rearranged when presented on acquisition viewer. | WIP15 |  |  | Subsystem | Acquisition Viewer | N/A | Change Proposal | bhawna_g | 0 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00623532 | Refresh | Closed | vinoliya_j | May 7, 2019 5:11:58 PM GMT+08:00 | BB refresh - Exam summary did not contain all the pages even though they are selected in preferences | BB_WIP6 | BB_WIP7 | NA | Subsystem | Exam Summary | Console-SRS.14285 | Fixed | sagar_v | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP7_SubCom | Brilliance Big Bore |  | BB_WIP7 | May 22, 2019 4:14:16 PM GMT+08:00
CTUDT00623530 | Refresh | Closed | gT_s | May 7, 2019 4:50:38 PM GMT+08:00 | Refresh : The message displayed does not match with the requirement, it contains extra "And" in message. | WIP 15 | Host SSRS DHF303211 Rev 11 WIP18 | NA | Subsystem | SSRS | Console-SRS.13859 | Fixed | kr_m | 0 | 4 | Console Software | Console4 | NO | Closed | WIP18SubCom | IQon1 |  | Host SSRS DHF303211 Rev 11 WIP18 | June 28, 2019 5:14:24 PM GMT+08:00
CTUDT00623525 | Refresh | Closed | shterner_a | May 7, 2019 4:03:22 PM GMT+08:00 | Pre-Val Refresh: Exam Summary shows wrong injection parameters | WIP15 |  | NA | Subsystem | Exam Summary | NA | Works As Specified / Not A Defect | jayakumar_p | 0 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00623525 | Refresh | Closed | shterner_a | May 7, 2019 4:03:22 PM GMT+08:00 | Pre-Val Refresh: Exam Summary shows wrong injection parameters | WIP15 |  | NA | Subsystem | Exam Summary | NA | Works As Specified / Not A Defect | jayakumar_p | 0 |  |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00623525 | Refresh | Closed | shterner_a | May 7, 2019 4:03:22 PM GMT+08:00 | Pre-Val Refresh: Exam Summary shows wrong injection parameters | WIP15 |  | NA | Subsystem | Exam Summary | NA | Works As Specified / Not A Defect | jayakumar_p | 0 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00623525 | Refresh | Closed | shterner_a | May 7, 2019 4:03:22 PM GMT+08:00 | Pre-Val Refresh: Exam Summary shows wrong injection parameters | WIP15 |  | NA | Subsystem | Exam Summary | NA | Works As Specified / Not A Defect | jayakumar_p | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00623525 | Refresh | Closed | shterner_a | May 7, 2019 4:03:22 PM GMT+08:00 | Pre-Val Refresh: Exam Summary shows wrong injection parameters | WIP15 |  | NA | Subsystem | Exam Summary | NA | Works As Specified / Not A Defect | jayakumar_p | 0 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00623520 | Refresh | Closed | Valeria_Kashtanov | May 7, 2019 3:15:37 PM GMT+08:00 | Refresh: Calibrations: Values of Tube data are not equal to values in Tube History Tool | BB_WIP5.1 | N/A |  | Subsystem | Calibrations | n/a | Fixed | yosfan_s | 0 | 3 | Test Design | DMS | NO | Closed | WIP18SubCom | Ovation |  | N/A | June 16, 2019 3:52:56 PM GMT+08:00
CTUDT00623520 | Refresh | Closed | Valeria_Kashtanov | May 7, 2019 3:15:37 PM GMT+08:00 | Refresh: Calibrations: Values of Tube data are not equal to values in Tube History Tool | BB_WIP5.1 | N/A |  | Subsystem | Calibrations | n/a | Fixed | yosfan_s | 0 | 3 | Test Design | DMS | NO | Closed | WIP18SubCom | Brilliance Big Bore |  | N/A | June 16, 2019 3:52:56 PM GMT+08:00
CTUDT00623517 | Refresh | Closed | nayak_r | May 7, 2019 3:33:54 AM GMT+08:00 | BB Refresh: Move to selected position button in Software IV did not move couch to specified Needle Position | BB_WIP5.1 | BB_WIP8 |  | Subsystem | IV Controls | NA | Fixed | goswami_s | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP8_SubCom | Brilliance Big Bore |  | BB_WIP8 | July 1, 2019 2:59:22 PM GMT+08:00
CTUDT00623515 | Refresh | Closed | shekhar_p | May 7, 2019 3:00:57 AM GMT+08:00 | During localization test system crashes. | WIP16 | WIP17 |  | Subsystem | Localization | N/A | Fixed | navale_s | 0 | 3 | Console Software | Console4 | NO | Closed | WIP17SegReq | Ovation |  | WIP17 | May 28, 2019 6:22:44 PM GMT+08:00
CTUDT00623515 | Refresh | Closed | shekhar_p | May 7, 2019 3:00:57 AM GMT+08:00 | During localization test system crashes. | WIP16 | WIP17 |  | Subsystem | Localization | N/A | Fixed | navale_s | 0 | 3 | Console Software | Console4 | NO | Closed | WIP17SegReq | Brilliance 64 |  | WIP17 | May 28, 2019 6:22:44 PM GMT+08:00
CTUDT00623515 | Refresh | Closed | shekhar_p | May 7, 2019 3:00:57 AM GMT+08:00 | During localization test system crashes. | WIP16 | WIP17 |  | Subsystem | Localization | N/A | Fixed | navale_s | 0 | 3 | Console Software | Console4 | NO | Closed | WIP17SegReq | Ingenuity 64 |  | WIP17 | May 28, 2019 6:22:44 PM GMT+08:00
CTUDT00623515 | Refresh | Closed | shekhar_p | May 7, 2019 3:00:57 AM GMT+08:00 | During localization test system crashes. | WIP16 | WIP17 |  | Subsystem | Localization | N/A | Fixed | navale_s | 0 | 3 | Console Software | Console4 | NO | Closed | WIP17SegReq | IQon1 |  | WIP17 | May 28, 2019 6:22:44 PM GMT+08:00
CTUDT00623511 | Refresh | Closed | shterner_a | May 6, 2019 8:28:22 PM GMT+08:00 | Pre-Val Refresh: Changing interface language is not properly implemented | WIP15 |  |  | Subsystem |  |  | Duplicate | prabhu_m | 0 |  |  |  |  | Closed |  | IQon1
CTUDT00623503 | Refresh | Closed | levy_s | May 6, 2019 7:32:32 PM GMT+08:00 | Refresh: “System.cfg” file is modified with no apparent reason and prevent the machine functionality | WIP16 | WIP17 |  | Subsystem | Installation, Blocker | N/A | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | NO | Closed | WIP17SubCom | Ovation |  | WIP17 | May 15, 2019 6:32:24 PM GMT+08:00
CTUDT00623503 | Refresh | Closed | levy_s | May 6, 2019 7:32:32 PM GMT+08:00 | Refresh: “System.cfg” file is modified with no apparent reason and prevent the machine functionality | WIP16 | WIP17 |  | Subsystem | Installation, Blocker | N/A | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | NO | Closed | WIP17SubCom | Brilliance 64 |  | WIP17 | May 15, 2019 6:32:24 PM GMT+08:00
CTUDT00623503 | Refresh | Closed | levy_s | May 6, 2019 7:32:32 PM GMT+08:00 | Refresh: “System.cfg” file is modified with no apparent reason and prevent the machine functionality | WIP16 | WIP17 |  | Subsystem | Installation, Blocker | N/A | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | NO | Closed | WIP17SubCom | Ingenuity 64 |  | WIP17 | May 15, 2019 6:32:24 PM GMT+08:00
CTUDT00623503 | Refresh | Closed | levy_s | May 6, 2019 7:32:32 PM GMT+08:00 | Refresh: “System.cfg” file is modified with no apparent reason and prevent the machine functionality | WIP16 | WIP17 |  | Subsystem | Installation, Blocker | N/A | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | NO | Closed | WIP17SubCom | IQon1 |  | WIP17 | May 15, 2019 6:32:24 PM GMT+08:00
CTUDT00623501 | Refresh | Closed | shterner_a | May 6, 2019 5:58:59 PM GMT+08:00 | Pre-Val Refresh Monitor setting should be locked and not available for User as changing Luminance will not conform to DIN-6868 standard | WIP15 |  |  | Subsystem |  |  | Change Proposal | marino | 0 |  |  |  |  | Closed |  | IQon1
CTUDT00623497 | Refresh | Closed | shterner_a | May 6, 2019 5:49:36 PM GMT+08:00 | Pre Val Refresh: Daily Visual Monitor Calibration is not carrying out any calibration | WIP15 |  |  | Subsystem | Monitor Calibration |  | Change Proposal | mistry_h | 0 |  |  | Support-System | NO | Closed |  | IQon1
CTUDT00623494 | Refresh | Closed | v_sumanth | May 6, 2019 5:32:56 PM GMT+08:00 | BBRefresh: The exceed cycle value in log view message is incorrect. | 4.8 | BB_WIP7 | There is no workaround. This issue doesn't cause any obstacle to perform the scan and it is purely concerned with the log message that is logged into log viewer. | Subsystem | Log Viewer | NA | Fixed | jayakumar_p | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP7_SubCom | Brilliance Big Bore |  | BB_WIP7 | May 24, 2019 8:10:03 PM GMT+08:00
CTUDT00623486 | Refresh | Closed | shterner_a | May 6, 2019 4:03:17 PM GMT+08:00 | Pre-Val Refresh: A Problem has occurred, please retry. If the problem persists ,contact service - message popped up when the scan was paused | WIP15 |  |  | Subsystem | System | N/A | Works As Specified / Not A Defect | tamari_o | 1 |  |  | Support-System | YES | Closed |  | IQon1
CTUDT00623484 | Refresh | Closed | Jose_D_Sergio_Almeida | May 6, 2019 3:54:13 PM GMT+08:00 | Refresh-Spectral viewer crashed when comparing spectral plots for SBI data. WIP15 | WIP15 | WIP18 | Create the spectral ROI when not in Compare mode. | Subsystem | Spectral Viewer | NA | Fixed | aviw | 0 | 3 | SDViewer.exe | Support_Spectral | NO | Closed | WIP17SegReq | IQon1 |  | WIP18 | May 24, 2019 4:34:42 AM GMT+08:00
CTUDT00623481 | Refresh | Closed | shterner_a | May 6, 2019 3:32:16 PM GMT+08:00 | Pre-Val Refresh: Log in User name is not displayed on left monitor | WIP15 |  |  | Subsystem | Login | Console-SRS.8910 | Duplicate | jose_s | 0 |  |  | Console4 | NO | Closed | WIP17SubCom | Ovation
CTUDT00623481 | Refresh | Closed | shterner_a | May 6, 2019 3:32:16 PM GMT+08:00 | Pre-Val Refresh: Log in User name is not displayed on left monitor | WIP15 |  |  | Subsystem | Login | Console-SRS.8910 | Duplicate | jose_s | 0 |  |  | Console4 | NO | Closed | WIP17SubCom | Brilliance 64
CTUDT00623481 | Refresh | Closed | shterner_a | May 6, 2019 3:32:16 PM GMT+08:00 | Pre-Val Refresh: Log in User name is not displayed on left monitor | WIP15 |  |  | Subsystem | Login | Console-SRS.8910 | Duplicate | jose_s | 0 |  |  | Console4 | NO | Closed | WIP17SubCom | Brilliance Big Bore
CTUDT00623481 | Refresh | Closed | shterner_a | May 6, 2019 3:32:16 PM GMT+08:00 | Pre-Val Refresh: Log in User name is not displayed on left monitor | WIP15 |  |  | Subsystem | Login | Console-SRS.8910 | Duplicate | jose_s | 0 |  |  | Console4 | NO | Closed | WIP17SubCom | Ingenuity 64
CTUDT00623481 | Refresh | Closed | shterner_a | May 6, 2019 3:32:16 PM GMT+08:00 | Pre-Val Refresh: Log in User name is not displayed on left monitor | WIP15 |  |  | Subsystem | Login | Console-SRS.8910 | Duplicate | jose_s | 0 |  |  | Console4 | NO | Closed | WIP17SubCom | IQon1
CTUDT00623478 | Refresh | Closed | urim | May 6, 2019 3:15:47 PM GMT+08:00 | Refresh Usability: IV SW: Work Position Radio-Button is not selected when typing in work position Value | BB_WIP4 |  |  | Subsystem | IV Controls |  | Change Proposal | mistry_h | 0 |  |  | Support-System | NO | Closed |  | Brilliance Big Bore
CTUDT00623475 | Refresh | Closed | shekhar_p | May 6, 2019 3:38:32 AM GMT+08:00 | Refresh: Not able to open and save the logs in the log viewer | WIP15 |  | Once the user encounters the Refresh After some time popup, Please wait for sometime for the logs to load and then click Refresh. Since Logviewer is been used by many subsystems, though there are no actions from your end, there are chances that other subsystems will push the logs. The same suggestion is applicable for Save scenario as well. | Subsystem | Log Viewer | NA | Duplicate | tamari_o | 1 |  |  | Support-System | NO | Closed |  | Ovation
CTUDT00623475 | Refresh | Closed | shekhar_p | May 6, 2019 3:38:32 AM GMT+08:00 | Refresh: Not able to open and save the logs in the log viewer | WIP15 |  | Once the user encounters the Refresh After some time popup, Please wait for sometime for the logs to load and then click Refresh. Since Logviewer is been used by many subsystems, though there are no actions from your end, there are chances that other subsystems will push the logs. The same suggestion is applicable for Save scenario as well. | Subsystem | Log Viewer | NA | Duplicate | tamari_o | 1 |  |  | Support-System | NO | Closed |  | Brilliance 64
CTUDT00623475 | Refresh | Closed | shekhar_p | May 6, 2019 3:38:32 AM GMT+08:00 | Refresh: Not able to open and save the logs in the log viewer | WIP15 |  | Once the user encounters the Refresh After some time popup, Please wait for sometime for the logs to load and then click Refresh. Since Logviewer is been used by many subsystems, though there are no actions from your end, there are chances that other subsystems will push the logs. The same suggestion is applicable for Save scenario as well. | Subsystem | Log Viewer | NA | Duplicate | tamari_o | 1 |  |  | Support-System | NO | Closed |  | Brilliance Big Bore
CTUDT00623475 | Refresh | Closed | shekhar_p | May 6, 2019 3:38:32 AM GMT+08:00 | Refresh: Not able to open and save the logs in the log viewer | WIP15 |  | Once the user encounters the Refresh After some time popup, Please wait for sometime for the logs to load and then click Refresh. Since Logviewer is been used by many subsystems, though there are no actions from your end, there are chances that other subsystems will push the logs. The same suggestion is applicable for Save scenario as well. | Subsystem | Log Viewer | NA | Duplicate | tamari_o | 1 |  |  | Support-System | NO | Closed |  | Ingenuity 64
CTUDT00623475 | Refresh | Closed | shekhar_p | May 6, 2019 3:38:32 AM GMT+08:00 | Refresh: Not able to open and save the logs in the log viewer | WIP15 |  | Once the user encounters the Refresh After some time popup, Please wait for sometime for the logs to load and then click Refresh. Since Logviewer is been used by many subsystems, though there are no actions from your end, there are chances that other subsystems will push the logs. The same suggestion is applicable for Save scenario as well. | Subsystem | Log Viewer | NA | Duplicate | tamari_o | 1 |  |  | Support-System | NO | Closed |  | IQon1
CTUDT00623471 | Refresh | Closed | 320053770 | May 5, 2019 9:58:47 PM GMT+08:00 | After finishing installation of WIP 16 HOST & GPC there is an un-programmable deviation related to the GPC. | WIP16 |  |  | Subsystem | GPC |  | Invalid/Uncontrolled Configuration | kline_w | 0 |  |  | GCX | NO | Closed |  | Ovation
CTUDT00623469 | Refresh | Closed | 320053770 | May 5, 2019 9:53:55 PM GMT+08:00 | During integration tests for WIP 16, the concurrent recon feature doesn’t work properly in IMR configuration. | WIP 14,15,16 |  |  | Subsystem | CIRS |  | Works As Specified / Not A Defect | daniel_l | 0 |  |  | Support-System | NO | Closed |  | Ovation
CTUDT00623464 | Refresh | Closed | finestine_m | May 5, 2019 7:22:46 PM GMT+08:00 | Refresh: Common Functions: Missing Tube Data in BDC results | WIP15 | WIP17 |  | Subsystem | Calibrations | N/A | Fixed | Kostya_Baklan | 0 | 2 | DMS Calibrations | DMS | NO | Closed | SubCom | Ovation |  | WIP17 | May 26, 2019 5:37:49 PM GMT+08:00
CTUDT00623464 | Refresh | Closed | finestine_m | May 5, 2019 7:22:46 PM GMT+08:00 | Refresh: Common Functions: Missing Tube Data in BDC results | WIP15 | WIP17 |  | Subsystem | Calibrations | N/A | Fixed | Kostya_Baklan | 0 | 2 | DMS Calibrations | DMS | NO | Closed | SubCom | Brilliance Big Bore |  | WIP17 | May 26, 2019 5:37:49 PM GMT+08:00
CTUDT00623464 | Refresh | Closed | finestine_m | May 5, 2019 7:22:46 PM GMT+08:00 | Refresh: Common Functions: Missing Tube Data in BDC results | WIP15 | WIP17 |  | Subsystem | Calibrations | N/A | Fixed | Kostya_Baklan | 0 | 2 | DMS Calibrations | DMS | NO | Closed | SubCom | Ingenuity 64 |  | WIP17 | May 26, 2019 5:37:49 PM GMT+08:00
CTUDT00623464 | Refresh | Closed | finestine_m | May 5, 2019 7:22:46 PM GMT+08:00 | Refresh: Common Functions: Missing Tube Data in BDC results | WIP15 | WIP17 |  | Subsystem | Calibrations | N/A | Fixed | Kostya_Baklan | 0 | 2 | DMS Calibrations | DMS | NO | Closed | SubCom | IQon1 |  | WIP17 | May 26, 2019 5:37:49 PM GMT+08:00
CTUDT00623460 | Refresh | Closed | wang_l | May 5, 2019 4:53:23 PM GMT+08:00 | Refresh:System can access any webset even if the webset is not added in URL Restriction list | WIP15 | BB_WIP9 WIP19 | No | Subsystem | Configuration | NA | Fixed | venkatachalam_p | 0 | 2 | Service Tools_Software | Support-Security | NO | Closed | SubNC | Ovation |  | BB_WIP9 WIP19 | July 25, 2019 1:50:04 PM GMT+08:00
CTUDT00623460 | Refresh | Closed | wang_l | May 5, 2019 4:53:23 PM GMT+08:00 | Refresh:System can access any webset even if the webset is not added in URL Restriction list | WIP15 | BB_WIP9 WIP19 | No | Subsystem | Configuration | NA | Fixed | venkatachalam_p | 0 | 2 | Service Tools_Software | Support-Security | NO | Closed | SubNC | Brilliance 64 |  | BB_WIP9 WIP19 | July 25, 2019 1:50:04 PM GMT+08:00
CTUDT00623460 | Refresh | Closed | wang_l | May 5, 2019 4:53:23 PM GMT+08:00 | Refresh:System can access any webset even if the webset is not added in URL Restriction list | WIP15 | BB_WIP9 WIP19 | No | Subsystem | Configuration | NA | Fixed | venkatachalam_p | 0 | 2 | Service Tools_Software | Support-Security | NO | Closed | SubNC | Brilliance Big Bore |  | BB_WIP9 WIP19 | July 25, 2019 1:50:04 PM GMT+08:00
CTUDT00623460 | Refresh | Closed | wang_l | May 5, 2019 4:53:23 PM GMT+08:00 | Refresh:System can access any webset even if the webset is not added in URL Restriction list | WIP15 | BB_WIP9 WIP19 | No | Subsystem | Configuration | NA | Fixed | venkatachalam_p | 0 | 2 | Service Tools_Software | Support-Security | NO | Closed | SubNC | Ingenuity 64 |  | BB_WIP9 WIP19 | July 25, 2019 1:50:04 PM GMT+08:00
CTUDT00623460 | Refresh | Closed | wang_l | May 5, 2019 4:53:23 PM GMT+08:00 | Refresh:System can access any webset even if the webset is not added in URL Restriction list | WIP15 | BB_WIP9 WIP19 | No | Subsystem | Configuration | NA | Fixed | venkatachalam_p | 0 | 2 | Service Tools_Software | Support-Security | NO | Closed | SubNC | IQon1 |  | BB_WIP9 WIP19 | July 25, 2019 1:50:04 PM GMT+08:00
CTUDT00623458 | Refresh | Closed | ding_f | May 5, 2019 10:55:53 AM GMT+08:00 | Some Reconstruction Filters are not available for selected scan type and Resolution. | WIP15 | ING, BR64 SRS DHF298554 Rev08 iCT, IQon SRS DHF298556 Rev.09 and DHF298290 Rev.07 |  | Subsystem | ECM;Blocker;Verified1 | IQ.1059 | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP20SegReq | Ovation |  | ING, BR64 SRS DHF298554 Rev08 iCT, IQon SRS DHF298556 Rev.09 and DHF298290 Rev.07 | November 5, 2019 2:49:23 PM GMT+08:00
CTUDT00623458 | Refresh | Closed | ding_f | May 5, 2019 10:55:53 AM GMT+08:00 | Some Reconstruction Filters are not available for selected scan type and Resolution. | WIP15 | ING, BR64 SRS DHF298554 Rev08 iCT, IQon SRS DHF298556 Rev.09 and DHF298290 Rev.07 |  | Subsystem | ECM;Blocker;Verified1 | IQ.1059 | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP20SegReq | Brilliance 64 |  | ING, BR64 SRS DHF298554 Rev08 iCT, IQon SRS DHF298556 Rev.09 and DHF298290 Rev.07 | November 5, 2019 2:49:23 PM GMT+08:00
CTUDT00623458 | Refresh | Closed | ding_f | May 5, 2019 10:55:53 AM GMT+08:00 | Some Reconstruction Filters are not available for selected scan type and Resolution. | WIP15 | ING, BR64 SRS DHF298554 Rev08 iCT, IQon SRS DHF298556 Rev.09 and DHF298290 Rev.07 |  | Subsystem | ECM;Blocker;Verified1 | IQ.1059 | Fixed | guy_sheaffer | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP20SegReq | Ingenuity 64 |  | ING, BR64 SRS DHF298554 Rev08 iCT, IQon SRS DHF298556 Rev.09 and DHF298290 Rev.07 | November 5, 2019 2:49:23 PM GMT+08:00
CTUDT00623455 | Refresh | Closed | dadon_e | May 4, 2019 7:41:22 PM GMT+08:00 | Refresh, Pre-Val (WIP15): ICT-SP, "no additional progress information to show for this procedure" - MISSING calibration status and calibration steps during calibration progress | WIP15 |  |  | Subsystem | Calibrations | n/a | Works As Specified / Not A Defect | karkour | 0 |  |  | GCX | NO | Closed |  | Ovation
CTUDT00623455 | Refresh | Closed | dadon_e | May 4, 2019 7:41:22 PM GMT+08:00 | Refresh, Pre-Val (WIP15): ICT-SP, "no additional progress information to show for this procedure" - MISSING calibration status and calibration steps during calibration progress | WIP15 |  |  | Subsystem | Calibrations | n/a | Works As Specified / Not A Defect | karkour | 0 |  |  | GCX | NO | Closed |  | IQon1
CTUDT00623453 | Refresh | Closed | dadon_e | May 4, 2019 7:32:19 PM GMT+08:00 | Refresh, Pre-Val (WIP15): ICT-SP, IQ Check calibration failed | WIP15 |  |  | Subsystem |  |  | Invalid/Uncontrolled Configuration | dadon_e | 1 |  |  | Support-System | NO | Closed |  | Brilliance 64
CTUDT00623453 | Refresh | Closed | dadon_e | May 4, 2019 7:32:19 PM GMT+08:00 | Refresh, Pre-Val (WIP15): ICT-SP, IQ Check calibration failed | WIP15 |  |  | Subsystem |  |  | Invalid/Uncontrolled Configuration | dadon_e | 1 |  |  | Support-System | NO | Closed |  | IQon1
CTUDT00623451 | Refresh | Closed | dadon_e | May 4, 2019 7:24:19 PM GMT+08:00 | Refresh, Pre-Val (WIP15): ICT-SP, Constancy Test calibration failed | WIP15 |  |  | Subsystem |  |  | Invalid/Uncontrolled Configuration | dadon_e | 1 |  |  | Support-System | NO | Closed |  | Brilliance 64
CTUDT00623451 | Refresh | Closed | dadon_e | May 4, 2019 7:24:19 PM GMT+08:00 | Refresh, Pre-Val (WIP15): ICT-SP, Constancy Test calibration failed | WIP15 |  |  | Subsystem |  |  | Invalid/Uncontrolled Configuration | dadon_e | 1 |  |  | Support-System | NO | Closed |  | IQon1
CTUDT00623449 | Refresh | Closed | dadon_e | May 4, 2019 7:07:26 PM GMT+08:00 | Refresh, Pre-Val (WIP15): ICT-SP, Acceptance Test calibration failed | WIP15 |  |  | Subsystem |  |  | Invalid/Uncontrolled Configuration | dadon_e | 1 |  |  | Support-System | NO | Closed |  | Brilliance 64
CTUDT00623449 | Refresh | Closed | dadon_e | May 4, 2019 7:07:26 PM GMT+08:00 | Refresh, Pre-Val (WIP15): ICT-SP, Acceptance Test calibration failed | WIP15 |  |  | Subsystem |  |  | Invalid/Uncontrolled Configuration | dadon_e | 1 |  |  | Support-System | NO | Closed |  | IQon1
CTUDT00623447 | Refresh | Closed | nayak_r | May 4, 2019 1:46:41 AM GMT+08:00 | BB Refresh: US-DoD Config and I2P Configuration Tool is not available in service user account | BB_WIP5.1 |  |  | Subsystem |  |  | Invalid/Uncontrolled Configuration | bollineni_d | 0 |  |  | Service Tools | NO | Closed |  | Brilliance Big Bore
CTUDT00623445 | Refresh | Closed | nayak_r | May 4, 2019 1:23:26 AM GMT+08:00 | BB Refresh: Software Distribution tool is not available in Olevel user account | BB_WIP5.1 | BB_WIP7 |  | Subsystem | Service Tools | STPS-SSRS-1183 | Fixed | jose_s | 0 | 2 | Console Software | Service Tools | NO | Closed | BB_WIP7_SubCom | Brilliance Big Bore |  | BB_WIP7 | May 10, 2019 5:18:09 PM GMT+08:00
CTUDT00623436 | Refresh | Closed | 320053770 | May 3, 2019 12:17:06 AM GMT+08:00 | During integration for WIP 15, in a physics analysis test, there was an error in the physics analysis tool. | WIP 15 & 16 | WIP19 |  | Subsystem | Physics Analysis | n/a | Fixed | yatziv_l | 0 | 2 | Service Tools_Software | DMS | NO | Closed | WIP18SubCom | Ovation |  | WIP19 | July 14, 2019 4:48:28 PM GMT+08:00
CTUDT00623436 | Refresh | Closed | 320053770 | May 3, 2019 12:17:06 AM GMT+08:00 | During integration for WIP 15, in a physics analysis test, there was an error in the physics analysis tool. | WIP 15 & 16 | WIP19 |  | Subsystem | Physics Analysis | n/a | Fixed | yatziv_l | 0 | 2 | Service Tools_Software | DMS | NO | Closed | WIP18SubCom | Brilliance Big Bore |  | WIP19 | July 14, 2019 4:48:28 PM GMT+08:00
CTUDT00623436 | Refresh | Closed | 320053770 | May 3, 2019 12:17:06 AM GMT+08:00 | During integration for WIP 15, in a physics analysis test, there was an error in the physics analysis tool. | WIP 15 & 16 | WIP19 |  | Subsystem | Physics Analysis | n/a | Fixed | yatziv_l | 0 | 2 | Service Tools_Software | DMS | NO | Closed | WIP18SubCom | Ingenuity 64 |  | WIP19 | July 14, 2019 4:48:28 PM GMT+08:00
CTUDT00623436 | Refresh | Closed | 320053770 | May 3, 2019 12:17:06 AM GMT+08:00 | During integration for WIP 15, in a physics analysis test, there was an error in the physics analysis tool. | WIP 15 & 16 | WIP19 |  | Subsystem | Physics Analysis | n/a | Fixed | yatziv_l | 0 | 2 | Service Tools_Software | DMS | NO | Closed | WIP18SubCom | IQon1 |  | WIP19 | July 14, 2019 4:48:28 PM GMT+08:00
CTUDT00623434 | Refresh | Closed | dadon_e | May 2, 2019 9:19:19 PM GMT+08:00 | Refresh, Pre-Val (WIP15): ICT-SP, performance infant calibration failed | WIP15 |  |  | Subsystem |  |  | Invalid/Uncontrolled Configuration | dadon_e | 1 |  |  | Support-System | NO | Closed |  | Brilliance 64
CTUDT00623432 | Refresh | Closed | dadon_e | May 2, 2019 8:57:53 PM GMT+08:00 | Refresh Pre-Val:(WIP15 - iCT , WIP 16 IQON) HCOR slice adult failed on case 84 | WIP16 | WIP18 | N/A | Subsystem | Calibrations | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP18SubCom | IQon1 |  | WIP18 | June 30, 2019 3:46:09 PM GMT+08:00
CTUDT00623430 | Refresh | Closed | 320053770 | May 2, 2019 8:27:26 PM GMT+08:00 | During integration for WIP 15, in a Visual Diagnostics test – X-Ray test failed. | WIP 15 & WIP 16 | WIP20 |  | Subsystem | Visual Diagnostics | NA | Fixed | Sandeep_Kumar | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP18SubCom | Ovation |  | WIP20 | August 30, 2019 7:50:57 PM GMT+08:00
CTUDT00623430 | Refresh | Closed | 320053770 | May 2, 2019 8:27:26 PM GMT+08:00 | During integration for WIP 15, in a Visual Diagnostics test – X-Ray test failed. | WIP 15 & WIP 16 | WIP20 |  | Subsystem | Visual Diagnostics | NA | Fixed | Sandeep_Kumar | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP18SubCom | IQon1 |  | WIP20 | August 30, 2019 7:50:57 PM GMT+08:00
CTUDT00623426 | Refresh | Closed | 320053770 | May 2, 2019 8:11:09 PM GMT+08:00 | During integration for WIP 15, in a Visual Diagnostics test – Couch test failed. ICT & IQon | WIP15 & WIP 16 |  |  | Subsystem | Installation |  | Duplicate | sanjeeva_a | 0 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00623426 | Refresh | Closed | 320053770 | May 2, 2019 8:11:09 PM GMT+08:00 | During integration for WIP 15, in a Visual Diagnostics test – Couch test failed. ICT & IQon | WIP15 & WIP 16 |  |  | Subsystem | Installation |  | Duplicate | sanjeeva_a | 0 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00623424 | Refresh | Closed | 320053770 | May 2, 2019 7:47:54 PM GMT+08:00 | During installation of WIP 16, The NTP server didn’t sync automatically the CIRS & GPC with the HOST. | WIP16 |  |  | Subsystem | Installation | NA | Duplicate | kr_m | 0 |  |  | Console4 | NO | Closed | WAS | Ovation
CTUDT00623424 | Refresh | Closed | 320053770 | May 2, 2019 7:47:54 PM GMT+08:00 | During installation of WIP 16, The NTP server didn’t sync automatically the CIRS & GPC with the HOST. | WIP16 |  |  | Subsystem | Installation | NA | Duplicate | kr_m | 0 |  |  | Console4 | NO | Closed | WAS | Ingenuity 64
CTUDT00623422 | Refresh | Closed | shterner_a | May 2, 2019 6:48:04 PM GMT+08:00 | Pre-Val Refresh: Failed to get session from the service manager message popped up when attempting to log in as third party service user | WIP15 |  |  | Subsystem | UI Performance |  | Duplicate | babu_s | 0 |  |  | Service Tools | NO | Closed |  | IQon1
CTUDT00623418 | Refresh | Closed | g_di | May 2, 2019 6:05:16 PM GMT+08:00 | Blank screen is appearing in the login page after restarting the system | WIP15 |  |  | Subsystem | Installation | NA | Duplicate | kapadia_v | 0 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00623418 | Refresh | Closed | g_di | May 2, 2019 6:05:16 PM GMT+08:00 | Blank screen is appearing in the login page after restarting the system | WIP15 |  |  | Subsystem | Installation | NA | Duplicate | kapadia_v | 0 |  |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00623418 | Refresh | Closed | g_di | May 2, 2019 6:05:16 PM GMT+08:00 | Blank screen is appearing in the login page after restarting the system | WIP15 |  |  | Subsystem | Installation | NA | Duplicate | kapadia_v | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00623418 | Refresh | Closed | g_di | May 2, 2019 6:05:16 PM GMT+08:00 | Blank screen is appearing in the login page after restarting the system | WIP15 |  |  | Subsystem | Installation | NA | Duplicate | kapadia_v | 0 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00623415 | Refresh | Closed | 320053770 | May 2, 2019 5:51:34 PM GMT+08:00 | WIP 16 GPC installation File doesn't work. | WIP15 | WIP18 |  | Subsystem | Installation | NA | Fixed | kline_w | 0 | 2 | Gantry Software | GCX | NO | Closed | WIP17SegReq | Ovation |  | WIP18 | July 9, 2019 11:13:34 PM GMT+08:00
CTUDT00623408 | Refresh | Closed | attali_s | May 2, 2019 2:00:13 PM GMT+08:00 | Refresh: Allowed Combinations: AllowedCombination cfgs in spec and in scanner are different | WIP14 | N/A - Test Design |  | Subsystem | Calibrations | n/a | Fixed | yosfan_s | 0 | 2 | Test Design | DMS | NO | Closed | WIP17SegReq | Ovation |  | N/A - Test Design | July 11, 2019 2:24:13 PM GMT+08:00
CTUDT00623406 | Refresh | Closed | swamypreethi | May 2, 2019 1:50:11 PM GMT+08:00 | IOCC_Refresh: Many new DICOM printers are missing in configuration tool. | WIP13 |  | Use current printers | Subsystem | requirement | NA | Change Proposal | bandil_j | 0 | 2 |  | Support-System | NO | Closed | SegRec | Ovation