CTUDT00623406 | Refresh | Closed | swamypreethi | May 2, 2019 1:50:11 PM GMT+08:00 | IOCC_Refresh: Many new DICOM printers are missing in configuration tool. | WIP13 |  | Use current printers | Subsystem | requirement | NA | Change Proposal | bandil_j | 0 | 2 |  | Support-System | NO | Closed | SegRec | Brilliance 64
CTUDT00623406 | Refresh | Closed | swamypreethi | May 2, 2019 1:50:11 PM GMT+08:00 | IOCC_Refresh: Many new DICOM printers are missing in configuration tool. | WIP13 |  | Use current printers | Subsystem | requirement | NA | Change Proposal | bandil_j | 0 | 2 |  | Support-System | NO | Closed | SegRec | Brilliance Big Bore
CTUDT00623406 | Refresh | Closed | swamypreethi | May 2, 2019 1:50:11 PM GMT+08:00 | IOCC_Refresh: Many new DICOM printers are missing in configuration tool. | WIP13 |  | Use current printers | Subsystem | requirement | NA | Change Proposal | bandil_j | 0 | 2 |  | Support-System | NO | Closed | SegRec | Ingenuity 64
CTUDT00623406 | Refresh | Closed | swamypreethi | May 2, 2019 1:50:11 PM GMT+08:00 | IOCC_Refresh: Many new DICOM printers are missing in configuration tool. | WIP13 |  | Use current printers | Subsystem | requirement | NA | Change Proposal | bandil_j | 0 | 2 |  | Support-System | NO | Closed | SegRec | IQon1
CTUDT00623403 | Refresh | Closed | nayak_r | May 2, 2019 3:05:00 AM GMT+08:00 | BB Refresh: FRU Help Options are disabled in Visual Diagnostics tool | BB_WIP5.1 | WIP18 |  | Subsystem | Service Tools | N/A | Fixed | bollineni_d | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WAS | Ovation |  | WIP18 | June 7, 2019 7:53:03 PM GMT+08:00
CTUDT00623403 | Refresh | Closed | nayak_r | May 2, 2019 3:05:00 AM GMT+08:00 | BB Refresh: FRU Help Options are disabled in Visual Diagnostics tool | BB_WIP5.1 | WIP18 |  | Subsystem | Service Tools | N/A | Fixed | bollineni_d | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WAS | Brilliance 64 |  | WIP18 | June 7, 2019 7:53:03 PM GMT+08:00
CTUDT00623403 | Refresh | Closed | nayak_r | May 2, 2019 3:05:00 AM GMT+08:00 | BB Refresh: FRU Help Options are disabled in Visual Diagnostics tool | BB_WIP5.1 | WIP18 |  | Subsystem | Service Tools | N/A | Fixed | bollineni_d | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WAS | Brilliance Big Bore |  | WIP18 | June 7, 2019 7:53:03 PM GMT+08:00
CTUDT00623403 | Refresh | Closed | nayak_r | May 2, 2019 3:05:00 AM GMT+08:00 | BB Refresh: FRU Help Options are disabled in Visual Diagnostics tool | BB_WIP5.1 | WIP18 |  | Subsystem | Service Tools | N/A | Fixed | bollineni_d | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WAS | Ingenuity 64 |  | WIP18 | June 7, 2019 7:53:03 PM GMT+08:00
CTUDT00623403 | Refresh | Closed | nayak_r | May 2, 2019 3:05:00 AM GMT+08:00 | BB Refresh: FRU Help Options are disabled in Visual Diagnostics tool | BB_WIP5.1 | WIP18 |  | Subsystem | Service Tools | N/A | Fixed | bollineni_d | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WAS | IQon1 |  | WIP18 | June 7, 2019 7:53:03 PM GMT+08:00
CTUDT00623401 | Refresh | Closed | nayak_r | May 2, 2019 2:20:43 AM GMT+08:00 | BB Refresh: When couch is at maximum IN/OUT position, corresponding IN/OUT increment button did not work | BB_WIP5.1 |  |  | Subsystem | Couch | NA | Works As Specified / Not A Defect | karkour | 0 |  |  | GCX | NO | Closed |  | Brilliance Big Bore
CTUDT00623399 | Refresh | Closed | levy_s | May 1, 2019 8:08:26 PM GMT+08:00 | Refresh: “Rotating Frame cannot comply. Please Retry" message appeared when trying to scan Head protocol with certain parameters | WIP15 |  |  | Subsystem | Gantry |  | Works As Specified / Not A Defect | tamari_o | 1 |  |  | Support-System | NO | Closed |  | Ovation
CTUDT00623397 | Refresh | Closed | nayak_r | May 1, 2019 3:16:42 AM GMT+08:00 | BB Refresh: EPHI informations like Content Time and Study ID were displayed in log viewer incorrectly | BB_WIP5 | DHF330449 Rev 08 Host verification Test protocol |  | Subsystem | Host | NA | Fixed | vinoliya_j | 0 | 2 | Test Design | Console4 | NO | Closed | BB_WIP7_SubCom | Brilliance Big Bore |  | DHF330449 Rev 08 Host verification Test protocol | June 27, 2019 9:20:37 PM GMT+08:00
CTUDT00623395 | Refresh | Closed | shekhar_p | May 1, 2019 2:46:11 AM GMT+08:00 | Refresh:Console is taking more than 180 seconds to burn 200 512x512 images | WIP15 | Host SSRS DHF303211 Rev 12 |  | Subsystem | Performance ;Blocker | NA | Fixed | s_sa | 0 | 4 | Host SSRS | iCAP | NO | Closed | WIP17SegReq | Ovation |  | Host SSRS DHF303211 Rev 12 | July 3, 2019 6:35:46 PM GMT+08:00
CTUDT00623395 | Refresh | Closed | shekhar_p | May 1, 2019 2:46:11 AM GMT+08:00 | Refresh:Console is taking more than 180 seconds to burn 200 512x512 images | WIP15 | Host SSRS DHF303211 Rev 12 |  | Subsystem | Performance ;Blocker | NA | Fixed | s_sa | 0 | 4 | Host SSRS | iCAP | NO | Closed | WIP17SegReq | Brilliance 64 |  | Host SSRS DHF303211 Rev 12 | July 3, 2019 6:35:46 PM GMT+08:00
CTUDT00623395 | Refresh | Closed | shekhar_p | May 1, 2019 2:46:11 AM GMT+08:00 | Refresh:Console is taking more than 180 seconds to burn 200 512x512 images | WIP15 | Host SSRS DHF303211 Rev 12 |  | Subsystem | Performance ;Blocker | NA | Fixed | s_sa | 0 | 4 | Host SSRS | iCAP | NO | Closed | WIP17SegReq | Brilliance Big Bore |  | Host SSRS DHF303211 Rev 12 | July 3, 2019 6:35:46 PM GMT+08:00
CTUDT00623395 | Refresh | Closed | shekhar_p | May 1, 2019 2:46:11 AM GMT+08:00 | Refresh:Console is taking more than 180 seconds to burn 200 512x512 images | WIP15 | Host SSRS DHF303211 Rev 12 |  | Subsystem | Performance ;Blocker | NA | Fixed | s_sa | 0 | 4 | Host SSRS | iCAP | NO | Closed | WIP17SegReq | Ingenuity 64 |  | Host SSRS DHF303211 Rev 12 | July 3, 2019 6:35:46 PM GMT+08:00
CTUDT00623395 | Refresh | Closed | shekhar_p | May 1, 2019 2:46:11 AM GMT+08:00 | Refresh:Console is taking more than 180 seconds to burn 200 512x512 images | WIP15 | Host SSRS DHF303211 Rev 12 |  | Subsystem | Performance ;Blocker | NA | Fixed | s_sa | 0 | 4 | Host SSRS | iCAP | NO | Closed | WIP17SegReq | IQon1 |  | Host SSRS DHF303211 Rev 12 | July 3, 2019 6:35:46 PM GMT+08:00
CTUDT00623393 | Refresh | Closed | bommiReddy_r | April 30, 2019 10:07:01 PM GMT+08:00 | Refresh: Spectral: KeV field is accepting decimal point for the Spectral result type ‘MonoE’ | WIP15 | WIP17 |  | Subsystem | ECM | N/A | Fixed | chokkalingam_r | 0 | 4 | Console Software | Console4 | NO | Closed | WIP17SegReq | Ovation |  | WIP17 | May 24, 2019 2:18:12 PM GMT+08:00
CTUDT00623393 | Refresh | Closed | bommiReddy_r | April 30, 2019 10:07:01 PM GMT+08:00 | Refresh: Spectral: KeV field is accepting decimal point for the Spectral result type ‘MonoE’ | WIP15 | WIP17 |  | Subsystem | ECM | N/A | Fixed | chokkalingam_r | 0 | 4 | Console Software | Console4 | NO | Closed | WIP17SegReq | Brilliance 64 |  | WIP17 | May 24, 2019 2:18:12 PM GMT+08:00
CTUDT00623393 | Refresh | Closed | bommiReddy_r | April 30, 2019 10:07:01 PM GMT+08:00 | Refresh: Spectral: KeV field is accepting decimal point for the Spectral result type ‘MonoE’ | WIP15 | WIP17 |  | Subsystem | ECM | N/A | Fixed | chokkalingam_r | 0 | 4 | Console Software | Console4 | NO | Closed | WIP17SegReq | Ingenuity 64 |  | WIP17 | May 24, 2019 2:18:12 PM GMT+08:00
CTUDT00623393 | Refresh | Closed | bommiReddy_r | April 30, 2019 10:07:01 PM GMT+08:00 | Refresh: Spectral: KeV field is accepting decimal point for the Spectral result type ‘MonoE’ | WIP15 | WIP17 |  | Subsystem | ECM | N/A | Fixed | chokkalingam_r | 0 | 4 | Console Software | Console4 | NO | Closed | WIP17SegReq | IQon1 |  | WIP17 | May 24, 2019 2:18:12 PM GMT+08:00
CTUDT00623390 | Refresh | Closed | Noy_Seelig | April 30, 2019 8:35:37 PM GMT+08:00 | Refresh: Bad Detectors Quick Check NVRAM log doesn't update Succ Pass column | WIP14 | WIP17 |  | Subsystem | Calibrations | NA | Fixed | Kostya_Baklan | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP17SegReq | Ovation |  | WIP17 | May 12, 2019 9:05:13 PM GMT+08:00
CTUDT00623390 | Refresh | Closed | Noy_Seelig | April 30, 2019 8:35:37 PM GMT+08:00 | Refresh: Bad Detectors Quick Check NVRAM log doesn't update Succ Pass column | WIP14 | WIP17 |  | Subsystem | Calibrations | NA | Fixed | Kostya_Baklan | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP17SegReq | IQon1 |  | WIP17 | May 12, 2019 9:05:13 PM GMT+08:00
CTUDT00623387 | Refresh | Closed | jotawar_d | April 30, 2019 7:48:47 PM GMT+08:00 | Application Error after performing EndExam for Calcium Scoring procedure | 4.8 | BB_WIP7 | User must wait for atleast 10 seconds after the scan is complete & wave has arrived & then the user can click on EndExam button | Subsystem | System | Console-SRS.17099 | Fixed | saha_d | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP7_SubCom | Brilliance Big Bore | PR5468874 | BB_WIP7 | May 10, 2019 5:40:29 PM GMT+08:00
CTUDT00623383 | Refresh | Closed | 320053770 | April 30, 2019 3:03:00 PM GMT+08:00 | Refresh - During integration for WIP 15, in a Visual Diagnostics test - GPC remote Remote viewer tool failed. | WIP15 |  |  | Subsystem | Remote Configuration |  | Duplicate | sanjeeva_a | 0 |  |  | Service Tools | NO | Closed |  | Ovation
CTUDT00623375 | Refresh | Closed | panchami_pradeep | April 30, 2019 1:35:27 PM GMT+08:00 | Refresh: PSC: Remote viewer tool displays user name as "Philips_service" while logged in from olevel user. | WIP15 | WIP17 BB_WIP7 |  | Subsystem | Remote Configuration | STPS.SRS-321 | Fixed | prasad_r | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP17SubCom | Ovation |  | WIP17 BB_WIP7 | May 25, 2019 1:16:53 PM GMT+08:00
CTUDT00623375 | Refresh | Closed | panchami_pradeep | April 30, 2019 1:35:27 PM GMT+08:00 | Refresh: PSC: Remote viewer tool displays user name as "Philips_service" while logged in from olevel user. | WIP15 | WIP17 BB_WIP7 |  | Subsystem | Remote Configuration | STPS.SRS-321 | Fixed | prasad_r | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP17SubCom | Brilliance 64 |  | WIP17 BB_WIP7 | May 25, 2019 1:16:53 PM GMT+08:00
CTUDT00623375 | Refresh | Closed | panchami_pradeep | April 30, 2019 1:35:27 PM GMT+08:00 | Refresh: PSC: Remote viewer tool displays user name as "Philips_service" while logged in from olevel user. | WIP15 | WIP17 BB_WIP7 |  | Subsystem | Remote Configuration | STPS.SRS-321 | Fixed | prasad_r | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP17SubCom | Brilliance Big Bore |  | WIP17 BB_WIP7 | May 25, 2019 1:16:53 PM GMT+08:00
CTUDT00623375 | Refresh | Closed | panchami_pradeep | April 30, 2019 1:35:27 PM GMT+08:00 | Refresh: PSC: Remote viewer tool displays user name as "Philips_service" while logged in from olevel user. | WIP15 | WIP17 BB_WIP7 |  | Subsystem | Remote Configuration | STPS.SRS-321 | Fixed | prasad_r | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP17SubCom | Ingenuity 64 |  | WIP17 BB_WIP7 | May 25, 2019 1:16:53 PM GMT+08:00
CTUDT00623375 | Refresh | Closed | panchami_pradeep | April 30, 2019 1:35:27 PM GMT+08:00 | Refresh: PSC: Remote viewer tool displays user name as "Philips_service" while logged in from olevel user. | WIP15 | WIP17 BB_WIP7 |  | Subsystem | Remote Configuration | STPS.SRS-321 | Fixed | prasad_r | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP17SubCom | IQon1 |  | WIP17 BB_WIP7 | May 25, 2019 1:16:53 PM GMT+08:00
CTUDT00623373 | Refresh | Closed | kapadia_v | April 30, 2019 1:33:09 PM GMT+08:00 | Recon Jobs in RQM were not deleted | BB_WIP3 | BB_WIP7 | No | Subsystem | Recon Queue Manager | N/A | Fixed | sagar_v | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP7_SegReq | Brilliance Big Bore |  | BB_WIP7 | May 16, 2019 4:46:15 PM GMT+08:00
CTUDT00623372 | Refresh | Closed | richards_b | April 30, 2019 2:38:36 AM GMT+08:00 | Cardiac first block not processed | 4.7.5 | WIP16 | N/A | Subsystem | Reconstruction | N/A | Fixed | rakic_k | 0 | 3 | CIRS_SW_ImgNodeSvr | CIRS | YES | Closed | WIP16SubCom | IQon1 | Escalation 0112696410 | WIP16 | May 6, 2019 9:56:09 AM GMT+08:00
CTUDT00623370 | Refresh | Closed | shekhar_p | April 30, 2019 12:00:53 AM GMT+08:00 | GPC is not connecting through service tools | WIP15 | WIP18 |  | Subsystem | Service Tools | STPS.SRS-321 | Fixed | prasad_r | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP18SubCom | IQon1 |  | WIP18 | June 26, 2019 3:08:10 PM GMT+08:00
CTUDT00623366 | Refresh | Closed | pandey_a | April 29, 2019 7:45:02 PM GMT+08:00 | Service manual change for Ing64, BR64, BB for Upgrade Instructions. | WIP15 |  |  | Subsystem | Service Manual; | N/A | Does Not Apply | yao_j | 0 |  |  | Support-System | NO | Closed |  | Ovation
CTUDT00623366 | Refresh | Closed | pandey_a | April 29, 2019 7:45:02 PM GMT+08:00 | Service manual change for Ing64, BR64, BB for Upgrade Instructions. | WIP15 |  |  | Subsystem | Service Manual; | N/A | Does Not Apply | yao_j | 0 |  |  | Support-System | NO | Closed |  | Brilliance 64
CTUDT00623366 | Refresh | Closed | pandey_a | April 29, 2019 7:45:02 PM GMT+08:00 | Service manual change for Ing64, BR64, BB for Upgrade Instructions. | WIP15 |  |  | Subsystem | Service Manual; | N/A | Does Not Apply | yao_j | 0 |  |  | Support-System | NO | Closed |  | Brilliance Big Bore
CTUDT00623366 | Refresh | Closed | pandey_a | April 29, 2019 7:45:02 PM GMT+08:00 | Service manual change for Ing64, BR64, BB for Upgrade Instructions. | WIP15 |  |  | Subsystem | Service Manual; | N/A | Does Not Apply | yao_j | 0 |  |  | Support-System | NO | Closed |  | Ingenuity 64
CTUDT00623366 | Refresh | Closed | pandey_a | April 29, 2019 7:45:02 PM GMT+08:00 | Service manual change for Ing64, BR64, BB for Upgrade Instructions. | WIP15 |  |  | Subsystem | Service Manual; | N/A | Does Not Apply | yao_j | 0 |  |  | Support-System | NO | Closed |  | IQon1
CTUDT00623363 | Refresh | Closed | panchami_pradeep | April 29, 2019 6:00:38 PM GMT+08:00 | Refresh: In ct user, Bug rep pop up with system error was displayed while surview and offline recon. | WIP16 | WIP17 | When this happens, System.cfg will be corrupted (cfg recovery mechanism also fails, there is already a defect CTUDT00623503 for this issue.) Manually replacing this file with default file will solve this issue. | Subsystem | System | NA | Fixed | ravindra_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP17SubCom | Ovation |  | WIP17 | May 24, 2019 2:45:49 PM GMT+08:00
CTUDT00623363 | Refresh | Closed | panchami_pradeep | April 29, 2019 6:00:38 PM GMT+08:00 | Refresh: In ct user, Bug rep pop up with system error was displayed while surview and offline recon. | WIP16 | WIP17 | When this happens, System.cfg will be corrupted (cfg recovery mechanism also fails, there is already a defect CTUDT00623503 for this issue.) Manually replacing this file with default file will solve this issue. | Subsystem | System | NA | Fixed | ravindra_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP17SubCom | Brilliance 64 |  | WIP17 | May 24, 2019 2:45:49 PM GMT+08:00
CTUDT00623363 | Refresh | Closed | panchami_pradeep | April 29, 2019 6:00:38 PM GMT+08:00 | Refresh: In ct user, Bug rep pop up with system error was displayed while surview and offline recon. | WIP16 | WIP17 | When this happens, System.cfg will be corrupted (cfg recovery mechanism also fails, there is already a defect CTUDT00623503 for this issue.) Manually replacing this file with default file will solve this issue. | Subsystem | System | NA | Fixed | ravindra_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP17SubCom | Ingenuity 64 |  | WIP17 | May 24, 2019 2:45:49 PM GMT+08:00
CTUDT00623363 | Refresh | Closed | panchami_pradeep | April 29, 2019 6:00:38 PM GMT+08:00 | Refresh: In ct user, Bug rep pop up with system error was displayed while surview and offline recon. | WIP16 | WIP17 | When this happens, System.cfg will be corrupted (cfg recovery mechanism also fails, there is already a defect CTUDT00623503 for this issue.) Manually replacing this file with default file will solve this issue. | Subsystem | System | NA | Fixed | ravindra_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP17SubCom | IQon1 |  | WIP17 | May 24, 2019 2:45:49 PM GMT+08:00
CTUDT00623360 | Refresh | Closed | gT_s | April 29, 2019 5:31:42 PM GMT+08:00 | Refresh : The solid core analyzer displays the error " please update the patterns file with valid structure and inputs" | WIP16 | WIP18 |  | Subsystem | Service Tools | NA | Fixed | venkatachalam_p | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP18SubCom | Ovation |  | WIP18 | June 6, 2019 5:45:06 PM GMT+08:00
CTUDT00623360 | Refresh | Closed | gT_s | April 29, 2019 5:31:42 PM GMT+08:00 | Refresh : The solid core analyzer displays the error " please update the patterns file with valid structure and inputs" | WIP16 | WIP18 |  | Subsystem | Service Tools | NA | Fixed | venkatachalam_p | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP18SubCom | Brilliance 64 |  | WIP18 | June 6, 2019 5:45:06 PM GMT+08:00
CTUDT00623360 | Refresh | Closed | gT_s | April 29, 2019 5:31:42 PM GMT+08:00 | Refresh : The solid core analyzer displays the error " please update the patterns file with valid structure and inputs" | WIP16 | WIP18 |  | Subsystem | Service Tools | NA | Fixed | venkatachalam_p | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP18SubCom | Brilliance Big Bore |  | WIP18 | June 6, 2019 5:45:06 PM GMT+08:00
CTUDT00623360 | Refresh | Closed | gT_s | April 29, 2019 5:31:42 PM GMT+08:00 | Refresh : The solid core analyzer displays the error " please update the patterns file with valid structure and inputs" | WIP16 | WIP18 |  | Subsystem | Service Tools | NA | Fixed | venkatachalam_p | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP18SubCom | Ingenuity 64 |  | WIP18 | June 6, 2019 5:45:06 PM GMT+08:00
CTUDT00623360 | Refresh | Closed | gT_s | April 29, 2019 5:31:42 PM GMT+08:00 | Refresh : The solid core analyzer displays the error " please update the patterns file with valid structure and inputs" | WIP16 | WIP18 |  | Subsystem | Service Tools | NA | Fixed | venkatachalam_p | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP18SubCom | IQon1 |  | WIP18 | June 6, 2019 5:45:06 PM GMT+08:00
CTUDT00623358 | Refresh | Closed | wei_f | April 29, 2019 4:58:51 PM GMT+08:00 | Patient study infomation showed blank in CDR folder when did copy raw data. | BB_WIP5.1 |  |  | Subsystem | Patient Directory | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | Ovation
CTUDT00623358 | Refresh | Closed | wei_f | April 29, 2019 4:58:51 PM GMT+08:00 | Patient study infomation showed blank in CDR folder when did copy raw data. | BB_WIP5.1 |  |  | Subsystem | Patient Directory | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00623358 | Refresh | Closed | wei_f | April 29, 2019 4:58:51 PM GMT+08:00 | Patient study infomation showed blank in CDR folder when did copy raw data. | BB_WIP5.1 |  |  | Subsystem | Patient Directory | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00623358 | Refresh | Closed | wei_f | April 29, 2019 4:58:51 PM GMT+08:00 | Patient study infomation showed blank in CDR folder when did copy raw data. | BB_WIP5.1 |  |  | Subsystem | Patient Directory | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00623358 | Refresh | Closed | wei_f | April 29, 2019 4:58:51 PM GMT+08:00 | Patient study infomation showed blank in CDR folder when did copy raw data. | BB_WIP5.1 |  |  | Subsystem | Patient Directory | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | IQon1
CTUDT00623356 | Refresh | Closed | lei_yang | April 29, 2019 3:55:57 PM GMT+08:00 | Scan delay time is greater than 2 seconds when executing a Resp Breath Hold EC. | BB_WIP5 | DHF298555 Rev08 |  | Subsystem | Performance;Blocker | CT-NM.SRS.Modes.520; GCX.SSRS.iCT.17140 | Fixed | wang_k | 0 | 2 | SRS | Support-System | NO | Closed | BB_WIP8_SubCom | Brilliance Big Bore |  | DHF298555 Rev08 | November 4, 2019 4:42:01 PM GMT+08:00
CTUDT00623354 | Refresh | Closed | lei_yang | April 29, 2019 3:42:35 PM GMT+08:00 | The CCT Single and CCT Continuous display "Scanning..." in the scan indication area after pressing the pedal. | BB_WIP5 | BB_WIP8 |  | Subsystem | CCT | NA | Fixed | veeraswamy_k | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP7_SubCom | Brilliance Big Bore |  | BB_WIP8 | June 11, 2019 8:27:50 PM GMT+08:00
CTUDT00623352 | Refresh | Closed | lei_yang | April 29, 2019 3:10:49 PM GMT+08:00 | The Options for "Reference Series" display "Surview" and "Clinical scan" during Clinical scan reconstructing and scanning. | BB_WIP5 | BB_WIP8 DHF316705 Rev11 |  | Subsystem | CCT | Console-SRS.10232 | Fixed | roy_r | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP8_SubCom | Brilliance Big Bore |  | BB_WIP8 DHF316705 Rev11 | June 17, 2019 5:29:23 PM GMT+08:00
CTUDT00623348 | Refresh | Closed | levy_s | April 29, 2019 3:10:33 AM GMT+08:00 | Refresh: Scanning Dual Surview is impossible in Gen5.1 | wip15 |  |  | Subsystem | CIRS |  | Invalid/Uncontrolled Configuration | richards_b | 0 |  |  | CIRS | NO | Closed |  | IQon1
CTUDT00623346 | Refresh | Closed | Valeria_Kashtanov | April 29, 2019 2:34:21 AM GMT+08:00 | Refresh: Allowed Combinations: Workflow initialization error when using wild card "*" in allowed combinations cfg | BB_WIP5 | N/A |  | Subsystem | DMS | n/a | Fixed | yosfan_s | 0 | 2 | Test Design | DMS | NO | Closed | BB_WIP8_SubCom | Brilliance Big Bore |  | N/A | June 11, 2019 1:54:56 PM GMT+08:00
CTUDT00623344 | Refresh | Closed | Valeria_Kashtanov | April 29, 2019 12:19:14 AM GMT+08:00 | Refresh: Physics Analysis: Reconstruction Stage of Bad Detectors / Pixel Correction doesn't appear in MDU after checking in prep tab | BB_WIP5 | WIP18 |  | Subsystem | Physics Analysis | n/a | Fixed | bikman | 0 | 3 | DMS Software | DMS | NO | Closed | WIP18SubCom | Ovation |  | WIP18 | June 13, 2019 2:55:43 PM GMT+08:00
CTUDT00623344 | Refresh | Closed | Valeria_Kashtanov | April 29, 2019 12:19:14 AM GMT+08:00 | Refresh: Physics Analysis: Reconstruction Stage of Bad Detectors / Pixel Correction doesn't appear in MDU after checking in prep tab | BB_WIP5 | WIP18 |  | Subsystem | Physics Analysis | n/a | Fixed | bikman | 0 | 3 | DMS Software | DMS | NO | Closed | WIP18SubCom | Brilliance Big Bore |  | WIP18 | June 13, 2019 2:55:43 PM GMT+08:00
CTUDT00623344 | Refresh | Closed | Valeria_Kashtanov | April 29, 2019 12:19:14 AM GMT+08:00 | Refresh: Physics Analysis: Reconstruction Stage of Bad Detectors / Pixel Correction doesn't appear in MDU after checking in prep tab | BB_WIP5 | WIP18 |  | Subsystem | Physics Analysis | n/a | Fixed | bikman | 0 | 3 | DMS Software | DMS | NO | Closed | WIP18SubCom | Ingenuity 64 |  | WIP18 | June 13, 2019 2:55:43 PM GMT+08:00
CTUDT00623344 | Refresh | Closed | Valeria_Kashtanov | April 29, 2019 12:19:14 AM GMT+08:00 | Refresh: Physics Analysis: Reconstruction Stage of Bad Detectors / Pixel Correction doesn't appear in MDU after checking in prep tab | BB_WIP5 | WIP18 |  | Subsystem | Physics Analysis | n/a | Fixed | bikman | 0 | 3 | DMS Software | DMS | NO | Closed | WIP18SubCom | IQon1 |  | WIP18 | June 13, 2019 2:55:43 PM GMT+08:00
CTUDT00623342 | Refresh | Closed | attali_s | April 28, 2019 9:51:22 PM GMT+08:00 | Refresh: FEE Advanced Mode: Rotation Angle of First Frame is not equal to Rotor Start Pos chosen in FEE Advanced Mode | WIP14 | WIP17 | n/a | Subsystem | Calibrations | n/a | Fixed | bikman | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP17SegReq | Ovation |  | WIP17 | May 26, 2019 10:28:10 PM GMT+08:00
CTUDT00623340 | Refresh | Closed | wang_l | April 28, 2019 7:27:34 PM GMT+08:00 | Refresh BB: Failed to create converter for Brilliance Big Bore 4.2.X during use Upgrade Utility tool in BB 4.8 | WIP5.1 | BB_WIP6 |  | Subsystem | Installation | NA | Fixed | navale_s | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | BB_WIP6_SubCom | Brilliance Big Bore |  | BB_WIP6 | May 3, 2019 7:47:29 PM GMT+08:00
CTUDT00623338 | Refresh | Closed | wang_l | April 28, 2019 6:17:34 PM GMT+08:00 | Refresh: User added voice language cannot be upgraded | WIP15 | WIP20 | Add a language and phrases again after upgarde | Subsystem | Installation; Blocker | NA | Fixed | chakraborty_d | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP17SubCom | Ovation |  | WIP20 | August 13, 2019 3:58:00 PM GMT+08:00
CTUDT00623338 | Refresh | Closed | wang_l | April 28, 2019 6:17:34 PM GMT+08:00 | Refresh: User added voice language cannot be upgraded | WIP15 | WIP20 | Add a language and phrases again after upgarde | Subsystem | Installation; Blocker | NA | Fixed | chakraborty_d | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP17SubCom | Brilliance 64 |  | WIP20 | August 13, 2019 3:58:00 PM GMT+08:00
CTUDT00623338 | Refresh | Closed | wang_l | April 28, 2019 6:17:34 PM GMT+08:00 | Refresh: User added voice language cannot be upgraded | WIP15 | WIP20 | Add a language and phrases again after upgarde | Subsystem | Installation; Blocker | NA | Fixed | chakraborty_d | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP17SubCom | Ingenuity 64 |  | WIP20 | August 13, 2019 3:58:00 PM GMT+08:00
CTUDT00623338 | Refresh | Closed | wang_l | April 28, 2019 6:17:34 PM GMT+08:00 | Refresh: User added voice language cannot be upgraded | WIP15 | WIP20 | Add a language and phrases again after upgarde | Subsystem | Installation; Blocker | NA | Fixed | chakraborty_d | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP17SubCom | IQon1 |  | WIP20 | August 13, 2019 3:58:00 PM GMT+08:00
CTUDT00623335 | Refresh | Closed | wei_f | April 27, 2019 4:51:00 PM GMT+08:00 | Olevel service user could not launch service tool and had the error ''Failed to get session from the service manager'' when Olevel service user logged in. | BB_WIP5 |  | Yes. User needs to restart the system. Once the system is restarted, the faulted services will be restart and communication will not be failed. | Subsystem | Test Design | NA | Works As Specified / Not A Defect | wei_f | 0 | 4 | Test Design | Support-System | NO | Closed | WIP18SubCom | Ovation |  |  | November 11, 2019 3:57:53 PM GMT+08:00
CTUDT00623335 | Refresh | Closed | wei_f | April 27, 2019 4:51:00 PM GMT+08:00 | Olevel service user could not launch service tool and had the error ''Failed to get session from the service manager'' when Olevel service user logged in. | BB_WIP5 |  | Yes. User needs to restart the system. Once the system is restarted, the faulted services will be restart and communication will not be failed. | Subsystem | Test Design | NA | Works As Specified / Not A Defect | wei_f | 0 | 4 | Test Design | Support-System | NO | Closed | WIP18SubCom | Brilliance 64 |  |  | November 11, 2019 3:57:53 PM GMT+08:00
CTUDT00623335 | Refresh | Closed | wei_f | April 27, 2019 4:51:00 PM GMT+08:00 | Olevel service user could not launch service tool and had the error ''Failed to get session from the service manager'' when Olevel service user logged in. | BB_WIP5 |  | Yes. User needs to restart the system. Once the system is restarted, the faulted services will be restart and communication will not be failed. | Subsystem | Test Design | NA | Works As Specified / Not A Defect | wei_f | 0 | 4 | Test Design | Support-System | NO | Closed | WIP18SubCom | Brilliance Big Bore |  |  | November 11, 2019 3:57:53 PM GMT+08:00
CTUDT00623335 | Refresh | Closed | wei_f | April 27, 2019 4:51:00 PM GMT+08:00 | Olevel service user could not launch service tool and had the error ''Failed to get session from the service manager'' when Olevel service user logged in. | BB_WIP5 |  | Yes. User needs to restart the system. Once the system is restarted, the faulted services will be restart and communication will not be failed. | Subsystem | Test Design | NA | Works As Specified / Not A Defect | wei_f | 0 | 4 | Test Design | Support-System | NO | Closed | WIP18SubCom | Ingenuity 64 |  |  | November 11, 2019 3:57:53 PM GMT+08:00
CTUDT00623335 | Refresh | Closed | wei_f | April 27, 2019 4:51:00 PM GMT+08:00 | Olevel service user could not launch service tool and had the error ''Failed to get session from the service manager'' when Olevel service user logged in. | BB_WIP5 |  | Yes. User needs to restart the system. Once the system is restarted, the faulted services will be restart and communication will not be failed. | Subsystem | Test Design | NA | Works As Specified / Not A Defect | wei_f | 0 | 4 | Test Design | Support-System | NO | Closed | WIP18SubCom | IQon1 |  |  | November 11, 2019 3:57:53 PM GMT+08:00
CTUDT00623333 | Refresh | Closed | wei_f | April 27, 2019 4:35:55 PM GMT+08:00 | Log Viewer was not responding and prompted 'The view is currently loading. Refresh after some time...' message. | BB_WIP5 |  | Please wait for the Refresh operation to end. Close and Reopen the logs. Logs remain intact without any loss of data. | Subsystem | Log Viewer | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | Ovation
CTUDT00623333 | Refresh | Closed | wei_f | April 27, 2019 4:35:55 PM GMT+08:00 | Log Viewer was not responding and prompted 'The view is currently loading. Refresh after some time...' message. | BB_WIP5 |  | Please wait for the Refresh operation to end. Close and Reopen the logs. Logs remain intact without any loss of data. | Subsystem | Log Viewer | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00623333 | Refresh | Closed | wei_f | April 27, 2019 4:35:55 PM GMT+08:00 | Log Viewer was not responding and prompted 'The view is currently loading. Refresh after some time...' message. | BB_WIP5 |  | Please wait for the Refresh operation to end. Close and Reopen the logs. Logs remain intact without any loss of data. | Subsystem | Log Viewer | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00623333 | Refresh | Closed | wei_f | April 27, 2019 4:35:55 PM GMT+08:00 | Log Viewer was not responding and prompted 'The view is currently loading. Refresh after some time...' message. | BB_WIP5 |  | Please wait for the Refresh operation to end. Close and Reopen the logs. Logs remain intact without any loss of data. | Subsystem | Log Viewer | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00623333 | Refresh | Closed | wei_f | April 27, 2019 4:35:55 PM GMT+08:00 | Log Viewer was not responding and prompted 'The view is currently loading. Refresh after some time...' message. | BB_WIP5 |  | Please wait for the Refresh operation to end. Close and Reopen the logs. Logs remain intact without any loss of data. | Subsystem | Log Viewer | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | IQon1
CTUDT00623331 | Refresh | Closed | wei_f | April 27, 2019 4:13:31 PM GMT+08:00 | 0008,0070 MANUFACTURER DICOM tag value was not blank in newly created study after de-Identification. | BB_WIP5 | WIP19 BB_WIP9 |  | Subsystem | DICOM | N/A | Fixed | ye_h | 0 | 3 | Requirements Documentation | Support-System | NO | Closed | WIP17SegReq | Ovation |  | WIP19 BB_WIP9 | September 3, 2019 2:50:34 PM GMT+08:00
CTUDT00623331 | Refresh | Closed | wei_f | April 27, 2019 4:13:31 PM GMT+08:00 | 0008,0070 MANUFACTURER DICOM tag value was not blank in newly created study after de-Identification. | BB_WIP5 | WIP19 BB_WIP9 |  | Subsystem | DICOM | N/A | Fixed | ye_h | 0 | 3 | Requirements Documentation | Support-System | NO | Closed | WIP17SegReq | Brilliance Big Bore |  | WIP19 BB_WIP9 | September 3, 2019 2:50:34 PM GMT+08:00
CTUDT00623331 | Refresh | Closed | wei_f | April 27, 2019 4:13:31 PM GMT+08:00 | 0008,0070 MANUFACTURER DICOM tag value was not blank in newly created study after de-Identification. | BB_WIP5 | WIP19 BB_WIP9 |  | Subsystem | DICOM | N/A | Fixed | ye_h | 0 | 3 | Requirements Documentation | Support-System | NO | Closed | WIP17SegReq | Ingenuity 64 |  | WIP19 BB_WIP9 | September 3, 2019 2:50:34 PM GMT+08:00
CTUDT00623331 | Refresh | Closed | wei_f | April 27, 2019 4:13:31 PM GMT+08:00 | 0008,0070 MANUFACTURER DICOM tag value was not blank in newly created study after de-Identification. | BB_WIP5 | WIP19 BB_WIP9 |  | Subsystem | DICOM | N/A | Fixed | ye_h | 0 | 3 | Requirements Documentation | Support-System | NO | Closed | WIP17SegReq | IQon1 |  | WIP19 BB_WIP9 | September 3, 2019 2:50:34 PM GMT+08:00
CTUDT00623327 | Refresh | Closed | nayak_r | April 27, 2019 2:04:33 AM GMT+08:00 | BB Refresh: Dose notification appeared in exam summary/dose log even after pausing scan immediately after the first X-ray on | BB_WIP5 |  |  | Subsystem | Exam Summary - Dose Notification | Console-SRS.15408 | Works As Specified / Not A Defect | saha_d | 0 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00623323 | Refresh | Closed | lakshmikanth_s | April 26, 2019 9:39:27 PM GMT+08:00 | BB Refresh: Unhandled exception error occurred and Restore button is disabled. | BB_WIP5 | BB_WIP8 WIP18 |  | Subsystem | Backup/Restore | NA | Fixed | venkatachalam_p | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP18SubCom | Ovation |  | BB_WIP8 WIP18 | June 11, 2019 5:14:05 PM GMT+08:00
CTUDT00623323 | Refresh | Closed | lakshmikanth_s | April 26, 2019 9:39:27 PM GMT+08:00 | BB Refresh: Unhandled exception error occurred and Restore button is disabled. | BB_WIP5 | BB_WIP8 WIP18 |  | Subsystem | Backup/Restore | NA | Fixed | venkatachalam_p | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP18SubCom | Brilliance 64 |  | BB_WIP8 WIP18 | June 11, 2019 5:14:05 PM GMT+08:00
CTUDT00623323 | Refresh | Closed | lakshmikanth_s | April 26, 2019 9:39:27 PM GMT+08:00 | BB Refresh: Unhandled exception error occurred and Restore button is disabled. | BB_WIP5 | BB_WIP8 WIP18 |  | Subsystem | Backup/Restore | NA | Fixed | venkatachalam_p | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP18SubCom | Brilliance Big Bore |  | BB_WIP8 WIP18 | June 11, 2019 5:14:05 PM GMT+08:00
CTUDT00623323 | Refresh | Closed | lakshmikanth_s | April 26, 2019 9:39:27 PM GMT+08:00 | BB Refresh: Unhandled exception error occurred and Restore button is disabled. | BB_WIP5 | BB_WIP8 WIP18 |  | Subsystem | Backup/Restore | NA | Fixed | venkatachalam_p | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP18SubCom | Ingenuity 64 |  | BB_WIP8 WIP18 | June 11, 2019 5:14:05 PM GMT+08:00
CTUDT00623323 | Refresh | Closed | lakshmikanth_s | April 26, 2019 9:39:27 PM GMT+08:00 | BB Refresh: Unhandled exception error occurred and Restore button is disabled. | BB_WIP5 | BB_WIP8 WIP18 |  | Subsystem | Backup/Restore | NA | Fixed | venkatachalam_p | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP18SubCom | IQon1 |  | BB_WIP8 WIP18 | June 11, 2019 5:14:05 PM GMT+08:00
CTUDT00623321 | Refresh | Closed | nayak_r | April 26, 2019 9:32:03 PM GMT+08:00 | BB Refresh: Few Raw data not getting copied to Rawdata tab from Reconstruction tab | BB_WIP5 |  |  | Subsystem | Host | NA | Works As Specified / Not A Defect | s_sa | 1 |  |  | iCAP | NO | Closed |  | Ovation
CTUDT00623321 | Refresh | Closed | nayak_r | April 26, 2019 9:32:03 PM GMT+08:00 | BB Refresh: Few Raw data not getting copied to Rawdata tab from Reconstruction tab | BB_WIP5 |  |  | Subsystem | Host | NA | Works As Specified / Not A Defect | s_sa | 1 |  |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00623321 | Refresh | Closed | nayak_r | April 26, 2019 9:32:03 PM GMT+08:00 | BB Refresh: Few Raw data not getting copied to Rawdata tab from Reconstruction tab | BB_WIP5 |  |  | Subsystem | Host | NA | Works As Specified / Not A Defect | s_sa | 1 |  |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00623321 | Refresh | Closed | nayak_r | April 26, 2019 9:32:03 PM GMT+08:00 | BB Refresh: Few Raw data not getting copied to Rawdata tab from Reconstruction tab | BB_WIP5 |  |  | Subsystem | Host | NA | Works As Specified / Not A Defect | s_sa | 1 |  |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00623321 | Refresh | Closed | nayak_r | April 26, 2019 9:32:03 PM GMT+08:00 | BB Refresh: Few Raw data not getting copied to Rawdata tab from Reconstruction tab | BB_WIP5 |  |  | Subsystem | Host | NA | Works As Specified / Not A Defect | s_sa | 1 |  |  | iCAP | NO | Closed |  | IQon1
CTUDT00623318 | Refresh | Closed | manju_kolkar | April 26, 2019 9:17:46 PM GMT+08:00 | Big Bore Refresh: "The optimal pitch and rotation time at this heart rate are: pitch XX, rotation time=xx" message doesn’t appear in status line. | BB_WIP5 | BB_WIP7 | For Issue# 2: After manually changing the Pitch/Rotation, manually deselect the option "Automatic Pitch and Rotation Time based on HR" | Subsystem | Performance | NA | Fixed | prabhu_m | 0 | 3 | Console Software | Console4 | YES | Closed | WIP17SubCom | Brilliance Big Bore |  | BB_WIP7 | August 2, 2019 7:41:57 PM GMT+08:00
CTUDT00623312 | Refresh | Closed | Sateesh_S | April 26, 2019 8:14:21 PM GMT+08:00 | BigBore Refresh: Certain Filter and resolution combinations are not listed in the IMRTHRESHOLDS cfg file | BB_WIP5 | DHF330449 Rev 07 HOST verification test protocol |  | Subsystem | ECM | NA | Fixed | vinoliya_j | 0 | 3 | Test Design | Console4 | NO | Closed | BB_WIP6_SubCom | Brilliance Big Bore |  | DHF330449 Rev 07 HOST verification test protocol | May 24, 2019 8:58:49 PM GMT+08:00
CTUDT00623307 | Refresh | Closed | nayak_r | April 26, 2019 7:49:43 PM GMT+08:00 | BB Refresh: Injection summary displays incorrect values upon performing a cancel injection operation | BB_WIP5 |  |  | Subsystem | Performance | Console-SRS.17148 | Duplicate | killada_s | 0 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00623307 | Refresh | Closed | nayak_r | April 26, 2019 7:49:43 PM GMT+08:00 | BB Refresh: Injection summary displays incorrect values upon performing a cancel injection operation | BB_WIP5 |  |  | Subsystem | Performance | Console-SRS.17148 | Duplicate | killada_s | 0 |  |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00623307 | Refresh | Closed | nayak_r | April 26, 2019 7:49:43 PM GMT+08:00 | BB Refresh: Injection summary displays incorrect values upon performing a cancel injection operation | BB_WIP5 |  |  | Subsystem | Performance | Console-SRS.17148 | Duplicate | killada_s | 0 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00623307 | Refresh | Closed | nayak_r | April 26, 2019 7:49:43 PM GMT+08:00 | BB Refresh: Injection summary displays incorrect values upon performing a cancel injection operation | BB_WIP5 |  |  | Subsystem | Performance | Console-SRS.17148 | Duplicate | killada_s | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00623307 | Refresh | Closed | nayak_r | April 26, 2019 7:49:43 PM GMT+08:00 | BB Refresh: Injection summary displays incorrect values upon performing a cancel injection operation | BB_WIP5 |  |  | Subsystem | Performance | Console-SRS.17148 | Duplicate | killada_s | 0 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00623305 | Refresh | Closed | nayak_r | April 26, 2019 6:29:14 PM GMT+08:00 | BB Refresh: "Injection Started" message did not appear in the logs | BB_WIP5 | BB_WIP7 | NA | Subsystem | Host | NA | Fixed | kr_m | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP7_SubCom | Brilliance Big Bore |  | BB_WIP7 | May 10, 2019 2:42:39 PM GMT+08:00
CTUDT00623300 | Refresh | Closed | wang_l | April 26, 2019 3:54:30 PM GMT+08:00 | Refresh: The Filament Current should not be displayed for Refresh system | WIP15 |  |  | Subsystem | Real Time Data Tool | NA | Duplicate | v_su | 0 |  |  | Service Tools | NO | Closed |  | Brilliance Big Bore
CTUDT00623300 | Refresh | Closed | wang_l | April 26, 2019 3:54:30 PM GMT+08:00 | Refresh: The Filament Current should not be displayed for Refresh system | WIP15 |  |  | Subsystem | Real Time Data Tool | NA | Duplicate | v_su | 0 |  |  | Service Tools | NO | Closed |  | Ingenuity 64
CTUDT00623298 | Refresh | Closed | wang_l | April 26, 2019 3:34:12 PM GMT+08:00 | Refresh: olevel service login DMS Temperature Monitor tool failed | WIP15 | WIP19 | NA | Subsystem | DSM Temperature Monitoring | NA | Fixed | s_shabeer | 0 | 2 | Console Software | Support-Security | NO | Closed | WIP18SubCom | Brilliance 64 |  | WIP19 | July 22, 2019 8:14:03 PM GMT+08:00
CTUDT00623298 | Refresh | Closed | wang_l | April 26, 2019 3:34:12 PM GMT+08:00 | Refresh: olevel service login DMS Temperature Monitor tool failed | WIP15 | WIP19 | NA | Subsystem | DSM Temperature Monitoring | NA | Fixed | s_shabeer | 0 | 2 | Console Software | Support-Security | NO | Closed | WIP18SubCom | Ingenuity 64 |  | WIP19 | July 22, 2019 8:14:03 PM GMT+08:00
CTUDT00623297 | Refresh | Closed | datta_i | April 26, 2019 11:47:50 AM GMT+08:00 | IV: Solo - Injection summary not present in exam summary | 4.8 | BB_WIP8 | Through additional testing, the workaround is to use the injector's stop button and not the injector's pause button to stop the injector before ending the exam to obtain the injection summary. | Subsystem | Exam Summary | NA | Fixed | datta_i | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP7_SubCom | Brilliance Big Bore |  | BB_WIP8 | June 21, 2019 4:39:07 PM GMT+08:00
CTUDT00623295 | Refresh | Closed | manjunath_c | April 25, 2019 9:52:12 PM GMT+08:00 | Recon Jobs in RQM were not deleted | WIP15 | WIP17 | No | Subsystem | RQM | Console-SRS.9976 | Fixed | sagar_v | 0 | 4 | Console Software | Console4 | NO | Closed | WIP16SegReq | Ovation |  | WIP17 | May 8, 2019 4:56:31 PM GMT+08:00
CTUDT00623295 | Refresh | Closed | manjunath_c | April 25, 2019 9:52:12 PM GMT+08:00 | Recon Jobs in RQM were not deleted | WIP15 | WIP17 | No | Subsystem | RQM | Console-SRS.9976 | Fixed | sagar_v | 0 | 4 | Console Software | Console4 | NO | Closed | WIP16SegReq | Brilliance 64 |  | WIP17 | May 8, 2019 4:56:31 PM GMT+08:00
CTUDT00623295 | Refresh | Closed | manjunath_c | April 25, 2019 9:52:12 PM GMT+08:00 | Recon Jobs in RQM were not deleted | WIP15 | WIP17 | No | Subsystem | RQM | Console-SRS.9976 | Fixed | sagar_v | 0 | 4 | Console Software | Console4 | NO | Closed | WIP16SegReq | Ingenuity 64 |  | WIP17 | May 8, 2019 4:56:31 PM GMT+08:00
CTUDT00623295 | Refresh | Closed | manjunath_c | April 25, 2019 9:52:12 PM GMT+08:00 | Recon Jobs in RQM were not deleted | WIP15 | WIP17 | No | Subsystem | RQM | Console-SRS.9976 | Fixed | sagar_v | 0 | 4 | Console Software | Console4 | NO | Closed | WIP16SegReq | IQon1 |  | WIP17 | May 8, 2019 4:56:31 PM GMT+08:00
CTUDT00623293 | Refresh | Closed | manjunath_c | April 25, 2019 9:46:31 PM GMT+08:00 | Original EC in Suggested Exam cards group didn't exist after editing | WIP15 | WIP18 | Delete the examcard used for editing from suggested examcard group and create the new examcard. Once the new examcard is created, again add the parent examcard to suggested examcard group. | Subsystem | ECM | N/A | Fixed | Prakash_Chand_Yadav | 0 | 3 | Console Software | Console4 | NO | Closed | WIP18SubCom | Ovation |  | WIP18 | June 27, 2019 11:57:58 PM GMT+08:00
CTUDT00623293 | Refresh | Closed | manjunath_c | April 25, 2019 9:46:31 PM GMT+08:00 | Original EC in Suggested Exam cards group didn't exist after editing | WIP15 | WIP18 | Delete the examcard used for editing from suggested examcard group and create the new examcard. Once the new examcard is created, again add the parent examcard to suggested examcard group. | Subsystem | ECM | N/A | Fixed | Prakash_Chand_Yadav | 0 | 3 | Console Software | Console4 | NO | Closed | WIP18SubCom | Brilliance 64 |  | WIP18 | June 27, 2019 11:57:58 PM GMT+08:00
CTUDT00623293 | Refresh | Closed | manjunath_c | April 25, 2019 9:46:31 PM GMT+08:00 | Original EC in Suggested Exam cards group didn't exist after editing | WIP15 | WIP18 | Delete the examcard used for editing from suggested examcard group and create the new examcard. Once the new examcard is created, again add the parent examcard to suggested examcard group. | Subsystem | ECM | N/A | Fixed | Prakash_Chand_Yadav | 0 | 3 | Console Software | Console4 | NO | Closed | WIP18SubCom | Ingenuity 64 |  | WIP18 | June 27, 2019 11:57:58 PM GMT+08:00
CTUDT00623293 | Refresh | Closed | manjunath_c | April 25, 2019 9:46:31 PM GMT+08:00 | Original EC in Suggested Exam cards group didn't exist after editing | WIP15 | WIP18 | Delete the examcard used for editing from suggested examcard group and create the new examcard. Once the new examcard is created, again add the parent examcard to suggested examcard group. | Subsystem | ECM | N/A | Fixed | Prakash_Chand_Yadav | 0 | 3 | Console Software | Console4 | NO | Closed | WIP18SubCom | IQon1 |  | WIP18 | June 27, 2019 11:57:58 PM GMT+08:00
CTUDT00623287 | Refresh | Closed | shekhar_p | April 25, 2019 8:26:10 PM GMT+08:00 | USP logs are missing in the FSE bugrep | WIP15 | WIP17 | Use full bugrep | Subsystem | Bugrep | STPS.SRS-1090 | Fixed | ranjithkumar_m | 0 | 3 | Service Tools_Software | Service Tools | NO | Closed | WIP17SubCom | IQon1 |  | WIP17 | May 20, 2019 7:34:13 PM GMT+08:00
CTUDT00623286 | Refresh | Closed | lakshmikanth_s | April 25, 2019 8:23:51 PM GMT+08:00 | BB_Refresh: Dilution factor value in logs is not matching with the calculated dilution factor value when thickness and increment are different from each other. | BB_WIP5 | DHF330449 Rev 07 Big Bore HOST Verification test protocol |  | Subsystem | IMR | NA | Fixed | vinoliya_j | 0 | 2 | Test Design | Console4 | NO | Closed | BB_WIP7_SubCom | Brilliance Big Bore |  | DHF330449 Rev 07 Big Bore HOST Verification test protocol | May 24, 2019 9:00:09 PM GMT+08:00
CTUDT00623284 | Refresh | Closed | nimma_p | April 25, 2019 7:05:31 PM GMT+08:00 | ING :-> Performance :-> Log viewer not responding. | WIP15 | WIP17 | Wait for the refresh operation to complete. | Subsystem | Log Viewer | NA | Fixed | s_sa | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP17SegReq | Ovation |  | WIP17 | May 27, 2019 1:53:55 PM GMT+08:00
CTUDT00623284 | Refresh | Closed | nimma_p | April 25, 2019 7:05:31 PM GMT+08:00 | ING :-> Performance :-> Log viewer not responding. | WIP15 | WIP17 | Wait for the refresh operation to complete. | Subsystem | Log Viewer | NA | Fixed | s_sa | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP17SegReq | Brilliance 64 |  | WIP17 | May 27, 2019 1:53:55 PM GMT+08:00
CTUDT00623284 | Refresh | Closed | nimma_p | April 25, 2019 7:05:31 PM GMT+08:00 | ING :-> Performance :-> Log viewer not responding. | WIP15 | WIP17 | Wait for the refresh operation to complete. | Subsystem | Log Viewer | NA | Fixed | s_sa | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP17SegReq | Brilliance Big Bore |  | WIP17 | May 27, 2019 1:53:55 PM GMT+08:00
CTUDT00623284 | Refresh | Closed | nimma_p | April 25, 2019 7:05:31 PM GMT+08:00 | ING :-> Performance :-> Log viewer not responding. | WIP15 | WIP17 | Wait for the refresh operation to complete. | Subsystem | Log Viewer | NA | Fixed | s_sa | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP17SegReq | Ingenuity 64 |  | WIP17 | May 27, 2019 1:53:55 PM GMT+08:00
CTUDT00623284 | Refresh | Closed | nimma_p | April 25, 2019 7:05:31 PM GMT+08:00 | ING :-> Performance :-> Log viewer not responding. | WIP15 | WIP17 | Wait for the refresh operation to complete. | Subsystem | Log Viewer | NA | Fixed | s_sa | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP17SegReq | IQon1 |  | WIP17 | May 27, 2019 1:53:55 PM GMT+08:00
CTUDT00623280 | Refresh | Closed | Sateesh_S | April 25, 2019 6:50:23 PM GMT+08:00 | BigBore Refresh: BugRep Viewer and Backup Restore options are displayed in the workflow utilities for Philips Service user | BB_WIP5 |  |  | Subsystem | Utilities | N/A | Works As Specified / Not A Defect | roy_r | 0 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00623280 | Refresh | Closed | Sateesh_S | April 25, 2019 6:50:23 PM GMT+08:00 | BigBore Refresh: BugRep Viewer and Backup Restore options are displayed in the workflow utilities for Philips Service user | BB_WIP5 |  |  | Subsystem | Utilities | N/A | Works As Specified / Not A Defect | roy_r | 0 |  |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00623280 | Refresh | Closed | Sateesh_S | April 25, 2019 6:50:23 PM GMT+08:00 | BigBore Refresh: BugRep Viewer and Backup Restore options are displayed in the workflow utilities for Philips Service user | BB_WIP5 |  |  | Subsystem | Utilities | N/A | Works As Specified / Not A Defect | roy_r | 0 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00623280 | Refresh | Closed | Sateesh_S | April 25, 2019 6:50:23 PM GMT+08:00 | BigBore Refresh: BugRep Viewer and Backup Restore options are displayed in the workflow utilities for Philips Service user | BB_WIP5 |  |  | Subsystem | Utilities | N/A | Works As Specified / Not A Defect | roy_r | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00623280 | Refresh | Closed | Sateesh_S | April 25, 2019 6:50:23 PM GMT+08:00 | BigBore Refresh: BugRep Viewer and Backup Restore options are displayed in the workflow utilities for Philips Service user | BB_WIP5 |  |  | Subsystem | Utilities | N/A | Works As Specified / Not A Defect | roy_r | 0 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00623279 | Refresh | Closed | kr_m | April 25, 2019 6:48:12 PM GMT+08:00 | Refresh: Localization not working for few fields | WIP15 | Test design update |  | Subsystem | Localization | NA | Fixed | nanjunda_s | 0 | 4 | Test Design | Console4 | NO | Closed | SegReq | Ovation |  | Test design update | May 24, 2019 6:21:59 PM GMT+08:00
CTUDT00623279 | Refresh | Closed | kr_m | April 25, 2019 6:48:12 PM GMT+08:00 | Refresh: Localization not working for few fields | WIP15 | Test design update |  | Subsystem | Localization | NA | Fixed | nanjunda_s | 0 | 4 | Test Design | Console4 | NO | Closed | SegReq | Brilliance 64 |  | Test design update | May 24, 2019 6:21:59 PM GMT+08:00
CTUDT00623279 | Refresh | Closed | kr_m | April 25, 2019 6:48:12 PM GMT+08:00 | Refresh: Localization not working for few fields | WIP15 | Test design update |  | Subsystem | Localization | NA | Fixed | nanjunda_s | 0 | 4 | Test Design | Console4 | NO | Closed | SegReq | Brilliance Big Bore |  | Test design update | May 24, 2019 6:21:59 PM GMT+08:00
CTUDT00623279 | Refresh | Closed | kr_m | April 25, 2019 6:48:12 PM GMT+08:00 | Refresh: Localization not working for few fields | WIP15 | Test design update |  | Subsystem | Localization | NA | Fixed | nanjunda_s | 0 | 4 | Test Design | Console4 | NO | Closed | SegReq | Ingenuity 64 |  | Test design update | May 24, 2019 6:21:59 PM GMT+08:00
CTUDT00623279 | Refresh | Closed | kr_m | April 25, 2019 6:48:12 PM GMT+08:00 | Refresh: Localization not working for few fields | WIP15 | Test design update |  | Subsystem | Localization | NA | Fixed | nanjunda_s | 0 | 4 | Test Design | Console4 | NO | Closed | SegReq | IQon1 |  | Test design update | May 24, 2019 6:21:59 PM GMT+08:00
CTUDT00623278 | Refresh | Closed | gT_s | April 25, 2019 6:46:01 PM GMT+08:00 | Refresh : Actual Latency for closing of Edit Results was more than 1000 milliseconds | WIP 15 | WIP19 | no workaround for this scenario | Subsystem | Edit Results | Console-SRS.10110 | Fixed | ravindra_a | 0 | 4 | Console Software | Console4 | NO | Closed | WIP17SegReq | Ovation |  | WIP19 | August 6, 2019 5:35:06 PM GMT+08:00
CTUDT00623278 | Refresh | Closed | gT_s | April 25, 2019 6:46:01 PM GMT+08:00 | Refresh : Actual Latency for closing of Edit Results was more than 1000 milliseconds | WIP 15 | WIP19 | no workaround for this scenario | Subsystem | Edit Results | Console-SRS.10110 | Fixed | ravindra_a | 0 | 4 | Console Software | Console4 | NO | Closed | WIP17SegReq | Brilliance 64 |  | WIP19 | August 6, 2019 5:35:06 PM GMT+08:00
CTUDT00623278 | Refresh | Closed | gT_s | April 25, 2019 6:46:01 PM GMT+08:00 | Refresh : Actual Latency for closing of Edit Results was more than 1000 milliseconds | WIP 15 | WIP19 | no workaround for this scenario | Subsystem | Edit Results | Console-SRS.10110 | Fixed | ravindra_a | 0 | 4 | Console Software | Console4 | NO | Closed | WIP17SegReq | Ingenuity 64 |  | WIP19 | August 6, 2019 5:35:06 PM GMT+08:00
CTUDT00623278 | Refresh | Closed | gT_s | April 25, 2019 6:46:01 PM GMT+08:00 | Refresh : Actual Latency for closing of Edit Results was more than 1000 milliseconds | WIP 15 | WIP19 | no workaround for this scenario | Subsystem | Edit Results | Console-SRS.10110 | Fixed | ravindra_a | 0 | 4 | Console Software | Console4 | NO | Closed | WIP17SegReq | IQon1 |  | WIP19 | August 6, 2019 5:35:06 PM GMT+08:00
CTUDT00623275 | Refresh | Closed | shekhar_p | April 25, 2019 6:31:56 PM GMT+08:00 | Console is not triggering the clinical scan within 1 second from the time the tracker image | WIP15 |  |  | Subsystem | Performance | NA | Works As Specified / Not A Defect | sowmya_k | 0 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00623271 | Refresh | Closed | lakshmikanth_s | April 25, 2019 3:56:08 PM GMT+08:00 | BB refresh : "OnImageReady(), Image arrived, INSTANCE_NUMBER XYZ, for result XXX, recon YYY, imageid ZZZ” Log message not available in Log viewer. | BB_WIP5 | DHF330449 Rev 08 Host verification Test protocol |  | Subsystem |  | Console.SRS-11572 | Fixed | vinoliya_j | 0 | 2 | Test Design | Console4 | NO | Closed | BB_WIP7_SubCom | Brilliance Big Bore |  | DHF330449 Rev 08 Host verification Test protocol | June 27, 2019 9:19:19 PM GMT+08:00
CTUDT00623269 | Refresh | Closed | bommiReddy_r | April 25, 2019 3:34:11 PM GMT+08:00 | Refresh: Security: During Encrypt Entire Drive, after 95% encryption System is being frozen | WIP15 |  | Restart system | Subsystem | Encryption | STPS-SSRS-1304 | No Plan to Fix | veeraswamy_k | 0 | 3 |  | Service Tools | NO | Closed |  | IQon1
CTUDT00623264 | Refresh | Closed | v_sumanth | April 25, 2019 12:33:25 PM GMT+08:00 | The result HR phase label displays cannot be refreshed in time in the Show All page. | 4.8 | BB_WIP7 |  | Subsystem | Planning | NA | Fixed | mY_h | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP7_SubCom | Brilliance Big Bore |  | BB_WIP7 | April 30, 2019 7:08:36 PM GMT+08:00
CTUDT00623263 | Refresh | Closed | panchami_pradeep | April 25, 2019 12:30:53 PM GMT+08:00 | Refresh: In CCT Fluro Aquisition view, actual CTDIvol and DLPvol values are displayed with ~ 14 decimal places. | WIP15 | WIP18 |  | Subsystem |  | Console-SRS.15972 | Fixed | navale_s | 0 | 3 | Console Software | Console4 | NO | Closed | WIP17SubCom | Ovation |  | WIP18 | June 6, 2019 6:18:08 PM GMT+08:00
CTUDT00623263 | Refresh | Closed | panchami_pradeep | April 25, 2019 12:30:53 PM GMT+08:00 | Refresh: In CCT Fluro Aquisition view, actual CTDIvol and DLPvol values are displayed with ~ 14 decimal places. | WIP15 | WIP18 |  | Subsystem |  | Console-SRS.15972 | Fixed | navale_s | 0 | 3 | Console Software | Console4 | NO | Closed | WIP17SubCom | Brilliance 64 |  | WIP18 | June 6, 2019 6:18:08 PM GMT+08:00
CTUDT00623263 | Refresh | Closed | panchami_pradeep | April 25, 2019 12:30:53 PM GMT+08:00 | Refresh: In CCT Fluro Aquisition view, actual CTDIvol and DLPvol values are displayed with ~ 14 decimal places. | WIP15 | WIP18 |  | Subsystem |  | Console-SRS.15972 | Fixed | navale_s | 0 | 3 | Console Software | Console4 | NO | Closed | WIP17SubCom | Ingenuity 64 |  | WIP18 | June 6, 2019 6:18:08 PM GMT+08:00
CTUDT00623263 | Refresh | Closed | panchami_pradeep | April 25, 2019 12:30:53 PM GMT+08:00 | Refresh: In CCT Fluro Aquisition view, actual CTDIvol and DLPvol values are displayed with ~ 14 decimal places. | WIP15 | WIP18 |  | Subsystem |  | Console-SRS.15972 | Fixed | navale_s | 0 | 3 | Console Software | Console4 | NO | Closed | WIP17SubCom | IQon1 |  | WIP18 | June 6, 2019 6:18:08 PM GMT+08:00
CTUDT00623261 | Refresh | Closed | nimma_p | April 25, 2019 12:25:05 PM GMT+08:00 | Refresh:->ING :-> Rules :-> Scan type :-> CDRH QA (Axial Body 2D) EC not displayed for Service, Olevel and R&D users. | WIP15 | Host SSRS DHF303211 Rev 11 |  | Subsystem |  | Console-SRS.16112,Console-SRS.17642 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | SubCom | Ovation |  | Host SSRS DHF303211 Rev 11 | June 7, 2019 8:02:41 PM GMT+08:00
CTUDT00623261 | Refresh | Closed | nimma_p | April 25, 2019 12:25:05 PM GMT+08:00 | Refresh:->ING :-> Rules :-> Scan type :-> CDRH QA (Axial Body 2D) EC not displayed for Service, Olevel and R&D users. | WIP15 | Host SSRS DHF303211 Rev 11 |  | Subsystem |  | Console-SRS.16112,Console-SRS.17642 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | SubCom | Brilliance 64 |  | Host SSRS DHF303211 Rev 11 | June 7, 2019 8:02:41 PM GMT+08:00
CTUDT00623261 | Refresh | Closed | nimma_p | April 25, 2019 12:25:05 PM GMT+08:00 | Refresh:->ING :-> Rules :-> Scan type :-> CDRH QA (Axial Body 2D) EC not displayed for Service, Olevel and R&D users. | WIP15 | Host SSRS DHF303211 Rev 11 |  | Subsystem |  | Console-SRS.16112,Console-SRS.17642 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | SubCom | Ingenuity 64 |  | Host SSRS DHF303211 Rev 11 | June 7, 2019 8:02:41 PM GMT+08:00
CTUDT00623261 | Refresh | Closed | nimma_p | April 25, 2019 12:25:05 PM GMT+08:00 | Refresh:->ING :-> Rules :-> Scan type :-> CDRH QA (Axial Body 2D) EC not displayed for Service, Olevel and R&D users. | WIP15 | Host SSRS DHF303211 Rev 11 |  | Subsystem |  | Console-SRS.16112,Console-SRS.17642 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | SubCom | IQon1 |  | Host SSRS DHF303211 Rev 11 | June 7, 2019 8:02:41 PM GMT+08:00
CTUDT00623258 | Refresh | Closed | bhawna_g | April 24, 2019 10:09:23 PM GMT+08:00 | Refresh: Protocol Group displayed blank in Exam summary | 4.8 | BB_WIP7 |  | Subsystem | DICOM | NA | Fixed | agrawal_ay | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP6_SubCom | Brilliance Big Bore |  | BB_WIP7 | May 10, 2019 12:28:32 PM GMT+08:00
CTUDT00623256 | Refresh | Closed | levy_s | April 24, 2019 10:01:44 PM GMT+08:00 | Refresh: “Gantry acquisition system cannot comply” error appear while scanning Body Perfusion | WIP15 |  |  | Subsystem | DMS | N/A | System Repaired | yatziv_l | 0 |  |  | DMS | YES | Closed |  | Ovation
CTUDT00623250 | Refresh | Closed | ravindra_a | April 23, 2019 7:18:02 PM GMT+08:00 | Refresh: Security: Improper Message is displayed during the Philips service authentication using incorrect entry of IST password | 4.8 | BB_WIP7 |  | Subsystem | Login | Na | Fixed | goswami_s | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP6_SubCom | Brilliance Big Bore |  | BB_WIP7 | July 19, 2019 8:03:06 PM GMT+08:00
CTUDT00623248 | Refresh | Closed | levy_s | April 23, 2019 6:18:43 PM GMT+08:00 | Refresh: While running performance scenarios “ExamManager has stopped working” error prevents the user from executing the scan. | WIP15 | WIP21 |  | System | ECM;Verified1 | NA | Fixed | levy_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP21SubCom | Ovation |  | WIP21 | October 12, 2019 7:01:01 PM GMT+08:00
CTUDT00623248 | Refresh | Closed | levy_s | April 23, 2019 6:18:43 PM GMT+08:00 | Refresh: While running performance scenarios “ExamManager has stopped working” error prevents the user from executing the scan. | WIP15 | WIP21 |  | System | ECM;Verified1 | NA | Fixed | levy_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP21SubCom | Brilliance 64 |  | WIP21 | October 12, 2019 7:01:01 PM GMT+08:00
CTUDT00623248 | Refresh | Closed | levy_s | April 23, 2019 6:18:43 PM GMT+08:00 | Refresh: While running performance scenarios “ExamManager has stopped working” error prevents the user from executing the scan. | WIP15 | WIP21 |  | System | ECM;Verified1 | NA | Fixed | levy_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP21SubCom | Ingenuity 64 |  | WIP21 | October 12, 2019 7:01:01 PM GMT+08:00
CTUDT00623248 | Refresh | Closed | levy_s | April 23, 2019 6:18:43 PM GMT+08:00 | Refresh: While running performance scenarios “ExamManager has stopped working” error prevents the user from executing the scan. | WIP15 | WIP21 |  | System | ECM;Verified1 | NA | Fixed | levy_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP21SubCom | IQon1 |  | WIP21 | October 12, 2019 7:01:01 PM GMT+08:00
CTUDT00623245 | Refresh | Closed | altshuller_mark | April 23, 2019 2:19:41 PM GMT+08:00 | Application Error after performing EndExam for Calcium Scoring procedure | 4.7.5 | WIP17 | User must wait for atleast 10 seconds after the scan is complete & wave has arrived & then the user can click on EndExam button | Subsystem | System | Console-SRS.17099 | Fixed | saha_d | 0 | 3 | Console Software | Console4 | NO | Closed | WIP16SubCom | Ovation | PR5468874 | WIP17 | May 2, 2019 7:05:59 PM GMT+08:00
CTUDT00623245 | Refresh | Closed | altshuller_mark | April 23, 2019 2:19:41 PM GMT+08:00 | Application Error after performing EndExam for Calcium Scoring procedure | 4.7.5 | WIP17 | User must wait for atleast 10 seconds after the scan is complete & wave has arrived & then the user can click on EndExam button | Subsystem | System | Console-SRS.17099 | Fixed | saha_d | 0 | 3 | Console Software | Console4 | NO | Closed | WIP16SubCom | Brilliance 64 | PR5468874 | WIP17 | May 2, 2019 7:05:59 PM GMT+08:00
CTUDT00623245 | Refresh | Closed | altshuller_mark | April 23, 2019 2:19:41 PM GMT+08:00 | Application Error after performing EndExam for Calcium Scoring procedure | 4.7.5 | WIP17 | User must wait for atleast 10 seconds after the scan is complete & wave has arrived & then the user can click on EndExam button | Subsystem | System | Console-SRS.17099 | Fixed | saha_d | 0 | 3 | Console Software | Console4 | NO | Closed | WIP16SubCom | Ingenuity 64 | PR5468874 | WIP17 | May 2, 2019 7:05:59 PM GMT+08:00
CTUDT00623245 | Refresh | Closed | altshuller_mark | April 23, 2019 2:19:41 PM GMT+08:00 | Application Error after performing EndExam for Calcium Scoring procedure | 4.7.5 | WIP17 | User must wait for atleast 10 seconds after the scan is complete & wave has arrived & then the user can click on EndExam button | Subsystem | System | Console-SRS.17099 | Fixed | saha_d | 0 | 3 | Console Software | Console4 | NO | Closed | WIP16SubCom | IQon1 | PR5468874 | WIP17 | May 2, 2019 7:05:59 PM GMT+08:00
CTUDT00623243 | Refresh | Closed | gT_s | April 23, 2019 2:06:38 PM GMT+08:00 | Refresh : The packages are not getting downloaded into the system from the PRS portal. | WIP 15 |  | NA | Subsystem | Software Distribution Tool | STPS-SSRS-1249 | Works As Specified / Not A Defect | jose_s | 0 |  |  | Service Tools | NO | Closed |  | Ingenuity 64
CTUDT00623241 | Refresh | Closed | bandil_j | April 23, 2019 1:02:04 PM GMT+08:00 | RfVD: Plan Box shrinks and moves on reducing FOV of the result | 4.7.5 | WIP17 | Since only one collimation is available for the Calcium Scoring or Step & Shoot Exam Card for which we allow non-full coverage scan so instead of selecting any collimation, user should enable and use auto-collimation. | Subsystem | Plan Viewer | None | Fixed | v_sumanth | 0 | 3 | Console Software | Console4 | NO | Closed | WIP17SubCom | Ovation | PR7390293 | WIP17 | April 29, 2019 1:47:49 PM GMT+08:00
CTUDT00623241 | Refresh | Closed | bandil_j | April 23, 2019 1:02:04 PM GMT+08:00 | RfVD: Plan Box shrinks and moves on reducing FOV of the result | 4.7.5 | WIP17 | Since only one collimation is available for the Calcium Scoring or Step & Shoot Exam Card for which we allow non-full coverage scan so instead of selecting any collimation, user should enable and use auto-collimation. | Subsystem | Plan Viewer | None | Fixed | v_sumanth | 0 | 3 | Console Software | Console4 | NO | Closed | WIP17SubCom | Brilliance 64 | PR7390293 | WIP17 | April 29, 2019 1:47:49 PM GMT+08:00
CTUDT00623241 | Refresh | Closed | bandil_j | April 23, 2019 1:02:04 PM GMT+08:00 | RfVD: Plan Box shrinks and moves on reducing FOV of the result | 4.7.5 | WIP17 | Since only one collimation is available for the Calcium Scoring or Step & Shoot Exam Card for which we allow non-full coverage scan so instead of selecting any collimation, user should enable and use auto-collimation. | Subsystem | Plan Viewer | None | Fixed | v_sumanth | 0 | 3 | Console Software | Console4 | NO | Closed | WIP17SubCom | Ingenuity 64 | PR7390293 | WIP17 | April 29, 2019 1:47:49 PM GMT+08:00
CTUDT00623241 | Refresh | Closed | bandil_j | April 23, 2019 1:02:04 PM GMT+08:00 | RfVD: Plan Box shrinks and moves on reducing FOV of the result | 4.7.5 | WIP17 | Since only one collimation is available for the Calcium Scoring or Step & Shoot Exam Card for which we allow non-full coverage scan so instead of selecting any collimation, user should enable and use auto-collimation. | Subsystem | Plan Viewer | None | Fixed | v_sumanth | 0 | 3 | Console Software | Console4 | NO | Closed | WIP17SubCom | IQon1 | PR7390293 | WIP17 | April 29, 2019 1:47:49 PM GMT+08:00
CTUDT00623231 | Refresh | Closed | shekhar_p | April 22, 2019 11:35:25 PM GMT+08:00 | While pressing go button during clinical scan Recon box error appears | WIP15 |  |  | Subsystem | CIRS |  | Duplicate | richards_b | 0 |  |  | CIRS | NO | Closed |  | IQon1
CTUDT00623228 | Refresh | Closed | ravindra_a | April 22, 2019 8:14:06 PM GMT+08:00 | Refresh : Auto Launch option does not retain the previous selection of the results to different viewer in Online & View2 | 4.8 | BB_WIP6 | User has to refresh the UI by clicking on some other result in the series. | Subsystem | UI Performance | NA | Fixed | kr_m | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP6_SubCom | Brilliance Big Bore |  | BB_WIP6 | April 26, 2019 6:53:22 PM GMT+08:00
CTUDT00623226 | Refresh | Closed | nimma_p | April 22, 2019 8:06:59 PM GMT+08:00 | Refresh :-> ING :-> Right Monitor :-> Raw data patient details not displayed under CDR. | WIP15 |  |  | Subsystem | Patient Directory | Console SRS: 2837 | Duplicate | s_sa | 0 |  |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00623225 | Refresh | Closed | kapadia_v | April 22, 2019 7:59:48 PM GMT+08:00 | Exam Card Manager causes Workflow to crash | BB_WIP3 | BB_WIP6 | NA | Subsystem | ECM | Console-SRS.2143 | Fixed | sagar_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP16SegReq | Ovation |  | BB_WIP6 | April 25, 2019 6:46:51 PM GMT+08:00
CTUDT00623225 | Refresh | Closed | kapadia_v | April 22, 2019 7:59:48 PM GMT+08:00 | Exam Card Manager causes Workflow to crash | BB_WIP3 | BB_WIP6 | NA | Subsystem | ECM | Console-SRS.2143 | Fixed | sagar_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP16SegReq | Brilliance 64 |  | BB_WIP6 | April 25, 2019 6:46:51 PM GMT+08:00
CTUDT00623225 | Refresh | Closed | kapadia_v | April 22, 2019 7:59:48 PM GMT+08:00 | Exam Card Manager causes Workflow to crash | BB_WIP3 | BB_WIP6 | NA | Subsystem | ECM | Console-SRS.2143 | Fixed | sagar_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP16SegReq | Brilliance Big Bore |  | BB_WIP6 | April 25, 2019 6:46:51 PM GMT+08:00
CTUDT00623225 | Refresh | Closed | kapadia_v | April 22, 2019 7:59:48 PM GMT+08:00 | Exam Card Manager causes Workflow to crash | BB_WIP3 | BB_WIP6 | NA | Subsystem | ECM | Console-SRS.2143 | Fixed | sagar_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP16SegReq | IQon1 |  | BB_WIP6 | April 25, 2019 6:46:51 PM GMT+08:00
CTUDT00623222 | Refresh | Closed | levy_s | April 22, 2019 7:45:25 PM GMT+08:00 | Refresh: While executing performance scenario “S_DMC_DATA_BUFFER_FULL” error prevents the user from executing the scan. | WIP14 |  |  | Subsystem | DMS | N/A | System Repaired | yatziv_l | 0 |  |  | DMS | NO | Closed |  | IQon1
CTUDT00623220 | Refresh | Closed | PremKumar_Rajendran | April 22, 2019 7:35:42 PM GMT+08:00 | Refresh - In Norwegian Localized strings (i). Yes/No buttons not localized in Delete Pop Up Messages | WIP15 | WIP17 |  | Subsystem | ECM | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP17SubCom | Ovation |  | WIP17 | May 27, 2019 1:54:22 PM GMT+08:00
CTUDT00623220 | Refresh | Closed | PremKumar_Rajendran | April 22, 2019 7:35:42 PM GMT+08:00 | Refresh - In Norwegian Localized strings (i). Yes/No buttons not localized in Delete Pop Up Messages | WIP15 | WIP17 |  | Subsystem | ECM | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP17SubCom | Brilliance 64 |  | WIP17 | May 27, 2019 1:54:22 PM GMT+08:00
CTUDT00623220 | Refresh | Closed | PremKumar_Rajendran | April 22, 2019 7:35:42 PM GMT+08:00 | Refresh - In Norwegian Localized strings (i). Yes/No buttons not localized in Delete Pop Up Messages | WIP15 | WIP17 |  | Subsystem | ECM | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP17SubCom | Brilliance Big Bore |  | WIP17 | May 27, 2019 1:54:22 PM GMT+08:00
CTUDT00623220 | Refresh | Closed | PremKumar_Rajendran | April 22, 2019 7:35:42 PM GMT+08:00 | Refresh - In Norwegian Localized strings (i). Yes/No buttons not localized in Delete Pop Up Messages | WIP15 | WIP17 |  | Subsystem | ECM | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP17SubCom | Ingenuity 64 |  | WIP17 | May 27, 2019 1:54:22 PM GMT+08:00
CTUDT00623220 | Refresh | Closed | PremKumar_Rajendran | April 22, 2019 7:35:42 PM GMT+08:00 | Refresh - In Norwegian Localized strings (i). Yes/No buttons not localized in Delete Pop Up Messages | WIP15 | WIP17 |  | Subsystem | ECM | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP17SubCom | IQon1 |  | WIP17 | May 27, 2019 1:54:22 PM GMT+08:00
CTUDT00623215 | Refresh | Closed | Sharmila_Vasireddy | April 22, 2019 6:43:27 PM GMT+08:00 | In Big Bore refresh : As per config file “MassCalcuimScoringApplicationMapping.cfg”, the values are not displaying in Dicom tag [0018, 9352] | BB_WIP5 | BB_WIP7 | N/A | Subsystem | DICOM | N/A | Fixed | naveen_k | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP7_SubCom | Brilliance Big Bore |  | BB_WIP7 | May 29, 2019 7:07:56 PM GMT+08:00
CTUDT00623210 | Refresh | Closed | sunilkumar_j | April 22, 2019 3:12:56 PM GMT+08:00 | Clear All and few fields were not localized as per selected language. | WIP15 | WIP18 | NA | Subsystem | Localization | NA | Fixed | saha_d | 0 | 2 | Console Software | Console4 | YES | Closed | WIP17SubCom | Ovation |  | WIP18 | May 31, 2019 5:24:11 PM GMT+08:00
CTUDT00623210 | Refresh | Closed | sunilkumar_j | April 22, 2019 3:12:56 PM GMT+08:00 | Clear All and few fields were not localized as per selected language. | WIP15 | WIP18 | NA | Subsystem | Localization | NA | Fixed | saha_d | 0 | 2 | Console Software | Console4 | YES | Closed | WIP17SubCom | Brilliance 64 |  | WIP18 | May 31, 2019 5:24:11 PM GMT+08:00
CTUDT00623210 | Refresh | Closed | sunilkumar_j | April 22, 2019 3:12:56 PM GMT+08:00 | Clear All and few fields were not localized as per selected language. | WIP15 | WIP18 | NA | Subsystem | Localization | NA | Fixed | saha_d | 0 | 2 | Console Software | Console4 | YES | Closed | WIP17SubCom | Ingenuity 64 |  | WIP18 | May 31, 2019 5:24:11 PM GMT+08:00
CTUDT00623210 | Refresh | Closed | sunilkumar_j | April 22, 2019 3:12:56 PM GMT+08:00 | Clear All and few fields were not localized as per selected language. | WIP15 | WIP18 | NA | Subsystem | Localization | NA | Fixed | saha_d | 0 | 2 | Console Software | Console4 | YES | Closed | WIP17SubCom | IQon1 |  | WIP18 | May 31, 2019 5:24:11 PM GMT+08:00
CTUDT00623208 | Refresh | Closed | wei_f | April 22, 2019 9:37:59 AM GMT+08:00 | System displayed different value after loaded from HISRIS. | BB_WIP5 |  |  | Subsystem | Connectivity | Console SRS-17643, Console SRS-3379 | No Plan to Fix | wei_f | 0 | 3 | Test Design | Support-System | NO | Closed | WIP17SegReq | Ovation |  |  | July 1, 2019 7:26:31 PM GMT+08:00
CTUDT00623208 | Refresh | Closed | wei_f | April 22, 2019 9:37:59 AM GMT+08:00 | System displayed different value after loaded from HISRIS. | BB_WIP5 |  |  | Subsystem | Connectivity | Console SRS-17643, Console SRS-3379 | No Plan to Fix | wei_f | 0 | 3 | Test Design | Support-System | NO | Closed | WIP17SegReq | Brilliance 64 |  |  | July 1, 2019 7:26:31 PM GMT+08:00
CTUDT00623208 | Refresh | Closed | wei_f | April 22, 2019 9:37:59 AM GMT+08:00 | System displayed different value after loaded from HISRIS. | BB_WIP5 |  |  | Subsystem | Connectivity | Console SRS-17643, Console SRS-3379 | No Plan to Fix | wei_f | 0 | 3 | Test Design | Support-System | NO | Closed | WIP17SegReq | Brilliance Big Bore |  |  | July 1, 2019 7:26:31 PM GMT+08:00
CTUDT00623208 | Refresh | Closed | wei_f | April 22, 2019 9:37:59 AM GMT+08:00 | System displayed different value after loaded from HISRIS. | BB_WIP5 |  |  | Subsystem | Connectivity | Console SRS-17643, Console SRS-3379 | No Plan to Fix | wei_f | 0 | 3 | Test Design | Support-System | NO | Closed | WIP17SegReq | Ingenuity 64 |  |  | July 1, 2019 7:26:31 PM GMT+08:00
CTUDT00623208 | Refresh | Closed | wei_f | April 22, 2019 9:37:59 AM GMT+08:00 | System displayed different value after loaded from HISRIS. | BB_WIP5 |  |  | Subsystem | Connectivity | Console SRS-17643, Console SRS-3379 | No Plan to Fix | wei_f | 0 | 3 | Test Design | Support-System | NO | Closed | WIP17SegReq | IQon1 |  |  | July 1, 2019 7:26:31 PM GMT+08:00
CTUDT00623206 | Refresh | Closed | zhang_t | April 22, 2019 8:51:18 AM GMT+08:00 | Refresh: CIRS S1 and S2 logs were missing from Bug report | WIP15 |  |  | Subsystem | BugRep | N/A | Invalid/Uncontrolled Configuration | oren_d | 0 |  |  | Support-System | NO | Closed |  | Ovation
CTUDT00623206 | Refresh | Closed | zhang_t | April 22, 2019 8:51:18 AM GMT+08:00 | Refresh: CIRS S1 and S2 logs were missing from Bug report | WIP15 |  |  | Subsystem | BugRep | N/A | Invalid/Uncontrolled Configuration | oren_d | 0 |  |  | Support-System | NO | Closed |  | Brilliance 64
CTUDT00623206 | Refresh | Closed | zhang_t | April 22, 2019 8:51:18 AM GMT+08:00 | Refresh: CIRS S1 and S2 logs were missing from Bug report | WIP15 |  |  | Subsystem | BugRep | N/A | Invalid/Uncontrolled Configuration | oren_d | 0 |  |  | Support-System | NO | Closed |  | Brilliance Big Bore
CTUDT00623206 | Refresh | Closed | zhang_t | April 22, 2019 8:51:18 AM GMT+08:00 | Refresh: CIRS S1 and S2 logs were missing from Bug report | WIP15 |  |  | Subsystem | BugRep | N/A | Invalid/Uncontrolled Configuration | oren_d | 0 |  |  | Support-System | NO | Closed |  | Ingenuity 64
CTUDT00623206 | Refresh | Closed | zhang_t | April 22, 2019 8:51:18 AM GMT+08:00 | Refresh: CIRS S1 and S2 logs were missing from Bug report | WIP15 |  |  | Subsystem | BugRep | N/A | Invalid/Uncontrolled Configuration | oren_d | 0 |  |  | Support-System | NO | Closed |  | IQon1
CTUDT00623201 | Refresh | Closed | beck_d | April 21, 2019 6:46:53 PM GMT+08:00 | Refresh: CTDI Test: Probe not found warning message is not according to spec | BB_WIP5 | WIP17 |  | Subsystem | Calibrations | N/A | Fixed | daibes_s | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP17SegReq | Ovation |  | WIP17 | May 26, 2019 3:24:34 PM GMT+08:00
CTUDT00623201 | Refresh | Closed | beck_d | April 21, 2019 6:46:53 PM GMT+08:00 | Refresh: CTDI Test: Probe not found warning message is not according to spec | BB_WIP5 | WIP17 |  | Subsystem | Calibrations | N/A | Fixed | daibes_s | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP17SegReq | Brilliance Big Bore |  | WIP17 | May 26, 2019 3:24:34 PM GMT+08:00
CTUDT00623201 | Refresh | Closed | beck_d | April 21, 2019 6:46:53 PM GMT+08:00 | Refresh: CTDI Test: Probe not found warning message is not according to spec | BB_WIP5 | WIP17 |  | Subsystem | Calibrations | N/A | Fixed | daibes_s | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP17SegReq | Ingenuity 64 |  | WIP17 | May 26, 2019 3:24:34 PM GMT+08:00
CTUDT00623201 | Refresh | Closed | beck_d | April 21, 2019 6:46:53 PM GMT+08:00 | Refresh: CTDI Test: Probe not found warning message is not according to spec | BB_WIP5 | WIP17 |  | Subsystem | Calibrations | N/A | Fixed | daibes_s | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP17SegReq | IQon1 |  | WIP17 | May 26, 2019 3:24:34 PM GMT+08:00
CTUDT00623198 | Refresh | Closed | shekhar_p | April 19, 2019 6:03:46 PM GMT+08:00 | View2 stopped working and crashed | WIP15 |  |  | Subsystem | View2 | NA | Duplicate | v_sumanth | 0 | 4 |  | Console4 | NO | Closed |  | Ovation
CTUDT00623198 | Refresh | Closed | shekhar_p | April 19, 2019 6:03:46 PM GMT+08:00 | View2 stopped working and crashed | WIP15 |  |  | Subsystem | View2 | NA | Duplicate | v_sumanth | 0 | 4 |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00623198 | Refresh | Closed | shekhar_p | April 19, 2019 6:03:46 PM GMT+08:00 | View2 stopped working and crashed | WIP15 |  |  | Subsystem | View2 | NA | Duplicate | v_sumanth | 0 | 4 |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00623198 | Refresh | Closed | shekhar_p | April 19, 2019 6:03:46 PM GMT+08:00 | View2 stopped working and crashed | WIP15 |  |  | Subsystem | View2 | NA | Duplicate | v_sumanth | 0 | 4 |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00623198 | Refresh | Closed | shekhar_p | April 19, 2019 6:03:46 PM GMT+08:00 | View2 stopped working and crashed | WIP15 |  |  | Subsystem | View2 | NA | Duplicate | v_sumanth | 0 | 4 |  | Console4 | NO | Closed |  | IQon1
CTUDT00623195 | Refresh | Closed | wei_f | April 19, 2019 4:44:35 PM GMT+08:00 | Cannot do scan after data encryption finished on BBwip5.1 | BB_WIP5.1 |  | Can card driver stopped is related to windows system kernel, restart host could enable Can card driver automatically. | Subsystem | Encryption | N/A | Works As Specified / Not A Defect | babu_s | 1 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00623189 | Refresh | Closed | zhang_t | April 19, 2019 10:56:03 AM GMT+08:00 | Refresh, BB: Recon actual image number not equal to expected number. | BB_WIP5.1 |  |  | Subsystem | Reconstruction | N/A | Invalid/Uncontrolled Configuration | richards_b | 0 |  |  | CIRS | NO | Closed |  | Brilliance Big Bore
CTUDT00623187 | Refresh | Closed | liu_j | April 18, 2019 9:28:10 PM GMT+08:00 | IVC First Image Couch Position Display Seems to be Wrong | BB_WIP5 |  |  | Subsystem | Couch Position |  | Forwarded to Different Database | reddy_g | 0 |  |  | Support-Simulator | NO | Closed |  | Brilliance Big Bore
CTUDT00623183 | Refresh | Closed | Dana_Kamyshev | April 18, 2019 6:03:37 PM GMT+08:00 | Refresh - System error during CCT Continuous and Fluoro | WIP15 |  |  | Subsystem | System | n/a | Invalid/Uncontrolled Configuration | yatziv_l | 0 |  |  | DMS | YES | Closed |  | IQon1
CTUDT00623174 | Refresh | Closed | gT_s | April 18, 2019 2:37:03 PM GMT+08:00 | Refresh : No error is displayed when tried to login using ct credentials after the one click installation is stopped at 85%. | WIP 15 | BB_WIP8 WIP18 |  | Subsystem | Login | NA | Fixed | kapadia_v | 0 | 4 | Console Software | Console4 | YES | Closed | WIP17SubCom | Ovation |  | BB_WIP8 WIP18 | June 26, 2019 4:51:24 PM GMT+08:00
CTUDT00623174 | Refresh | Closed | gT_s | April 18, 2019 2:37:03 PM GMT+08:00 | Refresh : No error is displayed when tried to login using ct credentials after the one click installation is stopped at 85%. | WIP 15 | BB_WIP8 WIP18 |  | Subsystem | Login | NA | Fixed | kapadia_v | 0 | 4 | Console Software | Console4 | YES | Closed | WIP17SubCom | Brilliance 64 |  | BB_WIP8 WIP18 | June 26, 2019 4:51:24 PM GMT+08:00
CTUDT00623174 | Refresh | Closed | gT_s | April 18, 2019 2:37:03 PM GMT+08:00 | Refresh : No error is displayed when tried to login using ct credentials after the one click installation is stopped at 85%. | WIP 15 | BB_WIP8 WIP18 |  | Subsystem | Login | NA | Fixed | kapadia_v | 0 | 4 | Console Software | Console4 | YES | Closed | WIP17SubCom | Brilliance Big Bore |  | BB_WIP8 WIP18 | June 26, 2019 4:51:24 PM GMT+08:00
CTUDT00623174 | Refresh | Closed | gT_s | April 18, 2019 2:37:03 PM GMT+08:00 | Refresh : No error is displayed when tried to login using ct credentials after the one click installation is stopped at 85%. | WIP 15 | BB_WIP8 WIP18 |  | Subsystem | Login | NA | Fixed | kapadia_v | 0 | 4 | Console Software | Console4 | YES | Closed | WIP17SubCom | Ingenuity 64 |  | BB_WIP8 WIP18 | June 26, 2019 4:51:24 PM GMT+08:00
CTUDT00623174 | Refresh | Closed | gT_s | April 18, 2019 2:37:03 PM GMT+08:00 | Refresh : No error is displayed when tried to login using ct credentials after the one click installation is stopped at 85%. | WIP 15 | BB_WIP8 WIP18 |  | Subsystem | Login | NA | Fixed | kapadia_v | 0 | 4 | Console Software | Console4 | YES | Closed | WIP17SubCom | IQon1 |  | BB_WIP8 WIP18 | June 26, 2019 4:51:24 PM GMT+08:00
CTUDT00623172 | Refresh | Closed | Noy_Seelig | April 18, 2019 1:42:02 PM GMT+08:00 | Refresh: Constancy test in CT user fails due to Object reference error | BB_WIP5 | WIP18 |  | Subsystem | Constancy Test; Verified1 | NA | Fixed | Kostya_Baklan | 0 | 2 | DMS Calibrations | DMS | NO | Closed | SubCom | Ovation |  | WIP18 | June 18, 2019 1:18:54 PM GMT+08:00
CTUDT00623172 | Refresh | Closed | Noy_Seelig | April 18, 2019 1:42:02 PM GMT+08:00 | Refresh: Constancy test in CT user fails due to Object reference error | BB_WIP5 | WIP18 |  | Subsystem | Constancy Test; Verified1 | NA | Fixed | Kostya_Baklan | 0 | 2 | DMS Calibrations | DMS | NO | Closed | SubCom | Brilliance Big Bore |  | WIP18 | June 18, 2019 1:18:54 PM GMT+08:00
CTUDT00623172 | Refresh | Closed | Noy_Seelig | April 18, 2019 1:42:02 PM GMT+08:00 | Refresh: Constancy test in CT user fails due to Object reference error | BB_WIP5 | WIP18 |  | Subsystem | Constancy Test; Verified1 | NA | Fixed | Kostya_Baklan | 0 | 2 | DMS Calibrations | DMS | NO | Closed | SubCom | Ingenuity 64 |  | WIP18 | June 18, 2019 1:18:54 PM GMT+08:00
CTUDT00623172 | Refresh | Closed | Noy_Seelig | April 18, 2019 1:42:02 PM GMT+08:00 | Refresh: Constancy test in CT user fails due to Object reference error | BB_WIP5 | WIP18 |  | Subsystem | Constancy Test; Verified1 | NA | Fixed | Kostya_Baklan | 0 | 2 | DMS Calibrations | DMS | NO | Closed | SubCom | IQon1 |  | WIP18 | June 18, 2019 1:18:54 PM GMT+08:00
CTUDT00623170 | Refresh | Closed | shekhar_p | April 18, 2019 4:08:20 AM GMT+08:00 | Refresh: Collimation with 32X0.625 helical scans are showing System error | WIP 15 |  |  | Subsystem | System | NaN | Duplicate | yatziv_l | 0 |  |  | DMS | YES | Closed |  | IQon1
CTUDT00623168 | Refresh | Closed | kapadia_v | April 17, 2019 11:22:28 PM GMT+08:00 | Refresh: When enabled Emergency user with privilege to access ECM, cannot login ECM with username Emergency and blank password while other user account requesting authorization. | BB_WIP5.1 | BB_WIP6 | 1. Emergency user's password can be set to a non-empty password and then he can login to ECM with that credential or 2. Unauthorized ECM user can ask any other clinical user who has ecm access and can there by use ECM or 3. Unauthorized ECM user can get access to ECM by requesting local hospital administrator | Subsystem | ECM | NA | Fixed | kapadia_v | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP6_SubCom | Brilliance Big Bore |  | BB_WIP6 | April 26, 2019 5:58:04 PM GMT+08:00
CTUDT00623162 | Refresh | Closed | manjunath_c | April 17, 2019 2:56:30 PM GMT+08:00 | Bottom part of Worklist window in patient page was greyed out | WIP14 | WIP17 | Hover mouse over the controls to reveal UI controls (background is still blocky as shown in the attachment) | Subsystem | Patient Scene | N/A | Fixed | naveen_k | 0 | 4 | Console Software | Console4 | NO | Closed | WIP17SubCom | Ovation |  | WIP17 | May 23, 2019 9:06:10 PM GMT+08:00
CTUDT00623162 | Refresh | Closed | manjunath_c | April 17, 2019 2:56:30 PM GMT+08:00 | Bottom part of Worklist window in patient page was greyed out | WIP14 | WIP17 | Hover mouse over the controls to reveal UI controls (background is still blocky as shown in the attachment) | Subsystem | Patient Scene | N/A | Fixed | naveen_k | 0 | 4 | Console Software | Console4 | NO | Closed | WIP17SubCom | Brilliance 64 |  | WIP17 | May 23, 2019 9:06:10 PM GMT+08:00
CTUDT00623162 | Refresh | Closed | manjunath_c | April 17, 2019 2:56:30 PM GMT+08:00 | Bottom part of Worklist window in patient page was greyed out | WIP14 | WIP17 | Hover mouse over the controls to reveal UI controls (background is still blocky as shown in the attachment) | Subsystem | Patient Scene | N/A | Fixed | naveen_k | 0 | 4 | Console Software | Console4 | NO | Closed | WIP17SubCom | Ingenuity 64 |  | WIP17 | May 23, 2019 9:06:10 PM GMT+08:00
CTUDT00623162 | Refresh | Closed | manjunath_c | April 17, 2019 2:56:30 PM GMT+08:00 | Bottom part of Worklist window in patient page was greyed out | WIP14 | WIP17 | Hover mouse over the controls to reveal UI controls (background is still blocky as shown in the attachment) | Subsystem | Patient Scene | N/A | Fixed | naveen_k | 0 | 4 | Console Software | Console4 | NO | Closed | WIP17SubCom | IQon1 |  | WIP17 | May 23, 2019 9:06:10 PM GMT+08:00
CTUDT00623160 | Refresh | Closed | v_sumanth | April 17, 2019 1:10:14 PM GMT+08:00 | Val: IQon HB selecting 'yes' to a system prompted BugRep collection prompts another BugReport | 4.8 | BB_WIP6 |  | Subsystem | Bugrep | NA | Fixed | rath_s | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP6_SubCom | Brilliance Big Bore |  | BB_WIP6 | April 26, 2019 5:57:46 PM GMT+08:00
CTUDT00623158 | Refresh | Closed | Noy_Seelig | April 17, 2019 12:43:50 AM GMT+08:00 | Refresh: No Acceptance test when Harmonized is selected in iCT SP | WIP14 | WIP17 | In file 'C:\pms\Config\Physics\Applications\AutomaticIQTest\ConstancyAcceptance_Scenario.cfg' Rename 'Acceptance_ScenarioOvation64HPhantom.xml' to 'Acceptance_ScenarioOvation64_HPhantom.xml'. | Subsystem | Calibrations | NA | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP15SegReq | Ovation |  | WIP17 | May 2, 2019 9:52:17 PM GMT+08:00
CTUDT00623156 | Refresh | Closed | 320053769 | April 16, 2019 10:19:24 PM GMT+08:00 | After removing retrieve service in lan config of iSite PACS and trying to copy a study from iSite PACS to local , remove option in transfer tab doesn’t work. | WIP14 | WIP17 | NA | Subsystem | Host | NA | Fixed | srushti_k | 0 | 4 | Console USP_EBW | iCAP | NO | Closed | WIP17SubCom | Ovation |  | WIP17 | June 4, 2019 1:53:33 PM GMT+08:00
CTUDT00623156 | Refresh | Closed | 320053769 | April 16, 2019 10:19:24 PM GMT+08:00 | After removing retrieve service in lan config of iSite PACS and trying to copy a study from iSite PACS to local , remove option in transfer tab doesn’t work. | WIP14 | WIP17 | NA | Subsystem | Host | NA | Fixed | srushti_k | 0 | 4 | Console USP_EBW | iCAP | NO | Closed | WIP17SubCom | Brilliance 64 |  | WIP17 | June 4, 2019 1:53:33 PM GMT+08:00
CTUDT00623156 | Refresh | Closed | 320053769 | April 16, 2019 10:19:24 PM GMT+08:00 | After removing retrieve service in lan config of iSite PACS and trying to copy a study from iSite PACS to local , remove option in transfer tab doesn’t work. | WIP14 | WIP17 | NA | Subsystem | Host | NA | Fixed | srushti_k | 0 | 4 | Console USP_EBW | iCAP | NO | Closed | WIP17SubCom | Brilliance Big Bore |  | WIP17 | June 4, 2019 1:53:33 PM GMT+08:00
CTUDT00623156 | Refresh | Closed | 320053769 | April 16, 2019 10:19:24 PM GMT+08:00 | After removing retrieve service in lan config of iSite PACS and trying to copy a study from iSite PACS to local , remove option in transfer tab doesn’t work. | WIP14 | WIP17 | NA | Subsystem | Host | NA | Fixed | srushti_k | 0 | 4 | Console USP_EBW | iCAP | NO | Closed | WIP17SubCom | Ingenuity 64 |  | WIP17 | June 4, 2019 1:53:33 PM GMT+08:00
CTUDT00623156 | Refresh | Closed | 320053769 | April 16, 2019 10:19:24 PM GMT+08:00 | After removing retrieve service in lan config of iSite PACS and trying to copy a study from iSite PACS to local , remove option in transfer tab doesn’t work. | WIP14 | WIP17 | NA | Subsystem | Host | NA | Fixed | srushti_k | 0 | 4 | Console USP_EBW | iCAP | NO | Closed | WIP17SubCom | IQon1 |  | WIP17 | June 4, 2019 1:53:33 PM GMT+08:00
CTUDT00623154 | Refresh | Closed | 320053769 | April 16, 2019 9:40:14 PM GMT+08:00 | When copying a study from PACS to Local the copy fails. | WIP14 |  |  | Subsystem | Test Design | NA | Works As Specified / Not A Defect | tamari_o | 0 | 2 | Console USP_EBW | Support-System | NO | Closed | WIP19SegReq | Ovation |  |  | September 19, 2019 7:14:56 PM GMT+08:00
CTUDT00623154 | Refresh | Closed | 320053769 | April 16, 2019 9:40:14 PM GMT+08:00 | When copying a study from PACS to Local the copy fails. | WIP14 |  |  | Subsystem | Test Design | NA | Works As Specified / Not A Defect | tamari_o | 0 | 2 | Console USP_EBW | Support-System | NO | Closed | WIP19SegReq | Brilliance 64 |  |  | September 19, 2019 7:14:56 PM GMT+08:00
CTUDT00623154 | Refresh | Closed | 320053769 | April 16, 2019 9:40:14 PM GMT+08:00 | When copying a study from PACS to Local the copy fails. | WIP14 |  |  | Subsystem | Test Design | NA | Works As Specified / Not A Defect | tamari_o | 0 | 2 | Console USP_EBW | Support-System | NO | Closed | WIP19SegReq | Brilliance Big Bore |  |  | September 19, 2019 7:14:56 PM GMT+08:00
CTUDT00623154 | Refresh | Closed | 320053769 | April 16, 2019 9:40:14 PM GMT+08:00 | When copying a study from PACS to Local the copy fails. | WIP14 |  |  | Subsystem | Test Design | NA | Works As Specified / Not A Defect | tamari_o | 0 | 2 | Console USP_EBW | Support-System | NO | Closed | WIP19SegReq | Ingenuity 64 |  |  | September 19, 2019 7:14:56 PM GMT+08:00
CTUDT00623154 | Refresh | Closed | 320053769 | April 16, 2019 9:40:14 PM GMT+08:00 | When copying a study from PACS to Local the copy fails. | WIP14 |  |  | Subsystem | Test Design | NA | Works As Specified / Not A Defect | tamari_o | 0 | 2 | Console USP_EBW | Support-System | NO | Closed | WIP19SegReq | IQon1 |  |  | September 19, 2019 7:14:56 PM GMT+08:00
CTUDT00623152 | Refresh | Closed | beck_d | April 16, 2019 9:20:10 PM GMT+08:00 | Refresh: Tube Current: Error during scan appears and the test stops | BB_WIP5 | N/A |  | Subsystem | Blocker | NaN | Fixed | braunstein_d | 0 | 2 | Design Documentation | Support-Physics | NO | Closed | WIP18SubCom | Brilliance Big Bore |  | N/A | July 4, 2019 8:08:54 PM GMT+08:00
CTUDT00623150 | Refresh | Closed | swamypreethi | April 16, 2019 7:02:22 PM GMT+08:00 | IOCC_Refresh : MWL worklist filter is not reflected in C_FIND_RQ | WIP15 |  | This Works as specified | Subsystem | Worklist | N/A | Works As Specified / Not A Defect | naveen_k | 0 |  |  | Console4 | NO | Closed |  | Ingenuity BigBore
CTUDT00623150 | Refresh | Closed | swamypreethi | April 16, 2019 7:02:22 PM GMT+08:00 | IOCC_Refresh : MWL worklist filter is not reflected in C_FIND_RQ | WIP15 |  | This Works as specified | Subsystem | Worklist | N/A | Works As Specified / Not A Defect | naveen_k | 0 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00623148 | Refresh | Closed | bommiReddy_r | April 16, 2019 6:34:24 PM GMT+08:00 | Refresh: Security: Improper Message is displayed during the Philips service authentication using incorrect entry of IST password | WIP15 | WIP17 |  | Subsystem | Login | NA | Fixed | goswami_s | 0 | 4 | Console Software | Console4 | NO | Closed | WIP17SubCom | Ingenuity 64 |  | WIP17 | April 29, 2019 9:07:49 PM GMT+08:00
CTUDT00623145 | Refresh | Closed | nimma_p | April 16, 2019 6:04:38 PM GMT+08:00 | Refresh->Rules: - User was not able to set the minimum and maximum slice thickness. | WIP14 | WIP17 | enter correct value for the thickness | Subsystem | ECM | Console SRS: - 15520 | Fixed | kumar-san | 0 | 3 | Console Software | Console4 | NO | Closed | WIP17SubCom | IQon1 |  | WIP17 | April 29, 2019 3:37:03 PM GMT+08:00
CTUDT00623144 | Refresh | Closed | Valeria_Kashtanov | April 16, 2019 6:03:06 PM GMT+08:00 | Refresh: Calibrations: UHR Comb Align: Wrong shift measurement is displayed when UHR Comb is shifted | WIP14 | N/A |  | Subsystem | Calibrations | n/a | Fixed | yosfan_s | 0 | 3 | Test Design | DMS | NO | Closed | WIP19SubCom | Ovation |  | N/A | August 11, 2019 9:37:15 PM GMT+08:00
CTUDT00623143 | Refresh | Closed | bhawna_g | April 16, 2019 5:34:47 PM GMT+08:00 | Console sent inconsistent rotation time data to Gantry and CIRS and resulted in ring artifact | 4.8 | BB_WIP6 | 1. Perform scan without AST 2. If scan is performed with AST, do not perform pan/zoom operation on online viewer. | Subsystem | System | N/A | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | YES | Closed | BB_WIP6_SubCom | Brilliance Big Bore | PR8536687 | BB_WIP6 | April 24, 2019 10:06:21 PM GMT+08:00
CTUDT00623138 | Refresh | Closed | sunilkumar_j | April 16, 2019 3:38:56 PM GMT+08:00 | Refresh: The Purpose of visit field was not editable | WIP14 | Host Verification Test Protocol DHF318051 Rev 13 |  | Subsystem | Service Tools | STPS-SSRS-1338 | Fixed | jindal_k | 0 | 3 | Test Design | Service Tools | NO | Closed | WIP16SubCom | Ingenuity 64 |  | Host Verification Test Protocol DHF318051 Rev 13 | May 3, 2019 1:38:51 PM GMT+08:00
CTUDT00623136 | Refresh | Closed | jindal_k | April 16, 2019 2:57:02 PM GMT+08:00 | No warning or error message displayed while upgrading from different software version | WIP14 | Host SSRS DHF303211 Rev 10 DHF316705 Rev10 |  | Subsystem | STPS | STPS-SSRS-899 | Fixed | bandil_j | 0 | 3 | Host SSRS | Console4 | NO | Closed | WIP17SegReq | Ovation |  | Host SSRS DHF303211 Rev 10 DHF316705 Rev10 | May 9, 2019 5:33:39 PM GMT+08:00
CTUDT00623136 | Refresh | Closed | jindal_k | April 16, 2019 2:57:02 PM GMT+08:00 | No warning or error message displayed while upgrading from different software version | WIP14 | Host SSRS DHF303211 Rev 10 DHF316705 Rev10 |  | Subsystem | STPS | STPS-SSRS-899 | Fixed | bandil_j | 0 | 3 | Host SSRS | Console4 | NO | Closed | WIP17SegReq | Brilliance 64 |  | Host SSRS DHF303211 Rev 10 DHF316705 Rev10 | May 9, 2019 5:33:39 PM GMT+08:00
CTUDT00623136 | Refresh | Closed | jindal_k | April 16, 2019 2:57:02 PM GMT+08:00 | No warning or error message displayed while upgrading from different software version | WIP14 | Host SSRS DHF303211 Rev 10 DHF316705 Rev10 |  | Subsystem | STPS | STPS-SSRS-899 | Fixed | bandil_j | 0 | 3 | Host SSRS | Console4 | NO | Closed | WIP17SegReq | Brilliance Big Bore |  | Host SSRS DHF303211 Rev 10 DHF316705 Rev10 | May 9, 2019 5:33:39 PM GMT+08:00
CTUDT00623136 | Refresh | Closed | jindal_k | April 16, 2019 2:57:02 PM GMT+08:00 | No warning or error message displayed while upgrading from different software version | WIP14 | Host SSRS DHF303211 Rev 10 DHF316705 Rev10 |  | Subsystem | STPS | STPS-SSRS-899 | Fixed | bandil_j | 0 | 3 | Host SSRS | Console4 | NO | Closed | WIP17SegReq | Ingenuity 64 |  | Host SSRS DHF303211 Rev 10 DHF316705 Rev10 | May 9, 2019 5:33:39 PM GMT+08:00
CTUDT00623136 | Refresh | Closed | jindal_k | April 16, 2019 2:57:02 PM GMT+08:00 | No warning or error message displayed while upgrading from different software version | WIP14 | Host SSRS DHF303211 Rev 10 DHF316705 Rev10 |  | Subsystem | STPS | STPS-SSRS-899 | Fixed | bandil_j | 0 | 3 | Host SSRS | Console4 | NO | Closed | WIP17SegReq | IQon1 |  | Host SSRS DHF303211 Rev 10 DHF316705 Rev10 | May 9, 2019 5:33:39 PM GMT+08:00
CTUDT00623135 | Refresh | Closed | ravindra_a | April 16, 2019 1:57:17 PM GMT+08:00 | Spectral Parent and Child can have different value for "Single Cycle Reconstruction" and Recon optimisation does not consider "Single Cycle Reconstruction" while optimizing | 4.8 | BB_WIP8 | User should assure both spectral and conventional result are with the same "single-cycle" / "Multi-cycle" parameters. | Subsystem | Image Quality | NA | Fixed | tejaswini_nidamanuri | 0 | 3 | Console Software | Console4 | YES | Closed | BB_WIP7_SubCom | Brilliance Big Bore |  | BB_WIP8 | June 5, 2019 5:10:20 PM GMT+08:00
CTUDT00623133 | Refresh | Closed | shiraguppi_k | April 16, 2019 12:55:22 PM GMT+08:00 | In Spectral CT Viewer Abdomen Windowing Preset values are not matching. | WIP13 | DHF298556 Rev.07 |  | Subsystem | Requirements | CT-NM.SRS.TA.935 | Fixed | tamari_o | 0 | 4 | Requirements Documentation | Support-System | NO | Closed | WIP17SegReq | Ovation |  | DHF298556 Rev.07 | August 8, 2019 4:41:04 PM GMT+08:00
CTUDT00623133 | Refresh | Closed | shiraguppi_k | April 16, 2019 12:55:22 PM GMT+08:00 | In Spectral CT Viewer Abdomen Windowing Preset values are not matching. | WIP13 | DHF298556 Rev.07 |  | Subsystem | Requirements | CT-NM.SRS.TA.935 | Fixed | tamari_o | 0 | 4 | Requirements Documentation | Support-System | NO | Closed | WIP17SegReq | Brilliance 64 |  | DHF298556 Rev.07 | August 8, 2019 4:41:04 PM GMT+08:00
CTUDT00623133 | Refresh | Closed | shiraguppi_k | April 16, 2019 12:55:22 PM GMT+08:00 | In Spectral CT Viewer Abdomen Windowing Preset values are not matching. | WIP13 | DHF298556 Rev.07 |  | Subsystem | Requirements | CT-NM.SRS.TA.935 | Fixed | tamari_o | 0 | 4 | Requirements Documentation | Support-System | NO | Closed | WIP17SegReq | Brilliance Big Bore |  | DHF298556 Rev.07 | August 8, 2019 4:41:04 PM GMT+08:00
CTUDT00623133 | Refresh | Closed | shiraguppi_k | April 16, 2019 12:55:22 PM GMT+08:00 | In Spectral CT Viewer Abdomen Windowing Preset values are not matching. | WIP13 | DHF298556 Rev.07 |  | Subsystem | Requirements | CT-NM.SRS.TA.935 | Fixed | tamari_o | 0 | 4 | Requirements Documentation | Support-System | NO | Closed | WIP17SegReq | Ingenuity 64 |  | DHF298556 Rev.07 | August 8, 2019 4:41:04 PM GMT+08:00
CTUDT00623133 | Refresh | Closed | shiraguppi_k | April 16, 2019 12:55:22 PM GMT+08:00 | In Spectral CT Viewer Abdomen Windowing Preset values are not matching. | WIP13 | DHF298556 Rev.07 |  | Subsystem | Requirements | CT-NM.SRS.TA.935 | Fixed | tamari_o | 0 | 4 | Requirements Documentation | Support-System | NO | Closed | WIP17SegReq | Ingenuity BigBore |  | DHF298556 Rev.07 | August 8, 2019 4:41:04 PM GMT+08:00
CTUDT00623133 | Refresh | Closed | shiraguppi_k | April 16, 2019 12:55:22 PM GMT+08:00 | In Spectral CT Viewer Abdomen Windowing Preset values are not matching. | WIP13 | DHF298556 Rev.07 |  | Subsystem | Requirements | CT-NM.SRS.TA.935 | Fixed | tamari_o | 0 | 4 | Requirements Documentation | Support-System | NO | Closed | WIP17SegReq | IQon1 |  | DHF298556 Rev.07 | August 8, 2019 4:41:04 PM GMT+08:00
CTUDT00623128 | Refresh | Closed | v_sumanth | April 15, 2019 5:51:00 PM GMT+08:00 | The display and operation of message "Maximal allowed number of groups for perfusion is 10" are out of sync. | 4.8 | BB_WIP6 |  | Subsystem | UI Performance | NA | Fixed | v_sumanth | 0 | 2 | Console Software | Console4 | YES | Closed | BB_WIP6_SubCom | Brilliance Big Bore |  | BB_WIP6 | April 23, 2019 2:32:16 PM GMT+08:00
CTUDT00623123 | Refresh | Closed | bhawna_g | April 15, 2019 3:05:51 PM GMT+08:00 | "Incorrect pixel spacing among images. Reselect data." message appears when data is loaded to HeartBeat Calcium Score | 4.8 | BB_WIP6 | Reducing scan length and changing to different phase-70% this will allow the user to load the images. | Subsystem | Cardiac | N/A | Fixed | v_su | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP6_SubCom | Brilliance Big Bore | 9448055 | BB_WIP6 | April 26, 2019 3:10:23 PM GMT+08:00
CTUDT00623120 | Refresh | Closed | g_di | April 15, 2019 1:51:26 PM GMT+08:00 | “Average planned mA” is not displayed in the acquisition viewer for Biopsy exam card. | WIP14 | Host SSRS DHF303211 Rev 10 DHF316705 Rev10 |  | Subsystem | Requirements | Console-SRS.16127 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | YES | Closed | WIP16SegReq | Ovation |  | Host SSRS DHF303211 Rev 10 DHF316705 Rev10 | May 9, 2019 4:45:28 PM GMT+08:00
CTUDT00623120 | Refresh | Closed | g_di | April 15, 2019 1:51:26 PM GMT+08:00 | “Average planned mA” is not displayed in the acquisition viewer for Biopsy exam card. | WIP14 | Host SSRS DHF303211 Rev 10 DHF316705 Rev10 |  | Subsystem | Requirements | Console-SRS.16127 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | YES | Closed | WIP16SegReq | Brilliance 64 |  | Host SSRS DHF303211 Rev 10 DHF316705 Rev10 | May 9, 2019 4:45:28 PM GMT+08:00
CTUDT00623120 | Refresh | Closed | g_di | April 15, 2019 1:51:26 PM GMT+08:00 | “Average planned mA” is not displayed in the acquisition viewer for Biopsy exam card. | WIP14 | Host SSRS DHF303211 Rev 10 DHF316705 Rev10 |  | Subsystem | Requirements | Console-SRS.16127 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | YES | Closed | WIP16SegReq | Brilliance Big Bore |  | Host SSRS DHF303211 Rev 10 DHF316705 Rev10 | May 9, 2019 4:45:28 PM GMT+08:00
CTUDT00623120 | Refresh | Closed | g_di | April 15, 2019 1:51:26 PM GMT+08:00 | “Average planned mA” is not displayed in the acquisition viewer for Biopsy exam card. | WIP14 | Host SSRS DHF303211 Rev 10 DHF316705 Rev10 |  | Subsystem | Requirements | Console-SRS.16127 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | YES | Closed | WIP16SegReq | Ingenuity 64 |  | Host SSRS DHF303211 Rev 10 DHF316705 Rev10 | May 9, 2019 4:45:28 PM GMT+08:00
CTUDT00623120 | Refresh | Closed | g_di | April 15, 2019 1:51:26 PM GMT+08:00 | “Average planned mA” is not displayed in the acquisition viewer for Biopsy exam card. | WIP14 | Host SSRS DHF303211 Rev 10 DHF316705 Rev10 |  | Subsystem | Requirements | Console-SRS.16127 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | YES | Closed | WIP16SegReq | IQon1 |  | Host SSRS DHF303211 Rev 10 DHF316705 Rev10 | May 9, 2019 4:45:28 PM GMT+08:00
CTUDT00623118 | Refresh | Closed | panchami_pradeep | April 15, 2019 1:37:57 PM GMT+08:00 | Refresh: PSC : X-ray trace viewer tool is not displaying the progress of operation. | WIP14 | WIP17 |  | Subsystem | Service Tools | NA | Fixed | chakraborty_d | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP17SubCom | Ovation |  | WIP17 | May 7, 2019 2:33:25 PM GMT+08:00
CTUDT00623118 | Refresh | Closed | panchami_pradeep | April 15, 2019 1:37:57 PM GMT+08:00 | Refresh: PSC : X-ray trace viewer tool is not displaying the progress of operation. | WIP14 | WIP17 |  | Subsystem | Service Tools | NA | Fixed | chakraborty_d | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP17SubCom | IQon1 |  | WIP17 | May 7, 2019 2:33:25 PM GMT+08:00
CTUDT00623117 | Refresh | Closed | kapadia_v | April 15, 2019 1:09:06 PM GMT+08:00 | User cannot change the priority of “Body Perfusion Helical” offline-recon | BB_WIP3 | BB_WIP6 | NA | Subsystem | Recon Priority | N/A | Fixed | sagar_v | 0 | 2 | Console Generic | Console4 | NO | Closed | BB_WIP6_SubCom | Brilliance Big Bore |  | BB_WIP6 | April 24, 2019 10:10:13 PM GMT+08:00
CTUDT00623114 | Refresh | Closed | kapadia_v | April 15, 2019 12:41:40 PM GMT+08:00 | Refresh : The clinical user is not alerted with the message when the new version of software is available for installation.on PRS portal for download. | BB_WIP5 | BB_WIP6 |  | Subsystem | Patch Management | NA | Fixed | kapadia_v | 0 | 4 | Console Software | Support-Simulator | NO | Closed | WIP16SubCom | Brilliance Big Bore |  | BB_WIP6 | April 26, 2019 6:10:55 PM GMT+08:00
CTUDT00623111 | Refresh | Closed | 320053769 | April 14, 2019 10:12:34 PM GMT+08:00 | Disk space warning massage doesn't pop until lowering the space in local to 16gb | WIP14 |  | NA | Subsystem | Disk Space | Console-SRS.2526 | Works As Specified / Not A Defect | veeraswamy_k | 0 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00623104 | Refresh | Closed | Sharmila_Vasireddy | April 13, 2019 2:44:56 PM GMT+08:00 | BB refresh: Multiple Dose SR reports are generated instead of one SR report | BB_WIP5 | BB_WIP7 | No, There is no workaround | Subsystem | Host | NA | Fixed | jayakumar_p | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP7_SubCom | Ovation |  | BB_WIP7 | May 31, 2019 2:44:15 PM GMT+08:00
CTUDT00623104 | Refresh | Closed | Sharmila_Vasireddy | April 13, 2019 2:44:56 PM GMT+08:00 | BB refresh: Multiple Dose SR reports are generated instead of one SR report | BB_WIP5 | BB_WIP7 | No, There is no workaround | Subsystem | Host | NA | Fixed | jayakumar_p | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP7_SubCom | Brilliance 64 |  | BB_WIP7 | May 31, 2019 2:44:15 PM GMT+08:00
CTUDT00623104 | Refresh | Closed | Sharmila_Vasireddy | April 13, 2019 2:44:56 PM GMT+08:00 | BB refresh: Multiple Dose SR reports are generated instead of one SR report | BB_WIP5 | BB_WIP7 | No, There is no workaround | Subsystem | Host | NA | Fixed | jayakumar_p | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP7_SubCom | Brilliance Big Bore |  | BB_WIP7 | May 31, 2019 2:44:15 PM GMT+08:00
CTUDT00623104 | Refresh | Closed | Sharmila_Vasireddy | April 13, 2019 2:44:56 PM GMT+08:00 | BB refresh: Multiple Dose SR reports are generated instead of one SR report | BB_WIP5 | BB_WIP7 | No, There is no workaround | Subsystem | Host | NA | Fixed | jayakumar_p | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP7_SubCom | Ingenuity 64 |  | BB_WIP7 | May 31, 2019 2:44:15 PM GMT+08:00
CTUDT00623104 | Refresh | Closed | Sharmila_Vasireddy | April 13, 2019 2:44:56 PM GMT+08:00 | BB refresh: Multiple Dose SR reports are generated instead of one SR report | BB_WIP5 | BB_WIP7 | No, There is no workaround | Subsystem | Host | NA | Fixed | jayakumar_p | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP7_SubCom | IQon1 |  | BB_WIP7 | May 31, 2019 2:44:15 PM GMT+08:00
CTUDT00623101 | Refresh | Closed | bommiReddy_r | April 12, 2019 8:32:19 PM GMT+08:00 | Refresh STPS: "Access to CIRS is not granted" warning message is displayed when tried to access MDU viewer Tool and Error message is displayed when the MDU file was loaded | WIP14 |  |  | Subsystem | MDU |  | Duplicate | yatziv_l | 0 |  |  | DMS | NO | Closed |  | Ovation
CTUDT00623099 | Refresh | Closed | manjunath_c | April 12, 2019 8:01:54 PM GMT+08:00 | End Exam continue working not working | WIP14 |  |  | Subsystem | Host | NA | Works As Specified / Not A Defect | babu_s | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00623091 | Refresh | Closed | Sharmila_Vasireddy | April 12, 2019 5:16:12 PM GMT+08:00 | Popup message was not displaying as per the message #450 while resetting the Philips service user’s password | BB_WIP5 | WIP17 BB_WIP7 | No. | Subsystem | Password Management | NA | Fixed | ravindra_a | 0 | 3 | Console Software | Console4 | NO | Closed | WIP16SubCom | Ovation |  | WIP17 BB_WIP7 | May 23, 2019 5:49:01 PM GMT+08:00
CTUDT00623091 | Refresh | Closed | Sharmila_Vasireddy | April 12, 2019 5:16:12 PM GMT+08:00 | Popup message was not displaying as per the message #450 while resetting the Philips service user’s password | BB_WIP5 | WIP17 BB_WIP7 | No. | Subsystem | Password Management | NA | Fixed | ravindra_a | 0 | 3 | Console Software | Console4 | NO | Closed | WIP16SubCom | Brilliance 64 |  | WIP17 BB_WIP7 | May 23, 2019 5:49:01 PM GMT+08:00
CTUDT00623091 | Refresh | Closed | Sharmila_Vasireddy | April 12, 2019 5:16:12 PM GMT+08:00 | Popup message was not displaying as per the message #450 while resetting the Philips service user’s password | BB_WIP5 | WIP17 BB_WIP7 | No. | Subsystem | Password Management | NA | Fixed | ravindra_a | 0 | 3 | Console Software | Console4 | NO | Closed | WIP16SubCom | Brilliance Big Bore |  | WIP17 BB_WIP7 | May 23, 2019 5:49:01 PM GMT+08:00
CTUDT00623091 | Refresh | Closed | Sharmila_Vasireddy | April 12, 2019 5:16:12 PM GMT+08:00 | Popup message was not displaying as per the message #450 while resetting the Philips service user’s password | BB_WIP5 | WIP17 BB_WIP7 | No. | Subsystem | Password Management | NA | Fixed | ravindra_a | 0 | 3 | Console Software | Console4 | NO | Closed | WIP16SubCom | Ingenuity 64 |  | WIP17 BB_WIP7 | May 23, 2019 5:49:01 PM GMT+08:00
CTUDT00623091 | Refresh | Closed | Sharmila_Vasireddy | April 12, 2019 5:16:12 PM GMT+08:00 | Popup message was not displaying as per the message #450 while resetting the Philips service user’s password | BB_WIP5 | WIP17 BB_WIP7 | No. | Subsystem | Password Management | NA | Fixed | ravindra_a | 0 | 3 | Console Software | Console4 | NO | Closed | WIP16SubCom | IQon1 |  | WIP17 BB_WIP7 | May 23, 2019 5:49:01 PM GMT+08:00
CTUDT00623081 | Refresh | Closed | gT_s | April 12, 2019 10:08:43 AM GMT+08:00 | Refresh : The remote connection request gets disconnected before the disclaimer in the host machine reaches zero second. | WIP 14 | BB_WIP11 WIP21 |  | Subsystem | Remote Configuration | N\A | Fixed | venkatachalam_p | 0 | 4 | Requirements Documentation | Service Tools | NO | Closed | WIP21SubCom | Ovation |  | BB_WIP11 WIP21 | October 11, 2019 8:45:01 PM GMT+08:00
CTUDT00623081 | Refresh | Closed | gT_s | April 12, 2019 10:08:43 AM GMT+08:00 | Refresh : The remote connection request gets disconnected before the disclaimer in the host machine reaches zero second. | WIP 14 | BB_WIP11 WIP21 |  | Subsystem | Remote Configuration | N\A | Fixed | venkatachalam_p | 0 | 4 | Requirements Documentation | Service Tools | NO | Closed | WIP21SubCom | Brilliance 64 |  | BB_WIP11 WIP21 | October 11, 2019 8:45:01 PM GMT+08:00
CTUDT00623081 | Refresh | Closed | gT_s | April 12, 2019 10:08:43 AM GMT+08:00 | Refresh : The remote connection request gets disconnected before the disclaimer in the host machine reaches zero second. | WIP 14 | BB_WIP11 WIP21 |  | Subsystem | Remote Configuration | N\A | Fixed | venkatachalam_p | 0 | 4 | Requirements Documentation | Service Tools | NO | Closed | WIP21SubCom | Brilliance Big Bore |  | BB_WIP11 WIP21 | October 11, 2019 8:45:01 PM GMT+08:00
CTUDT00623081 | Refresh | Closed | gT_s | April 12, 2019 10:08:43 AM GMT+08:00 | Refresh : The remote connection request gets disconnected before the disclaimer in the host machine reaches zero second. | WIP 14 | BB_WIP11 WIP21 |  | Subsystem | Remote Configuration | N\A | Fixed | venkatachalam_p | 0 | 4 | Requirements Documentation | Service Tools | NO | Closed | WIP21SubCom | Ingenuity 64 |  | BB_WIP11 WIP21 | October 11, 2019 8:45:01 PM GMT+08:00
CTUDT00623081 | Refresh | Closed | gT_s | April 12, 2019 10:08:43 AM GMT+08:00 | Refresh : The remote connection request gets disconnected before the disclaimer in the host machine reaches zero second. | WIP 14 | BB_WIP11 WIP21 |  | Subsystem | Remote Configuration | N\A | Fixed | venkatachalam_p | 0 | 4 | Requirements Documentation | Service Tools | NO | Closed | WIP21SubCom | IQon1 |  | BB_WIP11 WIP21 | October 11, 2019 8:45:01 PM GMT+08:00
CTUDT00623079 | Refresh | Closed | Valeria_Kashtanov | April 12, 2019 2:05:24 AM GMT+08:00 | Refresh: Calibrations: Performance Test: during full Performance, Application gets stuck on IMR stage, and QA stops responding. | WIP14 |  | N/A | Subsystem | Calibrations | N/A | Invalid/Uncontrolled Configuration | elviral | 0 |  |  | DMS | NO | Closed |  | Ovation
CTUDT00623067 | Refresh | Closed | richards_b | April 11, 2019 7:27:54 PM GMT+08:00 | Reconstruction stalling - not all images created | 4.7.5 | WIP15 | Offline recon | Subsystem |  | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_ImgNodeSvr | CIRS | NO | Closed | WIP15SubCom | IQon1 | 8934059 | WIP15 | April 26, 2019 4:53:39 AM GMT+08:00
CTUDT00623066 | Refresh | Closed | richards_b | April 11, 2019 7:23:59 PM GMT+08:00 | System appeared stuck after a cancel | 4.1.7 | WIP15 | In the situation that the study did get past intialization, the operator would have the ability to reconstruct this data without any additional dose to the patient being required. | Subsystem | CIRS | N/A | Fixed | rakic_k | 0 | 2 | CIRS_SW_Supervisor | CIRS | YES | Closed | WIP15SubCom | Ovation | 8687803 | WIP15 | April 30, 2019 10:05:12 PM GMT+08:00
CTUDT00623066 | Refresh | Closed | richards_b | April 11, 2019 7:23:59 PM GMT+08:00 | System appeared stuck after a cancel | 4.1.7 | WIP15 | In the situation that the study did get past intialization, the operator would have the ability to reconstruct this data without any additional dose to the patient being required. | Subsystem | CIRS | N/A | Fixed | rakic_k | 0 | 2 | CIRS_SW_Supervisor | CIRS | YES | Closed | WIP15SubCom | Brilliance 64 | 8687803 | WIP15 | April 30, 2019 10:05:12 PM GMT+08:00
CTUDT00623066 | Refresh | Closed | richards_b | April 11, 2019 7:23:59 PM GMT+08:00 | System appeared stuck after a cancel | 4.1.7 | WIP15 | In the situation that the study did get past intialization, the operator would have the ability to reconstruct this data without any additional dose to the patient being required. | Subsystem | CIRS | N/A | Fixed | rakic_k | 0 | 2 | CIRS_SW_Supervisor | CIRS | YES | Closed | WIP15SubCom | Brilliance Big Bore | 8687803 | WIP15 | April 30, 2019 10:05:12 PM GMT+08:00
CTUDT00623066 | Refresh | Closed | richards_b | April 11, 2019 7:23:59 PM GMT+08:00 | System appeared stuck after a cancel | 4.1.7 | WIP15 | In the situation that the study did get past intialization, the operator would have the ability to reconstruct this data without any additional dose to the patient being required. | Subsystem | CIRS | N/A | Fixed | rakic_k | 0 | 2 | CIRS_SW_Supervisor | CIRS | YES | Closed | WIP15SubCom | Ingenuity 64 | 8687803 | WIP15 | April 30, 2019 10:05:12 PM GMT+08:00
CTUDT00623065 | Refresh | Closed | richards_b | April 11, 2019 7:20:56 PM GMT+08:00 | SNUH: Dual Surview images did not appear | 4.7.5 | WIP15 | The operator would have the ability to reconstruct this data without any additional dose to the patient being required. | Subsystem | CIRS | N/A | Fixed | rakic_k | 0 | 2 | CIRS_SW_Supervisor | CIRS | YES | Closed | WIP15SubCom | Ovation | PR8735021 | WIP15 | April 30, 2019 10:05:38 PM GMT+08:00
CTUDT00623065 | Refresh | Closed | richards_b | April 11, 2019 7:20:56 PM GMT+08:00 | SNUH: Dual Surview images did not appear | 4.7.5 | WIP15 | The operator would have the ability to reconstruct this data without any additional dose to the patient being required. | Subsystem | CIRS | N/A | Fixed | rakic_k | 0 | 2 | CIRS_SW_Supervisor | CIRS | YES | Closed | WIP15SubCom | Brilliance 64 | PR8735021 | WIP15 | April 30, 2019 10:05:38 PM GMT+08:00
CTUDT00623065 | Refresh | Closed | richards_b | April 11, 2019 7:20:56 PM GMT+08:00 | SNUH: Dual Surview images did not appear | 4.7.5 | WIP15 | The operator would have the ability to reconstruct this data without any additional dose to the patient being required. | Subsystem | CIRS | N/A | Fixed | rakic_k | 0 | 2 | CIRS_SW_Supervisor | CIRS | YES | Closed | WIP15SubCom | Brilliance Big Bore | PR8735021 | WIP15 | April 30, 2019 10:05:38 PM GMT+08:00
CTUDT00623065 | Refresh | Closed | richards_b | April 11, 2019 7:20:56 PM GMT+08:00 | SNUH: Dual Surview images did not appear | 4.7.5 | WIP15 | The operator would have the ability to reconstruct this data without any additional dose to the patient being required. | Subsystem | CIRS | N/A | Fixed | rakic_k | 0 | 2 | CIRS_SW_Supervisor | CIRS | YES | Closed | WIP15SubCom | Ingenuity 64 | PR8735021 | WIP15 | April 30, 2019 10:05:38 PM GMT+08:00
CTUDT00623065 | Refresh | Closed | richards_b | April 11, 2019 7:20:56 PM GMT+08:00 | SNUH: Dual Surview images did not appear | 4.7.5 | WIP15 | The operator would have the ability to reconstruct this data without any additional dose to the patient being required. | Subsystem | CIRS | N/A | Fixed | rakic_k | 0 | 2 | CIRS_SW_Supervisor | CIRS | YES | Closed | WIP15SubCom | IQon1 | PR8735021 | WIP15 | April 30, 2019 10:05:38 PM GMT+08:00
CTUDT00623063 | Refresh | Closed | divya_n | April 11, 2019 6:03:06 PM GMT+08:00 | Refresh - Status of study shows as In Progress count juggles between Waiting to Failed when there is low disk space in remote system. | WIP14 |  | Since the studies are available in local, the user can re-try the transfer later , on ensuring that the remote system is in good state. | Subsystem | Console Infrastructure | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | Ovation
CTUDT00623063 | Refresh | Closed | divya_n | April 11, 2019 6:03:06 PM GMT+08:00 | Refresh - Status of study shows as In Progress count juggles between Waiting to Failed when there is low disk space in remote system. | WIP14 |  | Since the studies are available in local, the user can re-try the transfer later , on ensuring that the remote system is in good state. | Subsystem | Console Infrastructure | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00623063 | Refresh | Closed | divya_n | April 11, 2019 6:03:06 PM GMT+08:00 | Refresh - Status of study shows as In Progress count juggles between Waiting to Failed when there is low disk space in remote system. | WIP14 |  | Since the studies are available in local, the user can re-try the transfer later , on ensuring that the remote system is in good state. | Subsystem | Console Infrastructure | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00623063 | Refresh | Closed | divya_n | April 11, 2019 6:03:06 PM GMT+08:00 | Refresh - Status of study shows as In Progress count juggles between Waiting to Failed when there is low disk space in remote system. | WIP14 |  | Since the studies are available in local, the user can re-try the transfer later , on ensuring that the remote system is in good state. | Subsystem | Console Infrastructure | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00623063 | Refresh | Closed | divya_n | April 11, 2019 6:03:06 PM GMT+08:00 | Refresh - Status of study shows as In Progress count juggles between Waiting to Failed when there is low disk space in remote system. | WIP14 |  | Since the studies are available in local, the user can re-try the transfer later , on ensuring that the remote system is in good state. | Subsystem | Console Infrastructure | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | IQon1
CTUDT00623060 | Refresh | Closed | divya_n | April 11, 2019 5:56:51 PM GMT+08:00 | Refresh - Status of the study stucks at Paused for Priority and needs manual intervention to transfer the study even when there are no priority devices transfer in Queue Manager. | WIP14 |  | User can press Pause\Resume button to continue the transfer to PACS | Subsystem | USP | NA | No Plan to Fix | s_sa | 0 | 3 |  | iCAP | NO | Closed |  | Ovation
CTUDT00623060 | Refresh | Closed | divya_n | April 11, 2019 5:56:51 PM GMT+08:00 | Refresh - Status of the study stucks at Paused for Priority and needs manual intervention to transfer the study even when there are no priority devices transfer in Queue Manager. | WIP14 |  | User can press Pause\Resume button to continue the transfer to PACS | Subsystem | USP | NA | No Plan to Fix | s_sa | 0 | 3 |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00623060 | Refresh | Closed | divya_n | April 11, 2019 5:56:51 PM GMT+08:00 | Refresh - Status of the study stucks at Paused for Priority and needs manual intervention to transfer the study even when there are no priority devices transfer in Queue Manager. | WIP14 |  | User can press Pause\Resume button to continue the transfer to PACS | Subsystem | USP | NA | No Plan to Fix | s_sa | 0 | 3 |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00623060 | Refresh | Closed | divya_n | April 11, 2019 5:56:51 PM GMT+08:00 | Refresh - Status of the study stucks at Paused for Priority and needs manual intervention to transfer the study even when there are no priority devices transfer in Queue Manager. | WIP14 |  | User can press Pause\Resume button to continue the transfer to PACS | Subsystem | USP | NA | No Plan to Fix | s_sa | 0 | 3 |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00623060 | Refresh | Closed | divya_n | April 11, 2019 5:56:51 PM GMT+08:00 | Refresh - Status of the study stucks at Paused for Priority and needs manual intervention to transfer the study even when there are no priority devices transfer in Queue Manager. | WIP14 |  | User can press Pause\Resume button to continue the transfer to PACS | Subsystem | USP | NA | No Plan to Fix | s_sa | 0 | 3 |  | iCAP | NO | Closed |  | IQon1
CTUDT00623057 | Refresh | Closed | sunilkumar_j | April 11, 2019 5:44:16 PM GMT+08:00 | ICT & IQON [Refresh WIP 14] :ECM: Result's Center X and center Y values were not 0 , not as defined in Preferenses | WIP14 |  | There is workaround. User can save the Exam card and reopen the Exam card in Exam card manager. The value of Center X and Center Y is set to 0. | Subsystem | Host | NA | No Plan to Fix | jayakumar_p | 0 | 4 |  | Console4 | NO | Closed | SubNC | Ovation
CTUDT00623057 | Refresh | Closed | sunilkumar_j | April 11, 2019 5:44:16 PM GMT+08:00 | ICT & IQON [Refresh WIP 14] :ECM: Result's Center X and center Y values were not 0 , not as defined in Preferenses | WIP14 |  | There is workaround. User can save the Exam card and reopen the Exam card in Exam card manager. The value of Center X and Center Y is set to 0. | Subsystem | Host | NA | No Plan to Fix | jayakumar_p | 0 | 4 |  | Console4 | NO | Closed | SubNC | Brilliance 64
CTUDT00623057 | Refresh | Closed | sunilkumar_j | April 11, 2019 5:44:16 PM GMT+08:00 | ICT & IQON [Refresh WIP 14] :ECM: Result's Center X and center Y values were not 0 , not as defined in Preferenses | WIP14 |  | There is workaround. User can save the Exam card and reopen the Exam card in Exam card manager. The value of Center X and Center Y is set to 0. | Subsystem | Host | NA | No Plan to Fix | jayakumar_p | 0 | 4 |  | Console4 | NO | Closed | SubNC | Brilliance Big Bore
CTUDT00623057 | Refresh | Closed | sunilkumar_j | April 11, 2019 5:44:16 PM GMT+08:00 | ICT & IQON [Refresh WIP 14] :ECM: Result's Center X and center Y values were not 0 , not as defined in Preferenses | WIP14 |  | There is workaround. User can save the Exam card and reopen the Exam card in Exam card manager. The value of Center X and Center Y is set to 0. | Subsystem | Host | NA | No Plan to Fix | jayakumar_p | 0 | 4 |  | Console4 | NO | Closed | SubNC | Ingenuity 64
CTUDT00623057 | Refresh | Closed | sunilkumar_j | April 11, 2019 5:44:16 PM GMT+08:00 | ICT & IQON [Refresh WIP 14] :ECM: Result's Center X and center Y values were not 0 , not as defined in Preferenses | WIP14 |  | There is workaround. User can save the Exam card and reopen the Exam card in Exam card manager. The value of Center X and Center Y is set to 0. | Subsystem | Host | NA | No Plan to Fix | jayakumar_p | 0 | 4 |  | Console4 | NO | Closed | SubNC | IQon1
CTUDT00623056 | Refresh | Closed | divya_n | April 11, 2019 5:34:03 PM GMT+08:00 | Refresh - Status of the transaction in Queue Manager displays as In Progress on performing Auto-store with the combination of multiple High Priority devices + High priority series and the Priority device goes down. | WIP14 |  | Manual transfer after auto transfer was paused | Subsystem | Patient Directory | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | Ovation
CTUDT00623056 | Refresh | Closed | divya_n | April 11, 2019 5:34:03 PM GMT+08:00 | Refresh - Status of the transaction in Queue Manager displays as In Progress on performing Auto-store with the combination of multiple High Priority devices + High priority series and the Priority device goes down. | WIP14 |  | Manual transfer after auto transfer was paused | Subsystem | Patient Directory | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00623056 | Refresh | Closed | divya_n | April 11, 2019 5:34:03 PM GMT+08:00 | Refresh - Status of the transaction in Queue Manager displays as In Progress on performing Auto-store with the combination of multiple High Priority devices + High priority series and the Priority device goes down. | WIP14 |  | Manual transfer after auto transfer was paused | Subsystem | Patient Directory | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00623056 | Refresh | Closed | divya_n | April 11, 2019 5:34:03 PM GMT+08:00 | Refresh - Status of the transaction in Queue Manager displays as In Progress on performing Auto-store with the combination of multiple High Priority devices + High priority series and the Priority device goes down. | WIP14 |  | Manual transfer after auto transfer was paused | Subsystem | Patient Directory | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00623056 | Refresh | Closed | divya_n | April 11, 2019 5:34:03 PM GMT+08:00 | Refresh - Status of the transaction in Queue Manager displays as In Progress on performing Auto-store with the combination of multiple High Priority devices + High priority series and the Priority device goes down. | WIP14 |  | Manual transfer after auto transfer was paused | Subsystem | Patient Directory | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | IQon1
CTUDT00623054 | Refresh | Closed | Sharmila_Vasireddy | April 11, 2019 5:07:48 PM GMT+08:00 | BB refresh: System Error after completion of axial scan, when plan the clinical scan where the scan can't be extended further | BB_WIP5 |  |  | Subsystem | Plan | NA | Works As Specified / Not A Defect | babu_s | 1 |  |  | Console4 | YES | Closed |  | Brilliance Big Bore
CTUDT00623051 | Refresh | Closed | ding_s | April 11, 2019 2:51:29 PM GMT+08:00 | The minimum and maxmum mA allowed to set do not meet the requirement. | BB_WIP5 | DHF298555 Rev 06 - BB SRS DHF298554 Rev 06 - Ing/BR64 SRS | None | System | Host | Console-SRS.17016, Console-SRS.10060 | Fixed | mistry_h | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP17SubCom & WIP7SubCom | Brilliance Big Bore |  | DHF298555 Rev 06 - BB SRS DHF298554 Rev 06 - Ing/BR64 SRS | July 26, 2019 8:17:04 PM GMT+08:00
CTUDT00623051 | Refresh | Closed | ding_s | April 11, 2019 2:51:29 PM GMT+08:00 | The minimum and maxmum mA allowed to set do not meet the requirement. | BB_WIP5 | DHF298555 Rev 06 - BB SRS DHF298554 Rev 06 - Ing/BR64 SRS | None | System | Host | Console-SRS.17016, Console-SRS.10060 | Fixed | mistry_h | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP17SubCom & WIP7SubCom | Ingenuity 64 |  | DHF298555 Rev 06 - BB SRS DHF298554 Rev 06 - Ing/BR64 SRS | July 26, 2019 8:17:04 PM GMT+08:00
CTUDT00623048 | Refresh | Closed | swamypreethi | April 11, 2019 2:14:11 PM GMT+08:00 | IOCC_Refresh: Insufficient coverage of de-identification attributes as per Basic Application Level Confidentiality Profile defined in DICOM part 15(DICOM standard 2019a). | WIP12 |  |  | Subsystem | DICOM | N/A | No Plan to Fix | bandil_j | 0 | 3 |  | Console4 | NO | Closed | SubNC | Brilliance Big Bore
CTUDT00623048 | Refresh | Closed | swamypreethi | April 11, 2019 2:14:11 PM GMT+08:00 | IOCC_Refresh: Insufficient coverage of de-identification attributes as per Basic Application Level Confidentiality Profile defined in DICOM part 15(DICOM standard 2019a). | WIP12 |  |  | Subsystem | DICOM | N/A | No Plan to Fix | bandil_j | 0 | 3 |  | Console4 | NO | Closed | SubNC | Ingenuity 64
CTUDT00623048 | Refresh | Closed | swamypreethi | April 11, 2019 2:14:11 PM GMT+08:00 | IOCC_Refresh: Insufficient coverage of de-identification attributes as per Basic Application Level Confidentiality Profile defined in DICOM part 15(DICOM standard 2019a). | WIP12 |  |  | Subsystem | DICOM | N/A | No Plan to Fix | bandil_j | 0 | 3 |  | Console4 | NO | Closed | SubNC | IQon1
CTUDT00623033 | Refresh | Closed | v_sumanth | April 11, 2019 12:18:33 PM GMT+08:00 | Wrong user message was displayed | 4.8 | BB_WIP6 | No Workaround | Subsystem | Voice Manager | None | Fixed | Prakash_Chand_Yadav | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP6_SubCom | Brilliance Big Bore |  | BB_WIP6 | April 23, 2019 2:02:33 PM GMT+08:00
CTUDT00623030 | Refresh | Closed | cohen_y | April 11, 2019 12:01:13 AM GMT+08:00 | Refresh: Hcor Basic Calibration: Harmonized: Ratios for Infant Body cases are 1 even though Infant body cases are performed | WIP14 | WIP17 | N/A | Subsystem | Calibrations | N/A | Fixed | Kostya_Baklan | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP17SegReq | Ovation |  | WIP17 | May 26, 2019 5:41:22 PM GMT+08:00
CTUDT00623030 | Refresh | Closed | cohen_y | April 11, 2019 12:01:13 AM GMT+08:00 | Refresh: Hcor Basic Calibration: Harmonized: Ratios for Infant Body cases are 1 even though Infant body cases are performed | WIP14 | WIP17 | N/A | Subsystem | Calibrations | N/A | Fixed | Kostya_Baklan | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP17SegReq | Brilliance Big Bore |  | WIP17 | May 26, 2019 5:41:22 PM GMT+08:00
CTUDT00623030 | Refresh | Closed | cohen_y | April 11, 2019 12:01:13 AM GMT+08:00 | Refresh: Hcor Basic Calibration: Harmonized: Ratios for Infant Body cases are 1 even though Infant body cases are performed | WIP14 | WIP17 | N/A | Subsystem | Calibrations | N/A | Fixed | Kostya_Baklan | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP17SegReq | Ingenuity 64 |  | WIP17 | May 26, 2019 5:41:22 PM GMT+08:00
CTUDT00623030 | Refresh | Closed | cohen_y | April 11, 2019 12:01:13 AM GMT+08:00 | Refresh: Hcor Basic Calibration: Harmonized: Ratios for Infant Body cases are 1 even though Infant body cases are performed | WIP14 | WIP17 | N/A | Subsystem | Calibrations | N/A | Fixed | Kostya_Baklan | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP17SegReq | IQon1 |  | WIP17 | May 26, 2019 5:41:22 PM GMT+08:00
CTUDT00623028 | Refresh | Closed | manju_kolkar | April 10, 2019 8:58:43 PM GMT+08:00 | Big Bore Refresh - Result geometry, Number of Cycles and Scan Length were not set as per expectations. | BB_WIP5 |  |  | Subsystem | ECM | N/A | Duplicate | bhawna_g | 0 |  |  | Console4 | YES | Closed | BB_WIP6_SubCom | Brilliance Big Bore
CTUDT00623027 | Refresh | Closed | sowmya_k | April 10, 2019 8:10:20 PM GMT+08:00 | Refresh: Bugrep Error window comes Up on Ending the study, when tried to save an EC in Plan. | BB_WIP5.1 | BB_WIP6 |  | Subsystem | Bugrep | NA | Fixed | wilsonraja | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP6_SubCom | Brilliance Big Bore |  | BB_WIP6 | April 26, 2019 2:54:47 PM GMT+08:00
CTUDT00623024 | Refresh | Closed | Revital_Cohen | April 10, 2019 7:54:21 PM GMT+08:00 | Refresh - Emergency user login button should be the only enabled button once login to Emergency User. | WIP14 | WIP17 | NA | Subsystem | Login | N/A | Fixed | kapadia_v | 0 | 3 | Console Software | Service Tools | NO | Closed | WIP16SubCom | IQon1 |  | WIP17 | May 8, 2019 5:42:23 PM GMT+08:00
CTUDT00623021 | Refresh | Closed | Revital_Cohen | April 10, 2019 7:39:03 PM GMT+08:00 | Refresh- Emergency users shouldn’t be locked and when locked it appeared as Not Locked in IT user and can be released only by Reset Password. | WIP14 | WIP18 | Nor required. | Subsystem | Login | Console-SRS.15910 | Fixed | kapadia_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP16SubCom | Ovation |  | WIP18 | June 3, 2019 7:37:13 PM GMT+08:00
CTUDT00623021 | Refresh | Closed | Revital_Cohen | April 10, 2019 7:39:03 PM GMT+08:00 | Refresh- Emergency users shouldn’t be locked and when locked it appeared as Not Locked in IT user and can be released only by Reset Password. | WIP14 | WIP18 | Nor required. | Subsystem | Login | Console-SRS.15910 | Fixed | kapadia_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP16SubCom | Brilliance 64 |  | WIP18 | June 3, 2019 7:37:13 PM GMT+08:00
CTUDT00623021 | Refresh | Closed | Revital_Cohen | April 10, 2019 7:39:03 PM GMT+08:00 | Refresh- Emergency users shouldn’t be locked and when locked it appeared as Not Locked in IT user and can be released only by Reset Password. | WIP14 | WIP18 | Nor required. | Subsystem | Login | Console-SRS.15910 | Fixed | kapadia_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP16SubCom | Ingenuity 64 |  | WIP18 | June 3, 2019 7:37:13 PM GMT+08:00
CTUDT00623021 | Refresh | Closed | Revital_Cohen | April 10, 2019 7:39:03 PM GMT+08:00 | Refresh- Emergency users shouldn’t be locked and when locked it appeared as Not Locked in IT user and can be released only by Reset Password. | WIP14 | WIP18 | Nor required. | Subsystem | Login | Console-SRS.15910 | Fixed | kapadia_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP16SubCom | IQon1 |  | WIP18 | June 3, 2019 7:37:13 PM GMT+08:00
CTUDT00623019 | Refresh | Closed | PremKumar_Rajendran | April 10, 2019 7:34:29 PM GMT+08:00 | Refresh : Auto Launch option does not retain the previous selection of the results to different viewer in Online & View2 | WIP14 | WIP17 | User has to refresh the UI by clicking on some other result in the series. | Subsystem | UI Performance | NA | Fixed | kr_m | 0 | 3 | Console Software | Console4 | NO | Closed | WIP16SubCom | Ovation |  | WIP17 | April 29, 2019 2:32:00 PM GMT+08:00
CTUDT00623019 | Refresh | Closed | PremKumar_Rajendran | April 10, 2019 7:34:29 PM GMT+08:00 | Refresh : Auto Launch option does not retain the previous selection of the results to different viewer in Online & View2 | WIP14 | WIP17 | User has to refresh the UI by clicking on some other result in the series. | Subsystem | UI Performance | NA | Fixed | kr_m | 0 | 3 | Console Software | Console4 | NO | Closed | WIP16SubCom | Brilliance 64 |  | WIP17 | April 29, 2019 2:32:00 PM GMT+08:00
CTUDT00623019 | Refresh | Closed | PremKumar_Rajendran | April 10, 2019 7:34:29 PM GMT+08:00 | Refresh : Auto Launch option does not retain the previous selection of the results to different viewer in Online & View2 | WIP14 | WIP17 | User has to refresh the UI by clicking on some other result in the series. | Subsystem | UI Performance | NA | Fixed | kr_m | 0 | 3 | Console Software | Console4 | NO | Closed | WIP16SubCom | Ingenuity 64 |  | WIP17 | April 29, 2019 2:32:00 PM GMT+08:00
CTUDT00623019 | Refresh | Closed | PremKumar_Rajendran | April 10, 2019 7:34:29 PM GMT+08:00 | Refresh : Auto Launch option does not retain the previous selection of the results to different viewer in Online & View2 | WIP14 | WIP17 | User has to refresh the UI by clicking on some other result in the series. | Subsystem | UI Performance | NA | Fixed | kr_m | 0 | 3 | Console Software | Console4 | NO | Closed | WIP16SubCom | IQon1 |  | WIP17 | April 29, 2019 2:32:00 PM GMT+08:00
CTUDT00623017 | Refresh | Closed | sunilkumar_j | April 10, 2019 7:24:14 PM GMT+08:00 | Ingenuity [Refresh WIP 14] : Backup Restore option was not listed under utilities section | WIP14 | DHF303211_Rev09_Host_SSRS |  | Subsystem | Service Tools | Console-SSRS.2934 | Fixed | hatti_c | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | WIP16SubCom | Ovation |  | DHF303211_Rev09_Host_SSRS | May 28, 2019 8:00:50 PM GMT+08:00
CTUDT00623017 | Refresh | Closed | sunilkumar_j | April 10, 2019 7:24:14 PM GMT+08:00 | Ingenuity [Refresh WIP 14] : Backup Restore option was not listed under utilities section | WIP14 | DHF303211_Rev09_Host_SSRS |  | Subsystem | Service Tools | Console-SSRS.2934 | Fixed | hatti_c | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | WIP16SubCom | Brilliance 64 |  | DHF303211_Rev09_Host_SSRS | May 28, 2019 8:00:50 PM GMT+08:00
CTUDT00623017 | Refresh | Closed | sunilkumar_j | April 10, 2019 7:24:14 PM GMT+08:00 | Ingenuity [Refresh WIP 14] : Backup Restore option was not listed under utilities section | WIP14 | DHF303211_Rev09_Host_SSRS |  | Subsystem | Service Tools | Console-SSRS.2934 | Fixed | hatti_c | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | WIP16SubCom | Brilliance Big Bore |  | DHF303211_Rev09_Host_SSRS | May 28, 2019 8:00:50 PM GMT+08:00
CTUDT00623017 | Refresh | Closed | sunilkumar_j | April 10, 2019 7:24:14 PM GMT+08:00 | Ingenuity [Refresh WIP 14] : Backup Restore option was not listed under utilities section | WIP14 | DHF303211_Rev09_Host_SSRS |  | Subsystem | Service Tools | Console-SSRS.2934 | Fixed | hatti_c | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | WIP16SubCom | Ingenuity 64 |  | DHF303211_Rev09_Host_SSRS | May 28, 2019 8:00:50 PM GMT+08:00
CTUDT00623017 | Refresh | Closed | sunilkumar_j | April 10, 2019 7:24:14 PM GMT+08:00 | Ingenuity [Refresh WIP 14] : Backup Restore option was not listed under utilities section | WIP14 | DHF303211_Rev09_Host_SSRS |  | Subsystem | Service Tools | Console-SSRS.2934 | Fixed | hatti_c | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | WIP16SubCom | IQon1 |  | DHF303211_Rev09_Host_SSRS | May 28, 2019 8:00:50 PM GMT+08:00
CTUDT00623015 | Refresh | Closed | nanjunda_s | April 10, 2019 7:11:23 PM GMT+08:00 | Conversion report is not generated after upgrade operation completed. | WIP14 | WIP16 |  | Subsystem | Installation | Console-SRS.14729 | Fixed | navale_s | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP16SubCom | Ovation |  | WIP16 | April 20, 2019 12:53:06 AM GMT+08:00
CTUDT00623015 | Refresh | Closed | nanjunda_s | April 10, 2019 7:11:23 PM GMT+08:00 | Conversion report is not generated after upgrade operation completed. | WIP14 | WIP16 |  | Subsystem | Installation | Console-SRS.14729 | Fixed | navale_s | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP16SubCom | Brilliance 64 |  | WIP16 | April 20, 2019 12:53:06 AM GMT+08:00
CTUDT00623015 | Refresh | Closed | nanjunda_s | April 10, 2019 7:11:23 PM GMT+08:00 | Conversion report is not generated after upgrade operation completed. | WIP14 | WIP16 |  | Subsystem | Installation | Console-SRS.14729 | Fixed | navale_s | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP16SubCom | Ingenuity 64 |  | WIP16 | April 20, 2019 12:53:06 AM GMT+08:00
CTUDT00623015 | Refresh | Closed | nanjunda_s | April 10, 2019 7:11:23 PM GMT+08:00 | Conversion report is not generated after upgrade operation completed. | WIP14 | WIP16 |  | Subsystem | Installation | Console-SRS.14729 | Fixed | navale_s | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP16SubCom | IQon1 |  | WIP16 | April 20, 2019 12:53:06 AM GMT+08:00
CTUDT00623003 | Refresh | Closed | shekhar_p | April 9, 2019 8:11:18 PM GMT+08:00 | Refresh:Not able to launch the PSC tools | WIP15 Sanity | WIP16 BB_WIP6 |  | Subsystem | Service Tools | N/A | Fixed | navale_s | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP16SubCom | Ovation |  | WIP16 BB_WIP6 | April 22, 2019 8:46:06 PM GMT+08:00
CTUDT00623003 | Refresh | Closed | shekhar_p | April 9, 2019 8:11:18 PM GMT+08:00 | Refresh:Not able to launch the PSC tools | WIP15 Sanity | WIP16 BB_WIP6 |  | Subsystem | Service Tools | N/A | Fixed | navale_s | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP16SubCom | Brilliance 64 |  | WIP16 BB_WIP6 | April 22, 2019 8:46:06 PM GMT+08:00
CTUDT00623003 | Refresh | Closed | shekhar_p | April 9, 2019 8:11:18 PM GMT+08:00 | Refresh:Not able to launch the PSC tools | WIP15 Sanity | WIP16 BB_WIP6 |  | Subsystem | Service Tools | N/A | Fixed | navale_s | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP16SubCom | Brilliance Big Bore |  | WIP16 BB_WIP6 | April 22, 2019 8:46:06 PM GMT+08:00
CTUDT00623003 | Refresh | Closed | shekhar_p | April 9, 2019 8:11:18 PM GMT+08:00 | Refresh:Not able to launch the PSC tools | WIP15 Sanity | WIP16 BB_WIP6 |  | Subsystem | Service Tools | N/A | Fixed | navale_s | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP16SubCom | Ingenuity 64 |  | WIP16 BB_WIP6 | April 22, 2019 8:46:06 PM GMT+08:00
CTUDT00623003 | Refresh | Closed | shekhar_p | April 9, 2019 8:11:18 PM GMT+08:00 | Refresh:Not able to launch the PSC tools | WIP15 Sanity | WIP16 BB_WIP6 |  | Subsystem | Service Tools | N/A | Fixed | navale_s | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP16SubCom | IQon1 |  | WIP16 BB_WIP6 | April 22, 2019 8:46:06 PM GMT+08:00
CTUDT00623001 | Refresh | Closed | g_di | April 9, 2019 7:15:45 PM GMT+08:00 | Displaying wrong values for Default (Type in) Windowing preset value | WIP14 | WIP17 DHF303211 Rev16 DHF316705 Rev16 | N/A | Subsystem | Windowing | Console-SRS.17124, Console-SRS.16960 | Fixed | roy_r | 0 | 4 | Console USP_EBW | Console4 | NO | Closed | WIP16SegReq | Ovation |  | WIP17 DHF303211 Rev16 DHF316705 Rev16 | October 3, 2019 1:35:09 PM GMT+08:00
CTUDT00623001 | Refresh | Closed | g_di | April 9, 2019 7:15:45 PM GMT+08:00 | Displaying wrong values for Default (Type in) Windowing preset value | WIP14 | WIP17 DHF303211 Rev16 DHF316705 Rev16 | N/A | Subsystem | Windowing | Console-SRS.17124, Console-SRS.16960 | Fixed | roy_r | 0 | 4 | Console USP_EBW | Console4 | NO | Closed | WIP16SegReq | Brilliance Big Bore |  | WIP17 DHF303211 Rev16 DHF316705 Rev16 | October 3, 2019 1:35:09 PM GMT+08:00
CTUDT00623001 | Refresh | Closed | g_di | April 9, 2019 7:15:45 PM GMT+08:00 | Displaying wrong values for Default (Type in) Windowing preset value | WIP14 | WIP17 DHF303211 Rev16 DHF316705 Rev16 | N/A | Subsystem | Windowing | Console-SRS.17124, Console-SRS.16960 | Fixed | roy_r | 0 | 4 | Console USP_EBW | Console4 | NO | Closed | WIP16SegReq | Ingenuity 64 |  | WIP17 DHF303211 Rev16 DHF316705 Rev16 | October 3, 2019 1:35:09 PM GMT+08:00
CTUDT00623001 | Refresh | Closed | g_di | April 9, 2019 7:15:45 PM GMT+08:00 | Displaying wrong values for Default (Type in) Windowing preset value | WIP14 | WIP17 DHF303211 Rev16 DHF316705 Rev16 | N/A | Subsystem | Windowing | Console-SRS.17124, Console-SRS.16960 | Fixed | roy_r | 0 | 4 | Console USP_EBW | Console4 | NO | Closed | WIP16SegReq | IQon1 |  | WIP17 DHF303211 Rev16 DHF316705 Rev16 | October 3, 2019 1:35:09 PM GMT+08:00
CTUDT00622999 | Refresh | Closed | shekhar_p | April 9, 2019 6:22:48 PM GMT+08:00 | User is not able to launch the service tool from Philips service | WIP15 Sanity | WIP17 BB_WIP7 |  | Subsystem | Service Tools | N/A | Fixed | kapadia_v | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP16SubCom | Ovation |  | WIP17 BB_WIP7 | May 8, 2019 10:34:47 PM GMT+08:00
CTUDT00622999 | Refresh | Closed | shekhar_p | April 9, 2019 6:22:48 PM GMT+08:00 | User is not able to launch the service tool from Philips service | WIP15 Sanity | WIP17 BB_WIP7 |  | Subsystem | Service Tools | N/A | Fixed | kapadia_v | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP16SubCom | Brilliance 64 |  | WIP17 BB_WIP7 | May 8, 2019 10:34:47 PM GMT+08:00
CTUDT00622999 | Refresh | Closed | shekhar_p | April 9, 2019 6:22:48 PM GMT+08:00 | User is not able to launch the service tool from Philips service | WIP15 Sanity | WIP17 BB_WIP7 |  | Subsystem | Service Tools | N/A | Fixed | kapadia_v | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP16SubCom | Brilliance Big Bore |  | WIP17 BB_WIP7 | May 8, 2019 10:34:47 PM GMT+08:00
CTUDT00622999 | Refresh | Closed | shekhar_p | April 9, 2019 6:22:48 PM GMT+08:00 | User is not able to launch the service tool from Philips service | WIP15 Sanity | WIP17 BB_WIP7 |  | Subsystem | Service Tools | N/A | Fixed | kapadia_v | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP16SubCom | Ingenuity 64 |  | WIP17 BB_WIP7 | May 8, 2019 10:34:47 PM GMT+08:00
CTUDT00622999 | Refresh | Closed | shekhar_p | April 9, 2019 6:22:48 PM GMT+08:00 | User is not able to launch the service tool from Philips service | WIP15 Sanity | WIP17 BB_WIP7 |  | Subsystem | Service Tools | N/A | Fixed | kapadia_v | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP16SubCom | IQon1 |  | WIP17 BB_WIP7 | May 8, 2019 10:34:47 PM GMT+08:00
CTUDT00622997 | Refresh | Closed | shekhar_p | April 9, 2019 6:10:45 PM GMT+08:00 | Refresh:LogDB is missing in D:\PMS\log and system crash observed | WIP15 Sanity |  |  | Subsystem | 1-click Installation |  | Works As Specified / Not A Defect | babu_s | 1 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00622995 | Refresh | Closed | shekhar_p | April 9, 2019 5:50:55 PM GMT+08:00 | On right monitor the texts are hiding for the place where it displays free disk space availability of Local directory | WIP15 Sanity | WIP17 | NA | Subsystem | UI | NA | Fixed | srushti_k | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP17SegReq | Ovation |  | WIP17 | May 27, 2019 2:12:13 PM GMT+08:00
CTUDT00622995 | Refresh | Closed | shekhar_p | April 9, 2019 5:50:55 PM GMT+08:00 | On right monitor the texts are hiding for the place where it displays free disk space availability of Local directory | WIP15 Sanity | WIP17 | NA | Subsystem | UI | NA | Fixed | srushti_k | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP17SegReq | Brilliance 64 |  | WIP17 | May 27, 2019 2:12:13 PM GMT+08:00
CTUDT00622995 | Refresh | Closed | shekhar_p | April 9, 2019 5:50:55 PM GMT+08:00 | On right monitor the texts are hiding for the place where it displays free disk space availability of Local directory | WIP15 Sanity | WIP17 | NA | Subsystem | UI | NA | Fixed | srushti_k | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP17SegReq | Brilliance Big Bore |  | WIP17 | May 27, 2019 2:12:13 PM GMT+08:00
CTUDT00622995 | Refresh | Closed | shekhar_p | April 9, 2019 5:50:55 PM GMT+08:00 | On right monitor the texts are hiding for the place where it displays free disk space availability of Local directory | WIP15 Sanity | WIP17 | NA | Subsystem | UI | NA | Fixed | srushti_k | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP17SegReq | Ingenuity 64 |  | WIP17 | May 27, 2019 2:12:13 PM GMT+08:00
CTUDT00622995 | Refresh | Closed | shekhar_p | April 9, 2019 5:50:55 PM GMT+08:00 | On right monitor the texts are hiding for the place where it displays free disk space availability of Local directory | WIP15 Sanity | WIP17 | NA | Subsystem | UI | NA | Fixed | srushti_k | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP17SegReq | IQon1 |  | WIP17 | May 27, 2019 2:12:13 PM GMT+08:00
CTUDT00622990 | Refresh | Closed | shekhar_p | April 9, 2019 5:29:06 PM GMT+08:00 | Login user name is not displaying in ECM or workflow | WIP15 Sanity | WIP17 BB_WIP7 | User needs to login again to the system. | Subsystem | ECM | NA | Fixed | kapadia_v | 0 | 4 | Console Software | Console4 | NO | Closed | WIP17SubCom | Ovation |  | WIP17 BB_WIP7 | May 15, 2019 8:38:19 PM GMT+08:00
CTUDT00622990 | Refresh | Closed | shekhar_p | April 9, 2019 5:29:06 PM GMT+08:00 | Login user name is not displaying in ECM or workflow | WIP15 Sanity | WIP17 BB_WIP7 | User needs to login again to the system. | Subsystem | ECM | NA | Fixed | kapadia_v | 0 | 4 | Console Software | Console4 | NO | Closed | WIP17SubCom | Brilliance 64 |  | WIP17 BB_WIP7 | May 15, 2019 8:38:19 PM GMT+08:00
CTUDT00622990 | Refresh | Closed | shekhar_p | April 9, 2019 5:29:06 PM GMT+08:00 | Login user name is not displaying in ECM or workflow | WIP15 Sanity | WIP17 BB_WIP7 | User needs to login again to the system. | Subsystem | ECM | NA | Fixed | kapadia_v | 0 | 4 | Console Software | Console4 | NO | Closed | WIP17SubCom | Brilliance Big Bore |  | WIP17 BB_WIP7 | May 15, 2019 8:38:19 PM GMT+08:00
CTUDT00622990 | Refresh | Closed | shekhar_p | April 9, 2019 5:29:06 PM GMT+08:00 | Login user name is not displaying in ECM or workflow | WIP15 Sanity | WIP17 BB_WIP7 | User needs to login again to the system. | Subsystem | ECM | NA | Fixed | kapadia_v | 0 | 4 | Console Software | Console4 | NO | Closed | WIP17SubCom | Ingenuity 64 |  | WIP17 BB_WIP7 | May 15, 2019 8:38:19 PM GMT+08:00
CTUDT00622990 | Refresh | Closed | shekhar_p | April 9, 2019 5:29:06 PM GMT+08:00 | Login user name is not displaying in ECM or workflow | WIP15 Sanity | WIP17 BB_WIP7 | User needs to login again to the system. | Subsystem | ECM | NA | Fixed | kapadia_v | 0 | 4 | Console Software | Console4 | NO | Closed | WIP17SubCom | IQon1 |  | WIP17 BB_WIP7 | May 15, 2019 8:38:19 PM GMT+08:00
CTUDT00622988 | Refresh | Closed | shekhar_p | April 9, 2019 5:12:43 PM GMT+08:00 | Refresh: The appropriate message is not displaying after the bolus triggered gated cardiac scans | WIP15 Sanity | WIP16 |  | Subsystem | Bolus | NA | Fixed | datta_i | 0 | 2 | Console Software | Console4 | YES | Closed | WIP16SubCom | Ovation |  | WIP16 | April 15, 2019 8:18:38 PM GMT+08:00
CTUDT00622988 | Refresh | Closed | shekhar_p | April 9, 2019 5:12:43 PM GMT+08:00 | Refresh: The appropriate message is not displaying after the bolus triggered gated cardiac scans | WIP15 Sanity | WIP16 |  | Subsystem | Bolus | NA | Fixed | datta_i | 0 | 2 | Console Software | Console4 | YES | Closed | WIP16SubCom | Brilliance 64 |  | WIP16 | April 15, 2019 8:18:38 PM GMT+08:00
CTUDT00622988 | Refresh | Closed | shekhar_p | April 9, 2019 5:12:43 PM GMT+08:00 | Refresh: The appropriate message is not displaying after the bolus triggered gated cardiac scans | WIP15 Sanity | WIP16 |  | Subsystem | Bolus | NA | Fixed | datta_i | 0 | 2 | Console Software | Console4 | YES | Closed | WIP16SubCom | Ingenuity 64 |  | WIP16 | April 15, 2019 8:18:38 PM GMT+08:00
CTUDT00622988 | Refresh | Closed | shekhar_p | April 9, 2019 5:12:43 PM GMT+08:00 | Refresh: The appropriate message is not displaying after the bolus triggered gated cardiac scans | WIP15 Sanity | WIP16 |  | Subsystem | Bolus | NA | Fixed | datta_i | 0 | 2 | Console Software | Console4 | YES | Closed | WIP16SubCom | IQon1 |  | WIP16 | April 15, 2019 8:18:38 PM GMT+08:00
CTUDT00622981 | Refresh | Closed | lei_yang | April 9, 2019 10:25:27 AM GMT+08:00 | The display and operation of message "Maximal allowed number of groups for perfusion is 10" are out of sync. | WIP14 | WIP16 |  | Subsystem | UI Performance | NA | Fixed | v_sumanth | 0 | 2 | Console Software | Console4 | YES | Closed | WIP16SubCom | Ovation |  | WIP16 | April 19, 2019 4:26:18 PM GMT+08:00
CTUDT00622981 | Refresh | Closed | lei_yang | April 9, 2019 10:25:27 AM GMT+08:00 | The display and operation of message "Maximal allowed number of groups for perfusion is 10" are out of sync. | WIP14 | WIP16 |  | Subsystem | UI Performance | NA | Fixed | v_sumanth | 0 | 2 | Console Software | Console4 | YES | Closed | WIP16SubCom | Brilliance 64 |  | WIP16 | April 19, 2019 4:26:18 PM GMT+08:00
CTUDT00622981 | Refresh | Closed | lei_yang | April 9, 2019 10:25:27 AM GMT+08:00 | The display and operation of message "Maximal allowed number of groups for perfusion is 10" are out of sync. | WIP14 | WIP16 |  | Subsystem | UI Performance | NA | Fixed | v_sumanth | 0 | 2 | Console Software | Console4 | YES | Closed | WIP16SubCom | Ingenuity 64 |  | WIP16 | April 19, 2019 4:26:18 PM GMT+08:00
CTUDT00622981 | Refresh | Closed | lei_yang | April 9, 2019 10:25:27 AM GMT+08:00 | The display and operation of message "Maximal allowed number of groups for perfusion is 10" are out of sync. | WIP14 | WIP16 |  | Subsystem | UI Performance | NA | Fixed | v_sumanth | 0 | 2 | Console Software | Console4 | YES | Closed | WIP16SubCom | IQon1 |  | WIP16 | April 19, 2019 4:26:18 PM GMT+08:00
CTUDT00622977 | Refresh | Closed | morton_t | April 9, 2019 2:20:09 AM GMT+08:00 | REFRESH IQon: APlane opening diagnostic test shows incorrect collimations | WIP14 | WIP18 |  | Subsystem | Service Tools | n/a | Fixed | zhou_x | 0 | 2 | Gantry Software | GCX | NO | Closed | WIP17SegReq | Ovation |  | WIP18 | June 13, 2019 2:55:15 PM GMT+08:00
CTUDT00622977 | Refresh | Closed | morton_t | April 9, 2019 2:20:09 AM GMT+08:00 | REFRESH IQon: APlane opening diagnostic test shows incorrect collimations | WIP14 | WIP18 |  | Subsystem | Service Tools | n/a | Fixed | zhou_x | 0 | 2 | Gantry Software | GCX | NO | Closed | WIP17SegReq | IQon1 |  | WIP18 | June 13, 2019 2:55:15 PM GMT+08:00
CTUDT00622969 | Refresh | Closed | altshuller_mark | April 8, 2019 8:36:58 PM GMT+08:00 | IV:Reference Executed Surview did not appear in Exam Summary as expected. | 4.1.7 | BB_WIP6 | NA | Subsystem | Executed Surview | None | Fixed | Prakash_Chand_Yadav | 0 | 3 | Console Generic | Console4 | NO | Closed | BB_WIP6_SubCom | Brilliance Big Bore |  | BB_WIP6 | April 23, 2019 12:03:16 PM GMT+08:00
CTUDT00622966 | Refresh | Closed | manjunath_c | April 8, 2019 7:05:44 PM GMT+08:00 | Bottom part of Worklist window in patient page was greyed out | WIP14 |  | search by Patient ID or Accession Number | Subsystem | Planning | N/A | Works As Specified / Not A Defect | ramesh_t | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00622965 | Refresh | Closed | gT_s | April 8, 2019 7:03:07 PM GMT+08:00 | Refresh : The clinical user is not alerted with the message when the new version of software is available for installation.on PRS portal for download. | WIP 14 | WIP16 |  | Subsystem | Service Tools | NA | Fixed | kapadia_v | 0 | 3 | Console Software | Service Tools | NO | Closed | WIP16SubCom | Ovation |  | WIP16 | April 15, 2019 9:16:56 PM GMT+08:00
CTUDT00622965 | Refresh | Closed | gT_s | April 8, 2019 7:03:07 PM GMT+08:00 | Refresh : The clinical user is not alerted with the message when the new version of software is available for installation.on PRS portal for download. | WIP 14 | WIP16 |  | Subsystem | Service Tools | NA | Fixed | kapadia_v | 0 | 3 | Console Software | Service Tools | NO | Closed | WIP16SubCom | Brilliance 64 |  | WIP16 | April 15, 2019 9:16:56 PM GMT+08:00
CTUDT00622965 | Refresh | Closed | gT_s | April 8, 2019 7:03:07 PM GMT+08:00 | Refresh : The clinical user is not alerted with the message when the new version of software is available for installation.on PRS portal for download. | WIP 14 | WIP16 |  | Subsystem | Service Tools | NA | Fixed | kapadia_v | 0 | 3 | Console Software | Service Tools | NO | Closed | WIP16SubCom | Ingenuity 64 |  | WIP16 | April 15, 2019 9:16:56 PM GMT+08:00
CTUDT00622965 | Refresh | Closed | gT_s | April 8, 2019 7:03:07 PM GMT+08:00 | Refresh : The clinical user is not alerted with the message when the new version of software is available for installation.on PRS portal for download. | WIP 14 | WIP16 |  | Subsystem | Service Tools | NA | Fixed | kapadia_v | 0 | 3 | Console Software | Service Tools | NO | Closed | WIP16SubCom | IQon1 |  | WIP16 | April 15, 2019 9:16:56 PM GMT+08:00
CTUDT00622961 | Refresh | Closed | lei_yang | April 8, 2019 5:37:06 PM GMT+08:00 | The exceed cycle value in log view message is incorrect. | WIP14 | WIP17 | There is no workaround. This issue doesn't cause any obstacle to perform the scan and it is purely concerned with the log message that is logged into log viewer. | Subsystem | Log Viewer | NA | Fixed | jayakumar_p | 0 | 4 | Console Software | Console4 | NO | Closed | WIP16SegReq | Ovation |  | WIP17 | April 29, 2019 2:48:59 PM GMT+08:00
CTUDT00622961 | Refresh | Closed | lei_yang | April 8, 2019 5:37:06 PM GMT+08:00 | The exceed cycle value in log view message is incorrect. | WIP14 | WIP17 | There is no workaround. This issue doesn't cause any obstacle to perform the scan and it is purely concerned with the log message that is logged into log viewer. | Subsystem | Log Viewer | NA | Fixed | jayakumar_p | 0 | 4 | Console Software | Console4 | NO | Closed | WIP16SegReq | Brilliance 64 |  | WIP17 | April 29, 2019 2:48:59 PM GMT+08:00
CTUDT00622961 | Refresh | Closed | lei_yang | April 8, 2019 5:37:06 PM GMT+08:00 | The exceed cycle value in log view message is incorrect. | WIP14 | WIP17 | There is no workaround. This issue doesn't cause any obstacle to perform the scan and it is purely concerned with the log message that is logged into log viewer. | Subsystem | Log Viewer | NA | Fixed | jayakumar_p | 0 | 4 | Console Software | Console4 | NO | Closed | WIP16SegReq | Ingenuity 64 |  | WIP17 | April 29, 2019 2:48:59 PM GMT+08:00
CTUDT00622961 | Refresh | Closed | lei_yang | April 8, 2019 5:37:06 PM GMT+08:00 | The exceed cycle value in log view message is incorrect. | WIP14 | WIP17 | There is no workaround. This issue doesn't cause any obstacle to perform the scan and it is purely concerned with the log message that is logged into log viewer. | Subsystem | Log Viewer | NA | Fixed | jayakumar_p | 0 | 4 | Console Software | Console4 | NO | Closed | WIP16SegReq | IQon1 |  | WIP17 | April 29, 2019 2:48:59 PM GMT+08:00
CTUDT00622959 | Refresh | Closed | patra_r | April 8, 2019 5:22:39 PM GMT+08:00 | Poor image quality of MPR (non axial) IMR results | V4.7.5 | BB_WIP6 | Use offline recon | Subsystem | Image Quality | None | Fixed | patra_r | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP6_SubCom | Brilliance Big Bore |  | BB_WIP6 | April 25, 2019 7:14:12 PM GMT+08:00
CTUDT00622953 | Refresh | Closed | janakiraman_j | April 8, 2019 4:38:47 PM GMT+08:00 | Console sent inconsistent rotation time data to Gantry and CIRS and resulted in ring artifact | 4.1.7 | WIP16 | 1. Perform scan without AST 2. If scan is performed with AST, do not perform pan/zoom operation on online viewer. | System | System | N/A | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | YES | Closed | WIP16SubCom | Ovation | PR8536687 | WIP16 | April 16, 2019 5:52:41 PM GMT+08:00
CTUDT00622953 | Refresh | Closed | janakiraman_j | April 8, 2019 4:38:47 PM GMT+08:00 | Console sent inconsistent rotation time data to Gantry and CIRS and resulted in ring artifact | 4.1.7 | WIP16 | 1. Perform scan without AST 2. If scan is performed with AST, do not perform pan/zoom operation on online viewer. | System | System | N/A | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | YES | Closed | WIP16SubCom | Brilliance 64 | PR8536687 | WIP16 | April 16, 2019 5:52:41 PM GMT+08:00
CTUDT00622953 | Refresh | Closed | janakiraman_j | April 8, 2019 4:38:47 PM GMT+08:00 | Console sent inconsistent rotation time data to Gantry and CIRS and resulted in ring artifact | 4.1.7 | WIP16 | 1. Perform scan without AST 2. If scan is performed with AST, do not perform pan/zoom operation on online viewer. | System | System | N/A | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | YES | Closed | WIP16SubCom | Brilliance Big Bore | PR8536687 | WIP16 | April 16, 2019 5:52:41 PM GMT+08:00
CTUDT00622953 | Refresh | Closed | janakiraman_j | April 8, 2019 4:38:47 PM GMT+08:00 | Console sent inconsistent rotation time data to Gantry and CIRS and resulted in ring artifact | 4.1.7 | WIP16 | 1. Perform scan without AST 2. If scan is performed with AST, do not perform pan/zoom operation on online viewer. | System | System | N/A | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | YES | Closed | WIP16SubCom | Ingenuity 64 | PR8536687 | WIP16 | April 16, 2019 5:52:41 PM GMT+08:00
CTUDT00622953 | Refresh | Closed | janakiraman_j | April 8, 2019 4:38:47 PM GMT+08:00 | Console sent inconsistent rotation time data to Gantry and CIRS and resulted in ring artifact | 4.1.7 | WIP16 | 1. Perform scan without AST 2. If scan is performed with AST, do not perform pan/zoom operation on online viewer. | System | System | N/A | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | YES | Closed | WIP16SubCom | IQon1 | PR8536687 | WIP16 | April 16, 2019 5:52:41 PM GMT+08:00
CTUDT00622949 | Refresh | Closed | levy_s | April 8, 2019 3:07:21 PM GMT+08:00 | Refresh: Initializing state is presented to the user although Get-Ready request failed. | WIP14 |  |  | Subsystem | Automation Test | N/A | Invalid/Uncontrolled Configuration | tamari_o | 0 | 2 |  | Support-System | NO | Closed | SubNC | IQon1
CTUDT00622948 | Refresh | Closed | deshbhratar_m | April 8, 2019 3:03:41 PM GMT+08:00 | [Refresh]: Date format displayed next to Date of Birth field was not localized in Spanish language | 4.8 | BB_WIP6 | NA | Subsystem | Localization | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP6_SubCom | Brilliance Big Bore |  | BB_WIP6 | April 25, 2019 9:21:14 PM GMT+08:00
CTUDT00622947 | Refresh | Closed | altshuller_mark | April 8, 2019 3:01:00 PM GMT+08:00 | Poor image quality of MPR (non axial) IMR results | IQon HB1 | WIP16 | Use offline recon | Subsystem | Image Quality | None | Fixed | patra_r | 0 | 3 | Console Software | Console4 | NO | Closed | WIP16SubCom | Ovation |  | WIP16 | April 19, 2019 6:42:40 PM GMT+08:00
CTUDT00622947 | Refresh | Closed | altshuller_mark | April 8, 2019 3:01:00 PM GMT+08:00 | Poor image quality of MPR (non axial) IMR results | IQon HB1 | WIP16 | Use offline recon | Subsystem | Image Quality | None | Fixed | patra_r | 0 | 3 | Console Software | Console4 | NO | Closed | WIP16SubCom | Brilliance 64 |  | WIP16 | April 19, 2019 6:42:40 PM GMT+08:00
CTUDT00622947 | Refresh | Closed | altshuller_mark | April 8, 2019 3:01:00 PM GMT+08:00 | Poor image quality of MPR (non axial) IMR results | IQon HB1 | WIP16 | Use offline recon | Subsystem | Image Quality | None | Fixed | patra_r | 0 | 3 | Console Software | Console4 | NO | Closed | WIP16SubCom | Ingenuity 64 |  | WIP16 | April 19, 2019 6:42:40 PM GMT+08:00
CTUDT00622947 | Refresh | Closed | altshuller_mark | April 8, 2019 3:01:00 PM GMT+08:00 | Poor image quality of MPR (non axial) IMR results | IQon HB1 | WIP16 | Use offline recon | Subsystem | Image Quality | None | Fixed | patra_r | 0 | 3 | Console Software | Console4 | NO | Closed | WIP16SubCom | IQon1 |  | WIP16 | April 19, 2019 6:42:40 PM GMT+08:00
CTUDT00622944 | Refresh | Closed | beck_d | April 7, 2019 11:22:40 PM GMT+08:00 | Refresh: Tube Current: Collimation is not as defined in Spec. | WIP14 | WIP17 |  | Subsystem | Performance | n/a | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | SubCom | Ovation |  | WIP17 | May 26, 2019 4:36:40 PM GMT+08:00
CTUDT00622941 | Refresh | Closed | alex_nihaev | April 7, 2019 9:14:03 PM GMT+08:00 | Refresh: Visual Diagnostics: DMS option is not available for Visual Diagnostics | WIP14 | WIP17 |  | Subsystem | Service Tools | N/A | Fixed | franco_a | 0 | 2 | DMS Software | DMS | NO | Closed | SubCom | Ovation |  | WIP17 | May 26, 2019 7:41:12 PM GMT+08:00
CTUDT00622939 | Refresh | Closed | Ashish_Sharma | April 7, 2019 7:41:11 PM GMT+08:00 | In spectral CT viewer Application After loading study to spectral viewer, warning message is coming in status bar but couch is not removing from images. | WIP13 | WIP18 | No | Subsystem | Spectral CT Viewer | CT-NM.SRS.TA.356 to 439 | Works As Specified / Not A Defect | aviw | 0 | 4 | SDViewer.exe | Support_Spectral | NO | Closed | WIP18SegReq | IQon1 |  | WIP18 | June 26, 2019 10:34:26 PM GMT+08:00
CTUDT00622937 | Refresh | Closed | Ashish_Sharma | April 7, 2019 7:11:28 PM GMT+08:00 | Dicom information for rescale type,electron density was not showing correct value on saving image as original in spectral CT viewer Application. | WIP14 | WIP18 |  | Subsystem | sCTV ; Blocker | NA | Fixed | aviw | 0 | 2 | Dicom.dll | Support_Spectral | NO | Closed | WIP18SegReq | IQon1 |  | WIP18 | June 7, 2019 8:23:33 PM GMT+08:00
CTUDT00622935 | Refresh | Closed | Ashish_Sharma | April 7, 2019 7:00:46 PM GMT+08:00 | BMD application crashing while performing Pan, Zoom and swivel after changing orientation Axial to coronal or sagittal. | WIP14 | WIP18 |  | Subsystem | BMD ; | NA | No Plan to Fix | aswani_v | 0 | 3 | Console USP_EBW | iCAP | NO | Closed |  | Ovation |  | WIP18 | July 2, 2019 12:47:52 PM GMT+08:00
CTUDT00622935 | Refresh | Closed | Ashish_Sharma | April 7, 2019 7:00:46 PM GMT+08:00 | BMD application crashing while performing Pan, Zoom and swivel after changing orientation Axial to coronal or sagittal. | WIP14 | WIP18 |  | Subsystem | BMD ; | NA | No Plan to Fix | aswani_v | 0 | 3 | Console USP_EBW | iCAP | NO | Closed |  | Brilliance 64 |  | WIP18 | July 2, 2019 12:47:52 PM GMT+08:00
CTUDT00622935 | Refresh | Closed | Ashish_Sharma | April 7, 2019 7:00:46 PM GMT+08:00 | BMD application crashing while performing Pan, Zoom and swivel after changing orientation Axial to coronal or sagittal. | WIP14 | WIP18 |  | Subsystem | BMD ; | NA | No Plan to Fix | aswani_v | 0 | 3 | Console USP_EBW | iCAP | NO | Closed |  | Brilliance Big Bore |  | WIP18 | July 2, 2019 12:47:52 PM GMT+08:00
CTUDT00622935 | Refresh | Closed | Ashish_Sharma | April 7, 2019 7:00:46 PM GMT+08:00 | BMD application crashing while performing Pan, Zoom and swivel after changing orientation Axial to coronal or sagittal. | WIP14 | WIP18 |  | Subsystem | BMD ; | NA | No Plan to Fix | aswani_v | 0 | 3 | Console USP_EBW | iCAP | NO | Closed |  | Ingenuity 64 |  | WIP18 | July 2, 2019 12:47:52 PM GMT+08:00
CTUDT00622935 | Refresh | Closed | Ashish_Sharma | April 7, 2019 7:00:46 PM GMT+08:00 | BMD application crashing while performing Pan, Zoom and swivel after changing orientation Axial to coronal or sagittal. | WIP14 | WIP18 |  | Subsystem | BMD ; | NA | No Plan to Fix | aswani_v | 0 | 3 | Console USP_EBW | iCAP | NO | Closed |  | IQon1 |  | WIP18 | July 2, 2019 12:47:52 PM GMT+08:00
CTUDT00622931 | Refresh | Closed | manjunath_c | April 5, 2019 6:05:46 PM GMT+08:00 | Wrong user message was displayed | WIP14 | WIP16 | No Workaround | Subsystem | Voice Manager | None | Fixed | Prakash_Chand_Yadav | 0 | 4 | Console Software | Console4 | NO | Closed | WIP16SubCom | Ovation |  | WIP16 | April 11, 2019 9:22:20 PM GMT+08:00
CTUDT00622931 | Refresh | Closed | manjunath_c | April 5, 2019 6:05:46 PM GMT+08:00 | Wrong user message was displayed | WIP14 | WIP16 | No Workaround | Subsystem | Voice Manager | None | Fixed | Prakash_Chand_Yadav | 0 | 4 | Console Software | Console4 | NO | Closed | WIP16SubCom | Brilliance 64 |  | WIP16 | April 11, 2019 9:22:20 PM GMT+08:00
CTUDT00622931 | Refresh | Closed | manjunath_c | April 5, 2019 6:05:46 PM GMT+08:00 | Wrong user message was displayed | WIP14 | WIP16 | No Workaround | Subsystem | Voice Manager | None | Fixed | Prakash_Chand_Yadav | 0 | 4 | Console Software | Console4 | NO | Closed | WIP16SubCom | Ingenuity 64 |  | WIP16 | April 11, 2019 9:22:20 PM GMT+08:00
CTUDT00622931 | Refresh | Closed | manjunath_c | April 5, 2019 6:05:46 PM GMT+08:00 | Wrong user message was displayed | WIP14 | WIP16 | No Workaround | Subsystem | Voice Manager | None | Fixed | Prakash_Chand_Yadav | 0 | 4 | Console Software | Console4 | NO | Closed | WIP16SubCom | IQon1 |  | WIP16 | April 11, 2019 9:22:20 PM GMT+08:00
CTUDT00622929 | Refresh | Closed | manjunath_c | April 5, 2019 5:53:02 PM GMT+08:00 | Splash screen once workflow is launched, does not show OS version Number | WIP14 | WIP16 |  | Subsystem | Login | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP18SegReq | Ovation |  | WIP16 | May 14, 2019 4:56:10 PM GMT+08:00
CTUDT00622929 | Refresh | Closed | manjunath_c | April 5, 2019 5:53:02 PM GMT+08:00 | Splash screen once workflow is launched, does not show OS version Number | WIP14 | WIP16 |  | Subsystem | Login | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP18SegReq | Brilliance 64 |  | WIP16 | May 14, 2019 4:56:10 PM GMT+08:00
CTUDT00622929 | Refresh | Closed | manjunath_c | April 5, 2019 5:53:02 PM GMT+08:00 | Splash screen once workflow is launched, does not show OS version Number | WIP14 | WIP16 |  | Subsystem | Login | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP18SegReq | Brilliance Big Bore |  | WIP16 | May 14, 2019 4:56:10 PM GMT+08:00
CTUDT00622929 | Refresh | Closed | manjunath_c | April 5, 2019 5:53:02 PM GMT+08:00 | Splash screen once workflow is launched, does not show OS version Number | WIP14 | WIP16 |  | Subsystem | Login | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP18SegReq | Ingenuity 64 |  | WIP16 | May 14, 2019 4:56:10 PM GMT+08:00
CTUDT00622929 | Refresh | Closed | manjunath_c | April 5, 2019 5:53:02 PM GMT+08:00 | Splash screen once workflow is launched, does not show OS version Number | WIP14 | WIP16 |  | Subsystem | Login | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP18SegReq | IQon1 |  | WIP16 | May 14, 2019 4:56:10 PM GMT+08:00
CTUDT00622925 | Refresh | Closed | datta_i | April 5, 2019 4:28:32 PM GMT+08:00 | The message "Maximal allowed number of cycles for spiral perfusion is 15" still showing if deleting some of group to decrease the cycle number. | 4.8 | BB_WIP6 |  | Subsystem | Host | Na | Fixed | datta_i | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP16_SubCom | Brilliance Big Bore |  | BB_WIP6 | April 19, 2019 7:19:18 PM GMT+08:00
CTUDT00622923 | Refresh | Closed | gT_s | April 5, 2019 3:42:48 PM GMT+08:00 | Refresh : The release information and installation instructions are not displayed when software package is uploaded form the USB. | WIP 14 | WIP16 BB_WIP6 | None | Subsystem | Software Distribution Tool | STPS-SSRS-1246 | Fixed | jose_s | 0 | 4 | Console Software | Service Tools | NO | Closed | WIP16SubCom | Ovation |  | WIP16 BB_WIP6 | April 22, 2019 6:18:42 PM GMT+08:00
CTUDT00622923 | Refresh | Closed | gT_s | April 5, 2019 3:42:48 PM GMT+08:00 | Refresh : The release information and installation instructions are not displayed when software package is uploaded form the USB. | WIP 14 | WIP16 BB_WIP6 | None | Subsystem | Software Distribution Tool | STPS-SSRS-1246 | Fixed | jose_s | 0 | 4 | Console Software | Service Tools | NO | Closed | WIP16SubCom | Brilliance Big Bore |  | WIP16 BB_WIP6 | April 22, 2019 6:18:42 PM GMT+08:00
CTUDT00622923 | Refresh | Closed | gT_s | April 5, 2019 3:42:48 PM GMT+08:00 | Refresh : The release information and installation instructions are not displayed when software package is uploaded form the USB. | WIP 14 | WIP16 BB_WIP6 | None | Subsystem | Software Distribution Tool | STPS-SSRS-1246 | Fixed | jose_s | 0 | 4 | Console Software | Service Tools | NO | Closed | WIP16SubCom | Ingenuity 64 |  | WIP16 BB_WIP6 | April 22, 2019 6:18:42 PM GMT+08:00
CTUDT00622923 | Refresh | Closed | gT_s | April 5, 2019 3:42:48 PM GMT+08:00 | Refresh : The release information and installation instructions are not displayed when software package is uploaded form the USB. | WIP 14 | WIP16 BB_WIP6 | None | Subsystem | Software Distribution Tool | STPS-SSRS-1246 | Fixed | jose_s | 0 | 4 | Console Software | Service Tools | NO | Closed | WIP16SubCom | IQon1 |  | WIP16 BB_WIP6 | April 22, 2019 6:18:42 PM GMT+08:00
CTUDT00622920 | Refresh | Closed | PremKumar_Rajendran | April 5, 2019 2:12:40 PM GMT+08:00 | BB Refresh - Series description was appended with "E1" on Patient Directory in Cardiac result as “UnTagged E1” | BB_WIP5 | BB_WIP6 |  | Subsystem | Cardiac | Console-SRS.15321, Console-SRS.15322 | Fixed | agrawal_ay | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP6_SubCom | Brilliance Big Bore |  | BB_WIP6 | April 29, 2019 5:07:11 PM GMT+08:00
CTUDT00622913 | Refresh | Closed | bandil_j | April 5, 2019 1:48:50 PM GMT+08:00 | pitch too low in scan with auto-scan-time | WIP510 | Host SSRS DHF303211 Rev 10 WIP17 |  | System | System | Console-SRS.10645 | Fixed | killada_s | 0 | 2 | Host SSRS | Console4 | YES | Closed | WIP16SubCom | Ovation | PR8556251 | Host SSRS DHF303211 Rev 10 WIP17 | May 16, 2019 7:02:55 PM GMT+08:00
CTUDT00622913 | Refresh | Closed | bandil_j | April 5, 2019 1:48:50 PM GMT+08:00 | pitch too low in scan with auto-scan-time | WIP510 | Host SSRS DHF303211 Rev 10 WIP17 |  | System | System | Console-SRS.10645 | Fixed | killada_s | 0 | 2 | Host SSRS | Console4 | YES | Closed | WIP16SubCom | IQon1 | PR8556251 | Host SSRS DHF303211 Rev 10 WIP17 | May 16, 2019 7:02:55 PM GMT+08:00
CTUDT00622912 | Refresh | Closed | bandil_j | April 5, 2019 1:47:02 PM GMT+08:00 | "Incorrect pixel spacing among images. Reselect data." message appears when data is loaded to HeartBeat Calcium Score | V4.1.6 | WIP16 | Reducing scan length and changing to different phase-70% this will allow the user to load the images. | System | Cardiac;Verified1 | N/A | Fixed | bhawna_g | 0 | 3 | Console Software | Console4 | NO | Closed | WIP16SubCom | Ovation | 9448055 | WIP16 | April 16, 2019 3:31:24 PM GMT+08:00
CTUDT00622912 | Refresh | Closed | bandil_j | April 5, 2019 1:47:02 PM GMT+08:00 | "Incorrect pixel spacing among images. Reselect data." message appears when data is loaded to HeartBeat Calcium Score | V4.1.6 | WIP16 | Reducing scan length and changing to different phase-70% this will allow the user to load the images. | System | Cardiac;Verified1 | N/A | Fixed | bhawna_g | 0 | 3 | Console Software | Console4 | NO | Closed | WIP16SubCom | Brilliance 64 | 9448055 | WIP16 | April 16, 2019 3:31:24 PM GMT+08:00
CTUDT00622912 | Refresh | Closed | bandil_j | April 5, 2019 1:47:02 PM GMT+08:00 | "Incorrect pixel spacing among images. Reselect data." message appears when data is loaded to HeartBeat Calcium Score | V4.1.6 | WIP16 | Reducing scan length and changing to different phase-70% this will allow the user to load the images. | System | Cardiac;Verified1 | N/A | Fixed | bhawna_g | 0 | 3 | Console Software | Console4 | NO | Closed | WIP16SubCom | Ingenuity 64 | 9448055 | WIP16 | April 16, 2019 3:31:24 PM GMT+08:00
CTUDT00622912 | Refresh | Closed | bandil_j | April 5, 2019 1:47:02 PM GMT+08:00 | "Incorrect pixel spacing among images. Reselect data." message appears when data is loaded to HeartBeat Calcium Score | V4.1.6 | WIP16 | Reducing scan length and changing to different phase-70% this will allow the user to load the images. | System | Cardiac;Verified1 | N/A | Fixed | bhawna_g | 0 | 3 | Console Software | Console4 | NO | Closed | WIP16SubCom | IQon1 | 9448055 | WIP16 | April 16, 2019 3:31:24 PM GMT+08:00
CTUDT00622899 | Refresh | Closed | datta_i | April 5, 2019 12:40:10 PM GMT+08:00 | Refresh - ‘DoseRight’ check box was enabled even though ‘Enabled DoseRight’ option was unchecked from preferences | 4.8 | BB_WIP6 |  | Subsystem | ECM | NA | Fixed | datta_i | 0 | 2 | Console Software | Console4 | YES | Closed | WIP16SubCom | Brilliance Big Bore |  | BB_WIP6 | April 22, 2019 2:01:56 PM GMT+08:00
CTUDT00622895 | Refresh | Closed | g_di | April 4, 2019 8:05:36 PM GMT+08:00 | Spectral Annotation is overlapping in the image when the layout is changed to 4x4 in view tab | WIP14 |  |  | Subsystem | viewer |  | Duplicate | redrouthu_r | 0 |  |  | iCAP | NO | Closed |  | IQon1
CTUDT00622889 | Refresh | Closed | alex_nihaev | April 4, 2019 5:11:42 PM GMT+08:00 | Refresh:ICT: Invalid value inside DMC_NORMAL_ERROR_BKST appear in Cantrace. | WIP13 |  |  | Subsystem | CAN Trace | N/A | Works As Specified / Not A Defect | yatziv_l | 0 |  |  | DMS | NO | Closed |  | iCT2
CTUDT00622883 | Refresh | Closed | bhawna_g | April 4, 2019 2:36:55 PM GMT+08:00 | Duplicate Dose SR series are created for each study. | 4.8 | BB_WIP6 | N/A | Subsystem | Dose SR | N/A | Fixed | agrawal_ay | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP6_SubCom | Brilliance Big Bore |  | BB_WIP6 | April 12, 2019 7:52:48 PM GMT+08:00
CTUDT00622881 | Refresh | Closed | wang_l | April 4, 2019 2:15:30 PM GMT+08:00 | Refresh:Timed out error displayed in bugreport>CollectionErrorWarnings file | BB_WIP5 | WIP18 |  | Subsystem | Bugreport | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP18SegReq | Ovation |  | WIP18 | June 21, 2019 1:54:05 PM GMT+08:00
CTUDT00622881 | Refresh | Closed | wang_l | April 4, 2019 2:15:30 PM GMT+08:00 | Refresh:Timed out error displayed in bugreport>CollectionErrorWarnings file | BB_WIP5 | WIP18 |  | Subsystem | Bugreport | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP18SegReq | Brilliance 64 |  | WIP18 | June 21, 2019 1:54:05 PM GMT+08:00
CTUDT00622881 | Refresh | Closed | wang_l | April 4, 2019 2:15:30 PM GMT+08:00 | Refresh:Timed out error displayed in bugreport>CollectionErrorWarnings file | BB_WIP5 | WIP18 |  | Subsystem | Bugreport | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP18SegReq | Brilliance Big Bore |  | WIP18 | June 21, 2019 1:54:05 PM GMT+08:00
CTUDT00622881 | Refresh | Closed | wang_l | April 4, 2019 2:15:30 PM GMT+08:00 | Refresh:Timed out error displayed in bugreport>CollectionErrorWarnings file | BB_WIP5 | WIP18 |  | Subsystem | Bugreport | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP18SegReq | Ingenuity 64 |  | WIP18 | June 21, 2019 1:54:05 PM GMT+08:00
CTUDT00622881 | Refresh | Closed | wang_l | April 4, 2019 2:15:30 PM GMT+08:00 | Refresh:Timed out error displayed in bugreport>CollectionErrorWarnings file | BB_WIP5 | WIP18 |  | Subsystem | Bugreport | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP18SegReq | IQon1 |  | WIP18 | June 21, 2019 1:54:05 PM GMT+08:00
CTUDT00622880 | Refresh | Closed | patra_r | April 4, 2019 1:27:56 PM GMT+08:00 | v4.1.3 FOK – Scanner never gave the Continue with Manual Scan during this study | V4.1.3 | BB_WIP6 | Press the manual or auto scan button on CT box. | Subsystem | Bolus Tracking | NA | Fixed | Prakash_Chand_Yadav | 0 | 2 | Console Software | Console4 | YES | Closed | WIP16SubCom | Brilliance Big Bore | PR4214074 PR4393676 PR4499398 | BB_WIP6 | April 11, 2019 1:38:35 PM GMT+08:00
CTUDT00622878 | Refresh | Closed | lei_yang | April 4, 2019 11:14:41 AM GMT+08:00 | The result HR phase label displays cannot be refreshed in time in the Show All page. | WIP14 | WIP17 |  | Subsystem | planning | None | Fixed | mY_h | 0 | 2 | Console Software | Console4 | NO | Closed | WIP15SegReq | Ovation |  | WIP17 | April 29, 2019 3:12:01 PM GMT+08:00
CTUDT00622878 | Refresh | Closed | lei_yang | April 4, 2019 11:14:41 AM GMT+08:00 | The result HR phase label displays cannot be refreshed in time in the Show All page. | WIP14 | WIP17 |  | Subsystem | planning | None | Fixed | mY_h | 0 | 2 | Console Software | Console4 | NO | Closed | WIP15SegReq | Brilliance 64 |  | WIP17 | April 29, 2019 3:12:01 PM GMT+08:00
CTUDT00622878 | Refresh | Closed | lei_yang | April 4, 2019 11:14:41 AM GMT+08:00 | The result HR phase label displays cannot be refreshed in time in the Show All page. | WIP14 | WIP17 |  | Subsystem | planning | None | Fixed | mY_h | 0 | 2 | Console Software | Console4 | NO | Closed | WIP15SegReq | Ingenuity 64 |  | WIP17 | April 29, 2019 3:12:01 PM GMT+08:00
CTUDT00622878 | Refresh | Closed | lei_yang | April 4, 2019 11:14:41 AM GMT+08:00 | The result HR phase label displays cannot be refreshed in time in the Show All page. | WIP14 | WIP17 |  | Subsystem | planning | None | Fixed | mY_h | 0 | 2 | Console Software | Console4 | NO | Closed | WIP15SegReq | IQon1 |  | WIP17 | April 29, 2019 3:12:01 PM GMT+08:00
CTUDT00622876 | Refresh | Closed | gT_s | April 3, 2019 10:37:10 PM GMT+08:00 | Refresh : The message displayed when the user click on the install tab for the packages does not match the requirement. | WIP 14 | DHF303211_Rev09_Host_SSRS | NA | Subsystem | Software Distribution Tool | STPS SSRS-1264 | Fixed | hatti_c | 0 | 3 | Requirements Documentation | Service Tools | NO | Closed | WIP16SubCom | Ovation |  | DHF303211_Rev09_Host_SSRS | May 3, 2019 8:03:10 PM GMT+08:00
CTUDT00622876 | Refresh | Closed | gT_s | April 3, 2019 10:37:10 PM GMT+08:00 | Refresh : The message displayed when the user click on the install tab for the packages does not match the requirement. | WIP 14 | DHF303211_Rev09_Host_SSRS | NA | Subsystem | Software Distribution Tool | STPS SSRS-1264 | Fixed | hatti_c | 0 | 3 | Requirements Documentation | Service Tools | NO | Closed | WIP16SubCom | Brilliance Big Bore |  | DHF303211_Rev09_Host_SSRS | May 3, 2019 8:03:10 PM GMT+08:00
CTUDT00622876 | Refresh | Closed | gT_s | April 3, 2019 10:37:10 PM GMT+08:00 | Refresh : The message displayed when the user click on the install tab for the packages does not match the requirement. | WIP 14 | DHF303211_Rev09_Host_SSRS | NA | Subsystem | Software Distribution Tool | STPS SSRS-1264 | Fixed | hatti_c | 0 | 3 | Requirements Documentation | Service Tools | NO | Closed | WIP16SubCom | Ingenuity 64 |  | DHF303211_Rev09_Host_SSRS | May 3, 2019 8:03:10 PM GMT+08:00
CTUDT00622876 | Refresh | Closed | gT_s | April 3, 2019 10:37:10 PM GMT+08:00 | Refresh : The message displayed when the user click on the install tab for the packages does not match the requirement. | WIP 14 | DHF303211_Rev09_Host_SSRS | NA | Subsystem | Software Distribution Tool | STPS SSRS-1264 | Fixed | hatti_c | 0 | 3 | Requirements Documentation | Service Tools | NO | Closed | WIP16SubCom | IQon1 |  | DHF303211_Rev09_Host_SSRS | May 3, 2019 8:03:10 PM GMT+08:00
CTUDT00622874 | Refresh | Closed | shterner_a | April 3, 2019 10:04:01 PM GMT+08:00 | Refresh Pre- Val: System error popped up after enable SyncRight in Preferences | WIP14 | WIP17 |  | Subsystem | Preferences | NA | Fixed | wilsonraja | 0 | 4 | Console Software | Console4 | NO | Closed | WIP16SubCom | Ovation |  | WIP17 | May 3, 2019 2:23:18 PM GMT+08:00
CTUDT00622874 | Refresh | Closed | shterner_a | April 3, 2019 10:04:01 PM GMT+08:00 | Refresh Pre- Val: System error popped up after enable SyncRight in Preferences | WIP14 | WIP17 |  | Subsystem | Preferences | NA | Fixed | wilsonraja | 0 | 4 | Console Software | Console4 | NO | Closed | WIP16SubCom | Brilliance 64 |  | WIP17 | May 3, 2019 2:23:18 PM GMT+08:00
CTUDT00622874 | Refresh | Closed | shterner_a | April 3, 2019 10:04:01 PM GMT+08:00 | Refresh Pre- Val: System error popped up after enable SyncRight in Preferences | WIP14 | WIP17 |  | Subsystem | Preferences | NA | Fixed | wilsonraja | 0 | 4 | Console Software | Console4 | NO | Closed | WIP16SubCom | Ingenuity 64 |  | WIP17 | May 3, 2019 2:23:18 PM GMT+08:00
CTUDT00622874 | Refresh | Closed | shterner_a | April 3, 2019 10:04:01 PM GMT+08:00 | Refresh Pre- Val: System error popped up after enable SyncRight in Preferences | WIP14 | WIP17 |  | Subsystem | Preferences | NA | Fixed | wilsonraja | 0 | 4 | Console Software | Console4 | NO | Closed | WIP16SubCom | IQon1 |  | WIP17 | May 3, 2019 2:23:18 PM GMT+08:00
CTUDT00622872 | Refresh | Closed | bagdady_b | April 3, 2019 7:57:08 PM GMT+08:00 | Refresh- Test design issue for Tools and Applications module | WIP14 |  |  | Subsystem |  |  | Does Not Apply | bagdady_b | 0 |  |  |  |  | Closed |  | Ovation
CTUDT00622872 | Refresh | Closed | bagdady_b | April 3, 2019 7:57:08 PM GMT+08:00 | Refresh- Test design issue for Tools and Applications module | WIP14 |  |  | Subsystem |  |  | Does Not Apply | bagdady_b | 0 |  |  |  |  | Closed |  | IQon1
CTUDT00622871 | Refresh | Closed | bhawna_g | April 3, 2019 5:06:05 PM GMT+08:00 | Refresh - Voice Manager: User is unable to rename user added language. | 4.8 | BB_WIP6 |  | Subsystem | Voice Manager | NA | Fixed | v_sumanth | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP6_SubCom | Brilliance Big Bore |  | BB_WIP6 | April 10, 2019 2:39:54 PM GMT+08:00
CTUDT00622868 | Refresh | Closed | bhawna_g | April 3, 2019 1:55:17 PM GMT+08:00 | Number of Cycles increase on Copy geometry from MPR to axial results with tilt | BB_WIP5 |  |  | Subsystem | Geometry | N/A | No Plan to Fix | shaham_t | 0 | 3 |  | Console4 | YES | Closed |  | Brilliance Big Bore
CTUDT00622861 | Refresh | Closed | wang_l | April 3, 2019 9:27:58 AM GMT+08:00 | Refresh: Olevel user cannot Add and Delete printer | WIP14 | WIP19 | NA | System | Printer Config Tool; Verified1 | NA | Fixed | li_e | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP18SubCom | Ovation |  | WIP19 | July 26, 2019 9:07:36 PM GMT+08:00
CTUDT00622861 | Refresh | Closed | wang_l | April 3, 2019 9:27:58 AM GMT+08:00 | Refresh: Olevel user cannot Add and Delete printer | WIP14 | WIP19 | NA | System | Printer Config Tool; Verified1 | NA | Fixed | li_e | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP18SubCom | Brilliance 64 |  | WIP19 | July 26, 2019 9:07:36 PM GMT+08:00
CTUDT00622861 | Refresh | Closed | wang_l | April 3, 2019 9:27:58 AM GMT+08:00 | Refresh: Olevel user cannot Add and Delete printer | WIP14 | WIP19 | NA | System | Printer Config Tool; Verified1 | NA | Fixed | li_e | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP18SubCom | Ingenuity 64 |  | WIP19 | July 26, 2019 9:07:36 PM GMT+08:00
CTUDT00622861 | Refresh | Closed | wang_l | April 3, 2019 9:27:58 AM GMT+08:00 | Refresh: Olevel user cannot Add and Delete printer | WIP14 | WIP19 | NA | System | Printer Config Tool; Verified1 | NA | Fixed | li_e | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP18SubCom | IQon1 |  | WIP19 | July 26, 2019 9:07:36 PM GMT+08:00
CTUDT00622858 | Refresh | Closed | rakic_k | April 2, 2019 11:03:25 PM GMT+08:00 | ServiceTools Visual Diagnostics Recon Tests Issue | WIP14 | WIP15 |  | Subsystem | CIRS; Verified1 | CIRS.CS.290, CIRS.SSRS.4781, CIRS.SSRS.4856, and CIRS.SSRS.5362 | Fixed | richards_b | 0 | 2 | CIRS_SW_ServiceNode | CIRS | NO | Closed | WIP15SubCom | Ovation |  | WIP15 | April 12, 2019 8:38:04 PM GMT+08:00
CTUDT00622858 | Refresh | Closed | rakic_k | April 2, 2019 11:03:25 PM GMT+08:00 | ServiceTools Visual Diagnostics Recon Tests Issue | WIP14 | WIP15 |  | Subsystem | CIRS; Verified1 | CIRS.CS.290, CIRS.SSRS.4781, CIRS.SSRS.4856, and CIRS.SSRS.5362 | Fixed | richards_b | 0 | 2 | CIRS_SW_ServiceNode | CIRS | NO | Closed | WIP15SubCom | Brilliance 64 |  | WIP15 | April 12, 2019 8:38:04 PM GMT+08:00
CTUDT00622858 | Refresh | Closed | rakic_k | April 2, 2019 11:03:25 PM GMT+08:00 | ServiceTools Visual Diagnostics Recon Tests Issue | WIP14 | WIP15 |  | Subsystem | CIRS; Verified1 | CIRS.CS.290, CIRS.SSRS.4781, CIRS.SSRS.4856, and CIRS.SSRS.5362 | Fixed | richards_b | 0 | 2 | CIRS_SW_ServiceNode | CIRS | NO | Closed | WIP15SubCom | Brilliance Big Bore |  | WIP15 | April 12, 2019 8:38:04 PM GMT+08:00
CTUDT00622858 | Refresh | Closed | rakic_k | April 2, 2019 11:03:25 PM GMT+08:00 | ServiceTools Visual Diagnostics Recon Tests Issue | WIP14 | WIP15 |  | Subsystem | CIRS; Verified1 | CIRS.CS.290, CIRS.SSRS.4781, CIRS.SSRS.4856, and CIRS.SSRS.5362 | Fixed | richards_b | 0 | 2 | CIRS_SW_ServiceNode | CIRS | NO | Closed | WIP15SubCom | Ingenuity 64 |  | WIP15 | April 12, 2019 8:38:04 PM GMT+08:00
CTUDT00622858 | Refresh | Closed | rakic_k | April 2, 2019 11:03:25 PM GMT+08:00 | ServiceTools Visual Diagnostics Recon Tests Issue | WIP14 | WIP15 |  | Subsystem | CIRS; Verified1 | CIRS.CS.290, CIRS.SSRS.4781, CIRS.SSRS.4856, and CIRS.SSRS.5362 | Fixed | richards_b | 0 | 2 | CIRS_SW_ServiceNode | CIRS | NO | Closed | WIP15SubCom | IQon1 |  | WIP15 | April 12, 2019 8:38:04 PM GMT+08:00
CTUDT00622857 | Refresh | Closed | sowmya_k | April 2, 2019 8:48:51 PM GMT+08:00 | Out of range thickness read from a user CCT volume exam card treated as valid | 4.1.6 | BB_WIP7 |  | Subsystem | ECM | NA | Fixed | sowmya_k | 0 | 2 | Console Software | Console4 | YES | Closed | BB_WIP6_SubCom | Brilliance Big Bore | PR6113361 | BB_WIP7 | May 14, 2019 8:09:43 PM GMT+08:00
CTUDT00622855 | Refresh | Closed | sunilkumar_j | April 2, 2019 8:14:03 PM GMT+08:00 | [Refresh]: Date format displayed next to Date of Birth field was not localized in Spanish language | WIP14 | WIP16 | NA | Subsystem | Localization | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP16SubCom | Ovation |  | WIP16 | April 8, 2019 8:59:14 PM GMT+08:00
CTUDT00622855 | Refresh | Closed | sunilkumar_j | April 2, 2019 8:14:03 PM GMT+08:00 | [Refresh]: Date format displayed next to Date of Birth field was not localized in Spanish language | WIP14 | WIP16 | NA | Subsystem | Localization | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP16SubCom | Brilliance 64 |  | WIP16 | April 8, 2019 8:59:14 PM GMT+08:00
CTUDT00622855 | Refresh | Closed | sunilkumar_j | April 2, 2019 8:14:03 PM GMT+08:00 | [Refresh]: Date format displayed next to Date of Birth field was not localized in Spanish language | WIP14 | WIP16 | NA | Subsystem | Localization | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP16SubCom | Ingenuity 64 |  | WIP16 | April 8, 2019 8:59:14 PM GMT+08:00
CTUDT00622855 | Refresh | Closed | sunilkumar_j | April 2, 2019 8:14:03 PM GMT+08:00 | [Refresh]: Date format displayed next to Date of Birth field was not localized in Spanish language | WIP14 | WIP16 | NA | Subsystem | Localization | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP16SubCom | IQon1 |  | WIP16 | April 8, 2019 8:59:14 PM GMT+08:00
CTUDT00622852 | Refresh | Closed | g_di | April 2, 2019 7:36:32 PM GMT+08:00 | On extending an oblique result, extended result is not covering the plan box associated to the original result. | WIP14 |  |  | Subsystem | Host | NA | Works As Specified / Not A Defect | babu_s | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00622850 | Refresh | Closed | wang_l | April 2, 2019 6:45:33 PM GMT+08:00 | Refresh: the HU value of conventional images reconstructed from the generated partial raw data did not match with value of conventional images reconstructed from the full raw data | WIP14 |  |  | Subsystem | CT Viewer |  | Works As Specified / Not A Defect | richards_b | 0 |  |  | CIRS | NO | Closed |  | Ingenuity 64
CTUDT00622847 | Refresh | Closed | kr_m | April 2, 2019 3:52:08 PM GMT+08:00 | Refresh: console user messages localization mismatch | WIP13 | WIP17 | NA | Subsystem | Host | Console-SRS.11210 | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP16SubCom | Ovation |  | WIP17 | May 17, 2019 8:33:52 PM GMT+08:00
CTUDT00622847 | Refresh | Closed | kr_m | April 2, 2019 3:52:08 PM GMT+08:00 | Refresh: console user messages localization mismatch | WIP13 | WIP17 | NA | Subsystem | Host | Console-SRS.11210 | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP16SubCom | Brilliance 64 |  | WIP17 | May 17, 2019 8:33:52 PM GMT+08:00
CTUDT00622847 | Refresh | Closed | kr_m | April 2, 2019 3:52:08 PM GMT+08:00 | Refresh: console user messages localization mismatch | WIP13 | WIP17 | NA | Subsystem | Host | Console-SRS.11210 | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP16SubCom | Ingenuity 64 |  | WIP17 | May 17, 2019 8:33:52 PM GMT+08:00
CTUDT00622847 | Refresh | Closed | kr_m | April 2, 2019 3:52:08 PM GMT+08:00 | Refresh: console user messages localization mismatch | WIP13 | WIP17 | NA | Subsystem | Host | Console-SRS.11210 | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP16SubCom | IQon1 |  | WIP17 | May 17, 2019 8:33:52 PM GMT+08:00
CTUDT00622846 | Refresh | Closed | bhawna_g | April 2, 2019 3:05:57 PM GMT+08:00 | Refresh Pre-Val: BugRep and system error observed after end exam | V4.8 |  |  | Subsystem | EA Crash | N/A | Works As Specified / Not A Defect | blessily | 1 | 4 |  | Console4 | NO | Closed | BB_WIP6_SubCom | Brilliance Big Bore
CTUDT00622835 | Refresh | Closed | wei_f | April 2, 2019 10:32:02 AM GMT+08:00 | philips service user Re-Install Default Users failed | BB_WIP5 | BB_WIP6 WIP16 | NA. | Subsystem | Service Tools | NA | Fixed | kapadia_v | 0 | 4 | Couch Software | Console4 | NO | Closed | WIP16SubCom | Ovation |  | BB_WIP6 WIP16 | April 17, 2019 1:07:37 PM GMT+08:00
CTUDT00622835 | Refresh | Closed | wei_f | April 2, 2019 10:32:02 AM GMT+08:00 | philips service user Re-Install Default Users failed | BB_WIP5 | BB_WIP6 WIP16 | NA. | Subsystem | Service Tools | NA | Fixed | kapadia_v | 0 | 4 | Couch Software | Console4 | NO | Closed | WIP16SubCom | Brilliance 64 |  | BB_WIP6 WIP16 | April 17, 2019 1:07:37 PM GMT+08:00
CTUDT00622835 | Refresh | Closed | wei_f | April 2, 2019 10:32:02 AM GMT+08:00 | philips service user Re-Install Default Users failed | BB_WIP5 | BB_WIP6 WIP16 | NA. | Subsystem | Service Tools | NA | Fixed | kapadia_v | 0 | 4 | Couch Software | Console4 | NO | Closed | WIP16SubCom | Brilliance Big Bore |  | BB_WIP6 WIP16 | April 17, 2019 1:07:37 PM GMT+08:00
CTUDT00622835 | Refresh | Closed | wei_f | April 2, 2019 10:32:02 AM GMT+08:00 | philips service user Re-Install Default Users failed | BB_WIP5 | BB_WIP6 WIP16 | NA. | Subsystem | Service Tools | NA | Fixed | kapadia_v | 0 | 4 | Couch Software | Console4 | NO | Closed | WIP16SubCom | Ingenuity 64 |  | BB_WIP6 WIP16 | April 17, 2019 1:07:37 PM GMT+08:00
CTUDT00622835 | Refresh | Closed | wei_f | April 2, 2019 10:32:02 AM GMT+08:00 | philips service user Re-Install Default Users failed | BB_WIP5 | BB_WIP6 WIP16 | NA. | Subsystem | Service Tools | NA | Fixed | kapadia_v | 0 | 4 | Couch Software | Console4 | NO | Closed | WIP16SubCom | IQon1 |  | BB_WIP6 WIP16 | April 17, 2019 1:07:37 PM GMT+08:00
CTUDT00622833 | Refresh | Closed | lei_yang | April 2, 2019 10:16:50 AM GMT+08:00 | For CCT scan(CCT Single/CCT Continuous/CCT Fluoro/Biosy), the increment sometimes display, sometimes doesn't display after choose Coverage value | WIP14 |  |  | Subsystem | Planning | N/A | Works As Specified / Not A Defect | lei_yang | 0 | 2 |  | Support-System | NO | Closed | WIP16SubCom | Ingenuity 64
CTUDT00622831 | Refresh | Closed | lei_yang | April 2, 2019 9:56:53 AM GMT+08:00 | The "O-MAR" annotation is overlap with specific layout. | WIP14 |  |  | Subsystem | Host | NA | No Plan to Fix | aswani_v | 0 | 4 |  | iCAP | NO | Closed |  | Ovation
CTUDT00622831 | Refresh | Closed | lei_yang | April 2, 2019 9:56:53 AM GMT+08:00 | The "O-MAR" annotation is overlap with specific layout. | WIP14 |  |  | Subsystem | Host | NA | No Plan to Fix | aswani_v | 0 | 4 |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00622831 | Refresh | Closed | lei_yang | April 2, 2019 9:56:53 AM GMT+08:00 | The "O-MAR" annotation is overlap with specific layout. | WIP14 |  |  | Subsystem | Host | NA | No Plan to Fix | aswani_v | 0 | 4 |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00622831 | Refresh | Closed | lei_yang | April 2, 2019 9:56:53 AM GMT+08:00 | The "O-MAR" annotation is overlap with specific layout. | WIP14 |  |  | Subsystem | Host | NA | No Plan to Fix | aswani_v | 0 | 4 |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00622831 | Refresh | Closed | lei_yang | April 2, 2019 9:56:53 AM GMT+08:00 | The "O-MAR" annotation is overlap with specific layout. | WIP14 |  |  | Subsystem | Host | NA | No Plan to Fix | aswani_v | 0 | 4 |  | iCAP | NO | Closed |  | IQon1
CTUDT00622829 | Refresh | Closed | lei_yang | April 2, 2019 9:34:06 AM GMT+08:00 | There is no "Dental" scan type in the drop down list for Axial EC. | WIP14 | WIP19 BB_WIP9 | NA | Subsystem | Host | CT-NM.SRS.Modes.1211 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | SegReq | Ovation |  | WIP19 BB_WIP9 | July 25, 2019 4:46:08 PM GMT+08:00
CTUDT00622829 | Refresh | Closed | lei_yang | April 2, 2019 9:34:06 AM GMT+08:00 | There is no "Dental" scan type in the drop down list for Axial EC. | WIP14 | WIP19 BB_WIP9 | NA | Subsystem | Host | CT-NM.SRS.Modes.1211 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | SegReq | Brilliance 64 |  | WIP19 BB_WIP9 | July 25, 2019 4:46:08 PM GMT+08:00
CTUDT00622829 | Refresh | Closed | lei_yang | April 2, 2019 9:34:06 AM GMT+08:00 | There is no "Dental" scan type in the drop down list for Axial EC. | WIP14 | WIP19 BB_WIP9 | NA | Subsystem | Host | CT-NM.SRS.Modes.1211 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | SegReq | Brilliance Big Bore |  | WIP19 BB_WIP9 | July 25, 2019 4:46:08 PM GMT+08:00
CTUDT00622829 | Refresh | Closed | lei_yang | April 2, 2019 9:34:06 AM GMT+08:00 | There is no "Dental" scan type in the drop down list for Axial EC. | WIP14 | WIP19 BB_WIP9 | NA | Subsystem | Host | CT-NM.SRS.Modes.1211 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | SegReq | Ingenuity 64 |  | WIP19 BB_WIP9 | July 25, 2019 4:46:08 PM GMT+08:00
CTUDT00622829 | Refresh | Closed | lei_yang | April 2, 2019 9:34:06 AM GMT+08:00 | There is no "Dental" scan type in the drop down list for Axial EC. | WIP14 | WIP19 BB_WIP9 | NA | Subsystem | Host | CT-NM.SRS.Modes.1211 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | SegReq | IQon1 |  | WIP19 BB_WIP9 | July 25, 2019 4:46:08 PM GMT+08:00
CTUDT00622827 | Refresh | Closed | lei_yang | April 2, 2019 9:13:11 AM GMT+08:00 | The message "Maximal allowed number of cycles for spiral perfusion is 15" still showing if deleting some of group to decrease the cycle number. | WIP14 | WIP16 |  | Subsystem | Host | NA | Fixed | datta_i | 0 | 2 | Console Software | Console4 | NO | Closed | WIP16SegReq | Ovation |  | WIP16 | April 10, 2019 7:35:06 PM GMT+08:00
CTUDT00622827 | Refresh | Closed | lei_yang | April 2, 2019 9:13:11 AM GMT+08:00 | The message "Maximal allowed number of cycles for spiral perfusion is 15" still showing if deleting some of group to decrease the cycle number. | WIP14 | WIP16 |  | Subsystem | Host | NA | Fixed | datta_i | 0 | 2 | Console Software | Console4 | NO | Closed | WIP16SegReq | Brilliance 64 |  | WIP16 | April 10, 2019 7:35:06 PM GMT+08:00
CTUDT00622827 | Refresh | Closed | lei_yang | April 2, 2019 9:13:11 AM GMT+08:00 | The message "Maximal allowed number of cycles for spiral perfusion is 15" still showing if deleting some of group to decrease the cycle number. | WIP14 | WIP16 |  | Subsystem | Host | NA | Fixed | datta_i | 0 | 2 | Console Software | Console4 | NO | Closed | WIP16SegReq | Ingenuity 64 |  | WIP16 | April 10, 2019 7:35:06 PM GMT+08:00
CTUDT00622827 | Refresh | Closed | lei_yang | April 2, 2019 9:13:11 AM GMT+08:00 | The message "Maximal allowed number of cycles for spiral perfusion is 15" still showing if deleting some of group to decrease the cycle number. | WIP14 | WIP16 |  | Subsystem | Host | NA | Fixed | datta_i | 0 | 2 | Console Software | Console4 | NO | Closed | WIP16SegReq | IQon1 |  | WIP16 | April 10, 2019 7:35:06 PM GMT+08:00
CTUDT00622823 | Refresh | Closed | nimma_p | April 1, 2019 8:05:19 PM GMT+08:00 | Refresh:-“Visual Diagnostics Tools” are not similar when we compare with the config files. | WIP14 |  |  | Subsystem | Visual Diagnostics |  | Duplicate | shreekanth_manniyath | 0 |  |  | Service Tools | NO | Closed |  | Ingenuity 64
CTUDT00622818 | Refresh | Closed | lei_yang | April 1, 2019 4:33:23 PM GMT+08:00 | CCT Fluoro and Biopsy display field are not grey out although it cannot be editable. | WIP14 |  |  | Subsystem | Planning | NA | No Plan to Fix | wilsonraja | 0 | 4 |  | Console4 | NO | Closed |  | Ovation
CTUDT00622818 | Refresh | Closed | lei_yang | April 1, 2019 4:33:23 PM GMT+08:00 | CCT Fluoro and Biopsy display field are not grey out although it cannot be editable. | WIP14 |  |  | Subsystem | Planning | NA | No Plan to Fix | wilsonraja | 0 | 4 |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00622818 | Refresh | Closed | lei_yang | April 1, 2019 4:33:23 PM GMT+08:00 | CCT Fluoro and Biopsy display field are not grey out although it cannot be editable. | WIP14 |  |  | Subsystem | Planning | NA | No Plan to Fix | wilsonraja | 0 | 4 |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00622818 | Refresh | Closed | lei_yang | April 1, 2019 4:33:23 PM GMT+08:00 | CCT Fluoro and Biopsy display field are not grey out although it cannot be editable. | WIP14 |  |  | Subsystem | Planning | NA | No Plan to Fix | wilsonraja | 0 | 4 |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00622818 | Refresh | Closed | lei_yang | April 1, 2019 4:33:23 PM GMT+08:00 | CCT Fluoro and Biopsy display field are not grey out although it cannot be editable. | WIP14 |  |  | Subsystem | Planning | NA | No Plan to Fix | wilsonraja | 0 | 4 |  | Console4 | NO | Closed |  | IQon1
CTUDT00622817 | Refresh | Closed | wei_f | April 1, 2019 4:30:58 PM GMT+08:00 | Using the "Record" option to burn the ePHI data failed | BB_WIP5 | BB_WIP8 |  | Subsystem | CD-DVD operation | NA | Works As Specified / Not A Defect | wei_f | 0 |  | Verification Documentation | Support-System | NO | Closed |  | Brilliance Big Bore |  | BB_WIP8
CTUDT00622814 | Refresh | Closed | wei_f | April 1, 2019 3:57:21 PM GMT+08:00 | The system cannot identify the CD/DVD when insert a CD/DVD to the device. | BB_WIP5 | BB_WIP6 WIP16 | NA | Subsystem | CD-DVD operation | NA | Fixed | kapadia_v | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP16SubCom | Ovation |  | BB_WIP6 WIP16 | April 15, 2019 6:00:56 PM GMT+08:00
CTUDT00622814 | Refresh | Closed | wei_f | April 1, 2019 3:57:21 PM GMT+08:00 | The system cannot identify the CD/DVD when insert a CD/DVD to the device. | BB_WIP5 | BB_WIP6 WIP16 | NA | Subsystem | CD-DVD operation | NA | Fixed | kapadia_v | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP16SubCom | Brilliance 64 |  | BB_WIP6 WIP16 | April 15, 2019 6:00:56 PM GMT+08:00
CTUDT00622814 | Refresh | Closed | wei_f | April 1, 2019 3:57:21 PM GMT+08:00 | The system cannot identify the CD/DVD when insert a CD/DVD to the device. | BB_WIP5 | BB_WIP6 WIP16 | NA | Subsystem | CD-DVD operation | NA | Fixed | kapadia_v | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP16SubCom | Brilliance Big Bore |  | BB_WIP6 WIP16 | April 15, 2019 6:00:56 PM GMT+08:00
CTUDT00622814 | Refresh | Closed | wei_f | April 1, 2019 3:57:21 PM GMT+08:00 | The system cannot identify the CD/DVD when insert a CD/DVD to the device. | BB_WIP5 | BB_WIP6 WIP16 | NA | Subsystem | CD-DVD operation | NA | Fixed | kapadia_v | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP16SubCom | Ingenuity 64 |  | BB_WIP6 WIP16 | April 15, 2019 6:00:56 PM GMT+08:00
CTUDT00622814 | Refresh | Closed | wei_f | April 1, 2019 3:57:21 PM GMT+08:00 | The system cannot identify the CD/DVD when insert a CD/DVD to the device. | BB_WIP5 | BB_WIP6 WIP16 | NA | Subsystem | CD-DVD operation | NA | Fixed | kapadia_v | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP16SubCom | IQon1 |  | BB_WIP6 WIP16 | April 15, 2019 6:00:56 PM GMT+08:00
CTUDT00622813 | Refresh | Closed | 320053769 | April 1, 2019 3:54:51 PM GMT+08:00 | Refresh- search button of patient ID is disabled | WIP14 |  |  | Subsystem | Host | NA | Invalid/Uncontrolled Configuration | veerabhadraiah | 0 |  |  | iCAP | NO | Closed |  | Ovation
CTUDT00622813 | Refresh | Closed | 320053769 | April 1, 2019 3:54:51 PM GMT+08:00 | Refresh- search button of patient ID is disabled | WIP14 |  |  | Subsystem | Host | NA | Invalid/Uncontrolled Configuration | veerabhadraiah | 0 |  |  | iCAP | NO | Closed |  | IQon1
CTUDT00622811 | Refresh | Closed | wei_f | April 1, 2019 3:20:23 PM GMT+08:00 | Attempt to execute an unsigned executable file,Windows did not inform the user of insufficient permissions to run the executable. | BB_WIP5 |  |  | Subsystem | Service Tools |  | Works As Specified / Not A Defect | wei_f | 1 |  |  | Support-System | NO | Closed |  | Ingenuity 64
CTUDT00622811 | Refresh | Closed | wei_f | April 1, 2019 3:20:23 PM GMT+08:00 | Attempt to execute an unsigned executable file,Windows did not inform the user of insufficient permissions to run the executable. | BB_WIP5 |  |  | Subsystem | Service Tools |  | Works As Specified / Not A Defect | wei_f | 1 |  |  | Support-System | NO | Closed |  | Ingenuity BigBore
CTUDT00622807 | Refresh | Closed | chakraborty_d | April 1, 2019 1:18:03 PM GMT+08:00 | User is unable to view the complete result for Datapath Diagnostics. | WIP14 |  |  | Subsystem | Service Tools |  | Duplicate | shreekanth_manniyath | 0 |  |  | Service Tools | NO | Closed |  | Ovation
CTUDT00622805 | Refresh | Closed | shekhar_p | April 1, 2019 12:19:45 PM GMT+08:00 | Refresh:Wrong image numbers getting generated in the exam summary page . | WIP14 |  |  | Subsystem | Host | NA | No Plan to Fix | redrouthu_r | 0 | 4 |  | iCAP | NO | Closed |  | Ovation
CTUDT00622805 | Refresh | Closed | shekhar_p | April 1, 2019 12:19:45 PM GMT+08:00 | Refresh:Wrong image numbers getting generated in the exam summary page . | WIP14 |  |  | Subsystem | Host | NA | No Plan to Fix | redrouthu_r | 0 | 4 |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00622805 | Refresh | Closed | shekhar_p | April 1, 2019 12:19:45 PM GMT+08:00 | Refresh:Wrong image numbers getting generated in the exam summary page . | WIP14 |  |  | Subsystem | Host | NA | No Plan to Fix | redrouthu_r | 0 | 4 |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00622805 | Refresh | Closed | shekhar_p | April 1, 2019 12:19:45 PM GMT+08:00 | Refresh:Wrong image numbers getting generated in the exam summary page . | WIP14 |  |  | Subsystem | Host | NA | No Plan to Fix | redrouthu_r | 0 | 4 |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00622805 | Refresh | Closed | shekhar_p | April 1, 2019 12:19:45 PM GMT+08:00 | Refresh:Wrong image numbers getting generated in the exam summary page . | WIP14 |  |  | Subsystem | Host | NA | No Plan to Fix | redrouthu_r | 0 | 4 |  | iCAP | NO | Closed |  | IQon1
CTUDT00622804 | Refresh | Closed | sowmya_k | April 1, 2019 12:17:26 PM GMT+08:00 | DRI Field accepting integers other than 1 to 55 | WIP5 | BB_WIP6 |  | Subsystem | ECM | NA | Fixed | datta_i | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP6_SubCom | Brilliance Big Bore |  | BB_WIP6 | April 19, 2019 2:11:38 PM GMT+08:00
CTUDT00622802 | Refresh | Closed | Noy_Seelig | March 31, 2019 10:06:02 PM GMT+08:00 | Refresh: Performance- wrong ROI ED value | WIP14 | WIP17 | N/A | Subsystem | DMS Calibrations | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP17SegReq | Ovation |  | WIP17 | May 12, 2019 6:22:56 PM GMT+08:00
CTUDT00622802 | Refresh | Closed | Noy_Seelig | March 31, 2019 10:06:02 PM GMT+08:00 | Refresh: Performance- wrong ROI ED value | WIP14 | WIP17 | N/A | Subsystem | DMS Calibrations | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP17SegReq | Ingenuity 64 |  | WIP17 | May 12, 2019 6:22:56 PM GMT+08:00
CTUDT00622800 | Refresh | Closed | beck_d | March 31, 2019 6:59:02 PM GMT+08:00 | Refresh: Calibrations: Calibration Main Menu: Beam and DMS Checks appears under System tests instead of Subsystem Tests | WIP14 | WIP17 |  | Subsystem | Calibrations; Verified1 | NA | Fixed | Kostya_Baklan | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP15SubCom | Ovation |  | WIP17 | April 11, 2019 3:42:32 PM GMT+08:00
CTUDT00622800 | Refresh | Closed | beck_d | March 31, 2019 6:59:02 PM GMT+08:00 | Refresh: Calibrations: Calibration Main Menu: Beam and DMS Checks appears under System tests instead of Subsystem Tests | WIP14 | WIP17 |  | Subsystem | Calibrations; Verified1 | NA | Fixed | Kostya_Baklan | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP15SubCom | Brilliance Big Bore |  | WIP17 | April 11, 2019 3:42:32 PM GMT+08:00
CTUDT00622800 | Refresh | Closed | beck_d | March 31, 2019 6:59:02 PM GMT+08:00 | Refresh: Calibrations: Calibration Main Menu: Beam and DMS Checks appears under System tests instead of Subsystem Tests | WIP14 | WIP17 |  | Subsystem | Calibrations; Verified1 | NA | Fixed | Kostya_Baklan | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP15SubCom | Ingenuity 64 |  | WIP17 | April 11, 2019 3:42:32 PM GMT+08:00
CTUDT00622800 | Refresh | Closed | beck_d | March 31, 2019 6:59:02 PM GMT+08:00 | Refresh: Calibrations: Calibration Main Menu: Beam and DMS Checks appears under System tests instead of Subsystem Tests | WIP14 | WIP17 |  | Subsystem | Calibrations; Verified1 | NA | Fixed | Kostya_Baklan | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP15SubCom | IQon1 |  | WIP17 | April 11, 2019 3:42:32 PM GMT+08:00
CTUDT00622798 | Refresh | Closed | attali_s | March 31, 2019 6:23:07 PM GMT+08:00 | Refresh: HTML Report Header: Final Status field is not according to the spec | WIP14 |  |  | Subsystem |  |  | Duplicate | frider_o | 0 |  |  |  |  | Closed |  | Ingenuity 64
CTUDT00622795 | Refresh | Closed | alex_nihaev | March 31, 2019 5:37:47 PM GMT+08:00 | Refresh: Beam Quality, Tube Potential Test, HVL value is not measured/shown for 100 KV,120KV,140KV in RND user , when run was performed in MFG user | BB_WIP4 | N/A |  | Subsystem | DMS | NaN | Fixed | braunstein_d | 0 | 2 | Design Documentation | Support-Physics | NO | Closed | WIP16SegReq | Ovation |  | N/A | May 27, 2019 7:22:37 PM GMT+08:00
CTUDT00622795 | Refresh | Closed | alex_nihaev | March 31, 2019 5:37:47 PM GMT+08:00 | Refresh: Beam Quality, Tube Potential Test, HVL value is not measured/shown for 100 KV,120KV,140KV in RND user , when run was performed in MFG user | BB_WIP4 | N/A |  | Subsystem | DMS | NaN | Fixed | braunstein_d | 0 | 2 | Design Documentation | Support-Physics | NO | Closed | WIP16SegReq | Brilliance Big Bore |  | N/A | May 27, 2019 7:22:37 PM GMT+08:00
CTUDT00622795 | Refresh | Closed | alex_nihaev | March 31, 2019 5:37:47 PM GMT+08:00 | Refresh: Beam Quality, Tube Potential Test, HVL value is not measured/shown for 100 KV,120KV,140KV in RND user , when run was performed in MFG user | BB_WIP4 | N/A |  | Subsystem | DMS | NaN | Fixed | braunstein_d | 0 | 2 | Design Documentation | Support-Physics | NO | Closed | WIP16SegReq | Ingenuity 64 |  | N/A | May 27, 2019 7:22:37 PM GMT+08:00
CTUDT00622795 | Refresh | Closed | alex_nihaev | March 31, 2019 5:37:47 PM GMT+08:00 | Refresh: Beam Quality, Tube Potential Test, HVL value is not measured/shown for 100 KV,120KV,140KV in RND user , when run was performed in MFG user | BB_WIP4 | N/A |  | Subsystem | DMS | NaN | Fixed | braunstein_d | 0 | 2 | Design Documentation | Support-Physics | NO | Closed | WIP16SegReq | IQon1 |  | N/A | May 27, 2019 7:22:37 PM GMT+08:00
CTUDT00622794 | Refresh | Closed | beck_d | March 31, 2019 5:34:51 PM GMT+08:00 | Refresh: Calibrations: Acceptance and Constancy Test: Noise ROI is missing in some slices in sequence images | WIP14 |  |  | Subsystem | Calibrations | n/a | Duplicate | redrouthu_r | 0 |  |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00622792 | Refresh | Closed | dunitza_m | March 31, 2019 3:52:44 PM GMT+08:00 | High pitches in Extremity and Knee scan types are allowed and generates artifacts in the images | WIP15 |  |  | Subsystem | ECM | N/A | Change Proposal | dunitza_m | 0 |  |  | Support-System | NO | Closed | SubNC | Ovation
CTUDT00622792 | Refresh | Closed | dunitza_m | March 31, 2019 3:52:44 PM GMT+08:00 | High pitches in Extremity and Knee scan types are allowed and generates artifacts in the images | WIP15 |  |  | Subsystem | ECM | N/A | Change Proposal | dunitza_m | 0 |  |  | Support-System | NO | Closed | SubNC | Brilliance 64
CTUDT00622792 | Refresh | Closed | dunitza_m | March 31, 2019 3:52:44 PM GMT+08:00 | High pitches in Extremity and Knee scan types are allowed and generates artifacts in the images | WIP15 |  |  | Subsystem | ECM | N/A | Change Proposal | dunitza_m | 0 |  |  | Support-System | NO | Closed | SubNC | Brilliance Big Bore
CTUDT00622792 | Refresh | Closed | dunitza_m | March 31, 2019 3:52:44 PM GMT+08:00 | High pitches in Extremity and Knee scan types are allowed and generates artifacts in the images | WIP15 |  |  | Subsystem | ECM | N/A | Change Proposal | dunitza_m | 0 |  |  | Support-System | NO | Closed | SubNC | Ingenuity 64
CTUDT00622792 | Refresh | Closed | dunitza_m | March 31, 2019 3:52:44 PM GMT+08:00 | High pitches in Extremity and Knee scan types are allowed and generates artifacts in the images | WIP15 |  |  | Subsystem | ECM | N/A | Change Proposal | dunitza_m | 0 |  |  | Support-System | NO | Closed | SubNC | Ingenuity Flex
CTUDT00622792 | Refresh | Closed | dunitza_m | March 31, 2019 3:52:44 PM GMT+08:00 | High pitches in Extremity and Knee scan types are allowed and generates artifacts in the images | WIP15 |  |  | Subsystem | ECM | N/A | Change Proposal | dunitza_m | 0 |  |  | Support-System | NO | Closed | SubNC | IQon1
CTUDT00622789 | Refresh | Closed | chakraborty_d | March 29, 2019 6:42:31 PM GMT+08:00 | Console Status not displayed in System Status tool. | WIP14 | WIP19 |  | System | Service Tools | NA | Fixed | s_shabeer | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP16SubCom | Ovation |  | WIP19 | July 17, 2019 4:29:47 PM GMT+08:00
CTUDT00622789 | Refresh | Closed | chakraborty_d | March 29, 2019 6:42:31 PM GMT+08:00 | Console Status not displayed in System Status tool. | WIP14 | WIP19 |  | System | Service Tools | NA | Fixed | s_shabeer | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP16SubCom | IQon1 |  | WIP19 | July 17, 2019 4:29:47 PM GMT+08:00
CTUDT00622787 | Refresh | Closed | chakraborty_d | March 29, 2019 6:37:27 PM GMT+08:00 | Error thrown on applying selection in Universal Log Viewer. | WIP14 |  |  | Subsystem | Log Viewer | CT-NM.SRS.Service 2515 | Works As Specified / Not A Defect | shreekanth_manniyath | 0 |  |  | Service Tools | NO | Closed | WIP17SubCom | Ovation
CTUDT00622787 | Refresh | Closed | chakraborty_d | March 29, 2019 6:37:27 PM GMT+08:00 | Error thrown on applying selection in Universal Log Viewer. | WIP14 |  |  | Subsystem | Log Viewer | CT-NM.SRS.Service 2515 | Works As Specified / Not A Defect | shreekanth_manniyath | 0 |  |  | Service Tools | NO | Closed | WIP17SubCom | Brilliance 64
CTUDT00622787 | Refresh | Closed | chakraborty_d | March 29, 2019 6:37:27 PM GMT+08:00 | Error thrown on applying selection in Universal Log Viewer. | WIP14 |  |  | Subsystem | Log Viewer | CT-NM.SRS.Service 2515 | Works As Specified / Not A Defect | shreekanth_manniyath | 0 |  |  | Service Tools | NO | Closed | WIP17SubCom | Brilliance Big Bore
CTUDT00622787 | Refresh | Closed | chakraborty_d | March 29, 2019 6:37:27 PM GMT+08:00 | Error thrown on applying selection in Universal Log Viewer. | WIP14 |  |  | Subsystem | Log Viewer | CT-NM.SRS.Service 2515 | Works As Specified / Not A Defect | shreekanth_manniyath | 0 |  |  | Service Tools | NO | Closed | WIP17SubCom | Ingenuity 64
CTUDT00622787 | Refresh | Closed | chakraborty_d | March 29, 2019 6:37:27 PM GMT+08:00 | Error thrown on applying selection in Universal Log Viewer. | WIP14 |  |  | Subsystem | Log Viewer | CT-NM.SRS.Service 2515 | Works As Specified / Not A Defect | shreekanth_manniyath | 0 |  |  | Service Tools | NO | Closed | WIP17SubCom | IQon1
CTUDT00622785 | Refresh | Closed | chakraborty_d | March 29, 2019 6:24:15 PM GMT+08:00 | Settings tool shows registry key missing value | WIP14 | WIP18 |  | Subsystem | Service Tools | NA | Fixed | ranjithkumar_m | 0 | 3 | Service Tools_Software | Service Tools | NO | Closed | WIP17SubCom | Ovation |  | WIP18 | June 7, 2019 6:47:27 PM GMT+08:00
CTUDT00622785 | Refresh | Closed | chakraborty_d | March 29, 2019 6:24:15 PM GMT+08:00 | Settings tool shows registry key missing value | WIP14 | WIP18 |  | Subsystem | Service Tools | NA | Fixed | ranjithkumar_m | 0 | 3 | Service Tools_Software | Service Tools | NO | Closed | WIP17SubCom | Ingenuity 64 |  | WIP18 | June 7, 2019 6:47:27 PM GMT+08:00
CTUDT00622785 | Refresh | Closed | chakraborty_d | March 29, 2019 6:24:15 PM GMT+08:00 | Settings tool shows registry key missing value | WIP14 | WIP18 |  | Subsystem | Service Tools | NA | Fixed | ranjithkumar_m | 0 | 3 | Service Tools_Software | Service Tools | NO | Closed | WIP17SubCom | IQon1 |  | WIP18 | June 7, 2019 6:47:27 PM GMT+08:00
CTUDT00622782 | Refresh | Closed | chakraborty_d | March 29, 2019 6:02:41 PM GMT+08:00 | Script error is thrown in Universal Log Viewer | WIP14 | WIP17 BB_WIP7 |  | Subsystem | Log Viewer | NA | Fixed | shreekanth_manniyath | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP17SubCom | Ovation |  | WIP17 BB_WIP7 | May 6, 2019 4:47:35 PM GMT+08:00
CTUDT00622782 | Refresh | Closed | chakraborty_d | March 29, 2019 6:02:41 PM GMT+08:00 | Script error is thrown in Universal Log Viewer | WIP14 | WIP17 BB_WIP7 |  | Subsystem | Log Viewer | NA | Fixed | shreekanth_manniyath | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP17SubCom | Brilliance 64 |  | WIP17 BB_WIP7 | May 6, 2019 4:47:35 PM GMT+08:00
CTUDT00622782 | Refresh | Closed | chakraborty_d | March 29, 2019 6:02:41 PM GMT+08:00 | Script error is thrown in Universal Log Viewer | WIP14 | WIP17 BB_WIP7 |  | Subsystem | Log Viewer | NA | Fixed | shreekanth_manniyath | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP17SubCom | Brilliance Big Bore |  | WIP17 BB_WIP7 | May 6, 2019 4:47:35 PM GMT+08:00
CTUDT00622782 | Refresh | Closed | chakraborty_d | March 29, 2019 6:02:41 PM GMT+08:00 | Script error is thrown in Universal Log Viewer | WIP14 | WIP17 BB_WIP7 |  | Subsystem | Log Viewer | NA | Fixed | shreekanth_manniyath | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP17SubCom | Ingenuity 64 |  | WIP17 BB_WIP7 | May 6, 2019 4:47:35 PM GMT+08:00
CTUDT00622782 | Refresh | Closed | chakraborty_d | March 29, 2019 6:02:41 PM GMT+08:00 | Script error is thrown in Universal Log Viewer | WIP14 | WIP17 BB_WIP7 |  | Subsystem | Log Viewer | NA | Fixed | shreekanth_manniyath | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP17SubCom | IQon1 |  | WIP17 BB_WIP7 | May 6, 2019 4:47:35 PM GMT+08:00
CTUDT00622780 | Refresh | Closed | chakraborty_d | March 29, 2019 5:49:12 PM GMT+08:00 | Error is thrown on launching Remote Viewer tool. | WIP14 |  |  | Subsystem | Remote Configuration |  | Does Not Apply | prabhu_m | 0 |  |  | Service Tools | NO | Closed |  | Ingenuity 64
CTUDT00622778 | Refresh | Closed | chakraborty_d | March 29, 2019 5:44:19 PM GMT+08:00 | DateTime field appears as DataTime in Universal Log Viewer. | WIP14 |  |  | Subsystem | Log Viewer | NA | No Plan to Fix | chakraborty_d | 0 | 3 |  | Service Tools | NO | Closed | SubNC | Ovation
CTUDT00622778 | Refresh | Closed | chakraborty_d | March 29, 2019 5:44:19 PM GMT+08:00 | DateTime field appears as DataTime in Universal Log Viewer. | WIP14 |  |  | Subsystem | Log Viewer | NA | No Plan to Fix | chakraborty_d | 0 | 3 |  | Service Tools | NO | Closed | SubNC | Brilliance 64
CTUDT00622778 | Refresh | Closed | chakraborty_d | March 29, 2019 5:44:19 PM GMT+08:00 | DateTime field appears as DataTime in Universal Log Viewer. | WIP14 |  |  | Subsystem | Log Viewer | NA | No Plan to Fix | chakraborty_d | 0 | 3 |  | Service Tools | NO | Closed | SubNC | Brilliance Big Bore
CTUDT00622778 | Refresh | Closed | chakraborty_d | March 29, 2019 5:44:19 PM GMT+08:00 | DateTime field appears as DataTime in Universal Log Viewer. | WIP14 |  |  | Subsystem | Log Viewer | NA | No Plan to Fix | chakraborty_d | 0 | 3 |  | Service Tools | NO | Closed | SubNC | Ingenuity 64
CTUDT00622778 | Refresh | Closed | chakraborty_d | March 29, 2019 5:44:19 PM GMT+08:00 | DateTime field appears as DataTime in Universal Log Viewer. | WIP14 |  |  | Subsystem | Log Viewer | NA | No Plan to Fix | chakraborty_d | 0 | 3 |  | Service Tools | NO | Closed | SubNC | Veyron
CTUDT00622778 | Refresh | Closed | chakraborty_d | March 29, 2019 5:44:19 PM GMT+08:00 | DateTime field appears as DataTime in Universal Log Viewer. | WIP14 |  |  | Subsystem | Log Viewer | NA | No Plan to Fix | chakraborty_d | 0 | 3 |  | Service Tools | NO | Closed | SubNC | IQon1
CTUDT00622775 | Refresh | Closed | chakraborty_d | March 29, 2019 5:17:11 PM GMT+08:00 | Windows Command Processor stopped working in Copy A-Plane Calibration File tool, on selecting non-existent drive. | WIP14 |  | Re-launch the tool and select the desired drive letter. | Subsystem | Service Tools | NA | No Plan to Fix | chakraborty_d | 0 | 4 |  | Service Tools | NO | Closed | SubNC | Brilliance 64
CTUDT00622775 | Refresh | Closed | chakraborty_d | March 29, 2019 5:17:11 PM GMT+08:00 | Windows Command Processor stopped working in Copy A-Plane Calibration File tool, on selecting non-existent drive. | WIP14 |  | Re-launch the tool and select the desired drive letter. | Subsystem | Service Tools | NA | No Plan to Fix | chakraborty_d | 0 | 4 |  | Service Tools | NO | Closed | SubNC | Brilliance Big Bore
CTUDT00622775 | Refresh | Closed | chakraborty_d | March 29, 2019 5:17:11 PM GMT+08:00 | Windows Command Processor stopped working in Copy A-Plane Calibration File tool, on selecting non-existent drive. | WIP14 |  | Re-launch the tool and select the desired drive letter. | Subsystem | Service Tools | NA | No Plan to Fix | chakraborty_d | 0 | 4 |  | Service Tools | NO | Closed | SubNC | Ingenuity 64
CTUDT00622767 | Refresh | Closed | wang_l | March 29, 2019 3:17:27 PM GMT+08:00 | Refresh Upgrade: Refresh Keyboard language "Chinese" disappeared after upgrade | WIP14 | WIP16 |  | Subsystem | Localization | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP17SegReq | Ovation |  | WIP16 | May 14, 2019 4:55:32 PM GMT+08:00
CTUDT00622767 | Refresh | Closed | wang_l | March 29, 2019 3:17:27 PM GMT+08:00 | Refresh Upgrade: Refresh Keyboard language "Chinese" disappeared after upgrade | WIP14 | WIP16 |  | Subsystem | Localization | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP17SegReq | Brilliance 64 |  | WIP16 | May 14, 2019 4:55:32 PM GMT+08:00
CTUDT00622767 | Refresh | Closed | wang_l | March 29, 2019 3:17:27 PM GMT+08:00 | Refresh Upgrade: Refresh Keyboard language "Chinese" disappeared after upgrade | WIP14 | WIP16 |  | Subsystem | Localization | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP17SegReq | Brilliance Big Bore |  | WIP16 | May 14, 2019 4:55:32 PM GMT+08:00
CTUDT00622767 | Refresh | Closed | wang_l | March 29, 2019 3:17:27 PM GMT+08:00 | Refresh Upgrade: Refresh Keyboard language "Chinese" disappeared after upgrade | WIP14 | WIP16 |  | Subsystem | Localization | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP17SegReq | Ingenuity 64 |  | WIP16 | May 14, 2019 4:55:32 PM GMT+08:00
CTUDT00622767 | Refresh | Closed | wang_l | March 29, 2019 3:17:27 PM GMT+08:00 | Refresh Upgrade: Refresh Keyboard language "Chinese" disappeared after upgrade | WIP14 | WIP16 |  | Subsystem | Localization | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP17SegReq | IQon1 |  | WIP16 | May 14, 2019 4:55:32 PM GMT+08:00
CTUDT00622765 | Refresh | Closed | nimma_p | March 29, 2019 2:48:05 PM GMT+08:00 | Refresh:-“Refresh Failed” message displayed after performing the refresh from the Reconstruction Tab. | WIP14 | Host SSRS DHF303211 Rev 12 DHF316705 Rev11 |  | Subsystem | MDU | Console- SRS- 6672 | Fixed | roy_r | 0 | 3 | Host SSRS | Console4 | NO | Closed | SubCom | Ovation |  | Host SSRS DHF303211 Rev 12 DHF316705 Rev11 | June 17, 2019 5:28:56 PM GMT+08:00
CTUDT00622765 | Refresh | Closed | nimma_p | March 29, 2019 2:48:05 PM GMT+08:00 | Refresh:-“Refresh Failed” message displayed after performing the refresh from the Reconstruction Tab. | WIP14 | Host SSRS DHF303211 Rev 12 DHF316705 Rev11 |  | Subsystem | MDU | Console- SRS- 6672 | Fixed | roy_r | 0 | 3 | Host SSRS | Console4 | NO | Closed | SubCom | Brilliance 64 |  | Host SSRS DHF303211 Rev 12 DHF316705 Rev11 | June 17, 2019 5:28:56 PM GMT+08:00
CTUDT00622765 | Refresh | Closed | nimma_p | March 29, 2019 2:48:05 PM GMT+08:00 | Refresh:-“Refresh Failed” message displayed after performing the refresh from the Reconstruction Tab. | WIP14 | Host SSRS DHF303211 Rev 12 DHF316705 Rev11 |  | Subsystem | MDU | Console- SRS- 6672 | Fixed | roy_r | 0 | 3 | Host SSRS | Console4 | NO | Closed | SubCom | Brilliance Big Bore |  | Host SSRS DHF303211 Rev 12 DHF316705 Rev11 | June 17, 2019 5:28:56 PM GMT+08:00
CTUDT00622765 | Refresh | Closed | nimma_p | March 29, 2019 2:48:05 PM GMT+08:00 | Refresh:-“Refresh Failed” message displayed after performing the refresh from the Reconstruction Tab. | WIP14 | Host SSRS DHF303211 Rev 12 DHF316705 Rev11 |  | Subsystem | MDU | Console- SRS- 6672 | Fixed | roy_r | 0 | 3 | Host SSRS | Console4 | NO | Closed | SubCom | Ingenuity 64 |  | Host SSRS DHF303211 Rev 12 DHF316705 Rev11 | June 17, 2019 5:28:56 PM GMT+08:00
CTUDT00622765 | Refresh | Closed | nimma_p | March 29, 2019 2:48:05 PM GMT+08:00 | Refresh:-“Refresh Failed” message displayed after performing the refresh from the Reconstruction Tab. | WIP14 | Host SSRS DHF303211 Rev 12 DHF316705 Rev11 |  | Subsystem | MDU | Console- SRS- 6672 | Fixed | roy_r | 0 | 3 | Host SSRS | Console4 | NO | Closed | SubCom | IQon1 |  | Host SSRS DHF303211 Rev 12 DHF316705 Rev11 | June 17, 2019 5:28:56 PM GMT+08:00
CTUDT00622755 | Refresh | Closed | attali_s | March 28, 2019 9:49:49 PM GMT+08:00 | Refresh: Acceptance/Constancy: Wrong Limits in Constancy after Baseline from Acceptance was loaded | WIP14 | WIP17 | N/A | Subsystem | Calibrations | N/A | Fixed | elviral | 0 | 2 | DMS Software | DMS | NO | Closed | WIP16SegReq | Ovation |  | WIP17 | May 28, 2019 1:18:09 AM GMT+08:00
CTUDT00622755 | Refresh | Closed | attali_s | March 28, 2019 9:49:49 PM GMT+08:00 | Refresh: Acceptance/Constancy: Wrong Limits in Constancy after Baseline from Acceptance was loaded | WIP14 | WIP17 | N/A | Subsystem | Calibrations | N/A | Fixed | elviral | 0 | 2 | DMS Software | DMS | NO | Closed | WIP16SegReq | Ingenuity 64 |  | WIP17 | May 28, 2019 1:18:09 AM GMT+08:00
CTUDT00622755 | Refresh | Closed | attali_s | March 28, 2019 9:49:49 PM GMT+08:00 | Refresh: Acceptance/Constancy: Wrong Limits in Constancy after Baseline from Acceptance was loaded | WIP14 | WIP17 | N/A | Subsystem | Calibrations | N/A | Fixed | elviral | 0 | 2 | DMS Software | DMS | NO | Closed | WIP16SegReq | Ingenuity BigBore |  | WIP17 | May 28, 2019 1:18:09 AM GMT+08:00
CTUDT00622755 | Refresh | Closed | attali_s | March 28, 2019 9:49:49 PM GMT+08:00 | Refresh: Acceptance/Constancy: Wrong Limits in Constancy after Baseline from Acceptance was loaded | WIP14 | WIP17 | N/A | Subsystem | Calibrations | N/A | Fixed | elviral | 0 | 2 | DMS Software | DMS | NO | Closed | WIP16SegReq | IQon1 |  | WIP17 | May 28, 2019 1:18:09 AM GMT+08:00
CTUDT00622749 | Refresh | Closed | gT_s | March 28, 2019 7:57:15 PM GMT+08:00 | Refresh :The alerts are not being generated in the PRS portal for the IQON CONFIG & few are not getting generated for iCT-SP | WIP 14 | WIP18 |  | Subsystem | service tools | STPS-SSRS-1198 | Fixed | ranjithkumar_m | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP18SubCom | Ovation |  | WIP18 | June 6, 2019 5:31:14 PM GMT+08:00
CTUDT00622749 | Refresh | Closed | gT_s | March 28, 2019 7:57:15 PM GMT+08:00 | Refresh :The alerts are not being generated in the PRS portal for the IQON CONFIG & few are not getting generated for iCT-SP | WIP 14 | WIP18 |  | Subsystem | service tools | STPS-SSRS-1198 | Fixed | ranjithkumar_m | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP18SubCom | IQon1 |  | WIP18 | June 6, 2019 5:31:14 PM GMT+08:00
CTUDT00622748 | Refresh | Closed | PremKumar_Rajendran | March 28, 2019 7:56:24 PM GMT+08:00 | Refresh : CD-DVD tray-eject was not blocked during acquisition when ‘Go’ for the scan, scan was in ‘Ready’ status, scan is in ‘Progress’, Repeat Scan & Extend Scan operations. | WIP14 | v4.1.10 |  | Subsystem | IFU | NA | Fixed | charles_j | 0 | 3 | User Documentation | Support-System | YES | Closed | 6-Dec-2019 | Ovation |  | v4.1.10 | December 9, 2019 10:30:36 AM GMT+08:00
CTUDT00622748 | Refresh | Closed | PremKumar_Rajendran | March 28, 2019 7:56:24 PM GMT+08:00 | Refresh : CD-DVD tray-eject was not blocked during acquisition when ‘Go’ for the scan, scan was in ‘Ready’ status, scan is in ‘Progress’, Repeat Scan & Extend Scan operations. | WIP14 | v4.1.10 |  | Subsystem | IFU | NA | Fixed | charles_j | 0 | 3 | User Documentation | Support-System | YES | Closed | 6-Dec-2019 | Brilliance 64 |  | v4.1.10 | December 9, 2019 10:30:36 AM GMT+08:00
CTUDT00622748 | Refresh | Closed | PremKumar_Rajendran | March 28, 2019 7:56:24 PM GMT+08:00 | Refresh : CD-DVD tray-eject was not blocked during acquisition when ‘Go’ for the scan, scan was in ‘Ready’ status, scan is in ‘Progress’, Repeat Scan & Extend Scan operations. | WIP14 | v4.1.10 |  | Subsystem | IFU | NA | Fixed | charles_j | 0 | 3 | User Documentation | Support-System | YES | Closed | 6-Dec-2019 | Brilliance Big Bore |  | v4.1.10 | December 9, 2019 10:30:36 AM GMT+08:00
CTUDT00622748 | Refresh | Closed | PremKumar_Rajendran | March 28, 2019 7:56:24 PM GMT+08:00 | Refresh : CD-DVD tray-eject was not blocked during acquisition when ‘Go’ for the scan, scan was in ‘Ready’ status, scan is in ‘Progress’, Repeat Scan & Extend Scan operations. | WIP14 | v4.1.10 |  | Subsystem | IFU | NA | Fixed | charles_j | 0 | 3 | User Documentation | Support-System | YES | Closed | 6-Dec-2019 | Ingenuity 64 |  | v4.1.10 | December 9, 2019 10:30:36 AM GMT+08:00
CTUDT00622748 | Refresh | Closed | PremKumar_Rajendran | March 28, 2019 7:56:24 PM GMT+08:00 | Refresh : CD-DVD tray-eject was not blocked during acquisition when ‘Go’ for the scan, scan was in ‘Ready’ status, scan is in ‘Progress’, Repeat Scan & Extend Scan operations. | WIP14 | v4.1.10 |  | Subsystem | IFU | NA | Fixed | charles_j | 0 | 3 | User Documentation | Support-System | YES | Closed | 6-Dec-2019 | IQon1 |  | v4.1.10 | December 9, 2019 10:30:36 AM GMT+08:00
CTUDT00622732 | Refresh | Closed | patra_r | March 28, 2019 1:58:28 PM GMT+08:00 | Clear All option does not clear the entries from the drop down in the Patient window | 4.1.10 & 4.7.7 |  | Once user press tab or click outside (loses focus) the combobox, all items will be cleared | Subsystem | Patient Scene | N/A | Works As Specified / Not A Defect | roy_r | 0 | 3 |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00622730 | Refresh | Closed | swamypreethi | March 28, 2019 12:12:23 PM GMT+08:00 | IOCC_Refresh: During Modality Worklist query, received response shows invalid characters for patient name in console UI for ISO-IR 13, ISO-IR 87 | WIP12 | DCS |  | Subsystem | HOST;Verified1 | NA | Fixed | swamypreethi | 0 | 2 | Verification Documentation | Support-System | NO | Closed | WIP17SegReq | Brilliance 64 |  | DCS | June 3, 2019 4:04:46 PM GMT+08:00
CTUDT00622730 | Refresh | Closed | swamypreethi | March 28, 2019 12:12:23 PM GMT+08:00 | IOCC_Refresh: During Modality Worklist query, received response shows invalid characters for patient name in console UI for ISO-IR 13, ISO-IR 87 | WIP12 | DCS |  | Subsystem | HOST;Verified1 | NA | Fixed | swamypreethi | 0 | 2 | Verification Documentation | Support-System | NO | Closed | WIP17SegReq | Brilliance Big Bore |  | DCS | June 3, 2019 4:04:46 PM GMT+08:00
CTUDT00622730 | Refresh | Closed | swamypreethi | March 28, 2019 12:12:23 PM GMT+08:00 | IOCC_Refresh: During Modality Worklist query, received response shows invalid characters for patient name in console UI for ISO-IR 13, ISO-IR 87 | WIP12 | DCS |  | Subsystem | HOST;Verified1 | NA | Fixed | swamypreethi | 0 | 2 | Verification Documentation | Support-System | NO | Closed | WIP17SegReq | Ingenuity 64 |  | DCS | June 3, 2019 4:04:46 PM GMT+08:00
CTUDT00622721 | Refresh | Closed | Bat-El_Asaraf | March 27, 2019 10:00:12 PM GMT+08:00 | Refresh Pre-Val: Plan box extension appears in wrong location | WIP14 |  |  | Subsystem | Planning | NA | Works As Specified / Not A Defect | sowmya_k | 0 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00622717 | Refresh | Closed | swack_g | March 27, 2019 9:02:17 PM GMT+08:00 | Refresh: Bad Detectors Quick Check: BDQC does not log bad detectors log files in an incremental fashion. | WIP13 | N/A |  | Subsystem | Calibrations | n/a | Fixed | yatziv_l | 0 | 2 | Design Documentation | DMS | NO | Closed | WIP15SegReq | Ovation |  | N/A | May 16, 2019 5:41:09 PM GMT+08:00
CTUDT00622717 | Refresh | Closed | swack_g | March 27, 2019 9:02:17 PM GMT+08:00 | Refresh: Bad Detectors Quick Check: BDQC does not log bad detectors log files in an incremental fashion. | WIP13 | N/A |  | Subsystem | Calibrations | n/a | Fixed | yatziv_l | 0 | 2 | Design Documentation | DMS | NO | Closed | WIP15SegReq | Brilliance Big Bore |  | N/A | May 16, 2019 5:41:09 PM GMT+08:00
CTUDT00622717 | Refresh | Closed | swack_g | March 27, 2019 9:02:17 PM GMT+08:00 | Refresh: Bad Detectors Quick Check: BDQC does not log bad detectors log files in an incremental fashion. | WIP13 | N/A |  | Subsystem | Calibrations | n/a | Fixed | yatziv_l | 0 | 2 | Design Documentation | DMS | NO | Closed | WIP15SegReq | Ingenuity 64 |  | N/A | May 16, 2019 5:41:09 PM GMT+08:00
CTUDT00622717 | Refresh | Closed | swack_g | March 27, 2019 9:02:17 PM GMT+08:00 | Refresh: Bad Detectors Quick Check: BDQC does not log bad detectors log files in an incremental fashion. | WIP13 | N/A |  | Subsystem | Calibrations | n/a | Fixed | yatziv_l | 0 | 2 | Design Documentation | DMS | NO | Closed | WIP15SegReq | IQon1 |  | N/A | May 16, 2019 5:41:09 PM GMT+08:00
CTUDT00622712 | Refresh | Closed | g_di | March 27, 2019 8:17:40 PM GMT+08:00 | Refresh:Recon Thickness and Recon Increment values are not displayed as expected in the log viewer for an axial (MIP/MinIP), ultrahigh resolution and collimation other than 2x0.5 | WIP14 | Host SSRS DHF303211 Rev 10 |  | Subsystem | log viewer | Console-SRS.17449,Console-SRS.17497 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | WIP16SubCom | Ovation |  | Host SSRS DHF303211 Rev 10 | May 9, 2019 4:47:32 PM GMT+08:00
CTUDT00622712 | Refresh | Closed | g_di | March 27, 2019 8:17:40 PM GMT+08:00 | Refresh:Recon Thickness and Recon Increment values are not displayed as expected in the log viewer for an axial (MIP/MinIP), ultrahigh resolution and collimation other than 2x0.5 | WIP14 | Host SSRS DHF303211 Rev 10 |  | Subsystem | log viewer | Console-SRS.17449,Console-SRS.17497 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | WIP16SubCom | Brilliance 64 |  | Host SSRS DHF303211 Rev 10 | May 9, 2019 4:47:32 PM GMT+08:00
CTUDT00622712 | Refresh | Closed | g_di | March 27, 2019 8:17:40 PM GMT+08:00 | Refresh:Recon Thickness and Recon Increment values are not displayed as expected in the log viewer for an axial (MIP/MinIP), ultrahigh resolution and collimation other than 2x0.5 | WIP14 | Host SSRS DHF303211 Rev 10 |  | Subsystem | log viewer | Console-SRS.17449,Console-SRS.17497 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | WIP16SubCom | Ingenuity 64 |  | Host SSRS DHF303211 Rev 10 | May 9, 2019 4:47:32 PM GMT+08:00
CTUDT00622712 | Refresh | Closed | g_di | March 27, 2019 8:17:40 PM GMT+08:00 | Refresh:Recon Thickness and Recon Increment values are not displayed as expected in the log viewer for an axial (MIP/MinIP), ultrahigh resolution and collimation other than 2x0.5 | WIP14 | Host SSRS DHF303211 Rev 10 |  | Subsystem | log viewer | Console-SRS.17449,Console-SRS.17497 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | WIP16SubCom | IQon1 |  | Host SSRS DHF303211 Rev 10 | May 9, 2019 4:47:32 PM GMT+08:00
CTUDT00622709 | Refresh | Closed | shiraguppi_k | March 27, 2019 6:59:34 PM GMT+08:00 | In Spectral CT viewer in Slab Viewer After Changing Orientation to coronal , Rendering mode changing to Average mode and thickness chaging to default | WIP13 |  | manually set all parameters | Subsystem | viewer | NA | No Plan to Fix | juster_d | 0 | 3 |  | Support_Spectral | NO | Closed |  | IQon1
CTUDT00622706 | Refresh | Closed | wei_f | March 27, 2019 6:17:33 PM GMT+08:00 | IT user restore the user account failed. | WIP14 |  | Not required. | Subsystem |  | NA | Works As Specified / Not A Defect | kapadia_v | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00622704 | Refresh | Closed | manju_kolkar | March 27, 2019 5:37:31 PM GMT+08:00 | BigBore Refresh - Exam Summary is not created when ‘End Exam’ is pressed | BB_WIP4 | BB_WIP6 | No | Subsystem | Exam Summary | Console-SRS.14285 | Fixed | sagar_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP16SubCom | Brilliance Big Bore |  | BB_WIP6 | April 24, 2019 2:04:33 PM GMT+08:00
CTUDT00622693 | Refresh | Closed | alex_nihaev | March 26, 2019 8:47:55 PM GMT+08:00 | Refresh: BB Bay 136 with MRC60, Performance can not pass with MRC 60 | BB_WIP4 |  | None. If this test remains, then no 60 kW BigBore system can be released to a customer. | Subsystem | Performance | N/A | Works As Specified / Not A Defect | morton_t | 0 | 2 |  | Support-Physics | NO | Closed | BB_WIP7_SegReq | Brilliance Big Bore
CTUDT00622690 | Refresh | Closed | alex_nihaev | March 26, 2019 7:54:28 PM GMT+08:00 | Refresh: Calibrations: CTDI Free Air Spec should be updated | WIP14 | WIP18 | Do not connect the Xi dose meter to the USB port, and system should work in semi-automatic mode. | Subsystem | Blocker | (ID 23) of DHF170399, Rev 07 | Fixed | morton_t | 0 | 2 | Requirements Documentation | Support-Physics | NO | Closed | WIP17SegReq | Ovation |  | WIP18 | June 24, 2019 7:46:29 PM GMT+08:00
CTUDT00622690 | Refresh | Closed | alex_nihaev | March 26, 2019 7:54:28 PM GMT+08:00 | Refresh: Calibrations: CTDI Free Air Spec should be updated | WIP14 | WIP18 | Do not connect the Xi dose meter to the USB port, and system should work in semi-automatic mode. | Subsystem | Blocker | (ID 23) of DHF170399, Rev 07 | Fixed | morton_t | 0 | 2 | Requirements Documentation | Support-Physics | NO | Closed | WIP17SegReq | Brilliance Big Bore |  | WIP18 | June 24, 2019 7:46:29 PM GMT+08:00
CTUDT00622690 | Refresh | Closed | alex_nihaev | March 26, 2019 7:54:28 PM GMT+08:00 | Refresh: Calibrations: CTDI Free Air Spec should be updated | WIP14 | WIP18 | Do not connect the Xi dose meter to the USB port, and system should work in semi-automatic mode. | Subsystem | Blocker | (ID 23) of DHF170399, Rev 07 | Fixed | morton_t | 0 | 2 | Requirements Documentation | Support-Physics | NO | Closed | WIP17SegReq | Ingenuity 64 |  | WIP18 | June 24, 2019 7:46:29 PM GMT+08:00
CTUDT00622690 | Refresh | Closed | alex_nihaev | March 26, 2019 7:54:28 PM GMT+08:00 | Refresh: Calibrations: CTDI Free Air Spec should be updated | WIP14 | WIP18 | Do not connect the Xi dose meter to the USB port, and system should work in semi-automatic mode. | Subsystem | Blocker | (ID 23) of DHF170399, Rev 07 | Fixed | morton_t | 0 | 2 | Requirements Documentation | Support-Physics | NO | Closed | WIP17SegReq | IQon1 |  | WIP18 | June 24, 2019 7:46:29 PM GMT+08:00
CTUDT00622689 | Refresh | Closed | bommiReddy_r | March 26, 2019 7:52:38 PM GMT+08:00 | Refresh: console user messages localization mismatches | WIP13 | WIP17 | NA | Subsystem | Localization | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP15SegReq | Ovation |  | WIP17 | May 24, 2019 6:40:42 PM GMT+08:00
CTUDT00622689 | Refresh | Closed | bommiReddy_r | March 26, 2019 7:52:38 PM GMT+08:00 | Refresh: console user messages localization mismatches | WIP13 | WIP17 | NA | Subsystem | Localization | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP15SegReq | Brilliance 64 |  | WIP17 | May 24, 2019 6:40:42 PM GMT+08:00
CTUDT00622689 | Refresh | Closed | bommiReddy_r | March 26, 2019 7:52:38 PM GMT+08:00 | Refresh: console user messages localization mismatches | WIP13 | WIP17 | NA | Subsystem | Localization | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP15SegReq | Ingenuity 64 |  | WIP17 | May 24, 2019 6:40:42 PM GMT+08:00
CTUDT00622689 | Refresh | Closed | bommiReddy_r | March 26, 2019 7:52:38 PM GMT+08:00 | Refresh: console user messages localization mismatches | WIP13 | WIP17 | NA | Subsystem | Localization | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP15SegReq | IQon1 |  | WIP17 | May 24, 2019 6:40:42 PM GMT+08:00
CTUDT00622687 | Refresh | Closed | Revital_Cohen | March 26, 2019 7:33:34 PM GMT+08:00 | Refresh - Language buttons in License Agreement screen when opening WorkFlow appears as the last language before the change | WIP14 | WIP17 |  | Subsystem | Localization | NA | Fixed | s_sa | 0 | 4 | Console USP_EBW | iCAP | NO | Closed | WIP17SubCom | Ovation |  | WIP17 | May 27, 2019 1:53:25 PM GMT+08:00
CTUDT00622687 | Refresh | Closed | Revital_Cohen | March 26, 2019 7:33:34 PM GMT+08:00 | Refresh - Language buttons in License Agreement screen when opening WorkFlow appears as the last language before the change | WIP14 | WIP17 |  | Subsystem | Localization | NA | Fixed | s_sa | 0 | 4 | Console USP_EBW | iCAP | NO | Closed | WIP17SubCom | Brilliance 64 |  | WIP17 | May 27, 2019 1:53:25 PM GMT+08:00
CTUDT00622687 | Refresh | Closed | Revital_Cohen | March 26, 2019 7:33:34 PM GMT+08:00 | Refresh - Language buttons in License Agreement screen when opening WorkFlow appears as the last language before the change | WIP14 | WIP17 |  | Subsystem | Localization | NA | Fixed | s_sa | 0 | 4 | Console USP_EBW | iCAP | NO | Closed | WIP17SubCom | Brilliance Big Bore |  | WIP17 | May 27, 2019 1:53:25 PM GMT+08:00
CTUDT00622687 | Refresh | Closed | Revital_Cohen | March 26, 2019 7:33:34 PM GMT+08:00 | Refresh - Language buttons in License Agreement screen when opening WorkFlow appears as the last language before the change | WIP14 | WIP17 |  | Subsystem | Localization | NA | Fixed | s_sa | 0 | 4 | Console USP_EBW | iCAP | NO | Closed | WIP17SubCom | Ingenuity 64 |  | WIP17 | May 27, 2019 1:53:25 PM GMT+08:00
CTUDT00622687 | Refresh | Closed | Revital_Cohen | March 26, 2019 7:33:34 PM GMT+08:00 | Refresh - Language buttons in License Agreement screen when opening WorkFlow appears as the last language before the change | WIP14 | WIP17 |  | Subsystem | Localization | NA | Fixed | s_sa | 0 | 4 | Console USP_EBW | iCAP | NO | Closed | WIP17SubCom | IQon1 |  | WIP17 | May 27, 2019 1:53:25 PM GMT+08:00
CTUDT00622685 | Refresh | Closed | Revital_Cohen | March 26, 2019 7:19:12 PM GMT+08:00 | Refresh iQon - The buttons in Reconstruction Tab for “delete” and “now” appears with the word on top of the image | WIP14 |  |  | Subsystem | Localization | NA | No Plan to Fix | prasad_r | 0 | 4 |  | Console4 | NO | Closed |  | IQon1
CTUDT00622681 | Refresh | Closed | Revital_Cohen | March 26, 2019 6:04:21 PM GMT+08:00 | Refresh iQon Only - Although only the interface language is changed to French it seems that the special keys in the keyboard were changed too | WIP14 | WIP19 | Press Win+Space to change keyboard | Subsystem | Localization | NA | Fixed | s_sa | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP19SubCom | Ovation |  | WIP19 | July 18, 2019 2:11:53 PM GMT+08:00
CTUDT00622681 | Refresh | Closed | Revital_Cohen | March 26, 2019 6:04:21 PM GMT+08:00 | Refresh iQon Only - Although only the interface language is changed to French it seems that the special keys in the keyboard were changed too | WIP14 | WIP19 | Press Win+Space to change keyboard | Subsystem | Localization | NA | Fixed | s_sa | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP19SubCom | Brilliance 64 |  | WIP19 | July 18, 2019 2:11:53 PM GMT+08:00
CTUDT00622681 | Refresh | Closed | Revital_Cohen | March 26, 2019 6:04:21 PM GMT+08:00 | Refresh iQon Only - Although only the interface language is changed to French it seems that the special keys in the keyboard were changed too | WIP14 | WIP19 | Press Win+Space to change keyboard | Subsystem | Localization | NA | Fixed | s_sa | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP19SubCom | Brilliance Big Bore |  | WIP19 | July 18, 2019 2:11:53 PM GMT+08:00
CTUDT00622681 | Refresh | Closed | Revital_Cohen | March 26, 2019 6:04:21 PM GMT+08:00 | Refresh iQon Only - Although only the interface language is changed to French it seems that the special keys in the keyboard were changed too | WIP14 | WIP19 | Press Win+Space to change keyboard | Subsystem | Localization | NA | Fixed | s_sa | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP19SubCom | Ingenuity 64 |  | WIP19 | July 18, 2019 2:11:53 PM GMT+08:00
CTUDT00622681 | Refresh | Closed | Revital_Cohen | March 26, 2019 6:04:21 PM GMT+08:00 | Refresh iQon Only - Although only the interface language is changed to French it seems that the special keys in the keyboard were changed too | WIP14 | WIP19 | Press Win+Space to change keyboard | Subsystem | Localization | NA | Fixed | s_sa | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP19SubCom | IQon1 |  | WIP19 | July 18, 2019 2:11:53 PM GMT+08:00
CTUDT00622679 | Refresh | Closed | shiraguppi_k | March 26, 2019 5:44:55 PM GMT+08:00 | Spectral CT viewer application crashing with bug rep error after loading saved Bookmark of Fusion preset | WIP13 | WIP18 | Launch sCTV again with the original SBI/results. | Subsystem | Console | NA | Fixed | aviw | 0 | 3 | SDViewer.exe | Support_Spectral | NO | Closed | WIP17SegReq | IQon1 |  | WIP18 | May 30, 2019 11:16:39 PM GMT+08:00
CTUDT00622677 | Refresh | Closed | wang_l | March 26, 2019 4:00:13 PM GMT+08:00 | Tube configuration data restoredd into incorrect path for upgrade from 4.1.7 to 4.1.10 | WIP14 | WIP16 |  | Subsystem | Installation | NA | Fixed | chakraborty_d | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP16SubCom | Ovation |  | WIP16 | April 12, 2019 2:18:23 PM GMT+08:00
CTUDT00622677 | Refresh | Closed | wang_l | March 26, 2019 4:00:13 PM GMT+08:00 | Tube configuration data restoredd into incorrect path for upgrade from 4.1.7 to 4.1.10 | WIP14 | WIP16 |  | Subsystem | Installation | NA | Fixed | chakraborty_d | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP16SubCom | Brilliance 64 |  | WIP16 | April 12, 2019 2:18:23 PM GMT+08:00
CTUDT00622677 | Refresh | Closed | wang_l | March 26, 2019 4:00:13 PM GMT+08:00 | Tube configuration data restoredd into incorrect path for upgrade from 4.1.7 to 4.1.10 | WIP14 | WIP16 |  | Subsystem | Installation | NA | Fixed | chakraborty_d | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP16SubCom | Ingenuity 64 |  | WIP16 | April 12, 2019 2:18:23 PM GMT+08:00
CTUDT00622672 | Refresh | Closed | levy_s | March 26, 2019 3:42:08 PM GMT+08:00 | Refresh: System error appears when trying to execute different scenarios | WIP14 |  | DMS Workaround for Calibration: o Option 1: Run Phantom/hcor calibration… OR o Option 2: Update configuration file C:\Pms\Config\Physics\Config Files\SystemParams\AllowedCombinations_Phantom\PhantomAllowedCombinations<model>.cfg change DefaultTable=Phantom_Zero.tbl from None. Same for hcor slice under AllowedCombinations_HcoreSlice This shall re-enable zero table for integration. | Subsystem | System | NA | Invalid/Uncontrolled Configuration | yatziv_l | 0 | 2 |  | DMS | NO | Closed | WIP16SubCom | Ovation
CTUDT00622671 | Refresh | Closed | wang_l | March 26, 2019 3:16:15 PM GMT+08:00 | Refresh- Remote configuration not allowed by olevel service | WIP14 |  |  | Subsystem | Remote Configuration | NA | Invalid/Uncontrolled Configuration | navale_s | 0 |  |  | Service Tools | NO | Closed |  | Ingenuity 64
CTUDT00622668 | Refresh | Closed | bommiReddy_r | March 26, 2019 3:04:22 PM GMT+08:00 | Refresh: Security: Incorrect message displayed when user enters un supported character in emergency user name | WIP14 | WIP16 |  | Subsystem | ECM | NA | Fixed | goswami_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIp16Subcom | Ovation |  | WIP16 | April 9, 2019 6:54:51 PM GMT+08:00
CTUDT00622668 | Refresh | Closed | bommiReddy_r | March 26, 2019 3:04:22 PM GMT+08:00 | Refresh: Security: Incorrect message displayed when user enters un supported character in emergency user name | WIP14 | WIP16 |  | Subsystem | ECM | NA | Fixed | goswami_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIp16Subcom | Brilliance 64 |  | WIP16 | April 9, 2019 6:54:51 PM GMT+08:00
CTUDT00622668 | Refresh | Closed | bommiReddy_r | March 26, 2019 3:04:22 PM GMT+08:00 | Refresh: Security: Incorrect message displayed when user enters un supported character in emergency user name | WIP14 | WIP16 |  | Subsystem | ECM | NA | Fixed | goswami_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIp16Subcom | Ingenuity 64 |  | WIP16 | April 9, 2019 6:54:51 PM GMT+08:00
CTUDT00622668 | Refresh | Closed | bommiReddy_r | March 26, 2019 3:04:22 PM GMT+08:00 | Refresh: Security: Incorrect message displayed when user enters un supported character in emergency user name | WIP14 | WIP16 |  | Subsystem | ECM | NA | Fixed | goswami_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIp16Subcom | IQon1 |  | WIP16 | April 9, 2019 6:54:51 PM GMT+08:00
CTUDT00622667 | Refresh | Closed | jotawar_d | March 26, 2019 1:15:43 PM GMT+08:00 | Refresh - Wrong reconstruction time in acquisition tab | 4.8 | BB_WIP6 | NA | Subsystem | Patient Directory (UI) | Console-SRS 15999 | Fixed | jose_s | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP6_SubCom | Brilliance Big Bore |  | BB_WIP6 | April 23, 2019 7:24:18 PM GMT+08:00
CTUDT00622666 | Refresh | Closed | jotawar_d | March 26, 2019 1:11:39 PM GMT+08:00 | Refresh: Paste option should not be there for suggested EC group | 4.8 | BB_WIP5.1 | leave the page and go there again | Subsystem | Exam Cards | Console.SRS-834 | Fixed | datta_i | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP6_SubCom | Brilliance Big Bore |  | BB_WIP5.1 | April 1, 2019 4:33:19 PM GMT+08:00
CTUDT00622664 | Refresh | Closed | kadish_n | March 26, 2019 2:58:15 AM GMT+08:00 | Refresh: CTDI application started in Manual mode even though some of the dose meter's fields were blank | WIP14 | WIP17 |  | Subsystem | Calibrations | (ID 23) of DHF170399, Rev 07 | Fixed | daibes_s | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP16SubCom | Ovation |  | WIP17 | May 12, 2019 10:35:47 PM GMT+08:00
CTUDT00622664 | Refresh | Closed | kadish_n | March 26, 2019 2:58:15 AM GMT+08:00 | Refresh: CTDI application started in Manual mode even though some of the dose meter's fields were blank | WIP14 | WIP17 |  | Subsystem | Calibrations | (ID 23) of DHF170399, Rev 07 | Fixed | daibes_s | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP16SubCom | Brilliance Big Bore |  | WIP17 | May 12, 2019 10:35:47 PM GMT+08:00
CTUDT00622664 | Refresh | Closed | kadish_n | March 26, 2019 2:58:15 AM GMT+08:00 | Refresh: CTDI application started in Manual mode even though some of the dose meter's fields were blank | WIP14 | WIP17 |  | Subsystem | Calibrations | (ID 23) of DHF170399, Rev 07 | Fixed | daibes_s | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP16SubCom | Ingenuity 64 |  | WIP17 | May 12, 2019 10:35:47 PM GMT+08:00
CTUDT00622664 | Refresh | Closed | kadish_n | March 26, 2019 2:58:15 AM GMT+08:00 | Refresh: CTDI application started in Manual mode even though some of the dose meter's fields were blank | WIP14 | WIP17 |  | Subsystem | Calibrations | (ID 23) of DHF170399, Rev 07 | Fixed | daibes_s | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP16SubCom | IQon1 |  | WIP17 | May 12, 2019 10:35:47 PM GMT+08:00
CTUDT00622661 | Refresh | Closed | kadish_n | March 26, 2019 12:40:55 AM GMT+08:00 | Refresh: Use the wild card "*" is not supported in allowed combinations cfg file | WIP14 |  |  | Subsystem | DMS |  | Duplicate | frider_o | 0 |  |  | DMS | NO | Closed |  | Ingenuity 64
CTUDT00622660 | Refresh | Closed | djakovic_g | March 26, 2019 12:38:49 AM GMT+08:00 | CIRS RAID Diagnostics Test FAILS on Gen5.3 | WIP14 | WIP18 |  | Subsystem | CIRS; Verified1 | CIRS.SSRS.4856 and CIRS.SSRS.5362 | Fixed | richards_b | 0 | 2 | CIRS_SW_ServiceNode | CIRS | NO | Closed | WIP19SubCom | Ovation |  | WIP18 | July 12, 2019 12:54:40 AM GMT+08:00
CTUDT00622660 | Refresh | Closed | djakovic_g | March 26, 2019 12:38:49 AM GMT+08:00 | CIRS RAID Diagnostics Test FAILS on Gen5.3 | WIP14 | WIP18 |  | Subsystem | CIRS; Verified1 | CIRS.SSRS.4856 and CIRS.SSRS.5362 | Fixed | richards_b | 0 | 2 | CIRS_SW_ServiceNode | CIRS | NO | Closed | WIP19SubCom | Brilliance 64 |  | WIP18 | July 12, 2019 12:54:40 AM GMT+08:00
CTUDT00622660 | Refresh | Closed | djakovic_g | March 26, 2019 12:38:49 AM GMT+08:00 | CIRS RAID Diagnostics Test FAILS on Gen5.3 | WIP14 | WIP18 |  | Subsystem | CIRS; Verified1 | CIRS.SSRS.4856 and CIRS.SSRS.5362 | Fixed | richards_b | 0 | 2 | CIRS_SW_ServiceNode | CIRS | NO | Closed | WIP19SubCom | Brilliance Big Bore |  | WIP18 | July 12, 2019 12:54:40 AM GMT+08:00
CTUDT00622660 | Refresh | Closed | djakovic_g | March 26, 2019 12:38:49 AM GMT+08:00 | CIRS RAID Diagnostics Test FAILS on Gen5.3 | WIP14 | WIP18 |  | Subsystem | CIRS; Verified1 | CIRS.SSRS.4856 and CIRS.SSRS.5362 | Fixed | richards_b | 0 | 2 | CIRS_SW_ServiceNode | CIRS | NO | Closed | WIP19SubCom | Ingenuity 64 |  | WIP18 | July 12, 2019 12:54:40 AM GMT+08:00
CTUDT00622657 | Refresh | Closed | levy_s | March 25, 2019 10:05:13 PM GMT+08:00 | Refresh: Start final recon button was disabled when a bugRep is collected in the background | WIP14 |  | Perform scan before or after collect BugRep. | Subsystem | System | N/A | No Plan to Fix | Mark_Huang | 0 | 4 |  | GCX | NO | Closed |  | Ovation
CTUDT00622652 | Refresh | Closed | Noy_Seelig | March 25, 2019 9:26:21 PM GMT+08:00 | Refresh: Estop opens with countdown window | BB_WIP4 |  |  | Subsystem | Gantry | N/A | Works As Specified / Not A Defect | yatziv_l | 0 |  |  | DMS | NO | Closed |  | Ingenuity BigBore
CTUDT00622648 | Refresh | Closed | atin_s | March 25, 2019 4:35:53 PM GMT+08:00 | IQON 4.7.5:In the generated Dose report, the "Procedure Context" is described as "... with contrast media..." despite non-contrast procedure | 4.8 | BB_WIP6 |  | Subsystem | Patient Directory | N/A | Fixed | atin_s | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP6_SubCom | Brilliance Big Bore | PR8468706 | BB_WIP6 | April 15, 2019 3:20:54 PM GMT+08:00
CTUDT00622645 | Refresh | Closed | 320053767 | March 25, 2019 4:30:39 PM GMT+08:00 | Refresh- unable to add EC to Suggested group tab | WIP14 |  |  | Subsystem | ECM |  | Duplicate | sanjeeva_a | 0 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00622644 | Refresh | Closed | atin_s | March 25, 2019 4:26:21 PM GMT+08:00 | Thin Slice images are not stored on configured PACS devices. | 4.8 | BB_WIP6 | Recommendation/ workaround to User: Please ensure that all thin slices are appeared to PD before performing the end exam. | Subsystem |  | N/A | Fixed | atin_s | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP6_SubCom | Brilliance Big Bore | 8240184 | BB_WIP6 | April 15, 2019 5:12:13 PM GMT+08:00
CTUDT00622643 | Refresh | Closed | atin_s | March 25, 2019 4:24:16 PM GMT+08:00 | System allows to proceed to Plan without complete patient demographics and Application crashed at press of GO | 4.1.6 | BB_WIP6 | Press New Patient/Anonymous or fill in all the details before proceeding with the scan. Do the above step in case the patient info is being pulled from HIS/RIS | Subsystem | System | N/A | Fixed | atin_s | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP6_SubCom | Brilliance Big Bore | CFR# 9456433 | BB_WIP6 | April 15, 2019 3:35:19 PM GMT+08:00
CTUDT00622642 | Refresh | Closed | atin_s | March 25, 2019 4:16:02 PM GMT+08:00 | HB: FoK: Yokote: Exam manager instance is not getting closed even when user has exited the View2 window | 4.8 | BB_WIP6 | Modify the Exam Summary preferences .If any one of results,injection,Executed-surview etc option is selected along with the first two then this problem can be avoided | Subsystem | Exam Card Manager | N/A | Fixed | atin_s | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP6_SubCom | Brilliance Big Bore | IQon HB 6003368 | BB_WIP6 | April 15, 2019 9:03:33 PM GMT+08:00
CTUDT00622636 | Refresh | Closed | Noy_Seelig | March 25, 2019 1:52:28 PM GMT+08:00 | Refresh: During FEE test, error window appeared preventing test from completing | BB_WIP4 |  | N/A | Subsystem | Calibrations | N/A | Works As Specified / Not A Defect | yosfan_s | 0 |  |  | DMS | NO | Closed |  | Brilliance Big Bore
CTUDT00622633 | Refresh | Closed | cohen_y | March 25, 2019 1:11:09 AM GMT+08:00 | Refresh: HCOR Slice: Harmonized: Collimation mapping mismatch between Spec and HcorSliceAllowedCombinationsOrion_HPhantom.cfg file. | WIP14 | N/A |  | Subsystem | HCOR | CT-NM.SRS.IQ.3095, CT-NM.SRS.IQ.2638 | Fixed | braunstein_d | 0 | 2 | Design Documentation | Support-Physics | YES | Closed | WIP16SubCom | Ingenuity 64 |  | N/A | May 27, 2019 7:26:09 PM GMT+08:00
CTUDT00622631 | Refresh | Closed | Valeria_Kashtanov | March 25, 2019 12:31:32 AM GMT+08:00 | Refresh: Calibrations: Performance Test: Total scan seconds from Tube related data section is missing in HTML Report | WIP13 | WIP17 | See the "Total Scan Seconds" in XRay Tube Data - same report, but several rows downwards - it is a part of the summary as well! | Subsystem | Calibrations | DHF170396 | Fixed | bikman | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP17SegReq | Ovation |  | WIP17 | May 13, 2019 6:45:36 PM GMT+08:00
CTUDT00622629 | Refresh | Closed | attali_s | March 25, 2019 12:06:52 AM GMT+08:00 | Refresh: Performance: Wrong Position of ROI for case Body 13c with Blue Phantom | WIP14 | WIP17 | N/A | Subsystem | Performance Test | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP17SubCom | Ingenuity 64 |  | WIP17 | May 14, 2019 12:04:34 AM GMT+08:00
CTUDT00622626 | Refresh | Closed | beck_d | March 24, 2019 11:05:16 PM GMT+08:00 | Refresh: Calibrations: Constancy Test: Harmonized Phantom: For tolerances where two values are presented in spec (Percentage and absolute value) the percentage is applied and not the biggest value | WIP14 |  |  | Subsystem | Acceptance/Constancy |  | Duplicate | frider_o | 0 |  |  | DMS | YES | Closed |  | Ingenuity 64
CTUDT00622625 | Refresh | Closed | urim | March 24, 2019 11:04:28 PM GMT+08:00 | Refresh Pre-Val: Iso-Center Line on Lateral Surview shows same value for different couch heights | WIP14 |  |  | Subsystem |  |  | Not Enough Information | Prakash_Chand_Yadav | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00622625 | Refresh | Closed | urim | March 24, 2019 11:04:28 PM GMT+08:00 | Refresh Pre-Val: Iso-Center Line on Lateral Surview shows same value for different couch heights | WIP14 |  |  | Subsystem |  |  | Not Enough Information | Prakash_Chand_Yadav | 0 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00622620 | Refresh | Closed | beck_d | March 24, 2019 10:01:53 PM GMT+08:00 | Refresh: Calibrations: Constancy Test: In Uniformity, N/A appears under scale column in the result’s summary | WIP14 | WIP17 | To use the table with numerical results which exists and correct. | Subsystem | Constancy Test | N/A | Fixed | elviral | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP17SegReq | Ovation |  | WIP17 | May 7, 2019 10:24:54 PM GMT+08:00
CTUDT00622620 | Refresh | Closed | beck_d | March 24, 2019 10:01:53 PM GMT+08:00 | Refresh: Calibrations: Constancy Test: In Uniformity, N/A appears under scale column in the result’s summary | WIP14 | WIP17 | To use the table with numerical results which exists and correct. | Subsystem | Constancy Test | N/A | Fixed | elviral | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP17SegReq | Ingenuity 64 |  | WIP17 | May 7, 2019 10:24:54 PM GMT+08:00
CTUDT00622620 | Refresh | Closed | beck_d | March 24, 2019 10:01:53 PM GMT+08:00 | Refresh: Calibrations: Constancy Test: In Uniformity, N/A appears under scale column in the result’s summary | WIP14 | WIP17 | To use the table with numerical results which exists and correct. | Subsystem | Constancy Test | N/A | Fixed | elviral | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP17SegReq | Ingenuity BigBore |  | WIP17 | May 7, 2019 10:24:54 PM GMT+08:00
CTUDT00622620 | Refresh | Closed | beck_d | March 24, 2019 10:01:53 PM GMT+08:00 | Refresh: Calibrations: Constancy Test: In Uniformity, N/A appears under scale column in the result’s summary | WIP14 | WIP17 | To use the table with numerical results which exists and correct. | Subsystem | Constancy Test | N/A | Fixed | elviral | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP17SegReq | IQon1 |  | WIP17 | May 7, 2019 10:24:54 PM GMT+08:00
CTUDT00622619 | Refresh | Closed | kadish_n | March 24, 2019 9:55:00 PM GMT+08:00 | Refresh: Invalid collimation options in ECs | WIP14 | N/A | N/A | Subsystem | Invalid Parameter | N/A | Fixed | yosfan_s | 0 | 2 | Test Design | DMS | NO | Closed | WIP17SubCom | Ingenuity 64 |  | N/A | May 23, 2019 2:57:08 PM GMT+08:00
CTUDT00622617 | Refresh | Closed | urim | March 24, 2019 9:15:24 PM GMT+08:00 | Refresh Pre-Val: Emojies are allowed in Chinese keyboard for Patient Name | WIP14 | WIP18 |  | Subsystem | IFU | NA | Fixed | s_sa | 0 | 4 | Console USP_EBW | iCAP | NO | Closed | WIP18SubCom | Ovation |  | WIP18 | July 2, 2019 4:24:25 PM GMT+08:00
CTUDT00622617 | Refresh | Closed | urim | March 24, 2019 9:15:24 PM GMT+08:00 | Refresh Pre-Val: Emojies are allowed in Chinese keyboard for Patient Name | WIP14 | WIP18 |  | Subsystem | IFU | NA | Fixed | s_sa | 0 | 4 | Console USP_EBW | iCAP | NO | Closed | WIP18SubCom | Brilliance 64 |  | WIP18 | July 2, 2019 4:24:25 PM GMT+08:00
CTUDT00622617 | Refresh | Closed | urim | March 24, 2019 9:15:24 PM GMT+08:00 | Refresh Pre-Val: Emojies are allowed in Chinese keyboard for Patient Name | WIP14 | WIP18 |  | Subsystem | IFU | NA | Fixed | s_sa | 0 | 4 | Console USP_EBW | iCAP | NO | Closed | WIP18SubCom | Brilliance Big Bore |  | WIP18 | July 2, 2019 4:24:25 PM GMT+08:00
CTUDT00622617 | Refresh | Closed | urim | March 24, 2019 9:15:24 PM GMT+08:00 | Refresh Pre-Val: Emojies are allowed in Chinese keyboard for Patient Name | WIP14 | WIP18 |  | Subsystem | IFU | NA | Fixed | s_sa | 0 | 4 | Console USP_EBW | iCAP | NO | Closed | WIP18SubCom | Ingenuity 64 |  | WIP18 | July 2, 2019 4:24:25 PM GMT+08:00
CTUDT00622617 | Refresh | Closed | urim | March 24, 2019 9:15:24 PM GMT+08:00 | Refresh Pre-Val: Emojies are allowed in Chinese keyboard for Patient Name | WIP14 | WIP18 |  | Subsystem | IFU | NA | Fixed | s_sa | 0 | 4 | Console USP_EBW | iCAP | NO | Closed | WIP18SubCom | IQon1 |  | WIP18 | July 2, 2019 4:24:25 PM GMT+08:00
CTUDT00622615 | Refresh | Closed | attali_s | March 24, 2019 8:12:46 PM GMT+08:00 | Refresh: HCOR Slice Infant: Voltage section is missing in the customization window | WIP13 | N/A | There is no workaround since the software doesn't include a feature that is required by the specification. In this case the specification requirement is not correct so there is no workaround needed and the system is performing as expected. The specification needs to be revised to reconcile this issue. | Subsystem | HCOR | CT-NM.SRS.IQ.3095, CT-NM.SRS.IQ.2638 | Fixed | braunstein_d | 0 | 2 | Design Documentation | Support-Physics | NO | Closed | WIP16SubCom | Ingenuity 64 |  | N/A | May 27, 2019 7:32:10 PM GMT+08:00
CTUDT00622612 | Refresh | Closed | shekhar_p | March 22, 2019 9:19:22 PM GMT+08:00 | DRI Field accepting integers other than 1 to 55 | WIP14 | WIP16 |  | Subsystem | ECM | NA | Fixed | datta_i | 0 | 4 | Console Software | Console4 | NO | Closed | WIP16SubCom | Ovation |  | WIP16 | April 17, 2019 3:05:39 PM GMT+08:00
CTUDT00622612 | Refresh | Closed | shekhar_p | March 22, 2019 9:19:22 PM GMT+08:00 | DRI Field accepting integers other than 1 to 55 | WIP14 | WIP16 |  | Subsystem | ECM | NA | Fixed | datta_i | 0 | 4 | Console Software | Console4 | NO | Closed | WIP16SubCom | Brilliance 64 |  | WIP16 | April 17, 2019 3:05:39 PM GMT+08:00
CTUDT00622612 | Refresh | Closed | shekhar_p | March 22, 2019 9:19:22 PM GMT+08:00 | DRI Field accepting integers other than 1 to 55 | WIP14 | WIP16 |  | Subsystem | ECM | NA | Fixed | datta_i | 0 | 4 | Console Software | Console4 | NO | Closed | WIP16SubCom | Ingenuity 64 |  | WIP16 | April 17, 2019 3:05:39 PM GMT+08:00
CTUDT00622612 | Refresh | Closed | shekhar_p | March 22, 2019 9:19:22 PM GMT+08:00 | DRI Field accepting integers other than 1 to 55 | WIP14 | WIP16 |  | Subsystem | ECM | NA | Fixed | datta_i | 0 | 4 | Console Software | Console4 | NO | Closed | WIP16SubCom | IQon1 |  | WIP16 | April 17, 2019 3:05:39 PM GMT+08:00
CTUDT00622605 | Refresh | Closed | sowmya_k | March 22, 2019 5:42:04 PM GMT+08:00 | Refresh, IQon - Usability: Disable options are not display clear to the user | BB_WIP5 | BB_WIP5.1 |  | Subsystem | ECM | NA | Fixed | wilsonraja | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP5_SubCom | Brilliance Big Bore |  | BB_WIP5.1 | March 26, 2019 5:36:29 PM GMT+08:00
CTUDT00622600 | Refresh | Closed | atrash_f | March 21, 2019 10:50:52 PM GMT+08:00 | Refresh- N/A SSDE and patient size values | WIP14 | WIP15 |  | Subsystem | SSDE | N/A | Fixed | bhawna_g | 0 | 2 | Couch Software | Console4 | NO | Closed | WIP15SubCom | Ovation |  | WIP15 | March 28, 2019 2:01:13 PM GMT+08:00
CTUDT00622600 | Refresh | Closed | atrash_f | March 21, 2019 10:50:52 PM GMT+08:00 | Refresh- N/A SSDE and patient size values | WIP14 | WIP15 |  | Subsystem | SSDE | N/A | Fixed | bhawna_g | 0 | 2 | Couch Software | Console4 | NO | Closed | WIP15SubCom | Brilliance 64 |  | WIP15 | March 28, 2019 2:01:13 PM GMT+08:00
CTUDT00622600 | Refresh | Closed | atrash_f | March 21, 2019 10:50:52 PM GMT+08:00 | Refresh- N/A SSDE and patient size values | WIP14 | WIP15 |  | Subsystem | SSDE | N/A | Fixed | bhawna_g | 0 | 2 | Couch Software | Console4 | NO | Closed | WIP15SubCom | Ingenuity 64 |  | WIP15 | March 28, 2019 2:01:13 PM GMT+08:00
CTUDT00622600 | Refresh | Closed | atrash_f | March 21, 2019 10:50:52 PM GMT+08:00 | Refresh- N/A SSDE and patient size values | WIP14 | WIP15 |  | Subsystem | SSDE | N/A | Fixed | bhawna_g | 0 | 2 | Couch Software | Console4 | NO | Closed | WIP15SubCom | IQon1 |  | WIP15 | March 28, 2019 2:01:13 PM GMT+08:00
CTUDT00622598 | Refresh | Closed | atrash_f | March 21, 2019 10:41:26 PM GMT+08:00 | Refresh - Couch Average size is reported as 10cm when lateral surview is performed | WIP14 |  |  | Subsystem | SSDE |  | Works As Specified / Not A Defect | richards_b | 0 |  |  | CIRS | NO | Closed |  | Ovation
CTUDT00622598 | Refresh | Closed | atrash_f | March 21, 2019 10:41:26 PM GMT+08:00 | Refresh - Couch Average size is reported as 10cm when lateral surview is performed | WIP14 |  |  | Subsystem | SSDE |  | Works As Specified / Not A Defect | richards_b | 0 |  |  | CIRS | NO | Closed |  | IQon1
CTUDT00622594 | Refresh | Closed | atrash_f | March 21, 2019 10:34:38 PM GMT+08:00 | Refresh - SSDE values displayed in UI of Head scan use body exam (16cm) conversion factor parameters | WIP14 | WIP15 |  | Subsystem | SSDE | N/A | Fixed | bhawna_g | 0 | 2 | Couch Software | Console4 | NO | Closed | WIP15SubCom | Ovation |  | WIP15 | March 28, 2019 2:00:38 PM GMT+08:00
CTUDT00622594 | Refresh | Closed | atrash_f | March 21, 2019 10:34:38 PM GMT+08:00 | Refresh - SSDE values displayed in UI of Head scan use body exam (16cm) conversion factor parameters | WIP14 | WIP15 |  | Subsystem | SSDE | N/A | Fixed | bhawna_g | 0 | 2 | Couch Software | Console4 | NO | Closed | WIP15SubCom | Brilliance 64 |  | WIP15 | March 28, 2019 2:00:38 PM GMT+08:00
CTUDT00622594 | Refresh | Closed | atrash_f | March 21, 2019 10:34:38 PM GMT+08:00 | Refresh - SSDE values displayed in UI of Head scan use body exam (16cm) conversion factor parameters | WIP14 | WIP15 |  | Subsystem | SSDE | N/A | Fixed | bhawna_g | 0 | 2 | Couch Software | Console4 | NO | Closed | WIP15SubCom | Ingenuity 64 |  | WIP15 | March 28, 2019 2:00:38 PM GMT+08:00
CTUDT00622594 | Refresh | Closed | atrash_f | March 21, 2019 10:34:38 PM GMT+08:00 | Refresh - SSDE values displayed in UI of Head scan use body exam (16cm) conversion factor parameters | WIP14 | WIP15 |  | Subsystem | SSDE | N/A | Fixed | bhawna_g | 0 | 2 | Couch Software | Console4 | NO | Closed | WIP15SubCom | IQon1 |  | WIP15 | March 28, 2019 2:00:38 PM GMT+08:00
CTUDT00622593 | Refresh | Closed | alex_nihaev | March 21, 2019 10:23:36 PM GMT+08:00 | Refresh: BB: Slowness before clinical scan is started | BB_WIP4 |  |  | Subsystem | Performance |  | Invalid/Uncontrolled Configuration | rakic_k | 0 |  |  | CIRS | NO | Closed |  | Brilliance Big Bore
CTUDT00622593 | Refresh | Closed | alex_nihaev | March 21, 2019 10:23:36 PM GMT+08:00 | Refresh: BB: Slowness before clinical scan is started | BB_WIP4 |  |  | Subsystem | Performance |  | Invalid/Uncontrolled Configuration | rakic_k | 0 |  |  | CIRS | NO | Closed |  | Ingenuity 64
CTUDT00622587 | Refresh | Closed | Noy_Seelig | March 21, 2019 3:59:40 PM GMT+08:00 | Refresh: In DMS/DataPath Diagnostics there was no continue to Diagnostics button | BB_WIP4 | WIP15 | A workaround is available after you click 'DMS / DataPath Diagnostics' A user can check the check box in the pop-up window whose content is to skip the DMS status screen and navigate directly to the Data Path application at its next logon. | Subsystem | DMS/DataPath Diagnostic | N/A | Fixed | franco_a | 0 | 2 | DMS Software | DMS | NO | Closed | WIP15SubCom | Brilliance Big Bore |  | WIP15 | March 27, 2019 4:26:09 PM GMT+08:00
CTUDT00622587 | Refresh | Closed | Noy_Seelig | March 21, 2019 3:59:40 PM GMT+08:00 | Refresh: In DMS/DataPath Diagnostics there was no continue to Diagnostics button | BB_WIP4 | WIP15 | A workaround is available after you click 'DMS / DataPath Diagnostics' A user can check the check box in the pop-up window whose content is to skip the DMS status screen and navigate directly to the Data Path application at its next logon. | Subsystem | DMS/DataPath Diagnostic | N/A | Fixed | franco_a | 0 | 2 | DMS Software | DMS | NO | Closed | WIP15SubCom | Ingenuity 64 |  | WIP15 | March 27, 2019 4:26:09 PM GMT+08:00
CTUDT00622581 | Refresh | Closed | bandil_j | March 21, 2019 1:49:10 PM GMT+08:00 | Console-SRS.9064 and Console-SRS.9065 needs to be updated. | WIP 8.1 | DHF316705 Rev08 | N/A | Subsystem |  | Console-SRS.9065 | Fixed | roy_r | 0 | 2 | Host SSRS | Console4 | NO | Closed | BB_WIP6_SubCom | Brilliance Big Bore |  | DHF316705 Rev08 | April 5, 2019 4:34:27 PM GMT+08:00
CTUDT00622580 | Refresh | Closed | bandil_j | March 21, 2019 1:44:03 PM GMT+08:00 | Refresh: “Recon Queue Manager” is still accessible when the connection between the host and CIRS is lost. | WIP10 | DHF316705 Rev09 | None | Subsystem | SSRS | Console-SRS.9948 | Fixed | roy_r | 0 | 4 | Host SSRS | Console4 | NO | Closed | BB_WIP5.1_SegReq | Brilliance Big Bore |  | DHF316705 Rev09 | April 25, 2019 12:40:07 PM GMT+08:00
CTUDT00622575 | Refresh | Closed | gT_s | March 20, 2019 9:42:44 PM GMT+08:00 | Refresh :The blank screen is displayed after the login of the IT user, the IT tools are not getting displayed. | WIP 14 | DHF303211_Rev09_Host_SSRS |  | Subsystem |  | NA | Fixed | hatti_c | 0 | 3 | Requirements Documentation | Console4 | NO | Closed | WIp16Subcom | Ovation |  | DHF303211_Rev09_Host_SSRS | April 23, 2019 7:55:53 PM GMT+08:00
CTUDT00622575 | Refresh | Closed | gT_s | March 20, 2019 9:42:44 PM GMT+08:00 | Refresh :The blank screen is displayed after the login of the IT user, the IT tools are not getting displayed. | WIP 14 | DHF303211_Rev09_Host_SSRS |  | Subsystem |  | NA | Fixed | hatti_c | 0 | 3 | Requirements Documentation | Console4 | NO | Closed | WIp16Subcom | Brilliance 64 |  | DHF303211_Rev09_Host_SSRS | April 23, 2019 7:55:53 PM GMT+08:00
CTUDT00622575 | Refresh | Closed | gT_s | March 20, 2019 9:42:44 PM GMT+08:00 | Refresh :The blank screen is displayed after the login of the IT user, the IT tools are not getting displayed. | WIP 14 | DHF303211_Rev09_Host_SSRS |  | Subsystem |  | NA | Fixed | hatti_c | 0 | 3 | Requirements Documentation | Console4 | NO | Closed | WIp16Subcom | Ingenuity 64 |  | DHF303211_Rev09_Host_SSRS | April 23, 2019 7:55:53 PM GMT+08:00
CTUDT00622575 | Refresh | Closed | gT_s | March 20, 2019 9:42:44 PM GMT+08:00 | Refresh :The blank screen is displayed after the login of the IT user, the IT tools are not getting displayed. | WIP 14 | DHF303211_Rev09_Host_SSRS |  | Subsystem |  | NA | Fixed | hatti_c | 0 | 3 | Requirements Documentation | Console4 | NO | Closed | WIp16Subcom | IQon1 |  | DHF303211_Rev09_Host_SSRS | April 23, 2019 7:55:53 PM GMT+08:00
CTUDT00622570 | Refresh | Closed | katy_rubin | March 20, 2019 8:31:12 PM GMT+08:00 | Refresh: Exam Card Manager crashes when trying to choose Resolution: 'Detailed' in ECM with 'Step&Shoot' EC | WIP13 | WIP16 | Plan all the results before changing the resolution. | Subsystem | ECM | NA | Fixed | sowmya_k | 0 | 3 | Console Software | Console4 | NO | Closed | WIP16SegReq | Ovation |  | WIP16 | April 19, 2019 12:41:37 PM GMT+08:00
CTUDT00622570 | Refresh | Closed | katy_rubin | March 20, 2019 8:31:12 PM GMT+08:00 | Refresh: Exam Card Manager crashes when trying to choose Resolution: 'Detailed' in ECM with 'Step&Shoot' EC | WIP13 | WIP16 | Plan all the results before changing the resolution. | Subsystem | ECM | NA | Fixed | sowmya_k | 0 | 3 | Console Software | Console4 | NO | Closed | WIP16SegReq | Brilliance 64 |  | WIP16 | April 19, 2019 12:41:37 PM GMT+08:00
CTUDT00622570 | Refresh | Closed | katy_rubin | March 20, 2019 8:31:12 PM GMT+08:00 | Refresh: Exam Card Manager crashes when trying to choose Resolution: 'Detailed' in ECM with 'Step&Shoot' EC | WIP13 | WIP16 | Plan all the results before changing the resolution. | Subsystem | ECM | NA | Fixed | sowmya_k | 0 | 3 | Console Software | Console4 | NO | Closed | WIP16SegReq | Ingenuity 64 |  | WIP16 | April 19, 2019 12:41:37 PM GMT+08:00
CTUDT00622570 | Refresh | Closed | katy_rubin | March 20, 2019 8:31:12 PM GMT+08:00 | Refresh: Exam Card Manager crashes when trying to choose Resolution: 'Detailed' in ECM with 'Step&Shoot' EC | WIP13 | WIP16 | Plan all the results before changing the resolution. | Subsystem | ECM | NA | Fixed | sowmya_k | 0 | 3 | Console Software | Console4 | NO | Closed | WIP16SegReq | IQon1 |  | WIP16 | April 19, 2019 12:41:37 PM GMT+08:00
CTUDT00622568 | Refresh | Closed | manju_kolkar | March 20, 2019 8:13:00 PM GMT+08:00 | Refresh - ‘DoseRight’ check box was enabled even though ‘Enabled DoseRight’ option was unchecked from preferences | WIP14 | WIP16 |  | Subsystem | ECM | NA | Fixed | datta_i | 0 | 2 | Console Software | Console4 | YES | Closed | WIP16SubCom | Ovation |  | WIP16 | April 19, 2019 7:10:09 PM GMT+08:00
CTUDT00622568 | Refresh | Closed | manju_kolkar | March 20, 2019 8:13:00 PM GMT+08:00 | Refresh - ‘DoseRight’ check box was enabled even though ‘Enabled DoseRight’ option was unchecked from preferences | WIP14 | WIP16 |  | Subsystem | ECM | NA | Fixed | datta_i | 0 | 2 | Console Software | Console4 | YES | Closed | WIP16SubCom | Brilliance 64 |  | WIP16 | April 19, 2019 7:10:09 PM GMT+08:00
CTUDT00622568 | Refresh | Closed | manju_kolkar | March 20, 2019 8:13:00 PM GMT+08:00 | Refresh - ‘DoseRight’ check box was enabled even though ‘Enabled DoseRight’ option was unchecked from preferences | WIP14 | WIP16 |  | Subsystem | ECM | NA | Fixed | datta_i | 0 | 2 | Console Software | Console4 | YES | Closed | WIP16SubCom | Ingenuity 64 |  | WIP16 | April 19, 2019 7:10:09 PM GMT+08:00
CTUDT00622568 | Refresh | Closed | manju_kolkar | March 20, 2019 8:13:00 PM GMT+08:00 | Refresh - ‘DoseRight’ check box was enabled even though ‘Enabled DoseRight’ option was unchecked from preferences | WIP14 | WIP16 |  | Subsystem | ECM | NA | Fixed | datta_i | 0 | 2 | Console Software | Console4 | YES | Closed | WIP16SubCom | IQon1 |  | WIP16 | April 19, 2019 7:10:09 PM GMT+08:00
CTUDT00622553 | Refresh | Closed | altshuller_mark | March 20, 2019 5:59:02 PM GMT+08:00 | Refresh: Acceptance/Constancy: Center Y value of the Resolution cases is not according to Spec | WIP13 |  |  | Subsystem |  |  | Does Not Apply | yatziv_l | 0 | 2 |  | DMS | NO | Closed |  | Brilliance Big Bore
CTUDT00622549 | Refresh | Closed | Valeria_Kashtanov | March 20, 2019 5:12:51 PM GMT+08:00 | Bigbore Refresh: Calibrations: Performance Test: Missing iDose Body CT measurements in results, and IMR Noise check. | BB_WIP4 | WIP15 |  | Subsystem | Performance test; Verified1 | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | BB_WIP6_SubCom | Brilliance Big Bore |  | WIP15 | March 31, 2019 4:00:23 PM GMT+08:00
CTUDT00622548 | Refresh | Closed | jindal_k | March 20, 2019 5:08:56 PM GMT+08:00 | In X-ray Trace Viewer tool incorrect name of trace is displayed | WIP14 | WIP16 |  | Subsystem | Service Tools | NA | Fixed | chakraborty_d | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP16SubCom | Ovation |  | WIP16 | April 17, 2019 2:39:02 PM GMT+08:00
CTUDT00622548 | Refresh | Closed | jindal_k | March 20, 2019 5:08:56 PM GMT+08:00 | In X-ray Trace Viewer tool incorrect name of trace is displayed | WIP14 | WIP16 |  | Subsystem | Service Tools | NA | Fixed | chakraborty_d | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP16SubCom | Brilliance 64 |  | WIP16 | April 17, 2019 2:39:02 PM GMT+08:00
CTUDT00622548 | Refresh | Closed | jindal_k | March 20, 2019 5:08:56 PM GMT+08:00 | In X-ray Trace Viewer tool incorrect name of trace is displayed | WIP14 | WIP16 |  | Subsystem | Service Tools | NA | Fixed | chakraborty_d | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP16SubCom | Ingenuity 64 |  | WIP16 | April 17, 2019 2:39:02 PM GMT+08:00
CTUDT00622548 | Refresh | Closed | jindal_k | March 20, 2019 5:08:56 PM GMT+08:00 | In X-ray Trace Viewer tool incorrect name of trace is displayed | WIP14 | WIP16 |  | Subsystem | Service Tools | NA | Fixed | chakraborty_d | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP16SubCom | IQon1 |  | WIP16 | April 17, 2019 2:39:02 PM GMT+08:00
CTUDT00622546 | Refresh | Closed | manju_kolkar | March 20, 2019 4:55:11 PM GMT+08:00 | Big Bore Refresh - ‘Could not create list of printers’ error popped up while launching Printer Configuration tool. | BB_WIP4 |  |  | Subsystem | Printer Config Tool |  | Works As Specified / Not A Defect | blessily | 0 |  |  | Service Tools | NO | Closed |  | Brilliance Big Bore
CTUDT00622544 | Refresh | Closed | manjunath_mk | March 20, 2019 4:39:44 PM GMT+08:00 | Exam cards does not get added to Suggested Exam cards while selecting Add to Suggested Group option from context menu | BB_WIP4 |  |  | Subsystem | ECM |  | Duplicate | jotawar_d | 0 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00622544 | Refresh | Closed | manjunath_mk | March 20, 2019 4:39:44 PM GMT+08:00 | Exam cards does not get added to Suggested Exam cards while selecting Add to Suggested Group option from context menu | BB_WIP4 |  |  | Subsystem | ECM |  | Duplicate | jotawar_d | 0 |  |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00622544 | Refresh | Closed | manjunath_mk | March 20, 2019 4:39:44 PM GMT+08:00 | Exam cards does not get added to Suggested Exam cards while selecting Add to Suggested Group option from context menu | BB_WIP4 |  |  | Subsystem | ECM |  | Duplicate | jotawar_d | 0 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00622544 | Refresh | Closed | manjunath_mk | March 20, 2019 4:39:44 PM GMT+08:00 | Exam cards does not get added to Suggested Exam cards while selecting Add to Suggested Group option from context menu | BB_WIP4 |  |  | Subsystem | ECM |  | Duplicate | jotawar_d | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00622544 | Refresh | Closed | manjunath_mk | March 20, 2019 4:39:44 PM GMT+08:00 | Exam cards does not get added to Suggested Exam cards while selecting Add to Suggested Group option from context menu | BB_WIP4 |  |  | Subsystem | ECM |  | Duplicate | jotawar_d | 0 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00622534 | Refresh | Closed | zhang_t | March 20, 2019 2:32:14 PM GMT+08:00 | Refresh: The reconstruction images number not correct after Estop closed | WIP14 |  |  | Subsystem | Reconstruction | NA | No Plan to Fix | richards_b | 0 | 4 |  | CIRS | NO | Closed | WIP18SubCom | Ingenuity 64
CTUDT00622533 | Refresh | Closed | ding_s | March 20, 2019 2:28:01 PM GMT+08:00 | Refresh BigBore WIP4: sound was not heard at the CT Box speaker in Voice Manager. | BB_WIP4 |  |  | System |  |  | Duplicate | prabhu_m | 0 |  |  |  |  | Closed |  | Brilliance Big Bore
CTUDT00622533 | Refresh | Closed | ding_s | March 20, 2019 2:28:01 PM GMT+08:00 | Refresh BigBore WIP4: sound was not heard at the CT Box speaker in Voice Manager. | BB_WIP4 |  |  | System |  |  | Duplicate | prabhu_m | 0 |  |  |  |  | Closed |  | Ingenuity 64
CTUDT00622531 | Refresh | Closed | manju_kolkar | March 20, 2019 2:18:41 PM GMT+08:00 | Refresh - Voice Manager: User is unable to rename user added language. | WIP14 | WIP16 |  | Subsystem | Voice Manager | NA | Fixed | v_sumanth | 0 | 2 | Console Software | Console4 | NO | Closed | WIP15SegReq | Ovation |  | WIP16 | April 5, 2019 6:26:59 PM GMT+08:00
CTUDT00622531 | Refresh | Closed | manju_kolkar | March 20, 2019 2:18:41 PM GMT+08:00 | Refresh - Voice Manager: User is unable to rename user added language. | WIP14 | WIP16 |  | Subsystem | Voice Manager | NA | Fixed | v_sumanth | 0 | 2 | Console Software | Console4 | NO | Closed | WIP15SegReq | Brilliance 64 |  | WIP16 | April 5, 2019 6:26:59 PM GMT+08:00
CTUDT00622531 | Refresh | Closed | manju_kolkar | March 20, 2019 2:18:41 PM GMT+08:00 | Refresh - Voice Manager: User is unable to rename user added language. | WIP14 | WIP16 |  | Subsystem | Voice Manager | NA | Fixed | v_sumanth | 0 | 2 | Console Software | Console4 | NO | Closed | WIP15SegReq | Ingenuity 64 |  | WIP16 | April 5, 2019 6:26:59 PM GMT+08:00
CTUDT00622531 | Refresh | Closed | manju_kolkar | March 20, 2019 2:18:41 PM GMT+08:00 | Refresh - Voice Manager: User is unable to rename user added language. | WIP14 | WIP16 |  | Subsystem | Voice Manager | NA | Fixed | v_sumanth | 0 | 2 | Console Software | Console4 | NO | Closed | WIP15SegReq | IQon1 |  | WIP16 | April 5, 2019 6:26:59 PM GMT+08:00
CTUDT00622529 | Refresh | Closed | sun_q | March 20, 2019 1:24:38 PM GMT+08:00 | [Refresh-BB] : DFS Postion failed on Intensity Check(0.5,small,80) | BB_WIP4 |  |  | Subsystem | Calibrations |  | Invalid/Uncontrolled Configuration | karkour | 0 |  |  | GCX | NO | Closed |  | Brilliance Big Bore
CTUDT00622525 | Refresh | Closed | rakic_k | March 20, 2019 1:41:33 AM GMT+08:00 | Console Application closed and r2hImageReady could not send | WIP14 | WIP17 |  | Subsystem | System | NA | Fixed | navale_s | 0 | 2 | Console Software | Console4 | YES | Closed | WIP16SubCom | Ovation |  | WIP17 | May 14, 2019 1:28:53 PM GMT+08:00
CTUDT00622525 | Refresh | Closed | rakic_k | March 20, 2019 1:41:33 AM GMT+08:00 | Console Application closed and r2hImageReady could not send | WIP14 | WIP17 |  | Subsystem | System | NA | Fixed | navale_s | 0 | 2 | Console Software | Console4 | YES | Closed | WIP16SubCom | Brilliance 64 |  | WIP17 | May 14, 2019 1:28:53 PM GMT+08:00
CTUDT00622525 | Refresh | Closed | rakic_k | March 20, 2019 1:41:33 AM GMT+08:00 | Console Application closed and r2hImageReady could not send | WIP14 | WIP17 |  | Subsystem | System | NA | Fixed | navale_s | 0 | 2 | Console Software | Console4 | YES | Closed | WIP16SubCom | Ingenuity 64 |  | WIP17 | May 14, 2019 1:28:53 PM GMT+08:00
CTUDT00622525 | Refresh | Closed | rakic_k | March 20, 2019 1:41:33 AM GMT+08:00 | Console Application closed and r2hImageReady could not send | WIP14 | WIP17 |  | Subsystem | System | NA | Fixed | navale_s | 0 | 2 | Console Software | Console4 | YES | Closed | WIP16SubCom | IQon1 |  | WIP17 | May 14, 2019 1:28:53 PM GMT+08:00
CTUDT00622523 | Refresh | Closed | rakic_k | March 20, 2019 1:33:01 AM GMT+08:00 | CIRS Application Installation Errors | WIP14 | WIP15 | manually define a drive F. | Subsystem | CIRS; Verified1 | CIRS.SSRS.4060, CIRS.SSRS.5084, and CIRS.SSRS.5085 | Fixed | richards_b | 0 | 3 | CIRS_SW_Installation | CIRS | NO | Closed | WIP15SubCom | Ovation |  | WIP15 | April 9, 2019 5:06:14 AM GMT+08:00
CTUDT00622523 | Refresh | Closed | rakic_k | March 20, 2019 1:33:01 AM GMT+08:00 | CIRS Application Installation Errors | WIP14 | WIP15 | manually define a drive F. | Subsystem | CIRS; Verified1 | CIRS.SSRS.4060, CIRS.SSRS.5084, and CIRS.SSRS.5085 | Fixed | richards_b | 0 | 3 | CIRS_SW_Installation | CIRS | NO | Closed | WIP15SubCom | Brilliance 64 |  | WIP15 | April 9, 2019 5:06:14 AM GMT+08:00
CTUDT00622523 | Refresh | Closed | rakic_k | March 20, 2019 1:33:01 AM GMT+08:00 | CIRS Application Installation Errors | WIP14 | WIP15 | manually define a drive F. | Subsystem | CIRS; Verified1 | CIRS.SSRS.4060, CIRS.SSRS.5084, and CIRS.SSRS.5085 | Fixed | richards_b | 0 | 3 | CIRS_SW_Installation | CIRS | NO | Closed | WIP15SubCom | Brilliance Big Bore |  | WIP15 | April 9, 2019 5:06:14 AM GMT+08:00
CTUDT00622523 | Refresh | Closed | rakic_k | March 20, 2019 1:33:01 AM GMT+08:00 | CIRS Application Installation Errors | WIP14 | WIP15 | manually define a drive F. | Subsystem | CIRS; Verified1 | CIRS.SSRS.4060, CIRS.SSRS.5084, and CIRS.SSRS.5085 | Fixed | richards_b | 0 | 3 | CIRS_SW_Installation | CIRS | NO | Closed | WIP15SubCom | Ingenuity 64 |  | WIP15 | April 9, 2019 5:06:14 AM GMT+08:00
CTUDT00622523 | Refresh | Closed | rakic_k | March 20, 2019 1:33:01 AM GMT+08:00 | CIRS Application Installation Errors | WIP14 | WIP15 | manually define a drive F. | Subsystem | CIRS; Verified1 | CIRS.SSRS.4060, CIRS.SSRS.5084, and CIRS.SSRS.5085 | Fixed | richards_b | 0 | 3 | CIRS_SW_Installation | CIRS | NO | Closed | WIP15SubCom | IQon1 |  | WIP15 | April 9, 2019 5:06:14 AM GMT+08:00
CTUDT00622521 | Refresh | Closed | levy_s | March 19, 2019 11:35:19 PM GMT+08:00 | Refresh: Low GUI performance of the system login screen | WIP13 |  |  | Subsystem | Login | NA | Duplicate | kapadia_v | 0 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00622521 | Refresh | Closed | levy_s | March 19, 2019 11:35:19 PM GMT+08:00 | Refresh: Low GUI performance of the system login screen | WIP13 |  |  | Subsystem | Login | NA | Duplicate | kapadia_v | 0 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00622518 | Refresh | Closed | bommiReddy_r | March 19, 2019 9:44:48 PM GMT+08:00 | Refresh: Cardiac: Console not added any conventional coronary result for adult patient when cardiac helical(Non-Gated) exam loaded in ECM/Plan | WIP14 | WIP16 |  | Subsystem | ECM | NA | Fixed | chokkalingam_r | 0 | 3 | Console Software | Console4 | YES | Closed | WIP15SegReq | Ovation |  | WIP16 | April 9, 2019 6:27:04 PM GMT+08:00
CTUDT00622518 | Refresh | Closed | bommiReddy_r | March 19, 2019 9:44:48 PM GMT+08:00 | Refresh: Cardiac: Console not added any conventional coronary result for adult patient when cardiac helical(Non-Gated) exam loaded in ECM/Plan | WIP14 | WIP16 |  | Subsystem | ECM | NA | Fixed | chokkalingam_r | 0 | 3 | Console Software | Console4 | YES | Closed | WIP15SegReq | Brilliance 64 |  | WIP16 | April 9, 2019 6:27:04 PM GMT+08:00
CTUDT00622518 | Refresh | Closed | bommiReddy_r | March 19, 2019 9:44:48 PM GMT+08:00 | Refresh: Cardiac: Console not added any conventional coronary result for adult patient when cardiac helical(Non-Gated) exam loaded in ECM/Plan | WIP14 | WIP16 |  | Subsystem | ECM | NA | Fixed | chokkalingam_r | 0 | 3 | Console Software | Console4 | YES | Closed | WIP15SegReq | Ingenuity 64 |  | WIP16 | April 9, 2019 6:27:04 PM GMT+08:00
CTUDT00622516 | Refresh | Closed | Valeria_Kashtanov | March 19, 2019 9:38:18 PM GMT+08:00 | Bigbore Refresh: Calibrations: Performance: wrong collimation for Head water3 case | BB_WIP4 | WIP17 | N/A | Subsystem | Performance | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | BB_WIP6_SubCom | Ingenuity 64 |  | WIP17 | May 5, 2019 9:44:38 PM GMT+08:00
CTUDT00622516 | Refresh | Closed | Valeria_Kashtanov | March 19, 2019 9:38:18 PM GMT+08:00 | Bigbore Refresh: Calibrations: Performance: wrong collimation for Head water3 case | BB_WIP4 | WIP17 | N/A | Subsystem | Performance | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | BB_WIP6_SubCom | Ingenuity BigBore |  | WIP17 | May 5, 2019 9:44:38 PM GMT+08:00
CTUDT00622509 | Refresh | Closed | Sharmila_Vasireddy | March 19, 2019 8:48:27 PM GMT+08:00 | In BB refresh, in online viewer, the Heart rate is displaying instead of breathe rate | BB_WIP4 | BB_WIP7 |  | Subsystem | Respiratory | Console-SRS.11253,Console-SRS.11162 | Fixed | mY_h | 0 | 2 | Console Software | Console4 | YES | Closed | BB_WIP6_SubCom | Brilliance Big Bore |  | BB_WIP7 | May 16, 2019 12:43:52 PM GMT+08:00
CTUDT00622508 | Refresh | Closed | levy_s | March 19, 2019 8:30:10 PM GMT+08:00 | Refresh: Gantry cannot initialize error message during automation tests | WIP13 | WIP20 | User to press Go button again | Subsystem | Gantry | NA | Fixed | karkour | 0 | 4 | Gantry Software | GCX | YES | Closed | WIP20SubCom | Ovation |  | WIP20 | September 12, 2019 9:38:31 PM GMT+08:00
CTUDT00622507 | Refresh | Closed | jotawar_d | March 19, 2019 8:28:04 PM GMT+08:00 | Refresh/BB Refresh: Surview scan followed by multiple Add-Replace Operations throws bug rep error | 4.1.10 | WIP16 | Select the step (surview) before clicking add/replace button. | Subsystem | Exam Card | NA | Fixed | goswami_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP16SubCom | Ovation |  | WIP16 | April 5, 2019 6:28:12 PM GMT+08:00
CTUDT00622507 | Refresh | Closed | jotawar_d | March 19, 2019 8:28:04 PM GMT+08:00 | Refresh/BB Refresh: Surview scan followed by multiple Add-Replace Operations throws bug rep error | 4.1.10 | WIP16 | Select the step (surview) before clicking add/replace button. | Subsystem | Exam Card | NA | Fixed | goswami_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP16SubCom | Brilliance 64 |  | WIP16 | April 5, 2019 6:28:12 PM GMT+08:00
CTUDT00622507 | Refresh | Closed | jotawar_d | March 19, 2019 8:28:04 PM GMT+08:00 | Refresh/BB Refresh: Surview scan followed by multiple Add-Replace Operations throws bug rep error | 4.1.10 | WIP16 | Select the step (surview) before clicking add/replace button. | Subsystem | Exam Card | NA | Fixed | goswami_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP16SubCom | Ingenuity 64 |  | WIP16 | April 5, 2019 6:28:12 PM GMT+08:00
CTUDT00622507 | Refresh | Closed | jotawar_d | March 19, 2019 8:28:04 PM GMT+08:00 | Refresh/BB Refresh: Surview scan followed by multiple Add-Replace Operations throws bug rep error | 4.1.10 | WIP16 | Select the step (surview) before clicking add/replace button. | Subsystem | Exam Card | NA | Fixed | goswami_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP16SubCom | IQon1 |  | WIP16 | April 5, 2019 6:28:12 PM GMT+08:00
CTUDT00622490 | Refresh | Closed | Valeria_Kashtanov | March 19, 2019 5:26:04 PM GMT+08:00 | Bigbore Refresh: Calibrations: Constancy: Wrong cases in application | BB_WIP4 | WIP17 | N/A | Subsystem | Acceptance/Constancy | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP17SubCom | Brilliance Big Bore |  | WIP17 | May 12, 2019 6:10:48 PM GMT+08:00
CTUDT00622487 | Refresh | Closed | Valeria_Kashtanov | March 19, 2019 5:15:53 PM GMT+08:00 | Bigbore Refresh: Calibrations: Hcor Basic: Body ROIs diameter is not as defined in spec with Harmonized Phantom | BB_WIP4 | WIP15 |  | Subsystem | HCOR calibration | n/a | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | YES | Closed | BB_WIP6_SubCom | Brilliance Big Bore |  | WIP15 | March 27, 2019 3:50:14 PM GMT+08:00
CTUDT00622485 | Refresh | Closed | Valeria_Kashtanov | March 19, 2019 4:59:33 PM GMT+08:00 | Bigbore Refresh: Calibrations: Performance: Infant Body ROIs diameter is not as defined in spec with Harmonized Phantom | BB_WIP4 | WIP17 | N/A | Subsystem | Performance test | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | BB_WIP6_SubCom | Brilliance Big Bore |  | WIP17 | April 23, 2019 6:45:09 PM GMT+08:00
CTUDT00622485 | Refresh | Closed | Valeria_Kashtanov | March 19, 2019 4:59:33 PM GMT+08:00 | Bigbore Refresh: Calibrations: Performance: Infant Body ROIs diameter is not as defined in spec with Harmonized Phantom | BB_WIP4 | WIP17 | N/A | Subsystem | Performance test | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | BB_WIP6_SubCom | Ingenuity 64 |  | WIP17 | April 23, 2019 6:45:09 PM GMT+08:00
CTUDT00622483 | Refresh | Closed | krisher_i | March 19, 2019 4:00:39 PM GMT+08:00 | Exam Card Manager causes Workflow to crash | WIP13 | WIP17 | NA | Subsystem | ECM | Console-SRS.2143 | Fixed | sagar_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP17SubCom | Ovation |  | WIP17 | April 29, 2019 2:40:44 PM GMT+08:00
CTUDT00622483 | Refresh | Closed | krisher_i | March 19, 2019 4:00:39 PM GMT+08:00 | Exam Card Manager causes Workflow to crash | WIP13 | WIP17 | NA | Subsystem | ECM | Console-SRS.2143 | Fixed | sagar_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP17SubCom | Brilliance 64 |  | WIP17 | April 29, 2019 2:40:44 PM GMT+08:00
CTUDT00622483 | Refresh | Closed | krisher_i | March 19, 2019 4:00:39 PM GMT+08:00 | Exam Card Manager causes Workflow to crash | WIP13 | WIP17 | NA | Subsystem | ECM | Console-SRS.2143 | Fixed | sagar_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP17SubCom | Ingenuity 64 |  | WIP17 | April 29, 2019 2:40:44 PM GMT+08:00
CTUDT00622483 | Refresh | Closed | krisher_i | March 19, 2019 4:00:39 PM GMT+08:00 | Exam Card Manager causes Workflow to crash | WIP13 | WIP17 | NA | Subsystem | ECM | Console-SRS.2143 | Fixed | sagar_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP17SubCom | IQon1 |  | WIP17 | April 29, 2019 2:40:44 PM GMT+08:00
CTUDT00622479 | Refresh | Closed | datta_i | March 19, 2019 1:50:53 PM GMT+08:00 | Refresh : The Console does not prevent the user from performing a perfusion scan when the accumulated planned CTDI exceeds 500mGy. | 4.1.10 | WIP21 |  | Subsystem | Perfusion | NA | Fixed | pandey_a | 0 | 2 | Console Software | Console4 | YES | Closed | WIP20SegReq | Ovation |  | WIP21 | October 3, 2019 6:32:19 PM GMT+08:00
CTUDT00622479 | Refresh | Closed | datta_i | March 19, 2019 1:50:53 PM GMT+08:00 | Refresh : The Console does not prevent the user from performing a perfusion scan when the accumulated planned CTDI exceeds 500mGy. | 4.1.10 | WIP21 |  | Subsystem | Perfusion | NA | Fixed | pandey_a | 0 | 2 | Console Software | Console4 | YES | Closed | WIP20SegReq | Brilliance 64 |  | WIP21 | October 3, 2019 6:32:19 PM GMT+08:00
CTUDT00622479 | Refresh | Closed | datta_i | March 19, 2019 1:50:53 PM GMT+08:00 | Refresh : The Console does not prevent the user from performing a perfusion scan when the accumulated planned CTDI exceeds 500mGy. | 4.1.10 | WIP21 |  | Subsystem | Perfusion | NA | Fixed | pandey_a | 0 | 2 | Console Software | Console4 | YES | Closed | WIP20SegReq | Ingenuity 64 |  | WIP21 | October 3, 2019 6:32:19 PM GMT+08:00
CTUDT00622479 | Refresh | Closed | datta_i | March 19, 2019 1:50:53 PM GMT+08:00 | Refresh : The Console does not prevent the user from performing a perfusion scan when the accumulated planned CTDI exceeds 500mGy. | 4.1.10 | WIP21 |  | Subsystem | Perfusion | NA | Fixed | pandey_a | 0 | 2 | Console Software | Console4 | YES | Closed | WIP20SegReq | IQon1 |  | WIP21 | October 3, 2019 6:32:19 PM GMT+08:00
CTUDT00622473 | Refresh | Closed | mN_m | March 19, 2019 11:08:33 AM GMT+08:00 | BBRefresh: Unable to set minimum recon increment to 0.1mm and it is allowing only to set 10% of selected thickness. | BB_WIP4 | BB_WIP6 |  | Subsystem | Plan | NA | Fixed | v_sumanth | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP6_SubCom | Brilliance Big Bore |  | BB_WIP6 | April 8, 2019 8:45:47 PM GMT+08:00
CTUDT00622471 | Refresh | Closed | cherylli | March 19, 2019 10:28:26 AM GMT+08:00 | Refresh Pre-Val: when we check the license in WIP4,only find the 120KW | BB_WIP4 | WIP16 |  | Subsystem | License | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP16SegReq | Ovation |  | WIP16 | May 14, 2019 4:56:42 PM GMT+08:00
CTUDT00622471 | Refresh | Closed | cherylli | March 19, 2019 10:28:26 AM GMT+08:00 | Refresh Pre-Val: when we check the license in WIP4,only find the 120KW | BB_WIP4 | WIP16 |  | Subsystem | License | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP16SegReq | Brilliance 64 |  | WIP16 | May 14, 2019 4:56:42 PM GMT+08:00
CTUDT00622471 | Refresh | Closed | cherylli | March 19, 2019 10:28:26 AM GMT+08:00 | Refresh Pre-Val: when we check the license in WIP4,only find the 120KW | BB_WIP4 | WIP16 |  | Subsystem | License | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP16SegReq | Brilliance Big Bore |  | WIP16 | May 14, 2019 4:56:42 PM GMT+08:00
CTUDT00622471 | Refresh | Closed | cherylli | March 19, 2019 10:28:26 AM GMT+08:00 | Refresh Pre-Val: when we check the license in WIP4,only find the 120KW | BB_WIP4 | WIP16 |  | Subsystem | License | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP16SegReq | Ingenuity 64 |  | WIP16 | May 14, 2019 4:56:42 PM GMT+08:00
CTUDT00622471 | Refresh | Closed | cherylli | March 19, 2019 10:28:26 AM GMT+08:00 | Refresh Pre-Val: when we check the license in WIP4,only find the 120KW | BB_WIP4 | WIP16 |  | Subsystem | License | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP16SegReq | IQon1 |  | WIP16 | May 14, 2019 4:56:42 PM GMT+08:00
CTUDT00622467 | Refresh | Closed | Valeria_Kashtanov | March 19, 2019 3:52:34 AM GMT+08:00 | Refresh: Acceptance/Constancy, Harmonized Phantom: ROIs on Body cases are in wrong position | BB_WIP4 | WIP17 | N/A | Subsystem | Acceptance/Constancy | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | YES | Closed | BB_WIP7_SubCom | Brilliance Big Bore |  | WIP17 | May 21, 2019 10:01:08 PM GMT+08:00
CTUDT00622463 | Refresh | Closed | Valeria_Kashtanov | March 18, 2019 8:47:02 PM GMT+08:00 | Refresh: Calibrations: Beam Quality Tube Potential Test: Peak Potential Test noted as failed even though all cases passed | WIP13 | WIP15 |  | Subsystem | Beam Quality Tube Potential Test | NaN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP17SegReq | Ovation |  | WIP15 | April 7, 2019 2:50:07 PM GMT+08:00
CTUDT00622460 | Refresh | Closed | alex_nihaev | March 18, 2019 3:15:38 PM GMT+08:00 | Refresh:Physics Analysis: No MDU is loaded error after trying to load MDU | BB_WIP4 | N/A | Grant CIRS credentials | Subsystem | Test Design | N/A | Fixed | yosfan_s | 0 | 3 | Test Design | DMS | NO | Closed | WIP17SegReq | Ovation |  | N/A | April 11, 2019 8:46:03 PM GMT+08:00
CTUDT00622460 | Refresh | Closed | alex_nihaev | March 18, 2019 3:15:38 PM GMT+08:00 | Refresh:Physics Analysis: No MDU is loaded error after trying to load MDU | BB_WIP4 | N/A | Grant CIRS credentials | Subsystem | Test Design | N/A | Fixed | yosfan_s | 0 | 3 | Test Design | DMS | NO | Closed | WIP17SegReq | Brilliance Big Bore |  | N/A | April 11, 2019 8:46:03 PM GMT+08:00
CTUDT00622460 | Refresh | Closed | alex_nihaev | March 18, 2019 3:15:38 PM GMT+08:00 | Refresh:Physics Analysis: No MDU is loaded error after trying to load MDU | BB_WIP4 | N/A | Grant CIRS credentials | Subsystem | Test Design | N/A | Fixed | yosfan_s | 0 | 3 | Test Design | DMS | NO | Closed | WIP17SegReq | Ingenuity 64 |  | N/A | April 11, 2019 8:46:03 PM GMT+08:00
CTUDT00622460 | Refresh | Closed | alex_nihaev | March 18, 2019 3:15:38 PM GMT+08:00 | Refresh:Physics Analysis: No MDU is loaded error after trying to load MDU | BB_WIP4 | N/A | Grant CIRS credentials | Subsystem | Test Design | N/A | Fixed | yosfan_s | 0 | 3 | Test Design | DMS | NO | Closed | WIP17SegReq | IQon1 |  | N/A | April 11, 2019 8:46:03 PM GMT+08:00
CTUDT00622455 | Refresh | Closed | Valeria_Kashtanov | March 18, 2019 4:43:46 AM GMT+08:00 | Refresh: Slice Thickness: Slice Width for Surview images is not as defined in spec. | WIP13 | WIP15 | N/A | Subsystem | Slice Thickness | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP15SegReq | Ovation |  | WIP15 | April 2, 2019 5:02:46 PM GMT+08:00
CTUDT00622455 | Refresh | Closed | Valeria_Kashtanov | March 18, 2019 4:43:46 AM GMT+08:00 | Refresh: Slice Thickness: Slice Width for Surview images is not as defined in spec. | WIP13 | WIP15 | N/A | Subsystem | Slice Thickness | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP15SegReq | Brilliance Big Bore |  | WIP15 | April 2, 2019 5:02:46 PM GMT+08:00
CTUDT00622455 | Refresh | Closed | Valeria_Kashtanov | March 18, 2019 4:43:46 AM GMT+08:00 | Refresh: Slice Thickness: Slice Width for Surview images is not as defined in spec. | WIP13 | WIP15 | N/A | Subsystem | Slice Thickness | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP15SegReq | Ingenuity 64 |  | WIP15 | April 2, 2019 5:02:46 PM GMT+08:00
CTUDT00622455 | Refresh | Closed | Valeria_Kashtanov | March 18, 2019 4:43:46 AM GMT+08:00 | Refresh: Slice Thickness: Slice Width for Surview images is not as defined in spec. | WIP13 | WIP15 | N/A | Subsystem | Slice Thickness | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP15SegReq | IQon1 |  | WIP15 | April 2, 2019 5:02:46 PM GMT+08:00
CTUDT00622451 | Refresh | Closed | beck_d | March 17, 2019 8:18:45 PM GMT+08:00 | Refresh: Calibrations: ACR Test: Issues within the report header | WIP13 | WIP17 |  | Subsystem | ACR test | NA | Fixed | Kostya_Baklan | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP16SegReq | Ovation |  | WIP17 | April 21, 2019 2:43:59 PM GMT+08:00
CTUDT00622451 | Refresh | Closed | beck_d | March 17, 2019 8:18:45 PM GMT+08:00 | Refresh: Calibrations: ACR Test: Issues within the report header | WIP13 | WIP17 |  | Subsystem | ACR test | NA | Fixed | Kostya_Baklan | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP16SegReq | Brilliance Big Bore |  | WIP17 | April 21, 2019 2:43:59 PM GMT+08:00
CTUDT00622451 | Refresh | Closed | beck_d | March 17, 2019 8:18:45 PM GMT+08:00 | Refresh: Calibrations: ACR Test: Issues within the report header | WIP13 | WIP17 |  | Subsystem | ACR test | NA | Fixed | Kostya_Baklan | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP16SegReq | Ingenuity 64 |  | WIP17 | April 21, 2019 2:43:59 PM GMT+08:00
CTUDT00622451 | Refresh | Closed | beck_d | March 17, 2019 8:18:45 PM GMT+08:00 | Refresh: Calibrations: ACR Test: Issues within the report header | WIP13 | WIP17 |  | Subsystem | ACR test | NA | Fixed | Kostya_Baklan | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP16SegReq | IQon1 |  | WIP17 | April 21, 2019 2:43:59 PM GMT+08:00
CTUDT00622449 | Refresh | Closed | alex_nihaev | March 17, 2019 7:24:44 PM GMT+08:00 | Refresh, BB, Performance infant test , W value of the body cases is 100, instead of 200 | BB_WIP4 |  |  | Subsystem |  |  | Duplicate | frider_o | 0 |  |  |  |  | Closed |  | Brilliance Big Bore
CTUDT00622445 | Refresh | Closed | dadon_e | March 16, 2019 6:50:11 PM GMT+08:00 | Refresh, Pre-Val (WIP14): ICT-SP, performance test calibration failed - The performance test stuck on offline recon case45 | WIP14 |  |  | Subsystem | Service Tools | N/A | Works As Specified / Not A Defect | yatziv_l | 0 |  |  | DMS | NO | Closed |  | Ovation
CTUDT00622445 | Refresh | Closed | dadon_e | March 16, 2019 6:50:11 PM GMT+08:00 | Refresh, Pre-Val (WIP14): ICT-SP, performance test calibration failed - The performance test stuck on offline recon case45 | WIP14 |  |  | Subsystem | Service Tools | N/A | Works As Specified / Not A Defect | yatziv_l | 0 |  |  | DMS | NO | Closed |  | Ingenuity 64
CTUDT00622445 | Refresh | Closed | dadon_e | March 16, 2019 6:50:11 PM GMT+08:00 | Refresh, Pre-Val (WIP14): ICT-SP, performance test calibration failed - The performance test stuck on offline recon case45 | WIP14 |  |  | Subsystem | Service Tools | N/A | Works As Specified / Not A Defect | yatziv_l | 0 |  |  | DMS | NO | Closed |  | Ingenuity BigBore
CTUDT00622445 | Refresh | Closed | dadon_e | March 16, 2019 6:50:11 PM GMT+08:00 | Refresh, Pre-Val (WIP14): ICT-SP, performance test calibration failed - The performance test stuck on offline recon case45 | WIP14 |  |  | Subsystem | Service Tools | N/A | Works As Specified / Not A Defect | yatziv_l | 0 |  |  | DMS | NO | Closed |  | IQon1
CTUDT00622443 | Refresh | Closed | dadon_e | March 16, 2019 6:30:50 PM GMT+08:00 | Refresh, Pre-Val (WIP14): ICT-SP, rotor balance calibration failed - The calibration failed (cannot open MATLAB app ) | WIP14 | WIP16 | Assuming that the Rotor is approximately balanced, It is possible to perform to run a Legacy Calibration first. For BigBore and Ingenuity, A-Plane Skew is the fastest. For iCT and IQon, Filter Damage Check is the fastest. | Subsystem | Calibrations | CT-NM.SRS.Service.2560 | Fixed | liu_b | 0 | 2 | Gantry Software | GCX | NO | Closed | WIP16SegReq | Brilliance Big Bore |  | WIP16 | April 27, 2019 5:00:39 AM GMT+08:00
CTUDT00622443 | Refresh | Closed | dadon_e | March 16, 2019 6:30:50 PM GMT+08:00 | Refresh, Pre-Val (WIP14): ICT-SP, rotor balance calibration failed - The calibration failed (cannot open MATLAB app ) | WIP14 | WIP16 | Assuming that the Rotor is approximately balanced, It is possible to perform to run a Legacy Calibration first. For BigBore and Ingenuity, A-Plane Skew is the fastest. For iCT and IQon, Filter Damage Check is the fastest. | Subsystem | Calibrations | CT-NM.SRS.Service.2560 | Fixed | liu_b | 0 | 2 | Gantry Software | GCX | NO | Closed | WIP16SegReq | iCT4 |  | WIP16 | April 27, 2019 5:00:39 AM GMT+08:00
CTUDT00622443 | Refresh | Closed | dadon_e | March 16, 2019 6:30:50 PM GMT+08:00 | Refresh, Pre-Val (WIP14): ICT-SP, rotor balance calibration failed - The calibration failed (cannot open MATLAB app ) | WIP14 | WIP16 | Assuming that the Rotor is approximately balanced, It is possible to perform to run a Legacy Calibration first. For BigBore and Ingenuity, A-Plane Skew is the fastest. For iCT and IQon, Filter Damage Check is the fastest. | Subsystem | Calibrations | CT-NM.SRS.Service.2560 | Fixed | liu_b | 0 | 2 | Gantry Software | GCX | NO | Closed | WIP16SegReq | Ingenuity 64 |  | WIP16 | April 27, 2019 5:00:39 AM GMT+08:00
CTUDT00622443 | Refresh | Closed | dadon_e | March 16, 2019 6:30:50 PM GMT+08:00 | Refresh, Pre-Val (WIP14): ICT-SP, rotor balance calibration failed - The calibration failed (cannot open MATLAB app ) | WIP14 | WIP16 | Assuming that the Rotor is approximately balanced, It is possible to perform to run a Legacy Calibration first. For BigBore and Ingenuity, A-Plane Skew is the fastest. For iCT and IQon, Filter Damage Check is the fastest. | Subsystem | Calibrations | CT-NM.SRS.Service.2560 | Fixed | liu_b | 0 | 2 | Gantry Software | GCX | NO | Closed | WIP16SegReq | iCT2 |  | WIP16 | April 27, 2019 5:00:39 AM GMT+08:00
CTUDT00622443 | Refresh | Closed | dadon_e | March 16, 2019 6:30:50 PM GMT+08:00 | Refresh, Pre-Val (WIP14): ICT-SP, rotor balance calibration failed - The calibration failed (cannot open MATLAB app ) | WIP14 | WIP16 | Assuming that the Rotor is approximately balanced, It is possible to perform to run a Legacy Calibration first. For BigBore and Ingenuity, A-Plane Skew is the fastest. For iCT and IQon, Filter Damage Check is the fastest. | Subsystem | Calibrations | CT-NM.SRS.Service.2560 | Fixed | liu_b | 0 | 2 | Gantry Software | GCX | NO | Closed | WIP16SegReq | iCT SDCT |  | WIP16 | April 27, 2019 5:00:39 AM GMT+08:00
CTUDT00622443 | Refresh | Closed | dadon_e | March 16, 2019 6:30:50 PM GMT+08:00 | Refresh, Pre-Val (WIP14): ICT-SP, rotor balance calibration failed - The calibration failed (cannot open MATLAB app ) | WIP14 | WIP16 | Assuming that the Rotor is approximately balanced, It is possible to perform to run a Legacy Calibration first. For BigBore and Ingenuity, A-Plane Skew is the fastest. For iCT and IQon, Filter Damage Check is the fastest. | Subsystem | Calibrations | CT-NM.SRS.Service.2560 | Fixed | liu_b | 0 | 2 | Gantry Software | GCX | NO | Closed | WIP16SegReq | IQon1 |  | WIP16 | April 27, 2019 5:00:39 AM GMT+08:00
CTUDT00622441 | Refresh | Closed | dadon_e | March 16, 2019 6:08:48 PM GMT+08:00 | Refresh, Pre-Val (WIP14): ICT-SP, UHR X position calibration failed - The calibration failed on 100K (not allowed combination) also allowed to clinical user 100k option on UHR | WIP14 | WIP16 | Disable 100 kVp in Customize options until this is fixed. | Subsystem | Calibrations | DHF201223 - section 10.1 | Fixed | zhou_x | 0 | 2 | Gantry Calibrations | GCX | NO | Closed | WIP16SegReq | iCT4 |  | WIP16 | May 31, 2019 1:09:52 PM GMT+08:00
CTUDT00622441 | Refresh | Closed | dadon_e | March 16, 2019 6:08:48 PM GMT+08:00 | Refresh, Pre-Val (WIP14): ICT-SP, UHR X position calibration failed - The calibration failed on 100K (not allowed combination) also allowed to clinical user 100k option on UHR | WIP14 | WIP16 | Disable 100 kVp in Customize options until this is fixed. | Subsystem | Calibrations | DHF201223 - section 10.1 | Fixed | zhou_x | 0 | 2 | Gantry Calibrations | GCX | NO | Closed | WIP16SegReq | iCT2 |  | WIP16 | May 31, 2019 1:09:52 PM GMT+08:00
CTUDT00622441 | Refresh | Closed | dadon_e | March 16, 2019 6:08:48 PM GMT+08:00 | Refresh, Pre-Val (WIP14): ICT-SP, UHR X position calibration failed - The calibration failed on 100K (not allowed combination) also allowed to clinical user 100k option on UHR | WIP14 | WIP16 | Disable 100 kVp in Customize options until this is fixed. | Subsystem | Calibrations | DHF201223 - section 10.1 | Fixed | zhou_x | 0 | 2 | Gantry Calibrations | GCX | NO | Closed | WIP16SegReq | iCT SDCT |  | WIP16 | May 31, 2019 1:09:52 PM GMT+08:00
CTUDT00622433 | Refresh | Closed | Valeria_Kashtanov | March 15, 2019 3:14:56 AM GMT+08:00 | Refresh: Calibrations: Air Calibration: no data is displayed on pre-scan images in CT user | WIP13 | DHF170383 Rev13 |  | Subsystem | Physics Spec | n/a | Fixed | richards_b | 0 | 3 | Design Documentation | Support-Physics | NO | Closed | WIP15SegReq | Ovation |  | DHF170383 Rev13 | May 16, 2019 10:32:09 AM GMT+08:00
CTUDT00622431 | Refresh | Closed | cohen_y | March 14, 2019 11:10:32 PM GMT+08:00 | Refresh: Steps Test: Steps Z Surview length is not as defined in Spec. | BB_WIP4 | N/A |  | Subsystem | DMS | Appendix A of DHF213007 Rev 02 | Fixed | yatziv_l | 0 | 2 | Design Documentation | DMS | NO | Closed | BB_WIP6_SubCom | Brilliance Big Bore |  | N/A | August 25, 2019 3:17:37 PM GMT+08:00
CTUDT00622426 | Refresh | Closed | beck_d | March 14, 2019 7:39:50 PM GMT+08:00 | Refresh: Calibrations: Steps Test: X-ray filter of Steps Z scan and the Air scan is not the same | WIP13 | WIP17 | N/A | Subsystem | Steps test | DHF198870rev 09 ID017 | Fixed | yatziv_l | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP15SegReq | Ovation |  | WIP17 | August 1, 2019 8:30:50 PM GMT+08:00
CTUDT00622423 | Refresh | Closed | 320053770 | March 14, 2019 6:59:22 PM GMT+08:00 | REFRESH WIP 14 for Step and shoot clinical exam card, the clinical scan was paused when triggered by a bolus tracker timer. | WIP14 |  |  | Subsystem | Bolus |  | Duplicate | karkour | 0 |  |  | GCX | YES | Closed |  | Ovation
CTUDT00622408 | Refresh | Closed | divya_n | March 14, 2019 4:49:02 PM GMT+08:00 | Refresh: Film and report page going behind and not able to view. | WIP13 | WIP19 | The user can view the applications by pressing Alt+Tab and select the applications. This issue is a one time occurrence , once the workaround is applied, the issue will not be seen again. | Subsystem | USP | NA | Fixed | s_sa | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP19SubCom | Ovation |  | WIP19 | July 18, 2019 2:16:42 PM GMT+08:00
CTUDT00622408 | Refresh | Closed | divya_n | March 14, 2019 4:49:02 PM GMT+08:00 | Refresh: Film and report page going behind and not able to view. | WIP13 | WIP19 | The user can view the applications by pressing Alt+Tab and select the applications. This issue is a one time occurrence , once the workaround is applied, the issue will not be seen again. | Subsystem | USP | NA | Fixed | s_sa | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP19SubCom | Brilliance 64 |  | WIP19 | July 18, 2019 2:16:42 PM GMT+08:00
CTUDT00622408 | Refresh | Closed | divya_n | March 14, 2019 4:49:02 PM GMT+08:00 | Refresh: Film and report page going behind and not able to view. | WIP13 | WIP19 | The user can view the applications by pressing Alt+Tab and select the applications. This issue is a one time occurrence , once the workaround is applied, the issue will not be seen again. | Subsystem | USP | NA | Fixed | s_sa | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP19SubCom | Brilliance Big Bore |  | WIP19 | July 18, 2019 2:16:42 PM GMT+08:00
CTUDT00622408 | Refresh | Closed | divya_n | March 14, 2019 4:49:02 PM GMT+08:00 | Refresh: Film and report page going behind and not able to view. | WIP13 | WIP19 | The user can view the applications by pressing Alt+Tab and select the applications. This issue is a one time occurrence , once the workaround is applied, the issue will not be seen again. | Subsystem | USP | NA | Fixed | s_sa | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP19SubCom | Ingenuity 64 |  | WIP19 | July 18, 2019 2:16:42 PM GMT+08:00
CTUDT00622408 | Refresh | Closed | divya_n | March 14, 2019 4:49:02 PM GMT+08:00 | Refresh: Film and report page going behind and not able to view. | WIP13 | WIP19 | The user can view the applications by pressing Alt+Tab and select the applications. This issue is a one time occurrence , once the workaround is applied, the issue will not be seen again. | Subsystem | USP | NA | Fixed | s_sa | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP19SubCom | IQon1 |  | WIP19 | July 18, 2019 2:16:42 PM GMT+08:00
CTUDT00622406 | Refresh | Closed | divya_n | March 14, 2019 4:29:56 PM GMT+08:00 | Refresh: Patient Directory: Copy of data containing iDose Radiation information fails | WIP13 |  | attempt the copy operation again. | Subsystem | USP | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | Ovation
CTUDT00622406 | Refresh | Closed | divya_n | March 14, 2019 4:29:56 PM GMT+08:00 | Refresh: Patient Directory: Copy of data containing iDose Radiation information fails | WIP13 |  | attempt the copy operation again. | Subsystem | USP | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00622406 | Refresh | Closed | divya_n | March 14, 2019 4:29:56 PM GMT+08:00 | Refresh: Patient Directory: Copy of data containing iDose Radiation information fails | WIP13 |  | attempt the copy operation again. | Subsystem | USP | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00622406 | Refresh | Closed | divya_n | March 14, 2019 4:29:56 PM GMT+08:00 | Refresh: Patient Directory: Copy of data containing iDose Radiation information fails | WIP13 |  | attempt the copy operation again. | Subsystem | USP | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00622406 | Refresh | Closed | divya_n | March 14, 2019 4:29:56 PM GMT+08:00 | Refresh: Patient Directory: Copy of data containing iDose Radiation information fails | WIP13 |  | attempt the copy operation again. | Subsystem | USP | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | IQon1
CTUDT00622402 | Refresh | Closed | shiraguppi_k | March 14, 2019 3:46:27 PM GMT+08:00 | In Calcium score application when selecting Delete segmented ROI(s)' and "Delete current ROI" from left pane (from drop down) ROIs are not deleting from images. | BB_WIP3 |  |  | Subsystem | EBW | NA | Change Proposal | redrouthu_r | 0 | 4 |  | iCAP | NO | Closed | WIP15SegReq | Ovation
CTUDT00622402 | Refresh | Closed | shiraguppi_k | March 14, 2019 3:46:27 PM GMT+08:00 | In Calcium score application when selecting Delete segmented ROI(s)' and "Delete current ROI" from left pane (from drop down) ROIs are not deleting from images. | BB_WIP3 |  |  | Subsystem | EBW | NA | Change Proposal | redrouthu_r | 0 | 4 |  | iCAP | NO | Closed | WIP15SegReq | Brilliance 64
CTUDT00622402 | Refresh | Closed | shiraguppi_k | March 14, 2019 3:46:27 PM GMT+08:00 | In Calcium score application when selecting Delete segmented ROI(s)' and "Delete current ROI" from left pane (from drop down) ROIs are not deleting from images. | BB_WIP3 |  |  | Subsystem | EBW | NA | Change Proposal | redrouthu_r | 0 | 4 |  | iCAP | NO | Closed | WIP15SegReq | Brilliance Big Bore
CTUDT00622402 | Refresh | Closed | shiraguppi_k | March 14, 2019 3:46:27 PM GMT+08:00 | In Calcium score application when selecting Delete segmented ROI(s)' and "Delete current ROI" from left pane (from drop down) ROIs are not deleting from images. | BB_WIP3 |  |  | Subsystem | EBW | NA | Change Proposal | redrouthu_r | 0 | 4 |  | iCAP | NO | Closed | WIP15SegReq | Ingenuity 64
CTUDT00622402 | Refresh | Closed | shiraguppi_k | March 14, 2019 3:46:27 PM GMT+08:00 | In Calcium score application when selecting Delete segmented ROI(s)' and "Delete current ROI" from left pane (from drop down) ROIs are not deleting from images. | BB_WIP3 |  |  | Subsystem | EBW | NA | Change Proposal | redrouthu_r | 0 | 4 |  | iCAP | NO | Closed | WIP15SegReq | IQon1
CTUDT00622401 | Refresh | Closed | cohen_y | March 14, 2019 3:42:51 PM GMT+08:00 | Refresh: Shadow Test: Shadow per half-tile plot appears instead of Shadow per half-module | BB_WIP4 | WIP15 WIP14 |  | Subsystem | DMS Calibrations | NaN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | BB_WIP6_SubCom | Brilliance Big Bore |  | WIP15 WIP14 | March 31, 2019 3:57:30 PM GMT+08:00
CTUDT00622399 | Refresh | Closed | shiraguppi_k | March 14, 2019 3:28:45 PM GMT+08:00 | In Brain Perfsuion application in Parameters page (i) Series No displaying as 0 instaed of original series no of the study. | WIP13 | WIP16 |  | Subsystem | HOST;Verified1 | NA | Fixed | redrouthu_r | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP15SubCom | Ovation |  | WIP16 | April 30, 2019 6:19:14 PM GMT+08:00
CTUDT00622399 | Refresh | Closed | shiraguppi_k | March 14, 2019 3:28:45 PM GMT+08:00 | In Brain Perfsuion application in Parameters page (i) Series No displaying as 0 instaed of original series no of the study. | WIP13 | WIP16 |  | Subsystem | HOST;Verified1 | NA | Fixed | redrouthu_r | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP15SubCom | Brilliance 64 |  | WIP16 | April 30, 2019 6:19:14 PM GMT+08:00
CTUDT00622399 | Refresh | Closed | shiraguppi_k | March 14, 2019 3:28:45 PM GMT+08:00 | In Brain Perfsuion application in Parameters page (i) Series No displaying as 0 instaed of original series no of the study. | WIP13 | WIP16 |  | Subsystem | HOST;Verified1 | NA | Fixed | redrouthu_r | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP15SubCom | Brilliance Big Bore |  | WIP16 | April 30, 2019 6:19:14 PM GMT+08:00
CTUDT00622399 | Refresh | Closed | shiraguppi_k | March 14, 2019 3:28:45 PM GMT+08:00 | In Brain Perfsuion application in Parameters page (i) Series No displaying as 0 instaed of original series no of the study. | WIP13 | WIP16 |  | Subsystem | HOST;Verified1 | NA | Fixed | redrouthu_r | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP15SubCom | Ingenuity 64 |  | WIP16 | April 30, 2019 6:19:14 PM GMT+08:00
CTUDT00622399 | Refresh | Closed | shiraguppi_k | March 14, 2019 3:28:45 PM GMT+08:00 | In Brain Perfsuion application in Parameters page (i) Series No displaying as 0 instaed of original series no of the study. | WIP13 | WIP16 |  | Subsystem | HOST;Verified1 | NA | Fixed | redrouthu_r | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP15SubCom | IQon1 |  | WIP16 | April 30, 2019 6:19:14 PM GMT+08:00
CTUDT00622398 | Refresh | Closed | bandil_j | March 14, 2019 3:03:20 PM GMT+08:00 | SpectralSBIStoreDevice flag is not saved correctlly with changes in Auto Store devices. | 4.7.5 | WIP15 | for impacted Exam card please unchecked spectral result field , then explicitly check Save SBI with configured devices and then check back Spectral result Field. | Subsystem |  | N/A | Fixed | sheik_c | 0 | 3 | Console Software | Console4 | NO | Closed | WIP15SubCom | IQon1 | 8240184 | WIP15 | March 21, 2019 7:06:34 PM GMT+08:00
CTUDT00622397 | Refresh | Closed | bandil_j | March 14, 2019 3:01:36 PM GMT+08:00 | Thin Slice images are not stored on configured PACS devices. | 4.7.5 | WIP15 | Please ensure that all thin slices are appeared to PD before performing end exam. | Subsystem | Console | N/A | Fixed | sheik_c | 0 | 3 | Console Software | Console4 | NO | Closed | WIP15SubCom | Ovation | 8240184 | WIP15 | March 21, 2019 4:34:48 PM GMT+08:00
CTUDT00622397 | Refresh | Closed | bandil_j | March 14, 2019 3:01:36 PM GMT+08:00 | Thin Slice images are not stored on configured PACS devices. | 4.7.5 | WIP15 | Please ensure that all thin slices are appeared to PD before performing end exam. | Subsystem | Console | N/A | Fixed | sheik_c | 0 | 3 | Console Software | Console4 | NO | Closed | WIP15SubCom | Brilliance 64 | 8240184 | WIP15 | March 21, 2019 4:34:48 PM GMT+08:00
CTUDT00622397 | Refresh | Closed | bandil_j | March 14, 2019 3:01:36 PM GMT+08:00 | Thin Slice images are not stored on configured PACS devices. | 4.7.5 | WIP15 | Please ensure that all thin slices are appeared to PD before performing end exam. | Subsystem | Console | N/A | Fixed | sheik_c | 0 | 3 | Console Software | Console4 | NO | Closed | WIP15SubCom | Ingenuity 64 | 8240184 | WIP15 | March 21, 2019 4:34:48 PM GMT+08:00
CTUDT00622397 | Refresh | Closed | bandil_j | March 14, 2019 3:01:36 PM GMT+08:00 | Thin Slice images are not stored on configured PACS devices. | 4.7.5 | WIP15 | Please ensure that all thin slices are appeared to PD before performing end exam. | Subsystem | Console | N/A | Fixed | sheik_c | 0 | 3 | Console Software | Console4 | NO | Closed | WIP15SubCom | IQon1 | 8240184 | WIP15 | March 21, 2019 4:34:48 PM GMT+08:00
CTUDT00622396 | Refresh | Closed | bandil_j | March 14, 2019 3:00:11 PM GMT+08:00 | System allows to proceed to Plan without complete patient demographics and Application crashed at press of GO | 4.7.5 | WIP15 | Press New Patient/Anonymous or fill in all the details before proceeding with the scan. Do the above step in case the patient info is being pulled from HIS/RIS | System |  | N/A | Fixed | veeraswamy_k | 0 | 4 | Console Software | Console4 | NO | Closed | WIP15SubCom | Ovation | CFR# 9456433 | WIP15 | March 20, 2019 5:47:04 PM GMT+08:00
CTUDT00622396 | Refresh | Closed | bandil_j | March 14, 2019 3:00:11 PM GMT+08:00 | System allows to proceed to Plan without complete patient demographics and Application crashed at press of GO | 4.7.5 | WIP15 | Press New Patient/Anonymous or fill in all the details before proceeding with the scan. Do the above step in case the patient info is being pulled from HIS/RIS | System |  | N/A | Fixed | veeraswamy_k | 0 | 4 | Console Software | Console4 | NO | Closed | WIP15SubCom | Brilliance 64 | CFR# 9456433 | WIP15 | March 20, 2019 5:47:04 PM GMT+08:00
CTUDT00622396 | Refresh | Closed | bandil_j | March 14, 2019 3:00:11 PM GMT+08:00 | System allows to proceed to Plan without complete patient demographics and Application crashed at press of GO | 4.7.5 | WIP15 | Press New Patient/Anonymous or fill in all the details before proceeding with the scan. Do the above step in case the patient info is being pulled from HIS/RIS | System |  | N/A | Fixed | veeraswamy_k | 0 | 4 | Console Software | Console4 | NO | Closed | WIP15SubCom | Ingenuity 64 | CFR# 9456433 | WIP15 | March 20, 2019 5:47:04 PM GMT+08:00
CTUDT00622396 | Refresh | Closed | bandil_j | March 14, 2019 3:00:11 PM GMT+08:00 | System allows to proceed to Plan without complete patient demographics and Application crashed at press of GO | 4.7.5 | WIP15 | Press New Patient/Anonymous or fill in all the details before proceeding with the scan. Do the above step in case the patient info is being pulled from HIS/RIS | System |  | N/A | Fixed | veeraswamy_k | 0 | 4 | Console Software | Console4 | NO | Closed | WIP15SubCom | IQon1 | CFR# 9456433 | WIP15 | March 20, 2019 5:47:04 PM GMT+08:00
CTUDT00622394 | Refresh | Closed | bandil_j | March 14, 2019 2:55:10 PM GMT+08:00 | HB: FoK: Yokote: Exam manager instance is not getting closed even when user has exited the View2 window | 4.7.5 | WIP15 | Modify the Exam Summary preferences .If any one of results,injection,Executed-surview etc option is selected along with the first two then this problem can be avoided | Subsystem | Exam Manager | N/A | Fixed | veeraswamy_k | 0 | 4 | Console Software | Console4 | NO | Closed | WIP15SubCom | Ovation | IQon HB 6003368 | WIP15 | March 20, 2019 5:34:52 PM GMT+08:00
CTUDT00622394 | Refresh | Closed | bandil_j | March 14, 2019 2:55:10 PM GMT+08:00 | HB: FoK: Yokote: Exam manager instance is not getting closed even when user has exited the View2 window | 4.7.5 | WIP15 | Modify the Exam Summary preferences .If any one of results,injection,Executed-surview etc option is selected along with the first two then this problem can be avoided | Subsystem | Exam Manager | N/A | Fixed | veeraswamy_k | 0 | 4 | Console Software | Console4 | NO | Closed | WIP15SubCom | Brilliance 64 | IQon HB 6003368 | WIP15 | March 20, 2019 5:34:52 PM GMT+08:00
CTUDT00622394 | Refresh | Closed | bandil_j | March 14, 2019 2:55:10 PM GMT+08:00 | HB: FoK: Yokote: Exam manager instance is not getting closed even when user has exited the View2 window | 4.7.5 | WIP15 | Modify the Exam Summary preferences .If any one of results,injection,Executed-surview etc option is selected along with the first two then this problem can be avoided | Subsystem | Exam Manager | N/A | Fixed | veeraswamy_k | 0 | 4 | Console Software | Console4 | NO | Closed | WIP15SubCom | Ingenuity 64 | IQon HB 6003368 | WIP15 | March 20, 2019 5:34:52 PM GMT+08:00
CTUDT00622394 | Refresh | Closed | bandil_j | March 14, 2019 2:55:10 PM GMT+08:00 | HB: FoK: Yokote: Exam manager instance is not getting closed even when user has exited the View2 window | 4.7.5 | WIP15 | Modify the Exam Summary preferences .If any one of results,injection,Executed-surview etc option is selected along with the first two then this problem can be avoided | Subsystem | Exam Manager | N/A | Fixed | veeraswamy_k | 0 | 4 | Console Software | Console4 | NO | Closed | WIP15SubCom | IQon1 | IQon HB 6003368 | WIP15 | March 20, 2019 5:34:52 PM GMT+08:00
CTUDT00622392 | Refresh | Closed | manjunath_mk | March 14, 2019 2:44:07 PM GMT+08:00 | Clear All option does not clear the entries from the drop down in the Patient window | BB_WIP4 |  | Once user press tab or click outside (loses focus) the combobox, all items will be cleared | Subsystem | Patient Entry | NA | Works As Specified / Not A Defect | Prakash_Chand_Yadav | 0 | 3 |  | Console4 | NO | Closed | WIP16SubCom | Ovation
CTUDT00622392 | Refresh | Closed | manjunath_mk | March 14, 2019 2:44:07 PM GMT+08:00 | Clear All option does not clear the entries from the drop down in the Patient window | BB_WIP4 |  | Once user press tab or click outside (loses focus) the combobox, all items will be cleared | Subsystem | Patient Entry | NA | Works As Specified / Not A Defect | Prakash_Chand_Yadav | 0 | 3 |  | Console4 | NO | Closed | WIP16SubCom | Brilliance 64
CTUDT00622392 | Refresh | Closed | manjunath_mk | March 14, 2019 2:44:07 PM GMT+08:00 | Clear All option does not clear the entries from the drop down in the Patient window | BB_WIP4 |  | Once user press tab or click outside (loses focus) the combobox, all items will be cleared | Subsystem | Patient Entry | NA | Works As Specified / Not A Defect | Prakash_Chand_Yadav | 0 | 3 |  | Console4 | NO | Closed | WIP16SubCom | Brilliance Big Bore
CTUDT00622392 | Refresh | Closed | manjunath_mk | March 14, 2019 2:44:07 PM GMT+08:00 | Clear All option does not clear the entries from the drop down in the Patient window | BB_WIP4 |  | Once user press tab or click outside (loses focus) the combobox, all items will be cleared | Subsystem | Patient Entry | NA | Works As Specified / Not A Defect | Prakash_Chand_Yadav | 0 | 3 |  | Console4 | NO | Closed | WIP16SubCom | Ingenuity 64
CTUDT00622392 | Refresh | Closed | manjunath_mk | March 14, 2019 2:44:07 PM GMT+08:00 | Clear All option does not clear the entries from the drop down in the Patient window | BB_WIP4 |  | Once user press tab or click outside (loses focus) the combobox, all items will be cleared | Subsystem | Patient Entry | NA | Works As Specified / Not A Defect | Prakash_Chand_Yadav | 0 | 3 |  | Console4 | NO | Closed | WIP16SubCom | IQon1
CTUDT00622389 | Refresh | Closed | shekhar_p | March 14, 2019 2:34:36 PM GMT+08:00 | When Host -CIRS connection disconnected and the host is unable to remove the ePHI from MDU not displaying message properly | WIP13 | DHF318051 Rev 14 HOST Verification Test Protocol |  | Subsystem | Test Design | N/A | Fixed | ramesh_t | 0 | 3 | Test Design | Console4 | NO | Closed | WIP15SegReq | Ovation |  | DHF318051 Rev 14 HOST Verification Test Protocol | April 26, 2019 7:53:47 PM GMT+08:00
CTUDT00622389 | Refresh | Closed | shekhar_p | March 14, 2019 2:34:36 PM GMT+08:00 | When Host -CIRS connection disconnected and the host is unable to remove the ePHI from MDU not displaying message properly | WIP13 | DHF318051 Rev 14 HOST Verification Test Protocol |  | Subsystem | Test Design | N/A | Fixed | ramesh_t | 0 | 3 | Test Design | Console4 | NO | Closed | WIP15SegReq | Brilliance 64 |  | DHF318051 Rev 14 HOST Verification Test Protocol | April 26, 2019 7:53:47 PM GMT+08:00
CTUDT00622389 | Refresh | Closed | shekhar_p | March 14, 2019 2:34:36 PM GMT+08:00 | When Host -CIRS connection disconnected and the host is unable to remove the ePHI from MDU not displaying message properly | WIP13 | DHF318051 Rev 14 HOST Verification Test Protocol |  | Subsystem | Test Design | N/A | Fixed | ramesh_t | 0 | 3 | Test Design | Console4 | NO | Closed | WIP15SegReq | Ingenuity 64 |  | DHF318051 Rev 14 HOST Verification Test Protocol | April 26, 2019 7:53:47 PM GMT+08:00
CTUDT00622389 | Refresh | Closed | shekhar_p | March 14, 2019 2:34:36 PM GMT+08:00 | When Host -CIRS connection disconnected and the host is unable to remove the ePHI from MDU not displaying message properly | WIP13 | DHF318051 Rev 14 HOST Verification Test Protocol |  | Subsystem | Test Design | N/A | Fixed | ramesh_t | 0 | 3 | Test Design | Console4 | NO | Closed | WIP15SegReq | IQon1 |  | DHF318051 Rev 14 HOST Verification Test Protocol | April 26, 2019 7:53:47 PM GMT+08:00
CTUDT00622388 | Refresh | Closed | manjunath_mk | March 14, 2019 2:16:07 PM GMT+08:00 | Edit Phases window gets hidden behind the Show all window in View 2 while performing cardiac study | BB_WIP4 | BB_WIP6 |  | Subsystem | View2 | NA | Fixed | chokkalingam_r | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP6_SubCom | Brilliance Big Bore |  | BB_WIP6 | April 8, 2019 3:10:54 PM GMT+08:00
CTUDT00622384 | Refresh | Closed | cohen_y | March 14, 2019 1:10:59 AM GMT+08:00 | Refresh: Constancy Test: Harmonized: Slice Thickness case 2 ROIs are not placed on the images due to Baseline do not match error. | WIP13 |  | Repeat the test | Subsystem | DMS Calibrations | N/A | Works As Specified / Not A Defect | yosfan_s | 0 | 3 |  | DMS | NO | Closed | SegReq | Ovation
CTUDT00622382 | Refresh | Closed | cohen_y | March 14, 2019 12:34:16 AM GMT+08:00 | Refresh: Acceptance/Constancy Tests: Harmonized: Slice Thickness case 2 ROIs are not placed in the correct slice according to Spec. | WIP13 | DHF170395, Rev 09 |  | Subsystem |  | DHF170395 | Fixed | morton_t | 0 | 2 | Design Documentation | Support-Physics | NO | Closed | WIP16SegReq | Ovation |  | DHF170395, Rev 09 | May 1, 2019 3:41:40 AM GMT+08:00
CTUDT00622382 | Refresh | Closed | cohen_y | March 14, 2019 12:34:16 AM GMT+08:00 | Refresh: Acceptance/Constancy Tests: Harmonized: Slice Thickness case 2 ROIs are not placed in the correct slice according to Spec. | WIP13 | DHF170395, Rev 09 |  | Subsystem |  | DHF170395 | Fixed | morton_t | 0 | 2 | Design Documentation | Support-Physics | NO | Closed | WIP16SegReq | Brilliance Big Bore |  | DHF170395, Rev 09 | May 1, 2019 3:41:40 AM GMT+08:00
CTUDT00622382 | Refresh | Closed | cohen_y | March 14, 2019 12:34:16 AM GMT+08:00 | Refresh: Acceptance/Constancy Tests: Harmonized: Slice Thickness case 2 ROIs are not placed in the correct slice according to Spec. | WIP13 | DHF170395, Rev 09 |  | Subsystem |  | DHF170395 | Fixed | morton_t | 0 | 2 | Design Documentation | Support-Physics | NO | Closed | WIP16SegReq | Ingenuity 64 |  | DHF170395, Rev 09 | May 1, 2019 3:41:40 AM GMT+08:00
CTUDT00622382 | Refresh | Closed | cohen_y | March 14, 2019 12:34:16 AM GMT+08:00 | Refresh: Acceptance/Constancy Tests: Harmonized: Slice Thickness case 2 ROIs are not placed in the correct slice according to Spec. | WIP13 | DHF170395, Rev 09 |  | Subsystem |  | DHF170395 | Fixed | morton_t | 0 | 2 | Design Documentation | Support-Physics | NO | Closed | WIP16SegReq | IQon1 |  | DHF170395, Rev 09 | May 1, 2019 3:41:40 AM GMT+08:00
CTUDT00622380 | Refresh | Closed | cohen_y | March 13, 2019 11:53:32 PM GMT+08:00 | Refresh: Acceptance/Constancy Test: Case 4 (Body 80kV 128x0.625) with Harmonized Phantom failed due to missing Phantom table | WIP13 | WIP15 |  | Subsystem | DMS Calibrations | NaN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP15SegReq | Ovation |  | WIP15 | March 27, 2019 1:16:31 AM GMT+08:00
CTUDT00622377 | Refresh | Closed | cohen_y | March 13, 2019 10:31:08 PM GMT+08:00 | Refresh: Performance Test: IMR stages fail during initialization due to missing Axial Body 3D physics Exam Card | WIP14 | WIP18 |  | System | Exam Cards | n/a | Fixed | sowmya_k | 0 | 2 | DMS Calibrations | Console4 | NO | Closed | WIP16SubCom | Ingenuity 64 |  | WIP18 | June 10, 2019 1:49:58 PM GMT+08:00
CTUDT00622375 | Refresh | Closed | shterner_a | March 13, 2019 9:43:27 PM GMT+08:00 | Refresh Pre-Val: BugRep and system error observed after end exam | WIP13 |  |  | Subsystem | EA Crash | N/A | Works As Specified / Not A Defect | tamari_o | 1 | 4 |  | Support-System | NO | Closed | TBR | Ovation
CTUDT00622375 | Refresh | Closed | shterner_a | March 13, 2019 9:43:27 PM GMT+08:00 | Refresh Pre-Val: BugRep and system error observed after end exam | WIP13 |  |  | Subsystem | EA Crash | N/A | Works As Specified / Not A Defect | tamari_o | 1 | 4 |  | Support-System | NO | Closed | TBR | Brilliance 64
CTUDT00622375 | Refresh | Closed | shterner_a | March 13, 2019 9:43:27 PM GMT+08:00 | Refresh Pre-Val: BugRep and system error observed after end exam | WIP13 |  |  | Subsystem | EA Crash | N/A | Works As Specified / Not A Defect | tamari_o | 1 | 4 |  | Support-System | NO | Closed | TBR | Brilliance Big Bore
CTUDT00622375 | Refresh | Closed | shterner_a | March 13, 2019 9:43:27 PM GMT+08:00 | Refresh Pre-Val: BugRep and system error observed after end exam | WIP13 |  |  | Subsystem | EA Crash | N/A | Works As Specified / Not A Defect | tamari_o | 1 | 4 |  | Support-System | NO | Closed | TBR | Ingenuity 64
CTUDT00622375 | Refresh | Closed | shterner_a | March 13, 2019 9:43:27 PM GMT+08:00 | Refresh Pre-Val: BugRep and system error observed after end exam | WIP13 |  |  | Subsystem | EA Crash | N/A | Works As Specified / Not A Defect | tamari_o | 1 | 4 |  | Support-System | NO | Closed | TBR | IQon1
CTUDT00622372 | Refresh | Closed | shterner_a | March 13, 2019 9:32:32 PM GMT+08:00 | Refresh Pre-Val: Chest High Resolution Axial Supine and Prone images are having artifact and noisy images | WIP13 |  |  | Subsystem | Image Quality |  | Works As Specified / Not A Defect | richards_b | 0 |  |  | CIRS | NO | Closed |  | IQon1
CTUDT00622368 | Refresh | Closed | alex_nihaev | March 13, 2019 8:07:36 PM GMT+08:00 | Refresh,BB, Reasons User Access, the chosen procedures are not as defined in the spec. | BB_WIP4 | WIP15 |  | Subsystem | DMS Calibrations; Blocker | n/a | Fixed | Kostya_Baklan | 0 | 3 | DMS Calibrations | DMS | NO | Closed | BB_WIP6_SubCom | Ovation |  | WIP15 | March 28, 2019 4:17:27 PM GMT+08:00
CTUDT00622368 | Refresh | Closed | alex_nihaev | March 13, 2019 8:07:36 PM GMT+08:00 | Refresh,BB, Reasons User Access, the chosen procedures are not as defined in the spec. | BB_WIP4 | WIP15 |  | Subsystem | DMS Calibrations; Blocker | n/a | Fixed | Kostya_Baklan | 0 | 3 | DMS Calibrations | DMS | NO | Closed | BB_WIP6_SubCom | Brilliance Big Bore |  | WIP15 | March 28, 2019 4:17:27 PM GMT+08:00
CTUDT00622362 | Refresh | Closed | Sharmila_Vasireddy | March 13, 2019 4:48:57 PM GMT+08:00 | IRS_STAGE_WARP should not exist in the MDU files for axial 2D EC result | BB_WIP4 | BB_WIP6 |  | System | MDU | Console-SRS.17096 | Fixed | chokkalingam_r | 0 | 2 | Console Software | Console4 | YES | Closed | BB_WIP6_SubCom | Brilliance Big Bore |  | BB_WIP6 | April 24, 2019 1:33:53 PM GMT+08:00
CTUDT00622360 | Refresh | Closed | wang_l | March 13, 2019 3:08:37 PM GMT+08:00 | SFTP connection failed | WIP14 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | wang_l | 0 |  |  |  |  | Closed |  | Ingenuity 64
CTUDT00622352 | Refresh | Closed | wang_l | March 13, 2019 12:18:54 PM GMT+08:00 | The calibration items in Quality Assurance do not match with SRS requirement CT-NM.SRS.Service.2561 | WIP14 | WIP19 |  | Subsystem | Quality Assurance | Service.2561 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | BB_WIP8_SubCom | Brilliance 64 |  | WIP19 | July 18, 2019 5:01:58 PM GMT+08:00
CTUDT00622352 | Refresh | Closed | wang_l | March 13, 2019 12:18:54 PM GMT+08:00 | The calibration items in Quality Assurance do not match with SRS requirement CT-NM.SRS.Service.2561 | WIP14 | WIP19 |  | Subsystem | Quality Assurance | Service.2561 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | BB_WIP8_SubCom | Ingenuity 64 |  | WIP19 | July 18, 2019 5:01:58 PM GMT+08:00
CTUDT00622350 | Refresh | Closed | wang_l | March 13, 2019 12:06:32 PM GMT+08:00 | Properties value are not retrieved in PRS portal after remote to host | WIP13 | WIP19 |  | System | Remote connection;Blocker19;Verified1 | STPS-SSRS-273 | Fixed | Mark_Huang | 0 | 2 | Gantry Software | GCX | NO | Closed | WIP18SegReq | Ovation |  | WIP19 | August 1, 2019 10:10:17 AM GMT+08:00
CTUDT00622350 | Refresh | Closed | wang_l | March 13, 2019 12:06:32 PM GMT+08:00 | Properties value are not retrieved in PRS portal after remote to host | WIP13 | WIP19 |  | System | Remote connection;Blocker19;Verified1 | STPS-SSRS-273 | Fixed | Mark_Huang | 0 | 2 | Gantry Software | GCX | NO | Closed | WIP18SegReq | Brilliance 64 |  | WIP19 | August 1, 2019 10:10:17 AM GMT+08:00
CTUDT00622350 | Refresh | Closed | wang_l | March 13, 2019 12:06:32 PM GMT+08:00 | Properties value are not retrieved in PRS portal after remote to host | WIP13 | WIP19 |  | System | Remote connection;Blocker19;Verified1 | STPS-SSRS-273 | Fixed | Mark_Huang | 0 | 2 | Gantry Software | GCX | NO | Closed | WIP18SegReq | Brilliance Big Bore |  | WIP19 | August 1, 2019 10:10:17 AM GMT+08:00
CTUDT00622350 | Refresh | Closed | wang_l | March 13, 2019 12:06:32 PM GMT+08:00 | Properties value are not retrieved in PRS portal after remote to host | WIP13 | WIP19 |  | System | Remote connection;Blocker19;Verified1 | STPS-SSRS-273 | Fixed | Mark_Huang | 0 | 2 | Gantry Software | GCX | NO | Closed | WIP18SegReq | Ingenuity 64 |  | WIP19 | August 1, 2019 10:10:17 AM GMT+08:00
CTUDT00622350 | Refresh | Closed | wang_l | March 13, 2019 12:06:32 PM GMT+08:00 | Properties value are not retrieved in PRS portal after remote to host | WIP13 | WIP19 |  | System | Remote connection;Blocker19;Verified1 | STPS-SSRS-273 | Fixed | Mark_Huang | 0 | 2 | Gantry Software | GCX | NO | Closed | WIP18SegReq | IQon1 |  | WIP19 | August 1, 2019 10:10:17 AM GMT+08:00
CTUDT00622346 | Refresh | Closed | Noy_Seelig | March 13, 2019 3:53:33 AM GMT+08:00 | Refresh: Bigbore: In Hcor Slice Adult and Infant customize windows there are no voltages | BB_WIP4 |  |  | Subsystem |  |  | Duplicate | Noy_Seelig | 0 |  |  |  |  | Closed |  | Ingenuity BigBore
CTUDT00622343 | Refresh | Closed | Noy_Seelig | March 13, 2019 2:19:13 AM GMT+08:00 | Refresh: Bigbore: No Beam and DMS Check test in O-level user | BB_WIP4 | WIP15 |  | Subsystem | Service Tools; Verified1 | NA | Fixed | Kostya_Baklan | 0 | 2 | DMS Calibrations | DMS | NO | Closed | BB_WIP6_SubCom | Brilliance Big Bore |  | WIP15 | April 1, 2019 4:06:49 PM GMT+08:00
CTUDT00622339 | Refresh | Closed | krisher_i | March 12, 2019 11:06:13 PM GMT+08:00 | Low GUI performance in Patient scene loading | WIP14 |  |  | Subsystem |  |  | Duplicate | sanjeeva_a | 0 |  |  |  |  | Closed |  | IQon1
CTUDT00622332 | Refresh | Closed | Revital_Cohen | March 12, 2019 8:08:05 PM GMT+08:00 | WIP14 Refresh VDD test failed due to mismatch CIRS version details | WIP14 | WIP15 | NA | Subsystem | Verified1 | CIRS.SSRS.4060, CIRS.SSRS.5084, and CIRS.SSRS.5044 | Fixed | richards_b | 0 | 3 | CIRS_SW_Installation | CIRS | NO | Closed | WIP15SubCom | Ovation |  | WIP15 | April 9, 2019 5:05:35 AM GMT+08:00
CTUDT00622332 | Refresh | Closed | Revital_Cohen | March 12, 2019 8:08:05 PM GMT+08:00 | WIP14 Refresh VDD test failed due to mismatch CIRS version details | WIP14 | WIP15 | NA | Subsystem | Verified1 | CIRS.SSRS.4060, CIRS.SSRS.5084, and CIRS.SSRS.5044 | Fixed | richards_b | 0 | 3 | CIRS_SW_Installation | CIRS | NO | Closed | WIP15SubCom | Brilliance 64 |  | WIP15 | April 9, 2019 5:05:35 AM GMT+08:00
CTUDT00622332 | Refresh | Closed | Revital_Cohen | March 12, 2019 8:08:05 PM GMT+08:00 | WIP14 Refresh VDD test failed due to mismatch CIRS version details | WIP14 | WIP15 | NA | Subsystem | Verified1 | CIRS.SSRS.4060, CIRS.SSRS.5084, and CIRS.SSRS.5044 | Fixed | richards_b | 0 | 3 | CIRS_SW_Installation | CIRS | NO | Closed | WIP15SubCom | Brilliance Big Bore |  | WIP15 | April 9, 2019 5:05:35 AM GMT+08:00
CTUDT00622332 | Refresh | Closed | Revital_Cohen | March 12, 2019 8:08:05 PM GMT+08:00 | WIP14 Refresh VDD test failed due to mismatch CIRS version details | WIP14 | WIP15 | NA | Subsystem | Verified1 | CIRS.SSRS.4060, CIRS.SSRS.5084, and CIRS.SSRS.5044 | Fixed | richards_b | 0 | 3 | CIRS_SW_Installation | CIRS | NO | Closed | WIP15SubCom | Ingenuity 64 |  | WIP15 | April 9, 2019 5:05:35 AM GMT+08:00
CTUDT00622332 | Refresh | Closed | Revital_Cohen | March 12, 2019 8:08:05 PM GMT+08:00 | WIP14 Refresh VDD test failed due to mismatch CIRS version details | WIP14 | WIP15 | NA | Subsystem | Verified1 | CIRS.SSRS.4060, CIRS.SSRS.5084, and CIRS.SSRS.5044 | Fixed | richards_b | 0 | 3 | CIRS_SW_Installation | CIRS | NO | Closed | WIP15SubCom | IQon1 |  | WIP15 | April 9, 2019 5:05:35 AM GMT+08:00
CTUDT00622330 | Refresh | Closed | levy_s | March 12, 2019 7:41:14 PM GMT+08:00 | Refresh: The system total recon time for trauma protocol takes more than 12 minutes | WIP13 |  |  | Subsystem | Test Design | CT-NM.SRS.SysPerf.702 | Works As Specified / Not A Defect | tamari_o | 1 | 4 |  | Support-System | NO | Closed | WIP18SegReq | Ovation
CTUDT00622330 | Refresh | Closed | levy_s | March 12, 2019 7:41:14 PM GMT+08:00 | Refresh: The system total recon time for trauma protocol takes more than 12 minutes | WIP13 |  |  | Subsystem | Test Design | CT-NM.SRS.SysPerf.702 | Works As Specified / Not A Defect | tamari_o | 1 | 4 |  | Support-System | NO | Closed | WIP18SegReq | Brilliance 64
CTUDT00622330 | Refresh | Closed | levy_s | March 12, 2019 7:41:14 PM GMT+08:00 | Refresh: The system total recon time for trauma protocol takes more than 12 minutes | WIP13 |  |  | Subsystem | Test Design | CT-NM.SRS.SysPerf.702 | Works As Specified / Not A Defect | tamari_o | 1 | 4 |  | Support-System | NO | Closed | WIP18SegReq | Brilliance Big Bore
CTUDT00622330 | Refresh | Closed | levy_s | March 12, 2019 7:41:14 PM GMT+08:00 | Refresh: The system total recon time for trauma protocol takes more than 12 minutes | WIP13 |  |  | Subsystem | Test Design | CT-NM.SRS.SysPerf.702 | Works As Specified / Not A Defect | tamari_o | 1 | 4 |  | Support-System | NO | Closed | WIP18SegReq | Ingenuity 64
CTUDT00622330 | Refresh | Closed | levy_s | March 12, 2019 7:41:14 PM GMT+08:00 | Refresh: The system total recon time for trauma protocol takes more than 12 minutes | WIP13 |  |  | Subsystem | Test Design | CT-NM.SRS.SysPerf.702 | Works As Specified / Not A Defect | tamari_o | 1 | 4 |  | Support-System | NO | Closed | WIP18SegReq | IQon1
CTUDT00622311 | Refresh | Closed | Valeria_Kashtanov | March 11, 2019 12:20:29 AM GMT+08:00 | Refresh: Calibrations: Performance Test: Cardiac Body 1 Noise limits are not according to spec | WIP13 | DHF170396 Rev 18 |  | Subsystem | Performance Test | n/a | Fixed | richards_b | 0 | 2 | Design Documentation | Support-Physics | NO | Closed | WIP15SegReq | Ovation |  | DHF170396 Rev 18 | April 27, 2019 4:34:57 AM GMT+08:00
CTUDT00622309 | Refresh | Closed | Valeria_Kashtanov | March 10, 2019 11:14:41 PM GMT+08:00 | Refresh: Calibrations: Performance: Body CT11 Water case appears in CT Head results instead of Body results | WIP13 | WIP15 |  | Subsystem |  | N/A | Fixed | Kostya_Baklan | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP15SubCom | Ovation |  | WIP15 | March 26, 2019 3:45:14 PM GMT+08:00
CTUDT00622307 | Refresh | Closed | Valeria_Kashtanov | March 10, 2019 10:24:07 PM GMT+08:00 | Refresh: Calibrations: Performance Test: Water ROI diameter in Linearity pins cases is not as defined in spec | WIP13 | WIP15 | N/A | Subsystem | DMS Calibrations; Verified1 | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP15SubCom | Ovation |  | WIP15 | March 25, 2019 9:10:19 PM GMT+08:00
CTUDT00622307 | Refresh | Closed | Valeria_Kashtanov | March 10, 2019 10:24:07 PM GMT+08:00 | Refresh: Calibrations: Performance Test: Water ROI diameter in Linearity pins cases is not as defined in spec | WIP13 | WIP15 | N/A | Subsystem | DMS Calibrations; Verified1 | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP15SubCom | Brilliance 64 |  | WIP15 | March 25, 2019 9:10:19 PM GMT+08:00
CTUDT00622307 | Refresh | Closed | Valeria_Kashtanov | March 10, 2019 10:24:07 PM GMT+08:00 | Refresh: Calibrations: Performance Test: Water ROI diameter in Linearity pins cases is not as defined in spec | WIP13 | WIP15 | N/A | Subsystem | DMS Calibrations; Verified1 | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP15SubCom | Ingenuity 64 |  | WIP15 | March 25, 2019 9:10:19 PM GMT+08:00
CTUDT00622307 | Refresh | Closed | Valeria_Kashtanov | March 10, 2019 10:24:07 PM GMT+08:00 | Refresh: Calibrations: Performance Test: Water ROI diameter in Linearity pins cases is not as defined in spec | WIP13 | WIP15 | N/A | Subsystem | DMS Calibrations; Verified1 | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP15SubCom | IQon1 |  | WIP15 | March 25, 2019 9:10:19 PM GMT+08:00
CTUDT00622304 | Refresh | Closed | alex_nihaev | March 10, 2019 7:52:47 PM GMT+08:00 | Refresh-BB-DMS/DataPath Diagnostics-The fiber optic cable disconnection is not identified by Diagnostics. | BB_WIP3 | WIP15 |  | Subsystem | DMS Diagnostics | DMS-SRS-658 | Fixed | franco_a | 0 | 2 | DMS Software | DMS | NO | Closed | WIP15SubCom | Brilliance 64 |  | WIP15 | March 25, 2019 9:14:13 PM GMT+08:00
CTUDT00622304 | Refresh | Closed | alex_nihaev | March 10, 2019 7:52:47 PM GMT+08:00 | Refresh-BB-DMS/DataPath Diagnostics-The fiber optic cable disconnection is not identified by Diagnostics. | BB_WIP3 | WIP15 |  | Subsystem | DMS Diagnostics | DMS-SRS-658 | Fixed | franco_a | 0 | 2 | DMS Software | DMS | NO | Closed | WIP15SubCom | Brilliance Big Bore |  | WIP15 | March 25, 2019 9:14:13 PM GMT+08:00
CTUDT00622304 | Refresh | Closed | alex_nihaev | March 10, 2019 7:52:47 PM GMT+08:00 | Refresh-BB-DMS/DataPath Diagnostics-The fiber optic cable disconnection is not identified by Diagnostics. | BB_WIP3 | WIP15 |  | Subsystem | DMS Diagnostics | DMS-SRS-658 | Fixed | franco_a | 0 | 2 | DMS Software | DMS | NO | Closed | WIP15SubCom | Ingenuity 64 |  | WIP15 | March 25, 2019 9:14:13 PM GMT+08:00
CTUDT00622301 | Refresh | Closed | gushansky_m | March 10, 2019 2:39:04 PM GMT+08:00 | Remote Service IQON:Registration on PRS portal over internet connection failed | WIP13 |  |  | Subsystem | Remote Service | STPS-SSRS-1067 | Invalid/Uncontrolled Configuration | ranjithkumar_m | 0 | 2 |  | Service Tools | NO | Closed | WIP16SubCom | IQon1
CTUDT00622300 | Refresh | Closed | richards_b | March 9, 2019 3:41:52 AM GMT+08:00 | Wedge Order Check, collimator diagnostic tool, failed | WIP14 |  | The FSE can perform a visual inspection of the wedge order | Subsystem |  |  | Duplicate | richards_b | 0 | 3 |  | Support-Physics | NO | Closed |  | IQon1 | PR6174664
CTUDT00622297 | Refresh | Closed | babu_s | March 8, 2019 8:12:12 PM GMT+08:00 | Not getting E-STOP in Simulator environment | WIP14 | WIP15 |  | Subsystem | Simulator | NA | Fixed | prasad_r | 0 | 4 | Service Tools_Software | Service Tools | NO | Closed | WIP15SubCom | Ovation |  | WIP15 | March 27, 2019 6:54:57 PM GMT+08:00
CTUDT00622297 | Refresh | Closed | babu_s | March 8, 2019 8:12:12 PM GMT+08:00 | Not getting E-STOP in Simulator environment | WIP14 | WIP15 |  | Subsystem | Simulator | NA | Fixed | prasad_r | 0 | 4 | Service Tools_Software | Service Tools | NO | Closed | WIP15SubCom | Brilliance 64 |  | WIP15 | March 27, 2019 6:54:57 PM GMT+08:00
CTUDT00622297 | Refresh | Closed | babu_s | March 8, 2019 8:12:12 PM GMT+08:00 | Not getting E-STOP in Simulator environment | WIP14 | WIP15 |  | Subsystem | Simulator | NA | Fixed | prasad_r | 0 | 4 | Service Tools_Software | Service Tools | NO | Closed | WIP15SubCom | Ingenuity 64 |  | WIP15 | March 27, 2019 6:54:57 PM GMT+08:00
CTUDT00622297 | Refresh | Closed | babu_s | March 8, 2019 8:12:12 PM GMT+08:00 | Not getting E-STOP in Simulator environment | WIP14 | WIP15 |  | Subsystem | Simulator | NA | Fixed | prasad_r | 0 | 4 | Service Tools_Software | Service Tools | NO | Closed | WIP15SubCom | IQon1 |  | WIP15 | March 27, 2019 6:54:57 PM GMT+08:00
CTUDT00622294 | Refresh | Closed | mistry_h | March 8, 2019 5:01:04 PM GMT+08:00 | Service manual for Ing64, BR64, BB need correction for ULV. | WIP13 | 459801587191 Rev B 459801587121 Rev B |  | System | Service Manual | N/A | Fixed | shi_b | 0 | 3 | Service Documentation | Support-System | NO | Closed | SubNC | Brilliance 64 |  | 459801587191 Rev B 459801587121 Rev B | October 15, 2019 1:52:15 PM GMT+08:00
CTUDT00622294 | Refresh | Closed | mistry_h | March 8, 2019 5:01:04 PM GMT+08:00 | Service manual for Ing64, BR64, BB need correction for ULV. | WIP13 | 459801587191 Rev B 459801587121 Rev B |  | System | Service Manual | N/A | Fixed | shi_b | 0 | 3 | Service Documentation | Support-System | NO | Closed | SubNC | Brilliance Big Bore |  | 459801587191 Rev B 459801587121 Rev B | October 15, 2019 1:52:15 PM GMT+08:00
CTUDT00622294 | Refresh | Closed | mistry_h | March 8, 2019 5:01:04 PM GMT+08:00 | Service manual for Ing64, BR64, BB need correction for ULV. | WIP13 | 459801587191 Rev B 459801587121 Rev B |  | System | Service Manual | N/A | Fixed | shi_b | 0 | 3 | Service Documentation | Support-System | NO | Closed | SubNC | Ingenuity 64 |  | 459801587191 Rev B 459801587121 Rev B | October 15, 2019 1:52:15 PM GMT+08:00
CTUDT00622292 | Refresh | Closed | mistry_h | March 8, 2019 4:56:00 PM GMT+08:00 | Service manual for iCT/IQon need correction for ULV. | WIP13 |  |  | Subsystem | Service manual | N/A | Does Not Apply | hilas | 0 | 2 |  | Support-System | NO | Closed | WIP20SubCom | Ovation
CTUDT00622292 | Refresh | Closed | mistry_h | March 8, 2019 4:56:00 PM GMT+08:00 | Service manual for iCT/IQon need correction for ULV. | WIP13 |  |  | Subsystem | Service manual | N/A | Does Not Apply | hilas | 0 | 2 |  | Support-System | NO | Closed | WIP20SubCom | IQon1
CTUDT00622287 | Refresh | Closed | zhong_x | March 8, 2019 12:23:33 PM GMT+08:00 | DICOM Tag "Software_Versions" Still show 4.2 in Refresh instead of 4.8 | BB_WIP4 | BB_WIP5.1 |  | Subsystem | DICOM | NA | Fixed | v_sumanth | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP5.1_SegReq | Brilliance Big Bore |  | BB_WIP5.1 | March 26, 2019 5:35:03 PM GMT+08:00
CTUDT00622280 | Refresh | Closed | krisher_i | March 7, 2019 10:35:08 PM GMT+08:00 | Degradation in GUI performance in switch from Patient scene to Plan Scene in Chest scan | WIP14 |  |  | Subsystem | GUI performance |  | Duplicate | sanjeeva_a | 0 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00622278 | Refresh | Closed | krisher_i | March 7, 2019 10:00:46 PM GMT+08:00 | Degradation in GUI performance from previous WIP's in CCT scans | WIP14 | WIP15 | to verify connection to CRTC and connect properly | Subsystem | GUI performance | None | Fixed | patra_r | 0 | 3 | Console Software | Console4 | NO | Closed | WIP15SubCom | Ovation |  | WIP15 | March 22, 2019 2:36:20 PM GMT+08:00
CTUDT00622278 | Refresh | Closed | krisher_i | March 7, 2019 10:00:46 PM GMT+08:00 | Degradation in GUI performance from previous WIP's in CCT scans | WIP14 | WIP15 | to verify connection to CRTC and connect properly | Subsystem | GUI performance | None | Fixed | patra_r | 0 | 3 | Console Software | Console4 | NO | Closed | WIP15SubCom | IQon1 |  | WIP15 | March 22, 2019 2:36:20 PM GMT+08:00
CTUDT00622277 | Refresh | Closed | jotawar_d | March 7, 2019 9:38:18 PM GMT+08:00 | [Refresh]: Unable to add EC to Suggested Exam Cards group | 4.8 | BB_WIP7 |  | Subsystem | Exam Cards | NA | Fixed | datta_i | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP6_SubCom | Brilliance Big Bore |  | BB_WIP7 | July 19, 2019 6:51:35 PM GMT+08:00
CTUDT00622269 | Refresh | Closed | Kevin_ZHAO | March 7, 2019 7:29:22 PM GMT+08:00 | Refresh: Need reset Chinese IM again with double press WIN+SPACE after end each Exam | WIP13 | WIP21.1 BB_WIP11.1 | Shift is recommended to be pressed every time when the user wants to change between English and Chinese. This is the behavior from Microsoft Pinyin IME and cannot be overridden. | Subsystem | Localization ; Blocker21 | NA | Fixed | datta_i | 0 | 3 | Console Software | Console4 | NO | Closed | WIP21.1Subcom | Ovation |  | WIP21.1 BB_WIP11.1 | October 24, 2019 5:02:46 PM GMT+08:00
CTUDT00622269 | Refresh | Closed | Kevin_ZHAO | March 7, 2019 7:29:22 PM GMT+08:00 | Refresh: Need reset Chinese IM again with double press WIN+SPACE after end each Exam | WIP13 | WIP21.1 BB_WIP11.1 | Shift is recommended to be pressed every time when the user wants to change between English and Chinese. This is the behavior from Microsoft Pinyin IME and cannot be overridden. | Subsystem | Localization ; Blocker21 | NA | Fixed | datta_i | 0 | 3 | Console Software | Console4 | NO | Closed | WIP21.1Subcom | Brilliance 64 |  | WIP21.1 BB_WIP11.1 | October 24, 2019 5:02:46 PM GMT+08:00
CTUDT00622269 | Refresh | Closed | Kevin_ZHAO | March 7, 2019 7:29:22 PM GMT+08:00 | Refresh: Need reset Chinese IM again with double press WIN+SPACE after end each Exam | WIP13 | WIP21.1 BB_WIP11.1 | Shift is recommended to be pressed every time when the user wants to change between English and Chinese. This is the behavior from Microsoft Pinyin IME and cannot be overridden. | Subsystem | Localization ; Blocker21 | NA | Fixed | datta_i | 0 | 3 | Console Software | Console4 | NO | Closed | WIP21.1Subcom | Brilliance Big Bore |  | WIP21.1 BB_WIP11.1 | October 24, 2019 5:02:46 PM GMT+08:00
CTUDT00622269 | Refresh | Closed | Kevin_ZHAO | March 7, 2019 7:29:22 PM GMT+08:00 | Refresh: Need reset Chinese IM again with double press WIN+SPACE after end each Exam | WIP13 | WIP21.1 BB_WIP11.1 | Shift is recommended to be pressed every time when the user wants to change between English and Chinese. This is the behavior from Microsoft Pinyin IME and cannot be overridden. | Subsystem | Localization ; Blocker21 | NA | Fixed | datta_i | 0 | 3 | Console Software | Console4 | NO | Closed | WIP21.1Subcom | Ingenuity 64 |  | WIP21.1 BB_WIP11.1 | October 24, 2019 5:02:46 PM GMT+08:00
CTUDT00622269 | Refresh | Closed | Kevin_ZHAO | March 7, 2019 7:29:22 PM GMT+08:00 | Refresh: Need reset Chinese IM again with double press WIN+SPACE after end each Exam | WIP13 | WIP21.1 BB_WIP11.1 | Shift is recommended to be pressed every time when the user wants to change between English and Chinese. This is the behavior from Microsoft Pinyin IME and cannot be overridden. | Subsystem | Localization ; Blocker21 | NA | Fixed | datta_i | 0 | 3 | Console Software | Console4 | NO | Closed | WIP21.1Subcom | IQon1 |  | WIP21.1 BB_WIP11.1 | October 24, 2019 5:02:46 PM GMT+08:00
CTUDT00622267 | Refresh | Closed | Kevin_ZHAO | March 7, 2019 7:21:26 PM GMT+08:00 | Refresh:Chinese Character not the same Font, some characters display issue | WIP13 |  |  | Subsystem | Localization | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | Ovation
CTUDT00622267 | Refresh | Closed | Kevin_ZHAO | March 7, 2019 7:21:26 PM GMT+08:00 | Refresh:Chinese Character not the same Font, some characters display issue | WIP13 |  |  | Subsystem | Localization | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00622267 | Refresh | Closed | Kevin_ZHAO | March 7, 2019 7:21:26 PM GMT+08:00 | Refresh:Chinese Character not the same Font, some characters display issue | WIP13 |  |  | Subsystem | Localization | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00622267 | Refresh | Closed | Kevin_ZHAO | March 7, 2019 7:21:26 PM GMT+08:00 | Refresh:Chinese Character not the same Font, some characters display issue | WIP13 |  |  | Subsystem | Localization | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00622267 | Refresh | Closed | Kevin_ZHAO | March 7, 2019 7:21:26 PM GMT+08:00 | Refresh:Chinese Character not the same Font, some characters display issue | WIP13 |  |  | Subsystem | Localization | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | IQon1
CTUDT00622264 | Refresh | Closed | patra_r | March 7, 2019 6:19:46 PM GMT+08:00 | Incorrect values for REVOLUTION_TIME in DICOM files on Ingenuity CT 4.1.6. | V4.8 | BB_WIP8 | User has to treat 0.44 as 0.444 rotation time for manual calculations | Subsystem | DICOM | NA | Fixed | jose_s | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP8_SubCom | Brilliance Big Bore | NA | BB_WIP8 | July 17, 2019 6:51:07 PM GMT+08:00
CTUDT00622261 | Refresh | Closed | Kevin_ZHAO | March 7, 2019 4:17:38 PM GMT+08:00 | Refresh:Searching window popup automatically in Patient Page | WIP13 |  | NA | Subsystem | Patient Page | NA | Works As Specified / Not A Defect | killada_s | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00622255 | Refresh | Closed | shaham_t | March 7, 2019 3:46:40 PM GMT+08:00 | Demo Folder SRS Requirements Require Modification | Refresh SRS | WIP8 & wip18 |  | Subsystem | Requirement | SRS.Service.2820, SRS.Service.2822 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP15SegReq | Ovation |  | WIP8 & wip18 | July 29, 2019 9:09:45 AM GMT+08:00
CTUDT00622255 | Refresh | Closed | shaham_t | March 7, 2019 3:46:40 PM GMT+08:00 | Demo Folder SRS Requirements Require Modification | Refresh SRS | WIP8 & wip18 |  | Subsystem | Requirement | SRS.Service.2820, SRS.Service.2822 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP15SegReq | Brilliance 64 |  | WIP8 & wip18 | July 29, 2019 9:09:45 AM GMT+08:00
CTUDT00622255 | Refresh | Closed | shaham_t | March 7, 2019 3:46:40 PM GMT+08:00 | Demo Folder SRS Requirements Require Modification | Refresh SRS | WIP8 & wip18 |  | Subsystem | Requirement | SRS.Service.2820, SRS.Service.2822 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP15SegReq | Brilliance Big Bore |  | WIP8 & wip18 | July 29, 2019 9:09:45 AM GMT+08:00
CTUDT00622255 | Refresh | Closed | shaham_t | March 7, 2019 3:46:40 PM GMT+08:00 | Demo Folder SRS Requirements Require Modification | Refresh SRS | WIP8 & wip18 |  | Subsystem | Requirement | SRS.Service.2820, SRS.Service.2822 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP15SegReq | Ingenuity 64 |  | WIP8 & wip18 | July 29, 2019 9:09:45 AM GMT+08:00
CTUDT00622255 | Refresh | Closed | shaham_t | March 7, 2019 3:46:40 PM GMT+08:00 | Demo Folder SRS Requirements Require Modification | Refresh SRS | WIP8 & wip18 |  | Subsystem | Requirement | SRS.Service.2820, SRS.Service.2822 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP15SegReq | IQon1 |  | WIP8 & wip18 | July 29, 2019 9:09:45 AM GMT+08:00
CTUDT00622252 | Refresh | Closed | jotawar_d | March 7, 2019 1:19:24 PM GMT+08:00 | When pressing logout while there are reconstructions in queue, the wrong message appears for the user. | 4.8 | BB_WIP5.1 | Use offline reconstruction | Subsystem |  | NA | Fixed | v_sumanth | 0 | 3 | Console SSRS | Console4 | NO | Closed | BB_WIP5_SubCom | Brilliance Big Bore |  | BB_WIP5.1 | March 20, 2019 7:28:26 PM GMT+08:00
CTUDT00622245 | Refresh | Closed | Valeria_Kashtanov | March 6, 2019 9:42:41 PM GMT+08:00 | Refresh: Calibrations: HCOR: HCOR Basic scan resolutions are not performed as defined in spec | WIP13 | N/A |  | Subsystem | HCOR calibration | DHF170392 | Fixed | braunstein_d | 0 | 2 | Design Documentation | Support-Physics | NO | Closed | WIP15SegReq | Ovation |  | N/A | May 27, 2019 7:40:33 PM GMT+08:00
CTUDT00622237 | Refresh | Closed | Sateesh_S | March 6, 2019 6:49:54 PM GMT+08:00 | BB Refresh :When login as Clinical user without ECM and Dose alert permission then Wrong message and Bugrep displayed in ECM and Dose alert authorization for the locked SeniorTech user having access to ECM and Dose alert | BB_WIP4 | BB_WIP5 | Mo workaround. | Subsystem | User Authentication | NA | Fixed | Prakash_Chand_Yadav | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP6_SegReq | Brilliance Big Bore |  | BB_WIP5 | March 14, 2019 5:00:51 PM GMT+08:00
CTUDT00622235 | Refresh | Closed | jotawar_d | March 6, 2019 4:57:56 PM GMT+08:00 | Min mAs and mA s fields are not highlighted blue when Doseright option is checked and surview is performed. | 4.8 | BB_WIP5.1 | NO as it's a UI issue | Subsystem | Dose Right | N/A | Fixed | sagar_v | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP5_SubCom | Brilliance Big Bore |  | BB_WIP5.1 | March 19, 2019 5:29:07 PM GMT+08:00
CTUDT00622233 | Refresh | Closed | tiwari_g | March 6, 2019 3:19:56 PM GMT+08:00 | Dose Management was displayed as "CTDIvol XX.XX(0-2000 mG)" in plan page | 4.8 |  | NA | Subsystem | Plan Scene | NA | Works As Specified / Not A Defect | sagar_v | 0 | 2 | Console Software | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00622230 | Refresh | Closed | beck_d | March 6, 2019 3:20:00 AM GMT+08:00 | Refresh: Calibrations: CTDI Test: Institute address is not shown in the report header | WIP13 | WIP15 |  | Subsystem | DMS Calibrations | NaN | Fixed | daibes_s | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP15SubCom | Ovation |  | WIP15 | March 25, 2019 4:21:31 PM GMT+08:00
CTUDT00622230 | Refresh | Closed | beck_d | March 6, 2019 3:20:00 AM GMT+08:00 | Refresh: Calibrations: CTDI Test: Institute address is not shown in the report header | WIP13 | WIP15 |  | Subsystem | DMS Calibrations | NaN | Fixed | daibes_s | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP15SubCom | Brilliance Big Bore |  | WIP15 | March 25, 2019 4:21:31 PM GMT+08:00
CTUDT00622230 | Refresh | Closed | beck_d | March 6, 2019 3:20:00 AM GMT+08:00 | Refresh: Calibrations: CTDI Test: Institute address is not shown in the report header | WIP13 | WIP15 |  | Subsystem | DMS Calibrations | NaN | Fixed | daibes_s | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP15SubCom | Ingenuity 64 |  | WIP15 | March 25, 2019 4:21:31 PM GMT+08:00
CTUDT00622230 | Refresh | Closed | beck_d | March 6, 2019 3:20:00 AM GMT+08:00 | Refresh: Calibrations: CTDI Test: Institute address is not shown in the report header | WIP13 | WIP15 |  | Subsystem | DMS Calibrations | NaN | Fixed | daibes_s | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP15SubCom | IQon1 |  | WIP15 | March 25, 2019 4:21:31 PM GMT+08:00
CTUDT00622227 | Refresh | Closed | beck_d | March 6, 2019 12:55:13 AM GMT+08:00 | Refresh: Calibrations: CTDI Test: Cardiac scan fails | WIP13 | DHF170399 Rev08 WIP18 |  | Subsystem | CTDI test | N/A | Fixed | yatziv_l | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP18SegReq | Ovation |  | DHF170399 Rev08 WIP18 | July 16, 2019 10:17:05 PM GMT+08:00
CTUDT00622226 | Refresh | Closed | Noy_Seelig | March 6, 2019 12:43:49 AM GMT+08:00 | Refresh: Help pages not translated to other languages and only to English | BB_WIP3 | WIP15 |  | Subsystem | Verified1 | n/a | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP15SubCom | Ovation |  | WIP15 | March 25, 2019 4:49:24 PM GMT+08:00
CTUDT00622226 | Refresh | Closed | Noy_Seelig | March 6, 2019 12:43:49 AM GMT+08:00 | Refresh: Help pages not translated to other languages and only to English | BB_WIP3 | WIP15 |  | Subsystem | Verified1 | n/a | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP15SubCom | Brilliance Big Bore |  | WIP15 | March 25, 2019 4:49:24 PM GMT+08:00
CTUDT00622226 | Refresh | Closed | Noy_Seelig | March 6, 2019 12:43:49 AM GMT+08:00 | Refresh: Help pages not translated to other languages and only to English | BB_WIP3 | WIP15 |  | Subsystem | Verified1 | n/a | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP15SubCom | Ingenuity 64 |  | WIP15 | March 25, 2019 4:49:24 PM GMT+08:00
CTUDT00622226 | Refresh | Closed | Noy_Seelig | March 6, 2019 12:43:49 AM GMT+08:00 | Refresh: Help pages not translated to other languages and only to English | BB_WIP3 | WIP15 |  | Subsystem | Verified1 | n/a | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP15SubCom | IQon1 |  | WIP15 | March 25, 2019 4:49:24 PM GMT+08:00
CTUDT00622224 | Refresh | Closed | beck_d | March 5, 2019 11:47:18 PM GMT+08:00 | Refresh: Calibrations: Steps Test: Horizontal couch start position of Steps Z is not according to spec | WIP13 | WIP15 |  | Subsystem | DMS Calibrations | n/a | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP15SubCom | Ovation |  | WIP15 | March 25, 2019 3:44:28 PM GMT+08:00
CTUDT00622222 | Refresh | Closed | Sharmila_Vasireddy | March 5, 2019 8:40:57 PM GMT+08:00 | In BB refresh, in splitted studies case, the ECG series was created for only one study not for the 2nd one. | BB_WIP4 | BB_WIP5.1 | None | Subsystem | Split Study | N/A | Fixed | v_sumanth | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP6_SubCom | Brilliance Big Bore |  | BB_WIP5.1 | March 20, 2019 7:50:55 PM GMT+08:00
CTUDT00622218 | Refresh | Closed | attali_s | March 5, 2019 7:13:07 PM GMT+08:00 | Refresh: Beam Quality Tube Potential Test: In manual mode, test starts even when some of the Dose-meter Details window fields are blank | WIP13 | WIP17 |  | Subsystem | DMS Calibrations | NaN | Fixed | daibes_s | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP15SubCom | Ovation |  | WIP17 | May 12, 2019 10:35:34 PM GMT+08:00
CTUDT00622218 | Refresh | Closed | attali_s | March 5, 2019 7:13:07 PM GMT+08:00 | Refresh: Beam Quality Tube Potential Test: In manual mode, test starts even when some of the Dose-meter Details window fields are blank | WIP13 | WIP17 |  | Subsystem | DMS Calibrations | NaN | Fixed | daibes_s | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP15SubCom | Brilliance Big Bore |  | WIP17 | May 12, 2019 10:35:34 PM GMT+08:00
CTUDT00622218 | Refresh | Closed | attali_s | March 5, 2019 7:13:07 PM GMT+08:00 | Refresh: Beam Quality Tube Potential Test: In manual mode, test starts even when some of the Dose-meter Details window fields are blank | WIP13 | WIP17 |  | Subsystem | DMS Calibrations | NaN | Fixed | daibes_s | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP15SubCom | Ingenuity 64 |  | WIP17 | May 12, 2019 10:35:34 PM GMT+08:00
CTUDT00622218 | Refresh | Closed | attali_s | March 5, 2019 7:13:07 PM GMT+08:00 | Refresh: Beam Quality Tube Potential Test: In manual mode, test starts even when some of the Dose-meter Details window fields are blank | WIP13 | WIP17 |  | Subsystem | DMS Calibrations | NaN | Fixed | daibes_s | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP15SubCom | IQon1 |  | WIP17 | May 12, 2019 10:35:34 PM GMT+08:00
CTUDT00622217 | Refresh | Closed | sowmya_k | March 5, 2019 6:44:33 PM GMT+08:00 | Operator Name should not be displayed when clicked on Anonymous button | BB_WIP4 | DHF316705 Rev09 | User can manually delete the Operator name | Subsystem | Patient Scene | Console.SRS.2039 | Fixed | roy_r | 0 | 3 | HOST SSRS | Console4 | NO | Closed | BB_WIP5_SubCom | Brilliance Big Bore |  | DHF316705 Rev09 | April 25, 2019 12:30:29 PM GMT+08:00
CTUDT00622209 | Refresh | Closed | Sateesh_S | March 5, 2019 4:46:33 PM GMT+08:00 | Emergency user account was expired after changing the system date greater than the number of days set in the password Expiry after field. | BB_WIP4 | BB_WIP5 | No. | Subsystem | Login App | NA | Fixed | kapadia_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP15SubCom | Ovation |  | BB_WIP5 | April 2, 2019 7:57:46 PM GMT+08:00
CTUDT00622209 | Refresh | Closed | Sateesh_S | March 5, 2019 4:46:33 PM GMT+08:00 | Emergency user account was expired after changing the system date greater than the number of days set in the password Expiry after field. | BB_WIP4 | BB_WIP5 | No. | Subsystem | Login App | NA | Fixed | kapadia_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP15SubCom | Brilliance 64 |  | BB_WIP5 | April 2, 2019 7:57:46 PM GMT+08:00
CTUDT00622209 | Refresh | Closed | Sateesh_S | March 5, 2019 4:46:33 PM GMT+08:00 | Emergency user account was expired after changing the system date greater than the number of days set in the password Expiry after field. | BB_WIP4 | BB_WIP5 | No. | Subsystem | Login App | NA | Fixed | kapadia_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP15SubCom | Brilliance Big Bore |  | BB_WIP5 | April 2, 2019 7:57:46 PM GMT+08:00
CTUDT00622209 | Refresh | Closed | Sateesh_S | March 5, 2019 4:46:33 PM GMT+08:00 | Emergency user account was expired after changing the system date greater than the number of days set in the password Expiry after field. | BB_WIP4 | BB_WIP5 | No. | Subsystem | Login App | NA | Fixed | kapadia_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP15SubCom | Ingenuity 64 |  | BB_WIP5 | April 2, 2019 7:57:46 PM GMT+08:00
CTUDT00622209 | Refresh | Closed | Sateesh_S | March 5, 2019 4:46:33 PM GMT+08:00 | Emergency user account was expired after changing the system date greater than the number of days set in the password Expiry after field. | BB_WIP4 | BB_WIP5 | No. | Subsystem | Login App | NA | Fixed | kapadia_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP15SubCom | IQon1 |  | BB_WIP5 | April 2, 2019 7:57:46 PM GMT+08:00
CTUDT00622199 | Refresh | Closed | cohen_y | March 5, 2019 4:35:00 AM GMT+08:00 | Refresh: Acceptance/Constancy: iCT SP: In Head1, Thickness1 and Thickness5 cases two scans are performed instead of one. | WIP13 | WIP15 | N/A | Subsystem | DMS Calibrations | N/A | Fixed | elviral | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP16SubCom | Ovation |  | WIP15 | March 25, 2019 8:51:02 PM GMT+08:00
CTUDT00622197 | Refresh | Closed | attali_s | March 4, 2019 10:03:12 PM GMT+08:00 | Refresh: Performance Test: Wrong scan parameters in Cardiac Body scans on iCT SP | WIP13 | WIP18 |  | Subsystem | DMS Calibrations | n/a | Fixed | braunstein_d | 0 | 2 | DMS Calibrations | Support-Physics | NO | Closed | WIP21SubCom | Ovation |  | WIP18 | October 27, 2019 8:33:31 PM GMT+08:00
CTUDT00622194 | Refresh | Closed | patra_r | March 4, 2019 8:22:17 PM GMT+08:00 | Delete option is not available in edit results when review before save is checked | V4.8 | BB_WIP5.1 |  | Subsystem | Edit Results | Console-SRS.9081 | Fixed | Prakash_Chand_Yadav | 0 | 2 | Console SSRS | Console4 | NO | Closed | BB_WIP5_SubCom | Brilliance Big Bore |  | BB_WIP5.1 | March 15, 2019 11:35:08 AM GMT+08:00
CTUDT00622192 | Refresh | Closed | attali_s | March 4, 2019 8:12:09 PM GMT+08:00 | Refresh: HCOR Calibration: Wrong mapping of modes in HCOR for iCT and iCT SP | WIP13 | N/A |  | Subsystem | HCOR | N/A | Fixed | braunstein_d | 0 | 2 | Design Documentation | Support-Physics | NO | Closed | WIP15SegReq | Ovation |  | N/A | May 27, 2019 7:49:06 PM GMT+08:00
CTUDT00622190 | Refresh | Closed | beck_d | March 4, 2019 7:28:29 PM GMT+08:00 | Refresh: Calibrations: FEE Test: Half tile mAs linearity limits are not as defined in spec | WIP13 | WIP15 |  | Subsystem | DMS Calibrations | DHF302470 | Fixed | bikman | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP15SubCom | Ovation |  | WIP15 | March 27, 2019 5:25:01 PM GMT+08:00
CTUDT00622188 | Refresh | Closed | attali_s | March 4, 2019 5:19:22 PM GMT+08:00 | Refresh: Air Calibration: Wrong LogTables in TablesFollowup file for iCT SP | WIP13 | DHF170383 Rev13 WIP17 | N/A | Subsystem | Air Calibration | N/A | Fixed | elviral | 0 | 2 | Design Documentation | DMS | NO | Closed | WIP15SubCom | Ovation |  | DHF170383 Rev13 WIP17 | May 6, 2019 8:03:43 PM GMT+08:00
CTUDT00622186 | Refresh | Closed | attali_s | March 4, 2019 4:52:39 PM GMT+08:00 | Refresh: HCOR Calibration: Missing Table in HCOR Slice for iCT SP | WIP13 | WIP15 |  | Subsystem | DMS Calibrations | n/a | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP15SegReq | Ovation |  | WIP15 | March 17, 2019 9:58:48 PM GMT+08:00
CTUDT00622182 | Refresh | Closed | Valeria_Kashtanov | March 4, 2019 1:07:15 AM GMT+08:00 | Refresh: HCOR: wrong HcorConstFactorTable cfg file is used for Coefficients calculation | WIP13 |  | N/A | Subsystem | DMS Calibrations | N/A | Works As Specified / Not A Defect | yosfan_s | 0 |  |  | DMS | NO | Closed |  | Ingenuity 64
CTUDT00622170 | Refresh | Closed | manjunath_c | March 1, 2019 9:06:00 PM GMT+08:00 | ECM is crashing on creating a timed bolus EC | WIP13 | WIP16 |  | Subsystem | ECM | NA | Fixed | datta_i | 0 | 2 | Console Software | Console4 | NO | Closed | WIP16SubCom | Ingenuity 64 |  | WIP16 | April 10, 2019 8:00:05 PM GMT+08:00
CTUDT00622168 | Refresh | Closed | vinoliya_j | March 1, 2019 8:33:53 PM GMT+08:00 | When the interface language is set as Dutch, the number of login attempts in the user management tool does not follow the range 3 >= to =< 10 | WIP13 | BB_WIP6 WIP17 |  | Subsystem | Login | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP6_SubCom | Ovation |  | BB_WIP6 WIP17 | April 3, 2019 6:06:05 PM GMT+08:00
CTUDT00622168 | Refresh | Closed | vinoliya_j | March 1, 2019 8:33:53 PM GMT+08:00 | When the interface language is set as Dutch, the number of login attempts in the user management tool does not follow the range 3 >= to =< 10 | WIP13 | BB_WIP6 WIP17 |  | Subsystem | Login | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP6_SubCom | Brilliance 64 |  | BB_WIP6 WIP17 | April 3, 2019 6:06:05 PM GMT+08:00
CTUDT00622168 | Refresh | Closed | vinoliya_j | March 1, 2019 8:33:53 PM GMT+08:00 | When the interface language is set as Dutch, the number of login attempts in the user management tool does not follow the range 3 >= to =< 10 | WIP13 | BB_WIP6 WIP17 |  | Subsystem | Login | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP6_SubCom | Brilliance Big Bore |  | BB_WIP6 WIP17 | April 3, 2019 6:06:05 PM GMT+08:00
CTUDT00622168 | Refresh | Closed | vinoliya_j | March 1, 2019 8:33:53 PM GMT+08:00 | When the interface language is set as Dutch, the number of login attempts in the user management tool does not follow the range 3 >= to =< 10 | WIP13 | BB_WIP6 WIP17 |  | Subsystem | Login | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP6_SubCom | Ingenuity 64 |  | BB_WIP6 WIP17 | April 3, 2019 6:06:05 PM GMT+08:00
CTUDT00622168 | Refresh | Closed | vinoliya_j | March 1, 2019 8:33:53 PM GMT+08:00 | When the interface language is set as Dutch, the number of login attempts in the user management tool does not follow the range 3 >= to =< 10 | WIP13 | BB_WIP6 WIP17 |  | Subsystem | Login | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP6_SubCom | IQon1 |  | BB_WIP6 WIP17 | April 3, 2019 6:06:05 PM GMT+08:00
CTUDT00622158 | Refresh | Closed | cohen_y | March 1, 2019 3:29:41 AM GMT+08:00 | Refresh: Acceptance/Constancy: Window Center (Level) value of the Resolution cases with Harmonized Phantom is not according to Spec. | WIP13 | WIP15 |  | Subsystem | DMS Calibrations | N/A | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP15SubCom | Ingenuity 64 |  | WIP15 | March 21, 2019 9:16:31 PM GMT+08:00
CTUDT00622156 | Refresh | Closed | cohen_y | March 1, 2019 2:45:25 AM GMT+08:00 | Refresh: Constancy: No indication in report when Constancy uses Acceptance tolerances if baseline was not loaded. | WIP13 | WIP17 | N/A | Subsystem | DMS Calibrations | N/A | Fixed | elviral | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP15SubCom | Ovation |  | WIP17 | May 22, 2019 11:53:29 PM GMT+08:00
CTUDT00622156 | Refresh | Closed | cohen_y | March 1, 2019 2:45:25 AM GMT+08:00 | Refresh: Constancy: No indication in report when Constancy uses Acceptance tolerances if baseline was not loaded. | WIP13 | WIP17 | N/A | Subsystem | DMS Calibrations | N/A | Fixed | elviral | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP15SubCom | Brilliance Big Bore |  | WIP17 | May 22, 2019 11:53:29 PM GMT+08:00
CTUDT00622156 | Refresh | Closed | cohen_y | March 1, 2019 2:45:25 AM GMT+08:00 | Refresh: Constancy: No indication in report when Constancy uses Acceptance tolerances if baseline was not loaded. | WIP13 | WIP17 | N/A | Subsystem | DMS Calibrations | N/A | Fixed | elviral | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP15SubCom | Ingenuity 64 |  | WIP17 | May 22, 2019 11:53:29 PM GMT+08:00
CTUDT00622156 | Refresh | Closed | cohen_y | March 1, 2019 2:45:25 AM GMT+08:00 | Refresh: Constancy: No indication in report when Constancy uses Acceptance tolerances if baseline was not loaded. | WIP13 | WIP17 | N/A | Subsystem | DMS Calibrations | N/A | Fixed | elviral | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP15SubCom | IQon1 |  | WIP17 | May 22, 2019 11:53:29 PM GMT+08:00
CTUDT00622149 | Refresh | Closed | panchami_pradeep | February 28, 2019 6:29:57 PM GMT+08:00 | Refresh: PSC : Auto download of packages is happening in Software distribution tab even in non idle mode (philips_service mode). | WIP13 | HostSSRS DHF303211 Rev 08 |  | Subsystem | Service Tools | STPS-SSRS-1276 | Fixed | hatti_c | 0 | 4 | Requirements Documentation | Service Tools | NO | Closed | WIP15SegReq | Ovation |  | HostSSRS DHF303211 Rev 08 | March 25, 2019 6:30:57 PM GMT+08:00
CTUDT00622149 | Refresh | Closed | panchami_pradeep | February 28, 2019 6:29:57 PM GMT+08:00 | Refresh: PSC : Auto download of packages is happening in Software distribution tab even in non idle mode (philips_service mode). | WIP13 | HostSSRS DHF303211 Rev 08 |  | Subsystem | Service Tools | STPS-SSRS-1276 | Fixed | hatti_c | 0 | 4 | Requirements Documentation | Service Tools | NO | Closed | WIP15SegReq | Brilliance 64 |  | HostSSRS DHF303211 Rev 08 | March 25, 2019 6:30:57 PM GMT+08:00
CTUDT00622149 | Refresh | Closed | panchami_pradeep | February 28, 2019 6:29:57 PM GMT+08:00 | Refresh: PSC : Auto download of packages is happening in Software distribution tab even in non idle mode (philips_service mode). | WIP13 | HostSSRS DHF303211 Rev 08 |  | Subsystem | Service Tools | STPS-SSRS-1276 | Fixed | hatti_c | 0 | 4 | Requirements Documentation | Service Tools | NO | Closed | WIP15SegReq | Ingenuity 64 |  | HostSSRS DHF303211 Rev 08 | March 25, 2019 6:30:57 PM GMT+08:00
CTUDT00622149 | Refresh | Closed | panchami_pradeep | February 28, 2019 6:29:57 PM GMT+08:00 | Refresh: PSC : Auto download of packages is happening in Software distribution tab even in non idle mode (philips_service mode). | WIP13 | HostSSRS DHF303211 Rev 08 |  | Subsystem | Service Tools | STPS-SSRS-1276 | Fixed | hatti_c | 0 | 4 | Requirements Documentation | Service Tools | NO | Closed | WIP15SegReq | IQon1 |  | HostSSRS DHF303211 Rev 08 | March 25, 2019 6:30:57 PM GMT+08:00
CTUDT00622148 | Refresh | Closed | manju_kolkar | February 28, 2019 6:17:17 PM GMT+08:00 | Refresh - Pan and Zoom applied on preview images were not applied on both the acquisition results. | WIP13 |  |  | Subsystem | Pan & Zoom | NA | No Plan to Fix | redrouthu_r | 0 | 4 |  | iCAP | NO | Closed |  | Ovation
CTUDT00622148 | Refresh | Closed | manju_kolkar | February 28, 2019 6:17:17 PM GMT+08:00 | Refresh - Pan and Zoom applied on preview images were not applied on both the acquisition results. | WIP13 |  |  | Subsystem | Pan & Zoom | NA | No Plan to Fix | redrouthu_r | 0 | 4 |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00622148 | Refresh | Closed | manju_kolkar | February 28, 2019 6:17:17 PM GMT+08:00 | Refresh - Pan and Zoom applied on preview images were not applied on both the acquisition results. | WIP13 |  |  | Subsystem | Pan & Zoom | NA | No Plan to Fix | redrouthu_r | 0 | 4 |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00622148 | Refresh | Closed | manju_kolkar | February 28, 2019 6:17:17 PM GMT+08:00 | Refresh - Pan and Zoom applied on preview images were not applied on both the acquisition results. | WIP13 |  |  | Subsystem | Pan & Zoom | NA | No Plan to Fix | redrouthu_r | 0 | 4 |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00622148 | Refresh | Closed | manju_kolkar | February 28, 2019 6:17:17 PM GMT+08:00 | Refresh - Pan and Zoom applied on preview images were not applied on both the acquisition results. | WIP13 |  |  | Subsystem | Pan & Zoom | NA | No Plan to Fix | redrouthu_r | 0 | 4 |  | iCAP | NO | Closed |  | IQon1
CTUDT00622145 | Refresh | Closed | swack_g | February 28, 2019 5:17:22 PM GMT+08:00 | Refresh: Bad Detectors Quick Check: missing messages. | WIP13 | N/A |  | Subsystem | Test Design | n/a | Fixed | yosfan_s | 0 | 2 | Test Design | DMS | NO | Closed | WIP15SegReq | Ovation |  | N/A | April 8, 2019 1:15:55 PM GMT+08:00
CTUDT00622141 | Refresh | Closed | swack_g | February 28, 2019 3:43:41 PM GMT+08:00 | Refresh: Allowed Combinations: incorrect X-Ray Filter. | WIP13 |  |  | Subsystem | DMS Calibrations | n/a | Works As Specified / Not A Defect | yatziv_l | 0 |  |  | DMS | NO | Closed |  | Ovation
CTUDT00622139 | Refresh | Closed | panchami_pradeep | February 28, 2019 12:12:50 AM GMT+08:00 | Refresh: Console: Short tube conditioning message is displayed even before 12 hours of calibration | WIP13 | WIP17 |  | Subsystem | HOST | N/A | Fixed | pandey_a | 0 | 4 | Couch Software | Console4 | NO | Closed | WIP15SegReq | Ovation |  | WIP17 | May 15, 2019 8:36:29 PM GMT+08:00
CTUDT00622139 | Refresh | Closed | panchami_pradeep | February 28, 2019 12:12:50 AM GMT+08:00 | Refresh: Console: Short tube conditioning message is displayed even before 12 hours of calibration | WIP13 | WIP17 |  | Subsystem | HOST | N/A | Fixed | pandey_a | 0 | 4 | Couch Software | Console4 | NO | Closed | WIP15SegReq | Brilliance 64 |  | WIP17 | May 15, 2019 8:36:29 PM GMT+08:00
CTUDT00622139 | Refresh | Closed | panchami_pradeep | February 28, 2019 12:12:50 AM GMT+08:00 | Refresh: Console: Short tube conditioning message is displayed even before 12 hours of calibration | WIP13 | WIP17 |  | Subsystem | HOST | N/A | Fixed | pandey_a | 0 | 4 | Couch Software | Console4 | NO | Closed | WIP15SegReq | Ingenuity 64 |  | WIP17 | May 15, 2019 8:36:29 PM GMT+08:00
CTUDT00622139 | Refresh | Closed | panchami_pradeep | February 28, 2019 12:12:50 AM GMT+08:00 | Refresh: Console: Short tube conditioning message is displayed even before 12 hours of calibration | WIP13 | WIP17 |  | Subsystem | HOST | N/A | Fixed | pandey_a | 0 | 4 | Couch Software | Console4 | NO | Closed | WIP15SegReq | IQon1 |  | WIP17 | May 15, 2019 8:36:29 PM GMT+08:00
CTUDT00622137 | Refresh | Closed | panchami_pradeep | February 27, 2019 11:34:29 PM GMT+08:00 | Refresh: PSC: Software distribution tool is not launching when remote connection is through ISSL. | wip 13 |  |  | Subsystem | Service Tools | STPS SSRS-1085 | Works As Specified / Not A Defect | jindal_k | 0 | 2 |  | Service Tools | NO | Closed | WIP16SubCom | Ovation
CTUDT00622137 | Refresh | Closed | panchami_pradeep | February 27, 2019 11:34:29 PM GMT+08:00 | Refresh: PSC: Software distribution tool is not launching when remote connection is through ISSL. | wip 13 |  |  | Subsystem | Service Tools | STPS SSRS-1085 | Works As Specified / Not A Defect | jindal_k | 0 | 2 |  | Service Tools | NO | Closed | WIP16SubCom | Brilliance Big Bore
CTUDT00622137 | Refresh | Closed | panchami_pradeep | February 27, 2019 11:34:29 PM GMT+08:00 | Refresh: PSC: Software distribution tool is not launching when remote connection is through ISSL. | wip 13 |  |  | Subsystem | Service Tools | STPS SSRS-1085 | Works As Specified / Not A Defect | jindal_k | 0 | 2 |  | Service Tools | NO | Closed | WIP16SubCom | Ingenuity 64
CTUDT00622137 | Refresh | Closed | panchami_pradeep | February 27, 2019 11:34:29 PM GMT+08:00 | Refresh: PSC: Software distribution tool is not launching when remote connection is through ISSL. | wip 13 |  |  | Subsystem | Service Tools | STPS SSRS-1085 | Works As Specified / Not A Defect | jindal_k | 0 | 2 |  | Service Tools | NO | Closed | WIP16SubCom | IQon1
CTUDT00622135 | Refresh | Closed | zhong_x | February 27, 2019 10:17:12 PM GMT+08:00 | Refresh: Login screen switched to another monitor after restart the host | BB_WIP4 | WIP16 | reconnect displays manually | Subsystem | Login | N/A | Fixed | kr_m | 0 | 4 | Console Software | Console4 | NO | Closed | WIP15SubCom | Ovation |  | WIP16 | April 11, 2019 1:02:27 PM GMT+08:00
CTUDT00622135 | Refresh | Closed | zhong_x | February 27, 2019 10:17:12 PM GMT+08:00 | Refresh: Login screen switched to another monitor after restart the host | BB_WIP4 | WIP16 | reconnect displays manually | Subsystem | Login | N/A | Fixed | kr_m | 0 | 4 | Console Software | Console4 | NO | Closed | WIP15SubCom | Brilliance 64 |  | WIP16 | April 11, 2019 1:02:27 PM GMT+08:00
CTUDT00622135 | Refresh | Closed | zhong_x | February 27, 2019 10:17:12 PM GMT+08:00 | Refresh: Login screen switched to another monitor after restart the host | BB_WIP4 | WIP16 | reconnect displays manually | Subsystem | Login | N/A | Fixed | kr_m | 0 | 4 | Console Software | Console4 | NO | Closed | WIP15SubCom | Brilliance Big Bore |  | WIP16 | April 11, 2019 1:02:27 PM GMT+08:00
CTUDT00622135 | Refresh | Closed | zhong_x | February 27, 2019 10:17:12 PM GMT+08:00 | Refresh: Login screen switched to another monitor after restart the host | BB_WIP4 | WIP16 | reconnect displays manually | Subsystem | Login | N/A | Fixed | kr_m | 0 | 4 | Console Software | Console4 | NO | Closed | WIP15SubCom | Ingenuity 64 |  | WIP16 | April 11, 2019 1:02:27 PM GMT+08:00
CTUDT00622135 | Refresh | Closed | zhong_x | February 27, 2019 10:17:12 PM GMT+08:00 | Refresh: Login screen switched to another monitor after restart the host | BB_WIP4 | WIP16 | reconnect displays manually | Subsystem | Login | N/A | Fixed | kr_m | 0 | 4 | Console Software | Console4 | NO | Closed | WIP15SubCom | IQon1 |  | WIP16 | April 11, 2019 1:02:27 PM GMT+08:00
CTUDT00622133 | Refresh | Closed | zhong_x | February 27, 2019 9:58:38 PM GMT+08:00 | Refresh: Login screen displayed abnormally | BB_WIP4 | WIP18 BB_WIP8 |  | Subsystem | Login | NA | Fixed | kapadia_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP15SubCom | Ovation |  | WIP18 BB_WIP8 | June 17, 2019 1:07:35 PM GMT+08:00
CTUDT00622133 | Refresh | Closed | zhong_x | February 27, 2019 9:58:38 PM GMT+08:00 | Refresh: Login screen displayed abnormally | BB_WIP4 | WIP18 BB_WIP8 |  | Subsystem | Login | NA | Fixed | kapadia_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP15SubCom | Brilliance 64 |  | WIP18 BB_WIP8 | June 17, 2019 1:07:35 PM GMT+08:00
CTUDT00622133 | Refresh | Closed | zhong_x | February 27, 2019 9:58:38 PM GMT+08:00 | Refresh: Login screen displayed abnormally | BB_WIP4 | WIP18 BB_WIP8 |  | Subsystem | Login | NA | Fixed | kapadia_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP15SubCom | Brilliance Big Bore |  | WIP18 BB_WIP8 | June 17, 2019 1:07:35 PM GMT+08:00
CTUDT00622133 | Refresh | Closed | zhong_x | February 27, 2019 9:58:38 PM GMT+08:00 | Refresh: Login screen displayed abnormally | BB_WIP4 | WIP18 BB_WIP8 |  | Subsystem | Login | NA | Fixed | kapadia_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP15SubCom | Ingenuity 64 |  | WIP18 BB_WIP8 | June 17, 2019 1:07:35 PM GMT+08:00
CTUDT00622133 | Refresh | Closed | zhong_x | February 27, 2019 9:58:38 PM GMT+08:00 | Refresh: Login screen displayed abnormally | BB_WIP4 | WIP18 BB_WIP8 |  | Subsystem | Login | NA | Fixed | kapadia_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP15SubCom | IQon1 |  | WIP18 BB_WIP8 | June 17, 2019 1:07:35 PM GMT+08:00
CTUDT00622131 | Refresh | Closed | zhong_x | February 27, 2019 9:48:15 PM GMT+08:00 | Refresh: "System is still initializing. Please wait" displayed when try to load brain Helical EC | BB_WIP4 |  |  | Subsystem | Brian Helical | NA | Does Not Apply | babu_s | 1 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00622129 | Refresh | Closed | zhong_x | February 27, 2019 9:29:55 PM GMT+08:00 | Refresh: Lots of Drivers were not installed in HPZ440 device Manager. | BB_WIP4 | WIP18 BB_WIP8 |  | Subsystem | 1-Click | N/A | Fixed | dasgupta_s | 0 | 2 | Console Software | Support-Security | NO | Closed | WIP18SubCom | Ovation |  | WIP18 BB_WIP8 | June 19, 2019 4:39:40 PM GMT+08:00
CTUDT00622129 | Refresh | Closed | zhong_x | February 27, 2019 9:29:55 PM GMT+08:00 | Refresh: Lots of Drivers were not installed in HPZ440 device Manager. | BB_WIP4 | WIP18 BB_WIP8 |  | Subsystem | 1-Click | N/A | Fixed | dasgupta_s | 0 | 2 | Console Software | Support-Security | NO | Closed | WIP18SubCom | Brilliance 64 |  | WIP18 BB_WIP8 | June 19, 2019 4:39:40 PM GMT+08:00
CTUDT00622129 | Refresh | Closed | zhong_x | February 27, 2019 9:29:55 PM GMT+08:00 | Refresh: Lots of Drivers were not installed in HPZ440 device Manager. | BB_WIP4 | WIP18 BB_WIP8 |  | Subsystem | 1-Click | N/A | Fixed | dasgupta_s | 0 | 2 | Console Software | Support-Security | NO | Closed | WIP18SubCom | Orion P |  | WIP18 BB_WIP8 | June 19, 2019 4:39:40 PM GMT+08:00
CTUDT00622129 | Refresh | Closed | zhong_x | February 27, 2019 9:29:55 PM GMT+08:00 | Refresh: Lots of Drivers were not installed in HPZ440 device Manager. | BB_WIP4 | WIP18 BB_WIP8 |  | Subsystem | 1-Click | N/A | Fixed | dasgupta_s | 0 | 2 | Console Software | Support-Security | NO | Closed | WIP18SubCom | Brilliance Big Bore |  | WIP18 BB_WIP8 | June 19, 2019 4:39:40 PM GMT+08:00
CTUDT00622129 | Refresh | Closed | zhong_x | February 27, 2019 9:29:55 PM GMT+08:00 | Refresh: Lots of Drivers were not installed in HPZ440 device Manager. | BB_WIP4 | WIP18 BB_WIP8 |  | Subsystem | 1-Click | N/A | Fixed | dasgupta_s | 0 | 2 | Console Software | Support-Security | NO | Closed | WIP18SubCom | IQon1 |  | WIP18 BB_WIP8 | June 19, 2019 4:39:40 PM GMT+08:00
CTUDT00622127 | Refresh | Closed | zhong_x | February 27, 2019 8:59:48 PM GMT+08:00 | Refresh: No options display in Choose model tab when install the host | BB_WIP4 |  | NA | Subsystem | 1-Click | NA | Works As Specified / Not A Defect | bilal_m | 1 |  |  | DI | NO | Closed |  | Brilliance Big Bore
CTUDT00622118 | Refresh | Closed | Sateesh_S | February 27, 2019 7:18:38 PM GMT+08:00 | BB Refresh: Message number 133 and 134 are not displaying correctly for Axial jog and Non Axial Jog exam cards when the group limit value is exceeds. | BB_WIP3 | BB_WIP5.1 |  | Subsystem | Perfusion | None | Fixed | jose_s | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP5_SegReq | Brilliance Big Bore |  | BB_WIP5.1 | March 15, 2019 12:51:46 PM GMT+08:00
CTUDT00622111 | Refresh | Closed | manjunath_mk | February 27, 2019 6:20:57 PM GMT+08:00 | The console displays Critical error followed by Bugrep error and logs off | BB_WIP3 | BB_WIP6 |  | Subsystem | Workflow | NA | Fixed | agrawal_ay | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP6_SubCom | Brilliance Big Bore |  | BB_WIP6 | April 26, 2019 8:21:01 PM GMT+08:00
CTUDT00622096 | Refresh | Closed | swamypreethi | February 27, 2019 4:08:46 PM GMT+08:00 | IOCC_Refresh: SUT sends Abort message after receiving result/reason = 0,1,2,3,4 in Association accept message from from remote system. | WIP12 | WIP15 |  | Subsystem | USP | NA | Fixed | s_sa | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP15SubCom | Ovation |  | WIP15 | April 10, 2019 4:43:01 PM GMT+08:00
CTUDT00622096 | Refresh | Closed | swamypreethi | February 27, 2019 4:08:46 PM GMT+08:00 | IOCC_Refresh: SUT sends Abort message after receiving result/reason = 0,1,2,3,4 in Association accept message from from remote system. | WIP12 | WIP15 |  | Subsystem | USP | NA | Fixed | s_sa | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP15SubCom | Brilliance 64 |  | WIP15 | April 10, 2019 4:43:01 PM GMT+08:00
CTUDT00622096 | Refresh | Closed | swamypreethi | February 27, 2019 4:08:46 PM GMT+08:00 | IOCC_Refresh: SUT sends Abort message after receiving result/reason = 0,1,2,3,4 in Association accept message from from remote system. | WIP12 | WIP15 |  | Subsystem | USP | NA | Fixed | s_sa | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP15SubCom | Brilliance Big Bore |  | WIP15 | April 10, 2019 4:43:01 PM GMT+08:00
CTUDT00622096 | Refresh | Closed | swamypreethi | February 27, 2019 4:08:46 PM GMT+08:00 | IOCC_Refresh: SUT sends Abort message after receiving result/reason = 0,1,2,3,4 in Association accept message from from remote system. | WIP12 | WIP15 |  | Subsystem | USP | NA | Fixed | s_sa | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP15SubCom | Ingenuity 64 |  | WIP15 | April 10, 2019 4:43:01 PM GMT+08:00
CTUDT00622096 | Refresh | Closed | swamypreethi | February 27, 2019 4:08:46 PM GMT+08:00 | IOCC_Refresh: SUT sends Abort message after receiving result/reason = 0,1,2,3,4 in Association accept message from from remote system. | WIP12 | WIP15 |  | Subsystem | USP | NA | Fixed | s_sa | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP15SubCom | IQon1 |  | WIP15 | April 10, 2019 4:43:01 PM GMT+08:00
CTUDT00622091 | Refresh | Closed | bandil_j | February 27, 2019 2:56:00 PM GMT+08:00 | Console-SRS.1118 requires update | 4.8 | DHF316705 Rev08 | None | Subsystem | Requirements | Console-SRS.1118 | Fixed | vinoliya_j | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | BB_WIP5_SubCom | Brilliance Big Bore |  | DHF316705 Rev08 | March 15, 2019 5:03:25 PM GMT+08:00
CTUDT00622088 | Refresh | Closed | swamypreethi | February 27, 2019 2:41:39 PM GMT+08:00 | IOCC_Refresh: Incorrect count in related fields attributes in C-MOVE-RSP. | WIP12 |  |  | Subsystem | USP | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | Ovation
CTUDT00622088 | Refresh | Closed | swamypreethi | February 27, 2019 2:41:39 PM GMT+08:00 | IOCC_Refresh: Incorrect count in related fields attributes in C-MOVE-RSP. | WIP12 |  |  | Subsystem | USP | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00622088 | Refresh | Closed | swamypreethi | February 27, 2019 2:41:39 PM GMT+08:00 | IOCC_Refresh: Incorrect count in related fields attributes in C-MOVE-RSP. | WIP12 |  |  | Subsystem | USP | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00622088 | Refresh | Closed | swamypreethi | February 27, 2019 2:41:39 PM GMT+08:00 | IOCC_Refresh: Incorrect count in related fields attributes in C-MOVE-RSP. | WIP12 |  |  | Subsystem | USP | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00622088 | Refresh | Closed | swamypreethi | February 27, 2019 2:41:39 PM GMT+08:00 | IOCC_Refresh: Incorrect count in related fields attributes in C-MOVE-RSP. | WIP12 |  |  | Subsystem | USP | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | IQon1
CTUDT00622086 | Refresh | Closed | swamypreethi | February 27, 2019 2:01:08 PM GMT+08:00 | IOCC_Refresh: Retrieve fails at Image level when move request with List of UID containing one non-matching and one matching UID. | WIP12 |  | repeat request with matching UID | Subsystem | USP | NA | No Plan to Fix | s_sa | 0 | 3 |  | iCAP | NO | Closed |  | Ovation
CTUDT00622086 | Refresh | Closed | swamypreethi | February 27, 2019 2:01:08 PM GMT+08:00 | IOCC_Refresh: Retrieve fails at Image level when move request with List of UID containing one non-matching and one matching UID. | WIP12 |  | repeat request with matching UID | Subsystem | USP | NA | No Plan to Fix | s_sa | 0 | 3 |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00622086 | Refresh | Closed | swamypreethi | February 27, 2019 2:01:08 PM GMT+08:00 | IOCC_Refresh: Retrieve fails at Image level when move request with List of UID containing one non-matching and one matching UID. | WIP12 |  | repeat request with matching UID | Subsystem | USP | NA | No Plan to Fix | s_sa | 0 | 3 |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00622086 | Refresh | Closed | swamypreethi | February 27, 2019 2:01:08 PM GMT+08:00 | IOCC_Refresh: Retrieve fails at Image level when move request with List of UID containing one non-matching and one matching UID. | WIP12 |  | repeat request with matching UID | Subsystem | USP | NA | No Plan to Fix | s_sa | 0 | 3 |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00622086 | Refresh | Closed | swamypreethi | February 27, 2019 2:01:08 PM GMT+08:00 | IOCC_Refresh: Retrieve fails at Image level when move request with List of UID containing one non-matching and one matching UID. | WIP12 |  | repeat request with matching UID | Subsystem | USP | NA | No Plan to Fix | s_sa | 0 | 3 |  | iCAP | NO | Closed |  | IQon1
CTUDT00622085 | Refresh | Closed | datta_i | February 27, 2019 1:07:01 PM GMT+08:00 | Refresh: some user controls in UI cannot well accept Chinese Characters as input | 4.8 | BB_WIP8 |  | Subsystem | Chinese Keyboard | N/A | Fixed | wilsonraja | 0 | 2 | Console Software | Console4 | YES | Closed | BB_WIP5_SubCom | Brilliance Big Bore |  | BB_WIP8 | June 18, 2019 3:29:02 PM GMT+08:00
CTUDT00622083 | Refresh | Closed | swamypreethi | February 27, 2019 12:23:21 PM GMT+08:00 | IOCC_Refresh: SUT displays patient data in the patient directory which has more than 5 components. | WIP12 |  | NA | Subsystem | USP | NA | Works As Specified / Not A Defect | s_sa | 0 |  |  | iCAP | NO | Closed |  | Ovation
CTUDT00622083 | Refresh | Closed | swamypreethi | February 27, 2019 12:23:21 PM GMT+08:00 | IOCC_Refresh: SUT displays patient data in the patient directory which has more than 5 components. | WIP12 |  | NA | Subsystem | USP | NA | Works As Specified / Not A Defect | s_sa | 0 |  |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00622083 | Refresh | Closed | swamypreethi | February 27, 2019 12:23:21 PM GMT+08:00 | IOCC_Refresh: SUT displays patient data in the patient directory which has more than 5 components. | WIP12 |  | NA | Subsystem | USP | NA | Works As Specified / Not A Defect | s_sa | 0 |  |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00622083 | Refresh | Closed | swamypreethi | February 27, 2019 12:23:21 PM GMT+08:00 | IOCC_Refresh: SUT displays patient data in the patient directory which has more than 5 components. | WIP12 |  | NA | Subsystem | USP | NA | Works As Specified / Not A Defect | s_sa | 0 |  |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00622083 | Refresh | Closed | swamypreethi | February 27, 2019 12:23:21 PM GMT+08:00 | IOCC_Refresh: SUT displays patient data in the patient directory which has more than 5 components. | WIP12 |  | NA | Subsystem | USP | NA | Works As Specified / Not A Defect | s_sa | 0 |  |  | iCAP | NO | Closed |  | IQon1
CTUDT00622081 | Refresh | Closed | swamypreethi | February 27, 2019 12:09:20 PM GMT+08:00 | IOCC_Refresh: Export job in SUT is marked as failed when status B007H [ Warning, "Data Set does not match SOP Class" and 0107H [ Warning, "Attribute List Error"] is sent in C-STORE-RSP of remote system | WIP12 | WIP15 |  | Subsystem | Test Design | NA | Fixed | s_sa | 0 | 4 | Console USP_EBW | iCAP | NO | Closed | WIP15SegReq | Ovation |  | WIP15 | April 10, 2019 4:39:50 PM GMT+08:00
CTUDT00622081 | Refresh | Closed | swamypreethi | February 27, 2019 12:09:20 PM GMT+08:00 | IOCC_Refresh: Export job in SUT is marked as failed when status B007H [ Warning, "Data Set does not match SOP Class" and 0107H [ Warning, "Attribute List Error"] is sent in C-STORE-RSP of remote system | WIP12 | WIP15 |  | Subsystem | Test Design | NA | Fixed | s_sa | 0 | 4 | Console USP_EBW | iCAP | NO | Closed | WIP15SegReq | Brilliance 64 |  | WIP15 | April 10, 2019 4:39:50 PM GMT+08:00
CTUDT00622081 | Refresh | Closed | swamypreethi | February 27, 2019 12:09:20 PM GMT+08:00 | IOCC_Refresh: Export job in SUT is marked as failed when status B007H [ Warning, "Data Set does not match SOP Class" and 0107H [ Warning, "Attribute List Error"] is sent in C-STORE-RSP of remote system | WIP12 | WIP15 |  | Subsystem | Test Design | NA | Fixed | s_sa | 0 | 4 | Console USP_EBW | iCAP | NO | Closed | WIP15SegReq | Brilliance Big Bore |  | WIP15 | April 10, 2019 4:39:50 PM GMT+08:00
CTUDT00622081 | Refresh | Closed | swamypreethi | February 27, 2019 12:09:20 PM GMT+08:00 | IOCC_Refresh: Export job in SUT is marked as failed when status B007H [ Warning, "Data Set does not match SOP Class" and 0107H [ Warning, "Attribute List Error"] is sent in C-STORE-RSP of remote system | WIP12 | WIP15 |  | Subsystem | Test Design | NA | Fixed | s_sa | 0 | 4 | Console USP_EBW | iCAP | NO | Closed | WIP15SegReq | Ingenuity 64 |  | WIP15 | April 10, 2019 4:39:50 PM GMT+08:00
CTUDT00622081 | Refresh | Closed | swamypreethi | February 27, 2019 12:09:20 PM GMT+08:00 | IOCC_Refresh: Export job in SUT is marked as failed when status B007H [ Warning, "Data Set does not match SOP Class" and 0107H [ Warning, "Attribute List Error"] is sent in C-STORE-RSP of remote system | WIP12 | WIP15 |  | Subsystem | Test Design | NA | Fixed | s_sa | 0 | 4 | Console USP_EBW | iCAP | NO | Closed | WIP15SegReq | IQon1 |  | WIP15 | April 10, 2019 4:39:50 PM GMT+08:00
CTUDT00622079 | Refresh | Closed | attali_s | February 27, 2019 3:54:06 AM GMT+08:00 | Refresh: Acceptance/Constancy: No warning during System Setup when Physics H-Phantom is not attached | WIP13 | WIP15 | Attach a phantom | Subsystem | DMS calibrations | N/A | Fixed | daibes_s | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP15SegReq | Ovation |  | WIP15 | March 27, 2019 4:36:17 PM GMT+08:00
CTUDT00622079 | Refresh | Closed | attali_s | February 27, 2019 3:54:06 AM GMT+08:00 | Refresh: Acceptance/Constancy: No warning during System Setup when Physics H-Phantom is not attached | WIP13 | WIP15 | Attach a phantom | Subsystem | DMS calibrations | N/A | Fixed | daibes_s | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP15SegReq | Brilliance 64 |  | WIP15 | March 27, 2019 4:36:17 PM GMT+08:00
CTUDT00622079 | Refresh | Closed | attali_s | February 27, 2019 3:54:06 AM GMT+08:00 | Refresh: Acceptance/Constancy: No warning during System Setup when Physics H-Phantom is not attached | WIP13 | WIP15 | Attach a phantom | Subsystem | DMS calibrations | N/A | Fixed | daibes_s | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP15SegReq | Brilliance Big Bore |  | WIP15 | March 27, 2019 4:36:17 PM GMT+08:00
CTUDT00622079 | Refresh | Closed | attali_s | February 27, 2019 3:54:06 AM GMT+08:00 | Refresh: Acceptance/Constancy: No warning during System Setup when Physics H-Phantom is not attached | WIP13 | WIP15 | Attach a phantom | Subsystem | DMS calibrations | N/A | Fixed | daibes_s | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP15SegReq | Ingenuity 64 |  | WIP15 | March 27, 2019 4:36:17 PM GMT+08:00
CTUDT00622079 | Refresh | Closed | attali_s | February 27, 2019 3:54:06 AM GMT+08:00 | Refresh: Acceptance/Constancy: No warning during System Setup when Physics H-Phantom is not attached | WIP13 | WIP15 | Attach a phantom | Subsystem | DMS calibrations | N/A | Fixed | daibes_s | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP15SegReq | IQon1 |  | WIP15 | March 27, 2019 4:36:17 PM GMT+08:00
CTUDT00622073 | Refresh | Closed | swamypreethi | February 26, 2019 7:22:47 PM GMT+08:00 | IOCC_Refresh: SUT not sending C-FIND-RSP with status 0122H: Refused SOP Class not Supported for the Query request with non-negotiated SOP class-RQ sent from remote system | WIP12 |  |  | Subsystem | USP - Connectivity | NA | No Plan to Fix | s_sa | 0 | 3 |  | iCAP | NO | Closed |  | Ovation
CTUDT00622073 | Refresh | Closed | swamypreethi | February 26, 2019 7:22:47 PM GMT+08:00 | IOCC_Refresh: SUT not sending C-FIND-RSP with status 0122H: Refused SOP Class not Supported for the Query request with non-negotiated SOP class-RQ sent from remote system | WIP12 |  |  | Subsystem | USP - Connectivity | NA | No Plan to Fix | s_sa | 0 | 3 |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00622073 | Refresh | Closed | swamypreethi | February 26, 2019 7:22:47 PM GMT+08:00 | IOCC_Refresh: SUT not sending C-FIND-RSP with status 0122H: Refused SOP Class not Supported for the Query request with non-negotiated SOP class-RQ sent from remote system | WIP12 |  |  | Subsystem | USP - Connectivity | NA | No Plan to Fix | s_sa | 0 | 3 |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00622073 | Refresh | Closed | swamypreethi | February 26, 2019 7:22:47 PM GMT+08:00 | IOCC_Refresh: SUT not sending C-FIND-RSP with status 0122H: Refused SOP Class not Supported for the Query request with non-negotiated SOP class-RQ sent from remote system | WIP12 |  |  | Subsystem | USP - Connectivity | NA | No Plan to Fix | s_sa | 0 | 3 |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00622073 | Refresh | Closed | swamypreethi | February 26, 2019 7:22:47 PM GMT+08:00 | IOCC_Refresh: SUT not sending C-FIND-RSP with status 0122H: Refused SOP Class not Supported for the Query request with non-negotiated SOP class-RQ sent from remote system | WIP12 |  |  | Subsystem | USP - Connectivity | NA | No Plan to Fix | s_sa | 0 | 3 |  | iCAP | NO | Closed |  | IQon1
CTUDT00622068 | Refresh | Closed | finestine_m | February 26, 2019 5:22:50 PM GMT+08:00 | Refresh: Hcor: No Priority to Full Runs in Saving Ruls | WIP13 | WIP15 |  | Subsystem | DMS Calibrations | N/A | Fixed | Kostya_Baklan | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP15SubCom | Ovation |  | WIP15 | March 24, 2019 4:43:34 PM GMT+08:00
CTUDT00622068 | Refresh | Closed | finestine_m | February 26, 2019 5:22:50 PM GMT+08:00 | Refresh: Hcor: No Priority to Full Runs in Saving Ruls | WIP13 | WIP15 |  | Subsystem | DMS Calibrations | N/A | Fixed | Kostya_Baklan | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP15SubCom | Brilliance Big Bore |  | WIP15 | March 24, 2019 4:43:34 PM GMT+08:00
CTUDT00622068 | Refresh | Closed | finestine_m | February 26, 2019 5:22:50 PM GMT+08:00 | Refresh: Hcor: No Priority to Full Runs in Saving Ruls | WIP13 | WIP15 |  | Subsystem | DMS Calibrations | N/A | Fixed | Kostya_Baklan | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP15SubCom | Ingenuity 64 |  | WIP15 | March 24, 2019 4:43:34 PM GMT+08:00
CTUDT00622068 | Refresh | Closed | finestine_m | February 26, 2019 5:22:50 PM GMT+08:00 | Refresh: Hcor: No Priority to Full Runs in Saving Ruls | WIP13 | WIP15 |  | Subsystem | DMS Calibrations | N/A | Fixed | Kostya_Baklan | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP15SubCom | IQon1 |  | WIP15 | March 24, 2019 4:43:34 PM GMT+08:00
CTUDT00622065 | Refresh | Closed | swamypreethi | February 26, 2019 5:02:04 PM GMT+08:00 | IOCC_Refresh: Study Description (Procedure), Series Description (Description), Label, Study ID displays invalid characters in UI | WIP12 | WIP15 |  | Subsystem | USP-Connectivity | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP15SubCom | Ovation |  | WIP15 | April 10, 2019 4:41:43 PM GMT+08:00
CTUDT00622065 | Refresh | Closed | swamypreethi | February 26, 2019 5:02:04 PM GMT+08:00 | IOCC_Refresh: Study Description (Procedure), Series Description (Description), Label, Study ID displays invalid characters in UI | WIP12 | WIP15 |  | Subsystem | USP-Connectivity | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP15SubCom | Brilliance 64 |  | WIP15 | April 10, 2019 4:41:43 PM GMT+08:00
CTUDT00622065 | Refresh | Closed | swamypreethi | February 26, 2019 5:02:04 PM GMT+08:00 | IOCC_Refresh: Study Description (Procedure), Series Description (Description), Label, Study ID displays invalid characters in UI | WIP12 | WIP15 |  | Subsystem | USP-Connectivity | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP15SubCom | Brilliance Big Bore |  | WIP15 | April 10, 2019 4:41:43 PM GMT+08:00
CTUDT00622065 | Refresh | Closed | swamypreethi | February 26, 2019 5:02:04 PM GMT+08:00 | IOCC_Refresh: Study Description (Procedure), Series Description (Description), Label, Study ID displays invalid characters in UI | WIP12 | WIP15 |  | Subsystem | USP-Connectivity | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP15SubCom | Ingenuity 64 |  | WIP15 | April 10, 2019 4:41:43 PM GMT+08:00
CTUDT00622065 | Refresh | Closed | swamypreethi | February 26, 2019 5:02:04 PM GMT+08:00 | IOCC_Refresh: Study Description (Procedure), Series Description (Description), Label, Study ID displays invalid characters in UI | WIP12 | WIP15 |  | Subsystem | USP-Connectivity | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP15SubCom | IQon1 |  | WIP15 | April 10, 2019 4:41:43 PM GMT+08:00
CTUDT00622062 | Refresh | Closed | wang_l | February 26, 2019 4:38:44 PM GMT+08:00 | BB Refresh :an error occurred” was displayed for each calibration item once Harmonized Phantom is checked | BB_WIP4 |  |  | Subsystem |  |  | Duplicate | yatziv_l | 0 |  |  | DMS | NO | Closed |  | Brilliance Big Bore
CTUDT00622060 | Refresh | Closed | sowmya_k | February 26, 2019 4:12:12 PM GMT+08:00 | Refresh:-Recon:-User was unable to perform the reconstruction on Dual surview. | BB_WIP4 | BB_WIP5 |  | Subsystem | Reconstruction | NA | Fixed | v_sumanth | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP5_SubCom | Brilliance Big Bore |  | BB_WIP5 | March 4, 2019 5:28:45 PM GMT+08:00
CTUDT00622052 | Refresh | Closed | kapadia_v | February 26, 2019 3:48:06 PM GMT+08:00 | Refresh:Plan- After performing dual surview, changing the patient position and repeating the surview, message displays to repeat the surview again | BB_WIP1 | BB_WIP5 |  | Subsystem | Repeat Scan | N/A | Fixed | sagar_v | 0 | 2 | Console Software | Console4 | YES | Closed | BB_WIP5_SubCom | Brilliance Big Bore |  | BB_WIP5 | March 14, 2019 1:18:05 PM GMT+08:00
CTUDT00622050 | Refresh | Closed | roshan_r | February 26, 2019 3:23:34 PM GMT+08:00 | BB Refresh : An additional Threshold parameter named “Voltage_threshold” with value is available in the IMRThresholds CFG file. | BB_WIP3 | BB_WIP5.1 |  | Subsystem | IMR | None | Fixed | prasad_r | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP5_SubCom | Brilliance Big Bore |  | BB_WIP5.1 | March 15, 2019 4:50:47 PM GMT+08:00
CTUDT00622049 | Refresh | Closed | patra_r | February 26, 2019 2:47:19 PM GMT+08:00 | IV: Solo Cycle 2 - Brain result will not angle to the right | V4.8 | BB_WIP5 | For Mouse rotation from zero degrees(clockwise will not work) - So use keyboard For Keyboard typing rotation from zero degrees (anti clockwise will not work) - Use mouse for rotation. | Subsystem |  | N/A | Fixed | patra_r | 0 | 3 | Console Software | Console4 | YES | Closed | BB_WIP5_SubCom | Brilliance Big Bore |  | BB_WIP5 | March 7, 2019 6:35:33 PM GMT+08:00
CTUDT00622042 | Refresh | Closed | patra_r | February 26, 2019 1:24:03 PM GMT+08:00 | Refresh- User Name value in Dose Log Viewer is not recorded. | V4.8 | BB_WIP5.1 |  | System | Dose Log Viewer | N/A | Fixed | chokkalingam_r | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP5_SubCom | Brilliance Big Bore |  | BB_WIP5.1 | March 18, 2019 6:52:51 PM GMT+08:00
CTUDT00622040 | Refresh | Closed | beck_d | February 25, 2019 11:22:52 PM GMT+08:00 | Refresh: Module NVRAM Manager: Red zone BDs that failed during BDC do not appear in the non- compliant table | WIP13 | N/A |  | Subsystem | Bad Detectors | N/A | Fixed | yosfan_s | 0 | 2 | Test Design | DMS | NO | Closed | WIP15SegReq | Ovation |  | N/A | April 11, 2019 8:45:37 PM GMT+08:00
CTUDT00622038 | Refresh | Closed | Noy_Seelig | February 25, 2019 10:43:45 PM GMT+08:00 | Refresh: Wrong Title for Reasons: DMC Board Replacement | WIP13 | DHF232540 Rev11 |  | Subsystem | Reason Matrix | n/a | Fixed | richards_b | 0 | 2 | Design Documentation | Support-Physics | NO | Closed | WIP15SegReq | Ingenuity 64 |  | DHF232540 Rev11 | April 27, 2019 4:24:30 AM GMT+08:00
CTUDT00622036 | Refresh | Closed | attali_s | February 25, 2019 10:10:47 PM GMT+08:00 | Refresh: HCOR Calibration: Adult Head current factors are not copied from Infant Body current factors | WIP13 | WIP15 | N/A | Subsystem | DMS Calibrations | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP15SubCom | Brilliance Big Bore |  | WIP15 | March 25, 2019 8:13:03 PM GMT+08:00
CTUDT00622036 | Refresh | Closed | attali_s | February 25, 2019 10:10:47 PM GMT+08:00 | Refresh: HCOR Calibration: Adult Head current factors are not copied from Infant Body current factors | WIP13 | WIP15 | N/A | Subsystem | DMS Calibrations | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP15SubCom | Ingenuity 64 |  | WIP15 | March 25, 2019 8:13:03 PM GMT+08:00
CTUDT00622036 | Refresh | Closed | attali_s | February 25, 2019 10:10:47 PM GMT+08:00 | Refresh: HCOR Calibration: Adult Head current factors are not copied from Infant Body current factors | WIP13 | WIP15 | N/A | Subsystem | DMS Calibrations | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP15SubCom | IQon1 |  | WIP15 | March 25, 2019 8:13:03 PM GMT+08:00
CTUDT00622034 | Refresh | Closed | attali_s | February 25, 2019 9:45:04 PM GMT+08:00 | Refresh: HCOR Calibration: Wrong Scan Parameters in HCOR Slice | WIP13 | WIP17 |  | Subsystem | DMS Calibrations | n/a | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP15SubCom | Brilliance Big Bore |  | WIP17 | May 21, 2019 8:37:08 PM GMT+08:00
CTUDT00622034 | Refresh | Closed | attali_s | February 25, 2019 9:45:04 PM GMT+08:00 | Refresh: HCOR Calibration: Wrong Scan Parameters in HCOR Slice | WIP13 | WIP17 |  | Subsystem | DMS Calibrations | n/a | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP15SubCom | Ingenuity 64 |  | WIP17 | May 21, 2019 8:37:08 PM GMT+08:00
CTUDT00622027 | Refresh | Closed | patra_r | February 25, 2019 7:26:16 PM GMT+08:00 | BBRefresh: Less images generated than planned, when the result is loaded to slab viewer in view and view2 and performed recon. | V4.1.10 & V4.7.7 | WIP15 | Yes, User can typed-in the values in thickness and increment field in UI. | System | Slab Viewer | None | Fixed | sagar_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP15SegReq | Ovation |  | WIP15 | April 1, 2019 11:04:04 PM GMT+08:00
CTUDT00622027 | Refresh | Closed | patra_r | February 25, 2019 7:26:16 PM GMT+08:00 | BBRefresh: Less images generated than planned, when the result is loaded to slab viewer in view and view2 and performed recon. | V4.1.10 & V4.7.7 | WIP15 | Yes, User can typed-in the values in thickness and increment field in UI. | System | Slab Viewer | None | Fixed | sagar_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP15SegReq | Brilliance 64 |  | WIP15 | April 1, 2019 11:04:04 PM GMT+08:00
CTUDT00622027 | Refresh | Closed | patra_r | February 25, 2019 7:26:16 PM GMT+08:00 | BBRefresh: Less images generated than planned, when the result is loaded to slab viewer in view and view2 and performed recon. | V4.1.10 & V4.7.7 | WIP15 | Yes, User can typed-in the values in thickness and increment field in UI. | System | Slab Viewer | None | Fixed | sagar_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP15SegReq | Ingenuity 64 |  | WIP15 | April 1, 2019 11:04:04 PM GMT+08:00
CTUDT00622027 | Refresh | Closed | patra_r | February 25, 2019 7:26:16 PM GMT+08:00 | BBRefresh: Less images generated than planned, when the result is loaded to slab viewer in view and view2 and performed recon. | V4.1.10 & V4.7.7 | WIP15 | Yes, User can typed-in the values in thickness and increment field in UI. | System | Slab Viewer | None | Fixed | sagar_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP15SegReq | IQon1 |  | WIP15 | April 1, 2019 11:04:04 PM GMT+08:00
CTUDT00622024 | Refresh | Closed | Sharmila_Vasireddy | February 25, 2019 5:48:28 PM GMT+08:00 | System Error after completion of the 1st timed scan | BB_WIP3 |  |  | System | Simulator |  | Works As Specified / Not A Defect | blessily | 1 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00622024 | Refresh | Closed | Sharmila_Vasireddy | February 25, 2019 5:48:28 PM GMT+08:00 | System Error after completion of the 1st timed scan | BB_WIP3 |  |  | System | Simulator |  | Works As Specified / Not A Defect | blessily | 1 |  |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00622024 | Refresh | Closed | Sharmila_Vasireddy | February 25, 2019 5:48:28 PM GMT+08:00 | System Error after completion of the 1st timed scan | BB_WIP3 |  |  | System | Simulator |  | Works As Specified / Not A Defect | blessily | 1 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00622024 | Refresh | Closed | Sharmila_Vasireddy | February 25, 2019 5:48:28 PM GMT+08:00 | System Error after completion of the 1st timed scan | BB_WIP3 |  |  | System | Simulator |  | Works As Specified / Not A Defect | blessily | 1 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00622024 | Refresh | Closed | Sharmila_Vasireddy | February 25, 2019 5:48:28 PM GMT+08:00 | System Error after completion of the 1st timed scan | BB_WIP3 |  |  | System | Simulator |  | Works As Specified / Not A Defect | blessily | 1 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00622023 | Refresh | Closed | attali_s | February 25, 2019 5:41:22 PM GMT+08:00 | Refresh: Constancy/Acceptance: Wrong Uniformity ROIs' diameter for case Adult Body on Head | WIP13 | WIP15 |  | Subsystem | DMS Calibrations | N/A | Fixed | Kostya_Baklan | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP15SubCom | Ingenuity 64 |  | WIP15 | March 24, 2019 5:15:12 PM GMT+08:00
CTUDT00622021 | Refresh | Closed | finestine_m | February 25, 2019 4:24:36 PM GMT+08:00 | Refresh: Common Functions: No Zero.tbl in the Scan's MDU | WIP13 |  |  | Subsystem | Common Functions | n/a | Duplicate | richards_b | 0 |  |  | Support-Physics | NO | Closed |  | Ingenuity 64
CTUDT00622014 | Refresh | Closed | attali_s | February 25, 2019 3:55:28 PM GMT+08:00 | Refresh: HCOR Calibration: Wrong Customization Window for HCOR Infant | WIP13 | N/A - Test Design |  | Subsystem | HCOR | n/a | Fixed | gehtin_e | 0 | 4 | Test Design | DMS | NO | Closed | WIP15SubCom | Ingenuity 64 |  | N/A - Test Design | May 14, 2019 10:11:14 PM GMT+08:00
CTUDT00622007 | Refresh | Closed | patra_r | February 25, 2019 2:24:07 PM GMT+08:00 | Refresh: For Dual Surview, Auto filming not working as expected. | V4.8 | BB_WIP5.1 |  | System | Dual Surview | N/A | Fixed | kr_m | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP5_SubCom | Brilliance Big Bore |  | BB_WIP5.1 | April 1, 2019 8:45:09 PM GMT+08:00
CTUDT00622006 | Refresh | Closed | patra_r | February 25, 2019 2:15:02 PM GMT+08:00 | BB Refresh: In ECM,selection of specific scantype is not getting allowed | V4.1.10 | WIP15 | NA | Subsystem | ECM | N/A | Fixed | navale_s | 0 | 3 | Console Software | Console4 | NO | Closed | WIP15SubCom | Ovation |  | WIP15 | April 1, 2019 1:01:59 PM GMT+08:00
CTUDT00622006 | Refresh | Closed | patra_r | February 25, 2019 2:15:02 PM GMT+08:00 | BB Refresh: In ECM,selection of specific scantype is not getting allowed | V4.1.10 | WIP15 | NA | Subsystem | ECM | N/A | Fixed | navale_s | 0 | 3 | Console Software | Console4 | NO | Closed | WIP15SubCom | Brilliance 64 |  | WIP15 | April 1, 2019 1:01:59 PM GMT+08:00
CTUDT00622006 | Refresh | Closed | patra_r | February 25, 2019 2:15:02 PM GMT+08:00 | BB Refresh: In ECM,selection of specific scantype is not getting allowed | V4.1.10 | WIP15 | NA | Subsystem | ECM | N/A | Fixed | navale_s | 0 | 3 | Console Software | Console4 | NO | Closed | WIP15SubCom | Ingenuity 64 |  | WIP15 | April 1, 2019 1:01:59 PM GMT+08:00
CTUDT00622001 | Refresh | Closed | cohen_y | February 25, 2019 3:11:14 AM GMT+08:00 | Refresh: Acceptance/Constancy: Protocol type of Thickness cases is Head instead of Body. | WIP13 | WIP15 | N/A | Subsystem | Acceptance/Constancy | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP15SubCom | Ovation |  | WIP15 | March 25, 2019 8:53:16 PM GMT+08:00
CTUDT00621999 | Refresh | Closed | cohen_y | February 25, 2019 3:04:35 AM GMT+08:00 | Refresh: Acceptance/Constancy: “According To” heading does not appear in the test report summary. | WIP13 | WIP17 |  | Subsystem | DMS Calibrations | N/A | Fixed | Kostya_Baklan | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP15SubCom | Ovation |  | WIP17 | March 26, 2019 3:45:02 PM GMT+08:00
CTUDT00621999 | Refresh | Closed | cohen_y | February 25, 2019 3:04:35 AM GMT+08:00 | Refresh: Acceptance/Constancy: “According To” heading does not appear in the test report summary. | WIP13 | WIP17 |  | Subsystem | DMS Calibrations | N/A | Fixed | Kostya_Baklan | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP15SubCom | Brilliance Big Bore |  | WIP17 | March 26, 2019 3:45:02 PM GMT+08:00
CTUDT00621999 | Refresh | Closed | cohen_y | February 25, 2019 3:04:35 AM GMT+08:00 | Refresh: Acceptance/Constancy: “According To” heading does not appear in the test report summary. | WIP13 | WIP17 |  | Subsystem | DMS Calibrations | N/A | Fixed | Kostya_Baklan | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP15SubCom | Ingenuity 64 |  | WIP17 | March 26, 2019 3:45:02 PM GMT+08:00
CTUDT00621999 | Refresh | Closed | cohen_y | February 25, 2019 3:04:35 AM GMT+08:00 | Refresh: Acceptance/Constancy: “According To” heading does not appear in the test report summary. | WIP13 | WIP17 |  | Subsystem | DMS Calibrations | N/A | Fixed | Kostya_Baklan | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP15SubCom | IQon1 |  | WIP17 | March 26, 2019 3:45:02 PM GMT+08:00
CTUDT00621996 | Refresh | Closed | Valeria_Kashtanov | February 25, 2019 12:45:45 AM GMT+08:00 | Refresh: Air Calibration: Wrong scan Rotation Time when choosing 0.42 Rotation Time in Customization | WIP13 |  |  | Subsystem | Air Calib |  | Duplicate | frider_o | 0 |  |  | DMS | NO | Closed |  | Ingenuity 64
CTUDT00621994 | Refresh | Closed | cohen_y | February 25, 2019 12:38:16 AM GMT+08:00 | Refresh: Acceptance/Constancy: Center Y value of the Resolution cases is not according to Spec | WIP13 | WIP15 |  | Subsystem | DMS Calibrations | n/a | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP15SegReq | Ovation |  | WIP15 | March 21, 2019 8:39:49 PM GMT+08:00
CTUDT00621994 | Refresh | Closed | cohen_y | February 25, 2019 12:38:16 AM GMT+08:00 | Refresh: Acceptance/Constancy: Center Y value of the Resolution cases is not according to Spec | WIP13 | WIP15 |  | Subsystem | DMS Calibrations | n/a | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP15SegReq | Brilliance Big Bore |  | WIP15 | March 21, 2019 8:39:49 PM GMT+08:00
CTUDT00621991 | Refresh | Closed | Dana_Kamyshev | February 24, 2019 10:09:37 PM GMT+08:00 | Refresh - Stroke EC completed with more than 5.5 minutes with tolerance of 5% | WIP13 |  |  | Subsystem | Performance |  | Works As Specified / Not A Defect | tamari_o | 1 |  |  | Support-System | NO | Closed |  | IQon1
CTUDT00621987 | Refresh | Closed | bagdady_b | February 24, 2019 7:04:58 PM GMT+08:00 | Refresh- When performing import of IQon HB EC to IQon refresh some of the parameters are changed | WIP13 |  |  | Subsystem | Exam Prep | NA | Works As Specified / Not A Defect | sanjeeva_a | 0 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00621981 | Refresh | Closed | veeraswamy_k | February 22, 2019 6:05:18 PM GMT+08:00 | CD/DVD burned on scanner does not open in the Windows 10 laptop with version 10.0.17134 | 4.1.10 | WIP15 |  | Subsystem | Service Tools | NA | Fixed | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP15SubCom | Ovation | . | WIP15 | April 4, 2019 4:03:15 PM GMT+08:00
CTUDT00621981 | Refresh | Closed | veeraswamy_k | February 22, 2019 6:05:18 PM GMT+08:00 | CD/DVD burned on scanner does not open in the Windows 10 laptop with version 10.0.17134 | 4.1.10 | WIP15 |  | Subsystem | Service Tools | NA | Fixed | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP15SubCom | Brilliance 64 | . | WIP15 | April 4, 2019 4:03:15 PM GMT+08:00
CTUDT00621981 | Refresh | Closed | veeraswamy_k | February 22, 2019 6:05:18 PM GMT+08:00 | CD/DVD burned on scanner does not open in the Windows 10 laptop with version 10.0.17134 | 4.1.10 | WIP15 |  | Subsystem | Service Tools | NA | Fixed | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP15SubCom | Brilliance Big Bore | . | WIP15 | April 4, 2019 4:03:15 PM GMT+08:00
CTUDT00621981 | Refresh | Closed | veeraswamy_k | February 22, 2019 6:05:18 PM GMT+08:00 | CD/DVD burned on scanner does not open in the Windows 10 laptop with version 10.0.17134 | 4.1.10 | WIP15 |  | Subsystem | Service Tools | NA | Fixed | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP15SubCom | Ingenuity 64 | . | WIP15 | April 4, 2019 4:03:15 PM GMT+08:00
CTUDT00621981 | Refresh | Closed | veeraswamy_k | February 22, 2019 6:05:18 PM GMT+08:00 | CD/DVD burned on scanner does not open in the Windows 10 laptop with version 10.0.17134 | 4.1.10 | WIP15 |  | Subsystem | Service Tools | NA | Fixed | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP15SubCom | IQon1 | . | WIP15 | April 4, 2019 4:03:15 PM GMT+08:00
CTUDT00621975 | Refresh | Closed | patra_r | February 22, 2019 2:19:50 PM GMT+08:00 | Refresh - Study doesn’t appear in History Database list when user performs ‘End Exam and Continue Working’ | V4.8 | DHF316705 Rev08 |  | Subsystem | Host | Console-SRS.14661 | Fixed | roy_r | 0 | 2 | Host SSRS | Console4 | NO | Closed | BB_WIP5_SubCom | Brilliance Big Bore |  | DHF316705 Rev08 | April 5, 2019 4:35:14 PM GMT+08:00
CTUDT00621969 | Refresh | Closed | Sateesh_S | February 22, 2019 12:22:50 PM GMT+08:00 | BB Refresh :Selecting any configured devices in dose structured report in Preferences->Dose management tab, then local was not selected by default | BB_WIP3 | BB_WIP7 | None | Subsystem | HOST | Console-SRS.3405, Console-SRS.3406, Console-SRS.10681 | Fixed | mY_h | 0 | 4 | Console Software | Console4 | NO | Closed | SSRS | Brilliance Big Bore |  | BB_WIP7 | May 29, 2019 6:09:03 PM GMT+08:00
CTUDT00621967 | Refresh | Closed | roshan_r | February 22, 2019 12:40:12 AM GMT+08:00 | BB Refresh : The Console does not prevent the user from performing a perfusion scan when the accumulated planned CTDI exceeds 500mGy. | BB_WIP3 | BB_WIP5.1 |  | Subsystem | HOST | NA | Fixed | datta_i | 0 | 2 | Console Software | Console4 | YES | Closed | WIP15SubCom | Brilliance Big Bore |  | BB_WIP5.1 | April 3, 2019 6:42:29 PM GMT+08:00
CTUDT00621959 | Refresh | Closed | alex_nihaev | February 21, 2019 6:54:53 PM GMT+08:00 | Refresh: Phantom calibration, Error message after starting Phantom calibration | BB_WIP3 |  |  | Subsystem | DMS Calibrations | n/a | No Plan to Fix | yatziv_l | 0 | 4 |  | DMS | NO | Closed | BB_SubNC | Brilliance Big Bore
CTUDT00621956 | Refresh | Closed | nimma_p | February 21, 2019 3:39:20 PM GMT+08:00 | Refresh:-IQON:- SBI series not transferred after final Recon to newly added local folder & remote folder. | WIP13 |  |  | Subsystem | USP |  | Works As Specified / Not A Defect | veerabhadraiah | 1 |  |  | iCAP | NO | Closed |  | IQon1
CTUDT00621939 | Refresh | Closed | levy_s | February 20, 2019 8:29:31 PM GMT+08:00 | Refresh: ESTOP opens infrequently at the beginning of scan preparation | WIP13 |  |  | Subsystem | Gantry | None | No Plan to Fix | karkour | 0 | 3 |  | GCX | NO | Closed |  | Ovation
CTUDT00621939 | Refresh | Closed | levy_s | February 20, 2019 8:29:31 PM GMT+08:00 | Refresh: ESTOP opens infrequently at the beginning of scan preparation | WIP13 |  |  | Subsystem | Gantry | None | No Plan to Fix | karkour | 0 | 3 |  | GCX | NO | Closed |  | IQon1
CTUDT00621934 | Refresh | Closed | Noy_Seelig | February 20, 2019 7:48:32 PM GMT+08:00 | Refresh: Phantom calibration, wrong bed position in Instructions window prevents calibration from completing first case | BB_WIP4 |  |  | Subsystem | DMS Calibrations | n/a | Works As Specified / Not A Defect | yosfan_s | 1 |  |  | DMS | NO | Closed |  | Brilliance Big Bore
CTUDT00621933 | Refresh | Closed | shaham_t | February 20, 2019 7:46:39 PM GMT+08:00 | ANSAN:System error while drawing ROI on locator in bolus tracking scan | 4.7.5 | WIP14 |  | System | Verified1 | N/A | Fixed | redrouthu_r | 0 | 2 | Console USP_EBW | iCAP | YES | Closed | WIP14SegReq | Ovation | CFR#9472222 | WIP14 | March 14, 2019 5:35:28 PM GMT+08:00
CTUDT00621933 | Refresh | Closed | shaham_t | February 20, 2019 7:46:39 PM GMT+08:00 | ANSAN:System error while drawing ROI on locator in bolus tracking scan | 4.7.5 | WIP14 |  | System | Verified1 | N/A | Fixed | redrouthu_r | 0 | 2 | Console USP_EBW | iCAP | YES | Closed | WIP14SegReq | Brilliance 64 | CFR#9472222 | WIP14 | March 14, 2019 5:35:28 PM GMT+08:00
CTUDT00621933 | Refresh | Closed | shaham_t | February 20, 2019 7:46:39 PM GMT+08:00 | ANSAN:System error while drawing ROI on locator in bolus tracking scan | 4.7.5 | WIP14 |  | System | Verified1 | N/A | Fixed | redrouthu_r | 0 | 2 | Console USP_EBW | iCAP | YES | Closed | WIP14SegReq | Brilliance Big Bore | CFR#9472222 | WIP14 | March 14, 2019 5:35:28 PM GMT+08:00
CTUDT00621933 | Refresh | Closed | shaham_t | February 20, 2019 7:46:39 PM GMT+08:00 | ANSAN:System error while drawing ROI on locator in bolus tracking scan | 4.7.5 | WIP14 |  | System | Verified1 | N/A | Fixed | redrouthu_r | 0 | 2 | Console USP_EBW | iCAP | YES | Closed | WIP14SegReq | Ingenuity 64 | CFR#9472222 | WIP14 | March 14, 2019 5:35:28 PM GMT+08:00
CTUDT00621933 | Refresh | Closed | shaham_t | February 20, 2019 7:46:39 PM GMT+08:00 | ANSAN:System error while drawing ROI on locator in bolus tracking scan | 4.7.5 | WIP14 |  | System | Verified1 | N/A | Fixed | redrouthu_r | 0 | 2 | Console USP_EBW | iCAP | YES | Closed | WIP14SegReq | IQon1 | CFR#9472222 | WIP14 | March 14, 2019 5:35:28 PM GMT+08:00
CTUDT00621932 | Refresh | Closed | veerabhadraiah | February 20, 2019 7:20:38 PM GMT+08:00 | Failed transaction in Queue manager shows improper retry time, 553+ Hours | 4.7.5 | WIP14 | Use Manual Retry option from Queue Manager | Subsystem |  | N/A | Fixed | veerabhadraiah | 0 | 4 | Console USP_EBW | iCAP | NO | Closed | WIP15SubCom | Ovation | 8758725 | WIP14 | March 12, 2019 6:12:16 PM GMT+08:00
CTUDT00621932 | Refresh | Closed | veerabhadraiah | February 20, 2019 7:20:38 PM GMT+08:00 | Failed transaction in Queue manager shows improper retry time, 553+ Hours | 4.7.5 | WIP14 | Use Manual Retry option from Queue Manager | Subsystem |  | N/A | Fixed | veerabhadraiah | 0 | 4 | Console USP_EBW | iCAP | NO | Closed | WIP15SubCom | Brilliance 64 | 8758725 | WIP14 | March 12, 2019 6:12:16 PM GMT+08:00
CTUDT00621932 | Refresh | Closed | veerabhadraiah | February 20, 2019 7:20:38 PM GMT+08:00 | Failed transaction in Queue manager shows improper retry time, 553+ Hours | 4.7.5 | WIP14 | Use Manual Retry option from Queue Manager | Subsystem |  | N/A | Fixed | veerabhadraiah | 0 | 4 | Console USP_EBW | iCAP | NO | Closed | WIP15SubCom | Brilliance Big Bore | 8758725 | WIP14 | March 12, 2019 6:12:16 PM GMT+08:00
CTUDT00621932 | Refresh | Closed | veerabhadraiah | February 20, 2019 7:20:38 PM GMT+08:00 | Failed transaction in Queue manager shows improper retry time, 553+ Hours | 4.7.5 | WIP14 | Use Manual Retry option from Queue Manager | Subsystem |  | N/A | Fixed | veerabhadraiah | 0 | 4 | Console USP_EBW | iCAP | NO | Closed | WIP15SubCom | Ingenuity 64 | 8758725 | WIP14 | March 12, 2019 6:12:16 PM GMT+08:00
CTUDT00621932 | Refresh | Closed | veerabhadraiah | February 20, 2019 7:20:38 PM GMT+08:00 | Failed transaction in Queue manager shows improper retry time, 553+ Hours | 4.7.5 | WIP14 | Use Manual Retry option from Queue Manager | Subsystem |  | N/A | Fixed | veerabhadraiah | 0 | 4 | Console USP_EBW | iCAP | NO | Closed | WIP15SubCom | IQon1 | 8758725 | WIP14 | March 12, 2019 6:12:16 PM GMT+08:00
CTUDT00621931 | Refresh | Closed | veerabhadraiah | February 20, 2019 7:19:58 PM GMT+08:00 | IQoN-HB: User need to restart the system | 4.7.5 | WIP14 |  | Subsystem |  | N/A | Fixed | veerabhadraiah | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP14SegReq | Ovation | CFR #9473459 CFR #9458915 PR7513235 PR8084011 | WIP14 | March 12, 2019 6:11:57 PM GMT+08:00
CTUDT00621931 | Refresh | Closed | veerabhadraiah | February 20, 2019 7:19:58 PM GMT+08:00 | IQoN-HB: User need to restart the system | 4.7.5 | WIP14 |  | Subsystem |  | N/A | Fixed | veerabhadraiah | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP14SegReq | Brilliance 64 | CFR #9473459 CFR #9458915 PR7513235 PR8084011 | WIP14 | March 12, 2019 6:11:57 PM GMT+08:00
CTUDT00621931 | Refresh | Closed | veerabhadraiah | February 20, 2019 7:19:58 PM GMT+08:00 | IQoN-HB: User need to restart the system | 4.7.5 | WIP14 |  | Subsystem |  | N/A | Fixed | veerabhadraiah | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP14SegReq | Brilliance Big Bore | CFR #9473459 CFR #9458915 PR7513235 PR8084011 | WIP14 | March 12, 2019 6:11:57 PM GMT+08:00
CTUDT00621931 | Refresh | Closed | veerabhadraiah | February 20, 2019 7:19:58 PM GMT+08:00 | IQoN-HB: User need to restart the system | 4.7.5 | WIP14 |  | Subsystem |  | N/A | Fixed | veerabhadraiah | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP14SegReq | Ingenuity 64 | CFR #9473459 CFR #9458915 PR7513235 PR8084011 | WIP14 | March 12, 2019 6:11:57 PM GMT+08:00
CTUDT00621931 | Refresh | Closed | veerabhadraiah | February 20, 2019 7:19:58 PM GMT+08:00 | IQoN-HB: User need to restart the system | 4.7.5 | WIP14 |  | Subsystem |  | N/A | Fixed | veerabhadraiah | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP14SegReq | IQon1 | CFR #9473459 CFR #9458915 PR7513235 PR8084011 | WIP14 | March 12, 2019 6:11:57 PM GMT+08:00
CTUDT00621930 | Refresh | Closed | veerabhadraiah | February 20, 2019 7:18:33 PM GMT+08:00 | Image transfer to PACS is Failed | 4.7.5 | WIP14 | Delete the failed transaction from QM and go to PD. Manually send the images from local directory to desired nodes. | Subsystem | Connectivity | N/A | Fixed | veerabhadraiah | 0 | 4 | Console USP_EBW | iCAP | NO | Closed | WIP14SegReq | Ovation | 8758725 | WIP14 | March 12, 2019 6:12:55 PM GMT+08:00
CTUDT00621930 | Refresh | Closed | veerabhadraiah | February 20, 2019 7:18:33 PM GMT+08:00 | Image transfer to PACS is Failed | 4.7.5 | WIP14 | Delete the failed transaction from QM and go to PD. Manually send the images from local directory to desired nodes. | Subsystem | Connectivity | N/A | Fixed | veerabhadraiah | 0 | 4 | Console USP_EBW | iCAP | NO | Closed | WIP14SegReq | Brilliance 64 | 8758725 | WIP14 | March 12, 2019 6:12:55 PM GMT+08:00
CTUDT00621930 | Refresh | Closed | veerabhadraiah | February 20, 2019 7:18:33 PM GMT+08:00 | Image transfer to PACS is Failed | 4.7.5 | WIP14 | Delete the failed transaction from QM and go to PD. Manually send the images from local directory to desired nodes. | Subsystem | Connectivity | N/A | Fixed | veerabhadraiah | 0 | 4 | Console USP_EBW | iCAP | NO | Closed | WIP14SegReq | Brilliance Big Bore | 8758725 | WIP14 | March 12, 2019 6:12:55 PM GMT+08:00
CTUDT00621930 | Refresh | Closed | veerabhadraiah | February 20, 2019 7:18:33 PM GMT+08:00 | Image transfer to PACS is Failed | 4.7.5 | WIP14 | Delete the failed transaction from QM and go to PD. Manually send the images from local directory to desired nodes. | Subsystem | Connectivity | N/A | Fixed | veerabhadraiah | 0 | 4 | Console USP_EBW | iCAP | NO | Closed | WIP14SegReq | Ingenuity 64 | 8758725 | WIP14 | March 12, 2019 6:12:55 PM GMT+08:00
CTUDT00621930 | Refresh | Closed | veerabhadraiah | February 20, 2019 7:18:33 PM GMT+08:00 | Image transfer to PACS is Failed | 4.7.5 | WIP14 | Delete the failed transaction from QM and go to PD. Manually send the images from local directory to desired nodes. | Subsystem | Connectivity | N/A | Fixed | veerabhadraiah | 0 | 4 | Console USP_EBW | iCAP | NO | Closed | WIP14SegReq | IQon1 | 8758725 | WIP14 | March 12, 2019 6:12:55 PM GMT+08:00
CTUDT00621929 | Refresh | Closed | veerabhadraiah | February 20, 2019 7:17:49 PM GMT+08:00 | Sometimes while storing result images to remote device, PmsCT Remoting service throws exception, following which Console throws socket exception and it crashes. | 4.7.5 | WIP14 |  | Subsystem | PmsCT Remoting (Patch fix) | N/A | Fixed | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP14SegReq | Ovation | 9458915 PR10033730 | WIP14 | March 12, 2019 6:12:29 PM GMT+08:00
CTUDT00621929 | Refresh | Closed | veerabhadraiah | February 20, 2019 7:17:49 PM GMT+08:00 | Sometimes while storing result images to remote device, PmsCT Remoting service throws exception, following which Console throws socket exception and it crashes. | 4.7.5 | WIP14 |  | Subsystem | PmsCT Remoting (Patch fix) | N/A | Fixed | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP14SegReq | Brilliance 64 | 9458915 PR10033730 | WIP14 | March 12, 2019 6:12:29 PM GMT+08:00
CTUDT00621929 | Refresh | Closed | veerabhadraiah | February 20, 2019 7:17:49 PM GMT+08:00 | Sometimes while storing result images to remote device, PmsCT Remoting service throws exception, following which Console throws socket exception and it crashes. | 4.7.5 | WIP14 |  | Subsystem | PmsCT Remoting (Patch fix) | N/A | Fixed | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP14SegReq | Brilliance Big Bore | 9458915 PR10033730 | WIP14 | March 12, 2019 6:12:29 PM GMT+08:00
CTUDT00621929 | Refresh | Closed | veerabhadraiah | February 20, 2019 7:17:49 PM GMT+08:00 | Sometimes while storing result images to remote device, PmsCT Remoting service throws exception, following which Console throws socket exception and it crashes. | 4.7.5 | WIP14 |  | Subsystem | PmsCT Remoting (Patch fix) | N/A | Fixed | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP14SegReq | Ingenuity 64 | 9458915 PR10033730 | WIP14 | March 12, 2019 6:12:29 PM GMT+08:00
CTUDT00621929 | Refresh | Closed | veerabhadraiah | February 20, 2019 7:17:49 PM GMT+08:00 | Sometimes while storing result images to remote device, PmsCT Remoting service throws exception, following which Console throws socket exception and it crashes. | 4.7.5 | WIP14 |  | Subsystem | PmsCT Remoting (Patch fix) | N/A | Fixed | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP14SegReq | IQon1 | 9458915 PR10033730 | WIP14 | March 12, 2019 6:12:29 PM GMT+08:00
CTUDT00621928 | Refresh | Closed | veerabhadraiah | February 20, 2019 7:15:27 PM GMT+08:00 | When system is configured for auto image transfer to several destinations such as PACS, ISP9, mini PACS, sometimes images were not sent to PACS or ISP9. Host system restart is needed to resolve this. | 4.7.5 | WIP14 | Restart the machine | Subsystem | Auto Transfer (Patch Fix) | N/A | Fixed | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP14SegReq | Ovation | 8758725 | WIP14 | March 12, 2019 6:11:42 PM GMT+08:00
CTUDT00621928 | Refresh | Closed | veerabhadraiah | February 20, 2019 7:15:27 PM GMT+08:00 | When system is configured for auto image transfer to several destinations such as PACS, ISP9, mini PACS, sometimes images were not sent to PACS or ISP9. Host system restart is needed to resolve this. | 4.7.5 | WIP14 | Restart the machine | Subsystem | Auto Transfer (Patch Fix) | N/A | Fixed | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP14SegReq | Brilliance 64 | 8758725 | WIP14 | March 12, 2019 6:11:42 PM GMT+08:00
CTUDT00621928 | Refresh | Closed | veerabhadraiah | February 20, 2019 7:15:27 PM GMT+08:00 | When system is configured for auto image transfer to several destinations such as PACS, ISP9, mini PACS, sometimes images were not sent to PACS or ISP9. Host system restart is needed to resolve this. | 4.7.5 | WIP14 | Restart the machine | Subsystem | Auto Transfer (Patch Fix) | N/A | Fixed | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP14SegReq | Brilliance Big Bore | 8758725 | WIP14 | March 12, 2019 6:11:42 PM GMT+08:00
CTUDT00621928 | Refresh | Closed | veerabhadraiah | February 20, 2019 7:15:27 PM GMT+08:00 | When system is configured for auto image transfer to several destinations such as PACS, ISP9, mini PACS, sometimes images were not sent to PACS or ISP9. Host system restart is needed to resolve this. | 4.7.5 | WIP14 | Restart the machine | Subsystem | Auto Transfer (Patch Fix) | N/A | Fixed | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP14SegReq | Ingenuity 64 | 8758725 | WIP14 | March 12, 2019 6:11:42 PM GMT+08:00
CTUDT00621928 | Refresh | Closed | veerabhadraiah | February 20, 2019 7:15:27 PM GMT+08:00 | When system is configured for auto image transfer to several destinations such as PACS, ISP9, mini PACS, sometimes images were not sent to PACS or ISP9. Host system restart is needed to resolve this. | 4.7.5 | WIP14 | Restart the machine | Subsystem | Auto Transfer (Patch Fix) | N/A | Fixed | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP14SegReq | IQon1 | 8758725 | WIP14 | March 12, 2019 6:11:42 PM GMT+08:00
CTUDT00621920 | Refresh | Closed | Revital_Cohen | February 20, 2019 5:06:33 PM GMT+08:00 | iQon Refresh - Second Scan enter the Recon queue and the result for that Scan doesn’t exist in the study tab | WIP13 |  |  | Subsystem | Reconstruction | NA | Works As Specified / Not A Defect | tamari_o | 1 |  |  | Support-System | YES | Closed |  | IQon1
CTUDT00621917 | Refresh | Closed | dadon_e | February 20, 2019 4:38:43 PM GMT+08:00 | Tube alignment calibration failed - Error: "unexpected calculation error occurred" | WIP13 |  |  | Subsystem | Calibrations |  | Duplicate | karkour | 0 |  |  | GCX | NO | Closed |  | iCT2
CTUDT00621917 | Refresh | Closed | dadon_e | February 20, 2019 4:38:43 PM GMT+08:00 | Tube alignment calibration failed - Error: "unexpected calculation error occurred" | WIP13 |  |  | Subsystem | Calibrations |  | Duplicate | karkour | 0 |  |  | GCX | NO | Closed |  | IQon1
CTUDT00621916 | Refresh | Closed | Noy_Seelig | February 20, 2019 4:12:19 PM GMT+08:00 | Refresh: Filament test fails due to tube heat. | BB_WIP4 |  | Don't use optional license of 80kw on the bay of 60kw tube power. | Subsystem | Gantry Calibrations | n/a | Invalid/Uncontrolled Configuration | wang_c | 0 |  |  | GCX | NO | Closed |  | Brilliance Big Bore
CTUDT00621906 | Refresh | Closed | jotawar_d | February 19, 2019 7:21:25 PM GMT+08:00 | IMR CIRS connected but raw data missed in Reconstruction tab | BB_WIP2 | BB_WIP5 | N/A | Subsystem | Recon Directory | N/A | Fixed | kr_m | 0 | 2 | Console Software | Console4 | YES | Closed | BB_WIP5_SegReq | Brilliance Big Bore |  | BB_WIP5 | March 7, 2019 9:15:19 PM GMT+08:00
CTUDT00621902 | Refresh | Closed | nayak_r | February 19, 2019 4:48:29 PM GMT+08:00 | When any user is prompted to change the password for the first time login, after adding the password and clicking on “OK”, screen hangs for few secs and then logs in. | BB_WIP4 |  |  | Subsystem | Workflow | NA | No Plan to Fix | goswami_s | 0 | 4 |  | Console4 | NO | Closed |  | Ovation
CTUDT00621902 | Refresh | Closed | nayak_r | February 19, 2019 4:48:29 PM GMT+08:00 | When any user is prompted to change the password for the first time login, after adding the password and clicking on “OK”, screen hangs for few secs and then logs in. | BB_WIP4 |  |  | Subsystem | Workflow | NA | No Plan to Fix | goswami_s | 0 | 4 |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00621902 | Refresh | Closed | nayak_r | February 19, 2019 4:48:29 PM GMT+08:00 | When any user is prompted to change the password for the first time login, after adding the password and clicking on “OK”, screen hangs for few secs and then logs in. | BB_WIP4 |  |  | Subsystem | Workflow | NA | No Plan to Fix | goswami_s | 0 | 4 |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00621902 | Refresh | Closed | nayak_r | February 19, 2019 4:48:29 PM GMT+08:00 | When any user is prompted to change the password for the first time login, after adding the password and clicking on “OK”, screen hangs for few secs and then logs in. | BB_WIP4 |  |  | Subsystem | Workflow | NA | No Plan to Fix | goswami_s | 0 | 4 |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00621902 | Refresh | Closed | nayak_r | February 19, 2019 4:48:29 PM GMT+08:00 | When any user is prompted to change the password for the first time login, after adding the password and clicking on “OK”, screen hangs for few secs and then logs in. | BB_WIP4 |  |  | Subsystem | Workflow | NA | No Plan to Fix | goswami_s | 0 | 4 |  | Console4 | NO | Closed |  | IQon1
CTUDT00621900 | Refresh | Closed | manjunath_mk | February 19, 2019 4:40:08 PM GMT+08:00 | 'Resend Analyzer' utility is displayed in the List of utilities | BB_WIP3 | BB_WIP5 |  | Subsystem | Workflow | Console-SRS.16004 | Fixed | Prakash_Chand_Yadav | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP5_SubCom | Brilliance Big Bore |  | BB_WIP5 | March 8, 2019 5:22:41 PM GMT+08:00
CTUDT00621899 | Refresh | Closed | shaham_t | February 19, 2019 4:05:59 PM GMT+08:00 | IQON 4.7.5:In the generated Dose report, the "Procedure Context" is described as "... with contrast media..." despite non-contrast procedure | 4.7.5 | WIP15 |  | Subsystem |  | N/A | Fixed | atin_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP15SubCom | Ovation | PR8468706 | WIP15 | March 22, 2019 1:37:31 PM GMT+08:00
CTUDT00621899 | Refresh | Closed | shaham_t | February 19, 2019 4:05:59 PM GMT+08:00 | IQON 4.7.5:In the generated Dose report, the "Procedure Context" is described as "... with contrast media..." despite non-contrast procedure | 4.7.5 | WIP15 |  | Subsystem |  | N/A | Fixed | atin_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP15SubCom | Brilliance 64 | PR8468706 | WIP15 | March 22, 2019 1:37:31 PM GMT+08:00
CTUDT00621899 | Refresh | Closed | shaham_t | February 19, 2019 4:05:59 PM GMT+08:00 | IQON 4.7.5:In the generated Dose report, the "Procedure Context" is described as "... with contrast media..." despite non-contrast procedure | 4.7.5 | WIP15 |  | Subsystem |  | N/A | Fixed | atin_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP15SubCom | Ingenuity 64 | PR8468706 | WIP15 | March 22, 2019 1:37:31 PM GMT+08:00
CTUDT00621899 | Refresh | Closed | shaham_t | February 19, 2019 4:05:59 PM GMT+08:00 | IQON 4.7.5:In the generated Dose report, the "Procedure Context" is described as "... with contrast media..." despite non-contrast procedure | 4.7.5 | WIP15 |  | Subsystem |  | N/A | Fixed | atin_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP15SubCom | IQon1 | PR8468706 | WIP15 | March 22, 2019 1:37:31 PM GMT+08:00
CTUDT00621891 | Refresh | Closed | jotawar_d | February 18, 2019 10:04:33 PM GMT+08:00 | Unable to Access Device Manager from Olevel Service Tools | WIP8.1 | WIP18 BB_WIP8 |  | Subsystem | Olevel | STPS-SSRS-858, STPS-SSRS-1214 | Fixed | kanagaraj_n | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP17SubCom | Ovation |  | WIP18 BB_WIP8 | June 18, 2019 6:37:49 PM GMT+08:00
CTUDT00621891 | Refresh | Closed | jotawar_d | February 18, 2019 10:04:33 PM GMT+08:00 | Unable to Access Device Manager from Olevel Service Tools | WIP8.1 | WIP18 BB_WIP8 |  | Subsystem | Olevel | STPS-SSRS-858, STPS-SSRS-1214 | Fixed | kanagaraj_n | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP17SubCom | Brilliance 64 |  | WIP18 BB_WIP8 | June 18, 2019 6:37:49 PM GMT+08:00
CTUDT00621891 | Refresh | Closed | jotawar_d | February 18, 2019 10:04:33 PM GMT+08:00 | Unable to Access Device Manager from Olevel Service Tools | WIP8.1 | WIP18 BB_WIP8 |  | Subsystem | Olevel | STPS-SSRS-858, STPS-SSRS-1214 | Fixed | kanagaraj_n | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP17SubCom | Brilliance Big Bore |  | WIP18 BB_WIP8 | June 18, 2019 6:37:49 PM GMT+08:00
CTUDT00621891 | Refresh | Closed | jotawar_d | February 18, 2019 10:04:33 PM GMT+08:00 | Unable to Access Device Manager from Olevel Service Tools | WIP8.1 | WIP18 BB_WIP8 |  | Subsystem | Olevel | STPS-SSRS-858, STPS-SSRS-1214 | Fixed | kanagaraj_n | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP17SubCom | Ingenuity 64 |  | WIP18 BB_WIP8 | June 18, 2019 6:37:49 PM GMT+08:00
CTUDT00621891 | Refresh | Closed | jotawar_d | February 18, 2019 10:04:33 PM GMT+08:00 | Unable to Access Device Manager from Olevel Service Tools | WIP8.1 | WIP18 BB_WIP8 |  | Subsystem | Olevel | STPS-SSRS-858, STPS-SSRS-1214 | Fixed | kanagaraj_n | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP17SubCom | IQon1 |  | WIP18 BB_WIP8 | June 18, 2019 6:37:49 PM GMT+08:00
CTUDT00621884 | Refresh | Closed | nayak_r | February 18, 2019 8:07:32 PM GMT+08:00 | Refresh/BB Refresh: Surview scan followed by multiple Add-Replace Operations throws bug rep error | BB_WIP4 | BB_WIP6 | Select the step (surview) before clicking add/replace button. | Subsystem | Exam Card | NA | Fixed | goswami_s | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP5.1_SegReq | Brilliance Big Bore |  | BB_WIP6 | April 5, 2019 6:38:30 PM GMT+08:00
CTUDT00621880 | Refresh | Closed | tiwari_g | February 18, 2019 7:48:32 PM GMT+08:00 | Refresh : SMPTE Image for Overall Image Quality is not matching with the reference image document in both DIN and Standard monitor. | 4.8 | BB_WIP5 |  | Subsystem | Monitor Calibration | None | Fixed | Prakash_Chand_Yadav | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SubCom | Brilliance Big Bore |  | BB_WIP5 | February 25, 2019 6:42:17 PM GMT+08:00
CTUDT00621879 | Refresh | Closed | nayak_r | February 18, 2019 7:43:05 PM GMT+08:00 | Irrelevant message displayed on clicking ok on the failure warning message of “Enable remote access” tool | BB_WIP4 |  | No Workaround | Subsystem | Remote Access | CT-NM.SRS.Service.144 | No Plan to Fix | ranjithkumar_m | 0 | 4 |  | Service Tools | NO | Closed |  | Ovation
CTUDT00621879 | Refresh | Closed | nayak_r | February 18, 2019 7:43:05 PM GMT+08:00 | Irrelevant message displayed on clicking ok on the failure warning message of “Enable remote access” tool | BB_WIP4 |  | No Workaround | Subsystem | Remote Access | CT-NM.SRS.Service.144 | No Plan to Fix | ranjithkumar_m | 0 | 4 |  | Service Tools | NO | Closed |  | Brilliance 64
CTUDT00621879 | Refresh | Closed | nayak_r | February 18, 2019 7:43:05 PM GMT+08:00 | Irrelevant message displayed on clicking ok on the failure warning message of “Enable remote access” tool | BB_WIP4 |  | No Workaround | Subsystem | Remote Access | CT-NM.SRS.Service.144 | No Plan to Fix | ranjithkumar_m | 0 | 4 |  | Service Tools | NO | Closed |  | Brilliance Big Bore
CTUDT00621879 | Refresh | Closed | nayak_r | February 18, 2019 7:43:05 PM GMT+08:00 | Irrelevant message displayed on clicking ok on the failure warning message of “Enable remote access” tool | BB_WIP4 |  | No Workaround | Subsystem | Remote Access | CT-NM.SRS.Service.144 | No Plan to Fix | ranjithkumar_m | 0 | 4 |  | Service Tools | NO | Closed |  | Ingenuity 64
CTUDT00621879 | Refresh | Closed | nayak_r | February 18, 2019 7:43:05 PM GMT+08:00 | Irrelevant message displayed on clicking ok on the failure warning message of “Enable remote access” tool | BB_WIP4 |  | No Workaround | Subsystem | Remote Access | CT-NM.SRS.Service.144 | No Plan to Fix | ranjithkumar_m | 0 | 4 |  | Service Tools | NO | Closed |  | IQon1
CTUDT00621876 | Refresh | Closed | kumar_v | February 18, 2019 7:17:33 PM GMT+08:00 | Incorrect values for REVOLUTION_TIME in DICOM files on Ingenuity CT 4.1.6. | 4.1.6.25032 | WIP19 |  | Subsystem | DICOM | NA | Fixed | dunitza_m | 0 | 4 | Console Software | Support-System | NO | Closed | WIP14SubCom | Brilliance 64 | NA | WIP19 | October 25, 2019 2:41:06 AM GMT+08:00
CTUDT00621876 | Refresh | Closed | kumar_v | February 18, 2019 7:17:33 PM GMT+08:00 | Incorrect values for REVOLUTION_TIME in DICOM files on Ingenuity CT 4.1.6. | 4.1.6.25032 | WIP19 |  | Subsystem | DICOM | NA | Fixed | dunitza_m | 0 | 4 | Console Software | Support-System | NO | Closed | WIP14SubCom | Ingenuity 64 | NA | WIP19 | October 25, 2019 2:41:06 AM GMT+08:00
CTUDT00621875 | Refresh | Closed | Sateesh_S | February 18, 2019 7:04:55 PM GMT+08:00 | BB Refresh:Priority Transfer option in patient scene and High priority result option in ECM/plan are missing | BB_WIP3 | DHF330449 Rev 06 Big Bore Host verification test protocol |  | Subsystem | Test Design | N/A | Fixed | vinoliya_j | 0 | 3 | Test Design | Console4 | NO | Closed | BB_WIP5_SegReq | Brilliance Big Bore |  | DHF330449 Rev 06 Big Bore Host verification test protocol | May 2, 2019 3:22:56 PM GMT+08:00
CTUDT00621870 | Refresh | Closed | nayak_r | February 18, 2019 5:35:51 PM GMT+08:00 | Refresh/BB Refresh: When logged in as the new created clinical user, launching the “Quality assurance” tool showing the o-level user options incorrectly. | BB_WIP4 |  |  | Subsystem | Quality assurance | n/a | Works As Specified / Not A Defect | frider_o | 0 |  |  | DMS | NO | Closed |  | Ovation
CTUDT00621870 | Refresh | Closed | nayak_r | February 18, 2019 5:35:51 PM GMT+08:00 | Refresh/BB Refresh: When logged in as the new created clinical user, launching the “Quality assurance” tool showing the o-level user options incorrectly. | BB_WIP4 |  |  | Subsystem | Quality assurance | n/a | Works As Specified / Not A Defect | frider_o | 0 |  |  | DMS | NO | Closed |  | Brilliance 64
CTUDT00621870 | Refresh | Closed | nayak_r | February 18, 2019 5:35:51 PM GMT+08:00 | Refresh/BB Refresh: When logged in as the new created clinical user, launching the “Quality assurance” tool showing the o-level user options incorrectly. | BB_WIP4 |  |  | Subsystem | Quality assurance | n/a | Works As Specified / Not A Defect | frider_o | 0 |  |  | DMS | NO | Closed |  | Brilliance Big Bore
CTUDT00621870 | Refresh | Closed | nayak_r | February 18, 2019 5:35:51 PM GMT+08:00 | Refresh/BB Refresh: When logged in as the new created clinical user, launching the “Quality assurance” tool showing the o-level user options incorrectly. | BB_WIP4 |  |  | Subsystem | Quality assurance | n/a | Works As Specified / Not A Defect | frider_o | 0 |  |  | DMS | NO | Closed |  | Ingenuity 64
CTUDT00621870 | Refresh | Closed | nayak_r | February 18, 2019 5:35:51 PM GMT+08:00 | Refresh/BB Refresh: When logged in as the new created clinical user, launching the “Quality assurance” tool showing the o-level user options incorrectly. | BB_WIP4 |  |  | Subsystem | Quality assurance | n/a | Works As Specified / Not A Defect | frider_o | 0 |  |  | DMS | NO | Closed |  | IQon1
CTUDT00621866 | Refresh | Closed | Noy_Seelig | February 18, 2019 1:23:00 AM GMT+08:00 | Refresh: No plots in Slice thickness or Spatial resolution cases in Acceptance and Constancy tests. | WIP13 |  |  | Subsystem | DMS Calibrations | N/A | Works As Specified / Not A Defect | yosfan_s | 0 |  |  | DMS | NO | Closed |  | Ingenuity 64
CTUDT00621864 | Refresh | Closed | Noy_Seelig | February 17, 2019 11:09:36 PM GMT+08:00 | Refresh: No Performance test when Blue phantom is chosen. | WIP13 |  |  | Subsystem | DMS Calibrations | n/a | Duplicate | yatziv_l | 0 | 2 |  | DMS | NO | Closed | WIP14SegReq | Ingenuity 64
CTUDT00621861 | Refresh | Closed | finestine_m | February 17, 2019 7:12:45 PM GMT+08:00 | Refresh: Beam and DMS Check: Missing plots in the results. | WIP13 | WIP15 |  | Subsystem | DMS Check | n/a | Fixed | yosefk | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP15SubCom | Ovation |  | WIP15 | March 24, 2019 10:18:13 PM GMT+08:00
CTUDT00621861 | Refresh | Closed | finestine_m | February 17, 2019 7:12:45 PM GMT+08:00 | Refresh: Beam and DMS Check: Missing plots in the results. | WIP13 | WIP15 |  | Subsystem | DMS Check | n/a | Fixed | yosefk | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP15SubCom | Brilliance Big Bore |  | WIP15 | March 24, 2019 10:18:13 PM GMT+08:00
CTUDT00621861 | Refresh | Closed | finestine_m | February 17, 2019 7:12:45 PM GMT+08:00 | Refresh: Beam and DMS Check: Missing plots in the results. | WIP13 | WIP15 |  | Subsystem | DMS Check | n/a | Fixed | yosefk | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP15SubCom | Ingenuity 64 |  | WIP15 | March 24, 2019 10:18:13 PM GMT+08:00
CTUDT00621861 | Refresh | Closed | finestine_m | February 17, 2019 7:12:45 PM GMT+08:00 | Refresh: Beam and DMS Check: Missing plots in the results. | WIP13 | WIP15 |  | Subsystem | DMS Check | n/a | Fixed | yosefk | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP15SubCom | IQon1 |  | WIP15 | March 24, 2019 10:18:13 PM GMT+08:00
CTUDT00621856 | Refresh | Closed | mN_m | February 15, 2019 3:41:08 PM GMT+08:00 | BBRefresh: Less images generated than planned, when the result is loaded to slab viewer in view and view2 and performed recon. | BB_WIP4 | BB_WIP5.1 | User can typed-in the values in thickness and increment field in UI. | Subsystem | Recon Images | None | Fixed | sagar_v | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP5_SegReq | Brilliance Big Bore |  | BB_WIP5.1 | March 26, 2019 8:02:17 PM GMT+08:00
CTUDT00621854 | Refresh | Closed | mN_m | February 15, 2019 3:36:40 PM GMT+08:00 | BBRefresh: Result phases are not displayed on the offline ECG viewer by default. | BB_WIP4 | BB_WIP5.1 |  | Subsystem | ECG Viewer | None | Fixed | Prakash_Chand_Yadav | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP5_SubCom | Brilliance Big Bore |  | BB_WIP5.1 | March 18, 2019 4:37:18 PM GMT+08:00
CTUDT00621852 | Refresh | Closed | nayak_r | February 15, 2019 2:53:46 PM GMT+08:00 | BB Refresh: Setting Force center X,Y value to zero is not applied on mpr result When 'CT Simulation Exam Card' is checked in ECM | BB_WIP3 | DHF316705 Rev08 | None | System | CT Simulation | Console-SRS.14357 | Fixed | roy_r | 0 | 3 | Requirements Documentation | Console4 | NO | Closed | BB_WIP5_SegReq | Brilliance Big Bore |  | DHF316705 Rev08 | March 15, 2019 5:04:01 PM GMT+08:00
CTUDT00621840 | Refresh | Closed | manju_kolkar | February 14, 2019 6:53:58 PM GMT+08:00 | Refresh - ‘System Error’ message appears when user tries to Cut Paste timed step | WIP13 |  | lo start exam again and place all planned steps in desired order | Subsystem | Plan Viewer; | NA | No Plan to Fix | sowmya_k | 0 | 4 |  | Console4 | NO | Closed |  | Ovation
CTUDT00621840 | Refresh | Closed | manju_kolkar | February 14, 2019 6:53:58 PM GMT+08:00 | Refresh - ‘System Error’ message appears when user tries to Cut Paste timed step | WIP13 |  | lo start exam again and place all planned steps in desired order | Subsystem | Plan Viewer; | NA | No Plan to Fix | sowmya_k | 0 | 4 |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00621840 | Refresh | Closed | manju_kolkar | February 14, 2019 6:53:58 PM GMT+08:00 | Refresh - ‘System Error’ message appears when user tries to Cut Paste timed step | WIP13 |  | lo start exam again and place all planned steps in desired order | Subsystem | Plan Viewer; | NA | No Plan to Fix | sowmya_k | 0 | 4 |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00621840 | Refresh | Closed | manju_kolkar | February 14, 2019 6:53:58 PM GMT+08:00 | Refresh - ‘System Error’ message appears when user tries to Cut Paste timed step | WIP13 |  | lo start exam again and place all planned steps in desired order | Subsystem | Plan Viewer; | NA | No Plan to Fix | sowmya_k | 0 | 4 |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00621840 | Refresh | Closed | manju_kolkar | February 14, 2019 6:53:58 PM GMT+08:00 | Refresh - ‘System Error’ message appears when user tries to Cut Paste timed step | WIP13 |  | lo start exam again and place all planned steps in desired order | Subsystem | Plan Viewer; | NA | No Plan to Fix | sowmya_k | 0 | 4 |  | Console4 | NO | Closed |  | IQon1
CTUDT00621839 | Refresh | Closed | kr_m | February 14, 2019 6:41:09 PM GMT+08:00 | Refresh: For Dual Surview, Auto filming not working as expected. | WIP13 | WIP15 | to send to film second surview manually | Subsystem | Auto film Console | None | Fixed | kr_m | 0 | 3 | Console Software | Console4 | NO | Closed | WIP15SegReq | Ovation |  | WIP15 | April 1, 2019 12:51:40 PM GMT+08:00
CTUDT00621839 | Refresh | Closed | kr_m | February 14, 2019 6:41:09 PM GMT+08:00 | Refresh: For Dual Surview, Auto filming not working as expected. | WIP13 | WIP15 | to send to film second surview manually | Subsystem | Auto film Console | None | Fixed | kr_m | 0 | 3 | Console Software | Console4 | NO | Closed | WIP15SegReq | Brilliance 64 |  | WIP15 | April 1, 2019 12:51:40 PM GMT+08:00
CTUDT00621839 | Refresh | Closed | kr_m | February 14, 2019 6:41:09 PM GMT+08:00 | Refresh: For Dual Surview, Auto filming not working as expected. | WIP13 | WIP15 | to send to film second surview manually | Subsystem | Auto film Console | None | Fixed | kr_m | 0 | 3 | Console Software | Console4 | NO | Closed | WIP15SegReq | Ingenuity 64 |  | WIP15 | April 1, 2019 12:51:40 PM GMT+08:00
CTUDT00621839 | Refresh | Closed | kr_m | February 14, 2019 6:41:09 PM GMT+08:00 | Refresh: For Dual Surview, Auto filming not working as expected. | WIP13 | WIP15 | to send to film second surview manually | Subsystem | Auto film Console | None | Fixed | kr_m | 0 | 3 | Console Software | Console4 | NO | Closed | WIP15SegReq | IQon1 |  | WIP15 | April 1, 2019 12:51:40 PM GMT+08:00
CTUDT00621837 | Refresh | Closed | Sharmila_Vasireddy | February 14, 2019 6:19:48 PM GMT+08:00 | Refresh: Cut/Paste, Delete, Accession Number are not available in context menu of edited results. | BB_WIP3 | BB_WIP7 |  | Subsystem | Edit results | Console-SRS.17571 ,Console-SRS.17572 | Fixed | kumar-san | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP7_SubCom | Brilliance Big Bore |  | BB_WIP7 | May 14, 2019 6:41:30 PM GMT+08:00
CTUDT00621828 | Refresh | Closed | Noy_Seelig | February 14, 2019 5:43:40 PM GMT+08:00 | Refresh: Two rows for mode 140kV, small, 0.5 sec Rotation Time which fails 140kV cases in DFS Position calibration | WIP13 | WIP16 | Execute "load tube default" before DFS Position calibration. | Subsystem | Gantry Calibrations | n/a | Fixed | zhou_x | 0 | 2 | Gantry Calibrations | GCX | NO | Closed | WIP17SubCom | Ingenuity 64 |  | WIP16 | May 31, 2019 1:07:59 PM GMT+08:00
CTUDT00621825 | Refresh | Closed | Sharmila_Vasireddy | February 14, 2019 5:40:17 PM GMT+08:00 | Refresh: Paste option should not be there for suggested EC group | BB_WIP3 | WIP15 | leave the page and go there again | Subsystem | Exam Cards | NA | Fixed | datta_i | 0 | 3 | Console Software | Console4 | NO | Closed | WIP15SubCom | Ovation |  | WIP15 | March 28, 2019 8:35:13 PM GMT+08:00
CTUDT00621825 | Refresh | Closed | Sharmila_Vasireddy | February 14, 2019 5:40:17 PM GMT+08:00 | Refresh: Paste option should not be there for suggested EC group | BB_WIP3 | WIP15 | leave the page and go there again | Subsystem | Exam Cards | NA | Fixed | datta_i | 0 | 3 | Console Software | Console4 | NO | Closed | WIP15SubCom | Brilliance 64 |  | WIP15 | March 28, 2019 8:35:13 PM GMT+08:00
CTUDT00621825 | Refresh | Closed | Sharmila_Vasireddy | February 14, 2019 5:40:17 PM GMT+08:00 | Refresh: Paste option should not be there for suggested EC group | BB_WIP3 | WIP15 | leave the page and go there again | Subsystem | Exam Cards | NA | Fixed | datta_i | 0 | 3 | Console Software | Console4 | NO | Closed | WIP15SubCom | Ingenuity 64 |  | WIP15 | March 28, 2019 8:35:13 PM GMT+08:00
CTUDT00621825 | Refresh | Closed | Sharmila_Vasireddy | February 14, 2019 5:40:17 PM GMT+08:00 | Refresh: Paste option should not be there for suggested EC group | BB_WIP3 | WIP15 | leave the page and go there again | Subsystem | Exam Cards | NA | Fixed | datta_i | 0 | 3 | Console Software | Console4 | NO | Closed | WIP15SubCom | IQon1 |  | WIP15 | March 28, 2019 8:35:13 PM GMT+08:00
CTUDT00621808 | Refresh | Closed | jotawar_d | February 14, 2019 3:09:43 PM GMT+08:00 | DSR Report has Exposure Time in ms but unit shown as s | _Current Production HW_ | BB_WIP5 |  | Subsystem |  | NA | Fixed | sagar_v | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP5_SubCom | Brilliance Big Bore | Mercury ID: 9446164. | BB_WIP5 | March 1, 2019 7:35:13 PM GMT+08:00
CTUDT00621807 | Refresh | Closed | Sateesh_S | February 14, 2019 3:02:57 PM GMT+08:00 | The word perfusion is missing in the pop-up message (message#0095) for English and all non-English Languages. | BB_WIP3 | BB_WIP8 |  | System | Exam Prep | None | Fixed | v_sumanth | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP5_SubCom | Brilliance Big Bore |  | BB_WIP8 | July 24, 2019 1:32:34 PM GMT+08:00
CTUDT00621805 | Refresh | Closed | wang_l | February 14, 2019 1:39:36 PM GMT+08:00 | Olevel user do spooler restart and reset failed | WIP13 | WIP15 |  | Subsystem | Spooler,o-level | N/ | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP15SubCom | Ovation |  | WIP15 | April 4, 2019 4:09:13 PM GMT+08:00
CTUDT00621805 | Refresh | Closed | wang_l | February 14, 2019 1:39:36 PM GMT+08:00 | Olevel user do spooler restart and reset failed | WIP13 | WIP15 |  | Subsystem | Spooler,o-level | N/ | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP15SubCom | Brilliance 64 |  | WIP15 | April 4, 2019 4:09:13 PM GMT+08:00
CTUDT00621805 | Refresh | Closed | wang_l | February 14, 2019 1:39:36 PM GMT+08:00 | Olevel user do spooler restart and reset failed | WIP13 | WIP15 |  | Subsystem | Spooler,o-level | N/ | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP15SubCom | Brilliance Big Bore |  | WIP15 | April 4, 2019 4:09:13 PM GMT+08:00
CTUDT00621805 | Refresh | Closed | wang_l | February 14, 2019 1:39:36 PM GMT+08:00 | Olevel user do spooler restart and reset failed | WIP13 | WIP15 |  | Subsystem | Spooler,o-level | N/ | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP15SubCom | Ingenuity 64 |  | WIP15 | April 4, 2019 4:09:13 PM GMT+08:00
CTUDT00621805 | Refresh | Closed | wang_l | February 14, 2019 1:39:36 PM GMT+08:00 | Olevel user do spooler restart and reset failed | WIP13 | WIP15 |  | Subsystem | Spooler,o-level | N/ | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP15SubCom | IQon1 |  | WIP15 | April 4, 2019 4:09:13 PM GMT+08:00
CTUDT00621803 | Refresh | Closed | wang_l | February 14, 2019 1:23:23 PM GMT+08:00 | System error occurred for workflow if doing spooler restart or reset when transfer is in progress | WIP13 | WIP13 |  | System | Spooler-USP | N/A | Fixed | wang_l | 0 | 3 | Console USP_EBW | Support-System | NO | Closed | WIP14SegReq | Ingenuity 64 |  | WIP13 | June 3, 2019 6:03:06 PM GMT+08:00
CTUDT00621796 | Refresh | Closed | levy_s | February 13, 2019 6:59:37 PM GMT+08:00 | Refresh: Test execution cannot be performed since the system is stuck at “initializing” mode | WIP13 |  |  | Subsystem | CCT |  | Duplicate | levy_s | 0 |  |  | Support-System | NO | Closed |  | Ovation
CTUDT00621795 | Refresh | Closed | manju_kolkar | February 13, 2019 6:43:29 PM GMT+08:00 | Refresh- User Name value in Dose Log Viewer is not recorded. | WIP13 | WIP15 |  | Subsystem | Dose Log Viewer | None | Fixed | chokkalingam_r | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SegReq | Ovation |  | WIP15 | March 14, 2019 5:40:20 PM GMT+08:00
CTUDT00621795 | Refresh | Closed | manju_kolkar | February 13, 2019 6:43:29 PM GMT+08:00 | Refresh- User Name value in Dose Log Viewer is not recorded. | WIP13 | WIP15 |  | Subsystem | Dose Log Viewer | None | Fixed | chokkalingam_r | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SegReq | Brilliance 64 |  | WIP15 | March 14, 2019 5:40:20 PM GMT+08:00
CTUDT00621795 | Refresh | Closed | manju_kolkar | February 13, 2019 6:43:29 PM GMT+08:00 | Refresh- User Name value in Dose Log Viewer is not recorded. | WIP13 | WIP15 |  | Subsystem | Dose Log Viewer | None | Fixed | chokkalingam_r | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SegReq | Ingenuity 64 |  | WIP15 | March 14, 2019 5:40:20 PM GMT+08:00
CTUDT00621795 | Refresh | Closed | manju_kolkar | February 13, 2019 6:43:29 PM GMT+08:00 | Refresh- User Name value in Dose Log Viewer is not recorded. | WIP13 | WIP15 |  | Subsystem | Dose Log Viewer | None | Fixed | chokkalingam_r | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SegReq | IQon1 |  | WIP15 | March 14, 2019 5:40:20 PM GMT+08:00
CTUDT00621792 | Refresh | Closed | g_di | February 13, 2019 5:30:33 PM GMT+08:00 | In Emergency login “Login for emergency” pop-up message is displayed when user click on end exam of the first scan for new user name. | WIP13 | WIP15 |  | Subsystem | Emergency login | NA | Fixed | veerabhadraiah | 0 | 4 | Console USP_EBW | iCAP | NO | Closed | WIP15SubCom | Ovation |  | WIP15 | April 4, 2019 4:03:25 PM GMT+08:00
CTUDT00621792 | Refresh | Closed | g_di | February 13, 2019 5:30:33 PM GMT+08:00 | In Emergency login “Login for emergency” pop-up message is displayed when user click on end exam of the first scan for new user name. | WIP13 | WIP15 |  | Subsystem | Emergency login | NA | Fixed | veerabhadraiah | 0 | 4 | Console USP_EBW | iCAP | NO | Closed | WIP15SubCom | Brilliance 64 |  | WIP15 | April 4, 2019 4:03:25 PM GMT+08:00
CTUDT00621792 | Refresh | Closed | g_di | February 13, 2019 5:30:33 PM GMT+08:00 | In Emergency login “Login for emergency” pop-up message is displayed when user click on end exam of the first scan for new user name. | WIP13 | WIP15 |  | Subsystem | Emergency login | NA | Fixed | veerabhadraiah | 0 | 4 | Console USP_EBW | iCAP | NO | Closed | WIP15SubCom | Brilliance Big Bore |  | WIP15 | April 4, 2019 4:03:25 PM GMT+08:00
CTUDT00621792 | Refresh | Closed | g_di | February 13, 2019 5:30:33 PM GMT+08:00 | In Emergency login “Login for emergency” pop-up message is displayed when user click on end exam of the first scan for new user name. | WIP13 | WIP15 |  | Subsystem | Emergency login | NA | Fixed | veerabhadraiah | 0 | 4 | Console USP_EBW | iCAP | NO | Closed | WIP15SubCom | Ingenuity 64 |  | WIP15 | April 4, 2019 4:03:25 PM GMT+08:00
CTUDT00621792 | Refresh | Closed | g_di | February 13, 2019 5:30:33 PM GMT+08:00 | In Emergency login “Login for emergency” pop-up message is displayed when user click on end exam of the first scan for new user name. | WIP13 | WIP15 |  | Subsystem | Emergency login | NA | Fixed | veerabhadraiah | 0 | 4 | Console USP_EBW | iCAP | NO | Closed | WIP15SubCom | IQon1 |  | WIP15 | April 4, 2019 4:03:25 PM GMT+08:00
CTUDT00621791 | Refresh | Closed | g_di | February 13, 2019 5:29:48 PM GMT+08:00 | Emergency user login and logout are not displayed in “Audit trial” under “Security Alert” | WIP13 | BB_WIP9 WIP19 | No Workaround | Subsystem | Audit trail | N/A | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP19SubCom | Ovation |  | BB_WIP9 WIP19 | July 30, 2019 6:57:53 PM GMT+08:00
CTUDT00621791 | Refresh | Closed | g_di | February 13, 2019 5:29:48 PM GMT+08:00 | Emergency user login and logout are not displayed in “Audit trial” under “Security Alert” | WIP13 | BB_WIP9 WIP19 | No Workaround | Subsystem | Audit trail | N/A | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP19SubCom | Brilliance 64 |  | BB_WIP9 WIP19 | July 30, 2019 6:57:53 PM GMT+08:00
CTUDT00621791 | Refresh | Closed | g_di | February 13, 2019 5:29:48 PM GMT+08:00 | Emergency user login and logout are not displayed in “Audit trial” under “Security Alert” | WIP13 | BB_WIP9 WIP19 | No Workaround | Subsystem | Audit trail | N/A | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP19SubCom | Brilliance Big Bore |  | BB_WIP9 WIP19 | July 30, 2019 6:57:53 PM GMT+08:00
CTUDT00621791 | Refresh | Closed | g_di | February 13, 2019 5:29:48 PM GMT+08:00 | Emergency user login and logout are not displayed in “Audit trial” under “Security Alert” | WIP13 | BB_WIP9 WIP19 | No Workaround | Subsystem | Audit trail | N/A | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP19SubCom | Ingenuity 64 |  | BB_WIP9 WIP19 | July 30, 2019 6:57:53 PM GMT+08:00
CTUDT00621791 | Refresh | Closed | g_di | February 13, 2019 5:29:48 PM GMT+08:00 | Emergency user login and logout are not displayed in “Audit trial” under “Security Alert” | WIP13 | BB_WIP9 WIP19 | No Workaround | Subsystem | Audit trail | N/A | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP19SubCom | IQon1 |  | BB_WIP9 WIP19 | July 30, 2019 6:57:53 PM GMT+08:00
CTUDT00621789 | Refresh | Closed | naik_p | February 13, 2019 5:00:01 PM GMT+08:00 | Refresh_wip 13_Brain Perfusion scan type available for body perfusion Helical ExamCard | WIP13 | WIP15 |  | System | Scan types | None | Fixed | chokkalingam_r | 0 | 2 | Console Software | Console4 | YES | Closed | WIP14SegReq | Ovation |  | WIP15 | March 20, 2019 4:58:31 PM GMT+08:00
CTUDT00621789 | Refresh | Closed | naik_p | February 13, 2019 5:00:01 PM GMT+08:00 | Refresh_wip 13_Brain Perfusion scan type available for body perfusion Helical ExamCard | WIP13 | WIP15 |  | System | Scan types | None | Fixed | chokkalingam_r | 0 | 2 | Console Software | Console4 | YES | Closed | WIP14SegReq | Brilliance 64 |  | WIP15 | March 20, 2019 4:58:31 PM GMT+08:00
CTUDT00621789 | Refresh | Closed | naik_p | February 13, 2019 5:00:01 PM GMT+08:00 | Refresh_wip 13_Brain Perfusion scan type available for body perfusion Helical ExamCard | WIP13 | WIP15 |  | System | Scan types | None | Fixed | chokkalingam_r | 0 | 2 | Console Software | Console4 | YES | Closed | WIP14SegReq | Ingenuity 64 |  | WIP15 | March 20, 2019 4:58:31 PM GMT+08:00
CTUDT00621787 | Refresh | Closed | panchami_pradeep | February 13, 2019 2:43:04 PM GMT+08:00 | Refresh: PSC: In CT mode, Error is displayed while clicking on "Switch to Remote View". | WIP13 | WIP14 | NA | Subsystem | PSC | NA | Fixed | kapadia_v | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP14SubCom | Ovation |  | WIP14 | March 5, 2019 2:41:25 PM GMT+08:00
CTUDT00621787 | Refresh | Closed | panchami_pradeep | February 13, 2019 2:43:04 PM GMT+08:00 | Refresh: PSC: In CT mode, Error is displayed while clicking on "Switch to Remote View". | WIP13 | WIP14 | NA | Subsystem | PSC | NA | Fixed | kapadia_v | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP14SubCom | Brilliance 64 |  | WIP14 | March 5, 2019 2:41:25 PM GMT+08:00
CTUDT00621787 | Refresh | Closed | panchami_pradeep | February 13, 2019 2:43:04 PM GMT+08:00 | Refresh: PSC: In CT mode, Error is displayed while clicking on "Switch to Remote View". | WIP13 | WIP14 | NA | Subsystem | PSC | NA | Fixed | kapadia_v | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP14SubCom | Ingenuity 64 |  | WIP14 | March 5, 2019 2:41:25 PM GMT+08:00
CTUDT00621787 | Refresh | Closed | panchami_pradeep | February 13, 2019 2:43:04 PM GMT+08:00 | Refresh: PSC: In CT mode, Error is displayed while clicking on "Switch to Remote View". | WIP13 | WIP14 | NA | Subsystem | PSC | NA | Fixed | kapadia_v | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP14SubCom | IQon1 |  | WIP14 | March 5, 2019 2:41:25 PM GMT+08:00
CTUDT00621786 | Refresh | Closed | ravindra_a | February 13, 2019 1:27:33 PM GMT+08:00 | Refresh, IQon - Cannot restore UserAcounts | 4.8 | BB_WIP5 | Yes, user can create required accounts as and when he wants using IT tool. | Subsystem | Backup/Restore | N/A | Fixed | ravindra_a | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP5_SegReq | Brilliance Big Bore |  | BB_WIP5 | March 14, 2019 4:43:12 PM GMT+08:00
CTUDT00621778 | Refresh | Closed | alex_nihaev | February 12, 2019 9:38:33 PM GMT+08:00 | Bigbore Refresh: Physics Analysis: Crash after clicking on Matrix view checkbox | BB_WIP3 |  | M/A | Subsystem | Physics analysis | No such requirement | No Plan to Fix | yatziv_l | 0 | 4 |  | DMS | NO | Closed | BB_SubNC | Brilliance Big Bore
CTUDT00621776 | Refresh | Closed | hiremath_b | February 12, 2019 9:01:55 PM GMT+08:00 | [Refresh/IQON]: Bugrep error pops up on clicking GO for a clinical scan when 1st result is deleted from non-evolving scan. | WIP13 | WIP15 | Instead of clicking go -> cancel -> deleting first result, use previous surview and delete first result before clicking go, application will not crash. | Subsystem | Exam Application | NA | Fixed | datta_i | 0 | 3 | Console Software | Console4 | YES | Closed | WIP15SegReq | IQon1 |  | WIP15 | March 21, 2019 4:59:37 PM GMT+08:00
CTUDT00621773 | Refresh | Closed | g_di | February 12, 2019 7:28:09 PM GMT+08:00 | IRS_STAGE_WARP should not exist in the MDU files for axial 2D EC result and should exist for QA axial body 3d | WIP12 | WIP14 | No workaround | Subsystem | MDU | NA | Fixed | sagar_v | 0 | 2 | Console Software | Console4 | YES | Closed | BB_WIP5_SegReq | Ingenuity 64 |  | WIP14 | February 26, 2019 4:33:52 PM GMT+08:00
CTUDT00621756 | Refresh | Closed | alex_nihaev | February 11, 2019 10:32:16 PM GMT+08:00 | Bigbore Refresh: Calibrations: HCor cannot be performed with Harmonized Phantom | BB_WIP3 | WIP14 WIP15 |  | Subsystem | Calibrations | n/a | Fixed | frider_o | 0 | 2 | DMS Software | DMS | YES | Closed | BB_WIP5_SegReq | Brilliance Big Bore |  | WIP14 WIP15 | March 17, 2019 10:16:00 PM GMT+08:00
CTUDT00621743 | Refresh | Closed | hiremath_b | February 11, 2019 6:16:10 PM GMT+08:00 | BB refresh: Recon Error is displaying along with CIRS initialization message in patient page and in plan. | BB_WIP3 | BB_WIP5 | NA | Subsystem |  | NA | Fixed | chokkalingam_r | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP4_SegReq | Brilliance Big Bore |  | BB_WIP5 | March 1, 2019 1:27:45 PM GMT+08:00
CTUDT00621741 | Refresh | Closed | rokni_m | February 11, 2019 5:46:00 PM GMT+08:00 | REFRESH IQon: APlane opening diagnostic test crashes with an error | WIP13 | WIP15 | FSE can manually measure openings with gauge pins. | Subsystem | Gantry Calibrations | CT-NM.SRS.Service.2593 | Fixed | morton_t | 0 | 2 | Gantry Software | Support-Physics | NO | Closed | WIP15SubCom | IQon1 |  | WIP15 | April 9, 2019 8:41:42 PM GMT+08:00
CTUDT00621737 | Refresh | Closed | jotawar_d | February 11, 2019 3:09:52 PM GMT+08:00 | Under Preference in Connectivity tab "Patient's size " option gets unselected even after selecting and saving the changes. | BB_WIP3 | BB_WIP5 | the user can manually update patient size in the console | System | Preferences | NA | Fixed | jose_s | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP5_SegReq | Brilliance Big Bore | PR#8563081 | BB_WIP5 | March 5, 2019 7:32:11 PM GMT+08:00
CTUDT00621730 | Refresh | Closed | Valeria_Kashtanov | February 11, 2019 3:15:27 AM GMT+08:00 | Refresh: Calibrations: Performance Test: Scans limits are not as according to spec | WIP12 |  | N/A | Subsystem | Performance Test | N/A | Works As Specified / Not A Defect | yosfan_s | 1 | 2 |  | DMS | NO | Closed | WIP15SubCom | Ingenuity 64
CTUDT00621725 | Refresh | Closed | Noy_Seelig | February 10, 2019 7:56:13 PM GMT+08:00 | Refresh: Performance Test: Windowing not according to spec. | WIP12 | WIP15 | N/A | Subsystem | Performance test; Verified1 | N/A | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP14SegReq | Ovation |  | WIP15 | March 17, 2019 10:27:22 PM GMT+08:00
CTUDT00621725 | Refresh | Closed | Noy_Seelig | February 10, 2019 7:56:13 PM GMT+08:00 | Refresh: Performance Test: Windowing not according to spec. | WIP12 | WIP15 | N/A | Subsystem | Performance test; Verified1 | N/A | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP14SegReq | Brilliance Big Bore |  | WIP15 | March 17, 2019 10:27:22 PM GMT+08:00
CTUDT00621725 | Refresh | Closed | Noy_Seelig | February 10, 2019 7:56:13 PM GMT+08:00 | Refresh: Performance Test: Windowing not according to spec. | WIP12 | WIP15 | N/A | Subsystem | Performance test; Verified1 | N/A | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP14SegReq | Ingenuity 64 |  | WIP15 | March 17, 2019 10:27:22 PM GMT+08:00
CTUDT00621725 | Refresh | Closed | Noy_Seelig | February 10, 2019 7:56:13 PM GMT+08:00 | Refresh: Performance Test: Windowing not according to spec. | WIP12 | WIP15 | N/A | Subsystem | Performance test; Verified1 | N/A | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP14SegReq | IQon1 |  | WIP15 | March 17, 2019 10:27:22 PM GMT+08:00
CTUDT00621706 | Refresh | Closed | ravindra_a | February 8, 2019 10:26:22 PM GMT+08:00 | EC is getting duplicated after renaming with existing EC name | 4.8 | DHF316705 Rev08 | No work around | Subsystem | Exam Cards | Console-SRS.16027 | Fixed | roy_r | 0 | 3 | Requirements Documentation | Console4 | NO | Closed | BB_WIP5_SegReq | Brilliance Big Bore |  | DHF316705 Rev08 | April 1, 2019 7:49:30 PM GMT+08:00
CTUDT00621705 | Refresh | Closed | ravindra_a | February 8, 2019 8:35:40 PM GMT+08:00 | CIRS should not send unrelated images during Bolus tracker scan | 4.8 |  | No. User should avoid performing offline recons during tracker scans. | Subsystem | Performance improvement | N/A | Does Not Apply | patra_r | 0 | 3 |  | Console4 | YES | Closed | BB_WIP5.1_SegReq | Brilliance Big Bore | ESC#0110990827 PR 8567370 PR 8539682 Case # 0111665405
CTUDT00621699 | Refresh | Closed | ravindra_a | February 8, 2019 8:13:11 PM GMT+08:00 | Refresh: In olevel user, while log out application launched from PSC tool (By clicking on "Start Scanner Application") system is logging out from o level user itself. | 4.8 | BB_WIP5.1 | Olevel User can re-login and continue work on PSC tool. | Subsystem | System Log off | N/A | Fixed | kapadia_v | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP5_SegReq | Brilliance Big Bore |  | BB_WIP5.1 | March 27, 2019 12:58:28 PM GMT+08:00
CTUDT00621698 | Refresh | Closed | jotawar_d | February 8, 2019 4:11:31 PM GMT+08:00 | Refresh:For CCT Fluoro scan,Unit of Planned DLP in show all is mGy*cm/s where as during scan it is mGy*cm in acquisition scene. it should be consistent as per the requirement Console-SRS.15974 | BB_WIP3 | BB_WIP5 |  | Subsystem | CCT | NA | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP5_SegReq | Brilliance Big Bore |  | BB_WIP5 | March 14, 2019 7:05:11 PM GMT+08:00
CTUDT00621690 | Refresh | Closed | manju_kolkar | February 8, 2019 2:16:37 PM GMT+08:00 | Refresh - Study doesn’t appear in History Database list when user performs ‘End Exam and Continue Working’ | WIP12 | Host SSRS DHF303211 Rev 08 |  | Subsystem | History Database | N/A | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | WIP15SegReq | Ovation |  | Host SSRS DHF303211 Rev 08 | March 22, 2019 6:09:50 PM GMT+08:00
CTUDT00621690 | Refresh | Closed | manju_kolkar | February 8, 2019 2:16:37 PM GMT+08:00 | Refresh - Study doesn’t appear in History Database list when user performs ‘End Exam and Continue Working’ | WIP12 | Host SSRS DHF303211 Rev 08 |  | Subsystem | History Database | N/A | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | WIP15SegReq | Brilliance 64 |  | Host SSRS DHF303211 Rev 08 | March 22, 2019 6:09:50 PM GMT+08:00
CTUDT00621690 | Refresh | Closed | manju_kolkar | February 8, 2019 2:16:37 PM GMT+08:00 | Refresh - Study doesn’t appear in History Database list when user performs ‘End Exam and Continue Working’ | WIP12 | Host SSRS DHF303211 Rev 08 |  | Subsystem | History Database | N/A | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | WIP15SegReq | Ingenuity 64 |  | Host SSRS DHF303211 Rev 08 | March 22, 2019 6:09:50 PM GMT+08:00
CTUDT00621690 | Refresh | Closed | manju_kolkar | February 8, 2019 2:16:37 PM GMT+08:00 | Refresh - Study doesn’t appear in History Database list when user performs ‘End Exam and Continue Working’ | WIP12 | Host SSRS DHF303211 Rev 08 |  | Subsystem | History Database | N/A | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | WIP15SegReq | IQon1 |  | Host SSRS DHF303211 Rev 08 | March 22, 2019 6:09:50 PM GMT+08:00
CTUDT00621688 | Refresh | Closed | nayak_r | February 7, 2019 6:47:15 PM GMT+08:00 | Refresh/BB Refresh: In Lan config,editing of the Port and AE title values for each of the QUERY,STORE and MOVE services after configuring a CT Workstation is not getting allowed | BB_WIP3 |  |  | Subsystem | LAN Config |  | Works As Specified / Not A Defect | nayak_r | 0 |  |  |  |  | Closed |  | Ovation
CTUDT00621688 | Refresh | Closed | nayak_r | February 7, 2019 6:47:15 PM GMT+08:00 | Refresh/BB Refresh: In Lan config,editing of the Port and AE title values for each of the QUERY,STORE and MOVE services after configuring a CT Workstation is not getting allowed | BB_WIP3 |  |  | Subsystem | LAN Config |  | Works As Specified / Not A Defect | nayak_r | 0 |  |  |  |  | Closed |  | Brilliance 64
CTUDT00621688 | Refresh | Closed | nayak_r | February 7, 2019 6:47:15 PM GMT+08:00 | Refresh/BB Refresh: In Lan config,editing of the Port and AE title values for each of the QUERY,STORE and MOVE services after configuring a CT Workstation is not getting allowed | BB_WIP3 |  |  | Subsystem | LAN Config |  | Works As Specified / Not A Defect | nayak_r | 0 |  |  |  |  | Closed |  | Brilliance Big Bore
CTUDT00621688 | Refresh | Closed | nayak_r | February 7, 2019 6:47:15 PM GMT+08:00 | Refresh/BB Refresh: In Lan config,editing of the Port and AE title values for each of the QUERY,STORE and MOVE services after configuring a CT Workstation is not getting allowed | BB_WIP3 |  |  | Subsystem | LAN Config |  | Works As Specified / Not A Defect | nayak_r | 0 |  |  |  |  | Closed |  | Ingenuity 64
CTUDT00621688 | Refresh | Closed | nayak_r | February 7, 2019 6:47:15 PM GMT+08:00 | Refresh/BB Refresh: In Lan config,editing of the Port and AE title values for each of the QUERY,STORE and MOVE services after configuring a CT Workstation is not getting allowed | BB_WIP3 |  |  | Subsystem | LAN Config |  | Works As Specified / Not A Defect | nayak_r | 0 |  |  |  |  | Closed |  | IQon1
CTUDT00621676 | Refresh | Closed | nanjunda_s | February 7, 2019 1:46:17 PM GMT+08:00 | Refresh Sanity: Not able to perform DUAL energy scan ,getting an error message "Sytem error.Please end the exam and retry". | WIP13 |  |  | Subsystem | Exam App | n/a | Does Not Apply | altshuller_mark | 1 |  |  | Support-System | NO | Closed |  | Ovation
CTUDT00621674 | Refresh | Closed | levy_s | February 6, 2019 10:56:13 PM GMT+08:00 | Refresh: "A problem occurred" error appeared during a clinical scan and surview execution | WIP12 |  |  | Subsystem | DMS | NA | System Repaired | avril | 0 |  |  | Support-System | YES | Closed |  | Ovation
CTUDT00621674 | Refresh | Closed | levy_s | February 6, 2019 10:56:13 PM GMT+08:00 | Refresh: "A problem occurred" error appeared during a clinical scan and surview execution | WIP12 |  |  | Subsystem | DMS | NA | System Repaired | avril | 0 |  |  | Support-System | YES | Closed |  | IQon1
CTUDT00621666 | Refresh | Closed | roshan_r | February 6, 2019 8:09:02 PM GMT+08:00 | BB Refresh : Console takes more than 8 seconds to end an exam | BB_WIP3 | BB_WIP10 | No workaround. User has to wait till patient page comes up. | Subsystem | Performance | NA | Fixed | ravindra_a | 0 | 3 | Console Software | Console4 | NO | Closed | WIP16SegReq | Brilliance Big Bore |  | BB_WIP10 | September 16, 2019 1:18:58 PM GMT+08:00
CTUDT00621661 | Refresh | Closed | roshan_r | February 6, 2019 3:10:55 PM GMT+08:00 | Refresh : The Console takes more than 5 seconds to get the CIRS and Gantry ready to scan when go button is clicked for surview | BB_WIP3 | BB_WIP7 | NA | Subsystem | Performance | NA | Fixed | ravindra_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP15SegReq | Brilliance Big Bore |  | BB_WIP7 | May 10, 2019 8:42:49 PM GMT+08:00
CTUDT00621659 | Refresh | Closed | g_di | February 6, 2019 2:48:27 PM GMT+08:00 | Refresh: Emergency user logout log in the audit trial is not displayed as expected when logged in by giving Name, Password and Emergency user | WIP12 | WIP15 | No workaround. However user can check which emergency user last logged in and can make out which emergency user logged out. | Subsystem | Audit Trail | NA | Fixed | kr_m | 0 | 4 | Console Software | Console4 | NO | Closed | WIP14SegReq | Ovation |  | WIP15 | March 26, 2019 6:32:25 PM GMT+08:00
CTUDT00621659 | Refresh | Closed | g_di | February 6, 2019 2:48:27 PM GMT+08:00 | Refresh: Emergency user logout log in the audit trial is not displayed as expected when logged in by giving Name, Password and Emergency user | WIP12 | WIP15 | No workaround. However user can check which emergency user last logged in and can make out which emergency user logged out. | Subsystem | Audit Trail | NA | Fixed | kr_m | 0 | 4 | Console Software | Console4 | NO | Closed | WIP14SegReq | Brilliance 64 |  | WIP15 | March 26, 2019 6:32:25 PM GMT+08:00
CTUDT00621659 | Refresh | Closed | g_di | February 6, 2019 2:48:27 PM GMT+08:00 | Refresh: Emergency user logout log in the audit trial is not displayed as expected when logged in by giving Name, Password and Emergency user | WIP12 | WIP15 | No workaround. However user can check which emergency user last logged in and can make out which emergency user logged out. | Subsystem | Audit Trail | NA | Fixed | kr_m | 0 | 4 | Console Software | Console4 | NO | Closed | WIP14SegReq | Brilliance Big Bore |  | WIP15 | March 26, 2019 6:32:25 PM GMT+08:00
CTUDT00621659 | Refresh | Closed | g_di | February 6, 2019 2:48:27 PM GMT+08:00 | Refresh: Emergency user logout log in the audit trial is not displayed as expected when logged in by giving Name, Password and Emergency user | WIP12 | WIP15 | No workaround. However user can check which emergency user last logged in and can make out which emergency user logged out. | Subsystem | Audit Trail | NA | Fixed | kr_m | 0 | 4 | Console Software | Console4 | NO | Closed | WIP14SegReq | Ingenuity 64 |  | WIP15 | March 26, 2019 6:32:25 PM GMT+08:00
CTUDT00621659 | Refresh | Closed | g_di | February 6, 2019 2:48:27 PM GMT+08:00 | Refresh: Emergency user logout log in the audit trial is not displayed as expected when logged in by giving Name, Password and Emergency user | WIP12 | WIP15 | No workaround. However user can check which emergency user last logged in and can make out which emergency user logged out. | Subsystem | Audit Trail | NA | Fixed | kr_m | 0 | 4 | Console Software | Console4 | NO | Closed | WIP14SegReq | IQon1 |  | WIP15 | March 26, 2019 6:32:25 PM GMT+08:00
CTUDT00621652 | Refresh | Closed | nimma_p | February 6, 2019 1:54:51 PM GMT+08:00 | Refresh:-IQON:-Performance:-End exam Actual Latency was more than Expected Latency. | WIP12 |  |  | Subsystem | Performance |  | Duplicate | ravindra_a | 0 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00621646 | Refresh | Closed | roshan_r | February 5, 2019 11:13:52 PM GMT+08:00 | BB Refresh : Importing EC fails | BB_WIP3 |  | User can load the EC using load external exam card option | Subsystem | ECM | NA | Duplicate | ravindra_a | 0 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00621645 | Refresh | Closed | jotawar_d | February 5, 2019 9:58:43 PM GMT+08:00 | Autosend failure due to ExamManager exit before completion of saving of results. | 4.1.6 | BB_WIP5.1 | Workaround is to use the CopyTo function to send the results to the external device at a later time | Subsystem | ExamManager | NA | Fixed | Prakash_Chand_Yadav | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP5_SegReq | Brilliance Big Bore | PR 5771594 PR 5786361 PR 5824426 PR 5824428 PR 5824429 PR 5991017 PR 5991018 PR 5785859 | BB_WIP5.1 | March 19, 2019 2:15:55 PM GMT+08:00
CTUDT00621643 | Refresh | Closed | Valeria_Kashtanov | February 5, 2019 9:43:20 PM GMT+08:00 | Bigbore Refresh: Calibrations: there is no Performance Application in Quality Assurance when Harmonized Phantom is checked | BB_WIP3 | WIP14 |  | Subsystem | DMS Calibrations | NaN | Fixed | daibes_s | 0 | 2 | DMS Software | DMS | NO | Closed | BB_WIP5_SegReq | Brilliance Big Bore |  | WIP14 | February 17, 2019 4:14:43 PM GMT+08:00
CTUDT00621640 | Refresh | Closed | shekhar_p | February 5, 2019 9:28:20 PM GMT+08:00 | Refresh Sanity: Bugrep: System error | WIP13 |  |  | Subsystem | Bugrep | NA | Works As Specified / Not A Defect | pandey_a | 0 |  |  | Service Tools | NO | Closed |  | Ovation
CTUDT00621640 | Refresh | Closed | shekhar_p | February 5, 2019 9:28:20 PM GMT+08:00 | Refresh Sanity: Bugrep: System error | WIP13 |  |  | Subsystem | Bugrep | NA | Works As Specified / Not A Defect | pandey_a | 0 |  |  | Service Tools | NO | Closed |  | Ingenuity 64
CTUDT00621635 | Refresh | Closed | Sharmila_Vasireddy | February 5, 2019 7:45:02 PM GMT+08:00 | "Real Time Image for SOP ID xxxx added module: Recon Device" log was not appeared in the logs | BB_WIP3 | DHF330449 Rev06 Big Bore HOST Verification Test Protocol | No workaround | Subsystem | Logging | NA | Fixed | vinoliya_j | 0 | 4 | Test Design | Console4 | NO | Closed | BB_WIP5_SegReq | Brilliance Big Bore |  | DHF330449 Rev06 Big Bore HOST Verification Test Protocol | May 2, 2019 5:39:45 PM GMT+08:00
CTUDT00621625 | Refresh | Closed | Sharmila_Vasireddy | February 5, 2019 6:40:27 PM GMT+08:00 | BB Refresh : For P3T, automatically clinical phases (A, B) are not getting loading in injection parameter box | BB_WIP3 | BB_WIP5 | User can add phases manually. | Subsystem | Exam Prep | NA | Fixed | datta_i | 0 | 3 | Console Software | Console4 | NO | Closed | WIP15SubCom | Brilliance Big Bore |  | BB_WIP5 | March 5, 2019 7:41:01 PM GMT+08:00
CTUDT00621619 | Refresh | Closed | sowmya_k | February 5, 2019 12:26:22 PM GMT+08:00 | Refresh: The following HOW DO I movies-> plan Geometry limitations, Dose alert, Dose notification, Volume display, Offline Reconstruction are not displayed | BB_WIP3 | BB_WIP5.1 |  | Subsystem | HOW DO I | NA | Fixed | vishnu_devi | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP5_SegReq | Brilliance Big Bore |  | BB_WIP5.1 | March 28, 2019 5:59:06 PM GMT+08:00
CTUDT00621616 | Refresh | Closed | Valeria_Kashtanov | February 4, 2019 8:39:26 PM GMT+08:00 | Refresh: Calibrations: CTDI: In manual mode test starts even when some of the Dose-meter Details window fields are blank | BB_WIP3 |  |  | Subsystem | CTDI |  | Duplicate | frider_o | 0 |  |  | DMS | NO | Closed |  | Brilliance Big Bore
CTUDT00621608 | Refresh | Closed | cohen_y | February 4, 2019 4:38:03 PM GMT+08:00 | Refresh: Calibrations: Hcor Calibration: HcorSlice_ScanSortSettings.cfg maps to the wrong file in IQon. | WIP12 | WIP14 | N/A | Subsystem | Hcor Calibration | N/A | Fixed | elviral | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP14SegReq | IQon1 |  | WIP14 | February 21, 2019 12:28:34 AM GMT+08:00
CTUDT00621607 | Refresh | Closed | manjunath_mk | February 4, 2019 4:19:08 PM GMT+08:00 | Schedule button in Worklist gets displayed even though the 'Display Scheduled Button' checkbox is unchecked | BB_WIP3 | BB_WIP5 |  | Subsystem | Worklist | NA | Fixed | chokkalingam_r | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP5_SegReq | Brilliance Big Bore |  | BB_WIP5 | March 4, 2019 2:32:36 PM GMT+08:00
CTUDT00621594 | Refresh | Closed | Sateesh_S | February 4, 2019 2:09:53 PM GMT+08:00 | BB Refresh :While importing an EC when the same Exam Card name already exists then Rename and Cancel options are not displaying and Overwrite option is not working | BB_WIP3 | BB_WIP5 DHF316705 Rev.08 | Scenario1: User can load the EC using load external exam card option Scenario1: No workaround. | Subsystem | ECM | Console.SRS.16027 | Fixed | roy_r | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP5_SegReq | Brilliance Big Bore |  | BB_WIP5 DHF316705 Rev.08 | March 14, 2019 8:26:15 PM GMT+08:00
CTUDT00621591 | Refresh | Closed | levin_c | February 3, 2019 11:57:25 PM GMT+08:00 | Refresh - Coast Down on iCT and iCTSP is not exist in service tools | WIP12 | WIP18 & WIP8 |  | System | Service Tools | CT-NM.SRS.Service.2992 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP14SegReq | Ovation |  | WIP18 & WIP8 | July 29, 2019 10:43:49 AM GMT+08:00
CTUDT00621591 | Refresh | Closed | levin_c | February 3, 2019 11:57:25 PM GMT+08:00 | Refresh - Coast Down on iCT and iCTSP is not exist in service tools | WIP12 | WIP18 & WIP8 |  | System | Service Tools | CT-NM.SRS.Service.2992 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP14SegReq | Brilliance 64 |  | WIP18 & WIP8 | July 29, 2019 10:43:49 AM GMT+08:00
CTUDT00621591 | Refresh | Closed | levin_c | February 3, 2019 11:57:25 PM GMT+08:00 | Refresh - Coast Down on iCT and iCTSP is not exist in service tools | WIP12 | WIP18 & WIP8 |  | System | Service Tools | CT-NM.SRS.Service.2992 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP14SegReq | Brilliance Big Bore |  | WIP18 & WIP8 | July 29, 2019 10:43:49 AM GMT+08:00
CTUDT00621591 | Refresh | Closed | levin_c | February 3, 2019 11:57:25 PM GMT+08:00 | Refresh - Coast Down on iCT and iCTSP is not exist in service tools | WIP12 | WIP18 & WIP8 |  | System | Service Tools | CT-NM.SRS.Service.2992 | Fixed | lin_s | 0 | 2 | Requirements Documentation | Support-System | NO | Closed | WIP14SegReq | Ingenuity 64 |  | WIP18 & WIP8 | July 29, 2019 10:43:49 AM GMT+08:00
CTUDT00621589 | Refresh | Closed | oreg_o | February 3, 2019 7:55:15 PM GMT+08:00 | Refresh: Calibrations: CTDI: Tube Warm-up was not performed in CTDI test | WIP12 | WIP14 | the service user can run tube warm up separately or calibrate the system in case the value can be out of limit. | Subsystem | DMS Calibrations | n/a | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP14SegReq | Ovation |  | WIP14 | February 14, 2019 3:36:13 PM GMT+08:00
CTUDT00621589 | Refresh | Closed | oreg_o | February 3, 2019 7:55:15 PM GMT+08:00 | Refresh: Calibrations: CTDI: Tube Warm-up was not performed in CTDI test | WIP12 | WIP14 | the service user can run tube warm up separately or calibrate the system in case the value can be out of limit. | Subsystem | DMS Calibrations | n/a | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP14SegReq | Brilliance Big Bore |  | WIP14 | February 14, 2019 3:36:13 PM GMT+08:00
CTUDT00621589 | Refresh | Closed | oreg_o | February 3, 2019 7:55:15 PM GMT+08:00 | Refresh: Calibrations: CTDI: Tube Warm-up was not performed in CTDI test | WIP12 | WIP14 | the service user can run tube warm up separately or calibrate the system in case the value can be out of limit. | Subsystem | DMS Calibrations | n/a | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP14SegReq | Ingenuity 64 |  | WIP14 | February 14, 2019 3:36:13 PM GMT+08:00
CTUDT00621589 | Refresh | Closed | oreg_o | February 3, 2019 7:55:15 PM GMT+08:00 | Refresh: Calibrations: CTDI: Tube Warm-up was not performed in CTDI test | WIP12 | WIP14 | the service user can run tube warm up separately or calibrate the system in case the value can be out of limit. | Subsystem | DMS Calibrations | n/a | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP14SegReq | IQon1 |  | WIP14 | February 14, 2019 3:36:13 PM GMT+08:00
CTUDT00621586 | Refresh | Closed | zhang_t | February 3, 2019 10:17:06 AM GMT+08:00 | Refresh: Console-SRS.15914 and CT-NM.SRS.SecPri.245 conflicted, any local administrator users and Philips service user can Reset password for default Local administrator user (IT) | WIP12 | DHF298555 Rev 06 - BB SRS DHF298554 Rev 06 - Ing/BR64 SRS DHF298553 Rev 06 - iCT/iCT-SP SRS DHF298556 Rev 07 - IQon SRS |  | System | SRS | CT-NM.SRS.SecPri.245 | Fixed | mistry_h | 0 | 3 | Requirements Documentation | Support-System | NO | Closed | BB_WIP5_SegReq | Ovation |  | DHF298555 Rev 06 - BB SRS DHF298554 Rev 06 - Ing/BR64 SRS DHF298553 Rev 06 - iCT/iCT-SP SRS DHF298556 Rev 07 - IQon SRS | July 26, 2019 8:16:45 PM GMT+08:00
CTUDT00621586 | Refresh | Closed | zhang_t | February 3, 2019 10:17:06 AM GMT+08:00 | Refresh: Console-SRS.15914 and CT-NM.SRS.SecPri.245 conflicted, any local administrator users and Philips service user can Reset password for default Local administrator user (IT) | WIP12 | DHF298555 Rev 06 - BB SRS DHF298554 Rev 06 - Ing/BR64 SRS DHF298553 Rev 06 - iCT/iCT-SP SRS DHF298556 Rev 07 - IQon SRS |  | System | SRS | CT-NM.SRS.SecPri.245 | Fixed | mistry_h | 0 | 3 | Requirements Documentation | Support-System | NO | Closed | BB_WIP5_SegReq | Brilliance 64 |  | DHF298555 Rev 06 - BB SRS DHF298554 Rev 06 - Ing/BR64 SRS DHF298553 Rev 06 - iCT/iCT-SP SRS DHF298556 Rev 07 - IQon SRS | July 26, 2019 8:16:45 PM GMT+08:00
CTUDT00621586 | Refresh | Closed | zhang_t | February 3, 2019 10:17:06 AM GMT+08:00 | Refresh: Console-SRS.15914 and CT-NM.SRS.SecPri.245 conflicted, any local administrator users and Philips service user can Reset password for default Local administrator user (IT) | WIP12 | DHF298555 Rev 06 - BB SRS DHF298554 Rev 06 - Ing/BR64 SRS DHF298553 Rev 06 - iCT/iCT-SP SRS DHF298556 Rev 07 - IQon SRS |  | System | SRS | CT-NM.SRS.SecPri.245 | Fixed | mistry_h | 0 | 3 | Requirements Documentation | Support-System | NO | Closed | BB_WIP5_SegReq | Brilliance Big Bore |  | DHF298555 Rev 06 - BB SRS DHF298554 Rev 06 - Ing/BR64 SRS DHF298553 Rev 06 - iCT/iCT-SP SRS DHF298556 Rev 07 - IQon SRS | July 26, 2019 8:16:45 PM GMT+08:00
CTUDT00621586 | Refresh | Closed | zhang_t | February 3, 2019 10:17:06 AM GMT+08:00 | Refresh: Console-SRS.15914 and CT-NM.SRS.SecPri.245 conflicted, any local administrator users and Philips service user can Reset password for default Local administrator user (IT) | WIP12 | DHF298555 Rev 06 - BB SRS DHF298554 Rev 06 - Ing/BR64 SRS DHF298553 Rev 06 - iCT/iCT-SP SRS DHF298556 Rev 07 - IQon SRS |  | System | SRS | CT-NM.SRS.SecPri.245 | Fixed | mistry_h | 0 | 3 | Requirements Documentation | Support-System | NO | Closed | BB_WIP5_SegReq | Ingenuity 64 |  | DHF298555 Rev 06 - BB SRS DHF298554 Rev 06 - Ing/BR64 SRS DHF298553 Rev 06 - iCT/iCT-SP SRS DHF298556 Rev 07 - IQon SRS | July 26, 2019 8:16:45 PM GMT+08:00
CTUDT00621586 | Refresh | Closed | zhang_t | February 3, 2019 10:17:06 AM GMT+08:00 | Refresh: Console-SRS.15914 and CT-NM.SRS.SecPri.245 conflicted, any local administrator users and Philips service user can Reset password for default Local administrator user (IT) | WIP12 | DHF298555 Rev 06 - BB SRS DHF298554 Rev 06 - Ing/BR64 SRS DHF298553 Rev 06 - iCT/iCT-SP SRS DHF298556 Rev 07 - IQon SRS |  | System | SRS | CT-NM.SRS.SecPri.245 | Fixed | mistry_h | 0 | 3 | Requirements Documentation | Support-System | NO | Closed | BB_WIP5_SegReq | IQon1 |  | DHF298555 Rev 06 - BB SRS DHF298554 Rev 06 - Ing/BR64 SRS DHF298553 Rev 06 - iCT/iCT-SP SRS DHF298556 Rev 07 - IQon SRS | July 26, 2019 8:16:45 PM GMT+08:00
CTUDT00621584 | Refresh | Closed | zhang_t | February 3, 2019 10:08:31 AM GMT+08:00 | Refresh: Long emergency user name is cut off in workflow GUI and operator name cut off on image annotation | WIP12 |  | NA | Subsystem | Emergency | NA | Works As Specified / Not A Defect | ravindra_a | 0 |  |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00621584 | Refresh | Closed | zhang_t | February 3, 2019 10:08:31 AM GMT+08:00 | Refresh: Long emergency user name is cut off in workflow GUI and operator name cut off on image annotation | WIP12 |  | NA | Subsystem | Emergency | NA | Works As Specified / Not A Defect | ravindra_a | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00621582 | Refresh | Closed | lei_yang | February 3, 2019 9:59:03 AM GMT+08:00 | "The data processing has reach the maximum iteration limit." pops up when performing A-Plane Offset calibration. | BB_WIP3 |  |  | Subsystem | Gantry Calibrations |  | Invalid/Uncontrolled Configuration | morton_t | 0 |  |  | Support-Physics | NO | Closed |  | Brilliance Big Bore
CTUDT00621580 | Refresh | Closed | zhang_t | February 3, 2019 9:31:21 AM GMT+08:00 | Refresh: When password complexity policy is disabled, the password character length is still effective, and system still prompts not comply with password complexity policy if password length is small then the setting. | WIP12 |  | This is not an issue, works as specified | Subsystem | User Managemant Tool | NA | Works As Specified / Not A Defect | ravindra_a | 0 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00621580 | Refresh | Closed | zhang_t | February 3, 2019 9:31:21 AM GMT+08:00 | Refresh: When password complexity policy is disabled, the password character length is still effective, and system still prompts not comply with password complexity policy if password length is small then the setting. | WIP12 |  | This is not an issue, works as specified | Subsystem | User Managemant Tool | NA | Works As Specified / Not A Defect | ravindra_a | 0 |  |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00621580 | Refresh | Closed | zhang_t | February 3, 2019 9:31:21 AM GMT+08:00 | Refresh: When password complexity policy is disabled, the password character length is still effective, and system still prompts not comply with password complexity policy if password length is small then the setting. | WIP12 |  | This is not an issue, works as specified | Subsystem | User Managemant Tool | NA | Works As Specified / Not A Defect | ravindra_a | 0 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00621580 | Refresh | Closed | zhang_t | February 3, 2019 9:31:21 AM GMT+08:00 | Refresh: When password complexity policy is disabled, the password character length is still effective, and system still prompts not comply with password complexity policy if password length is small then the setting. | WIP12 |  | This is not an issue, works as specified | Subsystem | User Managemant Tool | NA | Works As Specified / Not A Defect | ravindra_a | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00621580 | Refresh | Closed | zhang_t | February 3, 2019 9:31:21 AM GMT+08:00 | Refresh: When password complexity policy is disabled, the password character length is still effective, and system still prompts not comply with password complexity policy if password length is small then the setting. | WIP12 |  | This is not an issue, works as specified | Subsystem | User Managemant Tool | NA | Works As Specified / Not A Defect | ravindra_a | 0 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00621578 | Refresh | Closed | zhang_t | February 3, 2019 9:26:27 AM GMT+08:00 | Refresh: Local Administrator user account cannot reset password of Philips service user account which against Console-SRS.16091. | WIP12 |  | No workaround available | Subsystem | User Management Tool | NA | Works As Specified / Not A Defect | zhang_t | 1 |  |  | Support-System | NO | Closed |  | Ovation
CTUDT00621578 | Refresh | Closed | zhang_t | February 3, 2019 9:26:27 AM GMT+08:00 | Refresh: Local Administrator user account cannot reset password of Philips service user account which against Console-SRS.16091. | WIP12 |  | No workaround available | Subsystem | User Management Tool | NA | Works As Specified / Not A Defect | zhang_t | 1 |  |  | Support-System | NO | Closed |  | Brilliance 64
CTUDT00621578 | Refresh | Closed | zhang_t | February 3, 2019 9:26:27 AM GMT+08:00 | Refresh: Local Administrator user account cannot reset password of Philips service user account which against Console-SRS.16091. | WIP12 |  | No workaround available | Subsystem | User Management Tool | NA | Works As Specified / Not A Defect | zhang_t | 1 |  |  | Support-System | NO | Closed |  | Brilliance Big Bore
CTUDT00621578 | Refresh | Closed | zhang_t | February 3, 2019 9:26:27 AM GMT+08:00 | Refresh: Local Administrator user account cannot reset password of Philips service user account which against Console-SRS.16091. | WIP12 |  | No workaround available | Subsystem | User Management Tool | NA | Works As Specified / Not A Defect | zhang_t | 1 |  |  | Support-System | NO | Closed |  | Ingenuity 64
CTUDT00621578 | Refresh | Closed | zhang_t | February 3, 2019 9:26:27 AM GMT+08:00 | Refresh: Local Administrator user account cannot reset password of Philips service user account which against Console-SRS.16091. | WIP12 |  | No workaround available | Subsystem | User Management Tool | NA | Works As Specified / Not A Defect | zhang_t | 1 |  |  | Support-System | NO | Closed |  | IQon1
CTUDT00621576 | Refresh | Closed | zhang_t | February 3, 2019 9:20:10 AM GMT+08:00 | Refresh: Other user accounts besides service user can also be able to access file system and modify system folder and files (copy, paste and delete) via Snagit tool. | WIP12 |  |  | Subsystem | User Management | N/A | Works As Specified / Not A Defect | dasgupta_s | 0 |  |  | Console4 | NO | Closed | WIP16SubCom | Ingenuity 64
CTUDT00621574 | Refresh | Closed | zhang_t | February 3, 2019 9:17:39 AM GMT+08:00 | Refresh: User created local administrator user cannot modify privilege to view ePHI logs for other local administrator (IT user) | WIP12 | HOST SSRS DHF303211Rev 08 |  | Subsystem | user management tool | Console-SRS.15914 | Fixed | hatti_c | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | WIP15SegReq | Ovation |  | HOST SSRS DHF303211Rev 08 | April 23, 2019 6:45:58 PM GMT+08:00
CTUDT00621574 | Refresh | Closed | zhang_t | February 3, 2019 9:17:39 AM GMT+08:00 | Refresh: User created local administrator user cannot modify privilege to view ePHI logs for other local administrator (IT user) | WIP12 | HOST SSRS DHF303211Rev 08 |  | Subsystem | user management tool | Console-SRS.15914 | Fixed | hatti_c | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | WIP15SegReq | Brilliance 64 |  | HOST SSRS DHF303211Rev 08 | April 23, 2019 6:45:58 PM GMT+08:00
CTUDT00621574 | Refresh | Closed | zhang_t | February 3, 2019 9:17:39 AM GMT+08:00 | Refresh: User created local administrator user cannot modify privilege to view ePHI logs for other local administrator (IT user) | WIP12 | HOST SSRS DHF303211Rev 08 |  | Subsystem | user management tool | Console-SRS.15914 | Fixed | hatti_c | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | WIP15SegReq | Brilliance Big Bore |  | HOST SSRS DHF303211Rev 08 | April 23, 2019 6:45:58 PM GMT+08:00
CTUDT00621574 | Refresh | Closed | zhang_t | February 3, 2019 9:17:39 AM GMT+08:00 | Refresh: User created local administrator user cannot modify privilege to view ePHI logs for other local administrator (IT user) | WIP12 | HOST SSRS DHF303211Rev 08 |  | Subsystem | user management tool | Console-SRS.15914 | Fixed | hatti_c | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | WIP15SegReq | Ingenuity 64 |  | HOST SSRS DHF303211Rev 08 | April 23, 2019 6:45:58 PM GMT+08:00
CTUDT00621574 | Refresh | Closed | zhang_t | February 3, 2019 9:17:39 AM GMT+08:00 | Refresh: User created local administrator user cannot modify privilege to view ePHI logs for other local administrator (IT user) | WIP12 | HOST SSRS DHF303211Rev 08 |  | Subsystem | user management tool | Console-SRS.15914 | Fixed | hatti_c | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | WIP15SegReq | IQon1 |  | HOST SSRS DHF303211Rev 08 | April 23, 2019 6:45:58 PM GMT+08:00
CTUDT00621571 | Refresh | Closed | lei_yang | February 2, 2019 4:59:51 PM GMT+08:00 | Stage Body 1 didn't pass by Acceptance Test. | BB_WIP3 | WIP15 | N/A | Subsystem | Acceptance Test | N/A | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | NO | Closed | BB_WIP6_SubCom | Brilliance Big Bore |  | WIP15 | March 31, 2019 8:20:46 PM GMT+08:00
CTUDT00621569 | Refresh | Closed | lei_yang | February 2, 2019 4:19:11 PM GMT+08:00 | Stage Spatial Resolution 2 didn't pass by Constancy Test. | BB_WIP3 |  |  | Subsystem | Blocker |  | Duplicate | lei_yang | 0 |  |  | Support-System | NO | Closed |  | Brilliance Big Bore
CTUDT00621567 | Refresh | Closed | lei_yang | February 2, 2019 3:15:26 PM GMT+08:00 | Beam and DMS Check calibration was failed. | BB_WIP3 |  |  | Subsystem | DMS Calibrations |  | Duplicate | berent_i | 0 |  |  | DMS | NO | Closed |  | Brilliance Big Bore
CTUDT00621565 | Refresh | Closed | lei_yang | February 2, 2019 2:53:41 PM GMT+08:00 | Shadow Test completed but failed. | BB_WIP3 |  |  | Subsystem | DMS Calibrations | NA | Works As Specified / Not A Defect | bBeilin_b | 0 |  |  | DMS | NO | Closed |  | Brilliance Big Bore
CTUDT00621563 | Refresh | Closed | lei_yang | February 2, 2019 2:18:19 PM GMT+08:00 | Stage Head pins lin1 didn't pass by Performance Test calibration. | BB_WIP3 |  |  | Subsystem | DMS Calibrations | N/A | Works As Specified / Not A Defect | lei_yang | 1 |  |  | Support-System | NO | Closed |  | Brilliance Big Bore
CTUDT00621561 | Refresh | Closed | lei_yang | February 2, 2019 10:44:24 AM GMT+08:00 | The instruction messages displayed by A-Plane Encoder calibration is incorrect. | BB_WIP3 | BB_WIP5 WIP14 | Bypass the prompt and continue the calibration | Subsystem | Gantry Calibrations | N/A | Fixed | zhou_x | 0 | 2 | Gantry Calibrations | GCX | NO | Closed | BB_WIP5_SubCom | Brilliance Big Bore |  | BB_WIP5 WIP14 | March 21, 2019 2:29:05 PM GMT+08:00
CTUDT00621561 | Refresh | Closed | lei_yang | February 2, 2019 10:44:24 AM GMT+08:00 | The instruction messages displayed by A-Plane Encoder calibration is incorrect. | BB_WIP3 | BB_WIP5 WIP14 | Bypass the prompt and continue the calibration | Subsystem | Gantry Calibrations | N/A | Fixed | zhou_x | 0 | 2 | Gantry Calibrations | GCX | NO | Closed | BB_WIP5_SubCom | Ingenuity 64 |  | BB_WIP5 WIP14 | March 21, 2019 2:29:05 PM GMT+08:00
CTUDT00621553 | Refresh | Closed | zhang_t | February 1, 2019 3:48:42 PM GMT+08:00 | Refresh: When enabled Emergency user with privilege to access ECM, cannot login ECM with username Emergency and blank password while other user account requesting authorization. | WIP12 | WIP16 | 1. Emergency user's password can be set to a non-empty password and then he can login to ECM with that credential or 2. Unauthorized ECM user can ask any other clinical user who has ecm access and can there by use ECM or 3. Unauthorized ECM user can get access to ECM by requesting local hospital administrator | Subsystem | ECM | NA | Fixed | kapadia_v | 0 | 4 | Console Software | Console4 | NO | Closed | WIP15SegReq | Ovation |  | WIP16 | April 19, 2019 8:42:21 PM GMT+08:00
CTUDT00621553 | Refresh | Closed | zhang_t | February 1, 2019 3:48:42 PM GMT+08:00 | Refresh: When enabled Emergency user with privilege to access ECM, cannot login ECM with username Emergency and blank password while other user account requesting authorization. | WIP12 | WIP16 | 1. Emergency user's password can be set to a non-empty password and then he can login to ECM with that credential or 2. Unauthorized ECM user can ask any other clinical user who has ecm access and can there by use ECM or 3. Unauthorized ECM user can get access to ECM by requesting local hospital administrator | Subsystem | ECM | NA | Fixed | kapadia_v | 0 | 4 | Console Software | Console4 | NO | Closed | WIP15SegReq | Brilliance 64 |  | WIP16 | April 19, 2019 8:42:21 PM GMT+08:00
CTUDT00621553 | Refresh | Closed | zhang_t | February 1, 2019 3:48:42 PM GMT+08:00 | Refresh: When enabled Emergency user with privilege to access ECM, cannot login ECM with username Emergency and blank password while other user account requesting authorization. | WIP12 | WIP16 | 1. Emergency user's password can be set to a non-empty password and then he can login to ECM with that credential or 2. Unauthorized ECM user can ask any other clinical user who has ecm access and can there by use ECM or 3. Unauthorized ECM user can get access to ECM by requesting local hospital administrator | Subsystem | ECM | NA | Fixed | kapadia_v | 0 | 4 | Console Software | Console4 | NO | Closed | WIP15SegReq | Ingenuity 64 |  | WIP16 | April 19, 2019 8:42:21 PM GMT+08:00
CTUDT00621553 | Refresh | Closed | zhang_t | February 1, 2019 3:48:42 PM GMT+08:00 | Refresh: When enabled Emergency user with privilege to access ECM, cannot login ECM with username Emergency and blank password while other user account requesting authorization. | WIP12 | WIP16 | 1. Emergency user's password can be set to a non-empty password and then he can login to ECM with that credential or 2. Unauthorized ECM user can ask any other clinical user who has ecm access and can there by use ECM or 3. Unauthorized ECM user can get access to ECM by requesting local hospital administrator | Subsystem | ECM | NA | Fixed | kapadia_v | 0 | 4 | Console Software | Console4 | NO | Closed | WIP15SegReq | IQon1 |  | WIP16 | April 19, 2019 8:42:21 PM GMT+08:00
CTUDT00621548 | Refresh | Closed | Noy_Seelig | February 1, 2019 12:04:37 AM GMT+08:00 | Refresh: In MainStatForm there is no FS1 and Average check boxes | WIP12 | DHF318940 Rev 14 | This is not defect, so, workaround not needed. | Subsystem | DMS Calibrations | N/A | Fixed | yosfan_s | 0 | 3 | Test Design | DMS | NO | Closed | WIP14SegReq | Ingenuity 64 |  | DHF318940 Rev 14 | February 21, 2019 8:17:18 PM GMT+08:00
CTUDT00621546 | Refresh | Closed | levy_s | January 31, 2019 10:51:38 PM GMT+08:00 | Refresh: ESTOP opens infrequently when executing performance scenarios. | WIP12 |  |  | Subsystem | ESTOP | NA | No Plan to Fix | karkour | 0 | 4 |  | GCX | NO | Closed |  | IQon1
CTUDT00621538 | Refresh | Closed | Ashish_Sharma | January 31, 2019 6:49:41 PM GMT+08:00 | In Functional CT Application, TLCT graph is not copying to filmer when using shortcut key "F" | WIP12 |  |  | Subsystem | EBW Apps | NA | No Plan to Fix | aswani_v | 0 | 4 | EBW | iCAP | NO | Closed |  | Ovation
CTUDT00621538 | Refresh | Closed | Ashish_Sharma | January 31, 2019 6:49:41 PM GMT+08:00 | In Functional CT Application, TLCT graph is not copying to filmer when using shortcut key "F" | WIP12 |  |  | Subsystem | EBW Apps | NA | No Plan to Fix | aswani_v | 0 | 4 | EBW | iCAP | NO | Closed |  | Brilliance 64
CTUDT00621538 | Refresh | Closed | Ashish_Sharma | January 31, 2019 6:49:41 PM GMT+08:00 | In Functional CT Application, TLCT graph is not copying to filmer when using shortcut key "F" | WIP12 |  |  | Subsystem | EBW Apps | NA | No Plan to Fix | aswani_v | 0 | 4 | EBW | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00621538 | Refresh | Closed | Ashish_Sharma | January 31, 2019 6:49:41 PM GMT+08:00 | In Functional CT Application, TLCT graph is not copying to filmer when using shortcut key "F" | WIP12 |  |  | Subsystem | EBW Apps | NA | No Plan to Fix | aswani_v | 0 | 4 | EBW | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00621538 | Refresh | Closed | Ashish_Sharma | January 31, 2019 6:49:41 PM GMT+08:00 | In Functional CT Application, TLCT graph is not copying to filmer when using shortcut key "F" | WIP12 |  |  | Subsystem | EBW Apps | NA | No Plan to Fix | aswani_v | 0 | 4 | EBW | iCAP | NO | Closed |  | IQon1
CTUDT00621536 | Refresh | Closed | bommiReddy_r | January 31, 2019 5:53:54 PM GMT+08:00 | Refresh: Performance -Console is taking more than 1sec in changing the surview mode from Frontal to Lateral in ECM | WIP12 | DHF318051 Rev 13 HOST Verification Test Protocol | NA | Subsystem | Test Design | Console-SRS.10083 | Fixed | ramesh_t | 0 | 2 | Test Design | Console4 | NO | Closed | WIP15SegReq | Ovation |  | DHF318051 Rev 13 HOST Verification Test Protocol | April 23, 2019 7:50:53 PM GMT+08:00
CTUDT00621536 | Refresh | Closed | bommiReddy_r | January 31, 2019 5:53:54 PM GMT+08:00 | Refresh: Performance -Console is taking more than 1sec in changing the surview mode from Frontal to Lateral in ECM | WIP12 | DHF318051 Rev 13 HOST Verification Test Protocol | NA | Subsystem | Test Design | Console-SRS.10083 | Fixed | ramesh_t | 0 | 2 | Test Design | Console4 | NO | Closed | WIP15SegReq | Brilliance 64 |  | DHF318051 Rev 13 HOST Verification Test Protocol | April 23, 2019 7:50:53 PM GMT+08:00
CTUDT00621536 | Refresh | Closed | bommiReddy_r | January 31, 2019 5:53:54 PM GMT+08:00 | Refresh: Performance -Console is taking more than 1sec in changing the surview mode from Frontal to Lateral in ECM | WIP12 | DHF318051 Rev 13 HOST Verification Test Protocol | NA | Subsystem | Test Design | Console-SRS.10083 | Fixed | ramesh_t | 0 | 2 | Test Design | Console4 | NO | Closed | WIP15SegReq | Ingenuity 64 |  | DHF318051 Rev 13 HOST Verification Test Protocol | April 23, 2019 7:50:53 PM GMT+08:00
CTUDT00621536 | Refresh | Closed | bommiReddy_r | January 31, 2019 5:53:54 PM GMT+08:00 | Refresh: Performance -Console is taking more than 1sec in changing the surview mode from Frontal to Lateral in ECM | WIP12 | DHF318051 Rev 13 HOST Verification Test Protocol | NA | Subsystem | Test Design | Console-SRS.10083 | Fixed | ramesh_t | 0 | 2 | Test Design | Console4 | NO | Closed | WIP15SegReq | IQon1 |  | DHF318051 Rev 13 HOST Verification Test Protocol | April 23, 2019 7:50:53 PM GMT+08:00
CTUDT00621534 | Refresh | Closed | bommiReddy_r | January 31, 2019 3:45:02 PM GMT+08:00 | Refresh: Exam Prep -Console not sent Dose Structured Report to CD/DVD storage device when “create Dose Structured Report” is checked in preferences and selected CD/DVD device | WIP12 | DHF318051 Rev 17 |  | Subsystem | Reports | NA | Fixed | jindal_k | 0 | 4 | Test Design | Console4 | NO | Closed | WIP17SubCom | Ovation |  | DHF318051 Rev 17 | June 21, 2019 2:09:52 PM GMT+08:00
CTUDT00621534 | Refresh | Closed | bommiReddy_r | January 31, 2019 3:45:02 PM GMT+08:00 | Refresh: Exam Prep -Console not sent Dose Structured Report to CD/DVD storage device when “create Dose Structured Report” is checked in preferences and selected CD/DVD device | WIP12 | DHF318051 Rev 17 |  | Subsystem | Reports | NA | Fixed | jindal_k | 0 | 4 | Test Design | Console4 | NO | Closed | WIP17SubCom | Brilliance 64 |  | DHF318051 Rev 17 | June 21, 2019 2:09:52 PM GMT+08:00
CTUDT00621534 | Refresh | Closed | bommiReddy_r | January 31, 2019 3:45:02 PM GMT+08:00 | Refresh: Exam Prep -Console not sent Dose Structured Report to CD/DVD storage device when “create Dose Structured Report” is checked in preferences and selected CD/DVD device | WIP12 | DHF318051 Rev 17 |  | Subsystem | Reports | NA | Fixed | jindal_k | 0 | 4 | Test Design | Console4 | NO | Closed | WIP17SubCom | Brilliance Big Bore |  | DHF318051 Rev 17 | June 21, 2019 2:09:52 PM GMT+08:00
CTUDT00621534 | Refresh | Closed | bommiReddy_r | January 31, 2019 3:45:02 PM GMT+08:00 | Refresh: Exam Prep -Console not sent Dose Structured Report to CD/DVD storage device when “create Dose Structured Report” is checked in preferences and selected CD/DVD device | WIP12 | DHF318051 Rev 17 |  | Subsystem | Reports | NA | Fixed | jindal_k | 0 | 4 | Test Design | Console4 | NO | Closed | WIP17SubCom | Ingenuity 64 |  | DHF318051 Rev 17 | June 21, 2019 2:09:52 PM GMT+08:00
CTUDT00621534 | Refresh | Closed | bommiReddy_r | January 31, 2019 3:45:02 PM GMT+08:00 | Refresh: Exam Prep -Console not sent Dose Structured Report to CD/DVD storage device when “create Dose Structured Report” is checked in preferences and selected CD/DVD device | WIP12 | DHF318051 Rev 17 |  | Subsystem | Reports | NA | Fixed | jindal_k | 0 | 4 | Test Design | Console4 | NO | Closed | WIP17SubCom | IQon1 |  | DHF318051 Rev 17 | June 21, 2019 2:09:52 PM GMT+08:00
CTUDT00621531 | Refresh | Closed | sowmya_k | January 31, 2019 1:52:14 PM GMT+08:00 | Incorrect values for REVOLUTION_TIME and TABLE_FEED_PER_ROTATION in DICOM files on Ingenuity CT 4.1.6. | 4.1.6.25032 | WIP14 |  | Subsystem | DICOM | NA | Fixed | kumar_v | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP5_SegReq | Brilliance Big Bore | PR6334765 | WIP14 | February 25, 2019 4:58:12 PM GMT+08:00
CTUDT00621527 | Refresh | Closed | Noy_Seelig | January 31, 2019 2:01:53 AM GMT+08:00 | Refresh: Instead of Modules Temperature table, there is "Error during" empty table | BB_WIP3 | WIP14 |  | Subsystem | DMS Calibrations | n/a | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | BB_WIP5_SegReq | Brilliance Big Bore |  | WIP14 | February 24, 2019 5:56:00 PM GMT+08:00
CTUDT00621525 | Refresh | Closed | oreg_o | January 30, 2019 11:24:18 PM GMT+08:00 | Refresh: Calibrations: CTDI: The Probe was not installed and the correct message did not display during centering. | WIP12 | WIP14 |  | System | DMS Calibrations | n/a | Fixed | daibes_s | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP14SegReq | Ovation |  | WIP14 | February 13, 2019 8:54:24 PM GMT+08:00
CTUDT00621525 | Refresh | Closed | oreg_o | January 30, 2019 11:24:18 PM GMT+08:00 | Refresh: Calibrations: CTDI: The Probe was not installed and the correct message did not display during centering. | WIP12 | WIP14 |  | System | DMS Calibrations | n/a | Fixed | daibes_s | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP14SegReq | Brilliance Big Bore |  | WIP14 | February 13, 2019 8:54:24 PM GMT+08:00
CTUDT00621525 | Refresh | Closed | oreg_o | January 30, 2019 11:24:18 PM GMT+08:00 | Refresh: Calibrations: CTDI: The Probe was not installed and the correct message did not display during centering. | WIP12 | WIP14 |  | System | DMS Calibrations | n/a | Fixed | daibes_s | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP14SegReq | Ingenuity 64 |  | WIP14 | February 13, 2019 8:54:24 PM GMT+08:00
CTUDT00621525 | Refresh | Closed | oreg_o | January 30, 2019 11:24:18 PM GMT+08:00 | Refresh: Calibrations: CTDI: The Probe was not installed and the correct message did not display during centering. | WIP12 | WIP14 |  | System | DMS Calibrations | n/a | Fixed | daibes_s | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP14SegReq | IQon1 |  | WIP14 | February 13, 2019 8:54:24 PM GMT+08:00
CTUDT00621520 | Refresh | Closed | Valeria_Kashtanov | January 30, 2019 9:32:35 PM GMT+08:00 | Refresh: Crosstalk Test: Test results only show failed cases in red - without showing pass cases in green | BB_WIP3 | N/A | N/A | Subsystem | DMS Calibrations | DHF302490 {V4} IND004 | Fixed | yosfan_s | 0 | 2 | Test Design | DMS | NO | Closed | WIP15SegReq | Ovation |  | N/A | March 27, 2019 3:52:14 PM GMT+08:00
CTUDT00621520 | Refresh | Closed | Valeria_Kashtanov | January 30, 2019 9:32:35 PM GMT+08:00 | Refresh: Crosstalk Test: Test results only show failed cases in red - without showing pass cases in green | BB_WIP3 | N/A | N/A | Subsystem | DMS Calibrations | DHF302490 {V4} IND004 | Fixed | yosfan_s | 0 | 2 | Test Design | DMS | NO | Closed | WIP15SegReq | Brilliance 64 |  | N/A | March 27, 2019 3:52:14 PM GMT+08:00
CTUDT00621520 | Refresh | Closed | Valeria_Kashtanov | January 30, 2019 9:32:35 PM GMT+08:00 | Refresh: Crosstalk Test: Test results only show failed cases in red - without showing pass cases in green | BB_WIP3 | N/A | N/A | Subsystem | DMS Calibrations | DHF302490 {V4} IND004 | Fixed | yosfan_s | 0 | 2 | Test Design | DMS | NO | Closed | WIP15SegReq | Brilliance Big Bore |  | N/A | March 27, 2019 3:52:14 PM GMT+08:00
CTUDT00621520 | Refresh | Closed | Valeria_Kashtanov | January 30, 2019 9:32:35 PM GMT+08:00 | Refresh: Crosstalk Test: Test results only show failed cases in red - without showing pass cases in green | BB_WIP3 | N/A | N/A | Subsystem | DMS Calibrations | DHF302490 {V4} IND004 | Fixed | yosfan_s | 0 | 2 | Test Design | DMS | NO | Closed | WIP15SegReq | IQon1 |  | N/A | March 27, 2019 3:52:14 PM GMT+08:00
CTUDT00621506 | Refresh | Closed | Dana_Kamyshev | January 30, 2019 3:45:12 PM GMT+08:00 | Refresh - Mandatory information in the Reports can be edited | WIP12 | WIP17 |  | Subsystem | Report;Verified1 | NA | Fixed | redrouthu_r | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP15SegReq | Ovation |  | WIP17 | June 10, 2019 4:54:03 PM GMT+08:00
CTUDT00621506 | Refresh | Closed | Dana_Kamyshev | January 30, 2019 3:45:12 PM GMT+08:00 | Refresh - Mandatory information in the Reports can be edited | WIP12 | WIP17 |  | Subsystem | Report;Verified1 | NA | Fixed | redrouthu_r | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP15SegReq | Brilliance 64 |  | WIP17 | June 10, 2019 4:54:03 PM GMT+08:00
CTUDT00621506 | Refresh | Closed | Dana_Kamyshev | January 30, 2019 3:45:12 PM GMT+08:00 | Refresh - Mandatory information in the Reports can be edited | WIP12 | WIP17 |  | Subsystem | Report;Verified1 | NA | Fixed | redrouthu_r | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP15SegReq | Brilliance Big Bore |  | WIP17 | June 10, 2019 4:54:03 PM GMT+08:00
CTUDT00621506 | Refresh | Closed | Dana_Kamyshev | January 30, 2019 3:45:12 PM GMT+08:00 | Refresh - Mandatory information in the Reports can be edited | WIP12 | WIP17 |  | Subsystem | Report;Verified1 | NA | Fixed | redrouthu_r | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP15SegReq | Ingenuity 64 |  | WIP17 | June 10, 2019 4:54:03 PM GMT+08:00
CTUDT00621506 | Refresh | Closed | Dana_Kamyshev | January 30, 2019 3:45:12 PM GMT+08:00 | Refresh - Mandatory information in the Reports can be edited | WIP12 | WIP17 |  | Subsystem | Report;Verified1 | NA | Fixed | redrouthu_r | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP15SegReq | IQon1 |  | WIP17 | June 10, 2019 4:54:03 PM GMT+08:00
CTUDT00621502 | Refresh | Closed | panchami_pradeep | January 30, 2019 2:11:14 PM GMT+08:00 | Refresh: CTDI value displayed in Exam Summary page is not within the expected range of calculated value.. | WIP12 | DHF318051 Rev 11 | NA | Subsystem | Exam Summary | Console-SRS.2168 | Fixed | ramesh_t | 0 | 2 | Test Design | Console4 | NO | Closed | WIP14SegReq | Ovation |  | DHF318051 Rev 11 | March 25, 2019 7:57:00 PM GMT+08:00
CTUDT00621502 | Refresh | Closed | panchami_pradeep | January 30, 2019 2:11:14 PM GMT+08:00 | Refresh: CTDI value displayed in Exam Summary page is not within the expected range of calculated value.. | WIP12 | DHF318051 Rev 11 | NA | Subsystem | Exam Summary | Console-SRS.2168 | Fixed | ramesh_t | 0 | 2 | Test Design | Console4 | NO | Closed | WIP14SegReq | Brilliance 64 |  | DHF318051 Rev 11 | March 25, 2019 7:57:00 PM GMT+08:00
CTUDT00621502 | Refresh | Closed | panchami_pradeep | January 30, 2019 2:11:14 PM GMT+08:00 | Refresh: CTDI value displayed in Exam Summary page is not within the expected range of calculated value.. | WIP12 | DHF318051 Rev 11 | NA | Subsystem | Exam Summary | Console-SRS.2168 | Fixed | ramesh_t | 0 | 2 | Test Design | Console4 | NO | Closed | WIP14SegReq | Ingenuity 64 |  | DHF318051 Rev 11 | March 25, 2019 7:57:00 PM GMT+08:00
CTUDT00621502 | Refresh | Closed | panchami_pradeep | January 30, 2019 2:11:14 PM GMT+08:00 | Refresh: CTDI value displayed in Exam Summary page is not within the expected range of calculated value.. | WIP12 | DHF318051 Rev 11 | NA | Subsystem | Exam Summary | Console-SRS.2168 | Fixed | ramesh_t | 0 | 2 | Test Design | Console4 | NO | Closed | WIP14SegReq | IQon1 |  | DHF318051 Rev 11 | March 25, 2019 7:57:00 PM GMT+08:00
CTUDT00621498 | Refresh | Closed | Valeria_Kashtanov | January 30, 2019 4:09:56 AM GMT+08:00 | Refresh: Crosstalk Test: Wrong plot names for half-module results | BB_WIP3 | WIP14 | Not needed for sane user | Subsystem | Crosstalk Test | N/A | Fixed | yosefk | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP14SegReq | Brilliance Big Bore |  | WIP14 | February 13, 2019 3:19:47 PM GMT+08:00
CTUDT00621496 | Refresh | Closed | golzban_v | January 30, 2019 3:12:17 AM GMT+08:00 | Refresh: Limit Plan Geometry EC behaves inconsistently when loaded to Plan along with EC with Limit Plan Geometry option disabled | WIP12 |  | NA | Subsystem | Plan | NA | Works As Specified / Not A Defect | ravindra_a | 0 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00621496 | Refresh | Closed | golzban_v | January 30, 2019 3:12:17 AM GMT+08:00 | Refresh: Limit Plan Geometry EC behaves inconsistently when loaded to Plan along with EC with Limit Plan Geometry option disabled | WIP12 |  | NA | Subsystem | Plan | NA | Works As Specified / Not A Defect | ravindra_a | 0 |  |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00621496 | Refresh | Closed | golzban_v | January 30, 2019 3:12:17 AM GMT+08:00 | Refresh: Limit Plan Geometry EC behaves inconsistently when loaded to Plan along with EC with Limit Plan Geometry option disabled | WIP12 |  | NA | Subsystem | Plan | NA | Works As Specified / Not A Defect | ravindra_a | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00621496 | Refresh | Closed | golzban_v | January 30, 2019 3:12:17 AM GMT+08:00 | Refresh: Limit Plan Geometry EC behaves inconsistently when loaded to Plan along with EC with Limit Plan Geometry option disabled | WIP12 |  | NA | Subsystem | Plan | NA | Works As Specified / Not A Defect | ravindra_a | 0 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00621494 | Refresh | Closed | levy_s | January 29, 2019 11:00:19 PM GMT+08:00 | Refresh: Recon Box cannot initialize error displayed to the user in several scenarios | WIP12 | WIP14 | NA | Subsystem | Rule System; Verified1 | N/A | Fixed | rakic_k | 0 | 2 | CIRS_SW_ImgNodeSvr | CIRS | YES | Closed | WIP14SubCom | Ovation |  | WIP14 | March 18, 2019 8:47:38 PM GMT+08:00
CTUDT00621494 | Refresh | Closed | levy_s | January 29, 2019 11:00:19 PM GMT+08:00 | Refresh: Recon Box cannot initialize error displayed to the user in several scenarios | WIP12 | WIP14 | NA | Subsystem | Rule System; Verified1 | N/A | Fixed | rakic_k | 0 | 2 | CIRS_SW_ImgNodeSvr | CIRS | YES | Closed | WIP14SubCom | IQon1 |  | WIP14 | March 18, 2019 8:47:38 PM GMT+08:00
CTUDT00621489 | Refresh | Closed | Valeria_Kashtanov | January 29, 2019 9:27:26 PM GMT+08:00 | Refresh: Beam and DMS Check show low signal in results when contrast material is detected | WIP12 | WIP14 | N/A | Subsystem | DMS Calibrations | N/A | Fixed | yosefk | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP14SegReq | Ingenuity 64 |  | WIP14 | February 27, 2019 2:48:30 PM GMT+08:00
CTUDT00621484 | Refresh | Closed | bommiReddy_r | January 29, 2019 5:37:44 PM GMT+08:00 | Refresh:Plan- After performing dual surview, changing the patient position and repeating the surview, message displays to repeat the surview again | WIP12 | WIP15 |  | Subsystem | dual surview | N/A | Fixed | sagar_v | 0 | 2 | Console Software | Console4 | YES | Closed | WIP14SegReq | Ovation |  | WIP15 | March 5, 2019 8:58:09 PM GMT+08:00
CTUDT00621484 | Refresh | Closed | bommiReddy_r | January 29, 2019 5:37:44 PM GMT+08:00 | Refresh:Plan- After performing dual surview, changing the patient position and repeating the surview, message displays to repeat the surview again | WIP12 | WIP15 |  | Subsystem | dual surview | N/A | Fixed | sagar_v | 0 | 2 | Console Software | Console4 | YES | Closed | WIP14SegReq | Brilliance 64 |  | WIP15 | March 5, 2019 8:58:09 PM GMT+08:00
CTUDT00621484 | Refresh | Closed | bommiReddy_r | January 29, 2019 5:37:44 PM GMT+08:00 | Refresh:Plan- After performing dual surview, changing the patient position and repeating the surview, message displays to repeat the surview again | WIP12 | WIP15 |  | Subsystem | dual surview | N/A | Fixed | sagar_v | 0 | 2 | Console Software | Console4 | YES | Closed | WIP14SegReq | Ingenuity 64 |  | WIP15 | March 5, 2019 8:58:09 PM GMT+08:00
CTUDT00621484 | Refresh | Closed | bommiReddy_r | January 29, 2019 5:37:44 PM GMT+08:00 | Refresh:Plan- After performing dual surview, changing the patient position and repeating the surview, message displays to repeat the surview again | WIP12 | WIP15 |  | Subsystem | dual surview | N/A | Fixed | sagar_v | 0 | 2 | Console Software | Console4 | YES | Closed | WIP14SegReq | IQon1 |  | WIP15 | March 5, 2019 8:58:09 PM GMT+08:00
CTUDT00621482 | Refresh | Closed | nayak_r | January 29, 2019 4:36:45 PM GMT+08:00 | BB Refresh: In ECM,selection of specific scantype is not getting allowed | BB_WIP3 | BB_WIP5.1 | NA | Subsystem | ECM | NA | Fixed | navale_s | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP5_SubCom | Brilliance Big Bore |  | BB_WIP5.1 | March 28, 2019 6:44:20 PM GMT+08:00
CTUDT00621475 | Refresh | Closed | hiremath_b | January 29, 2019 3:06:11 AM GMT+08:00 | Refresh ING: Can’t perform Head and Body scans having Detailed Resolution with any kV value. Gives system error message | WIP12 |  | NA as it's not a defect | Subsystem | Rules System | NA | Invalid/Uncontrolled Configuration | richards_b | 0 |  |  | CIRS | NO | Closed |  | Ingenuity 64
CTUDT00621473 | Refresh | Closed | hiremath_b | January 29, 2019 2:43:20 AM GMT+08:00 | Refresh: Bugrep Error window comes Up on Ending the study, when tried to save an EC in Plan. | WIP12 | WIP16 |  | Subsystem | Bugrep | NA | Fixed | wilsonraja | 0 | 4 | Console Software | Console4 | NO | Closed | WIP14SegReq | Ovation |  | WIP16 | April 11, 2019 5:33:01 PM GMT+08:00
CTUDT00621473 | Refresh | Closed | hiremath_b | January 29, 2019 2:43:20 AM GMT+08:00 | Refresh: Bugrep Error window comes Up on Ending the study, when tried to save an EC in Plan. | WIP12 | WIP16 |  | Subsystem | Bugrep | NA | Fixed | wilsonraja | 0 | 4 | Console Software | Console4 | NO | Closed | WIP14SegReq | Brilliance 64 |  | WIP16 | April 11, 2019 5:33:01 PM GMT+08:00
CTUDT00621473 | Refresh | Closed | hiremath_b | January 29, 2019 2:43:20 AM GMT+08:00 | Refresh: Bugrep Error window comes Up on Ending the study, when tried to save an EC in Plan. | WIP12 | WIP16 |  | Subsystem | Bugrep | NA | Fixed | wilsonraja | 0 | 4 | Console Software | Console4 | NO | Closed | WIP14SegReq | Ingenuity 64 |  | WIP16 | April 11, 2019 5:33:01 PM GMT+08:00
CTUDT00621473 | Refresh | Closed | hiremath_b | January 29, 2019 2:43:20 AM GMT+08:00 | Refresh: Bugrep Error window comes Up on Ending the study, when tried to save an EC in Plan. | WIP12 | WIP16 |  | Subsystem | Bugrep | NA | Fixed | wilsonraja | 0 | 4 | Console Software | Console4 | NO | Closed | WIP14SegReq | IQon1 |  | WIP16 | April 11, 2019 5:33:01 PM GMT+08:00
CTUDT00621469 | Refresh | Closed | naik_p | January 28, 2019 6:13:11 PM GMT+08:00 | Refresh_WIP12_Mininum FOV accepted for cardiac Axial was less than the given range. | WIP12 | WIP14 |  | Subsystem | Cardiac | NA | Fixed | datta_i | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SegReq | Ovation |  | WIP14 | February 20, 2019 5:37:44 PM GMT+08:00
CTUDT00621469 | Refresh | Closed | naik_p | January 28, 2019 6:13:11 PM GMT+08:00 | Refresh_WIP12_Mininum FOV accepted for cardiac Axial was less than the given range. | WIP12 | WIP14 |  | Subsystem | Cardiac | NA | Fixed | datta_i | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SegReq | Brilliance 64 |  | WIP14 | February 20, 2019 5:37:44 PM GMT+08:00
CTUDT00621469 | Refresh | Closed | naik_p | January 28, 2019 6:13:11 PM GMT+08:00 | Refresh_WIP12_Mininum FOV accepted for cardiac Axial was less than the given range. | WIP12 | WIP14 |  | Subsystem | Cardiac | NA | Fixed | datta_i | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SegReq | Ingenuity 64 |  | WIP14 | February 20, 2019 5:37:44 PM GMT+08:00
CTUDT00621469 | Refresh | Closed | naik_p | January 28, 2019 6:13:11 PM GMT+08:00 | Refresh_WIP12_Mininum FOV accepted for cardiac Axial was less than the given range. | WIP12 | WIP14 |  | Subsystem | Cardiac | NA | Fixed | datta_i | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SegReq | IQon1 |  | WIP14 | February 20, 2019 5:37:44 PM GMT+08:00
CTUDT00621467 | Refresh | Closed | panchami_pradeep | January 28, 2019 5:08:04 PM GMT+08:00 | Refresh: Default Rotation Time value for Head & Body Dual Energy scans displayed is not matching with the default value specified in the scan type document. | WIP12 | WIP14 | NA | Subsystem | scan type | NA | Fixed | kr_m | 0 | 4 | Console Software | Console4 | NO | Closed | WIP14SegReq | Brilliance 64 |  | WIP14 | February 11, 2019 7:22:33 PM GMT+08:00
CTUDT00621467 | Refresh | Closed | panchami_pradeep | January 28, 2019 5:08:04 PM GMT+08:00 | Refresh: Default Rotation Time value for Head & Body Dual Energy scans displayed is not matching with the default value specified in the scan type document. | WIP12 | WIP14 | NA | Subsystem | scan type | NA | Fixed | kr_m | 0 | 4 | Console Software | Console4 | NO | Closed | WIP14SegReq | Ingenuity 64 |  | WIP14 | February 11, 2019 7:22:33 PM GMT+08:00
CTUDT00621464 | Refresh | Closed | PremKumar_Rajendran | January 28, 2019 2:43:25 PM GMT+08:00 | BB Refresh: Issues on ‘Edit Results in Slab Viewer/Edit Results on Surview’ and ‘Slab’ button on tools area when loading results with Scan parameter ‘Axial Increment 0’ | BB_WIP2 | BB_WIP5 |  | Subsystem |  | NA | Fixed | datta_i | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP5_SegReq | Brilliance Big Bore |  | BB_WIP5 | March 1, 2019 7:35:24 PM GMT+08:00
CTUDT00621461 | Refresh | Closed | Valeria_Kashtanov | January 27, 2019 10:35:11 PM GMT+08:00 | Refresh: Calibrations: Tube Current: Results mA and mAs limits are not as defined in spec | WIP12 | N/A |  | Subsystem | Gantry Calibrations | DHF279250 Appendix 12.5 Ingenuity Scan and test parameters | Fixed | morton_t | 0 | 2 | DMS Calibrations | DMS | NO | Closed | BB_WIP5_SegReq | Ingenuity 64 |  | N/A | March 20, 2019 8:53:48 PM GMT+08:00
CTUDT00621458 | Refresh | Closed | swack_g | January 27, 2019 7:20:49 PM GMT+08:00 | Missing Results when Beam and DMS Check detects a crack in the wedge | WIP12, WIP 11 | WIP14 | N/A | Subsystem | DMS Calibrations | N/A | Fixed | yosefk | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP14SegReq | Ovation |  | WIP14 | February 17, 2019 5:56:52 PM GMT+08:00
CTUDT00621458 | Refresh | Closed | swack_g | January 27, 2019 7:20:49 PM GMT+08:00 | Missing Results when Beam and DMS Check detects a crack in the wedge | WIP12, WIP 11 | WIP14 | N/A | Subsystem | DMS Calibrations | N/A | Fixed | yosefk | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP14SegReq | Brilliance Big Bore |  | WIP14 | February 17, 2019 5:56:52 PM GMT+08:00
CTUDT00621458 | Refresh | Closed | swack_g | January 27, 2019 7:20:49 PM GMT+08:00 | Missing Results when Beam and DMS Check detects a crack in the wedge | WIP12, WIP 11 | WIP14 | N/A | Subsystem | DMS Calibrations | N/A | Fixed | yosefk | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP14SegReq | Ingenuity 64 |  | WIP14 | February 17, 2019 5:56:52 PM GMT+08:00
CTUDT00621458 | Refresh | Closed | swack_g | January 27, 2019 7:20:49 PM GMT+08:00 | Missing Results when Beam and DMS Check detects a crack in the wedge | WIP12, WIP 11 | WIP14 | N/A | Subsystem | DMS Calibrations | N/A | Fixed | yosefk | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP14SegReq | IQon1 |  | WIP14 | February 17, 2019 5:56:52 PM GMT+08:00
CTUDT00621454 | Refresh | Closed | swack_g | January 27, 2019 5:01:37 PM GMT+08:00 | Refresh: Missing sector number in the Messages window when detecting contrast material in Beam and DMS Check application | WIP11 | WIP14 | N/A | Subsystem | DMS Calibrations | N/A | Fixed | yosefk | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP14SegReq | Ovation |  | WIP14 | February 27, 2019 2:48:21 PM GMT+08:00
CTUDT00621454 | Refresh | Closed | swack_g | January 27, 2019 5:01:37 PM GMT+08:00 | Refresh: Missing sector number in the Messages window when detecting contrast material in Beam and DMS Check application | WIP11 | WIP14 | N/A | Subsystem | DMS Calibrations | N/A | Fixed | yosefk | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP14SegReq | Brilliance Big Bore |  | WIP14 | February 27, 2019 2:48:21 PM GMT+08:00
CTUDT00621454 | Refresh | Closed | swack_g | January 27, 2019 5:01:37 PM GMT+08:00 | Refresh: Missing sector number in the Messages window when detecting contrast material in Beam and DMS Check application | WIP11 | WIP14 | N/A | Subsystem | DMS Calibrations | N/A | Fixed | yosefk | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP14SegReq | Ingenuity 64 |  | WIP14 | February 27, 2019 2:48:21 PM GMT+08:00
CTUDT00621454 | Refresh | Closed | swack_g | January 27, 2019 5:01:37 PM GMT+08:00 | Refresh: Missing sector number in the Messages window when detecting contrast material in Beam and DMS Check application | WIP11 | WIP14 | N/A | Subsystem | DMS Calibrations | N/A | Fixed | yosefk | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP14SegReq | IQon1 |  | WIP14 | February 27, 2019 2:48:21 PM GMT+08:00
CTUDT00621452 | Refresh | Closed | shekhar_p | January 25, 2019 7:20:03 PM GMT+08:00 | Refresh: RCT:RCT is not working with Proxy settings | WIP12 | Test design update |  | Subsystem | Service Tools | N/A | Fixed | nanjunda_s | 0 | 3 | Test Design | Service Tools | NO | Closed | WAS | Ovation |  | Test design update | May 24, 2019 6:20:59 PM GMT+08:00
CTUDT00621452 | Refresh | Closed | shekhar_p | January 25, 2019 7:20:03 PM GMT+08:00 | Refresh: RCT:RCT is not working with Proxy settings | WIP12 | Test design update |  | Subsystem | Service Tools | N/A | Fixed | nanjunda_s | 0 | 3 | Test Design | Service Tools | NO | Closed | WAS | Brilliance 64 |  | Test design update | May 24, 2019 6:20:59 PM GMT+08:00
CTUDT00621452 | Refresh | Closed | shekhar_p | January 25, 2019 7:20:03 PM GMT+08:00 | Refresh: RCT:RCT is not working with Proxy settings | WIP12 | Test design update |  | Subsystem | Service Tools | N/A | Fixed | nanjunda_s | 0 | 3 | Test Design | Service Tools | NO | Closed | WAS | Brilliance Big Bore |  | Test design update | May 24, 2019 6:20:59 PM GMT+08:00
CTUDT00621452 | Refresh | Closed | shekhar_p | January 25, 2019 7:20:03 PM GMT+08:00 | Refresh: RCT:RCT is not working with Proxy settings | WIP12 | Test design update |  | Subsystem | Service Tools | N/A | Fixed | nanjunda_s | 0 | 3 | Test Design | Service Tools | NO | Closed | WAS | Ingenuity 64 |  | Test design update | May 24, 2019 6:20:59 PM GMT+08:00
CTUDT00621452 | Refresh | Closed | shekhar_p | January 25, 2019 7:20:03 PM GMT+08:00 | Refresh: RCT:RCT is not working with Proxy settings | WIP12 | Test design update |  | Subsystem | Service Tools | N/A | Fixed | nanjunda_s | 0 | 3 | Test Design | Service Tools | NO | Closed | WAS | IQon1 |  | Test design update | May 24, 2019 6:20:59 PM GMT+08:00
CTUDT00621450 | Refresh | Closed | bommiReddy_r | January 25, 2019 6:51:01 PM GMT+08:00 | Refresh: Performance-Console is taking more than 8 sec to end the exam | WIP12 |  |  | Subsystem | Performance |  | Duplicate | ravindra_a | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00621442 | Refresh | Closed | Valeria_Kashtanov | January 25, 2019 3:40:32 AM GMT+08:00 | Refresh: Calibrations: Tube current: Rotation time doesn't appear in description in archive for some scans | WIP12 | WIP15 |  | Subsystem | Gantry Calibrations | NaN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP15SegReq | Ovation |  | WIP15 | March 21, 2019 5:21:08 PM GMT+08:00
CTUDT00621442 | Refresh | Closed | Valeria_Kashtanov | January 25, 2019 3:40:32 AM GMT+08:00 | Refresh: Calibrations: Tube current: Rotation time doesn't appear in description in archive for some scans | WIP12 | WIP15 |  | Subsystem | Gantry Calibrations | NaN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP15SegReq | Brilliance Big Bore |  | WIP15 | March 21, 2019 5:21:08 PM GMT+08:00
CTUDT00621442 | Refresh | Closed | Valeria_Kashtanov | January 25, 2019 3:40:32 AM GMT+08:00 | Refresh: Calibrations: Tube current: Rotation time doesn't appear in description in archive for some scans | WIP12 | WIP15 |  | Subsystem | Gantry Calibrations | NaN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP15SegReq | Ingenuity 64 |  | WIP15 | March 21, 2019 5:21:08 PM GMT+08:00
CTUDT00621442 | Refresh | Closed | Valeria_Kashtanov | January 25, 2019 3:40:32 AM GMT+08:00 | Refresh: Calibrations: Tube current: Rotation time doesn't appear in description in archive for some scans | WIP12 | WIP15 |  | Subsystem | Gantry Calibrations | NaN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP15SegReq | IQon1 |  | WIP15 | March 21, 2019 5:21:08 PM GMT+08:00
CTUDT00601440 | Refresh | Closed | Valeria_Kashtanov | January 24, 2019 11:34:35 PM GMT+08:00 | Refresh: Calibrations: Acceptance fails due to calculation errors | WIP12 |  | N/A | Subsystem | DMS Calibrations | n/a | Duplicate | elviral | 0 |  |  | DMS | NO | Closed |  | Ingenuity 64
CTUDT00601438 | Refresh | Closed | attali_s | January 24, 2019 10:35:46 PM GMT+08:00 | Refresh: Physics Analysis: No option to select frame Indexes in Raw Tab of Physics Analysis | WIP12 | N/A - Test DEsign | N/A | Subsystem | DMS Calibrations | N/A | Fixed | gehtin_e | 0 | 2 | Test Design | DMS | NO | Closed | WIP17SubCom | Ingenuity 64 |  | N/A - Test DEsign | May 15, 2019 8:58:32 PM GMT+08:00
CTUDT00601436 | Refresh | Closed | Valeria_Kashtanov | January 24, 2019 9:46:28 PM GMT+08:00 | Refresh: Calibrations: Temperature Notification window closes automatically after time reaches to 0 | WIP12 | WIP14 | No Workaround. | Subsystem | Stability Tool | N/A | Fixed | kovalev_a | 0 | 4 | DMS Software | DMS | NO | Closed | WIP15SegReq | Ingenuity 64 |  | WIP14 | February 18, 2019 3:38:48 PM GMT+08:00
CTUDT00601434 | Refresh | Closed | bommiReddy_r | January 24, 2019 8:40:01 PM GMT+08:00 | Console-SRS.1118 requires update | WIP12 | Host SSRS DHF303211 Rev 08 |  | System | Console SRS | Console-SRS.1118,Console-SRS.13706 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | WIP14SegReq | Ovation |  | Host SSRS DHF303211 Rev 08 | March 22, 2019 6:09:22 PM GMT+08:00
CTUDT00601434 | Refresh | Closed | bommiReddy_r | January 24, 2019 8:40:01 PM GMT+08:00 | Console-SRS.1118 requires update | WIP12 | Host SSRS DHF303211 Rev 08 |  | System | Console SRS | Console-SRS.1118,Console-SRS.13706 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | WIP14SegReq | Brilliance 64 |  | Host SSRS DHF303211 Rev 08 | March 22, 2019 6:09:22 PM GMT+08:00
CTUDT00601434 | Refresh | Closed | bommiReddy_r | January 24, 2019 8:40:01 PM GMT+08:00 | Console-SRS.1118 requires update | WIP12 | Host SSRS DHF303211 Rev 08 |  | System | Console SRS | Console-SRS.1118,Console-SRS.13706 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | WIP14SegReq | Ingenuity 64 |  | Host SSRS DHF303211 Rev 08 | March 22, 2019 6:09:22 PM GMT+08:00
CTUDT00601431 | Refresh | Closed | attali_s | January 24, 2019 8:12:28 PM GMT+08:00 | Refresh: Common Functions: MTF Calculated from Slice Thickness does not appear in History of RnD User when performed with Service User | WIP12 | DHF318940 |  | Subsystem | DMS Calibrations | N/A | Fixed | yosfan_s | 0 | 2 | Test Design | DMS | NO | Closed | WIP14SegReq | Ovation |  | DHF318940 | February 27, 2019 4:12:42 PM GMT+08:00
CTUDT00601431 | Refresh | Closed | attali_s | January 24, 2019 8:12:28 PM GMT+08:00 | Refresh: Common Functions: MTF Calculated from Slice Thickness does not appear in History of RnD User when performed with Service User | WIP12 | DHF318940 |  | Subsystem | DMS Calibrations | N/A | Fixed | yosfan_s | 0 | 2 | Test Design | DMS | NO | Closed | WIP14SegReq | Brilliance Big Bore |  | DHF318940 | February 27, 2019 4:12:42 PM GMT+08:00
CTUDT00601431 | Refresh | Closed | attali_s | January 24, 2019 8:12:28 PM GMT+08:00 | Refresh: Common Functions: MTF Calculated from Slice Thickness does not appear in History of RnD User when performed with Service User | WIP12 | DHF318940 |  | Subsystem | DMS Calibrations | N/A | Fixed | yosfan_s | 0 | 2 | Test Design | DMS | NO | Closed | WIP14SegReq | Ingenuity 64 |  | DHF318940 | February 27, 2019 4:12:42 PM GMT+08:00
CTUDT00601431 | Refresh | Closed | attali_s | January 24, 2019 8:12:28 PM GMT+08:00 | Refresh: Common Functions: MTF Calculated from Slice Thickness does not appear in History of RnD User when performed with Service User | WIP12 | DHF318940 |  | Subsystem | DMS Calibrations | N/A | Fixed | yosfan_s | 0 | 2 | Test Design | DMS | NO | Closed | WIP14SegReq | IQon1 |  | DHF318940 | February 27, 2019 4:12:42 PM GMT+08:00
CTUDT00601430 | Refresh | Closed | singh_p | January 24, 2019 8:03:31 PM GMT+08:00 | STPS-Refresh Ing WIP 12 : Error popped up when BackUp file collection path was set to CD/DVD. | WIP12 | WIP17 BB_WIP7 |  | Subsystem | Backup/Restore | NA | Fixed | shreekanth_manniyath | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP17SubCom | Ovation |  | WIP17 BB_WIP7 | May 3, 2019 6:55:25 PM GMT+08:00
CTUDT00601430 | Refresh | Closed | singh_p | January 24, 2019 8:03:31 PM GMT+08:00 | STPS-Refresh Ing WIP 12 : Error popped up when BackUp file collection path was set to CD/DVD. | WIP12 | WIP17 BB_WIP7 |  | Subsystem | Backup/Restore | NA | Fixed | shreekanth_manniyath | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP17SubCom | Brilliance 64 |  | WIP17 BB_WIP7 | May 3, 2019 6:55:25 PM GMT+08:00
CTUDT00601430 | Refresh | Closed | singh_p | January 24, 2019 8:03:31 PM GMT+08:00 | STPS-Refresh Ing WIP 12 : Error popped up when BackUp file collection path was set to CD/DVD. | WIP12 | WIP17 BB_WIP7 |  | Subsystem | Backup/Restore | NA | Fixed | shreekanth_manniyath | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP17SubCom | Brilliance Big Bore |  | WIP17 BB_WIP7 | May 3, 2019 6:55:25 PM GMT+08:00
CTUDT00601430 | Refresh | Closed | singh_p | January 24, 2019 8:03:31 PM GMT+08:00 | STPS-Refresh Ing WIP 12 : Error popped up when BackUp file collection path was set to CD/DVD. | WIP12 | WIP17 BB_WIP7 |  | Subsystem | Backup/Restore | NA | Fixed | shreekanth_manniyath | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP17SubCom | Ingenuity 64 |  | WIP17 BB_WIP7 | May 3, 2019 6:55:25 PM GMT+08:00
CTUDT00601430 | Refresh | Closed | singh_p | January 24, 2019 8:03:31 PM GMT+08:00 | STPS-Refresh Ing WIP 12 : Error popped up when BackUp file collection path was set to CD/DVD. | WIP12 | WIP17 BB_WIP7 |  | Subsystem | Backup/Restore | NA | Fixed | shreekanth_manniyath | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP17SubCom | IQon1 |  | WIP17 BB_WIP7 | May 3, 2019 6:55:25 PM GMT+08:00
CTUDT00601425 | Refresh | Closed | jotawar_d | January 24, 2019 7:46:10 PM GMT+08:00 | Refresh [IQON WIP 8] : Voice Manager is taking more than 15 sec during launch and exit. | BB_WIP1 | BB_WIP4 |  | Subsystem | Voice Manager | NA | Fixed | Snehashis_Dey | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP4_SubCom | Brilliance Big Bore |  | BB_WIP4 | February 7, 2019 5:02:07 PM GMT+08:00
CTUDT00601414 | Refresh | Closed | attali_s | January 24, 2019 5:10:45 PM GMT+08:00 | Refresh: Common Functions: Wrong and missing units in HTML Report | WIP12 | WIP14 | N/A | Subsystem | DMS Calibrations | N/A | Fixed | yosefk | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP14SegReq | Ovation |  | WIP14 | February 20, 2019 10:58:38 PM GMT+08:00
CTUDT00601414 | Refresh | Closed | attali_s | January 24, 2019 5:10:45 PM GMT+08:00 | Refresh: Common Functions: Wrong and missing units in HTML Report | WIP12 | WIP14 | N/A | Subsystem | DMS Calibrations | N/A | Fixed | yosefk | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP14SegReq | Brilliance 64 |  | WIP14 | February 20, 2019 10:58:38 PM GMT+08:00
CTUDT00601414 | Refresh | Closed | attali_s | January 24, 2019 5:10:45 PM GMT+08:00 | Refresh: Common Functions: Wrong and missing units in HTML Report | WIP12 | WIP14 | N/A | Subsystem | DMS Calibrations | N/A | Fixed | yosefk | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP14SegReq | Brilliance Big Bore |  | WIP14 | February 20, 2019 10:58:38 PM GMT+08:00
CTUDT00601414 | Refresh | Closed | attali_s | January 24, 2019 5:10:45 PM GMT+08:00 | Refresh: Common Functions: Wrong and missing units in HTML Report | WIP12 | WIP14 | N/A | Subsystem | DMS Calibrations | N/A | Fixed | yosefk | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP14SegReq | Ingenuity 64 |  | WIP14 | February 20, 2019 10:58:38 PM GMT+08:00
CTUDT00601414 | Refresh | Closed | attali_s | January 24, 2019 5:10:45 PM GMT+08:00 | Refresh: Common Functions: Wrong and missing units in HTML Report | WIP12 | WIP14 | N/A | Subsystem | DMS Calibrations | N/A | Fixed | yosefk | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP14SegReq | IQon1 |  | WIP14 | February 20, 2019 10:58:38 PM GMT+08:00
CTUDT00601412 | Refresh | Closed | manjunath_mk | January 24, 2019 4:16:05 PM GMT+08:00 | Scan Number in RQM does not match with the number of scans performed with dual surview | BB_WIP3 | BB_WIP5 |  | Subsystem | RQM | None | Fixed | tejaswini_nidamanuri | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP5_SegReq | Brilliance Big Bore |  | BB_WIP5 | March 7, 2019 11:56:09 AM GMT+08:00
CTUDT00601408 | Refresh | Closed | patra_r | January 24, 2019 2:44:02 PM GMT+08:00 | BB Refresh: ECG and Pulmo signal is not detected in ECG viewer and Resp Viewer respectively. | V4.1.10 and V4.7.7 | WIP14 |  | Subsystem |  | None | Fixed | Prakash_Chand_Yadav | 0 | 2 | Console Software | Console4 | YES | Closed | WIP14SegReq | Ovation |  | WIP14 | February 22, 2019 5:09:34 PM GMT+08:00
CTUDT00601408 | Refresh | Closed | patra_r | January 24, 2019 2:44:02 PM GMT+08:00 | BB Refresh: ECG and Pulmo signal is not detected in ECG viewer and Resp Viewer respectively. | V4.1.10 and V4.7.7 | WIP14 |  | Subsystem |  | None | Fixed | Prakash_Chand_Yadav | 0 | 2 | Console Software | Console4 | YES | Closed | WIP14SegReq | Brilliance 64 |  | WIP14 | February 22, 2019 5:09:34 PM GMT+08:00
CTUDT00601408 | Refresh | Closed | patra_r | January 24, 2019 2:44:02 PM GMT+08:00 | BB Refresh: ECG and Pulmo signal is not detected in ECG viewer and Resp Viewer respectively. | V4.1.10 and V4.7.7 | WIP14 |  | Subsystem |  | None | Fixed | Prakash_Chand_Yadav | 0 | 2 | Console Software | Console4 | YES | Closed | WIP14SegReq | Ingenuity 64 |  | WIP14 | February 22, 2019 5:09:34 PM GMT+08:00
CTUDT00601408 | Refresh | Closed | patra_r | January 24, 2019 2:44:02 PM GMT+08:00 | BB Refresh: ECG and Pulmo signal is not detected in ECG viewer and Resp Viewer respectively. | V4.1.10 and V4.7.7 | WIP14 |  | Subsystem |  | None | Fixed | Prakash_Chand_Yadav | 0 | 2 | Console Software | Console4 | YES | Closed | WIP14SegReq | IQon1 |  | WIP14 | February 22, 2019 5:09:34 PM GMT+08:00
CTUDT00601406 | Refresh | Closed | zhong_x | January 24, 2019 1:24:51 PM GMT+08:00 | System Error Occurred after change thickness to 9. | BB_WIP3 | BB_WIP5 |  | Subsystem | Plan | N/A | Fixed | Prakash_Chand_Yadav | 0 | 2 | Console Software | Console4 | YES | Closed | BB_WIP5_SegReq | Brilliance Big Bore |  | BB_WIP5 | February 18, 2019 4:29:37 PM GMT+08:00
CTUDT00601404 | Refresh | Closed | shaham_t | January 23, 2019 7:17:01 PM GMT+08:00 | Image sets are grouped randomly for IMR Recon with Perfusion scans. | 4.7.7 | WIP22 BB_WIP12 | The workaround is not to reverse the direction instead keep scan and result directions to be the same. | Subsystem | Perfusion scans IMR | None | Fixed | kr_m | 0 | 3 | Console Software | Console4 | NO | Closed | WIP15SegReq | Brilliance Big Bore | TBF | WIP22 BB_WIP12 | November 6, 2019 8:46:12 PM GMT+08:00
CTUDT00601404 | Refresh | Closed | shaham_t | January 23, 2019 7:17:01 PM GMT+08:00 | Image sets are grouped randomly for IMR Recon with Perfusion scans. | 4.7.7 | WIP22 BB_WIP12 | The workaround is not to reverse the direction instead keep scan and result directions to be the same. | Subsystem | Perfusion scans IMR | None | Fixed | kr_m | 0 | 3 | Console Software | Console4 | NO | Closed | WIP15SegReq | Ingenuity 64 | TBF | WIP22 BB_WIP12 | November 6, 2019 8:46:12 PM GMT+08:00
CTUDT00601404 | Refresh | Closed | shaham_t | January 23, 2019 7:17:01 PM GMT+08:00 | Image sets are grouped randomly for IMR Recon with Perfusion scans. | 4.7.7 | WIP22 BB_WIP12 | The workaround is not to reverse the direction instead keep scan and result directions to be the same. | Subsystem | Perfusion scans IMR | None | Fixed | kr_m | 0 | 3 | Console Software | Console4 | NO | Closed | WIP15SegReq | IQon1 | TBF | WIP22 BB_WIP12 | November 6, 2019 8:46:12 PM GMT+08:00
CTUDT00601403 | Refresh | Closed | shaham_t | January 23, 2019 7:13:18 PM GMT+08:00 | Delay in Evolving Image display and Result Reconstruction start | 4.7.5 | WIP15 |  | Subsystem | Performance | None | Fixed | patra_r | 0 | 3 | Console Software | Console4 | YES | Closed | WIP15SubCom | IQon1 | CFR# 9472675 CFR# 9472676 CFR# 9474217 | WIP15 | March 20, 2019 8:26:17 PM GMT+08:00
CTUDT00601401 | Refresh | Closed | manjunath_mk | January 23, 2019 6:16:10 PM GMT+08:00 | Bugrep error when trying to access dose alert window with a disabled ct user | BB_WIP3 | BB_WIP5 | Valid user account can be used for Dose authentication. | Subsystem | User Authentication | N/A | Fixed | Prakash_Chand_Yadav | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP5_SegReq | Brilliance Big Bore |  | BB_WIP5 | March 5, 2019 6:51:16 PM GMT+08:00
CTUDT00601395 | Refresh | Closed | manjunath_mk | January 23, 2019 12:39:39 PM GMT+08:00 | 'UI has stopped working' pop up appears when "Enable Remote Access" tab is clicked in Tools menu | BB_WIP3 | BB_WIP5 | No Workaround | Subsystem | Remote Access | STPS-SRS-96 | Fixed | ranjithkumar_m | 0 | 2 | Console Software | Service Tools | NO | Closed | BB_WIP5_SegReq | Brilliance Big Bore |  | BB_WIP5 | February 19, 2019 2:16:51 PM GMT+08:00
CTUDT00601392 | Refresh | Closed | Valeria_Kashtanov | January 22, 2019 10:28:51 PM GMT+08:00 | Refresh: Calibrations: Error in attempt to perform only failed modes in calibration | WIP12 |  |  | Subsystem |  | n/a | Duplicate | frider_o | 0 |  |  | DMS | NO | Closed |  | Ingenuity 64
CTUDT00601387 | Refresh | Closed | Noy_Seelig | January 22, 2019 5:22:13 PM GMT+08:00 | Refresh: BB_WIP 3 installation fails during 1Click installation progress. | BB_WIP3 | BB_WIP5.1 |  | Subsystem | 1-Click | N/A | Fixed | venugopal_s | 0 | 2 | Service Tools_Software | DI | NO | Closed | BB_WIP5_SegReq | Brilliance Big Bore |  | BB_WIP5.1 | July 16, 2019 8:14:21 PM GMT+08:00
CTUDT00601382 | Refresh | Closed | Balaji_Subbaram | January 22, 2019 2:49:45 PM GMT+08:00 | [Refresh] [iQon]: Not all results get reconstructed when each result of different spectral type. | WIP12 |  |  | Subsystem | Recon |  | Works As Specified / Not A Defect | nanjunda_s | 0 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00601373 | Refresh | Closed | zhong_x | January 22, 2019 11:27:03 AM GMT+08:00 | After Artery was Drawn, “Save Results as…” button is still inactive in Functional CT Analysis. | BB_WIP3 | WIP14 |  | Subsystem | Functional CT;Verified1 | NA | Fixed | redrouthu_r | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP13SegReq | Ovation |  | WIP14 | March 14, 2019 5:29:44 PM GMT+08:00
CTUDT00601373 | Refresh | Closed | zhong_x | January 22, 2019 11:27:03 AM GMT+08:00 | After Artery was Drawn, “Save Results as…” button is still inactive in Functional CT Analysis. | BB_WIP3 | WIP14 |  | Subsystem | Functional CT;Verified1 | NA | Fixed | redrouthu_r | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP13SegReq | Brilliance 64 |  | WIP14 | March 14, 2019 5:29:44 PM GMT+08:00
CTUDT00601373 | Refresh | Closed | zhong_x | January 22, 2019 11:27:03 AM GMT+08:00 | After Artery was Drawn, “Save Results as…” button is still inactive in Functional CT Analysis. | BB_WIP3 | WIP14 |  | Subsystem | Functional CT;Verified1 | NA | Fixed | redrouthu_r | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP13SegReq | Brilliance Big Bore |  | WIP14 | March 14, 2019 5:29:44 PM GMT+08:00
CTUDT00601373 | Refresh | Closed | zhong_x | January 22, 2019 11:27:03 AM GMT+08:00 | After Artery was Drawn, “Save Results as…” button is still inactive in Functional CT Analysis. | BB_WIP3 | WIP14 |  | Subsystem | Functional CT;Verified1 | NA | Fixed | redrouthu_r | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP13SegReq | Ingenuity 64 |  | WIP14 | March 14, 2019 5:29:44 PM GMT+08:00
CTUDT00601373 | Refresh | Closed | zhong_x | January 22, 2019 11:27:03 AM GMT+08:00 | After Artery was Drawn, “Save Results as…” button is still inactive in Functional CT Analysis. | BB_WIP3 | WIP14 |  | Subsystem | Functional CT;Verified1 | NA | Fixed | redrouthu_r | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP13SegReq | IQon1 |  | WIP14 | March 14, 2019 5:29:44 PM GMT+08:00
CTUDT00601370 | Refresh | Closed | Jose_D_Sergio_Almeida | January 21, 2019 8:58:14 PM GMT+08:00 | When loading movies from Multimedia viewer patient directory crashed with Bug Rep error | WIP12 | WIP13 |  | System |  | NA | Fixed | dasgupta_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP13SegReq | Ovation |  | WIP13 | February 8, 2019 6:14:08 PM GMT+08:00
CTUDT00601368 | Refresh | Closed | attali_s | January 21, 2019 7:22:24 PM GMT+08:00 | Refresh: Common Functions: Wrong Bubble Detection for IQ applications | WIP11 | WIP17 |  | Subsystem | DMS Calibrations | NaN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP14SegReq | Ovation |  | WIP17 | May 28, 2019 7:53:12 PM GMT+08:00
CTUDT00601368 | Refresh | Closed | attali_s | January 21, 2019 7:22:24 PM GMT+08:00 | Refresh: Common Functions: Wrong Bubble Detection for IQ applications | WIP11 | WIP17 |  | Subsystem | DMS Calibrations | NaN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP14SegReq | Brilliance 64 |  | WIP17 | May 28, 2019 7:53:12 PM GMT+08:00
CTUDT00601368 | Refresh | Closed | attali_s | January 21, 2019 7:22:24 PM GMT+08:00 | Refresh: Common Functions: Wrong Bubble Detection for IQ applications | WIP11 | WIP17 |  | Subsystem | DMS Calibrations | NaN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP14SegReq | Brilliance Big Bore |  | WIP17 | May 28, 2019 7:53:12 PM GMT+08:00
CTUDT00601368 | Refresh | Closed | attali_s | January 21, 2019 7:22:24 PM GMT+08:00 | Refresh: Common Functions: Wrong Bubble Detection for IQ applications | WIP11 | WIP17 |  | Subsystem | DMS Calibrations | NaN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP14SegReq | Ingenuity 64 |  | WIP17 | May 28, 2019 7:53:12 PM GMT+08:00
CTUDT00601368 | Refresh | Closed | attali_s | January 21, 2019 7:22:24 PM GMT+08:00 | Refresh: Common Functions: Wrong Bubble Detection for IQ applications | WIP11 | WIP17 |  | Subsystem | DMS Calibrations | NaN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP14SegReq | IQon1 |  | WIP17 | May 28, 2019 7:53:12 PM GMT+08:00
CTUDT00601358 | Refresh | Closed | wang_l | January 21, 2019 9:43:52 AM GMT+08:00 | Service Tool crashed once click checkbox of items in Can Trace | WIP12 |  |  | Subsystem | CAN Trace |  | Duplicate | sanjeeva_a | 0 |  |  | Service Tools | NO | Closed |  | Ingenuity 64
CTUDT00601352 | Refresh | Closed | patra_r | January 18, 2019 9:21:55 PM GMT+08:00 | Refresh/UI display Issue :The value of the FOV is getting incorrectly displayed | V4.1.10 and V4.7.7 | WIP14 |  | Subsystem |  | NA | Fixed | v_sumanth | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SegReq | Ovation |  | WIP14 | February 20, 2019 4:31:14 PM GMT+08:00
CTUDT00601352 | Refresh | Closed | patra_r | January 18, 2019 9:21:55 PM GMT+08:00 | Refresh/UI display Issue :The value of the FOV is getting incorrectly displayed | V4.1.10 and V4.7.7 | WIP14 |  | Subsystem |  | NA | Fixed | v_sumanth | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SegReq | Brilliance 64 |  | WIP14 | February 20, 2019 4:31:14 PM GMT+08:00
CTUDT00601352 | Refresh | Closed | patra_r | January 18, 2019 9:21:55 PM GMT+08:00 | Refresh/UI display Issue :The value of the FOV is getting incorrectly displayed | V4.1.10 and V4.7.7 | WIP14 |  | Subsystem |  | NA | Fixed | v_sumanth | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SegReq | Ingenuity 64 |  | WIP14 | February 20, 2019 4:31:14 PM GMT+08:00
CTUDT00601352 | Refresh | Closed | patra_r | January 18, 2019 9:21:55 PM GMT+08:00 | Refresh/UI display Issue :The value of the FOV is getting incorrectly displayed | V4.1.10 and V4.7.7 | WIP14 |  | Subsystem |  | NA | Fixed | v_sumanth | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SegReq | IQon1 |  | WIP14 | February 20, 2019 4:31:14 PM GMT+08:00
CTUDT00601340 | Refresh | Closed | wei_f | January 18, 2019 1:53:26 PM GMT+08:00 | After copy the study contain chinese characters to PACS,switch the location between Local and PACS, pop up an error. | WIP12 | WIP13 |  | Subsystem | Chinese | NA | Fixed | s_sa | 0 | 2 | USP | iCAP | NO | Closed | WIP14SegReq | Ingenuity 64 |  | WIP13 | February 26, 2019 3:38:32 PM GMT+08:00
CTUDT00601336 | Refresh | Closed | PremKumar_Rajendran | January 18, 2019 1:20:59 PM GMT+08:00 | BB Refresh: Issues with ‘Edit Results in Slab Viewer/Edit Results on Surview’ and ‘Slab’ button on tools area for different kinds of exam cards. | BB_WIP2 |  |  | Subsystem | Slab Viewer |  | Not Enough Information | mistry_h | 0 |  |  |  |  | Closed |  | Brilliance Big Bore
CTUDT00601335 | Refresh | Closed | wei_f | January 18, 2019 1:18:26 PM GMT+08:00 | Institution Name is shown as ??? when Chinese language keyboard is used | WIP12 |  |  | Subsystem | Localization - Chinese characters | NA | No Plan to Fix | chokkalingam_r | 0 | 4 |  | Console4 | NO | Closed |  | Ovation
CTUDT00601335 | Refresh | Closed | wei_f | January 18, 2019 1:18:26 PM GMT+08:00 | Institution Name is shown as ??? when Chinese language keyboard is used | WIP12 |  |  | Subsystem | Localization - Chinese characters | NA | No Plan to Fix | chokkalingam_r | 0 | 4 |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00601335 | Refresh | Closed | wei_f | January 18, 2019 1:18:26 PM GMT+08:00 | Institution Name is shown as ??? when Chinese language keyboard is used | WIP12 |  |  | Subsystem | Localization - Chinese characters | NA | No Plan to Fix | chokkalingam_r | 0 | 4 |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00601335 | Refresh | Closed | wei_f | January 18, 2019 1:18:26 PM GMT+08:00 | Institution Name is shown as ??? when Chinese language keyboard is used | WIP12 |  |  | Subsystem | Localization - Chinese characters | NA | No Plan to Fix | chokkalingam_r | 0 | 4 |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00601335 | Refresh | Closed | wei_f | January 18, 2019 1:18:26 PM GMT+08:00 | Institution Name is shown as ??? when Chinese language keyboard is used | WIP12 |  |  | Subsystem | Localization - Chinese characters | NA | No Plan to Fix | chokkalingam_r | 0 | 4 |  | Console4 | NO | Closed |  | IQon1
CTUDT00601333 | Refresh | Closed | wang_l | January 18, 2019 1:15:19 PM GMT+08:00 | Refresh - Black Background occurred during launch system by service mode and then recovered quickly | BB_WIP3 |  | NA | Subsystem | Workflow | NA | Works As Specified / Not A Defect | ravindra_a | 0 |  |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00601333 | Refresh | Closed | wang_l | January 18, 2019 1:15:19 PM GMT+08:00 | Refresh - Black Background occurred during launch system by service mode and then recovered quickly | BB_WIP3 |  | NA | Subsystem | Workflow | NA | Works As Specified / Not A Defect | ravindra_a | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00601327 | Refresh | Closed | shiraguppi_k | January 18, 2019 12:40:43 PM GMT+08:00 | In Brain Perfusion application warning message is not translated to Norwegian language | WIP12 | WIP17 |  | Subsystem | Localization | NA | Fixed | redrouthu_r | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP16SegReq | Ovation |  | WIP17 | July 2, 2019 5:56:44 PM GMT+08:00
CTUDT00601327 | Refresh | Closed | shiraguppi_k | January 18, 2019 12:40:43 PM GMT+08:00 | In Brain Perfusion application warning message is not translated to Norwegian language | WIP12 | WIP17 |  | Subsystem | Localization | NA | Fixed | redrouthu_r | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP16SegReq | IQon1 |  | WIP17 | July 2, 2019 5:56:44 PM GMT+08:00
CTUDT00601326 | Refresh | Closed | shiraguppi_k | January 18, 2019 12:32:06 PM GMT+08:00 | In Brain Perfusion Application After creating ROI on right side first image of CBV perfusion image view port application crashes (error message: Brain Perfusion app has stopped working- close program) | WIP12 | WIP13 |  | Subsystem |  | NA | Fixed | aswani_v | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP13SegReq | Ovation |  | WIP13 | February 26, 2019 5:42:36 PM GMT+08:00
CTUDT00601326 | Refresh | Closed | shiraguppi_k | January 18, 2019 12:32:06 PM GMT+08:00 | In Brain Perfusion Application After creating ROI on right side first image of CBV perfusion image view port application crashes (error message: Brain Perfusion app has stopped working- close program) | WIP12 | WIP13 |  | Subsystem |  | NA | Fixed | aswani_v | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP13SegReq | IQon1 |  | WIP13 | February 26, 2019 5:42:36 PM GMT+08:00
CTUDT00601321 | Refresh | Closed | levy_s | January 17, 2019 10:43:43 PM GMT+08:00 | Refresh: Error appears infrequently while collecting BugRep of the previous day | WIP12 |  |  | Subsystem | BugRep | N/A | Works As Specified / Not A Defect | tamari_o | 1 | 2 |  | Support-System | NO | Closed | WIP16SubCom | Ovation
CTUDT00601321 | Refresh | Closed | levy_s | January 17, 2019 10:43:43 PM GMT+08:00 | Refresh: Error appears infrequently while collecting BugRep of the previous day | WIP12 |  |  | Subsystem | BugRep | N/A | Works As Specified / Not A Defect | tamari_o | 1 | 2 |  | Support-System | NO | Closed | WIP16SubCom | Brilliance 64
CTUDT00601321 | Refresh | Closed | levy_s | January 17, 2019 10:43:43 PM GMT+08:00 | Refresh: Error appears infrequently while collecting BugRep of the previous day | WIP12 |  |  | Subsystem | BugRep | N/A | Works As Specified / Not A Defect | tamari_o | 1 | 2 |  | Support-System | NO | Closed | WIP16SubCom | Brilliance Big Bore
CTUDT00601321 | Refresh | Closed | levy_s | January 17, 2019 10:43:43 PM GMT+08:00 | Refresh: Error appears infrequently while collecting BugRep of the previous day | WIP12 |  |  | Subsystem | BugRep | N/A | Works As Specified / Not A Defect | tamari_o | 1 | 2 |  | Support-System | NO | Closed | WIP16SubCom | Ingenuity 64
CTUDT00601321 | Refresh | Closed | levy_s | January 17, 2019 10:43:43 PM GMT+08:00 | Refresh: Error appears infrequently while collecting BugRep of the previous day | WIP12 |  |  | Subsystem | BugRep | N/A | Works As Specified / Not A Defect | tamari_o | 1 | 2 |  | Support-System | NO | Closed | WIP16SubCom | IQon1
CTUDT00601317 | Refresh | Closed | cohen_y | January 17, 2019 8:39:42 PM GMT+08:00 | Refresh: Tube Current: Collimation is not as defined in Spec. | WIP12 | WIP15 WIP14 |  | Subsystem | Tube Current | n/a | Fixed | frider_o | 0 | 2 | DMS Software | DMS | YES | Closed | WIP15SubCom | Ovation |  | WIP15 WIP14 | March 17, 2019 10:46:16 PM GMT+08:00
CTUDT00601317 | Refresh | Closed | cohen_y | January 17, 2019 8:39:42 PM GMT+08:00 | Refresh: Tube Current: Collimation is not as defined in Spec. | WIP12 | WIP15 WIP14 |  | Subsystem | Tube Current | n/a | Fixed | frider_o | 0 | 2 | DMS Software | DMS | YES | Closed | WIP15SubCom | Brilliance Big Bore |  | WIP15 WIP14 | March 17, 2019 10:46:16 PM GMT+08:00
CTUDT00601317 | Refresh | Closed | cohen_y | January 17, 2019 8:39:42 PM GMT+08:00 | Refresh: Tube Current: Collimation is not as defined in Spec. | WIP12 | WIP15 WIP14 |  | Subsystem | Tube Current | n/a | Fixed | frider_o | 0 | 2 | DMS Software | DMS | YES | Closed | WIP15SubCom | Ingenuity 64 |  | WIP15 WIP14 | March 17, 2019 10:46:16 PM GMT+08:00
CTUDT00601309 | Refresh | Closed | nayak_r | January 17, 2019 2:30:25 PM GMT+08:00 | Refresh\BB Refresh:User was not able to access the display settings and set the dual monitor changes | BB_WIP2 |  | Philips_service user can use NVIDIA control panel-> set up multiple displays to configure display settings. | Subsystem | Monitor Setting | NA | Duplicate | kapadia_v | 0 | 3 |  | Support-Security | NO | Closed | WIP14SegReq | Ovation
CTUDT00601309 | Refresh | Closed | nayak_r | January 17, 2019 2:30:25 PM GMT+08:00 | Refresh\BB Refresh:User was not able to access the display settings and set the dual monitor changes | BB_WIP2 |  | Philips_service user can use NVIDIA control panel-> set up multiple displays to configure display settings. | Subsystem | Monitor Setting | NA | Duplicate | kapadia_v | 0 | 3 |  | Support-Security | NO | Closed | WIP14SegReq | Brilliance 64
CTUDT00601309 | Refresh | Closed | nayak_r | January 17, 2019 2:30:25 PM GMT+08:00 | Refresh\BB Refresh:User was not able to access the display settings and set the dual monitor changes | BB_WIP2 |  | Philips_service user can use NVIDIA control panel-> set up multiple displays to configure display settings. | Subsystem | Monitor Setting | NA | Duplicate | kapadia_v | 0 | 3 |  | Support-Security | NO | Closed | WIP14SegReq | Brilliance Big Bore
CTUDT00601309 | Refresh | Closed | nayak_r | January 17, 2019 2:30:25 PM GMT+08:00 | Refresh\BB Refresh:User was not able to access the display settings and set the dual monitor changes | BB_WIP2 |  | Philips_service user can use NVIDIA control panel-> set up multiple displays to configure display settings. | Subsystem | Monitor Setting | NA | Duplicate | kapadia_v | 0 | 3 |  | Support-Security | NO | Closed | WIP14SegReq | Ingenuity 64
CTUDT00601309 | Refresh | Closed | nayak_r | January 17, 2019 2:30:25 PM GMT+08:00 | Refresh\BB Refresh:User was not able to access the display settings and set the dual monitor changes | BB_WIP2 |  | Philips_service user can use NVIDIA control panel-> set up multiple displays to configure display settings. | Subsystem | Monitor Setting | NA | Duplicate | kapadia_v | 0 | 3 |  | Support-Security | NO | Closed | WIP14SegReq | IQon1
CTUDT00601305 | Refresh | Closed | Dana_Kamyshev | January 16, 2019 8:43:21 PM GMT+08:00 | Refresh - Unable to reconstruct conventional result with saved SBI | WIP12 |  |  | Subsystem | Recon | NA | Duplicate | richards_b | 0 |  |  | CIRS | YES | Closed |  | Ovation
CTUDT00601305 | Refresh | Closed | Dana_Kamyshev | January 16, 2019 8:43:21 PM GMT+08:00 | Refresh - Unable to reconstruct conventional result with saved SBI | WIP12 |  |  | Subsystem | Recon | NA | Duplicate | richards_b | 0 |  |  | CIRS | YES | Closed |  | IQon1
CTUDT00601304 | Refresh | Closed | roshan_r | January 16, 2019 8:40:36 PM GMT+08:00 | BB Refresh : The Console does not disable and un-check "Review before save" option for any result selected for “merge series” | BB_WIP2 | BB_WIP5 | Results can be generated offline if original results are deleted. | Subsystem | Merge Series | None | Fixed | v_sumanth | 0 | 3 | Console Software | Console4 | YES | Closed | BB_WIP5_SegReq | Brilliance Big Bore |  | BB_WIP5 | March 7, 2019 3:35:11 PM GMT+08:00
CTUDT00601287 | Refresh | Closed | Valeria_Kashtanov | January 16, 2019 5:37:53 PM GMT+08:00 | Refresh: Air Calibration in CT user fails due to error after new host install | WIP12 | DHF170383 Rev 13 |  | Subsystem | Air calibration | n/a | Fixed | richards_b | 0 | 3 | Design Documentation | Support-Physics | NO | Closed | WIP14SegReq | Ovation |  | DHF170383 Rev 13 | April 27, 2019 4:07:53 AM GMT+08:00
CTUDT00601287 | Refresh | Closed | Valeria_Kashtanov | January 16, 2019 5:37:53 PM GMT+08:00 | Refresh: Air Calibration in CT user fails due to error after new host install | WIP12 | DHF170383 Rev 13 |  | Subsystem | Air calibration | n/a | Fixed | richards_b | 0 | 3 | Design Documentation | Support-Physics | NO | Closed | WIP14SegReq | Brilliance 64 |  | DHF170383 Rev 13 | April 27, 2019 4:07:53 AM GMT+08:00
CTUDT00601287 | Refresh | Closed | Valeria_Kashtanov | January 16, 2019 5:37:53 PM GMT+08:00 | Refresh: Air Calibration in CT user fails due to error after new host install | WIP12 | DHF170383 Rev 13 |  | Subsystem | Air calibration | n/a | Fixed | richards_b | 0 | 3 | Design Documentation | Support-Physics | NO | Closed | WIP14SegReq | Brilliance Big Bore |  | DHF170383 Rev 13 | April 27, 2019 4:07:53 AM GMT+08:00
CTUDT00601287 | Refresh | Closed | Valeria_Kashtanov | January 16, 2019 5:37:53 PM GMT+08:00 | Refresh: Air Calibration in CT user fails due to error after new host install | WIP12 | DHF170383 Rev 13 |  | Subsystem | Air calibration | n/a | Fixed | richards_b | 0 | 3 | Design Documentation | Support-Physics | NO | Closed | WIP14SegReq | Ingenuity 64 |  | DHF170383 Rev 13 | April 27, 2019 4:07:53 AM GMT+08:00
CTUDT00601287 | Refresh | Closed | Valeria_Kashtanov | January 16, 2019 5:37:53 PM GMT+08:00 | Refresh: Air Calibration in CT user fails due to error after new host install | WIP12 | DHF170383 Rev 13 |  | Subsystem | Air calibration | n/a | Fixed | richards_b | 0 | 3 | Design Documentation | Support-Physics | NO | Closed | WIP14SegReq | IQon1 |  | DHF170383 Rev 13 | April 27, 2019 4:07:53 AM GMT+08:00
CTUDT00601286 | Refresh | Closed | manjunath_mk | January 16, 2019 5:34:13 PM GMT+08:00 | The timestamp of the middle image in view tab does not match with the timestamp of slab viewer reference port when there are even number of images in the result | BB_WIP2 | DHF330449 Rev06 Big Bore HOST Verification Test Protocol |  | Subsystem | Slab Viewer | None | Fixed | vinoliya_j | 0 | 4 | Test Design | Console4 | NO | Closed | BB_WIP5_SegReq | Brilliance Big Bore |  | DHF330449 Rev06 Big Bore HOST Verification Test Protocol | May 2, 2019 3:22:29 PM GMT+08:00
CTUDT00601282 | Refresh | Closed | kadish_n | January 16, 2019 3:49:01 PM GMT+08:00 | Refresh: Service Tools: RCOM Components' versions do not appear in System Configuration. | WIP12 |  |  | Subsystem |  | N/A | Duplicate | Mark_Huang | 0 |  |  | GCX | NO | Closed |  | Ingenuity 64
CTUDT00601279 | Refresh | Closed | gT_s | January 16, 2019 3:07:03 PM GMT+08:00 | [Refresh]: iDose Recon mode does exist in the drop down for the CCT fluoro EC | WIP12 | DHF318051 Rev 11 |  | Subsystem | Test Design | N/A | Fixed | shekhar_p | 0 | 2 | Test Design | Console4 | NO | Closed | WIP14SegReq | IQon1 |  | DHF318051 Rev 11 | March 22, 2019 7:53:12 PM GMT+08:00
CTUDT00601276 | Refresh | Closed | ramesh_t | January 16, 2019 2:24:52 PM GMT+08:00 | Refresh:For CCT Fluoro scan,Unit of Planned DLP in show all is mGy*cm/s where as during scan it is mGy*cm in acquisition scene. it should be consistent as per the requirement Console-SRS.15974 | WIP12 | WIP14 |  | Subsystem | Show All | NA | Fixed | prasad_r | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SegReq | Ovation |  | WIP14 | February 14, 2019 2:30:11 PM GMT+08:00
CTUDT00601276 | Refresh | Closed | ramesh_t | January 16, 2019 2:24:52 PM GMT+08:00 | Refresh:For CCT Fluoro scan,Unit of Planned DLP in show all is mGy*cm/s where as during scan it is mGy*cm in acquisition scene. it should be consistent as per the requirement Console-SRS.15974 | WIP12 | WIP14 |  | Subsystem | Show All | NA | Fixed | prasad_r | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SegReq | Brilliance 64 |  | WIP14 | February 14, 2019 2:30:11 PM GMT+08:00
CTUDT00601276 | Refresh | Closed | ramesh_t | January 16, 2019 2:24:52 PM GMT+08:00 | Refresh:For CCT Fluoro scan,Unit of Planned DLP in show all is mGy*cm/s where as during scan it is mGy*cm in acquisition scene. it should be consistent as per the requirement Console-SRS.15974 | WIP12 | WIP14 |  | Subsystem | Show All | NA | Fixed | prasad_r | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SegReq | Ingenuity 64 |  | WIP14 | February 14, 2019 2:30:11 PM GMT+08:00
CTUDT00601276 | Refresh | Closed | ramesh_t | January 16, 2019 2:24:52 PM GMT+08:00 | Refresh:For CCT Fluoro scan,Unit of Planned DLP in show all is mGy*cm/s where as during scan it is mGy*cm in acquisition scene. it should be consistent as per the requirement Console-SRS.15974 | WIP12 | WIP14 |  | Subsystem | Show All | NA | Fixed | prasad_r | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SegReq | IQon1 |  | WIP14 | February 14, 2019 2:30:11 PM GMT+08:00
CTUDT00601274 | Refresh | Closed | nimma_p | January 16, 2019 1:49:36 PM GMT+08:00 | Refresh>IQON >Data not transferred after Unblocking the remote device by reconnecting the network cable | WIP12 | DHF318051 Rev 14 HOST Verification Test Protocol | Manually resend failed images | Subsystem | Test Design | NA | Fixed | ramesh_t | 0 | 3 | Test Design | Console4 | NO | Closed | WIP16SegReq | IQon1 |  | DHF318051 Rev 14 HOST Verification Test Protocol | May 2, 2019 8:46:04 PM GMT+08:00
CTUDT00601266 | Refresh | Closed | roshan_r | January 15, 2019 9:40:07 PM GMT+08:00 | BB Refresh : Console does not limit the number of MDUS sent to the CIRS in a get ready request to 30 | BB_WIP2 | BB_WIP5.1 | User can plan in such a way that more than 30 recons are not present. | Subsystem | Recon | None | Fixed | chokkalingam_r | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP5_SegReq | Brilliance Big Bore |  | BB_WIP5.1 | March 25, 2019 12:13:10 PM GMT+08:00
CTUDT00601263 | Refresh | Closed | Noy_Seelig | January 15, 2019 8:24:51 PM GMT+08:00 | Refresh: Exception error during Contrast in Air calibration | WIP12 | WIP15 |  | Subsystem | DMS calibrations | n/a | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP14SegReq | Ovation |  | WIP15 | March 17, 2019 11:04:38 PM GMT+08:00
CTUDT00601263 | Refresh | Closed | Noy_Seelig | January 15, 2019 8:24:51 PM GMT+08:00 | Refresh: Exception error during Contrast in Air calibration | WIP12 | WIP15 |  | Subsystem | DMS calibrations | n/a | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP14SegReq | Brilliance 64 |  | WIP15 | March 17, 2019 11:04:38 PM GMT+08:00
CTUDT00601263 | Refresh | Closed | Noy_Seelig | January 15, 2019 8:24:51 PM GMT+08:00 | Refresh: Exception error during Contrast in Air calibration | WIP12 | WIP15 |  | Subsystem | DMS calibrations | n/a | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP14SegReq | Ingenuity 64 |  | WIP15 | March 17, 2019 11:04:38 PM GMT+08:00
CTUDT00601263 | Refresh | Closed | Noy_Seelig | January 15, 2019 8:24:51 PM GMT+08:00 | Refresh: Exception error during Contrast in Air calibration | WIP12 | WIP15 |  | Subsystem | DMS calibrations | n/a | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP14SegReq | IQon1 |  | WIP15 | March 17, 2019 11:04:38 PM GMT+08:00
CTUDT00601261 | Refresh | Closed | mN_m | January 15, 2019 6:52:27 PM GMT+08:00 | BBRefresh: Bug rep error occurred when performing the last cycle of CCT scan followed by a system error. | BB_WIP3 | BB_WIP5.1 |  | Subsystem | CCT | NA | Fixed | pandey_a | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP5_SubCom | Brilliance Big Bore |  | BB_WIP5.1 | March 20, 2019 4:33:47 PM GMT+08:00
CTUDT00601259 | Refresh | Closed | mN_m | January 15, 2019 6:47:44 PM GMT+08:00 | BBRefresh: IV software controls on the scan ruler are disabled and not working. | BB_WIP3 | BB_WIP5 | NA | Subsystem |  | NA | Fixed | agrawal_ay | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP4_SubCom | Brilliance Big Bore |  | BB_WIP5 | February 28, 2019 2:14:28 PM GMT+08:00
CTUDT00601254 | Refresh | Closed | singh_p | January 15, 2019 5:34:45 PM GMT+08:00 | STPS Refresh WIP 12, PSC: History Information Tab was found empty when Gantry Subsystem was programmed under System Configuration tool. | WIP12 |  |  | Subsystem | Service Tools | STPS-SSRS-246 | Works As Specified / Not A Defect | Mark_Huang | 0 |  |  | GCX | NO | Closed |  | Ingenuity 64
CTUDT00601250 | Refresh | Closed | manjunath_mk | January 15, 2019 4:52:29 PM GMT+08:00 | Extend option for clinical scan is available even though the results are merged | BB_WIP2 | BB_WIP5.1 |  | Subsystem | Merge Series | None | Fixed | jose_s | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP5_SegReq | Brilliance Big Bore |  | BB_WIP5.1 | March 14, 2019 1:55:34 PM GMT+08:00
CTUDT00601249 | Refresh | Closed | singh_p | January 15, 2019 4:29:03 PM GMT+08:00 | STPS Refresh WIP 12, PSC: Error pop up thrown when RHOST was updated in Tube History Tools. | WIP12 |  |  | Subsystem | Tube History tool |  | Duplicate | karkour | 0 |  |  | GCX | NO | Closed |  | Ingenuity 64
CTUDT00601246 | Refresh | Closed | manjunath_c | January 15, 2019 3:46:14 PM GMT+08:00 | [Refresh] Ingenuity: Deviation in the default parameter values displayed in Exam card manager compared to what is mentioned in the "DHF251235 - Rev.01 Ingenuity Reference Exam Cards" Specification | WIP11.1 | WIP15 | The user can change the Exam card parameters and save the changes | Subsystem | Exam Card Manager | N/A | Fixed | dunitza_m | 0 | 2 | Design Documentation | Console4 | NO | Closed | SubCom | Ingenuity 64 |  | WIP15 | November 12, 2019 2:13:34 AM GMT+08:00
CTUDT00601244 | Refresh | Closed | manjunath_c | January 15, 2019 3:40:39 PM GMT+08:00 | [Refresh] iCTSP: Deviation in the default parameter values displayed in Exam card manager compared to what is mentioned in the "DHF316206 Rev.01 iCTSP Reference Exam Cards" Specification | WIP11.1 | WIP15 | The user can change the Exam card parameters and save the changes | Subsystem | Exam Card Manager | N/A | Fixed | dunitza_m | 0 | 2 | Design Documentation | Support-System | NO | Closed | SubCom | Ovation |  | WIP15 | May 19, 2019 9:29:13 PM GMT+08:00
CTUDT00601242 | Refresh | Closed | manjunath_c | January 15, 2019 3:35:04 PM GMT+08:00 | [Refresh] iCT: Deviation in the default parameter values displayed in Exam card manager compared to what is mentioned in the "DHF316204 - Rev.01 iCT Reference Exam Cards" Specification | WIP11.1 | WIP15 | User can create user ECs | Subsystem | Exam Card Manager | N/A | Fixed | dunitza_m | 0 | 2 | Design Documentation | Support-System | NO | Closed | SubCom | Ovation |  | WIP15 | May 19, 2019 9:37:25 PM GMT+08:00
CTUDT00601240 | Refresh | Closed | manjunath_c | January 15, 2019 3:30:26 PM GMT+08:00 | Deviation in the default parameter values displayed in Exam card manager compared to what is mentioned in the "DHF251245 - Rev01 Br64 Reference Exam Cards" Specification | WIP11.1 | WIP15 | user can create user ECs | Subsystem | Exam Card Manager | N/A | Fixed | dunitza_m | 0 | 2 | Design Documentation | Console4 | NO | Closed | SubCom | Brilliance 64 |  | WIP15 | May 21, 2019 8:50:30 PM GMT+08:00
CTUDT00601235 | Refresh | Closed | manjunath_mk | January 15, 2019 12:35:07 PM GMT+08:00 | Bugrep error when performing a clinical scan having both coronal and axial result | BB_WIP2 | BB_WIP4 | If user plans all the results manually before clicking go, the system does gives the error. | System |  | None | Fixed | patra_r | 0 | 3 | Console Software | Console4 | YES | Closed | BB_WIP4_SubCom | Brilliance Big Bore |  | BB_WIP4 | February 5, 2019 6:56:32 PM GMT+08:00
CTUDT00601225 | Refresh | Closed | Noy_Seelig | January 15, 2019 12:23:37 AM GMT+08:00 | Refresh: Beam Quality, Tube Potential Test, Deviation from Nominal appears in HTML report | WIP12 | DHF318940 Rev 12 |  | Subsystem |  | n/a | Fixed | yosfan_s | 0 | 2 | Test Design | DMS | NO | Closed | WIP14SegReq | Ovation |  | DHF318940 Rev 12 | February 24, 2019 8:40:56 PM GMT+08:00
CTUDT00601225 | Refresh | Closed | Noy_Seelig | January 15, 2019 12:23:37 AM GMT+08:00 | Refresh: Beam Quality, Tube Potential Test, Deviation from Nominal appears in HTML report | WIP12 | DHF318940 Rev 12 |  | Subsystem |  | n/a | Fixed | yosfan_s | 0 | 2 | Test Design | DMS | NO | Closed | WIP14SegReq | Brilliance Big Bore |  | DHF318940 Rev 12 | February 24, 2019 8:40:56 PM GMT+08:00
CTUDT00601225 | Refresh | Closed | Noy_Seelig | January 15, 2019 12:23:37 AM GMT+08:00 | Refresh: Beam Quality, Tube Potential Test, Deviation from Nominal appears in HTML report | WIP12 | DHF318940 Rev 12 |  | Subsystem |  | n/a | Fixed | yosfan_s | 0 | 2 | Test Design | DMS | NO | Closed | WIP14SegReq | Ingenuity 64 |  | DHF318940 Rev 12 | February 24, 2019 8:40:56 PM GMT+08:00
CTUDT00601225 | Refresh | Closed | Noy_Seelig | January 15, 2019 12:23:37 AM GMT+08:00 | Refresh: Beam Quality, Tube Potential Test, Deviation from Nominal appears in HTML report | WIP12 | DHF318940 Rev 12 |  | Subsystem |  | n/a | Fixed | yosfan_s | 0 | 2 | Test Design | DMS | NO | Closed | WIP14SegReq | IQon1 |  | DHF318940 Rev 12 | February 24, 2019 8:40:56 PM GMT+08:00
CTUDT00601222 | Refresh | Closed | manju_kolkar | January 14, 2019 6:56:19 PM GMT+08:00 | Difference in brightness was seen in both left and right monitors even though monitors are same and same configuration/settings been used. | BB_WIP2 |  | NA | Subsystem | Monitor Settings | NA | Invalid/Uncontrolled Configuration | manju_kolkar | 0 |  |  | Support-System | NO | Closed |  | Ovation
CTUDT00601222 | Refresh | Closed | manju_kolkar | January 14, 2019 6:56:19 PM GMT+08:00 | Difference in brightness was seen in both left and right monitors even though monitors are same and same configuration/settings been used. | BB_WIP2 |  | NA | Subsystem | Monitor Settings | NA | Invalid/Uncontrolled Configuration | manju_kolkar | 0 |  |  | Support-System | NO | Closed |  | Brilliance Big Bore
CTUDT00601222 | Refresh | Closed | manju_kolkar | January 14, 2019 6:56:19 PM GMT+08:00 | Difference in brightness was seen in both left and right monitors even though monitors are same and same configuration/settings been used. | BB_WIP2 |  | NA | Subsystem | Monitor Settings | NA | Invalid/Uncontrolled Configuration | manju_kolkar | 0 |  |  | Support-System | NO | Closed |  | Ingenuity 64
CTUDT00601222 | Refresh | Closed | manju_kolkar | January 14, 2019 6:56:19 PM GMT+08:00 | Difference in brightness was seen in both left and right monitors even though monitors are same and same configuration/settings been used. | BB_WIP2 |  | NA | Subsystem | Monitor Settings | NA | Invalid/Uncontrolled Configuration | manju_kolkar | 0 |  |  | Support-System | NO | Closed |  | IQon1
CTUDT00601218 | Refresh | Closed | ravindra_a | January 14, 2019 4:22:53 PM GMT+08:00 | Refresh WIP 11 : ADLDS is still configured with Refresh. | 4.8 | BB_WIP4 | No | Subsystem |  | NA | Fixed | ravindra_a | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP4_SubCom | Brilliance Big Bore |  | BB_WIP4 | January 30, 2019 5:08:48 PM GMT+08:00
CTUDT00601214 | Refresh | Closed | zhang_t | January 14, 2019 2:55:49 PM GMT+08:00 | Refresh BB: IFU needs to cover both the old and new panel | BB_WIP2 | WIP16 |  | Subsystem | IFU | N/A | Fixed | rajalakshmi_j | 0 | 3 | User Documentation | Support-System | NO | Closed | BB_WIP5_SegReq | Brilliance Big Bore |  | WIP16 | May 29, 2019 3:32:54 PM GMT+08:00
CTUDT00601207 | Refresh | Closed | Noy_Seelig | January 13, 2019 10:09:47 PM GMT+08:00 | Refresh: Values in Hcor Calibration do not match to the spec | WIP12 | WIP17 | N/A | Subsystem | Hcor calibration | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP17SubCom | Ovation |  | WIP17 | May 28, 2019 3:35:02 PM GMT+08:00
CTUDT00601207 | Refresh | Closed | Noy_Seelig | January 13, 2019 10:09:47 PM GMT+08:00 | Refresh: Values in Hcor Calibration do not match to the spec | WIP12 | WIP17 | N/A | Subsystem | Hcor calibration | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP17SubCom | Brilliance Big Bore |  | WIP17 | May 28, 2019 3:35:02 PM GMT+08:00
CTUDT00601207 | Refresh | Closed | Noy_Seelig | January 13, 2019 10:09:47 PM GMT+08:00 | Refresh: Values in Hcor Calibration do not match to the spec | WIP12 | WIP17 | N/A | Subsystem | Hcor calibration | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP17SubCom | Ingenuity 64 |  | WIP17 | May 28, 2019 3:35:02 PM GMT+08:00
CTUDT00601206 | Refresh | Closed | finestine_m | January 13, 2019 10:03:01 PM GMT+08:00 | Refresh: Calibrations: Hcor Calibration: Hcor Slice Infant cancelled due to high noise level | WIP11 |  |  | Subsystem | DMS Calibrations | n/a | Works As Specified / Not A Defect | gehtin_e | 0 |  |  | DMS | NO | Closed |  | iCT2
CTUDT00601201 | Refresh | Closed | nayak_r | January 11, 2019 7:54:01 PM GMT+08:00 | BB Refresh/UI display Issue :The value of the FOV is getting incorrectly displayed | BB_WIP2 | BB_WIP5 |  | Subsystem |  | None | Fixed | v_sumanth | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SegReq | Brilliance Big Bore |  | BB_WIP5 | February 25, 2019 2:26:39 PM GMT+08:00
CTUDT00601199 | Refresh | Closed | nayak_r | January 11, 2019 7:54:00 PM GMT+08:00 | BB Refresh: Launching of few user created ECs to plan gives bug rep error | BB_WIP2 | BB_WIP4 |  | Subsystem |  | None | Fixed | patra_r | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP4_SubCom | Brilliance Big Bore |  | BB_WIP4 | January 24, 2019 8:47:45 PM GMT+08:00
CTUDT00601187 | Refresh | Closed | nayak_r | January 10, 2019 8:47:53 PM GMT+08:00 | Refresh\BB Refresh: Injection symbols appearing in wrong positions or missing for the scans in the series list and scan ruler for few ECs loaded in Plan | BB_WIP2 | DHF330449 Rev06 Big Bore HOST Verification Test Protocol |  | Subsystem | Planning | N/A | Fixed | vinoliya_j | 0 | 3 | Test Design | Console4 | NO | Closed | WIP14SegReq | Ovation |  | DHF330449 Rev06 Big Bore HOST Verification Test Protocol | May 2, 2019 3:22:03 PM GMT+08:00
CTUDT00601187 | Refresh | Closed | nayak_r | January 10, 2019 8:47:53 PM GMT+08:00 | Refresh\BB Refresh: Injection symbols appearing in wrong positions or missing for the scans in the series list and scan ruler for few ECs loaded in Plan | BB_WIP2 | DHF330449 Rev06 Big Bore HOST Verification Test Protocol |  | Subsystem | Planning | N/A | Fixed | vinoliya_j | 0 | 3 | Test Design | Console4 | NO | Closed | WIP14SegReq | Brilliance 64 |  | DHF330449 Rev06 Big Bore HOST Verification Test Protocol | May 2, 2019 3:22:03 PM GMT+08:00
CTUDT00601187 | Refresh | Closed | nayak_r | January 10, 2019 8:47:53 PM GMT+08:00 | Refresh\BB Refresh: Injection symbols appearing in wrong positions or missing for the scans in the series list and scan ruler for few ECs loaded in Plan | BB_WIP2 | DHF330449 Rev06 Big Bore HOST Verification Test Protocol |  | Subsystem | Planning | N/A | Fixed | vinoliya_j | 0 | 3 | Test Design | Console4 | NO | Closed | WIP14SegReq | Brilliance Big Bore |  | DHF330449 Rev06 Big Bore HOST Verification Test Protocol | May 2, 2019 3:22:03 PM GMT+08:00
CTUDT00601187 | Refresh | Closed | nayak_r | January 10, 2019 8:47:53 PM GMT+08:00 | Refresh\BB Refresh: Injection symbols appearing in wrong positions or missing for the scans in the series list and scan ruler for few ECs loaded in Plan | BB_WIP2 | DHF330449 Rev06 Big Bore HOST Verification Test Protocol |  | Subsystem | Planning | N/A | Fixed | vinoliya_j | 0 | 3 | Test Design | Console4 | NO | Closed | WIP14SegReq | Ingenuity 64 |  | DHF330449 Rev06 Big Bore HOST Verification Test Protocol | May 2, 2019 3:22:03 PM GMT+08:00
CTUDT00601187 | Refresh | Closed | nayak_r | January 10, 2019 8:47:53 PM GMT+08:00 | Refresh\BB Refresh: Injection symbols appearing in wrong positions or missing for the scans in the series list and scan ruler for few ECs loaded in Plan | BB_WIP2 | DHF330449 Rev06 Big Bore HOST Verification Test Protocol |  | Subsystem | Planning | N/A | Fixed | vinoliya_j | 0 | 3 | Test Design | Console4 | NO | Closed | WIP14SegReq | IQon1 |  | DHF330449 Rev06 Big Bore HOST Verification Test Protocol | May 2, 2019 3:22:03 PM GMT+08:00
CTUDT00601186 | Refresh | Closed | nayak_r | January 10, 2019 8:37:41 PM GMT+08:00 | BB Refresh: In o-Level user login,the reset password tool should not display | BB_WIP2 |  |  | Subsystem |  |  | Duplicate | mistry_h | 0 |  |  | Service Tools |  | Closed |  | Brilliance Big Bore
CTUDT00601184 | Refresh | Closed | mN_m | January 10, 2019 8:24:15 PM GMT+08:00 | “Password Policy:” section in change password screen is misaligned. | BB_WIP2 |  |  | Subsystem | User Management |  | Change Proposal | sanjeeva_a | 0 |  |  | Console4 |  | Closed |  | Ovation
CTUDT00601184 | Refresh | Closed | mN_m | January 10, 2019 8:24:15 PM GMT+08:00 | “Password Policy:” section in change password screen is misaligned. | BB_WIP2 |  |  | Subsystem | User Management |  | Change Proposal | sanjeeva_a | 0 |  |  | Console4 |  | Closed |  | Brilliance Big Bore
CTUDT00601184 | Refresh | Closed | mN_m | January 10, 2019 8:24:15 PM GMT+08:00 | “Password Policy:” section in change password screen is misaligned. | BB_WIP2 |  |  | Subsystem | User Management |  | Change Proposal | sanjeeva_a | 0 |  |  | Console4 |  | Closed |  | Ingenuity 64
CTUDT00601184 | Refresh | Closed | mN_m | January 10, 2019 8:24:15 PM GMT+08:00 | “Password Policy:” section in change password screen is misaligned. | BB_WIP2 |  |  | Subsystem | User Management |  | Change Proposal | sanjeeva_a | 0 |  |  | Console4 |  | Closed |  | IQon1
CTUDT00601181 | Refresh | Closed | mN_m | January 10, 2019 8:17:31 PM GMT+08:00 | BBRefresh: In ECM, for sagittal results, it is showing Oas Oblique when you mouse hover on the exam cards. | BB_WIP2 | BB_WIP4 |  | Subsystem |  | None | Fixed | patra_r | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP4_SubCom | Brilliance Big Bore |  | BB_WIP4 | January 24, 2019 8:42:03 PM GMT+08:00
CTUDT00601179 | Refresh | Closed | mN_m | January 10, 2019 8:10:23 PM GMT+08:00 | BBRefresh: “Reset Password” is not working for Philips service user and the option is available for O-level user. | BB_WIP2 | BB_WIP5 | No. | Subsystem | Reset Password | NA | Fixed | kapadia_v | 0 | 2 | Console Software | Service Tools | NO | Closed | BB_WIP5_SegReq | Brilliance Big Bore |  | BB_WIP5 | February 19, 2019 8:06:54 PM GMT+08:00
CTUDT00601178 | Refresh | Closed | PremKumar_Rajendran | January 10, 2019 7:53:36 PM GMT+08:00 | BBRefresh:Emergency_operator name was missing in DICOM & ‘Operator Name’ field. | BB_WIP2 | BB_WIP4 |  | Subsystem |  | N/A | Fixed | Prakash_Chand_Yadav | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP4_SubCom | Brilliance Big Bore |  | BB_WIP4 | February 6, 2019 2:17:14 PM GMT+08:00
CTUDT00601175 | Refresh | Closed | singh_p | January 10, 2019 6:49:28 PM GMT+08:00 | Refresh PSC: Pre-recorded sound was not heard at the Gantry Speaker during Audio test | WIP12 |  |  | Subsystem | Audio Test tool |  | Duplicate | shekhar_p | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00601171 | Refresh | Closed | Noy_Seelig | January 10, 2019 4:38:48 PM GMT+08:00 | Refresh: Acceptance Test can not create Baseline due to ROI not drawn correctly | WIP12 | WIP14 | N/A | Subsystem |  | N/A | Fixed | elviral | 0 | 2 | DMS Software | DMS | NO | Closed | WIP14SegReq | Ovation |  | WIP14 | February 7, 2019 4:49:39 PM GMT+08:00
CTUDT00601171 | Refresh | Closed | Noy_Seelig | January 10, 2019 4:38:48 PM GMT+08:00 | Refresh: Acceptance Test can not create Baseline due to ROI not drawn correctly | WIP12 | WIP14 | N/A | Subsystem |  | N/A | Fixed | elviral | 0 | 2 | DMS Software | DMS | NO | Closed | WIP14SegReq | Brilliance 64 |  | WIP14 | February 7, 2019 4:49:39 PM GMT+08:00
CTUDT00601171 | Refresh | Closed | Noy_Seelig | January 10, 2019 4:38:48 PM GMT+08:00 | Refresh: Acceptance Test can not create Baseline due to ROI not drawn correctly | WIP12 | WIP14 | N/A | Subsystem |  | N/A | Fixed | elviral | 0 | 2 | DMS Software | DMS | NO | Closed | WIP14SegReq | Brilliance Big Bore |  | WIP14 | February 7, 2019 4:49:39 PM GMT+08:00
CTUDT00601171 | Refresh | Closed | Noy_Seelig | January 10, 2019 4:38:48 PM GMT+08:00 | Refresh: Acceptance Test can not create Baseline due to ROI not drawn correctly | WIP12 | WIP14 | N/A | Subsystem |  | N/A | Fixed | elviral | 0 | 2 | DMS Software | DMS | NO | Closed | WIP14SegReq | Ingenuity 64 |  | WIP14 | February 7, 2019 4:49:39 PM GMT+08:00
CTUDT00601171 | Refresh | Closed | Noy_Seelig | January 10, 2019 4:38:48 PM GMT+08:00 | Refresh: Acceptance Test can not create Baseline due to ROI not drawn correctly | WIP12 | WIP14 | N/A | Subsystem |  | N/A | Fixed | elviral | 0 | 2 | DMS Software | DMS | NO | Closed | WIP14SegReq | IQon1 |  | WIP14 | February 7, 2019 4:49:39 PM GMT+08:00
CTUDT00601166 | Refresh | Closed | wang_l | January 10, 2019 3:40:00 PM GMT+08:00 | System hanged when load to Plan window | WIP12 |  |  | Subsystem | Performance |  | Duplicate | singh_j | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00601165 | Refresh | Closed | finestine_m | January 10, 2019 3:33:27 PM GMT+08:00 | Calibrations: Hcor Adult Calibration: An error appears while performing Hcor Slice Adult and Basic Adult with H-Phantom | WIP12 |  | N/A | Subsystem | DMS Calibration | N/A | Works As Specified / Not A Defect | yosfan_s | 1 |  |  | DMS | NO | Closed |  | Ingenuity 64
CTUDT00601159 | Refresh | Closed | Dana_Kamyshev | January 9, 2019 9:22:18 PM GMT+08:00 | Refresh - Failed to send Chinese study to EBW | WIP12 |  |  | Subsystem | Chinese characters | NA | Works As Specified / Not A Defect | s_sa | 0 |  |  | iCAP | NO | Closed |  | Ovation
CTUDT00601159 | Refresh | Closed | Dana_Kamyshev | January 9, 2019 9:22:18 PM GMT+08:00 | Refresh - Failed to send Chinese study to EBW | WIP12 |  |  | Subsystem | Chinese characters | NA | Works As Specified / Not A Defect | s_sa | 0 |  |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00601159 | Refresh | Closed | Dana_Kamyshev | January 9, 2019 9:22:18 PM GMT+08:00 | Refresh - Failed to send Chinese study to EBW | WIP12 |  |  | Subsystem | Chinese characters | NA | Works As Specified / Not A Defect | s_sa | 0 |  |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00601159 | Refresh | Closed | Dana_Kamyshev | January 9, 2019 9:22:18 PM GMT+08:00 | Refresh - Failed to send Chinese study to EBW | WIP12 |  |  | Subsystem | Chinese characters | NA | Works As Specified / Not A Defect | s_sa | 0 |  |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00601159 | Refresh | Closed | Dana_Kamyshev | January 9, 2019 9:22:18 PM GMT+08:00 | Refresh - Failed to send Chinese study to EBW | WIP12 |  |  | Subsystem | Chinese characters | NA | Works As Specified / Not A Defect | s_sa | 0 |  |  | iCAP | NO | Closed |  | IQon1
CTUDT00601156 | Refresh | Closed | gushansky_m | January 9, 2019 9:14:24 PM GMT+08:00 | Refresh: Impossible copy Partial raw data for olevel and clinical users | WIP12 |  |  | Subsystem | Copy Partial raw data | NA | Works As Specified / Not A Defect | goswami_s | 0 |  |  | Console4 | NO | Closed | WIP15SubCom | IQon1
CTUDT00601155 | Refresh | Closed | Dana_Kamyshev | January 9, 2019 9:12:55 PM GMT+08:00 | Refresh - Chinese characters are not displayed in PACS | WIP12 |  | Please use iWatt as PACS which supports Chinese characters to do the testing | Subsystem | Chinese characters | NA | Works As Specified / Not A Defect | s_sa | 0 |  |  | iCAP | NO | Closed |  | Ovation
CTUDT00601155 | Refresh | Closed | Dana_Kamyshev | January 9, 2019 9:12:55 PM GMT+08:00 | Refresh - Chinese characters are not displayed in PACS | WIP12 |  | Please use iWatt as PACS which supports Chinese characters to do the testing | Subsystem | Chinese characters | NA | Works As Specified / Not A Defect | s_sa | 0 |  |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00601155 | Refresh | Closed | Dana_Kamyshev | January 9, 2019 9:12:55 PM GMT+08:00 | Refresh - Chinese characters are not displayed in PACS | WIP12 |  | Please use iWatt as PACS which supports Chinese characters to do the testing | Subsystem | Chinese characters | NA | Works As Specified / Not A Defect | s_sa | 0 |  |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00601155 | Refresh | Closed | Dana_Kamyshev | January 9, 2019 9:12:55 PM GMT+08:00 | Refresh - Chinese characters are not displayed in PACS | WIP12 |  | Please use iWatt as PACS which supports Chinese characters to do the testing | Subsystem | Chinese characters | NA | Works As Specified / Not A Defect | s_sa | 0 |  |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00601155 | Refresh | Closed | Dana_Kamyshev | January 9, 2019 9:12:55 PM GMT+08:00 | Refresh - Chinese characters are not displayed in PACS | WIP12 |  | Please use iWatt as PACS which supports Chinese characters to do the testing | Subsystem | Chinese characters | NA | Works As Specified / Not A Defect | s_sa | 0 |  |  | iCAP | NO | Closed |  | IQon1
CTUDT00601153 | Refresh | Closed | or_shmuel | January 9, 2019 7:40:41 PM GMT+08:00 | Refresh, IQon - Cannot connect to CIRS by ServiceTools --> Remote Viewer | WIP12 |  |  | Subsystem | Remote Connection | NA | Does Not Apply | prasad_r | 0 |  |  | Console4 | NO | Closed | WIP15SegReq | Ovation
CTUDT00601153 | Refresh | Closed | or_shmuel | January 9, 2019 7:40:41 PM GMT+08:00 | Refresh, IQon - Cannot connect to CIRS by ServiceTools --> Remote Viewer | WIP12 |  |  | Subsystem | Remote Connection | NA | Does Not Apply | prasad_r | 0 |  |  | Console4 | NO | Closed | WIP15SegReq | Brilliance 64
CTUDT00601153 | Refresh | Closed | or_shmuel | January 9, 2019 7:40:41 PM GMT+08:00 | Refresh, IQon - Cannot connect to CIRS by ServiceTools --> Remote Viewer | WIP12 |  |  | Subsystem | Remote Connection | NA | Does Not Apply | prasad_r | 0 |  |  | Console4 | NO | Closed | WIP15SegReq | Brilliance Big Bore
CTUDT00601153 | Refresh | Closed | or_shmuel | January 9, 2019 7:40:41 PM GMT+08:00 | Refresh, IQon - Cannot connect to CIRS by ServiceTools --> Remote Viewer | WIP12 |  |  | Subsystem | Remote Connection | NA | Does Not Apply | prasad_r | 0 |  |  | Console4 | NO | Closed | WIP15SegReq | Ingenuity 64
CTUDT00601153 | Refresh | Closed | or_shmuel | January 9, 2019 7:40:41 PM GMT+08:00 | Refresh, IQon - Cannot connect to CIRS by ServiceTools --> Remote Viewer | WIP12 |  |  | Subsystem | Remote Connection | NA | Does Not Apply | prasad_r | 0 |  |  | Console4 | NO | Closed | WIP15SegReq | IQon1
CTUDT00601150 | Refresh | Closed | gushansky_m | January 9, 2019 6:38:33 PM GMT+08:00 | refresh: Update and Get data from RHost failed | WIP12 |  |  | Subsystem | Tube History Tool |  | Duplicate | karkour | 0 |  |  | GCX | NO | Closed |  | IQon1
CTUDT00601148 | Refresh | Closed | wang_l | January 9, 2019 5:41:06 PM GMT+08:00 | The 'Pause" icon in system does not match with white cover design | WIP12 | WIP18 BB_WIP8 |  | Subsystem | White cover | NA | Fixed | Sandeep_Kumar | 0 | 3 | Console Software | Service Tools | NO | Closed | WIP15SegReq | Ovation |  | WIP18 BB_WIP8 | June 7, 2019 4:55:50 PM GMT+08:00
CTUDT00601148 | Refresh | Closed | wang_l | January 9, 2019 5:41:06 PM GMT+08:00 | The 'Pause" icon in system does not match with white cover design | WIP12 | WIP18 BB_WIP8 |  | Subsystem | White cover | NA | Fixed | Sandeep_Kumar | 0 | 3 | Console Software | Service Tools | NO | Closed | WIP15SegReq | Brilliance 64 |  | WIP18 BB_WIP8 | June 7, 2019 4:55:50 PM GMT+08:00
CTUDT00601148 | Refresh | Closed | wang_l | January 9, 2019 5:41:06 PM GMT+08:00 | The 'Pause" icon in system does not match with white cover design | WIP12 | WIP18 BB_WIP8 |  | Subsystem | White cover | NA | Fixed | Sandeep_Kumar | 0 | 3 | Console Software | Service Tools | NO | Closed | WIP15SegReq | Ingenuity 64 |  | WIP18 BB_WIP8 | June 7, 2019 4:55:50 PM GMT+08:00
CTUDT00601148 | Refresh | Closed | wang_l | January 9, 2019 5:41:06 PM GMT+08:00 | The 'Pause" icon in system does not match with white cover design | WIP12 | WIP18 BB_WIP8 |  | Subsystem | White cover | NA | Fixed | Sandeep_Kumar | 0 | 3 | Console Software | Service Tools | NO | Closed | WIP15SegReq | IQon1 |  | WIP18 BB_WIP8 | June 7, 2019 4:55:50 PM GMT+08:00
CTUDT00601144 | Refresh | Closed | panchami_pradeep | January 9, 2019 11:44:44 AM GMT+08:00 | Refresh: PSC: "Configuration" tab is missing in service tools UI, when the screen resolution is 800*600. | WIP12 | HOST Verification Test Protocol DHF318051 Rev 12 |  | Subsystem | Test Design | N/A | Fixed | jindal_k | 0 | 3 | Test Design | Service Tools | NO | Closed | TestDesign | Ingenuity 64 |  | HOST Verification Test Protocol DHF318051 Rev 12 | April 25, 2019 12:59:37 PM GMT+08:00
CTUDT00601140 | Refresh | Closed | klimovitsky_n | January 8, 2019 9:34:22 PM GMT+08:00 | Refresh - Series do not appear under the study in studies tab | WIP12 |  |  | Subsystem | Reconstruction |  | Works As Specified / Not A Defect | tamari_o | 1 |  |  | Support-System | NO | Closed |  | Ovation
CTUDT00601140 | Refresh | Closed | klimovitsky_n | January 8, 2019 9:34:22 PM GMT+08:00 | Refresh - Series do not appear under the study in studies tab | WIP12 |  |  | Subsystem | Reconstruction |  | Works As Specified / Not A Defect | tamari_o | 1 |  |  | Support-System | NO | Closed |  | IQon1
CTUDT00601133 | Refresh | Closed | gT_s | January 8, 2019 6:51:49 PM GMT+08:00 | [Refresh]:The Recon Mode, Recon level and filter fields is not highlighted when the spectral results option is unchecked. | WIP 12 | WIP15 |  | Subsystem | Spectral, Planning | NA | Fixed | sagar_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP15SubCom | IQon1 |  | WIP15 | March 18, 2019 10:13:37 PM GMT+08:00
CTUDT00601131 | Refresh | Closed | gT_s | January 8, 2019 6:40:57 PM GMT+08:00 | [Refresh] :The Recon Mode field is not highlighted when the spectral results option is selected for a result. | WIP 12 | WIP15 |  | Subsystem | Spectral, ECM | NA | Fixed | sheik_c | 0 | 2 | Console Software | Console4 | NO | Closed | WIP15SubCom | IQon1 |  | WIP15 | March 28, 2019 7:11:51 PM GMT+08:00
CTUDT00601121 | Refresh | Closed | Valeria_Kashtanov | January 8, 2019 5:15:15 AM GMT+08:00 | Refresh: Calibrations: FEE Test: mAs value of SNR scan is not as defined in spec | WIP12 | WIP14 |  | Subsystem | FEE Test | n/a | Fixed | yosefk | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP13SegReq | Brilliance 64 |  | WIP14 | February 3, 2019 10:23:11 PM GMT+08:00
CTUDT00601121 | Refresh | Closed | Valeria_Kashtanov | January 8, 2019 5:15:15 AM GMT+08:00 | Refresh: Calibrations: FEE Test: mAs value of SNR scan is not as defined in spec | WIP12 | WIP14 |  | Subsystem | FEE Test | n/a | Fixed | yosefk | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP13SegReq | Ingenuity 64 |  | WIP14 | February 3, 2019 10:23:11 PM GMT+08:00
CTUDT00601118 | Refresh | Closed | Revital_Cohen | January 7, 2019 10:32:37 PM GMT+08:00 | Refresh WIP12 CANTrace is closing in Service tools | WIP12 | WIP14 |  | Subsystem | CAN trace | N/A | Fixed | shreekanth_manniyath | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP14SubCom | Brilliance 64 |  | WIP14 | February 6, 2019 7:18:03 PM GMT+08:00
CTUDT00601118 | Refresh | Closed | Revital_Cohen | January 7, 2019 10:32:37 PM GMT+08:00 | Refresh WIP12 CANTrace is closing in Service tools | WIP12 | WIP14 |  | Subsystem | CAN trace | N/A | Fixed | shreekanth_manniyath | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP14SubCom | Brilliance Big Bore |  | WIP14 | February 6, 2019 7:18:03 PM GMT+08:00
CTUDT00601118 | Refresh | Closed | Revital_Cohen | January 7, 2019 10:32:37 PM GMT+08:00 | Refresh WIP12 CANTrace is closing in Service tools | WIP12 | WIP14 |  | Subsystem | CAN trace | N/A | Fixed | shreekanth_manniyath | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP14SubCom | Ingenuity 64 |  | WIP14 | February 6, 2019 7:18:03 PM GMT+08:00
CTUDT00601118 | Refresh | Closed | Revital_Cohen | January 7, 2019 10:32:37 PM GMT+08:00 | Refresh WIP12 CANTrace is closing in Service tools | WIP12 | WIP14 |  | Subsystem | CAN trace | N/A | Fixed | shreekanth_manniyath | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP14SubCom | IQon1 |  | WIP14 | February 6, 2019 7:18:03 PM GMT+08:00
CTUDT00601116 | Refresh | Closed | attali_s | January 7, 2019 9:49:19 PM GMT+08:00 | Refresh: Common Functions: No graphical representation in results of IQ Check | WIP12 | DHF318940 |  | Subsystem | Test Design | N/A | Fixed | yosfan_s | 0 | 2 | Test Design | DMS | NO | Closed | WIP14SegReq | Ingenuity 64 |  | DHF318940 | February 27, 2019 4:12:53 PM GMT+08:00
CTUDT00601113 | Refresh | Closed | Noy_Seelig | January 7, 2019 7:40:25 PM GMT+08:00 | Refresh calibrations: FEE test, instead of bias there is signal intensity (Y axis) in Bias of reference detectors plot | WIP12 | WIP13 | N/A | Subsystem |  | N/A | Fixed | yosefk | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | Ovation |  | WIP13 | January 20, 2019 3:41:00 PM GMT+08:00
CTUDT00601113 | Refresh | Closed | Noy_Seelig | January 7, 2019 7:40:25 PM GMT+08:00 | Refresh calibrations: FEE test, instead of bias there is signal intensity (Y axis) in Bias of reference detectors plot | WIP12 | WIP13 | N/A | Subsystem |  | N/A | Fixed | yosefk | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | Brilliance Big Bore |  | WIP13 | January 20, 2019 3:41:00 PM GMT+08:00
CTUDT00601113 | Refresh | Closed | Noy_Seelig | January 7, 2019 7:40:25 PM GMT+08:00 | Refresh calibrations: FEE test, instead of bias there is signal intensity (Y axis) in Bias of reference detectors plot | WIP12 | WIP13 | N/A | Subsystem |  | N/A | Fixed | yosefk | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | Ingenuity 64 |  | WIP13 | January 20, 2019 3:41:00 PM GMT+08:00
CTUDT00601113 | Refresh | Closed | Noy_Seelig | January 7, 2019 7:40:25 PM GMT+08:00 | Refresh calibrations: FEE test, instead of bias there is signal intensity (Y axis) in Bias of reference detectors plot | WIP12 | WIP13 | N/A | Subsystem |  | N/A | Fixed | yosefk | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | IQon1 |  | WIP13 | January 20, 2019 3:41:00 PM GMT+08:00
CTUDT00601108 | Refresh | Closed | jotawar_d | January 7, 2019 5:12:56 PM GMT+08:00 | Discard, overwrite and Rename options are not displaying in Import dialogue window After few exam cards are discarded | BB_WIP1 | BB_WIP4 |  | Subsystem |  | N/A | Fixed | prasad_r | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP4_SubCom | Brilliance Big Bore |  | BB_WIP4 | January 31, 2019 3:11:40 PM GMT+08:00
CTUDT00601107 | Refresh | Closed | attali_s | January 7, 2019 5:07:00 PM GMT+08:00 | Refresh: Performance Test: There is no measurement in air in Body CT13c case | WIP12 | WIP14 | N/A | Subsystem | Calibrations | N/A | Fixed | elviral | 0 | 2 | DMS Software | DMS | YES | Closed | WIP14SegReq | Ingenuity 64 |  | WIP14 | February 18, 2019 3:31:33 PM GMT+08:00
CTUDT00601107 | Refresh | Closed | attali_s | January 7, 2019 5:07:00 PM GMT+08:00 | Refresh: Performance Test: There is no measurement in air in Body CT13c case | WIP12 | WIP14 | N/A | Subsystem | Calibrations | N/A | Fixed | elviral | 0 | 2 | DMS Software | DMS | YES | Closed | WIP14SegReq | IQon1 |  | WIP14 | February 18, 2019 3:31:33 PM GMT+08:00
CTUDT00601106 | Refresh | Closed | dadon_e | January 7, 2019 4:57:28 PM GMT+08:00 | Refresh, IQon WIP12: CIRS Serviceability tests failed | WIP12 |  |  | Subsystem | CIRS Serviceability tests |  | Duplicate | richards_b | 0 |  |  | CIRS | NO | Closed |  | IQon1
CTUDT00601104 | Refresh | Closed | or_shmuel | January 7, 2019 2:11:02 AM GMT+08:00 | Refresh Connection to CIRS takes approximately 1 minute | WIP12 |  | N/A | Subsystem | Performance | N/A | No Plan to Fix | prasad_r | 0 | 4 |  | Console4 | NO | Closed |  | Ovation
CTUDT00601104 | Refresh | Closed | or_shmuel | January 7, 2019 2:11:02 AM GMT+08:00 | Refresh Connection to CIRS takes approximately 1 minute | WIP12 |  | N/A | Subsystem | Performance | N/A | No Plan to Fix | prasad_r | 0 | 4 |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00601104 | Refresh | Closed | or_shmuel | January 7, 2019 2:11:02 AM GMT+08:00 | Refresh Connection to CIRS takes approximately 1 minute | WIP12 |  | N/A | Subsystem | Performance | N/A | No Plan to Fix | prasad_r | 0 | 4 |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00601104 | Refresh | Closed | or_shmuel | January 7, 2019 2:11:02 AM GMT+08:00 | Refresh Connection to CIRS takes approximately 1 minute | WIP12 |  | N/A | Subsystem | Performance | N/A | No Plan to Fix | prasad_r | 0 | 4 |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00601104 | Refresh | Closed | or_shmuel | January 7, 2019 2:11:02 AM GMT+08:00 | Refresh Connection to CIRS takes approximately 1 minute | WIP12 |  | N/A | Subsystem | Performance | N/A | No Plan to Fix | prasad_r | 0 | 4 |  | Console4 | NO | Closed |  | IQon1
CTUDT00601102 | Refresh | Closed | beck_d | January 6, 2019 11:14:45 PM GMT+08:00 | Refresh: Calibrations: Harmonized Phantom centering is not stable | WIP12 | WIP15 WIP17 |  | Subsystem | Calibrations; Verified1 | n/a | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP15SegReq | Ovation |  | WIP15 WIP17 | May 26, 2019 10:55:12 PM GMT+08:00
CTUDT00601102 | Refresh | Closed | beck_d | January 6, 2019 11:14:45 PM GMT+08:00 | Refresh: Calibrations: Harmonized Phantom centering is not stable | WIP12 | WIP15 WIP17 |  | Subsystem | Calibrations; Verified1 | n/a | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP15SegReq | Brilliance Big Bore |  | WIP15 WIP17 | May 26, 2019 10:55:12 PM GMT+08:00
CTUDT00601102 | Refresh | Closed | beck_d | January 6, 2019 11:14:45 PM GMT+08:00 | Refresh: Calibrations: Harmonized Phantom centering is not stable | WIP12 | WIP15 WIP17 |  | Subsystem | Calibrations; Verified1 | n/a | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP15SegReq | Ingenuity 64 |  | WIP15 WIP17 | May 26, 2019 10:55:12 PM GMT+08:00
CTUDT00601102 | Refresh | Closed | beck_d | January 6, 2019 11:14:45 PM GMT+08:00 | Refresh: Calibrations: Harmonized Phantom centering is not stable | WIP12 | WIP15 WIP17 |  | Subsystem | Calibrations; Verified1 | n/a | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP15SegReq | IQon1 |  | WIP15 WIP17 | May 26, 2019 10:55:12 PM GMT+08:00
CTUDT00601099 | Refresh | Closed | beck_d | January 6, 2019 7:15:40 PM GMT+08:00 | Refresh: Calibrations: Hcor Calibration: Vertical bed position is too high | WIP12 | WIP14 | YES - Select minimum one head stage, or perform the full calibration/test. | Subsystem | Calibrations | N/A | Fixed | yatziv_l | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP14SegReq | Ovation |  | WIP14 | May 19, 2019 2:11:22 PM GMT+08:00
CTUDT00601099 | Refresh | Closed | beck_d | January 6, 2019 7:15:40 PM GMT+08:00 | Refresh: Calibrations: Hcor Calibration: Vertical bed position is too high | WIP12 | WIP14 | YES - Select minimum one head stage, or perform the full calibration/test. | Subsystem | Calibrations | N/A | Fixed | yatziv_l | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP14SegReq | Brilliance 64 |  | WIP14 | May 19, 2019 2:11:22 PM GMT+08:00
CTUDT00601099 | Refresh | Closed | beck_d | January 6, 2019 7:15:40 PM GMT+08:00 | Refresh: Calibrations: Hcor Calibration: Vertical bed position is too high | WIP12 | WIP14 | YES - Select minimum one head stage, or perform the full calibration/test. | Subsystem | Calibrations | N/A | Fixed | yatziv_l | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP14SegReq | Brilliance Big Bore |  | WIP14 | May 19, 2019 2:11:22 PM GMT+08:00
CTUDT00601099 | Refresh | Closed | beck_d | January 6, 2019 7:15:40 PM GMT+08:00 | Refresh: Calibrations: Hcor Calibration: Vertical bed position is too high | WIP12 | WIP14 | YES - Select minimum one head stage, or perform the full calibration/test. | Subsystem | Calibrations | N/A | Fixed | yatziv_l | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP14SegReq | Ingenuity 64 |  | WIP14 | May 19, 2019 2:11:22 PM GMT+08:00
CTUDT00601099 | Refresh | Closed | beck_d | January 6, 2019 7:15:40 PM GMT+08:00 | Refresh: Calibrations: Hcor Calibration: Vertical bed position is too high | WIP12 | WIP14 | YES - Select minimum one head stage, or perform the full calibration/test. | Subsystem | Calibrations | N/A | Fixed | yatziv_l | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP14SegReq | IQon1 |  | WIP14 | May 19, 2019 2:11:22 PM GMT+08:00
CTUDT00601097 | Refresh | Closed | beck_d | January 6, 2019 7:02:07 PM GMT+08:00 | Refresh: Calibrations: IQ Check: Noise limits are not as defined in spec | WIP12 | WIP14 |  | Subsystem | Calibrations | DHF170394-Rev.08 | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | NO | Closed | BB_WIP4_SegReq | Brilliance Big Bore |  | WIP14 | February 20, 2019 10:16:21 PM GMT+08:00
CTUDT00601097 | Refresh | Closed | beck_d | January 6, 2019 7:02:07 PM GMT+08:00 | Refresh: Calibrations: IQ Check: Noise limits are not as defined in spec | WIP12 | WIP14 |  | Subsystem | Calibrations | DHF170394-Rev.08 | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | NO | Closed | BB_WIP4_SegReq | Ingenuity 64 |  | WIP14 | February 20, 2019 10:16:21 PM GMT+08:00
CTUDT00601094 | Refresh | Closed | beck_d | January 6, 2019 4:35:20 PM GMT+08:00 | Refresh: Calibrations: CTDI: Error appears during the test | WIP12 | WIP15 |  | Subsystem | Calibrations | NaN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP15SubCom | Ovation |  | WIP15 | March 25, 2019 5:53:45 PM GMT+08:00
CTUDT00601094 | Refresh | Closed | beck_d | January 6, 2019 4:35:20 PM GMT+08:00 | Refresh: Calibrations: CTDI: Error appears during the test | WIP12 | WIP15 |  | Subsystem | Calibrations | NaN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP15SubCom | Brilliance 64 |  | WIP15 | March 25, 2019 5:53:45 PM GMT+08:00
CTUDT00601094 | Refresh | Closed | beck_d | January 6, 2019 4:35:20 PM GMT+08:00 | Refresh: Calibrations: CTDI: Error appears during the test | WIP12 | WIP15 |  | Subsystem | Calibrations | NaN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP15SubCom | Brilliance Big Bore |  | WIP15 | March 25, 2019 5:53:45 PM GMT+08:00
CTUDT00601094 | Refresh | Closed | beck_d | January 6, 2019 4:35:20 PM GMT+08:00 | Refresh: Calibrations: CTDI: Error appears during the test | WIP12 | WIP15 |  | Subsystem | Calibrations | NaN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP15SubCom | Ingenuity 64 |  | WIP15 | March 25, 2019 5:53:45 PM GMT+08:00
CTUDT00601094 | Refresh | Closed | beck_d | January 6, 2019 4:35:20 PM GMT+08:00 | Refresh: Calibrations: CTDI: Error appears during the test | WIP12 | WIP15 |  | Subsystem | Calibrations | NaN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP15SubCom | IQon1 |  | WIP15 | March 25, 2019 5:53:45 PM GMT+08:00
CTUDT00601091 | Refresh | Closed | beck_d | January 6, 2019 2:56:48 PM GMT+08:00 | Refresh: Calibrations: CTDI: Xi is not detected | WIP12 | WIP14 | The user can use manual mode if the probe is not detected | Subsystem | Calibrations | NaN | Fixed | daibes_s | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP14SegReq | Ovation |  | WIP14 | February 21, 2019 8:30:40 PM GMT+08:00
CTUDT00601091 | Refresh | Closed | beck_d | January 6, 2019 2:56:48 PM GMT+08:00 | Refresh: Calibrations: CTDI: Xi is not detected | WIP12 | WIP14 | The user can use manual mode if the probe is not detected | Subsystem | Calibrations | NaN | Fixed | daibes_s | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP14SegReq | Brilliance 64 |  | WIP14 | February 21, 2019 8:30:40 PM GMT+08:00
CTUDT00601091 | Refresh | Closed | beck_d | January 6, 2019 2:56:48 PM GMT+08:00 | Refresh: Calibrations: CTDI: Xi is not detected | WIP12 | WIP14 | The user can use manual mode if the probe is not detected | Subsystem | Calibrations | NaN | Fixed | daibes_s | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP14SegReq | Brilliance Big Bore |  | WIP14 | February 21, 2019 8:30:40 PM GMT+08:00
CTUDT00601091 | Refresh | Closed | beck_d | January 6, 2019 2:56:48 PM GMT+08:00 | Refresh: Calibrations: CTDI: Xi is not detected | WIP12 | WIP14 | The user can use manual mode if the probe is not detected | Subsystem | Calibrations | NaN | Fixed | daibes_s | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP14SegReq | Ingenuity 64 |  | WIP14 | February 21, 2019 8:30:40 PM GMT+08:00
CTUDT00601091 | Refresh | Closed | beck_d | January 6, 2019 2:56:48 PM GMT+08:00 | Refresh: Calibrations: CTDI: Xi is not detected | WIP12 | WIP14 | The user can use manual mode if the probe is not detected | Subsystem | Calibrations | NaN | Fixed | daibes_s | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP14SegReq | IQon1 |  | WIP14 | February 21, 2019 8:30:40 PM GMT+08:00
CTUDT00601084 | Refresh | Closed | navale_s | January 4, 2019 9:55:02 PM GMT+08:00 | Clicking on Pause button after selecting “Start Clinical Scan” which appears during the tracker scan does not bring back the focus to Plan tab | 4.8 | BB_WIP5 | No | Subsystem |  | 2243 | Fixed | goswami_s | 0 | 2 | Console Software | Console4 | YES | Closed | BB_WIP4_SegReq | Brilliance Big Bore |  | BB_WIP5 | February 15, 2019 5:11:32 PM GMT+08:00
CTUDT00601083 | Refresh | Closed | tiwari_g | January 4, 2019 4:48:19 PM GMT+08:00 | Center(C) value is not resetting to maximum value upon entering higher value than maximum in the Windowing field for any result. | 4.8 | DHF316705 Rev06 |  | Subsystem |  | Console. SRS.5164 | Fixed | roy_r | 0 | 4 | Requirements Documentation | Console4 | NO | Closed | BB_WIP4_SubCom | Brilliance Big Bore |  | DHF316705 Rev06 | February 8, 2019 7:39:26 PM GMT+08:00
CTUDT00601070 | Refresh | Closed | attali_s | January 3, 2019 10:37:21 PM GMT+08:00 | Refresh: Tube Current: Error about Unvalid FocalSpot Type (DFX) | WIP12 | N/A |  | Subsystem | Tube | N/A | Fixed | braunstein_d | 0 | 2 | Design Documentation | DMS | NO | Closed | WIP14SegReq | Ovation |  | N/A | February 20, 2019 7:27:54 PM GMT+08:00
CTUDT00601070 | Refresh | Closed | attali_s | January 3, 2019 10:37:21 PM GMT+08:00 | Refresh: Tube Current: Error about Unvalid FocalSpot Type (DFX) | WIP12 | N/A |  | Subsystem | Tube | N/A | Fixed | braunstein_d | 0 | 2 | Design Documentation | DMS | NO | Closed | WIP14SegReq | Brilliance Big Bore |  | N/A | February 20, 2019 7:27:54 PM GMT+08:00
CTUDT00601070 | Refresh | Closed | attali_s | January 3, 2019 10:37:21 PM GMT+08:00 | Refresh: Tube Current: Error about Unvalid FocalSpot Type (DFX) | WIP12 | N/A |  | Subsystem | Tube | N/A | Fixed | braunstein_d | 0 | 2 | Design Documentation | DMS | NO | Closed | WIP14SegReq | Ingenuity 64 |  | N/A | February 20, 2019 7:27:54 PM GMT+08:00
CTUDT00601062 | Refresh | Closed | gushansky_m | January 3, 2019 8:06:21 PM GMT+08:00 | Refresh: Wrong Visual Diagnostics Menu | WIP12 |  |  | Subsystem | Visual Diagnostics |  | Does Not Apply | Mark_Huang | 0 |  |  | GCX | NO | Closed | WIP14SegReq | IQon1
CTUDT00601058 | Refresh | Closed | gushansky_m | January 3, 2019 4:20:56 PM GMT+08:00 | Refresh: Couch Visual Diagnostic failed | WIP12 |  |  | Subsystem | Visual Diagnostics |  | Duplicate | pandey_a | 0 |  |  | Service Tools | NO | Closed | WIP14SegReq | IQon1
CTUDT00601055 | Refresh | Closed | manjunath_c | January 3, 2019 3:21:17 PM GMT+08:00 | [Refresh]: Unable to add EC to Suggested Exam Cards group | WIP12 | WIP17 |  | Subsystem | Exam Cards; Blocker | NA | Fixed | datta_i | 0 | 4 | Console Software | Console4 | NO | Closed | WIP17SegReq | Ovation |  | WIP17 | May 29, 2019 7:18:37 PM GMT+08:00
CTUDT00601055 | Refresh | Closed | manjunath_c | January 3, 2019 3:21:17 PM GMT+08:00 | [Refresh]: Unable to add EC to Suggested Exam Cards group | WIP12 | WIP17 |  | Subsystem | Exam Cards; Blocker | NA | Fixed | datta_i | 0 | 4 | Console Software | Console4 | NO | Closed | WIP17SegReq | Brilliance 64 |  | WIP17 | May 29, 2019 7:18:37 PM GMT+08:00
CTUDT00601055 | Refresh | Closed | manjunath_c | January 3, 2019 3:21:17 PM GMT+08:00 | [Refresh]: Unable to add EC to Suggested Exam Cards group | WIP12 | WIP17 |  | Subsystem | Exam Cards; Blocker | NA | Fixed | datta_i | 0 | 4 | Console Software | Console4 | NO | Closed | WIP17SegReq | Ingenuity 64 |  | WIP17 | May 29, 2019 7:18:37 PM GMT+08:00
CTUDT00601055 | Refresh | Closed | manjunath_c | January 3, 2019 3:21:17 PM GMT+08:00 | [Refresh]: Unable to add EC to Suggested Exam Cards group | WIP12 | WIP17 |  | Subsystem | Exam Cards; Blocker | NA | Fixed | datta_i | 0 | 4 | Console Software | Console4 | NO | Closed | WIP17SegReq | IQon1 |  | WIP17 | May 29, 2019 7:18:37 PM GMT+08:00
CTUDT00601051 | Refresh | Closed | manjunath_c | January 3, 2019 2:30:51 PM GMT+08:00 | In ECM for Axial scans having more than one cycle the CTDIvol value is displayed with units wrongly arranged. | WIP12 |  |  | Subsystem | Verification Test Design | NA | No Plan to Fix | bandil_j | 0 | 4 |  | Console4 | NO | Closed | SubNC | Ovation
CTUDT00601051 | Refresh | Closed | manjunath_c | January 3, 2019 2:30:51 PM GMT+08:00 | In ECM for Axial scans having more than one cycle the CTDIvol value is displayed with units wrongly arranged. | WIP12 |  |  | Subsystem | Verification Test Design | NA | No Plan to Fix | bandil_j | 0 | 4 |  | Console4 | NO | Closed | SubNC | Brilliance 64
CTUDT00601051 | Refresh | Closed | manjunath_c | January 3, 2019 2:30:51 PM GMT+08:00 | In ECM for Axial scans having more than one cycle the CTDIvol value is displayed with units wrongly arranged. | WIP12 |  |  | Subsystem | Verification Test Design | NA | No Plan to Fix | bandil_j | 0 | 4 |  | Console4 | NO | Closed | SubNC | Brilliance Big Bore
CTUDT00601051 | Refresh | Closed | manjunath_c | January 3, 2019 2:30:51 PM GMT+08:00 | In ECM for Axial scans having more than one cycle the CTDIvol value is displayed with units wrongly arranged. | WIP12 |  |  | Subsystem | Verification Test Design | NA | No Plan to Fix | bandil_j | 0 | 4 |  | Console4 | NO | Closed | SubNC | Ingenuity 64
CTUDT00601051 | Refresh | Closed | manjunath_c | January 3, 2019 2:30:51 PM GMT+08:00 | In ECM for Axial scans having more than one cycle the CTDIvol value is displayed with units wrongly arranged. | WIP12 |  |  | Subsystem | Verification Test Design | NA | No Plan to Fix | bandil_j | 0 | 4 |  | Console4 | NO | Closed | SubNC | IQon1
CTUDT00601049 | Refresh | Closed | bagdady_b | January 2, 2019 10:05:03 PM GMT+08:00 | Refresh - Thickness field is editable during CCT Continuous scan with display set to 3 | WIP12 | WIP14 |  | Subsystem | CCT Continuous | NA | Fixed | datta_i | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SegReq | Ovation |  | WIP14 | February 14, 2019 6:53:43 PM GMT+08:00
CTUDT00601049 | Refresh | Closed | bagdady_b | January 2, 2019 10:05:03 PM GMT+08:00 | Refresh - Thickness field is editable during CCT Continuous scan with display set to 3 | WIP12 | WIP14 |  | Subsystem | CCT Continuous | NA | Fixed | datta_i | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SegReq | Brilliance 64 |  | WIP14 | February 14, 2019 6:53:43 PM GMT+08:00
CTUDT00601049 | Refresh | Closed | bagdady_b | January 2, 2019 10:05:03 PM GMT+08:00 | Refresh - Thickness field is editable during CCT Continuous scan with display set to 3 | WIP12 | WIP14 |  | Subsystem | CCT Continuous | NA | Fixed | datta_i | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SegReq | Brilliance Big Bore |  | WIP14 | February 14, 2019 6:53:43 PM GMT+08:00
CTUDT00601049 | Refresh | Closed | bagdady_b | January 2, 2019 10:05:03 PM GMT+08:00 | Refresh - Thickness field is editable during CCT Continuous scan with display set to 3 | WIP12 | WIP14 |  | Subsystem | CCT Continuous | NA | Fixed | datta_i | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SegReq | Ingenuity 64 |  | WIP14 | February 14, 2019 6:53:43 PM GMT+08:00
CTUDT00601049 | Refresh | Closed | bagdady_b | January 2, 2019 10:05:03 PM GMT+08:00 | Refresh - Thickness field is editable during CCT Continuous scan with display set to 3 | WIP12 | WIP14 |  | Subsystem | CCT Continuous | NA | Fixed | datta_i | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SegReq | IQon1 |  | WIP14 | February 14, 2019 6:53:43 PM GMT+08:00
CTUDT00601045 | Refresh | Closed | Dana_Kamyshev | January 2, 2019 8:11:39 PM GMT+08:00 | Refresh - 1-Click installation GUI is not displayed | WIP12 |  |  | Subsystem | Installation | STPS-SSRS-1155 | Works As Specified / Not A Defect | tamari_o | 1 |  |  | DI | NO | Closed |  | Ovation
CTUDT00601045 | Refresh | Closed | Dana_Kamyshev | January 2, 2019 8:11:39 PM GMT+08:00 | Refresh - 1-Click installation GUI is not displayed | WIP12 |  |  | Subsystem | Installation | STPS-SSRS-1155 | Works As Specified / Not A Defect | tamari_o | 1 |  |  | DI | NO | Closed |  | IQon1
CTUDT00601044 | Refresh | Closed | levin_c | January 2, 2019 8:10:08 PM GMT+08:00 | Refresh - Wrong values in "Recon MDU" and "Scan MDU" dicom tags | WIP11.1 |  | NA | Subsystem |  | NA | Works As Specified / Not A Defect | goswami_s | 0 |  |  | Console4 | NO | Closed | WIP14SegReq | Ovation
CTUDT00601044 | Refresh | Closed | levin_c | January 2, 2019 8:10:08 PM GMT+08:00 | Refresh - Wrong values in "Recon MDU" and "Scan MDU" dicom tags | WIP11.1 |  | NA | Subsystem |  | NA | Works As Specified / Not A Defect | goswami_s | 0 |  |  | Console4 | NO | Closed | WIP14SegReq | IQon1
CTUDT00601039 | Refresh | Closed | g_di | January 2, 2019 6:54:10 PM GMT+08:00 | Adaptive Filter value is “YES(CARDIAC) for a cardiac helical EC with IMR Recon mode and Cardiac Adaptive filter “ON” in image parameters and MDU file | WIP12 | DHF318051 Rev 11 | NA | Subsystem | cardiac with IMR | N/A | Fixed | ramesh_t | 0 | 2 | Test Design | Console4 | NO | Closed | WIP14SegReq | Ovation |  | DHF318051 Rev 11 | March 25, 2019 7:56:19 PM GMT+08:00
CTUDT00601039 | Refresh | Closed | g_di | January 2, 2019 6:54:10 PM GMT+08:00 | Adaptive Filter value is “YES(CARDIAC) for a cardiac helical EC with IMR Recon mode and Cardiac Adaptive filter “ON” in image parameters and MDU file | WIP12 | DHF318051 Rev 11 | NA | Subsystem | cardiac with IMR | N/A | Fixed | ramesh_t | 0 | 2 | Test Design | Console4 | NO | Closed | WIP14SegReq | Brilliance 64 |  | DHF318051 Rev 11 | March 25, 2019 7:56:19 PM GMT+08:00
CTUDT00601039 | Refresh | Closed | g_di | January 2, 2019 6:54:10 PM GMT+08:00 | Adaptive Filter value is “YES(CARDIAC) for a cardiac helical EC with IMR Recon mode and Cardiac Adaptive filter “ON” in image parameters and MDU file | WIP12 | DHF318051 Rev 11 | NA | Subsystem | cardiac with IMR | N/A | Fixed | ramesh_t | 0 | 2 | Test Design | Console4 | NO | Closed | WIP14SegReq | Ingenuity 64 |  | DHF318051 Rev 11 | March 25, 2019 7:56:19 PM GMT+08:00
CTUDT00601039 | Refresh | Closed | g_di | January 2, 2019 6:54:10 PM GMT+08:00 | Adaptive Filter value is “YES(CARDIAC) for a cardiac helical EC with IMR Recon mode and Cardiac Adaptive filter “ON” in image parameters and MDU file | WIP12 | DHF318051 Rev 11 | NA | Subsystem | cardiac with IMR | N/A | Fixed | ramesh_t | 0 | 2 | Test Design | Console4 | NO | Closed | WIP14SegReq | IQon1 |  | DHF318051 Rev 11 | March 25, 2019 7:56:19 PM GMT+08:00
CTUDT00601037 | Refresh | Closed | g_di | January 2, 2019 3:23:08 PM GMT+08:00 | Operator Name should not be displayed when clicked on Anonymous button | WIP12 | Host SSRS DHF303211 rev 09 | the user can manually delete the Operator name | Subsystem | Patient Scene | NA | Fixed | hatti_c | 0 | 2 | Console Software | Console4 | NO | Closed | WIP15SegReq | Ovation |  | Host SSRS DHF303211 rev 09 | April 16, 2019 5:31:47 PM GMT+08:00
CTUDT00601037 | Refresh | Closed | g_di | January 2, 2019 3:23:08 PM GMT+08:00 | Operator Name should not be displayed when clicked on Anonymous button | WIP12 | Host SSRS DHF303211 rev 09 | the user can manually delete the Operator name | Subsystem | Patient Scene | NA | Fixed | hatti_c | 0 | 2 | Console Software | Console4 | NO | Closed | WIP15SegReq | Brilliance 64 |  | Host SSRS DHF303211 rev 09 | April 16, 2019 5:31:47 PM GMT+08:00
CTUDT00601037 | Refresh | Closed | g_di | January 2, 2019 3:23:08 PM GMT+08:00 | Operator Name should not be displayed when clicked on Anonymous button | WIP12 | Host SSRS DHF303211 rev 09 | the user can manually delete the Operator name | Subsystem | Patient Scene | NA | Fixed | hatti_c | 0 | 2 | Console Software | Console4 | NO | Closed | WIP15SegReq | Ingenuity 64 |  | Host SSRS DHF303211 rev 09 | April 16, 2019 5:31:47 PM GMT+08:00
CTUDT00601037 | Refresh | Closed | g_di | January 2, 2019 3:23:08 PM GMT+08:00 | Operator Name should not be displayed when clicked on Anonymous button | WIP12 | Host SSRS DHF303211 rev 09 | the user can manually delete the Operator name | Subsystem | Patient Scene | NA | Fixed | hatti_c | 0 | 2 | Console Software | Console4 | NO | Closed | WIP15SegReq | IQon1 |  | Host SSRS DHF303211 rev 09 | April 16, 2019 5:31:47 PM GMT+08:00
CTUDT00601035 | Refresh | Closed | Balaji_Subbaram | January 2, 2019 3:05:08 PM GMT+08:00 | EC is getting duplicated after renaming with existing EC name | WIP12 | WIP18 | the user can rename the imported EC in the ECM, and when loading to plan, the user will review the parameters and change to the clinical need or load another EC | Subsystem | Exam Cards | NA | Fixed | datta_i | 0 | 3 | Console Software | Console4 | NO | Closed | WIP14SegReq | Ovation |  | WIP18 | June 3, 2019 7:47:18 PM GMT+08:00
CTUDT00601035 | Refresh | Closed | Balaji_Subbaram | January 2, 2019 3:05:08 PM GMT+08:00 | EC is getting duplicated after renaming with existing EC name | WIP12 | WIP18 | the user can rename the imported EC in the ECM, and when loading to plan, the user will review the parameters and change to the clinical need or load another EC | Subsystem | Exam Cards | NA | Fixed | datta_i | 0 | 3 | Console Software | Console4 | NO | Closed | WIP14SegReq | Brilliance 64 |  | WIP18 | June 3, 2019 7:47:18 PM GMT+08:00
CTUDT00601035 | Refresh | Closed | Balaji_Subbaram | January 2, 2019 3:05:08 PM GMT+08:00 | EC is getting duplicated after renaming with existing EC name | WIP12 | WIP18 | the user can rename the imported EC in the ECM, and when loading to plan, the user will review the parameters and change to the clinical need or load another EC | Subsystem | Exam Cards | NA | Fixed | datta_i | 0 | 3 | Console Software | Console4 | NO | Closed | WIP14SegReq | Ingenuity 64 |  | WIP18 | June 3, 2019 7:47:18 PM GMT+08:00
CTUDT00601035 | Refresh | Closed | Balaji_Subbaram | January 2, 2019 3:05:08 PM GMT+08:00 | EC is getting duplicated after renaming with existing EC name | WIP12 | WIP18 | the user can rename the imported EC in the ECM, and when loading to plan, the user will review the parameters and change to the clinical need or load another EC | Subsystem | Exam Cards | NA | Fixed | datta_i | 0 | 3 | Console Software | Console4 | NO | Closed | WIP14SegReq | IQon1 |  | WIP18 | June 3, 2019 7:47:18 PM GMT+08:00
CTUDT00601032 | Refresh | Closed | Dana_Kamyshev | January 2, 2019 3:21:28 AM GMT+08:00 | Refresh - Performance issue, long time for 'patient screen to be displayed | WIP12 | WIP16 | Yes, On clicking home or wrench icon or any other UI operation will bring back the Patient page to the front. | Subsystem | Performance | NA | Fixed | ravindra_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SegReq | Ovation |  | WIP16 | April 29, 2019 2:30:00 PM GMT+08:00
CTUDT00601032 | Refresh | Closed | Dana_Kamyshev | January 2, 2019 3:21:28 AM GMT+08:00 | Refresh - Performance issue, long time for 'patient screen to be displayed | WIP12 | WIP16 | Yes, On clicking home or wrench icon or any other UI operation will bring back the Patient page to the front. | Subsystem | Performance | NA | Fixed | ravindra_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SegReq | Brilliance 64 |  | WIP16 | April 29, 2019 2:30:00 PM GMT+08:00
CTUDT00601032 | Refresh | Closed | Dana_Kamyshev | January 2, 2019 3:21:28 AM GMT+08:00 | Refresh - Performance issue, long time for 'patient screen to be displayed | WIP12 | WIP16 | Yes, On clicking home or wrench icon or any other UI operation will bring back the Patient page to the front. | Subsystem | Performance | NA | Fixed | ravindra_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SegReq | Ingenuity 64 |  | WIP16 | April 29, 2019 2:30:00 PM GMT+08:00
CTUDT00601032 | Refresh | Closed | Dana_Kamyshev | January 2, 2019 3:21:28 AM GMT+08:00 | Refresh - Performance issue, long time for 'patient screen to be displayed | WIP12 | WIP16 | Yes, On clicking home or wrench icon or any other UI operation will bring back the Patient page to the front. | Subsystem | Performance | NA | Fixed | ravindra_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SegReq | IQon1 |  | WIP16 | April 29, 2019 2:30:00 PM GMT+08:00
CTUDT00601029 | Refresh | Closed | krisher_i | January 1, 2019 11:08:40 PM GMT+08:00 | Left panel is frozen on the transition from Patient scene to Plan scene | WIP12 | WIP17 |  | System | Performance | n/a | Fixed | sahar_l | 0 | 3 | DMS Calibrations | Support-System | NO | Closed | WIP14SegReq | IQon1 |  | WIP17 | May 22, 2019 7:06:49 PM GMT+08:00
CTUDT00601026 | Refresh | Closed | oreg_o | January 1, 2019 3:40:44 PM GMT+08:00 | Refresh: Calibrations: Air: The sorting order is not as written in the spec | WIP11.1 | DHF318940 | N/A | Subsystem | Calibrations | N/A | Fixed | yosfan_s | 0 | 2 | Test Design | DMS | NO | Closed | WIP14SegReq | Ovation |  | DHF318940 | February 27, 2019 4:04:53 PM GMT+08:00
CTUDT00601026 | Refresh | Closed | oreg_o | January 1, 2019 3:40:44 PM GMT+08:00 | Refresh: Calibrations: Air: The sorting order is not as written in the spec | WIP11.1 | DHF318940 | N/A | Subsystem | Calibrations | N/A | Fixed | yosfan_s | 0 | 2 | Test Design | DMS | NO | Closed | WIP14SegReq | Brilliance Big Bore |  | DHF318940 | February 27, 2019 4:04:53 PM GMT+08:00
CTUDT00601026 | Refresh | Closed | oreg_o | January 1, 2019 3:40:44 PM GMT+08:00 | Refresh: Calibrations: Air: The sorting order is not as written in the spec | WIP11.1 | DHF318940 | N/A | Subsystem | Calibrations | N/A | Fixed | yosfan_s | 0 | 2 | Test Design | DMS | NO | Closed | WIP14SegReq | Ingenuity 64 |  | DHF318940 | February 27, 2019 4:04:53 PM GMT+08:00
CTUDT00601026 | Refresh | Closed | oreg_o | January 1, 2019 3:40:44 PM GMT+08:00 | Refresh: Calibrations: Air: The sorting order is not as written in the spec | WIP11.1 | DHF318940 | N/A | Subsystem | Calibrations | N/A | Fixed | yosfan_s | 0 | 2 | Test Design | DMS | NO | Closed | WIP14SegReq | IQon1 |  | DHF318940 | February 27, 2019 4:04:53 PM GMT+08:00
CTUDT00601024 | Refresh | Closed | Dana_Kamyshev | December 31, 2018 9:58:06 PM GMT+08:00 | Refresh - Performance issue, cannot login to CT in 30 seconds | WIP12 | DHF318051 Rev 14 HOST Verification Test Protocol | N/A | Subsystem | Performance | N/A | Fixed | ramesh_t | 0 | 3 | Test Design | Console4 | NO | Closed | WIP15SubCom | Ovation |  | DHF318051 Rev 14 HOST Verification Test Protocol | April 26, 2019 2:15:42 PM GMT+08:00
CTUDT00601024 | Refresh | Closed | Dana_Kamyshev | December 31, 2018 9:58:06 PM GMT+08:00 | Refresh - Performance issue, cannot login to CT in 30 seconds | WIP12 | DHF318051 Rev 14 HOST Verification Test Protocol | N/A | Subsystem | Performance | N/A | Fixed | ramesh_t | 0 | 3 | Test Design | Console4 | NO | Closed | WIP15SubCom | Brilliance 64 |  | DHF318051 Rev 14 HOST Verification Test Protocol | April 26, 2019 2:15:42 PM GMT+08:00
CTUDT00601024 | Refresh | Closed | Dana_Kamyshev | December 31, 2018 9:58:06 PM GMT+08:00 | Refresh - Performance issue, cannot login to CT in 30 seconds | WIP12 | DHF318051 Rev 14 HOST Verification Test Protocol | N/A | Subsystem | Performance | N/A | Fixed | ramesh_t | 0 | 3 | Test Design | Console4 | NO | Closed | WIP15SubCom | Brilliance Big Bore |  | DHF318051 Rev 14 HOST Verification Test Protocol | April 26, 2019 2:15:42 PM GMT+08:00
CTUDT00601024 | Refresh | Closed | Dana_Kamyshev | December 31, 2018 9:58:06 PM GMT+08:00 | Refresh - Performance issue, cannot login to CT in 30 seconds | WIP12 | DHF318051 Rev 14 HOST Verification Test Protocol | N/A | Subsystem | Performance | N/A | Fixed | ramesh_t | 0 | 3 | Test Design | Console4 | NO | Closed | WIP15SubCom | Ingenuity 64 |  | DHF318051 Rev 14 HOST Verification Test Protocol | April 26, 2019 2:15:42 PM GMT+08:00
CTUDT00601024 | Refresh | Closed | Dana_Kamyshev | December 31, 2018 9:58:06 PM GMT+08:00 | Refresh - Performance issue, cannot login to CT in 30 seconds | WIP12 | DHF318051 Rev 14 HOST Verification Test Protocol | N/A | Subsystem | Performance | N/A | Fixed | ramesh_t | 0 | 3 | Test Design | Console4 | NO | Closed | WIP15SubCom | IQon1 |  | DHF318051 Rev 14 HOST Verification Test Protocol | April 26, 2019 2:15:42 PM GMT+08:00
CTUDT00601018 | Refresh | Closed | beck_d | December 30, 2018 6:55:43 PM GMT+08:00 | Refresh: Calibrations: Performance Test: Parameters and limits are not according to spec | WIP11 | WIP14 |  | Subsystem |  | DHF170396 | Documentation Created/Updated | braunstein_d | 0 | 2 | Design Documentation | DMS | NO | Closed | WIP13SegReq | Ovation |  | WIP14 | February 3, 2019 3:53:13 PM GMT+08:00
CTUDT00601015 | Refresh | Closed | gushansky_m | December 30, 2018 3:16:03 PM GMT+08:00 | Remote Service in Service Tools doesn't work | WIP12 |  |  | Subsystem | Remote ConfigurationTool |  | Duplicate | Sandeep_Kumar | 0 |  |  | Service Tools | NO | Closed | WIP14SegReq | IQon1
CTUDT00601013 | Refresh | Closed | krisher_i | December 27, 2018 10:56:59 PM GMT+08:00 | Refresh - Burnt Label on Zoom and Pan is Not in the correct size | WIP12 |  |  | Subsystem | sCTV - Burnt Label | NA | No Plan to Fix | redrouthu_r | 0 | 4 |  | iCAP | NO | Closed |  | IQon1
CTUDT00601011 | Refresh | Closed | bagdady_b | December 27, 2018 9:36:10 PM GMT+08:00 | Refresh - When performing a perfusion scan a CIRS error appears | WIP12 |  | N/A | Subsystem | Perfusion | N/A | Works As Specified / Not A Defect | tamari_o | 1 |  |  | Support-System | NO | Closed |  | Ovation
CTUDT00601010 | Refresh | Closed | borise | December 27, 2018 9:05:37 PM GMT+08:00 | Easy memory leaks due to memory leaks in CanStreamControl client | BR3.5 |  |  | Subsystem | Performance | NA | No Plan to Fix | karkour | 0 | 4 | Gantry | GCX | NO | Closed |  | Ovation
CTUDT00601010 | Refresh | Closed | borise | December 27, 2018 9:05:37 PM GMT+08:00 | Easy memory leaks due to memory leaks in CanStreamControl client | BR3.5 |  |  | Subsystem | Performance | NA | No Plan to Fix | karkour | 0 | 4 | Gantry | GCX | NO | Closed |  | Brilliance 64
CTUDT00601010 | Refresh | Closed | borise | December 27, 2018 9:05:37 PM GMT+08:00 | Easy memory leaks due to memory leaks in CanStreamControl client | BR3.5 |  |  | Subsystem | Performance | NA | No Plan to Fix | karkour | 0 | 4 | Gantry | GCX | NO | Closed |  | Brilliance Big Bore
CTUDT00601010 | Refresh | Closed | borise | December 27, 2018 9:05:37 PM GMT+08:00 | Easy memory leaks due to memory leaks in CanStreamControl client | BR3.5 |  |  | Subsystem | Performance | NA | No Plan to Fix | karkour | 0 | 4 | Gantry | GCX | NO | Closed |  | Ingenuity 64
CTUDT00601010 | Refresh | Closed | borise | December 27, 2018 9:05:37 PM GMT+08:00 | Easy memory leaks due to memory leaks in CanStreamControl client | BR3.5 |  |  | Subsystem | Performance | NA | No Plan to Fix | karkour | 0 | 4 | Gantry | GCX | NO | Closed |  | IQon1
CTUDT00601007 | Refresh | Closed | or_shmuel | December 27, 2018 5:49:09 PM GMT+08:00 | Refresh, iCT - Windows login appeared after restarting the HOST | WIP12 | WIP13 | On hard reboot of Host System, Login page appeared. | Subsystem |  | NA | Fixed | kapadia_v | 0 | 3 | Console Software | Console4 | NO | Closed | WIP15SegReq | Ovation |  | WIP13 | January 29, 2019 6:24:18 PM GMT+08:00
CTUDT00601007 | Refresh | Closed | or_shmuel | December 27, 2018 5:49:09 PM GMT+08:00 | Refresh, iCT - Windows login appeared after restarting the HOST | WIP12 | WIP13 | On hard reboot of Host System, Login page appeared. | Subsystem |  | NA | Fixed | kapadia_v | 0 | 3 | Console Software | Console4 | NO | Closed | WIP15SegReq | Brilliance 64 |  | WIP13 | January 29, 2019 6:24:18 PM GMT+08:00
CTUDT00601007 | Refresh | Closed | or_shmuel | December 27, 2018 5:49:09 PM GMT+08:00 | Refresh, iCT - Windows login appeared after restarting the HOST | WIP12 | WIP13 | On hard reboot of Host System, Login page appeared. | Subsystem |  | NA | Fixed | kapadia_v | 0 | 3 | Console Software | Console4 | NO | Closed | WIP15SegReq | Brilliance Big Bore |  | WIP13 | January 29, 2019 6:24:18 PM GMT+08:00
CTUDT00601007 | Refresh | Closed | or_shmuel | December 27, 2018 5:49:09 PM GMT+08:00 | Refresh, iCT - Windows login appeared after restarting the HOST | WIP12 | WIP13 | On hard reboot of Host System, Login page appeared. | Subsystem |  | NA | Fixed | kapadia_v | 0 | 3 | Console Software | Console4 | NO | Closed | WIP15SegReq | Ingenuity 64 |  | WIP13 | January 29, 2019 6:24:18 PM GMT+08:00
CTUDT00601007 | Refresh | Closed | or_shmuel | December 27, 2018 5:49:09 PM GMT+08:00 | Refresh, iCT - Windows login appeared after restarting the HOST | WIP12 | WIP13 | On hard reboot of Host System, Login page appeared. | Subsystem |  | NA | Fixed | kapadia_v | 0 | 3 | Console Software | Console4 | NO | Closed | WIP15SegReq | IQon1 |  | WIP13 | January 29, 2019 6:24:18 PM GMT+08:00
CTUDT00601005 | Refresh | Closed | attali_s | December 27, 2018 5:22:37 PM GMT+08:00 | Refresh: Performance test: When disabling RT 0.272 in the Key, the scan Cardiac Body 1 (that uses RT 0.272) is erased from Customization Window | WIP11 | WIP14 | N/A | Subsystem |  | N/A | Fixed | elviral | 0 | 2 | DMS Software | DMS | YES | Closed | WIP14SegReq | Ovation |  | WIP14 | February 7, 2019 4:43:32 PM GMT+08:00
CTUDT00601000 | Refresh | Closed | Noy_Seelig | December 26, 2018 10:14:50 PM GMT+08:00 | Refresh: No CTDI Free Air values in Performance Infant Head cases | WIP11 | DHF218005 Rev05 WIP17 | n/a | System | Physics | DHF218005 | Fixed | mao_y | 0 | 2 | Physics Configurations | Support-Physics | NO | Closed | WIP14SegReq | Ovation |  | DHF218005 Rev05 WIP17 | May 24, 2019 10:34:16 PM GMT+08:00
CTUDT00600997 | Refresh | Closed | braunstein_d | December 26, 2018 9:12:14 PM GMT+08:00 | Implementing mapping from Adult Head to Infant Head and Body the mapping do not read the scaling factor from Adult Head | WIP12 | WIP14 | N/A | Subsystem | Calibrations | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP14SegReq | Brilliance Big Bore |  | WIP14 | February 21, 2019 12:05:07 AM GMT+08:00
CTUDT00600997 | Refresh | Closed | braunstein_d | December 26, 2018 9:12:14 PM GMT+08:00 | Implementing mapping from Adult Head to Infant Head and Body the mapping do not read the scaling factor from Adult Head | WIP12 | WIP14 | N/A | Subsystem | Calibrations | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP14SegReq | Ingenuity 64 |  | WIP14 | February 21, 2019 12:05:07 AM GMT+08:00
CTUDT00600997 | Refresh | Closed | braunstein_d | December 26, 2018 9:12:14 PM GMT+08:00 | Implementing mapping from Adult Head to Infant Head and Body the mapping do not read the scaling factor from Adult Head | WIP12 | WIP14 | N/A | Subsystem | Calibrations | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP14SegReq | IQon1 |  | WIP14 | February 21, 2019 12:05:07 AM GMT+08:00
CTUDT00600991 | Refresh | Closed | beck_d | December 24, 2018 10:00:15 PM GMT+08:00 | Refresh: Calibrations: Error appears when trying to re-perform only the failed cases of an IQ test | WIP11.1 | WIP15 | the user can run the entire IQ test again. | Subsystem | Calibrations | N/A | Fixed | elviral | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | Ovation |  | WIP15 | January 13, 2019 3:23:29 PM GMT+08:00
CTUDT00600991 | Refresh | Closed | beck_d | December 24, 2018 10:00:15 PM GMT+08:00 | Refresh: Calibrations: Error appears when trying to re-perform only the failed cases of an IQ test | WIP11.1 | WIP15 | the user can run the entire IQ test again. | Subsystem | Calibrations | N/A | Fixed | elviral | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | Brilliance 64 |  | WIP15 | January 13, 2019 3:23:29 PM GMT+08:00
CTUDT00600991 | Refresh | Closed | beck_d | December 24, 2018 10:00:15 PM GMT+08:00 | Refresh: Calibrations: Error appears when trying to re-perform only the failed cases of an IQ test | WIP11.1 | WIP15 | the user can run the entire IQ test again. | Subsystem | Calibrations | N/A | Fixed | elviral | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | Brilliance Big Bore |  | WIP15 | January 13, 2019 3:23:29 PM GMT+08:00
CTUDT00600991 | Refresh | Closed | beck_d | December 24, 2018 10:00:15 PM GMT+08:00 | Refresh: Calibrations: Error appears when trying to re-perform only the failed cases of an IQ test | WIP11.1 | WIP15 | the user can run the entire IQ test again. | Subsystem | Calibrations | N/A | Fixed | elviral | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | Ingenuity 64 |  | WIP15 | January 13, 2019 3:23:29 PM GMT+08:00
CTUDT00600991 | Refresh | Closed | beck_d | December 24, 2018 10:00:15 PM GMT+08:00 | Refresh: Calibrations: Error appears when trying to re-perform only the failed cases of an IQ test | WIP11.1 | WIP15 | the user can run the entire IQ test again. | Subsystem | Calibrations | N/A | Fixed | elviral | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | IQon1 |  | WIP15 | January 13, 2019 3:23:29 PM GMT+08:00
CTUDT00600988 | Refresh | Closed | oreg_o | December 23, 2018 5:25:54 PM GMT+08:00 | Refresh: Calibrations: Performance:in linearity case, the ROIs appeared on few slices and not on the two middle slices. | WIP11.1 |  |  | Subsystem |  |  | Duplicate | gehtin_e | 0 |  |  |  |  | Closed |  | Ingenuity 64
CTUDT00600976 | Refresh | Closed | vinoliya_j | December 20, 2018 8:23:22 PM GMT+08:00 | Refresh : When P3T is on, weight field shows wrong units. | WIP11.1 | WIP14 |  | System | Exam Cards | NA | Fixed | datta_i | 0 | 2 | Console Software | Console4 | NO | Closed | WIP15SegReq | Brilliance 64 |  | WIP14 | February 13, 2019 9:48:44 PM GMT+08:00
CTUDT00600976 | Refresh | Closed | vinoliya_j | December 20, 2018 8:23:22 PM GMT+08:00 | Refresh : When P3T is on, weight field shows wrong units. | WIP11.1 | WIP14 |  | System | Exam Cards | NA | Fixed | datta_i | 0 | 2 | Console Software | Console4 | NO | Closed | WIP15SegReq | Brilliance Big Bore |  | WIP14 | February 13, 2019 9:48:44 PM GMT+08:00
CTUDT00600976 | Refresh | Closed | vinoliya_j | December 20, 2018 8:23:22 PM GMT+08:00 | Refresh : When P3T is on, weight field shows wrong units. | WIP11.1 | WIP14 |  | System | Exam Cards | NA | Fixed | datta_i | 0 | 2 | Console Software | Console4 | NO | Closed | WIP15SegReq | Ingenuity 64 |  | WIP14 | February 13, 2019 9:48:44 PM GMT+08:00
CTUDT00600976 | Refresh | Closed | vinoliya_j | December 20, 2018 8:23:22 PM GMT+08:00 | Refresh : When P3T is on, weight field shows wrong units. | WIP11.1 | WIP14 |  | System | Exam Cards | NA | Fixed | datta_i | 0 | 2 | Console Software | Console4 | NO | Closed | WIP15SegReq | IQon1 |  | WIP14 | February 13, 2019 9:48:44 PM GMT+08:00
CTUDT00600971 | Refresh | Closed | dongping_w | December 20, 2018 7:49:01 PM GMT+08:00 | Typo correction: The FDA clause information in RegStand.349 is not correct. | 4.1.7 |  | NA | System |  |  | Does Not Apply | shaham_t | 0 | 5 |  | CIRS | NO | Closed |  | iCT4
CTUDT00600971 | Refresh | Closed | dongping_w | December 20, 2018 7:49:01 PM GMT+08:00 | Typo correction: The FDA clause information in RegStand.349 is not correct. | 4.1.7 |  | NA | System |  |  | Does Not Apply | shaham_t | 0 | 5 |  | CIRS | NO | Closed |  | iCT2
CTUDT00600971 | Refresh | Closed | dongping_w | December 20, 2018 7:49:01 PM GMT+08:00 | Typo correction: The FDA clause information in RegStand.349 is not correct. | 4.1.7 |  | NA | System |  |  | Does Not Apply | shaham_t | 0 | 5 |  | CIRS | NO | Closed |  | iCT SDCT
CTUDT00600971 | Refresh | Closed | dongping_w | December 20, 2018 7:49:01 PM GMT+08:00 | Typo correction: The FDA clause information in RegStand.349 is not correct. | 4.1.7 |  | NA | System |  |  | Does Not Apply | shaham_t | 0 | 5 |  | CIRS | NO | Closed |  | IQon1
CTUDT00600967 | Refresh | Closed | daniel_l | December 20, 2018 7:10:59 PM GMT+08:00 | External and Internal NTP Sync does not work at the same time | WIP11 | DMR356920 Rev.06 459801587361 459801587331 | N/A - Service documentation should be updated. | System | NPT server connection; | TBD | Fixed | li_e | 0 | 4 | Design Documentation | Support-Service | NO | Closed | SubCom-03-Dec-2019 | Ovation |  | DMR356920 Rev.06 459801587361 459801587331 | December 4, 2019 3:08:38 PM GMT+08:00
CTUDT00600967 | Refresh | Closed | daniel_l | December 20, 2018 7:10:59 PM GMT+08:00 | External and Internal NTP Sync does not work at the same time | WIP11 | DMR356920 Rev.06 459801587361 459801587331 | N/A - Service documentation should be updated. | System | NPT server connection; | TBD | Fixed | li_e | 0 | 4 | Design Documentation | Support-Service | NO | Closed | SubCom-03-Dec-2019 | Brilliance 64 |  | DMR356920 Rev.06 459801587361 459801587331 | December 4, 2019 3:08:38 PM GMT+08:00
CTUDT00600967 | Refresh | Closed | daniel_l | December 20, 2018 7:10:59 PM GMT+08:00 | External and Internal NTP Sync does not work at the same time | WIP11 | DMR356920 Rev.06 459801587361 459801587331 | N/A - Service documentation should be updated. | System | NPT server connection; | TBD | Fixed | li_e | 0 | 4 | Design Documentation | Support-Service | NO | Closed | SubCom-03-Dec-2019 | Brilliance Big Bore |  | DMR356920 Rev.06 459801587361 459801587331 | December 4, 2019 3:08:38 PM GMT+08:00
CTUDT00600967 | Refresh | Closed | daniel_l | December 20, 2018 7:10:59 PM GMT+08:00 | External and Internal NTP Sync does not work at the same time | WIP11 | DMR356920 Rev.06 459801587361 459801587331 | N/A - Service documentation should be updated. | System | NPT server connection; | TBD | Fixed | li_e | 0 | 4 | Design Documentation | Support-Service | NO | Closed | SubCom-03-Dec-2019 | Ingenuity 64 |  | DMR356920 Rev.06 459801587361 459801587331 | December 4, 2019 3:08:38 PM GMT+08:00
CTUDT00600967 | Refresh | Closed | daniel_l | December 20, 2018 7:10:59 PM GMT+08:00 | External and Internal NTP Sync does not work at the same time | WIP11 | DMR356920 Rev.06 459801587361 459801587331 | N/A - Service documentation should be updated. | System | NPT server connection; | TBD | Fixed | li_e | 0 | 4 | Design Documentation | Support-Service | NO | Closed | SubCom-03-Dec-2019 | IQon1 |  | DMR356920 Rev.06 459801587361 459801587331 | December 4, 2019 3:08:38 PM GMT+08:00
CTUDT00600965 | Refresh | Closed | vinoliya_j | December 20, 2018 6:36:06 PM GMT+08:00 | Refresh: Mapping of Drive V throws an error message. | WIP11.1 |  | 1. GO to Control panel -> Turn windows features on or off 2. Check the checkbox next to SMB 1.0/CIFS File Sharing Support and SMB Direct, click OK to accept the changes 3. Click restart now to reboot the system. 4. Retry mapping V drive 5. Map drive will be successfully mapped. | Subsystem | Map Drive | NA | Invalid/Uncontrolled Configuration | ramesh_t | 0 | 2 |  | Console4 | NO | Closed | SubCom | Brilliance 64
CTUDT00600965 | Refresh | Closed | vinoliya_j | December 20, 2018 6:36:06 PM GMT+08:00 | Refresh: Mapping of Drive V throws an error message. | WIP11.1 |  | 1. GO to Control panel -> Turn windows features on or off 2. Check the checkbox next to SMB 1.0/CIFS File Sharing Support and SMB Direct, click OK to accept the changes 3. Click restart now to reboot the system. 4. Retry mapping V drive 5. Map drive will be successfully mapped. | Subsystem | Map Drive | NA | Invalid/Uncontrolled Configuration | ramesh_t | 0 | 2 |  | Console4 | NO | Closed | SubCom | Brilliance Big Bore
CTUDT00600965 | Refresh | Closed | vinoliya_j | December 20, 2018 6:36:06 PM GMT+08:00 | Refresh: Mapping of Drive V throws an error message. | WIP11.1 |  | 1. GO to Control panel -> Turn windows features on or off 2. Check the checkbox next to SMB 1.0/CIFS File Sharing Support and SMB Direct, click OK to accept the changes 3. Click restart now to reboot the system. 4. Retry mapping V drive 5. Map drive will be successfully mapped. | Subsystem | Map Drive | NA | Invalid/Uncontrolled Configuration | ramesh_t | 0 | 2 |  | Console4 | NO | Closed | SubCom | Ingenuity 64
CTUDT00600960 | Refresh | Closed | vinoliya_j | December 20, 2018 1:42:42 AM GMT+08:00 | Refresh: Cosmix filter result is stored as 16 bit HU image but it has to be stored as 12 bit HU image. | WIP11.1 | WIP17 |  | Subsystem | Cardiac | NA | Fixed | rath_s | 0 | 4 | Console Software | Console4 | NO | Closed | WIP17SubCom | Ovation |  | WIP17 | May 22, 2019 5:06:07 PM GMT+08:00
CTUDT00600960 | Refresh | Closed | vinoliya_j | December 20, 2018 1:42:42 AM GMT+08:00 | Refresh: Cosmix filter result is stored as 16 bit HU image but it has to be stored as 12 bit HU image. | WIP11.1 | WIP17 |  | Subsystem | Cardiac | NA | Fixed | rath_s | 0 | 4 | Console Software | Console4 | NO | Closed | WIP17SubCom | Brilliance 64 |  | WIP17 | May 22, 2019 5:06:07 PM GMT+08:00
CTUDT00600960 | Refresh | Closed | vinoliya_j | December 20, 2018 1:42:42 AM GMT+08:00 | Refresh: Cosmix filter result is stored as 16 bit HU image but it has to be stored as 12 bit HU image. | WIP11.1 | WIP17 |  | Subsystem | Cardiac | NA | Fixed | rath_s | 0 | 4 | Console Software | Console4 | NO | Closed | WIP17SubCom | Brilliance Big Bore |  | WIP17 | May 22, 2019 5:06:07 PM GMT+08:00
CTUDT00600960 | Refresh | Closed | vinoliya_j | December 20, 2018 1:42:42 AM GMT+08:00 | Refresh: Cosmix filter result is stored as 16 bit HU image but it has to be stored as 12 bit HU image. | WIP11.1 | WIP17 |  | Subsystem | Cardiac | NA | Fixed | rath_s | 0 | 4 | Console Software | Console4 | NO | Closed | WIP17SubCom | Ingenuity 64 |  | WIP17 | May 22, 2019 5:06:07 PM GMT+08:00
CTUDT00600960 | Refresh | Closed | vinoliya_j | December 20, 2018 1:42:42 AM GMT+08:00 | Refresh: Cosmix filter result is stored as 16 bit HU image but it has to be stored as 12 bit HU image. | WIP11.1 | WIP17 |  | Subsystem | Cardiac | NA | Fixed | rath_s | 0 | 4 | Console Software | Console4 | NO | Closed | WIP17SubCom | IQon1 |  | WIP17 | May 22, 2019 5:06:07 PM GMT+08:00
CTUDT00600955 | Refresh | Closed | or_shmuel | December 19, 2018 10:53:12 PM GMT+08:00 | Refresh, IQon - Usability: Disable options are not display clear to the user | WIP11.1 | WIP15 |  | Subsystem | ECM | NA | Fixed | wilsonraja | 0 | 4 | Console Software | Console4 | NO | Closed | WIP14SegReq | Ovation |  | WIP15 | March 22, 2019 2:47:06 PM GMT+08:00
CTUDT00600955 | Refresh | Closed | or_shmuel | December 19, 2018 10:53:12 PM GMT+08:00 | Refresh, IQon - Usability: Disable options are not display clear to the user | WIP11.1 | WIP15 |  | Subsystem | ECM | NA | Fixed | wilsonraja | 0 | 4 | Console Software | Console4 | NO | Closed | WIP14SegReq | Brilliance 64 |  | WIP15 | March 22, 2019 2:47:06 PM GMT+08:00
CTUDT00600955 | Refresh | Closed | or_shmuel | December 19, 2018 10:53:12 PM GMT+08:00 | Refresh, IQon - Usability: Disable options are not display clear to the user | WIP11.1 | WIP15 |  | Subsystem | ECM | NA | Fixed | wilsonraja | 0 | 4 | Console Software | Console4 | NO | Closed | WIP14SegReq | Ingenuity 64 |  | WIP15 | March 22, 2019 2:47:06 PM GMT+08:00
CTUDT00600955 | Refresh | Closed | or_shmuel | December 19, 2018 10:53:12 PM GMT+08:00 | Refresh, IQon - Usability: Disable options are not display clear to the user | WIP11.1 | WIP15 |  | Subsystem | ECM | NA | Fixed | wilsonraja | 0 | 4 | Console Software | Console4 | NO | Closed | WIP14SegReq | IQon1 |  | WIP15 | March 22, 2019 2:47:06 PM GMT+08:00
CTUDT00600950 | Refresh | Closed | panchami_pradeep | December 19, 2018 6:40:26 PM GMT+08:00 | Refresh: PSC: Remote configuration: connection through ISSL and RSN networks are displaying an error message. | WIP12 | WIP14 | No Workaround | Subsystem | PSC - remote configuration | N/A | Fixed | ranjithkumar_m | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP13SubCom | Ovation |  | WIP14 | February 19, 2019 5:53:31 PM GMT+08:00
CTUDT00600950 | Refresh | Closed | panchami_pradeep | December 19, 2018 6:40:26 PM GMT+08:00 | Refresh: PSC: Remote configuration: connection through ISSL and RSN networks are displaying an error message. | WIP12 | WIP14 | No Workaround | Subsystem | PSC - remote configuration | N/A | Fixed | ranjithkumar_m | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP13SubCom | Brilliance 64 |  | WIP14 | February 19, 2019 5:53:31 PM GMT+08:00
CTUDT00600950 | Refresh | Closed | panchami_pradeep | December 19, 2018 6:40:26 PM GMT+08:00 | Refresh: PSC: Remote configuration: connection through ISSL and RSN networks are displaying an error message. | WIP12 | WIP14 | No Workaround | Subsystem | PSC - remote configuration | N/A | Fixed | ranjithkumar_m | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP13SubCom | Brilliance Big Bore |  | WIP14 | February 19, 2019 5:53:31 PM GMT+08:00
CTUDT00600950 | Refresh | Closed | panchami_pradeep | December 19, 2018 6:40:26 PM GMT+08:00 | Refresh: PSC: Remote configuration: connection through ISSL and RSN networks are displaying an error message. | WIP12 | WIP14 | No Workaround | Subsystem | PSC - remote configuration | N/A | Fixed | ranjithkumar_m | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP13SubCom | Ingenuity 64 |  | WIP14 | February 19, 2019 5:53:31 PM GMT+08:00
CTUDT00600950 | Refresh | Closed | panchami_pradeep | December 19, 2018 6:40:26 PM GMT+08:00 | Refresh: PSC: Remote configuration: connection through ISSL and RSN networks are displaying an error message. | WIP12 | WIP14 | No Workaround | Subsystem | PSC - remote configuration | N/A | Fixed | ranjithkumar_m | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP13SubCom | IQon1 |  | WIP14 | February 19, 2019 5:53:31 PM GMT+08:00
CTUDT00600940 | Refresh | Closed | or_shmuel | December 18, 2018 10:27:22 PM GMT+08:00 | Refresh, IQon - Bugrep cannot be collected in CT user | WIP11.1 | WIP19 |  | Subsystem | Bugrep Collection;Blocker | N/A | Fixed | v_sumanth | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP14SegReq | IQon1 |  | WIP19 | August 1, 2019 9:16:24 PM GMT+08:00
CTUDT00600938 | Refresh | Closed | Noy_Seelig | December 18, 2018 9:51:30 PM GMT+08:00 | Refresh:Physics Analysis: Error during Reconstruction stage asking to user to create a Bug Report | WIP11.1 |  | to reopen Physics analysis and start with desired case | Subsystem | Physics Analysis | N/A | No Plan to Fix | yosefk | 0 | 3 |  | DMS | NO | Closed |  | Ovation
CTUDT00600938 | Refresh | Closed | Noy_Seelig | December 18, 2018 9:51:30 PM GMT+08:00 | Refresh:Physics Analysis: Error during Reconstruction stage asking to user to create a Bug Report | WIP11.1 |  | to reopen Physics analysis and start with desired case | Subsystem | Physics Analysis | N/A | No Plan to Fix | yosefk | 0 | 3 |  | DMS | NO | Closed |  | Brilliance Big Bore
CTUDT00600938 | Refresh | Closed | Noy_Seelig | December 18, 2018 9:51:30 PM GMT+08:00 | Refresh:Physics Analysis: Error during Reconstruction stage asking to user to create a Bug Report | WIP11.1 |  | to reopen Physics analysis and start with desired case | Subsystem | Physics Analysis | N/A | No Plan to Fix | yosefk | 0 | 3 |  | DMS | NO | Closed |  | Ingenuity 64
CTUDT00600938 | Refresh | Closed | Noy_Seelig | December 18, 2018 9:51:30 PM GMT+08:00 | Refresh:Physics Analysis: Error during Reconstruction stage asking to user to create a Bug Report | WIP11.1 |  | to reopen Physics analysis and start with desired case | Subsystem | Physics Analysis | N/A | No Plan to Fix | yosefk | 0 | 3 |  | DMS | NO | Closed |  | IQon1
CTUDT00600928 | Refresh | Closed | Revital_Cohen | December 18, 2018 5:30:47 PM GMT+08:00 | Refresh - IQon - SBI image files are created with more than 1000KB | WIP11 | Host SSRS DHF303211 Rev 08 | none | Subsystem | SBI size (requirement update) | Console-SRS.13349 | Fixed | bandil_j | 0 | 2 | Host SSRS | Support_Spectral | NO | Closed | WIP15SegReq | IQon1 |  | Host SSRS DHF303211 Rev 08 | March 22, 2019 7:42:15 PM GMT+08:00
CTUDT00600926 | Refresh | Closed | gT_s | December 18, 2018 5:02:38 PM GMT+08:00 | Refresh: The following HOW DO I movies-> plan Geometry limitations, Dose alert, Dose notification, Volume display, Offline Reconstruction are not displayed | WIP11.1 | WIP15 |  | Subsystem | How Do I | NA | Fixed | vishnu_devi | 0 | 2 | Console Software | Console4 | NO | Closed | WIP15SegReq | Ovation |  | WIP15 | March 26, 2019 2:02:42 PM GMT+08:00
CTUDT00600926 | Refresh | Closed | gT_s | December 18, 2018 5:02:38 PM GMT+08:00 | Refresh: The following HOW DO I movies-> plan Geometry limitations, Dose alert, Dose notification, Volume display, Offline Reconstruction are not displayed | WIP11.1 | WIP15 |  | Subsystem | How Do I | NA | Fixed | vishnu_devi | 0 | 2 | Console Software | Console4 | NO | Closed | WIP15SegReq | Brilliance 64 |  | WIP15 | March 26, 2019 2:02:42 PM GMT+08:00
CTUDT00600926 | Refresh | Closed | gT_s | December 18, 2018 5:02:38 PM GMT+08:00 | Refresh: The following HOW DO I movies-> plan Geometry limitations, Dose alert, Dose notification, Volume display, Offline Reconstruction are not displayed | WIP11.1 | WIP15 |  | Subsystem | How Do I | NA | Fixed | vishnu_devi | 0 | 2 | Console Software | Console4 | NO | Closed | WIP15SegReq | Ingenuity 64 |  | WIP15 | March 26, 2019 2:02:42 PM GMT+08:00
CTUDT00600926 | Refresh | Closed | gT_s | December 18, 2018 5:02:38 PM GMT+08:00 | Refresh: The following HOW DO I movies-> plan Geometry limitations, Dose alert, Dose notification, Volume display, Offline Reconstruction are not displayed | WIP11.1 | WIP15 |  | Subsystem | How Do I | NA | Fixed | vishnu_devi | 0 | 2 | Console Software | Console4 | NO | Closed | WIP15SegReq | IQon1 |  | WIP15 | March 26, 2019 2:02:42 PM GMT+08:00
CTUDT00600924 | Refresh | Closed | ravindra_a | December 18, 2018 4:48:22 PM GMT+08:00 | Refresh: Bugrep error window pops up on entering valid credentials in Dose Alert window of any clinical user who has not logged in at least once after resetting password/user creation. | BB_WIP2 | BB_WIP5 | No workaround | Subsystem | User Authentication | NA | Fixed | Prakash_Chand_Yadav | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP5_SubCom | Brilliance Big Bore |  | BB_WIP5 | March 14, 2019 5:26:00 PM GMT+08:00
CTUDT00600921 | Refresh | Closed | Revital_Cohen | December 18, 2018 4:35:15 PM GMT+08:00 | Refresh - Iqon - Zeffective images are not reconstructed | WIP11.1 |  |  | Subsystem | Spectral Recon |  | Works As Specified / Not A Defect | tamari_o | 1 |  |  | Support-System | NO | Closed | WIP14SegReq | IQon1
CTUDT00600918 | Refresh | Closed | blessily | December 18, 2018 2:59:26 PM GMT+08:00 | BB Refresh : BugRep error thrown when we try to copy partial raw data | BB_WIP2 | BB_WIP5.1 |  | Subsystem | Copy Raw | None | Fixed | Prakash_Chand_Yadav | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP5_SegReq | Brilliance Big Bore |  | BB_WIP5.1 | March 20, 2019 7:20:06 PM GMT+08:00
CTUDT00600910 | Refresh | Closed | roshan_r | December 17, 2018 6:58:55 PM GMT+08:00 | Refresh : “Reference Noise” level is not available in result show all and preferences. | WIP11.1 | DHF318051 Rev 13 HOST Verification Test Protocol | NA | Subsystem | Exam Cards | NA | Fixed | ramesh_t | 0 | 2 | Test Design | Console4 | NO | Closed | WIP14SegReq | Ovation |  | DHF318051 Rev 13 HOST Verification Test Protocol | April 23, 2019 7:51:25 PM GMT+08:00
CTUDT00600910 | Refresh | Closed | roshan_r | December 17, 2018 6:58:55 PM GMT+08:00 | Refresh : “Reference Noise” level is not available in result show all and preferences. | WIP11.1 | DHF318051 Rev 13 HOST Verification Test Protocol | NA | Subsystem | Exam Cards | NA | Fixed | ramesh_t | 0 | 2 | Test Design | Console4 | NO | Closed | WIP14SegReq | Brilliance 64 |  | DHF318051 Rev 13 HOST Verification Test Protocol | April 23, 2019 7:51:25 PM GMT+08:00
CTUDT00600910 | Refresh | Closed | roshan_r | December 17, 2018 6:58:55 PM GMT+08:00 | Refresh : “Reference Noise” level is not available in result show all and preferences. | WIP11.1 | DHF318051 Rev 13 HOST Verification Test Protocol | NA | Subsystem | Exam Cards | NA | Fixed | ramesh_t | 0 | 2 | Test Design | Console4 | NO | Closed | WIP14SegReq | Ingenuity 64 |  | DHF318051 Rev 13 HOST Verification Test Protocol | April 23, 2019 7:51:25 PM GMT+08:00
CTUDT00600910 | Refresh | Closed | roshan_r | December 17, 2018 6:58:55 PM GMT+08:00 | Refresh : “Reference Noise” level is not available in result show all and preferences. | WIP11.1 | DHF318051 Rev 13 HOST Verification Test Protocol | NA | Subsystem | Exam Cards | NA | Fixed | ramesh_t | 0 | 2 | Test Design | Console4 | NO | Closed | WIP14SegReq | IQon1 |  | DHF318051 Rev 13 HOST Verification Test Protocol | April 23, 2019 7:51:25 PM GMT+08:00
CTUDT00600908 | Refresh | Closed | wang_l | December 17, 2018 5:47:02 PM GMT+08:00 | BB Refresh: The dropdown for scan parameter "mAs" changed to not editable after surview scan finished | BB_WIP2 |  |  | Subsystem |  | None | Works As Specified / Not A Defect | patra_r | 0 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00600906 | Refresh | Closed | wang_l | December 17, 2018 5:39:46 PM GMT+08:00 | BB Refresh: blank desktop was displayed once login CT and IT account for the first time | BB_WIP2 | BB_WIP5 | This issue occur only first time after fresh installation. On 2nd login onward the issue does not occur. | Subsystem | Login | NA | Fixed | kapadia_v | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP4_SegReq | Brilliance Big Bore |  | BB_WIP5 | February 19, 2019 8:07:08 PM GMT+08:00
CTUDT00600904 | Refresh | Closed | wang_l | December 17, 2018 5:30:25 PM GMT+08:00 | BB Refresh: Pulmo scan failed | BB_WIP2 |  |  | Subsystem |  | None | Duplicate | patra_r | 0 |  |  | Console4 | NO | Closed | BB_WIP4_SubCom | Brilliance Big Bore
CTUDT00600902 | Refresh | Closed | wang_l | December 17, 2018 5:22:40 PM GMT+08:00 | BB Refresh: Cardiac scan failed | BB_WIP2 |  |  | Subsystem |  | None | Duplicate | patra_r | 0 |  |  | Console4 | NO | Closed | BB_WIP4_SubCom | Brilliance Big Bore
CTUDT00600900 | Refresh | Closed | wang_l | December 17, 2018 5:05:07 PM GMT+08:00 | BB Refresh: View2 tab was not displayed once click End Exam and Continue Working | BB_WIP2 |  | Click PD, then click view2 or click View2 button directly for View2 to be displayed. | Subsystem | Offline Recon | N/A | Duplicate | s_sa | 0 | 4 |  | iCAP | NO | Closed | SubNC | Brilliance Big Bore
CTUDT00600898 | Refresh | Closed | wang_l | December 17, 2018 4:41:55 PM GMT+08:00 | BB Refresh: Head scan failed with error "System error. Please end exam and retry." | BB_WIP2 |  |  | Subsystem |  | n/a | Duplicate | frider_o | 0 |  |  | DMS | YES | Closed |  | Brilliance Big Bore
CTUDT00600895 | Refresh | Closed | zhang_t | December 17, 2018 4:06:59 PM GMT+08:00 | Refresh: Bookmarks were missing in right region after re-login, but displaying in study series list in PD, Bug report collector popup while trying to load bookmarks into Quick Review | WIP11.1 |  | 1. If the study is copied to other folder, the bookmarks are listed back and are intact in Bookmark Manager. 2. Try to rebuild the local database, the bookmark comes back to bookmark manager. | Subsystem | Bookmark - IFU | NA | No Plan to Fix | srushti_k | 0 | 3 |  | iCAP | NO | Closed |  | Ovation
CTUDT00600895 | Refresh | Closed | zhang_t | December 17, 2018 4:06:59 PM GMT+08:00 | Refresh: Bookmarks were missing in right region after re-login, but displaying in study series list in PD, Bug report collector popup while trying to load bookmarks into Quick Review | WIP11.1 |  | 1. If the study is copied to other folder, the bookmarks are listed back and are intact in Bookmark Manager. 2. Try to rebuild the local database, the bookmark comes back to bookmark manager. | Subsystem | Bookmark - IFU | NA | No Plan to Fix | srushti_k | 0 | 3 |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00600895 | Refresh | Closed | zhang_t | December 17, 2018 4:06:59 PM GMT+08:00 | Refresh: Bookmarks were missing in right region after re-login, but displaying in study series list in PD, Bug report collector popup while trying to load bookmarks into Quick Review | WIP11.1 |  | 1. If the study is copied to other folder, the bookmarks are listed back and are intact in Bookmark Manager. 2. Try to rebuild the local database, the bookmark comes back to bookmark manager. | Subsystem | Bookmark - IFU | NA | No Plan to Fix | srushti_k | 0 | 3 |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00600895 | Refresh | Closed | zhang_t | December 17, 2018 4:06:59 PM GMT+08:00 | Refresh: Bookmarks were missing in right region after re-login, but displaying in study series list in PD, Bug report collector popup while trying to load bookmarks into Quick Review | WIP11.1 |  | 1. If the study is copied to other folder, the bookmarks are listed back and are intact in Bookmark Manager. 2. Try to rebuild the local database, the bookmark comes back to bookmark manager. | Subsystem | Bookmark - IFU | NA | No Plan to Fix | srushti_k | 0 | 3 |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00600895 | Refresh | Closed | zhang_t | December 17, 2018 4:06:59 PM GMT+08:00 | Refresh: Bookmarks were missing in right region after re-login, but displaying in study series list in PD, Bug report collector popup while trying to load bookmarks into Quick Review | WIP11.1 |  | 1. If the study is copied to other folder, the bookmarks are listed back and are intact in Bookmark Manager. 2. Try to rebuild the local database, the bookmark comes back to bookmark manager. | Subsystem | Bookmark - IFU | NA | No Plan to Fix | srushti_k | 0 | 3 |  | iCAP | NO | Closed |  | IQon1
CTUDT00600893 | Refresh | Closed | zhang_t | December 17, 2018 3:57:13 PM GMT+08:00 | Refresh: Protocol Group displayed blank in Exam summary | WIP11.1 | WIP16 |  | Subsystem | Chinese DICOM | NA | Fixed | bhawna_g | 0 | 4 | Console Software | Console4 | NO | Closed | WIP14SegReq | Ovation |  | WIP16 | April 19, 2019 7:58:10 PM GMT+08:00
CTUDT00600893 | Refresh | Closed | zhang_t | December 17, 2018 3:57:13 PM GMT+08:00 | Refresh: Protocol Group displayed blank in Exam summary | WIP11.1 | WIP16 |  | Subsystem | Chinese DICOM | NA | Fixed | bhawna_g | 0 | 4 | Console Software | Console4 | NO | Closed | WIP14SegReq | Brilliance 64 |  | WIP16 | April 19, 2019 7:58:10 PM GMT+08:00
CTUDT00600893 | Refresh | Closed | zhang_t | December 17, 2018 3:57:13 PM GMT+08:00 | Refresh: Protocol Group displayed blank in Exam summary | WIP11.1 | WIP16 |  | Subsystem | Chinese DICOM | NA | Fixed | bhawna_g | 0 | 4 | Console Software | Console4 | NO | Closed | WIP14SegReq | Brilliance Big Bore |  | WIP16 | April 19, 2019 7:58:10 PM GMT+08:00
CTUDT00600893 | Refresh | Closed | zhang_t | December 17, 2018 3:57:13 PM GMT+08:00 | Refresh: Protocol Group displayed blank in Exam summary | WIP11.1 | WIP16 |  | Subsystem | Chinese DICOM | NA | Fixed | bhawna_g | 0 | 4 | Console Software | Console4 | NO | Closed | WIP14SegReq | Ingenuity 64 |  | WIP16 | April 19, 2019 7:58:10 PM GMT+08:00
CTUDT00600893 | Refresh | Closed | zhang_t | December 17, 2018 3:57:13 PM GMT+08:00 | Refresh: Protocol Group displayed blank in Exam summary | WIP11.1 | WIP16 |  | Subsystem | Chinese DICOM | NA | Fixed | bhawna_g | 0 | 4 | Console Software | Console4 | NO | Closed | WIP14SegReq | IQon1 |  | WIP16 | April 19, 2019 7:58:10 PM GMT+08:00
CTUDT00600891 | Refresh | Closed | zhang_t | December 17, 2018 2:56:58 PM GMT+08:00 | Refresh: Cannot fill Chinese characters in some user controls when IM is Chinese Language | WIP11.1 | WIP20 | 1. Select the Chinese keyboard from preference => regional settings. 2. Perform logoff & login two times. 3. Select the Chinese keyboard by using WIN+SPACE. 4. Press SHIFT once before typing anything in Chinese. This is needed because, Microsoft Pinyin Input profile has two input languages(Chinese and English). To toggle we need use SHIFT key. | Subsystem | IFU;Verified1 | NA | Fixed | charles_j | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | Before SVAL | Brilliance 64 |  | WIP20 | August 27, 2019 1:13:31 PM GMT+08:00
CTUDT00600891 | Refresh | Closed | zhang_t | December 17, 2018 2:56:58 PM GMT+08:00 | Refresh: Cannot fill Chinese characters in some user controls when IM is Chinese Language | WIP11.1 | WIP20 | 1. Select the Chinese keyboard from preference => regional settings. 2. Perform logoff & login two times. 3. Select the Chinese keyboard by using WIN+SPACE. 4. Press SHIFT once before typing anything in Chinese. This is needed because, Microsoft Pinyin Input profile has two input languages(Chinese and English). To toggle we need use SHIFT key. | Subsystem | IFU;Verified1 | NA | Fixed | charles_j | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | Before SVAL | Brilliance Big Bore |  | WIP20 | August 27, 2019 1:13:31 PM GMT+08:00
CTUDT00600891 | Refresh | Closed | zhang_t | December 17, 2018 2:56:58 PM GMT+08:00 | Refresh: Cannot fill Chinese characters in some user controls when IM is Chinese Language | WIP11.1 | WIP20 | 1. Select the Chinese keyboard from preference => regional settings. 2. Perform logoff & login two times. 3. Select the Chinese keyboard by using WIN+SPACE. 4. Press SHIFT once before typing anything in Chinese. This is needed because, Microsoft Pinyin Input profile has two input languages(Chinese and English). To toggle we need use SHIFT key. | Subsystem | IFU;Verified1 | NA | Fixed | charles_j | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | Before SVAL | Ingenuity 64 |  | WIP20 | August 27, 2019 1:13:31 PM GMT+08:00
CTUDT00600888 | Refresh | Closed | zhang_t | December 17, 2018 2:33:37 PM GMT+08:00 | Refresh: some user controls in UI cannot well accept Chinese Characters as input | WIP11.1 | WIP18 |  | Subsystem | Chinese DICOM | N/A | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | YES | Closed | WIP14SubCom | Ovation |  | WIP18 | October 3, 2019 2:49:14 PM GMT+08:00
CTUDT00600888 | Refresh | Closed | zhang_t | December 17, 2018 2:33:37 PM GMT+08:00 | Refresh: some user controls in UI cannot well accept Chinese Characters as input | WIP11.1 | WIP18 |  | Subsystem | Chinese DICOM | N/A | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | YES | Closed | WIP14SubCom | Brilliance 64 |  | WIP18 | October 3, 2019 2:49:14 PM GMT+08:00
CTUDT00600888 | Refresh | Closed | zhang_t | December 17, 2018 2:33:37 PM GMT+08:00 | Refresh: some user controls in UI cannot well accept Chinese Characters as input | WIP11.1 | WIP18 |  | Subsystem | Chinese DICOM | N/A | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | YES | Closed | WIP14SubCom | Ingenuity 64 |  | WIP18 | October 3, 2019 2:49:14 PM GMT+08:00
CTUDT00600888 | Refresh | Closed | zhang_t | December 17, 2018 2:33:37 PM GMT+08:00 | Refresh: some user controls in UI cannot well accept Chinese Characters as input | WIP11.1 | WIP18 |  | Subsystem | Chinese DICOM | N/A | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | YES | Closed | WIP14SubCom | IQon1 |  | WIP18 | October 3, 2019 2:49:14 PM GMT+08:00
CTUDT00600885 | Refresh | Closed | finestine_m | December 17, 2018 3:25:15 AM GMT+08:00 | Refresh: Calibrations: No Data in 'DMS/DataPath Diagnostics' in Service Tools | WIP11.1 | WIP14 | No workaround | Subsystem | Service Tools | DMS-SRS-658 | Fixed | franco_a | 0 | 2 | DMS Software | DMS | NO | Closed | WIP14SegReq | Brilliance 64 |  | WIP14 | February 24, 2019 7:24:48 PM GMT+08:00
CTUDT00600885 | Refresh | Closed | finestine_m | December 17, 2018 3:25:15 AM GMT+08:00 | Refresh: Calibrations: No Data in 'DMS/DataPath Diagnostics' in Service Tools | WIP11.1 | WIP14 | No workaround | Subsystem | Service Tools | DMS-SRS-658 | Fixed | franco_a | 0 | 2 | DMS Software | DMS | NO | Closed | WIP14SegReq | Ingenuity 64 |  | WIP14 | February 24, 2019 7:24:48 PM GMT+08:00
CTUDT00600881 | Refresh | Closed | shaham_t | December 16, 2018 5:09:39 PM GMT+08:00 | v4.1.7 service documents instruct the FSE to install the Telnet server. | 4.7.7 |  |  | Subsystem | Document | N/A | Does Not Apply | hilas | 0 | 2 | Documentation - N/A | Support-System | YES | Closed | WIP20SubCom | Ingenuity 64 | PR ID 6855415 |  | August 12, 2019 4:45:31 PM GMT+08:00
CTUDT00600881 | Refresh | Closed | shaham_t | December 16, 2018 5:09:39 PM GMT+08:00 | v4.1.7 service documents instruct the FSE to install the Telnet server. | 4.7.7 |  |  | Subsystem | Document | N/A | Does Not Apply | hilas | 0 | 2 | Documentation - N/A | Support-System | YES | Closed | WIP20SubCom | IQon1 | PR ID 6855415 |  | August 12, 2019 4:45:31 PM GMT+08:00
CTUDT00600866 | Refresh | Closed | manju_kolkar | December 13, 2018 6:24:06 PM GMT+08:00 | Refresh : SMPTE Image for Overall Image Quality is not matching with the reference image document in both DIN and Standard monitor. | WIP11.1 | WIP14 | No workaround. | Subsystem | DIN monitor | None | Fixed | Prakash_Chand_Yadav | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SubCom | Ovation |  | WIP14 | February 18, 2019 8:30:45 PM GMT+08:00
CTUDT00600866 | Refresh | Closed | manju_kolkar | December 13, 2018 6:24:06 PM GMT+08:00 | Refresh : SMPTE Image for Overall Image Quality is not matching with the reference image document in both DIN and Standard monitor. | WIP11.1 | WIP14 | No workaround. | Subsystem | DIN monitor | None | Fixed | Prakash_Chand_Yadav | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SubCom | Brilliance 64 |  | WIP14 | February 18, 2019 8:30:45 PM GMT+08:00
CTUDT00600866 | Refresh | Closed | manju_kolkar | December 13, 2018 6:24:06 PM GMT+08:00 | Refresh : SMPTE Image for Overall Image Quality is not matching with the reference image document in both DIN and Standard monitor. | WIP11.1 | WIP14 | No workaround. | Subsystem | DIN monitor | None | Fixed | Prakash_Chand_Yadav | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SubCom | Ingenuity 64 |  | WIP14 | February 18, 2019 8:30:45 PM GMT+08:00
CTUDT00600866 | Refresh | Closed | manju_kolkar | December 13, 2018 6:24:06 PM GMT+08:00 | Refresh : SMPTE Image for Overall Image Quality is not matching with the reference image document in both DIN and Standard monitor. | WIP11.1 | WIP14 | No workaround. | Subsystem | DIN monitor | None | Fixed | Prakash_Chand_Yadav | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SubCom | IQon1 |  | WIP14 | February 18, 2019 8:30:45 PM GMT+08:00
CTUDT00600862 | Refresh | Closed | Noy_Seelig | December 13, 2018 4:13:02 PM GMT+08:00 | Refresh: During Performance test Physics Quality Assurance error appear | WIP11.1 |  |  | Subsystem | Console Infrastructure | n/a | No Plan to Fix | redrouthu_r | 0 | 4 |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00600860 | Refresh | Closed | Noy_Seelig | December 13, 2018 3:27:04 PM GMT+08:00 | Refresh: No ROIs appear on the images in Quick Review and CT Viewer | WIP11.1 |  |  | Subsystem |  |  | Duplicate | frider_o | 0 |  |  | DMS | NO | Closed |  | Ingenuity 64
CTUDT00600857 | Refresh | Closed | zhang_t | December 13, 2018 11:16:53 AM GMT+08:00 | Refresh: several places where the Philips service password is hardcoded (service_only) on Host | WIP11.1 | WIP13 |  | Subsystem | Hardcoded passwords | N\A | Fixed | kaushik_u | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP15SegReq | Ovation |  | WIP13 | January 31, 2019 6:25:22 PM GMT+08:00
CTUDT00600857 | Refresh | Closed | zhang_t | December 13, 2018 11:16:53 AM GMT+08:00 | Refresh: several places where the Philips service password is hardcoded (service_only) on Host | WIP11.1 | WIP13 |  | Subsystem | Hardcoded passwords | N\A | Fixed | kaushik_u | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP15SegReq | Brilliance 64 |  | WIP13 | January 31, 2019 6:25:22 PM GMT+08:00
CTUDT00600857 | Refresh | Closed | zhang_t | December 13, 2018 11:16:53 AM GMT+08:00 | Refresh: several places where the Philips service password is hardcoded (service_only) on Host | WIP11.1 | WIP13 |  | Subsystem | Hardcoded passwords | N\A | Fixed | kaushik_u | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP15SegReq | Brilliance Big Bore |  | WIP13 | January 31, 2019 6:25:22 PM GMT+08:00
CTUDT00600857 | Refresh | Closed | zhang_t | December 13, 2018 11:16:53 AM GMT+08:00 | Refresh: several places where the Philips service password is hardcoded (service_only) on Host | WIP11.1 | WIP13 |  | Subsystem | Hardcoded passwords | N\A | Fixed | kaushik_u | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP15SegReq | Ingenuity 64 |  | WIP13 | January 31, 2019 6:25:22 PM GMT+08:00
CTUDT00600857 | Refresh | Closed | zhang_t | December 13, 2018 11:16:53 AM GMT+08:00 | Refresh: several places where the Philips service password is hardcoded (service_only) on Host | WIP11.1 | WIP13 |  | Subsystem | Hardcoded passwords | N\A | Fixed | kaushik_u | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP15SegReq | IQon1 |  | WIP13 | January 31, 2019 6:25:22 PM GMT+08:00
CTUDT00600853 | Refresh | Closed | wang_l | December 13, 2018 9:48:59 AM GMT+08:00 | BB-Refresh: Air calibration failed for CT user with message "Error using <Scan> Activity. Query I depend params table ’Current’ failed.” | BB_WIP2 | WIP15 |  | Subsystem | Air Calibrations | n/a | Fixed | zhou_x | 0 | 2 | Gantry Calibrations | GCX | YES | Closed | BB_WIP3_SegReq | Brilliance Big Bore |  | WIP15 | April 11, 2019 3:51:00 PM GMT+08:00
CTUDT00600851 | Refresh | Closed | Valeria_Kashtanov | December 12, 2018 10:49:21 PM GMT+08:00 | Refresh: Common Functions: Inconsistent behavior of temperature notification in different tests/scanners | WIP11 | N/A - Test Design | Test Design of 232510 should be updated. | Subsystem | Calibrations | N/A | Fixed | gehtin_e | 0 | 2 | Test Design | DMS | NO | Closed | WIP13SegReq | Ovation |  | N/A - Test Design | February 3, 2019 7:37:48 PM GMT+08:00
CTUDT00600851 | Refresh | Closed | Valeria_Kashtanov | December 12, 2018 10:49:21 PM GMT+08:00 | Refresh: Common Functions: Inconsistent behavior of temperature notification in different tests/scanners | WIP11 | N/A - Test Design | Test Design of 232510 should be updated. | Subsystem | Calibrations | N/A | Fixed | gehtin_e | 0 | 2 | Test Design | DMS | NO | Closed | WIP13SegReq | IQon1 |  | N/A - Test Design | February 3, 2019 7:37:48 PM GMT+08:00
CTUDT00600845 | Refresh | Closed | ding_f | December 12, 2018 3:30:41 PM GMT+08:00 | Scan was hang up during Phantom calibration & Hcor slice, Adult. | WIP11.1 |  |  | Subsystem | Calibrations | n/a | Duplicate | altshuller_mark | 0 |  |  | DMS | NO | Closed | WIP16SegReq | Ingenuity 64
CTUDT00600843 | Refresh | Closed | mN_m | December 12, 2018 11:53:25 AM GMT+08:00 | Refresh: “Recon Queue Manager” is still accessible when the connection between the host and CIRS is lost. | WIP10 | Host SSRS DHF303211 Rev 08 |  | System | Reconstruction | Console-SRS.9948 | Fixed | bandil_j | 0 | 4 | Host SSRS | Console4 | NO | Closed | WIP15SegReq | Ovation |  | Host SSRS DHF303211 Rev 08 | March 22, 2019 6:06:53 PM GMT+08:00
CTUDT00600843 | Refresh | Closed | mN_m | December 12, 2018 11:53:25 AM GMT+08:00 | Refresh: “Recon Queue Manager” is still accessible when the connection between the host and CIRS is lost. | WIP10 | Host SSRS DHF303211 Rev 08 |  | System | Reconstruction | Console-SRS.9948 | Fixed | bandil_j | 0 | 4 | Host SSRS | Console4 | NO | Closed | WIP15SegReq | Brilliance 64 |  | Host SSRS DHF303211 Rev 08 | March 22, 2019 6:06:53 PM GMT+08:00
CTUDT00600843 | Refresh | Closed | mN_m | December 12, 2018 11:53:25 AM GMT+08:00 | Refresh: “Recon Queue Manager” is still accessible when the connection between the host and CIRS is lost. | WIP10 | Host SSRS DHF303211 Rev 08 |  | System | Reconstruction | Console-SRS.9948 | Fixed | bandil_j | 0 | 4 | Host SSRS | Console4 | NO | Closed | WIP15SegReq | Ingenuity 64 |  | Host SSRS DHF303211 Rev 08 | March 22, 2019 6:06:53 PM GMT+08:00
CTUDT00600843 | Refresh | Closed | mN_m | December 12, 2018 11:53:25 AM GMT+08:00 | Refresh: “Recon Queue Manager” is still accessible when the connection between the host and CIRS is lost. | WIP10 | Host SSRS DHF303211 Rev 08 |  | System | Reconstruction | Console-SRS.9948 | Fixed | bandil_j | 0 | 4 | Host SSRS | Console4 | NO | Closed | WIP15SegReq | IQon1 |  | Host SSRS DHF303211 Rev 08 | March 22, 2019 6:06:53 PM GMT+08:00
CTUDT00600826 | Refresh | Closed | swack_g | December 11, 2018 6:51:32 PM GMT+08:00 | Refresh: Calibrations: ACR Test: Distance from pin geometrical center to the ROI center is more than 3mm | WIP11 | WIP14 |  | Subsystem |  | n/a | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP14SegReq | Ovation |  | WIP14 | February 13, 2019 7:06:09 PM GMT+08:00
CTUDT00600826 | Refresh | Closed | swack_g | December 11, 2018 6:51:32 PM GMT+08:00 | Refresh: Calibrations: ACR Test: Distance from pin geometrical center to the ROI center is more than 3mm | WIP11 | WIP14 |  | Subsystem |  | n/a | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP14SegReq | Brilliance Big Bore |  | WIP14 | February 13, 2019 7:06:09 PM GMT+08:00
CTUDT00600826 | Refresh | Closed | swack_g | December 11, 2018 6:51:32 PM GMT+08:00 | Refresh: Calibrations: ACR Test: Distance from pin geometrical center to the ROI center is more than 3mm | WIP11 | WIP14 |  | Subsystem |  | n/a | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP14SegReq | Ingenuity 64 |  | WIP14 | February 13, 2019 7:06:09 PM GMT+08:00
CTUDT00600826 | Refresh | Closed | swack_g | December 11, 2018 6:51:32 PM GMT+08:00 | Refresh: Calibrations: ACR Test: Distance from pin geometrical center to the ROI center is more than 3mm | WIP11 | WIP14 |  | Subsystem |  | n/a | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP14SegReq | IQon1 |  | WIP14 | February 13, 2019 7:06:09 PM GMT+08:00
CTUDT00600816 | Refresh | Closed | Valeria_Kashtanov | December 10, 2018 10:03:23 PM GMT+08:00 | Refresh: Stability: XRayFilter parameter in Stability scan is not as defined in Spec | WIP11 | WIP13 | N/A | Subsystem |  | N/A | Fixed | yosefk | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP13SegReq | Ovation |  | WIP13 | January 20, 2019 4:05:24 PM GMT+08:00
CTUDT00600815 | Refresh | Closed | attali_s | December 10, 2018 9:54:37 PM GMT+08:00 | Refresh: FEE Test: In RnD Advanced Mode, default parameters are scanned instead of chosen parameters | WIP11 | WIP13 | The RnD user can override in sequence XML files | Subsystem | Calibrations | n/a | Fixed | bikman | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | Ovation |  | WIP13 | January 13, 2019 3:51:01 PM GMT+08:00
CTUDT00600815 | Refresh | Closed | attali_s | December 10, 2018 9:54:37 PM GMT+08:00 | Refresh: FEE Test: In RnD Advanced Mode, default parameters are scanned instead of chosen parameters | WIP11 | WIP13 | The RnD user can override in sequence XML files | Subsystem | Calibrations | n/a | Fixed | bikman | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | Brilliance Big Bore |  | WIP13 | January 13, 2019 3:51:01 PM GMT+08:00
CTUDT00600815 | Refresh | Closed | attali_s | December 10, 2018 9:54:37 PM GMT+08:00 | Refresh: FEE Test: In RnD Advanced Mode, default parameters are scanned instead of chosen parameters | WIP11 | WIP13 | The RnD user can override in sequence XML files | Subsystem | Calibrations | n/a | Fixed | bikman | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | Ingenuity 64 |  | WIP13 | January 13, 2019 3:51:01 PM GMT+08:00
CTUDT00600815 | Refresh | Closed | attali_s | December 10, 2018 9:54:37 PM GMT+08:00 | Refresh: FEE Test: In RnD Advanced Mode, default parameters are scanned instead of chosen parameters | WIP11 | WIP13 | The RnD user can override in sequence XML files | Subsystem | Calibrations | n/a | Fixed | bikman | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | IQon1 |  | WIP13 | January 13, 2019 3:51:01 PM GMT+08:00
CTUDT00600813 | Refresh | Closed | jotawar_d | December 10, 2018 8:43:08 PM GMT+08:00 | Refresh_Suggested Exam Cards group had more than one linking | 4.8 | BB_WIP4 | The user is able to change the name or parameters of the EC and resave it | Subsystem | Exam Card | N/A | Fixed | prasad_r | 0 | 3 | Console Software | Console4 | YES | Closed | BB_WIP4_SubCom | Brilliance Big Bore |  | BB_WIP4 | February 4, 2019 9:43:23 PM GMT+08:00
CTUDT00600809 | Refresh | Closed | or_shmuel | December 10, 2018 7:57:52 PM GMT+08:00 | Refresh, IQon - Cannot restore UserAcounts | WIP11.1 | WIP14 | Yes, user can create required accounts as and when he wants using IT tool. | Subsystem | Restore User Accounts | NA | Fixed | ravindra_a | 0 | 3 | Console Software | Console4 | NO | Closed | WIP13SubCom | Ovation |  | WIP14 | March 6, 2019 7:43:51 PM GMT+08:00
CTUDT00600809 | Refresh | Closed | or_shmuel | December 10, 2018 7:57:52 PM GMT+08:00 | Refresh, IQon - Cannot restore UserAcounts | WIP11.1 | WIP14 | Yes, user can create required accounts as and when he wants using IT tool. | Subsystem | Restore User Accounts | NA | Fixed | ravindra_a | 0 | 3 | Console Software | Console4 | NO | Closed | WIP13SubCom | Brilliance 64 |  | WIP14 | March 6, 2019 7:43:51 PM GMT+08:00
CTUDT00600809 | Refresh | Closed | or_shmuel | December 10, 2018 7:57:52 PM GMT+08:00 | Refresh, IQon - Cannot restore UserAcounts | WIP11.1 | WIP14 | Yes, user can create required accounts as and when he wants using IT tool. | Subsystem | Restore User Accounts | NA | Fixed | ravindra_a | 0 | 3 | Console Software | Console4 | NO | Closed | WIP13SubCom | Ingenuity 64 |  | WIP14 | March 6, 2019 7:43:51 PM GMT+08:00
CTUDT00600809 | Refresh | Closed | or_shmuel | December 10, 2018 7:57:52 PM GMT+08:00 | Refresh, IQon - Cannot restore UserAcounts | WIP11.1 | WIP14 | Yes, user can create required accounts as and when he wants using IT tool. | Subsystem | Restore User Accounts | NA | Fixed | ravindra_a | 0 | 3 | Console Software | Console4 | NO | Closed | WIP13SubCom | IQon1 |  | WIP14 | March 6, 2019 7:43:51 PM GMT+08:00
CTUDT00600806 | Refresh | Closed | pB_s | December 10, 2018 7:36:40 PM GMT+08:00 | Able to check the "Account Disabled" checkbox for all Hospital Administrator users including logged in IT user | WIP11.1 | WIP13 | By again clicking on any other checkbox the UI issue gets resolved. | Subsystem | User Management | NA | Fixed | ravindra_a | 0 | 4 | Console Software | Console4 | NO | Closed | WIP13SubCom | Ovation |  | WIP13 | January 16, 2019 9:23:46 PM GMT+08:00
CTUDT00600806 | Refresh | Closed | pB_s | December 10, 2018 7:36:40 PM GMT+08:00 | Able to check the "Account Disabled" checkbox for all Hospital Administrator users including logged in IT user | WIP11.1 | WIP13 | By again clicking on any other checkbox the UI issue gets resolved. | Subsystem | User Management | NA | Fixed | ravindra_a | 0 | 4 | Console Software | Console4 | NO | Closed | WIP13SubCom | Brilliance 64 |  | WIP13 | January 16, 2019 9:23:46 PM GMT+08:00
CTUDT00600806 | Refresh | Closed | pB_s | December 10, 2018 7:36:40 PM GMT+08:00 | Able to check the "Account Disabled" checkbox for all Hospital Administrator users including logged in IT user | WIP11.1 | WIP13 | By again clicking on any other checkbox the UI issue gets resolved. | Subsystem | User Management | NA | Fixed | ravindra_a | 0 | 4 | Console Software | Console4 | NO | Closed | WIP13SubCom | Brilliance Big Bore |  | WIP13 | January 16, 2019 9:23:46 PM GMT+08:00
CTUDT00600806 | Refresh | Closed | pB_s | December 10, 2018 7:36:40 PM GMT+08:00 | Able to check the "Account Disabled" checkbox for all Hospital Administrator users including logged in IT user | WIP11.1 | WIP13 | By again clicking on any other checkbox the UI issue gets resolved. | Subsystem | User Management | NA | Fixed | ravindra_a | 0 | 4 | Console Software | Console4 | NO | Closed | WIP13SubCom | Ingenuity 64 |  | WIP13 | January 16, 2019 9:23:46 PM GMT+08:00
CTUDT00600806 | Refresh | Closed | pB_s | December 10, 2018 7:36:40 PM GMT+08:00 | Able to check the "Account Disabled" checkbox for all Hospital Administrator users including logged in IT user | WIP11.1 | WIP13 | By again clicking on any other checkbox the UI issue gets resolved. | Subsystem | User Management | NA | Fixed | ravindra_a | 0 | 4 | Console Software | Console4 | NO | Closed | WIP13SubCom | IQon1 |  | WIP13 | January 16, 2019 9:23:46 PM GMT+08:00
CTUDT00600801 | Refresh | Closed | or_shmuel | December 10, 2018 5:23:30 PM GMT+08:00 | Refresh, IQon - Cannot fill Chinese characters in mandatory fields | WIP11.1 |  | Suggested method to use Chinese keyboard will be : 1. Select the keyboard from preference and login/logoff two times. 2. Press SHIFT before typing anything in Chinese. This is needed because, Microsoft Pinyin Input profile has two input languages(Chinese and English). To toggle we need use SHIFT key. | Subsystem | Chinese keyboard | NA | Duplicate | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP17SegReq | Ovation |  |  | May 27, 2019 1:54:47 PM GMT+08:00
CTUDT00600801 | Refresh | Closed | or_shmuel | December 10, 2018 5:23:30 PM GMT+08:00 | Refresh, IQon - Cannot fill Chinese characters in mandatory fields | WIP11.1 |  | Suggested method to use Chinese keyboard will be : 1. Select the keyboard from preference and login/logoff two times. 2. Press SHIFT before typing anything in Chinese. This is needed because, Microsoft Pinyin Input profile has two input languages(Chinese and English). To toggle we need use SHIFT key. | Subsystem | Chinese keyboard | NA | Duplicate | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP17SegReq | Brilliance 64 |  |  | May 27, 2019 1:54:47 PM GMT+08:00
CTUDT00600801 | Refresh | Closed | or_shmuel | December 10, 2018 5:23:30 PM GMT+08:00 | Refresh, IQon - Cannot fill Chinese characters in mandatory fields | WIP11.1 |  | Suggested method to use Chinese keyboard will be : 1. Select the keyboard from preference and login/logoff two times. 2. Press SHIFT before typing anything in Chinese. This is needed because, Microsoft Pinyin Input profile has two input languages(Chinese and English). To toggle we need use SHIFT key. | Subsystem | Chinese keyboard | NA | Duplicate | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP17SegReq | Brilliance Big Bore |  |  | May 27, 2019 1:54:47 PM GMT+08:00
CTUDT00600801 | Refresh | Closed | or_shmuel | December 10, 2018 5:23:30 PM GMT+08:00 | Refresh, IQon - Cannot fill Chinese characters in mandatory fields | WIP11.1 |  | Suggested method to use Chinese keyboard will be : 1. Select the keyboard from preference and login/logoff two times. 2. Press SHIFT before typing anything in Chinese. This is needed because, Microsoft Pinyin Input profile has two input languages(Chinese and English). To toggle we need use SHIFT key. | Subsystem | Chinese keyboard | NA | Duplicate | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP17SegReq | Ingenuity 64 |  |  | May 27, 2019 1:54:47 PM GMT+08:00
CTUDT00600801 | Refresh | Closed | or_shmuel | December 10, 2018 5:23:30 PM GMT+08:00 | Refresh, IQon - Cannot fill Chinese characters in mandatory fields | WIP11.1 |  | Suggested method to use Chinese keyboard will be : 1. Select the keyboard from preference and login/logoff two times. 2. Press SHIFT before typing anything in Chinese. This is needed because, Microsoft Pinyin Input profile has two input languages(Chinese and English). To toggle we need use SHIFT key. | Subsystem | Chinese keyboard | NA | Duplicate | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP17SegReq | IQon1 |  |  | May 27, 2019 1:54:47 PM GMT+08:00
CTUDT00600798 | Refresh | Closed | or_shmuel | December 10, 2018 5:04:39 PM GMT+08:00 | Refresh, IQon - After Changing the keyboard language, the previous language remain in the language options | WIP11.1 |  |  | Subsystem | Keyboard Language | NA | Works As Specified / Not A Defect | s_sa | 0 |  |  | iCAP | NO | Closed | WIP15SegReq | IQon1
CTUDT00600793 | Refresh | Closed | Dana_Kamyshev | December 10, 2018 3:25:29 PM GMT+08:00 | Refresh - Wrong reconstruction time in acquisition tab | WIP11.1 | WIP16 |  | Subsystem | Patient Directory (UI) | NA | Fixed | jose_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SegReq | Ovation |  | WIP16 | April 9, 2019 12:02:40 PM GMT+08:00
CTUDT00600793 | Refresh | Closed | Dana_Kamyshev | December 10, 2018 3:25:29 PM GMT+08:00 | Refresh - Wrong reconstruction time in acquisition tab | WIP11.1 | WIP16 |  | Subsystem | Patient Directory (UI) | NA | Fixed | jose_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SegReq | Brilliance 64 |  | WIP16 | April 9, 2019 12:02:40 PM GMT+08:00
CTUDT00600793 | Refresh | Closed | Dana_Kamyshev | December 10, 2018 3:25:29 PM GMT+08:00 | Refresh - Wrong reconstruction time in acquisition tab | WIP11.1 | WIP16 |  | Subsystem | Patient Directory (UI) | NA | Fixed | jose_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SegReq | Ingenuity 64 |  | WIP16 | April 9, 2019 12:02:40 PM GMT+08:00
CTUDT00600793 | Refresh | Closed | Dana_Kamyshev | December 10, 2018 3:25:29 PM GMT+08:00 | Refresh - Wrong reconstruction time in acquisition tab | WIP11.1 | WIP16 |  | Subsystem | Patient Directory (UI) | NA | Fixed | jose_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SegReq | IQon1 |  | WIP16 | April 9, 2019 12:02:40 PM GMT+08:00
CTUDT00600788 | Refresh | Closed | cohen_y | December 9, 2018 11:02:20 PM GMT+08:00 | Refresh: Backup/Restore Tool: The Backup/Restore DMS Profile xml file is taken from the wrong branch. | WIP11 | WIP15 |  | System | Backup/Restore;Verified1 | N/A | Fixed | venugopal_s | 0 | 2 | DMS Software | DI | NO | Closed | WIP14SegReq | Ovation |  | WIP15 | March 11, 2019 4:19:56 PM GMT+08:00
CTUDT00600788 | Refresh | Closed | cohen_y | December 9, 2018 11:02:20 PM GMT+08:00 | Refresh: Backup/Restore Tool: The Backup/Restore DMS Profile xml file is taken from the wrong branch. | WIP11 | WIP15 |  | System | Backup/Restore;Verified1 | N/A | Fixed | venugopal_s | 0 | 2 | DMS Software | DI | NO | Closed | WIP14SegReq | Brilliance 64 |  | WIP15 | March 11, 2019 4:19:56 PM GMT+08:00
CTUDT00600788 | Refresh | Closed | cohen_y | December 9, 2018 11:02:20 PM GMT+08:00 | Refresh: Backup/Restore Tool: The Backup/Restore DMS Profile xml file is taken from the wrong branch. | WIP11 | WIP15 |  | System | Backup/Restore;Verified1 | N/A | Fixed | venugopal_s | 0 | 2 | DMS Software | DI | NO | Closed | WIP14SegReq | Brilliance Big Bore |  | WIP15 | March 11, 2019 4:19:56 PM GMT+08:00
CTUDT00600788 | Refresh | Closed | cohen_y | December 9, 2018 11:02:20 PM GMT+08:00 | Refresh: Backup/Restore Tool: The Backup/Restore DMS Profile xml file is taken from the wrong branch. | WIP11 | WIP15 |  | System | Backup/Restore;Verified1 | N/A | Fixed | venugopal_s | 0 | 2 | DMS Software | DI | NO | Closed | WIP14SegReq | Ingenuity 64 |  | WIP15 | March 11, 2019 4:19:56 PM GMT+08:00
CTUDT00600788 | Refresh | Closed | cohen_y | December 9, 2018 11:02:20 PM GMT+08:00 | Refresh: Backup/Restore Tool: The Backup/Restore DMS Profile xml file is taken from the wrong branch. | WIP11 | WIP15 |  | System | Backup/Restore;Verified1 | N/A | Fixed | venugopal_s | 0 | 2 | DMS Software | DI | NO | Closed | WIP14SegReq | IQon1 |  | WIP15 | March 11, 2019 4:19:56 PM GMT+08:00
CTUDT00600785 | Refresh | Closed | beck_d | December 9, 2018 2:27:47 PM GMT+08:00 | Refresh: Calibrations: FEE Test: High/Low indication does not appear in the Table of failures for Bias/Noise | WIP11 | WIP13 | N/A | Subsystem |  | N/A | Fixed | bikman | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | IQon1 |  | WIP13 | January 13, 2019 3:55:17 PM GMT+08:00
CTUDT00600782 | Refresh | Closed | nimma_p | December 7, 2018 7:34:44 PM GMT+08:00 | Refresh-IQON--No warning message appeared when disk space was between60 to 80GB. | WIP11.1 |  |  | Subsystem |  | NA | Invalid/Uncontrolled Configuration | v_sumanth | 0 |  |  | Console4 | NO | Closed | WIP14SegReq | IQon1
CTUDT00600781 | Refresh | Closed | nimma_p | December 7, 2018 7:28:12 PM GMT+08:00 | Refresh-IQON--No warning message appeared when disk space was less than 60GB.. | WIP11.1 |  |  | Subsystem |  | NA | Invalid/Uncontrolled Configuration | v_sumanth | 0 |  |  | Console4 | NO | Closed | WIP14SegReq | IQon1
CTUDT00600779 | Refresh | Closed | hiremath_b | December 7, 2018 4:25:10 PM GMT+08:00 | Refresh: Cancel option should NOT be available in change password confirmation window in Login screen. | WIP11.1 | WIP15 | N/A | Subsystem | User management | N/A | Fixed | kapadia_v | 0 | 3 | Console Software | Console4 | NO | Closed | WIP14SegReq | Ovation |  | WIP15 | March 7, 2019 7:17:18 PM GMT+08:00
CTUDT00600779 | Refresh | Closed | hiremath_b | December 7, 2018 4:25:10 PM GMT+08:00 | Refresh: Cancel option should NOT be available in change password confirmation window in Login screen. | WIP11.1 | WIP15 | N/A | Subsystem | User management | N/A | Fixed | kapadia_v | 0 | 3 | Console Software | Console4 | NO | Closed | WIP14SegReq | Brilliance 64 |  | WIP15 | March 7, 2019 7:17:18 PM GMT+08:00
CTUDT00600779 | Refresh | Closed | hiremath_b | December 7, 2018 4:25:10 PM GMT+08:00 | Refresh: Cancel option should NOT be available in change password confirmation window in Login screen. | WIP11.1 | WIP15 | N/A | Subsystem | User management | N/A | Fixed | kapadia_v | 0 | 3 | Console Software | Console4 | NO | Closed | WIP14SegReq | Brilliance Big Bore |  | WIP15 | March 7, 2019 7:17:18 PM GMT+08:00
CTUDT00600779 | Refresh | Closed | hiremath_b | December 7, 2018 4:25:10 PM GMT+08:00 | Refresh: Cancel option should NOT be available in change password confirmation window in Login screen. | WIP11.1 | WIP15 | N/A | Subsystem | User management | N/A | Fixed | kapadia_v | 0 | 3 | Console Software | Console4 | NO | Closed | WIP14SegReq | Ingenuity 64 |  | WIP15 | March 7, 2019 7:17:18 PM GMT+08:00
CTUDT00600779 | Refresh | Closed | hiremath_b | December 7, 2018 4:25:10 PM GMT+08:00 | Refresh: Cancel option should NOT be available in change password confirmation window in Login screen. | WIP11.1 | WIP15 | N/A | Subsystem | User management | N/A | Fixed | kapadia_v | 0 | 3 | Console Software | Console4 | NO | Closed | WIP14SegReq | IQon1 |  | WIP15 | March 7, 2019 7:17:18 PM GMT+08:00
CTUDT00600776 | Refresh | Closed | hiremath_b | December 7, 2018 3:08:46 PM GMT+08:00 | Refresh:Messages #412 and #413 display when switching between ACS and Manual mA control in ECM. As per requirement messages should display only during planning. | WIP11.1 | WIP13 |  | Subsystem | ECM | N/A | Fixed | prasad_r | 0 | 4 | Console Software | Console4 | NO | Closed | WIP13SubCom | Ovation |  | WIP13 | January 16, 2019 3:10:15 PM GMT+08:00
CTUDT00600776 | Refresh | Closed | hiremath_b | December 7, 2018 3:08:46 PM GMT+08:00 | Refresh:Messages #412 and #413 display when switching between ACS and Manual mA control in ECM. As per requirement messages should display only during planning. | WIP11.1 | WIP13 |  | Subsystem | ECM | N/A | Fixed | prasad_r | 0 | 4 | Console Software | Console4 | NO | Closed | WIP13SubCom | Brilliance 64 |  | WIP13 | January 16, 2019 3:10:15 PM GMT+08:00
CTUDT00600776 | Refresh | Closed | hiremath_b | December 7, 2018 3:08:46 PM GMT+08:00 | Refresh:Messages #412 and #413 display when switching between ACS and Manual mA control in ECM. As per requirement messages should display only during planning. | WIP11.1 | WIP13 |  | Subsystem | ECM | N/A | Fixed | prasad_r | 0 | 4 | Console Software | Console4 | NO | Closed | WIP13SubCom | Brilliance Big Bore |  | WIP13 | January 16, 2019 3:10:15 PM GMT+08:00
CTUDT00600776 | Refresh | Closed | hiremath_b | December 7, 2018 3:08:46 PM GMT+08:00 | Refresh:Messages #412 and #413 display when switching between ACS and Manual mA control in ECM. As per requirement messages should display only during planning. | WIP11.1 | WIP13 |  | Subsystem | ECM | N/A | Fixed | prasad_r | 0 | 4 | Console Software | Console4 | NO | Closed | WIP13SubCom | Ingenuity 64 |  | WIP13 | January 16, 2019 3:10:15 PM GMT+08:00
CTUDT00600776 | Refresh | Closed | hiremath_b | December 7, 2018 3:08:46 PM GMT+08:00 | Refresh:Messages #412 and #413 display when switching between ACS and Manual mA control in ECM. As per requirement messages should display only during planning. | WIP11.1 | WIP13 |  | Subsystem | ECM | N/A | Fixed | prasad_r | 0 | 4 | Console Software | Console4 | NO | Closed | WIP13SubCom | IQon1 |  | WIP13 | January 16, 2019 3:10:15 PM GMT+08:00
CTUDT00600774 | Refresh | Closed | divya_n | December 7, 2018 2:20:00 PM GMT+08:00 | Login page is not displayed after restart of SUT | WIP11 | WIP13 | On hard reboot the login page appeared. | Subsystem | Login | NA | Fixed | kapadia_v | 0 | 3 | Console Software | Console4 | NO | Closed | WIP13SubCom | Ovation |  | WIP13 | January 29, 2019 6:26:21 PM GMT+08:00
CTUDT00600774 | Refresh | Closed | divya_n | December 7, 2018 2:20:00 PM GMT+08:00 | Login page is not displayed after restart of SUT | WIP11 | WIP13 | On hard reboot the login page appeared. | Subsystem | Login | NA | Fixed | kapadia_v | 0 | 3 | Console Software | Console4 | NO | Closed | WIP13SubCom | Brilliance 64 |  | WIP13 | January 29, 2019 6:26:21 PM GMT+08:00
CTUDT00600774 | Refresh | Closed | divya_n | December 7, 2018 2:20:00 PM GMT+08:00 | Login page is not displayed after restart of SUT | WIP11 | WIP13 | On hard reboot the login page appeared. | Subsystem | Login | NA | Fixed | kapadia_v | 0 | 3 | Console Software | Console4 | NO | Closed | WIP13SubCom | Brilliance Big Bore |  | WIP13 | January 29, 2019 6:26:21 PM GMT+08:00
CTUDT00600774 | Refresh | Closed | divya_n | December 7, 2018 2:20:00 PM GMT+08:00 | Login page is not displayed after restart of SUT | WIP11 | WIP13 | On hard reboot the login page appeared. | Subsystem | Login | NA | Fixed | kapadia_v | 0 | 3 | Console Software | Console4 | NO | Closed | WIP13SubCom | Ingenuity 64 |  | WIP13 | January 29, 2019 6:26:21 PM GMT+08:00
CTUDT00600774 | Refresh | Closed | divya_n | December 7, 2018 2:20:00 PM GMT+08:00 | Login page is not displayed after restart of SUT | WIP11 | WIP13 | On hard reboot the login page appeared. | Subsystem | Login | NA | Fixed | kapadia_v | 0 | 3 | Console Software | Console4 | NO | Closed | WIP13SubCom | IQon1 |  | WIP13 | January 29, 2019 6:26:21 PM GMT+08:00
CTUDT00600772 | Refresh | Closed | mN_m | December 7, 2018 11:12:22 AM GMT+08:00 | BB Refresh: Study crashes on loading user created EC (Chest + Resp 4D CT) to plan. | BB_WIP2 | BB_WIP3 |  | Subsystem |  | NA | Fixed | sowmya_k | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP3_SegReq | Brilliance Big Bore |  | BB_WIP3 | January 11, 2019 7:03:03 PM GMT+08:00
CTUDT00600768 | Refresh | Closed | mN_m | December 6, 2018 8:17:01 PM GMT+08:00 | BB Refresh: Screen goes off for a sec and come back when you login as service user | BB_WIP2 | BB_WIP5 |  | Subsystem | Login | NA | Fixed | kapadia_v | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP4_SegReq | Brilliance Big Bore |  | BB_WIP5 | February 19, 2019 8:07:44 PM GMT+08:00
CTUDT00600766 | Refresh | Closed | mN_m | December 6, 2018 7:30:17 PM GMT+08:00 | BB Refresh: Performance issue: Taking 4 to 5 minutes to load an EC to plan and sometimes it get hang. | BB_WIP2 | BB_WIP5 |  | Subsystem | Planning | N/A | Fixed | madamsetti_k | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP3_SegReq | Brilliance Big Bore |  | BB_WIP5 | March 15, 2019 8:12:46 PM GMT+08:00
CTUDT00600764 | Refresh | Closed | mN_m | December 6, 2018 7:14:09 PM GMT+08:00 | BB Refresh: Cardiac scan can be performed without ECG signal Open ECG\Open Resp links on the scan ruler is not working for one of the scenario. | BB_WIP2 |  |  | Subsystem | Planning | NA | Does Not Apply | babu_s | 1 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00600761 | Refresh | Closed | mN_m | December 6, 2018 6:51:05 PM GMT+08:00 | BB Refresh: “Unexpected error during workflow creation” critical error displayed on launching the workflow in service mode after the fresh installation. | BB_WIP2 |  |  | Subsystem | Workflow |  | Duplicate | karkour | 0 |  |  | GCX | NO | Closed |  | Brilliance Big Bore
CTUDT00600759 | Refresh | Closed | mN_m | December 6, 2018 6:44:54 PM GMT+08:00 | BB Refresh: When logged in as emergency user, username is missing on the image titles of online and offline result images. | BB_WIP2 | BB_WIP4 |  | Subsystem |  | N/A | Fixed | Prakash_Chand_Yadav | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP4_SubCom | Brilliance Big Bore |  | BB_WIP4 | February 6, 2019 2:18:11 PM GMT+08:00
CTUDT00600757 | Refresh | Closed | mN_m | December 6, 2018 6:39:53 PM GMT+08:00 | BB Refresh: “PmsWindowsPriniting.exe has stopped working” error occurred on logging in as CT user and Emergency user | BB_WIP2 | WIP4 | No. | Subsystem |  | NA | Fixed | kapadia_v | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP4_SubCom | Brilliance Big Bore |  | WIP4 | February 4, 2019 4:37:53 PM GMT+08:00
CTUDT00600746 | Refresh | Closed | navale_s | December 6, 2018 1:50:49 PM GMT+08:00 | Refresh- Wrong version appears for Console.SW.Host in System Configuration | 4.8 | BB_WIP3 |  | Subsystem |  | N/A | Fixed | navale_s | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | BB_WIP3_SegReq | Brilliance Big Bore |  | BB_WIP3 | December 19, 2018 4:54:48 PM GMT+08:00
CTUDT00600745 | Refresh | Closed | ding_s | December 6, 2018 1:34:29 PM GMT+08:00 | Performance Test failed with calculation errors | WIP11.1 |  |  | Subsystem | Performance |  | Duplicate | elviral | 0 |  |  | DMS | NO | Closed |  | Ingenuity 64
CTUDT00600743 | Refresh | Closed | ding_s | December 6, 2018 11:12:23 AM GMT+08:00 | Calibration abort due to high noise levels | WIP11.1 |  |  | Subsystem | Calibrations | n/a | Works As Specified / Not A Defect | ding_f | 1 |  |  | Support-System | YES | Closed |  | Ingenuity 64
CTUDT00600739 | Refresh | Closed | kollarits_s | December 5, 2018 11:12:48 PM GMT+08:00 | Per CAPA PR7437217: Refresh IFUs need to be updated to include a NOTE to check the intercom volume prior to a patient exam | N/A | WIP16 |  | Subsystem | Documentation | N/A | Fixed | rajalakshmi_j | 0 | 2 | User Documentation | Support-System | YES | Closed | WIP14SegReq | Ovation |  | WIP16 | May 29, 2019 2:05:29 PM GMT+08:00
CTUDT00600739 | Refresh | Closed | kollarits_s | December 5, 2018 11:12:48 PM GMT+08:00 | Per CAPA PR7437217: Refresh IFUs need to be updated to include a NOTE to check the intercom volume prior to a patient exam | N/A | WIP16 |  | Subsystem | Documentation | N/A | Fixed | rajalakshmi_j | 0 | 2 | User Documentation | Support-System | YES | Closed | WIP14SegReq | Brilliance Big Bore |  | WIP16 | May 29, 2019 2:05:29 PM GMT+08:00
CTUDT00600739 | Refresh | Closed | kollarits_s | December 5, 2018 11:12:48 PM GMT+08:00 | Per CAPA PR7437217: Refresh IFUs need to be updated to include a NOTE to check the intercom volume prior to a patient exam | N/A | WIP16 |  | Subsystem | Documentation | N/A | Fixed | rajalakshmi_j | 0 | 2 | User Documentation | Support-System | YES | Closed | WIP14SegReq | Ingenuity 64 |  | WIP16 | May 29, 2019 2:05:29 PM GMT+08:00
CTUDT00600739 | Refresh | Closed | kollarits_s | December 5, 2018 11:12:48 PM GMT+08:00 | Per CAPA PR7437217: Refresh IFUs need to be updated to include a NOTE to check the intercom volume prior to a patient exam | N/A | WIP16 |  | Subsystem | Documentation | N/A | Fixed | rajalakshmi_j | 0 | 2 | User Documentation | Support-System | YES | Closed | WIP14SegReq | IQon1 |  | WIP16 | May 29, 2019 2:05:29 PM GMT+08:00
CTUDT00600731 | Refresh | Closed | jotawar_d | December 5, 2018 5:57:40 PM GMT+08:00 | Refresh: Truncated information in Acquisition viewer and word Phantom appears twice in show all | 4.8 | BB_WIP4 |  | Subsystem | Acq Viewer | Console SRS 15962 | Fixed | prasad_r | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP4_SubCom | Brilliance Big Bore |  | BB_WIP4 | February 6, 2019 3:12:43 PM GMT+08:00
CTUDT00600729 | Refresh | Closed | jotawar_d | December 5, 2018 5:47:24 PM GMT+08:00 | Exam cards are getting deleted after performing copy paste operation and relaunching the ECM | 4.8 | BB_WIP3 | performing restore with the backup/restore tool to restore the deleted Exam Cards | System |  | N/A | Fixed | prasad_r | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP3_SegReq | Brilliance Big Bore | 9467908 | BB_WIP3 | January 9, 2019 1:45:53 PM GMT+08:00
CTUDT00600725 | Refresh | Closed | katy_rubin | December 5, 2018 5:25:42 PM GMT+08:00 | Refresh: The Exam cards Brain Perfusion Jog and Brain Perfusion Non-Jog appear with a space in the end | WIP11.1 | WIP17 | A possible workaround is to change the string which calls to these exam cards in the automation test by adding a space to the exam card name. | System | Exam cards | N/A | Fixed | kr_m | 0 | 3 | Console Software | Console4 | NO | Closed | WIP17SubCom | IQon1 |  | WIP17 | May 23, 2019 7:28:03 PM GMT+08:00
CTUDT00600720 | Refresh | Closed | datta_i | December 5, 2018 3:02:48 PM GMT+08:00 | The message in log viewer is not reasonable if group is exceed the limitation | 4.8 | BB_WIP3 |  | Subsystem |  | NA | Fixed | kumar_v | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP3_SegReq | Brilliance Big Bore |  | BB_WIP3 | December 11, 2018 5:58:05 PM GMT+08:00
CTUDT00600715 | Refresh | Closed | golzban_v | December 4, 2018 11:06:59 PM GMT+08:00 | Refresh: Not saved "review before save" result is deleted from series list along with reconstructed and saved result | WIP11 | WIP14 |  | Subsystem | Review before Save | NA | Fixed | sowmya_k | 0 | 3 | Console Software | Console4 | NO | Closed | WIP14SegReq | Ovation |  | WIP14 | February 13, 2019 12:05:51 PM GMT+08:00
CTUDT00600715 | Refresh | Closed | golzban_v | December 4, 2018 11:06:59 PM GMT+08:00 | Refresh: Not saved "review before save" result is deleted from series list along with reconstructed and saved result | WIP11 | WIP14 |  | Subsystem | Review before Save | NA | Fixed | sowmya_k | 0 | 3 | Console Software | Console4 | NO | Closed | WIP14SegReq | Brilliance 64 |  | WIP14 | February 13, 2019 12:05:51 PM GMT+08:00
CTUDT00600715 | Refresh | Closed | golzban_v | December 4, 2018 11:06:59 PM GMT+08:00 | Refresh: Not saved "review before save" result is deleted from series list along with reconstructed and saved result | WIP11 | WIP14 |  | Subsystem | Review before Save | NA | Fixed | sowmya_k | 0 | 3 | Console Software | Console4 | NO | Closed | WIP14SegReq | Ingenuity 64 |  | WIP14 | February 13, 2019 12:05:51 PM GMT+08:00
CTUDT00600715 | Refresh | Closed | golzban_v | December 4, 2018 11:06:59 PM GMT+08:00 | Refresh: Not saved "review before save" result is deleted from series list along with reconstructed and saved result | WIP11 | WIP14 |  | Subsystem | Review before Save | NA | Fixed | sowmya_k | 0 | 3 | Console Software | Console4 | NO | Closed | WIP14SegReq | IQon1 |  | WIP14 | February 13, 2019 12:05:51 PM GMT+08:00
CTUDT00600710 | Refresh | Closed | navale_s | December 4, 2018 8:15:21 PM GMT+08:00 | Refresh:Worklist-In Local Schedule DB “Accession Number” column is not displayed by default | 4.8 | BB_WIP3 | the Accession Number column can be added manually | Subsystem |  | N/A | Fixed | chakraborty_d | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP3_SegReq | Brilliance Big Bore |  | BB_WIP3 | December 13, 2018 12:40:18 PM GMT+08:00
CTUDT00600709 | Refresh | Closed | jotawar_d | December 4, 2018 8:04:38 PM GMT+08:00 | Cut operation of review before save result crashed the exam application | 4.2 | BB_WIP3 | The user is able to perform an offline reconstruction for the results. | Subsystem |  | Console-SRS.9081 | Fixed | prasad_r | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP2_SubCom | Brilliance Big Bore | CFR#9470203 | BB_WIP3 | January 9, 2019 1:43:17 PM GMT+08:00
CTUDT00600706 | Refresh | Closed | mN_m | December 4, 2018 6:48:16 PM GMT+08:00 | BB Refresh: On Pressing Go for clinical scan, “System Error, Please End Exam and Retry” error displayed for head and ear group exam cards. | BB_WIP2 | BB_WIP3 |  | Subsystem |  | n/a | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | YES | Closed | BB_WIP3_SegReq | Brilliance Big Bore |  | BB_WIP3 | December 24, 2018 7:47:33 PM GMT+08:00
CTUDT00600702 | Refresh | Closed | mN_m | December 4, 2018 6:31:03 PM GMT+08:00 | BB Refresh: ECG and Pulmo signal is not detected in ECG viewer and Resp Viewer respectively. | BB_WIP2 | BB_WIP4 |  | Subsystem |  | None | Fixed | patra_r | 0 | 2 | Console Software | Console4 | YES | Closed | BB_WIP4_SubCom | Brilliance Big Bore |  | BB_WIP4 | February 5, 2019 6:57:38 PM GMT+08:00
CTUDT00600700 | Refresh | Closed | finestine_m | December 4, 2018 5:39:28 PM GMT+08:00 | Refresh: Stability Test: Time Interval Between Each Temp Sample Is Not Equal To The Spec | WIP11 | N/A - Test Design |  | Subsystem |  | n/a | Fixed | gehtin_e | 0 | 2 | Test Design | DMS | NO | Closed | WIP14SegReq | Ovation |  | N/A - Test Design | February 3, 2019 7:38:26 PM GMT+08:00
CTUDT00600700 | Refresh | Closed | finestine_m | December 4, 2018 5:39:28 PM GMT+08:00 | Refresh: Stability Test: Time Interval Between Each Temp Sample Is Not Equal To The Spec | WIP11 | N/A - Test Design |  | Subsystem |  | n/a | Fixed | gehtin_e | 0 | 2 | Test Design | DMS | NO | Closed | WIP14SegReq | Ingenuity 64 |  | N/A - Test Design | February 3, 2019 7:38:26 PM GMT+08:00
CTUDT00600698 | Refresh | Closed | datta_i | December 4, 2018 2:34:38 PM GMT+08:00 | Surview present in Exam summary series is not of the current patient. | V4.1.6 | BB_WIP2 |  | Subsystem |  | NA | Fixed | datta_i | 0 | 2 | Console-Console4 | Console4 | NO | Closed | BB_WIP3_SegReq | Brilliance Big Bore | PR7252944 ESC# 0109806936 ESC# 0109673916 | BB_WIP2 | December 13, 2018 8:33:34 PM GMT+08:00
CTUDT00600695 | Refresh | Closed | beck_d | December 3, 2018 11:03:50 PM GMT+08:00 | Refresh: Calibrations: Performance Test: Number of scans is not according to spec | WIP11 | WIP13 | N/A | System |  | N/A | Fixed | elviral | 0 | 4 | DMS Software | DMS | NO | Closed | WIP13SegReq | Ovation |  | WIP13 | January 8, 2019 8:13:12 PM GMT+08:00
CTUDT00600695 | Refresh | Closed | beck_d | December 3, 2018 11:03:50 PM GMT+08:00 | Refresh: Calibrations: Performance Test: Number of scans is not according to spec | WIP11 | WIP13 | N/A | System |  | N/A | Fixed | elviral | 0 | 4 | DMS Software | DMS | NO | Closed | WIP13SegReq | Ingenuity 64 |  | WIP13 | January 8, 2019 8:13:12 PM GMT+08:00
CTUDT00600693 | Refresh | Closed | beck_d | December 3, 2018 10:45:59 PM GMT+08:00 | Refresh: Calibrations: Performance Test: Scans order is not according to Spec | WIP11 | WIP13 |  | Subsystem |  | n/a | Fixed | elviral | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | Ovation |  | WIP13 | January 6, 2019 4:50:37 PM GMT+08:00
CTUDT00600693 | Refresh | Closed | beck_d | December 3, 2018 10:45:59 PM GMT+08:00 | Refresh: Calibrations: Performance Test: Scans order is not according to Spec | WIP11 | WIP13 |  | Subsystem |  | n/a | Fixed | elviral | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | Brilliance 64 |  | WIP13 | January 6, 2019 4:50:37 PM GMT+08:00
CTUDT00600693 | Refresh | Closed | beck_d | December 3, 2018 10:45:59 PM GMT+08:00 | Refresh: Calibrations: Performance Test: Scans order is not according to Spec | WIP11 | WIP13 |  | Subsystem |  | n/a | Fixed | elviral | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | Ingenuity 64 |  | WIP13 | January 6, 2019 4:50:37 PM GMT+08:00
CTUDT00600691 | Refresh | Closed | finestine_m | December 3, 2018 10:05:09 PM GMT+08:00 | Refresh: Shadow Test: XrayFilter has wrong value ("TiTef" instead of "TiHalf") | WIP11 | WIP13 |  | Subsystem |  | N/A | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP13SegReq | Ovation |  | WIP13 | January 3, 2019 9:33:23 PM GMT+08:00
CTUDT00600689 | Refresh | Closed | beck_d | December 3, 2018 9:50:28 PM GMT+08:00 | Refresh: Calibrations: N/A is not noted under measured data in the report when a scan fails due to gantry | WIP11 | N/A |  | Subsystem | Calibrations | N/A | Fixed | braunstein_d | 0 | 2 | Design Documentation | DMS | NO | Closed | WIP14SegReq | Ovation |  | N/A | February 20, 2019 7:23:02 PM GMT+08:00
CTUDT00600689 | Refresh | Closed | beck_d | December 3, 2018 9:50:28 PM GMT+08:00 | Refresh: Calibrations: N/A is not noted under measured data in the report when a scan fails due to gantry | WIP11 | N/A |  | Subsystem | Calibrations | N/A | Fixed | braunstein_d | 0 | 2 | Design Documentation | DMS | NO | Closed | WIP14SegReq | Brilliance Big Bore |  | N/A | February 20, 2019 7:23:02 PM GMT+08:00
CTUDT00600689 | Refresh | Closed | beck_d | December 3, 2018 9:50:28 PM GMT+08:00 | Refresh: Calibrations: N/A is not noted under measured data in the report when a scan fails due to gantry | WIP11 | N/A |  | Subsystem | Calibrations | N/A | Fixed | braunstein_d | 0 | 2 | Design Documentation | DMS | NO | Closed | WIP14SegReq | Ingenuity 64 |  | N/A | February 20, 2019 7:23:02 PM GMT+08:00
CTUDT00600689 | Refresh | Closed | beck_d | December 3, 2018 9:50:28 PM GMT+08:00 | Refresh: Calibrations: N/A is not noted under measured data in the report when a scan fails due to gantry | WIP11 | N/A |  | Subsystem | Calibrations | N/A | Fixed | braunstein_d | 0 | 2 | Design Documentation | DMS | NO | Closed | WIP14SegReq | IQon1 |  | N/A | February 20, 2019 7:23:02 PM GMT+08:00
CTUDT00600688 | Refresh | Closed | datta_i | December 3, 2018 7:51:52 PM GMT+08:00 | Cannot modify "display" value if pause one CCT scan then duplicate the un-completed CCT scan | 4.8 | BB_WIP5.1 |  | Subsystem |  | NA | Fixed | datta_i | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP3_SegReq | Brilliance Big Bore |  | BB_WIP5.1 | March 26, 2019 6:21:04 PM GMT+08:00
CTUDT00600677 | Refresh | Closed | beck_d | December 3, 2018 5:53:38 PM GMT+08:00 | Refresh: Calibrations: User can start an IQ test even though the User Input fields are empty | WIP11 | WIP13 WIP15 |  | Subsystem | Calibrations (IQ test) | N/A | Fixed | bikman | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | Ovation |  | WIP13 WIP15 | March 28, 2019 7:23:42 PM GMT+08:00
CTUDT00600677 | Refresh | Closed | beck_d | December 3, 2018 5:53:38 PM GMT+08:00 | Refresh: Calibrations: User can start an IQ test even though the User Input fields are empty | WIP11 | WIP13 WIP15 |  | Subsystem | Calibrations (IQ test) | N/A | Fixed | bikman | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | Brilliance Big Bore |  | WIP13 WIP15 | March 28, 2019 7:23:42 PM GMT+08:00
CTUDT00600677 | Refresh | Closed | beck_d | December 3, 2018 5:53:38 PM GMT+08:00 | Refresh: Calibrations: User can start an IQ test even though the User Input fields are empty | WIP11 | WIP13 WIP15 |  | Subsystem | Calibrations (IQ test) | N/A | Fixed | bikman | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | Ingenuity 64 |  | WIP13 WIP15 | March 28, 2019 7:23:42 PM GMT+08:00
CTUDT00600677 | Refresh | Closed | beck_d | December 3, 2018 5:53:38 PM GMT+08:00 | Refresh: Calibrations: User can start an IQ test even though the User Input fields are empty | WIP11 | WIP13 WIP15 |  | Subsystem | Calibrations (IQ test) | N/A | Fixed | bikman | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | IQon1 |  | WIP13 WIP15 | March 28, 2019 7:23:42 PM GMT+08:00
CTUDT00600671 | Refresh | Closed | wang_l | December 3, 2018 2:46:17 PM GMT+08:00 | Can not open saved xml file in universal log viewer | WIP11.1 | WIP14 |  | Subsystem | Log Viewer | N/A | Fixed | zhang_l | 0 | 3 | Test Design | Support-System | NO | Closed | WIP14SegReq | Ovation |  | WIP14 | June 3, 2019 6:03:40 PM GMT+08:00
CTUDT00600671 | Refresh | Closed | wang_l | December 3, 2018 2:46:17 PM GMT+08:00 | Can not open saved xml file in universal log viewer | WIP11.1 | WIP14 |  | Subsystem | Log Viewer | N/A | Fixed | zhang_l | 0 | 3 | Test Design | Support-System | NO | Closed | WIP14SegReq | Brilliance 64 |  | WIP14 | June 3, 2019 6:03:40 PM GMT+08:00
CTUDT00600671 | Refresh | Closed | wang_l | December 3, 2018 2:46:17 PM GMT+08:00 | Can not open saved xml file in universal log viewer | WIP11.1 | WIP14 |  | Subsystem | Log Viewer | N/A | Fixed | zhang_l | 0 | 3 | Test Design | Support-System | NO | Closed | WIP14SegReq | Brilliance Big Bore |  | WIP14 | June 3, 2019 6:03:40 PM GMT+08:00
CTUDT00600671 | Refresh | Closed | wang_l | December 3, 2018 2:46:17 PM GMT+08:00 | Can not open saved xml file in universal log viewer | WIP11.1 | WIP14 |  | Subsystem | Log Viewer | N/A | Fixed | zhang_l | 0 | 3 | Test Design | Support-System | NO | Closed | WIP14SegReq | Ingenuity 64 |  | WIP14 | June 3, 2019 6:03:40 PM GMT+08:00
CTUDT00600671 | Refresh | Closed | wang_l | December 3, 2018 2:46:17 PM GMT+08:00 | Can not open saved xml file in universal log viewer | WIP11.1 | WIP14 |  | Subsystem | Log Viewer | N/A | Fixed | zhang_l | 0 | 3 | Test Design | Support-System | NO | Closed | WIP14SegReq | IQon1 |  | WIP14 | June 3, 2019 6:03:40 PM GMT+08:00
CTUDT00600670 | Refresh | Closed | goswami_s | December 3, 2018 2:41:02 PM GMT+08:00 | Requirement update for IMR: Result plan box and FOV is not aligned when matrix is changed | WIP 9.1 | Host SSRS DHF303211 Rev 06 |  | Subsystem |  | Console-SRS.8958 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | WIP12SubCom | Ovation |  | Host SSRS DHF303211 Rev 06 | December 12, 2018 10:26:07 PM GMT+08:00
CTUDT00600670 | Refresh | Closed | goswami_s | December 3, 2018 2:41:02 PM GMT+08:00 | Requirement update for IMR: Result plan box and FOV is not aligned when matrix is changed | WIP 9.1 | Host SSRS DHF303211 Rev 06 |  | Subsystem |  | Console-SRS.8958 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | WIP12SubCom | Brilliance 64 |  | Host SSRS DHF303211 Rev 06 | December 12, 2018 10:26:07 PM GMT+08:00
CTUDT00600670 | Refresh | Closed | goswami_s | December 3, 2018 2:41:02 PM GMT+08:00 | Requirement update for IMR: Result plan box and FOV is not aligned when matrix is changed | WIP 9.1 | Host SSRS DHF303211 Rev 06 |  | Subsystem |  | Console-SRS.8958 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | WIP12SubCom | Brilliance Big Bore |  | Host SSRS DHF303211 Rev 06 | December 12, 2018 10:26:07 PM GMT+08:00
CTUDT00600670 | Refresh | Closed | goswami_s | December 3, 2018 2:41:02 PM GMT+08:00 | Requirement update for IMR: Result plan box and FOV is not aligned when matrix is changed | WIP 9.1 | Host SSRS DHF303211 Rev 06 |  | Subsystem |  | Console-SRS.8958 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | WIP12SubCom | Ingenuity 64 |  | Host SSRS DHF303211 Rev 06 | December 12, 2018 10:26:07 PM GMT+08:00
CTUDT00600670 | Refresh | Closed | goswami_s | December 3, 2018 2:41:02 PM GMT+08:00 | Requirement update for IMR: Result plan box and FOV is not aligned when matrix is changed | WIP 9.1 | Host SSRS DHF303211 Rev 06 |  | Subsystem |  | Console-SRS.8958 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | WIP12SubCom | IQon1 |  | Host SSRS DHF303211 Rev 06 | December 12, 2018 10:26:07 PM GMT+08:00
CTUDT00600666 | Refresh | Closed | wang_l | December 3, 2018 9:53:52 AM GMT+08:00 | Backup and restore behavior incorrect for one click installation on Refresh Ingenuity | WIP11.1 |  |  | Subsystem | 1-Click Installation | N/A | Documentation Created/Updated | wang_l | 1 | 2 | Console Software | Support-System | NO | Closed | WIP14SegReq | Ovation |  |  | November 7, 2019 11:52:15 AM GMT+08:00
CTUDT00600666 | Refresh | Closed | wang_l | December 3, 2018 9:53:52 AM GMT+08:00 | Backup and restore behavior incorrect for one click installation on Refresh Ingenuity | WIP11.1 |  |  | Subsystem | 1-Click Installation | N/A | Documentation Created/Updated | wang_l | 1 | 2 | Console Software | Support-System | NO | Closed | WIP14SegReq | Brilliance 64 |  |  | November 7, 2019 11:52:15 AM GMT+08:00
CTUDT00600666 | Refresh | Closed | wang_l | December 3, 2018 9:53:52 AM GMT+08:00 | Backup and restore behavior incorrect for one click installation on Refresh Ingenuity | WIP11.1 |  |  | Subsystem | 1-Click Installation | N/A | Documentation Created/Updated | wang_l | 1 | 2 | Console Software | Support-System | NO | Closed | WIP14SegReq | Brilliance Big Bore |  |  | November 7, 2019 11:52:15 AM GMT+08:00
CTUDT00600666 | Refresh | Closed | wang_l | December 3, 2018 9:53:52 AM GMT+08:00 | Backup and restore behavior incorrect for one click installation on Refresh Ingenuity | WIP11.1 |  |  | Subsystem | 1-Click Installation | N/A | Documentation Created/Updated | wang_l | 1 | 2 | Console Software | Support-System | NO | Closed | WIP14SegReq | Ingenuity 64 |  |  | November 7, 2019 11:52:15 AM GMT+08:00
CTUDT00600666 | Refresh | Closed | wang_l | December 3, 2018 9:53:52 AM GMT+08:00 | Backup and restore behavior incorrect for one click installation on Refresh Ingenuity | WIP11.1 |  |  | Subsystem | 1-Click Installation | N/A | Documentation Created/Updated | wang_l | 1 | 2 | Console Software | Support-System | NO | Closed | WIP14SegReq | IQon1 |  |  | November 7, 2019 11:52:15 AM GMT+08:00
CTUDT00600663 | Refresh | Closed | Dana_Kamyshev | December 2, 2018 5:27:59 PM GMT+08:00 | Refresh - CIRS Event logs are missing in FSE bugrep | WIP11.1 | WIP18 |  | Subsystem | BugRep | N/A | Fixed | s_shabeer | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP12SegReq | IQon1 |  | WIP18 | June 26, 2019 2:31:06 PM GMT+08:00
CTUDT00600653 | Refresh | Closed | sowmya_k | November 30, 2018 5:14:02 PM GMT+08:00 | During Resp Single phase, a wrong message is displayed and the scan is not allowed. | BB_WIP2 | BB_WIP3 | User should press 'Start Leaning' and wait till the learning is complete to initiate the scan. | Subsystem | Resp4D CT | Console-SRS.16093 | Fixed | sowmya_k | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP3_SubCom | Brilliance Big Bore |  | BB_WIP3 | December 11, 2018 7:13:21 PM GMT+08:00
CTUDT00600652 | Refresh | Closed | zhou_x | November 30, 2018 4:40:44 PM GMT+08:00 | When execute A-plane encoder in Quality Assurance application, the iDms Utility can't be opened automatically. iDms Utility wasn't integrated in BB Refresh. | BB_WIP1.1 | WIP12 |  | Subsystem |  | N/A | Fixed | franco_a | 0 | 2 | DMS Software | DMS | NO | Closed | WIP12SegReq | Brilliance Big Bore |  | WIP12 | December 13, 2018 9:02:54 PM GMT+08:00
CTUDT00600646 | Refresh | Closed | vinoliya_j | November 29, 2018 10:58:34 PM GMT+08:00 | Refresh : Xray system cannot comply error pop ups and doesn't allow to scan | WIP11 |  | This is a HW fault. Logs do not go further than 15:44; no sure what the recovery was. | Subsystem | Scan Initialization X-Ray Error | NA | System Repaired | wang_c | 0 |  |  | GCX | NO | Closed |  | Ingenuity 64
CTUDT00600644 | Refresh | Closed | bommiReddy_r | November 29, 2018 8:47:03 PM GMT+08:00 | Refresh: Dose Management-DICOM- In Dicom file the Tag “ELSCINTCT_AVG_PATIENT_SIZE_CM (01F1, 1055)” is available instead of “ELSCINTCT_AVG_PATIENT_SIZE_CM (01F1, 1056)" | WIP11 | Host SSRS DHF303211 Rev 07 |  | Subsystem | DICOM tags | Console-SRS.15407 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | WIP13SubCom | Ovation |  | Host SSRS DHF303211 Rev 07 | February 28, 2019 1:33:41 PM GMT+08:00
CTUDT00600644 | Refresh | Closed | bommiReddy_r | November 29, 2018 8:47:03 PM GMT+08:00 | Refresh: Dose Management-DICOM- In Dicom file the Tag “ELSCINTCT_AVG_PATIENT_SIZE_CM (01F1, 1055)” is available instead of “ELSCINTCT_AVG_PATIENT_SIZE_CM (01F1, 1056)" | WIP11 | Host SSRS DHF303211 Rev 07 |  | Subsystem | DICOM tags | Console-SRS.15407 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | WIP13SubCom | Brilliance 64 |  | Host SSRS DHF303211 Rev 07 | February 28, 2019 1:33:41 PM GMT+08:00
CTUDT00600644 | Refresh | Closed | bommiReddy_r | November 29, 2018 8:47:03 PM GMT+08:00 | Refresh: Dose Management-DICOM- In Dicom file the Tag “ELSCINTCT_AVG_PATIENT_SIZE_CM (01F1, 1055)” is available instead of “ELSCINTCT_AVG_PATIENT_SIZE_CM (01F1, 1056)" | WIP11 | Host SSRS DHF303211 Rev 07 |  | Subsystem | DICOM tags | Console-SRS.15407 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | WIP13SubCom | Brilliance Big Bore |  | Host SSRS DHF303211 Rev 07 | February 28, 2019 1:33:41 PM GMT+08:00
CTUDT00600644 | Refresh | Closed | bommiReddy_r | November 29, 2018 8:47:03 PM GMT+08:00 | Refresh: Dose Management-DICOM- In Dicom file the Tag “ELSCINTCT_AVG_PATIENT_SIZE_CM (01F1, 1055)” is available instead of “ELSCINTCT_AVG_PATIENT_SIZE_CM (01F1, 1056)" | WIP11 | Host SSRS DHF303211 Rev 07 |  | Subsystem | DICOM tags | Console-SRS.15407 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | WIP13SubCom | Ingenuity 64 |  | Host SSRS DHF303211 Rev 07 | February 28, 2019 1:33:41 PM GMT+08:00
CTUDT00600644 | Refresh | Closed | bommiReddy_r | November 29, 2018 8:47:03 PM GMT+08:00 | Refresh: Dose Management-DICOM- In Dicom file the Tag “ELSCINTCT_AVG_PATIENT_SIZE_CM (01F1, 1055)” is available instead of “ELSCINTCT_AVG_PATIENT_SIZE_CM (01F1, 1056)" | WIP11 | Host SSRS DHF303211 Rev 07 |  | Subsystem | DICOM tags | Console-SRS.15407 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | WIP13SubCom | IQon1 |  | Host SSRS DHF303211 Rev 07 | February 28, 2019 1:33:41 PM GMT+08:00
CTUDT00600642 | Refresh | Closed | gT_s | November 29, 2018 8:08:50 PM GMT+08:00 | Refresh : The field displayed for the DICOM tag [0002:0013] is “ EBW4.8 for ICT”. | WIP 11 | DHF318051 Rev13 Host Verification Test Protocol |  | Subsystem | Test Design | NA | Fixed | ramesh_t | 0 | 3 | Test Design | Console4 | NO | Closed | WIP16SegReq | Ovation |  | DHF318051 Rev13 Host Verification Test Protocol | May 2, 2019 8:45:41 PM GMT+08:00
CTUDT00600642 | Refresh | Closed | gT_s | November 29, 2018 8:08:50 PM GMT+08:00 | Refresh : The field displayed for the DICOM tag [0002:0013] is “ EBW4.8 for ICT”. | WIP 11 | DHF318051 Rev13 Host Verification Test Protocol |  | Subsystem | Test Design | NA | Fixed | ramesh_t | 0 | 3 | Test Design | Console4 | NO | Closed | WIP16SegReq | Brilliance 64 |  | DHF318051 Rev13 Host Verification Test Protocol | May 2, 2019 8:45:41 PM GMT+08:00
CTUDT00600642 | Refresh | Closed | gT_s | November 29, 2018 8:08:50 PM GMT+08:00 | Refresh : The field displayed for the DICOM tag [0002:0013] is “ EBW4.8 for ICT”. | WIP 11 | DHF318051 Rev13 Host Verification Test Protocol |  | Subsystem | Test Design | NA | Fixed | ramesh_t | 0 | 3 | Test Design | Console4 | NO | Closed | WIP16SegReq | Brilliance Big Bore |  | DHF318051 Rev13 Host Verification Test Protocol | May 2, 2019 8:45:41 PM GMT+08:00
CTUDT00600642 | Refresh | Closed | gT_s | November 29, 2018 8:08:50 PM GMT+08:00 | Refresh : The field displayed for the DICOM tag [0002:0013] is “ EBW4.8 for ICT”. | WIP 11 | DHF318051 Rev13 Host Verification Test Protocol |  | Subsystem | Test Design | NA | Fixed | ramesh_t | 0 | 3 | Test Design | Console4 | NO | Closed | WIP16SegReq | Ingenuity 64 |  | DHF318051 Rev13 Host Verification Test Protocol | May 2, 2019 8:45:41 PM GMT+08:00
CTUDT00600642 | Refresh | Closed | gT_s | November 29, 2018 8:08:50 PM GMT+08:00 | Refresh : The field displayed for the DICOM tag [0002:0013] is “ EBW4.8 for ICT”. | WIP 11 | DHF318051 Rev13 Host Verification Test Protocol |  | Subsystem | Test Design | NA | Fixed | ramesh_t | 0 | 3 | Test Design | Console4 | NO | Closed | WIP16SegReq | IQon1 |  | DHF318051 Rev13 Host Verification Test Protocol | May 2, 2019 8:45:41 PM GMT+08:00
CTUDT00600639 | Refresh | Closed | divya_n | November 29, 2018 6:57:59 PM GMT+08:00 | Annotation Values are not shown and application crashes on loading CT RT study to Quick Review/CTV application. | WIP11 | WIP13 |  | Subsystem | CT Viewer;Verified1 | NA | Fixed | redrouthu_r | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP14SegReq | Brilliance 64 |  | WIP13 | February 27, 2019 12:11:31 PM GMT+08:00
CTUDT00600639 | Refresh | Closed | divya_n | November 29, 2018 6:57:59 PM GMT+08:00 | Annotation Values are not shown and application crashes on loading CT RT study to Quick Review/CTV application. | WIP11 | WIP13 |  | Subsystem | CT Viewer;Verified1 | NA | Fixed | redrouthu_r | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP14SegReq | Brilliance Big Bore |  | WIP13 | February 27, 2019 12:11:31 PM GMT+08:00
CTUDT00600639 | Refresh | Closed | divya_n | November 29, 2018 6:57:59 PM GMT+08:00 | Annotation Values are not shown and application crashes on loading CT RT study to Quick Review/CTV application. | WIP11 | WIP13 |  | Subsystem | CT Viewer;Verified1 | NA | Fixed | redrouthu_r | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP14SegReq | Ingenuity 64 |  | WIP13 | February 27, 2019 12:11:31 PM GMT+08:00
CTUDT00600629 | Refresh | Closed | kr_m | November 29, 2018 6:00:57 PM GMT+08:00 | [Refresh - Sequential Dual Energy ]: "Injection parameters" dialog button is not grayed out for 2nd Acquisition | WIP9.1 | WIP12 |  | Subsystem | Injection Parameters DE | N/A | Fixed | prasad_r | 0 | 2 | Console Software | Console4 | YES | Closed | WIP12SubCom | Ingenuity 64 |  | WIP12 | December 10, 2018 7:07:45 PM GMT+08:00
CTUDT00600626 | Refresh | Closed | s_ra | November 29, 2018 4:29:45 PM GMT+08:00 | Patient age getting one year less, only in the morning on birthday. | 4.1.7 | BB_WIP3 | N/A | Subsystem |  | N/A | Fixed | veeraswamy_k | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP2_SubCom | Brilliance Big Bore | PR8805817 | BB_WIP3 | January 4, 2019 7:55:38 PM GMT+08:00
CTUDT00600623 | Refresh | Closed | attali_s | November 29, 2018 2:51:33 PM GMT+08:00 | Refresh: Tube Current: Failed scans are noted as 'Pass' in red color in the Detailed Table in the Results | WIP11 | WIP13 |  | Subsystem | Calibrations | NaN | Fixed | daibes_s | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | Ovation |  | WIP13 | January 28, 2019 5:22:18 PM GMT+08:00
CTUDT00600623 | Refresh | Closed | attali_s | November 29, 2018 2:51:33 PM GMT+08:00 | Refresh: Tube Current: Failed scans are noted as 'Pass' in red color in the Detailed Table in the Results | WIP11 | WIP13 |  | Subsystem | Calibrations | NaN | Fixed | daibes_s | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | Brilliance 64 |  | WIP13 | January 28, 2019 5:22:18 PM GMT+08:00
CTUDT00600623 | Refresh | Closed | attali_s | November 29, 2018 2:51:33 PM GMT+08:00 | Refresh: Tube Current: Failed scans are noted as 'Pass' in red color in the Detailed Table in the Results | WIP11 | WIP13 |  | Subsystem | Calibrations | NaN | Fixed | daibes_s | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | Brilliance Big Bore |  | WIP13 | January 28, 2019 5:22:18 PM GMT+08:00
CTUDT00600623 | Refresh | Closed | attali_s | November 29, 2018 2:51:33 PM GMT+08:00 | Refresh: Tube Current: Failed scans are noted as 'Pass' in red color in the Detailed Table in the Results | WIP11 | WIP13 |  | Subsystem | Calibrations | NaN | Fixed | daibes_s | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | Ingenuity 64 |  | WIP13 | January 28, 2019 5:22:18 PM GMT+08:00
CTUDT00600623 | Refresh | Closed | attali_s | November 29, 2018 2:51:33 PM GMT+08:00 | Refresh: Tube Current: Failed scans are noted as 'Pass' in red color in the Detailed Table in the Results | WIP11 | WIP13 |  | Subsystem | Calibrations | NaN | Fixed | daibes_s | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | IQon1 |  | WIP13 | January 28, 2019 5:22:18 PM GMT+08:00
CTUDT00600618 | Refresh | Closed | divya_n | November 28, 2018 8:42:46 PM GMT+08:00 | Refresh/WIP11/Patient Directory: Copy to CIRS fails and no studies are displayed in the Reconstruction tab. | WIP11 |  |  | Subsystem | CopytoCIRS |  | Duplicate | richards_b | 0 |  |  | CIRS | NO | Closed |  | Brilliance 64
CTUDT00600618 | Refresh | Closed | divya_n | November 28, 2018 8:42:46 PM GMT+08:00 | Refresh/WIP11/Patient Directory: Copy to CIRS fails and no studies are displayed in the Reconstruction tab. | WIP11 |  |  | Subsystem | CopytoCIRS |  | Duplicate | richards_b | 0 |  |  | CIRS | NO | Closed |  | Ingenuity 16
CTUDT00600618 | Refresh | Closed | divya_n | November 28, 2018 8:42:46 PM GMT+08:00 | Refresh/WIP11/Patient Directory: Copy to CIRS fails and no studies are displayed in the Reconstruction tab. | WIP11 |  |  | Subsystem | CopytoCIRS |  | Duplicate | richards_b | 0 |  |  | CIRS | NO | Closed |  | IQon1
CTUDT00600616 | Refresh | Closed | divya_n | November 28, 2018 8:27:35 PM GMT+08:00 | Refresh-uCT-Unable to launch Rawdata | WIP 9 |  |  | Subsystem | Recon Tab | NA | Works As Specified / Not A Defect | divya_n | 0 |  |  | Console4 | NO | Closed | WIP13SegReq | Brilliance 64
CTUDT00600616 | Refresh | Closed | divya_n | November 28, 2018 8:27:35 PM GMT+08:00 | Refresh-uCT-Unable to launch Rawdata | WIP 9 |  |  | Subsystem | Recon Tab | NA | Works As Specified / Not A Defect | divya_n | 0 |  |  | Console4 | NO | Closed | WIP13SegReq | Ingenuity 16
CTUDT00600616 | Refresh | Closed | divya_n | November 28, 2018 8:27:35 PM GMT+08:00 | Refresh-uCT-Unable to launch Rawdata | WIP 9 |  |  | Subsystem | Recon Tab | NA | Works As Specified / Not A Defect | divya_n | 0 |  |  | Console4 | NO | Closed | WIP13SegReq | IQon1
CTUDT00600613 | Refresh | Closed | divya_n | November 28, 2018 7:09:11 PM GMT+08:00 | Refresh: Screen blanking functionality is not working | WIP9 | WIP13 | the user can manually blank or unblank the screen | Subsystem | Automatic Screen Blanking | N/A | Fixed | dasgupta_s | 0 | 3 | Console Software | Console4 | NO | Closed | WIP13SubCom | Ovation |  | WIP13 | February 8, 2019 5:45:44 PM GMT+08:00
CTUDT00600613 | Refresh | Closed | divya_n | November 28, 2018 7:09:11 PM GMT+08:00 | Refresh: Screen blanking functionality is not working | WIP9 | WIP13 | the user can manually blank or unblank the screen | Subsystem | Automatic Screen Blanking | N/A | Fixed | dasgupta_s | 0 | 3 | Console Software | Console4 | NO | Closed | WIP13SubCom | Brilliance 64 |  | WIP13 | February 8, 2019 5:45:44 PM GMT+08:00
CTUDT00600613 | Refresh | Closed | divya_n | November 28, 2018 7:09:11 PM GMT+08:00 | Refresh: Screen blanking functionality is not working | WIP9 | WIP13 | the user can manually blank or unblank the screen | Subsystem | Automatic Screen Blanking | N/A | Fixed | dasgupta_s | 0 | 3 | Console Software | Console4 | NO | Closed | WIP13SubCom | Brilliance Big Bore |  | WIP13 | February 8, 2019 5:45:44 PM GMT+08:00
CTUDT00600613 | Refresh | Closed | divya_n | November 28, 2018 7:09:11 PM GMT+08:00 | Refresh: Screen blanking functionality is not working | WIP9 | WIP13 | the user can manually blank or unblank the screen | Subsystem | Automatic Screen Blanking | N/A | Fixed | dasgupta_s | 0 | 3 | Console Software | Console4 | NO | Closed | WIP13SubCom | Ingenuity 64 |  | WIP13 | February 8, 2019 5:45:44 PM GMT+08:00
CTUDT00600613 | Refresh | Closed | divya_n | November 28, 2018 7:09:11 PM GMT+08:00 | Refresh: Screen blanking functionality is not working | WIP9 | WIP13 | the user can manually blank or unblank the screen | Subsystem | Automatic Screen Blanking | N/A | Fixed | dasgupta_s | 0 | 3 | Console Software | Console4 | NO | Closed | WIP13SubCom | IQon1 |  | WIP13 | February 8, 2019 5:45:44 PM GMT+08:00
CTUDT00600605 | Refresh | Closed | datta_i | November 28, 2018 3:09:20 PM GMT+08:00 | Not all DICOM Structured Reports are Accepted at Destination Node due to Missing TextValue (0040,a160) | 4.1.6 | WIP12 |  | Subsystem |  | N/A | Fixed | v_sumanth | 0 | 2 | Console Software | Console4 | NO | Closed | WIP12SegReq | Ovation | CFR9420595/PR6189708 | WIP12 | December 4, 2018 7:36:39 PM GMT+08:00
CTUDT00600605 | Refresh | Closed | datta_i | November 28, 2018 3:09:20 PM GMT+08:00 | Not all DICOM Structured Reports are Accepted at Destination Node due to Missing TextValue (0040,a160) | 4.1.6 | WIP12 |  | Subsystem |  | N/A | Fixed | v_sumanth | 0 | 2 | Console Software | Console4 | NO | Closed | WIP12SegReq | Brilliance 64 | CFR9420595/PR6189708 | WIP12 | December 4, 2018 7:36:39 PM GMT+08:00
CTUDT00600605 | Refresh | Closed | datta_i | November 28, 2018 3:09:20 PM GMT+08:00 | Not all DICOM Structured Reports are Accepted at Destination Node due to Missing TextValue (0040,a160) | 4.1.6 | WIP12 |  | Subsystem |  | N/A | Fixed | v_sumanth | 0 | 2 | Console Software | Console4 | NO | Closed | WIP12SegReq | Ingenuity 64 | CFR9420595/PR6189708 | WIP12 | December 4, 2018 7:36:39 PM GMT+08:00
CTUDT00600605 | Refresh | Closed | datta_i | November 28, 2018 3:09:20 PM GMT+08:00 | Not all DICOM Structured Reports are Accepted at Destination Node due to Missing TextValue (0040,a160) | 4.1.6 | WIP12 |  | Subsystem |  | N/A | Fixed | v_sumanth | 0 | 2 | Console Software | Console4 | NO | Closed | WIP12SegReq | IQon1 | CFR9420595/PR6189708 | WIP12 | December 4, 2018 7:36:39 PM GMT+08:00
CTUDT00600601 | Refresh | Closed | naik_p | November 28, 2018 2:24:40 PM GMT+08:00 | Refresh_WIP 11.1 Sanity_System error. | WIP11.1 |  |  | Subsystem | System Error | NA | Invalid/Uncontrolled Configuration | shekhar_p | 0 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00600601 | Refresh | Closed | naik_p | November 28, 2018 2:24:40 PM GMT+08:00 | Refresh_WIP 11.1 Sanity_System error. | WIP11.1 |  |  | Subsystem | System Error | NA | Invalid/Uncontrolled Configuration | shekhar_p | 0 |  |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00600601 | Refresh | Closed | naik_p | November 28, 2018 2:24:40 PM GMT+08:00 | Refresh_WIP 11.1 Sanity_System error. | WIP11.1 |  |  | Subsystem | System Error | NA | Invalid/Uncontrolled Configuration | shekhar_p | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00600601 | Refresh | Closed | naik_p | November 28, 2018 2:24:40 PM GMT+08:00 | Refresh_WIP 11.1 Sanity_System error. | WIP11.1 |  |  | Subsystem | System Error | NA | Invalid/Uncontrolled Configuration | shekhar_p | 0 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00600600 | Refresh | Closed | sowmya_k | November 28, 2018 12:39:25 PM GMT+08:00 | V4.2: Pulmo Wave Peak clipped by the maximum threshold line for Amplitude binning. | 4.2.1.17001 | BB_WIP3 |  | Subsystem |  | NA | Fixed | sowmya_k | 0 | 2 | Console Software | Console4 | YES | Closed | BB_WIP3_SubCom | Brilliance Big Bore | PR7779384 | BB_WIP3 | December 11, 2018 4:29:50 PM GMT+08:00
CTUDT00600594 | Refresh | Closed | singh_p | November 27, 2018 9:07:43 PM GMT+08:00 | Refresh WIP 11 : ADLDS is still configured with Refresh. | WIP11 | WIP13 | No | Subsystem | Login | NA | Fixed | ravindra_a | 0 | 3 | Console Software | Console4 | NO | Closed | WIP13SubCom | Ovation |  | WIP13 | January 14, 2019 7:04:16 PM GMT+08:00
CTUDT00600594 | Refresh | Closed | singh_p | November 27, 2018 9:07:43 PM GMT+08:00 | Refresh WIP 11 : ADLDS is still configured with Refresh. | WIP11 | WIP13 | No | Subsystem | Login | NA | Fixed | ravindra_a | 0 | 3 | Console Software | Console4 | NO | Closed | WIP13SubCom | Brilliance 64 |  | WIP13 | January 14, 2019 7:04:16 PM GMT+08:00
CTUDT00600594 | Refresh | Closed | singh_p | November 27, 2018 9:07:43 PM GMT+08:00 | Refresh WIP 11 : ADLDS is still configured with Refresh. | WIP11 | WIP13 | No | Subsystem | Login | NA | Fixed | ravindra_a | 0 | 3 | Console Software | Console4 | NO | Closed | WIP13SubCom | Ingenuity 64 |  | WIP13 | January 14, 2019 7:04:16 PM GMT+08:00
CTUDT00600594 | Refresh | Closed | singh_p | November 27, 2018 9:07:43 PM GMT+08:00 | Refresh WIP 11 : ADLDS is still configured with Refresh. | WIP11 | WIP13 | No | Subsystem | Login | NA | Fixed | ravindra_a | 0 | 3 | Console Software | Console4 | NO | Closed | WIP13SubCom | IQon1 |  | WIP13 | January 14, 2019 7:04:16 PM GMT+08:00
CTUDT00600592 | Refresh | Closed | zhang_t | November 27, 2018 4:18:41 PM GMT+08:00 | Refresh: cannot check CPM Application and Boot Versions for all Controller items via service tool-> system configuration. | WIP11.1 | WIP13 |  | Subsystem | Service Tools_CPM | N/A | Fixed | chakraborty_d | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP13SubCom | Ovation |  | WIP13 | January 23, 2019 4:50:05 PM GMT+08:00
CTUDT00600592 | Refresh | Closed | zhang_t | November 27, 2018 4:18:41 PM GMT+08:00 | Refresh: cannot check CPM Application and Boot Versions for all Controller items via service tool-> system configuration. | WIP11.1 | WIP13 |  | Subsystem | Service Tools_CPM | N/A | Fixed | chakraborty_d | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP13SubCom | Brilliance 64 |  | WIP13 | January 23, 2019 4:50:05 PM GMT+08:00
CTUDT00600592 | Refresh | Closed | zhang_t | November 27, 2018 4:18:41 PM GMT+08:00 | Refresh: cannot check CPM Application and Boot Versions for all Controller items via service tool-> system configuration. | WIP11.1 | WIP13 |  | Subsystem | Service Tools_CPM | N/A | Fixed | chakraborty_d | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP13SubCom | Brilliance Big Bore |  | WIP13 | January 23, 2019 4:50:05 PM GMT+08:00
CTUDT00600592 | Refresh | Closed | zhang_t | November 27, 2018 4:18:41 PM GMT+08:00 | Refresh: cannot check CPM Application and Boot Versions for all Controller items via service tool-> system configuration. | WIP11.1 | WIP13 |  | Subsystem | Service Tools_CPM | N/A | Fixed | chakraborty_d | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP13SubCom | Ingenuity 64 |  | WIP13 | January 23, 2019 4:50:05 PM GMT+08:00
CTUDT00600592 | Refresh | Closed | zhang_t | November 27, 2018 4:18:41 PM GMT+08:00 | Refresh: cannot check CPM Application and Boot Versions for all Controller items via service tool-> system configuration. | WIP11.1 | WIP13 |  | Subsystem | Service Tools_CPM | N/A | Fixed | chakraborty_d | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP13SubCom | IQon1 |  | WIP13 | January 23, 2019 4:50:05 PM GMT+08:00
CTUDT00600588 | Refresh | Closed | beck_d | November 26, 2018 8:28:42 PM GMT+08:00 | Refresh: Calibrations: ACR: Scans parameters not as defined in spec | WIP11 | WIP17 |  | Subsystem | Calibrations; Verified1 | n/a | Fixed | elviral | 0 | 2 | DMS Software | DMS | NO | Closed | BB_WIP4_SegReq | Ovation |  | WIP17 | May 26, 2019 5:56:44 PM GMT+08:00
CTUDT00600588 | Refresh | Closed | beck_d | November 26, 2018 8:28:42 PM GMT+08:00 | Refresh: Calibrations: ACR: Scans parameters not as defined in spec | WIP11 | WIP17 |  | Subsystem | Calibrations; Verified1 | n/a | Fixed | elviral | 0 | 2 | DMS Software | DMS | NO | Closed | BB_WIP4_SegReq | Brilliance Big Bore |  | WIP17 | May 26, 2019 5:56:44 PM GMT+08:00
CTUDT00600588 | Refresh | Closed | beck_d | November 26, 2018 8:28:42 PM GMT+08:00 | Refresh: Calibrations: ACR: Scans parameters not as defined in spec | WIP11 | WIP17 |  | Subsystem | Calibrations; Verified1 | n/a | Fixed | elviral | 0 | 2 | DMS Software | DMS | NO | Closed | BB_WIP4_SegReq | Ingenuity 64 |  | WIP17 | May 26, 2019 5:56:44 PM GMT+08:00
CTUDT00600587 | Refresh | Closed | jotawar_d | November 26, 2018 8:15:08 PM GMT+08:00 | Exam cards are getting deleted after performing copy paste operation and relaunching the ECM | WIP11 | WIP12 | performing restore with the backup/restore tool to restore the deleted Exam Cards | System | ECM | Console-SRS.834 | Fixed | prasad_r | 0 | 4 | Console Software | Console4 | NO | Closed | WIP12SubCom | Ovation | 9467908 | WIP12 | February 4, 2019 4:38:06 PM GMT+08:00
CTUDT00600587 | Refresh | Closed | jotawar_d | November 26, 2018 8:15:08 PM GMT+08:00 | Exam cards are getting deleted after performing copy paste operation and relaunching the ECM | WIP11 | WIP12 | performing restore with the backup/restore tool to restore the deleted Exam Cards | System | ECM | Console-SRS.834 | Fixed | prasad_r | 0 | 4 | Console Software | Console4 | NO | Closed | WIP12SubCom | Brilliance 64 | 9467908 | WIP12 | February 4, 2019 4:38:06 PM GMT+08:00
CTUDT00600587 | Refresh | Closed | jotawar_d | November 26, 2018 8:15:08 PM GMT+08:00 | Exam cards are getting deleted after performing copy paste operation and relaunching the ECM | WIP11 | WIP12 | performing restore with the backup/restore tool to restore the deleted Exam Cards | System | ECM | Console-SRS.834 | Fixed | prasad_r | 0 | 4 | Console Software | Console4 | NO | Closed | WIP12SubCom | Ingenuity 64 | 9467908 | WIP12 | February 4, 2019 4:38:06 PM GMT+08:00
CTUDT00600587 | Refresh | Closed | jotawar_d | November 26, 2018 8:15:08 PM GMT+08:00 | Exam cards are getting deleted after performing copy paste operation and relaunching the ECM | WIP11 | WIP12 | performing restore with the backup/restore tool to restore the deleted Exam Cards | System | ECM | Console-SRS.834 | Fixed | prasad_r | 0 | 4 | Console Software | Console4 | NO | Closed | WIP12SubCom | IQon1 | 9467908 | WIP12 | February 4, 2019 4:38:06 PM GMT+08:00
CTUDT00600585 | Refresh | Closed | panchami_pradeep | November 26, 2018 12:32:36 PM GMT+08:00 | Refresh: In olevel user, while log out application launched from PSC tool (By clicking on "Start Scanner Application") system is logging out from o level user itself. | WIP11 | WIP15 | Olevel User can re-login and continue work on PSC tool. | Subsystem | System Log Off | NA | Fixed | kapadia_v | 0 | 4 | Couch Software | Console4 | NO | Closed | WIP12SegReq | Ovation |  | WIP15 | March 22, 2019 4:28:21 PM GMT+08:00
CTUDT00600585 | Refresh | Closed | panchami_pradeep | November 26, 2018 12:32:36 PM GMT+08:00 | Refresh: In olevel user, while log out application launched from PSC tool (By clicking on "Start Scanner Application") system is logging out from o level user itself. | WIP11 | WIP15 | Olevel User can re-login and continue work on PSC tool. | Subsystem | System Log Off | NA | Fixed | kapadia_v | 0 | 4 | Couch Software | Console4 | NO | Closed | WIP12SegReq | Brilliance 64 |  | WIP15 | March 22, 2019 4:28:21 PM GMT+08:00
CTUDT00600585 | Refresh | Closed | panchami_pradeep | November 26, 2018 12:32:36 PM GMT+08:00 | Refresh: In olevel user, while log out application launched from PSC tool (By clicking on "Start Scanner Application") system is logging out from o level user itself. | WIP11 | WIP15 | Olevel User can re-login and continue work on PSC tool. | Subsystem | System Log Off | NA | Fixed | kapadia_v | 0 | 4 | Couch Software | Console4 | NO | Closed | WIP12SegReq | Ingenuity 64 |  | WIP15 | March 22, 2019 4:28:21 PM GMT+08:00
CTUDT00600585 | Refresh | Closed | panchami_pradeep | November 26, 2018 12:32:36 PM GMT+08:00 | Refresh: In olevel user, while log out application launched from PSC tool (By clicking on "Start Scanner Application") system is logging out from o level user itself. | WIP11 | WIP15 | Olevel User can re-login and continue work on PSC tool. | Subsystem | System Log Off | NA | Fixed | kapadia_v | 0 | 4 | Couch Software | Console4 | NO | Closed | WIP12SegReq | IQon1 |  | WIP15 | March 22, 2019 4:28:21 PM GMT+08:00
CTUDT00600582 | Refresh | Closed | beck_d | November 25, 2018 11:30:15 PM GMT+08:00 | Refresh: Calibrations: Acceptance and Constancy: Resolution cases failed because ROIs are not drawn on the images | WIP11 | WIP14 | N/A | Subsystem | Calibrations: Acceptance and Constancy | N/A | Works As Specified / Not A Defect | yosfan_s | 0 |  | DMS Calibrations | DMS | NO | Closed |  | Brilliance Big Bore |  | WIP14
CTUDT00600582 | Refresh | Closed | beck_d | November 25, 2018 11:30:15 PM GMT+08:00 | Refresh: Calibrations: Acceptance and Constancy: Resolution cases failed because ROIs are not drawn on the images | WIP11 | WIP14 | N/A | Subsystem | Calibrations: Acceptance and Constancy | N/A | Works As Specified / Not A Defect | yosfan_s | 0 |  | DMS Calibrations | DMS | NO | Closed |  | Ingenuity 64 |  | WIP14
CTUDT00600580 | Refresh | Closed | cohen_y | November 25, 2018 8:46:52 PM GMT+08:00 | Refresh: Beam Quality: Deviation from Nominal of Beam Quality that was performed in user other than RnD is not shown only in RnD user. | WIP11 |  |  | Subsystem | Beam Quality test |  | Duplicate | yatziv_l | 0 |  |  | DMS | NO | Closed |  | Ingenuity 64
CTUDT00600576 | Refresh | Closed | cohen_y | November 25, 2018 7:58:02 PM GMT+08:00 | Refresh: Beam Quality: Series Description in Archive is not as described in spec. | WIP11 | N/A | David will change Beam Quality SPEC | Subsystem | Beam Quality test - Documentation | NAN | Fixed | braunstein_d | 0 | 4 | Design Documentation | DMS | NO | Closed | WIP12SegReq | Ovation |  | N/A | January 17, 2019 5:03:13 PM GMT+08:00
CTUDT00600576 | Refresh | Closed | cohen_y | November 25, 2018 7:58:02 PM GMT+08:00 | Refresh: Beam Quality: Series Description in Archive is not as described in spec. | WIP11 | N/A | David will change Beam Quality SPEC | Subsystem | Beam Quality test - Documentation | NAN | Fixed | braunstein_d | 0 | 4 | Design Documentation | DMS | NO | Closed | WIP12SegReq | Brilliance 64 |  | N/A | January 17, 2019 5:03:13 PM GMT+08:00
CTUDT00600576 | Refresh | Closed | cohen_y | November 25, 2018 7:58:02 PM GMT+08:00 | Refresh: Beam Quality: Series Description in Archive is not as described in spec. | WIP11 | N/A | David will change Beam Quality SPEC | Subsystem | Beam Quality test - Documentation | NAN | Fixed | braunstein_d | 0 | 4 | Design Documentation | DMS | NO | Closed | WIP12SegReq | Brilliance Big Bore |  | N/A | January 17, 2019 5:03:13 PM GMT+08:00
CTUDT00600576 | Refresh | Closed | cohen_y | November 25, 2018 7:58:02 PM GMT+08:00 | Refresh: Beam Quality: Series Description in Archive is not as described in spec. | WIP11 | N/A | David will change Beam Quality SPEC | Subsystem | Beam Quality test - Documentation | NAN | Fixed | braunstein_d | 0 | 4 | Design Documentation | DMS | NO | Closed | WIP12SegReq | Ingenuity 64 |  | N/A | January 17, 2019 5:03:13 PM GMT+08:00
CTUDT00600576 | Refresh | Closed | cohen_y | November 25, 2018 7:58:02 PM GMT+08:00 | Refresh: Beam Quality: Series Description in Archive is not as described in spec. | WIP11 | N/A | David will change Beam Quality SPEC | Subsystem | Beam Quality test - Documentation | NAN | Fixed | braunstein_d | 0 | 4 | Design Documentation | DMS | NO | Closed | WIP12SegReq | IQon1 |  | N/A | January 17, 2019 5:03:13 PM GMT+08:00
CTUDT00600574 | Refresh | Closed | sowmya_k | November 24, 2018 1:47:19 PM GMT+08:00 | Refresh Ingenuity WIP 9.1 : System Error pop up appeared when bolus scan was paused. | WIP10 | BB_WIP3 |  | Subsystem |  | NA | Fixed | chokkalingam_r | 0 | 2 | Console Software | Console4 | YES | Closed | BB_WIP3_SegReq | Brilliance Big Bore |  | BB_WIP3 | December 4, 2018 7:19:12 PM GMT+08:00
CTUDT00600573 | Refresh | Closed | sowmya_k | November 24, 2018 1:42:59 PM GMT+08:00 | Refresh :Inconsistency in display of message “mAs for average {patient=Adult, Child or Infant}” in ECM/plan and also before and after surview. And | WIP10 | BB_WIP3 |  | Subsystem |  | NA | Fixed | chokkalingam_r | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP3_SegReq | Brilliance Big Bore |  | BB_WIP3 | December 4, 2018 6:20:47 PM GMT+08:00
CTUDT00600571 | Refresh | Closed | singh_p | November 23, 2018 9:49:11 PM GMT+08:00 | Refresh Ingenuity WIP 11 : Images did not appear in View tab post reconstruction. | WIP11 |  |  | Subsystem | View Tab | NA | Works As Specified / Not A Defect | babu_s | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00600567 | Refresh | Closed | navale_s | November 23, 2018 1:50:57 PM GMT+08:00 | Result Length is not getting Discretized within the surview boundaries on replacing exam card | 4.8 | BB_WIP2 | the result length change happens during planning and the user can adjust and plan the result as necessary | Subsystem |  | N/A | Fixed | navale_s | 0 | 3 | Console Software | Console4 | YES | Closed | BB_WIP3_SegReq | Brilliance Big Bore |  | BB_WIP2 | December 11, 2018 1:18:48 PM GMT+08:00
CTUDT00600565 | Refresh | Closed | pB_s | November 23, 2018 1:43:29 PM GMT+08:00 | WIP11: Screen blanking is not happening in console according to the given/set values | WIP11 | WIP13 | the user can manually blank or unblank the screen | Subsystem | Screen Blanking | N/A | Fixed | dasgupta_s | 0 | 3 | Console Software | Console4 | NO | Closed | WIP13SubCom | Ovation |  | WIP13 | February 7, 2019 12:54:58 PM GMT+08:00
CTUDT00600565 | Refresh | Closed | pB_s | November 23, 2018 1:43:29 PM GMT+08:00 | WIP11: Screen blanking is not happening in console according to the given/set values | WIP11 | WIP13 | the user can manually blank or unblank the screen | Subsystem | Screen Blanking | N/A | Fixed | dasgupta_s | 0 | 3 | Console Software | Console4 | NO | Closed | WIP13SubCom | Brilliance 64 |  | WIP13 | February 7, 2019 12:54:58 PM GMT+08:00
CTUDT00600565 | Refresh | Closed | pB_s | November 23, 2018 1:43:29 PM GMT+08:00 | WIP11: Screen blanking is not happening in console according to the given/set values | WIP11 | WIP13 | the user can manually blank or unblank the screen | Subsystem | Screen Blanking | N/A | Fixed | dasgupta_s | 0 | 3 | Console Software | Console4 | NO | Closed | WIP13SubCom | Brilliance Big Bore |  | WIP13 | February 7, 2019 12:54:58 PM GMT+08:00
CTUDT00600565 | Refresh | Closed | pB_s | November 23, 2018 1:43:29 PM GMT+08:00 | WIP11: Screen blanking is not happening in console according to the given/set values | WIP11 | WIP13 | the user can manually blank or unblank the screen | Subsystem | Screen Blanking | N/A | Fixed | dasgupta_s | 0 | 3 | Console Software | Console4 | NO | Closed | WIP13SubCom | Ingenuity 64 |  | WIP13 | February 7, 2019 12:54:58 PM GMT+08:00
CTUDT00600565 | Refresh | Closed | pB_s | November 23, 2018 1:43:29 PM GMT+08:00 | WIP11: Screen blanking is not happening in console according to the given/set values | WIP11 | WIP13 | the user can manually blank or unblank the screen | Subsystem | Screen Blanking | N/A | Fixed | dasgupta_s | 0 | 3 | Console Software | Console4 | NO | Closed | WIP13SubCom | IQon1 |  | WIP13 | February 7, 2019 12:54:58 PM GMT+08:00
CTUDT00600562 | Refresh | Closed | cohen_y | November 23, 2018 3:37:04 AM GMT+08:00 | Refresh: Beam Quality: Probe Serial Number is not printed correctly in the report. | WIP11 | WIP13 |  | Subsystem |  | NaN | Fixed | daibes_s | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | Ovation |  | WIP13 | January 3, 2019 5:18:23 PM GMT+08:00
CTUDT00600562 | Refresh | Closed | cohen_y | November 23, 2018 3:37:04 AM GMT+08:00 | Refresh: Beam Quality: Probe Serial Number is not printed correctly in the report. | WIP11 | WIP13 |  | Subsystem |  | NaN | Fixed | daibes_s | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | Brilliance 64 |  | WIP13 | January 3, 2019 5:18:23 PM GMT+08:00
CTUDT00600562 | Refresh | Closed | cohen_y | November 23, 2018 3:37:04 AM GMT+08:00 | Refresh: Beam Quality: Probe Serial Number is not printed correctly in the report. | WIP11 | WIP13 |  | Subsystem |  | NaN | Fixed | daibes_s | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | Brilliance Big Bore |  | WIP13 | January 3, 2019 5:18:23 PM GMT+08:00
CTUDT00600562 | Refresh | Closed | cohen_y | November 23, 2018 3:37:04 AM GMT+08:00 | Refresh: Beam Quality: Probe Serial Number is not printed correctly in the report. | WIP11 | WIP13 |  | Subsystem |  | NaN | Fixed | daibes_s | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | Ingenuity 64 |  | WIP13 | January 3, 2019 5:18:23 PM GMT+08:00
CTUDT00600562 | Refresh | Closed | cohen_y | November 23, 2018 3:37:04 AM GMT+08:00 | Refresh: Beam Quality: Probe Serial Number is not printed correctly in the report. | WIP11 | WIP13 |  | Subsystem |  | NaN | Fixed | daibes_s | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | IQon1 |  | WIP13 | January 3, 2019 5:18:23 PM GMT+08:00
CTUDT00600560 | Refresh | Closed | kadish_n | November 23, 2018 2:39:50 AM GMT+08:00 | Refresh: WIP 11: Missing Tube data in the results of some applications | WIP11 | WIP15 |  | Subsystem | Tube | n/a | Fixed | bikman | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP15SegReq | Ovation |  | WIP15 | March 27, 2019 6:37:39 PM GMT+08:00
CTUDT00600560 | Refresh | Closed | kadish_n | November 23, 2018 2:39:50 AM GMT+08:00 | Refresh: WIP 11: Missing Tube data in the results of some applications | WIP11 | WIP15 |  | Subsystem | Tube | n/a | Fixed | bikman | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP15SegReq | Brilliance Big Bore |  | WIP15 | March 27, 2019 6:37:39 PM GMT+08:00
CTUDT00600560 | Refresh | Closed | kadish_n | November 23, 2018 2:39:50 AM GMT+08:00 | Refresh: WIP 11: Missing Tube data in the results of some applications | WIP11 | WIP15 |  | Subsystem | Tube | n/a | Fixed | bikman | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP15SegReq | Ingenuity 64 |  | WIP15 | March 27, 2019 6:37:39 PM GMT+08:00
CTUDT00600560 | Refresh | Closed | kadish_n | November 23, 2018 2:39:50 AM GMT+08:00 | Refresh: WIP 11: Missing Tube data in the results of some applications | WIP11 | WIP15 |  | Subsystem | Tube | n/a | Fixed | bikman | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP15SegReq | IQon1 |  | WIP15 | March 27, 2019 6:37:39 PM GMT+08:00
CTUDT00600559 | Refresh | Closed | patra_r | November 22, 2018 7:36:06 PM GMT+08:00 | FoK: UCL: Exam Application crash during the acquisition | V4.8 | BB_WIP3 |  | Subsystem |  | NA | Fixed | v_sumanth | 0 | 3 | Console Software | Console4 | YES | Closed | BB_WIP3_SegReq | Brilliance Big Bore | 9413694 | BB_WIP3 | December 4, 2018 6:24:22 PM GMT+08:00
CTUDT00600557 | Refresh | Closed | pB_s | November 22, 2018 6:28:10 PM GMT+08:00 | Result is not launching automatically in review/analysis tab for edited results and System is not logging any logs under log viewer for auto launch. | WIP11 |  | Use loading images to application instead autolaunch | Subsystem | Auto Launch | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | Ovation
CTUDT00600557 | Refresh | Closed | pB_s | November 22, 2018 6:28:10 PM GMT+08:00 | Result is not launching automatically in review/analysis tab for edited results and System is not logging any logs under log viewer for auto launch. | WIP11 |  | Use loading images to application instead autolaunch | Subsystem | Auto Launch | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00600557 | Refresh | Closed | pB_s | November 22, 2018 6:28:10 PM GMT+08:00 | Result is not launching automatically in review/analysis tab for edited results and System is not logging any logs under log viewer for auto launch. | WIP11 |  | Use loading images to application instead autolaunch | Subsystem | Auto Launch | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00600557 | Refresh | Closed | pB_s | November 22, 2018 6:28:10 PM GMT+08:00 | Result is not launching automatically in review/analysis tab for edited results and System is not logging any logs under log viewer for auto launch. | WIP11 |  | Use loading images to application instead autolaunch | Subsystem | Auto Launch | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00600557 | Refresh | Closed | pB_s | November 22, 2018 6:28:10 PM GMT+08:00 | Result is not launching automatically in review/analysis tab for edited results and System is not logging any logs under log viewer for auto launch. | WIP11 |  | Use loading images to application instead autolaunch | Subsystem | Auto Launch | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | IQon1
CTUDT00600546 | Refresh | Closed | kapadia_v | November 22, 2018 4:18:50 PM GMT+08:00 | Repeat option missing for dual surview after changing the patient position. | BB_WIP1 | BB_WIP3 | User can repeat the dual surview scan by right clicking on surview step either on the scan ruler or on the series list. | Subsystem |  | NA | Fixed | ravindra_a | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP2_SubCom | Brilliance Big Bore |  | BB_WIP3 | January 3, 2019 4:59:14 PM GMT+08:00
CTUDT00600543 | Refresh | Closed | navale_s | November 22, 2018 1:43:16 PM GMT+08:00 | Actual CTDIvol and DLP is displayed incorrectly in Acquisition Tab during Perfusion and Interventional scan | 4.8 | BB_WIP2 |  | Subsystem | High_1 | N/A | Fixed | navale_s | 0 | 2 | Console Software | Console4 | YES | Closed | BB_WIP2_SubCom | Brilliance Big Bore |  | BB_WIP2 | December 11, 2018 1:21:14 PM GMT+08:00
CTUDT00600540 | Refresh | Closed | sowmya_k | November 22, 2018 12:02:07 PM GMT+08:00 | when using offline reconstruction, the resulting images intermittently have one less image than the original | V4.2 | BB_WIP3 |  | Subsystem |  | N/A | Fixed | sowmya_k | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP3_SubCom | Brilliance Big Bore | PR 8104549 | BB_WIP3 | December 3, 2018 4:40:50 PM GMT+08:00
CTUDT00600537 | Refresh | Closed | krisher_i | November 21, 2018 11:07:09 PM GMT+08:00 | Not all images were created in Image Orientation test. | WIP11 |  |  | Subsystem |  |  | Duplicate | richards_b | 0 |  |  | CIRS | NO | Closed |  | Ovation
CTUDT00600537 | Refresh | Closed | krisher_i | November 21, 2018 11:07:09 PM GMT+08:00 | Not all images were created in Image Orientation test. | WIP11 |  |  | Subsystem |  |  | Duplicate | richards_b | 0 |  |  | CIRS | NO | Closed |  | IQon1
CTUDT00600535 | Refresh | Closed | Valeria_Kashtanov | November 21, 2018 10:29:30 PM GMT+08:00 | Wrong time in calibration results and html report | WIP11 | WIP13 |  | Subsystem | Air Clibration | n/a | Fixed | bikman | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | Ovation |  | WIP13 | January 3, 2019 9:58:30 PM GMT+08:00
CTUDT00600535 | Refresh | Closed | Valeria_Kashtanov | November 21, 2018 10:29:30 PM GMT+08:00 | Wrong time in calibration results and html report | WIP11 | WIP13 |  | Subsystem | Air Clibration | n/a | Fixed | bikman | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | Brilliance 64 |  | WIP13 | January 3, 2019 9:58:30 PM GMT+08:00
CTUDT00600535 | Refresh | Closed | Valeria_Kashtanov | November 21, 2018 10:29:30 PM GMT+08:00 | Wrong time in calibration results and html report | WIP11 | WIP13 |  | Subsystem | Air Clibration | n/a | Fixed | bikman | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | Brilliance Big Bore |  | WIP13 | January 3, 2019 9:58:30 PM GMT+08:00
CTUDT00600535 | Refresh | Closed | Valeria_Kashtanov | November 21, 2018 10:29:30 PM GMT+08:00 | Wrong time in calibration results and html report | WIP11 | WIP13 |  | Subsystem | Air Clibration | n/a | Fixed | bikman | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | Ingenuity 64 |  | WIP13 | January 3, 2019 9:58:30 PM GMT+08:00
CTUDT00600535 | Refresh | Closed | Valeria_Kashtanov | November 21, 2018 10:29:30 PM GMT+08:00 | Wrong time in calibration results and html report | WIP11 | WIP13 |  | Subsystem | Air Clibration | n/a | Fixed | bikman | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | IQon1 |  | WIP13 | January 3, 2019 9:58:30 PM GMT+08:00
CTUDT00600532 | Refresh | Closed | fitusi_r | November 21, 2018 9:25:16 PM GMT+08:00 | Refresh, IQon-Philips Service IST expires. | WIP11 | BB_WIP4 WIP14 | No. | Subsystem | Login- IST | NA | Fixed | kapadia_v | 0 | 2 | Couch Software | Console4 | NO | Closed | WIP12SegReq | Ovation |  | BB_WIP4 WIP14 | February 20, 2019 2:57:05 PM GMT+08:00
CTUDT00600532 | Refresh | Closed | fitusi_r | November 21, 2018 9:25:16 PM GMT+08:00 | Refresh, IQon-Philips Service IST expires. | WIP11 | BB_WIP4 WIP14 | No. | Subsystem | Login- IST | NA | Fixed | kapadia_v | 0 | 2 | Couch Software | Console4 | NO | Closed | WIP12SegReq | Brilliance 64 |  | BB_WIP4 WIP14 | February 20, 2019 2:57:05 PM GMT+08:00
CTUDT00600532 | Refresh | Closed | fitusi_r | November 21, 2018 9:25:16 PM GMT+08:00 | Refresh, IQon-Philips Service IST expires. | WIP11 | BB_WIP4 WIP14 | No. | Subsystem | Login- IST | NA | Fixed | kapadia_v | 0 | 2 | Couch Software | Console4 | NO | Closed | WIP12SegReq | Brilliance Big Bore |  | BB_WIP4 WIP14 | February 20, 2019 2:57:05 PM GMT+08:00
CTUDT00600532 | Refresh | Closed | fitusi_r | November 21, 2018 9:25:16 PM GMT+08:00 | Refresh, IQon-Philips Service IST expires. | WIP11 | BB_WIP4 WIP14 | No. | Subsystem | Login- IST | NA | Fixed | kapadia_v | 0 | 2 | Couch Software | Console4 | NO | Closed | WIP12SegReq | Ingenuity 64 |  | BB_WIP4 WIP14 | February 20, 2019 2:57:05 PM GMT+08:00
CTUDT00600532 | Refresh | Closed | fitusi_r | November 21, 2018 9:25:16 PM GMT+08:00 | Refresh, IQon-Philips Service IST expires. | WIP11 | BB_WIP4 WIP14 | No. | Subsystem | Login- IST | NA | Fixed | kapadia_v | 0 | 2 | Couch Software | Console4 | NO | Closed | WIP12SegReq | IQon1 |  | BB_WIP4 WIP14 | February 20, 2019 2:57:05 PM GMT+08:00
CTUDT00600531 | Refresh | Closed | mistry_h | November 21, 2018 6:45:27 PM GMT+08:00 | Application throws error when user ends exam having combination of ECG and non-ECG steps. | v4.1.7 | WIP12 | After the final step( Non cardiac Step) scan is over ,wait for at least 10 seconds before clicking on End Exam . | Subsystem | High_5 | N/A | Fixed | naveen_k | 0 | 2 | Console Software | Console4 | NO | Closed | WIP12SubCom | Ovation | 9461389 | WIP12 | December 14, 2018 5:54:12 PM GMT+08:00
CTUDT00600531 | Refresh | Closed | mistry_h | November 21, 2018 6:45:27 PM GMT+08:00 | Application throws error when user ends exam having combination of ECG and non-ECG steps. | v4.1.7 | WIP12 | After the final step( Non cardiac Step) scan is over ,wait for at least 10 seconds before clicking on End Exam . | Subsystem | High_5 | N/A | Fixed | naveen_k | 0 | 2 | Console Software | Console4 | NO | Closed | WIP12SubCom | Brilliance 64 | 9461389 | WIP12 | December 14, 2018 5:54:12 PM GMT+08:00
CTUDT00600531 | Refresh | Closed | mistry_h | November 21, 2018 6:45:27 PM GMT+08:00 | Application throws error when user ends exam having combination of ECG and non-ECG steps. | v4.1.7 | WIP12 | After the final step( Non cardiac Step) scan is over ,wait for at least 10 seconds before clicking on End Exam . | Subsystem | High_5 | N/A | Fixed | naveen_k | 0 | 2 | Console Software | Console4 | NO | Closed | WIP12SubCom | Brilliance Big Bore | 9461389 | WIP12 | December 14, 2018 5:54:12 PM GMT+08:00
CTUDT00600531 | Refresh | Closed | mistry_h | November 21, 2018 6:45:27 PM GMT+08:00 | Application throws error when user ends exam having combination of ECG and non-ECG steps. | v4.1.7 | WIP12 | After the final step( Non cardiac Step) scan is over ,wait for at least 10 seconds before clicking on End Exam . | Subsystem | High_5 | N/A | Fixed | naveen_k | 0 | 2 | Console Software | Console4 | NO | Closed | WIP12SubCom | Ingenuity 64 | 9461389 | WIP12 | December 14, 2018 5:54:12 PM GMT+08:00
CTUDT00600521 | Refresh | Closed | gT_s | November 20, 2018 9:36:19 PM GMT+08:00 | Refresh :The philips service workflow is not getting launched. | WIP 11 |  |  | Subsystem | Service workflow |  | Works As Specified / Not A Defect | babu_s | 0 |  |  | Console4 | NO | Closed | WIP12SegReq | Ovation
CTUDT00600519 | Refresh | Closed | mohit_p | November 20, 2018 8:32:30 PM GMT+08:00 | Kiosk hardening bypass in HOST system. | WIP11 | WIP14 |  | Subsystem | Security - Hardening | None | Fixed | Prabhakar_Priyank_Muppidi | 0 | 2 | Console Software | Console4 | YES | Closed | WIP14SubCom | Ovation |  | WIP14 | March 13, 2019 2:25:41 PM GMT+08:00
CTUDT00600519 | Refresh | Closed | mohit_p | November 20, 2018 8:32:30 PM GMT+08:00 | Kiosk hardening bypass in HOST system. | WIP11 | WIP14 |  | Subsystem | Security - Hardening | None | Fixed | Prabhakar_Priyank_Muppidi | 0 | 2 | Console Software | Console4 | YES | Closed | WIP14SubCom | Brilliance 64 |  | WIP14 | March 13, 2019 2:25:41 PM GMT+08:00
CTUDT00600519 | Refresh | Closed | mohit_p | November 20, 2018 8:32:30 PM GMT+08:00 | Kiosk hardening bypass in HOST system. | WIP11 | WIP14 |  | Subsystem | Security - Hardening | None | Fixed | Prabhakar_Priyank_Muppidi | 0 | 2 | Console Software | Console4 | YES | Closed | WIP14SubCom | Brilliance Big Bore |  | WIP14 | March 13, 2019 2:25:41 PM GMT+08:00
CTUDT00600519 | Refresh | Closed | mohit_p | November 20, 2018 8:32:30 PM GMT+08:00 | Kiosk hardening bypass in HOST system. | WIP11 | WIP14 |  | Subsystem | Security - Hardening | None | Fixed | Prabhakar_Priyank_Muppidi | 0 | 2 | Console Software | Console4 | YES | Closed | WIP14SubCom | Ingenuity 64 |  | WIP14 | March 13, 2019 2:25:41 PM GMT+08:00
CTUDT00600519 | Refresh | Closed | mohit_p | November 20, 2018 8:32:30 PM GMT+08:00 | Kiosk hardening bypass in HOST system. | WIP11 | WIP14 |  | Subsystem | Security - Hardening | None | Fixed | Prabhakar_Priyank_Muppidi | 0 | 2 | Console Software | Console4 | YES | Closed | WIP14SubCom | IQon1 |  | WIP14 | March 13, 2019 2:25:41 PM GMT+08:00
CTUDT00600515 | Refresh | Closed | mohit_p | November 20, 2018 7:49:14 PM GMT+08:00 | Insecure Service Permissions on HOST system. | WIP11 |  |  | Subsystem | Security | None | No Plan to Fix | shaham_t | 0 | 3 |  | Console4 | YES | Closed | SubNC | Ovation
CTUDT00600515 | Refresh | Closed | mohit_p | November 20, 2018 7:49:14 PM GMT+08:00 | Insecure Service Permissions on HOST system. | WIP11 |  |  | Subsystem | Security | None | No Plan to Fix | shaham_t | 0 | 3 |  | Console4 | YES | Closed | SubNC | Brilliance 64
CTUDT00600515 | Refresh | Closed | mohit_p | November 20, 2018 7:49:14 PM GMT+08:00 | Insecure Service Permissions on HOST system. | WIP11 |  |  | Subsystem | Security | None | No Plan to Fix | shaham_t | 0 | 3 |  | Console4 | YES | Closed | SubNC | Brilliance Big Bore
CTUDT00600515 | Refresh | Closed | mohit_p | November 20, 2018 7:49:14 PM GMT+08:00 | Insecure Service Permissions on HOST system. | WIP11 |  |  | Subsystem | Security | None | No Plan to Fix | shaham_t | 0 | 3 |  | Console4 | YES | Closed | SubNC | Ingenuity 64
CTUDT00600515 | Refresh | Closed | mohit_p | November 20, 2018 7:49:14 PM GMT+08:00 | Insecure Service Permissions on HOST system. | WIP11 |  |  | Subsystem | Security | None | No Plan to Fix | shaham_t | 0 | 3 |  | Console4 | YES | Closed | SubNC | IQon1
CTUDT00600513 | Refresh | Closed | gT_s | November 20, 2018 7:48:42 PM GMT+08:00 | Gantry GetReady For Scan: The Actual Latency is more than 5000 milliseconds. | WIP11 | WIP17 | NA | Subsystem | Performance | NA | Fixed | ravindra_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP12SegReq | Ovation |  | WIP17 | May 10, 2019 8:42:22 PM GMT+08:00
CTUDT00600513 | Refresh | Closed | gT_s | November 20, 2018 7:48:42 PM GMT+08:00 | Gantry GetReady For Scan: The Actual Latency is more than 5000 milliseconds. | WIP11 | WIP17 | NA | Subsystem | Performance | NA | Fixed | ravindra_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP12SegReq | Brilliance 64 |  | WIP17 | May 10, 2019 8:42:22 PM GMT+08:00
CTUDT00600513 | Refresh | Closed | gT_s | November 20, 2018 7:48:42 PM GMT+08:00 | Gantry GetReady For Scan: The Actual Latency is more than 5000 milliseconds. | WIP11 | WIP17 | NA | Subsystem | Performance | NA | Fixed | ravindra_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP12SegReq | Ingenuity 64 |  | WIP17 | May 10, 2019 8:42:22 PM GMT+08:00
CTUDT00600513 | Refresh | Closed | gT_s | November 20, 2018 7:48:42 PM GMT+08:00 | Gantry GetReady For Scan: The Actual Latency is more than 5000 milliseconds. | WIP11 | WIP17 | NA | Subsystem | Performance | NA | Fixed | ravindra_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP12SegReq | IQon1 |  | WIP17 | May 10, 2019 8:42:22 PM GMT+08:00
CTUDT00600512 | Refresh | Closed | mohit_p | November 20, 2018 7:44:08 PM GMT+08:00 | Microsoft Windows Unquoted Service Path Enumeration | WIP11 | WIP16 |  | Subsystem | Security - Nessus | None | Fixed | Prabhakar_Priyank_Muppidi | 0 | 2 | Console Software | Console4 | YES | Closed | WIP13SegReq | Ovation |  | WIP16 | April 30, 2019 3:01:54 PM GMT+08:00
CTUDT00600512 | Refresh | Closed | mohit_p | November 20, 2018 7:44:08 PM GMT+08:00 | Microsoft Windows Unquoted Service Path Enumeration | WIP11 | WIP16 |  | Subsystem | Security - Nessus | None | Fixed | Prabhakar_Priyank_Muppidi | 0 | 2 | Console Software | Console4 | YES | Closed | WIP13SegReq | Brilliance 64 |  | WIP16 | April 30, 2019 3:01:54 PM GMT+08:00
CTUDT00600512 | Refresh | Closed | mohit_p | November 20, 2018 7:44:08 PM GMT+08:00 | Microsoft Windows Unquoted Service Path Enumeration | WIP11 | WIP16 |  | Subsystem | Security - Nessus | None | Fixed | Prabhakar_Priyank_Muppidi | 0 | 2 | Console Software | Console4 | YES | Closed | WIP13SegReq | Ingenuity 64 |  | WIP16 | April 30, 2019 3:01:54 PM GMT+08:00
CTUDT00600512 | Refresh | Closed | mohit_p | November 20, 2018 7:44:08 PM GMT+08:00 | Microsoft Windows Unquoted Service Path Enumeration | WIP11 | WIP16 |  | Subsystem | Security - Nessus | None | Fixed | Prabhakar_Priyank_Muppidi | 0 | 2 | Console Software | Console4 | YES | Closed | WIP13SegReq | IQon1 |  | WIP16 | April 30, 2019 3:01:54 PM GMT+08:00
CTUDT00600511 | Refresh | Closed | oreg_o | November 20, 2018 7:42:37 PM GMT+08:00 | Refresh - Crosstalk Measurement Tool, wrong plot name | WIP11 | WIP13 |  | Subsystem |  | n/a | Fixed | yosefk | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | Ovation |  | WIP13 | February 4, 2019 4:16:37 PM GMT+08:00
CTUDT00600511 | Refresh | Closed | oreg_o | November 20, 2018 7:42:37 PM GMT+08:00 | Refresh - Crosstalk Measurement Tool, wrong plot name | WIP11 | WIP13 |  | Subsystem |  | n/a | Fixed | yosefk | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | Brilliance Big Bore |  | WIP13 | February 4, 2019 4:16:37 PM GMT+08:00
CTUDT00600511 | Refresh | Closed | oreg_o | November 20, 2018 7:42:37 PM GMT+08:00 | Refresh - Crosstalk Measurement Tool, wrong plot name | WIP11 | WIP13 |  | Subsystem |  | n/a | Fixed | yosefk | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | Ingenuity 64 |  | WIP13 | February 4, 2019 4:16:37 PM GMT+08:00
CTUDT00600511 | Refresh | Closed | oreg_o | November 20, 2018 7:42:37 PM GMT+08:00 | Refresh - Crosstalk Measurement Tool, wrong plot name | WIP11 | WIP13 |  | Subsystem |  | n/a | Fixed | yosefk | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | IQon1 |  | WIP13 | February 4, 2019 4:16:37 PM GMT+08:00
CTUDT00600509 | Refresh | Closed | mohit_p | November 20, 2018 7:34:39 PM GMT+08:00 | Improper Validation of Incoming DICOM Packets. | WIP11 | WIP14 |  | Subsystem | Security - DICOM communication | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | YES | Closed | WIP14SegReq | Ovation |  | WIP14 | February 26, 2019 3:38:08 PM GMT+08:00
CTUDT00600509 | Refresh | Closed | mohit_p | November 20, 2018 7:34:39 PM GMT+08:00 | Improper Validation of Incoming DICOM Packets. | WIP11 | WIP14 |  | Subsystem | Security - DICOM communication | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | YES | Closed | WIP14SegReq | Brilliance 64 |  | WIP14 | February 26, 2019 3:38:08 PM GMT+08:00
CTUDT00600509 | Refresh | Closed | mohit_p | November 20, 2018 7:34:39 PM GMT+08:00 | Improper Validation of Incoming DICOM Packets. | WIP11 | WIP14 |  | Subsystem | Security - DICOM communication | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | YES | Closed | WIP14SegReq | Ingenuity 64 |  | WIP14 | February 26, 2019 3:38:08 PM GMT+08:00
CTUDT00600509 | Refresh | Closed | mohit_p | November 20, 2018 7:34:39 PM GMT+08:00 | Improper Validation of Incoming DICOM Packets. | WIP11 | WIP14 |  | Subsystem | Security - DICOM communication | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | YES | Closed | WIP14SegReq | IQon1 |  | WIP14 | February 26, 2019 3:38:08 PM GMT+08:00
CTUDT00600507 | Refresh | Closed | mohit_p | November 20, 2018 7:08:50 PM GMT+08:00 | Valid Certificate was missing for DICOM communication | WIP11 |  |  | Subsystem | Security - DICOM communication | NA | Does Not Apply | s_sa | 0 | 2 |  | iCAP | NO | Closed | WIP13SubCom | Ovation
CTUDT00600507 | Refresh | Closed | mohit_p | November 20, 2018 7:08:50 PM GMT+08:00 | Valid Certificate was missing for DICOM communication | WIP11 |  |  | Subsystem | Security - DICOM communication | NA | Does Not Apply | s_sa | 0 | 2 |  | iCAP | NO | Closed | WIP13SubCom | Brilliance 64
CTUDT00600507 | Refresh | Closed | mohit_p | November 20, 2018 7:08:50 PM GMT+08:00 | Valid Certificate was missing for DICOM communication | WIP11 |  |  | Subsystem | Security - DICOM communication | NA | Does Not Apply | s_sa | 0 | 2 |  | iCAP | NO | Closed | WIP13SubCom | Ingenuity 64
CTUDT00600507 | Refresh | Closed | mohit_p | November 20, 2018 7:08:50 PM GMT+08:00 | Valid Certificate was missing for DICOM communication | WIP11 |  |  | Subsystem | Security - DICOM communication | NA | Does Not Apply | s_sa | 0 | 2 |  | iCAP | NO | Closed | WIP13SubCom | IQon1
CTUDT00600497 | Refresh | Closed | oren_d | November 20, 2018 2:25:20 PM GMT+08:00 | Cut operation of review before save result crashed the exam application | 4.7.5 | WIP12 | The user is able to perform an offline reconstruction for the results. | Subsystem |  | Console-SRS.9081 | Fixed | prasad_r | 0 | 4 | Console Software | Console4 | NO | Closed | WIP11.1SegReq | Ovation | CFR#9470203 | WIP12 | November 26, 2018 9:04:44 PM GMT+08:00
CTUDT00600497 | Refresh | Closed | oren_d | November 20, 2018 2:25:20 PM GMT+08:00 | Cut operation of review before save result crashed the exam application | 4.7.5 | WIP12 | The user is able to perform an offline reconstruction for the results. | Subsystem |  | Console-SRS.9081 | Fixed | prasad_r | 0 | 4 | Console Software | Console4 | NO | Closed | WIP11.1SegReq | Brilliance 64 | CFR#9470203 | WIP12 | November 26, 2018 9:04:44 PM GMT+08:00
CTUDT00600497 | Refresh | Closed | oren_d | November 20, 2018 2:25:20 PM GMT+08:00 | Cut operation of review before save result crashed the exam application | 4.7.5 | WIP12 | The user is able to perform an offline reconstruction for the results. | Subsystem |  | Console-SRS.9081 | Fixed | prasad_r | 0 | 4 | Console Software | Console4 | NO | Closed | WIP11.1SegReq | iCT4 | CFR#9470203 | WIP12 | November 26, 2018 9:04:44 PM GMT+08:00
CTUDT00600497 | Refresh | Closed | oren_d | November 20, 2018 2:25:20 PM GMT+08:00 | Cut operation of review before save result crashed the exam application | 4.7.5 | WIP12 | The user is able to perform an offline reconstruction for the results. | Subsystem |  | Console-SRS.9081 | Fixed | prasad_r | 0 | 4 | Console Software | Console4 | NO | Closed | WIP11.1SegReq | Ingenuity 64 | CFR#9470203 | WIP12 | November 26, 2018 9:04:44 PM GMT+08:00
CTUDT00600497 | Refresh | Closed | oren_d | November 20, 2018 2:25:20 PM GMT+08:00 | Cut operation of review before save result crashed the exam application | 4.7.5 | WIP12 | The user is able to perform an offline reconstruction for the results. | Subsystem |  | Console-SRS.9081 | Fixed | prasad_r | 0 | 4 | Console Software | Console4 | NO | Closed | WIP11.1SegReq | IQon1 | CFR#9470203 | WIP12 | November 26, 2018 9:04:44 PM GMT+08:00
CTUDT00600491 | Refresh | Closed | Valeria_Kashtanov | November 19, 2018 11:15:41 PM GMT+08:00 | Wrong slice numbers of Bad Detectors in Mux2 in MBD NVRAM Write text file | WIP11 | WIP13 |  | Subsystem |  | N/A | Fixed | yosefk | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | Ovation |  | WIP13 | January 20, 2019 3:33:39 PM GMT+08:00
CTUDT00600483 | Refresh | Closed | rakic_k | November 19, 2018 8:32:34 PM GMT+08:00 | Refresh- Wedge scatter correction stage incorrect for head IQ integration | WIP10 | WIP14 |  | Subsystem | IQ | NA | Fixed | kumar_v | 0 | 2 | Console Software | Console4 | YES | Closed | WIP13SubCom | Ovation |  | WIP14 | January 23, 2019 6:00:13 PM GMT+08:00
CTUDT00600483 | Refresh | Closed | rakic_k | November 19, 2018 8:32:34 PM GMT+08:00 | Refresh- Wedge scatter correction stage incorrect for head IQ integration | WIP10 | WIP14 |  | Subsystem | IQ | NA | Fixed | kumar_v | 0 | 2 | Console Software | Console4 | YES | Closed | WIP13SubCom | Brilliance 64 |  | WIP14 | January 23, 2019 6:00:13 PM GMT+08:00
CTUDT00600483 | Refresh | Closed | rakic_k | November 19, 2018 8:32:34 PM GMT+08:00 | Refresh- Wedge scatter correction stage incorrect for head IQ integration | WIP10 | WIP14 |  | Subsystem | IQ | NA | Fixed | kumar_v | 0 | 2 | Console Software | Console4 | YES | Closed | WIP13SubCom | Ingenuity 64 |  | WIP14 | January 23, 2019 6:00:13 PM GMT+08:00
CTUDT00600480 | Refresh | Closed | rakic_k | November 19, 2018 8:10:51 PM GMT+08:00 | Refresh- Scatter correction table incorrect for head IQ integration | WIP10 | WIP14 |  | Subsystem |  | n/a | Fixed | yosefk | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP12SegReq | Ovation | NA | WIP14 | December 25, 2018 10:20:43 PM GMT+08:00
CTUDT00600476 | Refresh | Closed | sowmya_k | November 19, 2018 4:55:11 PM GMT+08:00 | Result length keeps reducing when toggling thickness in Axial 3D scans without tilt and Helical scans | WIP11 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | kumar_v | 0 |  |  | Console4 | NO | Closed | WIP12SegReq | Brilliance 64
CTUDT00600476 | Refresh | Closed | sowmya_k | November 19, 2018 4:55:11 PM GMT+08:00 | Result length keeps reducing when toggling thickness in Axial 3D scans without tilt and Helical scans | WIP11 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | kumar_v | 0 |  |  | Console4 | NO | Closed | WIP12SegReq | Ingenuity 64
CTUDT00600470 | Refresh | Closed | Noy_Seelig | November 18, 2018 10:41:40 PM GMT+08:00 | Refresh: Acceptance Test: Reoccurring error that prevents the test from passing. | WIP11 |  | N/A | Subsystem | Calibrations | N/A | Works As Specified / Not A Defect | gehtin_e | 0 |  |  | DMS | NO | Closed |  | Ingenuity 64
CTUDT00600470 | Refresh | Closed | Noy_Seelig | November 18, 2018 10:41:40 PM GMT+08:00 | Refresh: Acceptance Test: Reoccurring error that prevents the test from passing. | WIP11 |  | N/A | Subsystem | Calibrations | N/A | Works As Specified / Not A Defect | gehtin_e | 0 |  |  | DMS | NO | Closed |  | IQon1
CTUDT00600463 | Refresh | Closed | beck_d | November 18, 2018 8:20:53 PM GMT+08:00 | Refresh: Calibrations: Beam Quality, Tube Potential Test: kV in 120 kV, WOpen / Ti-Tef and 80 kV, WSmall / Ti-Half scans are measured and not according to spec | WIP11 | WIP13 |  | Subsystem |  | NaN | Fixed | daibes_s | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | Ovation |  | WIP13 | January 6, 2019 10:55:50 PM GMT+08:00
CTUDT00600461 | Refresh | Closed | beck_d | November 18, 2018 4:32:53 PM GMT+08:00 | Refresh: Calibrations: Beam Quality, Tube Potential Test: HVL in 120 kV, WOpen / Ti-Tef scan is not measured in Ops user | WIP11 | WIP13 |  | Subsystem |  | NaN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP13SegReq | Ovation |  | WIP13 | January 13, 2019 3:34:25 PM GMT+08:00
CTUDT00600457 | Refresh | Closed | beck_d | November 18, 2018 3:48:37 PM GMT+08:00 | Refresh: Calibrations: Tube Current: Limits files are missing | WIP11 | WIP13 |  | Subsystem |  | NaN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | Ovation |  | WIP13 | January 3, 2019 4:55:01 PM GMT+08:00
CTUDT00600457 | Refresh | Closed | beck_d | November 18, 2018 3:48:37 PM GMT+08:00 | Refresh: Calibrations: Tube Current: Limits files are missing | WIP11 | WIP13 |  | Subsystem |  | NaN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | Brilliance 64 |  | WIP13 | January 3, 2019 4:55:01 PM GMT+08:00
CTUDT00600457 | Refresh | Closed | beck_d | November 18, 2018 3:48:37 PM GMT+08:00 | Refresh: Calibrations: Tube Current: Limits files are missing | WIP11 | WIP13 |  | Subsystem |  | NaN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | Brilliance Big Bore |  | WIP13 | January 3, 2019 4:55:01 PM GMT+08:00
CTUDT00600457 | Refresh | Closed | beck_d | November 18, 2018 3:48:37 PM GMT+08:00 | Refresh: Calibrations: Tube Current: Limits files are missing | WIP11 | WIP13 |  | Subsystem |  | NaN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | Ingenuity 64 |  | WIP13 | January 3, 2019 4:55:01 PM GMT+08:00
CTUDT00600455 | Refresh | Closed | pB_s | November 16, 2018 7:58:06 PM GMT+08:00 | Backup Date and Time label is displayed incorrectly as "Backup Date amp;amp; Time" in IT tools. | WIP11 | WIP17 BB_WIP7 | No workaround | Subsystem | High_3 | NA | Fixed | ravindra_a | 0 | 4 | Console Software | Console4 | NO | Closed | WIP12SubCom | Ovation |  | WIP17 BB_WIP7 | May 8, 2019 5:01:34 PM GMT+08:00
CTUDT00600455 | Refresh | Closed | pB_s | November 16, 2018 7:58:06 PM GMT+08:00 | Backup Date and Time label is displayed incorrectly as "Backup Date amp;amp; Time" in IT tools. | WIP11 | WIP17 BB_WIP7 | No workaround | Subsystem | High_3 | NA | Fixed | ravindra_a | 0 | 4 | Console Software | Console4 | NO | Closed | WIP12SubCom | Brilliance 64 |  | WIP17 BB_WIP7 | May 8, 2019 5:01:34 PM GMT+08:00
CTUDT00600455 | Refresh | Closed | pB_s | November 16, 2018 7:58:06 PM GMT+08:00 | Backup Date and Time label is displayed incorrectly as "Backup Date amp;amp; Time" in IT tools. | WIP11 | WIP17 BB_WIP7 | No workaround | Subsystem | High_3 | NA | Fixed | ravindra_a | 0 | 4 | Console Software | Console4 | NO | Closed | WIP12SubCom | Brilliance Big Bore |  | WIP17 BB_WIP7 | May 8, 2019 5:01:34 PM GMT+08:00
CTUDT00600455 | Refresh | Closed | pB_s | November 16, 2018 7:58:06 PM GMT+08:00 | Backup Date and Time label is displayed incorrectly as "Backup Date amp;amp; Time" in IT tools. | WIP11 | WIP17 BB_WIP7 | No workaround | Subsystem | High_3 | NA | Fixed | ravindra_a | 0 | 4 | Console Software | Console4 | NO | Closed | WIP12SubCom | Ingenuity 64 |  | WIP17 BB_WIP7 | May 8, 2019 5:01:34 PM GMT+08:00
CTUDT00600455 | Refresh | Closed | pB_s | November 16, 2018 7:58:06 PM GMT+08:00 | Backup Date and Time label is displayed incorrectly as "Backup Date amp;amp; Time" in IT tools. | WIP11 | WIP17 BB_WIP7 | No workaround | Subsystem | High_3 | NA | Fixed | ravindra_a | 0 | 4 | Console Software | Console4 | NO | Closed | WIP12SubCom | IQon1 |  | WIP17 BB_WIP7 | May 8, 2019 5:01:34 PM GMT+08:00
CTUDT00600451 | Refresh | Closed | roshan_r | November 16, 2018 6:46:20 PM GMT+08:00 | Refresh : Error in saving logs in log viewer | WIP11 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | veerabhadraiah | 0 |  |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00600448 | Refresh | Closed | sowmya_k | November 16, 2018 6:33:06 PM GMT+08:00 | when using offline reconstruction, the resulting images intermittently have one less image than the original | WIP11 | WIP11.1 |  | Subsystem |  | N/A | Fixed | sowmya_k | 0 | 2 | Console Software | Console4 | NO | Closed | WIP11.1SubCom | Ovation | PR 8104549 | WIP11.1 | November 22, 2018 2:58:48 PM GMT+08:00
CTUDT00600448 | Refresh | Closed | sowmya_k | November 16, 2018 6:33:06 PM GMT+08:00 | when using offline reconstruction, the resulting images intermittently have one less image than the original | WIP11 | WIP11.1 |  | Subsystem |  | N/A | Fixed | sowmya_k | 0 | 2 | Console Software | Console4 | NO | Closed | WIP11.1SubCom | Brilliance 64 | PR 8104549 | WIP11.1 | November 22, 2018 2:58:48 PM GMT+08:00
CTUDT00600448 | Refresh | Closed | sowmya_k | November 16, 2018 6:33:06 PM GMT+08:00 | when using offline reconstruction, the resulting images intermittently have one less image than the original | WIP11 | WIP11.1 |  | Subsystem |  | N/A | Fixed | sowmya_k | 0 | 2 | Console Software | Console4 | NO | Closed | WIP11.1SubCom | Ingenuity 64 | PR 8104549 | WIP11.1 | November 22, 2018 2:58:48 PM GMT+08:00
CTUDT00600448 | Refresh | Closed | sowmya_k | November 16, 2018 6:33:06 PM GMT+08:00 | when using offline reconstruction, the resulting images intermittently have one less image than the original | WIP11 | WIP11.1 |  | Subsystem |  | N/A | Fixed | sowmya_k | 0 | 2 | Console Software | Console4 | NO | Closed | WIP11.1SubCom | IQon1 | PR 8104549 | WIP11.1 | November 22, 2018 2:58:48 PM GMT+08:00
CTUDT00600447 | Refresh | Closed | jotawar_d | November 16, 2018 5:53:20 PM GMT+08:00 | Language name displayed in Voice manager is incorrect for "Chinese-standard mandarin","Slovak" and also for "Portuguese" | BB_WIP1 | BB_WIP3 | NA | Subsystem |  | NA | Fixed | ravindra_a | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP3_SegReq | Brilliance Big Bore |  | BB_WIP3 | January 3, 2019 6:53:45 PM GMT+08:00
CTUDT00600446 | Refresh | Closed | patra_r | November 16, 2018 12:49:19 PM GMT+08:00 | FoK: UCL: Exam Application crash during the acquisition | V4.7.7 and V4.1.10 | WIP12 |  | System | Verified1 | NA | Fixed | v_sumanth | 0 | 3 | Console Software | Console4 | YES | Closed | WIP11.1SegReq | Brilliance 64 | 9413694 | WIP12 | November 23, 2018 4:58:56 PM GMT+08:00
CTUDT00600446 | Refresh | Closed | patra_r | November 16, 2018 12:49:19 PM GMT+08:00 | FoK: UCL: Exam Application crash during the acquisition | V4.7.7 and V4.1.10 | WIP12 |  | System | Verified1 | NA | Fixed | v_sumanth | 0 | 3 | Console Software | Console4 | YES | Closed | WIP11.1SegReq | Ingenuity 64 | 9413694 | WIP12 | November 23, 2018 4:58:56 PM GMT+08:00
CTUDT00600443 | Refresh | Closed | karkour | November 16, 2018 1:13:43 AM GMT+08:00 | Rotor positioning errors during surviews | WIP9.1 | WIP13 | In most cases the operator can re-plan a surview or the second scan of a dual surview, and continue without a patient re-scan. In cases where a surview exposure is interrupted, the operator may need to re-execute the surview. | GCX | Rotor positions | GCX.SSRS.iCT.15696 | Fixed | karkour | 0 | 2 | Gantry Software | GCX | YES | Closed | WIP13SegReq | Ingenuity 64 | 5421199 6394347 6579569 4946811 | WIP13 | May 7, 2019 11:35:15 PM GMT+08:00
CTUDT00600437 | Refresh | Closed | wang_l | November 15, 2018 3:11:21 PM GMT+08:00 | Error detected during File Comparison for two different files in control utilities >diagnostics | WIP11 | WIP16 |  | Subsystem | Service Tools | N/A | Fixed | Mark_Huang | 0 | 2 | Service Tools_Software | GCX | NO | Closed | WIP14SegReq | Ovation |  | WIP16 | April 18, 2019 3:28:17 PM GMT+08:00
CTUDT00600437 | Refresh | Closed | wang_l | November 15, 2018 3:11:21 PM GMT+08:00 | Error detected during File Comparison for two different files in control utilities >diagnostics | WIP11 | WIP16 |  | Subsystem | Service Tools | N/A | Fixed | Mark_Huang | 0 | 2 | Service Tools_Software | GCX | NO | Closed | WIP14SegReq | Brilliance 64 |  | WIP16 | April 18, 2019 3:28:17 PM GMT+08:00
CTUDT00600437 | Refresh | Closed | wang_l | November 15, 2018 3:11:21 PM GMT+08:00 | Error detected during File Comparison for two different files in control utilities >diagnostics | WIP11 | WIP16 |  | Subsystem | Service Tools | N/A | Fixed | Mark_Huang | 0 | 2 | Service Tools_Software | GCX | NO | Closed | WIP14SegReq | Brilliance Big Bore |  | WIP16 | April 18, 2019 3:28:17 PM GMT+08:00
CTUDT00600437 | Refresh | Closed | wang_l | November 15, 2018 3:11:21 PM GMT+08:00 | Error detected during File Comparison for two different files in control utilities >diagnostics | WIP11 | WIP16 |  | Subsystem | Service Tools | N/A | Fixed | Mark_Huang | 0 | 2 | Service Tools_Software | GCX | NO | Closed | WIP14SegReq | Ingenuity 64 |  | WIP16 | April 18, 2019 3:28:17 PM GMT+08:00
CTUDT00600437 | Refresh | Closed | wang_l | November 15, 2018 3:11:21 PM GMT+08:00 | Error detected during File Comparison for two different files in control utilities >diagnostics | WIP11 | WIP16 |  | Subsystem | Service Tools | N/A | Fixed | Mark_Huang | 0 | 2 | Service Tools_Software | GCX | NO | Closed | WIP14SegReq | IQon1 |  | WIP16 | April 18, 2019 3:28:17 PM GMT+08:00
CTUDT00600434 | Refresh | Closed | pB_s | November 15, 2018 2:26:24 PM GMT+08:00 | WIP11: Incorrect message is getting displayed for emergency login under audit trail viewer in IT tools. | WIP11 | WIP12 | NA | Subsystem |  | N/A | Fixed | goswami_s | 0 | 4 | Console Software | Console4 | NO | Closed | WIP12SubCom | Ovation |  | WIP12 | December 4, 2018 7:49:13 PM GMT+08:00
CTUDT00600434 | Refresh | Closed | pB_s | November 15, 2018 2:26:24 PM GMT+08:00 | WIP11: Incorrect message is getting displayed for emergency login under audit trail viewer in IT tools. | WIP11 | WIP12 | NA | Subsystem |  | N/A | Fixed | goswami_s | 0 | 4 | Console Software | Console4 | NO | Closed | WIP12SubCom | Brilliance 64 |  | WIP12 | December 4, 2018 7:49:13 PM GMT+08:00
CTUDT00600434 | Refresh | Closed | pB_s | November 15, 2018 2:26:24 PM GMT+08:00 | WIP11: Incorrect message is getting displayed for emergency login under audit trail viewer in IT tools. | WIP11 | WIP12 | NA | Subsystem |  | N/A | Fixed | goswami_s | 0 | 4 | Console Software | Console4 | NO | Closed | WIP12SubCom | Brilliance Big Bore |  | WIP12 | December 4, 2018 7:49:13 PM GMT+08:00
CTUDT00600434 | Refresh | Closed | pB_s | November 15, 2018 2:26:24 PM GMT+08:00 | WIP11: Incorrect message is getting displayed for emergency login under audit trail viewer in IT tools. | WIP11 | WIP12 | NA | Subsystem |  | N/A | Fixed | goswami_s | 0 | 4 | Console Software | Console4 | NO | Closed | WIP12SubCom | Ingenuity 64 |  | WIP12 | December 4, 2018 7:49:13 PM GMT+08:00
CTUDT00600434 | Refresh | Closed | pB_s | November 15, 2018 2:26:24 PM GMT+08:00 | WIP11: Incorrect message is getting displayed for emergency login under audit trail viewer in IT tools. | WIP11 | WIP12 | NA | Subsystem |  | N/A | Fixed | goswami_s | 0 | 4 | Console Software | Console4 | NO | Closed | WIP12SubCom | IQon1 |  | WIP12 | December 4, 2018 7:49:13 PM GMT+08:00
CTUDT00600431 | Refresh | Closed | wang_l | November 15, 2018 1:59:32 PM GMT+08:00 | Invalid message was displayed in User Feedback section during CPM progress in Control Utilities | WIP11 | WIP16 |  | Subsystem | Service Tools CPM | N/A | Fixed | Mark_Huang | 0 | 4 | Service Tools_Software | GCX | NO | Closed | WIP14SegReq | Ovation |  | WIP16 | April 18, 2019 3:24:44 PM GMT+08:00
CTUDT00600431 | Refresh | Closed | wang_l | November 15, 2018 1:59:32 PM GMT+08:00 | Invalid message was displayed in User Feedback section during CPM progress in Control Utilities | WIP11 | WIP16 |  | Subsystem | Service Tools CPM | N/A | Fixed | Mark_Huang | 0 | 4 | Service Tools_Software | GCX | NO | Closed | WIP14SegReq | Brilliance 64 |  | WIP16 | April 18, 2019 3:24:44 PM GMT+08:00
CTUDT00600431 | Refresh | Closed | wang_l | November 15, 2018 1:59:32 PM GMT+08:00 | Invalid message was displayed in User Feedback section during CPM progress in Control Utilities | WIP11 | WIP16 |  | Subsystem | Service Tools CPM | N/A | Fixed | Mark_Huang | 0 | 4 | Service Tools_Software | GCX | NO | Closed | WIP14SegReq | Brilliance Big Bore |  | WIP16 | April 18, 2019 3:24:44 PM GMT+08:00
CTUDT00600431 | Refresh | Closed | wang_l | November 15, 2018 1:59:32 PM GMT+08:00 | Invalid message was displayed in User Feedback section during CPM progress in Control Utilities | WIP11 | WIP16 |  | Subsystem | Service Tools CPM | N/A | Fixed | Mark_Huang | 0 | 4 | Service Tools_Software | GCX | NO | Closed | WIP14SegReq | Ingenuity 64 |  | WIP16 | April 18, 2019 3:24:44 PM GMT+08:00
CTUDT00600431 | Refresh | Closed | wang_l | November 15, 2018 1:59:32 PM GMT+08:00 | Invalid message was displayed in User Feedback section during CPM progress in Control Utilities | WIP11 | WIP16 |  | Subsystem | Service Tools CPM | N/A | Fixed | Mark_Huang | 0 | 4 | Service Tools_Software | GCX | NO | Closed | WIP14SegReq | IQon1 |  | WIP16 | April 18, 2019 3:24:44 PM GMT+08:00
CTUDT00600428 | Refresh | Closed | dadon_e | November 15, 2018 12:05:55 AM GMT+08:00 | Refresh, IQon - "1-click installation" with USB device is stuck - cannot be installed in WIP11 | WIP11 |  |  | Subsystem | 1-Click installation | N/A | Works As Specified / Not A Defect | tamari_o | 1 |  |  | Support-System | NO | Closed |  | iCT2
CTUDT00600428 | Refresh | Closed | dadon_e | November 15, 2018 12:05:55 AM GMT+08:00 | Refresh, IQon - "1-click installation" with USB device is stuck - cannot be installed in WIP11 | WIP11 |  |  | Subsystem | 1-Click installation | N/A | Works As Specified / Not A Defect | tamari_o | 1 |  |  | Support-System | NO | Closed |  | IQon1
CTUDT00600425 | Refresh | Closed | roshan_r | November 14, 2018 10:41:39 PM GMT+08:00 | Refresh : Anonymous patient numbering continuity missed after software upgrade. | WIP11 | DHF318051 Rev 10 HOST Verification Test Protocol |  | Subsystem | Test Design | N/A | Fixed | shekhar_p | 0 | 3 | Test Design | Console4 | NO | Closed | WIP12SegReq | Ingenuity 64 |  | DHF318051 Rev 10 HOST Verification Test Protocol | March 20, 2019 12:28:00 PM GMT+08:00
CTUDT00600423 | Refresh | Closed | or_shmuel | November 14, 2018 10:15:16 PM GMT+08:00 | Refresh- Installed Versions Failed | WIP11 |  |  | Subsystem |  |  | Duplicate | venugopal_s | 0 |  |  | DI | NO | Closed | WIP12SegReq | Ovation
CTUDT00600423 | Refresh | Closed | or_shmuel | November 14, 2018 10:15:16 PM GMT+08:00 | Refresh- Installed Versions Failed | WIP11 |  |  | Subsystem |  |  | Duplicate | venugopal_s | 0 |  |  | DI | NO | Closed | WIP12SegReq | IQon1
CTUDT00600416 | Refresh | Closed | roshan_r | November 14, 2018 6:09:50 PM GMT+08:00 | Wrong series number getting generated in the exam summary page | WIP10 | WIP14 |  | Subsystem | Medium_1 | NA | Fixed | chokkalingam_r | 0 | 2 | Console Software | Console4 | NO | Closed | WIP13SegReq | Ovation |  | WIP14 | February 1, 2019 5:24:51 PM GMT+08:00
CTUDT00600416 | Refresh | Closed | roshan_r | November 14, 2018 6:09:50 PM GMT+08:00 | Wrong series number getting generated in the exam summary page | WIP10 | WIP14 |  | Subsystem | Medium_1 | NA | Fixed | chokkalingam_r | 0 | 2 | Console Software | Console4 | NO | Closed | WIP13SegReq | Brilliance 64 |  | WIP14 | February 1, 2019 5:24:51 PM GMT+08:00
CTUDT00600416 | Refresh | Closed | roshan_r | November 14, 2018 6:09:50 PM GMT+08:00 | Wrong series number getting generated in the exam summary page | WIP10 | WIP14 |  | Subsystem | Medium_1 | NA | Fixed | chokkalingam_r | 0 | 2 | Console Software | Console4 | NO | Closed | WIP13SegReq | Ingenuity 64 |  | WIP14 | February 1, 2019 5:24:51 PM GMT+08:00
CTUDT00600416 | Refresh | Closed | roshan_r | November 14, 2018 6:09:50 PM GMT+08:00 | Wrong series number getting generated in the exam summary page | WIP10 | WIP14 |  | Subsystem | Medium_1 | NA | Fixed | chokkalingam_r | 0 | 2 | Console Software | Console4 | NO | Closed | WIP13SegReq | IQon1 |  | WIP14 | February 1, 2019 5:24:51 PM GMT+08:00
CTUDT00600412 | Refresh | Closed | s_ra | November 14, 2018 5:37:57 PM GMT+08:00 | Under Preference in Connectivity tab "Patient's size " option gets unselected even after selecting and saving the changes. | 4.1.7 | WIP14 | the user can manually update patient size in the console | System | Medium_1 | N/A | Fixed | jose_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP13SegReq | Ovation | PR#8563081 | WIP14 | February 21, 2019 8:06:23 PM GMT+08:00
CTUDT00600412 | Refresh | Closed | s_ra | November 14, 2018 5:37:57 PM GMT+08:00 | Under Preference in Connectivity tab "Patient's size " option gets unselected even after selecting and saving the changes. | 4.1.7 | WIP14 | the user can manually update patient size in the console | System | Medium_1 | N/A | Fixed | jose_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP13SegReq | Brilliance 64 | PR#8563081 | WIP14 | February 21, 2019 8:06:23 PM GMT+08:00
CTUDT00600412 | Refresh | Closed | s_ra | November 14, 2018 5:37:57 PM GMT+08:00 | Under Preference in Connectivity tab "Patient's size " option gets unselected even after selecting and saving the changes. | 4.1.7 | WIP14 | the user can manually update patient size in the console | System | Medium_1 | N/A | Fixed | jose_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP13SegReq | Ingenuity 64 | PR#8563081 | WIP14 | February 21, 2019 8:06:23 PM GMT+08:00
CTUDT00600412 | Refresh | Closed | s_ra | November 14, 2018 5:37:57 PM GMT+08:00 | Under Preference in Connectivity tab "Patient's size " option gets unselected even after selecting and saving the changes. | 4.1.7 | WIP14 | the user can manually update patient size in the console | System | Medium_1 | N/A | Fixed | jose_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP13SegReq | IQon1 | PR#8563081 | WIP14 | February 21, 2019 8:06:23 PM GMT+08:00
CTUDT00600408 | Refresh | Closed | zhong_x | November 14, 2018 2:38:54 PM GMT+08:00 | Recon box cannot initialized for IMR configuration | BB_WIP1.1 | BB_WIP4 |  | Subsystem | IMR reconstruction | NA | Fixed | v_sumanth | 0 | 2 | Console Software | Console4 | YES | Closed | BB_WIP4_SubCom | Brilliance Big Bore |  | BB_WIP4 | February 6, 2019 2:25:53 PM GMT+08:00
CTUDT00600405 | Refresh | Closed | roshan_r | November 14, 2018 1:58:28 PM GMT+08:00 | Refresh : Dose structured report cannot be sent to DVD/CD , DVD-RAM. while trying to record in CDR folder, Error pops up “CD/DVD drive is locked, because scanning is in progress!” even though left monitor is idle and no scan is in progress . | WIP10 | WIP12 | the user can save the DR to another media such as USB, or the IT user can reconfigure the drives. | Subsystem | Dose SR | NA | Fixed | veerabhadraiah | 0 | 3 | USP | iCAP | NO | Closed | WIP12SegReq | Ovation |  | WIP12 | December 14, 2018 5:15:50 PM GMT+08:00
CTUDT00600405 | Refresh | Closed | roshan_r | November 14, 2018 1:58:28 PM GMT+08:00 | Refresh : Dose structured report cannot be sent to DVD/CD , DVD-RAM. while trying to record in CDR folder, Error pops up “CD/DVD drive is locked, because scanning is in progress!” even though left monitor is idle and no scan is in progress . | WIP10 | WIP12 | the user can save the DR to another media such as USB, or the IT user can reconfigure the drives. | Subsystem | Dose SR | NA | Fixed | veerabhadraiah | 0 | 3 | USP | iCAP | NO | Closed | WIP12SegReq | Brilliance 64 |  | WIP12 | December 14, 2018 5:15:50 PM GMT+08:00
CTUDT00600405 | Refresh | Closed | roshan_r | November 14, 2018 1:58:28 PM GMT+08:00 | Refresh : Dose structured report cannot be sent to DVD/CD , DVD-RAM. while trying to record in CDR folder, Error pops up “CD/DVD drive is locked, because scanning is in progress!” even though left monitor is idle and no scan is in progress . | WIP10 | WIP12 | the user can save the DR to another media such as USB, or the IT user can reconfigure the drives. | Subsystem | Dose SR | NA | Fixed | veerabhadraiah | 0 | 3 | USP | iCAP | NO | Closed | WIP12SegReq | Brilliance Big Bore |  | WIP12 | December 14, 2018 5:15:50 PM GMT+08:00
CTUDT00600405 | Refresh | Closed | roshan_r | November 14, 2018 1:58:28 PM GMT+08:00 | Refresh : Dose structured report cannot be sent to DVD/CD , DVD-RAM. while trying to record in CDR folder, Error pops up “CD/DVD drive is locked, because scanning is in progress!” even though left monitor is idle and no scan is in progress . | WIP10 | WIP12 | the user can save the DR to another media such as USB, or the IT user can reconfigure the drives. | Subsystem | Dose SR | NA | Fixed | veerabhadraiah | 0 | 3 | USP | iCAP | NO | Closed | WIP12SegReq | Ingenuity 64 |  | WIP12 | December 14, 2018 5:15:50 PM GMT+08:00
CTUDT00600405 | Refresh | Closed | roshan_r | November 14, 2018 1:58:28 PM GMT+08:00 | Refresh : Dose structured report cannot be sent to DVD/CD , DVD-RAM. while trying to record in CDR folder, Error pops up “CD/DVD drive is locked, because scanning is in progress!” even though left monitor is idle and no scan is in progress . | WIP10 | WIP12 | the user can save the DR to another media such as USB, or the IT user can reconfigure the drives. | Subsystem | Dose SR | NA | Fixed | veerabhadraiah | 0 | 3 | USP | iCAP | NO | Closed | WIP12SegReq | IQon1 |  | WIP12 | December 14, 2018 5:15:50 PM GMT+08:00
CTUDT00600404 | Refresh | Closed | pB_s | November 14, 2018 1:55:30 PM GMT+08:00 | WIP11 : Incorrect message is getting displayed while performing the password reset during first time login for newly created user | WIP11 | WIP15 | No workaround. To avoid this user has to enter valid credentials which is according to complexity requirements. | Subsystem | High_2-Password Management | NA | Fixed | kapadia_v | 0 | 4 | Console Software | Console4 | NO | Closed | WIP12SubCom | Ingenuity 64 |  | WIP15 | April 4, 2019 8:14:35 PM GMT+08:00
CTUDT00600401 | Refresh | Closed | pB_s | November 14, 2018 1:14:38 PM GMT+08:00 | WIP11: While creating user in IT tools entered password is not following password complexity rules and the value set in password length field is resetting back to 8 instead of retaining the value set by the user. | WIP11 | WIP12 | No workaround | Subsystem | Password Management | NA | Fixed | ravindra_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP12SubCom | Ovation |  | WIP12 | December 5, 2018 6:41:30 PM GMT+08:00
CTUDT00600401 | Refresh | Closed | pB_s | November 14, 2018 1:14:38 PM GMT+08:00 | WIP11: While creating user in IT tools entered password is not following password complexity rules and the value set in password length field is resetting back to 8 instead of retaining the value set by the user. | WIP11 | WIP12 | No workaround | Subsystem | Password Management | NA | Fixed | ravindra_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP12SubCom | Brilliance 64 |  | WIP12 | December 5, 2018 6:41:30 PM GMT+08:00
CTUDT00600401 | Refresh | Closed | pB_s | November 14, 2018 1:14:38 PM GMT+08:00 | WIP11: While creating user in IT tools entered password is not following password complexity rules and the value set in password length field is resetting back to 8 instead of retaining the value set by the user. | WIP11 | WIP12 | No workaround | Subsystem | Password Management | NA | Fixed | ravindra_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP12SubCom | Brilliance Big Bore |  | WIP12 | December 5, 2018 6:41:30 PM GMT+08:00
CTUDT00600401 | Refresh | Closed | pB_s | November 14, 2018 1:14:38 PM GMT+08:00 | WIP11: While creating user in IT tools entered password is not following password complexity rules and the value set in password length field is resetting back to 8 instead of retaining the value set by the user. | WIP11 | WIP12 | No workaround | Subsystem | Password Management | NA | Fixed | ravindra_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP12SubCom | Ingenuity 64 |  | WIP12 | December 5, 2018 6:41:30 PM GMT+08:00
CTUDT00600401 | Refresh | Closed | pB_s | November 14, 2018 1:14:38 PM GMT+08:00 | WIP11: While creating user in IT tools entered password is not following password complexity rules and the value set in password length field is resetting back to 8 instead of retaining the value set by the user. | WIP11 | WIP12 | No workaround | Subsystem | Password Management | NA | Fixed | ravindra_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP12SubCom | IQon1 |  | WIP12 | December 5, 2018 6:41:30 PM GMT+08:00
CTUDT00600399 | Refresh | Closed | zhang_t | November 14, 2018 1:08:35 PM GMT+08:00 | Refresh: ING, Body Helical Perfusion scans getting stuck while clinical scan is going on. | WIP11 | WIP12 |  | Subsystem | Body Helical Perfusion | NA | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | YES | Closed | WIP12SubCom | Ovation |  | WIP12 | December 11, 2018 6:46:51 PM GMT+08:00
CTUDT00600399 | Refresh | Closed | zhang_t | November 14, 2018 1:08:35 PM GMT+08:00 | Refresh: ING, Body Helical Perfusion scans getting stuck while clinical scan is going on. | WIP11 | WIP12 |  | Subsystem | Body Helical Perfusion | NA | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | YES | Closed | WIP12SubCom | Brilliance 64 |  | WIP12 | December 11, 2018 6:46:51 PM GMT+08:00
CTUDT00600399 | Refresh | Closed | zhang_t | November 14, 2018 1:08:35 PM GMT+08:00 | Refresh: ING, Body Helical Perfusion scans getting stuck while clinical scan is going on. | WIP11 | WIP12 |  | Subsystem | Body Helical Perfusion | NA | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | YES | Closed | WIP12SubCom | Ingenuity 64 |  | WIP12 | December 11, 2018 6:46:51 PM GMT+08:00
CTUDT00600399 | Refresh | Closed | zhang_t | November 14, 2018 1:08:35 PM GMT+08:00 | Refresh: ING, Body Helical Perfusion scans getting stuck while clinical scan is going on. | WIP11 | WIP12 |  | Subsystem | Body Helical Perfusion | NA | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | YES | Closed | WIP12SubCom | IQon1 |  | WIP12 | December 11, 2018 6:46:51 PM GMT+08:00
CTUDT00600396 | Refresh | Closed | zhang_t | November 14, 2018 12:34:38 PM GMT+08:00 | Refresh: ING, Resp.4D CT EC is displayed in EC group but the selected exam cards cannot be used when loading it to plan while Basic license option Key is used. | WIP11 | WIP14 | No | Subsystem | 4D CT Scan | N/A | Fixed | prasad_r | 0 | 4 | Console Software | Console4 | NO | Closed | WIP14SubCom | Brilliance 64 |  | WIP14 | February 12, 2019 5:49:26 PM GMT+08:00
CTUDT00600396 | Refresh | Closed | zhang_t | November 14, 2018 12:34:38 PM GMT+08:00 | Refresh: ING, Resp.4D CT EC is displayed in EC group but the selected exam cards cannot be used when loading it to plan while Basic license option Key is used. | WIP11 | WIP14 | No | Subsystem | 4D CT Scan | N/A | Fixed | prasad_r | 0 | 4 | Console Software | Console4 | NO | Closed | WIP14SubCom | Ingenuity 64 |  | WIP14 | February 12, 2019 5:49:26 PM GMT+08:00
CTUDT00600395 | Refresh | Closed | pB_s | November 14, 2018 12:33:22 PM GMT+08:00 | WIP11: Message displayed while creating username with set of special characters in IT tools is mismatching. | WIP11 | DHF303211 Rev 06 |  | Subsystem | Password Management | console-SRS.16193 | Fixed | hatti_c | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | WIP12SegReq | Ovation |  | DHF303211 Rev 06 | January 30, 2019 7:18:37 PM GMT+08:00
CTUDT00600395 | Refresh | Closed | pB_s | November 14, 2018 12:33:22 PM GMT+08:00 | WIP11: Message displayed while creating username with set of special characters in IT tools is mismatching. | WIP11 | DHF303211 Rev 06 |  | Subsystem | Password Management | console-SRS.16193 | Fixed | hatti_c | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | WIP12SegReq | Brilliance 64 |  | DHF303211 Rev 06 | January 30, 2019 7:18:37 PM GMT+08:00
CTUDT00600395 | Refresh | Closed | pB_s | November 14, 2018 12:33:22 PM GMT+08:00 | WIP11: Message displayed while creating username with set of special characters in IT tools is mismatching. | WIP11 | DHF303211 Rev 06 |  | Subsystem | Password Management | console-SRS.16193 | Fixed | hatti_c | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | WIP12SegReq | Brilliance Big Bore |  | DHF303211 Rev 06 | January 30, 2019 7:18:37 PM GMT+08:00
CTUDT00600395 | Refresh | Closed | pB_s | November 14, 2018 12:33:22 PM GMT+08:00 | WIP11: Message displayed while creating username with set of special characters in IT tools is mismatching. | WIP11 | DHF303211 Rev 06 |  | Subsystem | Password Management | console-SRS.16193 | Fixed | hatti_c | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | WIP12SegReq | Ingenuity 64 |  | DHF303211 Rev 06 | January 30, 2019 7:18:37 PM GMT+08:00
CTUDT00600395 | Refresh | Closed | pB_s | November 14, 2018 12:33:22 PM GMT+08:00 | WIP11: Message displayed while creating username with set of special characters in IT tools is mismatching. | WIP11 | DHF303211 Rev 06 |  | Subsystem | Password Management | console-SRS.16193 | Fixed | hatti_c | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | WIP12SegReq | IQon1 |  | DHF303211 Rev 06 | January 30, 2019 7:18:37 PM GMT+08:00
CTUDT00600389 | Refresh | Closed | manjunath_c | November 13, 2018 8:37:41 PM GMT+08:00 | Philips_Service user account is locked | WIP11 | WIP11.1 | NA | Subsystem | High_1-User login | NA | Fixed | kapadia_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP11.1SubCom | Ovation |  | WIP11.1 | December 4, 2018 12:24:51 PM GMT+08:00
CTUDT00600389 | Refresh | Closed | manjunath_c | November 13, 2018 8:37:41 PM GMT+08:00 | Philips_Service user account is locked | WIP11 | WIP11.1 | NA | Subsystem | High_1-User login | NA | Fixed | kapadia_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP11.1SubCom | Brilliance 64 |  | WIP11.1 | December 4, 2018 12:24:51 PM GMT+08:00
CTUDT00600389 | Refresh | Closed | manjunath_c | November 13, 2018 8:37:41 PM GMT+08:00 | Philips_Service user account is locked | WIP11 | WIP11.1 | NA | Subsystem | High_1-User login | NA | Fixed | kapadia_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP11.1SubCom | Ingenuity 64 |  | WIP11.1 | December 4, 2018 12:24:51 PM GMT+08:00
CTUDT00600389 | Refresh | Closed | manjunath_c | November 13, 2018 8:37:41 PM GMT+08:00 | Philips_Service user account is locked | WIP11 | WIP11.1 | NA | Subsystem | High_1-User login | NA | Fixed | kapadia_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP11.1SubCom | IQon1 |  | WIP11.1 | December 4, 2018 12:24:51 PM GMT+08:00
CTUDT00600386 | Refresh | Closed | pB_s | November 13, 2018 7:55:04 PM GMT+08:00 | WIP11: Not following any password complexity rules while resetting password and Enable password complexity" field and password length field values are not retaining their original value set. | WIP11 | WIP12 | No workaround | Subsystem | High_1- Password Management | NA | Fixed | ravindra_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP12SubCom | Ovation |  | WIP12 | December 5, 2018 9:15:14 PM GMT+08:00
CTUDT00600386 | Refresh | Closed | pB_s | November 13, 2018 7:55:04 PM GMT+08:00 | WIP11: Not following any password complexity rules while resetting password and Enable password complexity" field and password length field values are not retaining their original value set. | WIP11 | WIP12 | No workaround | Subsystem | High_1- Password Management | NA | Fixed | ravindra_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP12SubCom | Brilliance 64 |  | WIP12 | December 5, 2018 9:15:14 PM GMT+08:00
CTUDT00600386 | Refresh | Closed | pB_s | November 13, 2018 7:55:04 PM GMT+08:00 | WIP11: Not following any password complexity rules while resetting password and Enable password complexity" field and password length field values are not retaining their original value set. | WIP11 | WIP12 | No workaround | Subsystem | High_1- Password Management | NA | Fixed | ravindra_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP12SubCom | Brilliance Big Bore |  | WIP12 | December 5, 2018 9:15:14 PM GMT+08:00
CTUDT00600386 | Refresh | Closed | pB_s | November 13, 2018 7:55:04 PM GMT+08:00 | WIP11: Not following any password complexity rules while resetting password and Enable password complexity" field and password length field values are not retaining their original value set. | WIP11 | WIP12 | No workaround | Subsystem | High_1- Password Management | NA | Fixed | ravindra_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP12SubCom | Ingenuity 64 |  | WIP12 | December 5, 2018 9:15:14 PM GMT+08:00
CTUDT00600386 | Refresh | Closed | pB_s | November 13, 2018 7:55:04 PM GMT+08:00 | WIP11: Not following any password complexity rules while resetting password and Enable password complexity" field and password length field values are not retaining their original value set. | WIP11 | WIP12 | No workaround | Subsystem | High_1- Password Management | NA | Fixed | ravindra_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP12SubCom | IQon1 |  | WIP12 | December 5, 2018 9:15:14 PM GMT+08:00
CTUDT00600384 | Refresh | Closed | pB_s | November 13, 2018 7:43:07 PM GMT+08:00 | WIP11: User is not able to launch User account backup and restore tool in service tools | WIP11 | WIP12 | Users can be backed up/ restored using IT tool. | Subsystem | User Account Backup/Restore | NA | Fixed | ravindra_a | 0 | 3 | Console Software | Console4 | NO | Closed | WIP12SubCom | Ovation |  | WIP12 | December 6, 2018 9:09:25 PM GMT+08:00
CTUDT00600384 | Refresh | Closed | pB_s | November 13, 2018 7:43:07 PM GMT+08:00 | WIP11: User is not able to launch User account backup and restore tool in service tools | WIP11 | WIP12 | Users can be backed up/ restored using IT tool. | Subsystem | User Account Backup/Restore | NA | Fixed | ravindra_a | 0 | 3 | Console Software | Console4 | NO | Closed | WIP12SubCom | Brilliance 64 |  | WIP12 | December 6, 2018 9:09:25 PM GMT+08:00
CTUDT00600384 | Refresh | Closed | pB_s | November 13, 2018 7:43:07 PM GMT+08:00 | WIP11: User is not able to launch User account backup and restore tool in service tools | WIP11 | WIP12 | Users can be backed up/ restored using IT tool. | Subsystem | User Account Backup/Restore | NA | Fixed | ravindra_a | 0 | 3 | Console Software | Console4 | NO | Closed | WIP12SubCom | Brilliance Big Bore |  | WIP12 | December 6, 2018 9:09:25 PM GMT+08:00
CTUDT00600384 | Refresh | Closed | pB_s | November 13, 2018 7:43:07 PM GMT+08:00 | WIP11: User is not able to launch User account backup and restore tool in service tools | WIP11 | WIP12 | Users can be backed up/ restored using IT tool. | Subsystem | User Account Backup/Restore | NA | Fixed | ravindra_a | 0 | 3 | Console Software | Console4 | NO | Closed | WIP12SubCom | Ingenuity 64 |  | WIP12 | December 6, 2018 9:09:25 PM GMT+08:00
CTUDT00600384 | Refresh | Closed | pB_s | November 13, 2018 7:43:07 PM GMT+08:00 | WIP11: User is not able to launch User account backup and restore tool in service tools | WIP11 | WIP12 | Users can be backed up/ restored using IT tool. | Subsystem | User Account Backup/Restore | NA | Fixed | ravindra_a | 0 | 3 | Console Software | Console4 | NO | Closed | WIP12SubCom | IQon1 |  | WIP12 | December 6, 2018 9:09:25 PM GMT+08:00
CTUDT00600377 | Refresh | Closed | zhang_t | November 13, 2018 2:13:02 PM GMT+08:00 | Refresh: CirsServiceablity stop working and non standard installation displayed in Installation Validation tool | WIP11 | WIP14 | See CIRS Release Notes DHF303418 for workaround regarding setting the number of nodes/servers during installation. Other workaround do not exist. | Subsystem | Installation Validation Tool; Verified1 | N/A | Fixed | rakic_k | 0 | 2 | CIRS_SW_ServiceNode | CIRS | NO | Closed | WIP14SubCom | Ovation |  | WIP14 | March 11, 2019 9:32:26 PM GMT+08:00
CTUDT00600377 | Refresh | Closed | zhang_t | November 13, 2018 2:13:02 PM GMT+08:00 | Refresh: CirsServiceablity stop working and non standard installation displayed in Installation Validation tool | WIP11 | WIP14 | See CIRS Release Notes DHF303418 for workaround regarding setting the number of nodes/servers during installation. Other workaround do not exist. | Subsystem | Installation Validation Tool; Verified1 | N/A | Fixed | rakic_k | 0 | 2 | CIRS_SW_ServiceNode | CIRS | NO | Closed | WIP14SubCom | Brilliance 64 |  | WIP14 | March 11, 2019 9:32:26 PM GMT+08:00
CTUDT00600377 | Refresh | Closed | zhang_t | November 13, 2018 2:13:02 PM GMT+08:00 | Refresh: CirsServiceablity stop working and non standard installation displayed in Installation Validation tool | WIP11 | WIP14 | See CIRS Release Notes DHF303418 for workaround regarding setting the number of nodes/servers during installation. Other workaround do not exist. | Subsystem | Installation Validation Tool; Verified1 | N/A | Fixed | rakic_k | 0 | 2 | CIRS_SW_ServiceNode | CIRS | NO | Closed | WIP14SubCom | Brilliance Big Bore |  | WIP14 | March 11, 2019 9:32:26 PM GMT+08:00
CTUDT00600377 | Refresh | Closed | zhang_t | November 13, 2018 2:13:02 PM GMT+08:00 | Refresh: CirsServiceablity stop working and non standard installation displayed in Installation Validation tool | WIP11 | WIP14 | See CIRS Release Notes DHF303418 for workaround regarding setting the number of nodes/servers during installation. Other workaround do not exist. | Subsystem | Installation Validation Tool; Verified1 | N/A | Fixed | rakic_k | 0 | 2 | CIRS_SW_ServiceNode | CIRS | NO | Closed | WIP14SubCom | Ingenuity 64 |  | WIP14 | March 11, 2019 9:32:26 PM GMT+08:00
CTUDT00600377 | Refresh | Closed | zhang_t | November 13, 2018 2:13:02 PM GMT+08:00 | Refresh: CirsServiceablity stop working and non standard installation displayed in Installation Validation tool | WIP11 | WIP14 | See CIRS Release Notes DHF303418 for workaround regarding setting the number of nodes/servers during installation. Other workaround do not exist. | Subsystem | Installation Validation Tool; Verified1 | N/A | Fixed | rakic_k | 0 | 2 | CIRS_SW_ServiceNode | CIRS | NO | Closed | WIP14SubCom | IQon1 |  | WIP14 | March 11, 2019 9:32:26 PM GMT+08:00
CTUDT00600371 | Refresh | Closed | naik_p | November 12, 2018 6:06:13 PM GMT+08:00 | Refresh_Suggested Exam Cards group had more than one linking | WIP10 | WIP12 | The user is able to change the name or parameters of the EC and resave it | Subsystem | ECM | N/A | Fixed | prasad_r | 0 | 3 | Console Software | Console4 | YES | Closed | WIP12SubCom | Ovation |  | WIP12 | December 10, 2018 8:44:16 PM GMT+08:00
CTUDT00600371 | Refresh | Closed | naik_p | November 12, 2018 6:06:13 PM GMT+08:00 | Refresh_Suggested Exam Cards group had more than one linking | WIP10 | WIP12 | The user is able to change the name or parameters of the EC and resave it | Subsystem | ECM | N/A | Fixed | prasad_r | 0 | 3 | Console Software | Console4 | YES | Closed | WIP12SubCom | Brilliance 64 |  | WIP12 | December 10, 2018 8:44:16 PM GMT+08:00
CTUDT00600371 | Refresh | Closed | naik_p | November 12, 2018 6:06:13 PM GMT+08:00 | Refresh_Suggested Exam Cards group had more than one linking | WIP10 | WIP12 | The user is able to change the name or parameters of the EC and resave it | Subsystem | ECM | N/A | Fixed | prasad_r | 0 | 3 | Console Software | Console4 | YES | Closed | WIP12SubCom | Ingenuity 64 |  | WIP12 | December 10, 2018 8:44:16 PM GMT+08:00
CTUDT00600366 | Refresh | Closed | datta_i | November 12, 2018 4:28:06 PM GMT+08:00 | Refresh [Scanner viewer]:The application is getting crashed and displayed "system error", On re-launching the same result where the main viewport & reference viewport images are being switched in slab viewer | 4.8 | BB_WIP3 |  | Subsystem | Slab Viewer | NA | Fixed | datta_i | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP3_SegReq | Brilliance Big Bore |  | BB_WIP3 | December 13, 2018 8:16:04 PM GMT+08:00
CTUDT00600363 | Refresh | Closed | zhu_je | November 12, 2018 2:34:50 PM GMT+08:00 | Cannot modify "display" value if pause one CCT scan then duplicate the un-completed CCT scan | WIP11 | WIP15 | the user can change the display through Show All | Subsystem | CCT | NA | Fixed | datta_i | 0 | 4 | Console Software | Console4 | NO | Closed | WIP13SubCom | Ovation |  | WIP15 | March 15, 2019 9:15:22 PM GMT+08:00
CTUDT00600363 | Refresh | Closed | zhu_je | November 12, 2018 2:34:50 PM GMT+08:00 | Cannot modify "display" value if pause one CCT scan then duplicate the un-completed CCT scan | WIP11 | WIP15 | the user can change the display through Show All | Subsystem | CCT | NA | Fixed | datta_i | 0 | 4 | Console Software | Console4 | NO | Closed | WIP13SubCom | Brilliance 64 |  | WIP15 | March 15, 2019 9:15:22 PM GMT+08:00
CTUDT00600363 | Refresh | Closed | zhu_je | November 12, 2018 2:34:50 PM GMT+08:00 | Cannot modify "display" value if pause one CCT scan then duplicate the un-completed CCT scan | WIP11 | WIP15 | the user can change the display through Show All | Subsystem | CCT | NA | Fixed | datta_i | 0 | 4 | Console Software | Console4 | NO | Closed | WIP13SubCom | Ingenuity 64 |  | WIP15 | March 15, 2019 9:15:22 PM GMT+08:00
CTUDT00600363 | Refresh | Closed | zhu_je | November 12, 2018 2:34:50 PM GMT+08:00 | Cannot modify "display" value if pause one CCT scan then duplicate the un-completed CCT scan | WIP11 | WIP15 | the user can change the display through Show All | Subsystem | CCT | NA | Fixed | datta_i | 0 | 4 | Console Software | Console4 | NO | Closed | WIP13SubCom | IQon1 |  | WIP15 | March 15, 2019 9:15:22 PM GMT+08:00
CTUDT00600361 | Refresh | Closed | zhu_je | November 12, 2018 2:27:02 PM GMT+08:00 | The added non-Jog perfusion scan will be changed to jog perfusion automatically after first jog perfusion's surview scan finished | WIP11 | WIP12 |  | Subsystem | Jog/Non Jog axial - Log Viewer | NA | Fixed | kumar_v | 0 | 2 | Console Software | Console4 | YES | Closed | WIP12SubCom | Ovation |  | WIP12 | December 7, 2018 3:05:22 PM GMT+08:00
CTUDT00600361 | Refresh | Closed | zhu_je | November 12, 2018 2:27:02 PM GMT+08:00 | The added non-Jog perfusion scan will be changed to jog perfusion automatically after first jog perfusion's surview scan finished | WIP11 | WIP12 |  | Subsystem | Jog/Non Jog axial - Log Viewer | NA | Fixed | kumar_v | 0 | 2 | Console Software | Console4 | YES | Closed | WIP12SubCom | Brilliance 64 |  | WIP12 | December 7, 2018 3:05:22 PM GMT+08:00
CTUDT00600361 | Refresh | Closed | zhu_je | November 12, 2018 2:27:02 PM GMT+08:00 | The added non-Jog perfusion scan will be changed to jog perfusion automatically after first jog perfusion's surview scan finished | WIP11 | WIP12 |  | Subsystem | Jog/Non Jog axial - Log Viewer | NA | Fixed | kumar_v | 0 | 2 | Console Software | Console4 | YES | Closed | WIP12SubCom | Ingenuity 64 |  | WIP12 | December 7, 2018 3:05:22 PM GMT+08:00
CTUDT00600361 | Refresh | Closed | zhu_je | November 12, 2018 2:27:02 PM GMT+08:00 | The added non-Jog perfusion scan will be changed to jog perfusion automatically after first jog perfusion's surview scan finished | WIP11 | WIP12 |  | Subsystem | Jog/Non Jog axial - Log Viewer | NA | Fixed | kumar_v | 0 | 2 | Console Software | Console4 | YES | Closed | WIP12SubCom | IQon1 |  | WIP12 | December 7, 2018 3:05:22 PM GMT+08:00
CTUDT00600357 | Refresh | Closed | zhang_t | November 12, 2018 1:42:22 PM GMT+08:00 | Refresh: ING, Recon Box cannot initialize error when performing Helical Brain EC | WIP11 | WIP12 |  | Subsystem | Reconstruction - Brain Helical | NA | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP12SubCom | Ovation |  | WIP12 | December 17, 2018 4:48:01 PM GMT+08:00
CTUDT00600357 | Refresh | Closed | zhang_t | November 12, 2018 1:42:22 PM GMT+08:00 | Refresh: ING, Recon Box cannot initialize error when performing Helical Brain EC | WIP11 | WIP12 |  | Subsystem | Reconstruction - Brain Helical | NA | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP12SubCom | Brilliance 64 |  | WIP12 | December 17, 2018 4:48:01 PM GMT+08:00
CTUDT00600357 | Refresh | Closed | zhang_t | November 12, 2018 1:42:22 PM GMT+08:00 | Refresh: ING, Recon Box cannot initialize error when performing Helical Brain EC | WIP11 | WIP12 |  | Subsystem | Reconstruction - Brain Helical | NA | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP12SubCom | Ingenuity 64 |  | WIP12 | December 17, 2018 4:48:01 PM GMT+08:00
CTUDT00600354 | Refresh | Closed | zhang_t | November 12, 2018 1:30:48 PM GMT+08:00 | Refresh: While 1 click installation, "Windows Firewall has blocked some features of server MFC application on all public and private networks." popup at 50% installation progress | WIP11 | WIP13 |  | Subsystem | 1-click installation | None | Fixed | jotawar_d | 0 | 3 | Console Software | Console4 | NO | Closed | WIP13SubCom | Ovation |  | WIP13 | February 14, 2019 8:17:25 PM GMT+08:00
CTUDT00600354 | Refresh | Closed | zhang_t | November 12, 2018 1:30:48 PM GMT+08:00 | Refresh: While 1 click installation, "Windows Firewall has blocked some features of server MFC application on all public and private networks." popup at 50% installation progress | WIP11 | WIP13 |  | Subsystem | 1-click installation | None | Fixed | jotawar_d | 0 | 3 | Console Software | Console4 | NO | Closed | WIP13SubCom | Brilliance 64 |  | WIP13 | February 14, 2019 8:17:25 PM GMT+08:00
CTUDT00600354 | Refresh | Closed | zhang_t | November 12, 2018 1:30:48 PM GMT+08:00 | Refresh: While 1 click installation, "Windows Firewall has blocked some features of server MFC application on all public and private networks." popup at 50% installation progress | WIP11 | WIP13 |  | Subsystem | 1-click installation | None | Fixed | jotawar_d | 0 | 3 | Console Software | Console4 | NO | Closed | WIP13SubCom | Brilliance Big Bore |  | WIP13 | February 14, 2019 8:17:25 PM GMT+08:00
CTUDT00600354 | Refresh | Closed | zhang_t | November 12, 2018 1:30:48 PM GMT+08:00 | Refresh: While 1 click installation, "Windows Firewall has blocked some features of server MFC application on all public and private networks." popup at 50% installation progress | WIP11 | WIP13 |  | Subsystem | 1-click installation | None | Fixed | jotawar_d | 0 | 3 | Console Software | Console4 | NO | Closed | WIP13SubCom | Ingenuity 64 |  | WIP13 | February 14, 2019 8:17:25 PM GMT+08:00
CTUDT00600354 | Refresh | Closed | zhang_t | November 12, 2018 1:30:48 PM GMT+08:00 | Refresh: While 1 click installation, "Windows Firewall has blocked some features of server MFC application on all public and private networks." popup at 50% installation progress | WIP11 | WIP13 |  | Subsystem | 1-click installation | None | Fixed | jotawar_d | 0 | 3 | Console Software | Console4 | NO | Closed | WIP13SubCom | IQon1 |  | WIP13 | February 14, 2019 8:17:25 PM GMT+08:00
CTUDT00600351 | Refresh | Closed | zhang_t | November 12, 2018 1:28:01 PM GMT+08:00 | Refresh: While 1 click installation, "Security settings Setup has stopped working” occurred at 85% progress. | WIP11 | WIP15 | Yes - No action is required from the user and the installation would continue. Once the system restart this alert would disapear. | Subsystem | 1-click installation | NA | Fixed | franco_a | 0 | 3 | DMS Software | DMS | NO | Closed | WIP14SubCom | Ovation |  | WIP15 | March 28, 2019 7:41:10 PM GMT+08:00
CTUDT00600351 | Refresh | Closed | zhang_t | November 12, 2018 1:28:01 PM GMT+08:00 | Refresh: While 1 click installation, "Security settings Setup has stopped working” occurred at 85% progress. | WIP11 | WIP15 | Yes - No action is required from the user and the installation would continue. Once the system restart this alert would disapear. | Subsystem | 1-click installation | NA | Fixed | franco_a | 0 | 3 | DMS Software | DMS | NO | Closed | WIP14SubCom | Brilliance 64 |  | WIP15 | March 28, 2019 7:41:10 PM GMT+08:00
CTUDT00600351 | Refresh | Closed | zhang_t | November 12, 2018 1:28:01 PM GMT+08:00 | Refresh: While 1 click installation, "Security settings Setup has stopped working” occurred at 85% progress. | WIP11 | WIP15 | Yes - No action is required from the user and the installation would continue. Once the system restart this alert would disapear. | Subsystem | 1-click installation | NA | Fixed | franco_a | 0 | 3 | DMS Software | DMS | NO | Closed | WIP14SubCom | Brilliance Big Bore |  | WIP15 | March 28, 2019 7:41:10 PM GMT+08:00
CTUDT00600351 | Refresh | Closed | zhang_t | November 12, 2018 1:28:01 PM GMT+08:00 | Refresh: While 1 click installation, "Security settings Setup has stopped working” occurred at 85% progress. | WIP11 | WIP15 | Yes - No action is required from the user and the installation would continue. Once the system restart this alert would disapear. | Subsystem | 1-click installation | NA | Fixed | franco_a | 0 | 3 | DMS Software | DMS | NO | Closed | WIP14SubCom | Ingenuity 64 |  | WIP15 | March 28, 2019 7:41:10 PM GMT+08:00
CTUDT00600351 | Refresh | Closed | zhang_t | November 12, 2018 1:28:01 PM GMT+08:00 | Refresh: While 1 click installation, "Security settings Setup has stopped working” occurred at 85% progress. | WIP11 | WIP15 | Yes - No action is required from the user and the installation would continue. Once the system restart this alert would disapear. | Subsystem | 1-click installation | NA | Fixed | franco_a | 0 | 3 | DMS Software | DMS | NO | Closed | WIP14SubCom | IQon1 |  | WIP15 | March 28, 2019 7:41:10 PM GMT+08:00
CTUDT00600347 | Refresh | Closed | zhu_je | November 12, 2018 11:13:02 AM GMT+08:00 | There is error message popping after click go button with resolution=Ultra High | WIP11 |  |  | Subsystem | Ultra High Resolution |  | Works As Specified / Not A Defect | prasad_r | 1 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00600344 | Refresh | Closed | zhu_je | November 12, 2018 10:37:41 AM GMT+08:00 | The message in log viewer is not reasonable if group is exceed the limitation | WIP11 | WIP12 |  | Subsystem | Jog/Non Jog axial - Log Viewer | NA | Fixed | kumar_v | 0 | 4 | Console Software | Console4 | NO | Closed | WIP12SubCom | Ovation |  | WIP12 | December 6, 2018 5:58:50 PM GMT+08:00
CTUDT00600344 | Refresh | Closed | zhu_je | November 12, 2018 10:37:41 AM GMT+08:00 | The message in log viewer is not reasonable if group is exceed the limitation | WIP11 | WIP12 |  | Subsystem | Jog/Non Jog axial - Log Viewer | NA | Fixed | kumar_v | 0 | 4 | Console Software | Console4 | NO | Closed | WIP12SubCom | Brilliance 64 |  | WIP12 | December 6, 2018 5:58:50 PM GMT+08:00
CTUDT00600344 | Refresh | Closed | zhu_je | November 12, 2018 10:37:41 AM GMT+08:00 | The message in log viewer is not reasonable if group is exceed the limitation | WIP11 | WIP12 |  | Subsystem | Jog/Non Jog axial - Log Viewer | NA | Fixed | kumar_v | 0 | 4 | Console Software | Console4 | NO | Closed | WIP12SubCom | Ingenuity 64 |  | WIP12 | December 6, 2018 5:58:50 PM GMT+08:00
CTUDT00600344 | Refresh | Closed | zhu_je | November 12, 2018 10:37:41 AM GMT+08:00 | The message in log viewer is not reasonable if group is exceed the limitation | WIP11 | WIP12 |  | Subsystem | Jog/Non Jog axial - Log Viewer | NA | Fixed | kumar_v | 0 | 4 | Console Software | Console4 | NO | Closed | WIP12SubCom | IQon1 |  | WIP12 | December 6, 2018 5:58:50 PM GMT+08:00
CTUDT00600336 | Refresh | Closed | panchami_pradeep | November 9, 2018 6:40:00 PM GMT+08:00 | Refresh: While 1 click installation, selecting "Format drive C: drive only" option is throwing error if the system is encrypted. | WIP11 | WIP14 |  | Subsystem | 1-click installation | STPS-SSRS-1302 | Fixed | venugopal_s | 0 | 2 | Requirements Documentation | DI | NO | Closed | WIP14SegReq | Ovation |  | WIP14 | March 11, 2019 4:20:22 PM GMT+08:00
CTUDT00600336 | Refresh | Closed | panchami_pradeep | November 9, 2018 6:40:00 PM GMT+08:00 | Refresh: While 1 click installation, selecting "Format drive C: drive only" option is throwing error if the system is encrypted. | WIP11 | WIP14 |  | Subsystem | 1-click installation | STPS-SSRS-1302 | Fixed | venugopal_s | 0 | 2 | Requirements Documentation | DI | NO | Closed | WIP14SegReq | Brilliance 64 |  | WIP14 | March 11, 2019 4:20:22 PM GMT+08:00
CTUDT00600336 | Refresh | Closed | panchami_pradeep | November 9, 2018 6:40:00 PM GMT+08:00 | Refresh: While 1 click installation, selecting "Format drive C: drive only" option is throwing error if the system is encrypted. | WIP11 | WIP14 |  | Subsystem | 1-click installation | STPS-SSRS-1302 | Fixed | venugopal_s | 0 | 2 | Requirements Documentation | DI | NO | Closed | WIP14SegReq | Brilliance Big Bore |  | WIP14 | March 11, 2019 4:20:22 PM GMT+08:00
CTUDT00600336 | Refresh | Closed | panchami_pradeep | November 9, 2018 6:40:00 PM GMT+08:00 | Refresh: While 1 click installation, selecting "Format drive C: drive only" option is throwing error if the system is encrypted. | WIP11 | WIP14 |  | Subsystem | 1-click installation | STPS-SSRS-1302 | Fixed | venugopal_s | 0 | 2 | Requirements Documentation | DI | NO | Closed | WIP14SegReq | Ingenuity 64 |  | WIP14 | March 11, 2019 4:20:22 PM GMT+08:00
CTUDT00600336 | Refresh | Closed | panchami_pradeep | November 9, 2018 6:40:00 PM GMT+08:00 | Refresh: While 1 click installation, selecting "Format drive C: drive only" option is throwing error if the system is encrypted. | WIP11 | WIP14 |  | Subsystem | 1-click installation | STPS-SSRS-1302 | Fixed | venugopal_s | 0 | 2 | Requirements Documentation | DI | NO | Closed | WIP14SegReq | IQon1 |  | WIP14 | March 11, 2019 4:20:22 PM GMT+08:00
CTUDT00600334 | Refresh | Closed | wang_l | November 9, 2018 11:14:41 AM GMT+08:00 | unexpected message displayed during operations in Printer Configuration tool | WIP11 |  |  | Subsystem | Printer Config Tool |  | Duplicate | Sunil_Kumar_B | 0 |  |  |  |  | Closed |  | Ingenuity 64
CTUDT00600332 | Refresh | Closed | wang_l | November 9, 2018 11:05:56 AM GMT+08:00 | Error message displayed once save lan config by olevel user and then unexpected error to open workflow | WIP10 | WIP12 |  | Subsystem | Lan Config;Verified1 | NA | Fixed | veerabhadraiah | 0 | 2 | USP | iCAP | NO | Closed | WIP12SegReq | Ovation |  | WIP12 | December 14, 2018 5:15:40 PM GMT+08:00
CTUDT00600332 | Refresh | Closed | wang_l | November 9, 2018 11:05:56 AM GMT+08:00 | Error message displayed once save lan config by olevel user and then unexpected error to open workflow | WIP10 | WIP12 |  | Subsystem | Lan Config;Verified1 | NA | Fixed | veerabhadraiah | 0 | 2 | USP | iCAP | NO | Closed | WIP12SegReq | Brilliance 64 |  | WIP12 | December 14, 2018 5:15:40 PM GMT+08:00
CTUDT00600332 | Refresh | Closed | wang_l | November 9, 2018 11:05:56 AM GMT+08:00 | Error message displayed once save lan config by olevel user and then unexpected error to open workflow | WIP10 | WIP12 |  | Subsystem | Lan Config;Verified1 | NA | Fixed | veerabhadraiah | 0 | 2 | USP | iCAP | NO | Closed | WIP12SegReq | Brilliance Big Bore |  | WIP12 | December 14, 2018 5:15:40 PM GMT+08:00
CTUDT00600332 | Refresh | Closed | wang_l | November 9, 2018 11:05:56 AM GMT+08:00 | Error message displayed once save lan config by olevel user and then unexpected error to open workflow | WIP10 | WIP12 |  | Subsystem | Lan Config;Verified1 | NA | Fixed | veerabhadraiah | 0 | 2 | USP | iCAP | NO | Closed | WIP12SegReq | Ingenuity 64 |  | WIP12 | December 14, 2018 5:15:40 PM GMT+08:00
CTUDT00600332 | Refresh | Closed | wang_l | November 9, 2018 11:05:56 AM GMT+08:00 | Error message displayed once save lan config by olevel user and then unexpected error to open workflow | WIP10 | WIP12 |  | Subsystem | Lan Config;Verified1 | NA | Fixed | veerabhadraiah | 0 | 2 | USP | iCAP | NO | Closed | WIP12SegReq | IQon1 |  | WIP12 | December 14, 2018 5:15:40 PM GMT+08:00
CTUDT00600328 | Refresh | Closed | wang_l | November 9, 2018 10:50:53 AM GMT+08:00 | Software Distribution is not found in Local Administrator account | WIP11 |  |  | Subsystem | Software Distribution Tool |  | Obsolete Function | shaham_t | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00600326 | Refresh | Closed | wei_f | November 9, 2018 10:46:54 AM GMT+08:00 | 'Can not find the log-in and log off log in Audit Trail Viewer tool. | WIP10 |  |  | Subsystem |  |  | Duplicate | wei_f | 0 |  |  |  |  | Closed |  | Ingenuity 64
CTUDT00600325 | Refresh | Closed | wang_l | November 9, 2018 10:42:07 AM GMT+08:00 | TubeHeat tool always display "Not available" even if gantry is power on for olevel user | WIP11 | BB_WIP8 WIP18 |  | Subsystem | Tube Heat Tool | NA | Fixed | venkatachalam_p | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP18SubCom | Ovation |  | BB_WIP8 WIP18 | July 5, 2019 8:46:48 PM GMT+08:00
CTUDT00600325 | Refresh | Closed | wang_l | November 9, 2018 10:42:07 AM GMT+08:00 | TubeHeat tool always display "Not available" even if gantry is power on for olevel user | WIP11 | BB_WIP8 WIP18 |  | Subsystem | Tube Heat Tool | NA | Fixed | venkatachalam_p | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP18SubCom | Brilliance 64 |  | BB_WIP8 WIP18 | July 5, 2019 8:46:48 PM GMT+08:00
CTUDT00600325 | Refresh | Closed | wang_l | November 9, 2018 10:42:07 AM GMT+08:00 | TubeHeat tool always display "Not available" even if gantry is power on for olevel user | WIP11 | BB_WIP8 WIP18 |  | Subsystem | Tube Heat Tool | NA | Fixed | venkatachalam_p | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP18SubCom | Brilliance Big Bore |  | BB_WIP8 WIP18 | July 5, 2019 8:46:48 PM GMT+08:00
CTUDT00600325 | Refresh | Closed | wang_l | November 9, 2018 10:42:07 AM GMT+08:00 | TubeHeat tool always display "Not available" even if gantry is power on for olevel user | WIP11 | BB_WIP8 WIP18 |  | Subsystem | Tube Heat Tool | NA | Fixed | venkatachalam_p | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP18SubCom | Ingenuity 64 |  | BB_WIP8 WIP18 | July 5, 2019 8:46:48 PM GMT+08:00
CTUDT00600325 | Refresh | Closed | wang_l | November 9, 2018 10:42:07 AM GMT+08:00 | TubeHeat tool always display "Not available" even if gantry is power on for olevel user | WIP11 | BB_WIP8 WIP18 |  | Subsystem | Tube Heat Tool | NA | Fixed | venkatachalam_p | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP18SubCom | IQon1 |  | BB_WIP8 WIP18 | July 5, 2019 8:46:48 PM GMT+08:00
CTUDT00600323 | Refresh | Closed | wei_f | November 9, 2018 10:35:49 AM GMT+08:00 | "Save the Batch as Movie" to USB drive failed | WIP10 |  | Once the movie file saves to Local, user can try saving it from Local to USB. | Subsystem | USP | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | Ovation
CTUDT00600323 | Refresh | Closed | wei_f | November 9, 2018 10:35:49 AM GMT+08:00 | "Save the Batch as Movie" to USB drive failed | WIP10 |  | Once the movie file saves to Local, user can try saving it from Local to USB. | Subsystem | USP | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00600323 | Refresh | Closed | wei_f | November 9, 2018 10:35:49 AM GMT+08:00 | "Save the Batch as Movie" to USB drive failed | WIP10 |  | Once the movie file saves to Local, user can try saving it from Local to USB. | Subsystem | USP | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00600323 | Refresh | Closed | wei_f | November 9, 2018 10:35:49 AM GMT+08:00 | "Save the Batch as Movie" to USB drive failed | WIP10 |  | Once the movie file saves to Local, user can try saving it from Local to USB. | Subsystem | USP | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00600323 | Refresh | Closed | wei_f | November 9, 2018 10:35:49 AM GMT+08:00 | "Save the Batch as Movie" to USB drive failed | WIP10 |  | Once the movie file saves to Local, user can try saving it from Local to USB. | Subsystem | USP | NA | No Plan to Fix | s_sa | 0 | 4 |  | iCAP | NO | Closed |  | IQon1
CTUDT00600321 | Refresh | Closed | wang_l | November 9, 2018 10:24:08 AM GMT+08:00 | Color Scheme Selection Tool not found in olevel user | WIP11 |  |  | Subsystem | "White Covers" Color scheme tool |  | Invalid/Uncontrolled Configuration | sanjeeva_a | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00600320 | Refresh | Closed | wei_f | November 9, 2018 10:23:51 AM GMT+08:00 | 'Send the display to film failed | WIP10 |  |  | Subsystem | Film |  | Works As Specified / Not A Defect | wei_f | 1 |  |  | Support-System | NO | Closed |  | Ingenuity 64
CTUDT00600314 | Refresh | Closed | Dana_Kamyshev | November 8, 2018 10:15:08 PM GMT+08:00 | Refresh - No information displayed in help page | WIP11 | WIP11.1 |  | Subsystem | Help Page | None | Fixed | Prabhakar_Priyank_Muppidi | 0 | 4 | Console Software | Console4 | NO | Closed | WIP11.1SubCom | Ovation |  | WIP11.1 | November 30, 2018 1:07:16 PM GMT+08:00
CTUDT00600314 | Refresh | Closed | Dana_Kamyshev | November 8, 2018 10:15:08 PM GMT+08:00 | Refresh - No information displayed in help page | WIP11 | WIP11.1 |  | Subsystem | Help Page | None | Fixed | Prabhakar_Priyank_Muppidi | 0 | 4 | Console Software | Console4 | NO | Closed | WIP11.1SubCom | Ingenuity 64 |  | WIP11.1 | November 30, 2018 1:07:16 PM GMT+08:00
CTUDT00600314 | Refresh | Closed | Dana_Kamyshev | November 8, 2018 10:15:08 PM GMT+08:00 | Refresh - No information displayed in help page | WIP11 | WIP11.1 |  | Subsystem | Help Page | None | Fixed | Prabhakar_Priyank_Muppidi | 0 | 4 | Console Software | Console4 | NO | Closed | WIP11.1SubCom | IQon1 |  | WIP11.1 | November 30, 2018 1:07:16 PM GMT+08:00
CTUDT00600313 | Refresh | Closed | Dana_Kamyshev | November 8, 2018 10:00:38 PM GMT+08:00 | Refresh - Unable to perform calibrations in CT user | WIP11 | WIP12 |  | Subsystem | Calibrations (CT User)-Quality Assurance - Calibra | n/a | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP12SegReq | Ovation |  | WIP12 | December 6, 2018 8:30:47 PM GMT+08:00
CTUDT00600313 | Refresh | Closed | Dana_Kamyshev | November 8, 2018 10:00:38 PM GMT+08:00 | Refresh - Unable to perform calibrations in CT user | WIP11 | WIP12 |  | Subsystem | Calibrations (CT User)-Quality Assurance - Calibra | n/a | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP12SegReq | Brilliance 64 |  | WIP12 | December 6, 2018 8:30:47 PM GMT+08:00
CTUDT00600313 | Refresh | Closed | Dana_Kamyshev | November 8, 2018 10:00:38 PM GMT+08:00 | Refresh - Unable to perform calibrations in CT user | WIP11 | WIP12 |  | Subsystem | Calibrations (CT User)-Quality Assurance - Calibra | n/a | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP12SegReq | Brilliance Big Bore |  | WIP12 | December 6, 2018 8:30:47 PM GMT+08:00
CTUDT00600313 | Refresh | Closed | Dana_Kamyshev | November 8, 2018 10:00:38 PM GMT+08:00 | Refresh - Unable to perform calibrations in CT user | WIP11 | WIP12 |  | Subsystem | Calibrations (CT User)-Quality Assurance - Calibra | n/a | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP12SegReq | Ingenuity 64 |  | WIP12 | December 6, 2018 8:30:47 PM GMT+08:00
CTUDT00600313 | Refresh | Closed | Dana_Kamyshev | November 8, 2018 10:00:38 PM GMT+08:00 | Refresh - Unable to perform calibrations in CT user | WIP11 | WIP12 |  | Subsystem | Calibrations (CT User)-Quality Assurance - Calibra | n/a | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP12SegReq | IQon1 |  | WIP12 | December 6, 2018 8:30:47 PM GMT+08:00
CTUDT00600305 | Refresh | Closed | ding_f | November 8, 2018 2:57:23 PM GMT+08:00 | Refresh - Harmonized phantom is not detected when running Phantom Centering. | WIP11 |  |  | Subsystem | Harmonized Phantom | n/a | Duplicate | ding_f | 0 |  |  | Support-System | NO | Closed |  | Ingenuity 64
CTUDT00600304 | Refresh | Closed | sowmya_k | November 8, 2018 2:18:26 PM GMT+08:00 | V4.2: Pulmo Wave Peak clipped by the maximum threshold line for Amplitude binning. | 4.1.7 | WIP12 |  | Subsystem | Pulmo - Amplitude Binning-Pulmo - Clipped Wave pea | N/A | Fixed | sowmya_k | 0 | 2 | Console Software | Console4 | YES | Closed | WIP11.1SegReq | Brilliance 64 | PR7779384 | WIP12 | December 3, 2018 7:44:37 PM GMT+08:00
CTUDT00600304 | Refresh | Closed | sowmya_k | November 8, 2018 2:18:26 PM GMT+08:00 | V4.2: Pulmo Wave Peak clipped by the maximum threshold line for Amplitude binning. | 4.1.7 | WIP12 |  | Subsystem | Pulmo - Amplitude Binning-Pulmo - Clipped Wave pea | N/A | Fixed | sowmya_k | 0 | 2 | Console Software | Console4 | YES | Closed | WIP11.1SegReq | Ingenuity 64 | PR7779384 | WIP12 | December 3, 2018 7:44:37 PM GMT+08:00
CTUDT00600300 | Refresh | Closed | zhong_x | November 7, 2018 12:46:03 PM GMT+08:00 | System was hang when loading EC. | BB_WIP1.1 | BB_WIP5 | NA | Subsystem | System Performance | Console-SRS.17098 | Fixed | madamsetti_k | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP6_SubCom | Brilliance Big Bore |  | BB_WIP5 | March 15, 2019 8:12:27 PM GMT+08:00
CTUDT00600297 | Refresh | Closed | zhong_x | November 7, 2018 9:54:22 AM GMT+08:00 | MA Framework exception occurred during Short tube conditioning and air calibration | BB_WIP1.1 | BB_WIP2 | No | System | Calibrations - Short Tube / Air Calibration | N/A | Fixed | madamsetti_k | 0 | 2 | Console Software | Console4 | YES | Closed | BB_WIP2_SubCom | Brilliance Big Bore |  | BB_WIP2 | December 13, 2018 8:31:21 PM GMT+08:00
CTUDT00600294 | Refresh | Closed | panchami_pradeep | November 6, 2018 9:23:35 PM GMT+08:00 | Refresh: PSC: Error was displayed while clicking on "Start Analysis" in Solid core Log analyzer page. | WIP11 | WIP11.1 |  | Subsystem | PSC - Log analyzer-Service Tools - Solid Core | N/A | Fixed | Sunil_Kumar_B | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP11.1SubCom | Ovation |  | WIP11.1 | November 28, 2018 8:06:36 PM GMT+08:00
CTUDT00600294 | Refresh | Closed | panchami_pradeep | November 6, 2018 9:23:35 PM GMT+08:00 | Refresh: PSC: Error was displayed while clicking on "Start Analysis" in Solid core Log analyzer page. | WIP11 | WIP11.1 |  | Subsystem | PSC - Log analyzer-Service Tools - Solid Core | N/A | Fixed | Sunil_Kumar_B | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP11.1SubCom | Brilliance 64 |  | WIP11.1 | November 28, 2018 8:06:36 PM GMT+08:00
CTUDT00600294 | Refresh | Closed | panchami_pradeep | November 6, 2018 9:23:35 PM GMT+08:00 | Refresh: PSC: Error was displayed while clicking on "Start Analysis" in Solid core Log analyzer page. | WIP11 | WIP11.1 |  | Subsystem | PSC - Log analyzer-Service Tools - Solid Core | N/A | Fixed | Sunil_Kumar_B | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP11.1SubCom | Ingenuity 64 |  | WIP11.1 | November 28, 2018 8:06:36 PM GMT+08:00
CTUDT00600294 | Refresh | Closed | panchami_pradeep | November 6, 2018 9:23:35 PM GMT+08:00 | Refresh: PSC: Error was displayed while clicking on "Start Analysis" in Solid core Log analyzer page. | WIP11 | WIP11.1 |  | Subsystem | PSC - Log analyzer-Service Tools - Solid Core | N/A | Fixed | Sunil_Kumar_B | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP11.1SubCom | IQon1 |  | WIP11.1 | November 28, 2018 8:06:36 PM GMT+08:00
CTUDT00600292 | Refresh | Closed | panchami_pradeep | November 6, 2018 8:45:19 PM GMT+08:00 | Refresh: PSC: Windows security warning error while launching tools like No Load test, Tilt profile, Couch vertical calibration, View Logs with markers etc. | WIP11 | WIP11.1 |  | Subsystem | PSC-OS messages | None | Fixed | Prabhakar_Priyank_Muppidi | 0 | 4 | Console Software | Console4 | NO | Closed | WIP11.1SubCom | Ovation |  | WIP11.1 | November 30, 2018 1:10:02 PM GMT+08:00
CTUDT00600292 | Refresh | Closed | panchami_pradeep | November 6, 2018 8:45:19 PM GMT+08:00 | Refresh: PSC: Windows security warning error while launching tools like No Load test, Tilt profile, Couch vertical calibration, View Logs with markers etc. | WIP11 | WIP11.1 |  | Subsystem | PSC-OS messages | None | Fixed | Prabhakar_Priyank_Muppidi | 0 | 4 | Console Software | Console4 | NO | Closed | WIP11.1SubCom | Brilliance 64 |  | WIP11.1 | November 30, 2018 1:10:02 PM GMT+08:00
CTUDT00600292 | Refresh | Closed | panchami_pradeep | November 6, 2018 8:45:19 PM GMT+08:00 | Refresh: PSC: Windows security warning error while launching tools like No Load test, Tilt profile, Couch vertical calibration, View Logs with markers etc. | WIP11 | WIP11.1 |  | Subsystem | PSC-OS messages | None | Fixed | Prabhakar_Priyank_Muppidi | 0 | 4 | Console Software | Console4 | NO | Closed | WIP11.1SubCom | Ingenuity 64 |  | WIP11.1 | November 30, 2018 1:10:02 PM GMT+08:00
CTUDT00600292 | Refresh | Closed | panchami_pradeep | November 6, 2018 8:45:19 PM GMT+08:00 | Refresh: PSC: Windows security warning error while launching tools like No Load test, Tilt profile, Couch vertical calibration, View Logs with markers etc. | WIP11 | WIP11.1 |  | Subsystem | PSC-OS messages | None | Fixed | Prabhakar_Priyank_Muppidi | 0 | 4 | Console Software | Console4 | NO | Closed | WIP11.1SubCom | IQon1 |  | WIP11.1 | November 30, 2018 1:10:02 PM GMT+08:00
CTUDT00600290 | Refresh | Closed | panchami_pradeep | November 6, 2018 8:34:56 PM GMT+08:00 | Refresh: PSC: Service Tools is not launching. | WIP11 | WIP11.1 | NA | Subsystem | PSC - Service Tools-Service Tools - Not launch | NA | Fixed | kapadia_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP11.1SubCom | Ovation |  | WIP11.1 | November 27, 2018 6:23:38 PM GMT+08:00
CTUDT00600290 | Refresh | Closed | panchami_pradeep | November 6, 2018 8:34:56 PM GMT+08:00 | Refresh: PSC: Service Tools is not launching. | WIP11 | WIP11.1 | NA | Subsystem | PSC - Service Tools-Service Tools - Not launch | NA | Fixed | kapadia_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP11.1SubCom | Brilliance 64 |  | WIP11.1 | November 27, 2018 6:23:38 PM GMT+08:00
CTUDT00600290 | Refresh | Closed | panchami_pradeep | November 6, 2018 8:34:56 PM GMT+08:00 | Refresh: PSC: Service Tools is not launching. | WIP11 | WIP11.1 | NA | Subsystem | PSC - Service Tools-Service Tools - Not launch | NA | Fixed | kapadia_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP11.1SubCom | Ingenuity 64 |  | WIP11.1 | November 27, 2018 6:23:38 PM GMT+08:00
CTUDT00600290 | Refresh | Closed | panchami_pradeep | November 6, 2018 8:34:56 PM GMT+08:00 | Refresh: PSC: Service Tools is not launching. | WIP11 | WIP11.1 | NA | Subsystem | PSC - Service Tools-Service Tools - Not launch | NA | Fixed | kapadia_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP11.1SubCom | IQon1 |  | WIP11.1 | November 27, 2018 6:23:38 PM GMT+08:00
CTUDT00600285 | Refresh | Closed | beck_d | November 5, 2018 9:56:35 PM GMT+08:00 | Refresh: Calibrations: Stability Test: The user is not notified if T-kit could not stabilize the temperature | WIP10 | WIP13 |  | Subsystem | Stability-Stability test | n/a | Fixed | kovalev_a | 0 | 2 | DMS Software | DMS | NO | Closed | WIP13SegReq | Ovation |  | WIP13 | January 16, 2019 8:52:36 PM GMT+08:00
CTUDT00600285 | Refresh | Closed | beck_d | November 5, 2018 9:56:35 PM GMT+08:00 | Refresh: Calibrations: Stability Test: The user is not notified if T-kit could not stabilize the temperature | WIP10 | WIP13 |  | Subsystem | Stability-Stability test | n/a | Fixed | kovalev_a | 0 | 2 | DMS Software | DMS | NO | Closed | WIP13SegReq | Brilliance 64 |  | WIP13 | January 16, 2019 8:52:36 PM GMT+08:00
CTUDT00600285 | Refresh | Closed | beck_d | November 5, 2018 9:56:35 PM GMT+08:00 | Refresh: Calibrations: Stability Test: The user is not notified if T-kit could not stabilize the temperature | WIP10 | WIP13 |  | Subsystem | Stability-Stability test | n/a | Fixed | kovalev_a | 0 | 2 | DMS Software | DMS | NO | Closed | WIP13SegReq | Ingenuity 64 |  | WIP13 | January 16, 2019 8:52:36 PM GMT+08:00
CTUDT00600285 | Refresh | Closed | beck_d | November 5, 2018 9:56:35 PM GMT+08:00 | Refresh: Calibrations: Stability Test: The user is not notified if T-kit could not stabilize the temperature | WIP10 | WIP13 |  | Subsystem | Stability-Stability test | n/a | Fixed | kovalev_a | 0 | 2 | DMS Software | DMS | NO | Closed | WIP13SegReq | IQon1 |  | WIP13 | January 16, 2019 8:52:36 PM GMT+08:00
CTUDT00600283 | Refresh | Closed | beck_d | November 5, 2018 9:53:33 PM GMT+08:00 | Refresh: Calibrations: Stability Test: The user is not notified if setting to nominal condition failed | WIP10 | N/A |  | Subsystem | Test design | n/a | Fixed | yosfan_s | 0 | 2 | Test Design | DMS | NO | Closed | WIP15SegReq | Ovation |  | N/A | March 27, 2019 3:52:51 PM GMT+08:00
CTUDT00600283 | Refresh | Closed | beck_d | November 5, 2018 9:53:33 PM GMT+08:00 | Refresh: Calibrations: Stability Test: The user is not notified if setting to nominal condition failed | WIP10 | N/A |  | Subsystem | Test design | n/a | Fixed | yosfan_s | 0 | 2 | Test Design | DMS | NO | Closed | WIP15SegReq | Ingenuity 64 |  | N/A | March 27, 2019 3:52:51 PM GMT+08:00
CTUDT00600283 | Refresh | Closed | beck_d | November 5, 2018 9:53:33 PM GMT+08:00 | Refresh: Calibrations: Stability Test: The user is not notified if setting to nominal condition failed | WIP10 | N/A |  | Subsystem | Test design | n/a | Fixed | yosfan_s | 0 | 2 | Test Design | DMS | NO | Closed | WIP15SegReq | IQon1 |  | N/A | March 27, 2019 3:52:51 PM GMT+08:00
CTUDT00600281 | Refresh | Closed | s_ra | November 5, 2018 4:23:32 PM GMT+08:00 | The result images ( plan ) are shifted and scan cycle reduces while planning a BRAIN AXIAL exam card WITH GANTRY TILT and TABLE direction INWARD | 4.1.7 | WIP11.1 | use the same result direction in case Tilt is used | Subsystem | Planning w Tilt-Plan - Brain Axial | None | Fixed | patra_r | 0 | 3 | Console Software | Console4 | YES | Closed | WIP12SegReq | Brilliance 64 |  | WIP11.1 | November 20, 2018 3:30:33 PM GMT+08:00
CTUDT00600281 | Refresh | Closed | s_ra | November 5, 2018 4:23:32 PM GMT+08:00 | The result images ( plan ) are shifted and scan cycle reduces while planning a BRAIN AXIAL exam card WITH GANTRY TILT and TABLE direction INWARD | 4.1.7 | WIP11.1 | use the same result direction in case Tilt is used | Subsystem | Planning w Tilt-Plan - Brain Axial | None | Fixed | patra_r | 0 | 3 | Console Software | Console4 | YES | Closed | WIP12SegReq | Ingenuity 64 |  | WIP11.1 | November 20, 2018 3:30:33 PM GMT+08:00
CTUDT00600277 | Refresh | Closed | beck_d | November 4, 2018 10:36:04 PM GMT+08:00 | Refresh: Embedded: APDM version is not read | WIP10 | WIP17 |  | Subsystem | System Configuration-Service Tools - System Config | N/A | Fixed | ivanf | 0 | 3 | DMS Software | DMS | NO | Closed | WIP13SegReq | Brilliance 64 |  | WIP17 | May 26, 2019 11:05:02 PM GMT+08:00
CTUDT00600277 | Refresh | Closed | beck_d | November 4, 2018 10:36:04 PM GMT+08:00 | Refresh: Embedded: APDM version is not read | WIP10 | WIP17 |  | Subsystem | System Configuration-Service Tools - System Config | N/A | Fixed | ivanf | 0 | 3 | DMS Software | DMS | NO | Closed | WIP13SegReq | Ingenuity 64 |  | WIP17 | May 26, 2019 11:05:02 PM GMT+08:00
CTUDT00600270 | Refresh | Closed | finestine_m | November 1, 2018 8:13:11 PM GMT+08:00 | Refresh: IQ Switch: Wrong value of BCorrectionMethodType parameter in cfg IQSwitchCommon_no_ABC | WIP9.1 | WIP17 |  | Subsystem | Calibrations - IQ switch | n/a | Fixed | Kostya_Baklan | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP12SegReq | Ovation |  | WIP17 | May 21, 2019 4:26:11 PM GMT+08:00
CTUDT00600270 | Refresh | Closed | finestine_m | November 1, 2018 8:13:11 PM GMT+08:00 | Refresh: IQ Switch: Wrong value of BCorrectionMethodType parameter in cfg IQSwitchCommon_no_ABC | WIP9.1 | WIP17 |  | Subsystem | Calibrations - IQ switch | n/a | Fixed | Kostya_Baklan | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP12SegReq | Ingenuity 64 |  | WIP17 | May 21, 2019 4:26:11 PM GMT+08:00
CTUDT00600268 | Refresh | Closed | Dana_Kamyshev | November 1, 2018 5:33:32 PM GMT+08:00 | Refresh - performing brain helical scan with 4 slices and save SBI not all images were reconstructed | WIP10 |  |  | Subsystem | Brain Helical scan reconstruction |  | Duplicate | richards_b | 0 |  |  | CIRS | NO | Closed |  | IQon1
CTUDT00600265 | Refresh | Closed | Dana_Kamyshev | November 1, 2018 5:02:43 PM GMT+08:00 | Refresh - sCTV When saving a new result the text isn't save | WIP10 |  |  | Subsystem | sCTV-sCTV - save operation | NA | Works As Specified / Not A Defect | redrouthu_r | 0 |  |  | iCAP | NO | Closed |  | Ovation
CTUDT00600265 | Refresh | Closed | Dana_Kamyshev | November 1, 2018 5:02:43 PM GMT+08:00 | Refresh - sCTV When saving a new result the text isn't save | WIP10 |  |  | Subsystem | sCTV-sCTV - save operation | NA | Works As Specified / Not A Defect | redrouthu_r | 0 |  |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00600265 | Refresh | Closed | Dana_Kamyshev | November 1, 2018 5:02:43 PM GMT+08:00 | Refresh - sCTV When saving a new result the text isn't save | WIP10 |  |  | Subsystem | sCTV-sCTV - save operation | NA | Works As Specified / Not A Defect | redrouthu_r | 0 |  |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00600265 | Refresh | Closed | Dana_Kamyshev | November 1, 2018 5:02:43 PM GMT+08:00 | Refresh - sCTV When saving a new result the text isn't save | WIP10 |  |  | Subsystem | sCTV-sCTV - save operation | NA | Works As Specified / Not A Defect | redrouthu_r | 0 |  |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00600265 | Refresh | Closed | Dana_Kamyshev | November 1, 2018 5:02:43 PM GMT+08:00 | Refresh - sCTV When saving a new result the text isn't save | WIP10 |  |  | Subsystem | sCTV-sCTV - save operation | NA | Works As Specified / Not A Defect | redrouthu_r | 0 |  |  | iCAP | NO | Closed |  | IQon1
CTUDT00600261 | Refresh | Closed | zhu_je | November 1, 2018 3:24:48 PM GMT+08:00 | There is no any message indicate user the system doesn't allow to set identical voltage values in sequential dual energy scans | WIP10 | WIP14 | N/A | Subsystem | Squential dual energy - messages | NA | Fixed | prasad_r | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SegReq | Ovation |  | WIP14 | February 21, 2019 5:40:30 PM GMT+08:00
CTUDT00600261 | Refresh | Closed | zhu_je | November 1, 2018 3:24:48 PM GMT+08:00 | There is no any message indicate user the system doesn't allow to set identical voltage values in sequential dual energy scans | WIP10 | WIP14 | N/A | Subsystem | Squential dual energy - messages | NA | Fixed | prasad_r | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SegReq | Brilliance 64 |  | WIP14 | February 21, 2019 5:40:30 PM GMT+08:00
CTUDT00600261 | Refresh | Closed | zhu_je | November 1, 2018 3:24:48 PM GMT+08:00 | There is no any message indicate user the system doesn't allow to set identical voltage values in sequential dual energy scans | WIP10 | WIP14 | N/A | Subsystem | Squential dual energy - messages | NA | Fixed | prasad_r | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SegReq | Ingenuity 64 |  | WIP14 | February 21, 2019 5:40:30 PM GMT+08:00
CTUDT00600250 | Refresh | Closed | ramesh_t | October 31, 2018 6:23:49 PM GMT+08:00 | Refresh: Truncated information in Acquisition viewer and word Phantom appears twice in show all | WIP10 | WIP12 |  | Subsystem | Acquision viewer | N/A | Fixed | prasad_r | 0 | 3 | Console Software | Console4 | NO | Closed | WIP12SubCom | Ovation |  | WIP12 | December 3, 2018 8:44:18 PM GMT+08:00
CTUDT00600250 | Refresh | Closed | ramesh_t | October 31, 2018 6:23:49 PM GMT+08:00 | Refresh: Truncated information in Acquisition viewer and word Phantom appears twice in show all | WIP10 | WIP12 |  | Subsystem | Acquision viewer | N/A | Fixed | prasad_r | 0 | 3 | Console Software | Console4 | NO | Closed | WIP12SubCom | Brilliance 64 |  | WIP12 | December 3, 2018 8:44:18 PM GMT+08:00
CTUDT00600250 | Refresh | Closed | ramesh_t | October 31, 2018 6:23:49 PM GMT+08:00 | Refresh: Truncated information in Acquisition viewer and word Phantom appears twice in show all | WIP10 | WIP12 |  | Subsystem | Acquision viewer | N/A | Fixed | prasad_r | 0 | 3 | Console Software | Console4 | NO | Closed | WIP12SubCom | Ingenuity 64 |  | WIP12 | December 3, 2018 8:44:18 PM GMT+08:00
CTUDT00600250 | Refresh | Closed | ramesh_t | October 31, 2018 6:23:49 PM GMT+08:00 | Refresh: Truncated information in Acquisition viewer and word Phantom appears twice in show all | WIP10 | WIP12 |  | Subsystem | Acquision viewer | N/A | Fixed | prasad_r | 0 | 3 | Console Software | Console4 | NO | Closed | WIP12SubCom | IQon1 |  | WIP12 | December 3, 2018 8:44:18 PM GMT+08:00
CTUDT00600248 | Refresh | Closed | ramesh_t | October 31, 2018 6:00:24 PM GMT+08:00 | Refresh: Harmonized Phantom will get enabled only after the restart of the system | WIP10 | WIP13 |  | Subsystem | Harmonized Phantom - enabled | n/a | Fixed | yosefk | 0 | 4 | DMS Calibrations | DMS | YES | Closed | WIP12SegReq | Ovation |  | WIP13 | January 20, 2019 3:16:37 PM GMT+08:00
CTUDT00600248 | Refresh | Closed | ramesh_t | October 31, 2018 6:00:24 PM GMT+08:00 | Refresh: Harmonized Phantom will get enabled only after the restart of the system | WIP10 | WIP13 |  | Subsystem | Harmonized Phantom - enabled | n/a | Fixed | yosefk | 0 | 4 | DMS Calibrations | DMS | YES | Closed | WIP12SegReq | Ingenuity 64 |  | WIP13 | January 20, 2019 3:16:37 PM GMT+08:00
CTUDT00600246 | Refresh | Closed | braunstein_d | October 31, 2018 2:07:14 PM GMT+08:00 | there is no collimation 2x0.625 in High res STD QA protocols iCT as per DHF192909 | WIP10 | WIP14 | NA | Subsystem | EC parameters - High Resolution | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | YES | Closed | WIP14SegReq | Ovation |  | WIP14 | February 15, 2019 12:43:12 PM GMT+08:00
CTUDT00600244 | Refresh | Closed | attali_s | October 29, 2018 9:21:31 PM GMT+08:00 | Refresh: Error Message about rotation time 0.27 sec in Tube Current Test | WIP10 | WIP12 |  | Subsystem | Calibrations - Tube Current | NAN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP12SegReq | Brilliance 64 |  | WIP12 | November 25, 2018 4:56:22 PM GMT+08:00
CTUDT00600244 | Refresh | Closed | attali_s | October 29, 2018 9:21:31 PM GMT+08:00 | Refresh: Error Message about rotation time 0.27 sec in Tube Current Test | WIP10 | WIP12 |  | Subsystem | Calibrations - Tube Current | NAN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP12SegReq | Brilliance Big Bore |  | WIP12 | November 25, 2018 4:56:22 PM GMT+08:00
CTUDT00600244 | Refresh | Closed | attali_s | October 29, 2018 9:21:31 PM GMT+08:00 | Refresh: Error Message about rotation time 0.27 sec in Tube Current Test | WIP10 | WIP12 |  | Subsystem | Calibrations - Tube Current | NAN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP12SegReq | Ingenuity 64 |  | WIP12 | November 25, 2018 4:56:22 PM GMT+08:00
CTUDT00600241 | Refresh | Closed | attali_s | October 29, 2018 8:46:34 PM GMT+08:00 | Refresh: Physics Analysis: System Vectors from Archive Physics Analysis use BD from Raw Data and not from D: | WIP 10 | WIP12 | Exit from Physics Analysis. Activate Physics Analysis by double click from c:\Pms\System\Physics.Analysis.exe and go to "System Vectors". Bad Detectors table will be used properly. | Subsystem | Physics Analysis | N/A | Fixed | yosefk | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP11.1SegReq | Ovation |  | WIP12 | November 22, 2018 9:09:54 PM GMT+08:00
CTUDT00600241 | Refresh | Closed | attali_s | October 29, 2018 8:46:34 PM GMT+08:00 | Refresh: Physics Analysis: System Vectors from Archive Physics Analysis use BD from Raw Data and not from D: | WIP 10 | WIP12 | Exit from Physics Analysis. Activate Physics Analysis by double click from c:\Pms\System\Physics.Analysis.exe and go to "System Vectors". Bad Detectors table will be used properly. | Subsystem | Physics Analysis | N/A | Fixed | yosefk | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP11.1SegReq | Brilliance 64 |  | WIP12 | November 22, 2018 9:09:54 PM GMT+08:00
CTUDT00600241 | Refresh | Closed | attali_s | October 29, 2018 8:46:34 PM GMT+08:00 | Refresh: Physics Analysis: System Vectors from Archive Physics Analysis use BD from Raw Data and not from D: | WIP 10 | WIP12 | Exit from Physics Analysis. Activate Physics Analysis by double click from c:\Pms\System\Physics.Analysis.exe and go to "System Vectors". Bad Detectors table will be used properly. | Subsystem | Physics Analysis | N/A | Fixed | yosefk | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP11.1SegReq | Brilliance Big Bore |  | WIP12 | November 22, 2018 9:09:54 PM GMT+08:00
CTUDT00600241 | Refresh | Closed | attali_s | October 29, 2018 8:46:34 PM GMT+08:00 | Refresh: Physics Analysis: System Vectors from Archive Physics Analysis use BD from Raw Data and not from D: | WIP 10 | WIP12 | Exit from Physics Analysis. Activate Physics Analysis by double click from c:\Pms\System\Physics.Analysis.exe and go to "System Vectors". Bad Detectors table will be used properly. | Subsystem | Physics Analysis | N/A | Fixed | yosefk | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP11.1SegReq | Ingenuity 64 |  | WIP12 | November 22, 2018 9:09:54 PM GMT+08:00
CTUDT00600241 | Refresh | Closed | attali_s | October 29, 2018 8:46:34 PM GMT+08:00 | Refresh: Physics Analysis: System Vectors from Archive Physics Analysis use BD from Raw Data and not from D: | WIP 10 | WIP12 | Exit from Physics Analysis. Activate Physics Analysis by double click from c:\Pms\System\Physics.Analysis.exe and go to "System Vectors". Bad Detectors table will be used properly. | Subsystem | Physics Analysis | N/A | Fixed | yosefk | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP11.1SegReq | IQon1 |  | WIP12 | November 22, 2018 9:09:54 PM GMT+08:00
CTUDT00600239 | Refresh | Closed | attali_s | October 29, 2018 7:45:11 PM GMT+08:00 | Refresh: IQ Check: Last Full Air Calibration date and time missing in HTML Report | WIP9.1 | WIP13 | N/A | Subsystem | Calibrations - Air Calibration | N/A | Fixed | elviral | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | Ovation |  | WIP13 | January 2, 2019 5:43:13 PM GMT+08:00
CTUDT00600239 | Refresh | Closed | attali_s | October 29, 2018 7:45:11 PM GMT+08:00 | Refresh: IQ Check: Last Full Air Calibration date and time missing in HTML Report | WIP9.1 | WIP13 | N/A | Subsystem | Calibrations - Air Calibration | N/A | Fixed | elviral | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | Brilliance Big Bore |  | WIP13 | January 2, 2019 5:43:13 PM GMT+08:00
CTUDT00600239 | Refresh | Closed | attali_s | October 29, 2018 7:45:11 PM GMT+08:00 | Refresh: IQ Check: Last Full Air Calibration date and time missing in HTML Report | WIP9.1 | WIP13 | N/A | Subsystem | Calibrations - Air Calibration | N/A | Fixed | elviral | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | Ingenuity 64 |  | WIP13 | January 2, 2019 5:43:13 PM GMT+08:00
CTUDT00600239 | Refresh | Closed | attali_s | October 29, 2018 7:45:11 PM GMT+08:00 | Refresh: IQ Check: Last Full Air Calibration date and time missing in HTML Report | WIP9.1 | WIP13 | N/A | Subsystem | Calibrations - Air Calibration | N/A | Fixed | elviral | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | IQon1 |  | WIP13 | January 2, 2019 5:43:13 PM GMT+08:00
CTUDT00600231 | Refresh | Closed | or_shmuel | October 29, 2018 3:58:42 PM GMT+08:00 | Refresh, Iqon - Missing option of change password | WIP10 |  |  | Subsystem | Dose Management - change password | Console.SRS.15920 | Works As Specified / Not A Defect | sanjeeva_a | 0 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00600229 | Refresh | Closed | or_shmuel | October 29, 2018 3:49:55 PM GMT+08:00 | Refresh, Iqon - Getting Critical error while trying to Open Workflow application Using Olevel User. | WIP10 |  |  | Subsystem | User Login - O-level |  | Duplicate | rooks | 0 |  |  | GCX | NO | Closed |  | IQon1
CTUDT00600216 | Refresh | Closed | golzban_v | October 28, 2018 3:56:35 PM GMT+08:00 | Refresh: Tube Alignment procedure fails with calculation error | WIP10 | WIP14 | Veritical couch position can be raised 20mm. | System | Calibrations - Tube Alignment | CT-NM.SRS.Service.2560 | Fixed | morton_t | 0 | 2 | Gantry Software | GCX | NO | Closed | WIP14SubCom | Ovation |  | WIP14 | March 30, 2019 2:53:23 AM GMT+08:00
CTUDT00600216 | Refresh | Closed | golzban_v | October 28, 2018 3:56:35 PM GMT+08:00 | Refresh: Tube Alignment procedure fails with calculation error | WIP10 | WIP14 | Veritical couch position can be raised 20mm. | System | Calibrations - Tube Alignment | CT-NM.SRS.Service.2560 | Fixed | morton_t | 0 | 2 | Gantry Software | GCX | NO | Closed | WIP14SubCom | IQon1 |  | WIP14 | March 30, 2019 2:53:23 AM GMT+08:00
CTUDT00600212 | Refresh | Closed | s_ra | October 26, 2018 4:36:35 PM GMT+08:00 | service data collection run in busy mode, causing system load | WIP510 | WIP16 BB_WIP6 |  | System | History Review; Verified1 | N/A | Fixed | sheik_c | 0 | 2 | Service Tools_Software | Console4 | YES | Closed | WIP14SegReq | Ovation | Escalation Case 0111410518 | WIP16 BB_WIP6 | April 25, 2019 4:15:51 PM GMT+08:00
CTUDT00600212 | Refresh | Closed | s_ra | October 26, 2018 4:36:35 PM GMT+08:00 | service data collection run in busy mode, causing system load | WIP510 | WIP16 BB_WIP6 |  | System | History Review; Verified1 | N/A | Fixed | sheik_c | 0 | 2 | Service Tools_Software | Console4 | YES | Closed | WIP14SegReq | Brilliance 64 | Escalation Case 0111410518 | WIP16 BB_WIP6 | April 25, 2019 4:15:51 PM GMT+08:00
CTUDT00600212 | Refresh | Closed | s_ra | October 26, 2018 4:36:35 PM GMT+08:00 | service data collection run in busy mode, causing system load | WIP510 | WIP16 BB_WIP6 |  | System | History Review; Verified1 | N/A | Fixed | sheik_c | 0 | 2 | Service Tools_Software | Console4 | YES | Closed | WIP14SegReq | Brilliance Big Bore | Escalation Case 0111410518 | WIP16 BB_WIP6 | April 25, 2019 4:15:51 PM GMT+08:00
CTUDT00600212 | Refresh | Closed | s_ra | October 26, 2018 4:36:35 PM GMT+08:00 | service data collection run in busy mode, causing system load | WIP510 | WIP16 BB_WIP6 |  | System | History Review; Verified1 | N/A | Fixed | sheik_c | 0 | 2 | Service Tools_Software | Console4 | YES | Closed | WIP14SegReq | Ingenuity 64 | Escalation Case 0111410518 | WIP16 BB_WIP6 | April 25, 2019 4:15:51 PM GMT+08:00
CTUDT00600212 | Refresh | Closed | s_ra | October 26, 2018 4:36:35 PM GMT+08:00 | service data collection run in busy mode, causing system load | WIP510 | WIP16 BB_WIP6 |  | System | History Review; Verified1 | N/A | Fixed | sheik_c | 0 | 2 | Service Tools_Software | Console4 | YES | Closed | WIP14SegReq | IQon1 | Escalation Case 0111410518 | WIP16 BB_WIP6 | April 25, 2019 4:15:51 PM GMT+08:00
CTUDT00600207 | Refresh | Closed | nanjunda_s | October 26, 2018 3:07:26 PM GMT+08:00 | Discard, overwrite and Rename options are not displaying in Import dialogue window After few exam cards are discarded | WIP10 | WIP13 |  | Subsystem | ECM - Import - Discard , Overview and Rename | Console-SRS.16027 | Fixed | prasad_r | 0 | 2 | Console Software | Console4 | NO | Closed | WIP13SubCom | Brilliance 64 |  | WIP13 | January 17, 2019 4:06:00 PM GMT+08:00
CTUDT00600207 | Refresh | Closed | nanjunda_s | October 26, 2018 3:07:26 PM GMT+08:00 | Discard, overwrite and Rename options are not displaying in Import dialogue window After few exam cards are discarded | WIP10 | WIP13 |  | Subsystem | ECM - Import - Discard , Overview and Rename | Console-SRS.16027 | Fixed | prasad_r | 0 | 2 | Console Software | Console4 | NO | Closed | WIP13SubCom | iCT4 |  | WIP13 | January 17, 2019 4:06:00 PM GMT+08:00
CTUDT00600207 | Refresh | Closed | nanjunda_s | October 26, 2018 3:07:26 PM GMT+08:00 | Discard, overwrite and Rename options are not displaying in Import dialogue window After few exam cards are discarded | WIP10 | WIP13 |  | Subsystem | ECM - Import - Discard , Overview and Rename | Console-SRS.16027 | Fixed | prasad_r | 0 | 2 | Console Software | Console4 | NO | Closed | WIP13SubCom | Ingenuity 64 |  | WIP13 | January 17, 2019 4:06:00 PM GMT+08:00
CTUDT00600205 | Refresh | Closed | singh_p | October 26, 2018 11:26:31 AM GMT+08:00 | Requirement STPS Refresh Ingenuity WIP 10 : Deviation in details provided by service log book when PSC tool closes | WIP10 | DHF303211 Rev06 |  | Subsystem | Documentation | STPS-SSRS-400 | Fixed | hatti_c | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | WIP13SegReq | Ovation |  | DHF303211 Rev06 | January 30, 2019 6:17:33 PM GMT+08:00
CTUDT00600205 | Refresh | Closed | singh_p | October 26, 2018 11:26:31 AM GMT+08:00 | Requirement STPS Refresh Ingenuity WIP 10 : Deviation in details provided by service log book when PSC tool closes | WIP10 | DHF303211 Rev06 |  | Subsystem | Documentation | STPS-SSRS-400 | Fixed | hatti_c | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | WIP13SegReq | Brilliance 64 |  | DHF303211 Rev06 | January 30, 2019 6:17:33 PM GMT+08:00
CTUDT00600205 | Refresh | Closed | singh_p | October 26, 2018 11:26:31 AM GMT+08:00 | Requirement STPS Refresh Ingenuity WIP 10 : Deviation in details provided by service log book when PSC tool closes | WIP10 | DHF303211 Rev06 |  | Subsystem | Documentation | STPS-SSRS-400 | Fixed | hatti_c | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | WIP13SegReq | Brilliance Big Bore |  | DHF303211 Rev06 | January 30, 2019 6:17:33 PM GMT+08:00
CTUDT00600205 | Refresh | Closed | singh_p | October 26, 2018 11:26:31 AM GMT+08:00 | Requirement STPS Refresh Ingenuity WIP 10 : Deviation in details provided by service log book when PSC tool closes | WIP10 | DHF303211 Rev06 |  | Subsystem | Documentation | STPS-SSRS-400 | Fixed | hatti_c | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | WIP13SegReq | Ingenuity 64 |  | DHF303211 Rev06 | January 30, 2019 6:17:33 PM GMT+08:00
CTUDT00600205 | Refresh | Closed | singh_p | October 26, 2018 11:26:31 AM GMT+08:00 | Requirement STPS Refresh Ingenuity WIP 10 : Deviation in details provided by service log book when PSC tool closes | WIP10 | DHF303211 Rev06 |  | Subsystem | Documentation | STPS-SSRS-400 | Fixed | hatti_c | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | WIP13SegReq | IQon1 |  | DHF303211 Rev06 | January 30, 2019 6:17:33 PM GMT+08:00
CTUDT00600203 | Refresh | Closed | rakic_k | October 26, 2018 3:05:43 AM GMT+08:00 | Unable to Scan SeriesProcess Blocked | WIP10 | WIP11 |  | Subsystem | Fixed; Verified1 | N/A | Fixed | rakic_k | 0 | 2 | CIRS_SW_Installation | CIRS | NO | Closed | WIP11SegReq | Ovation |  | WIP11 | November 16, 2018 4:08:30 AM GMT+08:00
CTUDT00600203 | Refresh | Closed | rakic_k | October 26, 2018 3:05:43 AM GMT+08:00 | Unable to Scan SeriesProcess Blocked | WIP10 | WIP11 |  | Subsystem | Fixed; Verified1 | N/A | Fixed | rakic_k | 0 | 2 | CIRS_SW_Installation | CIRS | NO | Closed | WIP11SegReq | Brilliance 64 |  | WIP11 | November 16, 2018 4:08:30 AM GMT+08:00
CTUDT00600203 | Refresh | Closed | rakic_k | October 26, 2018 3:05:43 AM GMT+08:00 | Unable to Scan SeriesProcess Blocked | WIP10 | WIP11 |  | Subsystem | Fixed; Verified1 | N/A | Fixed | rakic_k | 0 | 2 | CIRS_SW_Installation | CIRS | NO | Closed | WIP11SegReq | Ingenuity 64 |  | WIP11 | November 16, 2018 4:08:30 AM GMT+08:00
CTUDT00600203 | Refresh | Closed | rakic_k | October 26, 2018 3:05:43 AM GMT+08:00 | Unable to Scan SeriesProcess Blocked | WIP10 | WIP11 |  | Subsystem | Fixed; Verified1 | N/A | Fixed | rakic_k | 0 | 2 | CIRS_SW_Installation | CIRS | NO | Closed | WIP11SegReq | IQon1 |  | WIP11 | November 16, 2018 4:08:30 AM GMT+08:00
CTUDT00600198 | Refresh | Closed | finestine_m | October 25, 2018 6:06:18 PM GMT+08:00 | Refresh: Bad Detectors Quick Check: Missing Message regarding Found BD | WIP9.1 | WIP12 | Not available | Subsystem | Calibrations - Bad Detector | N/A | Fixed | yosefk | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP11.1SegReq | Ovation |  | WIP12 | December 12, 2018 10:14:50 PM GMT+08:00
CTUDT00600198 | Refresh | Closed | finestine_m | October 25, 2018 6:06:18 PM GMT+08:00 | Refresh: Bad Detectors Quick Check: Missing Message regarding Found BD | WIP9.1 | WIP12 | Not available | Subsystem | Calibrations - Bad Detector | N/A | Fixed | yosefk | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP11.1SegReq | Brilliance 64 |  | WIP12 | December 12, 2018 10:14:50 PM GMT+08:00
CTUDT00600198 | Refresh | Closed | finestine_m | October 25, 2018 6:06:18 PM GMT+08:00 | Refresh: Bad Detectors Quick Check: Missing Message regarding Found BD | WIP9.1 | WIP12 | Not available | Subsystem | Calibrations - Bad Detector | N/A | Fixed | yosefk | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP11.1SegReq | Brilliance Big Bore |  | WIP12 | December 12, 2018 10:14:50 PM GMT+08:00
CTUDT00600198 | Refresh | Closed | finestine_m | October 25, 2018 6:06:18 PM GMT+08:00 | Refresh: Bad Detectors Quick Check: Missing Message regarding Found BD | WIP9.1 | WIP12 | Not available | Subsystem | Calibrations - Bad Detector | N/A | Fixed | yosefk | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP11.1SegReq | Ingenuity 64 |  | WIP12 | December 12, 2018 10:14:50 PM GMT+08:00
CTUDT00600198 | Refresh | Closed | finestine_m | October 25, 2018 6:06:18 PM GMT+08:00 | Refresh: Bad Detectors Quick Check: Missing Message regarding Found BD | WIP9.1 | WIP12 | Not available | Subsystem | Calibrations - Bad Detector | N/A | Fixed | yosefk | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP11.1SegReq | IQon1 |  | WIP12 | December 12, 2018 10:14:50 PM GMT+08:00
CTUDT00600197 | Refresh | Closed | patra_r | October 25, 2018 5:20:40 PM GMT+08:00 | Scan box moves 15 mm toward the gantry opening immediately before scan starts | V4.1.7 | WIP11.1 | 1) In case tilted scan, user should not have overlap results in the Plan. Overlap results can be requested offline via “Edit Result”. [Recommended] 2) User should have only one result in the plan so that copy geometry is not initiated during plan. Other results can be requested offline via “Edit result”. [Recommended] 3)User should plan all the results by selecting them and then tilt the scan. [Recommended] | Subsystem | Plan scene - scan box moves | None | Fixed | patra_r | 0 | 3 | Console Software | Console4 | YES | Closed | WIP12SegReq | Brilliance 64 | PR7744385 | WIP11.1 | November 20, 2018 3:29:34 PM GMT+08:00
CTUDT00600197 | Refresh | Closed | patra_r | October 25, 2018 5:20:40 PM GMT+08:00 | Scan box moves 15 mm toward the gantry opening immediately before scan starts | V4.1.7 | WIP11.1 | 1) In case tilted scan, user should not have overlap results in the Plan. Overlap results can be requested offline via “Edit Result”. [Recommended] 2) User should have only one result in the plan so that copy geometry is not initiated during plan. Other results can be requested offline via “Edit result”. [Recommended] 3)User should plan all the results by selecting them and then tilt the scan. [Recommended] | Subsystem | Plan scene - scan box moves | None | Fixed | patra_r | 0 | 3 | Console Software | Console4 | YES | Closed | WIP12SegReq | Ingenuity 64 | PR7744385 | WIP11.1 | November 20, 2018 3:29:34 PM GMT+08:00
CTUDT00600192 | Refresh | Closed | panchami_pradeep | October 25, 2018 11:39:02 AM GMT+08:00 | Refresh: PSC: BugRep tool is displaying a command prompt error while cancelling Bug report collection. | WIP10 |  |  | Subsystem | Bugrep - Cancelling Bugrep | N/A | No Plan to Fix | hatti_c | 0 | 4 |  | Service Tools | NO | Closed |  | Ovation
CTUDT00600192 | Refresh | Closed | panchami_pradeep | October 25, 2018 11:39:02 AM GMT+08:00 | Refresh: PSC: BugRep tool is displaying a command prompt error while cancelling Bug report collection. | WIP10 |  |  | Subsystem | Bugrep - Cancelling Bugrep | N/A | No Plan to Fix | hatti_c | 0 | 4 |  | Service Tools | NO | Closed |  | Brilliance 64
CTUDT00600192 | Refresh | Closed | panchami_pradeep | October 25, 2018 11:39:02 AM GMT+08:00 | Refresh: PSC: BugRep tool is displaying a command prompt error while cancelling Bug report collection. | WIP10 |  |  | Subsystem | Bugrep - Cancelling Bugrep | N/A | No Plan to Fix | hatti_c | 0 | 4 |  | Service Tools | NO | Closed |  | Brilliance Big Bore
CTUDT00600192 | Refresh | Closed | panchami_pradeep | October 25, 2018 11:39:02 AM GMT+08:00 | Refresh: PSC: BugRep tool is displaying a command prompt error while cancelling Bug report collection. | WIP10 |  |  | Subsystem | Bugrep - Cancelling Bugrep | N/A | No Plan to Fix | hatti_c | 0 | 4 |  | Service Tools | NO | Closed |  | Ingenuity 64
CTUDT00600192 | Refresh | Closed | panchami_pradeep | October 25, 2018 11:39:02 AM GMT+08:00 | Refresh: PSC: BugRep tool is displaying a command prompt error while cancelling Bug report collection. | WIP10 |  |  | Subsystem | Bugrep - Cancelling Bugrep | N/A | No Plan to Fix | hatti_c | 0 | 4 |  | Service Tools | NO | Closed |  | IQon1
CTUDT00600190 | Refresh | Closed | jotawar_d | October 24, 2018 4:59:01 PM GMT+08:00 | iBrain filters are not available for "Head" scan types. | BB_WIP1 |  | NA | Subsystem |  | NA | Does Not Apply | kumar_v | 0 | 2 | Console Software | Console4 | NO | Closed |  | Ingenuity BigBore
CTUDT00600188 | Refresh | Closed | sowmya_k | October 24, 2018 2:25:36 PM GMT+08:00 | Refresh: No images are reconstructed when IMR Recon Mode is enabled. | BB_WIP1 | BB_WIP2 |  | Subsystem |  | Console-SRS.17079 | Fixed | kumar_v | 0 | 2 | Console Software | Console4 | YES | Closed | BB_WIP2_SubCom | Brilliance Big Bore |  | BB_WIP2 | October 31, 2018 1:48:01 PM GMT+08:00
CTUDT00600184 | Refresh | Closed | bagdady_b | October 23, 2018 9:08:42 PM GMT+08:00 | Refresh - When changing collimation from auto to the specific collimation mentioned in auto more parameters are manipulated | WIP10 |  | Change the desirable parameters before clicking on Save. Or after clicking Save, open the Show All parameters and change the values. | Subsystem | EC parameters - Auto colimation | N/A | No Plan to Fix | dunitza_m | 0 | 4 |  | Support-System | NO | Closed |  | Ovation
CTUDT00600184 | Refresh | Closed | bagdady_b | October 23, 2018 9:08:42 PM GMT+08:00 | Refresh - When changing collimation from auto to the specific collimation mentioned in auto more parameters are manipulated | WIP10 |  | Change the desirable parameters before clicking on Save. Or after clicking Save, open the Show All parameters and change the values. | Subsystem | EC parameters - Auto colimation | N/A | No Plan to Fix | dunitza_m | 0 | 4 |  | Support-System | NO | Closed |  | Brilliance 64
CTUDT00600184 | Refresh | Closed | bagdady_b | October 23, 2018 9:08:42 PM GMT+08:00 | Refresh - When changing collimation from auto to the specific collimation mentioned in auto more parameters are manipulated | WIP10 |  | Change the desirable parameters before clicking on Save. Or after clicking Save, open the Show All parameters and change the values. | Subsystem | EC parameters - Auto colimation | N/A | No Plan to Fix | dunitza_m | 0 | 4 |  | Support-System | NO | Closed |  | Ingenuity 64
CTUDT00600184 | Refresh | Closed | bagdady_b | October 23, 2018 9:08:42 PM GMT+08:00 | Refresh - When changing collimation from auto to the specific collimation mentioned in auto more parameters are manipulated | WIP10 |  | Change the desirable parameters before clicking on Save. Or after clicking Save, open the Show All parameters and change the values. | Subsystem | EC parameters - Auto colimation | N/A | No Plan to Fix | dunitza_m | 0 | 4 |  | Support-System | NO | Closed |  | IQon1
CTUDT00600181 | Refresh | Closed | nanjunda_s | October 23, 2018 6:30:03 PM GMT+08:00 | Workflow is not launched with CT user login. After accepted the license agreement reverting to Login Screen. | WIP10 | WIP11 | User can login to Philips_service user and start Pms remoting and pms logger servies manually. Other option is, If user restart the system, these services will be start by OS and work flow will be launched. | Subsystem | User Login - Workflow launch | NA | Fixed | kapadia_v | 0 | 4 | Console Software | Console4 | NO | Closed | WIP11.1SubCom | Ovation |  | WIP11 | November 27, 2018 8:25:29 PM GMT+08:00
CTUDT00600181 | Refresh | Closed | nanjunda_s | October 23, 2018 6:30:03 PM GMT+08:00 | Workflow is not launched with CT user login. After accepted the license agreement reverting to Login Screen. | WIP10 | WIP11 | User can login to Philips_service user and start Pms remoting and pms logger servies manually. Other option is, If user restart the system, these services will be start by OS and work flow will be launched. | Subsystem | User Login - Workflow launch | NA | Fixed | kapadia_v | 0 | 4 | Console Software | Console4 | NO | Closed | WIP11.1SubCom | Brilliance 64 |  | WIP11 | November 27, 2018 8:25:29 PM GMT+08:00
CTUDT00600181 | Refresh | Closed | nanjunda_s | October 23, 2018 6:30:03 PM GMT+08:00 | Workflow is not launched with CT user login. After accepted the license agreement reverting to Login Screen. | WIP10 | WIP11 | User can login to Philips_service user and start Pms remoting and pms logger servies manually. Other option is, If user restart the system, these services will be start by OS and work flow will be launched. | Subsystem | User Login - Workflow launch | NA | Fixed | kapadia_v | 0 | 4 | Console Software | Console4 | NO | Closed | WIP11.1SubCom | Ingenuity 64 |  | WIP11 | November 27, 2018 8:25:29 PM GMT+08:00
CTUDT00600181 | Refresh | Closed | nanjunda_s | October 23, 2018 6:30:03 PM GMT+08:00 | Workflow is not launched with CT user login. After accepted the license agreement reverting to Login Screen. | WIP10 | WIP11 | User can login to Philips_service user and start Pms remoting and pms logger servies manually. Other option is, If user restart the system, these services will be start by OS and work flow will be launched. | Subsystem | User Login - Workflow launch | NA | Fixed | kapadia_v | 0 | 4 | Console Software | Console4 | NO | Closed | WIP11.1SubCom | IQon1 |  | WIP11 | November 27, 2018 8:25:29 PM GMT+08:00
CTUDT00600178 | Refresh | Closed | datta_i | October 23, 2018 4:51:45 PM GMT+08:00 | [Refresh WIP 8.1] [Exam Card In Plan Viewer]: Phase value and filming value(result step) is not displaying in the Orientation tab | 4.8 | BB_WIP2 |  | Subsystem | 4D CT Scan - Phase value is not displayed | NA | Fixed | datta_i | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP2_SegReq | Brilliance Big Bore |  | BB_WIP2 | November 23, 2018 1:56:45 PM GMT+08:00
CTUDT00600172 | Refresh | Closed | shaham_t | October 22, 2018 9:07:39 PM GMT+08:00 | Refresh: Need to update DHF302477 revision 02, DHF302475 revision 02 and DHF302490 revision 02 | WIP8 | N/A |  | Subsystem | Documentation | N/A | Fixed | bBeilin_b | 0 | 2 | Requirements Documentation | DMS | NO | Closed | WIP11SegReq | IQon1 |  | N/A | December 30, 2018 3:51:27 PM GMT+08:00
CTUDT00600167 | Refresh | Closed | roshan_r | October 22, 2018 6:39:56 PM GMT+08:00 | Refresh : ECG viewer - The planned irradiation period/phase on the ECG wave/graph in online ECG viewer is not marked in blue color when step and shoot EC was selected. | WIP10 | Test design update |  | Subsystem | ECG Viewer - Offline ECG waveform | NA | Fixed | nanjunda_s | 0 | 4 | Test Design | Support-Simulator | NO | Closed | WIP13SegReq | Ingenuity 64 |  | Test design update | May 24, 2019 5:22:29 PM GMT+08:00
CTUDT00600157 | Refresh | Closed | wei_f | October 19, 2018 10:18:23 AM GMT+08:00 | The Subtraction of the ROI's appears on the image show ***,not the right value. | WIP10 |  |  | System | ROI - Calculations | NA | No Plan to Fix | wei_f | 0 | 2 | Test Design | Support-System | NO | Closed | WIP13SegReq | Ingenuity 64 |  |  | June 20, 2019 4:26:54 PM GMT+08:00
CTUDT00600153 | Refresh | Closed | basavaraju_p | October 18, 2018 5:21:14 PM GMT+08:00 | Refresh-Ingenuity: Keyboard Shortcuts for windowing presets is not working as expected | WIP10 | WIP15 | the user may change the windowing in multiple different ways, such as type in, use mouse controls or in the Plan, CT Presets dropdown in the tool bar can be used to select different presets. | System | Window & Center - keyboard shortcut; Verified1 | Console-SRS.2359 | Fixed | redrouthu_r | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP15SubCom | Ovation |  | WIP15 | April 4, 2019 4:15:10 PM GMT+08:00
CTUDT00600153 | Refresh | Closed | basavaraju_p | October 18, 2018 5:21:14 PM GMT+08:00 | Refresh-Ingenuity: Keyboard Shortcuts for windowing presets is not working as expected | WIP10 | WIP15 | the user may change the windowing in multiple different ways, such as type in, use mouse controls or in the Plan, CT Presets dropdown in the tool bar can be used to select different presets. | System | Window & Center - keyboard shortcut; Verified1 | Console-SRS.2359 | Fixed | redrouthu_r | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP15SubCom | Brilliance 64 |  | WIP15 | April 4, 2019 4:15:10 PM GMT+08:00
CTUDT00600153 | Refresh | Closed | basavaraju_p | October 18, 2018 5:21:14 PM GMT+08:00 | Refresh-Ingenuity: Keyboard Shortcuts for windowing presets is not working as expected | WIP10 | WIP15 | the user may change the windowing in multiple different ways, such as type in, use mouse controls or in the Plan, CT Presets dropdown in the tool bar can be used to select different presets. | System | Window & Center - keyboard shortcut; Verified1 | Console-SRS.2359 | Fixed | redrouthu_r | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP15SubCom | Brilliance Big Bore |  | WIP15 | April 4, 2019 4:15:10 PM GMT+08:00
CTUDT00600153 | Refresh | Closed | basavaraju_p | October 18, 2018 5:21:14 PM GMT+08:00 | Refresh-Ingenuity: Keyboard Shortcuts for windowing presets is not working as expected | WIP10 | WIP15 | the user may change the windowing in multiple different ways, such as type in, use mouse controls or in the Plan, CT Presets dropdown in the tool bar can be used to select different presets. | System | Window & Center - keyboard shortcut; Verified1 | Console-SRS.2359 | Fixed | redrouthu_r | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP15SubCom | Ingenuity 64 |  | WIP15 | April 4, 2019 4:15:10 PM GMT+08:00
CTUDT00600153 | Refresh | Closed | basavaraju_p | October 18, 2018 5:21:14 PM GMT+08:00 | Refresh-Ingenuity: Keyboard Shortcuts for windowing presets is not working as expected | WIP10 | WIP15 | the user may change the windowing in multiple different ways, such as type in, use mouse controls or in the Plan, CT Presets dropdown in the tool bar can be used to select different presets. | System | Window & Center - keyboard shortcut; Verified1 | Console-SRS.2359 | Fixed | redrouthu_r | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP15SubCom | IQon1 |  | WIP15 | April 4, 2019 4:15:10 PM GMT+08:00
CTUDT00600151 | Refresh | Closed | mistry_h | October 18, 2018 4:34:54 PM GMT+08:00 | V4.1.7: Dose Log Viewer is blank. | v4.1.7.19026 | BB_WIP2 |  | Subsystem |  | N/A | Fixed | atin_s | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP2_SegReq | Brilliance Big Bore | PR7741791 | BB_WIP2 | November 12, 2018 8:28:32 PM GMT+08:00
CTUDT00600142 | Refresh | Closed | bommiReddy_r | October 17, 2018 5:35:10 PM GMT+08:00 | Refresh: Exam Prep-Worklist-“Study Date and Time” column and Modality column filters are not working | WIP10 | Host SSRS DHF303211 Rev 11 |  | Subsystem | Worklist - column filters | Console-SRS.3378 | Fixed | bandil_j | 0 | 3 | Host SSRS | Console4 | NO | Closed | WIP17SegReq | Ingenuity 64 |  | Host SSRS DHF303211 Rev 11 | May 20, 2019 5:29:12 PM GMT+08:00
CTUDT00600121 | Refresh | Closed | attali_s | October 15, 2018 8:19:34 PM GMT+08:00 | Refresh: Calibrations: Disabling 0.42 RT in key is not reflected in Quality Assurance | WIP9.1 | N/A | Restarting the gantry will resolve the issue. Moreover, this is test design issue. | Subsystem | Calibrations - Air Calibration | N/A | Fixed | gehtin_e | 0 | 3 | Test Design | DMS | NO | Closed | WIP13SegReq | Ingenuity 64 |  | N/A | January 16, 2019 3:36:28 PM GMT+08:00
CTUDT00600115 | Refresh | Closed | panchami_pradeep | October 15, 2018 2:38:04 PM GMT+08:00 | In ct user, Bug rep pop up is displayed while clicking on "film display" option. | WIP10 | WIP12 |  | Subsystem | Film - fim display option | NA | Fixed | redrouthu_r | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP12SegReq | Ingenuity 64 |  | WIP12 | February 7, 2019 4:50:51 PM GMT+08:00
CTUDT00600113 | Refresh | Closed | attali_s | October 15, 2018 1:20:16 PM GMT+08:00 | Refresh: Stability: No message in the Temp Notification Log when the T-Kit cannot stabilized | WIP9.1 |  |  | Subsystem | Stability test | n/a | Duplicate | franco_a | 0 |  |  | DMS | NO | Closed |  | Ingenuity 64
CTUDT00600104 | Refresh | Closed | panchami_pradeep | October 12, 2018 3:15:04 PM GMT+08:00 | Refresh: In IT user, under password policy tab "Enable password complexity" is not checked after clicking on "Restore Factory Settings" | WIP10 | WIP11 | User can logout and login after restoring factory settings, the UI gets updated. | Subsystem |  | NA | Fixed | ravindra_a | 0 | 4 | Console Software | Console4 | NO | Closed | WIP11SegReq | Ovation |  | WIP11 | October 29, 2018 5:31:07 PM GMT+08:00
CTUDT00600104 | Refresh | Closed | panchami_pradeep | October 12, 2018 3:15:04 PM GMT+08:00 | Refresh: In IT user, under password policy tab "Enable password complexity" is not checked after clicking on "Restore Factory Settings" | WIP10 | WIP11 | User can logout and login after restoring factory settings, the UI gets updated. | Subsystem |  | NA | Fixed | ravindra_a | 0 | 4 | Console Software | Console4 | NO | Closed | WIP11SegReq | Brilliance 64 |  | WIP11 | October 29, 2018 5:31:07 PM GMT+08:00
CTUDT00600104 | Refresh | Closed | panchami_pradeep | October 12, 2018 3:15:04 PM GMT+08:00 | Refresh: In IT user, under password policy tab "Enable password complexity" is not checked after clicking on "Restore Factory Settings" | WIP10 | WIP11 | User can logout and login after restoring factory settings, the UI gets updated. | Subsystem |  | NA | Fixed | ravindra_a | 0 | 4 | Console Software | Console4 | NO | Closed | WIP11SegReq | Brilliance Big Bore |  | WIP11 | October 29, 2018 5:31:07 PM GMT+08:00
CTUDT00600104 | Refresh | Closed | panchami_pradeep | October 12, 2018 3:15:04 PM GMT+08:00 | Refresh: In IT user, under password policy tab "Enable password complexity" is not checked after clicking on "Restore Factory Settings" | WIP10 | WIP11 | User can logout and login after restoring factory settings, the UI gets updated. | Subsystem |  | NA | Fixed | ravindra_a | 0 | 4 | Console Software | Console4 | NO | Closed | WIP11SegReq | Ingenuity 64 |  | WIP11 | October 29, 2018 5:31:07 PM GMT+08:00
CTUDT00600104 | Refresh | Closed | panchami_pradeep | October 12, 2018 3:15:04 PM GMT+08:00 | Refresh: In IT user, under password policy tab "Enable password complexity" is not checked after clicking on "Restore Factory Settings" | WIP10 | WIP11 | User can logout and login after restoring factory settings, the UI gets updated. | Subsystem |  | NA | Fixed | ravindra_a | 0 | 4 | Console Software | Console4 | NO | Closed | WIP11SegReq | IQon1 |  | WIP11 | October 29, 2018 5:31:07 PM GMT+08:00
CTUDT00600102 | Refresh | Closed | wei_f | October 12, 2018 2:06:02 PM GMT+08:00 | Calculate the area ROI's sum of the following parameters: Area (Ar), Average (Av), SD, the results are all error | WIP9.1 |  | NA | Subsystem | ROI - Calculations | NA | No Plan to Fix | wei_f | 0 | 2 | Test Design | Support-System | NO | Closed | WIP11.1SegReq | Ingenuity 64 |  |  | June 20, 2019 4:42:21 PM GMT+08:00
CTUDT00600098 | Refresh | Closed | zhang_t | October 12, 2018 1:17:09 PM GMT+08:00 | Refresh WIP10: Gantry acquisition error when perform Helical Scan with long scan length | WIP10 |  |  | Subsystem | Gantry SW/HW | N/A | Works As Specified / Not A Defect | zhang_l | 1 | 4 |  | Support-System | YES | Closed | SubNC | Brilliance 64
CTUDT00600098 | Refresh | Closed | zhang_t | October 12, 2018 1:17:09 PM GMT+08:00 | Refresh WIP10: Gantry acquisition error when perform Helical Scan with long scan length | WIP10 |  |  | Subsystem | Gantry SW/HW | N/A | Works As Specified / Not A Defect | zhang_l | 1 | 4 |  | Support-System | YES | Closed | SubNC | Ingenuity 64
CTUDT00600097 | Refresh | Closed | wei_f | October 12, 2018 1:02:53 PM GMT+08:00 | CT user enters the wrong password to the maximum number of times, the account is not disabled in user management. | WIP9.1 | WIP14 | NA | Subsystem | Password Management - | NA | Fixed | wei_f | 0 | 2 | Design Documentation | Support-System | NO | Closed | WIP11SegReq | Brilliance 64 |  | WIP14 | June 20, 2019 4:41:16 PM GMT+08:00
CTUDT00600097 | Refresh | Closed | wei_f | October 12, 2018 1:02:53 PM GMT+08:00 | CT user enters the wrong password to the maximum number of times, the account is not disabled in user management. | WIP9.1 | WIP14 | NA | Subsystem | Password Management - | NA | Fixed | wei_f | 0 | 2 | Design Documentation | Support-System | NO | Closed | WIP11SegReq | Ingenuity 64 |  | WIP14 | June 20, 2019 4:41:16 PM GMT+08:00
CTUDT00600095 | Refresh | Closed | wei_f | October 12, 2018 9:58:08 AM GMT+08:00 | After one login attempt failed, the system did not display the number of login attempts left | WIP9.1 | WIP18 |  | Subsystem | User Login - login attempts | CT-NM.SRS.SecPri.308 | Fixed | lin_s | 0 | 2 | Design Documentation | Support-System | NO | Closed | WIP12SegReq | Ovation |  | WIP18 | July 25, 2019 8:50:30 PM GMT+08:00
CTUDT00600095 | Refresh | Closed | wei_f | October 12, 2018 9:58:08 AM GMT+08:00 | After one login attempt failed, the system did not display the number of login attempts left | WIP9.1 | WIP18 |  | Subsystem | User Login - login attempts | CT-NM.SRS.SecPri.308 | Fixed | lin_s | 0 | 2 | Design Documentation | Support-System | NO | Closed | WIP12SegReq | Brilliance 64 |  | WIP18 | July 25, 2019 8:50:30 PM GMT+08:00
CTUDT00600095 | Refresh | Closed | wei_f | October 12, 2018 9:58:08 AM GMT+08:00 | After one login attempt failed, the system did not display the number of login attempts left | WIP9.1 | WIP18 |  | Subsystem | User Login - login attempts | CT-NM.SRS.SecPri.308 | Fixed | lin_s | 0 | 2 | Design Documentation | Support-System | NO | Closed | WIP12SegReq | Brilliance Big Bore |  | WIP18 | July 25, 2019 8:50:30 PM GMT+08:00
CTUDT00600095 | Refresh | Closed | wei_f | October 12, 2018 9:58:08 AM GMT+08:00 | After one login attempt failed, the system did not display the number of login attempts left | WIP9.1 | WIP18 |  | Subsystem | User Login - login attempts | CT-NM.SRS.SecPri.308 | Fixed | lin_s | 0 | 2 | Design Documentation | Support-System | NO | Closed | WIP12SegReq | Ingenuity 64 |  | WIP18 | July 25, 2019 8:50:30 PM GMT+08:00
CTUDT00600095 | Refresh | Closed | wei_f | October 12, 2018 9:58:08 AM GMT+08:00 | After one login attempt failed, the system did not display the number of login attempts left | WIP9.1 | WIP18 |  | Subsystem | User Login - login attempts | CT-NM.SRS.SecPri.308 | Fixed | lin_s | 0 | 2 | Design Documentation | Support-System | NO | Closed | WIP12SegReq | IQon1 |  | WIP18 | July 25, 2019 8:50:30 PM GMT+08:00
CTUDT00600088 | Refresh | Closed | panchami_pradeep | October 11, 2018 4:40:44 PM GMT+08:00 | Refresh: Offline ECG waveform is not scrolling along with the reconstructed image in the offline ECG viewer. | WIP10 |  |  | Subsystem | ECG Viewer - Simulator | Console-SRS.11236 | Forwarded to Different Database | babu_s | 0 |  |  | Support-Simulator | NO | Closed |  | Ingenuity 64
CTUDT00600088 | Refresh | Closed | panchami_pradeep | October 11, 2018 4:40:44 PM GMT+08:00 | Refresh: Offline ECG waveform is not scrolling along with the reconstructed image in the offline ECG viewer. | WIP10 |  |  | Subsystem | ECG Viewer - Simulator | Console-SRS.11236 | Forwarded to Different Database | babu_s | 0 |  |  | Support-Simulator | NO | Closed |  | IQon1
CTUDT00600085 | Refresh | Closed | lin_s | October 11, 2018 3:08:55 PM GMT+08:00 | Refresh, IFU: No clear description of feature: Pause button is disabled during clinical scan initialization. | WIP8 | WIP16 |  | Subsystem | Documentation | N/A | Fixed | rajalakshmi_j | 0 | 3 | User Documentation | Support-System | NO | Closed | WIP12SegReq | Ovation |  | WIP16 | May 29, 2019 2:05:05 PM GMT+08:00
CTUDT00600085 | Refresh | Closed | lin_s | October 11, 2018 3:08:55 PM GMT+08:00 | Refresh, IFU: No clear description of feature: Pause button is disabled during clinical scan initialization. | WIP8 | WIP16 |  | Subsystem | Documentation | N/A | Fixed | rajalakshmi_j | 0 | 3 | User Documentation | Support-System | NO | Closed | WIP12SegReq | Brilliance 64 |  | WIP16 | May 29, 2019 2:05:05 PM GMT+08:00
CTUDT00600085 | Refresh | Closed | lin_s | October 11, 2018 3:08:55 PM GMT+08:00 | Refresh, IFU: No clear description of feature: Pause button is disabled during clinical scan initialization. | WIP8 | WIP16 |  | Subsystem | Documentation | N/A | Fixed | rajalakshmi_j | 0 | 3 | User Documentation | Support-System | NO | Closed | WIP12SegReq | Ingenuity 64 |  | WIP16 | May 29, 2019 2:05:05 PM GMT+08:00
CTUDT00600085 | Refresh | Closed | lin_s | October 11, 2018 3:08:55 PM GMT+08:00 | Refresh, IFU: No clear description of feature: Pause button is disabled during clinical scan initialization. | WIP8 | WIP16 |  | Subsystem | Documentation | N/A | Fixed | rajalakshmi_j | 0 | 3 | User Documentation | Support-System | NO | Closed | WIP12SegReq | IQon1 |  | WIP16 | May 29, 2019 2:05:05 PM GMT+08:00
CTUDT00600082 | Refresh | Closed | panchami_pradeep | October 11, 2018 2:59:07 PM GMT+08:00 | Refresh: Perfusion scans (Cardiac perfusion and brain perfusion) getting stuck while clinical scan is going on. | WIP10 |  |  | Subsystem | Perfusion Scan - Performance | NA | Duplicate | babu_s | 0 |  |  | Console4 | YES | Closed | WIP12SegReq | IQon1
CTUDT00600080 | Refresh | Closed | panchami_pradeep | October 11, 2018 1:30:24 PM GMT+08:00 | Refresh: PSC: Reregister libraries tool is displaying command prompt error. | WIP10 | WIP12 |  | Subsystem | Reregister tool | N/A | Fixed | Sunil_Kumar_B | 0 | 3 | Service Tools_Software | Service Tools | NO | Closed | WIP12SubCom | Ovation |  | WIP12 | December 6, 2018 5:08:25 PM GMT+08:00
CTUDT00600080 | Refresh | Closed | panchami_pradeep | October 11, 2018 1:30:24 PM GMT+08:00 | Refresh: PSC: Reregister libraries tool is displaying command prompt error. | WIP10 | WIP12 |  | Subsystem | Reregister tool | N/A | Fixed | Sunil_Kumar_B | 0 | 3 | Service Tools_Software | Service Tools | NO | Closed | WIP12SubCom | Brilliance 64 |  | WIP12 | December 6, 2018 5:08:25 PM GMT+08:00
CTUDT00600080 | Refresh | Closed | panchami_pradeep | October 11, 2018 1:30:24 PM GMT+08:00 | Refresh: PSC: Reregister libraries tool is displaying command prompt error. | WIP10 | WIP12 |  | Subsystem | Reregister tool | N/A | Fixed | Sunil_Kumar_B | 0 | 3 | Service Tools_Software | Service Tools | NO | Closed | WIP12SubCom | Brilliance Big Bore |  | WIP12 | December 6, 2018 5:08:25 PM GMT+08:00
CTUDT00600080 | Refresh | Closed | panchami_pradeep | October 11, 2018 1:30:24 PM GMT+08:00 | Refresh: PSC: Reregister libraries tool is displaying command prompt error. | WIP10 | WIP12 |  | Subsystem | Reregister tool | N/A | Fixed | Sunil_Kumar_B | 0 | 3 | Service Tools_Software | Service Tools | NO | Closed | WIP12SubCom | Ingenuity 64 |  | WIP12 | December 6, 2018 5:08:25 PM GMT+08:00
CTUDT00600080 | Refresh | Closed | panchami_pradeep | October 11, 2018 1:30:24 PM GMT+08:00 | Refresh: PSC: Reregister libraries tool is displaying command prompt error. | WIP10 | WIP12 |  | Subsystem | Reregister tool | N/A | Fixed | Sunil_Kumar_B | 0 | 3 | Service Tools_Software | Service Tools | NO | Closed | WIP12SubCom | IQon1 |  | WIP12 | December 6, 2018 5:08:25 PM GMT+08:00
CTUDT00600078 | Refresh | Closed | panchami_pradeep | October 11, 2018 1:18:17 PM GMT+08:00 | Refresh: PSC: Mapdrive tool is displaying script error while trying to map any drive. | WIP10 | WIP11.1 |  | Subsystem | Service Tools - Mapdrive tool | N/A | Fixed | Sunil_Kumar_B | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP11.1SubCom | Ovation |  | WIP11.1 | November 29, 2018 8:14:17 PM GMT+08:00
CTUDT00600078 | Refresh | Closed | panchami_pradeep | October 11, 2018 1:18:17 PM GMT+08:00 | Refresh: PSC: Mapdrive tool is displaying script error while trying to map any drive. | WIP10 | WIP11.1 |  | Subsystem | Service Tools - Mapdrive tool | N/A | Fixed | Sunil_Kumar_B | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP11.1SubCom | Brilliance 64 |  | WIP11.1 | November 29, 2018 8:14:17 PM GMT+08:00
CTUDT00600078 | Refresh | Closed | panchami_pradeep | October 11, 2018 1:18:17 PM GMT+08:00 | Refresh: PSC: Mapdrive tool is displaying script error while trying to map any drive. | WIP10 | WIP11.1 |  | Subsystem | Service Tools - Mapdrive tool | N/A | Fixed | Sunil_Kumar_B | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP11.1SubCom | Brilliance Big Bore |  | WIP11.1 | November 29, 2018 8:14:17 PM GMT+08:00
CTUDT00600078 | Refresh | Closed | panchami_pradeep | October 11, 2018 1:18:17 PM GMT+08:00 | Refresh: PSC: Mapdrive tool is displaying script error while trying to map any drive. | WIP10 | WIP11.1 |  | Subsystem | Service Tools - Mapdrive tool | N/A | Fixed | Sunil_Kumar_B | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP11.1SubCom | Ingenuity 64 |  | WIP11.1 | November 29, 2018 8:14:17 PM GMT+08:00
CTUDT00600078 | Refresh | Closed | panchami_pradeep | October 11, 2018 1:18:17 PM GMT+08:00 | Refresh: PSC: Mapdrive tool is displaying script error while trying to map any drive. | WIP10 | WIP11.1 |  | Subsystem | Service Tools - Mapdrive tool | N/A | Fixed | Sunil_Kumar_B | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP11.1SubCom | IQon1 |  | WIP11.1 | November 29, 2018 8:14:17 PM GMT+08:00
CTUDT00600074 | Refresh | Closed | sukumaran_s | October 10, 2018 9:38:15 PM GMT+08:00 | The User Manual pdf file is not opening after Visual Diagnostics analysis. | WIP8 | WIP18 | The service personnel can refer to User Manual available in the inCenter. | Subsystem | Adobe Reader - Visual Diagnostic | N/A | Fixed | bollineni_d | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP14SegReq | Ovation |  | WIP18 | June 7, 2019 7:52:40 PM GMT+08:00
CTUDT00600074 | Refresh | Closed | sukumaran_s | October 10, 2018 9:38:15 PM GMT+08:00 | The User Manual pdf file is not opening after Visual Diagnostics analysis. | WIP8 | WIP18 | The service personnel can refer to User Manual available in the inCenter. | Subsystem | Adobe Reader - Visual Diagnostic | N/A | Fixed | bollineni_d | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP14SegReq | Brilliance 64 |  | WIP18 | June 7, 2019 7:52:40 PM GMT+08:00
CTUDT00600074 | Refresh | Closed | sukumaran_s | October 10, 2018 9:38:15 PM GMT+08:00 | The User Manual pdf file is not opening after Visual Diagnostics analysis. | WIP8 | WIP18 | The service personnel can refer to User Manual available in the inCenter. | Subsystem | Adobe Reader - Visual Diagnostic | N/A | Fixed | bollineni_d | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP14SegReq | Brilliance Big Bore |  | WIP18 | June 7, 2019 7:52:40 PM GMT+08:00
CTUDT00600074 | Refresh | Closed | sukumaran_s | October 10, 2018 9:38:15 PM GMT+08:00 | The User Manual pdf file is not opening after Visual Diagnostics analysis. | WIP8 | WIP18 | The service personnel can refer to User Manual available in the inCenter. | Subsystem | Adobe Reader - Visual Diagnostic | N/A | Fixed | bollineni_d | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP14SegReq | Ingenuity 64 |  | WIP18 | June 7, 2019 7:52:40 PM GMT+08:00
CTUDT00600074 | Refresh | Closed | sukumaran_s | October 10, 2018 9:38:15 PM GMT+08:00 | The User Manual pdf file is not opening after Visual Diagnostics analysis. | WIP8 | WIP18 | The service personnel can refer to User Manual available in the inCenter. | Subsystem | Adobe Reader - Visual Diagnostic | N/A | Fixed | bollineni_d | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP14SegReq | IQon1 |  | WIP18 | June 7, 2019 7:52:40 PM GMT+08:00
CTUDT00600066 | Refresh | Closed | singh_p | October 10, 2018 8:06:04 PM GMT+08:00 | Refresh Ingenuity WIP 10 : Error pop up appeared on launch of Printer Configuration tool. | WIP10 | WIP12 |  | Subsystem | Printer Config Tool | N/A | Fixed | Sunil_Kumar_B | 0 | 3 | Service Tools_Software | Service Tools | NO | Closed | WIP12SegReq | Ovation |  | WIP12 | December 6, 2018 5:05:40 PM GMT+08:00
CTUDT00600066 | Refresh | Closed | singh_p | October 10, 2018 8:06:04 PM GMT+08:00 | Refresh Ingenuity WIP 10 : Error pop up appeared on launch of Printer Configuration tool. | WIP10 | WIP12 |  | Subsystem | Printer Config Tool | N/A | Fixed | Sunil_Kumar_B | 0 | 3 | Service Tools_Software | Service Tools | NO | Closed | WIP12SegReq | Brilliance 64 |  | WIP12 | December 6, 2018 5:05:40 PM GMT+08:00
CTUDT00600066 | Refresh | Closed | singh_p | October 10, 2018 8:06:04 PM GMT+08:00 | Refresh Ingenuity WIP 10 : Error pop up appeared on launch of Printer Configuration tool. | WIP10 | WIP12 |  | Subsystem | Printer Config Tool | N/A | Fixed | Sunil_Kumar_B | 0 | 3 | Service Tools_Software | Service Tools | NO | Closed | WIP12SegReq | Brilliance Big Bore |  | WIP12 | December 6, 2018 5:05:40 PM GMT+08:00
CTUDT00600066 | Refresh | Closed | singh_p | October 10, 2018 8:06:04 PM GMT+08:00 | Refresh Ingenuity WIP 10 : Error pop up appeared on launch of Printer Configuration tool. | WIP10 | WIP12 |  | Subsystem | Printer Config Tool | N/A | Fixed | Sunil_Kumar_B | 0 | 3 | Service Tools_Software | Service Tools | NO | Closed | WIP12SegReq | Ingenuity 64 |  | WIP12 | December 6, 2018 5:05:40 PM GMT+08:00
CTUDT00600066 | Refresh | Closed | singh_p | October 10, 2018 8:06:04 PM GMT+08:00 | Refresh Ingenuity WIP 10 : Error pop up appeared on launch of Printer Configuration tool. | WIP10 | WIP12 |  | Subsystem | Printer Config Tool | N/A | Fixed | Sunil_Kumar_B | 0 | 3 | Service Tools_Software | Service Tools | NO | Closed | WIP12SegReq | IQon1 |  | WIP12 | December 6, 2018 5:05:40 PM GMT+08:00
CTUDT00600064 | Refresh | Closed | hiremath_b | October 10, 2018 7:58:37 PM GMT+08:00 | Refresh: IST username is not displayed in Login screen even when IST dongle is detected and able to login to the system. | WIP10 Sanity | WIP11 | N/A | Subsystem |  | N/A | Fixed | singh_j | 0 | 2 | Console Software | Console4 | NO | Closed | WIP11SegReq | Ovation |  | WIP11 | November 12, 2018 8:09:20 PM GMT+08:00
CTUDT00600064 | Refresh | Closed | hiremath_b | October 10, 2018 7:58:37 PM GMT+08:00 | Refresh: IST username is not displayed in Login screen even when IST dongle is detected and able to login to the system. | WIP10 Sanity | WIP11 | N/A | Subsystem |  | N/A | Fixed | singh_j | 0 | 2 | Console Software | Console4 | NO | Closed | WIP11SegReq | Brilliance 64 |  | WIP11 | November 12, 2018 8:09:20 PM GMT+08:00
CTUDT00600064 | Refresh | Closed | hiremath_b | October 10, 2018 7:58:37 PM GMT+08:00 | Refresh: IST username is not displayed in Login screen even when IST dongle is detected and able to login to the system. | WIP10 Sanity | WIP11 | N/A | Subsystem |  | N/A | Fixed | singh_j | 0 | 2 | Console Software | Console4 | NO | Closed | WIP11SegReq | Brilliance Big Bore |  | WIP11 | November 12, 2018 8:09:20 PM GMT+08:00
CTUDT00600064 | Refresh | Closed | hiremath_b | October 10, 2018 7:58:37 PM GMT+08:00 | Refresh: IST username is not displayed in Login screen even when IST dongle is detected and able to login to the system. | WIP10 Sanity | WIP11 | N/A | Subsystem |  | N/A | Fixed | singh_j | 0 | 2 | Console Software | Console4 | NO | Closed | WIP11SegReq | Ingenuity 64 |  | WIP11 | November 12, 2018 8:09:20 PM GMT+08:00
CTUDT00600064 | Refresh | Closed | hiremath_b | October 10, 2018 7:58:37 PM GMT+08:00 | Refresh: IST username is not displayed in Login screen even when IST dongle is detected and able to login to the system. | WIP10 Sanity | WIP11 | N/A | Subsystem |  | N/A | Fixed | singh_j | 0 | 2 | Console Software | Console4 | NO | Closed | WIP11SegReq | IQon1 |  | WIP11 | November 12, 2018 8:09:20 PM GMT+08:00
CTUDT00600060 | Refresh | Closed | hiremath_b | October 10, 2018 6:45:13 PM GMT+08:00 | Refresh: E-Stop (gantry not ready error) does not come up on login as any Clinical user after password is reset. | WIP10 |  |  | Subsystem | E-stop window - User management |  | Works As Specified / Not A Defect | v_sumanth | 1 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00600058 | Refresh | Closed | hiremath_b | October 10, 2018 6:35:47 PM GMT+08:00 | Refresh: “Configuration was saved \r\n, could not restart the service. Please restart the system “ message pops up on saving changes in LAN configuration tool when launched from IT admin tool. | WIP10 | WIP12 |  | Subsystem | Lan Config - saving changes | NA | Fixed | veerabhadraiah | 0 | 2 | USP | iCAP | NO | Closed | WIP12SegReq | Ovation |  | WIP12 | February 25, 2019 4:31:26 PM GMT+08:00
CTUDT00600058 | Refresh | Closed | hiremath_b | October 10, 2018 6:35:47 PM GMT+08:00 | Refresh: “Configuration was saved \r\n, could not restart the service. Please restart the system “ message pops up on saving changes in LAN configuration tool when launched from IT admin tool. | WIP10 | WIP12 |  | Subsystem | Lan Config - saving changes | NA | Fixed | veerabhadraiah | 0 | 2 | USP | iCAP | NO | Closed | WIP12SegReq | Brilliance 64 |  | WIP12 | February 25, 2019 4:31:26 PM GMT+08:00
CTUDT00600058 | Refresh | Closed | hiremath_b | October 10, 2018 6:35:47 PM GMT+08:00 | Refresh: “Configuration was saved \r\n, could not restart the service. Please restart the system “ message pops up on saving changes in LAN configuration tool when launched from IT admin tool. | WIP10 | WIP12 |  | Subsystem | Lan Config - saving changes | NA | Fixed | veerabhadraiah | 0 | 2 | USP | iCAP | NO | Closed | WIP12SegReq | Brilliance Big Bore |  | WIP12 | February 25, 2019 4:31:26 PM GMT+08:00
CTUDT00600058 | Refresh | Closed | hiremath_b | October 10, 2018 6:35:47 PM GMT+08:00 | Refresh: “Configuration was saved \r\n, could not restart the service. Please restart the system “ message pops up on saving changes in LAN configuration tool when launched from IT admin tool. | WIP10 | WIP12 |  | Subsystem | Lan Config - saving changes | NA | Fixed | veerabhadraiah | 0 | 2 | USP | iCAP | NO | Closed | WIP12SegReq | Ingenuity 64 |  | WIP12 | February 25, 2019 4:31:26 PM GMT+08:00
CTUDT00600058 | Refresh | Closed | hiremath_b | October 10, 2018 6:35:47 PM GMT+08:00 | Refresh: “Configuration was saved \r\n, could not restart the service. Please restart the system “ message pops up on saving changes in LAN configuration tool when launched from IT admin tool. | WIP10 | WIP12 |  | Subsystem | Lan Config - saving changes | NA | Fixed | veerabhadraiah | 0 | 2 | USP | iCAP | NO | Closed | WIP12SegReq | IQon1 |  | WIP12 | February 25, 2019 4:31:26 PM GMT+08:00
CTUDT00600055 | Refresh | Closed | hiremath_b | October 10, 2018 6:05:09 PM GMT+08:00 | Refresh: Bugrep error window pops up on entering Invalid credentials in Dose Alert window of any user who has access to dose management. | WIP10 |  | No workaround | Subsystem |  | NA | Duplicate | ravindra_a | 0 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00600055 | Refresh | Closed | hiremath_b | October 10, 2018 6:05:09 PM GMT+08:00 | Refresh: Bugrep error window pops up on entering Invalid credentials in Dose Alert window of any user who has access to dose management. | WIP10 |  | No workaround | Subsystem |  | NA | Duplicate | ravindra_a | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00600053 | Refresh | Closed | hiremath_b | October 10, 2018 5:47:28 PM GMT+08:00 | Refresh: Bugrep error window pops up on entering valid credentials in Dose Alert window of any clinical user who has not logged in at least once after resetting password/user creation. | WIP10 | WIP12 | No workaround | Subsystem | Dose Management - Dose Alert window | NA | Fixed | ravindra_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP12SubCom | Ovation |  | WIP12 | December 18, 2018 4:32:44 PM GMT+08:00
CTUDT00600053 | Refresh | Closed | hiremath_b | October 10, 2018 5:47:28 PM GMT+08:00 | Refresh: Bugrep error window pops up on entering valid credentials in Dose Alert window of any clinical user who has not logged in at least once after resetting password/user creation. | WIP10 | WIP12 | No workaround | Subsystem | Dose Management - Dose Alert window | NA | Fixed | ravindra_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP12SubCom | Ingenuity 64 |  | WIP12 | December 18, 2018 4:32:44 PM GMT+08:00
CTUDT00600049 | Refresh | Closed | pB_s | October 10, 2018 5:17:50 PM GMT+08:00 | [Refresh] : Restrictions popup is getting displayed every time while logging into console every time | WIP10 | WIP11 | User can close the dialog and continue working with the system. | Subsystem |  | NA | Fixed | mohit_p | 0 | 3 | Console Software | Console4 | NO | Closed | WIP11SegReq | Ovation |  | WIP11 | October 31, 2018 7:28:24 PM GMT+08:00
CTUDT00600049 | Refresh | Closed | pB_s | October 10, 2018 5:17:50 PM GMT+08:00 | [Refresh] : Restrictions popup is getting displayed every time while logging into console every time | WIP10 | WIP11 | User can close the dialog and continue working with the system. | Subsystem |  | NA | Fixed | mohit_p | 0 | 3 | Console Software | Console4 | NO | Closed | WIP11SegReq | Brilliance 64 |  | WIP11 | October 31, 2018 7:28:24 PM GMT+08:00
CTUDT00600049 | Refresh | Closed | pB_s | October 10, 2018 5:17:50 PM GMT+08:00 | [Refresh] : Restrictions popup is getting displayed every time while logging into console every time | WIP10 | WIP11 | User can close the dialog and continue working with the system. | Subsystem |  | NA | Fixed | mohit_p | 0 | 3 | Console Software | Console4 | NO | Closed | WIP11SegReq | Ingenuity 64 |  | WIP11 | October 31, 2018 7:28:24 PM GMT+08:00
CTUDT00600049 | Refresh | Closed | pB_s | October 10, 2018 5:17:50 PM GMT+08:00 | [Refresh] : Restrictions popup is getting displayed every time while logging into console every time | WIP10 | WIP11 | User can close the dialog and continue working with the system. | Subsystem |  | NA | Fixed | mohit_p | 0 | 3 | Console Software | Console4 | NO | Closed | WIP11SegReq | IQon1 |  | WIP11 | October 31, 2018 7:28:24 PM GMT+08:00
CTUDT00600047 | Refresh | Closed | pB_s | October 10, 2018 5:06:11 PM GMT+08:00 | [Refresh] : Error/Alert popup "ms-settings:display" is getting displayed in service mode while opening Display settings | WIP10 | WIP14 | None | Subsystem |  | None | Fixed | kapadia_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP11SegReq | Ovation |  | WIP14 | February 20, 2019 4:14:17 PM GMT+08:00
CTUDT00600047 | Refresh | Closed | pB_s | October 10, 2018 5:06:11 PM GMT+08:00 | [Refresh] : Error/Alert popup "ms-settings:display" is getting displayed in service mode while opening Display settings | WIP10 | WIP14 | None | Subsystem |  | None | Fixed | kapadia_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP11SegReq | Brilliance Big Bore |  | WIP14 | February 20, 2019 4:14:17 PM GMT+08:00
CTUDT00600047 | Refresh | Closed | pB_s | October 10, 2018 5:06:11 PM GMT+08:00 | [Refresh] : Error/Alert popup "ms-settings:display" is getting displayed in service mode while opening Display settings | WIP10 | WIP14 | None | Subsystem |  | None | Fixed | kapadia_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP11SegReq | Ingenuity 64 |  | WIP14 | February 20, 2019 4:14:17 PM GMT+08:00
CTUDT00600047 | Refresh | Closed | pB_s | October 10, 2018 5:06:11 PM GMT+08:00 | [Refresh] : Error/Alert popup "ms-settings:display" is getting displayed in service mode while opening Display settings | WIP10 | WIP14 | None | Subsystem |  | None | Fixed | kapadia_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP11SegReq | IQon1 |  | WIP14 | February 20, 2019 4:14:17 PM GMT+08:00
CTUDT00600046 | Refresh | Closed | singh_p | October 10, 2018 4:58:46 PM GMT+08:00 | Refresh Ingenuity WIP 10 : Temp Device folder is not cleaning up post restart of machine. | WIP10 |  |  | Subsystem | Temp Device - cleanup after restart | NA | Works As Specified / Not A Defect | kapadia_v | 0 | 2 |  | Console4 | NO | Closed |  | Ovation
CTUDT00600046 | Refresh | Closed | singh_p | October 10, 2018 4:58:46 PM GMT+08:00 | Refresh Ingenuity WIP 10 : Temp Device folder is not cleaning up post restart of machine. | WIP10 |  |  | Subsystem | Temp Device - cleanup after restart | NA | Works As Specified / Not A Defect | kapadia_v | 0 | 2 |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00600046 | Refresh | Closed | singh_p | October 10, 2018 4:58:46 PM GMT+08:00 | Refresh Ingenuity WIP 10 : Temp Device folder is not cleaning up post restart of machine. | WIP10 |  |  | Subsystem | Temp Device - cleanup after restart | NA | Works As Specified / Not A Defect | kapadia_v | 0 | 2 |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00600046 | Refresh | Closed | singh_p | October 10, 2018 4:58:46 PM GMT+08:00 | Refresh Ingenuity WIP 10 : Temp Device folder is not cleaning up post restart of machine. | WIP10 |  |  | Subsystem | Temp Device - cleanup after restart | NA | Works As Specified / Not A Defect | kapadia_v | 0 | 2 |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00600046 | Refresh | Closed | singh_p | October 10, 2018 4:58:46 PM GMT+08:00 | Refresh Ingenuity WIP 10 : Temp Device folder is not cleaning up post restart of machine. | WIP10 |  |  | Subsystem | Temp Device - cleanup after restart | NA | Works As Specified / Not A Defect | kapadia_v | 0 | 2 |  | Console4 | NO | Closed |  | IQon1
CTUDT00600044 | Refresh | Closed | beck_d | October 10, 2018 4:56:42 PM GMT+08:00 | Refresh: Calibrations: Acceptance and Constancy Test: Graphic presentation is missing in the result | WIP9.1 |  |  | Subsystem |  |  | Duplicate | yosefk | 0 |  |  |  |  | Closed |  | Brilliance 64
CTUDT00600039 | Refresh | Closed | singh_p | October 10, 2018 3:18:42 PM GMT+08:00 | Refresh Ingenuity WIP 10 : There was NO audio when recorded voice was played in voice manager. | WIP10 |  | • Login in Service Mode • Open up Service Tools > Configurations > System > Adjust Audio Properties • Recording (Ensure the Headphones + Mic are connected) • Levels – Set it to 80-100% | Subsystem | Voice Manager - No audio; Blocker | N/A | Duplicate | hilas | 0 | 2 |  | Support-System | NO | Closed | SubCom | Ovation
CTUDT00600039 | Refresh | Closed | singh_p | October 10, 2018 3:18:42 PM GMT+08:00 | Refresh Ingenuity WIP 10 : There was NO audio when recorded voice was played in voice manager. | WIP10 |  | • Login in Service Mode • Open up Service Tools > Configurations > System > Adjust Audio Properties • Recording (Ensure the Headphones + Mic are connected) • Levels – Set it to 80-100% | Subsystem | Voice Manager - No audio; Blocker | N/A | Duplicate | hilas | 0 | 2 |  | Support-System | NO | Closed | SubCom | Brilliance 64
CTUDT00600039 | Refresh | Closed | singh_p | October 10, 2018 3:18:42 PM GMT+08:00 | Refresh Ingenuity WIP 10 : There was NO audio when recorded voice was played in voice manager. | WIP10 |  | • Login in Service Mode • Open up Service Tools > Configurations > System > Adjust Audio Properties • Recording (Ensure the Headphones + Mic are connected) • Levels – Set it to 80-100% | Subsystem | Voice Manager - No audio; Blocker | N/A | Duplicate | hilas | 0 | 2 |  | Support-System | NO | Closed | SubCom | Brilliance Big Bore
CTUDT00600039 | Refresh | Closed | singh_p | October 10, 2018 3:18:42 PM GMT+08:00 | Refresh Ingenuity WIP 10 : There was NO audio when recorded voice was played in voice manager. | WIP10 |  | • Login in Service Mode • Open up Service Tools > Configurations > System > Adjust Audio Properties • Recording (Ensure the Headphones + Mic are connected) • Levels – Set it to 80-100% | Subsystem | Voice Manager - No audio; Blocker | N/A | Duplicate | hilas | 0 | 2 |  | Support-System | NO | Closed | SubCom | Ingenuity 64
CTUDT00600039 | Refresh | Closed | singh_p | October 10, 2018 3:18:42 PM GMT+08:00 | Refresh Ingenuity WIP 10 : There was NO audio when recorded voice was played in voice manager. | WIP10 |  | • Login in Service Mode • Open up Service Tools > Configurations > System > Adjust Audio Properties • Recording (Ensure the Headphones + Mic are connected) • Levels – Set it to 80-100% | Subsystem | Voice Manager - No audio; Blocker | N/A | Duplicate | hilas | 0 | 2 |  | Support-System | NO | Closed | SubCom | IQon1
CTUDT00600037 | Refresh | Closed | naik_p | October 10, 2018 2:24:27 PM GMT+08:00 | Refresh:Critical error for any clinical user login for the first time. | WIP10 |  |  | Subsystem | User Login - First time login |  | Duplicate | karkour | 0 |  |  | GCX | NO | Closed |  | Ovation
CTUDT00600037 | Refresh | Closed | naik_p | October 10, 2018 2:24:27 PM GMT+08:00 | Refresh:Critical error for any clinical user login for the first time. | WIP10 |  |  | Subsystem | User Login - First time login |  | Duplicate | karkour | 0 |  |  | GCX | NO | Closed |  | Ingenuity 64
CTUDT00600037 | Refresh | Closed | naik_p | October 10, 2018 2:24:27 PM GMT+08:00 | Refresh:Critical error for any clinical user login for the first time. | WIP10 |  |  | Subsystem | User Login - First time login |  | Duplicate | karkour | 0 |  |  | GCX | NO | Closed |  | IQon1
CTUDT00600031 | Refresh | Closed | vinoliya_j | October 9, 2018 9:22:43 PM GMT+08:00 | Refresh : Preferences ->Patient directory page throws error. | WIP 10 | WIP12 |  | Subsystem | Preferences - Patient Directory | NA | Fixed | veerabhadraiah | 0 | 2 | USP | iCAP | NO | Closed | WIP12SegReq | Ovation |  | WIP12 | December 14, 2018 5:15:17 PM GMT+08:00
CTUDT00600031 | Refresh | Closed | vinoliya_j | October 9, 2018 9:22:43 PM GMT+08:00 | Refresh : Preferences ->Patient directory page throws error. | WIP 10 | WIP12 |  | Subsystem | Preferences - Patient Directory | NA | Fixed | veerabhadraiah | 0 | 2 | USP | iCAP | NO | Closed | WIP12SegReq | Brilliance 64 |  | WIP12 | December 14, 2018 5:15:17 PM GMT+08:00
CTUDT00600031 | Refresh | Closed | vinoliya_j | October 9, 2018 9:22:43 PM GMT+08:00 | Refresh : Preferences ->Patient directory page throws error. | WIP 10 | WIP12 |  | Subsystem | Preferences - Patient Directory | NA | Fixed | veerabhadraiah | 0 | 2 | USP | iCAP | NO | Closed | WIP12SegReq | Brilliance Big Bore |  | WIP12 | December 14, 2018 5:15:17 PM GMT+08:00
CTUDT00600031 | Refresh | Closed | vinoliya_j | October 9, 2018 9:22:43 PM GMT+08:00 | Refresh : Preferences ->Patient directory page throws error. | WIP 10 | WIP12 |  | Subsystem | Preferences - Patient Directory | NA | Fixed | veerabhadraiah | 0 | 2 | USP | iCAP | NO | Closed | WIP12SegReq | Ingenuity 64 |  | WIP12 | December 14, 2018 5:15:17 PM GMT+08:00
CTUDT00600031 | Refresh | Closed | vinoliya_j | October 9, 2018 9:22:43 PM GMT+08:00 | Refresh : Preferences ->Patient directory page throws error. | WIP 10 | WIP12 |  | Subsystem | Preferences - Patient Directory | NA | Fixed | veerabhadraiah | 0 | 2 | USP | iCAP | NO | Closed | WIP12SegReq | IQon1 |  | WIP12 | December 14, 2018 5:15:17 PM GMT+08:00
CTUDT00600029 | Refresh | Closed | attali_s | October 9, 2018 8:24:52 PM GMT+08:00 | Refresh: Stability: No Table of Failures with Service User | WIP9.1 | WIP13 |  | Subsystem | Stability test | N/A | Fixed | yosefk | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | Ovation |  | WIP13 | January 20, 2019 3:58:34 PM GMT+08:00
CTUDT00600029 | Refresh | Closed | attali_s | October 9, 2018 8:24:52 PM GMT+08:00 | Refresh: Stability: No Table of Failures with Service User | WIP9.1 | WIP13 |  | Subsystem | Stability test | N/A | Fixed | yosefk | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | Ingenuity 64 |  | WIP13 | January 20, 2019 3:58:34 PM GMT+08:00
CTUDT00600026 | Refresh | Closed | beck_d | October 9, 2018 6:44:48 PM GMT+08:00 | Refresh: Calibrations: Automatic phantom centering does not detect the tungsten edges of the CSOF phantom correctly | WIP9.1 | WIP18 |  | Subsystem | Calibrations - Automatoic Centering | NaN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | Ingenuity 64 |  | WIP18 | June 24, 2019 10:05:58 PM GMT+08:00
CTUDT00600026 | Refresh | Closed | beck_d | October 9, 2018 6:44:48 PM GMT+08:00 | Refresh: Calibrations: Automatic phantom centering does not detect the tungsten edges of the CSOF phantom correctly | WIP9.1 | WIP18 |  | Subsystem | Calibrations - Automatoic Centering | NaN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | IQon1 |  | WIP18 | June 24, 2019 10:05:58 PM GMT+08:00
CTUDT00600023 | Refresh | Closed | naik_p | October 9, 2018 12:30:22 PM GMT+08:00 | Req:2929_Quality Assurance tool displays pop up stating " GantryCalibWCF Host has stopped working"_WIP10. | WIP10 | WIP11.1 |  | Subsystem |  | N/A | Fixed | kapadia_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP11SegReq | Ovation |  | WIP11.1 | November 22, 2018 1:13:47 PM GMT+08:00
CTUDT00600023 | Refresh | Closed | naik_p | October 9, 2018 12:30:22 PM GMT+08:00 | Req:2929_Quality Assurance tool displays pop up stating " GantryCalibWCF Host has stopped working"_WIP10. | WIP10 | WIP11.1 |  | Subsystem |  | N/A | Fixed | kapadia_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP11SegReq | Brilliance 64 |  | WIP11.1 | November 22, 2018 1:13:47 PM GMT+08:00
CTUDT00600023 | Refresh | Closed | naik_p | October 9, 2018 12:30:22 PM GMT+08:00 | Req:2929_Quality Assurance tool displays pop up stating " GantryCalibWCF Host has stopped working"_WIP10. | WIP10 | WIP11.1 |  | Subsystem |  | N/A | Fixed | kapadia_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP11SegReq | Ingenuity 64 |  | WIP11.1 | November 22, 2018 1:13:47 PM GMT+08:00
CTUDT00600023 | Refresh | Closed | naik_p | October 9, 2018 12:30:22 PM GMT+08:00 | Req:2929_Quality Assurance tool displays pop up stating " GantryCalibWCF Host has stopped working"_WIP10. | WIP10 | WIP11.1 |  | Subsystem |  | N/A | Fixed | kapadia_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP11SegReq | IQon1 |  | WIP11.1 | November 22, 2018 1:13:47 PM GMT+08:00
CTUDT00600020 | Refresh | Closed | naik_p | October 8, 2018 6:44:46 PM GMT+08:00 | Repeat option missing for dual surview after changing the patient position. | WIP9.1 | WIP12 | User can repeat the dual surview scan by right clicking on surview step either on the scan ruler or on the series list. | Subsystem | EC parameters - Repeat option is missing | NA | Fixed | ravindra_a | 0 | 3 | Console Software | Console4 | NO | Closed | WIP12SubCom | Ovation |  | WIP12 | December 18, 2018 11:54:45 PM GMT+08:00
CTUDT00600020 | Refresh | Closed | naik_p | October 8, 2018 6:44:46 PM GMT+08:00 | Repeat option missing for dual surview after changing the patient position. | WIP9.1 | WIP12 | User can repeat the dual surview scan by right clicking on surview step either on the scan ruler or on the series list. | Subsystem | EC parameters - Repeat option is missing | NA | Fixed | ravindra_a | 0 | 3 | Console Software | Console4 | NO | Closed | WIP12SubCom | Brilliance 64 |  | WIP12 | December 18, 2018 11:54:45 PM GMT+08:00
CTUDT00600020 | Refresh | Closed | naik_p | October 8, 2018 6:44:46 PM GMT+08:00 | Repeat option missing for dual surview after changing the patient position. | WIP9.1 | WIP12 | User can repeat the dual surview scan by right clicking on surview step either on the scan ruler or on the series list. | Subsystem | EC parameters - Repeat option is missing | NA | Fixed | ravindra_a | 0 | 3 | Console Software | Console4 | NO | Closed | WIP12SubCom | Ingenuity 64 |  | WIP12 | December 18, 2018 11:54:45 PM GMT+08:00
CTUDT00600020 | Refresh | Closed | naik_p | October 8, 2018 6:44:46 PM GMT+08:00 | Repeat option missing for dual surview after changing the patient position. | WIP9.1 | WIP12 | User can repeat the dual surview scan by right clicking on surview step either on the scan ruler or on the series list. | Subsystem | EC parameters - Repeat option is missing | NA | Fixed | ravindra_a | 0 | 3 | Console Software | Console4 | NO | Closed | WIP12SubCom | IQon1 |  | WIP12 | December 18, 2018 11:54:45 PM GMT+08:00
CTUDT00600017 | Refresh | Closed | finestine_m | October 8, 2018 4:38:08 PM GMT+08:00 | Calibrations: Bad Detectors: Missing scan parameters from the spec in Debug mode | WIP9.1 | WIP12 | If somebody crave to know: "WHAT IS THE Focal Spot Type in THIS SCAN !!!" he(she) can look in MDU (by means of MDU viewer). Anyway, no one some (head ill person) was not detected in last 10 years... | Subsystem | Calibrations - Bad Detector | N/A | Fixed | yosefk | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP12SegReq | Ovation |  | WIP12 | January 20, 2019 3:21:42 PM GMT+08:00
CTUDT00600017 | Refresh | Closed | finestine_m | October 8, 2018 4:38:08 PM GMT+08:00 | Calibrations: Bad Detectors: Missing scan parameters from the spec in Debug mode | WIP9.1 | WIP12 | If somebody crave to know: "WHAT IS THE Focal Spot Type in THIS SCAN !!!" he(she) can look in MDU (by means of MDU viewer). Anyway, no one some (head ill person) was not detected in last 10 years... | Subsystem | Calibrations - Bad Detector | N/A | Fixed | yosefk | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP12SegReq | Brilliance 64 |  | WIP12 | January 20, 2019 3:21:42 PM GMT+08:00
CTUDT00600017 | Refresh | Closed | finestine_m | October 8, 2018 4:38:08 PM GMT+08:00 | Calibrations: Bad Detectors: Missing scan parameters from the spec in Debug mode | WIP9.1 | WIP12 | If somebody crave to know: "WHAT IS THE Focal Spot Type in THIS SCAN !!!" he(she) can look in MDU (by means of MDU viewer). Anyway, no one some (head ill person) was not detected in last 10 years... | Subsystem | Calibrations - Bad Detector | N/A | Fixed | yosefk | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP12SegReq | Brilliance Big Bore |  | WIP12 | January 20, 2019 3:21:42 PM GMT+08:00
CTUDT00600017 | Refresh | Closed | finestine_m | October 8, 2018 4:38:08 PM GMT+08:00 | Calibrations: Bad Detectors: Missing scan parameters from the spec in Debug mode | WIP9.1 | WIP12 | If somebody crave to know: "WHAT IS THE Focal Spot Type in THIS SCAN !!!" he(she) can look in MDU (by means of MDU viewer). Anyway, no one some (head ill person) was not detected in last 10 years... | Subsystem | Calibrations - Bad Detector | N/A | Fixed | yosefk | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP12SegReq | Ingenuity 64 |  | WIP12 | January 20, 2019 3:21:42 PM GMT+08:00
CTUDT00600017 | Refresh | Closed | finestine_m | October 8, 2018 4:38:08 PM GMT+08:00 | Calibrations: Bad Detectors: Missing scan parameters from the spec in Debug mode | WIP9.1 | WIP12 | If somebody crave to know: "WHAT IS THE Focal Spot Type in THIS SCAN !!!" he(she) can look in MDU (by means of MDU viewer). Anyway, no one some (head ill person) was not detected in last 10 years... | Subsystem | Calibrations - Bad Detector | N/A | Fixed | yosefk | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP12SegReq | IQon1 |  | WIP12 | January 20, 2019 3:21:42 PM GMT+08:00
CTUDT00600014 | Refresh | Closed | singh_p | October 8, 2018 3:23:08 PM GMT+08:00 | Refresh Ingenuity WIP 9.1 : System Error pop up appeared when bolus scan was paused. | WIP9.1 | WIP12 |  | System | Pause Scan - Bolus | N/A | Fixed | chokkalingam_r | 0 | 2 | Console Software | Console4 | YES | Closed | WIP12SubCom | Ovation |  | WIP12 | December 3, 2018 12:31:27 PM GMT+08:00
CTUDT00600014 | Refresh | Closed | singh_p | October 8, 2018 3:23:08 PM GMT+08:00 | Refresh Ingenuity WIP 9.1 : System Error pop up appeared when bolus scan was paused. | WIP9.1 | WIP12 |  | System | Pause Scan - Bolus | N/A | Fixed | chokkalingam_r | 0 | 2 | Console Software | Console4 | YES | Closed | WIP12SubCom | Brilliance 64 |  | WIP12 | December 3, 2018 12:31:27 PM GMT+08:00
CTUDT00600014 | Refresh | Closed | singh_p | October 8, 2018 3:23:08 PM GMT+08:00 | Refresh Ingenuity WIP 9.1 : System Error pop up appeared when bolus scan was paused. | WIP9.1 | WIP12 |  | System | Pause Scan - Bolus | N/A | Fixed | chokkalingam_r | 0 | 2 | Console Software | Console4 | YES | Closed | WIP12SubCom | Ingenuity 64 |  | WIP12 | December 3, 2018 12:31:27 PM GMT+08:00
CTUDT00600014 | Refresh | Closed | singh_p | October 8, 2018 3:23:08 PM GMT+08:00 | Refresh Ingenuity WIP 9.1 : System Error pop up appeared when bolus scan was paused. | WIP9.1 | WIP12 |  | System | Pause Scan - Bolus | N/A | Fixed | chokkalingam_r | 0 | 2 | Console Software | Console4 | YES | Closed | WIP12SubCom | IQon1 |  | WIP12 | December 3, 2018 12:31:27 PM GMT+08:00
CTUDT00600008 | Refresh | Closed | nimma_p | October 5, 2018 5:12:55 PM GMT+08:00 | Refresh:-Recon:-Surview being displayed under View 2, when we press recon by selecting the whole study from studies tab from the PD. | WIP9.1 | DHF318051 Rev 14 HOST Verification Test Protocol | Choose results to recon | Subsystem | Test Design | N/A | Fixed | ramesh_t | 0 | 4 | Test Design | Console4 | NO | Closed | WIP16SegReq | Ovation |  | DHF318051 Rev 14 HOST Verification Test Protocol | May 2, 2019 8:45:15 PM GMT+08:00
CTUDT00600008 | Refresh | Closed | nimma_p | October 5, 2018 5:12:55 PM GMT+08:00 | Refresh:-Recon:-Surview being displayed under View 2, when we press recon by selecting the whole study from studies tab from the PD. | WIP9.1 | DHF318051 Rev 14 HOST Verification Test Protocol | Choose results to recon | Subsystem | Test Design | N/A | Fixed | ramesh_t | 0 | 4 | Test Design | Console4 | NO | Closed | WIP16SegReq | Brilliance 64 |  | DHF318051 Rev 14 HOST Verification Test Protocol | May 2, 2019 8:45:15 PM GMT+08:00
CTUDT00600008 | Refresh | Closed | nimma_p | October 5, 2018 5:12:55 PM GMT+08:00 | Refresh:-Recon:-Surview being displayed under View 2, when we press recon by selecting the whole study from studies tab from the PD. | WIP9.1 | DHF318051 Rev 14 HOST Verification Test Protocol | Choose results to recon | Subsystem | Test Design | N/A | Fixed | ramesh_t | 0 | 4 | Test Design | Console4 | NO | Closed | WIP16SegReq | Brilliance Big Bore |  | DHF318051 Rev 14 HOST Verification Test Protocol | May 2, 2019 8:45:15 PM GMT+08:00
CTUDT00600008 | Refresh | Closed | nimma_p | October 5, 2018 5:12:55 PM GMT+08:00 | Refresh:-Recon:-Surview being displayed under View 2, when we press recon by selecting the whole study from studies tab from the PD. | WIP9.1 | DHF318051 Rev 14 HOST Verification Test Protocol | Choose results to recon | Subsystem | Test Design | N/A | Fixed | ramesh_t | 0 | 4 | Test Design | Console4 | NO | Closed | WIP16SegReq | Ingenuity 64 |  | DHF318051 Rev 14 HOST Verification Test Protocol | May 2, 2019 8:45:15 PM GMT+08:00
CTUDT00600008 | Refresh | Closed | nimma_p | October 5, 2018 5:12:55 PM GMT+08:00 | Refresh:-Recon:-Surview being displayed under View 2, when we press recon by selecting the whole study from studies tab from the PD. | WIP9.1 | DHF318051 Rev 14 HOST Verification Test Protocol | Choose results to recon | Subsystem | Test Design | N/A | Fixed | ramesh_t | 0 | 4 | Test Design | Console4 | NO | Closed | WIP16SegReq | IQon1 |  | DHF318051 Rev 14 HOST Verification Test Protocol | May 2, 2019 8:45:15 PM GMT+08:00
CTUDT00600005 | Refresh | Closed | ramesh_t | October 4, 2018 8:07:47 PM GMT+08:00 | Refresh:Rules: Max pitch for non-Ultra High Resolution scans was more than the cut off value 1.38 | WIP9.1 | WIP11.1 |  | Subsystem |  | N/A | Fixed | kumar_l | 0 | 2 | Console Software | Console4 | YES | Closed | WIP11SegReq | Ovation |  | WIP11.1 | November 7, 2018 7:34:01 PM GMT+08:00
CTUDT00600005 | Refresh | Closed | ramesh_t | October 4, 2018 8:07:47 PM GMT+08:00 | Refresh:Rules: Max pitch for non-Ultra High Resolution scans was more than the cut off value 1.38 | WIP9.1 | WIP11.1 |  | Subsystem |  | N/A | Fixed | kumar_l | 0 | 2 | Console Software | Console4 | YES | Closed | WIP11SegReq | Brilliance 64 |  | WIP11.1 | November 7, 2018 7:34:01 PM GMT+08:00
CTUDT00600005 | Refresh | Closed | ramesh_t | October 4, 2018 8:07:47 PM GMT+08:00 | Refresh:Rules: Max pitch for non-Ultra High Resolution scans was more than the cut off value 1.38 | WIP9.1 | WIP11.1 |  | Subsystem |  | N/A | Fixed | kumar_l | 0 | 2 | Console Software | Console4 | YES | Closed | WIP11SegReq | Orion P |  | WIP11.1 | November 7, 2018 7:34:01 PM GMT+08:00
CTUDT00600005 | Refresh | Closed | ramesh_t | October 4, 2018 8:07:47 PM GMT+08:00 | Refresh:Rules: Max pitch for non-Ultra High Resolution scans was more than the cut off value 1.38 | WIP9.1 | WIP11.1 |  | Subsystem |  | N/A | Fixed | kumar_l | 0 | 2 | Console Software | Console4 | YES | Closed | WIP11SegReq | Ingenuity 64 |  | WIP11.1 | November 7, 2018 7:34:01 PM GMT+08:00
CTUDT00600005 | Refresh | Closed | ramesh_t | October 4, 2018 8:07:47 PM GMT+08:00 | Refresh:Rules: Max pitch for non-Ultra High Resolution scans was more than the cut off value 1.38 | WIP9.1 | WIP11.1 |  | Subsystem |  | N/A | Fixed | kumar_l | 0 | 2 | Console Software | Console4 | YES | Closed | WIP11SegReq | IQon1 |  | WIP11.1 | November 7, 2018 7:34:01 PM GMT+08:00
CTUDT00599998 | Refresh | Closed | s_ra | October 3, 2018 5:32:27 PM GMT+08:00 | HB: FoK: After upgrade KeV setting for the Spectral Results/MonoE changed to the default selection 70KeV | WIP510 - HB FoK 4.7.5 | WIP15 | YES 1. Copy the User exam card files in the Backup from (backup taken before the 1- click started) Backup file: Root-> Package->Console-> Scanner-> iCT2_64_DL ExamCardsRepository_User to anywhere desired on the D: drive 2. Go into Exam Card Manager 3. Import the impacted exam cards (Select overwrite during import) from the path to which the exam cards were copied in Step 1 4. Verify imported exam cards are correct | Subsystem | History Review | NA | Fixed | shreekanth_manniyath | 0 | 3 | Console Software | Console4 | NO | Closed | WIP14SubCom | IQon1 |  | WIP15 | March 21, 2019 8:11:04 PM GMT+08:00
CTUDT00599996 | Refresh | Closed | s_ra | October 3, 2018 5:21:52 PM GMT+08:00 | Patient age getting one year less, only in the morning on birthday. | 4.1.7 | WIP11.1 | N/A | System | Patient Scene - Pt. Age | N/A | Fixed | veeraswamy_k | 0 | 2 | Console Software | Console4 | NO | Closed | WIP11SegReq | Brilliance 64 | SubCase 0111365370 | WIP11.1 | November 22, 2018 5:22:54 PM GMT+08:00
CTUDT00599996 | Refresh | Closed | s_ra | October 3, 2018 5:21:52 PM GMT+08:00 | Patient age getting one year less, only in the morning on birthday. | 4.1.7 | WIP11.1 | N/A | System | Patient Scene - Pt. Age | N/A | Fixed | veeraswamy_k | 0 | 2 | Console Software | Console4 | NO | Closed | WIP11SegReq | Ingenuity 64 | SubCase 0111365370 | WIP11.1 | November 22, 2018 5:22:54 PM GMT+08:00
CTUDT00599995 | Refresh | Closed | s_ra | October 3, 2018 5:10:40 PM GMT+08:00 | V4.1.7: Dose Log Viewer is blank. | v4.1.7.19026 | WIP11 |  | System |  | N/A | Fixed | atin_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP11SegReq | Ovation | PR7741791 | WIP11 | October 18, 2018 5:03:07 PM GMT+08:00
CTUDT00599995 | Refresh | Closed | s_ra | October 3, 2018 5:10:40 PM GMT+08:00 | V4.1.7: Dose Log Viewer is blank. | v4.1.7.19026 | WIP11 |  | System |  | N/A | Fixed | atin_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP11SegReq | Brilliance 64 | PR7741791 | WIP11 | October 18, 2018 5:03:07 PM GMT+08:00
CTUDT00599995 | Refresh | Closed | s_ra | October 3, 2018 5:10:40 PM GMT+08:00 | V4.1.7: Dose Log Viewer is blank. | v4.1.7.19026 | WIP11 |  | System |  | N/A | Fixed | atin_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP11SegReq | Ingenuity 64 | PR7741791 | WIP11 | October 18, 2018 5:03:07 PM GMT+08:00
CTUDT00599995 | Refresh | Closed | s_ra | October 3, 2018 5:10:40 PM GMT+08:00 | V4.1.7: Dose Log Viewer is blank. | v4.1.7.19026 | WIP11 |  | System |  | N/A | Fixed | atin_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP11SegReq | IQon1 | PR7741791 | WIP11 | October 18, 2018 5:03:07 PM GMT+08:00
CTUDT00599992 | Refresh | Closed | attali_s | October 3, 2018 2:15:04 AM GMT+08:00 | Refresh: Air Calibration: Wrong max number of Delta Tables | WIP9.1 | WIP12 | N/A | Subsystem | Calibrations - Air Calibration | N/A | Fixed | elviral | 0 | 4 | DMS Software | DMS | NO | Closed | WIP12SegReq | Ovation |  | WIP12 | December 13, 2018 3:23:53 PM GMT+08:00
CTUDT00599992 | Refresh | Closed | attali_s | October 3, 2018 2:15:04 AM GMT+08:00 | Refresh: Air Calibration: Wrong max number of Delta Tables | WIP9.1 | WIP12 | N/A | Subsystem | Calibrations - Air Calibration | N/A | Fixed | elviral | 0 | 4 | DMS Software | DMS | NO | Closed | WIP12SegReq | IQon1 |  | WIP12 | December 13, 2018 3:23:53 PM GMT+08:00
CTUDT00599988 | Refresh | Closed | bagdady_b | October 2, 2018 7:49:25 PM GMT+08:00 | Refresh- Unable to scan Head Axial protocol with collimations of 96x0.625 or 112x0.625 | WIP9.1 |  | Re-try with a valid collimation (Not 112x or 96x) | Subsystem | Rule System | See Investigation Rationale Tab | Does Not Apply | sanjeeva_a | 0 | 3 |  | Console4 | NO | Closed | WIP16SegReq | Ovation
CTUDT00599986 | Refresh | Closed | fitusi_r | September 30, 2018 6:38:17 PM GMT+08:00 | Refresh, ICT -Login User is empty. | WIP9.1 | WIP11 |  | Subsystem |  | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP11SegReq | Ovation |  | WIP11 | November 12, 2018 6:06:19 PM GMT+08:00
CTUDT00599986 | Refresh | Closed | fitusi_r | September 30, 2018 6:38:17 PM GMT+08:00 | Refresh, ICT -Login User is empty. | WIP9.1 | WIP11 |  | Subsystem |  | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP11SegReq | Brilliance 64 |  | WIP11 | November 12, 2018 6:06:19 PM GMT+08:00
CTUDT00599986 | Refresh | Closed | fitusi_r | September 30, 2018 6:38:17 PM GMT+08:00 | Refresh, ICT -Login User is empty. | WIP9.1 | WIP11 |  | Subsystem |  | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP11SegReq | Brilliance Big Bore |  | WIP11 | November 12, 2018 6:06:19 PM GMT+08:00
CTUDT00599986 | Refresh | Closed | fitusi_r | September 30, 2018 6:38:17 PM GMT+08:00 | Refresh, ICT -Login User is empty. | WIP9.1 | WIP11 |  | Subsystem |  | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP11SegReq | Ingenuity 64 |  | WIP11 | November 12, 2018 6:06:19 PM GMT+08:00
CTUDT00599986 | Refresh | Closed | fitusi_r | September 30, 2018 6:38:17 PM GMT+08:00 | Refresh, ICT -Login User is empty. | WIP9.1 | WIP11 |  | Subsystem |  | NA | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP11SegReq | IQon1 |  | WIP11 | November 12, 2018 6:06:19 PM GMT+08:00
CTUDT00599974 | Refresh | Closed | cohen_y | September 26, 2018 9:27:34 PM GMT+08:00 | Refresh: Air Calibration: Air Calibration Table names are not as defined in Air Calibration DHF170383 rev 11. | WIP9.1 | N/A |  | Subsystem | Documentation | N/A | Fixed | braunstein_d | 0 | 2 | Design Documentation | DMS | NO | Closed | WIP14SegReq | Ovation |  | N/A | January 23, 2019 3:34:59 PM GMT+08:00
CTUDT00599974 | Refresh | Closed | cohen_y | September 26, 2018 9:27:34 PM GMT+08:00 | Refresh: Air Calibration: Air Calibration Table names are not as defined in Air Calibration DHF170383 rev 11. | WIP9.1 | N/A |  | Subsystem | Documentation | N/A | Fixed | braunstein_d | 0 | 2 | Design Documentation | DMS | NO | Closed | WIP14SegReq | IQon1 |  | N/A | January 23, 2019 3:34:59 PM GMT+08:00
CTUDT00599972 | Refresh | Closed | attali_s | September 26, 2018 8:41:48 PM GMT+08:00 | Refresh - HCor Basic: the Noise limit of Head 80 kV is too low | WIP9.1 | WIP12 | N/A | Subsystem | Calibrations - HCOR basic | N/A | Fixed | elviral | 0 | 2 | DMS Software | DMS | NO | Closed | WIP12SegReq | Ingenuity 64 |  | WIP12 | December 13, 2018 2:49:19 PM GMT+08:00
CTUDT00599970 | Refresh | Closed | attali_s | September 26, 2018 7:07:03 PM GMT+08:00 | Refresh - Air Calibration: Delta Tables are not erased when the number of the delta tables reaches its maximum | WIP9.1 | N/A |  | Subsystem | Calibrations - Air Calibration | N/A | Fixed | gehtin_e | 0 | 2 | Test Design | DMS | NO | Closed | WIP13SegReq | iCT4 |  | N/A | December 19, 2018 4:35:25 PM GMT+08:00
CTUDT00599970 | Refresh | Closed | attali_s | September 26, 2018 7:07:03 PM GMT+08:00 | Refresh - Air Calibration: Delta Tables are not erased when the number of the delta tables reaches its maximum | WIP9.1 | N/A |  | Subsystem | Calibrations - Air Calibration | N/A | Fixed | gehtin_e | 0 | 2 | Test Design | DMS | NO | Closed | WIP13SegReq | Ingenuity 64 |  | N/A | December 19, 2018 4:35:25 PM GMT+08:00
CTUDT00599970 | Refresh | Closed | attali_s | September 26, 2018 7:07:03 PM GMT+08:00 | Refresh - Air Calibration: Delta Tables are not erased when the number of the delta tables reaches its maximum | WIP9.1 | N/A |  | Subsystem | Calibrations - Air Calibration | N/A | Fixed | gehtin_e | 0 | 2 | Test Design | DMS | NO | Closed | WIP13SegReq | Ingenuity BigBore |  | N/A | December 19, 2018 4:35:25 PM GMT+08:00
CTUDT00599970 | Refresh | Closed | attali_s | September 26, 2018 7:07:03 PM GMT+08:00 | Refresh - Air Calibration: Delta Tables are not erased when the number of the delta tables reaches its maximum | WIP9.1 | N/A |  | Subsystem | Calibrations - Air Calibration | N/A | Fixed | gehtin_e | 0 | 2 | Test Design | DMS | NO | Closed | WIP13SegReq | IQon1 |  | N/A | December 19, 2018 4:35:25 PM GMT+08:00
CTUDT00599958 | Refresh | Closed | nanjunda_s | September 25, 2018 6:16:32 PM GMT+08:00 | Requirement :Console-SRS.832- message given in the requirement is not correct for deletion of user created group. | WIP9.1 | Host SSRS DHF303211 Rev05 |  | Subsystem | Documentation | N/A | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | WIP12SegReq | Ovation |  | Host SSRS DHF303211 Rev05 | November 23, 2018 1:12:58 PM GMT+08:00
CTUDT00599958 | Refresh | Closed | nanjunda_s | September 25, 2018 6:16:32 PM GMT+08:00 | Requirement :Console-SRS.832- message given in the requirement is not correct for deletion of user created group. | WIP9.1 | Host SSRS DHF303211 Rev05 |  | Subsystem | Documentation | N/A | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | WIP12SegReq | Brilliance 64 |  | Host SSRS DHF303211 Rev05 | November 23, 2018 1:12:58 PM GMT+08:00
CTUDT00599958 | Refresh | Closed | nanjunda_s | September 25, 2018 6:16:32 PM GMT+08:00 | Requirement :Console-SRS.832- message given in the requirement is not correct for deletion of user created group. | WIP9.1 | Host SSRS DHF303211 Rev05 |  | Subsystem | Documentation | N/A | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | WIP12SegReq | Brilliance Big Bore |  | Host SSRS DHF303211 Rev05 | November 23, 2018 1:12:58 PM GMT+08:00
CTUDT00599958 | Refresh | Closed | nanjunda_s | September 25, 2018 6:16:32 PM GMT+08:00 | Requirement :Console-SRS.832- message given in the requirement is not correct for deletion of user created group. | WIP9.1 | Host SSRS DHF303211 Rev05 |  | Subsystem | Documentation | N/A | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | WIP12SegReq | Ingenuity 64 |  | Host SSRS DHF303211 Rev05 | November 23, 2018 1:12:58 PM GMT+08:00
CTUDT00599958 | Refresh | Closed | nanjunda_s | September 25, 2018 6:16:32 PM GMT+08:00 | Requirement :Console-SRS.832- message given in the requirement is not correct for deletion of user created group. | WIP9.1 | Host SSRS DHF303211 Rev05 |  | Subsystem | Documentation | N/A | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | WIP12SegReq | IQon1 |  | Host SSRS DHF303211 Rev05 | November 23, 2018 1:12:58 PM GMT+08:00
CTUDT00599956 | Refresh | Closed | pB_s | September 25, 2018 4:49:28 PM GMT+08:00 | A new line "Entering idle mode due to keyboard/mouse inactivity" is not added to the ConsoleIdleDetectionWindow with the new increased time. | WIP9.1 |  |  | Subsystem | ConsoleDetectionWindow |  | Works As Specified / Not A Defect | sanjeeva_a | 1 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00599949 | Refresh | Closed | nimma_p | September 21, 2018 3:24:01 PM GMT+08:00 | Refresh:-Recon:-User was unable to perform the reconstruction on Dual surview. | WIP9.1 | WIP14 |  | Subsystem | Dual Surview - Reconstruction | NA | Fixed | v_sumanth | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SegReq | Ovation |  | WIP14 | February 27, 2019 7:57:46 PM GMT+08:00
CTUDT00599949 | Refresh | Closed | nimma_p | September 21, 2018 3:24:01 PM GMT+08:00 | Refresh:-Recon:-User was unable to perform the reconstruction on Dual surview. | WIP9.1 | WIP14 |  | Subsystem | Dual Surview - Reconstruction | NA | Fixed | v_sumanth | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SegReq | Brilliance 64 |  | WIP14 | February 27, 2019 7:57:46 PM GMT+08:00
CTUDT00599949 | Refresh | Closed | nimma_p | September 21, 2018 3:24:01 PM GMT+08:00 | Refresh:-Recon:-User was unable to perform the reconstruction on Dual surview. | WIP9.1 | WIP14 |  | Subsystem | Dual Surview - Reconstruction | NA | Fixed | v_sumanth | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SegReq | Ingenuity 64 |  | WIP14 | February 27, 2019 7:57:46 PM GMT+08:00
CTUDT00599949 | Refresh | Closed | nimma_p | September 21, 2018 3:24:01 PM GMT+08:00 | Refresh:-Recon:-User was unable to perform the reconstruction on Dual surview. | WIP9.1 | WIP14 |  | Subsystem | Dual Surview - Reconstruction | NA | Fixed | v_sumanth | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SegReq | IQon1 |  | WIP14 | February 27, 2019 7:57:46 PM GMT+08:00
CTUDT00599946 | Refresh | Closed | hiremath_b | September 21, 2018 12:34:29 PM GMT+08:00 | Refresh :Inconsistency in display of message “mAs for average {patient=Adult, Child or Infant}” in ECM/plan and also before and after surview. And | WIP9.1 | WIP12 |  | Subsystem | Plan - messages to the user | NA | Fixed | chokkalingam_r | 0 | 4 | Console Software | Console4 | NO | Closed | WIP12SubCom | Ovation |  | WIP12 | November 27, 2018 5:19:31 PM GMT+08:00
CTUDT00599946 | Refresh | Closed | hiremath_b | September 21, 2018 12:34:29 PM GMT+08:00 | Refresh :Inconsistency in display of message “mAs for average {patient=Adult, Child or Infant}” in ECM/plan and also before and after surview. And | WIP9.1 | WIP12 |  | Subsystem | Plan - messages to the user | NA | Fixed | chokkalingam_r | 0 | 4 | Console Software | Console4 | NO | Closed | WIP12SubCom | Brilliance 64 |  | WIP12 | November 27, 2018 5:19:31 PM GMT+08:00
CTUDT00599946 | Refresh | Closed | hiremath_b | September 21, 2018 12:34:29 PM GMT+08:00 | Refresh :Inconsistency in display of message “mAs for average {patient=Adult, Child or Infant}” in ECM/plan and also before and after surview. And | WIP9.1 | WIP12 |  | Subsystem | Plan - messages to the user | NA | Fixed | chokkalingam_r | 0 | 4 | Console Software | Console4 | NO | Closed | WIP12SubCom | Ingenuity 64 |  | WIP12 | November 27, 2018 5:19:31 PM GMT+08:00
CTUDT00599946 | Refresh | Closed | hiremath_b | September 21, 2018 12:34:29 PM GMT+08:00 | Refresh :Inconsistency in display of message “mAs for average {patient=Adult, Child or Infant}” in ECM/plan and also before and after surview. And | WIP9.1 | WIP12 |  | Subsystem | Plan - messages to the user | NA | Fixed | chokkalingam_r | 0 | 4 | Console Software | Console4 | NO | Closed | WIP12SubCom | IQon1 |  | WIP12 | November 27, 2018 5:19:31 PM GMT+08:00
CTUDT00599943 | Refresh | Closed | hiremath_b | September 20, 2018 9:43:13 PM GMT+08:00 | Refresh : Can’t plan all axial results of the same ACQ with same FOV when "Limit Plan Geometry" is enabled. | WIP9.1 | Host SSRS DHF303211 Rev 05 |  | Subsystem | Limit Plan Geometry - same FOV | Console-SRS.15219 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | WIP11SegReq | Ovation |  | Host SSRS DHF303211 Rev 05 | November 23, 2018 1:13:09 PM GMT+08:00
CTUDT00599943 | Refresh | Closed | hiremath_b | September 20, 2018 9:43:13 PM GMT+08:00 | Refresh : Can’t plan all axial results of the same ACQ with same FOV when "Limit Plan Geometry" is enabled. | WIP9.1 | Host SSRS DHF303211 Rev 05 |  | Subsystem | Limit Plan Geometry - same FOV | Console-SRS.15219 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | WIP11SegReq | Brilliance 64 |  | Host SSRS DHF303211 Rev 05 | November 23, 2018 1:13:09 PM GMT+08:00
CTUDT00599943 | Refresh | Closed | hiremath_b | September 20, 2018 9:43:13 PM GMT+08:00 | Refresh : Can’t plan all axial results of the same ACQ with same FOV when "Limit Plan Geometry" is enabled. | WIP9.1 | Host SSRS DHF303211 Rev 05 |  | Subsystem | Limit Plan Geometry - same FOV | Console-SRS.15219 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | WIP11SegReq | Brilliance Big Bore |  | Host SSRS DHF303211 Rev 05 | November 23, 2018 1:13:09 PM GMT+08:00
CTUDT00599943 | Refresh | Closed | hiremath_b | September 20, 2018 9:43:13 PM GMT+08:00 | Refresh : Can’t plan all axial results of the same ACQ with same FOV when "Limit Plan Geometry" is enabled. | WIP9.1 | Host SSRS DHF303211 Rev 05 |  | Subsystem | Limit Plan Geometry - same FOV | Console-SRS.15219 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | WIP11SegReq | Ingenuity 64 |  | Host SSRS DHF303211 Rev 05 | November 23, 2018 1:13:09 PM GMT+08:00
CTUDT00599943 | Refresh | Closed | hiremath_b | September 20, 2018 9:43:13 PM GMT+08:00 | Refresh : Can’t plan all axial results of the same ACQ with same FOV when "Limit Plan Geometry" is enabled. | WIP9.1 | Host SSRS DHF303211 Rev 05 |  | Subsystem | Limit Plan Geometry - same FOV | Console-SRS.15219 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | WIP11SegReq | IQon1 |  | Host SSRS DHF303211 Rev 05 | November 23, 2018 1:13:09 PM GMT+08:00
CTUDT00599938 | Refresh | Closed | krisher_i | September 20, 2018 2:12:42 PM GMT+08:00 | Refresh-Images in the View2 and the images in the pictorial view disappear in offline recon | WIP9.1 |  |  | Subsystem | View2 | NA | No Plan to Fix | v_sumanth | 0 | 4 |  | Console4 | NO | Closed |  | IQon1
CTUDT00599931 | Refresh | Closed | levy_s | September 17, 2018 9:29:34 PM GMT+08:00 | Refresh - Gantry cannot initialize Error appears during performance 100 patient per day test | WIP9.1 |  |  | Subsystem | Gantry HW |  | Duplicate | wang_c | 0 |  |  | GCX | NO | Closed |  | IQon1
CTUDT00599929 | Refresh | Closed | zhang_t | September 17, 2018 6:50:34 PM GMT+08:00 | When login with philips_service, input the wrong IST dongle password, it logins successful | WIP9.1 |  |  | Subsystem | IST-User Login - IST component | N/A | Works As Specified / Not A Defect | bandil_j | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00599925 | Refresh | Closed | g_di | September 17, 2018 6:23:04 PM GMT+08:00 | Center(C) value is not resetting to maximum value upon entering higher value than maximum in the Windowing field for any result. | WIP8.1 | DHF303211 Rev07 |  | Subsystem | Window & Center - max values;Verified1 | Console. SRS.5164 | Fixed | aswani_v | 0 | 2 | Requirements Documentation | iCAP | NO | Closed | WIP13SubCom | Ovation |  | DHF303211 Rev07 | June 13, 2019 1:02:28 PM GMT+08:00
CTUDT00599925 | Refresh | Closed | g_di | September 17, 2018 6:23:04 PM GMT+08:00 | Center(C) value is not resetting to maximum value upon entering higher value than maximum in the Windowing field for any result. | WIP8.1 | DHF303211 Rev07 |  | Subsystem | Window & Center - max values;Verified1 | Console. SRS.5164 | Fixed | aswani_v | 0 | 2 | Requirements Documentation | iCAP | NO | Closed | WIP13SubCom | Brilliance 64 |  | DHF303211 Rev07 | June 13, 2019 1:02:28 PM GMT+08:00
CTUDT00599925 | Refresh | Closed | g_di | September 17, 2018 6:23:04 PM GMT+08:00 | Center(C) value is not resetting to maximum value upon entering higher value than maximum in the Windowing field for any result. | WIP8.1 | DHF303211 Rev07 |  | Subsystem | Window & Center - max values;Verified1 | Console. SRS.5164 | Fixed | aswani_v | 0 | 2 | Requirements Documentation | iCAP | NO | Closed | WIP13SubCom | Ingenuity 64 |  | DHF303211 Rev07 | June 13, 2019 1:02:28 PM GMT+08:00
CTUDT00599925 | Refresh | Closed | g_di | September 17, 2018 6:23:04 PM GMT+08:00 | Center(C) value is not resetting to maximum value upon entering higher value than maximum in the Windowing field for any result. | WIP8.1 | DHF303211 Rev07 |  | Subsystem | Window & Center - max values;Verified1 | Console. SRS.5164 | Fixed | aswani_v | 0 | 2 | Requirements Documentation | iCAP | NO | Closed | WIP13SubCom | IQon1 |  | DHF303211 Rev07 | June 13, 2019 1:02:28 PM GMT+08:00
CTUDT00599917 | Refresh | Closed | fitusi_r | September 16, 2018 9:02:02 PM GMT+08:00 | Refresh, IQon -Creating a new report causes the PD to not respond and get stuck. | WIP9.1 |  |  | Subsystem | UI Performance - report creation |  | Works As Specified / Not A Defect | redrouthu_r | 1 |  |  | iCAP | NO | Closed |  | IQon1
CTUDT00599915 | Refresh | Closed | bagdady_b | September 16, 2018 8:20:31 PM GMT+08:00 | Refresh- Monitors are not properly configured after fresh Host installation | WIP9.1 |  |  | Subsystem | Installation - Monitor configuration | NA | Works As Specified / Not A Defect | goswami_s | 0 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00599915 | Refresh | Closed | bagdady_b | September 16, 2018 8:20:31 PM GMT+08:00 | Refresh- Monitors are not properly configured after fresh Host installation | WIP9.1 |  |  | Subsystem | Installation - Monitor configuration | NA | Works As Specified / Not A Defect | goswami_s | 0 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00599909 | Refresh | Closed | attali_s | September 13, 2018 10:13:08 PM GMT+08:00 | Calibrations: For Body Linearity Image Quality cases in Performance, ROIs are measured on 8 slices instead of on the 2 middle slices only | WIP9.1 | WIP13 | N/A | Subsystem | Calibrations - Performance | N/A | Fixed | elviral | 0 | 4 | DMS Software | DMS | NO | Closed | WIP13SegReq | Ingenuity 64 |  | WIP13 | January 6, 2019 4:50:46 PM GMT+08:00
CTUDT00599906 | Refresh | Closed | attali_s | September 13, 2018 9:41:05 PM GMT+08:00 | Calibrations - After a gantry failure during an IQ Application, the following error appears after each case: "Object reference not set to an instance of an object." | WIP9.1 | WIP14 |  | Subsystem | IQ Application | N/A | Works As Specified / Not A Defect | yatziv_l | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP15SegReq | Ovation |  | WIP14
CTUDT00599906 | Refresh | Closed | attali_s | September 13, 2018 9:41:05 PM GMT+08:00 | Calibrations - After a gantry failure during an IQ Application, the following error appears after each case: "Object reference not set to an instance of an object." | WIP9.1 | WIP14 |  | Subsystem | IQ Application | N/A | Works As Specified / Not A Defect | yatziv_l | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP15SegReq | Ingenuity 64 |  | WIP14
CTUDT00599906 | Refresh | Closed | attali_s | September 13, 2018 9:41:05 PM GMT+08:00 | Calibrations - After a gantry failure during an IQ Application, the following error appears after each case: "Object reference not set to an instance of an object." | WIP9.1 | WIP14 |  | Subsystem | IQ Application | N/A | Works As Specified / Not A Defect | yatziv_l | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP15SegReq | IQon1 |  | WIP14
CTUDT00599903 | Refresh | Closed | bagdady_b | September 13, 2018 8:25:57 PM GMT+08:00 | Refresh- Fresh 1click installation doesn't show the completion of 100% | WIP9.1 |  |  | Subsystem | 1-Click Installation |  | Works As Specified / Not A Defect | sanjeeva_a | 0 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00599903 | Refresh | Closed | bagdady_b | September 13, 2018 8:25:57 PM GMT+08:00 | Refresh- Fresh 1click installation doesn't show the completion of 100% | WIP9.1 |  |  | Subsystem | 1-Click Installation |  | Works As Specified / Not A Defect | sanjeeva_a | 0 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00599901 | Refresh | Closed | bagdady_b | September 13, 2018 8:14:59 PM GMT+08:00 | Refresh- Collimations are missing in QA exam cards | WIP9.1 | WIP14 |  | Subsystem | QA Ecs - paramertes | N/A | Fixed | kr_m | 0 | 2 | Console Software | Console4 | YES | Closed | WIP14SubCom | Ovation |  | WIP14 | March 7, 2019 5:01:31 PM GMT+08:00
CTUDT00599899 | Refresh | Closed | bagdady_b | September 13, 2018 8:10:11 PM GMT+08:00 | Refresh- Scan is stuck in initializing when firewall is enabled | WIP9.1 | WIP11 | Yes. User needs to manually disable the firewall. | Subsystem |  | NA | Fixed | kapadia_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP11SegReq | Ovation |  | WIP11 | November 5, 2018 12:40:43 PM GMT+08:00
CTUDT00599899 | Refresh | Closed | bagdady_b | September 13, 2018 8:10:11 PM GMT+08:00 | Refresh- Scan is stuck in initializing when firewall is enabled | WIP9.1 | WIP11 | Yes. User needs to manually disable the firewall. | Subsystem |  | NA | Fixed | kapadia_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP11SegReq | Brilliance 64 |  | WIP11 | November 5, 2018 12:40:43 PM GMT+08:00
CTUDT00599899 | Refresh | Closed | bagdady_b | September 13, 2018 8:10:11 PM GMT+08:00 | Refresh- Scan is stuck in initializing when firewall is enabled | WIP9.1 | WIP11 | Yes. User needs to manually disable the firewall. | Subsystem |  | NA | Fixed | kapadia_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP11SegReq | Ingenuity 64 |  | WIP11 | November 5, 2018 12:40:43 PM GMT+08:00
CTUDT00599899 | Refresh | Closed | bagdady_b | September 13, 2018 8:10:11 PM GMT+08:00 | Refresh- Scan is stuck in initializing when firewall is enabled | WIP9.1 | WIP11 | Yes. User needs to manually disable the firewall. | Subsystem |  | NA | Fixed | kapadia_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP11SegReq | IQon1 |  | WIP11 | November 5, 2018 12:40:43 PM GMT+08:00
CTUDT00599897 | Refresh | Closed | attali_s | September 13, 2018 7:29:31 PM GMT+08:00 | Calibrations: iDose Level 0 Reduction limits are missing in Performance's Results | WIP9.1 | WIP13 | N/A | Subsystem | Calibrations - Performance | N/A | Fixed | elviral | 0 | 2 | DMS Software | DMS | NO | Closed | WIP13SegReq | Ingenuity 64 |  | WIP13 | January 6, 2019 4:50:27 PM GMT+08:00
CTUDT00599889 | Refresh | Closed | attali_s | September 12, 2018 8:39:09 PM GMT+08:00 | Calibrations: Error occurs during Ultra High scans if phantom calibration was not performed | WIP9.1 | WIP12 | Perform Phantom Calibration application. All needed tables will be recreated and requested scan will be performed properly. | Subsystem | Ultra High Resolution | N/A | Fixed | yosefk | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP12SubCom | Brilliance 64 |  | WIP12 | November 27, 2018 9:27:05 PM GMT+08:00
CTUDT00599889 | Refresh | Closed | attali_s | September 12, 2018 8:39:09 PM GMT+08:00 | Calibrations: Error occurs during Ultra High scans if phantom calibration was not performed | WIP9.1 | WIP12 | Perform Phantom Calibration application. All needed tables will be recreated and requested scan will be performed properly. | Subsystem | Ultra High Resolution | N/A | Fixed | yosefk | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP12SubCom | Ingenuity 64 |  | WIP12 | November 27, 2018 9:27:05 PM GMT+08:00
CTUDT00599885 | Refresh | Closed | ramesh_t | September 12, 2018 1:20:38 PM GMT+08:00 | Refresh:There were two different versions of MDU Editor got installed in 1 click. | WIP9 | WIP14 |  | Subsystem | 1-click installation - multiple MDU Editor version | N/A | Fixed | venugopal_s | 0 | 2 | Console Software | DI | NO | Closed | WIP14SubCom | Ovation |  | WIP14 | February 12, 2019 2:34:21 PM GMT+08:00
CTUDT00599885 | Refresh | Closed | ramesh_t | September 12, 2018 1:20:38 PM GMT+08:00 | Refresh:There were two different versions of MDU Editor got installed in 1 click. | WIP9 | WIP14 |  | Subsystem | 1-click installation - multiple MDU Editor version | N/A | Fixed | venugopal_s | 0 | 2 | Console Software | DI | NO | Closed | WIP14SubCom | Brilliance 64 |  | WIP14 | February 12, 2019 2:34:21 PM GMT+08:00
CTUDT00599885 | Refresh | Closed | ramesh_t | September 12, 2018 1:20:38 PM GMT+08:00 | Refresh:There were two different versions of MDU Editor got installed in 1 click. | WIP9 | WIP14 |  | Subsystem | 1-click installation - multiple MDU Editor version | N/A | Fixed | venugopal_s | 0 | 2 | Console Software | DI | NO | Closed | WIP14SubCom | Ingenuity 64 |  | WIP14 | February 12, 2019 2:34:21 PM GMT+08:00
CTUDT00599885 | Refresh | Closed | ramesh_t | September 12, 2018 1:20:38 PM GMT+08:00 | Refresh:There were two different versions of MDU Editor got installed in 1 click. | WIP9 | WIP14 |  | Subsystem | 1-click installation - multiple MDU Editor version | N/A | Fixed | venugopal_s | 0 | 2 | Console Software | DI | NO | Closed | WIP14SubCom | IQon1 |  | WIP14 | February 12, 2019 2:34:21 PM GMT+08:00
CTUDT00599881 | Refresh | Closed | shekhar_p | September 11, 2018 8:30:16 PM GMT+08:00 | IMR: Result plan box is not aligned when matrix is changed | WIP9.1 | WIP12 | NO | Subsystem | Plan - Plan box for IMR | Console-SRS.8958 | Fixed | goswami_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP12SubCom | Ovation |  | WIP12 | December 5, 2018 5:16:11 PM GMT+08:00
CTUDT00599881 | Refresh | Closed | shekhar_p | September 11, 2018 8:30:16 PM GMT+08:00 | IMR: Result plan box is not aligned when matrix is changed | WIP9.1 | WIP12 | NO | Subsystem | Plan - Plan box for IMR | Console-SRS.8958 | Fixed | goswami_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP12SubCom | Brilliance 64 |  | WIP12 | December 5, 2018 5:16:11 PM GMT+08:00
CTUDT00599881 | Refresh | Closed | shekhar_p | September 11, 2018 8:30:16 PM GMT+08:00 | IMR: Result plan box is not aligned when matrix is changed | WIP9.1 | WIP12 | NO | Subsystem | Plan - Plan box for IMR | Console-SRS.8958 | Fixed | goswami_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP12SubCom | Ingenuity 64 |  | WIP12 | December 5, 2018 5:16:11 PM GMT+08:00
CTUDT00599881 | Refresh | Closed | shekhar_p | September 11, 2018 8:30:16 PM GMT+08:00 | IMR: Result plan box is not aligned when matrix is changed | WIP9.1 | WIP12 | NO | Subsystem | Plan - Plan box for IMR | Console-SRS.8958 | Fixed | goswami_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP12SubCom | IQon1 |  | WIP12 | December 5, 2018 5:16:11 PM GMT+08:00
CTUDT00599879 | Refresh | Closed | sowmya_k | September 11, 2018 4:39:13 PM GMT+08:00 | End Exam takes more than 7 secs. | WIP9 |  | No. User has to wait for few more seconds than expected to get the patient page. | Subsystem | UI Performance - End Exam | NA | Duplicate | ravindra_a | 0 | 3 |  | Console4 | NO | Closed | WIP14SegReq | Ovation
CTUDT00599879 | Refresh | Closed | sowmya_k | September 11, 2018 4:39:13 PM GMT+08:00 | End Exam takes more than 7 secs. | WIP9 |  | No. User has to wait for few more seconds than expected to get the patient page. | Subsystem | UI Performance - End Exam | NA | Duplicate | ravindra_a | 0 | 3 |  | Console4 | NO | Closed | WIP14SegReq | Brilliance 64
CTUDT00599879 | Refresh | Closed | sowmya_k | September 11, 2018 4:39:13 PM GMT+08:00 | End Exam takes more than 7 secs. | WIP9 |  | No. User has to wait for few more seconds than expected to get the patient page. | Subsystem | UI Performance - End Exam | NA | Duplicate | ravindra_a | 0 | 3 |  | Console4 | NO | Closed | WIP14SegReq | Brilliance Big Bore
CTUDT00599879 | Refresh | Closed | sowmya_k | September 11, 2018 4:39:13 PM GMT+08:00 | End Exam takes more than 7 secs. | WIP9 |  | No. User has to wait for few more seconds than expected to get the patient page. | Subsystem | UI Performance - End Exam | NA | Duplicate | ravindra_a | 0 | 3 |  | Console4 | NO | Closed | WIP14SegReq | Ingenuity 64
CTUDT00599879 | Refresh | Closed | sowmya_k | September 11, 2018 4:39:13 PM GMT+08:00 | End Exam takes more than 7 secs. | WIP9 |  | No. User has to wait for few more seconds than expected to get the patient page. | Subsystem | UI Performance - End Exam | NA | Duplicate | ravindra_a | 0 | 3 |  | Console4 | NO | Closed | WIP14SegReq | IQon1
CTUDT00599877 | Refresh | Closed | sowmya_k | September 11, 2018 4:24:04 PM GMT+08:00 | Loading exam card takes more time | WIP9 | WIP19 | No workaround | Subsystem | UI Performance - Loading Ecs | Console-SRS.10087 | Fixed | ravindra_a | 0 | 4 | Console Software | Console4 | NO | Closed | WIP18SubCom | Ovation |  | WIP19 | August 5, 2019 2:12:25 PM GMT+08:00
CTUDT00599877 | Refresh | Closed | sowmya_k | September 11, 2018 4:24:04 PM GMT+08:00 | Loading exam card takes more time | WIP9 | WIP19 | No workaround | Subsystem | UI Performance - Loading Ecs | Console-SRS.10087 | Fixed | ravindra_a | 0 | 4 | Console Software | Console4 | NO | Closed | WIP18SubCom | Brilliance 64 |  | WIP19 | August 5, 2019 2:12:25 PM GMT+08:00
CTUDT00599877 | Refresh | Closed | sowmya_k | September 11, 2018 4:24:04 PM GMT+08:00 | Loading exam card takes more time | WIP9 | WIP19 | No workaround | Subsystem | UI Performance - Loading Ecs | Console-SRS.10087 | Fixed | ravindra_a | 0 | 4 | Console Software | Console4 | NO | Closed | WIP18SubCom | Ingenuity 64 |  | WIP19 | August 5, 2019 2:12:25 PM GMT+08:00
CTUDT00599877 | Refresh | Closed | sowmya_k | September 11, 2018 4:24:04 PM GMT+08:00 | Loading exam card takes more time | WIP9 | WIP19 | No workaround | Subsystem | UI Performance - Loading Ecs | Console-SRS.10087 | Fixed | ravindra_a | 0 | 4 | Console Software | Console4 | NO | Closed | WIP18SubCom | IQon1 |  | WIP19 | August 5, 2019 2:12:25 PM GMT+08:00
CTUDT00599875 | Refresh | Closed | shekhar_p | September 11, 2018 3:22:02 PM GMT+08:00 | Security: Windows Security center popup is appearing | WIP9.1 | WIP18 BB_WIP8 |  | Subsystem | OS messages | None | Fixed | dasgupta_s | 0 | 3 | Console Software | Support-Security | NO | Closed | WIP14SubCom | Ovation |  | WIP18 BB_WIP8 | June 14, 2019 8:35:47 PM GMT+08:00
CTUDT00599875 | Refresh | Closed | shekhar_p | September 11, 2018 3:22:02 PM GMT+08:00 | Security: Windows Security center popup is appearing | WIP9.1 | WIP18 BB_WIP8 |  | Subsystem | OS messages | None | Fixed | dasgupta_s | 0 | 3 | Console Software | Support-Security | NO | Closed | WIP14SubCom | Brilliance Big Bore |  | WIP18 BB_WIP8 | June 14, 2019 8:35:47 PM GMT+08:00
CTUDT00599875 | Refresh | Closed | shekhar_p | September 11, 2018 3:22:02 PM GMT+08:00 | Security: Windows Security center popup is appearing | WIP9.1 | WIP18 BB_WIP8 |  | Subsystem | OS messages | None | Fixed | dasgupta_s | 0 | 3 | Console Software | Support-Security | NO | Closed | WIP14SubCom | Ingenuity 64 |  | WIP18 BB_WIP8 | June 14, 2019 8:35:47 PM GMT+08:00
CTUDT00599875 | Refresh | Closed | shekhar_p | September 11, 2018 3:22:02 PM GMT+08:00 | Security: Windows Security center popup is appearing | WIP9.1 | WIP18 BB_WIP8 |  | Subsystem | OS messages | None | Fixed | dasgupta_s | 0 | 3 | Console Software | Support-Security | NO | Closed | WIP14SubCom | IQon1 |  | WIP18 BB_WIP8 | June 14, 2019 8:35:47 PM GMT+08:00
CTUDT00599868 | Refresh | Closed | nimma_p | September 10, 2018 5:15:42 PM GMT+08:00 | Refresh:-Cardiac:-In Preference,“Cancel Cardiac Dose Right henceforth” option not available. | WIP9.1 | WIP11 |  | Subsystem |  | NA | Fixed | kumar_v | 0 | 2 | Console Software | Console4 | YES | Closed | WIP11SegReq | Ovation |  | WIP11 | October 18, 2018 12:33:03 PM GMT+08:00
CTUDT00599868 | Refresh | Closed | nimma_p | September 10, 2018 5:15:42 PM GMT+08:00 | Refresh:-Cardiac:-In Preference,“Cancel Cardiac Dose Right henceforth” option not available. | WIP9.1 | WIP11 |  | Subsystem |  | NA | Fixed | kumar_v | 0 | 2 | Console Software | Console4 | YES | Closed | WIP11SegReq | Brilliance 64 |  | WIP11 | October 18, 2018 12:33:03 PM GMT+08:00
CTUDT00599868 | Refresh | Closed | nimma_p | September 10, 2018 5:15:42 PM GMT+08:00 | Refresh:-Cardiac:-In Preference,“Cancel Cardiac Dose Right henceforth” option not available. | WIP9.1 | WIP11 |  | Subsystem |  | NA | Fixed | kumar_v | 0 | 2 | Console Software | Console4 | YES | Closed | WIP11SegReq | Ingenuity 64 |  | WIP11 | October 18, 2018 12:33:03 PM GMT+08:00
CTUDT00599868 | Refresh | Closed | nimma_p | September 10, 2018 5:15:42 PM GMT+08:00 | Refresh:-Cardiac:-In Preference,“Cancel Cardiac Dose Right henceforth” option not available. | WIP9.1 | WIP11 |  | Subsystem |  | NA | Fixed | kumar_v | 0 | 2 | Console Software | Console4 | YES | Closed | WIP11SegReq | IQon1 |  | WIP11 | October 18, 2018 12:33:03 PM GMT+08:00
CTUDT00599866 | Refresh | Closed | navale_s | September 7, 2018 3:38:11 PM GMT+08:00 | during a tilted acquisition, the plan box moves up when changing the scan FOV | 4.8 | BB_WIP5.1 |  | Subsystem | Plan - plan box moves | N/A | Fixed | bhawna_g | 0 | 2 | Couch Software | Console4 | YES | Closed | BB_WIP1.1_SubCom | Brilliance Big Bore |  | BB_WIP5.1 | April 2, 2019 6:35:51 PM GMT+08:00
CTUDT00599850 | Refresh | Closed | bommiReddy_r | September 6, 2018 7:58:35 PM GMT+08:00 | Refresh:Worklist-In Local Schedule DB “Accession Number” column is not displayed by default | WIP9.1 | WIP12 | the Accession Number column can be added manually | System | Woklist - Accession Number | NA | Fixed | chakraborty_d | 0 | 2 | Console Software | Console4 | NO | Closed | WIP12SubCom | Ovation |  | WIP12 | December 4, 2018 6:27:04 PM GMT+08:00
CTUDT00599850 | Refresh | Closed | bommiReddy_r | September 6, 2018 7:58:35 PM GMT+08:00 | Refresh:Worklist-In Local Schedule DB “Accession Number” column is not displayed by default | WIP9.1 | WIP12 | the Accession Number column can be added manually | System | Woklist - Accession Number | NA | Fixed | chakraborty_d | 0 | 2 | Console Software | Console4 | NO | Closed | WIP12SubCom | Brilliance 64 |  | WIP12 | December 4, 2018 6:27:04 PM GMT+08:00
CTUDT00599850 | Refresh | Closed | bommiReddy_r | September 6, 2018 7:58:35 PM GMT+08:00 | Refresh:Worklist-In Local Schedule DB “Accession Number” column is not displayed by default | WIP9.1 | WIP12 | the Accession Number column can be added manually | System | Woklist - Accession Number | NA | Fixed | chakraborty_d | 0 | 2 | Console Software | Console4 | NO | Closed | WIP12SubCom | Ingenuity 64 |  | WIP12 | December 4, 2018 6:27:04 PM GMT+08:00
CTUDT00599850 | Refresh | Closed | bommiReddy_r | September 6, 2018 7:58:35 PM GMT+08:00 | Refresh:Worklist-In Local Schedule DB “Accession Number” column is not displayed by default | WIP9.1 | WIP12 | the Accession Number column can be added manually | System | Woklist - Accession Number | NA | Fixed | chakraborty_d | 0 | 2 | Console Software | Console4 | NO | Closed | WIP12SubCom | IQon1 |  | WIP12 | December 4, 2018 6:27:04 PM GMT+08:00
CTUDT00599848 | Refresh | Closed | pB_s | September 6, 2018 7:03:20 PM GMT+08:00 | [Refresh - Sequential Dual Energy ] : For Dual energy exam cards , in the context menu "Add Injection" option is not displayed for 1st acquisition | WIP9.1 | WIP14 | User will still be able to add Injection from Show all | Subsystem | EC parameters - Dual Energy | NA | Fixed | prasad_r | 0 | 3 | Console Software | Console4 | NO | Closed | WIP14SubCom | Ovation |  | WIP14 | February 12, 2019 5:39:34 PM GMT+08:00
CTUDT00599848 | Refresh | Closed | pB_s | September 6, 2018 7:03:20 PM GMT+08:00 | [Refresh - Sequential Dual Energy ] : For Dual energy exam cards , in the context menu "Add Injection" option is not displayed for 1st acquisition | WIP9.1 | WIP14 | User will still be able to add Injection from Show all | Subsystem | EC parameters - Dual Energy | NA | Fixed | prasad_r | 0 | 3 | Console Software | Console4 | NO | Closed | WIP14SubCom | Brilliance 64 |  | WIP14 | February 12, 2019 5:39:34 PM GMT+08:00
CTUDT00599848 | Refresh | Closed | pB_s | September 6, 2018 7:03:20 PM GMT+08:00 | [Refresh - Sequential Dual Energy ] : For Dual energy exam cards , in the context menu "Add Injection" option is not displayed for 1st acquisition | WIP9.1 | WIP14 | User will still be able to add Injection from Show all | Subsystem | EC parameters - Dual Energy | NA | Fixed | prasad_r | 0 | 3 | Console Software | Console4 | NO | Closed | WIP14SubCom | Ingenuity 64 |  | WIP14 | February 12, 2019 5:39:34 PM GMT+08:00
CTUDT00599846 | Refresh | Closed | manjunath_c | September 6, 2018 6:28:20 PM GMT+08:00 | ECM crashing while importing ECs | WIP9.1 | WIP11.1 |  | Subsystem |  | NA | Fixed | kumar_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP11SegReq | Ovation |  | WIP11.1 | October 30, 2018 7:14:21 PM GMT+08:00
CTUDT00599846 | Refresh | Closed | manjunath_c | September 6, 2018 6:28:20 PM GMT+08:00 | ECM crashing while importing ECs | WIP9.1 | WIP11.1 |  | Subsystem |  | NA | Fixed | kumar_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP11SegReq | Brilliance 64 |  | WIP11.1 | October 30, 2018 7:14:21 PM GMT+08:00
CTUDT00599846 | Refresh | Closed | manjunath_c | September 6, 2018 6:28:20 PM GMT+08:00 | ECM crashing while importing ECs | WIP9.1 | WIP11.1 |  | Subsystem |  | NA | Fixed | kumar_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP11SegReq | Ingenuity 64 |  | WIP11.1 | October 30, 2018 7:14:21 PM GMT+08:00
CTUDT00599846 | Refresh | Closed | manjunath_c | September 6, 2018 6:28:20 PM GMT+08:00 | ECM crashing while importing ECs | WIP9.1 | WIP11.1 |  | Subsystem |  | NA | Fixed | kumar_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP11SegReq | IQon1 |  | WIP11.1 | October 30, 2018 7:14:21 PM GMT+08:00
CTUDT00599842 | Refresh | Closed | manjunath_c | September 6, 2018 5:55:03 PM GMT+08:00 | ELSCINTCT_IDOSE_LEVEL tag missing from DICOM INFO | WIP9.1 | WIP11.1 |  | Subsystem |  | NA | Fixed | kumar_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP11SegReq | Ovation |  | WIP11.1 | October 30, 2018 7:19:23 PM GMT+08:00
CTUDT00599842 | Refresh | Closed | manjunath_c | September 6, 2018 5:55:03 PM GMT+08:00 | ELSCINTCT_IDOSE_LEVEL tag missing from DICOM INFO | WIP9.1 | WIP11.1 |  | Subsystem |  | NA | Fixed | kumar_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP11SegReq | Brilliance 64 |  | WIP11.1 | October 30, 2018 7:19:23 PM GMT+08:00
CTUDT00599842 | Refresh | Closed | manjunath_c | September 6, 2018 5:55:03 PM GMT+08:00 | ELSCINTCT_IDOSE_LEVEL tag missing from DICOM INFO | WIP9.1 | WIP11.1 |  | Subsystem |  | NA | Fixed | kumar_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP11SegReq | Ingenuity 64 |  | WIP11.1 | October 30, 2018 7:19:23 PM GMT+08:00
CTUDT00599842 | Refresh | Closed | manjunath_c | September 6, 2018 5:55:03 PM GMT+08:00 | ELSCINTCT_IDOSE_LEVEL tag missing from DICOM INFO | WIP9.1 | WIP11.1 |  | Subsystem |  | NA | Fixed | kumar_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP11SegReq | IQon1 |  | WIP11.1 | October 30, 2018 7:19:23 PM GMT+08:00
CTUDT00599840 | Refresh | Closed | ramesh_t | September 6, 2018 5:33:33 PM GMT+08:00 | Refresh:ExamCards: FOV field is not available in result show all after the surview is scanned | WIP9.1 | WIP12 | The user can change the FOV using view port controller | Subsystem | EC parameters - FOV | N/A | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | NO | Closed | WIP12SubCom | Ovation |  | WIP12 | November 28, 2018 4:30:51 PM GMT+08:00
CTUDT00599840 | Refresh | Closed | ramesh_t | September 6, 2018 5:33:33 PM GMT+08:00 | Refresh:ExamCards: FOV field is not available in result show all after the surview is scanned | WIP9.1 | WIP12 | The user can change the FOV using view port controller | Subsystem | EC parameters - FOV | N/A | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | NO | Closed | WIP12SubCom | Brilliance 64 |  | WIP12 | November 28, 2018 4:30:51 PM GMT+08:00
CTUDT00599840 | Refresh | Closed | ramesh_t | September 6, 2018 5:33:33 PM GMT+08:00 | Refresh:ExamCards: FOV field is not available in result show all after the surview is scanned | WIP9.1 | WIP12 | The user can change the FOV using view port controller | Subsystem | EC parameters - FOV | N/A | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | NO | Closed | WIP12SubCom | Ingenuity 64 |  | WIP12 | November 28, 2018 4:30:51 PM GMT+08:00
CTUDT00599840 | Refresh | Closed | ramesh_t | September 6, 2018 5:33:33 PM GMT+08:00 | Refresh:ExamCards: FOV field is not available in result show all after the surview is scanned | WIP9.1 | WIP12 | The user can change the FOV using view port controller | Subsystem | EC parameters - FOV | N/A | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | NO | Closed | WIP12SubCom | IQon1 |  | WIP12 | November 28, 2018 4:30:51 PM GMT+08:00
CTUDT00599837 | Refresh | Closed | nimma_p | September 6, 2018 12:29:48 PM GMT+08:00 | INGENUITY: Recon: In Preference, some of tabs have spectral related information in Ingenuity model. | WIP9.1 | WIP13 |  | Subsystem | Preferences - wrong tables for Ing | N/A | Fixed | pandey_a | 0 | 4 | Console Software | Console4 | NO | Closed | WIP13SubCom | Ovation |  | WIP13 | January 11, 2019 7:41:00 PM GMT+08:00
CTUDT00599837 | Refresh | Closed | nimma_p | September 6, 2018 12:29:48 PM GMT+08:00 | INGENUITY: Recon: In Preference, some of tabs have spectral related information in Ingenuity model. | WIP9.1 | WIP13 |  | Subsystem | Preferences - wrong tables for Ing | N/A | Fixed | pandey_a | 0 | 4 | Console Software | Console4 | NO | Closed | WIP13SubCom | Ingenuity 64 |  | WIP13 | January 11, 2019 7:41:00 PM GMT+08:00
CTUDT00599833 | Refresh | Closed | cohen_y | September 6, 2018 12:44:45 AM GMT+08:00 | Refresh: Calibrations: Beam and DMS check in CT user detects contrast spill even though in the standalone application it doesn’t. | WIP9.1 | WIP17 |  | Subsystem |  | N/A | Fixed | Kostya_Baklan | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP11SegReq | Ovation |  | WIP17 | May 13, 2019 3:27:18 PM GMT+08:00
CTUDT00599833 | Refresh | Closed | cohen_y | September 6, 2018 12:44:45 AM GMT+08:00 | Refresh: Calibrations: Beam and DMS check in CT user detects contrast spill even though in the standalone application it doesn’t. | WIP9.1 | WIP17 |  | Subsystem |  | N/A | Fixed | Kostya_Baklan | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP11SegReq | Brilliance 64 |  | WIP17 | May 13, 2019 3:27:18 PM GMT+08:00
CTUDT00599833 | Refresh | Closed | cohen_y | September 6, 2018 12:44:45 AM GMT+08:00 | Refresh: Calibrations: Beam and DMS check in CT user detects contrast spill even though in the standalone application it doesn’t. | WIP9.1 | WIP17 |  | Subsystem |  | N/A | Fixed | Kostya_Baklan | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP11SegReq | Brilliance Big Bore |  | WIP17 | May 13, 2019 3:27:18 PM GMT+08:00
CTUDT00599833 | Refresh | Closed | cohen_y | September 6, 2018 12:44:45 AM GMT+08:00 | Refresh: Calibrations: Beam and DMS check in CT user detects contrast spill even though in the standalone application it doesn’t. | WIP9.1 | WIP17 |  | Subsystem |  | N/A | Fixed | Kostya_Baklan | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP11SegReq | Ingenuity 64 |  | WIP17 | May 13, 2019 3:27:18 PM GMT+08:00
CTUDT00599833 | Refresh | Closed | cohen_y | September 6, 2018 12:44:45 AM GMT+08:00 | Refresh: Calibrations: Beam and DMS check in CT user detects contrast spill even though in the standalone application it doesn’t. | WIP9.1 | WIP17 |  | Subsystem |  | N/A | Fixed | Kostya_Baklan | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP11SegReq | IQon1 |  | WIP17 | May 13, 2019 3:27:18 PM GMT+08:00
CTUDT00599831 | Refresh | Closed | fitusi_r | September 5, 2018 11:25:19 PM GMT+08:00 | Refresh, IQon -The rows in the table in the recon tab are blank. | WIP9.1 |  |  | Subsystem | RQM - rows are blank |  | Duplicate | richards_b | 0 |  |  | CIRS | NO | Closed |  | IQon1
CTUDT00599828 | Refresh | Closed | cohen_y | September 5, 2018 10:48:25 PM GMT+08:00 | Refresh: Service Tools: TDMC Components do not appear in System Configuration. | WIP9.1 | WIP12 |  | Subsystem | System Configuration - TDMC component | N/A | Fixed | franco_a | 0 | 2 | DMS Software | DMS | NO | Closed | WIP12SegReq | Ovation |  | WIP12 | December 10, 2018 2:22:10 PM GMT+08:00
CTUDT00599826 | Refresh | Closed | cohen_y | September 5, 2018 9:54:27 PM GMT+08:00 | Refresh: Default button in Preferences does not affect the Harmonized Phantom option. | WIP9.1 |  |  | Subsystem |  | None | Works As Specified / Not A Defect | sanjeeva_a | 0 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00599826 | Refresh | Closed | cohen_y | September 5, 2018 9:54:27 PM GMT+08:00 | Refresh: Default button in Preferences does not affect the Harmonized Phantom option. | WIP9.1 |  |  | Subsystem |  | None | Works As Specified / Not A Defect | sanjeeva_a | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00599826 | Refresh | Closed | cohen_y | September 5, 2018 9:54:27 PM GMT+08:00 | Refresh: Default button in Preferences does not affect the Harmonized Phantom option. | WIP9.1 |  |  | Subsystem |  | None | Works As Specified / Not A Defect | sanjeeva_a | 0 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00599824 | Refresh | Closed | bagdady_b | September 5, 2018 9:29:40 PM GMT+08:00 | Refresh- Network Devices are not properly configured during 1Click installation | WIP9.1 | WIP11.1 |  | Subsystem |  | NA | Fixed | goswami_s | 0 | 3 | Console Software | Console4 | NO | Closed | WIP11SubCom | Ovation |  | WIP11.1 | November 12, 2018 6:01:48 PM GMT+08:00
CTUDT00599824 | Refresh | Closed | bagdady_b | September 5, 2018 9:29:40 PM GMT+08:00 | Refresh- Network Devices are not properly configured during 1Click installation | WIP9.1 | WIP11.1 |  | Subsystem |  | NA | Fixed | goswami_s | 0 | 3 | Console Software | Console4 | NO | Closed | WIP11SubCom | Brilliance 64 |  | WIP11.1 | November 12, 2018 6:01:48 PM GMT+08:00
CTUDT00599824 | Refresh | Closed | bagdady_b | September 5, 2018 9:29:40 PM GMT+08:00 | Refresh- Network Devices are not properly configured during 1Click installation | WIP9.1 | WIP11.1 |  | Subsystem |  | NA | Fixed | goswami_s | 0 | 3 | Console Software | Console4 | NO | Closed | WIP11SubCom | Brilliance Big Bore |  | WIP11.1 | November 12, 2018 6:01:48 PM GMT+08:00
CTUDT00599824 | Refresh | Closed | bagdady_b | September 5, 2018 9:29:40 PM GMT+08:00 | Refresh- Network Devices are not properly configured during 1Click installation | WIP9.1 | WIP11.1 |  | Subsystem |  | NA | Fixed | goswami_s | 0 | 3 | Console Software | Console4 | NO | Closed | WIP11SubCom | Ingenuity 64 |  | WIP11.1 | November 12, 2018 6:01:48 PM GMT+08:00
CTUDT00599824 | Refresh | Closed | bagdady_b | September 5, 2018 9:29:40 PM GMT+08:00 | Refresh- Network Devices are not properly configured during 1Click installation | WIP9.1 | WIP11.1 |  | Subsystem |  | NA | Fixed | goswami_s | 0 | 3 | Console Software | Console4 | NO | Closed | WIP11SubCom | IQon1 |  | WIP11.1 | November 12, 2018 6:01:48 PM GMT+08:00
CTUDT00599822 | Refresh | Closed | cohen_y | September 5, 2018 7:58:23 PM GMT+08:00 | Refresh: Workflow crashes during startup in users other than Service user and CT user. | WIP9.1 | WIP12 | Yes. User can start GantryWCFhostService from c:\usr\diamond.root\bin\gantryservers manually before launching the workflow. | Subsystem | User Login | NA | Fixed | halamek | 0 | 2 | Gantry Software | GCX | NO | Closed | WIP11SegReq | Ovation |  | WIP12 | January 14, 2019 8:58:13 PM GMT+08:00
CTUDT00599822 | Refresh | Closed | cohen_y | September 5, 2018 7:58:23 PM GMT+08:00 | Refresh: Workflow crashes during startup in users other than Service user and CT user. | WIP9.1 | WIP12 | Yes. User can start GantryWCFhostService from c:\usr\diamond.root\bin\gantryservers manually before launching the workflow. | Subsystem | User Login | NA | Fixed | halamek | 0 | 2 | Gantry Software | GCX | NO | Closed | WIP11SegReq | Brilliance 64 |  | WIP12 | January 14, 2019 8:58:13 PM GMT+08:00
CTUDT00599822 | Refresh | Closed | cohen_y | September 5, 2018 7:58:23 PM GMT+08:00 | Refresh: Workflow crashes during startup in users other than Service user and CT user. | WIP9.1 | WIP12 | Yes. User can start GantryWCFhostService from c:\usr\diamond.root\bin\gantryservers manually before launching the workflow. | Subsystem | User Login | NA | Fixed | halamek | 0 | 2 | Gantry Software | GCX | NO | Closed | WIP11SegReq | Ingenuity 64 |  | WIP12 | January 14, 2019 8:58:13 PM GMT+08:00
CTUDT00599822 | Refresh | Closed | cohen_y | September 5, 2018 7:58:23 PM GMT+08:00 | Refresh: Workflow crashes during startup in users other than Service user and CT user. | WIP9.1 | WIP12 | Yes. User can start GantryWCFhostService from c:\usr\diamond.root\bin\gantryservers manually before launching the workflow. | Subsystem | User Login | NA | Fixed | halamek | 0 | 2 | Gantry Software | GCX | NO | Closed | WIP11SegReq | IQon1 |  | WIP12 | January 14, 2019 8:58:13 PM GMT+08:00
CTUDT00599819 | Refresh | Closed | beck_d | September 5, 2018 7:31:15 PM GMT+08:00 | Refresh: Calibrations: Beam Quality, Tube Potential Test: Dose meter details in semi-automatic mode are not saved | WIP9.1 | N/A |  | Subsystem | Documentation | N/A | Fixed | braunstein_d | 0 | 2 | Design Documentation | DMS | NO | Closed | WIP14SegReq | Ovation |  | N/A | February 3, 2019 3:57:25 PM GMT+08:00
CTUDT00599819 | Refresh | Closed | beck_d | September 5, 2018 7:31:15 PM GMT+08:00 | Refresh: Calibrations: Beam Quality, Tube Potential Test: Dose meter details in semi-automatic mode are not saved | WIP9.1 | N/A |  | Subsystem | Documentation | N/A | Fixed | braunstein_d | 0 | 2 | Design Documentation | DMS | NO | Closed | WIP14SegReq | Brilliance Big Bore |  | N/A | February 3, 2019 3:57:25 PM GMT+08:00
CTUDT00599819 | Refresh | Closed | beck_d | September 5, 2018 7:31:15 PM GMT+08:00 | Refresh: Calibrations: Beam Quality, Tube Potential Test: Dose meter details in semi-automatic mode are not saved | WIP9.1 | N/A |  | Subsystem | Documentation | N/A | Fixed | braunstein_d | 0 | 2 | Design Documentation | DMS | NO | Closed | WIP14SegReq | Ingenuity 64 |  | N/A | February 3, 2019 3:57:25 PM GMT+08:00
CTUDT00599819 | Refresh | Closed | beck_d | September 5, 2018 7:31:15 PM GMT+08:00 | Refresh: Calibrations: Beam Quality, Tube Potential Test: Dose meter details in semi-automatic mode are not saved | WIP9.1 | N/A |  | Subsystem | Documentation | N/A | Fixed | braunstein_d | 0 | 2 | Design Documentation | DMS | NO | Closed | WIP14SegReq | IQon1 |  | N/A | February 3, 2019 3:57:25 PM GMT+08:00
CTUDT00599813 | Refresh | Closed | singh_p | September 5, 2018 5:36:06 PM GMT+08:00 | Refresh Ingenuity WIP 9.1 : System Error popped up during clinical scan. | WIP9.1 |  |  | Subsystem | Ultra High Resolution |  | Duplicate | frider_o | 0 |  |  | DMS | NO | Closed |  | Ingenuity 64
CTUDT00599810 | Refresh | Closed | bagdady_b | September 5, 2018 3:19:59 PM GMT+08:00 | Refresh- Unable to perform GPC installation on WIP9.1 | WIP9.1 |  |  | Subsystem | GPC Installation | N/A | Does Not Apply | yang_h | 0 |  |  | GCX | NO | Closed |  | Ovation
CTUDT00599808 | Refresh | Closed | beck_d | September 5, 2018 2:23:16 PM GMT+08:00 | Refresh: Calibrations: Beam Quality, Tube Potential Test: Institute address and Reason for test do not appear in the HTML report | WIP9.1 | WIP11 |  | Subsystem |  | NAN | Fixed | daibes_s | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP11SegReq | Ovation |  | WIP11 | October 15, 2018 4:29:11 PM GMT+08:00
CTUDT00599808 | Refresh | Closed | beck_d | September 5, 2018 2:23:16 PM GMT+08:00 | Refresh: Calibrations: Beam Quality, Tube Potential Test: Institute address and Reason for test do not appear in the HTML report | WIP9.1 | WIP11 |  | Subsystem |  | NAN | Fixed | daibes_s | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP11SegReq | Brilliance 64 |  | WIP11 | October 15, 2018 4:29:11 PM GMT+08:00
CTUDT00599808 | Refresh | Closed | beck_d | September 5, 2018 2:23:16 PM GMT+08:00 | Refresh: Calibrations: Beam Quality, Tube Potential Test: Institute address and Reason for test do not appear in the HTML report | WIP9.1 | WIP11 |  | Subsystem |  | NAN | Fixed | daibes_s | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP11SegReq | Brilliance Big Bore |  | WIP11 | October 15, 2018 4:29:11 PM GMT+08:00
CTUDT00599808 | Refresh | Closed | beck_d | September 5, 2018 2:23:16 PM GMT+08:00 | Refresh: Calibrations: Beam Quality, Tube Potential Test: Institute address and Reason for test do not appear in the HTML report | WIP9.1 | WIP11 |  | Subsystem |  | NAN | Fixed | daibes_s | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP11SegReq | Ingenuity 64 |  | WIP11 | October 15, 2018 4:29:11 PM GMT+08:00
CTUDT00599808 | Refresh | Closed | beck_d | September 5, 2018 2:23:16 PM GMT+08:00 | Refresh: Calibrations: Beam Quality, Tube Potential Test: Institute address and Reason for test do not appear in the HTML report | WIP9.1 | WIP11 |  | Subsystem |  | NAN | Fixed | daibes_s | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP11SegReq | IQon1 |  | WIP11 | October 15, 2018 4:29:11 PM GMT+08:00
CTUDT00599806 | Refresh | Closed | braunstein_d | September 5, 2018 1:04:57 PM GMT+08:00 | During Automatic IQ applications ROI is not always saved with the image in Archive. | WIP9.2 |  |  | Subsystem |  | n/a | No Plan to Fix | yatziv_l | 0 | 4 | DMS Software | DMS | NO | Closed | WIP21SegReq | Ovation
CTUDT00599806 | Refresh | Closed | braunstein_d | September 5, 2018 1:04:57 PM GMT+08:00 | During Automatic IQ applications ROI is not always saved with the image in Archive. | WIP9.2 |  |  | Subsystem |  | n/a | No Plan to Fix | yatziv_l | 0 | 4 | DMS Software | DMS | NO | Closed | WIP21SegReq | Brilliance 64
CTUDT00599806 | Refresh | Closed | braunstein_d | September 5, 2018 1:04:57 PM GMT+08:00 | During Automatic IQ applications ROI is not always saved with the image in Archive. | WIP9.2 |  |  | Subsystem |  | n/a | No Plan to Fix | yatziv_l | 0 | 4 | DMS Software | DMS | NO | Closed | WIP21SegReq | Brilliance Big Bore
CTUDT00599806 | Refresh | Closed | braunstein_d | September 5, 2018 1:04:57 PM GMT+08:00 | During Automatic IQ applications ROI is not always saved with the image in Archive. | WIP9.2 |  |  | Subsystem |  | n/a | No Plan to Fix | yatziv_l | 0 | 4 | DMS Software | DMS | NO | Closed | WIP21SegReq | Ingenuity 64
CTUDT00599806 | Refresh | Closed | braunstein_d | September 5, 2018 1:04:57 PM GMT+08:00 | During Automatic IQ applications ROI is not always saved with the image in Archive. | WIP9.2 |  |  | Subsystem |  | n/a | No Plan to Fix | yatziv_l | 0 | 4 | DMS Software | DMS | NO | Closed | WIP21SegReq | IQon1
CTUDT00599804 | Refresh | Closed | braunstein_d | September 5, 2018 12:53:29 PM GMT+08:00 | CTDI free air application Calculating the “Deviation from nominal (%)” | WIP9.2 | DHF170399, Rev 07 |  | Subsystem | Documentation | DHF170399 | Fixed | morton_t | 0 | 4 | Design Documentation | Support-Physics | NO | Closed | WIP15SegReq | Ovation |  | DHF170399, Rev 07 | May 7, 2019 8:12:22 PM GMT+08:00
CTUDT00599804 | Refresh | Closed | braunstein_d | September 5, 2018 12:53:29 PM GMT+08:00 | CTDI free air application Calculating the “Deviation from nominal (%)” | WIP9.2 | DHF170399, Rev 07 |  | Subsystem | Documentation | DHF170399 | Fixed | morton_t | 0 | 4 | Design Documentation | Support-Physics | NO | Closed | WIP15SegReq | Brilliance 64 |  | DHF170399, Rev 07 | May 7, 2019 8:12:22 PM GMT+08:00
CTUDT00599804 | Refresh | Closed | braunstein_d | September 5, 2018 12:53:29 PM GMT+08:00 | CTDI free air application Calculating the “Deviation from nominal (%)” | WIP9.2 | DHF170399, Rev 07 |  | Subsystem | Documentation | DHF170399 | Fixed | morton_t | 0 | 4 | Design Documentation | Support-Physics | NO | Closed | WIP15SegReq | Brilliance Big Bore |  | DHF170399, Rev 07 | May 7, 2019 8:12:22 PM GMT+08:00
CTUDT00599804 | Refresh | Closed | braunstein_d | September 5, 2018 12:53:29 PM GMT+08:00 | CTDI free air application Calculating the “Deviation from nominal (%)” | WIP9.2 | DHF170399, Rev 07 |  | Subsystem | Documentation | DHF170399 | Fixed | morton_t | 0 | 4 | Design Documentation | Support-Physics | NO | Closed | WIP15SegReq | Ingenuity 64 |  | DHF170399, Rev 07 | May 7, 2019 8:12:22 PM GMT+08:00
CTUDT00599804 | Refresh | Closed | braunstein_d | September 5, 2018 12:53:29 PM GMT+08:00 | CTDI free air application Calculating the “Deviation from nominal (%)” | WIP9.2 | DHF170399, Rev 07 |  | Subsystem | Documentation | DHF170399 | Fixed | morton_t | 0 | 4 | Design Documentation | Support-Physics | NO | Closed | WIP15SegReq | IQon1 |  | DHF170399, Rev 07 | May 7, 2019 8:12:22 PM GMT+08:00
CTUDT00599802 | Refresh | Closed | wang_l | September 5, 2018 10:55:48 AM GMT+08:00 | Scan with system error for IAC exam cards | WIP9.1 |  |  | Subsystem | IAC/Ear | n/a | Duplicate | frider_o | 0 |  |  | DMS | YES | Closed |  | Ingenuity 64
CTUDT00599797 | Refresh | Closed | azagury_d | September 4, 2018 9:41:03 PM GMT+08:00 | Refresh: Calibrations: Beam Quality, Tube potential: Deviation from Nominal value is not presented in R&D when performed in mfg user. | WIP9.1 | WIP12 |  | Subsystem |  | NAN | Fixed | daibes_s | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP11.1SegReq | Ovation |  | WIP12 | November 12, 2018 9:54:56 PM GMT+08:00
CTUDT00599797 | Refresh | Closed | azagury_d | September 4, 2018 9:41:03 PM GMT+08:00 | Refresh: Calibrations: Beam Quality, Tube potential: Deviation from Nominal value is not presented in R&D when performed in mfg user. | WIP9.1 | WIP12 |  | Subsystem |  | NAN | Fixed | daibes_s | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP11.1SegReq | Brilliance Big Bore |  | WIP12 | November 12, 2018 9:54:56 PM GMT+08:00
CTUDT00599797 | Refresh | Closed | azagury_d | September 4, 2018 9:41:03 PM GMT+08:00 | Refresh: Calibrations: Beam Quality, Tube potential: Deviation from Nominal value is not presented in R&D when performed in mfg user. | WIP9.1 | WIP12 |  | Subsystem |  | NAN | Fixed | daibes_s | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP11.1SegReq | Ingenuity 64 |  | WIP12 | November 12, 2018 9:54:56 PM GMT+08:00
CTUDT00599797 | Refresh | Closed | azagury_d | September 4, 2018 9:41:03 PM GMT+08:00 | Refresh: Calibrations: Beam Quality, Tube potential: Deviation from Nominal value is not presented in R&D when performed in mfg user. | WIP9.1 | WIP12 |  | Subsystem |  | NAN | Fixed | daibes_s | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP11.1SegReq | IQon1 |  | WIP12 | November 12, 2018 9:54:56 PM GMT+08:00
CTUDT00599796 | Refresh | Closed | shekhar_p | September 4, 2018 9:40:05 PM GMT+08:00 | Refresh: Logout: System crashing during Logout process. | WIP9.1 |  |  | Subsystem | User Login - crash |  | Works As Specified / Not A Defect | sanjeeva_a | 1 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00599793 | Refresh | Closed | azagury_d | September 4, 2018 8:41:26 PM GMT+08:00 | Refresh: Calibrations: Beam Quality, Tube potential test: HVL measurements are Automatically entered when Manual mode is selected and the base is connected | WIP9.1 | WIP11 |  | Subsystem |  | NA | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP11SegReq | Ovation |  | WIP11 | October 15, 2018 7:07:58 PM GMT+08:00
CTUDT00599793 | Refresh | Closed | azagury_d | September 4, 2018 8:41:26 PM GMT+08:00 | Refresh: Calibrations: Beam Quality, Tube potential test: HVL measurements are Automatically entered when Manual mode is selected and the base is connected | WIP9.1 | WIP11 |  | Subsystem |  | NA | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP11SegReq | Brilliance 64 |  | WIP11 | October 15, 2018 7:07:58 PM GMT+08:00
CTUDT00599793 | Refresh | Closed | azagury_d | September 4, 2018 8:41:26 PM GMT+08:00 | Refresh: Calibrations: Beam Quality, Tube potential test: HVL measurements are Automatically entered when Manual mode is selected and the base is connected | WIP9.1 | WIP11 |  | Subsystem |  | NA | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP11SegReq | Brilliance Big Bore |  | WIP11 | October 15, 2018 7:07:58 PM GMT+08:00
CTUDT00599793 | Refresh | Closed | azagury_d | September 4, 2018 8:41:26 PM GMT+08:00 | Refresh: Calibrations: Beam Quality, Tube potential test: HVL measurements are Automatically entered when Manual mode is selected and the base is connected | WIP9.1 | WIP11 |  | Subsystem |  | NA | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP11SegReq | Ingenuity 64 |  | WIP11 | October 15, 2018 7:07:58 PM GMT+08:00
CTUDT00599793 | Refresh | Closed | azagury_d | September 4, 2018 8:41:26 PM GMT+08:00 | Refresh: Calibrations: Beam Quality, Tube potential test: HVL measurements are Automatically entered when Manual mode is selected and the base is connected | WIP9.1 | WIP11 |  | Subsystem |  | NA | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP11SegReq | IQon1 |  | WIP11 | October 15, 2018 7:07:58 PM GMT+08:00
CTUDT00599790 | Refresh | Closed | azagury_d | September 4, 2018 8:00:58 PM GMT+08:00 | Refresh: Calibrations: Can't run HVL due to missing tables. | WIP9.1 | WIP11 |  | Subsystem |  | NAN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP11SegReq | Ovation |  | WIP11 | October 15, 2018 9:29:40 PM GMT+08:00
CTUDT00599790 | Refresh | Closed | azagury_d | September 4, 2018 8:00:58 PM GMT+08:00 | Refresh: Calibrations: Can't run HVL due to missing tables. | WIP9.1 | WIP11 |  | Subsystem |  | NAN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP11SegReq | Brilliance Big Bore |  | WIP11 | October 15, 2018 9:29:40 PM GMT+08:00
CTUDT00599790 | Refresh | Closed | azagury_d | September 4, 2018 8:00:58 PM GMT+08:00 | Refresh: Calibrations: Can't run HVL due to missing tables. | WIP9.1 | WIP11 |  | Subsystem |  | NAN | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP11SegReq | Ingenuity 64 |  | WIP11 | October 15, 2018 9:29:40 PM GMT+08:00
CTUDT00599787 | Refresh | Closed | azagury_d | September 4, 2018 7:04:24 PM GMT+08:00 | Refresh: Calibrations: HVL: Collimation parameter is not as defined in spec | WIP9.1 | WIP10 |  | Subsystem |  | N/A | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP11SegReq | Ovation |  | WIP10 | October 11, 2018 8:34:28 PM GMT+08:00
CTUDT00599787 | Refresh | Closed | azagury_d | September 4, 2018 7:04:24 PM GMT+08:00 | Refresh: Calibrations: HVL: Collimation parameter is not as defined in spec | WIP9.1 | WIP10 |  | Subsystem |  | N/A | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP11SegReq | Brilliance 64 |  | WIP10 | October 11, 2018 8:34:28 PM GMT+08:00
CTUDT00599787 | Refresh | Closed | azagury_d | September 4, 2018 7:04:24 PM GMT+08:00 | Refresh: Calibrations: HVL: Collimation parameter is not as defined in spec | WIP9.1 | WIP10 |  | Subsystem |  | N/A | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP11SegReq | Brilliance Big Bore |  | WIP10 | October 11, 2018 8:34:28 PM GMT+08:00
CTUDT00599787 | Refresh | Closed | azagury_d | September 4, 2018 7:04:24 PM GMT+08:00 | Refresh: Calibrations: HVL: Collimation parameter is not as defined in spec | WIP9.1 | WIP10 |  | Subsystem |  | N/A | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP11SegReq | Ingenuity 64 |  | WIP10 | October 11, 2018 8:34:28 PM GMT+08:00
CTUDT00599787 | Refresh | Closed | azagury_d | September 4, 2018 7:04:24 PM GMT+08:00 | Refresh: Calibrations: HVL: Collimation parameter is not as defined in spec | WIP9.1 | WIP10 |  | Subsystem |  | N/A | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP11SegReq | IQon1 |  | WIP10 | October 11, 2018 8:34:28 PM GMT+08:00
CTUDT00599784 | Refresh | Closed | gT_s | September 4, 2018 3:55:54 PM GMT+08:00 | Refresh [Scanner viewer]The exam card with the axial result of the detailed resolution is not accepting the maximum FOV value of 500 | WIP 9.1 | Host SSRS DHF303211 Rev 06 |  | Subsystem | EC parameters | Console-SRS.8922 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | WIP12SubCom | Ovation |  | Host SSRS DHF303211 Rev 06 | December 12, 2018 10:28:37 PM GMT+08:00
CTUDT00599784 | Refresh | Closed | gT_s | September 4, 2018 3:55:54 PM GMT+08:00 | Refresh [Scanner viewer]The exam card with the axial result of the detailed resolution is not accepting the maximum FOV value of 500 | WIP 9.1 | Host SSRS DHF303211 Rev 06 |  | Subsystem | EC parameters | Console-SRS.8922 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | WIP12SubCom | Brilliance 64 |  | Host SSRS DHF303211 Rev 06 | December 12, 2018 10:28:37 PM GMT+08:00
CTUDT00599784 | Refresh | Closed | gT_s | September 4, 2018 3:55:54 PM GMT+08:00 | Refresh [Scanner viewer]The exam card with the axial result of the detailed resolution is not accepting the maximum FOV value of 500 | WIP 9.1 | Host SSRS DHF303211 Rev 06 |  | Subsystem | EC parameters | Console-SRS.8922 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | WIP12SubCom | Ingenuity 64 |  | Host SSRS DHF303211 Rev 06 | December 12, 2018 10:28:37 PM GMT+08:00
CTUDT00599784 | Refresh | Closed | gT_s | September 4, 2018 3:55:54 PM GMT+08:00 | Refresh [Scanner viewer]The exam card with the axial result of the detailed resolution is not accepting the maximum FOV value of 500 | WIP 9.1 | Host SSRS DHF303211 Rev 06 |  | Subsystem | EC parameters | Console-SRS.8922 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | WIP12SubCom | IQon1 |  | Host SSRS DHF303211 Rev 06 | December 12, 2018 10:28:37 PM GMT+08:00
CTUDT00599781 | Refresh | Closed | wang_l | September 4, 2018 1:42:01 PM GMT+08:00 | Review button was enabled but Review screen was not launched once selecting CT viewer from PD or check Auto Launch for online recon | WIP9.1 |  | Click PD, then click review. | Subsystem | Auto Launch | NA | Works As Specified / Not A Defect | veerabhadraiah | 1 |  |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00599779 | Refresh | Closed | wang_l | September 4, 2018 12:37:29 PM GMT+08:00 | Not all images generated for reconstruction | WIP9.1 | WIP15 |  | Subsystem | Brain Helical; Verified1 | NA | Fixed | richards_b | 0 | 2 | CIRS_SW_ImgNodeSvr | CIRS | NO | Closed | WIP15SubCom | Ovation |  | WIP15 | April 9, 2019 4:57:06 AM GMT+08:00
CTUDT00599779 | Refresh | Closed | wang_l | September 4, 2018 12:37:29 PM GMT+08:00 | Not all images generated for reconstruction | WIP9.1 | WIP15 |  | Subsystem | Brain Helical; Verified1 | NA | Fixed | richards_b | 0 | 2 | CIRS_SW_ImgNodeSvr | CIRS | NO | Closed | WIP15SubCom | Brilliance 64 |  | WIP15 | April 9, 2019 4:57:06 AM GMT+08:00
CTUDT00599779 | Refresh | Closed | wang_l | September 4, 2018 12:37:29 PM GMT+08:00 | Not all images generated for reconstruction | WIP9.1 | WIP15 |  | Subsystem | Brain Helical; Verified1 | NA | Fixed | richards_b | 0 | 2 | CIRS_SW_ImgNodeSvr | CIRS | NO | Closed | WIP15SubCom | Brilliance Big Bore |  | WIP15 | April 9, 2019 4:57:06 AM GMT+08:00
CTUDT00599779 | Refresh | Closed | wang_l | September 4, 2018 12:37:29 PM GMT+08:00 | Not all images generated for reconstruction | WIP9.1 | WIP15 |  | Subsystem | Brain Helical; Verified1 | NA | Fixed | richards_b | 0 | 2 | CIRS_SW_ImgNodeSvr | CIRS | NO | Closed | WIP15SubCom | Ingenuity 64 |  | WIP15 | April 9, 2019 4:57:06 AM GMT+08:00
CTUDT00599779 | Refresh | Closed | wang_l | September 4, 2018 12:37:29 PM GMT+08:00 | Not all images generated for reconstruction | WIP9.1 | WIP15 |  | Subsystem | Brain Helical; Verified1 | NA | Fixed | richards_b | 0 | 2 | CIRS_SW_ImgNodeSvr | CIRS | NO | Closed | WIP15SubCom | IQon1 |  | WIP15 | April 9, 2019 4:57:06 AM GMT+08:00
CTUDT00599778 | Refresh | Closed | jotawar_d | September 3, 2018 9:49:33 PM GMT+08:00 | Delete Option not available for paused scan result | V4.8 | DHF316705 Rev 04 | NA | Subsystem |  | Console-SRS.2294 | Fixed | balasubramaian_m | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | BB_WIP1.1_SubCom | Brilliance Big Bore |  | DHF316705 Rev 04 | November 13, 2018 1:44:36 PM GMT+08:00
CTUDT00599777 | Refresh | Closed | jotawar_d | September 3, 2018 9:47:15 PM GMT+08:00 | Paste option not available in series list for completed scan result | V4.8 | DHF316705 Rev 04 | NA | Subsystem |  | Console-SRS.2294 | Fixed | balasubramaian_m | 0 | 4 | Requirements Documentation | Console4 | NO | Closed | BB_WIP1.1_SubCom | Brilliance Big Bore |  | DHF316705 Rev 04 | November 13, 2018 1:44:05 PM GMT+08:00
CTUDT00599776 | Refresh | Closed | jotawar_d | September 3, 2018 9:42:40 PM GMT+08:00 | Error message displayed once scan for EC Resp 4D CT without surview | V4.8 | BB_WIP1 |  | Subsystem |  | None | Fixed | datta_i | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP2_SubCom | Brilliance Big Bore |  | BB_WIP1 | September 11, 2018 7:43:23 PM GMT+08:00
CTUDT00599769 | Refresh | Closed | jotawar_d | September 3, 2018 9:32:23 PM GMT+08:00 | Assign accession number available for offline result | V4.8 | BB_WIP3 | NA | Subsystem | Offline Recon -accession number | N/A | Fixed | chokkalingam_r | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP3_SegReq | Brilliance Big Bore |  | BB_WIP3 | December 4, 2018 6:56:22 PM GMT+08:00
CTUDT00599765 | Refresh | Closed | kadish_n | September 3, 2018 9:22:17 PM GMT+08:00 | The application indicates that NVRAM information is not read even though it is | WIP9.1 | WIP10 | N/A | Subsystem |  | N/A | Fixed | yosefk | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP11SegReq | Ovation |  | WIP10 | October 11, 2018 3:49:24 PM GMT+08:00
CTUDT00599765 | Refresh | Closed | kadish_n | September 3, 2018 9:22:17 PM GMT+08:00 | The application indicates that NVRAM information is not read even though it is | WIP9.1 | WIP10 | N/A | Subsystem |  | N/A | Fixed | yosefk | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP11SegReq | IQon1 |  | WIP10 | October 11, 2018 3:49:24 PM GMT+08:00
CTUDT00599761 | Refresh | Closed | jotawar_d | September 3, 2018 9:19:44 PM GMT+08:00 | The exam application crashes sporadically after end of xray with error “StepEnded XXX arrived unexpectedly” | V4.1 | BB_WIP4 |  | Subsystem | History Review | NA | Fixed | v_sumanth | 0 | 2 | Console Software | Console4 | YES | Closed | BB_WIP4_SubCom | Brilliance Big Bore |  | BB_WIP4 | January 30, 2019 7:40:22 PM GMT+08:00
CTUDT00599760 | Refresh | Closed | jotawar_d | September 3, 2018 9:17:58 PM GMT+08:00 | Unable to pause the perfusion scan between the groups in perfusion scan when 1st group scan is completed | V4.8 | BB_WIP3 | User can pause the scan through CT Box. | Subsystem | Pause Scan | NA | Fixed | kumar_v | 0 | 2 | Console Software | Console4 | YES | Closed | BB_WIP3_SegReq | Brilliance Big Bore |  | BB_WIP3 | December 11, 2018 5:53:53 PM GMT+08:00
CTUDT00599759 | Refresh | Closed | mN_m | September 3, 2018 9:15:34 PM GMT+08:00 | Two recon requests generated for Dual energy scan with OMAR instead of one. | WIP9.1 | DHF303211 Rev07 |  | Subsystem | O-MAR - two additional recons were generated | Console-SRS.15065 | Fixed | hatti_c | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | WIP13SubCom | Ovation |  | DHF303211 Rev07 | January 25, 2019 6:44:19 PM GMT+08:00
CTUDT00599759 | Refresh | Closed | mN_m | September 3, 2018 9:15:34 PM GMT+08:00 | Two recon requests generated for Dual energy scan with OMAR instead of one. | WIP9.1 | DHF303211 Rev07 |  | Subsystem | O-MAR - two additional recons were generated | Console-SRS.15065 | Fixed | hatti_c | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | WIP13SubCom | Brilliance 64 |  | DHF303211 Rev07 | January 25, 2019 6:44:19 PM GMT+08:00
CTUDT00599759 | Refresh | Closed | mN_m | September 3, 2018 9:15:34 PM GMT+08:00 | Two recon requests generated for Dual energy scan with OMAR instead of one. | WIP9.1 | DHF303211 Rev07 |  | Subsystem | O-MAR - two additional recons were generated | Console-SRS.15065 | Fixed | hatti_c | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | WIP13SubCom | Ingenuity 64 |  | DHF303211 Rev07 | January 25, 2019 6:44:19 PM GMT+08:00
CTUDT00599758 | Refresh | Closed | jotawar_d | September 3, 2018 9:08:08 PM GMT+08:00 | For every offline recon of non gated results, new pulmo and cardiac series 299 is getting generated even though phase or any parameter is not changed. | 4.8 |  | No | Subsystem | Respiratory 4D CT - new series generated | NA | No Plan to Fix | sagar_v | 0 | 4 |  | Console4 | NO | Closed | SubNC | Brilliance Big Bore
CTUDT00599757 | Refresh | Closed | jotawar_d | September 3, 2018 9:05:26 PM GMT+08:00 | FOK: When LastScannedER file is corrupted the exam application crashes | V4.8 | BB_WIP7 |  | Subsystem | History Review | NA | Fixed | agrawal_ay | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP5_SegReq | Brilliance Big Bore | SWO 47764096 | BB_WIP7 | May 23, 2019 11:44:56 AM GMT+08:00
CTUDT00599756 | Refresh | Closed | jotawar_d | September 3, 2018 8:57:29 PM GMT+08:00 | Refresh: VIPR tag is turned ON in STANDARD reconstruction and also Low Signal (streak) correction reconstruction sub-stage(IRS_SUBSTAGE_PREP_LOW_SIGNAL) tag is missing. | V4.8 | BB_WIP1.1 | NA | Subsystem |  | Console-SRS.8955 | Fixed | sachan_a | 0 | 2 | Console Software | Console4 | YES | Closed | BB_WIP2_SubCom | Brilliance Big Bore |  | BB_WIP1.1 | September 21, 2018 3:00:53 PM GMT+08:00
CTUDT00599755 | Refresh | Closed | jotawar_d | September 3, 2018 8:52:21 PM GMT+08:00 | Refresh : Application is allowing usage of EC from different scanner model through “External Exam Card” option while loading a study on the home page. | V4.8 | BB_WIP2 | No workaround is required since user is knowingly loading EC from different scanner model. | Subsystem |  | N/A | Fixed | kumar_v | 0 | 3 | Console Software | Console4 | YES | Closed | BB_WIP1.1_SubCom | Brilliance Big Bore |  | BB_WIP2 | October 30, 2018 6:54:47 PM GMT+08:00
CTUDT00599751 | Refresh | Closed | vinoliya_j | September 3, 2018 5:58:09 PM GMT+08:00 | Console-SRS.9064 and Console-SRS.9065 needs to be updated. | WIP 8.1 | Host SSRS DHF303211 Rev 08 |  | Subsystem | Documentation | Console-SRS.9065 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | WIP14SegReq | Ovation |  | Host SSRS DHF303211 Rev 08 | March 22, 2019 6:07:48 PM GMT+08:00
CTUDT00599751 | Refresh | Closed | vinoliya_j | September 3, 2018 5:58:09 PM GMT+08:00 | Console-SRS.9064 and Console-SRS.9065 needs to be updated. | WIP 8.1 | Host SSRS DHF303211 Rev 08 |  | Subsystem | Documentation | Console-SRS.9065 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | WIP14SegReq | Brilliance 64 |  | Host SSRS DHF303211 Rev 08 | March 22, 2019 6:07:48 PM GMT+08:00
CTUDT00599751 | Refresh | Closed | vinoliya_j | September 3, 2018 5:58:09 PM GMT+08:00 | Console-SRS.9064 and Console-SRS.9065 needs to be updated. | WIP 8.1 | Host SSRS DHF303211 Rev 08 |  | Subsystem | Documentation | Console-SRS.9065 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | WIP14SegReq | Ingenuity 64 |  | Host SSRS DHF303211 Rev 08 | March 22, 2019 6:07:48 PM GMT+08:00
CTUDT00599751 | Refresh | Closed | vinoliya_j | September 3, 2018 5:58:09 PM GMT+08:00 | Console-SRS.9064 and Console-SRS.9065 needs to be updated. | WIP 8.1 | Host SSRS DHF303211 Rev 08 |  | Subsystem | Documentation | Console-SRS.9065 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | WIP14SegReq | IQon1 |  | Host SSRS DHF303211 Rev 08 | March 22, 2019 6:07:48 PM GMT+08:00
CTUDT00599746 | Refresh | Closed | lei_yang | September 3, 2018 1:53:02 PM GMT+08:00 | Options were not checked in Quality Assurance with Resolver Replacement. | WIP9.1 | WIP15 | The user (service) may run this separately | Subsystem | Quality Assurance - Resolver replacement | n/a | Fixed | frider_o | 0 | 2 | DMS Software | DMS | NO | Closed | WIP14SegReq | Ingenuity 64 |  | WIP15 | March 17, 2019 11:57:27 PM GMT+08:00
CTUDT00599743 | Refresh | Closed | fitusi_r | September 3, 2018 12:08:07 AM GMT+08:00 | Refresh, ICT -Axial scan perfusion fails. | WIP 9.1 | WIP13 |  | Subsystem | Axial Perfusion | n/a | Fixed | daibes_s | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP13SubCom | Ovation |  | WIP13 | January 17, 2019 4:53:43 PM GMT+08:00
CTUDT00599741 | Refresh | Closed | bagdady_b | September 2, 2018 11:09:59 PM GMT+08:00 | Refresh- BIG XRAY_ON message arrives to logviewer with delay of couple of minutes | WIP9.1 | WIP15 |  | Subsystem | Log - Performance | None | Fixed | ravindra_a | 0 | 4 | Console Software | Console4 | NO | Closed | WIP14SegReq | Ovation |  | WIP15 | March 14, 2019 7:49:30 PM GMT+08:00
CTUDT00599741 | Refresh | Closed | bagdady_b | September 2, 2018 11:09:59 PM GMT+08:00 | Refresh- BIG XRAY_ON message arrives to logviewer with delay of couple of minutes | WIP9.1 | WIP15 |  | Subsystem | Log - Performance | None | Fixed | ravindra_a | 0 | 4 | Console Software | Console4 | NO | Closed | WIP14SegReq | Brilliance 64 |  | WIP15 | March 14, 2019 7:49:30 PM GMT+08:00
CTUDT00599741 | Refresh | Closed | bagdady_b | September 2, 2018 11:09:59 PM GMT+08:00 | Refresh- BIG XRAY_ON message arrives to logviewer with delay of couple of minutes | WIP9.1 | WIP15 |  | Subsystem | Log - Performance | None | Fixed | ravindra_a | 0 | 4 | Console Software | Console4 | NO | Closed | WIP14SegReq | Ingenuity 64 |  | WIP15 | March 14, 2019 7:49:30 PM GMT+08:00
CTUDT00599741 | Refresh | Closed | bagdady_b | September 2, 2018 11:09:59 PM GMT+08:00 | Refresh- BIG XRAY_ON message arrives to logviewer with delay of couple of minutes | WIP9.1 | WIP15 |  | Subsystem | Log - Performance | None | Fixed | ravindra_a | 0 | 4 | Console Software | Console4 | NO | Closed | WIP14SegReq | IQon1 |  | WIP15 | March 14, 2019 7:49:30 PM GMT+08:00
CTUDT00599739 | Refresh | Closed | cohen_y | September 2, 2018 10:46:47 PM GMT+08:00 | Refresh: No images are reconstructed when IMR Recon Mode is enabled. | WIP9.1 | Host SSRS DHF303211 Rev 10 DHF316705 Rev10 |  | Subsystem |  | Console-SRS.17079 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | YES | Closed | WIP10SegReq | Ovation |  | Host SSRS DHF303211 Rev 10 DHF316705 Rev10 | May 9, 2019 4:46:52 PM GMT+08:00
CTUDT00599739 | Refresh | Closed | cohen_y | September 2, 2018 10:46:47 PM GMT+08:00 | Refresh: No images are reconstructed when IMR Recon Mode is enabled. | WIP9.1 | Host SSRS DHF303211 Rev 10 DHF316705 Rev10 |  | Subsystem |  | Console-SRS.17079 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | YES | Closed | WIP10SegReq | Brilliance 64 |  | Host SSRS DHF303211 Rev 10 DHF316705 Rev10 | May 9, 2019 4:46:52 PM GMT+08:00
CTUDT00599739 | Refresh | Closed | cohen_y | September 2, 2018 10:46:47 PM GMT+08:00 | Refresh: No images are reconstructed when IMR Recon Mode is enabled. | WIP9.1 | Host SSRS DHF303211 Rev 10 DHF316705 Rev10 |  | Subsystem |  | Console-SRS.17079 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | YES | Closed | WIP10SegReq | Brilliance Big Bore |  | Host SSRS DHF303211 Rev 10 DHF316705 Rev10 | May 9, 2019 4:46:52 PM GMT+08:00
CTUDT00599739 | Refresh | Closed | cohen_y | September 2, 2018 10:46:47 PM GMT+08:00 | Refresh: No images are reconstructed when IMR Recon Mode is enabled. | WIP9.1 | Host SSRS DHF303211 Rev 10 DHF316705 Rev10 |  | Subsystem |  | Console-SRS.17079 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | YES | Closed | WIP10SegReq | Ingenuity 64 |  | Host SSRS DHF303211 Rev 10 DHF316705 Rev10 | May 9, 2019 4:46:52 PM GMT+08:00
CTUDT00599737 | Refresh | Closed | attali_s | September 2, 2018 9:35:32 PM GMT+08:00 | Calibrations: Error in ConstancyAcceptance_ScenarioOrion.xml | WIP9.1 | WIP12 |  | Subsystem | Acceptence Test | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP11.1SegReq | Brilliance Big Bore |  | WIP12 | November 18, 2018 9:49:58 PM GMT+08:00
CTUDT00599737 | Refresh | Closed | attali_s | September 2, 2018 9:35:32 PM GMT+08:00 | Calibrations: Error in ConstancyAcceptance_ScenarioOrion.xml | WIP9.1 | WIP12 |  | Subsystem | Acceptence Test | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP11.1SegReq | Ingenuity 64 |  | WIP12 | November 18, 2018 9:49:58 PM GMT+08:00
CTUDT00599734 | Refresh | Closed | attali_s | September 2, 2018 7:15:28 PM GMT+08:00 | Calibrations: Wrong rotation time in HCor Slice Infant on Ingenuity | WIP9.1 | N/A |  | Subsystem | Calibrations - HCOR slice | N/A | Fixed | braunstein_d | 0 | 2 | Design Documentation | DMS | NO | Closed | BB_WIP4_SegReq | Ingenuity 64 |  | N/A | February 3, 2019 4:00:27 PM GMT+08:00
CTUDT00599719 | Refresh | Closed | sun_q | August 31, 2018 3:01:00 PM GMT+08:00 | Reliability Test is not existed in Quality Assurance with Service User | WIP9.1 | N/A |  | Subsystem |  | n/a | Fixed | braunstein_d | 0 | 2 | Design Documentation | DMS | NO | Closed | WIP13SegReq | Ovation |  | N/A | January 17, 2019 4:57:21 PM GMT+08:00
CTUDT00599719 | Refresh | Closed | sun_q | August 31, 2018 3:01:00 PM GMT+08:00 | Reliability Test is not existed in Quality Assurance with Service User | WIP9.1 | N/A |  | Subsystem |  | n/a | Fixed | braunstein_d | 0 | 2 | Design Documentation | DMS | NO | Closed | WIP13SegReq | Brilliance Big Bore |  | N/A | January 17, 2019 4:57:21 PM GMT+08:00
CTUDT00599719 | Refresh | Closed | sun_q | August 31, 2018 3:01:00 PM GMT+08:00 | Reliability Test is not existed in Quality Assurance with Service User | WIP9.1 | N/A |  | Subsystem |  | n/a | Fixed | braunstein_d | 0 | 2 | Design Documentation | DMS | NO | Closed | WIP13SegReq | Ingenuity 64 |  | N/A | January 17, 2019 4:57:21 PM GMT+08:00
CTUDT00599719 | Refresh | Closed | sun_q | August 31, 2018 3:01:00 PM GMT+08:00 | Reliability Test is not existed in Quality Assurance with Service User | WIP9.1 | N/A |  | Subsystem |  | n/a | Fixed | braunstein_d | 0 | 2 | Design Documentation | DMS | NO | Closed | WIP13SegReq | IQon1 |  | N/A | January 17, 2019 4:57:21 PM GMT+08:00
CTUDT00599716 | Refresh | Closed | g_di | August 31, 2018 1:48:10 PM GMT+08:00 | [Refresh WIP 8.1] [Exam Card In Plan Viewer]: Phase value and filming value (result step) is not displaying in the Orientation tab | WIP8.1 | WIP11 |  | System |  | NA | Fixed | datta_i | 0 | 2 | Console Software | Console4 | NO | Closed | WIP11SegReq | Ovation |  | WIP11 | October 24, 2018 7:51:16 PM GMT+08:00
CTUDT00599716 | Refresh | Closed | g_di | August 31, 2018 1:48:10 PM GMT+08:00 | [Refresh WIP 8.1] [Exam Card In Plan Viewer]: Phase value and filming value (result step) is not displaying in the Orientation tab | WIP8.1 | WIP11 |  | System |  | NA | Fixed | datta_i | 0 | 2 | Console Software | Console4 | NO | Closed | WIP11SegReq | Brilliance 64 |  | WIP11 | October 24, 2018 7:51:16 PM GMT+08:00
CTUDT00599716 | Refresh | Closed | g_di | August 31, 2018 1:48:10 PM GMT+08:00 | [Refresh WIP 8.1] [Exam Card In Plan Viewer]: Phase value and filming value (result step) is not displaying in the Orientation tab | WIP8.1 | WIP11 |  | System |  | NA | Fixed | datta_i | 0 | 2 | Console Software | Console4 | NO | Closed | WIP11SegReq | Ingenuity 64 |  | WIP11 | October 24, 2018 7:51:16 PM GMT+08:00
CTUDT00599716 | Refresh | Closed | g_di | August 31, 2018 1:48:10 PM GMT+08:00 | [Refresh WIP 8.1] [Exam Card In Plan Viewer]: Phase value and filming value (result step) is not displaying in the Orientation tab | WIP8.1 | WIP11 |  | System |  | NA | Fixed | datta_i | 0 | 2 | Console Software | Console4 | NO | Closed | WIP11SegReq | IQon1 |  | WIP11 | October 24, 2018 7:51:16 PM GMT+08:00
CTUDT00599713 | Refresh | Closed | gT_s | August 31, 2018 3:24:50 AM GMT+08:00 | Refresh [Scanner viewer]:The application is getting crashed and displayed "system error", On re-launching the same result where the main viewport & reference viewport images are being switched in slab viewer | WIP8.1 | WIP11.1 | the user can re-launch the result to a different viewer | Subsystem | Slab Viewer | NA | Fixed | datta_i | 0 | 4 | Console Software | Console4 | NO | Closed | WIP11.1SegReq | Ovation |  | WIP11.1 | November 22, 2018 4:27:54 PM GMT+08:00
CTUDT00599713 | Refresh | Closed | gT_s | August 31, 2018 3:24:50 AM GMT+08:00 | Refresh [Scanner viewer]:The application is getting crashed and displayed "system error", On re-launching the same result where the main viewport & reference viewport images are being switched in slab viewer | WIP8.1 | WIP11.1 | the user can re-launch the result to a different viewer | Subsystem | Slab Viewer | NA | Fixed | datta_i | 0 | 4 | Console Software | Console4 | NO | Closed | WIP11.1SegReq | Brilliance 64 |  | WIP11.1 | November 22, 2018 4:27:54 PM GMT+08:00
CTUDT00599713 | Refresh | Closed | gT_s | August 31, 2018 3:24:50 AM GMT+08:00 | Refresh [Scanner viewer]:The application is getting crashed and displayed "system error", On re-launching the same result where the main viewport & reference viewport images are being switched in slab viewer | WIP8.1 | WIP11.1 | the user can re-launch the result to a different viewer | Subsystem | Slab Viewer | NA | Fixed | datta_i | 0 | 4 | Console Software | Console4 | NO | Closed | WIP11.1SegReq | Ingenuity 64 |  | WIP11.1 | November 22, 2018 4:27:54 PM GMT+08:00
CTUDT00599713 | Refresh | Closed | gT_s | August 31, 2018 3:24:50 AM GMT+08:00 | Refresh [Scanner viewer]:The application is getting crashed and displayed "system error", On re-launching the same result where the main viewport & reference viewport images are being switched in slab viewer | WIP8.1 | WIP11.1 | the user can re-launch the result to a different viewer | Subsystem | Slab Viewer | NA | Fixed | datta_i | 0 | 4 | Console Software | Console4 | NO | Closed | WIP11.1SegReq | IQon1 |  | WIP11.1 | November 22, 2018 4:27:54 PM GMT+08:00
CTUDT00599706 | Refresh | Closed | bagdady_b | August 30, 2018 6:15:30 PM GMT+08:00 | Refresh- Unable to perfrom CPM Burn from System Configuration | WIP 9.1 |  |  | Subsystem | System Installation - CPM burn | N/A | Duplicate | Mark_Huang | 0 |  |  | Service Tools | NO | Closed |  | Ovation
CTUDT00599704 | Refresh | Closed | roshan_r | August 30, 2018 3:49:14 PM GMT+08:00 | Refresh > Exam Prep > patient Information : "Previous Exam" EC in the exam card selection is not displayed when "Continue previous exam" option in patient page is selected, after EC is modified in ECM. | WIP 8.1 |  |  | Subsystem | Continue previous exam - |  | Works As Specified / Not A Defect | babu_s | 1 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00599699 | Refresh | Closed | rakic_k | August 30, 2018 12:33:25 AM GMT+08:00 | Unknown VR Error Message when executing DMS calibrations | WIP8.0 | WIP9.1 |  | Subsystem | Verified1 | CIRS.CS.1138 and CIRS.SSRS.4929 | Fixed | rakic_k | 0 | 2 | CIRS_SW_ImgNodeSvr | CIRS | YES | Closed | WIP9.1SubCom | Ovation |  | WIP9.1 | September 18, 2018 8:46:50 PM GMT+08:00
CTUDT00599699 | Refresh | Closed | rakic_k | August 30, 2018 12:33:25 AM GMT+08:00 | Unknown VR Error Message when executing DMS calibrations | WIP8.0 | WIP9.1 |  | Subsystem | Verified1 | CIRS.CS.1138 and CIRS.SSRS.4929 | Fixed | rakic_k | 0 | 2 | CIRS_SW_ImgNodeSvr | CIRS | YES | Closed | WIP9.1SubCom | Brilliance 64 |  | WIP9.1 | September 18, 2018 8:46:50 PM GMT+08:00
CTUDT00599699 | Refresh | Closed | rakic_k | August 30, 2018 12:33:25 AM GMT+08:00 | Unknown VR Error Message when executing DMS calibrations | WIP8.0 | WIP9.1 |  | Subsystem | Verified1 | CIRS.CS.1138 and CIRS.SSRS.4929 | Fixed | rakic_k | 0 | 2 | CIRS_SW_ImgNodeSvr | CIRS | YES | Closed | WIP9.1SubCom | Ingenuity 64 |  | WIP9.1 | September 18, 2018 8:46:50 PM GMT+08:00
CTUDT00599699 | Refresh | Closed | rakic_k | August 30, 2018 12:33:25 AM GMT+08:00 | Unknown VR Error Message when executing DMS calibrations | WIP8.0 | WIP9.1 |  | Subsystem | Verified1 | CIRS.CS.1138 and CIRS.SSRS.4929 | Fixed | rakic_k | 0 | 2 | CIRS_SW_ImgNodeSvr | CIRS | YES | Closed | WIP9.1SubCom | IQon1 |  | WIP9.1 | September 18, 2018 8:46:50 PM GMT+08:00
CTUDT00599696 | Refresh | Closed | shekhar_p | August 29, 2018 7:08:27 PM GMT+08:00 | Refresh: User is able to access C and D drives in CT, IT and Olevel users | WIP9.1 | WIP10 |  | Subsystem |  | N/A | Fixed | mohit_p | 0 | 2 | Console Software | Console4 | YES | Closed | WIP10SegReq | Ovation |  | WIP10 | October 11, 2018 2:34:10 PM GMT+08:00
CTUDT00599696 | Refresh | Closed | shekhar_p | August 29, 2018 7:08:27 PM GMT+08:00 | Refresh: User is able to access C and D drives in CT, IT and Olevel users | WIP9.1 | WIP10 |  | Subsystem |  | N/A | Fixed | mohit_p | 0 | 2 | Console Software | Console4 | YES | Closed | WIP10SegReq | Brilliance 64 |  | WIP10 | October 11, 2018 2:34:10 PM GMT+08:00
CTUDT00599696 | Refresh | Closed | shekhar_p | August 29, 2018 7:08:27 PM GMT+08:00 | Refresh: User is able to access C and D drives in CT, IT and Olevel users | WIP9.1 | WIP10 |  | Subsystem |  | N/A | Fixed | mohit_p | 0 | 2 | Console Software | Console4 | YES | Closed | WIP10SegReq | Brilliance Big Bore |  | WIP10 | October 11, 2018 2:34:10 PM GMT+08:00
CTUDT00599696 | Refresh | Closed | shekhar_p | August 29, 2018 7:08:27 PM GMT+08:00 | Refresh: User is able to access C and D drives in CT, IT and Olevel users | WIP9.1 | WIP10 |  | Subsystem |  | N/A | Fixed | mohit_p | 0 | 2 | Console Software | Console4 | YES | Closed | WIP10SegReq | Ingenuity 64 |  | WIP10 | October 11, 2018 2:34:10 PM GMT+08:00
CTUDT00599696 | Refresh | Closed | shekhar_p | August 29, 2018 7:08:27 PM GMT+08:00 | Refresh: User is able to access C and D drives in CT, IT and Olevel users | WIP9.1 | WIP10 |  | Subsystem |  | N/A | Fixed | mohit_p | 0 | 2 | Console Software | Console4 | YES | Closed | WIP10SegReq | IQon1 |  | WIP10 | October 11, 2018 2:34:10 PM GMT+08:00
CTUDT00599694 | Refresh | Closed | shekhar_p | August 29, 2018 6:39:20 PM GMT+08:00 | Refresh: Login screen is frozen and UI is not displaying after a restart | WIP9.1 |  |  | Subsystem | UI Performance |  | Works As Specified / Not A Defect | sanjeeva_a | 1 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00599692 | Refresh | Closed | shekhar_p | August 29, 2018 5:59:15 PM GMT+08:00 | Refresh: Performance: Left monitor is freeze for 30 to 40secs after clicking on end exam. | WIP9.1 |  |  | Subsystem | UI Performance |  | Duplicate | ravindra_a | 0 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00599692 | Refresh | Closed | shekhar_p | August 29, 2018 5:59:15 PM GMT+08:00 | Refresh: Performance: Left monitor is freeze for 30 to 40secs after clicking on end exam. | WIP9.1 |  |  | Subsystem | UI Performance |  | Duplicate | ravindra_a | 0 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00599690 | Refresh | Closed | cohen_y | August 29, 2018 3:23:59 PM GMT+08:00 | Calibrations: CTDI: file not found error after each scan in CTDI when performing CTDI after a clean installation. | WIP9 | WIP11 | There is workaround: Create sub-folder D:\Pms\Physics\CTDI\ manually. | Subsystem |  | N/A | Fixed | yosefk | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP11SegReq | Ovation |  | WIP11 | October 23, 2018 9:21:32 PM GMT+08:00
CTUDT00599690 | Refresh | Closed | cohen_y | August 29, 2018 3:23:59 PM GMT+08:00 | Calibrations: CTDI: file not found error after each scan in CTDI when performing CTDI after a clean installation. | WIP9 | WIP11 | There is workaround: Create sub-folder D:\Pms\Physics\CTDI\ manually. | Subsystem |  | N/A | Fixed | yosefk | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP11SegReq | Brilliance Big Bore |  | WIP11 | October 23, 2018 9:21:32 PM GMT+08:00
CTUDT00599690 | Refresh | Closed | cohen_y | August 29, 2018 3:23:59 PM GMT+08:00 | Calibrations: CTDI: file not found error after each scan in CTDI when performing CTDI after a clean installation. | WIP9 | WIP11 | There is workaround: Create sub-folder D:\Pms\Physics\CTDI\ manually. | Subsystem |  | N/A | Fixed | yosefk | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP11SegReq | Ingenuity 64 |  | WIP11 | October 23, 2018 9:21:32 PM GMT+08:00
CTUDT00599690 | Refresh | Closed | cohen_y | August 29, 2018 3:23:59 PM GMT+08:00 | Calibrations: CTDI: file not found error after each scan in CTDI when performing CTDI after a clean installation. | WIP9 | WIP11 | There is workaround: Create sub-folder D:\Pms\Physics\CTDI\ manually. | Subsystem |  | N/A | Fixed | yosefk | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP11SegReq | IQon1 |  | WIP11 | October 23, 2018 9:21:32 PM GMT+08:00
CTUDT00599685 | Refresh | Closed | attali_s | August 28, 2018 8:53:09 PM GMT+08:00 | Calibrations: Error while performing Surview scans of system centering due to missing phantom calibration table | Wip 9.1 | WIP10 |  | Subsystem |  | N/A | Fixed | daibes_s | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP11SegReq | Ovation |  | WIP10 | October 11, 2018 7:22:39 PM GMT+08:00
CTUDT00599685 | Refresh | Closed | attali_s | August 28, 2018 8:53:09 PM GMT+08:00 | Calibrations: Error while performing Surview scans of system centering due to missing phantom calibration table | Wip 9.1 | WIP10 |  | Subsystem |  | N/A | Fixed | daibes_s | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP11SegReq | Brilliance 64 |  | WIP10 | October 11, 2018 7:22:39 PM GMT+08:00
CTUDT00599685 | Refresh | Closed | attali_s | August 28, 2018 8:53:09 PM GMT+08:00 | Calibrations: Error while performing Surview scans of system centering due to missing phantom calibration table | Wip 9.1 | WIP10 |  | Subsystem |  | N/A | Fixed | daibes_s | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP11SegReq | Brilliance Big Bore |  | WIP10 | October 11, 2018 7:22:39 PM GMT+08:00
CTUDT00599685 | Refresh | Closed | attali_s | August 28, 2018 8:53:09 PM GMT+08:00 | Calibrations: Error while performing Surview scans of system centering due to missing phantom calibration table | Wip 9.1 | WIP10 |  | Subsystem |  | N/A | Fixed | daibes_s | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP11SegReq | Ingenuity 64 |  | WIP10 | October 11, 2018 7:22:39 PM GMT+08:00
CTUDT00599685 | Refresh | Closed | attali_s | August 28, 2018 8:53:09 PM GMT+08:00 | Calibrations: Error while performing Surview scans of system centering due to missing phantom calibration table | Wip 9.1 | WIP10 |  | Subsystem |  | N/A | Fixed | daibes_s | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP11SegReq | IQon1 |  | WIP10 | October 11, 2018 7:22:39 PM GMT+08:00
CTUDT00599680 | Refresh | Closed | bagdady_b | August 28, 2018 4:14:52 PM GMT+08:00 | Refresh- Number of images reconstructed offline does not match the online | WIP8.1 |  |  | Subsystem | Spectral results planning |  | Duplicate | richards_b | 0 |  |  | CIRS | NO | Closed |  | IQon1
CTUDT00599676 | Refresh | Closed | singh_p | August 28, 2018 2:43:08 PM GMT+08:00 | Refresh Ingenuity WIP 8.1 : Deviation in parameters in UI Console from Exam Card Spec Ingenuity DHF192909. | WIP 8.1 | WIP13 |  | Subsystem | EC deviations | N/A | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP10SubCom | Brilliance 64 |  | WIP13 | December 14, 2018 7:15:45 PM GMT+08:00
CTUDT00599676 | Refresh | Closed | singh_p | August 28, 2018 2:43:08 PM GMT+08:00 | Refresh Ingenuity WIP 8.1 : Deviation in parameters in UI Console from Exam Card Spec Ingenuity DHF192909. | WIP 8.1 | WIP13 |  | Subsystem | EC deviations | N/A | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP10SubCom | Ingenuity 64 |  | WIP13 | December 14, 2018 7:15:45 PM GMT+08:00
CTUDT00599675 | Refresh | Closed | nanjunda_s | August 28, 2018 2:25:25 PM GMT+08:00 | Plan Viewer:Last saved setting for "Show in-active Results" and "Show in-active acquisitions" Context-Menu options not updated when logged out and login again | WIP8.1 | WIP13 | User can check\uncheck the "Show Inactive Acquisitions" and "Show Inactive Results" options in the context menu. | Subsystem | Plan Viewer - Context-Menu option | NA | Fixed | v_sumanth | 0 | 4 | Console Software | Console4 | NO | Closed | WIP13SubCom | Ovation |  | WIP13 | January 10, 2019 4:17:07 PM GMT+08:00
CTUDT00599675 | Refresh | Closed | nanjunda_s | August 28, 2018 2:25:25 PM GMT+08:00 | Plan Viewer:Last saved setting for "Show in-active Results" and "Show in-active acquisitions" Context-Menu options not updated when logged out and login again | WIP8.1 | WIP13 | User can check\uncheck the "Show Inactive Acquisitions" and "Show Inactive Results" options in the context menu. | Subsystem | Plan Viewer - Context-Menu option | NA | Fixed | v_sumanth | 0 | 4 | Console Software | Console4 | NO | Closed | WIP13SubCom | Brilliance 64 |  | WIP13 | January 10, 2019 4:17:07 PM GMT+08:00
CTUDT00599675 | Refresh | Closed | nanjunda_s | August 28, 2018 2:25:25 PM GMT+08:00 | Plan Viewer:Last saved setting for "Show in-active Results" and "Show in-active acquisitions" Context-Menu options not updated when logged out and login again | WIP8.1 | WIP13 | User can check\uncheck the "Show Inactive Acquisitions" and "Show Inactive Results" options in the context menu. | Subsystem | Plan Viewer - Context-Menu option | NA | Fixed | v_sumanth | 0 | 4 | Console Software | Console4 | NO | Closed | WIP13SubCom | Ingenuity 64 |  | WIP13 | January 10, 2019 4:17:07 PM GMT+08:00
CTUDT00599675 | Refresh | Closed | nanjunda_s | August 28, 2018 2:25:25 PM GMT+08:00 | Plan Viewer:Last saved setting for "Show in-active Results" and "Show in-active acquisitions" Context-Menu options not updated when logged out and login again | WIP8.1 | WIP13 | User can check\uncheck the "Show Inactive Acquisitions" and "Show Inactive Results" options in the context menu. | Subsystem | Plan Viewer - Context-Menu option | NA | Fixed | v_sumanth | 0 | 4 | Console Software | Console4 | NO | Closed | WIP13SubCom | IQon1 |  | WIP13 | January 10, 2019 4:17:07 PM GMT+08:00
CTUDT00599670 | Refresh | Closed | singh_p | August 27, 2018 6:02:14 PM GMT+08:00 | Ingenuity Refresh WIP 8.1 : Requirement Update - Texts "Scan type" and "Brain" need to be updated. | WIP 8.1 | Host SSRS DHF303211 Rev 08 |  | Subsystem | Documentation | Console-SRS-10461 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | WIP14SegReq | Ovation |  | Host SSRS DHF303211 Rev 08 | March 21, 2019 1:04:18 PM GMT+08:00
CTUDT00599670 | Refresh | Closed | singh_p | August 27, 2018 6:02:14 PM GMT+08:00 | Ingenuity Refresh WIP 8.1 : Requirement Update - Texts "Scan type" and "Brain" need to be updated. | WIP 8.1 | Host SSRS DHF303211 Rev 08 |  | Subsystem | Documentation | Console-SRS-10461 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | WIP14SegReq | Brilliance 64 |  | Host SSRS DHF303211 Rev 08 | March 21, 2019 1:04:18 PM GMT+08:00
CTUDT00599670 | Refresh | Closed | singh_p | August 27, 2018 6:02:14 PM GMT+08:00 | Ingenuity Refresh WIP 8.1 : Requirement Update - Texts "Scan type" and "Brain" need to be updated. | WIP 8.1 | Host SSRS DHF303211 Rev 08 |  | Subsystem | Documentation | Console-SRS-10461 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | WIP14SegReq | Ingenuity 64 |  | Host SSRS DHF303211 Rev 08 | March 21, 2019 1:04:18 PM GMT+08:00
CTUDT00599670 | Refresh | Closed | singh_p | August 27, 2018 6:02:14 PM GMT+08:00 | Ingenuity Refresh WIP 8.1 : Requirement Update - Texts "Scan type" and "Brain" need to be updated. | WIP 8.1 | Host SSRS DHF303211 Rev 08 |  | Subsystem | Documentation | Console-SRS-10461 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | WIP14SegReq | IQon1 |  | Host SSRS DHF303211 Rev 08 | March 21, 2019 1:04:18 PM GMT+08:00
CTUDT00599666 | Refresh | Closed | singh_p | August 27, 2018 2:56:52 PM GMT+08:00 | Ingenuity Refresh WIP 8.1 : Detailed Resolution was not available for Infant age group. | WIP 8.1 | Host SSRS DHF303211 Rev 05 |  | Subsystem | EC parameters - detailed resolution missing | Console-SRS-10052 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | WIP11SegReq | Brilliance 64 |  | Host SSRS DHF303211 Rev 05 | November 23, 2018 1:12:25 PM GMT+08:00
CTUDT00599666 | Refresh | Closed | singh_p | August 27, 2018 2:56:52 PM GMT+08:00 | Ingenuity Refresh WIP 8.1 : Detailed Resolution was not available for Infant age group. | WIP 8.1 | Host SSRS DHF303211 Rev 05 |  | Subsystem | EC parameters - detailed resolution missing | Console-SRS-10052 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | WIP11SegReq | Ingenuity 64 |  | Host SSRS DHF303211 Rev 05 | November 23, 2018 1:12:25 PM GMT+08:00
CTUDT00599661 | Refresh | Closed | zhong_x | August 27, 2018 9:05:40 AM GMT+08:00 | The Default Path is "G" in Lan Config, which is not available. | WIP9 | WIP12 | Change the settings using LANConfig. | Subsystem | Lan Config - Hard Drive letters | NA | Fixed | veerabhadraiah | 0 | 3 | USP | iCAP | NO | Closed | WIP12SegReq | Ovation |  | WIP12 | December 14, 2018 5:16:22 PM GMT+08:00
CTUDT00599661 | Refresh | Closed | zhong_x | August 27, 2018 9:05:40 AM GMT+08:00 | The Default Path is "G" in Lan Config, which is not available. | WIP9 | WIP12 | Change the settings using LANConfig. | Subsystem | Lan Config - Hard Drive letters | NA | Fixed | veerabhadraiah | 0 | 3 | USP | iCAP | NO | Closed | WIP12SegReq | Brilliance 64 |  | WIP12 | December 14, 2018 5:16:22 PM GMT+08:00
CTUDT00599661 | Refresh | Closed | zhong_x | August 27, 2018 9:05:40 AM GMT+08:00 | The Default Path is "G" in Lan Config, which is not available. | WIP9 | WIP12 | Change the settings using LANConfig. | Subsystem | Lan Config - Hard Drive letters | NA | Fixed | veerabhadraiah | 0 | 3 | USP | iCAP | NO | Closed | WIP12SegReq | Brilliance Big Bore |  | WIP12 | December 14, 2018 5:16:22 PM GMT+08:00
CTUDT00599661 | Refresh | Closed | zhong_x | August 27, 2018 9:05:40 AM GMT+08:00 | The Default Path is "G" in Lan Config, which is not available. | WIP9 | WIP12 | Change the settings using LANConfig. | Subsystem | Lan Config - Hard Drive letters | NA | Fixed | veerabhadraiah | 0 | 3 | USP | iCAP | NO | Closed | WIP12SegReq | Ingenuity 64 |  | WIP12 | December 14, 2018 5:16:22 PM GMT+08:00
CTUDT00599661 | Refresh | Closed | zhong_x | August 27, 2018 9:05:40 AM GMT+08:00 | The Default Path is "G" in Lan Config, which is not available. | WIP9 | WIP12 | Change the settings using LANConfig. | Subsystem | Lan Config - Hard Drive letters | NA | Fixed | veerabhadraiah | 0 | 3 | USP | iCAP | NO | Closed | WIP12SegReq | IQon1 |  | WIP12 | December 14, 2018 5:16:22 PM GMT+08:00
CTUDT00599658 | Refresh | Closed | mistry_h | August 24, 2018 7:40:56 PM GMT+08:00 | When pressing logout while there are reconstructions in queue, the wrong message appears for the user. | 4.1.7 | WIP15 | Use offline reconstruction | Subsystem | Login | NA | Fixed | v_sumanth | 0 | 3 | Console SSRS | Console4 | NO | Closed | WIP14SegReq | Ovation |  | WIP15 | April 2, 2019 7:12:22 PM GMT+08:00
CTUDT00599658 | Refresh | Closed | mistry_h | August 24, 2018 7:40:56 PM GMT+08:00 | When pressing logout while there are reconstructions in queue, the wrong message appears for the user. | 4.1.7 | WIP15 | Use offline reconstruction | Subsystem | Login | NA | Fixed | v_sumanth | 0 | 3 | Console SSRS | Console4 | NO | Closed | WIP14SegReq | Brilliance 64 |  | WIP15 | April 2, 2019 7:12:22 PM GMT+08:00
CTUDT00599658 | Refresh | Closed | mistry_h | August 24, 2018 7:40:56 PM GMT+08:00 | When pressing logout while there are reconstructions in queue, the wrong message appears for the user. | 4.1.7 | WIP15 | Use offline reconstruction | Subsystem | Login | NA | Fixed | v_sumanth | 0 | 3 | Console SSRS | Console4 | NO | Closed | WIP14SegReq | Ingenuity 64 |  | WIP15 | April 2, 2019 7:12:22 PM GMT+08:00
CTUDT00599658 | Refresh | Closed | mistry_h | August 24, 2018 7:40:56 PM GMT+08:00 | When pressing logout while there are reconstructions in queue, the wrong message appears for the user. | 4.1.7 | WIP15 | Use offline reconstruction | Subsystem | Login | NA | Fixed | v_sumanth | 0 | 3 | Console SSRS | Console4 | NO | Closed | WIP14SegReq | IQon1 |  | WIP15 | April 2, 2019 7:12:22 PM GMT+08:00
CTUDT00599657 | Refresh | Closed | mistry_h | August 24, 2018 7:38:19 PM GMT+08:00 | IQon HB: Duplicated Result(of ReviewBeforeSave result) does not have "Delete" option in context menu | 4.1.7 |  |  | Subsystem |  | None | Duplicate | Prakash_Chand_Yadav | 0 | 2 | Console SSRS | Console4 | NO | Closed |  | Ovation
CTUDT00599657 | Refresh | Closed | mistry_h | August 24, 2018 7:38:19 PM GMT+08:00 | IQon HB: Duplicated Result(of ReviewBeforeSave result) does not have "Delete" option in context menu | 4.1.7 |  |  | Subsystem |  | None | Duplicate | Prakash_Chand_Yadav | 0 | 2 | Console SSRS | Console4 | NO | Closed |  | Brilliance 64
CTUDT00599657 | Refresh | Closed | mistry_h | August 24, 2018 7:38:19 PM GMT+08:00 | IQon HB: Duplicated Result(of ReviewBeforeSave result) does not have "Delete" option in context menu | 4.1.7 |  |  | Subsystem |  | None | Duplicate | Prakash_Chand_Yadav | 0 | 2 | Console SSRS | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00599657 | Refresh | Closed | mistry_h | August 24, 2018 7:38:19 PM GMT+08:00 | IQon HB: Duplicated Result(of ReviewBeforeSave result) does not have "Delete" option in context menu | 4.1.7 |  |  | Subsystem |  | None | Duplicate | Prakash_Chand_Yadav | 0 | 2 | Console SSRS | Console4 | NO | Closed |  | IQon1
CTUDT00599656 | Refresh | Closed | mistry_h | August 24, 2018 7:34:25 PM GMT+08:00 | In ECM IQ Test group is available in User Set Exam cards | 4.1.7 |  |  | Subsystem | ECM | N/A | Works As Specified / Not A Defect | bandil_j | 0 | 2 | Console SSRS | Console4 | NO | Closed |  | Ovation
CTUDT00599656 | Refresh | Closed | mistry_h | August 24, 2018 7:34:25 PM GMT+08:00 | In ECM IQ Test group is available in User Set Exam cards | 4.1.7 |  |  | Subsystem | ECM | N/A | Works As Specified / Not A Defect | bandil_j | 0 | 2 | Console SSRS | Console4 | NO | Closed |  | Brilliance 64
CTUDT00599656 | Refresh | Closed | mistry_h | August 24, 2018 7:34:25 PM GMT+08:00 | In ECM IQ Test group is available in User Set Exam cards | 4.1.7 |  |  | Subsystem | ECM | N/A | Works As Specified / Not A Defect | bandil_j | 0 | 2 | Console SSRS | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00599655 | Refresh | Closed | mistry_h | August 24, 2018 7:32:33 PM GMT+08:00 | Delete option is not available in edit results when review before save is checked | 4.1.7 | WIP14 |  | Subsystem | Review before Save | None | Fixed | Prakash_Chand_Yadav | 0 | 2 | Console SSRS | Console4 | NO | Closed | WIP14SubCom | Ovation |  | WIP14 | February 26, 2019 4:20:20 PM GMT+08:00
CTUDT00599655 | Refresh | Closed | mistry_h | August 24, 2018 7:32:33 PM GMT+08:00 | Delete option is not available in edit results when review before save is checked | 4.1.7 | WIP14 |  | Subsystem | Review before Save | None | Fixed | Prakash_Chand_Yadav | 0 | 2 | Console SSRS | Console4 | NO | Closed | WIP14SubCom | Brilliance 64 |  | WIP14 | February 26, 2019 4:20:20 PM GMT+08:00
CTUDT00599655 | Refresh | Closed | mistry_h | August 24, 2018 7:32:33 PM GMT+08:00 | Delete option is not available in edit results when review before save is checked | 4.1.7 | WIP14 |  | Subsystem | Review before Save | None | Fixed | Prakash_Chand_Yadav | 0 | 2 | Console SSRS | Console4 | NO | Closed | WIP14SubCom | Ingenuity 64 |  | WIP14 | February 26, 2019 4:20:20 PM GMT+08:00
CTUDT00599655 | Refresh | Closed | mistry_h | August 24, 2018 7:32:33 PM GMT+08:00 | Delete option is not available in edit results when review before save is checked | 4.1.7 | WIP14 |  | Subsystem | Review before Save | None | Fixed | Prakash_Chand_Yadav | 0 | 2 | Console SSRS | Console4 | NO | Closed | WIP14SubCom | IQon1 |  | WIP14 | February 26, 2019 4:20:20 PM GMT+08:00
CTUDT00599654 | Refresh | Closed | mistry_h | August 24, 2018 7:30:58 PM GMT+08:00 | Found By IOCC:Procedure code sequence and Performed Protocol Code Sequence are not present in N-CREATE and N-SET -RQ. | 4.1.7 | WIP13 |  | System | Interop | N/A | Fixed | prasad_r | 0 | 4 | Console SSRS | Console4 | NO | Closed | WIP13SubCom | Ovation |  | WIP13 | January 24, 2019 7:25:02 PM GMT+08:00
CTUDT00599654 | Refresh | Closed | mistry_h | August 24, 2018 7:30:58 PM GMT+08:00 | Found By IOCC:Procedure code sequence and Performed Protocol Code Sequence are not present in N-CREATE and N-SET -RQ. | 4.1.7 | WIP13 |  | System | Interop | N/A | Fixed | prasad_r | 0 | 4 | Console SSRS | Console4 | NO | Closed | WIP13SubCom | Brilliance 64 |  | WIP13 | January 24, 2019 7:25:02 PM GMT+08:00
CTUDT00599654 | Refresh | Closed | mistry_h | August 24, 2018 7:30:58 PM GMT+08:00 | Found By IOCC:Procedure code sequence and Performed Protocol Code Sequence are not present in N-CREATE and N-SET -RQ. | 4.1.7 | WIP13 |  | System | Interop | N/A | Fixed | prasad_r | 0 | 4 | Console SSRS | Console4 | NO | Closed | WIP13SubCom | Brilliance Big Bore |  | WIP13 | January 24, 2019 7:25:02 PM GMT+08:00
CTUDT00599654 | Refresh | Closed | mistry_h | August 24, 2018 7:30:58 PM GMT+08:00 | Found By IOCC:Procedure code sequence and Performed Protocol Code Sequence are not present in N-CREATE and N-SET -RQ. | 4.1.7 | WIP13 |  | System | Interop | N/A | Fixed | prasad_r | 0 | 4 | Console SSRS | Console4 | NO | Closed | WIP13SubCom | Ingenuity 64 |  | WIP13 | January 24, 2019 7:25:02 PM GMT+08:00
CTUDT00599654 | Refresh | Closed | mistry_h | August 24, 2018 7:30:58 PM GMT+08:00 | Found By IOCC:Procedure code sequence and Performed Protocol Code Sequence are not present in N-CREATE and N-SET -RQ. | 4.1.7 | WIP13 |  | System | Interop | N/A | Fixed | prasad_r | 0 | 4 | Console SSRS | Console4 | NO | Closed | WIP13SubCom | IQon1 |  | WIP13 | January 24, 2019 7:25:02 PM GMT+08:00
CTUDT00599653 | Refresh | Closed | mistry_h | August 24, 2018 7:29:28 PM GMT+08:00 | found by IOCC:Referenced Non-Image Composite SOP Instance Sequence (0040, 0220) attribute not present in the N-SET request for both Scheduled and Unscheduled SWF case. | 4.1.7 |  | N/A | Subsystem | IOCC | N/A | Works As Specified / Not A Defect | bandil_j | 0 | 4 | Console SSRS | Console4 | NO | Closed | SubNC | Ovation
CTUDT00599653 | Refresh | Closed | mistry_h | August 24, 2018 7:29:28 PM GMT+08:00 | found by IOCC:Referenced Non-Image Composite SOP Instance Sequence (0040, 0220) attribute not present in the N-SET request for both Scheduled and Unscheduled SWF case. | 4.1.7 |  | N/A | Subsystem | IOCC | N/A | Works As Specified / Not A Defect | bandil_j | 0 | 4 | Console SSRS | Console4 | NO | Closed | SubNC | Brilliance 64
CTUDT00599653 | Refresh | Closed | mistry_h | August 24, 2018 7:29:28 PM GMT+08:00 | found by IOCC:Referenced Non-Image Composite SOP Instance Sequence (0040, 0220) attribute not present in the N-SET request for both Scheduled and Unscheduled SWF case. | 4.1.7 |  | N/A | Subsystem | IOCC | N/A | Works As Specified / Not A Defect | bandil_j | 0 | 4 | Console SSRS | Console4 | NO | Closed | SubNC | Brilliance Big Bore
CTUDT00599653 | Refresh | Closed | mistry_h | August 24, 2018 7:29:28 PM GMT+08:00 | found by IOCC:Referenced Non-Image Composite SOP Instance Sequence (0040, 0220) attribute not present in the N-SET request for both Scheduled and Unscheduled SWF case. | 4.1.7 |  | N/A | Subsystem | IOCC | N/A | Works As Specified / Not A Defect | bandil_j | 0 | 4 | Console SSRS | Console4 | NO | Closed | SubNC | Ingenuity 64
CTUDT00599653 | Refresh | Closed | mistry_h | August 24, 2018 7:29:28 PM GMT+08:00 | found by IOCC:Referenced Non-Image Composite SOP Instance Sequence (0040, 0220) attribute not present in the N-SET request for both Scheduled and Unscheduled SWF case. | 4.1.7 |  | N/A | Subsystem | IOCC | N/A | Works As Specified / Not A Defect | bandil_j | 0 | 4 | Console SSRS | Console4 | NO | Closed | SubNC | IQon1
CTUDT00599652 | Refresh | Closed | mistry_h | August 24, 2018 7:27:06 PM GMT+08:00 | Some of the reserved keywords FDOM, ZDOM, X ray are getting translated on changing the UI (interface) language to other languages. | 4.1.7 |  | NA | Subsystem | Localization | N/A | Works As Specified / Not A Defect | bandil_j | 0 | 2 | Requirements Documentation | Console4 | NO | Closed |  | Ovation
CTUDT00599652 | Refresh | Closed | mistry_h | August 24, 2018 7:27:06 PM GMT+08:00 | Some of the reserved keywords FDOM, ZDOM, X ray are getting translated on changing the UI (interface) language to other languages. | 4.1.7 |  | NA | Subsystem | Localization | N/A | Works As Specified / Not A Defect | bandil_j | 0 | 2 | Requirements Documentation | Console4 | NO | Closed |  | Brilliance 64
CTUDT00599652 | Refresh | Closed | mistry_h | August 24, 2018 7:27:06 PM GMT+08:00 | Some of the reserved keywords FDOM, ZDOM, X ray are getting translated on changing the UI (interface) language to other languages. | 4.1.7 |  | NA | Subsystem | Localization | N/A | Works As Specified / Not A Defect | bandil_j | 0 | 2 | Requirements Documentation | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00599651 | Refresh | Closed | mistry_h | August 24, 2018 6:29:08 PM GMT+08:00 | Iqon HB - Post threshold delay's accepted tolerance value is not met | 4.1.7 |  |  | Subsystem |  | N/A | Works As Specified / Not A Defect | bandil_j | 0 | 2 | Console SSRS | Console4 | NO | Closed |  | Ovation
CTUDT00599651 | Refresh | Closed | mistry_h | August 24, 2018 6:29:08 PM GMT+08:00 | Iqon HB - Post threshold delay's accepted tolerance value is not met | 4.1.7 |  |  | Subsystem |  | N/A | Works As Specified / Not A Defect | bandil_j | 0 | 2 | Console SSRS | Console4 | NO | Closed |  | Brilliance 64
CTUDT00599651 | Refresh | Closed | mistry_h | August 24, 2018 6:29:08 PM GMT+08:00 | Iqon HB - Post threshold delay's accepted tolerance value is not met | 4.1.7 |  |  | Subsystem |  | N/A | Works As Specified / Not A Defect | bandil_j | 0 | 2 | Console SSRS | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00599650 | Refresh | Closed | mistry_h | August 24, 2018 6:20:14 PM GMT+08:00 | Display of the first image in the Viewer is taking more than 1.5 second from X-ray On in CCT Scan | 4.1.7 |  |  | Subsystem | Performance | N/A | Works As Specified / Not A Defect | bandil_j | 0 | 2 | Console SSRS | Console4 | NO | Closed |  | Ovation
CTUDT00599650 | Refresh | Closed | mistry_h | August 24, 2018 6:20:14 PM GMT+08:00 | Display of the first image in the Viewer is taking more than 1.5 second from X-ray On in CCT Scan | 4.1.7 |  |  | Subsystem | Performance | N/A | Works As Specified / Not A Defect | bandil_j | 0 | 2 | Console SSRS | Console4 | NO | Closed |  | Brilliance 64
CTUDT00599650 | Refresh | Closed | mistry_h | August 24, 2018 6:20:14 PM GMT+08:00 | Display of the first image in the Viewer is taking more than 1.5 second from X-ray On in CCT Scan | 4.1.7 |  |  | Subsystem | Performance | N/A | Works As Specified / Not A Defect | bandil_j | 0 | 2 | Console SSRS | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00599649 | Refresh | Closed | mistry_h | August 24, 2018 6:19:07 PM GMT+08:00 | bugrep could be opened but bugrep.xml is absent | v4.7.5 | WIP14 | Manually save logs from logviewer. | Subsystem | Bugrep | N/A | Fixed | pandey_a | 0 | 2 | CIRS_SW_Generic | Console4 | NO | Closed | WIP16SubCom | Ovation |  | WIP14 | April 19, 2019 11:18:34 PM GMT+08:00
CTUDT00599649 | Refresh | Closed | mistry_h | August 24, 2018 6:19:07 PM GMT+08:00 | bugrep could be opened but bugrep.xml is absent | v4.7.5 | WIP14 | Manually save logs from logviewer. | Subsystem | Bugrep | N/A | Fixed | pandey_a | 0 | 2 | CIRS_SW_Generic | Console4 | NO | Closed | WIP16SubCom | Brilliance 64 |  | WIP14 | April 19, 2019 11:18:34 PM GMT+08:00
CTUDT00599649 | Refresh | Closed | mistry_h | August 24, 2018 6:19:07 PM GMT+08:00 | bugrep could be opened but bugrep.xml is absent | v4.7.5 | WIP14 | Manually save logs from logviewer. | Subsystem | Bugrep | N/A | Fixed | pandey_a | 0 | 2 | CIRS_SW_Generic | Console4 | NO | Closed | WIP16SubCom | Brilliance Big Bore |  | WIP14 | April 19, 2019 11:18:34 PM GMT+08:00
CTUDT00599649 | Refresh | Closed | mistry_h | August 24, 2018 6:19:07 PM GMT+08:00 | bugrep could be opened but bugrep.xml is absent | v4.7.5 | WIP14 | Manually save logs from logviewer. | Subsystem | Bugrep | N/A | Fixed | pandey_a | 0 | 2 | CIRS_SW_Generic | Console4 | NO | Closed | WIP16SubCom | Ingenuity 64 |  | WIP14 | April 19, 2019 11:18:34 PM GMT+08:00
CTUDT00599649 | Refresh | Closed | mistry_h | August 24, 2018 6:19:07 PM GMT+08:00 | bugrep could be opened but bugrep.xml is absent | v4.7.5 | WIP14 | Manually save logs from logviewer. | Subsystem | Bugrep | N/A | Fixed | pandey_a | 0 | 2 | CIRS_SW_Generic | Console4 | NO | Closed | WIP16SubCom | IQon1 |  | WIP14 | April 19, 2019 11:18:34 PM GMT+08:00
CTUDT00599648 | Refresh | Closed | mistry_h | August 24, 2018 6:17:12 PM GMT+08:00 | Val: Serviceability: Datapath test failed | 4.1.7 | WIP19 |  | System |  | NA | Fixed | chakraborty_d | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP17SubCom | Ovation |  | WIP19 | July 18, 2019 6:22:25 PM GMT+08:00
CTUDT00599647 | Refresh | Closed | mistry_h | August 24, 2018 6:13:54 PM GMT+08:00 | IQonHB: Perfusion word is missing in the message #0095 in all Non-English languages. | 4.1.7 | WIP17 | NA | System | Perfusion | N/A | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | MergeFrom4.x | Ovation |  | WIP17 | May 17, 2019 8:41:05 PM GMT+08:00
CTUDT00599647 | Refresh | Closed | mistry_h | August 24, 2018 6:13:54 PM GMT+08:00 | IQonHB: Perfusion word is missing in the message #0095 in all Non-English languages. | 4.1.7 | WIP17 | NA | System | Perfusion | N/A | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | MergeFrom4.x | Brilliance 64 |  | WIP17 | May 17, 2019 8:41:05 PM GMT+08:00
CTUDT00599647 | Refresh | Closed | mistry_h | August 24, 2018 6:13:54 PM GMT+08:00 | IQonHB: Perfusion word is missing in the message #0095 in all Non-English languages. | 4.1.7 | WIP17 | NA | System | Perfusion | N/A | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | MergeFrom4.x | Ingenuity 64 |  | WIP17 | May 17, 2019 8:41:05 PM GMT+08:00
CTUDT00599646 | Refresh | Closed | mistry_h | August 24, 2018 6:12:17 PM GMT+08:00 | The scan number in the series list doesn't match the recon queue manager | 4.1.7 | WIP14 | The highlighted scan provides indication on the scan number, although the recon series number does not match | System | RQM | N/A | Fixed | Snehashis_Dey | 0 | 2 | Console Software | Console4 | NO | Closed | MergeFrom4.x | Ovation |  | WIP14 | February 21, 2019 8:17:13 PM GMT+08:00
CTUDT00599646 | Refresh | Closed | mistry_h | August 24, 2018 6:12:17 PM GMT+08:00 | The scan number in the series list doesn't match the recon queue manager | 4.1.7 | WIP14 | The highlighted scan provides indication on the scan number, although the recon series number does not match | System | RQM | N/A | Fixed | Snehashis_Dey | 0 | 2 | Console Software | Console4 | NO | Closed | MergeFrom4.x | Brilliance 64 |  | WIP14 | February 21, 2019 8:17:13 PM GMT+08:00
CTUDT00599646 | Refresh | Closed | mistry_h | August 24, 2018 6:12:17 PM GMT+08:00 | The scan number in the series list doesn't match the recon queue manager | 4.1.7 | WIP14 | The highlighted scan provides indication on the scan number, although the recon series number does not match | System | RQM | N/A | Fixed | Snehashis_Dey | 0 | 2 | Console Software | Console4 | NO | Closed | MergeFrom4.x | Brilliance Big Bore |  | WIP14 | February 21, 2019 8:17:13 PM GMT+08:00
CTUDT00599646 | Refresh | Closed | mistry_h | August 24, 2018 6:12:17 PM GMT+08:00 | The scan number in the series list doesn't match the recon queue manager | 4.1.7 | WIP14 | The highlighted scan provides indication on the scan number, although the recon series number does not match | System | RQM | N/A | Fixed | Snehashis_Dey | 0 | 2 | Console Software | Console4 | NO | Closed | MergeFrom4.x | Ingenuity 64 |  | WIP14 | February 21, 2019 8:17:13 PM GMT+08:00
CTUDT00599645 | Refresh | Closed | mistry_h | August 24, 2018 6:10:45 PM GMT+08:00 | IqonHB: Localization testing translated string mismatch and also translated string is getting truncated. | 4.1.7 | WIP14 | Another language can be set to check the real information (in case of truncated string). But for the wrong strings, there is no workaround. | System | Localization | N/A | Fixed | Snehashis_Dey | 0 | 3 | Console Software | Console4 | NO | Closed | MergeFrom4.x | Ovation |  | WIP14 | February 26, 2019 2:57:10 PM GMT+08:00
CTUDT00599645 | Refresh | Closed | mistry_h | August 24, 2018 6:10:45 PM GMT+08:00 | IqonHB: Localization testing translated string mismatch and also translated string is getting truncated. | 4.1.7 | WIP14 | Another language can be set to check the real information (in case of truncated string). But for the wrong strings, there is no workaround. | System | Localization | N/A | Fixed | Snehashis_Dey | 0 | 3 | Console Software | Console4 | NO | Closed | MergeFrom4.x | Brilliance 64 |  | WIP14 | February 26, 2019 2:57:10 PM GMT+08:00
CTUDT00599645 | Refresh | Closed | mistry_h | August 24, 2018 6:10:45 PM GMT+08:00 | IqonHB: Localization testing translated string mismatch and also translated string is getting truncated. | 4.1.7 | WIP14 | Another language can be set to check the real information (in case of truncated string). But for the wrong strings, there is no workaround. | System | Localization | N/A | Fixed | Snehashis_Dey | 0 | 3 | Console Software | Console4 | NO | Closed | MergeFrom4.x | Brilliance Big Bore |  | WIP14 | February 26, 2019 2:57:10 PM GMT+08:00
CTUDT00599645 | Refresh | Closed | mistry_h | August 24, 2018 6:10:45 PM GMT+08:00 | IqonHB: Localization testing translated string mismatch and also translated string is getting truncated. | 4.1.7 | WIP14 | Another language can be set to check the real information (in case of truncated string). But for the wrong strings, there is no workaround. | System | Localization | N/A | Fixed | Snehashis_Dey | 0 | 3 | Console Software | Console4 | NO | Closed | MergeFrom4.x | Ingenuity 64 |  | WIP14 | February 26, 2019 2:57:10 PM GMT+08:00
CTUDT00599644 | Refresh | Closed | mistry_h | August 24, 2018 6:07:39 PM GMT+08:00 | Val: System Error during Acquisition Duplication in Active Scan | 4.1.7 | WIP14 | instead of cutting and pasting the clinical step, user can cut and paste the uncompleted surview. | System |  | N/A | Fixed | Snehashis_Dey | 0 | 4 | Console Software | Console4 | NO | Closed | MergeFrom4.x | Ovation |  | WIP14 | February 21, 2019 8:20:42 PM GMT+08:00
CTUDT00599644 | Refresh | Closed | mistry_h | August 24, 2018 6:07:39 PM GMT+08:00 | Val: System Error during Acquisition Duplication in Active Scan | 4.1.7 | WIP14 | instead of cutting and pasting the clinical step, user can cut and paste the uncompleted surview. | System |  | N/A | Fixed | Snehashis_Dey | 0 | 4 | Console Software | Console4 | NO | Closed | MergeFrom4.x | Brilliance 64 |  | WIP14 | February 21, 2019 8:20:42 PM GMT+08:00
CTUDT00599644 | Refresh | Closed | mistry_h | August 24, 2018 6:07:39 PM GMT+08:00 | Val: System Error during Acquisition Duplication in Active Scan | 4.1.7 | WIP14 | instead of cutting and pasting the clinical step, user can cut and paste the uncompleted surview. | System |  | N/A | Fixed | Snehashis_Dey | 0 | 4 | Console Software | Console4 | NO | Closed | MergeFrom4.x | Brilliance Big Bore |  | WIP14 | February 21, 2019 8:20:42 PM GMT+08:00
CTUDT00599644 | Refresh | Closed | mistry_h | August 24, 2018 6:07:39 PM GMT+08:00 | Val: System Error during Acquisition Duplication in Active Scan | 4.1.7 | WIP14 | instead of cutting and pasting the clinical step, user can cut and paste the uncompleted surview. | System |  | N/A | Fixed | Snehashis_Dey | 0 | 4 | Console Software | Console4 | NO | Closed | MergeFrom4.x | Ingenuity 64 |  | WIP14 | February 21, 2019 8:20:42 PM GMT+08:00
CTUDT00599643 | Refresh | Closed | mistry_h | August 24, 2018 5:52:39 PM GMT+08:00 | IQON HB: Number of images are showing more in PD and result viewer than planned when merge series is selected. | 4.1.7 | WIP16 | NA | Subsystem |  | NA | Fixed | jose_s | 0 | 2 | Console Software | Console4 | YES | Closed | MergeFrom4.x | Ovation |  | WIP16 | April 3, 2019 3:16:21 PM GMT+08:00
CTUDT00599643 | Refresh | Closed | mistry_h | August 24, 2018 5:52:39 PM GMT+08:00 | IQON HB: Number of images are showing more in PD and result viewer than planned when merge series is selected. | 4.1.7 | WIP16 | NA | Subsystem |  | NA | Fixed | jose_s | 0 | 2 | Console Software | Console4 | YES | Closed | MergeFrom4.x | Brilliance 64 |  | WIP16 | April 3, 2019 3:16:21 PM GMT+08:00
CTUDT00599643 | Refresh | Closed | mistry_h | August 24, 2018 5:52:39 PM GMT+08:00 | IQON HB: Number of images are showing more in PD and result viewer than planned when merge series is selected. | 4.1.7 | WIP16 | NA | Subsystem |  | NA | Fixed | jose_s | 0 | 2 | Console Software | Console4 | YES | Closed | MergeFrom4.x | Ingenuity 64 |  | WIP16 | April 3, 2019 3:16:21 PM GMT+08:00
CTUDT00599642 | Refresh | Closed | mistry_h | August 24, 2018 5:51:37 PM GMT+08:00 | Correct text of pop-up message is not displayed when exited from Voice Manager Application without saving changes | 4.1.7 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | kumar_l | 1 | 4 | Console Software | Console4 | NO | Closed | MergeFrom4.x | Ovation
CTUDT00599641 | Refresh | Closed | mistry_h | August 24, 2018 5:49:53 PM GMT+08:00 | "CopyToCIRS" operation fails to copy partial raw to CIRS. | 4.1.7 | WIP13 | Use the UTS Browser and import the raw | System | CopytoCIRS;Verified1 | NA | Fixed | Snehashis_Dey | 0 | 2 | Console Software | Console4 | NO | Closed | MergeFrom4.x | Ovation |  | WIP13 | February 4, 2019 4:08:21 PM GMT+08:00
CTUDT00599641 | Refresh | Closed | mistry_h | August 24, 2018 5:49:53 PM GMT+08:00 | "CopyToCIRS" operation fails to copy partial raw to CIRS. | 4.1.7 | WIP13 | Use the UTS Browser and import the raw | System | CopytoCIRS;Verified1 | NA | Fixed | Snehashis_Dey | 0 | 2 | Console Software | Console4 | NO | Closed | MergeFrom4.x | Brilliance 64 |  | WIP13 | February 4, 2019 4:08:21 PM GMT+08:00
CTUDT00599641 | Refresh | Closed | mistry_h | August 24, 2018 5:49:53 PM GMT+08:00 | "CopyToCIRS" operation fails to copy partial raw to CIRS. | 4.1.7 | WIP13 | Use the UTS Browser and import the raw | System | CopytoCIRS;Verified1 | NA | Fixed | Snehashis_Dey | 0 | 2 | Console Software | Console4 | NO | Closed | MergeFrom4.x | Ingenuity 64 |  | WIP13 | February 4, 2019 4:08:21 PM GMT+08:00
CTUDT00599640 | Refresh | Closed | mistry_h | August 24, 2018 5:48:23 PM GMT+08:00 | IQon HB:Exam Application crashes and Bugrep Error pops up on selecting Edit Auto Voice/BL from scan ruler when Auto Voice language is set to None in preference. | 4.1.7 | WIP13 | 1.Add Auto Voice language in preferences 2.Choose Previous patient in Patient Data 3. If surview was scanned, shoose Use Previous Surview 4. Continue | System |  | NA | Fixed | Snehashis_Dey | 0 | 4 | Console Software | Console4 | NO | Closed | MergeFrom4.x | Ovation |  | WIP13 | February 4, 2019 4:08:45 PM GMT+08:00
CTUDT00599640 | Refresh | Closed | mistry_h | August 24, 2018 5:48:23 PM GMT+08:00 | IQon HB:Exam Application crashes and Bugrep Error pops up on selecting Edit Auto Voice/BL from scan ruler when Auto Voice language is set to None in preference. | 4.1.7 | WIP13 | 1.Add Auto Voice language in preferences 2.Choose Previous patient in Patient Data 3. If surview was scanned, shoose Use Previous Surview 4. Continue | System |  | NA | Fixed | Snehashis_Dey | 0 | 4 | Console Software | Console4 | NO | Closed | MergeFrom4.x | Brilliance 64 |  | WIP13 | February 4, 2019 4:08:45 PM GMT+08:00
CTUDT00599640 | Refresh | Closed | mistry_h | August 24, 2018 5:48:23 PM GMT+08:00 | IQon HB:Exam Application crashes and Bugrep Error pops up on selecting Edit Auto Voice/BL from scan ruler when Auto Voice language is set to None in preference. | 4.1.7 | WIP13 | 1.Add Auto Voice language in preferences 2.Choose Previous patient in Patient Data 3. If surview was scanned, shoose Use Previous Surview 4. Continue | System |  | NA | Fixed | Snehashis_Dey | 0 | 4 | Console Software | Console4 | NO | Closed | MergeFrom4.x | Ingenuity 64 |  | WIP13 | February 4, 2019 4:08:45 PM GMT+08:00
CTUDT00599639 | Refresh | Closed | mistry_h | August 24, 2018 5:47:01 PM GMT+08:00 | RfVD: Hadassah: Images not sent to PACS if using "Review before save" and view2 is forced to exit | 4.1.7 | WIP13 | 1. manually copy to PACS 2. exit view2 only by using its ‘exit’ button, before clicking on “end exam and continue working” on the left monitor. | System |  | Console-SRS.7779 | Fixed | Snehashis_Dey | 0 | 4 | Console Software | Console4 | NO | Closed | MergeFrom4.x | Ovation | 9434185 | WIP13 | February 4, 2019 4:09:09 PM GMT+08:00
CTUDT00599639 | Refresh | Closed | mistry_h | August 24, 2018 5:47:01 PM GMT+08:00 | RfVD: Hadassah: Images not sent to PACS if using "Review before save" and view2 is forced to exit | 4.1.7 | WIP13 | 1. manually copy to PACS 2. exit view2 only by using its ‘exit’ button, before clicking on “end exam and continue working” on the left monitor. | System |  | Console-SRS.7779 | Fixed | Snehashis_Dey | 0 | 4 | Console Software | Console4 | NO | Closed | MergeFrom4.x | Brilliance 64 | 9434185 | WIP13 | February 4, 2019 4:09:09 PM GMT+08:00
CTUDT00599639 | Refresh | Closed | mistry_h | August 24, 2018 5:47:01 PM GMT+08:00 | RfVD: Hadassah: Images not sent to PACS if using "Review before save" and view2 is forced to exit | 4.1.7 | WIP13 | 1. manually copy to PACS 2. exit view2 only by using its ‘exit’ button, before clicking on “end exam and continue working” on the left monitor. | System |  | Console-SRS.7779 | Fixed | Snehashis_Dey | 0 | 4 | Console Software | Console4 | NO | Closed | MergeFrom4.x | Ingenuity 64 | 9434185 | WIP13 | February 4, 2019 4:09:09 PM GMT+08:00
CTUDT00599638 | Refresh | Closed | mistry_h | August 24, 2018 5:45:42 PM GMT+08:00 | Exam crashes after setting threshold below -100HU in Bolus sequence | 4.1.7 | WIP17 | no workaround for this scenario. | Subsystem | Bolus | NA | Fixed | Snehashis_Dey | 0 | 3 | Console Software | Console4 | YES | Closed | MergeFrom4.x | Ovation |  | WIP17 | May 13, 2019 6:47:27 PM GMT+08:00
CTUDT00599638 | Refresh | Closed | mistry_h | August 24, 2018 5:45:42 PM GMT+08:00 | Exam crashes after setting threshold below -100HU in Bolus sequence | 4.1.7 | WIP17 | no workaround for this scenario. | Subsystem | Bolus | NA | Fixed | Snehashis_Dey | 0 | 3 | Console Software | Console4 | YES | Closed | MergeFrom4.x | Brilliance 64 |  | WIP17 | May 13, 2019 6:47:27 PM GMT+08:00
CTUDT00599638 | Refresh | Closed | mistry_h | August 24, 2018 5:45:42 PM GMT+08:00 | Exam crashes after setting threshold below -100HU in Bolus sequence | 4.1.7 | WIP17 | no workaround for this scenario. | Subsystem | Bolus | NA | Fixed | Snehashis_Dey | 0 | 3 | Console Software | Console4 | YES | Closed | MergeFrom4.x | Ingenuity 64 |  | WIP17 | May 13, 2019 6:47:27 PM GMT+08:00
CTUDT00599637 | Refresh | Closed | mistry_h | August 24, 2018 5:44:14 PM GMT+08:00 | RfVD: Lahey: SYSTEM ERROR TRYING TO EXTEND SCAN | 4.1.7 | WIP13 | In the new Exam Application coming after crash make "Previous patient" and "Use previous surview" | Subsystem |  | NA | Fixed | sagar_v | 0 | 3 | Console Software | Console4 | YES | Closed | MergeFrom4.x | Ovation | CFR9438893 | WIP13 | February 4, 2019 4:09:56 PM GMT+08:00
CTUDT00599637 | Refresh | Closed | mistry_h | August 24, 2018 5:44:14 PM GMT+08:00 | RfVD: Lahey: SYSTEM ERROR TRYING TO EXTEND SCAN | 4.1.7 | WIP13 | In the new Exam Application coming after crash make "Previous patient" and "Use previous surview" | Subsystem |  | NA | Fixed | sagar_v | 0 | 3 | Console Software | Console4 | YES | Closed | MergeFrom4.x | Brilliance 64 | CFR9438893 | WIP13 | February 4, 2019 4:09:56 PM GMT+08:00
CTUDT00599637 | Refresh | Closed | mistry_h | August 24, 2018 5:44:14 PM GMT+08:00 | RfVD: Lahey: SYSTEM ERROR TRYING TO EXTEND SCAN | 4.1.7 | WIP13 | In the new Exam Application coming after crash make "Previous patient" and "Use previous surview" | Subsystem |  | NA | Fixed | sagar_v | 0 | 3 | Console Software | Console4 | YES | Closed | MergeFrom4.x | Ingenuity 64 | CFR9438893 | WIP13 | February 4, 2019 4:09:56 PM GMT+08:00
CTUDT00599636 | Refresh | Closed | mistry_h | August 24, 2018 5:29:59 PM GMT+08:00 | Reconstructions aren't appearing in the right order in view scene with Dual Rack | 4.1.7 | WIP13 | NA | Subsystem | RQM | NA | Fixed | Snehashis_Dey | 0 | 2 | Console Software | Console4 | NO | Closed | MergeFrom4.x | Ovation |  | WIP13 | February 4, 2019 4:10:15 PM GMT+08:00
CTUDT00599636 | Refresh | Closed | mistry_h | August 24, 2018 5:29:59 PM GMT+08:00 | Reconstructions aren't appearing in the right order in view scene with Dual Rack | 4.1.7 | WIP13 | NA | Subsystem | RQM | NA | Fixed | Snehashis_Dey | 0 | 2 | Console Software | Console4 | NO | Closed | MergeFrom4.x | Brilliance 64 |  | WIP13 | February 4, 2019 4:10:15 PM GMT+08:00
CTUDT00599636 | Refresh | Closed | mistry_h | August 24, 2018 5:29:59 PM GMT+08:00 | Reconstructions aren't appearing in the right order in view scene with Dual Rack | 4.1.7 | WIP13 | NA | Subsystem | RQM | NA | Fixed | Snehashis_Dey | 0 | 2 | Console Software | Console4 | NO | Closed | MergeFrom4.x | Ingenuity 64 |  | WIP13 | February 4, 2019 4:10:15 PM GMT+08:00
CTUDT00599635 | Refresh | Closed | mistry_h | August 24, 2018 5:28:15 PM GMT+08:00 | IQON HB: even though the orientation is different, the parameters like recon mode and level are getting updated for other results on selecting “all conventional phases” or “ all conventional coronary phases” | 4.1.7 | WIP13 | User can go and check each result to change manually | System |  | N/A | Fixed | Snehashis_Dey | 0 | 2 | Console Software | Console4 | YES | Closed | MergeFrom4.x | Ovation |  | WIP13 | February 4, 2019 4:10:41 PM GMT+08:00
CTUDT00599635 | Refresh | Closed | mistry_h | August 24, 2018 5:28:15 PM GMT+08:00 | IQON HB: even though the orientation is different, the parameters like recon mode and level are getting updated for other results on selecting “all conventional phases” or “ all conventional coronary phases” | 4.1.7 | WIP13 | User can go and check each result to change manually | System |  | N/A | Fixed | Snehashis_Dey | 0 | 2 | Console Software | Console4 | YES | Closed | MergeFrom4.x | Brilliance 64 |  | WIP13 | February 4, 2019 4:10:41 PM GMT+08:00
CTUDT00599635 | Refresh | Closed | mistry_h | August 24, 2018 5:28:15 PM GMT+08:00 | IQON HB: even though the orientation is different, the parameters like recon mode and level are getting updated for other results on selecting “all conventional phases” or “ all conventional coronary phases” | 4.1.7 | WIP13 | User can go and check each result to change manually | System |  | N/A | Fixed | Snehashis_Dey | 0 | 2 | Console Software | Console4 | YES | Closed | MergeFrom4.x | Ingenuity 64 |  | WIP13 | February 4, 2019 4:10:41 PM GMT+08:00
CTUDT00599634 | Refresh | Closed | mistry_h | August 24, 2018 5:26:48 PM GMT+08:00 | “ELSCINTCT_ECG_TRIGGER_DELAY” tag is missing for the new result generated on offline recon. | 4.1.7 | WIP13 |  | Subsystem | DICOM Tags | NA | Fixed | Snehashis_Dey | 0 | 2 | Console Software | Console4 | YES | Closed | MergeFrom4.x | Ovation |  | WIP13 | February 4, 2019 4:11:07 PM GMT+08:00
CTUDT00599634 | Refresh | Closed | mistry_h | August 24, 2018 5:26:48 PM GMT+08:00 | “ELSCINTCT_ECG_TRIGGER_DELAY” tag is missing for the new result generated on offline recon. | 4.1.7 | WIP13 |  | Subsystem | DICOM Tags | NA | Fixed | Snehashis_Dey | 0 | 2 | Console Software | Console4 | YES | Closed | MergeFrom4.x | Brilliance 64 |  | WIP13 | February 4, 2019 4:11:07 PM GMT+08:00
CTUDT00599634 | Refresh | Closed | mistry_h | August 24, 2018 5:26:48 PM GMT+08:00 | “ELSCINTCT_ECG_TRIGGER_DELAY” tag is missing for the new result generated on offline recon. | 4.1.7 | WIP13 |  | Subsystem | DICOM Tags | NA | Fixed | Snehashis_Dey | 0 | 2 | Console Software | Console4 | YES | Closed | MergeFrom4.x | Ingenuity 64 |  | WIP13 | February 4, 2019 4:11:07 PM GMT+08:00
CTUDT00599634 | Refresh | Closed | mistry_h | August 24, 2018 5:26:48 PM GMT+08:00 | “ELSCINTCT_ECG_TRIGGER_DELAY” tag is missing for the new result generated on offline recon. | 4.1.7 | WIP13 |  | Subsystem | DICOM Tags | NA | Fixed | Snehashis_Dey | 0 | 2 | Console Software | Console4 | YES | Closed | MergeFrom4.x | IQon1 |  | WIP13 | February 4, 2019 4:11:07 PM GMT+08:00
CTUDT00599633 | Refresh | Closed | mistry_h | August 24, 2018 5:10:40 PM GMT+08:00 | 4.1.3 FOK No thin slices generated + view2 crash | 4.2.1 | WIP14 | The user may view the images using off-line recon, potentially after logout-login | Subsystem |  | N/A | Fixed | Snehashis_Dey | 0 | 3 | Console Software | Console4 | YES | Closed | MergeFrom4.x | Ovation | CFR9376892 | WIP14 | February 21, 2019 8:14:14 PM GMT+08:00
CTUDT00599633 | Refresh | Closed | mistry_h | August 24, 2018 5:10:40 PM GMT+08:00 | 4.1.3 FOK No thin slices generated + view2 crash | 4.2.1 | WIP14 | The user may view the images using off-line recon, potentially after logout-login | Subsystem |  | N/A | Fixed | Snehashis_Dey | 0 | 3 | Console Software | Console4 | YES | Closed | MergeFrom4.x | Brilliance 64 | CFR9376892 | WIP14 | February 21, 2019 8:14:14 PM GMT+08:00
CTUDT00599633 | Refresh | Closed | mistry_h | August 24, 2018 5:10:40 PM GMT+08:00 | 4.1.3 FOK No thin slices generated + view2 crash | 4.2.1 | WIP14 | The user may view the images using off-line recon, potentially after logout-login | Subsystem |  | N/A | Fixed | Snehashis_Dey | 0 | 3 | Console Software | Console4 | YES | Closed | MergeFrom4.x | Brilliance Big Bore | CFR9376892 | WIP14 | February 21, 2019 8:14:14 PM GMT+08:00
CTUDT00599633 | Refresh | Closed | mistry_h | August 24, 2018 5:10:40 PM GMT+08:00 | 4.1.3 FOK No thin slices generated + view2 crash | 4.2.1 | WIP14 | The user may view the images using off-line recon, potentially after logout-login | Subsystem |  | N/A | Fixed | Snehashis_Dey | 0 | 3 | Console Software | Console4 | YES | Closed | MergeFrom4.x | Ingenuity 64 | CFR9376892 | WIP14 | February 21, 2019 8:14:14 PM GMT+08:00
CTUDT00599633 | Refresh | Closed | mistry_h | August 24, 2018 5:10:40 PM GMT+08:00 | 4.1.3 FOK No thin slices generated + view2 crash | 4.2.1 | WIP14 | The user may view the images using off-line recon, potentially after logout-login | Subsystem |  | N/A | Fixed | Snehashis_Dey | 0 | 3 | Console Software | Console4 | YES | Closed | MergeFrom4.x | IQon1 | CFR9376892 | WIP14 | February 21, 2019 8:14:14 PM GMT+08:00
CTUDT00599632 | Refresh | Closed | mistry_h | August 24, 2018 5:05:39 PM GMT+08:00 | IV: System ended exam before completion | Ingenuity_TF64-128 - RfVD WIP 109 | WIP14 |  | System |  | N/A | Fixed | Snehashis_Dey | 0 | 2 | PET-CT Host | Console4 | YES | Closed | MergeFrom4.x | Ovation |  | WIP14 | February 26, 2019 3:00:02 PM GMT+08:00
CTUDT00599632 | Refresh | Closed | mistry_h | August 24, 2018 5:05:39 PM GMT+08:00 | IV: System ended exam before completion | Ingenuity_TF64-128 - RfVD WIP 109 | WIP14 |  | System |  | N/A | Fixed | Snehashis_Dey | 0 | 2 | PET-CT Host | Console4 | YES | Closed | MergeFrom4.x | Brilliance 64 |  | WIP14 | February 26, 2019 3:00:02 PM GMT+08:00
CTUDT00599632 | Refresh | Closed | mistry_h | August 24, 2018 5:05:39 PM GMT+08:00 | IV: System ended exam before completion | Ingenuity_TF64-128 - RfVD WIP 109 | WIP14 |  | System |  | N/A | Fixed | Snehashis_Dey | 0 | 2 | PET-CT Host | Console4 | YES | Closed | MergeFrom4.x | Brilliance Big Bore |  | WIP14 | February 26, 2019 3:00:02 PM GMT+08:00
CTUDT00599632 | Refresh | Closed | mistry_h | August 24, 2018 5:05:39 PM GMT+08:00 | IV: System ended exam before completion | Ingenuity_TF64-128 - RfVD WIP 109 | WIP14 |  | System |  | N/A | Fixed | Snehashis_Dey | 0 | 2 | PET-CT Host | Console4 | YES | Closed | MergeFrom4.x | Ingenuity 64 |  | WIP14 | February 26, 2019 3:00:02 PM GMT+08:00
CTUDT00599632 | Refresh | Closed | mistry_h | August 24, 2018 5:05:39 PM GMT+08:00 | IV: System ended exam before completion | Ingenuity_TF64-128 - RfVD WIP 109 | WIP14 |  | System |  | N/A | Fixed | Snehashis_Dey | 0 | 2 | PET-CT Host | Console4 | YES | Closed | MergeFrom4.x | IQon1 |  | WIP14 | February 26, 2019 3:00:02 PM GMT+08:00
CTUDT00599631 | Refresh | Closed | mistry_h | August 24, 2018 4:50:47 PM GMT+08:00 | Cardiac Signal not found in Workflow | WIP5 |  |  | Subsystem |  |  | Does Not Apply | bazargani_n | 0 | 2 | Console-Console4 | Console4 | YES | Closed |  | IQon1
CTUDT00599630 | Refresh | Closed | mistry_h | August 24, 2018 4:45:35 PM GMT+08:00 | Fast Get Ready label in Workflow Preferences is confusing | Veyron_Pre_RfV | WIP14 |  | Subsystem |  | NA | Fixed | Snehashis_Dey | 0 | 3 | Console-Console4 | Console4 | NO | Closed | MergeFrom4.x | Ovation |  | WIP14 | February 26, 2019 2:59:25 PM GMT+08:00
CTUDT00599630 | Refresh | Closed | mistry_h | August 24, 2018 4:45:35 PM GMT+08:00 | Fast Get Ready label in Workflow Preferences is confusing | Veyron_Pre_RfV | WIP14 |  | Subsystem |  | NA | Fixed | Snehashis_Dey | 0 | 3 | Console-Console4 | Console4 | NO | Closed | MergeFrom4.x | Brilliance 64 |  | WIP14 | February 26, 2019 2:59:25 PM GMT+08:00
CTUDT00599630 | Refresh | Closed | mistry_h | August 24, 2018 4:45:35 PM GMT+08:00 | Fast Get Ready label in Workflow Preferences is confusing | Veyron_Pre_RfV | WIP14 |  | Subsystem |  | NA | Fixed | Snehashis_Dey | 0 | 3 | Console-Console4 | Console4 | NO | Closed | MergeFrom4.x | Brilliance Big Bore |  | WIP14 | February 26, 2019 2:59:25 PM GMT+08:00
CTUDT00599630 | Refresh | Closed | mistry_h | August 24, 2018 4:45:35 PM GMT+08:00 | Fast Get Ready label in Workflow Preferences is confusing | Veyron_Pre_RfV | WIP14 |  | Subsystem |  | NA | Fixed | Snehashis_Dey | 0 | 3 | Console-Console4 | Console4 | NO | Closed | MergeFrom4.x | Ingenuity 64 |  | WIP14 | February 26, 2019 2:59:25 PM GMT+08:00
CTUDT00599630 | Refresh | Closed | mistry_h | August 24, 2018 4:45:35 PM GMT+08:00 | Fast Get Ready label in Workflow Preferences is confusing | Veyron_Pre_RfV | WIP14 |  | Subsystem |  | NA | Fixed | Snehashis_Dey | 0 | 3 | Console-Console4 | Console4 | NO | Closed | MergeFrom4.x | IQon1 |  | WIP14 | February 26, 2019 2:59:25 PM GMT+08:00
CTUDT00599628 | Refresh | Closed | mistry_h | August 24, 2018 4:23:21 PM GMT+08:00 | IV: CCT: Pan & Zoom functions were not available on active CCT scan with 1 display mode. | 4.1.0.XX259-WIP212 | WIP14 | Switch to 3 image display if pan/zoom is required. | Subsystem | History Review | NA | Fixed | Snehashis_Dey | 0 | 3 | Console Generic | Console4 | NO | Closed | MergeFrom4.x | Ovation |  | WIP14 | February 26, 2019 2:54:45 PM GMT+08:00
CTUDT00599628 | Refresh | Closed | mistry_h | August 24, 2018 4:23:21 PM GMT+08:00 | IV: CCT: Pan & Zoom functions were not available on active CCT scan with 1 display mode. | 4.1.0.XX259-WIP212 | WIP14 | Switch to 3 image display if pan/zoom is required. | Subsystem | History Review | NA | Fixed | Snehashis_Dey | 0 | 3 | Console Generic | Console4 | NO | Closed | MergeFrom4.x | Brilliance 64 |  | WIP14 | February 26, 2019 2:54:45 PM GMT+08:00
CTUDT00599628 | Refresh | Closed | mistry_h | August 24, 2018 4:23:21 PM GMT+08:00 | IV: CCT: Pan & Zoom functions were not available on active CCT scan with 1 display mode. | 4.1.0.XX259-WIP212 | WIP14 | Switch to 3 image display if pan/zoom is required. | Subsystem | History Review | NA | Fixed | Snehashis_Dey | 0 | 3 | Console Generic | Console4 | NO | Closed | MergeFrom4.x | Ingenuity 64 |  | WIP14 | February 26, 2019 2:54:45 PM GMT+08:00
CTUDT00599628 | Refresh | Closed | mistry_h | August 24, 2018 4:23:21 PM GMT+08:00 | IV: CCT: Pan & Zoom functions were not available on active CCT scan with 1 display mode. | 4.1.0.XX259-WIP212 | WIP14 | Switch to 3 image display if pan/zoom is required. | Subsystem | History Review | NA | Fixed | Snehashis_Dey | 0 | 3 | Console Generic | Console4 | NO | Closed | MergeFrom4.x | IQon1 |  | WIP14 | February 26, 2019 2:54:45 PM GMT+08:00
CTUDT00599627 | Refresh | Closed | mistry_h | August 24, 2018 4:17:50 PM GMT+08:00 | RFC: IV: EC Sharing: Auto voice was not imported when using export\import ECs. | 4.1.0.XX241-WIP211 | BB_WIP6 | Turn On Auto Voice | Subsystem | Auto Voice - import | NA | Fixed | datta_i | 0 | 3 | Console Generic | Console4 | NO | Closed | WIP6.1 | IQon1 |  | BB_WIP6 | January 10, 2019 4:21:54 PM GMT+08:00
CTUDT00599619 | Refresh | Closed | levin_c | August 23, 2018 10:45:06 PM GMT+08:00 | when you choose to change language in preferences the localization does not apply. | WIP9 |  |  | Subsystem |  |  | Duplicate | levin_c | 0 |  |  |  |  | Closed |  | IQon1
CTUDT00599591 | Refresh | Closed | jotawar_d | August 23, 2018 2:22:08 PM GMT+08:00 | Refresh [Ingenuity] - Bosnian, Bulgarian, Croatian, Estonian, Indonesian, Macedonian, Serbian, Slovak, Ukrainian languages are missing in voice manager and Auto voice option. | WIP6.1 | BB_WIP1.1 | NA | Subsystem |  | N/A | Fixed | sachan_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP1_SegReq | Brilliance Big Bore |  | BB_WIP1.1 | September 6, 2018 8:14:08 PM GMT+08:00
CTUDT00599584 | Refresh | Closed | manjunath_c | August 22, 2018 9:30:09 PM GMT+08:00 | [Refresh] [Ingenuity]: Console is not saving EC, If EC name contains single quote | Wip8.1 |  |  | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 |  | Console4 | NO | Closed |  | Ovation
CTUDT00599584 | Refresh | Closed | manjunath_c | August 22, 2018 9:30:09 PM GMT+08:00 | [Refresh] [Ingenuity]: Console is not saving EC, If EC name contains single quote | Wip8.1 |  |  | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00599584 | Refresh | Closed | manjunath_c | August 22, 2018 9:30:09 PM GMT+08:00 | [Refresh] [Ingenuity]: Console is not saving EC, If EC name contains single quote | Wip8.1 |  |  | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00599584 | Refresh | Closed | manjunath_c | August 22, 2018 9:30:09 PM GMT+08:00 | [Refresh] [Ingenuity]: Console is not saving EC, If EC name contains single quote | Wip8.1 |  |  | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00599584 | Refresh | Closed | manjunath_c | August 22, 2018 9:30:09 PM GMT+08:00 | [Refresh] [Ingenuity]: Console is not saving EC, If EC name contains single quote | Wip8.1 |  |  | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 |  | Console4 | NO | Closed |  | IQon1
CTUDT00599551 | Refresh | Closed | rathore_cs | August 20, 2018 5:18:50 PM GMT+08:00 | ECG device/Bellow Device not connection displayed in Console UI while the device was connecting. | 4.8 | BB_WIP4 | NA | Subsystem | Bellow Device - Connection display | NA | Fixed | patra_r | 0 | 2 | Console Software | Console4 | NO | Closed | BB_WIP4_SubCom | Brilliance Big Bore |  | BB_WIP4 | February 5, 2019 7:37:28 PM GMT+08:00
CTUDT00599536 | Refresh | Closed | bagdady_b | August 19, 2018 7:21:07 PM GMT+08:00 | Refresh- Not all images created error when performing extend of MPR results | WIP9 | WIP15 |  | Subsystem | Extended Scan - MPR; Verified1 | NA | Fixed | richards_b | 0 | 2 | CIRS_SW_ImgNodeSvr | CIRS | NO | Closed | WIP15SubCom | Ovation |  | WIP15 | April 9, 2019 4:56:06 AM GMT+08:00
CTUDT00599536 | Refresh | Closed | bagdady_b | August 19, 2018 7:21:07 PM GMT+08:00 | Refresh- Not all images created error when performing extend of MPR results | WIP9 | WIP15 |  | Subsystem | Extended Scan - MPR; Verified1 | NA | Fixed | richards_b | 0 | 2 | CIRS_SW_ImgNodeSvr | CIRS | NO | Closed | WIP15SubCom | Brilliance 64 |  | WIP15 | April 9, 2019 4:56:06 AM GMT+08:00
CTUDT00599536 | Refresh | Closed | bagdady_b | August 19, 2018 7:21:07 PM GMT+08:00 | Refresh- Not all images created error when performing extend of MPR results | WIP9 | WIP15 |  | Subsystem | Extended Scan - MPR; Verified1 | NA | Fixed | richards_b | 0 | 2 | CIRS_SW_ImgNodeSvr | CIRS | NO | Closed | WIP15SubCom | Brilliance Big Bore |  | WIP15 | April 9, 2019 4:56:06 AM GMT+08:00
CTUDT00599536 | Refresh | Closed | bagdady_b | August 19, 2018 7:21:07 PM GMT+08:00 | Refresh- Not all images created error when performing extend of MPR results | WIP9 | WIP15 |  | Subsystem | Extended Scan - MPR; Verified1 | NA | Fixed | richards_b | 0 | 2 | CIRS_SW_ImgNodeSvr | CIRS | NO | Closed | WIP15SubCom | Ingenuity 64 |  | WIP15 | April 9, 2019 4:56:06 AM GMT+08:00
CTUDT00599536 | Refresh | Closed | bagdady_b | August 19, 2018 7:21:07 PM GMT+08:00 | Refresh- Not all images created error when performing extend of MPR results | WIP9 | WIP15 |  | Subsystem | Extended Scan - MPR; Verified1 | NA | Fixed | richards_b | 0 | 2 | CIRS_SW_ImgNodeSvr | CIRS | NO | Closed | WIP15SubCom | IQon1 |  | WIP15 | April 9, 2019 4:56:06 AM GMT+08:00
CTUDT00599529 | Refresh | Closed | gT_s | August 17, 2018 5:41:52 PM GMT+08:00 | Refresh [Scanner viewer]: Pictorial directory is not getting displayed in result viewer when clicked on start final recon in edit results or while switching to result viewer | WIP8.1 |  |  | Subsystem | Result Viewer - Pictorial directory |  | Works As Specified / Not A Defect | sanjeeva_a | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00599519 | Refresh | Closed | wang_l | August 17, 2018 1:50:01 PM GMT+08:00 | Acquisition for series with bolus can't display completely | WIP9 |  |  | Subsystem |  | Console.SRS.9801 | Works As Specified / Not A Defect | sanjeeva_a | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00599516 | Refresh | Closed | navale_s | August 17, 2018 1:38:09 PM GMT+08:00 | Auto Voice checkbox is getting rechecked after clicking on “Apply to All Scans” for Locator and tracker scans in “Auto Voice Parameters” window. | V4.8 | BB_WIP1.1 |  | Subsystem |  | NA | Fixed | navale_s | 0 | 3 | Console Software | Console4 | NO | Closed | BB_WIP1.1_SubCom | Brilliance Big Bore |  | BB_WIP1.1 | October 3, 2018 6:07:59 PM GMT+08:00
CTUDT00599512 | Refresh | Closed | singh_p | August 16, 2018 8:24:21 PM GMT+08:00 | Requirement Update : Refresh [Ingenuity] : Wrong range value of mA mentioned in requirement. | WIP8.1 | Host SSRS DHF303211 Rev 05 |  | Subsystem | Documentation | N/A | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | DHF303211 Rev04 | Ovation |  | Host SSRS DHF303211 Rev 05 | November 23, 2018 1:12:43 PM GMT+08:00
CTUDT00599512 | Refresh | Closed | singh_p | August 16, 2018 8:24:21 PM GMT+08:00 | Requirement Update : Refresh [Ingenuity] : Wrong range value of mA mentioned in requirement. | WIP8.1 | Host SSRS DHF303211 Rev 05 |  | Subsystem | Documentation | N/A | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | DHF303211 Rev04 | Brilliance 64 |  | Host SSRS DHF303211 Rev 05 | November 23, 2018 1:12:43 PM GMT+08:00
CTUDT00599512 | Refresh | Closed | singh_p | August 16, 2018 8:24:21 PM GMT+08:00 | Requirement Update : Refresh [Ingenuity] : Wrong range value of mA mentioned in requirement. | WIP8.1 | Host SSRS DHF303211 Rev 05 |  | Subsystem | Documentation | N/A | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | DHF303211 Rev04 | Ingenuity 64 |  | Host SSRS DHF303211 Rev 05 | November 23, 2018 1:12:43 PM GMT+08:00
CTUDT00599510 | Refresh | Closed | nimma_p | August 16, 2018 8:02:34 PM GMT+08:00 | After performing offline recon bed position is not matching with planning. | WIP 8.1 |  |  | Subsystem | Offline Recon - bed position |  | Works As Specified / Not A Defect | sanjeeva_a | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00599497 | Refresh | Closed | bagdady_b | August 15, 2018 6:26:33 PM GMT+08:00 | Refresh- Keyswitch is prompt when performing a restart to the host | WIP9 |  |  | Subsystem | System Reliability | NA | No Plan to Fix | karkour | 0 | 4 |  | GCX | NO | Closed |  | IQon1
CTUDT00599493 | Refresh | Closed | nimma_p | August 14, 2018 9:41:22 PM GMT+08:00 | User could not able perform the extended scan. | WIP 8.1 |  |  | Subsystem | Extended Scan |  | Works As Specified / Not A Defect | babu_s | 1 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00599489 | Refresh | Closed | golzban_v | August 14, 2018 7:18:21 PM GMT+08:00 | Refresh: IQon: "Set Phase Point" CIRS Event Viewer entries are missing for several phases | WIP9 |  |  | Subsystem | Coronary CTA | N/A | Works As Specified / Not A Defect | richards_b | 0 |  |  | CIRS | NO | Closed | SubNC | Ovation
CTUDT00599489 | Refresh | Closed | golzban_v | August 14, 2018 7:18:21 PM GMT+08:00 | Refresh: IQon: "Set Phase Point" CIRS Event Viewer entries are missing for several phases | WIP9 |  |  | Subsystem | Coronary CTA | N/A | Works As Specified / Not A Defect | richards_b | 0 |  |  | CIRS | NO | Closed | SubNC | IQon1
CTUDT00599487 | Refresh | Closed | golzban_v | August 14, 2018 7:02:12 PM GMT+08:00 | Refresh: IQon: Cardiac DoseRight enabled during planning, but no actual dose modulation takes place | WIP9 |  |  | Subsystem | Dose Right | NA | System Repaired | karkour | 0 |  |  | GCX | NO | Closed |  | Ovation
CTUDT00599487 | Refresh | Closed | golzban_v | August 14, 2018 7:02:12 PM GMT+08:00 | Refresh: IQon: Cardiac DoseRight enabled during planning, but no actual dose modulation takes place | WIP9 |  |  | Subsystem | Dose Right | NA | System Repaired | karkour | 0 |  |  | GCX | NO | Closed |  | IQon1
CTUDT00599485 | Refresh | Closed | golzban_v | August 14, 2018 4:31:33 PM GMT+08:00 | Refresh: IQon: Axial cardiac scans cannot be performed | WIP9 |  |  | Subsystem | Cardiac Axial scan | NA | System Repaired | karkour | 0 |  |  | GCX | NO | Closed |  | IQon1
CTUDT00599474 | Refresh | Closed | patra_r | August 13, 2018 2:31:09 PM GMT+08:00 | Maximum Buffer overflow length can be exceeded than the calculated value for selected Pitch and RT | V4.1.10 & V4.7.7 | WIP10 |  | Subsystem |  | None | Fixed | patra_r | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP1.1_SubCom | Ovation |  | WIP10 | September 14, 2018 9:55:01 PM GMT+08:00
CTUDT00599474 | Refresh | Closed | patra_r | August 13, 2018 2:31:09 PM GMT+08:00 | Maximum Buffer overflow length can be exceeded than the calculated value for selected Pitch and RT | V4.1.10 & V4.7.7 | WIP10 |  | Subsystem |  | None | Fixed | patra_r | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP1.1_SubCom | IQon1 |  | WIP10 | September 14, 2018 9:55:01 PM GMT+08:00
CTUDT00599472 | Refresh | Closed | wang_l | August 13, 2018 11:48:14 AM GMT+08:00 | Time value displayed in Directory>Study incorrect | WIP9 | WIP9.1 |  | Subsystem |  | N/A | Fixed | navale_s | 0 | 4 | Console Software | Console4 | NO | Closed | WIP9SegReq | Ingenuity 64 |  | WIP9.1 | January 10, 2019 4:22:37 PM GMT+08:00
CTUDT00599470 | Refresh | Closed | wang_l | August 13, 2018 11:30:43 AM GMT+08:00 | IMR CIRS connected but raw data missed in Reconstruction tab | WIP9 | WIP14 | N/A | Subsystem | ePHI Removal | N/A | Fixed | kr_m | 0 | 2 | Console Software | Console4 | YES | Closed | WIP14SubCom | Ovation |  | WIP14 | February 19, 2019 7:56:54 PM GMT+08:00
CTUDT00599470 | Refresh | Closed | wang_l | August 13, 2018 11:30:43 AM GMT+08:00 | IMR CIRS connected but raw data missed in Reconstruction tab | WIP9 | WIP14 | N/A | Subsystem | ePHI Removal | N/A | Fixed | kr_m | 0 | 2 | Console Software | Console4 | YES | Closed | WIP14SubCom | Brilliance 64 |  | WIP14 | February 19, 2019 7:56:54 PM GMT+08:00
CTUDT00599470 | Refresh | Closed | wang_l | August 13, 2018 11:30:43 AM GMT+08:00 | IMR CIRS connected but raw data missed in Reconstruction tab | WIP9 | WIP14 | N/A | Subsystem | ePHI Removal | N/A | Fixed | kr_m | 0 | 2 | Console Software | Console4 | YES | Closed | WIP14SubCom | Ingenuity 64 |  | WIP14 | February 19, 2019 7:56:54 PM GMT+08:00
CTUDT00599470 | Refresh | Closed | wang_l | August 13, 2018 11:30:43 AM GMT+08:00 | IMR CIRS connected but raw data missed in Reconstruction tab | WIP9 | WIP14 | N/A | Subsystem | ePHI Removal | N/A | Fixed | kr_m | 0 | 2 | Console Software | Console4 | YES | Closed | WIP14SubCom | IQon1 |  | WIP14 | February 19, 2019 7:56:54 PM GMT+08:00
CTUDT00599468 | Refresh | Closed | wang_l | August 13, 2018 11:18:58 AM GMT+08:00 | Reconstructioned failed with recon mode IMR | WIP9 | WIP9.1 |  | System | Verified1 | CIRS.SSRS.4873, CIRS.SSRS.5023, CIRS.SSRS.5022, CIRS.SSRS.4870, CIRS.SSRS.5019 | Fixed | rakic_k | 0 | 2 | CIRS_SW_ImgNodeSvr | CIRS | YES | Closed | WIP9.1SubCom | Ovation |  | WIP9.1 | September 18, 2018 9:00:44 PM GMT+08:00
CTUDT00599468 | Refresh | Closed | wang_l | August 13, 2018 11:18:58 AM GMT+08:00 | Reconstructioned failed with recon mode IMR | WIP9 | WIP9.1 |  | System | Verified1 | CIRS.SSRS.4873, CIRS.SSRS.5023, CIRS.SSRS.5022, CIRS.SSRS.4870, CIRS.SSRS.5019 | Fixed | rakic_k | 0 | 2 | CIRS_SW_ImgNodeSvr | CIRS | YES | Closed | WIP9.1SubCom | Brilliance 64 |  | WIP9.1 | September 18, 2018 9:00:44 PM GMT+08:00
CTUDT00599468 | Refresh | Closed | wang_l | August 13, 2018 11:18:58 AM GMT+08:00 | Reconstructioned failed with recon mode IMR | WIP9 | WIP9.1 |  | System | Verified1 | CIRS.SSRS.4873, CIRS.SSRS.5023, CIRS.SSRS.5022, CIRS.SSRS.4870, CIRS.SSRS.5019 | Fixed | rakic_k | 0 | 2 | CIRS_SW_ImgNodeSvr | CIRS | YES | Closed | WIP9.1SubCom | Ingenuity 64 |  | WIP9.1 | September 18, 2018 9:00:44 PM GMT+08:00
CTUDT00599466 | Refresh | Closed | wang_l | August 13, 2018 10:53:15 AM GMT+08:00 | Cannot continue scan after press red Estop button in CT box | WIP9 |  | For first bugrep: Check for stuck or faulty ESTOP button Verify Main Drive Contactor is closed If issue is persistent, HW team should investigate. For 2nd bugrep: Open the estop, wait ~ 20 seconds, then close the estop. | Subsystem | Gantry HW - CT - Box | N/A | Works As Specified / Not A Defect | wang_c | 0 |  |  | GCX | NO | Closed |  | Ingenuity 64
CTUDT00599464 | Refresh | Closed | wang_l | August 13, 2018 10:33:09 AM GMT+08:00 | Reconstruction images amount displayed in Reconstruction Jobs did not matched with the expected value in Plan Show All for result type Cor, Sag and Oblique; | WIP9 |  |  | System |  | N/A | Works As Specified / Not A Defect | bandil_j | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00599462 | Refresh | Closed | wang_l | August 13, 2018 10:17:00 AM GMT+08:00 | View2 tab was not displayed once click Recon button or End Exam and Continue Working | WIP9 | WIP10 | Click PD, then click view2. | Subsystem |  | NA | Fixed | s_sa | 0 | 4 | Console USP_EBW | iCAP | NO | Closed | WIP10SegReq | Ovation |  | WIP10 | November 27, 2019 3:17:29 PM GMT+08:00
CTUDT00599462 | Refresh | Closed | wang_l | August 13, 2018 10:17:00 AM GMT+08:00 | View2 tab was not displayed once click Recon button or End Exam and Continue Working | WIP9 | WIP10 | Click PD, then click view2. | Subsystem |  | NA | Fixed | s_sa | 0 | 4 | Console USP_EBW | iCAP | NO | Closed | WIP10SegReq | Brilliance 64 |  | WIP10 | November 27, 2019 3:17:29 PM GMT+08:00
CTUDT00599462 | Refresh | Closed | wang_l | August 13, 2018 10:17:00 AM GMT+08:00 | View2 tab was not displayed once click Recon button or End Exam and Continue Working | WIP9 | WIP10 | Click PD, then click view2. | Subsystem |  | NA | Fixed | s_sa | 0 | 4 | Console USP_EBW | iCAP | NO | Closed | WIP10SegReq | Brilliance Big Bore |  | WIP10 | November 27, 2019 3:17:29 PM GMT+08:00
CTUDT00599462 | Refresh | Closed | wang_l | August 13, 2018 10:17:00 AM GMT+08:00 | View2 tab was not displayed once click Recon button or End Exam and Continue Working | WIP9 | WIP10 | Click PD, then click view2. | Subsystem |  | NA | Fixed | s_sa | 0 | 4 | Console USP_EBW | iCAP | NO | Closed | WIP10SegReq | Ingenuity 64 |  | WIP10 | November 27, 2019 3:17:29 PM GMT+08:00
CTUDT00599462 | Refresh | Closed | wang_l | August 13, 2018 10:17:00 AM GMT+08:00 | View2 tab was not displayed once click Recon button or End Exam and Continue Working | WIP9 | WIP10 | Click PD, then click view2. | Subsystem |  | NA | Fixed | s_sa | 0 | 4 | Console USP_EBW | iCAP | NO | Closed | WIP10SegReq | IQon1 |  | WIP10 | November 27, 2019 3:17:29 PM GMT+08:00
CTUDT00599459 | Refresh | Closed | singh_p | August 10, 2018 5:56:32 PM GMT+08:00 | Ingenuity [Refresh WIP 8.1] : Small Focus Diameter has been found in MDUs series structure of clinical scan for Detailed Resolution. | WIP8.1 | Host SSRS DHF303211 Rev 05 |  | Subsystem | Documentation | Console-SRS-10754 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | WIP11SegReq | Ingenuity 64 |  | Host SSRS DHF303211 Rev 05 | November 22, 2018 6:12:45 PM GMT+08:00
CTUDT00599450 | Refresh | Closed | richards_b | August 9, 2018 11:11:47 AM GMT+08:00 | Update Console HW SSRS DHF272407 | Console HW SSRS | rev09 |  | Subsystem | Documentation | Performance CIRS.SSRS | Fixed | tysl_k | 0 | 4 | Design Documentation | CIRS | NO | Closed | SubCom | Ovation |  | rev09 | February 13, 2019 5:52:09 AM GMT+08:00
CTUDT00599450 | Refresh | Closed | richards_b | August 9, 2018 11:11:47 AM GMT+08:00 | Update Console HW SSRS DHF272407 | Console HW SSRS | rev09 |  | Subsystem | Documentation | Performance CIRS.SSRS | Fixed | tysl_k | 0 | 4 | Design Documentation | CIRS | NO | Closed | SubCom | Brilliance 64 |  | rev09 | February 13, 2019 5:52:09 AM GMT+08:00
CTUDT00599450 | Refresh | Closed | richards_b | August 9, 2018 11:11:47 AM GMT+08:00 | Update Console HW SSRS DHF272407 | Console HW SSRS | rev09 |  | Subsystem | Documentation | Performance CIRS.SSRS | Fixed | tysl_k | 0 | 4 | Design Documentation | CIRS | NO | Closed | SubCom | Brilliance Big Bore |  | rev09 | February 13, 2019 5:52:09 AM GMT+08:00
CTUDT00599450 | Refresh | Closed | richards_b | August 9, 2018 11:11:47 AM GMT+08:00 | Update Console HW SSRS DHF272407 | Console HW SSRS | rev09 |  | Subsystem | Documentation | Performance CIRS.SSRS | Fixed | tysl_k | 0 | 4 | Design Documentation | CIRS | NO | Closed | SubCom | Ingenuity 64 |  | rev09 | February 13, 2019 5:52:09 AM GMT+08:00
CTUDT00599450 | Refresh | Closed | richards_b | August 9, 2018 11:11:47 AM GMT+08:00 | Update Console HW SSRS DHF272407 | Console HW SSRS | rev09 |  | Subsystem | Documentation | Performance CIRS.SSRS | Fixed | tysl_k | 0 | 4 | Design Documentation | CIRS | NO | Closed | SubCom | IQon1 |  | rev09 | February 13, 2019 5:52:09 AM GMT+08:00
CTUDT00599447 | Refresh | Closed | roshan_r | August 8, 2018 9:46:28 PM GMT+08:00 | Refresh : DICOM tag (0032:1060 Requested Procedure Description) is not available in View DICOM Info in the Local folder | WIP8.1 | DHF318051 Rev 13 HOST Verification Test Protocol |  | Subsystem | DICOM | NA | Fixed | ramesh_t | 0 | 2 | Test Design | Console4 | NO | Closed | WIP15SegReq | Ovation |  | DHF318051 Rev 13 HOST Verification Test Protocol | April 23, 2019 7:51:49 PM GMT+08:00
CTUDT00599447 | Refresh | Closed | roshan_r | August 8, 2018 9:46:28 PM GMT+08:00 | Refresh : DICOM tag (0032:1060 Requested Procedure Description) is not available in View DICOM Info in the Local folder | WIP8.1 | DHF318051 Rev 13 HOST Verification Test Protocol |  | Subsystem | DICOM | NA | Fixed | ramesh_t | 0 | 2 | Test Design | Console4 | NO | Closed | WIP15SegReq | Brilliance 64 |  | DHF318051 Rev 13 HOST Verification Test Protocol | April 23, 2019 7:51:49 PM GMT+08:00
CTUDT00599447 | Refresh | Closed | roshan_r | August 8, 2018 9:46:28 PM GMT+08:00 | Refresh : DICOM tag (0032:1060 Requested Procedure Description) is not available in View DICOM Info in the Local folder | WIP8.1 | DHF318051 Rev 13 HOST Verification Test Protocol |  | Subsystem | DICOM | NA | Fixed | ramesh_t | 0 | 2 | Test Design | Console4 | NO | Closed | WIP15SegReq | Ingenuity 64 |  | DHF318051 Rev 13 HOST Verification Test Protocol | April 23, 2019 7:51:49 PM GMT+08:00
CTUDT00599447 | Refresh | Closed | roshan_r | August 8, 2018 9:46:28 PM GMT+08:00 | Refresh : DICOM tag (0032:1060 Requested Procedure Description) is not available in View DICOM Info in the Local folder | WIP8.1 | DHF318051 Rev 13 HOST Verification Test Protocol |  | Subsystem | DICOM | NA | Fixed | ramesh_t | 0 | 2 | Test Design | Console4 | NO | Closed | WIP15SegReq | IQon1 |  | DHF318051 Rev 13 HOST Verification Test Protocol | April 23, 2019 7:51:49 PM GMT+08:00
CTUDT00599445 | Refresh | Closed | g_di | August 8, 2018 5:28:34 PM GMT+08:00 | Refresh : Second Reference series was not displayed for second clinical scan in the acquisition tab. | WIP 8.1 | WIP11.1 |  | Subsystem | CCT - Series creation | Console-SRS.10232 | Fixed | hiremath_b | 0 | 4 | Test Design | Console4 | NO | Closed | WIP11SubCom | Ingenuity 64 |  | WIP11.1 | November 29, 2018 8:28:58 PM GMT+08:00
CTUDT00599433 | Refresh | Assigned | roshan_r | August 7, 2018 9:11:18 PM GMT+08:00 | Refresh :When a study is split with review before save option, the Console does not attach the thin slice series to all the studies produced by the split. | wip8.1 | WIP21 |  | Subsystem | SP2_Postpone | N/A |  | prasad_r | 0 | 2 | Console Software | Console4 | YES | Assigned | SP2 | Ovation |  | WIP21 | September 19, 2019 7:14:13 PM GMT+08:00
CTUDT00599433 | Refresh | Assigned | roshan_r | August 7, 2018 9:11:18 PM GMT+08:00 | Refresh :When a study is split with review before save option, the Console does not attach the thin slice series to all the studies produced by the split. | wip8.1 | WIP21 |  | Subsystem | SP2_Postpone | N/A |  | prasad_r | 0 | 2 | Console Software | Console4 | YES | Assigned | SP2 | Brilliance 64 |  | WIP21 | September 19, 2019 7:14:13 PM GMT+08:00
CTUDT00599433 | Refresh | Assigned | roshan_r | August 7, 2018 9:11:18 PM GMT+08:00 | Refresh :When a study is split with review before save option, the Console does not attach the thin slice series to all the studies produced by the split. | wip8.1 | WIP21 |  | Subsystem | SP2_Postpone | N/A |  | prasad_r | 0 | 2 | Console Software | Console4 | YES | Assigned | SP2 | Ingenuity 64 |  | WIP21 | September 19, 2019 7:14:13 PM GMT+08:00
CTUDT00599433 | Refresh | Assigned | roshan_r | August 7, 2018 9:11:18 PM GMT+08:00 | Refresh :When a study is split with review before save option, the Console does not attach the thin slice series to all the studies produced by the split. | wip8.1 | WIP21 |  | Subsystem | SP2_Postpone | N/A |  | prasad_r | 0 | 2 | Console Software | Console4 | YES | Assigned | SP2 | IQon1 |  | WIP21 | September 19, 2019 7:14:13 PM GMT+08:00
CTUDT00599429 | Refresh | Closed | sowmya_k | August 7, 2018 7:44:52 PM GMT+08:00 | If all Respiratory steps are removed from the series list and no Pulmo signal is detected and replaced by an Abdomen scan, still the Console displays the Resp Online Viewer. | 4.8 |  | User can manually minimise the Resp Viewer using the minimise button | Subsystem | Respiratory 4D CT | N/A | No Plan to Fix | bandil_j | 0 | 4 |  | Console4 | NO | Closed | WIP11SegReq | Brilliance Big Bore
CTUDT00599425 | Refresh | Closed | patra_r | August 7, 2018 6:53:05 PM GMT+08:00 | V4.1.3 & V4.1.6 When adjusting the countdown setting or when opening injection or voice balloon, the Exam Application crashed | 4.1.6 | BB_WIP2 | The user can restart the application, reselect the patient and use previous surview to recover. Using previous surview is discussed in the IFU. | Subsystem |  | Console-SRS.974, Console-SRS.2451 | Fixed | datta_i | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP1.1_SubCom | Brilliance Big Bore | CFR# 9418026 / PR# 6127996 PR5413542 | BB_WIP2 | October 29, 2018 1:50:30 PM GMT+08:00
CTUDT00599421 | Refresh | Closed | hiremath_b | August 7, 2018 5:46:11 PM GMT+08:00 | Refresh :Logon and Logoff events are not logged in Audit trail viewer as part of security events | WIP 8.1 | WIP11.1 |  | Subsystem |  | Console-SRS.15927 | Fixed | goswami_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP11SubCom | Ovation |  | WIP11.1 | November 8, 2018 3:06:15 PM GMT+08:00
CTUDT00599421 | Refresh | Closed | hiremath_b | August 7, 2018 5:46:11 PM GMT+08:00 | Refresh :Logon and Logoff events are not logged in Audit trail viewer as part of security events | WIP 8.1 | WIP11.1 |  | Subsystem |  | Console-SRS.15927 | Fixed | goswami_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP11SubCom | Brilliance 64 |  | WIP11.1 | November 8, 2018 3:06:15 PM GMT+08:00
CTUDT00599421 | Refresh | Closed | hiremath_b | August 7, 2018 5:46:11 PM GMT+08:00 | Refresh :Logon and Logoff events are not logged in Audit trail viewer as part of security events | WIP 8.1 | WIP11.1 |  | Subsystem |  | Console-SRS.15927 | Fixed | goswami_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP11SubCom | Ingenuity 64 |  | WIP11.1 | November 8, 2018 3:06:15 PM GMT+08:00
CTUDT00599421 | Refresh | Closed | hiremath_b | August 7, 2018 5:46:11 PM GMT+08:00 | Refresh :Logon and Logoff events are not logged in Audit trail viewer as part of security events | WIP 8.1 | WIP11.1 |  | Subsystem |  | Console-SRS.15927 | Fixed | goswami_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP11SubCom | IQon1 |  | WIP11.1 | November 8, 2018 3:06:15 PM GMT+08:00
CTUDT00599419 | Refresh | Closed | hiremath_b | August 7, 2018 5:26:52 PM GMT+08:00 | Refresh : Non Clinical user (IT admin) is able to launch clinical workflow. Duplicate username is created for 2 different roles. | WIP 8.1 | WIP10 | NA | Subsystem |  | NA | Fixed | ravindra_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP9.2SubCom | Ovation |  | WIP10 | October 8, 2018 6:51:42 PM GMT+08:00
CTUDT00599419 | Refresh | Closed | hiremath_b | August 7, 2018 5:26:52 PM GMT+08:00 | Refresh : Non Clinical user (IT admin) is able to launch clinical workflow. Duplicate username is created for 2 different roles. | WIP 8.1 | WIP10 | NA | Subsystem |  | NA | Fixed | ravindra_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP9.2SubCom | Brilliance 64 |  | WIP10 | October 8, 2018 6:51:42 PM GMT+08:00
CTUDT00599419 | Refresh | Closed | hiremath_b | August 7, 2018 5:26:52 PM GMT+08:00 | Refresh : Non Clinical user (IT admin) is able to launch clinical workflow. Duplicate username is created for 2 different roles. | WIP 8.1 | WIP10 | NA | Subsystem |  | NA | Fixed | ravindra_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP9.2SubCom | Brilliance Big Bore |  | WIP10 | October 8, 2018 6:51:42 PM GMT+08:00
CTUDT00599419 | Refresh | Closed | hiremath_b | August 7, 2018 5:26:52 PM GMT+08:00 | Refresh : Non Clinical user (IT admin) is able to launch clinical workflow. Duplicate username is created for 2 different roles. | WIP 8.1 | WIP10 | NA | Subsystem |  | NA | Fixed | ravindra_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP9.2SubCom | Ingenuity 64 |  | WIP10 | October 8, 2018 6:51:42 PM GMT+08:00
CTUDT00599419 | Refresh | Closed | hiremath_b | August 7, 2018 5:26:52 PM GMT+08:00 | Refresh : Non Clinical user (IT admin) is able to launch clinical workflow. Duplicate username is created for 2 different roles. | WIP 8.1 | WIP10 | NA | Subsystem |  | NA | Fixed | ravindra_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP9.2SubCom | IQon1 |  | WIP10 | October 8, 2018 6:51:42 PM GMT+08:00
CTUDT00599416 | Refresh | Closed | nanjunda_s | August 7, 2018 2:46:37 PM GMT+08:00 | Length is different between the results when changed start/end position less than 15mm | WIP 8.1 | WIP 11 |  | Subsystem | Plan - Scan Length | NA | Fixed | jindal_k | 0 | 2 | Test Design | Console4 | NO | Closed | WIP11SegReq | Ingenuity 64 |  | WIP 11 | December 5, 2018 3:16:02 PM GMT+08:00
CTUDT00599410 | Refresh | Closed | kadish_n | August 6, 2018 9:44:53 PM GMT+08:00 | In preferences the Infant Phantom switch is not on by default | WIP 8.1 | WIP15 |  | Subsystem | Perefrences - Infant Phantom | Console-SRS.15388 | Fixed | jose_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SegReq | Ovation |  | WIP15 | March 21, 2019 7:09:53 PM GMT+08:00
CTUDT00599407 | Refresh | Closed | nimma_p | August 6, 2018 9:11:45 PM GMT+08:00 | User could not able to enlarge the “Patient Name” section in Recon Tab under the “Recon Queue Manager”. | WIP8.1 |  |  | Subsystem | Not reproduceable | NA | Works As Specified / Not A Defect | babu_s | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00599405 | Refresh | Closed | fitusi_r | August 6, 2018 8:38:55 PM GMT+08:00 | Refresh, IQon-Rendering axial-MPR issue | WIP 8.1 | v4.1.10 | to rebuild MPR with different parameters | Subsystem | IFU | NA | Fixed | charles_j | 0 | 3 | User Documentation | Support-System | YES | Closed | 6-Dec-2019 | Brilliance 64 |  | v4.1.10 | December 9, 2019 10:29:59 AM GMT+08:00
CTUDT00599405 | Refresh | Closed | fitusi_r | August 6, 2018 8:38:55 PM GMT+08:00 | Refresh, IQon-Rendering axial-MPR issue | WIP 8.1 | v4.1.10 | to rebuild MPR with different parameters | Subsystem | IFU | NA | Fixed | charles_j | 0 | 3 | User Documentation | Support-System | YES | Closed | 6-Dec-2019 | Brilliance Big Bore |  | v4.1.10 | December 9, 2019 10:29:59 AM GMT+08:00
CTUDT00599405 | Refresh | Closed | fitusi_r | August 6, 2018 8:38:55 PM GMT+08:00 | Refresh, IQon-Rendering axial-MPR issue | WIP 8.1 | v4.1.10 | to rebuild MPR with different parameters | Subsystem | IFU | NA | Fixed | charles_j | 0 | 3 | User Documentation | Support-System | YES | Closed | 6-Dec-2019 | Ingenuity 64 |  | v4.1.10 | December 9, 2019 10:29:59 AM GMT+08:00
CTUDT00599405 | Refresh | Closed | fitusi_r | August 6, 2018 8:38:55 PM GMT+08:00 | Refresh, IQon-Rendering axial-MPR issue | WIP 8.1 | v4.1.10 | to rebuild MPR with different parameters | Subsystem | IFU | NA | Fixed | charles_j | 0 | 3 | User Documentation | Support-System | YES | Closed | 6-Dec-2019 | iCT2 |  | v4.1.10 | December 9, 2019 10:29:59 AM GMT+08:00
CTUDT00599405 | Refresh | Closed | fitusi_r | August 6, 2018 8:38:55 PM GMT+08:00 | Refresh, IQon-Rendering axial-MPR issue | WIP 8.1 | v4.1.10 | to rebuild MPR with different parameters | Subsystem | IFU | NA | Fixed | charles_j | 0 | 3 | User Documentation | Support-System | YES | Closed | 6-Dec-2019 | IQon1 |  | v4.1.10 | December 9, 2019 10:29:59 AM GMT+08:00
CTUDT00599387 | Refresh | Closed | attali_s | August 5, 2018 9:01:34 PM GMT+08:00 | Stability Test - No warning window displayed after cancelling Stability | Wip 8.1 | WIP14 |  | Subsystem | Stability test | n/a | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP14SegReq | Ovation |  | WIP14 | February 7, 2019 3:20:25 PM GMT+08:00
CTUDT00599387 | Refresh | Closed | attali_s | August 5, 2018 9:01:34 PM GMT+08:00 | Stability Test - No warning window displayed after cancelling Stability | Wip 8.1 | WIP14 |  | Subsystem | Stability test | n/a | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP14SegReq | Ingenuity 64 |  | WIP14 | February 7, 2019 3:20:25 PM GMT+08:00
CTUDT00599387 | Refresh | Closed | attali_s | August 5, 2018 9:01:34 PM GMT+08:00 | Stability Test - No warning window displayed after cancelling Stability | Wip 8.1 | WIP14 |  | Subsystem | Stability test | n/a | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP14SegReq | IQon1 |  | WIP14 | February 7, 2019 3:20:25 PM GMT+08:00
CTUDT00599386 | Refresh | Closed | sowmya_k | August 3, 2018 7:51:26 PM GMT+08:00 | v4.1.10 : delay between two timed axial scans (end of the first scan to start of the 2nd) is NOT 1000 seconds.(REQ 5146) | 4.8 | BB_WIP1.1 |  | Subsystem |  | NA | Fixed | kumar_v | 0 | 2 | Console Software | Console4 | YES | Closed | WIP9.1SegReq | Brilliance Big Bore |  | BB_WIP1.1 | August 17, 2018 5:51:14 PM GMT+08:00
CTUDT00599385 | Refresh | Closed | sowmya_k | August 3, 2018 7:49:56 PM GMT+08:00 | Ok button is not visible in Exam card manager | 4.8 | BB_WIP1.1 |  | Subsystem |  | NA | Fixed | kumar_v | 0 | 3 | Console Software | Console4 | NO | Closed | WIP9.1SegReq | Brilliance Big Bore |  | BB_WIP1.1 | August 17, 2018 5:39:35 PM GMT+08:00
CTUDT00599379 | Refresh | Closed | patra_r | August 3, 2018 2:40:47 PM GMT+08:00 | Tooltip is missing on the perfusion scans once the scan is initiated and completed. | 4.8 | BB_WIP1 |  | Subsystem |  | N/A | Fixed | navale_s | 0 | 3 | Console Software | Console4 | YES | Closed | WIP1SegReq | Brilliance Big Bore |  | BB_WIP1 | August 8, 2018 8:02:03 PM GMT+08:00
CTUDT00599368 | Refresh | Closed | bommiReddy_r | August 2, 2018 8:35:01 PM GMT+08:00 | Dose Notifications are not appeared in Exam summary and Dose Log viewer for the step & Shoot exam card | WIP8.1 | WIP11.1 |  | Subsystem | Exam Summary - Dose Notification | NA | Fixed | hiremath_b | 0 | 2 | Test Design | Console4 | NO | Closed | WIP9.1SegReq | Ingenuity 64 |  | WIP11.1 | November 30, 2018 5:35:48 PM GMT+08:00
CTUDT00599351 | Refresh | Closed | wang_l | August 2, 2018 10:26:33 AM GMT+08:00 | Sync Right option is checked by default in Preference | WIP8.1 | WIP13 |  | Subsystem | Preferences - SyncRight | N/A | Fixed | chakraborty_d | 0 | 2 | Console Software | Console4 | NO | Closed | WIP13SubCom | Brilliance 64 |  | WIP13 | January 14, 2019 5:42:06 PM GMT+08:00
CTUDT00599351 | Refresh | Closed | wang_l | August 2, 2018 10:26:33 AM GMT+08:00 | Sync Right option is checked by default in Preference | WIP8.1 | WIP13 |  | Subsystem | Preferences - SyncRight | N/A | Fixed | chakraborty_d | 0 | 2 | Console Software | Console4 | NO | Closed | WIP13SubCom | Ingenuity 64 |  | WIP13 | January 14, 2019 5:42:06 PM GMT+08:00
CTUDT00599348 | Refresh | Closed | shaham_t | August 1, 2018 10:30:23 PM GMT+08:00 | Paste option not available in series list for completed scan result | WIP6.1 | Host SSRS DHF303211 Rev 04 | NA | Subsystem |  | Console-SRS.2294 | Fixed | bandil_j | 0 | 4 | Host SSRS | Console4 | NO | Closed | DHF303211 Rev04 | Ovation |  | Host SSRS DHF303211 Rev 04 | September 18, 2018 6:48:13 PM GMT+08:00
CTUDT00599348 | Refresh | Closed | shaham_t | August 1, 2018 10:30:23 PM GMT+08:00 | Paste option not available in series list for completed scan result | WIP6.1 | Host SSRS DHF303211 Rev 04 | NA | Subsystem |  | Console-SRS.2294 | Fixed | bandil_j | 0 | 4 | Host SSRS | Console4 | NO | Closed | DHF303211 Rev04 | Brilliance 64 |  | Host SSRS DHF303211 Rev 04 | September 18, 2018 6:48:13 PM GMT+08:00
CTUDT00599348 | Refresh | Closed | shaham_t | August 1, 2018 10:30:23 PM GMT+08:00 | Paste option not available in series list for completed scan result | WIP6.1 | Host SSRS DHF303211 Rev 04 | NA | Subsystem |  | Console-SRS.2294 | Fixed | bandil_j | 0 | 4 | Host SSRS | Console4 | NO | Closed | DHF303211 Rev04 | Ingenuity 64 |  | Host SSRS DHF303211 Rev 04 | September 18, 2018 6:48:13 PM GMT+08:00
CTUDT00599348 | Refresh | Closed | shaham_t | August 1, 2018 10:30:23 PM GMT+08:00 | Paste option not available in series list for completed scan result | WIP6.1 | Host SSRS DHF303211 Rev 04 | NA | Subsystem |  | Console-SRS.2294 | Fixed | bandil_j | 0 | 4 | Host SSRS | Console4 | NO | Closed | DHF303211 Rev04 | IQon1 |  | Host SSRS DHF303211 Rev 04 | September 18, 2018 6:48:13 PM GMT+08:00
CTUDT00599346 | Refresh | Closed | shaham_t | August 1, 2018 10:24:14 PM GMT+08:00 | Delete Option not available for paused scan result | WIP6.1 | Host SSRS DHF303211 Rev04 | NA | Subsystem |  | Console-SRS.2294 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | DHF303211 Rev04 | Ovation |  | Host SSRS DHF303211 Rev04 | September 18, 2018 6:52:48 PM GMT+08:00
CTUDT00599346 | Refresh | Closed | shaham_t | August 1, 2018 10:24:14 PM GMT+08:00 | Delete Option not available for paused scan result | WIP6.1 | Host SSRS DHF303211 Rev04 | NA | Subsystem |  | Console-SRS.2294 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | DHF303211 Rev04 | Brilliance 64 |  | Host SSRS DHF303211 Rev04 | September 18, 2018 6:52:48 PM GMT+08:00
CTUDT00599346 | Refresh | Closed | shaham_t | August 1, 2018 10:24:14 PM GMT+08:00 | Delete Option not available for paused scan result | WIP6.1 | Host SSRS DHF303211 Rev04 | NA | Subsystem |  | Console-SRS.2294 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | DHF303211 Rev04 | Ingenuity 64 |  | Host SSRS DHF303211 Rev04 | September 18, 2018 6:52:48 PM GMT+08:00
CTUDT00599346 | Refresh | Closed | shaham_t | August 1, 2018 10:24:14 PM GMT+08:00 | Delete Option not available for paused scan result | WIP6.1 | Host SSRS DHF303211 Rev04 | NA | Subsystem |  | Console-SRS.2294 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | DHF303211 Rev04 | IQon1 |  | Host SSRS DHF303211 Rev04 | September 18, 2018 6:52:48 PM GMT+08:00
CTUDT00599344 | Refresh | Closed | azagury_d | August 1, 2018 9:03:01 PM GMT+08:00 | Refresh: Calibrations: Message warning about Air Bubble in phantom not translated. | WIP 8.1 | WIP11 |  | Subsystem | Verified1 | N/A | Fixed | trushin_d | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP10SegReq | Ovation |  | WIP11 | October 22, 2018 8:46:12 PM GMT+08:00
CTUDT00599344 | Refresh | Closed | azagury_d | August 1, 2018 9:03:01 PM GMT+08:00 | Refresh: Calibrations: Message warning about Air Bubble in phantom not translated. | WIP 8.1 | WIP11 |  | Subsystem | Verified1 | N/A | Fixed | trushin_d | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP10SegReq | Brilliance 64 |  | WIP11 | October 22, 2018 8:46:12 PM GMT+08:00
CTUDT00599344 | Refresh | Closed | azagury_d | August 1, 2018 9:03:01 PM GMT+08:00 | Refresh: Calibrations: Message warning about Air Bubble in phantom not translated. | WIP 8.1 | WIP11 |  | Subsystem | Verified1 | N/A | Fixed | trushin_d | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP10SegReq | Brilliance Big Bore |  | WIP11 | October 22, 2018 8:46:12 PM GMT+08:00
CTUDT00599344 | Refresh | Closed | azagury_d | August 1, 2018 9:03:01 PM GMT+08:00 | Refresh: Calibrations: Message warning about Air Bubble in phantom not translated. | WIP 8.1 | WIP11 |  | Subsystem | Verified1 | N/A | Fixed | trushin_d | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP10SegReq | Ingenuity 64 |  | WIP11 | October 22, 2018 8:46:12 PM GMT+08:00
CTUDT00599344 | Refresh | Closed | azagury_d | August 1, 2018 9:03:01 PM GMT+08:00 | Refresh: Calibrations: Message warning about Air Bubble in phantom not translated. | WIP 8.1 | WIP11 |  | Subsystem | Verified1 | N/A | Fixed | trushin_d | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP10SegReq | IQon1 |  | WIP11 | October 22, 2018 8:46:12 PM GMT+08:00
CTUDT00599341 | Refresh | Closed | kadish_n | August 1, 2018 7:12:48 PM GMT+08:00 | Crosstalk test fails on results of known BDs | WIP8.1 | WIP9 |  | Subsystem |  | N/A | Fixed | yosefk | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP10SegReq | Ovation |  | WIP9 | December 17, 2018 2:51:21 PM GMT+08:00
CTUDT00599341 | Refresh | Closed | kadish_n | August 1, 2018 7:12:48 PM GMT+08:00 | Crosstalk test fails on results of known BDs | WIP8.1 | WIP9 |  | Subsystem |  | N/A | Fixed | yosefk | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP10SegReq | IQon1 |  | WIP9 | December 17, 2018 2:51:21 PM GMT+08:00
CTUDT00599339 | Refresh | Closed | hiremath_b | August 1, 2018 5:55:17 PM GMT+08:00 | Refresh : Cant login as Philips Service user with IST and also user accounts are not enforced to change the password for the 1st time login after accounts restore. | WIP 8.1 | WIP12 | No workaround | Subsystem | User management-User login | NA | Fixed | ravindra_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP12SubCom | Brilliance 64 |  | WIP12 | November 27, 2018 5:44:30 PM GMT+08:00
CTUDT00599335 | Refresh | Closed | krisher_i | August 1, 2018 1:44:22 PM GMT+08:00 | Refresh - Message box appears on unchecking of Dose Right | WIP8.1 |  |  | Subsystem |  | N/A | Works As Specified / Not A Defect | kaushik_u | 0 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00599335 | Refresh | Closed | krisher_i | August 1, 2018 1:44:22 PM GMT+08:00 | Refresh - Message box appears on unchecking of Dose Right | WIP8.1 |  |  | Subsystem |  | N/A | Works As Specified / Not A Defect | kaushik_u | 0 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00599334 | Refresh | Closed | nanjunda_s | August 1, 2018 1:39:13 PM GMT+08:00 | #0095 Message displayed in the console and in the requirement is not same. | WIP8.1 | Host SSRS DHF303211 Rev 08 |  | Subsystem | Documentation | NA | Fixed | bandil_j | 0 | 4 | Host SSRS | Console4 | NO | Closed | WIP13SubCom | Ovation |  | Host SSRS DHF303211 Rev 08 | March 21, 2019 1:05:48 PM GMT+08:00
CTUDT00599334 | Refresh | Closed | nanjunda_s | August 1, 2018 1:39:13 PM GMT+08:00 | #0095 Message displayed in the console and in the requirement is not same. | WIP8.1 | Host SSRS DHF303211 Rev 08 |  | Subsystem | Documentation | NA | Fixed | bandil_j | 0 | 4 | Host SSRS | Console4 | NO | Closed | WIP13SubCom | Ingenuity 64 |  | Host SSRS DHF303211 Rev 08 | March 21, 2019 1:05:48 PM GMT+08:00
CTUDT00599334 | Refresh | Closed | nanjunda_s | August 1, 2018 1:39:13 PM GMT+08:00 | #0095 Message displayed in the console and in the requirement is not same. | WIP8.1 | Host SSRS DHF303211 Rev 08 |  | Subsystem | Documentation | NA | Fixed | bandil_j | 0 | 4 | Host SSRS | Console4 | NO | Closed | WIP13SubCom | IQon1 |  | Host SSRS DHF303211 Rev 08 | March 21, 2019 1:05:48 PM GMT+08:00
CTUDT00599324 | Refresh | Closed | wang_l | July 31, 2018 10:49:50 AM GMT+08:00 | Error message displayed once scan for EC Resp 4D CT without surview | WIP8.1 | WIP10 |  | Subsystem |  | None | Fixed | datta_i | 0 | 4 | Console Software | Console4 | NO | Closed | WIP9.1SegReq | Ovation |  | WIP10 | September 3, 2018 4:54:48 PM GMT+08:00
CTUDT00599324 | Refresh | Closed | wang_l | July 31, 2018 10:49:50 AM GMT+08:00 | Error message displayed once scan for EC Resp 4D CT without surview | WIP8.1 | WIP10 |  | Subsystem |  | None | Fixed | datta_i | 0 | 4 | Console Software | Console4 | NO | Closed | WIP9.1SegReq | Brilliance 64 |  | WIP10 | September 3, 2018 4:54:48 PM GMT+08:00
CTUDT00599324 | Refresh | Closed | wang_l | July 31, 2018 10:49:50 AM GMT+08:00 | Error message displayed once scan for EC Resp 4D CT without surview | WIP8.1 | WIP10 |  | Subsystem |  | None | Fixed | datta_i | 0 | 4 | Console Software | Console4 | NO | Closed | WIP9.1SegReq | Ingenuity 64 |  | WIP10 | September 3, 2018 4:54:48 PM GMT+08:00
CTUDT00599324 | Refresh | Closed | wang_l | July 31, 2018 10:49:50 AM GMT+08:00 | Error message displayed once scan for EC Resp 4D CT without surview | WIP8.1 | WIP10 |  | Subsystem |  | None | Fixed | datta_i | 0 | 4 | Console Software | Console4 | NO | Closed | WIP9.1SegReq | IQon1 |  | WIP10 | September 3, 2018 4:54:48 PM GMT+08:00
CTUDT00599322 | Refresh | Closed | wang_l | July 31, 2018 10:26:34 AM GMT+08:00 | Script Error message prompted once click "Easy Mode" in Service Tool>Diagnostics>Tube>Manual Exposure | WIP8.1 | WIP13 |  | Subsystem | Service Tools - Manual exposure | N/A | Fixed | Prakash_Chand_Yadav | 0 | 4 | Service Tools_Software | Service Tools | NO | Closed | WIP13SubCom | Ovation |  | WIP13 | January 11, 2019 2:34:01 PM GMT+08:00
CTUDT00599322 | Refresh | Closed | wang_l | July 31, 2018 10:26:34 AM GMT+08:00 | Script Error message prompted once click "Easy Mode" in Service Tool>Diagnostics>Tube>Manual Exposure | WIP8.1 | WIP13 |  | Subsystem | Service Tools - Manual exposure | N/A | Fixed | Prakash_Chand_Yadav | 0 | 4 | Service Tools_Software | Service Tools | NO | Closed | WIP13SubCom | Brilliance 64 |  | WIP13 | January 11, 2019 2:34:01 PM GMT+08:00
CTUDT00599322 | Refresh | Closed | wang_l | July 31, 2018 10:26:34 AM GMT+08:00 | Script Error message prompted once click "Easy Mode" in Service Tool>Diagnostics>Tube>Manual Exposure | WIP8.1 | WIP13 |  | Subsystem | Service Tools - Manual exposure | N/A | Fixed | Prakash_Chand_Yadav | 0 | 4 | Service Tools_Software | Service Tools | NO | Closed | WIP13SubCom | Brilliance Big Bore |  | WIP13 | January 11, 2019 2:34:01 PM GMT+08:00
CTUDT00599322 | Refresh | Closed | wang_l | July 31, 2018 10:26:34 AM GMT+08:00 | Script Error message prompted once click "Easy Mode" in Service Tool>Diagnostics>Tube>Manual Exposure | WIP8.1 | WIP13 |  | Subsystem | Service Tools - Manual exposure | N/A | Fixed | Prakash_Chand_Yadav | 0 | 4 | Service Tools_Software | Service Tools | NO | Closed | WIP13SubCom | Ingenuity 64 |  | WIP13 | January 11, 2019 2:34:01 PM GMT+08:00
CTUDT00599322 | Refresh | Closed | wang_l | July 31, 2018 10:26:34 AM GMT+08:00 | Script Error message prompted once click "Easy Mode" in Service Tool>Diagnostics>Tube>Manual Exposure | WIP8.1 | WIP13 |  | Subsystem | Service Tools - Manual exposure | N/A | Fixed | Prakash_Chand_Yadav | 0 | 4 | Service Tools_Software | Service Tools | NO | Closed | WIP13SubCom | IQon1 |  | WIP13 | January 11, 2019 2:34:01 PM GMT+08:00
CTUDT00599320 | Refresh | Closed | wang_l | July 31, 2018 10:06:17 AM GMT+08:00 | Service Tool crashed once click View Logs With Markers or click checkbox of items in Can Trace | WIP8.1 | WIP13 |  | Subsystem | Service Tools - Can Trace | N/A | Fixed | shreekanth_manniyath | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP13SubCom | Ingenuity 64 |  | WIP13 | January 30, 2019 7:52:29 PM GMT+08:00
CTUDT00599317 | Refresh | Closed | kadish_n | July 30, 2018 9:18:21 PM GMT+08:00 | Shadow Test: Plots' titles do not include the layer(s) | WIP8.1 |  |  | Subsystem |  |  | Duplicate | gehtin_e | 0 |  |  |  |  | Closed |  | IQon1
CTUDT00599313 | Refresh | Closed | gT_s | July 30, 2018 7:20:03 PM GMT+08:00 | Refresh : The application is getting crashed when the username is entered in the prompt message box and click on continue after the 5th exam for the emergency user | WIP8.1 | WIP11 |  | Subsystem |  | None | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP11SegReq | Ovation |  | WIP11 | November 12, 2018 6:05:35 PM GMT+08:00
CTUDT00599313 | Refresh | Closed | gT_s | July 30, 2018 7:20:03 PM GMT+08:00 | Refresh : The application is getting crashed when the username is entered in the prompt message box and click on continue after the 5th exam for the emergency user | WIP8.1 | WIP11 |  | Subsystem |  | None | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP11SegReq | Brilliance 64 |  | WIP11 | November 12, 2018 6:05:35 PM GMT+08:00
CTUDT00599313 | Refresh | Closed | gT_s | July 30, 2018 7:20:03 PM GMT+08:00 | Refresh : The application is getting crashed when the username is entered in the prompt message box and click on continue after the 5th exam for the emergency user | WIP8.1 | WIP11 |  | Subsystem |  | None | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP11SegReq | Brilliance Big Bore |  | WIP11 | November 12, 2018 6:05:35 PM GMT+08:00
CTUDT00599313 | Refresh | Closed | gT_s | July 30, 2018 7:20:03 PM GMT+08:00 | Refresh : The application is getting crashed when the username is entered in the prompt message box and click on continue after the 5th exam for the emergency user | WIP8.1 | WIP11 |  | Subsystem |  | None | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP11SegReq | Ingenuity 64 |  | WIP11 | November 12, 2018 6:05:35 PM GMT+08:00
CTUDT00599313 | Refresh | Closed | gT_s | July 30, 2018 7:20:03 PM GMT+08:00 | Refresh : The application is getting crashed when the username is entered in the prompt message box and click on continue after the 5th exam for the emergency user | WIP8.1 | WIP11 |  | Subsystem |  | None | Fixed | s_sa | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP11SegReq | IQon1 |  | WIP11 | November 12, 2018 6:05:35 PM GMT+08:00
CTUDT00599310 | Refresh | Closed | hiremath_b | July 30, 2018 7:14:23 PM GMT+08:00 | Refresh: Incorrect message is thrown to the user when user whos account is disabled try to login to the host. | WIP 8.1 | WIP10 | N/A | Subsystem |  | N/A | Fixed | singh_j | 0 | 4 | Console Software | Console4 | NO | Closed | WIP11SegReq | Ovation |  | WIP10 | October 23, 2018 1:51:41 PM GMT+08:00
CTUDT00599310 | Refresh | Closed | hiremath_b | July 30, 2018 7:14:23 PM GMT+08:00 | Refresh: Incorrect message is thrown to the user when user whos account is disabled try to login to the host. | WIP 8.1 | WIP10 | N/A | Subsystem |  | N/A | Fixed | singh_j | 0 | 4 | Console Software | Console4 | NO | Closed | WIP11SegReq | Brilliance 64 |  | WIP10 | October 23, 2018 1:51:41 PM GMT+08:00
CTUDT00599310 | Refresh | Closed | hiremath_b | July 30, 2018 7:14:23 PM GMT+08:00 | Refresh: Incorrect message is thrown to the user when user whos account is disabled try to login to the host. | WIP 8.1 | WIP10 | N/A | Subsystem |  | N/A | Fixed | singh_j | 0 | 4 | Console Software | Console4 | NO | Closed | WIP11SegReq | Ingenuity 64 |  | WIP10 | October 23, 2018 1:51:41 PM GMT+08:00
CTUDT00599310 | Refresh | Closed | hiremath_b | July 30, 2018 7:14:23 PM GMT+08:00 | Refresh: Incorrect message is thrown to the user when user whos account is disabled try to login to the host. | WIP 8.1 | WIP10 | N/A | Subsystem |  | N/A | Fixed | singh_j | 0 | 4 | Console Software | Console4 | NO | Closed | WIP11SegReq | IQon1 |  | WIP10 | October 23, 2018 1:51:41 PM GMT+08:00
CTUDT00599305 | Refresh | Closed | wang_l | July 30, 2018 9:58:57 AM GMT+08:00 | Keyswitch prompt when tilt ganry from -1.5 to 0 | WIP8.1 |  |  | Subsystem | HW Reliability |  | System Repaired | wang_k | 1 |  |  | GCX | NO | Closed |  | Ingenuity 64
CTUDT00599295 | Refresh | Closed | vinoliya_j | July 27, 2018 7:56:16 PM GMT+08:00 | during a tilted acquisition, the plan box moves up when changing the scan FOV | WIP 8.1 | WIP12 |  | Subsystem | Plan - Plan box moving | N/A | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | YES | Closed | WIP12SubCom | Brilliance 64 |  | WIP12 | December 5, 2018 5:06:04 PM GMT+08:00
CTUDT00599295 | Refresh | Closed | vinoliya_j | July 27, 2018 7:56:16 PM GMT+08:00 | during a tilted acquisition, the plan box moves up when changing the scan FOV | WIP 8.1 | WIP12 |  | Subsystem | Plan - Plan box moving | N/A | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | YES | Closed | WIP12SubCom | Ingenuity 64 |  | WIP12 | December 5, 2018 5:06:04 PM GMT+08:00
CTUDT00599293 | Refresh | Closed | vinoliya_j | July 27, 2018 7:18:27 PM GMT+08:00 | Start learning option is displayed for all Pulmo EC's | WIP 8.1 | WIP13 |  | Subsystem | Start Learning option | NA | Fixed | chokkalingam_r | 0 | 4 | Console Software | Console4 | NO | Closed | WIP13SubCom | Brilliance 64 |  | WIP13 | January 10, 2019 12:13:10 PM GMT+08:00
CTUDT00599293 | Refresh | Closed | vinoliya_j | July 27, 2018 7:18:27 PM GMT+08:00 | Start learning option is displayed for all Pulmo EC's | WIP 8.1 | WIP13 |  | Subsystem | Start Learning option | NA | Fixed | chokkalingam_r | 0 | 4 | Console Software | Console4 | NO | Closed | WIP13SubCom | Ingenuity 64 |  | WIP13 | January 10, 2019 12:13:10 PM GMT+08:00
CTUDT00599291 | Refresh | Closed | nimma_p | July 27, 2018 5:29:19 PM GMT+08:00 | Series description is not available for IMR results in RQM. | WIP8.1 | HOST Verification Test Protocol DHF318051 Rev 09 |  | Subsystem | RQM - IMR Series description | N/A | Fixed | shekhar_p | 0 | 2 | Test Design | Console4 | NO | Closed | WIP13SegReq | Ingenuity 64 |  | HOST Verification Test Protocol DHF318051 Rev 09 | February 13, 2019 5:07:51 PM GMT+08:00
CTUDT00599285 | Refresh | Closed | gT_s | July 26, 2018 7:09:52 PM GMT+08:00 | Refresh :Operator name is not displayed fully as Emergency_OperatorName,it is getting truncated when the username is large number of characters | WIP8.1 | WIP10 | Type emergency user name with less than 53 characters | Subsystem |  | N/A | Fixed | singh_j | 0 | 3 | Console Software | Console4 | NO | Closed | WIP11SegReq | Ovation |  | WIP10 | October 11, 2018 6:04:12 PM GMT+08:00
CTUDT00599285 | Refresh | Closed | gT_s | July 26, 2018 7:09:52 PM GMT+08:00 | Refresh :Operator name is not displayed fully as Emergency_OperatorName,it is getting truncated when the username is large number of characters | WIP8.1 | WIP10 | Type emergency user name with less than 53 characters | Subsystem |  | N/A | Fixed | singh_j | 0 | 3 | Console Software | Console4 | NO | Closed | WIP11SegReq | Brilliance 64 |  | WIP10 | October 11, 2018 6:04:12 PM GMT+08:00
CTUDT00599285 | Refresh | Closed | gT_s | July 26, 2018 7:09:52 PM GMT+08:00 | Refresh :Operator name is not displayed fully as Emergency_OperatorName,it is getting truncated when the username is large number of characters | WIP8.1 | WIP10 | Type emergency user name with less than 53 characters | Subsystem |  | N/A | Fixed | singh_j | 0 | 3 | Console Software | Console4 | NO | Closed | WIP11SegReq | Brilliance Big Bore |  | WIP10 | October 11, 2018 6:04:12 PM GMT+08:00
CTUDT00599285 | Refresh | Closed | gT_s | July 26, 2018 7:09:52 PM GMT+08:00 | Refresh :Operator name is not displayed fully as Emergency_OperatorName,it is getting truncated when the username is large number of characters | WIP8.1 | WIP10 | Type emergency user name with less than 53 characters | Subsystem |  | N/A | Fixed | singh_j | 0 | 3 | Console Software | Console4 | NO | Closed | WIP11SegReq | Ingenuity 64 |  | WIP10 | October 11, 2018 6:04:12 PM GMT+08:00
CTUDT00599285 | Refresh | Closed | gT_s | July 26, 2018 7:09:52 PM GMT+08:00 | Refresh :Operator name is not displayed fully as Emergency_OperatorName,it is getting truncated when the username is large number of characters | WIP8.1 | WIP10 | Type emergency user name with less than 53 characters | Subsystem |  | N/A | Fixed | singh_j | 0 | 3 | Console Software | Console4 | NO | Closed | WIP11SegReq | IQon1 |  | WIP10 | October 11, 2018 6:04:12 PM GMT+08:00
CTUDT00599280 | Refresh | Closed | wang_l | July 26, 2018 1:44:36 PM GMT+08:00 | Edit result with Slab Viewer with errors during recon | WIP8.1 |  |  | Subsystem | Brain Helical - Performance | N/A | Works As Specified / Not A Defect | Prakash_Chand_Yadav | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00599274 | Refresh | Closed | hiremath_b | July 25, 2018 7:41:04 PM GMT+08:00 | Refresh:Not able to launch some of O-level service tools like Date & time and Audjust audio properties. | WIP 8.1 | WIP14 |  | Subsystem | Service Tools_Olevel | N/A | Fixed | Prabhakar_Priyank_Muppidi | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP14SubCom | Ovation |  | WIP14 | May 23, 2019 2:42:52 PM GMT+08:00
CTUDT00599274 | Refresh | Closed | hiremath_b | July 25, 2018 7:41:04 PM GMT+08:00 | Refresh:Not able to launch some of O-level service tools like Date & time and Audjust audio properties. | WIP 8.1 | WIP14 |  | Subsystem | Service Tools_Olevel | N/A | Fixed | Prabhakar_Priyank_Muppidi | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP14SubCom | Brilliance 64 |  | WIP14 | May 23, 2019 2:42:52 PM GMT+08:00
CTUDT00599274 | Refresh | Closed | hiremath_b | July 25, 2018 7:41:04 PM GMT+08:00 | Refresh:Not able to launch some of O-level service tools like Date & time and Audjust audio properties. | WIP 8.1 | WIP14 |  | Subsystem | Service Tools_Olevel | N/A | Fixed | Prabhakar_Priyank_Muppidi | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP14SubCom | Ingenuity 64 |  | WIP14 | May 23, 2019 2:42:52 PM GMT+08:00
CTUDT00599274 | Refresh | Closed | hiremath_b | July 25, 2018 7:41:04 PM GMT+08:00 | Refresh:Not able to launch some of O-level service tools like Date & time and Audjust audio properties. | WIP 8.1 | WIP14 |  | Subsystem | Service Tools_Olevel | N/A | Fixed | Prabhakar_Priyank_Muppidi | 0 | 2 | Console Software | Service Tools | NO | Closed | WIP14SubCom | IQon1 |  | WIP14 | May 23, 2019 2:42:52 PM GMT+08:00
CTUDT00599272 | Refresh | Closed | hiremath_b | July 25, 2018 5:49:45 PM GMT+08:00 | Refresh : User account gets locked at random no of attempts but not as per password policy settings. | WIP 8.1 | WIP10 |  | Subsystem |  | None | Fixed | singh_j | 0 | 2 | Console Software | Console4 | NO | Closed | WIP10SegReq | Ovation |  | WIP10 | October 11, 2018 4:54:48 PM GMT+08:00
CTUDT00599272 | Refresh | Closed | hiremath_b | July 25, 2018 5:49:45 PM GMT+08:00 | Refresh : User account gets locked at random no of attempts but not as per password policy settings. | WIP 8.1 | WIP10 |  | Subsystem |  | None | Fixed | singh_j | 0 | 2 | Console Software | Console4 | NO | Closed | WIP10SegReq | Brilliance 64 |  | WIP10 | October 11, 2018 4:54:48 PM GMT+08:00
CTUDT00599272 | Refresh | Closed | hiremath_b | July 25, 2018 5:49:45 PM GMT+08:00 | Refresh : User account gets locked at random no of attempts but not as per password policy settings. | WIP 8.1 | WIP10 |  | Subsystem |  | None | Fixed | singh_j | 0 | 2 | Console Software | Console4 | NO | Closed | WIP10SegReq | Ingenuity 64 |  | WIP10 | October 11, 2018 4:54:48 PM GMT+08:00
CTUDT00599272 | Refresh | Closed | hiremath_b | July 25, 2018 5:49:45 PM GMT+08:00 | Refresh : User account gets locked at random no of attempts but not as per password policy settings. | WIP 8.1 | WIP10 |  | Subsystem |  | None | Fixed | singh_j | 0 | 2 | Console Software | Console4 | NO | Closed | WIP10SegReq | IQon1 |  | WIP10 | October 11, 2018 4:54:48 PM GMT+08:00
CTUDT00599270 | Refresh | Closed | wang_l | July 25, 2018 3:52:47 PM GMT+08:00 | Air calibration failed by CT user if selected air calibration combination not succeeded by philips_service | WIP8.1 | WIP9.1 |  | Subsystem |  | N/A | Fixed | wang_l | 0 | 2 | Gantry Calibrations | Support-System | NO | Closed | WIP9.1SubCom | Ovation |  | WIP9.1 | September 18, 2018 9:42:33 AM GMT+08:00
CTUDT00599270 | Refresh | Closed | wang_l | July 25, 2018 3:52:47 PM GMT+08:00 | Air calibration failed by CT user if selected air calibration combination not succeeded by philips_service | WIP8.1 | WIP9.1 |  | Subsystem |  | N/A | Fixed | wang_l | 0 | 2 | Gantry Calibrations | Support-System | NO | Closed | WIP9.1SubCom | Ingenuity 64 |  | WIP9.1 | September 18, 2018 9:42:33 AM GMT+08:00
CTUDT00599266 | Refresh | Closed | roshan_r | July 24, 2018 10:09:16 PM GMT+08:00 | Refresh: BIG messages are not getting logged in the log viewer as expected. | wip 8.1 |  |  | Subsystem | Log - Performance | Console.SRS-15239 | Works As Specified / Not A Defect | vinoliya_j | 0 | 3 |  | Console4 | NO | Closed | WIP13SubCom | Ovation
CTUDT00599264 | Refresh | Closed | roshan_r | July 24, 2018 9:32:41 PM GMT+08:00 | Refresh : Application is allowing usage of EC from different scanner model through “External Exam Card” option while loading a study on the home page. | WIP 8.1 | WIP10 | No workaround is required since user is knowingly loading EC from different scanner model. | Subsystem |  | N/A | Fixed | harish_n | 0 | 3 | Console Software | Console4 | YES | Closed | WIP10SegReq | Ovation |  | WIP10 | September 19, 2018 1:49:04 PM GMT+08:00
CTUDT00599264 | Refresh | Closed | roshan_r | July 24, 2018 9:32:41 PM GMT+08:00 | Refresh : Application is allowing usage of EC from different scanner model through “External Exam Card” option while loading a study on the home page. | WIP 8.1 | WIP10 | No workaround is required since user is knowingly loading EC from different scanner model. | Subsystem |  | N/A | Fixed | harish_n | 0 | 3 | Console Software | Console4 | YES | Closed | WIP10SegReq | Brilliance 64 |  | WIP10 | September 19, 2018 1:49:04 PM GMT+08:00
CTUDT00599264 | Refresh | Closed | roshan_r | July 24, 2018 9:32:41 PM GMT+08:00 | Refresh : Application is allowing usage of EC from different scanner model through “External Exam Card” option while loading a study on the home page. | WIP 8.1 | WIP10 | No workaround is required since user is knowingly loading EC from different scanner model. | Subsystem |  | N/A | Fixed | harish_n | 0 | 3 | Console Software | Console4 | YES | Closed | WIP10SegReq | Ingenuity 64 |  | WIP10 | September 19, 2018 1:49:04 PM GMT+08:00
CTUDT00599264 | Refresh | Closed | roshan_r | July 24, 2018 9:32:41 PM GMT+08:00 | Refresh : Application is allowing usage of EC from different scanner model through “External Exam Card” option while loading a study on the home page. | WIP 8.1 | WIP10 | No workaround is required since user is knowingly loading EC from different scanner model. | Subsystem |  | N/A | Fixed | harish_n | 0 | 3 | Console Software | Console4 | YES | Closed | WIP10SegReq | IQon1 |  | WIP10 | September 19, 2018 1:49:04 PM GMT+08:00
CTUDT00599262 | Refresh | Closed | roshan_r | July 24, 2018 9:01:48 PM GMT+08:00 | Refresh : Application crashes when new language and phrases are added and saved in the Voice manager. | WIP 8.1 | WIP14 | the user can attempt saving the voice again while waiting longer until save operation is completed | System | Voice Manager - Edit | NA | Fixed | Snehashis_Dey | 0 | 3 | Console Software | Console4 | NO | Closed | WIP13SubCom | Ovation |  | WIP14 | March 4, 2019 7:13:23 PM GMT+08:00
CTUDT00599262 | Refresh | Closed | roshan_r | July 24, 2018 9:01:48 PM GMT+08:00 | Refresh : Application crashes when new language and phrases are added and saved in the Voice manager. | WIP 8.1 | WIP14 | the user can attempt saving the voice again while waiting longer until save operation is completed | System | Voice Manager - Edit | NA | Fixed | Snehashis_Dey | 0 | 3 | Console Software | Console4 | NO | Closed | WIP13SubCom | Brilliance 64 |  | WIP14 | March 4, 2019 7:13:23 PM GMT+08:00
CTUDT00599262 | Refresh | Closed | roshan_r | July 24, 2018 9:01:48 PM GMT+08:00 | Refresh : Application crashes when new language and phrases are added and saved in the Voice manager. | WIP 8.1 | WIP14 | the user can attempt saving the voice again while waiting longer until save operation is completed | System | Voice Manager - Edit | NA | Fixed | Snehashis_Dey | 0 | 3 | Console Software | Console4 | NO | Closed | WIP13SubCom | Brilliance Big Bore |  | WIP14 | March 4, 2019 7:13:23 PM GMT+08:00
CTUDT00599262 | Refresh | Closed | roshan_r | July 24, 2018 9:01:48 PM GMT+08:00 | Refresh : Application crashes when new language and phrases are added and saved in the Voice manager. | WIP 8.1 | WIP14 | the user can attempt saving the voice again while waiting longer until save operation is completed | System | Voice Manager - Edit | NA | Fixed | Snehashis_Dey | 0 | 3 | Console Software | Console4 | NO | Closed | WIP13SubCom | Ingenuity 64 |  | WIP14 | March 4, 2019 7:13:23 PM GMT+08:00
CTUDT00599262 | Refresh | Closed | roshan_r | July 24, 2018 9:01:48 PM GMT+08:00 | Refresh : Application crashes when new language and phrases are added and saved in the Voice manager. | WIP 8.1 | WIP14 | the user can attempt saving the voice again while waiting longer until save operation is completed | System | Voice Manager - Edit | NA | Fixed | Snehashis_Dey | 0 | 3 | Console Software | Console4 | NO | Closed | WIP13SubCom | IQon1 |  | WIP14 | March 4, 2019 7:13:23 PM GMT+08:00
CTUDT00599257 | Refresh | Closed | hiremath_b | July 24, 2018 6:18:31 PM GMT+08:00 | Refresh:Miss match in EC parameters values between show-all and Print & text window | WIP8.1 | WIP14 | NA | System | EC parameters - Mismatch | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SegReq | Ovation |  | WIP14 | February 27, 2019 7:57:29 PM GMT+08:00
CTUDT00599257 | Refresh | Closed | hiremath_b | July 24, 2018 6:18:31 PM GMT+08:00 | Refresh:Miss match in EC parameters values between show-all and Print & text window | WIP8.1 | WIP14 | NA | System | EC parameters - Mismatch | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SegReq | Brilliance 64 |  | WIP14 | February 27, 2019 7:57:29 PM GMT+08:00
CTUDT00599257 | Refresh | Closed | hiremath_b | July 24, 2018 6:18:31 PM GMT+08:00 | Refresh:Miss match in EC parameters values between show-all and Print & text window | WIP8.1 | WIP14 | NA | System | EC parameters - Mismatch | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SegReq | Brilliance Big Bore |  | WIP14 | February 27, 2019 7:57:29 PM GMT+08:00
CTUDT00599257 | Refresh | Closed | hiremath_b | July 24, 2018 6:18:31 PM GMT+08:00 | Refresh:Miss match in EC parameters values between show-all and Print & text window | WIP8.1 | WIP14 | NA | System | EC parameters - Mismatch | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SegReq | Ingenuity 64 |  | WIP14 | February 27, 2019 7:57:29 PM GMT+08:00
CTUDT00599257 | Refresh | Closed | hiremath_b | July 24, 2018 6:18:31 PM GMT+08:00 | Refresh:Miss match in EC parameters values between show-all and Print & text window | WIP8.1 | WIP14 | NA | System | EC parameters - Mismatch | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SegReq | IQon1 |  | WIP14 | February 27, 2019 7:57:29 PM GMT+08:00
CTUDT00599254 | Refresh | Closed | gT_s | July 24, 2018 3:02:23 PM GMT+08:00 | Refresh: The Logged in user name is not displayed in all the ECM tabs for Philips service user. | WIP8.1 |  |  | Subsystem | ECM - Annotations |  | Works As Specified / Not A Defect | sanjeeva_a | 1 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00599254 | Refresh | Closed | gT_s | July 24, 2018 3:02:23 PM GMT+08:00 | Refresh: The Logged in user name is not displayed in all the ECM tabs for Philips service user. | WIP8.1 |  |  | Subsystem | ECM - Annotations |  | Works As Specified / Not A Defect | sanjeeva_a | 1 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00599254 | Refresh | Closed | gT_s | July 24, 2018 3:02:23 PM GMT+08:00 | Refresh: The Logged in user name is not displayed in all the ECM tabs for Philips service user. | WIP8.1 |  |  | Subsystem | ECM - Annotations |  | Works As Specified / Not A Defect | sanjeeva_a | 1 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00599252 | Refresh | Closed | oren_d | July 24, 2018 1:41:22 PM GMT+08:00 | CTDI free air test is failing on the 2x0.625 mm collimation mode | 4.1.7 |  |  | Subsystem | History Review | N/A | Duplicate | Lior_Embon | 0 | 3 |  | Support-System | NO | Closed | see CTUDT00627825 | Ovation
CTUDT00599252 | Refresh | Closed | oren_d | July 24, 2018 1:41:22 PM GMT+08:00 | CTDI free air test is failing on the 2x0.625 mm collimation mode | 4.1.7 |  |  | Subsystem | History Review | N/A | Duplicate | Lior_Embon | 0 | 3 |  | Support-System | NO | Closed | see CTUDT00627825 | IQon1
CTUDT00599244 | Refresh | Closed | pB_s | July 23, 2018 4:15:01 PM GMT+08:00 | Emergency user name is not displayed in the application | WIP8.1 | WIP9.1 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 4 | Console USP_EBW | iCAP | NO | Closed | WIP10SegReq | Ovation |  | WIP9.1 | September 12, 2018 4:47:44 PM GMT+08:00
CTUDT00599244 | Refresh | Closed | pB_s | July 23, 2018 4:15:01 PM GMT+08:00 | Emergency user name is not displayed in the application | WIP8.1 | WIP9.1 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 4 | Console USP_EBW | iCAP | NO | Closed | WIP10SegReq | Ingenuity 64 |  | WIP9.1 | September 12, 2018 4:47:44 PM GMT+08:00
CTUDT00599244 | Refresh | Closed | pB_s | July 23, 2018 4:15:01 PM GMT+08:00 | Emergency user name is not displayed in the application | WIP8.1 | WIP9.1 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 4 | Console USP_EBW | iCAP | NO | Closed | WIP10SegReq | IQon1 |  | WIP9.1 | September 12, 2018 4:47:44 PM GMT+08:00
CTUDT00599239 | Refresh | Closed | bagdady_b | July 22, 2018 10:49:43 PM GMT+08:00 | Refresh- CCT Fluoro displays a warning during scan current performance cannot be used for real time data | WIP8.1 | WIP18 |  | Subsystem | CCT Flouro - Performance; Verified1 | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_ImgNodeSvr | CIRS | YES | Closed | WIP18SubCom | Ovation |  | WIP18 | July 12, 2019 1:07:36 AM GMT+08:00
CTUDT00599239 | Refresh | Closed | bagdady_b | July 22, 2018 10:49:43 PM GMT+08:00 | Refresh- CCT Fluoro displays a warning during scan current performance cannot be used for real time data | WIP8.1 | WIP18 |  | Subsystem | CCT Flouro - Performance; Verified1 | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_ImgNodeSvr | CIRS | YES | Closed | WIP18SubCom | Brilliance 64 |  | WIP18 | July 12, 2019 1:07:36 AM GMT+08:00
CTUDT00599239 | Refresh | Closed | bagdady_b | July 22, 2018 10:49:43 PM GMT+08:00 | Refresh- CCT Fluoro displays a warning during scan current performance cannot be used for real time data | WIP8.1 | WIP18 |  | Subsystem | CCT Flouro - Performance; Verified1 | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_ImgNodeSvr | CIRS | YES | Closed | WIP18SubCom | Brilliance Big Bore |  | WIP18 | July 12, 2019 1:07:36 AM GMT+08:00
CTUDT00599239 | Refresh | Closed | bagdady_b | July 22, 2018 10:49:43 PM GMT+08:00 | Refresh- CCT Fluoro displays a warning during scan current performance cannot be used for real time data | WIP8.1 | WIP18 |  | Subsystem | CCT Flouro - Performance; Verified1 | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_ImgNodeSvr | CIRS | YES | Closed | WIP18SubCom | Ingenuity 64 |  | WIP18 | July 12, 2019 1:07:36 AM GMT+08:00
CTUDT00599239 | Refresh | Closed | bagdady_b | July 22, 2018 10:49:43 PM GMT+08:00 | Refresh- CCT Fluoro displays a warning during scan current performance cannot be used for real time data | WIP8.1 | WIP18 |  | Subsystem | CCT Flouro - Performance; Verified1 | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_ImgNodeSvr | CIRS | YES | Closed | WIP18SubCom | IQon1 |  | WIP18 | July 12, 2019 1:07:36 AM GMT+08:00
CTUDT00599223 | Refresh | Closed | mN_m | July 20, 2018 8:52:20 PM GMT+08:00 | User name is incorrectly displayed on the image title of the online and offline edit result when emergency user is logged in. | WIP8.1 |  |  | Subsystem |  | NA | Duplicate | s_sa | 0 |  |  | iCAP | NO | Closed |  | Ovation
CTUDT00599223 | Refresh | Closed | mN_m | July 20, 2018 8:52:20 PM GMT+08:00 | User name is incorrectly displayed on the image title of the online and offline edit result when emergency user is logged in. | WIP8.1 |  |  | Subsystem |  | NA | Duplicate | s_sa | 0 |  |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00599223 | Refresh | Closed | mN_m | July 20, 2018 8:52:20 PM GMT+08:00 | User name is incorrectly displayed on the image title of the online and offline edit result when emergency user is logged in. | WIP8.1 |  |  | Subsystem |  | NA | Duplicate | s_sa | 0 |  |  | iCAP | NO | Closed |  | IQon1
CTUDT00599217 | Refresh | Closed | nanjunda_s | July 20, 2018 6:58:22 PM GMT+08:00 | Body helical Exam card has the resolution field in the Scan show all of Plan scene | WIP8.1 | Host SSRS DHF303211 Rev 06 | NA | Subsystem | Plan Scene | Console-SRS.15902 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | WIP12SubCom | Ovation |  | Host SSRS DHF303211 Rev 06 | December 12, 2018 10:25:34 PM GMT+08:00
CTUDT00599217 | Refresh | Closed | nanjunda_s | July 20, 2018 6:58:22 PM GMT+08:00 | Body helical Exam card has the resolution field in the Scan show all of Plan scene | WIP8.1 | Host SSRS DHF303211 Rev 06 | NA | Subsystem | Plan Scene | Console-SRS.15902 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | WIP12SubCom | Ingenuity 64 |  | Host SSRS DHF303211 Rev 06 | December 12, 2018 10:25:34 PM GMT+08:00
CTUDT00599215 | Refresh | Closed | pB_s | July 20, 2018 5:55:12 PM GMT+08:00 | 'Edit auto voice' option is available for second acquisition in scan ruler for Dual energy Exam cards. | WIP8.1 | WIP12 |  | Subsystem | Auto Voice | N/A | Fixed | prasad_r | 0 | 2 | Console Software | Console4 | YES | Closed | WIP11.1SegReq | Ingenuity 64 |  | WIP12 | November 26, 2018 9:12:25 PM GMT+08:00
CTUDT00599213 | Refresh | Closed | vinoliya_j | July 19, 2018 9:21:23 PM GMT+08:00 | Refresh - Console.SRS-11256 needs to be updated. | WIP 8 | Host SSRS DHF303211 Rev04 |  | Subsystem |  | N/A | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | DHF303211 Rev04 | Brilliance 64 |  | Host SSRS DHF303211 Rev04 | September 18, 2018 6:42:21 PM GMT+08:00
CTUDT00599213 | Refresh | Closed | vinoliya_j | July 19, 2018 9:21:23 PM GMT+08:00 | Refresh - Console.SRS-11256 needs to be updated. | WIP 8 | Host SSRS DHF303211 Rev04 |  | Subsystem |  | N/A | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | DHF303211 Rev04 | Brilliance Big Bore |  | Host SSRS DHF303211 Rev04 | September 18, 2018 6:42:21 PM GMT+08:00
CTUDT00599213 | Refresh | Closed | vinoliya_j | July 19, 2018 9:21:23 PM GMT+08:00 | Refresh - Console.SRS-11256 needs to be updated. | WIP 8 | Host SSRS DHF303211 Rev04 |  | Subsystem |  | N/A | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | DHF303211 Rev04 | Ingenuity 64 |  | Host SSRS DHF303211 Rev04 | September 18, 2018 6:42:21 PM GMT+08:00
CTUDT00599209 | Refresh | Closed | bommiReddy_r | July 19, 2018 7:48:14 PM GMT+08:00 | Auto launch is not working for "Brain perfusion”. Displaying error message as "BrainPerfusionApp.exe has stopped working”. | WIP8.1 | WIP13 | the workaround is to launch BP application manually. | Subsystem | Brain Perfursion - Auto Launch | NA | Fixed | redrouthu_r | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP14SegReq | Ovation |  | WIP13 | June 11, 2019 5:45:04 PM GMT+08:00
CTUDT00599207 | Refresh | Closed | nimma_p | July 19, 2018 5:53:59 PM GMT+08:00 | Refresh: Console.SRS-15324 needs to be updated | WIP 8 | Host SSRS DHF303211 Rev04 |  | Subsystem |  | N/A | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | DHF303211 Rev04 | Ovation |  | Host SSRS DHF303211 Rev04 | September 18, 2018 1:22:41 PM GMT+08:00
CTUDT00599207 | Refresh | Closed | nimma_p | July 19, 2018 5:53:59 PM GMT+08:00 | Refresh: Console.SRS-15324 needs to be updated | WIP 8 | Host SSRS DHF303211 Rev04 |  | Subsystem |  | N/A | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | DHF303211 Rev04 | Brilliance 64 |  | Host SSRS DHF303211 Rev04 | September 18, 2018 1:22:41 PM GMT+08:00
CTUDT00599207 | Refresh | Closed | nimma_p | July 19, 2018 5:53:59 PM GMT+08:00 | Refresh: Console.SRS-15324 needs to be updated | WIP 8 | Host SSRS DHF303211 Rev04 |  | Subsystem |  | N/A | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | DHF303211 Rev04 | Brilliance Big Bore |  | Host SSRS DHF303211 Rev04 | September 18, 2018 1:22:41 PM GMT+08:00
CTUDT00599207 | Refresh | Closed | nimma_p | July 19, 2018 5:53:59 PM GMT+08:00 | Refresh: Console.SRS-15324 needs to be updated | WIP 8 | Host SSRS DHF303211 Rev04 |  | Subsystem |  | N/A | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | DHF303211 Rev04 | Ingenuity 64 |  | Host SSRS DHF303211 Rev04 | September 18, 2018 1:22:41 PM GMT+08:00
CTUDT00599207 | Refresh | Closed | nimma_p | July 19, 2018 5:53:59 PM GMT+08:00 | Refresh: Console.SRS-15324 needs to be updated | WIP 8 | Host SSRS DHF303211 Rev04 |  | Subsystem |  | N/A | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | DHF303211 Rev04 | IQon1 |  | Host SSRS DHF303211 Rev04 | September 18, 2018 1:22:41 PM GMT+08:00
CTUDT00599202 | Refresh | Closed | beck_d | July 19, 2018 1:53:19 PM GMT+08:00 | Refresh: Calibrations: Disabling 0.27 RT in key is not reflected in Quality Assurance | WIP8 | N/A |  | Subsystem | Calibrations - FEE test | n/a | Fixed | gehtin_e | 0 | 2 | Test Design | DMS | NO | Closed | WIP13SegReq | IQon1 |  | N/A | January 3, 2019 4:24:49 PM GMT+08:00
CTUDT00599191 | Refresh | Closed | singh_p | July 18, 2018 10:00:13 PM GMT+08:00 | Refresh [iCT] : Wrong version displayed as 4.7 for Dose log viewer. | WIP 8.1 |  | NA | Subsystem |  | NA | Duplicate | kr_m | 0 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00599190 | Refresh | Closed | singh_p | July 18, 2018 9:57:53 PM GMT+08:00 | Refresh [iCT] : Wrong default values displayed for CTDI and DLP in Dose Management under Preferences | WIP 8.1 |  | User can manually change the default value by logging into Service mode so that Dose Alert message won't come at 500. | Subsystem |  | NA | Works As Specified / Not A Defect | sachan_a | 0 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00599187 | Refresh | Closed | beck_d | July 18, 2018 9:03:09 PM GMT+08:00 | Refresh: Calibrations: ACR Test: System setup completes successfully with System Head phantom instead of the ACR phantom | WIP8 | N/A |  | Subsystem | ACR Test Test - System Phantom | N/A | Fixed | braunstein_d | 0 | 2 | Design Documentation | DMS | NO | Closed | WIP14SegReq | Ovation |  | N/A | February 20, 2019 7:20:05 PM GMT+08:00
CTUDT00599187 | Refresh | Closed | beck_d | July 18, 2018 9:03:09 PM GMT+08:00 | Refresh: Calibrations: ACR Test: System setup completes successfully with System Head phantom instead of the ACR phantom | WIP8 | N/A |  | Subsystem | ACR Test Test - System Phantom | N/A | Fixed | braunstein_d | 0 | 2 | Design Documentation | DMS | NO | Closed | WIP14SegReq | Brilliance Big Bore |  | N/A | February 20, 2019 7:20:05 PM GMT+08:00
CTUDT00599187 | Refresh | Closed | beck_d | July 18, 2018 9:03:09 PM GMT+08:00 | Refresh: Calibrations: ACR Test: System setup completes successfully with System Head phantom instead of the ACR phantom | WIP8 | N/A |  | Subsystem | ACR Test Test - System Phantom | N/A | Fixed | braunstein_d | 0 | 2 | Design Documentation | DMS | NO | Closed | WIP14SegReq | Ingenuity 64 |  | N/A | February 20, 2019 7:20:05 PM GMT+08:00
CTUDT00599187 | Refresh | Closed | beck_d | July 18, 2018 9:03:09 PM GMT+08:00 | Refresh: Calibrations: ACR Test: System setup completes successfully with System Head phantom instead of the ACR phantom | WIP8 | N/A |  | Subsystem | ACR Test Test - System Phantom | N/A | Fixed | braunstein_d | 0 | 2 | Design Documentation | DMS | NO | Closed | WIP14SegReq | IQon1 |  | N/A | February 20, 2019 7:20:05 PM GMT+08:00
CTUDT00599186 | Refresh | Closed | pB_s | July 18, 2018 8:52:21 PM GMT+08:00 | Auto logoff is not happening for the clinical user. | WIP8.1 |  |  | Subsystem |  | NA | Duplicate | sanjeeva_a | 0 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00599186 | Refresh | Closed | pB_s | July 18, 2018 8:52:21 PM GMT+08:00 | Auto logoff is not happening for the clinical user. | WIP8.1 |  |  | Subsystem |  | NA | Duplicate | sanjeeva_a | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00599186 | Refresh | Closed | pB_s | July 18, 2018 8:52:21 PM GMT+08:00 | Auto logoff is not happening for the clinical user. | WIP8.1 |  |  | Subsystem |  | NA | Duplicate | sanjeeva_a | 0 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00599183 | Refresh | Closed | beck_d | July 18, 2018 7:51:35 PM GMT+08:00 | Refresh: Calibrations: IQ Check: The Uniformity max value does not appear in the Summary table | WIP8 | WIP15 | N/A | Subsystem | IQ Check - Uniformity max value | N/A | Fixed | elviral | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP12SegReq | Ovation |  | WIP15 | March 26, 2019 8:34:57 PM GMT+08:00
CTUDT00599183 | Refresh | Closed | beck_d | July 18, 2018 7:51:35 PM GMT+08:00 | Refresh: Calibrations: IQ Check: The Uniformity max value does not appear in the Summary table | WIP8 | WIP15 | N/A | Subsystem | IQ Check - Uniformity max value | N/A | Fixed | elviral | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP12SegReq | Brilliance 64 |  | WIP15 | March 26, 2019 8:34:57 PM GMT+08:00
CTUDT00599183 | Refresh | Closed | beck_d | July 18, 2018 7:51:35 PM GMT+08:00 | Refresh: Calibrations: IQ Check: The Uniformity max value does not appear in the Summary table | WIP8 | WIP15 | N/A | Subsystem | IQ Check - Uniformity max value | N/A | Fixed | elviral | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP12SegReq | Brilliance Big Bore |  | WIP15 | March 26, 2019 8:34:57 PM GMT+08:00
CTUDT00599183 | Refresh | Closed | beck_d | July 18, 2018 7:51:35 PM GMT+08:00 | Refresh: Calibrations: IQ Check: The Uniformity max value does not appear in the Summary table | WIP8 | WIP15 | N/A | Subsystem | IQ Check - Uniformity max value | N/A | Fixed | elviral | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP12SegReq | Ingenuity 64 |  | WIP15 | March 26, 2019 8:34:57 PM GMT+08:00
CTUDT00599183 | Refresh | Closed | beck_d | July 18, 2018 7:51:35 PM GMT+08:00 | Refresh: Calibrations: IQ Check: The Uniformity max value does not appear in the Summary table | WIP8 | WIP15 | N/A | Subsystem | IQ Check - Uniformity max value | N/A | Fixed | elviral | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP12SegReq | IQon1 |  | WIP15 | March 26, 2019 8:34:57 PM GMT+08:00
CTUDT00599180 | Refresh | Closed | beck_d | July 18, 2018 7:33:23 PM GMT+08:00 | Refresh: Calibrations: ACR Test: System setup doesn’t recognize a phantom swivel | WIP8 | N/A |  | Subsystem | ACR Test Test - Phantom Swivel | N/A | Fixed | braunstein_d | 0 | 2 | Design Documentation | Support-Physics | NO | Closed | WIP14SegReq | Ovation |  | N/A | June 13, 2019 3:23:37 PM GMT+08:00
CTUDT00599180 | Refresh | Closed | beck_d | July 18, 2018 7:33:23 PM GMT+08:00 | Refresh: Calibrations: ACR Test: System setup doesn’t recognize a phantom swivel | WIP8 | N/A |  | Subsystem | ACR Test Test - Phantom Swivel | N/A | Fixed | braunstein_d | 0 | 2 | Design Documentation | Support-Physics | NO | Closed | WIP14SegReq | Brilliance Big Bore |  | N/A | June 13, 2019 3:23:37 PM GMT+08:00
CTUDT00599180 | Refresh | Closed | beck_d | July 18, 2018 7:33:23 PM GMT+08:00 | Refresh: Calibrations: ACR Test: System setup doesn’t recognize a phantom swivel | WIP8 | N/A |  | Subsystem | ACR Test Test - Phantom Swivel | N/A | Fixed | braunstein_d | 0 | 2 | Design Documentation | Support-Physics | NO | Closed | WIP14SegReq | Ingenuity 64 |  | N/A | June 13, 2019 3:23:37 PM GMT+08:00
CTUDT00599180 | Refresh | Closed | beck_d | July 18, 2018 7:33:23 PM GMT+08:00 | Refresh: Calibrations: ACR Test: System setup doesn’t recognize a phantom swivel | WIP8 | N/A |  | Subsystem | ACR Test Test - Phantom Swivel | N/A | Fixed | braunstein_d | 0 | 2 | Design Documentation | Support-Physics | NO | Closed | WIP14SegReq | IQon1 |  | N/A | June 13, 2019 3:23:37 PM GMT+08:00
CTUDT00599178 | Refresh | Closed | naik_p | July 18, 2018 6:37:15 PM GMT+08:00 | After entering username and password and clicking on "Use remote IST, certificate password is not cleared. | WIP8.1 | WIP10 | N/A | Subsystem |  | Console-SRS.15925 | Fixed | singh_j | 0 | 2 | Couch Software | Console4 | NO | Closed | WIP10SubCom | Ovation |  | WIP10 | October 23, 2018 1:52:48 PM GMT+08:00
CTUDT00599178 | Refresh | Closed | naik_p | July 18, 2018 6:37:15 PM GMT+08:00 | After entering username and password and clicking on "Use remote IST, certificate password is not cleared. | WIP8.1 | WIP10 | N/A | Subsystem |  | Console-SRS.15925 | Fixed | singh_j | 0 | 2 | Couch Software | Console4 | NO | Closed | WIP10SubCom | Brilliance 64 |  | WIP10 | October 23, 2018 1:52:48 PM GMT+08:00
CTUDT00599178 | Refresh | Closed | naik_p | July 18, 2018 6:37:15 PM GMT+08:00 | After entering username and password and clicking on "Use remote IST, certificate password is not cleared. | WIP8.1 | WIP10 | N/A | Subsystem |  | Console-SRS.15925 | Fixed | singh_j | 0 | 2 | Couch Software | Console4 | NO | Closed | WIP10SubCom | Ingenuity 64 |  | WIP10 | October 23, 2018 1:52:48 PM GMT+08:00
CTUDT00599178 | Refresh | Closed | naik_p | July 18, 2018 6:37:15 PM GMT+08:00 | After entering username and password and clicking on "Use remote IST, certificate password is not cleared. | WIP8.1 | WIP10 | N/A | Subsystem |  | Console-SRS.15925 | Fixed | singh_j | 0 | 2 | Couch Software | Console4 | NO | Closed | WIP10SubCom | IQon1 |  | WIP10 | October 23, 2018 1:52:48 PM GMT+08:00
CTUDT00599176 | Refresh | Closed | beck_d | July 18, 2018 12:30:20 AM GMT+08:00 | Refresh: Calibrations: Performance Test: Report Summary does not appear in the top of the Report and a message is missing | WIP 8 | WIP15 | N/A | Subsystem | Performance Test - Report Summary | N/A | Fixed | elviral | 0 | 2 | DMS Software | DMS | NO | Closed | WIP13SegReq | Ovation |  | WIP15 | April 2, 2019 7:29:04 PM GMT+08:00
CTUDT00599176 | Refresh | Closed | beck_d | July 18, 2018 12:30:20 AM GMT+08:00 | Refresh: Calibrations: Performance Test: Report Summary does not appear in the top of the Report and a message is missing | WIP 8 | WIP15 | N/A | Subsystem | Performance Test - Report Summary | N/A | Fixed | elviral | 0 | 2 | DMS Software | DMS | NO | Closed | WIP13SegReq | Brilliance 64 |  | WIP15 | April 2, 2019 7:29:04 PM GMT+08:00
CTUDT00599176 | Refresh | Closed | beck_d | July 18, 2018 12:30:20 AM GMT+08:00 | Refresh: Calibrations: Performance Test: Report Summary does not appear in the top of the Report and a message is missing | WIP 8 | WIP15 | N/A | Subsystem | Performance Test - Report Summary | N/A | Fixed | elviral | 0 | 2 | DMS Software | DMS | NO | Closed | WIP13SegReq | Brilliance Big Bore |  | WIP15 | April 2, 2019 7:29:04 PM GMT+08:00
CTUDT00599176 | Refresh | Closed | beck_d | July 18, 2018 12:30:20 AM GMT+08:00 | Refresh: Calibrations: Performance Test: Report Summary does not appear in the top of the Report and a message is missing | WIP 8 | WIP15 | N/A | Subsystem | Performance Test - Report Summary | N/A | Fixed | elviral | 0 | 2 | DMS Software | DMS | NO | Closed | WIP13SegReq | Ingenuity 64 |  | WIP15 | April 2, 2019 7:29:04 PM GMT+08:00
CTUDT00599176 | Refresh | Closed | beck_d | July 18, 2018 12:30:20 AM GMT+08:00 | Refresh: Calibrations: Performance Test: Report Summary does not appear in the top of the Report and a message is missing | WIP 8 | WIP15 | N/A | Subsystem | Performance Test - Report Summary | N/A | Fixed | elviral | 0 | 2 | DMS Software | DMS | NO | Closed | WIP13SegReq | IQon1 |  | WIP15 | April 2, 2019 7:29:04 PM GMT+08:00
CTUDT00599174 | Refresh | Closed | cohen_y | July 17, 2018 10:47:46 PM GMT+08:00 | Refresh: Calibrations: Performance: Head Linearity cases fail due to calculation errors | WIP8 | WIP13 | N/A | Subsystem | Performance Test - calculation error | N/A | Fixed | elviral | 0 | 2 | DMS Software | DMS | NO | Closed | WIP13SegReq | Ovation |  | WIP13 | January 10, 2019 6:21:25 PM GMT+08:00
CTUDT00599174 | Refresh | Closed | cohen_y | July 17, 2018 10:47:46 PM GMT+08:00 | Refresh: Calibrations: Performance: Head Linearity cases fail due to calculation errors | WIP8 | WIP13 | N/A | Subsystem | Performance Test - calculation error | N/A | Fixed | elviral | 0 | 2 | DMS Software | DMS | NO | Closed | WIP13SegReq | Brilliance 64 |  | WIP13 | January 10, 2019 6:21:25 PM GMT+08:00
CTUDT00599174 | Refresh | Closed | cohen_y | July 17, 2018 10:47:46 PM GMT+08:00 | Refresh: Calibrations: Performance: Head Linearity cases fail due to calculation errors | WIP8 | WIP13 | N/A | Subsystem | Performance Test - calculation error | N/A | Fixed | elviral | 0 | 2 | DMS Software | DMS | NO | Closed | WIP13SegReq | Brilliance Big Bore |  | WIP13 | January 10, 2019 6:21:25 PM GMT+08:00
CTUDT00599174 | Refresh | Closed | cohen_y | July 17, 2018 10:47:46 PM GMT+08:00 | Refresh: Calibrations: Performance: Head Linearity cases fail due to calculation errors | WIP8 | WIP13 | N/A | Subsystem | Performance Test - calculation error | N/A | Fixed | elviral | 0 | 2 | DMS Software | DMS | NO | Closed | WIP13SegReq | Ingenuity 64 |  | WIP13 | January 10, 2019 6:21:25 PM GMT+08:00
CTUDT00599174 | Refresh | Closed | cohen_y | July 17, 2018 10:47:46 PM GMT+08:00 | Refresh: Calibrations: Performance: Head Linearity cases fail due to calculation errors | WIP8 | WIP13 | N/A | Subsystem | Performance Test - calculation error | N/A | Fixed | elviral | 0 | 2 | DMS Software | DMS | NO | Closed | WIP13SegReq | IQon1 |  | WIP13 | January 10, 2019 6:21:25 PM GMT+08:00
CTUDT00599172 | Refresh | Closed | cohen_y | July 17, 2018 10:05:39 PM GMT+08:00 | Refresh: Calibrations: HCOR: HCOR sorting order is not as defined in spec | WIP8 | WIP12 |  | Subsystem | Documentation | N/A | Fixed | elviral | 0 | 4 | DMS Software | DMS | NO | Closed | WIP12SegReq | Ovation |  | WIP12 | December 13, 2018 3:51:09 PM GMT+08:00
CTUDT00599172 | Refresh | Closed | cohen_y | July 17, 2018 10:05:39 PM GMT+08:00 | Refresh: Calibrations: HCOR: HCOR sorting order is not as defined in spec | WIP8 | WIP12 |  | Subsystem | Documentation | N/A | Fixed | elviral | 0 | 4 | DMS Software | DMS | NO | Closed | WIP12SegReq | Brilliance 64 |  | WIP12 | December 13, 2018 3:51:09 PM GMT+08:00
CTUDT00599172 | Refresh | Closed | cohen_y | July 17, 2018 10:05:39 PM GMT+08:00 | Refresh: Calibrations: HCOR: HCOR sorting order is not as defined in spec | WIP8 | WIP12 |  | Subsystem | Documentation | N/A | Fixed | elviral | 0 | 4 | DMS Software | DMS | NO | Closed | WIP12SegReq | Brilliance Big Bore |  | WIP12 | December 13, 2018 3:51:09 PM GMT+08:00
CTUDT00599172 | Refresh | Closed | cohen_y | July 17, 2018 10:05:39 PM GMT+08:00 | Refresh: Calibrations: HCOR: HCOR sorting order is not as defined in spec | WIP8 | WIP12 |  | Subsystem | Documentation | N/A | Fixed | elviral | 0 | 4 | DMS Software | DMS | NO | Closed | WIP12SegReq | Ingenuity 64 |  | WIP12 | December 13, 2018 3:51:09 PM GMT+08:00
CTUDT00599172 | Refresh | Closed | cohen_y | July 17, 2018 10:05:39 PM GMT+08:00 | Refresh: Calibrations: HCOR: HCOR sorting order is not as defined in spec | WIP8 | WIP12 |  | Subsystem | Documentation | N/A | Fixed | elviral | 0 | 4 | DMS Software | DMS | NO | Closed | WIP12SegReq | IQon1 |  | WIP12 | December 13, 2018 3:51:09 PM GMT+08:00
CTUDT00599170 | Refresh | Closed | beck_d | July 17, 2018 9:34:21 PM GMT+08:00 | Refresh: Calibrations: Performance Test: Surview Body scans length is not as defined in the spec | WIP 8 | DHF170396 Rev18 |  | Subsystem | Documentation | N/A | Fixed | richards_b | 0 | 2 | Design Documentation | Support-Physics | NO | Closed | WIP12SegReq | Ovation |  | DHF170396 Rev18 | April 24, 2019 2:24:14 AM GMT+08:00
CTUDT00599170 | Refresh | Closed | beck_d | July 17, 2018 9:34:21 PM GMT+08:00 | Refresh: Calibrations: Performance Test: Surview Body scans length is not as defined in the spec | WIP 8 | DHF170396 Rev18 |  | Subsystem | Documentation | N/A | Fixed | richards_b | 0 | 2 | Design Documentation | Support-Physics | NO | Closed | WIP12SegReq | Brilliance Big Bore |  | DHF170396 Rev18 | April 24, 2019 2:24:14 AM GMT+08:00
CTUDT00599170 | Refresh | Closed | beck_d | July 17, 2018 9:34:21 PM GMT+08:00 | Refresh: Calibrations: Performance Test: Surview Body scans length is not as defined in the spec | WIP 8 | DHF170396 Rev18 |  | Subsystem | Documentation | N/A | Fixed | richards_b | 0 | 2 | Design Documentation | Support-Physics | NO | Closed | WIP12SegReq | Ingenuity 64 |  | DHF170396 Rev18 | April 24, 2019 2:24:14 AM GMT+08:00
CTUDT00599170 | Refresh | Closed | beck_d | July 17, 2018 9:34:21 PM GMT+08:00 | Refresh: Calibrations: Performance Test: Surview Body scans length is not as defined in the spec | WIP 8 | DHF170396 Rev18 |  | Subsystem | Documentation | N/A | Fixed | richards_b | 0 | 2 | Design Documentation | Support-Physics | NO | Closed | WIP12SegReq | IQon1 |  | DHF170396 Rev18 | April 24, 2019 2:24:14 AM GMT+08:00
CTUDT00599168 | Refresh | Closed | manjunath_c | July 17, 2018 5:51:43 PM GMT+08:00 | [Refresh] [iCT, iCT-SP, Ingenuity, Br64, iQon]: Console is crashing, when a result edited in slab viewer | WIP8.0 |  |  | Subsystem | Result Edit | None | Duplicate | patra_r | 0 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00599168 | Refresh | Closed | manjunath_c | July 17, 2018 5:51:43 PM GMT+08:00 | [Refresh] [iCT, iCT-SP, Ingenuity, Br64, iQon]: Console is crashing, when a result edited in slab viewer | WIP8.0 |  |  | Subsystem | Result Edit | None | Duplicate | patra_r | 0 |  |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00599168 | Refresh | Closed | manjunath_c | July 17, 2018 5:51:43 PM GMT+08:00 | [Refresh] [iCT, iCT-SP, Ingenuity, Br64, iQon]: Console is crashing, when a result edited in slab viewer | WIP8.0 |  |  | Subsystem | Result Edit | None | Duplicate | patra_r | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00599168 | Refresh | Closed | manjunath_c | July 17, 2018 5:51:43 PM GMT+08:00 | [Refresh] [iCT, iCT-SP, Ingenuity, Br64, iQon]: Console is crashing, when a result edited in slab viewer | WIP8.0 |  |  | Subsystem | Result Edit | None | Duplicate | patra_r | 0 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00599166 | Refresh | Closed | nanjunda_s | July 17, 2018 5:26:29 PM GMT+08:00 | Field Of View field value accepted greater than 250mm for Ultra High resolution. | WIP8 | WIP10 |  | Subsystem |  | N/A | Fixed | ramesh_t | 0 | 2 | Test Design | Console4 | NO | Closed | WIP10SegReq | Ovation |  | WIP10 | November 9, 2018 6:11:23 PM GMT+08:00
CTUDT00599166 | Refresh | Closed | nanjunda_s | July 17, 2018 5:26:29 PM GMT+08:00 | Field Of View field value accepted greater than 250mm for Ultra High resolution. | WIP8 | WIP10 |  | Subsystem |  | N/A | Fixed | ramesh_t | 0 | 2 | Test Design | Console4 | NO | Closed | WIP10SegReq | Brilliance 64 |  | WIP10 | November 9, 2018 6:11:23 PM GMT+08:00
CTUDT00599166 | Refresh | Closed | nanjunda_s | July 17, 2018 5:26:29 PM GMT+08:00 | Field Of View field value accepted greater than 250mm for Ultra High resolution. | WIP8 | WIP10 |  | Subsystem |  | N/A | Fixed | ramesh_t | 0 | 2 | Test Design | Console4 | NO | Closed | WIP10SegReq | Ingenuity 64 |  | WIP10 | November 9, 2018 6:11:23 PM GMT+08:00
CTUDT00599163 | Refresh | Closed | azagury_d | July 17, 2018 3:36:57 AM GMT+08:00 | Refresh: Phantom Calibration: WIP 8: A warning message alerting the user that the phantom is misaligned happens on the second case as well. | WIP 8 | N/A |  | Subsystem |  | N/A | Fixed | gehtin_e | 0 | 2 | Test Design | DMS | NO | Closed | WIP11SegReq | IQon1 |  | N/A | October 24, 2018 10:39:43 PM GMT+08:00
CTUDT00599161 | Refresh | Closed | cohen_y | July 16, 2018 10:59:23 PM GMT+08:00 | Refresh: Calibrations: Acceptance: Adult Body images in the HTML report are not correct | WIP8 | WIP12 | N/A | Subsystem | Acceptance Test - HTML report | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP11SegReq | Ovation |  | WIP12 | January 8, 2019 3:25:20 PM GMT+08:00
CTUDT00599161 | Refresh | Closed | cohen_y | July 16, 2018 10:59:23 PM GMT+08:00 | Refresh: Calibrations: Acceptance: Adult Body images in the HTML report are not correct | WIP8 | WIP12 | N/A | Subsystem | Acceptance Test - HTML report | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP11SegReq | Brilliance 64 |  | WIP12 | January 8, 2019 3:25:20 PM GMT+08:00
CTUDT00599161 | Refresh | Closed | cohen_y | July 16, 2018 10:59:23 PM GMT+08:00 | Refresh: Calibrations: Acceptance: Adult Body images in the HTML report are not correct | WIP8 | WIP12 | N/A | Subsystem | Acceptance Test - HTML report | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP11SegReq | Brilliance Big Bore |  | WIP12 | January 8, 2019 3:25:20 PM GMT+08:00
CTUDT00599161 | Refresh | Closed | cohen_y | July 16, 2018 10:59:23 PM GMT+08:00 | Refresh: Calibrations: Acceptance: Adult Body images in the HTML report are not correct | WIP8 | WIP12 | N/A | Subsystem | Acceptance Test - HTML report | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP11SegReq | Ingenuity 64 |  | WIP12 | January 8, 2019 3:25:20 PM GMT+08:00
CTUDT00599161 | Refresh | Closed | cohen_y | July 16, 2018 10:59:23 PM GMT+08:00 | Refresh: Calibrations: Acceptance: Adult Body images in the HTML report are not correct | WIP8 | WIP12 | N/A | Subsystem | Acceptance Test - HTML report | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP11SegReq | IQon1 |  | WIP12 | January 8, 2019 3:25:20 PM GMT+08:00
CTUDT00599159 | Refresh | Closed | cohen_y | July 16, 2018 10:41:50 PM GMT+08:00 | Refresh: Calibrations: Acceptance/Constancy: No graphical representation of the measurements in the Report | WIP8 | WIP12 | N/A | Subsystem | Acceptance Test - Graphical representation | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP11SegReq | Ovation |  | WIP12 | December 17, 2018 10:03:44 PM GMT+08:00
CTUDT00599159 | Refresh | Closed | cohen_y | July 16, 2018 10:41:50 PM GMT+08:00 | Refresh: Calibrations: Acceptance/Constancy: No graphical representation of the measurements in the Report | WIP8 | WIP12 | N/A | Subsystem | Acceptance Test - Graphical representation | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP11SegReq | Brilliance 64 |  | WIP12 | December 17, 2018 10:03:44 PM GMT+08:00
CTUDT00599159 | Refresh | Closed | cohen_y | July 16, 2018 10:41:50 PM GMT+08:00 | Refresh: Calibrations: Acceptance/Constancy: No graphical representation of the measurements in the Report | WIP8 | WIP12 | N/A | Subsystem | Acceptance Test - Graphical representation | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP11SegReq | Brilliance Big Bore |  | WIP12 | December 17, 2018 10:03:44 PM GMT+08:00
CTUDT00599159 | Refresh | Closed | cohen_y | July 16, 2018 10:41:50 PM GMT+08:00 | Refresh: Calibrations: Acceptance/Constancy: No graphical representation of the measurements in the Report | WIP8 | WIP12 | N/A | Subsystem | Acceptance Test - Graphical representation | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP11SegReq | Ingenuity 64 |  | WIP12 | December 17, 2018 10:03:44 PM GMT+08:00
CTUDT00599159 | Refresh | Closed | cohen_y | July 16, 2018 10:41:50 PM GMT+08:00 | Refresh: Calibrations: Acceptance/Constancy: No graphical representation of the measurements in the Report | WIP8 | WIP12 | N/A | Subsystem | Acceptance Test - Graphical representation | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP11SegReq | IQon1 |  | WIP12 | December 17, 2018 10:03:44 PM GMT+08:00
CTUDT00599154 | Refresh | Closed | pB_s | July 16, 2018 8:50:46 PM GMT+08:00 | In Olevel login, when trying to launch the workflow first time, it was throwing an error stating "Unexpected error during workflow creation" | WIP8.1 |  | NA. | Subsystem |  | NA | Duplicate | kapadia_v | 0 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00599154 | Refresh | Closed | pB_s | July 16, 2018 8:50:46 PM GMT+08:00 | In Olevel login, when trying to launch the workflow first time, it was throwing an error stating "Unexpected error during workflow creation" | WIP8.1 |  | NA. | Subsystem |  | NA | Duplicate | kapadia_v | 0 |  |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00599154 | Refresh | Closed | pB_s | July 16, 2018 8:50:46 PM GMT+08:00 | In Olevel login, when trying to launch the workflow first time, it was throwing an error stating "Unexpected error during workflow creation" | WIP8.1 |  | NA. | Subsystem |  | NA | Duplicate | kapadia_v | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00599154 | Refresh | Closed | pB_s | July 16, 2018 8:50:46 PM GMT+08:00 | In Olevel login, when trying to launch the workflow first time, it was throwing an error stating "Unexpected error during workflow creation" | WIP8.1 |  | NA. | Subsystem |  | NA | Duplicate | kapadia_v | 0 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00599153 | Refresh | Closed | singh_p | July 16, 2018 8:43:51 PM GMT+08:00 | IQON WIP8.1:All Spectral Results were not displayed to Spectral CT viewer with Autolaunch option as Spectral CT viewer. | WIP8.1 | Test design update |  | Subsystem | Auto Lunch - Spectral Results | N/A | Fixed | nanjunda_s | 0 | 2 | Test Design | Console4 | NO | Closed | WIP11SegReq | IQon1 |  | Test design update | May 24, 2019 6:23:34 PM GMT+08:00
CTUDT00599151 | Refresh | Closed | naik_p | July 16, 2018 8:19:20 PM GMT+08:00 | The iBrain recon filters as options to the user only for head scan types is missing. | WIP8 |  |  | Subsystem |  | N/A | Duplicate | kumar_v | 0 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00599151 | Refresh | Closed | naik_p | July 16, 2018 8:19:20 PM GMT+08:00 | The iBrain recon filters as options to the user only for head scan types is missing. | WIP8 |  |  | Subsystem |  | N/A | Duplicate | kumar_v | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00599146 | Refresh | Closed | cohen_y | July 16, 2018 1:46:22 AM GMT+08:00 | Refresh: Calibrations: Acceptance: Baseline is created although Acceptance test failed. | WIP8 | WIP12 | N/A | Subsystem | Acceptance Test - Baseline created | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP11SegReq | Ovation |  | WIP12 | December 17, 2018 10:03:08 PM GMT+08:00
CTUDT00599146 | Refresh | Closed | cohen_y | July 16, 2018 1:46:22 AM GMT+08:00 | Refresh: Calibrations: Acceptance: Baseline is created although Acceptance test failed. | WIP8 | WIP12 | N/A | Subsystem | Acceptance Test - Baseline created | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP11SegReq | Brilliance 64 |  | WIP12 | December 17, 2018 10:03:08 PM GMT+08:00
CTUDT00599146 | Refresh | Closed | cohen_y | July 16, 2018 1:46:22 AM GMT+08:00 | Refresh: Calibrations: Acceptance: Baseline is created although Acceptance test failed. | WIP8 | WIP12 | N/A | Subsystem | Acceptance Test - Baseline created | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP11SegReq | Brilliance Big Bore |  | WIP12 | December 17, 2018 10:03:08 PM GMT+08:00
CTUDT00599146 | Refresh | Closed | cohen_y | July 16, 2018 1:46:22 AM GMT+08:00 | Refresh: Calibrations: Acceptance: Baseline is created although Acceptance test failed. | WIP8 | WIP12 | N/A | Subsystem | Acceptance Test - Baseline created | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP11SegReq | Ingenuity 64 |  | WIP12 | December 17, 2018 10:03:08 PM GMT+08:00
CTUDT00599146 | Refresh | Closed | cohen_y | July 16, 2018 1:46:22 AM GMT+08:00 | Refresh: Calibrations: Acceptance: Baseline is created although Acceptance test failed. | WIP8 | WIP12 | N/A | Subsystem | Acceptance Test - Baseline created | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | YES | Closed | WIP11SegReq | IQon1 |  | WIP12 | December 17, 2018 10:03:08 PM GMT+08:00
CTUDT00599144 | Refresh | Closed | cohen_y | July 16, 2018 1:41:49 AM GMT+08:00 | Refresh: Calibrations: IQ Check: CTDIVol and CTDI-free-air of each case does not appear with scan parameters in the HTML report | WIP8 | WIP11 | N/A | Subsystem |  | Common Functions SPEC DHF179202 ID126 | Fixed | yosefk | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP11SegReq | Ovation |  | WIP11 | October 23, 2018 8:29:56 PM GMT+08:00
CTUDT00599144 | Refresh | Closed | cohen_y | July 16, 2018 1:41:49 AM GMT+08:00 | Refresh: Calibrations: IQ Check: CTDIVol and CTDI-free-air of each case does not appear with scan parameters in the HTML report | WIP8 | WIP11 | N/A | Subsystem |  | Common Functions SPEC DHF179202 ID126 | Fixed | yosefk | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP11SegReq | Brilliance 64 |  | WIP11 | October 23, 2018 8:29:56 PM GMT+08:00
CTUDT00599144 | Refresh | Closed | cohen_y | July 16, 2018 1:41:49 AM GMT+08:00 | Refresh: Calibrations: IQ Check: CTDIVol and CTDI-free-air of each case does not appear with scan parameters in the HTML report | WIP8 | WIP11 | N/A | Subsystem |  | Common Functions SPEC DHF179202 ID126 | Fixed | yosefk | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP11SegReq | Brilliance Big Bore |  | WIP11 | October 23, 2018 8:29:56 PM GMT+08:00
CTUDT00599144 | Refresh | Closed | cohen_y | July 16, 2018 1:41:49 AM GMT+08:00 | Refresh: Calibrations: IQ Check: CTDIVol and CTDI-free-air of each case does not appear with scan parameters in the HTML report | WIP8 | WIP11 | N/A | Subsystem |  | Common Functions SPEC DHF179202 ID126 | Fixed | yosefk | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP11SegReq | Ingenuity 64 |  | WIP11 | October 23, 2018 8:29:56 PM GMT+08:00
CTUDT00599144 | Refresh | Closed | cohen_y | July 16, 2018 1:41:49 AM GMT+08:00 | Refresh: Calibrations: IQ Check: CTDIVol and CTDI-free-air of each case does not appear with scan parameters in the HTML report | WIP8 | WIP11 | N/A | Subsystem |  | Common Functions SPEC DHF179202 ID126 | Fixed | yosefk | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP11SegReq | IQon1 |  | WIP11 | October 23, 2018 8:29:56 PM GMT+08:00
CTUDT00599142 | Refresh | Closed | azagury_d | July 15, 2018 8:52:55 PM GMT+08:00 | Refresh: WIP 8: Common Functions: Scale limits in color plots are not as they should be | WIP 8 | N/A | User always can set any desired scale limits manually. | Subsystem | Documentation | Common Functions SPEC DHF179202 ID43-47 | Fixed | braunstein_d | 0 | 2 | Requirements Documentation | DMS | NO | Closed | DHF179202 Rev06 | IQon1 |  | N/A | January 17, 2019 4:44:56 PM GMT+08:00
CTUDT00599139 | Refresh | Closed | azagury_d | July 15, 2018 8:43:13 PM GMT+08:00 | Refresh: WIP 8: Common Functions: There is no Tube information in the results tab. | WIP 8 | WIP11 WIP14 |  | Subsystem |  | N/A | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP9SegReq | Ovation |  | WIP11 WIP14 | February 7, 2019 3:39:11 PM GMT+08:00
CTUDT00599139 | Refresh | Closed | azagury_d | July 15, 2018 8:43:13 PM GMT+08:00 | Refresh: WIP 8: Common Functions: There is no Tube information in the results tab. | WIP 8 | WIP11 WIP14 |  | Subsystem |  | N/A | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP9SegReq | Brilliance 64 |  | WIP11 WIP14 | February 7, 2019 3:39:11 PM GMT+08:00
CTUDT00599139 | Refresh | Closed | azagury_d | July 15, 2018 8:43:13 PM GMT+08:00 | Refresh: WIP 8: Common Functions: There is no Tube information in the results tab. | WIP 8 | WIP11 WIP14 |  | Subsystem |  | N/A | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP9SegReq | Brilliance Big Bore |  | WIP11 WIP14 | February 7, 2019 3:39:11 PM GMT+08:00
CTUDT00599139 | Refresh | Closed | azagury_d | July 15, 2018 8:43:13 PM GMT+08:00 | Refresh: WIP 8: Common Functions: There is no Tube information in the results tab. | WIP 8 | WIP11 WIP14 |  | Subsystem |  | N/A | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP9SegReq | Ingenuity 64 |  | WIP11 WIP14 | February 7, 2019 3:39:11 PM GMT+08:00
CTUDT00599139 | Refresh | Closed | azagury_d | July 15, 2018 8:43:13 PM GMT+08:00 | Refresh: WIP 8: Common Functions: There is no Tube information in the results tab. | WIP 8 | WIP11 WIP14 |  | Subsystem |  | N/A | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP9SegReq | IQon1 |  | WIP11 WIP14 | February 7, 2019 3:39:11 PM GMT+08:00
CTUDT00599138 | Refresh | Closed | azagury_d | July 15, 2018 8:42:24 PM GMT+08:00 | Refresh: WIP 8: An error message appears when trying to scan several applications . | WIP 8 |  |  | Subsystem |  |  | Duplicate | frider_o | 0 |  |  | DMS | NO | Closed |  | Ingenuity 64
CTUDT00599134 | Refresh | Closed | azagury_d | July 15, 2018 1:20:47 PM GMT+08:00 | Refresh:WIP 8: Tube warm-up scans are not performed | WIP 8 | WIP11 | Run short tube conditioning before running air calibration. | Subsystem |  | n/a | Fixed | yosefk | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP10SubCom | Brilliance 64 |  | WIP11 | October 11, 2018 8:50:07 PM GMT+08:00
CTUDT00599134 | Refresh | Closed | azagury_d | July 15, 2018 1:20:47 PM GMT+08:00 | Refresh:WIP 8: Tube warm-up scans are not performed | WIP 8 | WIP11 | Run short tube conditioning before running air calibration. | Subsystem |  | n/a | Fixed | yosefk | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP10SubCom | Brilliance Big Bore |  | WIP11 | October 11, 2018 8:50:07 PM GMT+08:00
CTUDT00599134 | Refresh | Closed | azagury_d | July 15, 2018 1:20:47 PM GMT+08:00 | Refresh:WIP 8: Tube warm-up scans are not performed | WIP 8 | WIP11 | Run short tube conditioning before running air calibration. | Subsystem |  | n/a | Fixed | yosefk | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP10SubCom | Ingenuity 64 |  | WIP11 | October 11, 2018 8:50:07 PM GMT+08:00
CTUDT00599122 | Refresh | Closed | ramesh_t | July 12, 2018 6:57:11 PM GMT+08:00 | Refresh: Pause button is disabled during clinical scan initialization. | WIP8 | Host SSRS DHF303211 Rev 05 | NA | Subsystem | Pause Scan | Console-SRS.2949 | Fixed | bandil_j | 0 | 3 | Host SSRS | Console4 | NO | Closed | WIP11SegReq | Ovation |  | Host SSRS DHF303211 Rev 05 | November 23, 2018 1:13:21 PM GMT+08:00
CTUDT00599122 | Refresh | Closed | ramesh_t | July 12, 2018 6:57:11 PM GMT+08:00 | Refresh: Pause button is disabled during clinical scan initialization. | WIP8 | Host SSRS DHF303211 Rev 05 | NA | Subsystem | Pause Scan | Console-SRS.2949 | Fixed | bandil_j | 0 | 3 | Host SSRS | Console4 | NO | Closed | WIP11SegReq | Brilliance 64 |  | Host SSRS DHF303211 Rev 05 | November 23, 2018 1:13:21 PM GMT+08:00
CTUDT00599122 | Refresh | Closed | ramesh_t | July 12, 2018 6:57:11 PM GMT+08:00 | Refresh: Pause button is disabled during clinical scan initialization. | WIP8 | Host SSRS DHF303211 Rev 05 | NA | Subsystem | Pause Scan | Console-SRS.2949 | Fixed | bandil_j | 0 | 3 | Host SSRS | Console4 | NO | Closed | WIP11SegReq | Ingenuity 64 |  | Host SSRS DHF303211 Rev 05 | November 23, 2018 1:13:21 PM GMT+08:00
CTUDT00599107 | Refresh | Closed | nanjunda_s | July 11, 2018 8:08:14 PM GMT+08:00 | Auto logout is not happening for the clinical user. | WIP 8 |  |  | Subsystem |  | NA | Duplicate | kapadia_v | 0 |  |  | Console4 | NO | Closed | WIP9SegReq | Ovation
CTUDT00599107 | Refresh | Closed | nanjunda_s | July 11, 2018 8:08:14 PM GMT+08:00 | Auto logout is not happening for the clinical user. | WIP 8 |  |  | Subsystem |  | NA | Duplicate | kapadia_v | 0 |  |  | Console4 | NO | Closed | WIP9SegReq | Ingenuity 64
CTUDT00599107 | Refresh | Closed | nanjunda_s | July 11, 2018 8:08:14 PM GMT+08:00 | Auto logout is not happening for the clinical user. | WIP 8 |  |  | Subsystem |  | NA | Duplicate | kapadia_v | 0 |  |  | Console4 | NO | Closed | WIP9SegReq | IQon1
CTUDT00599089 | Refresh | Closed | azagury_d | July 10, 2018 10:03:19 PM GMT+08:00 | Refresh: Calibrations:In DC User when results are saved ‘Name’ does not appear in Result title | WIP 8 | N/A - Test Design Issue |  | Subsystem |  | N/A | Fixed | gehtin_e | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP10SegReq | Ovation |  | N/A - Test Design Issue | January 27, 2019 4:54:15 PM GMT+08:00
CTUDT00599089 | Refresh | Closed | azagury_d | July 10, 2018 10:03:19 PM GMT+08:00 | Refresh: Calibrations:In DC User when results are saved ‘Name’ does not appear in Result title | WIP 8 | N/A - Test Design Issue |  | Subsystem |  | N/A | Fixed | gehtin_e | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP10SegReq | Brilliance Big Bore |  | N/A - Test Design Issue | January 27, 2019 4:54:15 PM GMT+08:00
CTUDT00599089 | Refresh | Closed | azagury_d | July 10, 2018 10:03:19 PM GMT+08:00 | Refresh: Calibrations:In DC User when results are saved ‘Name’ does not appear in Result title | WIP 8 | N/A - Test Design Issue |  | Subsystem |  | N/A | Fixed | gehtin_e | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP10SegReq | Ingenuity 64 |  | N/A - Test Design Issue | January 27, 2019 4:54:15 PM GMT+08:00
CTUDT00599089 | Refresh | Closed | azagury_d | July 10, 2018 10:03:19 PM GMT+08:00 | Refresh: Calibrations:In DC User when results are saved ‘Name’ does not appear in Result title | WIP 8 | N/A - Test Design Issue |  | Subsystem |  | N/A | Fixed | gehtin_e | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP10SegReq | IQon1 |  | N/A - Test Design Issue | January 27, 2019 4:54:15 PM GMT+08:00
CTUDT00599084 | Refresh | Closed | ramesh_t | July 10, 2018 2:34:50 PM GMT+08:00 | Refresh: When keyboard language is changed to "Chinese" in preferences and Chinese keyboard is used, Data cannot be type in Chinese instead it is taking the characters in English itself. | WIP8 | WIP12 |  | Subsystem | Chinese characters;Verified1 | N/A | Fixed | veerabhadraiah | 0 | 2 | USP | iCAP | NO | Closed | WIP9SegReq | Brilliance 64 |  | WIP12 | December 14, 2018 5:16:08 PM GMT+08:00
CTUDT00599084 | Refresh | Closed | ramesh_t | July 10, 2018 2:34:50 PM GMT+08:00 | Refresh: When keyboard language is changed to "Chinese" in preferences and Chinese keyboard is used, Data cannot be type in Chinese instead it is taking the characters in English itself. | WIP8 | WIP12 |  | Subsystem | Chinese characters;Verified1 | N/A | Fixed | veerabhadraiah | 0 | 2 | USP | iCAP | NO | Closed | WIP9SegReq | Brilliance Big Bore |  | WIP12 | December 14, 2018 5:16:08 PM GMT+08:00
CTUDT00599084 | Refresh | Closed | ramesh_t | July 10, 2018 2:34:50 PM GMT+08:00 | Refresh: When keyboard language is changed to "Chinese" in preferences and Chinese keyboard is used, Data cannot be type in Chinese instead it is taking the characters in English itself. | WIP8 | WIP12 |  | Subsystem | Chinese characters;Verified1 | N/A | Fixed | veerabhadraiah | 0 | 2 | USP | iCAP | NO | Closed | WIP9SegReq | Ingenuity 64 |  | WIP12 | December 14, 2018 5:16:08 PM GMT+08:00
CTUDT00599081 | Refresh | Closed | hiremath_b | July 9, 2018 7:46:34 PM GMT+08:00 | Refresh: VIPR tag is turned ON in STANDARD reconstruction and also Low Signal (streak) correction reconstruction sub-stage(IRS_SUBSTAGE_PREP_LOW_SIGNAL) tag is missing. | WIP8 | WIP9 | NA | Subsystem |  | Console-SRS.8955 | Fixed | sachan_a | 0 | 2 | Console Software | Console4 | YES | Closed | WIP9SegReq | Ovation |  | WIP9 | August 7, 2018 1:52:42 PM GMT+08:00
CTUDT00599081 | Refresh | Closed | hiremath_b | July 9, 2018 7:46:34 PM GMT+08:00 | Refresh: VIPR tag is turned ON in STANDARD reconstruction and also Low Signal (streak) correction reconstruction sub-stage(IRS_SUBSTAGE_PREP_LOW_SIGNAL) tag is missing. | WIP8 | WIP9 | NA | Subsystem |  | Console-SRS.8955 | Fixed | sachan_a | 0 | 2 | Console Software | Console4 | YES | Closed | WIP9SegReq | Brilliance 64 |  | WIP9 | August 7, 2018 1:52:42 PM GMT+08:00
CTUDT00599081 | Refresh | Closed | hiremath_b | July 9, 2018 7:46:34 PM GMT+08:00 | Refresh: VIPR tag is turned ON in STANDARD reconstruction and also Low Signal (streak) correction reconstruction sub-stage(IRS_SUBSTAGE_PREP_LOW_SIGNAL) tag is missing. | WIP8 | WIP9 | NA | Subsystem |  | Console-SRS.8955 | Fixed | sachan_a | 0 | 2 | Console Software | Console4 | YES | Closed | WIP9SegReq | Ingenuity 64 |  | WIP9 | August 7, 2018 1:52:42 PM GMT+08:00
CTUDT00599079 | Refresh | Closed | beck_d | July 9, 2018 7:30:29 PM GMT+08:00 | Refresh: Calibrations: Bad Detectors Calibration and FEE Test: TubeCollimator parameter in Series Structure is o | WIP 8 | N/A |  | Subsystem |  | N/A | Fixed | gehtin_e | 0 | 2 | Test Design | DMS | NO | Closed | WIP10SegReq | IQon1 |  | N/A | November 5, 2018 4:24:40 PM GMT+08:00
CTUDT00599076 | Refresh | Closed | mN_m | July 9, 2018 6:50:53 PM GMT+08:00 | Recon mode is reset back to Non IMR on performing edit result of IMR result of Cardiac Perfusion scan. | WIP8 | WIP9.1 |  | Subsystem |  | N/A | Fixed | datta_i | 0 | 3 | Console Software | Console4 | YES | Closed | WIP10SegReq | Ovation |  | WIP9.1 | August 9, 2018 3:33:44 PM GMT+08:00
CTUDT00599076 | Refresh | Closed | mN_m | July 9, 2018 6:50:53 PM GMT+08:00 | Recon mode is reset back to Non IMR on performing edit result of IMR result of Cardiac Perfusion scan. | WIP8 | WIP9.1 |  | Subsystem |  | N/A | Fixed | datta_i | 0 | 3 | Console Software | Console4 | YES | Closed | WIP10SegReq | IQon1 |  | WIP9.1 | August 9, 2018 3:33:44 PM GMT+08:00
CTUDT00599072 | Refresh | Closed | beck_d | July 9, 2018 6:16:52 PM GMT+08:00 | Refresh: Calibrations: Quality Assurance gets stuck while creating or opening xres files | WIP 8 |  |  | Subsystem | Calibration History | N/A | Works As Specified / Not A Defect | tamari_o | 1 |  |  | Support-System | NO | Closed |  | IQon1
CTUDT00599070 | Refresh | Closed | blessily | July 9, 2018 5:39:41 PM GMT+08:00 | Monitor settings option is available for olevel users and fails req Console.SRS.15897 | WIP8 | Host SSRS DHF303211 Rev 04 |  | Subsystem |  | Console.SRS.15897 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | DHF303211 Rev04 | Ovation |  | Host SSRS DHF303211 Rev 04 | September 18, 2018 6:53:13 PM GMT+08:00
CTUDT00599070 | Refresh | Closed | blessily | July 9, 2018 5:39:41 PM GMT+08:00 | Monitor settings option is available for olevel users and fails req Console.SRS.15897 | WIP8 | Host SSRS DHF303211 Rev 04 |  | Subsystem |  | Console.SRS.15897 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | DHF303211 Rev04 | Brilliance 64 |  | Host SSRS DHF303211 Rev 04 | September 18, 2018 6:53:13 PM GMT+08:00
CTUDT00599070 | Refresh | Closed | blessily | July 9, 2018 5:39:41 PM GMT+08:00 | Monitor settings option is available for olevel users and fails req Console.SRS.15897 | WIP8 | Host SSRS DHF303211 Rev 04 |  | Subsystem |  | Console.SRS.15897 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | DHF303211 Rev04 | Brilliance Big Bore |  | Host SSRS DHF303211 Rev 04 | September 18, 2018 6:53:13 PM GMT+08:00
CTUDT00599070 | Refresh | Closed | blessily | July 9, 2018 5:39:41 PM GMT+08:00 | Monitor settings option is available for olevel users and fails req Console.SRS.15897 | WIP8 | Host SSRS DHF303211 Rev 04 |  | Subsystem |  | Console.SRS.15897 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | DHF303211 Rev04 | Ingenuity 64 |  | Host SSRS DHF303211 Rev 04 | September 18, 2018 6:53:13 PM GMT+08:00
CTUDT00599070 | Refresh | Closed | blessily | July 9, 2018 5:39:41 PM GMT+08:00 | Monitor settings option is available for olevel users and fails req Console.SRS.15897 | WIP8 | Host SSRS DHF303211 Rev 04 |  | Subsystem |  | Console.SRS.15897 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | DHF303211 Rev04 | IQon1 |  | Host SSRS DHF303211 Rev 04 | September 18, 2018 6:53:13 PM GMT+08:00
CTUDT00599067 | Refresh | Closed | beck_d | July 9, 2018 12:53:55 PM GMT+08:00 | Refresh: Calibrations: Plots titles do not contain the corresponding Low\High energy indication | WIP 8 | WIP9.1 |  | Subsystem |  | N/A | Fixed | frider_o | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP11SegReq | IQon1 |  | WIP9.1 | October 23, 2018 8:20:12 PM GMT+08:00
CTUDT00599065 | Refresh | Closed | azagury_d | July 9, 2018 2:17:35 AM GMT+08:00 | Refresh: WIP 8: Stability: Mean Tile TempCo plot has wrong x-axis label (TempCo) instead of module #. | WIP 8 |  |  | Subsystem | Stability test | N/A | Works As Specified / Not A Defect | gehtin_e | 1 |  |  | DMS | NO | Closed |  | IQon1
CTUDT00599053 | Refresh | Closed | kadish_n | July 5, 2018 8:58:43 PM GMT+08:00 | Error while loading raw data in Physics Analysis | WIP 8 | WIP11 |  | Subsystem |  | N/A | Fixed | yosefk | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP11SegReq | Ovation |  | WIP11 | October 23, 2018 8:43:24 PM GMT+08:00
CTUDT00599053 | Refresh | Closed | kadish_n | July 5, 2018 8:58:43 PM GMT+08:00 | Error while loading raw data in Physics Analysis | WIP 8 | WIP11 |  | Subsystem |  | N/A | Fixed | yosefk | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP11SegReq | Brilliance 64 |  | WIP11 | October 23, 2018 8:43:24 PM GMT+08:00
CTUDT00599053 | Refresh | Closed | kadish_n | July 5, 2018 8:58:43 PM GMT+08:00 | Error while loading raw data in Physics Analysis | WIP 8 | WIP11 |  | Subsystem |  | N/A | Fixed | yosefk | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP11SegReq | Ingenuity 64 |  | WIP11 | October 23, 2018 8:43:24 PM GMT+08:00
CTUDT00599053 | Refresh | Closed | kadish_n | July 5, 2018 8:58:43 PM GMT+08:00 | Error while loading raw data in Physics Analysis | WIP 8 | WIP11 |  | Subsystem |  | N/A | Fixed | yosefk | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP11SegReq | IQon1 |  | WIP11 | October 23, 2018 8:43:24 PM GMT+08:00
CTUDT00599043 | Refresh | Closed | kadish_n | July 4, 2018 10:15:37 PM GMT+08:00 | The limits of some plots generated during FEE test do not match the spec | WIP 8 | WIP10 |  | Subsystem | Verified1 | N/A | Fixed | yosefk | 0 | 2 | DMS Software | DMS | NO | Closed | WIP10SegReq | Ovation |  | WIP10 | October 8, 2018 10:01:58 PM GMT+08:00
CTUDT00599043 | Refresh | Closed | kadish_n | July 4, 2018 10:15:37 PM GMT+08:00 | The limits of some plots generated during FEE test do not match the spec | WIP 8 | WIP10 |  | Subsystem | Verified1 | N/A | Fixed | yosefk | 0 | 2 | DMS Software | DMS | NO | Closed | WIP10SegReq | Brilliance 64 |  | WIP10 | October 8, 2018 10:01:58 PM GMT+08:00
CTUDT00599043 | Refresh | Closed | kadish_n | July 4, 2018 10:15:37 PM GMT+08:00 | The limits of some plots generated during FEE test do not match the spec | WIP 8 | WIP10 |  | Subsystem | Verified1 | N/A | Fixed | yosefk | 0 | 2 | DMS Software | DMS | NO | Closed | WIP10SegReq | Brilliance Big Bore |  | WIP10 | October 8, 2018 10:01:58 PM GMT+08:00
CTUDT00599043 | Refresh | Closed | kadish_n | July 4, 2018 10:15:37 PM GMT+08:00 | The limits of some plots generated during FEE test do not match the spec | WIP 8 | WIP10 |  | Subsystem | Verified1 | N/A | Fixed | yosefk | 0 | 2 | DMS Software | DMS | NO | Closed | WIP10SegReq | Ingenuity 64 |  | WIP10 | October 8, 2018 10:01:58 PM GMT+08:00
CTUDT00599043 | Refresh | Closed | kadish_n | July 4, 2018 10:15:37 PM GMT+08:00 | The limits of some plots generated during FEE test do not match the spec | WIP 8 | WIP10 |  | Subsystem | Verified1 | N/A | Fixed | yosefk | 0 | 2 | DMS Software | DMS | NO | Closed | WIP10SegReq | IQon1 |  | WIP10 | October 8, 2018 10:01:58 PM GMT+08:00
CTUDT00599028 | Refresh | Closed | kadish_n | July 4, 2018 3:41:11 AM GMT+08:00 | MBD - an error appears when trying to set 'channel' | WIP 8 | WIP10 | For the relevant input application will reflect with correct output. | Subsystem |  | N/A | Fixed | yosefk | 0 | 4 | DMS Software | DMS | NO | Closed | WIP9.1 SegReq | Ovation |  | WIP10 | August 20, 2018 8:21:19 PM GMT+08:00
CTUDT00599028 | Refresh | Closed | kadish_n | July 4, 2018 3:41:11 AM GMT+08:00 | MBD - an error appears when trying to set 'channel' | WIP 8 | WIP10 | For the relevant input application will reflect with correct output. | Subsystem |  | N/A | Fixed | yosefk | 0 | 4 | DMS Software | DMS | NO | Closed | WIP9.1 SegReq | Brilliance 64 |  | WIP10 | August 20, 2018 8:21:19 PM GMT+08:00
CTUDT00599028 | Refresh | Closed | kadish_n | July 4, 2018 3:41:11 AM GMT+08:00 | MBD - an error appears when trying to set 'channel' | WIP 8 | WIP10 | For the relevant input application will reflect with correct output. | Subsystem |  | N/A | Fixed | yosefk | 0 | 4 | DMS Software | DMS | NO | Closed | WIP9.1 SegReq | Brilliance Big Bore |  | WIP10 | August 20, 2018 8:21:19 PM GMT+08:00
CTUDT00599028 | Refresh | Closed | kadish_n | July 4, 2018 3:41:11 AM GMT+08:00 | MBD - an error appears when trying to set 'channel' | WIP 8 | WIP10 | For the relevant input application will reflect with correct output. | Subsystem |  | N/A | Fixed | yosefk | 0 | 4 | DMS Software | DMS | NO | Closed | WIP9.1 SegReq | Ingenuity 64 |  | WIP10 | August 20, 2018 8:21:19 PM GMT+08:00
CTUDT00599028 | Refresh | Closed | kadish_n | July 4, 2018 3:41:11 AM GMT+08:00 | MBD - an error appears when trying to set 'channel' | WIP 8 | WIP10 | For the relevant input application will reflect with correct output. | Subsystem |  | N/A | Fixed | yosefk | 0 | 4 | DMS Software | DMS | NO | Closed | WIP9.1 SegReq | IQon1 |  | WIP10 | August 20, 2018 8:21:19 PM GMT+08:00
CTUDT00599026 | Refresh | Closed | manjunath_c | July 3, 2018 10:59:39 PM GMT+08:00 | New series not created with entered description in Acquisitions save as window | WIP6.1 | WIP12 |  | Subsystem | CCT - Series creation | N/A | Fixed | Prakash_Chand_Yadav | 0 | 3 | Console Software | Console4 | NO | Closed | WIP12SegReq | IQon1 |  | WIP12 | December 4, 2018 5:58:01 PM GMT+08:00
CTUDT00599024 | Refresh | Closed | kadish_n | July 3, 2018 10:05:42 PM GMT+08:00 | Wrong status for added Bad Detectors in MBD | WIP 8 | WIP12 |  | Subsystem | Bad Detector | N/A | Fixed | yosefk | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP11.1SegReq | Ovation |  | WIP12 | November 25, 2018 6:28:23 PM GMT+08:00
CTUDT00599024 | Refresh | Closed | kadish_n | July 3, 2018 10:05:42 PM GMT+08:00 | Wrong status for added Bad Detectors in MBD | WIP 8 | WIP12 |  | Subsystem | Bad Detector | N/A | Fixed | yosefk | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP11.1SegReq | IQon1 |  | WIP12 | November 25, 2018 6:28:23 PM GMT+08:00
CTUDT00599018 | Refresh | Closed | juster_d | July 3, 2018 7:42:07 PM GMT+08:00 | CIRS should not send unrelated images during Bolus tracker scan | 4.7.5 |  |  | Subsystem | Performance improvement | N/A | Duplicate | patra_r | 0 | 2 |  | Console4 | YES | Closed | WIP15SubCom | IQon1 | ESC#0110990827 PR 8567370 PR 8539682 Case # 0111665405
CTUDT00599017 | Refresh | Closed | manjunath_c | July 3, 2018 7:33:26 PM GMT+08:00 | Use Previous surview option exists in context menu | WIP6.1 |  |  | Subsystem | Use Previous Surview |  | Works As Specified / Not A Defect | sanjeeva_a | 1 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00599017 | Refresh | Closed | manjunath_c | July 3, 2018 7:33:26 PM GMT+08:00 | Use Previous surview option exists in context menu | WIP6.1 |  |  | Subsystem | Use Previous Surview |  | Works As Specified / Not A Defect | sanjeeva_a | 1 |  |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00599017 | Refresh | Closed | manjunath_c | July 3, 2018 7:33:26 PM GMT+08:00 | Use Previous surview option exists in context menu | WIP6.1 |  |  | Subsystem | Use Previous Surview |  | Works As Specified / Not A Defect | sanjeeva_a | 1 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00599017 | Refresh | Closed | manjunath_c | July 3, 2018 7:33:26 PM GMT+08:00 | Use Previous surview option exists in context menu | WIP6.1 |  |  | Subsystem | Use Previous Surview |  | Works As Specified / Not A Defect | sanjeeva_a | 1 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00599006 | Refresh | Closed | naik_p | July 3, 2018 4:43:00 PM GMT+08:00 | Locator tracker and MPR results include extra drop down option. | WIP 8 | WIP10 | No | Subsystem |  | N/A | Fixed | kumar_v | 0 | 2 | Console Software | Console4 | YES | Closed | WIP10SegReq | Ovation |  | WIP10 | September 19, 2018 6:16:11 PM GMT+08:00
CTUDT00599006 | Refresh | Closed | naik_p | July 3, 2018 4:43:00 PM GMT+08:00 | Locator tracker and MPR results include extra drop down option. | WIP 8 | WIP10 | No | Subsystem |  | N/A | Fixed | kumar_v | 0 | 2 | Console Software | Console4 | YES | Closed | WIP10SegReq | Brilliance 64 |  | WIP10 | September 19, 2018 6:16:11 PM GMT+08:00
CTUDT00599006 | Refresh | Closed | naik_p | July 3, 2018 4:43:00 PM GMT+08:00 | Locator tracker and MPR results include extra drop down option. | WIP 8 | WIP10 | No | Subsystem |  | N/A | Fixed | kumar_v | 0 | 2 | Console Software | Console4 | YES | Closed | WIP10SegReq | Ingenuity 64 |  | WIP10 | September 19, 2018 6:16:11 PM GMT+08:00
CTUDT00599006 | Refresh | Closed | naik_p | July 3, 2018 4:43:00 PM GMT+08:00 | Locator tracker and MPR results include extra drop down option. | WIP 8 | WIP10 | No | Subsystem |  | N/A | Fixed | kumar_v | 0 | 2 | Console Software | Console4 | YES | Closed | WIP10SegReq | IQon1 |  | WIP10 | September 19, 2018 6:16:11 PM GMT+08:00
CTUDT00598998 | Refresh | Closed | azagury_d | July 2, 2018 9:06:57 PM GMT+08:00 | Refresh: Calibraitions:Shadow Test: Shadow per sticks plot the X axis title is detector # instead of stick # | WIP 8 | WIP10 |  | Subsystem |  | N/A | Fixed | yosefk | 0 | 2 | DMS Software | DMS | NO | Closed | WIP9.1 SegReq | IQon1 |  | WIP10 | August 20, 2018 7:43:04 PM GMT+08:00
CTUDT00598985 | Refresh | Closed | juster_d | July 2, 2018 2:44:17 PM GMT+08:00 | The exam application crashes sporadically after end of xray with error “StepEnded XXX arrived unexpectedly” | 4.1 | WIP12 |  | Subsystem | History Review | N/A | Fixed | v_sumanth | 0 | 2 | Console Software | Console4 | YES | Closed | WIP10SegReq | Ovation | PR10110186 | WIP12 | December 7, 2018 8:00:11 PM GMT+08:00
CTUDT00598985 | Refresh | Closed | juster_d | July 2, 2018 2:44:17 PM GMT+08:00 | The exam application crashes sporadically after end of xray with error “StepEnded XXX arrived unexpectedly” | 4.1 | WIP12 |  | Subsystem | History Review | N/A | Fixed | v_sumanth | 0 | 2 | Console Software | Console4 | YES | Closed | WIP10SegReq | Brilliance 64 | PR10110186 | WIP12 | December 7, 2018 8:00:11 PM GMT+08:00
CTUDT00598985 | Refresh | Closed | juster_d | July 2, 2018 2:44:17 PM GMT+08:00 | The exam application crashes sporadically after end of xray with error “StepEnded XXX arrived unexpectedly” | 4.1 | WIP12 |  | Subsystem | History Review | N/A | Fixed | v_sumanth | 0 | 2 | Console Software | Console4 | YES | Closed | WIP10SegReq | Ingenuity 64 | PR10110186 | WIP12 | December 7, 2018 8:00:11 PM GMT+08:00
CTUDT00598985 | Refresh | Closed | juster_d | July 2, 2018 2:44:17 PM GMT+08:00 | The exam application crashes sporadically after end of xray with error “StepEnded XXX arrived unexpectedly” | 4.1 | WIP12 |  | Subsystem | History Review | N/A | Fixed | v_sumanth | 0 | 2 | Console Software | Console4 | YES | Closed | WIP10SegReq | IQon1 | PR10110186 | WIP12 | December 7, 2018 8:00:11 PM GMT+08:00
CTUDT00598983 | Refresh | Closed | golzban_v | July 1, 2018 9:16:02 PM GMT+08:00 | HB: Refresh: Failure of Spectral checkbox while planning a Cardiac result | WIP8 | WIP9 | In order to recover the Spectral phase, the faulted phase must be deleted and another Conventional phase must be duplicated and turned Spectral. | Subsystem |  | NA | Fixed | bhawna_g | 0 | 2 | Console Software | Console4 | NO | Closed | WIP9SegReq | IQon1 |  | WIP9 | August 1, 2018 8:59:00 PM GMT+08:00
CTUDT00598980 | Refresh | Closed | golzban_v | July 1, 2018 7:05:15 PM GMT+08:00 | Refresh: GPC requires a password after restart or log out/log in | WIP8 |  |  | Subsystem | GPC Installation | N/A | Does Not Apply | karkour | 0 |  |  | GCX | NO | Closed |  | IQon1
CTUDT00598971 | Refresh | Closed | nayak_r | June 29, 2018 8:09:31 PM GMT+08:00 | During the initialization of a tracker scan ,clicking on the pause button in status monitor module is not working properly | WIP 6.1 |  | NA | Subsystem |  | Console-SRS.2949 | Duplicate | sachan_a | 0 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00598971 | Refresh | Closed | nayak_r | June 29, 2018 8:09:31 PM GMT+08:00 | During the initialization of a tracker scan ,clicking on the pause button in status monitor module is not working properly | WIP 6.1 |  | NA | Subsystem |  | Console-SRS.2949 | Duplicate | sachan_a | 0 |  |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00598971 | Refresh | Closed | nayak_r | June 29, 2018 8:09:31 PM GMT+08:00 | During the initialization of a tracker scan ,clicking on the pause button in status monitor module is not working properly | WIP 6.1 |  | NA | Subsystem |  | Console-SRS.2949 | Duplicate | sachan_a | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00598967 | Refresh | Closed | nayak_r | June 29, 2018 7:13:27 PM GMT+08:00 | Clicking on Pause button after selecting “Start Clinical Scan” which appears during the tracker scan does not bring back the focus to Plan tab | WIP6.1 | WIP13 |  | Subsystem | Bolus Tracking - Pause | N/A | Fixed | goswami_s | 0 | 2 | Console Software | Console4 | YES | Closed | WIP11SegReq | Ovation |  | WIP13 | January 23, 2019 5:28:55 PM GMT+08:00
CTUDT00598967 | Refresh | Closed | nayak_r | June 29, 2018 7:13:27 PM GMT+08:00 | Clicking on Pause button after selecting “Start Clinical Scan” which appears during the tracker scan does not bring back the focus to Plan tab | WIP6.1 | WIP13 |  | Subsystem | Bolus Tracking - Pause | N/A | Fixed | goswami_s | 0 | 2 | Console Software | Console4 | YES | Closed | WIP11SegReq | Brilliance 64 |  | WIP13 | January 23, 2019 5:28:55 PM GMT+08:00
CTUDT00598967 | Refresh | Closed | nayak_r | June 29, 2018 7:13:27 PM GMT+08:00 | Clicking on Pause button after selecting “Start Clinical Scan” which appears during the tracker scan does not bring back the focus to Plan tab | WIP6.1 | WIP13 |  | Subsystem | Bolus Tracking - Pause | N/A | Fixed | goswami_s | 0 | 2 | Console Software | Console4 | YES | Closed | WIP11SegReq | Ingenuity 64 |  | WIP13 | January 23, 2019 5:28:55 PM GMT+08:00
CTUDT00598967 | Refresh | Closed | nayak_r | June 29, 2018 7:13:27 PM GMT+08:00 | Clicking on Pause button after selecting “Start Clinical Scan” which appears during the tracker scan does not bring back the focus to Plan tab | WIP6.1 | WIP13 |  | Subsystem | Bolus Tracking - Pause | N/A | Fixed | goswami_s | 0 | 2 | Console Software | Console4 | YES | Closed | WIP11SegReq | IQon1 |  | WIP13 | January 23, 2019 5:28:55 PM GMT+08:00
CTUDT00598963 | Refresh | Closed | nanjunda_s | June 29, 2018 3:37:42 PM GMT+08:00 | Changed the monitor setting in preferences from Standard to DIN (Visa versa),without re-launch of application changes getting reflected. | WIP 8 |  |  | Subsystem |  | None | No Plan to Fix | sanjeeva_a | 0 | 4 |  | Console4 | NO | Closed |  | Ovation
CTUDT00598963 | Refresh | Closed | nanjunda_s | June 29, 2018 3:37:42 PM GMT+08:00 | Changed the monitor setting in preferences from Standard to DIN (Visa versa),without re-launch of application changes getting reflected. | WIP 8 |  |  | Subsystem |  | None | No Plan to Fix | sanjeeva_a | 0 | 4 |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00598963 | Refresh | Closed | nanjunda_s | June 29, 2018 3:37:42 PM GMT+08:00 | Changed the monitor setting in preferences from Standard to DIN (Visa versa),without re-launch of application changes getting reflected. | WIP 8 |  |  | Subsystem |  | None | No Plan to Fix | sanjeeva_a | 0 | 4 |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00598963 | Refresh | Closed | nanjunda_s | June 29, 2018 3:37:42 PM GMT+08:00 | Changed the monitor setting in preferences from Standard to DIN (Visa versa),without re-launch of application changes getting reflected. | WIP 8 |  |  | Subsystem |  | None | No Plan to Fix | sanjeeva_a | 0 | 4 |  | Console4 | NO | Closed |  | IQon1
CTUDT00598954 | Refresh | Closed | vinoliya_j | June 28, 2018 10:02:14 PM GMT+08:00 | Refresh Ing: Gantry is not ready error shows up at various stages of scanning. Reported by T Zhang, 25-Jul. | WIP 8 |  |  | Subsystem | UI Performance - screen freeze |  | Does Not Apply | cumpson | 0 |  |  | GCX | NO | Closed |  | Ingenuity 64
CTUDT00598949 | Refresh | Closed | bagdady_b | June 28, 2018 8:49:00 PM GMT+08:00 | Refresh- Preferences are cut and not all of the pages are localized | WIP8 | WIP9 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP9SegReq | Ovation |  | WIP9 | August 23, 2018 1:55:38 PM GMT+08:00
CTUDT00598949 | Refresh | Closed | bagdady_b | June 28, 2018 8:49:00 PM GMT+08:00 | Refresh- Preferences are cut and not all of the pages are localized | WIP8 | WIP9 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP9SegReq | Brilliance 64 |  | WIP9 | August 23, 2018 1:55:38 PM GMT+08:00
CTUDT00598949 | Refresh | Closed | bagdady_b | June 28, 2018 8:49:00 PM GMT+08:00 | Refresh- Preferences are cut and not all of the pages are localized | WIP8 | WIP9 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP9SegReq | Brilliance Big Bore |  | WIP9 | August 23, 2018 1:55:38 PM GMT+08:00
CTUDT00598949 | Refresh | Closed | bagdady_b | June 28, 2018 8:49:00 PM GMT+08:00 | Refresh- Preferences are cut and not all of the pages are localized | WIP8 | WIP9 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP9SegReq | Ingenuity 64 |  | WIP9 | August 23, 2018 1:55:38 PM GMT+08:00
CTUDT00598949 | Refresh | Closed | bagdady_b | June 28, 2018 8:49:00 PM GMT+08:00 | Refresh- Preferences are cut and not all of the pages are localized | WIP8 | WIP9 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP9SegReq | IQon1 |  | WIP9 | August 23, 2018 1:55:38 PM GMT+08:00
CTUDT00598947 | Refresh | Closed | bagdady_b | June 28, 2018 8:47:14 PM GMT+08:00 | Refresh- The left monitor is not displayed when using Interface as Chinese in CT user | WIP8 | WIP9 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP9SegReq | Ovation |  | WIP9 | August 3, 2018 11:44:04 AM GMT+08:00
CTUDT00598947 | Refresh | Closed | bagdady_b | June 28, 2018 8:47:14 PM GMT+08:00 | Refresh- The left monitor is not displayed when using Interface as Chinese in CT user | WIP8 | WIP9 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP9SegReq | Brilliance 64 |  | WIP9 | August 3, 2018 11:44:04 AM GMT+08:00
CTUDT00598947 | Refresh | Closed | bagdady_b | June 28, 2018 8:47:14 PM GMT+08:00 | Refresh- The left monitor is not displayed when using Interface as Chinese in CT user | WIP8 | WIP9 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP9SegReq | Brilliance Big Bore |  | WIP9 | August 3, 2018 11:44:04 AM GMT+08:00
CTUDT00598947 | Refresh | Closed | bagdady_b | June 28, 2018 8:47:14 PM GMT+08:00 | Refresh- The left monitor is not displayed when using Interface as Chinese in CT user | WIP8 | WIP9 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP9SegReq | Ingenuity 64 |  | WIP9 | August 3, 2018 11:44:04 AM GMT+08:00
CTUDT00598947 | Refresh | Closed | bagdady_b | June 28, 2018 8:47:14 PM GMT+08:00 | Refresh- The left monitor is not displayed when using Interface as Chinese in CT user | WIP8 | WIP9 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP9SegReq | IQon1 |  | WIP9 | August 3, 2018 11:44:04 AM GMT+08:00
CTUDT00598946 | Refresh | Closed | pB_s | June 28, 2018 8:39:26 PM GMT+08:00 | Bug rep is launching after logging into CT mode and not able to proceed further. | WIP8 |  |  | Subsystem | BugRep |  | Works As Specified / Not A Defect | sanjeeva_a | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00598939 | Refresh | Closed | fitusi_r | June 28, 2018 7:41:22 PM GMT+08:00 | Refresh, IQon –User IT can delete user Philips_service. | WIP 8 |  | NA | Subsystem |  | NA | Works As Specified / Not A Defect | ravindra_a | 0 |  |  | Console4 | NO | Closed | WIP9SegReq | IQon1
CTUDT00598938 | Refresh | Closed | kadish_n | June 28, 2018 7:39:19 PM GMT+08:00 | Couch Position issues | WIP6.1 |  |  | Subsystem | Couch Position |  | Works As Specified / Not A Defect | gehtin_e | 0 |  |  | DMS | NO | Closed |  | IQon1
CTUDT00598934 | Refresh | Closed | mN_m | June 28, 2018 5:43:53 PM GMT+08:00 | Refresh Ing: Unexpected Bugrep error occurs once the result is launched to Pulmonary viewer. | WIP8 |  | user can load study manually from PD | Subsystem | Auto Launch | Console SRS - 13060 | No Plan to Fix | aswani_v | 0 | 3 |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00598934 | Refresh | Closed | mN_m | June 28, 2018 5:43:53 PM GMT+08:00 | Refresh Ing: Unexpected Bugrep error occurs once the result is launched to Pulmonary viewer. | WIP8 |  | user can load study manually from PD | Subsystem | Auto Launch | Console SRS - 13060 | No Plan to Fix | aswani_v | 0 | 3 |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00598934 | Refresh | Closed | mN_m | June 28, 2018 5:43:53 PM GMT+08:00 | Refresh Ing: Unexpected Bugrep error occurs once the result is launched to Pulmonary viewer. | WIP8 |  | user can load study manually from PD | Subsystem | Auto Launch | Console SRS - 13060 | No Plan to Fix | aswani_v | 0 | 3 |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00598932 | Refresh | Closed | mN_m | June 28, 2018 5:37:40 PM GMT+08:00 | Auto Voice checkbox is getting rechecked after clicking on “Apply to All Scans” for Locator and tracker scans in “Auto Voice Parameters” window. | WIP8 | WIP9.1 |  | Subsystem |  | NA | Fixed | navale_s | 0 | 3 | Console Software | Console4 | NO | Closed | WIP10SegReq | Ovation |  | WIP9.1 | August 8, 2018 8:14:17 PM GMT+08:00
CTUDT00598932 | Refresh | Closed | mN_m | June 28, 2018 5:37:40 PM GMT+08:00 | Auto Voice checkbox is getting rechecked after clicking on “Apply to All Scans” for Locator and tracker scans in “Auto Voice Parameters” window. | WIP8 | WIP9.1 |  | Subsystem |  | NA | Fixed | navale_s | 0 | 3 | Console Software | Console4 | NO | Closed | WIP10SegReq | Brilliance 64 |  | WIP9.1 | August 8, 2018 8:14:17 PM GMT+08:00
CTUDT00598932 | Refresh | Closed | mN_m | June 28, 2018 5:37:40 PM GMT+08:00 | Auto Voice checkbox is getting rechecked after clicking on “Apply to All Scans” for Locator and tracker scans in “Auto Voice Parameters” window. | WIP8 | WIP9.1 |  | Subsystem |  | NA | Fixed | navale_s | 0 | 3 | Console Software | Console4 | NO | Closed | WIP10SegReq | Ingenuity 64 |  | WIP9.1 | August 8, 2018 8:14:17 PM GMT+08:00
CTUDT00598932 | Refresh | Closed | mN_m | June 28, 2018 5:37:40 PM GMT+08:00 | Auto Voice checkbox is getting rechecked after clicking on “Apply to All Scans” for Locator and tracker scans in “Auto Voice Parameters” window. | WIP8 | WIP9.1 |  | Subsystem |  | NA | Fixed | navale_s | 0 | 3 | Console Software | Console4 | NO | Closed | WIP10SegReq | IQon1 |  | WIP9.1 | August 8, 2018 8:14:17 PM GMT+08:00
CTUDT00598930 | Refresh | Closed | mN_m | June 28, 2018 5:32:06 PM GMT+08:00 | Tooltip is missing on the perfusion scans once the scan is initiated and completed. | WIP8 | WIP9.1 |  | Subsystem |  | None | Fixed | navale_s | 0 | 3 | Console Software | Console4 | YES | Closed | WIP9SegReq | Ovation |  | WIP9.1 | August 1, 2018 8:20:37 PM GMT+08:00
CTUDT00598930 | Refresh | Closed | mN_m | June 28, 2018 5:32:06 PM GMT+08:00 | Tooltip is missing on the perfusion scans once the scan is initiated and completed. | WIP8 | WIP9.1 |  | Subsystem |  | None | Fixed | navale_s | 0 | 3 | Console Software | Console4 | YES | Closed | WIP9SegReq | Brilliance 64 |  | WIP9.1 | August 1, 2018 8:20:37 PM GMT+08:00
CTUDT00598930 | Refresh | Closed | mN_m | June 28, 2018 5:32:06 PM GMT+08:00 | Tooltip is missing on the perfusion scans once the scan is initiated and completed. | WIP8 | WIP9.1 |  | Subsystem |  | None | Fixed | navale_s | 0 | 3 | Console Software | Console4 | YES | Closed | WIP9SegReq | Ingenuity 64 |  | WIP9.1 | August 1, 2018 8:20:37 PM GMT+08:00
CTUDT00598930 | Refresh | Closed | mN_m | June 28, 2018 5:32:06 PM GMT+08:00 | Tooltip is missing on the perfusion scans once the scan is initiated and completed. | WIP8 | WIP9.1 |  | Subsystem |  | None | Fixed | navale_s | 0 | 3 | Console Software | Console4 | YES | Closed | WIP9SegReq | IQon1 |  | WIP9.1 | August 1, 2018 8:20:37 PM GMT+08:00
CTUDT00598928 | Refresh | Closed | mN_m | June 28, 2018 5:28:31 PM GMT+08:00 | The minimize button in the ECG viewer is disabled when there is no unperformed scan in the study. | WIP8 |  | NA | Subsystem |  | NA | Works As Specified / Not A Defect | madamsetti_k | 0 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00598928 | Refresh | Closed | mN_m | June 28, 2018 5:28:31 PM GMT+08:00 | The minimize button in the ECG viewer is disabled when there is no unperformed scan in the study. | WIP8 |  | NA | Subsystem |  | NA | Works As Specified / Not A Defect | madamsetti_k | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00598928 | Refresh | Closed | mN_m | June 28, 2018 5:28:31 PM GMT+08:00 | The minimize button in the ECG viewer is disabled when there is no unperformed scan in the study. | WIP8 |  | NA | Subsystem |  | NA | Works As Specified / Not A Defect | madamsetti_k | 0 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00598926 | Refresh | Closed | mN_m | June 28, 2018 5:25:10 PM GMT+08:00 | Unable to pause the perfusion scan between the groups in perfusion scan when 1st group scan is completed | WIP8 | WIP11.1 | User can pause the scan through CT Box. | Subsystem | Perfusion Scan - Pause | NA | Fixed | kumar_v | 0 | 2 | Console Software | Console4 | YES | Closed | WIP10SegReq | Ovation |  | WIP11.1 | November 19, 2018 12:56:24 PM GMT+08:00
CTUDT00598926 | Refresh | Closed | mN_m | June 28, 2018 5:25:10 PM GMT+08:00 | Unable to pause the perfusion scan between the groups in perfusion scan when 1st group scan is completed | WIP8 | WIP11.1 | User can pause the scan through CT Box. | Subsystem | Perfusion Scan - Pause | NA | Fixed | kumar_v | 0 | 2 | Console Software | Console4 | YES | Closed | WIP10SegReq | Brilliance 64 |  | WIP11.1 | November 19, 2018 12:56:24 PM GMT+08:00
CTUDT00598926 | Refresh | Closed | mN_m | June 28, 2018 5:25:10 PM GMT+08:00 | Unable to pause the perfusion scan between the groups in perfusion scan when 1st group scan is completed | WIP8 | WIP11.1 | User can pause the scan through CT Box. | Subsystem | Perfusion Scan - Pause | NA | Fixed | kumar_v | 0 | 2 | Console Software | Console4 | YES | Closed | WIP10SegReq | Ingenuity 64 |  | WIP11.1 | November 19, 2018 12:56:24 PM GMT+08:00
CTUDT00598926 | Refresh | Closed | mN_m | June 28, 2018 5:25:10 PM GMT+08:00 | Unable to pause the perfusion scan between the groups in perfusion scan when 1st group scan is completed | WIP8 | WIP11.1 | User can pause the scan through CT Box. | Subsystem | Perfusion Scan - Pause | NA | Fixed | kumar_v | 0 | 2 | Console Software | Console4 | YES | Closed | WIP10SegReq | IQon1 |  | WIP11.1 | November 19, 2018 12:56:24 PM GMT+08:00
CTUDT00598924 | Refresh | Closed | mN_m | June 28, 2018 5:17:49 PM GMT+08:00 | Not able to select mpr orientation for results in both ECM and plan for “High Res Axial” EC | WIP8 | WIP13 |  | Subsystem | Plan | N/A | Fixed | vinoliya_j | 0 | 2 | Console Software | Console4 | NO | Closed | WIP11SegReq | Ovation |  | WIP13 | February 12, 2019 6:49:27 PM GMT+08:00
CTUDT00598924 | Refresh | Closed | mN_m | June 28, 2018 5:17:49 PM GMT+08:00 | Not able to select mpr orientation for results in both ECM and plan for “High Res Axial” EC | WIP8 | WIP13 |  | Subsystem | Plan | N/A | Fixed | vinoliya_j | 0 | 2 | Console Software | Console4 | NO | Closed | WIP11SegReq | Ingenuity 64 |  | WIP13 | February 12, 2019 6:49:27 PM GMT+08:00
CTUDT00598924 | Refresh | Closed | mN_m | June 28, 2018 5:17:49 PM GMT+08:00 | Not able to select mpr orientation for results in both ECM and plan for “High Res Axial” EC | WIP8 | WIP13 |  | Subsystem | Plan | N/A | Fixed | vinoliya_j | 0 | 2 | Console Software | Console4 | NO | Closed | WIP11SegReq | IQon1 |  | WIP13 | February 12, 2019 6:49:27 PM GMT+08:00
CTUDT00598922 | Refresh | Closed | mN_m | June 28, 2018 5:02:11 PM GMT+08:00 | Study crashes on starting the clinical scan again after pausing the bolus scan during post threshold delay. | WIP8 | WIP16 | None | Subsystem | Bolus Tracking - Pause | N/A | Fixed | nanjunda_s | 0 | 4 | Test Design | Support-Simulator | YES | Closed | WIP10SegReq | Simulators |  | WIP16 | May 24, 2019 5:07:20 PM GMT+08:00
CTUDT00598911 | Refresh | Closed | hiremath_b | June 27, 2018 8:03:16 PM GMT+08:00 | v4.1.10: System Crash (Bugrep Error) on resizing the columns in Recon Queue Manager | WIP 8 | WIP9.1 |  | Subsystem |  | NA | Fixed | navale_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP10SegReq | IQon1 |  | WIP9.1 | August 10, 2018 6:55:16 PM GMT+08:00
CTUDT00598908 | Refresh | Closed | hiremath_b | June 27, 2018 7:43:41 PM GMT+08:00 | v4.1.10: iCT->default option is not selected in Plan Geometry Limitations section in preferences(REQ 15218) | WIP 8 | WIP9 |  | Subsystem |  | NA | Fixed | bhawna_g | 0 | 3 | Console Software | Console4 | NO | Closed | WIP9SegReq | Ovation |  | WIP9 | July 24, 2018 3:38:57 PM GMT+08:00
CTUDT00598906 | Refresh | Closed | fitusi_r | June 27, 2018 7:02:35 PM GMT+08:00 | Refresh, IQon -Monitor Calibration issue | WIP 8 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | tamari_o | 0 |  |  | Console4 | NO | Closed | WIP9SegReq | IQon1
CTUDT00598904 | Refresh | Closed | fitusi_r | June 27, 2018 6:48:51 PM GMT+08:00 | Refresh, IQon, ICT-cancel a Calibration in the middle of the process takes too long | WIP 8 |  |  | Subsystem | Quality Assurance - Calibrattion | n/a | Duplicate | frider_o | 0 |  |  | DMS | NO | Closed |  | Ovation
CTUDT00598904 | Refresh | Closed | fitusi_r | June 27, 2018 6:48:51 PM GMT+08:00 | Refresh, IQon, ICT-cancel a Calibration in the middle of the process takes too long | WIP 8 |  |  | Subsystem | Quality Assurance - Calibrattion | n/a | Duplicate | frider_o | 0 |  |  | DMS | NO | Closed |  | IQon1
CTUDT00598895 | Refresh | Closed | beck | June 27, 2018 2:30:16 AM GMT+08:00 | End Exam leads to Workflow application error | V4.1.10.xx025 | WIP9 |  | Subsystem |  | N/A | Fixed | rathore_cs | 0 | 3 | Console Software | Console4 | NO | Closed | WIP9SegReq | Ovation |  | WIP9 | September 21, 2018 4:41:39 PM GMT+08:00
CTUDT00598895 | Refresh | Closed | beck | June 27, 2018 2:30:16 AM GMT+08:00 | End Exam leads to Workflow application error | V4.1.10.xx025 | WIP9 |  | Subsystem |  | N/A | Fixed | rathore_cs | 0 | 3 | Console Software | Console4 | NO | Closed | WIP9SegReq | Brilliance 64 |  | WIP9 | September 21, 2018 4:41:39 PM GMT+08:00
CTUDT00598895 | Refresh | Closed | beck | June 27, 2018 2:30:16 AM GMT+08:00 | End Exam leads to Workflow application error | V4.1.10.xx025 | WIP9 |  | Subsystem |  | N/A | Fixed | rathore_cs | 0 | 3 | Console Software | Console4 | NO | Closed | WIP9SegReq | Ingenuity 64 |  | WIP9 | September 21, 2018 4:41:39 PM GMT+08:00
CTUDT00598895 | Refresh | Closed | beck | June 27, 2018 2:30:16 AM GMT+08:00 | End Exam leads to Workflow application error | V4.1.10.xx025 | WIP9 |  | Subsystem |  | N/A | Fixed | rathore_cs | 0 | 3 | Console Software | Console4 | NO | Closed | WIP9SegReq | IQon1 |  | WIP9 | September 21, 2018 4:41:39 PM GMT+08:00
CTUDT00598893 | Refresh | Closed | fitusi_r | June 26, 2018 10:47:17 PM GMT+08:00 | Refresh, IQon -Air calibration Via user CT | WIP 8 |  |  | Subsystem | Quality Assurance - Air Calibrattion | n/a | Works As Specified / Not A Defect | tamari_o | 1 |  |  | Support-System | NO | Closed |  | IQon1
CTUDT00598886 | Refresh | Closed | juster_d | June 26, 2018 2:39:35 PM GMT+08:00 | IQon, Cannot save ROIs in the "heartbeat-cs" application | 4.7.5 | WIP9.1 | No Workaround | Subsystem |  | NA | Fixed | redrouthu_r | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP9.1SegReq | IQon1 | PR7817485 | WIP9.1 | August 28, 2018 6:34:34 PM GMT+08:00
CTUDT00598884 | Refresh | Closed | zhou_x | June 26, 2018 5:14:46 AM GMT+08:00 | A-plane Skew hanged up after scanning. | WIP8 | WIP9.1 | None | Subsystem | Calibrations | GCX.SSRS.iCT.11989 | Fixed | zhou_x | 0 | 2 | Gantry Calibrations | GCX | NO | Closed | WIP9.1SubCom | Ovation |  | WIP9.1 | April 17, 2019 3:11:28 PM GMT+08:00
CTUDT00598884 | Refresh | Closed | zhou_x | June 26, 2018 5:14:46 AM GMT+08:00 | A-plane Skew hanged up after scanning. | WIP8 | WIP9.1 | None | Subsystem | Calibrations | GCX.SSRS.iCT.11989 | Fixed | zhou_x | 0 | 2 | Gantry Calibrations | GCX | NO | Closed | WIP9.1SubCom | Brilliance 64 |  | WIP9.1 | April 17, 2019 3:11:28 PM GMT+08:00
CTUDT00598884 | Refresh | Closed | zhou_x | June 26, 2018 5:14:46 AM GMT+08:00 | A-plane Skew hanged up after scanning. | WIP8 | WIP9.1 | None | Subsystem | Calibrations | GCX.SSRS.iCT.11989 | Fixed | zhou_x | 0 | 2 | Gantry Calibrations | GCX | NO | Closed | WIP9.1SubCom | Ingenuity 64 |  | WIP9.1 | April 17, 2019 3:11:28 PM GMT+08:00
CTUDT00598884 | Refresh | Closed | zhou_x | June 26, 2018 5:14:46 AM GMT+08:00 | A-plane Skew hanged up after scanning. | WIP8 | WIP9.1 | None | Subsystem | Calibrations | GCX.SSRS.iCT.11989 | Fixed | zhou_x | 0 | 2 | Gantry Calibrations | GCX | NO | Closed | WIP9.1SubCom | IQon1 |  | WIP9.1 | April 17, 2019 3:11:28 PM GMT+08:00
CTUDT00598882 | Refresh | Closed | zhou_x | June 26, 2018 4:41:39 AM GMT+08:00 | A-Plane Skew stopped due to invalid combination error. | WIP8 | WIP9.1 | None | Subsystem |  | GCX.SSRS.iCT.11989 | Fixed | zhou_x | 0 | 2 | Gantry Calibrations | GCX | NO | Closed | WIP9.1SubCom | Brilliance 64 |  | WIP9.1 | September 20, 2018 1:33:18 PM GMT+08:00
CTUDT00598882 | Refresh | Closed | zhou_x | June 26, 2018 4:41:39 AM GMT+08:00 | A-Plane Skew stopped due to invalid combination error. | WIP8 | WIP9.1 | None | Subsystem |  | GCX.SSRS.iCT.11989 | Fixed | zhou_x | 0 | 2 | Gantry Calibrations | GCX | NO | Closed | WIP9.1SubCom | Ingenuity 64 |  | WIP9.1 | September 20, 2018 1:33:18 PM GMT+08:00
CTUDT00598870 | Refresh | Closed | levy_s | June 25, 2018 8:00:04 PM GMT+08:00 | Refresh - Left monitor freeze during standard operation | WIP8 |  | NA | Subsystem | UI Performance | NA | Works As Specified / Not A Defect | tamari_o | 1 | 2 |  | Support-System | YES | Closed | SegReq_WIP9 | Ovation
CTUDT00598870 | Refresh | Closed | levy_s | June 25, 2018 8:00:04 PM GMT+08:00 | Refresh - Left monitor freeze during standard operation | WIP8 |  | NA | Subsystem | UI Performance | NA | Works As Specified / Not A Defect | tamari_o | 1 | 2 |  | Support-System | YES | Closed | SegReq_WIP9 | Brilliance 64
CTUDT00598870 | Refresh | Closed | levy_s | June 25, 2018 8:00:04 PM GMT+08:00 | Refresh - Left monitor freeze during standard operation | WIP8 |  | NA | Subsystem | UI Performance | NA | Works As Specified / Not A Defect | tamari_o | 1 | 2 |  | Support-System | YES | Closed | SegReq_WIP9 | Ingenuity 64
CTUDT00598870 | Refresh | Closed | levy_s | June 25, 2018 8:00:04 PM GMT+08:00 | Refresh - Left monitor freeze during standard operation | WIP8 |  | NA | Subsystem | UI Performance | NA | Works As Specified / Not A Defect | tamari_o | 1 | 2 |  | Support-System | YES | Closed | SegReq_WIP9 | IQon1
CTUDT00598868 | Refresh | Closed | bagdady_b | June 25, 2018 7:48:17 PM GMT+08:00 | Refresh- Bugrep is not collecting all the log data | WIP8 |  |  | Subsystem | BugRep |  | Works As Specified / Not A Defect | tamari_o | 1 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00598866 | Refresh | Closed | nanjunda_s | June 25, 2018 6:53:51 PM GMT+08:00 | How Do I has Iqon videos in ingenuity model. | WIP8 | WIP9 |  | Subsystem |  | Console-SRS.12590;Console-SRS.15521 | Fixed | datta_i | 0 | 2 | Console Software | Console4 | NO | Closed | WIP9SegReq | Ovation |  | WIP9 | August 3, 2018 2:51:56 PM GMT+08:00
CTUDT00598866 | Refresh | Closed | nanjunda_s | June 25, 2018 6:53:51 PM GMT+08:00 | How Do I has Iqon videos in ingenuity model. | WIP8 | WIP9 |  | Subsystem |  | Console-SRS.12590;Console-SRS.15521 | Fixed | datta_i | 0 | 2 | Console Software | Console4 | NO | Closed | WIP9SegReq | Brilliance 64 |  | WIP9 | August 3, 2018 2:51:56 PM GMT+08:00
CTUDT00598866 | Refresh | Closed | nanjunda_s | June 25, 2018 6:53:51 PM GMT+08:00 | How Do I has Iqon videos in ingenuity model. | WIP8 | WIP9 |  | Subsystem |  | Console-SRS.12590;Console-SRS.15521 | Fixed | datta_i | 0 | 2 | Console Software | Console4 | NO | Closed | WIP9SegReq | Ingenuity 64 |  | WIP9 | August 3, 2018 2:51:56 PM GMT+08:00
CTUDT00598862 | Refresh | Closed | bagdady_b | June 25, 2018 3:54:28 PM GMT+08:00 | Refresh- ABC Bone Correction table does not exist and causing the reconstruction to fail | WIP8 | WIP9.1 | N/A | Subsystem |  | N/A | Fixed | frider_o | 0 | 2 | CIRS_SW_Installation | DMS | YES | Closed | WIP8.1SubCom | Ovation |  | WIP9.1 | September 13, 2018 1:39:05 PM GMT+08:00
CTUDT00598852 | Refresh | Closed | naik_p | June 22, 2018 6:25:23 PM GMT+08:00 | Auto launch is not working, but Review tab is enabled with no data displayed AND Review tab is not exiting on clicking Exit button and throws Bugrep error on switching between the tabs. | WIP8 |  |  | Subsystem | Auto Launch - |  | Works As Specified / Not A Defect | redrouthu_r | 1 |  |  | iCAP | NO | Closed | WIP9SegReq | Ovation
CTUDT00598852 | Refresh | Closed | naik_p | June 22, 2018 6:25:23 PM GMT+08:00 | Auto launch is not working, but Review tab is enabled with no data displayed AND Review tab is not exiting on clicking Exit button and throws Bugrep error on switching between the tabs. | WIP8 |  |  | Subsystem | Auto Launch - |  | Works As Specified / Not A Defect | redrouthu_r | 1 |  |  | iCAP | NO | Closed | WIP9SegReq | Ingenuity 64
CTUDT00598852 | Refresh | Closed | naik_p | June 22, 2018 6:25:23 PM GMT+08:00 | Auto launch is not working, but Review tab is enabled with no data displayed AND Review tab is not exiting on clicking Exit button and throws Bugrep error on switching between the tabs. | WIP8 |  |  | Subsystem | Auto Launch - |  | Works As Specified / Not A Defect | redrouthu_r | 1 |  |  | iCAP | NO | Closed | WIP9SegReq | IQon1
CTUDT00598849 | Refresh | Closed | naik_p | June 22, 2018 5:56:21 PM GMT+08:00 | The Keyboard language dropdown has “Chinese PRC” instead of “Chinese Simplified” . | WIP8 | WIP9 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP9SubCom | Ovation |  | WIP9 | August 3, 2018 11:45:18 AM GMT+08:00
CTUDT00598849 | Refresh | Closed | naik_p | June 22, 2018 5:56:21 PM GMT+08:00 | The Keyboard language dropdown has “Chinese PRC” instead of “Chinese Simplified” . | WIP8 | WIP9 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP9SubCom | Brilliance 64 |  | WIP9 | August 3, 2018 11:45:18 AM GMT+08:00
CTUDT00598849 | Refresh | Closed | naik_p | June 22, 2018 5:56:21 PM GMT+08:00 | The Keyboard language dropdown has “Chinese PRC” instead of “Chinese Simplified” . | WIP8 | WIP9 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP9SubCom | Brilliance Big Bore |  | WIP9 | August 3, 2018 11:45:18 AM GMT+08:00
CTUDT00598849 | Refresh | Closed | naik_p | June 22, 2018 5:56:21 PM GMT+08:00 | The Keyboard language dropdown has “Chinese PRC” instead of “Chinese Simplified” . | WIP8 | WIP9 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP9SubCom | Ingenuity 64 |  | WIP9 | August 3, 2018 11:45:18 AM GMT+08:00
CTUDT00598849 | Refresh | Closed | naik_p | June 22, 2018 5:56:21 PM GMT+08:00 | The Keyboard language dropdown has “Chinese PRC” instead of “Chinese Simplified” . | WIP8 | WIP9 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP9SubCom | IQon1 |  | WIP9 | August 3, 2018 11:45:18 AM GMT+08:00
CTUDT00598848 | Refresh | Closed | singh_p | June 22, 2018 5:56:01 PM GMT+08:00 | Refresh [IQON WIP 8] : All spectral results are not opening with Auto Launch as Spectral CT Viewer | WIP8 |  |  | Subsystem | Auto Launch - Spectral results |  | Change Proposal | redrouthu_r | 0 |  |  | Support_Spectral | NO | Closed |  | IQon1
CTUDT00598847 | Refresh | Closed | singh_p | June 22, 2018 5:40:00 PM GMT+08:00 | Refresh [IQON WIP 8] : Texts are overlapping in result viewer for spectral result. | WIP8 | WIP14 | the user can change the layout or ensure using the spectral burn annotations. | Subsystem | Spectral CT Viewer | NA | Fixed | kumar_v | 0 | 3 | Console Software | Console4 | YES | Closed | WIP14SubCom | IQon1 |  | WIP14 | February 19, 2019 6:55:27 PM GMT+08:00
CTUDT00598841 | Refresh | Closed | naik_p | June 22, 2018 4:47:34 PM GMT+08:00 | In Olevel user, clicking on close “X”, system did not logged off. | WIP8 |  |  | Subsystem |  | NA | Duplicate | kapadia_v | 0 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00598841 | Refresh | Closed | naik_p | June 22, 2018 4:47:34 PM GMT+08:00 | In Olevel user, clicking on close “X”, system did not logged off. | WIP8 |  |  | Subsystem |  | NA | Duplicate | kapadia_v | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00598841 | Refresh | Closed | naik_p | June 22, 2018 4:47:34 PM GMT+08:00 | In Olevel user, clicking on close “X”, system did not logged off. | WIP8 |  |  | Subsystem |  | NA | Duplicate | kapadia_v | 0 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00598840 | Refresh | Closed | singh_p | June 22, 2018 4:41:02 PM GMT+08:00 | Refresh [IQON and Ingenuity WIP 8] : Log Viewer is crashing. Not able to filter any module | WIP8 | WIP9 |  | Subsystem |  | NA | Duplicate | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP9SubCom | Ovation |  | WIP9 | August 3, 2018 12:01:16 PM GMT+08:00
CTUDT00598840 | Refresh | Closed | singh_p | June 22, 2018 4:41:02 PM GMT+08:00 | Refresh [IQON and Ingenuity WIP 8] : Log Viewer is crashing. Not able to filter any module | WIP8 | WIP9 |  | Subsystem |  | NA | Duplicate | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP9SubCom | Brilliance 64 |  | WIP9 | August 3, 2018 12:01:16 PM GMT+08:00
CTUDT00598840 | Refresh | Closed | singh_p | June 22, 2018 4:41:02 PM GMT+08:00 | Refresh [IQON and Ingenuity WIP 8] : Log Viewer is crashing. Not able to filter any module | WIP8 | WIP9 |  | Subsystem |  | NA | Duplicate | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP9SubCom | Brilliance Big Bore |  | WIP9 | August 3, 2018 12:01:16 PM GMT+08:00
CTUDT00598840 | Refresh | Closed | singh_p | June 22, 2018 4:41:02 PM GMT+08:00 | Refresh [IQON and Ingenuity WIP 8] : Log Viewer is crashing. Not able to filter any module | WIP8 | WIP9 |  | Subsystem |  | NA | Duplicate | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP9SubCom | Ingenuity 64 |  | WIP9 | August 3, 2018 12:01:16 PM GMT+08:00
CTUDT00598840 | Refresh | Closed | singh_p | June 22, 2018 4:41:02 PM GMT+08:00 | Refresh [IQON and Ingenuity WIP 8] : Log Viewer is crashing. Not able to filter any module | WIP8 | WIP9 |  | Subsystem |  | NA | Duplicate | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP9SubCom | IQon1 |  | WIP9 | August 3, 2018 12:01:16 PM GMT+08:00
CTUDT00598838 | Refresh | Closed | singh_p | June 22, 2018 4:16:17 PM GMT+08:00 | Refresh [IQON WIP 8] : Voice Manager is taking more than 15 sec during launch and exit. | WIP8 | WIP13 |  | Subsystem | Voice Manager - UI Performance | NA | Fixed | Snehashis_Dey | 0 | 2 | Console Software | Console4 | NO | Closed | WIP13SubCom | IQon1 |  | WIP13 | January 23, 2019 6:12:23 PM GMT+08:00
CTUDT00598836 | Refresh | Closed | singh_p | June 22, 2018 3:56:40 PM GMT+08:00 | Dose Management was displayed as "CTDIvol XX.XX(0-2000 mG)" in plan page | WIP8 | WIP8 | NA | Subsystem |  | NA | Fixed | harish_n | 0 | 2 | Console Software | Console4 | NO | Closed | WIP10SegReq | Ovation |  | WIP8 | July 30, 2018 8:48:54 PM GMT+08:00
CTUDT00598836 | Refresh | Closed | singh_p | June 22, 2018 3:56:40 PM GMT+08:00 | Dose Management was displayed as "CTDIvol XX.XX(0-2000 mG)" in plan page | WIP8 | WIP8 | NA | Subsystem |  | NA | Fixed | harish_n | 0 | 2 | Console Software | Console4 | NO | Closed | WIP10SegReq | Brilliance 64 |  | WIP8 | July 30, 2018 8:48:54 PM GMT+08:00
CTUDT00598836 | Refresh | Closed | singh_p | June 22, 2018 3:56:40 PM GMT+08:00 | Dose Management was displayed as "CTDIvol XX.XX(0-2000 mG)" in plan page | WIP8 | WIP8 | NA | Subsystem |  | NA | Fixed | harish_n | 0 | 2 | Console Software | Console4 | NO | Closed | WIP10SegReq | Ingenuity 64 |  | WIP8 | July 30, 2018 8:48:54 PM GMT+08:00
CTUDT00598836 | Refresh | Closed | singh_p | June 22, 2018 3:56:40 PM GMT+08:00 | Dose Management was displayed as "CTDIvol XX.XX(0-2000 mG)" in plan page | WIP8 | WIP8 | NA | Subsystem |  | NA | Fixed | harish_n | 0 | 2 | Console Software | Console4 | NO | Closed | WIP10SegReq | IQon1 |  | WIP8 | July 30, 2018 8:48:54 PM GMT+08:00
CTUDT00598832 | Refresh | Closed | naik_p | June 22, 2018 1:04:23 PM GMT+08:00 | In Olevel user, clicking on close “X”, system did not logged off. | WIP 8 |  |  | Subsystem | System Log Off | NA | Works As Specified / Not A Defect | kapadia_v | 0 | 3 |  | Console4 | NO | Closed | WIP9SegReq | Ovation
CTUDT00598832 | Refresh | Closed | naik_p | June 22, 2018 1:04:23 PM GMT+08:00 | In Olevel user, clicking on close “X”, system did not logged off. | WIP 8 |  |  | Subsystem | System Log Off | NA | Works As Specified / Not A Defect | kapadia_v | 0 | 3 |  | Console4 | NO | Closed | WIP9SegReq | Brilliance 64
CTUDT00598832 | Refresh | Closed | naik_p | June 22, 2018 1:04:23 PM GMT+08:00 | In Olevel user, clicking on close “X”, system did not logged off. | WIP 8 |  |  | Subsystem | System Log Off | NA | Works As Specified / Not A Defect | kapadia_v | 0 | 3 |  | Console4 | NO | Closed | WIP9SegReq | Ingenuity 64
CTUDT00598832 | Refresh | Closed | naik_p | June 22, 2018 1:04:23 PM GMT+08:00 | In Olevel user, clicking on close “X”, system did not logged off. | WIP 8 |  |  | Subsystem | System Log Off | NA | Works As Specified / Not A Defect | kapadia_v | 0 | 3 |  | Console4 | NO | Closed | WIP9SegReq | IQon1
CTUDT00598830 | Refresh | Closed | naik_p | June 22, 2018 1:00:39 PM GMT+08:00 | Duplicate Dose SR series are created for each study. | WIP8 | WIP11 | N/A | Subsystem |  | Console-SRS.2171 | Fixed | bhawna_g | 0 | 4 | Console Software | Console4 | NO | Closed | WIP10SegReq | Ovation |  | WIP11 | October 30, 2018 4:51:14 PM GMT+08:00
CTUDT00598830 | Refresh | Closed | naik_p | June 22, 2018 1:00:39 PM GMT+08:00 | Duplicate Dose SR series are created for each study. | WIP8 | WIP11 | N/A | Subsystem |  | Console-SRS.2171 | Fixed | bhawna_g | 0 | 4 | Console Software | Console4 | NO | Closed | WIP10SegReq | Ingenuity 64 |  | WIP11 | October 30, 2018 4:51:14 PM GMT+08:00
CTUDT00598830 | Refresh | Closed | naik_p | June 22, 2018 1:00:39 PM GMT+08:00 | Duplicate Dose SR series are created for each study. | WIP8 | WIP11 | N/A | Subsystem |  | Console-SRS.2171 | Fixed | bhawna_g | 0 | 4 | Console Software | Console4 | NO | Closed | WIP10SegReq | IQon1 |  | WIP11 | October 30, 2018 4:51:14 PM GMT+08:00
CTUDT00598826 | Refresh | Closed | bagdady_b | June 21, 2018 10:45:04 PM GMT+08:00 | Refresh- Logviewer is not responding and crashes after pressing on refresh button | WIP8 | WIP9 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP9SegReq | Ovation |  | WIP9 | August 3, 2018 11:55:45 AM GMT+08:00
CTUDT00598826 | Refresh | Closed | bagdady_b | June 21, 2018 10:45:04 PM GMT+08:00 | Refresh- Logviewer is not responding and crashes after pressing on refresh button | WIP8 | WIP9 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP9SegReq | Brilliance 64 |  | WIP9 | August 3, 2018 11:55:45 AM GMT+08:00
CTUDT00598826 | Refresh | Closed | bagdady_b | June 21, 2018 10:45:04 PM GMT+08:00 | Refresh- Logviewer is not responding and crashes after pressing on refresh button | WIP8 | WIP9 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP9SegReq | Brilliance Big Bore |  | WIP9 | August 3, 2018 11:55:45 AM GMT+08:00
CTUDT00598826 | Refresh | Closed | bagdady_b | June 21, 2018 10:45:04 PM GMT+08:00 | Refresh- Logviewer is not responding and crashes after pressing on refresh button | WIP8 | WIP9 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP9SegReq | Ingenuity 64 |  | WIP9 | August 3, 2018 11:55:45 AM GMT+08:00
CTUDT00598826 | Refresh | Closed | bagdady_b | June 21, 2018 10:45:04 PM GMT+08:00 | Refresh- Logviewer is not responding and crashes after pressing on refresh button | WIP8 | WIP9 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP9SegReq | IQon1 |  | WIP9 | August 3, 2018 11:55:45 AM GMT+08:00
CTUDT00598824 | Refresh | Closed | naik_p | June 21, 2018 9:13:30 PM GMT+08:00 | Cardiac Perfusion- 2 ECG wave series are getting created when EditBEforeFinalRecon is unchecked | WIP8 | WIP17 |  | Subsystem | Cardiac Perfusion | NA | Fixed | tejaswini_nidamanuri | 0 | 4 | Console Software | Console4 | NO | Closed | WIP14SubCom | Ovation |  | WIP17 | April 29, 2019 7:39:00 PM GMT+08:00
CTUDT00598824 | Refresh | Closed | naik_p | June 21, 2018 9:13:30 PM GMT+08:00 | Cardiac Perfusion- 2 ECG wave series are getting created when EditBEforeFinalRecon is unchecked | WIP8 | WIP17 |  | Subsystem | Cardiac Perfusion | NA | Fixed | tejaswini_nidamanuri | 0 | 4 | Console Software | Console4 | NO | Closed | WIP14SubCom | IQon1 |  | WIP17 | April 29, 2019 7:39:00 PM GMT+08:00
CTUDT00598820 | Refresh | Closed | naik_p | June 21, 2018 8:30:57 PM GMT+08:00 | Console screen goes blank when ending exam, exiting from view2, EndExamAndContinue. Basically it is taking lot of time to refresh (~30 sec). | WIP8 |  |  | Subsystem | End Exam Continue Working - UI Performance |  | Works As Specified / Not A Defect | sachan_a | 0 |  |  | Console4 | NO | Closed | WIP9SegReq | Ovation
CTUDT00598820 | Refresh | Closed | naik_p | June 21, 2018 8:30:57 PM GMT+08:00 | Console screen goes blank when ending exam, exiting from view2, EndExamAndContinue. Basically it is taking lot of time to refresh (~30 sec). | WIP8 |  |  | Subsystem | End Exam Continue Working - UI Performance |  | Works As Specified / Not A Defect | sachan_a | 0 |  |  | Console4 | NO | Closed | WIP9SegReq | Ingenuity 64
CTUDT00598820 | Refresh | Closed | naik_p | June 21, 2018 8:30:57 PM GMT+08:00 | Console screen goes blank when ending exam, exiting from view2, EndExamAndContinue. Basically it is taking lot of time to refresh (~30 sec). | WIP8 |  |  | Subsystem | End Exam Continue Working - UI Performance |  | Works As Specified / Not A Defect | sachan_a | 0 |  |  | Console4 | NO | Closed | WIP9SegReq | IQon1
CTUDT00598818 | Refresh | Closed | naik_p | June 21, 2018 8:07:19 PM GMT+08:00 | In Clinical User, clicking on logout, system did not logged off . | WIP8 | WIP9 |  | Subsystem |  | NA | Fixed | kapadia_v | 0 | 3 | Console Software | Console4 | NO | Closed | WIP9SegReq | Ovation |  | WIP9 | August 7, 2018 3:08:51 PM GMT+08:00
CTUDT00598818 | Refresh | Closed | naik_p | June 21, 2018 8:07:19 PM GMT+08:00 | In Clinical User, clicking on logout, system did not logged off . | WIP8 | WIP9 |  | Subsystem |  | NA | Fixed | kapadia_v | 0 | 3 | Console Software | Console4 | NO | Closed | WIP9SegReq | Brilliance 64 |  | WIP9 | August 7, 2018 3:08:51 PM GMT+08:00
CTUDT00598818 | Refresh | Closed | naik_p | June 21, 2018 8:07:19 PM GMT+08:00 | In Clinical User, clicking on logout, system did not logged off . | WIP8 | WIP9 |  | Subsystem |  | NA | Fixed | kapadia_v | 0 | 3 | Console Software | Console4 | NO | Closed | WIP9SegReq | Ingenuity 64 |  | WIP9 | August 7, 2018 3:08:51 PM GMT+08:00
CTUDT00598818 | Refresh | Closed | naik_p | June 21, 2018 8:07:19 PM GMT+08:00 | In Clinical User, clicking on logout, system did not logged off . | WIP8 | WIP9 |  | Subsystem |  | NA | Fixed | kapadia_v | 0 | 3 | Console Software | Console4 | NO | Closed | WIP9SegReq | IQon1 |  | WIP9 | August 7, 2018 3:08:51 PM GMT+08:00
CTUDT00598813 | Refresh | Closed | naik_p | June 21, 2018 7:08:18 PM GMT+08:00 | Copyright version is still 2011 not changed to 2018 | WIP 8 | WIP9 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP9SubCom | Ovation |  | WIP9 | August 3, 2018 12:00:19 PM GMT+08:00
CTUDT00598813 | Refresh | Closed | naik_p | June 21, 2018 7:08:18 PM GMT+08:00 | Copyright version is still 2011 not changed to 2018 | WIP 8 | WIP9 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP9SubCom | Brilliance 64 |  | WIP9 | August 3, 2018 12:00:19 PM GMT+08:00
CTUDT00598813 | Refresh | Closed | naik_p | June 21, 2018 7:08:18 PM GMT+08:00 | Copyright version is still 2011 not changed to 2018 | WIP 8 | WIP9 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP9SubCom | Brilliance Big Bore |  | WIP9 | August 3, 2018 12:00:19 PM GMT+08:00
CTUDT00598813 | Refresh | Closed | naik_p | June 21, 2018 7:08:18 PM GMT+08:00 | Copyright version is still 2011 not changed to 2018 | WIP 8 | WIP9 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP9SubCom | Ingenuity 64 |  | WIP9 | August 3, 2018 12:00:19 PM GMT+08:00
CTUDT00598813 | Refresh | Closed | naik_p | June 21, 2018 7:08:18 PM GMT+08:00 | Copyright version is still 2011 not changed to 2018 | WIP 8 | WIP9 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP9SubCom | IQon1 |  | WIP9 | August 3, 2018 12:00:19 PM GMT+08:00
CTUDT00598810 | Refresh | Closed | naik_p | June 21, 2018 6:53:41 PM GMT+08:00 | Logging in to new 3rd party user gives error stating: "Failed to get session from the service manager" : | WIP 8 | WIP9 |  | Subsystem |  | N/A | Fixed | navale_s | 0 | 2 | Service Tools_Software | Console4 | NO | Closed | WIP9SegReq | Ovation |  | WIP9 | August 2, 2018 12:32:33 PM GMT+08:00
CTUDT00598810 | Refresh | Closed | naik_p | June 21, 2018 6:53:41 PM GMT+08:00 | Logging in to new 3rd party user gives error stating: "Failed to get session from the service manager" : | WIP 8 | WIP9 |  | Subsystem |  | N/A | Fixed | navale_s | 0 | 2 | Service Tools_Software | Console4 | NO | Closed | WIP9SegReq | Brilliance 64 |  | WIP9 | August 2, 2018 12:32:33 PM GMT+08:00
CTUDT00598810 | Refresh | Closed | naik_p | June 21, 2018 6:53:41 PM GMT+08:00 | Logging in to new 3rd party user gives error stating: "Failed to get session from the service manager" : | WIP 8 | WIP9 |  | Subsystem |  | N/A | Fixed | navale_s | 0 | 2 | Service Tools_Software | Console4 | NO | Closed | WIP9SegReq | Ingenuity 64 |  | WIP9 | August 2, 2018 12:32:33 PM GMT+08:00
CTUDT00598810 | Refresh | Closed | naik_p | June 21, 2018 6:53:41 PM GMT+08:00 | Logging in to new 3rd party user gives error stating: "Failed to get session from the service manager" : | WIP 8 | WIP9 |  | Subsystem |  | N/A | Fixed | navale_s | 0 | 2 | Service Tools_Software | Console4 | NO | Closed | WIP9SegReq | IQon1 |  | WIP9 | August 2, 2018 12:32:33 PM GMT+08:00
CTUDT00598808 | Refresh | Closed | naik_p | June 21, 2018 6:45:08 PM GMT+08:00 | Moving to next option using “Tab” is not working in Dose alert messages and in create user window. | WIP 8 | WIP9 | No | Subsystem |  | Console-SRS.8866 | Fixed | sowmya_k | 0 | 4 | Console Software | Console4 | NO | Closed | WIP10SegReq | Ovation |  | WIP9 | July 24, 2018 3:28:59 PM GMT+08:00
CTUDT00598808 | Refresh | Closed | naik_p | June 21, 2018 6:45:08 PM GMT+08:00 | Moving to next option using “Tab” is not working in Dose alert messages and in create user window. | WIP 8 | WIP9 | No | Subsystem |  | Console-SRS.8866 | Fixed | sowmya_k | 0 | 4 | Console Software | Console4 | NO | Closed | WIP10SegReq | Brilliance 64 |  | WIP9 | July 24, 2018 3:28:59 PM GMT+08:00
CTUDT00598808 | Refresh | Closed | naik_p | June 21, 2018 6:45:08 PM GMT+08:00 | Moving to next option using “Tab” is not working in Dose alert messages and in create user window. | WIP 8 | WIP9 | No | Subsystem |  | Console-SRS.8866 | Fixed | sowmya_k | 0 | 4 | Console Software | Console4 | NO | Closed | WIP10SegReq | Ingenuity 64 |  | WIP9 | July 24, 2018 3:28:59 PM GMT+08:00
CTUDT00598808 | Refresh | Closed | naik_p | June 21, 2018 6:45:08 PM GMT+08:00 | Moving to next option using “Tab” is not working in Dose alert messages and in create user window. | WIP 8 | WIP9 | No | Subsystem |  | Console-SRS.8866 | Fixed | sowmya_k | 0 | 4 | Console Software | Console4 | NO | Closed | WIP10SegReq | IQon1 |  | WIP9 | July 24, 2018 3:28:59 PM GMT+08:00
CTUDT00598806 | Refresh | Closed | naik_p | June 21, 2018 6:22:54 PM GMT+08:00 | When login as clinical user without ECM permission, then tried to give senior tech UN & PW it gave “UN and PW is incorrect” message instead of giving user understandable message. | WIP 8 | WIP9 | No | Subsystem |  | N/A | Fixed | sowmya_k | 0 | 2 | Console Software | Console4 | NO | Closed | WIP9SegReq | Ovation |  | WIP9 | July 24, 2018 3:27:26 PM GMT+08:00
CTUDT00598806 | Refresh | Closed | naik_p | June 21, 2018 6:22:54 PM GMT+08:00 | When login as clinical user without ECM permission, then tried to give senior tech UN & PW it gave “UN and PW is incorrect” message instead of giving user understandable message. | WIP 8 | WIP9 | No | Subsystem |  | N/A | Fixed | sowmya_k | 0 | 2 | Console Software | Console4 | NO | Closed | WIP9SegReq | Brilliance 64 |  | WIP9 | July 24, 2018 3:27:26 PM GMT+08:00
CTUDT00598806 | Refresh | Closed | naik_p | June 21, 2018 6:22:54 PM GMT+08:00 | When login as clinical user without ECM permission, then tried to give senior tech UN & PW it gave “UN and PW is incorrect” message instead of giving user understandable message. | WIP 8 | WIP9 | No | Subsystem |  | N/A | Fixed | sowmya_k | 0 | 2 | Console Software | Console4 | NO | Closed | WIP9SegReq | Ingenuity 64 |  | WIP9 | July 24, 2018 3:27:26 PM GMT+08:00
CTUDT00598806 | Refresh | Closed | naik_p | June 21, 2018 6:22:54 PM GMT+08:00 | When login as clinical user without ECM permission, then tried to give senior tech UN & PW it gave “UN and PW is incorrect” message instead of giving user understandable message. | WIP 8 | WIP9 | No | Subsystem |  | N/A | Fixed | sowmya_k | 0 | 2 | Console Software | Console4 | NO | Closed | WIP9SegReq | IQon1 |  | WIP9 | July 24, 2018 3:27:26 PM GMT+08:00
CTUDT00598804 | Refresh | Closed | naik_p | June 21, 2018 6:17:58 PM GMT+08:00 | For every offline recon of non gated results, new pulmo and cardiac series 299 is getting generated even though phase or any of the parameters are not changed. | WIP 8 |  |  | Subsystem |  |  | Duplicate | sanjeeva_a | 0 |  |  |  |  | Closed |  | Ingenuity 64
CTUDT00598802 | Refresh | Closed | naik_p | June 21, 2018 6:06:36 PM GMT+08:00 | For every offline recon of non gated results, new pulmo and cardiac series 299 is getting generated even though phase or any parameter is not changed. | WIP 8 |  | No | System | 4D CT Scan | NA | No Plan to Fix | sagar_v | 0 | 4 |  | Console4 | NO | Closed | SubNC | Ovation
CTUDT00598802 | Refresh | Closed | naik_p | June 21, 2018 6:06:36 PM GMT+08:00 | For every offline recon of non gated results, new pulmo and cardiac series 299 is getting generated even though phase or any parameter is not changed. | WIP 8 |  | No | System | 4D CT Scan | NA | No Plan to Fix | sagar_v | 0 | 4 |  | Console4 | NO | Closed | SubNC | Brilliance 64
CTUDT00598802 | Refresh | Closed | naik_p | June 21, 2018 6:06:36 PM GMT+08:00 | For every offline recon of non gated results, new pulmo and cardiac series 299 is getting generated even though phase or any parameter is not changed. | WIP 8 |  | No | System | 4D CT Scan | NA | No Plan to Fix | sagar_v | 0 | 4 |  | Console4 | NO | Closed | SubNC | Ingenuity 64
CTUDT00598802 | Refresh | Closed | naik_p | June 21, 2018 6:06:36 PM GMT+08:00 | For every offline recon of non gated results, new pulmo and cardiac series 299 is getting generated even though phase or any parameter is not changed. | WIP 8 |  | No | System | 4D CT Scan | NA | No Plan to Fix | sagar_v | 0 | 4 |  | Console4 | NO | Closed | SubNC | IQon1
CTUDT00598800 | Refresh | Closed | naik_p | June 21, 2018 5:33:31 PM GMT+08:00 | Workflow is crashing in clinical user always, when login as clinical user, should wait for 20sec to launch workflow. | WIP 8 | WIP9 | Yes. User needs to wait for some time before clicking OK button on license agreement window. | Subsystem |  | NA | Fixed | goswami_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP9SegReq | Ovation |  | WIP9 | August 2, 2018 2:37:37 PM GMT+08:00
CTUDT00598800 | Refresh | Closed | naik_p | June 21, 2018 5:33:31 PM GMT+08:00 | Workflow is crashing in clinical user always, when login as clinical user, should wait for 20sec to launch workflow. | WIP 8 | WIP9 | Yes. User needs to wait for some time before clicking OK button on license agreement window. | Subsystem |  | NA | Fixed | goswami_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP9SegReq | Brilliance 64 |  | WIP9 | August 2, 2018 2:37:37 PM GMT+08:00
CTUDT00598800 | Refresh | Closed | naik_p | June 21, 2018 5:33:31 PM GMT+08:00 | Workflow is crashing in clinical user always, when login as clinical user, should wait for 20sec to launch workflow. | WIP 8 | WIP9 | Yes. User needs to wait for some time before clicking OK button on license agreement window. | Subsystem |  | NA | Fixed | goswami_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP9SegReq | Ingenuity 64 |  | WIP9 | August 2, 2018 2:37:37 PM GMT+08:00
CTUDT00598800 | Refresh | Closed | naik_p | June 21, 2018 5:33:31 PM GMT+08:00 | Workflow is crashing in clinical user always, when login as clinical user, should wait for 20sec to launch workflow. | WIP 8 | WIP9 | Yes. User needs to wait for some time before clicking OK button on license agreement window. | Subsystem |  | NA | Fixed | goswami_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP9SegReq | IQon1 |  | WIP9 | August 2, 2018 2:37:37 PM GMT+08:00
CTUDT00598798 | Refresh | Closed | naik_p | June 21, 2018 5:23:00 PM GMT+08:00 | Workflow is crashing in Philips service user. On every launch of workflow, getting critical error which is blocking the testing in service user. | WIP 8 | WIP8.1 | Yes. User needs to wait for some time to launch the workflow. | Subsystem |  | NA | Fixed | kapadia_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP9SegReq | Ovation |  | WIP8.1 | August 2, 2018 2:39:03 PM GMT+08:00
CTUDT00598798 | Refresh | Closed | naik_p | June 21, 2018 5:23:00 PM GMT+08:00 | Workflow is crashing in Philips service user. On every launch of workflow, getting critical error which is blocking the testing in service user. | WIP 8 | WIP8.1 | Yes. User needs to wait for some time to launch the workflow. | Subsystem |  | NA | Fixed | kapadia_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP9SegReq | Brilliance 64 |  | WIP8.1 | August 2, 2018 2:39:03 PM GMT+08:00
CTUDT00598798 | Refresh | Closed | naik_p | June 21, 2018 5:23:00 PM GMT+08:00 | Workflow is crashing in Philips service user. On every launch of workflow, getting critical error which is blocking the testing in service user. | WIP 8 | WIP8.1 | Yes. User needs to wait for some time to launch the workflow. | Subsystem |  | NA | Fixed | kapadia_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP9SegReq | Ingenuity 64 |  | WIP8.1 | August 2, 2018 2:39:03 PM GMT+08:00
CTUDT00598798 | Refresh | Closed | naik_p | June 21, 2018 5:23:00 PM GMT+08:00 | Workflow is crashing in Philips service user. On every launch of workflow, getting critical error which is blocking the testing in service user. | WIP 8 | WIP8.1 | Yes. User needs to wait for some time to launch the workflow. | Subsystem |  | NA | Fixed | kapadia_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP9SegReq | IQon1 |  | WIP8.1 | August 2, 2018 2:39:03 PM GMT+08:00
CTUDT00598796 | Refresh | Closed | manjunath_c | June 21, 2018 12:59:27 PM GMT+08:00 | Assign accession number available for offline result | WIP6.1 | WIP12 | NA | Subsystem | Result deletion | N/A | Fixed | chokkalingam_r | 0 | 2 | Console Software | Console4 | NO | Closed | WIP12SegReq | Ovation |  | WIP12 | November 28, 2018 3:00:15 PM GMT+08:00
CTUDT00598796 | Refresh | Closed | manjunath_c | June 21, 2018 12:59:27 PM GMT+08:00 | Assign accession number available for offline result | WIP6.1 | WIP12 | NA | Subsystem | Result deletion | N/A | Fixed | chokkalingam_r | 0 | 2 | Console Software | Console4 | NO | Closed | WIP12SegReq | Brilliance 64 |  | WIP12 | November 28, 2018 3:00:15 PM GMT+08:00
CTUDT00598796 | Refresh | Closed | manjunath_c | June 21, 2018 12:59:27 PM GMT+08:00 | Assign accession number available for offline result | WIP6.1 | WIP12 | NA | Subsystem | Result deletion | N/A | Fixed | chokkalingam_r | 0 | 2 | Console Software | Console4 | NO | Closed | WIP12SegReq | Ingenuity 64 |  | WIP12 | November 28, 2018 3:00:15 PM GMT+08:00
CTUDT00598796 | Refresh | Closed | manjunath_c | June 21, 2018 12:59:27 PM GMT+08:00 | Assign accession number available for offline result | WIP6.1 | WIP12 | NA | Subsystem | Result deletion | N/A | Fixed | chokkalingam_r | 0 | 2 | Console Software | Console4 | NO | Closed | WIP12SegReq | IQon1 |  | WIP12 | November 28, 2018 3:00:15 PM GMT+08:00
CTUDT00598794 | Refresh | Closed | bagdady_b | June 21, 2018 1:19:10 AM GMT+08:00 | Refresh- Importing large quantity of exam cards result in constant Exam application crash | WIP8 | Host SSRS DHF303211 Rev 04 |  | Subsystem |  | NA | Fixed | bandil_j | 0 | 3 | Host SSRS | Console4 | NO | Closed | WIP11SegReq | IQon1 |  | Host SSRS DHF303211 Rev 04 | October 22, 2018 4:29:41 PM GMT+08:00
CTUDT00598792 | Refresh | Closed | bagdady_b | June 21, 2018 1:12:53 AM GMT+08:00 | Refresh- Wrong indication of Total failed items while performing restore | WIP8 | WIP17 |  | Subsystem | Backup / Restore - failed items list | N/A | Fixed | chakraborty_d | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP9_SegReq | Ovation |  | WIP17 | April 29, 2019 2:31:14 PM GMT+08:00
CTUDT00598792 | Refresh | Closed | bagdady_b | June 21, 2018 1:12:53 AM GMT+08:00 | Refresh- Wrong indication of Total failed items while performing restore | WIP8 | WIP17 |  | Subsystem | Backup / Restore - failed items list | N/A | Fixed | chakraborty_d | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP9_SegReq | Brilliance 64 |  | WIP17 | April 29, 2019 2:31:14 PM GMT+08:00
CTUDT00598792 | Refresh | Closed | bagdady_b | June 21, 2018 1:12:53 AM GMT+08:00 | Refresh- Wrong indication of Total failed items while performing restore | WIP8 | WIP17 |  | Subsystem | Backup / Restore - failed items list | N/A | Fixed | chakraborty_d | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP9_SegReq | Ingenuity 64 |  | WIP17 | April 29, 2019 2:31:14 PM GMT+08:00
CTUDT00598792 | Refresh | Closed | bagdady_b | June 21, 2018 1:12:53 AM GMT+08:00 | Refresh- Wrong indication of Total failed items while performing restore | WIP8 | WIP17 |  | Subsystem | Backup / Restore - failed items list | N/A | Fixed | chakraborty_d | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP9_SegReq | IQon1 |  | WIP17 | April 29, 2019 2:31:14 PM GMT+08:00
CTUDT00598790 | Refresh | Closed | bagdady_b | June 21, 2018 12:47:01 AM GMT+08:00 | Refresh- Nonstandard installation appears in Component Validation | WIP8 | WIP8.1 |  | Subsystem | Verified1 | CT-NM.SRS.Service.1724 | Fixed | venugopal_s | 0 | 2 | Service Tools_Software | DI | NO | Closed | WIP9SegReq | Ovation |  | WIP8.1 | August 30, 2018 7:57:33 PM GMT+08:00
CTUDT00598790 | Refresh | Closed | bagdady_b | June 21, 2018 12:47:01 AM GMT+08:00 | Refresh- Nonstandard installation appears in Component Validation | WIP8 | WIP8.1 |  | Subsystem | Verified1 | CT-NM.SRS.Service.1724 | Fixed | venugopal_s | 0 | 2 | Service Tools_Software | DI | NO | Closed | WIP9SegReq | Brilliance 64 |  | WIP8.1 | August 30, 2018 7:57:33 PM GMT+08:00
CTUDT00598790 | Refresh | Closed | bagdady_b | June 21, 2018 12:47:01 AM GMT+08:00 | Refresh- Nonstandard installation appears in Component Validation | WIP8 | WIP8.1 |  | Subsystem | Verified1 | CT-NM.SRS.Service.1724 | Fixed | venugopal_s | 0 | 2 | Service Tools_Software | DI | NO | Closed | WIP9SegReq | Ingenuity 64 |  | WIP8.1 | August 30, 2018 7:57:33 PM GMT+08:00
CTUDT00598790 | Refresh | Closed | bagdady_b | June 21, 2018 12:47:01 AM GMT+08:00 | Refresh- Nonstandard installation appears in Component Validation | WIP8 | WIP8.1 |  | Subsystem | Verified1 | CT-NM.SRS.Service.1724 | Fixed | venugopal_s | 0 | 2 | Service Tools_Software | DI | NO | Closed | WIP9SegReq | IQon1 |  | WIP8.1 | August 30, 2018 7:57:33 PM GMT+08:00
CTUDT00598788 | Refresh | Closed | bagdady_b | June 21, 2018 12:44:31 AM GMT+08:00 | Refresh- Eagle name appears for the IQON product during installation | WIP8 | WIP8.1 |  | Subsystem |  | NA | Fixed | venugopal_s | 0 | 4 | Service Tools_Software | DI | NO | Closed | WIP10SegReq | IQon1 |  | WIP8.1 | September 5, 2018 5:21:56 PM GMT+08:00
CTUDT00598786 | Refresh | Closed | bagdady_b | June 21, 2018 12:40:07 AM GMT+08:00 | Refresh- Wrong version appears for Console.SW.Host in System Configuration | WIP8 | WIP8.1 |  | Subsystem | Verified1 | Console-SRS.6105 | Fixed | venugopal_s | 0 | 2 | Service Tools_Software | DI | NO | Closed | SubCom_WIP8.1 | Ovation |  | WIP8.1 | July 18, 2018 1:16:44 PM GMT+08:00
CTUDT00598786 | Refresh | Closed | bagdady_b | June 21, 2018 12:40:07 AM GMT+08:00 | Refresh- Wrong version appears for Console.SW.Host in System Configuration | WIP8 | WIP8.1 |  | Subsystem | Verified1 | Console-SRS.6105 | Fixed | venugopal_s | 0 | 2 | Service Tools_Software | DI | NO | Closed | SubCom_WIP8.1 | Brilliance 64 |  | WIP8.1 | July 18, 2018 1:16:44 PM GMT+08:00
CTUDT00598786 | Refresh | Closed | bagdady_b | June 21, 2018 12:40:07 AM GMT+08:00 | Refresh- Wrong version appears for Console.SW.Host in System Configuration | WIP8 | WIP8.1 |  | Subsystem | Verified1 | Console-SRS.6105 | Fixed | venugopal_s | 0 | 2 | Service Tools_Software | DI | NO | Closed | SubCom_WIP8.1 | Ingenuity 64 |  | WIP8.1 | July 18, 2018 1:16:44 PM GMT+08:00
CTUDT00598786 | Refresh | Closed | bagdady_b | June 21, 2018 12:40:07 AM GMT+08:00 | Refresh- Wrong version appears for Console.SW.Host in System Configuration | WIP8 | WIP8.1 |  | Subsystem | Verified1 | Console-SRS.6105 | Fixed | venugopal_s | 0 | 2 | Service Tools_Software | DI | NO | Closed | SubCom_WIP8.1 | IQon1 |  | WIP8.1 | July 18, 2018 1:16:44 PM GMT+08:00
CTUDT00598784 | Refresh | Closed | bagdady_b | June 21, 2018 12:35:30 AM GMT+08:00 | Refresh- Full air calibration in CT user is failing | WIP8 | WIP9 |  | Subsystem |  | N/A | Fixed | yosefk | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP9SegReq | IQon1 |  | WIP9 | November 6, 2018 9:41:42 PM GMT+08:00
CTUDT00598782 | Refresh | Closed | bagdady_b | June 21, 2018 12:26:18 AM GMT+08:00 | Refresh- Not all images created pop up when multiple conventional and spectral results | WIP8 |  |  | Subsystem | Spectral reconstruction | NA | Duplicate | richards_b | 0 |  |  | CIRS | NO | Closed |  | IQon1
CTUDT00598780 | Refresh | Closed | beck_d | June 20, 2018 10:37:03 PM GMT+08:00 | Calibrations: Phantom Calibration: Calibration is stuck in the first Ultra High case | WIP 8 |  |  | Subsystem |  | N/A | Duplicate | frider_o | 0 |  |  | DMS | NO | Closed |  | Ingenuity 64
CTUDT00598771 | Refresh | Closed | manjunath_c | June 20, 2018 6:27:06 PM GMT+08:00 | Deviation in the default parameter values displayed in Exam card manager compared to what is mentioned in the "DHF251245 Rev01 Brilliance64 v4.8 Reference Exam Cards" Specification | WIP6.1 | DHF318051 Rev 10 HOST Verification Test Protocol |  | Subsystem | Reference Ecs | N/A | Fixed | shekhar_p | 0 | 3 | Test Design | Console4 | NO | Closed | BB_WIP5_SegReq | Brilliance 64 |  | DHF318051 Rev 10 HOST Verification Test Protocol | March 20, 2019 12:27:20 PM GMT+08:00
CTUDT00598770 | Refresh | Closed | manjunath_c | June 20, 2018 6:23:50 PM GMT+08:00 | Deviation in the default parameter values displayed in Exam card manager compared to what is mentioned in the "DHF251235 Rev01 Ingenuity v4.8 Reference Exam Cards" Specification | WIP6.1 | DHF318051 Rev 10 HOST Verification Test Protocol |  | Subsystem | Reference Ecs | N/A | Fixed | shekhar_p | 0 | 3 | Test Design | Console4 | NO | Closed | BB_WIP5_SegReq | Ingenuity 64 |  | DHF318051 Rev 10 HOST Verification Test Protocol | March 20, 2019 12:26:36 PM GMT+08:00
CTUDT00598769 | Refresh | Closed | manjunath_c | June 20, 2018 6:18:15 PM GMT+08:00 | Deviation in the default parameter values displayed in Exam card manager compared to what is mentioned in the "DHF316206 Rev00 iCT SP v4.8 Reference Exam Cards" Specification | WIP6.1 | DHF318051 Rev 10 HOST Verification Test Protocol |  | Subsystem | Reference Ecs | N/A | Fixed | shekhar_p | 0 | 2 | Test Design | Console4 | NO | Closed | WIP14SegReq | Ovation |  | DHF318051 Rev 10 HOST Verification Test Protocol | March 20, 2019 12:25:56 PM GMT+08:00
CTUDT00598768 | Refresh | Closed | manjunath_c | June 20, 2018 6:13:17 PM GMT+08:00 | Deviation in the default parameter values displayed in Exam card manager compared to what is mentioned in the "DHF316204 Rev00 iCT v4.8 Reference Exam Cards" Specification | WIP6.1 | DHF318051 Rev 10 HOST Verification Test Protocol |  | Subsystem | Reference Ecs | N/A | Fixed | shekhar_p | 0 | 2 | Test Design | Console4 | NO | Closed | WIP14SegReq | Ovation |  | DHF318051 Rev 10 HOST Verification Test Protocol | March 20, 2019 12:25:08 PM GMT+08:00
CTUDT00598764 | Refresh | Closed | beck | June 20, 2018 4:09:43 AM GMT+08:00 | Unable to perform Final Reconstruction | v4.1.10.xx025 | WIP8.1 | NA | Subsystem |  | NA | Fixed | ravindra_a | 0 | 3 | Console Software | Console4 | NO | Closed | WIP9_SegReq | Brilliance 64 |  | WIP8.1 | July 17, 2018 4:04:41 PM GMT+08:00
CTUDT00598751 | Refresh | Closed | beck_d | June 18, 2018 9:06:30 PM GMT+08:00 | Refresh: Preferences: Image Balance, Slice Normalization and Phantom Calibration check boxes do not appear under Reconstruction in the Preferences | WIP 7 | WIP8.1 | No | Subsystem |  | N/A | Fixed | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP9SegReq | Ovation |  | WIP8.1 | August 23, 2018 1:54:02 PM GMT+08:00
CTUDT00598751 | Refresh | Closed | beck_d | June 18, 2018 9:06:30 PM GMT+08:00 | Refresh: Preferences: Image Balance, Slice Normalization and Phantom Calibration check boxes do not appear under Reconstruction in the Preferences | WIP 7 | WIP8.1 | No | Subsystem |  | N/A | Fixed | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP9SegReq | Brilliance 64 |  | WIP8.1 | August 23, 2018 1:54:02 PM GMT+08:00
CTUDT00598751 | Refresh | Closed | beck_d | June 18, 2018 9:06:30 PM GMT+08:00 | Refresh: Preferences: Image Balance, Slice Normalization and Phantom Calibration check boxes do not appear under Reconstruction in the Preferences | WIP 7 | WIP8.1 | No | Subsystem |  | N/A | Fixed | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP9SegReq | Brilliance Big Bore |  | WIP8.1 | August 23, 2018 1:54:02 PM GMT+08:00
CTUDT00598751 | Refresh | Closed | beck_d | June 18, 2018 9:06:30 PM GMT+08:00 | Refresh: Preferences: Image Balance, Slice Normalization and Phantom Calibration check boxes do not appear under Reconstruction in the Preferences | WIP 7 | WIP8.1 | No | Subsystem |  | N/A | Fixed | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP9SegReq | Ingenuity 64 |  | WIP8.1 | August 23, 2018 1:54:02 PM GMT+08:00
CTUDT00598751 | Refresh | Closed | beck_d | June 18, 2018 9:06:30 PM GMT+08:00 | Refresh: Preferences: Image Balance, Slice Normalization and Phantom Calibration check boxes do not appear under Reconstruction in the Preferences | WIP 7 | WIP8.1 | No | Subsystem |  | N/A | Fixed | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP9SegReq | IQon1 |  | WIP8.1 | August 23, 2018 1:54:02 PM GMT+08:00
CTUDT00598747 | Refresh | Closed | jeffr | June 18, 2018 1:22:22 PM GMT+08:00 | iCT collimator CPM to baseboard communication failure | 4.1.7 | WIP18 | cycle power on the gantry | Subsystem |  | NA | Fixed | chen_ho | 0 | 3 | Gantry Software | GCX | YES | Closed | WIP15SegReq | Ovation |  | WIP18 | June 18, 2019 3:16:52 PM GMT+08:00
CTUDT00598746 | Refresh | Closed | rodan_d | June 17, 2018 6:26:04 PM GMT+08:00 | Spelling error in Pulmonary Edit Phases | Ingenuity_TF64-128 - RfVD WIP 108 |  |  | Subsystem |  |  | Change Proposal | sanjeeva_a | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00598745 | Refresh | Closed | rodan_d | June 17, 2018 6:24:58 PM GMT+08:00 | Multiple dose SR generated for a study | Ingenuity_TF64-128 - RfVD WIP 413 |  |  | Subsystem |  | NA | Duplicate | sowmya_k | 0 | 2 |  | Console4 | NO | Closed | WIP15SegReq | Ovation
CTUDT00598745 | Refresh | Closed | rodan_d | June 17, 2018 6:24:58 PM GMT+08:00 | Multiple dose SR generated for a study | Ingenuity_TF64-128 - RfVD WIP 413 |  |  | Subsystem |  | NA | Duplicate | sowmya_k | 0 | 2 |  | Console4 | NO | Closed | WIP15SegReq | Brilliance 64
CTUDT00598745 | Refresh | Closed | rodan_d | June 17, 2018 6:24:58 PM GMT+08:00 | Multiple dose SR generated for a study | Ingenuity_TF64-128 - RfVD WIP 413 |  |  | Subsystem |  | NA | Duplicate | sowmya_k | 0 | 2 |  | Console4 | NO | Closed | WIP15SegReq | Brilliance Big Bore
CTUDT00598745 | Refresh | Closed | rodan_d | June 17, 2018 6:24:58 PM GMT+08:00 | Multiple dose SR generated for a study | Ingenuity_TF64-128 - RfVD WIP 413 |  |  | Subsystem |  | NA | Duplicate | sowmya_k | 0 | 2 |  | Console4 | NO | Closed | WIP15SegReq | Ingenuity 64
CTUDT00598745 | Refresh | Closed | rodan_d | June 17, 2018 6:24:58 PM GMT+08:00 | Multiple dose SR generated for a study | Ingenuity_TF64-128 - RfVD WIP 413 |  |  | Subsystem |  | NA | Duplicate | sowmya_k | 0 | 2 |  | Console4 | NO | Closed | WIP15SegReq | IQon1
CTUDT00598744 | Refresh | Closed | rodan_d | June 17, 2018 6:23:31 PM GMT+08:00 | IV: Ingenuity TF PET/CT - No error message after loss of waveform during rest PET cardiac exam | Ingenuity_TF64-128 - RfVD WIP 414 |  |  | Subsystem |  |  | Change Proposal | oren_d | 0 |  |  | Support-System | NO | Closed |  | Ingenuity 64
CTUDT00598743 | Refresh | Closed | rodan_d | June 17, 2018 6:22:05 PM GMT+08:00 | Dose Alert is not notified when having multiple Surviews in the Study | Ingenuity_TF64-128 - 4.0.3 |  | NA | Subsystem |  | Console-SRS-15995 | Does Not Apply | sagar_v | 0 | 2 |  | Console4 | YES | Closed |  | Ovation
CTUDT00598743 | Refresh | Closed | rodan_d | June 17, 2018 6:22:05 PM GMT+08:00 | Dose Alert is not notified when having multiple Surviews in the Study | Ingenuity_TF64-128 - 4.0.3 |  | NA | Subsystem |  | Console-SRS-15995 | Does Not Apply | sagar_v | 0 | 2 |  | Console4 | YES | Closed |  | Brilliance 64
CTUDT00598743 | Refresh | Closed | rodan_d | June 17, 2018 6:22:05 PM GMT+08:00 | Dose Alert is not notified when having multiple Surviews in the Study | Ingenuity_TF64-128 - 4.0.3 |  | NA | Subsystem |  | Console-SRS-15995 | Does Not Apply | sagar_v | 0 | 2 |  | Console4 | YES | Closed |  | Brilliance Big Bore
CTUDT00598743 | Refresh | Closed | rodan_d | June 17, 2018 6:22:05 PM GMT+08:00 | Dose Alert is not notified when having multiple Surviews in the Study | Ingenuity_TF64-128 - 4.0.3 |  | NA | Subsystem |  | Console-SRS-15995 | Does Not Apply | sagar_v | 0 | 2 |  | Console4 | YES | Closed |  | Ingenuity 64
CTUDT00598743 | Refresh | Closed | rodan_d | June 17, 2018 6:22:05 PM GMT+08:00 | Dose Alert is not notified when having multiple Surviews in the Study | Ingenuity_TF64-128 - 4.0.3 |  | NA | Subsystem |  | Console-SRS-15995 | Does Not Apply | sagar_v | 0 | 2 |  | Console4 | YES | Closed |  | IQon1
CTUDT00598742 | Refresh | Closed | rodan_d | June 17, 2018 6:14:29 PM GMT+08:00 | Exam Card Manager Group Sorting Changes Not Being Applied for PET EC Group | V4.0.3 |  |  | Subsystem |  |  | Duplicate | sanjeeva_a | 0 | 3 |  | Console4 | NO | Closed |  | PET - Ingenuity_TF64/128 | 6800964
CTUDT00598741 | Refresh | Closed | rodan_d | June 17, 2018 6:06:00 PM GMT+08:00 | Exam Card Manager Group Sorting Changes Not Being Applied for PET EC Group | V4.0.3 | WIP16 BB_WIP6 |  | Subsystem |  | NA | Fixed | kapadia_v | 0 | 3 | Console Software | Console4 | NO | Closed | WIP16SubCom | Ovation | 6800964 | WIP16 BB_WIP6 | April 19, 2019 6:51:01 PM GMT+08:00
CTUDT00598741 | Refresh | Closed | rodan_d | June 17, 2018 6:06:00 PM GMT+08:00 | Exam Card Manager Group Sorting Changes Not Being Applied for PET EC Group | V4.0.3 | WIP16 BB_WIP6 |  | Subsystem |  | NA | Fixed | kapadia_v | 0 | 3 | Console Software | Console4 | NO | Closed | WIP16SubCom | Brilliance 64 | 6800964 | WIP16 BB_WIP6 | April 19, 2019 6:51:01 PM GMT+08:00
CTUDT00598741 | Refresh | Closed | rodan_d | June 17, 2018 6:06:00 PM GMT+08:00 | Exam Card Manager Group Sorting Changes Not Being Applied for PET EC Group | V4.0.3 | WIP16 BB_WIP6 |  | Subsystem |  | NA | Fixed | kapadia_v | 0 | 3 | Console Software | Console4 | NO | Closed | WIP16SubCom | Brilliance Big Bore | 6800964 | WIP16 BB_WIP6 | April 19, 2019 6:51:01 PM GMT+08:00
CTUDT00598741 | Refresh | Closed | rodan_d | June 17, 2018 6:06:00 PM GMT+08:00 | Exam Card Manager Group Sorting Changes Not Being Applied for PET EC Group | V4.0.3 | WIP16 BB_WIP6 |  | Subsystem |  | NA | Fixed | kapadia_v | 0 | 3 | Console Software | Console4 | NO | Closed | WIP16SubCom | Ingenuity 64 | 6800964 | WIP16 BB_WIP6 | April 19, 2019 6:51:01 PM GMT+08:00
CTUDT00598741 | Refresh | Closed | rodan_d | June 17, 2018 6:06:00 PM GMT+08:00 | Exam Card Manager Group Sorting Changes Not Being Applied for PET EC Group | V4.0.3 | WIP16 BB_WIP6 |  | Subsystem |  | NA | Fixed | kapadia_v | 0 | 3 | Console Software | Console4 | NO | Closed | WIP16SubCom | IQon1 | 6800964 | WIP16 BB_WIP6 | April 19, 2019 6:51:01 PM GMT+08:00
CTUDT00598740 | Refresh | Closed | rodan_d | June 17, 2018 5:57:43 PM GMT+08:00 | IV: Ingenuity TF - Post Respiratory Wave missing at top of final images | Ingenuity_TF64-128 - RfVD WIP 413 | WIP16 | The wave can be seen when using "Edit Results" and while loading the CT image in Pulmo Viewer. | Subsystem |  | Console-SRS.15323 | Fixed | Prakash_Chand_Yadav | 0 | 2 | Documentation-User Manuals | Console4 | NO | Closed | WIP16SubCom | Brilliance 64 |  | WIP16 | April 17, 2019 12:50:48 PM GMT+08:00
CTUDT00598740 | Refresh | Closed | rodan_d | June 17, 2018 5:57:43 PM GMT+08:00 | IV: Ingenuity TF - Post Respiratory Wave missing at top of final images | Ingenuity_TF64-128 - RfVD WIP 413 | WIP16 | The wave can be seen when using "Edit Results" and while loading the CT image in Pulmo Viewer. | Subsystem |  | Console-SRS.15323 | Fixed | Prakash_Chand_Yadav | 0 | 2 | Documentation-User Manuals | Console4 | NO | Closed | WIP16SubCom | Ingenuity 64 |  | WIP16 | April 17, 2019 12:50:48 PM GMT+08:00
CTUDT00598739 | Refresh | Closed | rodan_d | June 17, 2018 5:51:34 PM GMT+08:00 | IV: Ingenuity TF - Post scan ECG wave mising at the top of final images | Ingenuity_TF64-128 - RfVD WIP 413 | WIP16 | The wave can be seen when using "Edit Results" and while loading the CT image in Cardiac Viewer. | Subsystem |  | NA | Fixed | Prakash_Chand_Yadav | 0 | 3 | Documentation-User Manuals | Console4 | NO | Closed | WIP16SubCom | Ovation |  | WIP16 | April 9, 2019 5:24:57 PM GMT+08:00
CTUDT00598739 | Refresh | Closed | rodan_d | June 17, 2018 5:51:34 PM GMT+08:00 | IV: Ingenuity TF - Post scan ECG wave mising at the top of final images | Ingenuity_TF64-128 - RfVD WIP 413 | WIP16 | The wave can be seen when using "Edit Results" and while loading the CT image in Cardiac Viewer. | Subsystem |  | NA | Fixed | Prakash_Chand_Yadav | 0 | 3 | Documentation-User Manuals | Console4 | NO | Closed | WIP16SubCom | Brilliance 64 |  | WIP16 | April 9, 2019 5:24:57 PM GMT+08:00
CTUDT00598739 | Refresh | Closed | rodan_d | June 17, 2018 5:51:34 PM GMT+08:00 | IV: Ingenuity TF - Post scan ECG wave mising at the top of final images | Ingenuity_TF64-128 - RfVD WIP 413 | WIP16 | The wave can be seen when using "Edit Results" and while loading the CT image in Cardiac Viewer. | Subsystem |  | NA | Fixed | Prakash_Chand_Yadav | 0 | 3 | Documentation-User Manuals | Console4 | NO | Closed | WIP16SubCom | Ingenuity 64 |  | WIP16 | April 9, 2019 5:24:57 PM GMT+08:00
CTUDT00598736 | Refresh | Closed | rodan_d | June 17, 2018 3:55:10 PM GMT+08:00 | IV: Ingenuity TF - Printed exam card DoseRight and Coronal result length settings inconsistent with GUI settings | Ingenuity_TF64-128 - RfVD WIP 413 |  |  | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 | Documentation-User Manuals | Console4 | NO | Closed |  | Ovation
CTUDT00598736 | Refresh | Closed | rodan_d | June 17, 2018 3:55:10 PM GMT+08:00 | IV: Ingenuity TF - Printed exam card DoseRight and Coronal result length settings inconsistent with GUI settings | Ingenuity_TF64-128 - RfVD WIP 413 |  |  | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 | Documentation-User Manuals | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00598736 | Refresh | Closed | rodan_d | June 17, 2018 3:55:10 PM GMT+08:00 | IV: Ingenuity TF - Printed exam card DoseRight and Coronal result length settings inconsistent with GUI settings | Ingenuity_TF64-128 - RfVD WIP 413 |  |  | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 | Documentation-User Manuals | Console4 | NO | Closed |  | IQon1
CTUDT00598735 | Refresh | Closed | rodan_d | June 17, 2018 3:53:49 PM GMT+08:00 | IV: Ingenuity TF PET/CT - Offline reconstruction error message | Ingenuity_TF64-128 - RfVD WIP 414 |  | Log out, Log In and retry the offline reconstruction. | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 3 | Documentation-User Manuals | Console4 | NO | Closed |  | PET - Ingenuity_TF64/128
CTUDT00598734 | Refresh | Closed | rodan_d | June 17, 2018 3:52:25 PM GMT+08:00 | IV: Solo - CT Acquisition started but did not finish | Ingenuity_TF64-128 - RfVD WIP 109 | WIP16 | During a CT or PET respiratory exam, an error message may be displayed "Resp. signal was lost..." After clicking "OK" the scan may continue and complete. Images and wave form should be reviewed to determine if a rescan is required. If the scan does not continue, then the exam will need to be re-scanned. | Subsystem |  | NA | Fixed | sowmya_k | 0 | 2 | Documentation-User Manuals | Console4 | YES | Closed | WIP16SubCom | Brilliance 64 |  | WIP16 | April 19, 2019 6:46:13 PM GMT+08:00
CTUDT00598734 | Refresh | Closed | rodan_d | June 17, 2018 3:52:25 PM GMT+08:00 | IV: Solo - CT Acquisition started but did not finish | Ingenuity_TF64-128 - RfVD WIP 109 | WIP16 | During a CT or PET respiratory exam, an error message may be displayed "Resp. signal was lost..." After clicking "OK" the scan may continue and complete. Images and wave form should be reviewed to determine if a rescan is required. If the scan does not continue, then the exam will need to be re-scanned. | Subsystem |  | NA | Fixed | sowmya_k | 0 | 2 | Documentation-User Manuals | Console4 | YES | Closed | WIP16SubCom | Ingenuity 64 |  | WIP16 | April 19, 2019 6:46:13 PM GMT+08:00
CTUDT00598733 | Refresh | Closed | rodan_d | June 17, 2018 3:51:03 PM GMT+08:00 | IV: Solo – Dark images of reference Chest PE exam card | Ingenuity_TF64-128 - RfVD WIP 109 |  | The operator can manually change the window settings. | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 3 |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00598732 | Refresh | Closed | rodan_d | June 17, 2018 3:49:24 PM GMT+08:00 | IV: Solo - Images will not change between results in View tab | Ingenuity_TF64-128 - RfVD WIP 413 |  | It is recommended to wait until all images have transferred from CIRS prior to reviewing images in View Image or Quick Review. | Subsystem |  | Console.SRS.3323 | Does Not Apply | sagar_v | 0 | 3 | Documentation-User Manuals | Console4 | NO | Closed | WIP16SubCom | Ovation
CTUDT00598732 | Refresh | Closed | rodan_d | June 17, 2018 3:49:24 PM GMT+08:00 | IV: Solo - Images will not change between results in View tab | Ingenuity_TF64-128 - RfVD WIP 413 |  | It is recommended to wait until all images have transferred from CIRS prior to reviewing images in View Image or Quick Review. | Subsystem |  | Console.SRS.3323 | Does Not Apply | sagar_v | 0 | 3 | Documentation-User Manuals | Console4 | NO | Closed | WIP16SubCom | Brilliance 64
CTUDT00598732 | Refresh | Closed | rodan_d | June 17, 2018 3:49:24 PM GMT+08:00 | IV: Solo - Images will not change between results in View tab | Ingenuity_TF64-128 - RfVD WIP 413 |  | It is recommended to wait until all images have transferred from CIRS prior to reviewing images in View Image or Quick Review. | Subsystem |  | Console.SRS.3323 | Does Not Apply | sagar_v | 0 | 3 | Documentation-User Manuals | Console4 | NO | Closed | WIP16SubCom | Brilliance Big Bore
CTUDT00598732 | Refresh | Closed | rodan_d | June 17, 2018 3:49:24 PM GMT+08:00 | IV: Solo - Images will not change between results in View tab | Ingenuity_TF64-128 - RfVD WIP 413 |  | It is recommended to wait until all images have transferred from CIRS prior to reviewing images in View Image or Quick Review. | Subsystem |  | Console.SRS.3323 | Does Not Apply | sagar_v | 0 | 3 | Documentation-User Manuals | Console4 | NO | Closed | WIP16SubCom | Ingenuity 64
CTUDT00598732 | Refresh | Closed | rodan_d | June 17, 2018 3:49:24 PM GMT+08:00 | IV: Solo - Images will not change between results in View tab | Ingenuity_TF64-128 - RfVD WIP 413 |  | It is recommended to wait until all images have transferred from CIRS prior to reviewing images in View Image or Quick Review. | Subsystem |  | Console.SRS.3323 | Does Not Apply | sagar_v | 0 | 3 | Documentation-User Manuals | Console4 | NO | Closed | WIP16SubCom | IQon1
CTUDT00598731 | Refresh | Closed | rodan_d | June 17, 2018 3:47:55 PM GMT+08:00 | IV: Solo - Injection summary not present in exam summary | Ingenuity_TF64-128 - RfVD WIP 413 | WIP18 | Through additional testing, the workaround is to use the injector's stop button and not the injector's pause button to stop the injector before ending the exam to obtain the injection summary. | Subsystem |  | None | Fixed | datta_i | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SubCom | Ovation |  | WIP18 | June 7, 2019 7:46:15 PM GMT+08:00
CTUDT00598731 | Refresh | Closed | rodan_d | June 17, 2018 3:47:55 PM GMT+08:00 | IV: Solo - Injection summary not present in exam summary | Ingenuity_TF64-128 - RfVD WIP 413 | WIP18 | Through additional testing, the workaround is to use the injector's stop button and not the injector's pause button to stop the injector before ending the exam to obtain the injection summary. | Subsystem |  | None | Fixed | datta_i | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SubCom | Brilliance 64 |  | WIP18 | June 7, 2019 7:46:15 PM GMT+08:00
CTUDT00598731 | Refresh | Closed | rodan_d | June 17, 2018 3:47:55 PM GMT+08:00 | IV: Solo - Injection summary not present in exam summary | Ingenuity_TF64-128 - RfVD WIP 413 | WIP18 | Through additional testing, the workaround is to use the injector's stop button and not the injector's pause button to stop the injector before ending the exam to obtain the injection summary. | Subsystem |  | None | Fixed | datta_i | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SubCom | Ingenuity 64 |  | WIP18 | June 7, 2019 7:46:15 PM GMT+08:00
CTUDT00598731 | Refresh | Closed | rodan_d | June 17, 2018 3:47:55 PM GMT+08:00 | IV: Solo - Injection summary not present in exam summary | Ingenuity_TF64-128 - RfVD WIP 413 | WIP18 | Through additional testing, the workaround is to use the injector's stop button and not the injector's pause button to stop the injector before ending the exam to obtain the injection summary. | Subsystem |  | None | Fixed | datta_i | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SubCom | IQon1 |  | WIP18 | June 7, 2019 7:46:15 PM GMT+08:00
CTUDT00598730 | Refresh | Closed | rodan_d | June 17, 2018 3:39:55 PM GMT+08:00 | IV: Solo - Printed version of exam card did not contain some of the parameters configured in ECM. | Ingenuity_TF64-128 - RfVD WIP 109 |  |  | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 | Documentation-User Manuals | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00598728 | Refresh | Closed | rodan_d | June 17, 2018 3:38:37 PM GMT+08:00 | IV: Solo Cycle 2 - Brain result will not angle to the right | Ingenuity_TF64-128 - RfVD WIP 413 | WIP14 | For Mouse rotation from zero degrees(clockwise will not work) - So use keyboard For Keyboard typing rotation from zero degrees (anti clockwise will not work) - Use mouse for rotation. | System |  | None | Fixed | patra_r | 0 | 2 | Console Software | Console4 | YES | Closed | WIP14SegReq | Ovation |  | WIP14 | February 27, 2019 6:18:16 PM GMT+08:00
CTUDT00598728 | Refresh | Closed | rodan_d | June 17, 2018 3:38:37 PM GMT+08:00 | IV: Solo Cycle 2 - Brain result will not angle to the right | Ingenuity_TF64-128 - RfVD WIP 413 | WIP14 | For Mouse rotation from zero degrees(clockwise will not work) - So use keyboard For Keyboard typing rotation from zero degrees (anti clockwise will not work) - Use mouse for rotation. | System |  | None | Fixed | patra_r | 0 | 2 | Console Software | Console4 | YES | Closed | WIP14SegReq | Brilliance 64 |  | WIP14 | February 27, 2019 6:18:16 PM GMT+08:00
CTUDT00598728 | Refresh | Closed | rodan_d | June 17, 2018 3:38:37 PM GMT+08:00 | IV: Solo Cycle 2 - Brain result will not angle to the right | Ingenuity_TF64-128 - RfVD WIP 413 | WIP14 | For Mouse rotation from zero degrees(clockwise will not work) - So use keyboard For Keyboard typing rotation from zero degrees (anti clockwise will not work) - Use mouse for rotation. | System |  | None | Fixed | patra_r | 0 | 2 | Console Software | Console4 | YES | Closed | WIP14SegReq | Ingenuity 64 |  | WIP14 | February 27, 2019 6:18:16 PM GMT+08:00
CTUDT00598728 | Refresh | Closed | rodan_d | June 17, 2018 3:38:37 PM GMT+08:00 | IV: Solo Cycle 2 - Brain result will not angle to the right | Ingenuity_TF64-128 - RfVD WIP 413 | WIP14 | For Mouse rotation from zero degrees(clockwise will not work) - So use keyboard For Keyboard typing rotation from zero degrees (anti clockwise will not work) - Use mouse for rotation. | System |  | None | Fixed | patra_r | 0 | 2 | Console Software | Console4 | YES | Closed | WIP14SegReq | IQon1 |  | WIP14 | February 27, 2019 6:18:16 PM GMT+08:00
CTUDT00598727 | Refresh | Closed | rodan_d | June 17, 2018 3:31:17 PM GMT+08:00 | Message #0077 referred in requirement Console-SRS.6213, 6214, 6215, 6216 is incorrect | Ingenuity_TF64-128 - RfVD WIP 414 |  |  | Subsystem |  | Console-SRS.6213, 6214, 6215, 6216 | Does Not Apply | bandil_j | 0 | 2 | Requirements | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00598726 | Refresh | Closed | rodan_d | June 17, 2018 3:29:30 PM GMT+08:00 | System error starting tracker scan | Ingenuity_TF64-128 - RfVD WIP 114 |  | The work around is provided by the system by displaying an error message instructing the user to end exam and retry. | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 3 |  | Console4 | YES | Closed |  | PET - Ingenuity_TF64/128
CTUDT00598725 | Refresh | Closed | rodan_d | June 17, 2018 3:15:32 PM GMT+08:00 | Work flow crashed asking user to collect BugReport | Ingenuity_TF64-128 - RfVD WIP 413 | WIP16 |  | Subsystem |  | None | Fixed | patra_r | 0 | 2 | Console Software | Console4 | NO | Closed | WIP16SubCom | Ovation |  | WIP16 | April 25, 2019 8:23:04 PM GMT+08:00
CTUDT00598725 | Refresh | Closed | rodan_d | June 17, 2018 3:15:32 PM GMT+08:00 | Work flow crashed asking user to collect BugReport | Ingenuity_TF64-128 - RfVD WIP 413 | WIP16 |  | Subsystem |  | None | Fixed | patra_r | 0 | 2 | Console Software | Console4 | NO | Closed | WIP16SubCom | Brilliance 64 |  | WIP16 | April 25, 2019 8:23:04 PM GMT+08:00
CTUDT00598725 | Refresh | Closed | rodan_d | June 17, 2018 3:15:32 PM GMT+08:00 | Work flow crashed asking user to collect BugReport | Ingenuity_TF64-128 - RfVD WIP 413 | WIP16 |  | Subsystem |  | None | Fixed | patra_r | 0 | 2 | Console Software | Console4 | NO | Closed | WIP16SubCom | Brilliance Big Bore |  | WIP16 | April 25, 2019 8:23:04 PM GMT+08:00
CTUDT00598725 | Refresh | Closed | rodan_d | June 17, 2018 3:15:32 PM GMT+08:00 | Work flow crashed asking user to collect BugReport | Ingenuity_TF64-128 - RfVD WIP 413 | WIP16 |  | Subsystem |  | None | Fixed | patra_r | 0 | 2 | Console Software | Console4 | NO | Closed | WIP16SubCom | Ingenuity 64 |  | WIP16 | April 25, 2019 8:23:04 PM GMT+08:00
CTUDT00598725 | Refresh | Closed | rodan_d | June 17, 2018 3:15:32 PM GMT+08:00 | Work flow crashed asking user to collect BugReport | Ingenuity_TF64-128 - RfVD WIP 413 | WIP16 |  | Subsystem |  | None | Fixed | patra_r | 0 | 2 | Console Software | Console4 | NO | Closed | WIP16SubCom | IQon1 |  | WIP16 | April 25, 2019 8:23:04 PM GMT+08:00
CTUDT00598706 | Refresh | Closed | juster_d | June 14, 2018 3:21:57 PM GMT+08:00 | CHost FPGA firmware revision B is not burned on the CHost | 4.7.5 | WIP12 |  | Subsystem | HW Reliability | N/A | Fixed | chen_ho | 0 | 3 | Service Tools_Software | GCX | YES | Closed | WIP12SegReq | IQon1 |  | WIP12 | January 18, 2019 9:58:29 AM GMT+08:00
CTUDT00598694 | Refresh | Closed | cohen_y | June 14, 2018 12:28:12 AM GMT+08:00 | Phantom Calibration error due to Phantom Air Table being cut in CIRS due to table name size | WIP7 | WIP12 |  | Subsystem |  | N/A | Fixed | elviral | 0 | 2 | DMS Calibrations | DMS | NO | Closed | WIP12SubCom | Ovation |  | WIP12 | November 6, 2018 11:34:15 PM GMT+08:00
CTUDT00598683 | Refresh | Closed | swack_g | June 12, 2018 2:41:26 PM GMT+08:00 | Refresh: Calibrations: Air Calibration: Air Calibration freezes at 90% | WIP 7 |  |  | Subsystem | Quality Assurance - Air Calibrattion | N/A | Works As Specified / Not A Defect | gehtin_e | 0 |  |  | DMS | NO | Closed |  | Ovation
CTUDT00598674 | Refresh | Closed | attali_s | June 10, 2018 9:39:21 PM GMT+08:00 | Refresh - Error message with CT User in Air Calibration | Wip 7 |  |  | Subsystem | Quality Assurance - Air Calibrattion | N/A | Works As Specified / Not A Defect | gehtin_e | 0 |  |  | DMS | NO | Closed |  | IQon1
CTUDT00598664 | Refresh | Closed | rodan_d | June 6, 2018 10:19:22 PM GMT+08:00 | Prep - Physics Analysis: Prep fails and crashes when using it on some scans (above 4GB Size) | 4.2.0.xx0xx-WIP109 | WIP12 |  | Subsystem |  | n/a | Fixed | frider_o | 0 | 4 | System | DMS | NO | Closed | WIP12SegReq | Ovation |  | WIP12 | December 10, 2018 11:16:57 PM GMT+08:00
CTUDT00598664 | Refresh | Closed | rodan_d | June 6, 2018 10:19:22 PM GMT+08:00 | Prep - Physics Analysis: Prep fails and crashes when using it on some scans (above 4GB Size) | 4.2.0.xx0xx-WIP109 | WIP12 |  | Subsystem |  | n/a | Fixed | frider_o | 0 | 4 | System | DMS | NO | Closed | WIP12SegReq | Brilliance Big Bore |  | WIP12 | December 10, 2018 11:16:57 PM GMT+08:00
CTUDT00598664 | Refresh | Closed | rodan_d | June 6, 2018 10:19:22 PM GMT+08:00 | Prep - Physics Analysis: Prep fails and crashes when using it on some scans (above 4GB Size) | 4.2.0.xx0xx-WIP109 | WIP12 |  | Subsystem |  | n/a | Fixed | frider_o | 0 | 4 | System | DMS | NO | Closed | WIP12SegReq | Ingenuity 64 |  | WIP12 | December 10, 2018 11:16:57 PM GMT+08:00
CTUDT00598664 | Refresh | Closed | rodan_d | June 6, 2018 10:19:22 PM GMT+08:00 | Prep - Physics Analysis: Prep fails and crashes when using it on some scans (above 4GB Size) | 4.2.0.xx0xx-WIP109 | WIP12 |  | Subsystem |  | n/a | Fixed | frider_o | 0 | 4 | System | DMS | NO | Closed | WIP12SegReq | IQon1 |  | WIP12 | December 10, 2018 11:16:57 PM GMT+08:00
CTUDT00598663 | Refresh | Closed | rodan_d | June 6, 2018 10:17:04 PM GMT+08:00 | FOK 10-20s delay after End exam click to input new patient info. after implemention of fco72800675(v4.1.7) | 4.1.7 | WIP16 | No workaround. User has to wait till patient page comes up. | Subsystem | History Review | NA | Fixed | ravindra_a | 0 | 4 | Console Software | Console4 | NO | Closed | WIP17SubCom | Ovation | PR7608269 | WIP16 | April 29, 2019 5:28:30 PM GMT+08:00
CTUDT00598663 | Refresh | Closed | rodan_d | June 6, 2018 10:17:04 PM GMT+08:00 | FOK 10-20s delay after End exam click to input new patient info. after implemention of fco72800675(v4.1.7) | 4.1.7 | WIP16 | No workaround. User has to wait till patient page comes up. | Subsystem | History Review | NA | Fixed | ravindra_a | 0 | 4 | Console Software | Console4 | NO | Closed | WIP17SubCom | Brilliance 64 | PR7608269 | WIP16 | April 29, 2019 5:28:30 PM GMT+08:00
CTUDT00598663 | Refresh | Closed | rodan_d | June 6, 2018 10:17:04 PM GMT+08:00 | FOK 10-20s delay after End exam click to input new patient info. after implemention of fco72800675(v4.1.7) | 4.1.7 | WIP16 | No workaround. User has to wait till patient page comes up. | Subsystem | History Review | NA | Fixed | ravindra_a | 0 | 4 | Console Software | Console4 | NO | Closed | WIP17SubCom | Brilliance Big Bore | PR7608269 | WIP16 | April 29, 2019 5:28:30 PM GMT+08:00
CTUDT00598663 | Refresh | Closed | rodan_d | June 6, 2018 10:17:04 PM GMT+08:00 | FOK 10-20s delay after End exam click to input new patient info. after implemention of fco72800675(v4.1.7) | 4.1.7 | WIP16 | No workaround. User has to wait till patient page comes up. | Subsystem | History Review | NA | Fixed | ravindra_a | 0 | 4 | Console Software | Console4 | NO | Closed | WIP17SubCom | Ingenuity 64 | PR7608269 | WIP16 | April 29, 2019 5:28:30 PM GMT+08:00
CTUDT00598663 | Refresh | Closed | rodan_d | June 6, 2018 10:17:04 PM GMT+08:00 | FOK 10-20s delay after End exam click to input new patient info. after implemention of fco72800675(v4.1.7) | 4.1.7 | WIP16 | No workaround. User has to wait till patient page comes up. | Subsystem | History Review | NA | Fixed | ravindra_a | 0 | 4 | Console Software | Console4 | NO | Closed | WIP17SubCom | IQon1 | PR7608269 | WIP16 | April 29, 2019 5:28:30 PM GMT+08:00
CTUDT00598662 | Refresh | Closed | rodan_d | June 6, 2018 10:15:20 PM GMT+08:00 | No-Load test not working after 4.1.7 software load | v4.1.7.22026 | WIP9.1 |  | GCX |  | NA | Fixed | karkour | 0 | 2 | Gantry Software | GCX | NO | Closed | WIP9.1SubCom | Brilliance 64 | PR7669865 | WIP9.1 | September 19, 2018 8:59:28 PM GMT+08:00
CTUDT00598662 | Refresh | Closed | rodan_d | June 6, 2018 10:15:20 PM GMT+08:00 | No-Load test not working after 4.1.7 software load | v4.1.7.22026 | WIP9.1 |  | GCX |  | NA | Fixed | karkour | 0 | 2 | Gantry Software | GCX | NO | Closed | WIP9.1SubCom | Ingenuity 64 | PR7669865 | WIP9.1 | September 19, 2018 8:59:28 PM GMT+08:00
CTUDT00598661 | Refresh | Closed | rodan_d | June 6, 2018 10:10:48 PM GMT+08:00 | 4.1.6 BR64 – Trim edges option doesn't work | 4.1.6WIP104 |  | The user has a visual indication that the plan box has changed size and has the ability to adjust the size of the result by dragging the plan box. the IFU in the "plan on Surview" states that the user can adjust the size and position of the plan box: "Click and drag: – If the cursor is in the middle of the plan box, the box is dragged with the cursor. – If the cursor is over one of the box boundaries the boundary will drag with the cursor. " | Subsystem |  | Console.SRS-2384 | Works As Specified / Not A Defect | prasad_r | 0 | 4 |  | Console4 | NO | Closed | WIP16SubCom | Brilliance 64
CTUDT00598659 | Refresh | Closed | rodan_d | June 6, 2018 10:09:30 PM GMT+08:00 | 4.1.6 FOK Dose report did not generate | V4.1.6.25030 | WIP15 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP15SegReq | Ovation | PR 6199746 (CFM 9421304) | WIP15 | April 11, 2019 12:16:18 PM GMT+08:00
CTUDT00598659 | Refresh | Closed | rodan_d | June 6, 2018 10:09:30 PM GMT+08:00 | 4.1.6 FOK Dose report did not generate | V4.1.6.25030 | WIP15 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP15SegReq | Brilliance 64 | PR 6199746 (CFM 9421304) | WIP15 | April 11, 2019 12:16:18 PM GMT+08:00
CTUDT00598659 | Refresh | Closed | rodan_d | June 6, 2018 10:09:30 PM GMT+08:00 | 4.1.6 FOK Dose report did not generate | V4.1.6.25030 | WIP15 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP15SegReq | Brilliance Big Bore | PR 6199746 (CFM 9421304) | WIP15 | April 11, 2019 12:16:18 PM GMT+08:00
CTUDT00598659 | Refresh | Closed | rodan_d | June 6, 2018 10:09:30 PM GMT+08:00 | 4.1.6 FOK Dose report did not generate | V4.1.6.25030 | WIP15 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP15SegReq | Ingenuity 64 | PR 6199746 (CFM 9421304) | WIP15 | April 11, 2019 12:16:18 PM GMT+08:00
CTUDT00598658 | Refresh | Closed | rodan_d | June 6, 2018 10:08:30 PM GMT+08:00 | 4.1.6 release notes has incorrect verbiage. See description | 4.1.6 |  |  | Subsystem |  | N/A | Does Not Apply | rajalakshmi_j | 0 | 3 |  | Support-Service | NO | Closed | WIP14SegReq | Ovation
CTUDT00598658 | Refresh | Closed | rodan_d | June 6, 2018 10:08:30 PM GMT+08:00 | 4.1.6 release notes has incorrect verbiage. See description | 4.1.6 |  |  | Subsystem |  | N/A | Does Not Apply | rajalakshmi_j | 0 | 3 |  | Support-Service | NO | Closed | WIP14SegReq | Brilliance 64
CTUDT00598658 | Refresh | Closed | rodan_d | June 6, 2018 10:08:30 PM GMT+08:00 | 4.1.6 release notes has incorrect verbiage. See description | 4.1.6 |  |  | Subsystem |  | N/A | Does Not Apply | rajalakshmi_j | 0 | 3 |  | Support-Service | NO | Closed | WIP14SegReq | Ingenuity 64
CTUDT00598657 | Refresh | Closed | rodan_d | June 6, 2018 10:06:33 PM GMT+08:00 | Autosend failure due to ExamManager exit before completion of saving of results. | V 4.1.6 | WIP14 | Workaround is to use the CopyTo function to send the results to the external device at a later time | Subsystem |  | Console-SRS.9098, Console-SRS.9101 | Fixed | sagar_v | 0 | 3 | Console Software | Console4 | NO | Closed | WIP14SegReq | Ovation | PR 5771594 PR 5786361 PR 5824426 PR 5824428 PR 5824429 PR 5991017 PR 5991018 PR 5785859 | WIP14 | February 6, 2019 8:47:11 PM GMT+08:00
CTUDT00598657 | Refresh | Closed | rodan_d | June 6, 2018 10:06:33 PM GMT+08:00 | Autosend failure due to ExamManager exit before completion of saving of results. | V 4.1.6 | WIP14 | Workaround is to use the CopyTo function to send the results to the external device at a later time | Subsystem |  | Console-SRS.9098, Console-SRS.9101 | Fixed | sagar_v | 0 | 3 | Console Software | Console4 | NO | Closed | WIP14SegReq | Brilliance 64 | PR 5771594 PR 5786361 PR 5824426 PR 5824428 PR 5824429 PR 5991017 PR 5991018 PR 5785859 | WIP14 | February 6, 2019 8:47:11 PM GMT+08:00
CTUDT00598657 | Refresh | Closed | rodan_d | June 6, 2018 10:06:33 PM GMT+08:00 | Autosend failure due to ExamManager exit before completion of saving of results. | V 4.1.6 | WIP14 | Workaround is to use the CopyTo function to send the results to the external device at a later time | Subsystem |  | Console-SRS.9098, Console-SRS.9101 | Fixed | sagar_v | 0 | 3 | Console Software | Console4 | NO | Closed | WIP14SegReq | Brilliance Big Bore | PR 5771594 PR 5786361 PR 5824426 PR 5824428 PR 5824429 PR 5991017 PR 5991018 PR 5785859 | WIP14 | February 6, 2019 8:47:11 PM GMT+08:00
CTUDT00598657 | Refresh | Closed | rodan_d | June 6, 2018 10:06:33 PM GMT+08:00 | Autosend failure due to ExamManager exit before completion of saving of results. | V 4.1.6 | WIP14 | Workaround is to use the CopyTo function to send the results to the external device at a later time | Subsystem |  | Console-SRS.9098, Console-SRS.9101 | Fixed | sagar_v | 0 | 3 | Console Software | Console4 | NO | Closed | WIP14SegReq | Ingenuity 64 | PR 5771594 PR 5786361 PR 5824426 PR 5824428 PR 5824429 PR 5991017 PR 5991018 PR 5785859 | WIP14 | February 6, 2019 8:47:11 PM GMT+08:00
CTUDT00598657 | Refresh | Closed | rodan_d | June 6, 2018 10:06:33 PM GMT+08:00 | Autosend failure due to ExamManager exit before completion of saving of results. | V 4.1.6 | WIP14 | Workaround is to use the CopyTo function to send the results to the external device at a later time | Subsystem |  | Console-SRS.9098, Console-SRS.9101 | Fixed | sagar_v | 0 | 3 | Console Software | Console4 | NO | Closed | WIP14SegReq | IQon1 | PR 5771594 PR 5786361 PR 5824426 PR 5824428 PR 5824429 PR 5991017 PR 5991018 PR 5785859 | WIP14 | February 6, 2019 8:47:11 PM GMT+08:00
CTUDT00598656 | Refresh | Closed | rodan_d | June 6, 2018 10:03:55 PM GMT+08:00 | BB42 FOK: 4DCT Phases not reconstructed when bellows signal goes low in the middle of the scan | V4.1.6 | WIP16 | User can recover from this by performing offline reconstruction after editing the waveform. | System |  | NA | Fixed | wilsonraja | 0 | 3 | Console Software | Console4 | YES | Closed | BB_WIP6_SubCom | Brilliance 64 | CFR9418816 | WIP16 | April 5, 2019 7:44:12 PM GMT+08:00
CTUDT00598656 | Refresh | Closed | rodan_d | June 6, 2018 10:03:55 PM GMT+08:00 | BB42 FOK: 4DCT Phases not reconstructed when bellows signal goes low in the middle of the scan | V4.1.6 | WIP16 | User can recover from this by performing offline reconstruction after editing the waveform. | System |  | NA | Fixed | wilsonraja | 0 | 3 | Console Software | Console4 | YES | Closed | BB_WIP6_SubCom | Ingenuity 64 | CFR9418816 | WIP16 | April 5, 2019 7:44:12 PM GMT+08:00
CTUDT00598655 | Refresh | Closed | rodan_d | June 6, 2018 9:59:41 PM GMT+08:00 | BB42 FOK: Resp 4DCT Linear binning produces images at the same phase point. | v4.1.6 | WIP18 | User can recover from the error by doing offline reconstruction. | Subsystem |  | NA | Fixed | pandey_a | 0 | 3 | Console Software | Console4 | YES | Closed | BB_WIP6_SubCom | Brilliance 64 | CFR9418198 | WIP18 | June 4, 2019 2:45:05 PM GMT+08:00
CTUDT00598655 | Refresh | Closed | rodan_d | June 6, 2018 9:59:41 PM GMT+08:00 | BB42 FOK: Resp 4DCT Linear binning produces images at the same phase point. | v4.1.6 | WIP18 | User can recover from the error by doing offline reconstruction. | Subsystem |  | NA | Fixed | pandey_a | 0 | 3 | Console Software | Console4 | YES | Closed | BB_WIP6_SubCom | Ingenuity 64 | CFR9418198 | WIP18 | June 4, 2019 2:45:05 PM GMT+08:00
CTUDT00598654 | Refresh | Closed | rodan_d | June 6, 2018 9:51:06 PM GMT+08:00 | Calibrations: Hcor Slice Calibration: Coefficient factors are not being updated when running it a second time without host restart. | WIP 106 | WIP 504.2 | Restart Console or Host between different HCOR runs. | Subsystem |  | n/a | Fixed | elviral | 0 | 3 | DMS Calibrations | DMS | YES | Closed | WIP14SegReq | Ovation | PR #9007291 | WIP 504.2 | February 14, 2019 12:53:27 AM GMT+08:00
CTUDT00598654 | Refresh | Closed | rodan_d | June 6, 2018 9:51:06 PM GMT+08:00 | Calibrations: Hcor Slice Calibration: Coefficient factors are not being updated when running it a second time without host restart. | WIP 106 | WIP 504.2 | Restart Console or Host between different HCOR runs. | Subsystem |  | n/a | Fixed | elviral | 0 | 3 | DMS Calibrations | DMS | YES | Closed | WIP14SegReq | Brilliance 64 | PR #9007291 | WIP 504.2 | February 14, 2019 12:53:27 AM GMT+08:00
CTUDT00598654 | Refresh | Closed | rodan_d | June 6, 2018 9:51:06 PM GMT+08:00 | Calibrations: Hcor Slice Calibration: Coefficient factors are not being updated when running it a second time without host restart. | WIP 106 | WIP 504.2 | Restart Console or Host between different HCOR runs. | Subsystem |  | n/a | Fixed | elviral | 0 | 3 | DMS Calibrations | DMS | YES | Closed | WIP14SegReq | Brilliance Big Bore | PR #9007291 | WIP 504.2 | February 14, 2019 12:53:27 AM GMT+08:00
CTUDT00598654 | Refresh | Closed | rodan_d | June 6, 2018 9:51:06 PM GMT+08:00 | Calibrations: Hcor Slice Calibration: Coefficient factors are not being updated when running it a second time without host restart. | WIP 106 | WIP 504.2 | Restart Console or Host between different HCOR runs. | Subsystem |  | n/a | Fixed | elviral | 0 | 3 | DMS Calibrations | DMS | YES | Closed | WIP14SegReq | Ingenuity 64 | PR #9007291 | WIP 504.2 | February 14, 2019 12:53:27 AM GMT+08:00
CTUDT00598654 | Refresh | Closed | rodan_d | June 6, 2018 9:51:06 PM GMT+08:00 | Calibrations: Hcor Slice Calibration: Coefficient factors are not being updated when running it a second time without host restart. | WIP 106 | WIP 504.2 | Restart Console or Host between different HCOR runs. | Subsystem |  | n/a | Fixed | elviral | 0 | 3 | DMS Calibrations | DMS | YES | Closed | WIP14SegReq | IQon1 | PR #9007291 | WIP 504.2 | February 14, 2019 12:53:27 AM GMT+08:00
CTUDT00598653 | Refresh | Closed | rodan_d | June 6, 2018 9:48:32 PM GMT+08:00 | CCT: CCT Single display is blank and grayed out | 4.1.0.XX241-WIP211 |  |  | Subsystem |  | Console-SRS.10180 | No Plan to Fix | sanjeeva_a | 0 | 3 | Console Generic | Console4 | NO | Closed |  | Ovation
CTUDT00598653 | Refresh | Closed | rodan_d | June 6, 2018 9:48:32 PM GMT+08:00 | CCT: CCT Single display is blank and grayed out | 4.1.0.XX241-WIP211 |  |  | Subsystem |  | Console-SRS.10180 | No Plan to Fix | sanjeeva_a | 0 | 3 | Console Generic | Console4 | NO | Closed |  | Brilliance 64
CTUDT00598653 | Refresh | Closed | rodan_d | June 6, 2018 9:48:32 PM GMT+08:00 | CCT: CCT Single display is blank and grayed out | 4.1.0.XX241-WIP211 |  |  | Subsystem |  | Console-SRS.10180 | No Plan to Fix | sanjeeva_a | 0 | 3 | Console Generic | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00598652 | Refresh | Closed | rodan_d | June 6, 2018 9:47:24 PM GMT+08:00 | Crescent artifact in pediatric brain with Ingenuity v4.1.4 FOK | v4.1.4 FOK | WIP17 | Use of High Resolution for brain imaging | System | Verified1 | CT-NM.SRS.IQ.1408 | Fixed | richards_b | 0 | 2 | CIRS_SW_Installation | Support-Physics | YES | Closed | WIP17SubCom | Brilliance 64 | CFR 9367697 PR 4918461 | WIP17 | June 17, 2019 8:54:44 PM GMT+08:00
CTUDT00598652 | Refresh | Closed | rodan_d | June 6, 2018 9:47:24 PM GMT+08:00 | Crescent artifact in pediatric brain with Ingenuity v4.1.4 FOK | v4.1.4 FOK | WIP17 | Use of High Resolution for brain imaging | System | Verified1 | CT-NM.SRS.IQ.1408 | Fixed | richards_b | 0 | 2 | CIRS_SW_Installation | Support-Physics | YES | Closed | WIP17SubCom | Ingenuity 64 | CFR 9367697 PR 4918461 | WIP17 | June 17, 2019 8:54:44 PM GMT+08:00
CTUDT00598651 | Refresh | Closed | rodan_d | June 6, 2018 9:46:05 PM GMT+08:00 | CT-NM.SRS.Interop.182 requirement (The system shall control the operation of hard ejection during CCT and Bolus tracking scans) wording is confusing, | v4.1.6 |  | None Required. Not a hazard. | Subsystem |  |  | Change Proposal | oren_d | 0 | 3 |  | Support-System | NO | Closed |  | Ovation
CTUDT00598651 | Refresh | Closed | rodan_d | June 6, 2018 9:46:05 PM GMT+08:00 | CT-NM.SRS.Interop.182 requirement (The system shall control the operation of hard ejection during CCT and Bolus tracking scans) wording is confusing, | v4.1.6 |  | None Required. Not a hazard. | Subsystem |  |  | Change Proposal | oren_d | 0 | 3 |  | Support-System | NO | Closed |  | Brilliance 64
CTUDT00598651 | Refresh | Closed | rodan_d | June 6, 2018 9:46:05 PM GMT+08:00 | CT-NM.SRS.Interop.182 requirement (The system shall control the operation of hard ejection during CCT and Bolus tracking scans) wording is confusing, | v4.1.6 |  | None Required. Not a hazard. | Subsystem |  |  | Change Proposal | oren_d | 0 | 3 |  | Support-System | NO | Closed |  | Brilliance Big Bore
CTUDT00598651 | Refresh | Closed | rodan_d | June 6, 2018 9:46:05 PM GMT+08:00 | CT-NM.SRS.Interop.182 requirement (The system shall control the operation of hard ejection during CCT and Bolus tracking scans) wording is confusing, | v4.1.6 |  | None Required. Not a hazard. | Subsystem |  |  | Change Proposal | oren_d | 0 | 3 |  | Support-System | NO | Closed |  | Ingenuity 64
CTUDT00598651 | Refresh | Closed | rodan_d | June 6, 2018 9:46:05 PM GMT+08:00 | CT-NM.SRS.Interop.182 requirement (The system shall control the operation of hard ejection during CCT and Bolus tracking scans) wording is confusing, | v4.1.6 |  | None Required. Not a hazard. | Subsystem |  |  | Change Proposal | oren_d | 0 | 3 |  | Support-System | NO | Closed |  | IQon1
CTUDT00598650 | Refresh | Closed | rodan_d | June 6, 2018 9:43:32 PM GMT+08:00 | Deviations in parameter values during Scan Type Verification for V4.1.6 ING/BR64 WIP5 | 4.1.6.25017 |  |  | Subsystem |  | Console.SRS-14481 | No Plan to Fix | sanjeeva_a | 0 | 4 |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00598650 | Refresh | Closed | rodan_d | June 6, 2018 9:43:32 PM GMT+08:00 | Deviations in parameter values during Scan Type Verification for V4.1.6 ING/BR64 WIP5 | 4.1.6.25017 |  |  | Subsystem |  | Console.SRS-14481 | No Plan to Fix | sanjeeva_a | 0 | 4 |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00598649 | Refresh | Closed | rodan_d | June 6, 2018 9:42:20 PM GMT+08:00 | Direction column in the Overview tab of Show All still appear after uncheck it | 4.1.2.XX026-WIP307 |  |  | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 | Console Generic | Console4 | NO | Closed |  | Brilliance 64
CTUDT00598646 | Refresh | Closed | rodan_d | June 6, 2018 9:39:15 PM GMT+08:00 | DOM check section in the IQ and dose testing need to be updated | 4.1.3.XX022 |  |  | Subsystem |  | N/A | No Plan to Fix | yao_j | 0 | 3 |  | Support-Service | NO | Closed |  | Brilliance 64 | 9385963 9387699 9387698 9387697
CTUDT00598646 | Refresh | Closed | rodan_d | June 6, 2018 9:39:15 PM GMT+08:00 | DOM check section in the IQ and dose testing need to be updated | 4.1.3.XX022 |  |  | Subsystem |  | N/A | No Plan to Fix | yao_j | 0 | 3 |  | Support-Service | NO | Closed |  | Ingenuity 64 | 9385963 9387699 9387698 9387697
CTUDT00598645 | Refresh | Closed | rodan_d | June 6, 2018 9:38:00 PM GMT+08:00 | Dose Alert does not trigger for repeat scans | 4.1.4.XX001 |  |  | Subsystem |  | Console-SRS-15995 | Works As Specified / Not A Defect | sagar_v | 0 | 2 |  | Console4 | YES | Closed | WIP17SubCom | Ovation | CFR#9414426 CFR#9414427 CFR#9414428 PR6042337
CTUDT00598645 | Refresh | Closed | rodan_d | June 6, 2018 9:38:00 PM GMT+08:00 | Dose Alert does not trigger for repeat scans | 4.1.4.XX001 |  |  | Subsystem |  | Console-SRS-15995 | Works As Specified / Not A Defect | sagar_v | 0 | 2 |  | Console4 | YES | Closed | WIP17SubCom | Brilliance 64 | CFR#9414426 CFR#9414427 CFR#9414428 PR6042337
CTUDT00598645 | Refresh | Closed | rodan_d | June 6, 2018 9:38:00 PM GMT+08:00 | Dose Alert does not trigger for repeat scans | 4.1.4.XX001 |  |  | Subsystem |  | Console-SRS-15995 | Works As Specified / Not A Defect | sagar_v | 0 | 2 |  | Console4 | YES | Closed | WIP17SubCom | Brilliance Big Bore | CFR#9414426 CFR#9414427 CFR#9414428 PR6042337
CTUDT00598645 | Refresh | Closed | rodan_d | June 6, 2018 9:38:00 PM GMT+08:00 | Dose Alert does not trigger for repeat scans | 4.1.4.XX001 |  |  | Subsystem |  | Console-SRS-15995 | Works As Specified / Not A Defect | sagar_v | 0 | 2 |  | Console4 | YES | Closed | WIP17SubCom | Ingenuity 64 | CFR#9414426 CFR#9414427 CFR#9414428 PR6042337
CTUDT00598645 | Refresh | Closed | rodan_d | June 6, 2018 9:38:00 PM GMT+08:00 | Dose Alert does not trigger for repeat scans | 4.1.4.XX001 |  |  | Subsystem |  | Console-SRS-15995 | Works As Specified / Not A Defect | sagar_v | 0 | 2 |  | Console4 | YES | Closed | WIP17SubCom | IQon1 | CFR#9414426 CFR#9414427 CFR#9414428 PR6042337
CTUDT00598644 | Refresh | Closed | rodan_d | June 6, 2018 9:36:21 PM GMT+08:00 | DSR Report has Exposure Time in ms but unit shown as s | _Current Production HW_ | WIP14 |  | System |  | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SubCom | Ovation | Mercury ID: 9446164. | WIP14 | July 19, 2019 7:32:34 PM GMT+08:00
CTUDT00598644 | Refresh | Closed | rodan_d | June 6, 2018 9:36:21 PM GMT+08:00 | DSR Report has Exposure Time in ms but unit shown as s | _Current Production HW_ | WIP14 |  | System |  | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SubCom | Brilliance 64 | Mercury ID: 9446164. | WIP14 | July 19, 2019 7:32:34 PM GMT+08:00
CTUDT00598644 | Refresh | Closed | rodan_d | June 6, 2018 9:36:21 PM GMT+08:00 | DSR Report has Exposure Time in ms but unit shown as s | _Current Production HW_ | WIP14 |  | System |  | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SubCom | Ingenuity 64 | Mercury ID: 9446164. | WIP14 | July 19, 2019 7:32:34 PM GMT+08:00
CTUDT00598644 | Refresh | Closed | rodan_d | June 6, 2018 9:36:21 PM GMT+08:00 | DSR Report has Exposure Time in ms but unit shown as s | _Current Production HW_ | WIP14 |  | System |  | NA | Fixed | kr_m | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SubCom | IQon1 | Mercury ID: 9446164. | WIP14 | July 19, 2019 7:32:34 PM GMT+08:00
CTUDT00598643 | Refresh | Closed | rodan_d | June 6, 2018 9:29:45 PM GMT+08:00 | ECG file can't be associated to a corresponding Cardiac study after that study is anonymized. | 4.1.3.XX022 | WIP12 | The original patient data may be used on the customer site. | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 3 | USP | iCAP | NO | Closed | WIP12SegReq | Ovation |  | WIP12 | December 14, 2018 5:14:53 PM GMT+08:00
CTUDT00598643 | Refresh | Closed | rodan_d | June 6, 2018 9:29:45 PM GMT+08:00 | ECG file can't be associated to a corresponding Cardiac study after that study is anonymized. | 4.1.3.XX022 | WIP12 | The original patient data may be used on the customer site. | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 3 | USP | iCAP | NO | Closed | WIP12SegReq | Brilliance 64 |  | WIP12 | December 14, 2018 5:14:53 PM GMT+08:00
CTUDT00598643 | Refresh | Closed | rodan_d | June 6, 2018 9:29:45 PM GMT+08:00 | ECG file can't be associated to a corresponding Cardiac study after that study is anonymized. | 4.1.3.XX022 | WIP12 | The original patient data may be used on the customer site. | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 3 | USP | iCAP | NO | Closed | WIP12SegReq | Brilliance Big Bore |  | WIP12 | December 14, 2018 5:14:53 PM GMT+08:00
CTUDT00598643 | Refresh | Closed | rodan_d | June 6, 2018 9:29:45 PM GMT+08:00 | ECG file can't be associated to a corresponding Cardiac study after that study is anonymized. | 4.1.3.XX022 | WIP12 | The original patient data may be used on the customer site. | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 3 | USP | iCAP | NO | Closed | WIP12SegReq | Ingenuity 64 |  | WIP12 | December 14, 2018 5:14:53 PM GMT+08:00
CTUDT00598643 | Refresh | Closed | rodan_d | June 6, 2018 9:29:45 PM GMT+08:00 | ECG file can't be associated to a corresponding Cardiac study after that study is anonymized. | 4.1.3.XX022 | WIP12 | The original patient data may be used on the customer site. | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 3 | USP | iCAP | NO | Closed | WIP12SegReq | IQon1 |  | WIP12 | December 14, 2018 5:14:53 PM GMT+08:00
CTUDT00598642 | Refresh | Closed | rodan_d | June 6, 2018 9:26:07 PM GMT+08:00 | Exam Card collimation parameter was changed after V416 to V417 upgrade | V4.1.7xx014 WIP-4 | WIP16 |  | Subsystem |  | N/A | Fixed | navale_s | 0 | 3 | Console Software | Service Tools | NO | Closed | WIPSubCom16 | Ovation |  | WIP16 | April 19, 2019 10:49:20 PM GMT+08:00
CTUDT00598642 | Refresh | Closed | rodan_d | June 6, 2018 9:26:07 PM GMT+08:00 | Exam Card collimation parameter was changed after V416 to V417 upgrade | V4.1.7xx014 WIP-4 | WIP16 |  | Subsystem |  | N/A | Fixed | navale_s | 0 | 3 | Console Software | Service Tools | NO | Closed | WIPSubCom16 | Brilliance 64 |  | WIP16 | April 19, 2019 10:49:20 PM GMT+08:00
CTUDT00598642 | Refresh | Closed | rodan_d | June 6, 2018 9:26:07 PM GMT+08:00 | Exam Card collimation parameter was changed after V416 to V417 upgrade | V4.1.7xx014 WIP-4 | WIP16 |  | Subsystem |  | N/A | Fixed | navale_s | 0 | 3 | Console Software | Service Tools | NO | Closed | WIPSubCom16 | Ingenuity 64 |  | WIP16 | April 19, 2019 10:49:20 PM GMT+08:00
CTUDT00598640 | Refresh | Closed | rodan_d | June 6, 2018 9:24:41 PM GMT+08:00 | Exam Card Manager \ Import: Console checking the free space in the source device and not the local D:\. if there is no free space in the source device ‘not enough free space exists’ message pops–up | 4.1.0.XX229-WIP210 |  | Use source device with enough memory which may involve deleting files on the source drive or create the exam cards manually. | Subsystem |  | None | No Plan to Fix | sanjeeva_a | 0 | 3 | Console Generic | Console4 | NO | Closed |  | Ovation
CTUDT00598640 | Refresh | Closed | rodan_d | June 6, 2018 9:24:41 PM GMT+08:00 | Exam Card Manager \ Import: Console checking the free space in the source device and not the local D:\. if there is no free space in the source device ‘not enough free space exists’ message pops–up | 4.1.0.XX229-WIP210 |  | Use source device with enough memory which may involve deleting files on the source drive or create the exam cards manually. | Subsystem |  | None | No Plan to Fix | sanjeeva_a | 0 | 3 | Console Generic | Console4 | NO | Closed |  | Brilliance 64
CTUDT00598640 | Refresh | Closed | rodan_d | June 6, 2018 9:24:41 PM GMT+08:00 | Exam Card Manager \ Import: Console checking the free space in the source device and not the local D:\. if there is no free space in the source device ‘not enough free space exists’ message pops–up | 4.1.0.XX229-WIP210 |  | Use source device with enough memory which may involve deleting files on the source drive or create the exam cards manually. | Subsystem |  | None | No Plan to Fix | sanjeeva_a | 0 | 3 | Console Generic | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00598637 | Refresh | Closed | rodan_d | June 6, 2018 9:14:15 PM GMT+08:00 | Exam cards with foreign characters in the name cause exam application crash after backup and restore | V4.1.7-WIP-5.1 |  | Export the Exam Card with foreign character before taking the backup. Delete the Exam Card before taking backup. To use that Examcard, open ECM and import it directly from the located you exported it in instead of restoring from backup. | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 3 |  | Console4 | NO | Closed |  | Ovation
CTUDT00598637 | Refresh | Closed | rodan_d | June 6, 2018 9:14:15 PM GMT+08:00 | Exam cards with foreign characters in the name cause exam application crash after backup and restore | V4.1.7-WIP-5.1 |  | Export the Exam Card with foreign character before taking the backup. Delete the Exam Card before taking backup. To use that Examcard, open ECM and import it directly from the located you exported it in instead of restoring from backup. | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 3 |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00598637 | Refresh | Closed | rodan_d | June 6, 2018 9:14:15 PM GMT+08:00 | Exam cards with foreign characters in the name cause exam application crash after backup and restore | V4.1.7-WIP-5.1 |  | Export the Exam Card with foreign character before taking the backup. Delete the Exam Card before taking backup. To use that Examcard, open ECM and import it directly from the located you exported it in instead of restoring from backup. | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 3 |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00598637 | Refresh | Closed | rodan_d | June 6, 2018 9:14:15 PM GMT+08:00 | Exam cards with foreign characters in the name cause exam application crash after backup and restore | V4.1.7-WIP-5.1 |  | Export the Exam Card with foreign character before taking the backup. Delete the Exam Card before taking backup. To use that Examcard, open ECM and import it directly from the located you exported it in instead of restoring from backup. | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 3 |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00598636 | Refresh | Closed | rodan_d | June 6, 2018 9:12:10 PM GMT+08:00 | GCX Scan Mode GCX.SSRS.iCT.12520 "RTag Processing" fails in code review | 4.1.3.XX022 |  |  | Subsystem |  | N/A | No Plan to Fix | rooks | 0 | 3 | SRS SSRS | GCX | YES | Closed |  | Brilliance 64
CTUDT00598636 | Refresh | Closed | rodan_d | June 6, 2018 9:12:10 PM GMT+08:00 | GCX Scan Mode GCX.SSRS.iCT.12520 "RTag Processing" fails in code review | 4.1.3.XX022 |  |  | Subsystem |  | N/A | No Plan to Fix | rooks | 0 | 3 | SRS SSRS | GCX | YES | Closed |  | Brilliance Big Bore
CTUDT00598636 | Refresh | Closed | rodan_d | June 6, 2018 9:12:10 PM GMT+08:00 | GCX Scan Mode GCX.SSRS.iCT.12520 "RTag Processing" fails in code review | 4.1.3.XX022 |  |  | Subsystem |  | N/A | No Plan to Fix | rooks | 0 | 3 | SRS SSRS | GCX | YES | Closed |  | Ingenuity 64
CTUDT00598636 | Refresh | Closed | rodan_d | June 6, 2018 9:12:10 PM GMT+08:00 | GCX Scan Mode GCX.SSRS.iCT.12520 "RTag Processing" fails in code review | 4.1.3.XX022 |  |  | Subsystem |  | N/A | No Plan to Fix | rooks | 0 | 3 | SRS SSRS | GCX | YES | Closed |  | IQon1
CTUDT00598635 | Refresh | Closed | rodan_d | June 6, 2018 9:10:26 PM GMT+08:00 | Head Axial EC needs to be removed from reference EC | 4.1.6 | WIP15 |  | Subsystem |  | N/A | Fixed | dunitza_m | 0 | 2 | Console Software | Support-System | YES | Closed | SubCom | Ovation | NA | WIP15 | June 5, 2019 6:49:00 PM GMT+08:00
CTUDT00598635 | Refresh | Closed | rodan_d | June 6, 2018 9:10:26 PM GMT+08:00 | Head Axial EC needs to be removed from reference EC | 4.1.6 | WIP15 |  | Subsystem |  | N/A | Fixed | dunitza_m | 0 | 2 | Console Software | Support-System | YES | Closed | SubCom | Brilliance 64 | NA | WIP15 | June 5, 2019 6:49:00 PM GMT+08:00
CTUDT00598635 | Refresh | Closed | rodan_d | June 6, 2018 9:10:26 PM GMT+08:00 | Head Axial EC needs to be removed from reference EC | 4.1.6 | WIP15 |  | Subsystem |  | N/A | Fixed | dunitza_m | 0 | 2 | Console Software | Support-System | YES | Closed | SubCom | Ingenuity 64 | NA | WIP15 | June 5, 2019 6:49:00 PM GMT+08:00
CTUDT00598634 | Refresh | Closed | rodan_d | June 6, 2018 9:06:33 PM GMT+08:00 | Host computer crash with Blue Screen | V4.1.7.xx009_WIP-3 |  | The user must turn off the Host computer and turn the Host computer back on. | Subsystem |  |  | Does Not Apply | sanjeeva_a | 0 | 3 |  | Console4 | NO | Closed |  | Ovation | PR7646575
CTUDT00598634 | Refresh | Closed | rodan_d | June 6, 2018 9:06:33 PM GMT+08:00 | Host computer crash with Blue Screen | V4.1.7.xx009_WIP-3 |  | The user must turn off the Host computer and turn the Host computer back on. | Subsystem |  |  | Does Not Apply | sanjeeva_a | 0 | 3 |  | Console4 | NO | Closed |  | Brilliance 64 | PR7646575
CTUDT00598634 | Refresh | Closed | rodan_d | June 6, 2018 9:06:33 PM GMT+08:00 | Host computer crash with Blue Screen | V4.1.7.xx009_WIP-3 |  | The user must turn off the Host computer and turn the Host computer back on. | Subsystem |  |  | Does Not Apply | sanjeeva_a | 0 | 3 |  | Console4 | NO | Closed |  | Brilliance Big Bore | PR7646575
CTUDT00598634 | Refresh | Closed | rodan_d | June 6, 2018 9:06:33 PM GMT+08:00 | Host computer crash with Blue Screen | V4.1.7.xx009_WIP-3 |  | The user must turn off the Host computer and turn the Host computer back on. | Subsystem |  |  | Does Not Apply | sanjeeva_a | 0 | 3 |  | Console4 | NO | Closed |  | Ingenuity 64 | PR7646575
CTUDT00598634 | Refresh | Closed | rodan_d | June 6, 2018 9:06:33 PM GMT+08:00 | Host computer crash with Blue Screen | V4.1.7.xx009_WIP-3 |  | The user must turn off the Host computer and turn the Host computer back on. | Subsystem |  |  | Does Not Apply | sanjeeva_a | 0 | 3 |  | Console4 | NO | Closed |  | IQon1 | PR7646575
CTUDT00598633 | Refresh | Closed | rodan_d | June 6, 2018 9:01:29 PM GMT+08:00 | In voice manager some of the phrases were not translated for all the languages in the voice manager UI. | V4.1.7_WIP-3 | WIP12 | If user does not understand the string in English because it is not translated he can disable voice command and provide instructions to user manually. If necessary user could open the voice manager and play the original message and save it as a new custom voice command and store it with a name in the localized language. | System |  | NA | Fixed | kr_m | 0 | 3 | Console Software | Console4 | NO | Closed | WIP11.1SegReq | Ovation |  | WIP12 | December 13, 2018 6:13:41 PM GMT+08:00
CTUDT00598633 | Refresh | Closed | rodan_d | June 6, 2018 9:01:29 PM GMT+08:00 | In voice manager some of the phrases were not translated for all the languages in the voice manager UI. | V4.1.7_WIP-3 | WIP12 | If user does not understand the string in English because it is not translated he can disable voice command and provide instructions to user manually. If necessary user could open the voice manager and play the original message and save it as a new custom voice command and store it with a name in the localized language. | System |  | NA | Fixed | kr_m | 0 | 3 | Console Software | Console4 | NO | Closed | WIP11.1SegReq | Brilliance 64 |  | WIP12 | December 13, 2018 6:13:41 PM GMT+08:00
CTUDT00598633 | Refresh | Closed | rodan_d | June 6, 2018 9:01:29 PM GMT+08:00 | In voice manager some of the phrases were not translated for all the languages in the voice manager UI. | V4.1.7_WIP-3 | WIP12 | If user does not understand the string in English because it is not translated he can disable voice command and provide instructions to user manually. If necessary user could open the voice manager and play the original message and save it as a new custom voice command and store it with a name in the localized language. | System |  | NA | Fixed | kr_m | 0 | 3 | Console Software | Console4 | NO | Closed | WIP11.1SegReq | Ingenuity 64 |  | WIP12 | December 13, 2018 6:13:41 PM GMT+08:00
CTUDT00598633 | Refresh | Closed | rodan_d | June 6, 2018 9:01:29 PM GMT+08:00 | In voice manager some of the phrases were not translated for all the languages in the voice manager UI. | V4.1.7_WIP-3 | WIP12 | If user does not understand the string in English because it is not translated he can disable voice command and provide instructions to user manually. If necessary user could open the voice manager and play the original message and save it as a new custom voice command and store it with a name in the localized language. | System |  | NA | Fixed | kr_m | 0 | 3 | Console Software | Console4 | NO | Closed | WIP11.1SegReq | IQon1 |  | WIP12 | December 13, 2018 6:13:41 PM GMT+08:00
CTUDT00598632 | Refresh | Closed | rodan_d | June 6, 2018 8:58:48 PM GMT+08:00 | Incorrect values for TABLE_FEED_PER_ROTATION in DICOM files on Ingenuity CT 4.1.6. | 4.1.6.25032 | WIP15 |  | Subsystem |  | NA | Fixed | sowmya_k | 0 | 3 | Console Software | Console4 | NO | Closed | WIP15SubCom | Ovation | PR6334765 | WIP15 | March 4, 2019 8:11:33 PM GMT+08:00
CTUDT00598632 | Refresh | Closed | rodan_d | June 6, 2018 8:58:48 PM GMT+08:00 | Incorrect values for TABLE_FEED_PER_ROTATION in DICOM files on Ingenuity CT 4.1.6. | 4.1.6.25032 | WIP15 |  | Subsystem |  | NA | Fixed | sowmya_k | 0 | 3 | Console Software | Console4 | NO | Closed | WIP15SubCom | Brilliance 64 | PR6334765 | WIP15 | March 4, 2019 8:11:33 PM GMT+08:00
CTUDT00598632 | Refresh | Closed | rodan_d | June 6, 2018 8:58:48 PM GMT+08:00 | Incorrect values for TABLE_FEED_PER_ROTATION in DICOM files on Ingenuity CT 4.1.6. | 4.1.6.25032 | WIP15 |  | Subsystem |  | NA | Fixed | sowmya_k | 0 | 3 | Console Software | Console4 | NO | Closed | WIP15SubCom | Ingenuity 64 | PR6334765 | WIP15 | March 4, 2019 8:11:33 PM GMT+08:00
CTUDT00598632 | Refresh | Closed | rodan_d | June 6, 2018 8:58:48 PM GMT+08:00 | Incorrect values for TABLE_FEED_PER_ROTATION in DICOM files on Ingenuity CT 4.1.6. | 4.1.6.25032 | WIP15 |  | Subsystem |  | NA | Fixed | sowmya_k | 0 | 3 | Console Software | Console4 | NO | Closed | WIP15SubCom | IQon1 | PR6334765 | WIP15 | March 4, 2019 8:11:33 PM GMT+08:00
CTUDT00598630 | Refresh | Closed | rodan_d | June 6, 2018 8:54:34 PM GMT+08:00 | IV: 4.1.6 - Unable to review results in view scene | WIP 106.4 v4.1.6.xxx32 | WIP15 | Manually load images to CTViewer or Quick Review | Subsystem |  | NA | Fixed | sowmya_k | 0 | 3 | Console Software | Console4 | NO | Closed | WIP15SegReq | Ovation |  | WIP15 | April 4, 2019 4:37:25 PM GMT+08:00
CTUDT00598630 | Refresh | Closed | rodan_d | June 6, 2018 8:54:34 PM GMT+08:00 | IV: 4.1.6 - Unable to review results in view scene | WIP 106.4 v4.1.6.xxx32 | WIP15 | Manually load images to CTViewer or Quick Review | Subsystem |  | NA | Fixed | sowmya_k | 0 | 3 | Console Software | Console4 | NO | Closed | WIP15SegReq | Brilliance 64 |  | WIP15 | April 4, 2019 4:37:25 PM GMT+08:00
CTUDT00598630 | Refresh | Closed | rodan_d | June 6, 2018 8:54:34 PM GMT+08:00 | IV: 4.1.6 - Unable to review results in view scene | WIP 106.4 v4.1.6.xxx32 | WIP15 | Manually load images to CTViewer or Quick Review | Subsystem |  | NA | Fixed | sowmya_k | 0 | 3 | Console Software | Console4 | NO | Closed | WIP15SegReq | Ingenuity 64 |  | WIP15 | April 4, 2019 4:37:25 PM GMT+08:00
CTUDT00598629 | Refresh | Closed | rodan_d | June 6, 2018 8:52:44 PM GMT+08:00 | IV: 4.1.6 Serviceability - Critical error occurs upon launch of workflow after 4.1.6 patch update | 4.1.6.22030 |  |  | Subsystem |  |  | Duplicate | karkour | 0 | 4 |  | GCX | NO | Closed | WIP16SubCom | Brilliance 64
CTUDT00598627 | Refresh | Closed | rodan_d | June 6, 2018 8:44:55 PM GMT+08:00 | IV: HDI: Incorrect instructions in the "Rename EC" How Do I video | v4.0.0.xx621 WIP 609 | Wip 15 |  | Subsystem |  | N/A | Fixed | burke_c | 0 | 4 | Clinical Solutions | Support-System | NO | Closed | WIP16SubCom | Brilliance 64 |  | Wip 15 | May 10, 2019 2:23:40 AM GMT+08:00
CTUDT00598627 | Refresh | Closed | rodan_d | June 6, 2018 8:44:55 PM GMT+08:00 | IV: HDI: Incorrect instructions in the "Rename EC" How Do I video | v4.0.0.xx621 WIP 609 | Wip 15 |  | Subsystem |  | N/A | Fixed | burke_c | 0 | 4 | Clinical Solutions | Support-System | NO | Closed | WIP16SubCom | Ingenuity 64 |  | Wip 15 | May 10, 2019 2:23:40 AM GMT+08:00
CTUDT00598626 | Refresh | Closed | rodan_d | June 6, 2018 8:43:34 PM GMT+08:00 | IV: MITA: Delay in dose log entry following dose notification | 4.1.0.XX241-WIP211 |  | There is no need for a workaround. | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 | Console Generic | Console4 | NO | Closed |  | Ovation
CTUDT00598626 | Refresh | Closed | rodan_d | June 6, 2018 8:43:34 PM GMT+08:00 | IV: MITA: Delay in dose log entry following dose notification | 4.1.0.XX241-WIP211 |  | There is no need for a workaround. | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 | Console Generic | Console4 | NO | Closed |  | Brilliance 64
CTUDT00598626 | Refresh | Closed | rodan_d | June 6, 2018 8:43:34 PM GMT+08:00 | IV: MITA: Delay in dose log entry following dose notification | 4.1.0.XX241-WIP211 |  | There is no need for a workaround. | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 | Console Generic | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00598625 | Refresh | Closed | rodan_d | June 6, 2018 8:42:29 PM GMT+08:00 | IV: Pulmonary: "Scan XX is invalid and cannot be loaded. In order to fix the Exam Card.." message appeared afer a surview and free breathing EC was performed. | 4.1.3.XX005-WIP309 |  | If the message listed in this record is seen during a pulmo exam, try to decrease the scan time by decreasing the length or by increasing the pitch. | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 | Console Generic | Console4 | NO | Closed |  | Brilliance 64
CTUDT00598625 | Refresh | Closed | rodan_d | June 6, 2018 8:42:29 PM GMT+08:00 | IV: Pulmonary: "Scan XX is invalid and cannot be loaded. In order to fix the Exam Card.." message appeared afer a surview and free breathing EC was performed. | 4.1.3.XX005-WIP309 |  | If the message listed in this record is seen during a pulmo exam, try to decrease the scan time by decreasing the length or by increasing the pitch. | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 | Console Generic | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00598623 | Refresh | Closed | rodan_d | June 6, 2018 8:29:52 PM GMT+08:00 | Jog perfusion protocols do not work after upgrade to 4.1.6 | 4.1.6 |  | User has to create a new/custom perfusion exam card from perfusion exam card present in factory. | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 3 |  | Console4 | NO | Closed |  | Brilliance 64 | PR6090743 PR6077307 PR6189709
CTUDT00598623 | Refresh | Closed | rodan_d | June 6, 2018 8:29:52 PM GMT+08:00 | Jog perfusion protocols do not work after upgrade to 4.1.6 | 4.1.6 |  | User has to create a new/custom perfusion exam card from perfusion exam card present in factory. | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 3 |  | Console4 | NO | Closed |  | Ingenuity 64 | PR6090743 PR6077307 PR6189709
CTUDT00598622 | Refresh | Closed | rodan_d | June 6, 2018 8:24:54 PM GMT+08:00 | McAfee Application Control should be updated with new version. | V4.1.7-WIP-4 |  |  | Subsystem |  |  | Does Not Apply | sanjeeva_a | 0 | 2 |  | Console4 | NO | Closed |  | Ovation
CTUDT00598622 | Refresh | Closed | rodan_d | June 6, 2018 8:24:54 PM GMT+08:00 | McAfee Application Control should be updated with new version. | V4.1.7-WIP-4 |  |  | Subsystem |  |  | Does Not Apply | sanjeeva_a | 0 | 2 |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00598622 | Refresh | Closed | rodan_d | June 6, 2018 8:24:54 PM GMT+08:00 | McAfee Application Control should be updated with new version. | V4.1.7-WIP-4 |  |  | Subsystem |  |  | Does Not Apply | sanjeeva_a | 0 | 2 |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00598621 | Refresh | Closed | rodan_d | June 6, 2018 8:23:06 PM GMT+08:00 | Multi shot Axial may crash after resuming a reconstruction | V4.1.6 | WIP6 | The user will need to perform an offline reconstruction to recover. | Subsystem |  | N/A | Fixed | rakic_k | 0 | 3 | CIRS_SW_ImgNodeSvr | CIRS | YES | Closed | WIP6SubCom | Ovation | 5748832 | WIP6 | September 18, 2018 8:45:40 PM GMT+08:00
CTUDT00598621 | Refresh | Closed | rodan_d | June 6, 2018 8:23:06 PM GMT+08:00 | Multi shot Axial may crash after resuming a reconstruction | V4.1.6 | WIP6 | The user will need to perform an offline reconstruction to recover. | Subsystem |  | N/A | Fixed | rakic_k | 0 | 3 | CIRS_SW_ImgNodeSvr | CIRS | YES | Closed | WIP6SubCom | Brilliance 64 | 5748832 | WIP6 | September 18, 2018 8:45:40 PM GMT+08:00
CTUDT00598621 | Refresh | Closed | rodan_d | June 6, 2018 8:23:06 PM GMT+08:00 | Multi shot Axial may crash after resuming a reconstruction | V4.1.6 | WIP6 | The user will need to perform an offline reconstruction to recover. | Subsystem |  | N/A | Fixed | rakic_k | 0 | 3 | CIRS_SW_ImgNodeSvr | CIRS | YES | Closed | WIP6SubCom | Brilliance Big Bore | 5748832 | WIP6 | September 18, 2018 8:45:40 PM GMT+08:00
CTUDT00598621 | Refresh | Closed | rodan_d | June 6, 2018 8:23:06 PM GMT+08:00 | Multi shot Axial may crash after resuming a reconstruction | V4.1.6 | WIP6 | The user will need to perform an offline reconstruction to recover. | Subsystem |  | N/A | Fixed | rakic_k | 0 | 3 | CIRS_SW_ImgNodeSvr | CIRS | YES | Closed | WIP6SubCom | Ingenuity 64 | 5748832 | WIP6 | September 18, 2018 8:45:40 PM GMT+08:00
CTUDT00598620 | Refresh | Closed | rodan_d | June 6, 2018 8:20:45 PM GMT+08:00 | Negative encoder value can cause A-Plane Opening to fail | 4.1.3.XX011-WIP309.1 |  |  | Subsystem | SP1 | NA | No Plan to Fix | lin_s | 0 | 2 |  | GCX | NO | Closed | WIP18SubCom | Brilliance 64
CTUDT00598620 | Refresh | Closed | rodan_d | June 6, 2018 8:20:45 PM GMT+08:00 | Negative encoder value can cause A-Plane Opening to fail | 4.1.3.XX011-WIP309.1 |  |  | Subsystem | SP1 | NA | No Plan to Fix | lin_s | 0 | 2 |  | GCX | NO | Closed | WIP18SubCom | Brilliance Big Bore
CTUDT00598620 | Refresh | Closed | rodan_d | June 6, 2018 8:20:45 PM GMT+08:00 | Negative encoder value can cause A-Plane Opening to fail | 4.1.3.XX011-WIP309.1 |  |  | Subsystem | SP1 | NA | No Plan to Fix | lin_s | 0 | 2 |  | GCX | NO | Closed | WIP18SubCom | Ingenuity 64
CTUDT00598620 | Refresh | Closed | rodan_d | June 6, 2018 8:20:45 PM GMT+08:00 | Negative encoder value can cause A-Plane Opening to fail | 4.1.3.XX011-WIP309.1 |  |  | Subsystem | SP1 | NA | No Plan to Fix | lin_s | 0 | 2 |  | GCX | NO | Closed | WIP18SubCom | Ingenuity TF 64
CTUDT00598620 | Refresh | Closed | rodan_d | June 6, 2018 8:20:45 PM GMT+08:00 | Negative encoder value can cause A-Plane Opening to fail | 4.1.3.XX011-WIP309.1 |  |  | Subsystem | SP1 | NA | No Plan to Fix | lin_s | 0 | 2 |  | GCX | NO | Closed | WIP18SubCom | Veyron
CTUDT00598619 | Refresh | Closed | rodan_d | June 6, 2018 7:09:43 PM GMT+08:00 | Pediatric Head image quality exhibits halo artifact at the bone brain interface | 4.1.4.XX001 | WIP17 | Use of High Resolution for brain imaging | System | Verified1 | CT-NM.SRS.IQ.1408 | Fixed | richards_b | 0 | 2 | CIRS_SW_Installation | Support-Physics | YES | Closed | WIP17SubCom | Brilliance 64 | PR3934861 PR3960567 | WIP17 | June 17, 2019 8:53:39 PM GMT+08:00
CTUDT00598619 | Refresh | Closed | rodan_d | June 6, 2018 7:09:43 PM GMT+08:00 | Pediatric Head image quality exhibits halo artifact at the bone brain interface | 4.1.4.XX001 | WIP17 | Use of High Resolution for brain imaging | System | Verified1 | CT-NM.SRS.IQ.1408 | Fixed | richards_b | 0 | 2 | CIRS_SW_Installation | Support-Physics | YES | Closed | WIP17SubCom | Ingenuity 64 | PR3934861 PR3960567 | WIP17 | June 17, 2019 8:53:39 PM GMT+08:00
CTUDT00598618 | Refresh | Closed | rodan_d | June 6, 2018 7:01:25 PM GMT+08:00 | Periodic Noise bands observed in reformats along the Z-directions for low heart-rate, high pitch gated scans with Cardiac dose right on. | 4.1.3 | WIP19 | The user can minimize the noise bands with the following settings: • Pitch value = 0.2 or less or • Cardiac DoseRight is off | Subsystem | IFU | N/A | Fixed | charles_j | 0 | 2 | User Documentation | Support-System | YES | Closed | PostVerification | Ovation | PR4693125 CFR 9357647 | WIP19 | July 31, 2019 5:00:04 PM GMT+08:00
CTUDT00598617 | Refresh | Closed | rodan_d | June 6, 2018 6:59:25 PM GMT+08:00 | Plan box jumps while planning Axial 2D Brain with tilt | 4.1.6 | WIP11.1 | Yes before pressing Go, User can re-plan and modify the plan box as per his need. | Subsystem |  | None | Fixed | patra_r | 0 | 3 | Console Software | Console4 | YES | Closed | WIP11.1SegReq | Brilliance 64 | PR6445025 | WIP11.1 | November 12, 2018 6:08:34 PM GMT+08:00
CTUDT00598617 | Refresh | Closed | rodan_d | June 6, 2018 6:59:25 PM GMT+08:00 | Plan box jumps while planning Axial 2D Brain with tilt | 4.1.6 | WIP11.1 | Yes before pressing Go, User can re-plan and modify the plan box as per his need. | Subsystem |  | None | Fixed | patra_r | 0 | 3 | Console Software | Console4 | YES | Closed | WIP11.1SegReq | Ingenuity 64 | PR6445025 | WIP11.1 | November 12, 2018 6:08:34 PM GMT+08:00
CTUDT00598616 | Refresh | Closed | rodan_d | June 6, 2018 6:57:44 PM GMT+08:00 | Plan Scene/Acquisition scene: When performing Dual Surview Auto does not scan both surviews/ Auto is shown instead of Manual. | 4.1.2.XX017-WIP305.3 |  | Press Auto button as prompted and Next afterwards | Subsystem |  |  | Does Not Apply | sanjeeva_a | 0 | 4 | System | Console4 | NO | Closed |  | Ovation
CTUDT00598616 | Refresh | Closed | rodan_d | June 6, 2018 6:57:44 PM GMT+08:00 | Plan Scene/Acquisition scene: When performing Dual Surview Auto does not scan both surviews/ Auto is shown instead of Manual. | 4.1.2.XX017-WIP305.3 |  | Press Auto button as prompted and Next afterwards | Subsystem |  |  | Does Not Apply | sanjeeva_a | 0 | 4 | System | Console4 | NO | Closed |  | Brilliance 64
CTUDT00598616 | Refresh | Closed | rodan_d | June 6, 2018 6:57:44 PM GMT+08:00 | Plan Scene/Acquisition scene: When performing Dual Surview Auto does not scan both surviews/ Auto is shown instead of Manual. | 4.1.2.XX017-WIP305.3 |  | Press Auto button as prompted and Next afterwards | Subsystem |  |  | Does Not Apply | sanjeeva_a | 0 | 4 | System | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00598616 | Refresh | Closed | rodan_d | June 6, 2018 6:57:44 PM GMT+08:00 | Plan Scene/Acquisition scene: When performing Dual Surview Auto does not scan both surviews/ Auto is shown instead of Manual. | 4.1.2.XX017-WIP305.3 |  | Press Auto button as prompted and Next afterwards | Subsystem |  |  | Does Not Apply | sanjeeva_a | 0 | 4 | System | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00598615 | Refresh | Closed | rodan_d | June 6, 2018 6:54:36 PM GMT+08:00 | Real Time Data Tool Filament Current always reads 65535 | 4.1.4.XX001 | DHF317363 Rev02 | Service engineers do not use the filament current Real Time Data tool instead they use the kV and mA when analyzing a problem | Subsystem |  | NA | Fixed | karkour | 0 | 3 | GCX-GCX | GCX | NO | Closed | BB_WIP5_SegReq | Brilliance 64 |  | DHF317363 Rev02 | May 7, 2019 11:49:52 PM GMT+08:00
CTUDT00598615 | Refresh | Closed | rodan_d | June 6, 2018 6:54:36 PM GMT+08:00 | Real Time Data Tool Filament Current always reads 65535 | 4.1.4.XX001 | DHF317363 Rev02 | Service engineers do not use the filament current Real Time Data tool instead they use the kV and mA when analyzing a problem | Subsystem |  | NA | Fixed | karkour | 0 | 3 | GCX-GCX | GCX | NO | Closed | BB_WIP5_SegReq | Ingenuity 64 |  | DHF317363 Rev02 | May 7, 2019 11:49:52 PM GMT+08:00
CTUDT00598606 | Refresh | Closed | rodan_d | June 6, 2018 4:16:54 PM GMT+08:00 | Reconstruction produces image block artifacts | V4.1.6 | WIP12 | Site can perform offline reconstruction to recover | Subsystem |  | N/A | Fixed | richards_b | 0 | 3 | CIRS_SW_ImageProcess | CIRS | YES | Closed | WIP12SubCom | Ovation | CFR9419363 CFR9426695 PR6537334/CFR9431366 PR7953031 | WIP12 | December 19, 2018 4:42:05 AM GMT+08:00
CTUDT00598606 | Refresh | Closed | rodan_d | June 6, 2018 4:16:54 PM GMT+08:00 | Reconstruction produces image block artifacts | V4.1.6 | WIP12 | Site can perform offline reconstruction to recover | Subsystem |  | N/A | Fixed | richards_b | 0 | 3 | CIRS_SW_ImageProcess | CIRS | YES | Closed | WIP12SubCom | Brilliance 64 | CFR9419363 CFR9426695 PR6537334/CFR9431366 PR7953031 | WIP12 | December 19, 2018 4:42:05 AM GMT+08:00
CTUDT00598606 | Refresh | Closed | rodan_d | June 6, 2018 4:16:54 PM GMT+08:00 | Reconstruction produces image block artifacts | V4.1.6 | WIP12 | Site can perform offline reconstruction to recover | Subsystem |  | N/A | Fixed | richards_b | 0 | 3 | CIRS_SW_ImageProcess | CIRS | YES | Closed | WIP12SubCom | Brilliance Big Bore | CFR9419363 CFR9426695 PR6537334/CFR9431366 PR7953031 | WIP12 | December 19, 2018 4:42:05 AM GMT+08:00
CTUDT00598606 | Refresh | Closed | rodan_d | June 6, 2018 4:16:54 PM GMT+08:00 | Reconstruction produces image block artifacts | V4.1.6 | WIP12 | Site can perform offline reconstruction to recover | Subsystem |  | N/A | Fixed | richards_b | 0 | 3 | CIRS_SW_ImageProcess | CIRS | YES | Closed | WIP12SubCom | Ingenuity 64 | CFR9419363 CFR9426695 PR6537334/CFR9431366 PR7953031 | WIP12 | December 19, 2018 4:42:05 AM GMT+08:00
CTUDT00598605 | Refresh | Closed | rodan_d | June 6, 2018 4:05:21 PM GMT+08:00 | Remote Service Network (RSN) registration cannot be completed on Ingenuity Core 128 728324 China L4L, Ingenuity Core (728321) and Ingenuity Core 128 (728323). | v4.1.6 |  |  | Subsystem |  |  | Does Not Apply | sanjeeva_a | 0 | 2 |  | Service Tools | NO | Closed |  | Ingenuity 64
CTUDT00598604 | Refresh | Closed | rodan_d | June 6, 2018 4:04:04 PM GMT+08:00 | Scan ruler not updated in CCT Fluoro scan | Eagle_WIP202 | BB_WIP6 |  | Subsystem |  | NA | Fixed | agrawal_ay | 0 | 2 | Console Generic | Console4 | NO | Closed | BB_WIP6_SegReq | Brilliance Big Bore |  | BB_WIP6 | April 12, 2019 7:29:45 PM GMT+08:00
CTUDT00598603 | Refresh | Closed | rodan_d | June 6, 2018 4:00:48 PM GMT+08:00 | Upgrade utility does not deploy user exam cards when some exam cards names contain foreign characters | V4.1.7-WIP-5.1 | WIP16 BB_WIP6 |  | Subsystem |  | N/A | Fixed | pandey_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP17SubCom | Ovation |  | WIP16 BB_WIP6 | April 29, 2019 9:17:49 PM GMT+08:00
CTUDT00598603 | Refresh | Closed | rodan_d | June 6, 2018 4:00:48 PM GMT+08:00 | Upgrade utility does not deploy user exam cards when some exam cards names contain foreign characters | V4.1.7-WIP-5.1 | WIP16 BB_WIP6 |  | Subsystem |  | N/A | Fixed | pandey_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP17SubCom | Brilliance 64 |  | WIP16 BB_WIP6 | April 29, 2019 9:17:49 PM GMT+08:00
CTUDT00598603 | Refresh | Closed | rodan_d | June 6, 2018 4:00:48 PM GMT+08:00 | Upgrade utility does not deploy user exam cards when some exam cards names contain foreign characters | V4.1.7-WIP-5.1 | WIP16 BB_WIP6 |  | Subsystem |  | N/A | Fixed | pandey_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP17SubCom | Brilliance Big Bore |  | WIP16 BB_WIP6 | April 29, 2019 9:17:49 PM GMT+08:00
CTUDT00598603 | Refresh | Closed | rodan_d | June 6, 2018 4:00:48 PM GMT+08:00 | Upgrade utility does not deploy user exam cards when some exam cards names contain foreign characters | V4.1.7-WIP-5.1 | WIP16 BB_WIP6 |  | Subsystem |  | N/A | Fixed | pandey_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP17SubCom | Ingenuity 64 |  | WIP16 BB_WIP6 | April 29, 2019 9:17:49 PM GMT+08:00
CTUDT00598603 | Refresh | Closed | rodan_d | June 6, 2018 4:00:48 PM GMT+08:00 | Upgrade utility does not deploy user exam cards when some exam cards names contain foreign characters | V4.1.7-WIP-5.1 | WIP16 BB_WIP6 |  | Subsystem |  | N/A | Fixed | pandey_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP17SubCom | IQon1 |  | WIP16 BB_WIP6 | April 29, 2019 9:17:49 PM GMT+08:00
CTUDT00598602 | Refresh | Closed | rodan_d | June 6, 2018 3:55:14 PM GMT+08:00 | v4.1.3 FOK : Short Tube Conditioning Needed message appears while scanning | v4.1.3 |  | Run short tube conditioning | Subsystem |  |  | Does Not Apply | sanjeeva_a | 0 | 4 |  | Console4 | NO | Closed |  | Ovation | PR 5172281 PR5377699/CFM9386975
CTUDT00598602 | Refresh | Closed | rodan_d | June 6, 2018 3:55:14 PM GMT+08:00 | v4.1.3 FOK : Short Tube Conditioning Needed message appears while scanning | v4.1.3 |  | Run short tube conditioning | Subsystem |  |  | Does Not Apply | sanjeeva_a | 0 | 4 |  | Console4 | NO | Closed |  | Brilliance 64 | PR 5172281 PR5377699/CFM9386975
CTUDT00598602 | Refresh | Closed | rodan_d | June 6, 2018 3:55:14 PM GMT+08:00 | v4.1.3 FOK : Short Tube Conditioning Needed message appears while scanning | v4.1.3 |  | Run short tube conditioning | Subsystem |  |  | Does Not Apply | sanjeeva_a | 0 | 4 |  | Console4 | NO | Closed |  | Brilliance Big Bore | PR 5172281 PR5377699/CFM9386975
CTUDT00598602 | Refresh | Closed | rodan_d | June 6, 2018 3:55:14 PM GMT+08:00 | v4.1.3 FOK : Short Tube Conditioning Needed message appears while scanning | v4.1.3 |  | Run short tube conditioning | Subsystem |  |  | Does Not Apply | sanjeeva_a | 0 | 4 |  | Console4 | NO | Closed |  | Ingenuity 64 | PR 5172281 PR5377699/CFM9386975
CTUDT00598602 | Refresh | Closed | rodan_d | June 6, 2018 3:55:14 PM GMT+08:00 | v4.1.3 FOK : Short Tube Conditioning Needed message appears while scanning | v4.1.3 |  | Run short tube conditioning | Subsystem |  |  | Does Not Apply | sanjeeva_a | 0 | 4 |  | Console4 | NO | Closed |  | IQon1 | PR 5172281 PR5377699/CFM9386975
CTUDT00598601 | Refresh | Closed | rodan_d | June 6, 2018 3:41:03 PM GMT+08:00 | v4.1.7:No of cycles display in results plan box and Acquisition plan box differs when table height is changed. | V4.1.7-WIP-5 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | babu_s | 1 | 4 |  | Console4 | NO | Closed |  | Ovation
CTUDT00598601 | Refresh | Closed | rodan_d | June 6, 2018 3:41:03 PM GMT+08:00 | v4.1.7:No of cycles display in results plan box and Acquisition plan box differs when table height is changed. | V4.1.7-WIP-5 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | babu_s | 1 | 4 |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00598601 | Refresh | Closed | rodan_d | June 6, 2018 3:41:03 PM GMT+08:00 | v4.1.7:No of cycles display in results plan box and Acquisition plan box differs when table height is changed. | V4.1.7-WIP-5 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | babu_s | 1 | 4 |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00598601 | Refresh | Closed | rodan_d | June 6, 2018 3:41:03 PM GMT+08:00 | v4.1.7:No of cycles display in results plan box and Acquisition plan box differs when table height is changed. | V4.1.7-WIP-5 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | babu_s | 1 | 4 |  | Console4 | NO | Closed |  | IQon1
CTUDT00598599 | Refresh | Closed | rodan_d | June 6, 2018 3:31:18 PM GMT+08:00 | Val: 4.1.7 - "ExamCardManager has stopped working" after selecting ECM in utilities | 4.1.7.XX021 WIP 5.1 |  | user need to logoff and login the workflow or restart the system to clean up any memory instability caused because of previous exception. This issue has been identified in the Release Notes 459801063602 with a workaround. | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 3 |  | Console4 | NO | Closed |  | Ovation
CTUDT00598599 | Refresh | Closed | rodan_d | June 6, 2018 3:31:18 PM GMT+08:00 | Val: 4.1.7 - "ExamCardManager has stopped working" after selecting ECM in utilities | 4.1.7.XX021 WIP 5.1 |  | user need to logoff and login the workflow or restart the system to clean up any memory instability caused because of previous exception. This issue has been identified in the Release Notes 459801063602 with a workaround. | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 3 |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00598599 | Refresh | Closed | rodan_d | June 6, 2018 3:31:18 PM GMT+08:00 | Val: 4.1.7 - "ExamCardManager has stopped working" after selecting ECM in utilities | 4.1.7.XX021 WIP 5.1 |  | user need to logoff and login the workflow or restart the system to clean up any memory instability caused because of previous exception. This issue has been identified in the Release Notes 459801063602 with a workaround. | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 3 |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00598599 | Refresh | Closed | rodan_d | June 6, 2018 3:31:18 PM GMT+08:00 | Val: 4.1.7 - "ExamCardManager has stopped working" after selecting ECM in utilities | 4.1.7.XX021 WIP 5.1 |  | user need to logoff and login the workflow or restart the system to clean up any memory instability caused because of previous exception. This issue has been identified in the Release Notes 459801063602 with a workaround. | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 3 |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00598599 | Refresh | Closed | rodan_d | June 6, 2018 3:31:18 PM GMT+08:00 | Val: 4.1.7 - "ExamCardManager has stopped working" after selecting ECM in utilities | 4.1.7.XX021 WIP 5.1 |  | user need to logoff and login the workflow or restart the system to clean up any memory instability caused because of previous exception. This issue has been identified in the Release Notes 459801063602 with a workaround. | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 3 |  | Console4 | NO | Closed |  | IQon1
CTUDT00598598 | Refresh | Closed | rodan_d | June 6, 2018 3:29:18 PM GMT+08:00 | Val: 4.1.7 - Injection summary does not come over once injection is stopped | V4.1.7-WIP-5.1 |  | N/A | Subsystem |  |  | Duplicate | Prakash_Chand_Yadav | 0 | 2 |  | Console4 | NO | Closed |  | Ovation
CTUDT00598598 | Refresh | Closed | rodan_d | June 6, 2018 3:29:18 PM GMT+08:00 | Val: 4.1.7 - Injection summary does not come over once injection is stopped | V4.1.7-WIP-5.1 |  | N/A | Subsystem |  |  | Duplicate | Prakash_Chand_Yadav | 0 | 2 |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00598598 | Refresh | Closed | rodan_d | June 6, 2018 3:29:18 PM GMT+08:00 | Val: 4.1.7 - Injection summary does not come over once injection is stopped | V4.1.7-WIP-5.1 |  | N/A | Subsystem |  |  | Duplicate | Prakash_Chand_Yadav | 0 | 2 |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00598597 | Refresh | Closed | rodan_d | June 6, 2018 3:23:40 PM GMT+08:00 | Val: 4.1.7 - 'Start learning' should not be available for a 4D scan | 4.1.7.XX021 WIP 5.1 |  |  | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00598597 | Refresh | Closed | rodan_d | June 6, 2018 3:23:40 PM GMT+08:00 | Val: 4.1.7 - 'Start learning' should not be available for a 4D scan | 4.1.7.XX021 WIP 5.1 |  |  | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00598596 | Refresh | Closed | rodan_d | June 6, 2018 3:20:14 PM GMT+08:00 | Val: 4.1.7 - Wrong message to user, states Bellow is not connected to the patient. | V4.1.7-WIP-5.1 | WIP16 | N/A | Subsystem |  | Console-SRS.11271, Console-SRS.11269 | Fixed | pandey_a | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP5_SegReq | Brilliance 64 |  | WIP16 | April 22, 2019 8:32:06 PM GMT+08:00
CTUDT00598596 | Refresh | Closed | rodan_d | June 6, 2018 3:20:14 PM GMT+08:00 | Val: 4.1.7 - Wrong message to user, states Bellow is not connected to the patient. | V4.1.7-WIP-5.1 | WIP16 | N/A | Subsystem |  | Console-SRS.11271, Console-SRS.11269 | Fixed | pandey_a | 0 | 4 | Console Software | Console4 | NO | Closed | BB_WIP5_SegReq | Ingenuity 64 |  | WIP16 | April 22, 2019 8:32:06 PM GMT+08:00
CTUDT00598595 | Refresh | Closed | rodan_d | June 6, 2018 3:18:46 PM GMT+08:00 | Val: 4.1.7 Service - "A problem occurred, please retry..." error appeared during clinical scan test | V4.1.7-WIP-5.1 | WIP16 | user has to click “OK”, then press “GO” again. | Subsystem |  | NA | Fixed | karkour | 0 | 2 | Gantry Software | GCX | NO | Closed | WIP16SubCom | Ovation |  | WIP16 | June 12, 2019 3:17:40 AM GMT+08:00
CTUDT00598595 | Refresh | Closed | rodan_d | June 6, 2018 3:18:46 PM GMT+08:00 | Val: 4.1.7 Service - "A problem occurred, please retry..." error appeared during clinical scan test | V4.1.7-WIP-5.1 | WIP16 | user has to click “OK”, then press “GO” again. | Subsystem |  | NA | Fixed | karkour | 0 | 2 | Gantry Software | GCX | NO | Closed | WIP16SubCom | Brilliance 64 |  | WIP16 | June 12, 2019 3:17:40 AM GMT+08:00
CTUDT00598595 | Refresh | Closed | rodan_d | June 6, 2018 3:18:46 PM GMT+08:00 | Val: 4.1.7 Service - "A problem occurred, please retry..." error appeared during clinical scan test | V4.1.7-WIP-5.1 | WIP16 | user has to click “OK”, then press “GO” again. | Subsystem |  | NA | Fixed | karkour | 0 | 2 | Gantry Software | GCX | NO | Closed | WIP16SubCom | Ingenuity 64 |  | WIP16 | June 12, 2019 3:17:40 AM GMT+08:00
CTUDT00598594 | Refresh | Closed | rodan_d | June 6, 2018 3:14:01 PM GMT+08:00 | Varian error message "WARNING: No Respiratory Signal !" | V3.6.8 |  | Before Scan: Wait for warning message to clear and start scan During retrospective pulmo gated exam: Ignore message and continue normally. During prospective pulmo gated exam: Scan is cancelled. Re-scan patient. | Subsystem |  | CT-NM.SRS.Modes.1712 | No Plan to Fix | zhu_w | 0 | 4 |  | GCX | YES | Closed |  | Brilliance 64
CTUDT00598594 | Refresh | Closed | rodan_d | June 6, 2018 3:14:01 PM GMT+08:00 | Varian error message "WARNING: No Respiratory Signal !" | V3.6.8 |  | Before Scan: Wait for warning message to clear and start scan During retrospective pulmo gated exam: Ignore message and continue normally. During prospective pulmo gated exam: Scan is cancelled. Re-scan patient. | Subsystem |  | CT-NM.SRS.Modes.1712 | No Plan to Fix | zhu_w | 0 | 4 |  | GCX | YES | Closed |  | Brilliance Big Bore
CTUDT00598594 | Refresh | Closed | rodan_d | June 6, 2018 3:14:01 PM GMT+08:00 | Varian error message "WARNING: No Respiratory Signal !" | V3.6.8 |  | Before Scan: Wait for warning message to clear and start scan During retrospective pulmo gated exam: Ignore message and continue normally. During prospective pulmo gated exam: Scan is cancelled. Re-scan patient. | Subsystem |  | CT-NM.SRS.Modes.1712 | No Plan to Fix | zhu_w | 0 | 4 |  | GCX | YES | Closed |  | Ingenuity 64
CTUDT00598592 | Refresh | Closed | bagdady_b | June 6, 2018 2:53:50 PM GMT+08:00 | Refresh- CIRS Recon box cannot initialize due to time exceeded | WIP7.1 | WIP18 | The user would need to select Go again to proceed with the clinical scan. | Subsystem | Spine Axial EC | N/A | Fixed | richards_b | 0 | 3 | CIRS_SW_Supervisor | CIRS | NO | Closed | WIP18SubCom | IQon1 |  | WIP18 | July 12, 2019 12:54:05 AM GMT+08:00
CTUDT00598584 | Refresh | Closed | rodan_d | June 5, 2018 9:51:20 PM GMT+08:00 | RFC-BR64 Technical Reference Guide for edits for IQ section of the infant phantom | 4.1.3.XX008-WIP310 | WIP16 |  | Subsystem | Documentation | N/A | Fixed | rajalakshmi_j | 0 | 3 | User Documentation | Support-System | NO | Closed | WIP12SegReq | Ovation |  | WIP16 | May 29, 2019 2:04:32 PM GMT+08:00
CTUDT00598584 | Refresh | Closed | rodan_d | June 5, 2018 9:51:20 PM GMT+08:00 | RFC-BR64 Technical Reference Guide for edits for IQ section of the infant phantom | 4.1.3.XX008-WIP310 | WIP16 |  | Subsystem | Documentation | N/A | Fixed | rajalakshmi_j | 0 | 3 | User Documentation | Support-System | NO | Closed | WIP12SegReq | Brilliance 64 |  | WIP16 | May 29, 2019 2:04:32 PM GMT+08:00
CTUDT00598584 | Refresh | Closed | rodan_d | June 5, 2018 9:51:20 PM GMT+08:00 | RFC-BR64 Technical Reference Guide for edits for IQ section of the infant phantom | 4.1.3.XX008-WIP310 | WIP16 |  | Subsystem | Documentation | N/A | Fixed | rajalakshmi_j | 0 | 3 | User Documentation | Support-System | NO | Closed | WIP12SegReq | Ingenuity 64 |  | WIP16 | May 29, 2019 2:04:32 PM GMT+08:00
CTUDT00598584 | Refresh | Closed | rodan_d | June 5, 2018 9:51:20 PM GMT+08:00 | RFC-BR64 Technical Reference Guide for edits for IQ section of the infant phantom | 4.1.3.XX008-WIP310 | WIP16 |  | Subsystem | Documentation | N/A | Fixed | rajalakshmi_j | 0 | 3 | User Documentation | Support-System | NO | Closed | WIP12SegReq | IQon1 |  | WIP16 | May 29, 2019 2:04:32 PM GMT+08:00
CTUDT00598583 | Refresh | Closed | rodan_d | June 5, 2018 9:49:53 PM GMT+08:00 | Report - Send findings from CS to report using Mass Score protocol , in the impression section 2nd paragraph the comments are displayed for all the scores | BR BB 4.2 WIP 104.2 | WIP17 |  | Subsystem | Verified1 | NA | Fixed | redrouthu_r | 0 | 2 | Console USP_EBW | iCAP | YES | Closed | WIP14SegReq | Ovation | PR#5940478 PR#6123359 PR#6393965 | WIP17 | June 13, 2019 1:35:55 PM GMT+08:00
CTUDT00598583 | Refresh | Closed | rodan_d | June 5, 2018 9:49:53 PM GMT+08:00 | Report - Send findings from CS to report using Mass Score protocol , in the impression section 2nd paragraph the comments are displayed for all the scores | BR BB 4.2 WIP 104.2 | WIP17 |  | Subsystem | Verified1 | NA | Fixed | redrouthu_r | 0 | 2 | Console USP_EBW | iCAP | YES | Closed | WIP14SegReq | Brilliance 64 | PR#5940478 PR#6123359 PR#6393965 | WIP17 | June 13, 2019 1:35:55 PM GMT+08:00
CTUDT00598583 | Refresh | Closed | rodan_d | June 5, 2018 9:49:53 PM GMT+08:00 | Report - Send findings from CS to report using Mass Score protocol , in the impression section 2nd paragraph the comments are displayed for all the scores | BR BB 4.2 WIP 104.2 | WIP17 |  | Subsystem | Verified1 | NA | Fixed | redrouthu_r | 0 | 2 | Console USP_EBW | iCAP | YES | Closed | WIP14SegReq | Brilliance Big Bore | PR#5940478 PR#6123359 PR#6393965 | WIP17 | June 13, 2019 1:35:55 PM GMT+08:00
CTUDT00598583 | Refresh | Closed | rodan_d | June 5, 2018 9:49:53 PM GMT+08:00 | Report - Send findings from CS to report using Mass Score protocol , in the impression section 2nd paragraph the comments are displayed for all the scores | BR BB 4.2 WIP 104.2 | WIP17 |  | Subsystem | Verified1 | NA | Fixed | redrouthu_r | 0 | 2 | Console USP_EBW | iCAP | YES | Closed | WIP14SegReq | Ingenuity 64 | PR#5940478 PR#6123359 PR#6393965 | WIP17 | June 13, 2019 1:35:55 PM GMT+08:00
CTUDT00598583 | Refresh | Closed | rodan_d | June 5, 2018 9:49:53 PM GMT+08:00 | Report - Send findings from CS to report using Mass Score protocol , in the impression section 2nd paragraph the comments are displayed for all the scores | BR BB 4.2 WIP 104.2 | WIP17 |  | Subsystem | Verified1 | NA | Fixed | redrouthu_r | 0 | 2 | Console USP_EBW | iCAP | YES | Closed | WIP14SegReq | IQon1 | PR#5940478 PR#6123359 PR#6393965 | WIP17 | June 13, 2019 1:35:55 PM GMT+08:00
CTUDT00598582 | Refresh | Closed | rodan_d | June 5, 2018 9:48:19 PM GMT+08:00 | IV: Ingenuity Techncal Reference Guide (Pg. 51) | 4.1.3.XX008-WIP310 |  |  | Subsystem | Documentation |  | Change Proposal | rajalakshmi_j | 0 | 5 | IFU | Support-System | NO | Closed |  | Brilliance 64
CTUDT00598582 | Refresh | Closed | rodan_d | June 5, 2018 9:48:19 PM GMT+08:00 | IV: Ingenuity Techncal Reference Guide (Pg. 51) | 4.1.3.XX008-WIP310 |  |  | Subsystem | Documentation |  | Change Proposal | rajalakshmi_j | 0 | 5 | IFU | Support-System | NO | Closed |  | Ingenuity 64
CTUDT00598581 | Refresh | Closed | rodan_d | June 5, 2018 9:46:42 PM GMT+08:00 | IV: Ingenuity Technical Reference Guide (Pg. 51) | 4.1.3.XX008-WIP310 | WIP16 |  | Subsystem | Documentation | N/A | Fixed | rajalakshmi_j | 0 | 3 | User Documentation | Support-System | NO | Closed | WIP9.2SubCom | Ovation |  | WIP16 | May 29, 2019 2:04:05 PM GMT+08:00
CTUDT00598581 | Refresh | Closed | rodan_d | June 5, 2018 9:46:42 PM GMT+08:00 | IV: Ingenuity Technical Reference Guide (Pg. 51) | 4.1.3.XX008-WIP310 | WIP16 |  | Subsystem | Documentation | N/A | Fixed | rajalakshmi_j | 0 | 3 | User Documentation | Support-System | NO | Closed | WIP9.2SubCom | Brilliance 64 |  | WIP16 | May 29, 2019 2:04:05 PM GMT+08:00
CTUDT00598581 | Refresh | Closed | rodan_d | June 5, 2018 9:46:42 PM GMT+08:00 | IV: Ingenuity Technical Reference Guide (Pg. 51) | 4.1.3.XX008-WIP310 | WIP16 |  | Subsystem | Documentation | N/A | Fixed | rajalakshmi_j | 0 | 3 | User Documentation | Support-System | NO | Closed | WIP9.2SubCom | Brilliance Big Bore |  | WIP16 | May 29, 2019 2:04:05 PM GMT+08:00
CTUDT00598581 | Refresh | Closed | rodan_d | June 5, 2018 9:46:42 PM GMT+08:00 | IV: Ingenuity Technical Reference Guide (Pg. 51) | 4.1.3.XX008-WIP310 | WIP16 |  | Subsystem | Documentation | N/A | Fixed | rajalakshmi_j | 0 | 3 | User Documentation | Support-System | NO | Closed | WIP9.2SubCom | Ingenuity 64 |  | WIP16 | May 29, 2019 2:04:05 PM GMT+08:00
CTUDT00598581 | Refresh | Closed | rodan_d | June 5, 2018 9:46:42 PM GMT+08:00 | IV: Ingenuity Technical Reference Guide (Pg. 51) | 4.1.3.XX008-WIP310 | WIP16 |  | Subsystem | Documentation | N/A | Fixed | rajalakshmi_j | 0 | 3 | User Documentation | Support-System | NO | Closed | WIP9.2SubCom | IQon1 |  | WIP16 | May 29, 2019 2:04:05 PM GMT+08:00
CTUDT00598580 | Refresh | Closed | rodan_d | June 5, 2018 9:45:19 PM GMT+08:00 | iPatient BB- Pulmo w RPM: Export wave message: When User clicks the ‘Continue without the wave’ dialogue, they have no way of recovering from this error | 4.2.0.xx020-WIP104 |  |  | Subsystem | History Review |  | Change Proposal | sanjeeva_a | 0 | 5 |  | Console4 | NO | Closed | BB_SubNC | Brilliance Big Bore
CTUDT00598579 | Refresh | Closed | rodan_d | June 5, 2018 9:44:03 PM GMT+08:00 | iPatient BB: Pulmo:Edit phases dialog: slow performance when pressing Delete All and Apply | 4.2.0.xx020-WIP104 |  |  | Subsystem |  | None | Change Proposal | sanjeeva_a | 0 | 4 |  | Console4 | NO | Closed | BB_SubNC | Brilliance Big Bore
CTUDT00598578 | Refresh | Closed | rodan_d | June 5, 2018 9:42:23 PM GMT+08:00 | Cycle time for Jog Perfusion - please refer to CTUDT00568295 - RFC postponed to Ingenuity BB | 4.2.0.xx030-WIP104.3 |  |  | Subsystem | History Review |  | Change Proposal | sanjeeva_a | 0 | 5 |  | Console4 | NO | Closed | BB_SubNC | Brilliance Big Bore
CTUDT00598577 | Refresh | Closed | rodan_d | June 5, 2018 9:40:59 PM GMT+08:00 | print history of film cannot sort by print date | 4.2.0.xx030-WIP104.3 |  |  | Subsystem | Film | NA | No Plan to Fix | veerabhadraiah | 0 | 4 |  | iCAP | NO | Closed |  | Ovation
CTUDT00598577 | Refresh | Closed | rodan_d | June 5, 2018 9:40:59 PM GMT+08:00 | print history of film cannot sort by print date | 4.2.0.xx030-WIP104.3 |  |  | Subsystem | Film | NA | No Plan to Fix | veerabhadraiah | 0 | 4 |  | iCAP | NO | Closed |  | Brilliance 64
CTUDT00598577 | Refresh | Closed | rodan_d | June 5, 2018 9:40:59 PM GMT+08:00 | print history of film cannot sort by print date | 4.2.0.xx030-WIP104.3 |  |  | Subsystem | Film | NA | No Plan to Fix | veerabhadraiah | 0 | 4 |  | iCAP | NO | Closed |  | Brilliance Big Bore
CTUDT00598577 | Refresh | Closed | rodan_d | June 5, 2018 9:40:59 PM GMT+08:00 | print history of film cannot sort by print date | 4.2.0.xx030-WIP104.3 |  |  | Subsystem | Film | NA | No Plan to Fix | veerabhadraiah | 0 | 4 |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00598577 | Refresh | Closed | rodan_d | June 5, 2018 9:40:59 PM GMT+08:00 | print history of film cannot sort by print date | 4.2.0.xx030-WIP104.3 |  |  | Subsystem | Film | NA | No Plan to Fix | veerabhadraiah | 0 | 4 |  | iCAP | NO | Closed |  | IQon1
CTUDT00598576 | Refresh | Closed | rodan_d | June 5, 2018 9:39:32 PM GMT+08:00 | Pulmo Viewer - Slab scene - Orthogonal Planes selection drop down | 4.2.0.xx0xx-WIP107 |  |  | Subsystem |  |  | Change Proposal | sanjeeva_a | 0 |  |  | Console4 | NO | Closed | BB_SubNC | Brilliance Big Bore
CTUDT00598575 | Refresh | Closed | rodan_d | June 5, 2018 9:37:53 PM GMT+08:00 | BB42 FOK: RFC: No cursor on axial images | Post-RfLD-4.2.0.xx057-WIP111 |  |  | Subsystem |  |  | Change Proposal | sanjeeva_a | 0 | 3 |  | Support-System | NO | Closed |  | Brilliance Big Bore | CFR 9406807 PR5848223
CTUDT00598573 | Refresh | Closed | rodan_d | June 5, 2018 9:36:24 PM GMT+08:00 | FOK BB V4.2- Very difficult to foot pedal to make work correctly | Post-RfLD-4.2.0.xx057-WIP111 |  |  | Subsystem |  |  | Change Proposal | cumpson | 0 |  |  | GCX | NO | Closed |  | Brilliance Big Bore | CFR9407823 PR5876956
CTUDT00598571 | Refresh | Closed | rodan_d | June 5, 2018 9:34:41 PM GMT+08:00 | FOK BB V4.2- mAs in the ECM | Post-RfLD-4.2.0.xx057-WIP111 |  |  | Subsystem | History Review |  | Change Proposal | sanjeeva_a | 0 | 5 |  | Console4 | NO | Closed | BB_SubNC | Brilliance Big Bore
CTUDT00598570 | Refresh | Closed | rodan_d | June 5, 2018 9:32:56 PM GMT+08:00 | BB42 FOK RFC: Customer requests that the option of a full-length cursor be added to the graphic properties | Post-RfLD-4.2.0.xx057-WIP111 |  |  | Subsystem | History Review |  | Change Proposal | sanjeeva_a | 0 | 5 |  | Console4 | NO | Closed | BB_SubNC | Brilliance Big Bore | CFR 9407613 PR5876350 CFR 9407616 PR5876306
CTUDT00598569 | Refresh | Closed | bagdady_b | June 5, 2018 9:32:45 PM GMT+08:00 | FOK: When LastScannedER file is corrupted the exam application crashes | WIP7.1 | WIP17 |  | Subsystem | Plan | NA | Fixed | agrawal_ay | 0 | 3 | Console Software | Console4 | NO | Closed | WIP17SubCom | Ovation | SWO 47764096 | WIP17 | May 22, 2019 4:54:16 PM GMT+08:00
CTUDT00598569 | Refresh | Closed | bagdady_b | June 5, 2018 9:32:45 PM GMT+08:00 | FOK: When LastScannedER file is corrupted the exam application crashes | WIP7.1 | WIP17 |  | Subsystem | Plan | NA | Fixed | agrawal_ay | 0 | 3 | Console Software | Console4 | NO | Closed | WIP17SubCom | Brilliance 64 | SWO 47764096 | WIP17 | May 22, 2019 4:54:16 PM GMT+08:00
CTUDT00598569 | Refresh | Closed | bagdady_b | June 5, 2018 9:32:45 PM GMT+08:00 | FOK: When LastScannedER file is corrupted the exam application crashes | WIP7.1 | WIP17 |  | Subsystem | Plan | NA | Fixed | agrawal_ay | 0 | 3 | Console Software | Console4 | NO | Closed | WIP17SubCom | Ingenuity 64 | SWO 47764096 | WIP17 | May 22, 2019 4:54:16 PM GMT+08:00
CTUDT00598569 | Refresh | Closed | bagdady_b | June 5, 2018 9:32:45 PM GMT+08:00 | FOK: When LastScannedER file is corrupted the exam application crashes | WIP7.1 | WIP17 |  | Subsystem | Plan | NA | Fixed | agrawal_ay | 0 | 3 | Console Software | Console4 | NO | Closed | WIP17SubCom | IQon1 | SWO 47764096 | WIP17 | May 22, 2019 4:54:16 PM GMT+08:00
CTUDT00598568 | Refresh | Closed | rodan_d | June 5, 2018 9:31:11 PM GMT+08:00 | BB42 FOK RFC: user would like to have 4DCT pitch and breathing rate prompts | Post-RfLD-4.2.0.xx057-WIP111 |  |  | Subsystem |  |  | Change Proposal | sanjeeva_a | 0 |  |  | Console4 | NO | Closed | BB_SubNC | Brilliance Big Bore | CFR9410313 PR5929750
CTUDT00598566 | Refresh | Closed | rodan_d | June 5, 2018 9:28:16 PM GMT+08:00 | BB42 FOK: User has to scroll pulmo wave to get to view the irradiated portion prior to doing Start Final Recon. | Post-RfLD-4.2.0.xx057-WIP111 |  |  | Subsystem |  |  | Change Proposal | sanjeeva_a | 0 |  |  | Console4 | NO | Closed | BB_SubNC | Brilliance Big Bore | CFR9414841
CTUDT00598565 | Refresh | Closed | rodan_d | June 5, 2018 9:26:56 PM GMT+08:00 | Resp Breath Hold Scan: Incorrect message displayed to the user. | RfLD-4.2.0.xx057-WIP111 |  | No | Subsystem |  | Console-SRS-13501 | No Plan to Fix | hatti_c | 0 | 4 |  | Console4 | NO | Closed | SubNC | Brilliance 64
CTUDT00598565 | Refresh | Closed | rodan_d | June 5, 2018 9:26:56 PM GMT+08:00 | Resp Breath Hold Scan: Incorrect message displayed to the user. | RfLD-4.2.0.xx057-WIP111 |  | No | Subsystem |  | Console-SRS-13501 | No Plan to Fix | hatti_c | 0 | 4 |  | Console4 | NO | Closed | SubNC | Brilliance Big Bore
CTUDT00598565 | Refresh | Closed | rodan_d | June 5, 2018 9:26:56 PM GMT+08:00 | Resp Breath Hold Scan: Incorrect message displayed to the user. | RfLD-4.2.0.xx057-WIP111 |  | No | Subsystem |  | Console-SRS-13501 | No Plan to Fix | hatti_c | 0 | 4 |  | Console4 | NO | Closed | SubNC | Ingenuity 64
CTUDT00598564 | Refresh | Closed | rodan_d | June 5, 2018 9:25:09 PM GMT+08:00 | RFC against IFU for all CT products: We need specific warnings regarding defibrillation equipment use while patient is on the scanner. | _Current Production HW_ | WIP16 |  | Subsystem | Documentation | NA | Fixed | rajalakshmi_j | 0 | 3 | User Documentation | GCX | NO | Closed | WIP10SubCom | Ovation |  | WIP16 | May 29, 2019 2:03:36 PM GMT+08:00
CTUDT00598564 | Refresh | Closed | rodan_d | June 5, 2018 9:25:09 PM GMT+08:00 | RFC against IFU for all CT products: We need specific warnings regarding defibrillation equipment use while patient is on the scanner. | _Current Production HW_ | WIP16 |  | Subsystem | Documentation | NA | Fixed | rajalakshmi_j | 0 | 3 | User Documentation | GCX | NO | Closed | WIP10SubCom | Brilliance 64 |  | WIP16 | May 29, 2019 2:03:36 PM GMT+08:00
CTUDT00598564 | Refresh | Closed | rodan_d | June 5, 2018 9:25:09 PM GMT+08:00 | RFC against IFU for all CT products: We need specific warnings regarding defibrillation equipment use while patient is on the scanner. | _Current Production HW_ | WIP16 |  | Subsystem | Documentation | NA | Fixed | rajalakshmi_j | 0 | 3 | User Documentation | GCX | NO | Closed | WIP10SubCom | Brilliance Big Bore |  | WIP16 | May 29, 2019 2:03:36 PM GMT+08:00
CTUDT00598564 | Refresh | Closed | rodan_d | June 5, 2018 9:25:09 PM GMT+08:00 | RFC against IFU for all CT products: We need specific warnings regarding defibrillation equipment use while patient is on the scanner. | _Current Production HW_ | WIP16 |  | Subsystem | Documentation | NA | Fixed | rajalakshmi_j | 0 | 3 | User Documentation | GCX | NO | Closed | WIP10SubCom | Ingenuity 64 |  | WIP16 | May 29, 2019 2:03:36 PM GMT+08:00
CTUDT00598564 | Refresh | Closed | rodan_d | June 5, 2018 9:25:09 PM GMT+08:00 | RFC against IFU for all CT products: We need specific warnings regarding defibrillation equipment use while patient is on the scanner. | _Current Production HW_ | WIP16 |  | Subsystem | Documentation | NA | Fixed | rajalakshmi_j | 0 | 3 | User Documentation | GCX | NO | Closed | WIP10SubCom | IQon1 |  | WIP16 | May 29, 2019 2:03:36 PM GMT+08:00
CTUDT00598563 | Refresh | Closed | rodan_d | June 5, 2018 9:23:19 PM GMT+08:00 | Calibrations: Phantom Calibration graphs has misleading Title when viewed from "View EP" (EP out of limit) | 4.1.4.XX001 |  |  | Subsystem |  | N/A | Works As Specified / Not A Defect | gehtin_e | 0 |  | DMS Calibrations | DMS | NO | Closed |  | Ovation
CTUDT00598563 | Refresh | Closed | rodan_d | June 5, 2018 9:23:19 PM GMT+08:00 | Calibrations: Phantom Calibration graphs has misleading Title when viewed from "View EP" (EP out of limit) | 4.1.4.XX001 |  |  | Subsystem |  | N/A | Works As Specified / Not A Defect | gehtin_e | 0 |  | DMS Calibrations | DMS | NO | Closed |  | Brilliance 64
CTUDT00598563 | Refresh | Closed | rodan_d | June 5, 2018 9:23:19 PM GMT+08:00 | Calibrations: Phantom Calibration graphs has misleading Title when viewed from "View EP" (EP out of limit) | 4.1.4.XX001 |  |  | Subsystem |  | N/A | Works As Specified / Not A Defect | gehtin_e | 0 |  | DMS Calibrations | DMS | NO | Closed |  | Ingenuity 64
CTUDT00598562 | Refresh | Closed | rodan_d | June 5, 2018 9:17:52 PM GMT+08:00 | Pre-IV: Ingenuity - View 2 crashed after selecting End Exam and continue working | 4.1.6.xx021 WIP 105.1 |  |  | Subsystem |  | None | No Plan to Fix | sanjeeva_a | 0 | 4 |  | Console4 | NO | Closed |  | Ovation
CTUDT00598562 | Refresh | Closed | rodan_d | June 5, 2018 9:17:52 PM GMT+08:00 | Pre-IV: Ingenuity - View 2 crashed after selecting End Exam and continue working | 4.1.6.xx021 WIP 105.1 |  |  | Subsystem |  | None | No Plan to Fix | sanjeeva_a | 0 | 4 |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00598562 | Refresh | Closed | rodan_d | June 5, 2018 9:17:52 PM GMT+08:00 | Pre-IV: Ingenuity - View 2 crashed after selecting End Exam and continue working | 4.1.6.xx021 WIP 105.1 |  |  | Subsystem |  | None | No Plan to Fix | sanjeeva_a | 0 | 4 |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00598562 | Refresh | Closed | rodan_d | June 5, 2018 9:17:52 PM GMT+08:00 | Pre-IV: Ingenuity - View 2 crashed after selecting End Exam and continue working | 4.1.6.xx021 WIP 105.1 |  |  | Subsystem |  | None | No Plan to Fix | sanjeeva_a | 0 | 4 |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00598559 | Refresh | Closed | rodan_d | June 5, 2018 9:09:58 PM GMT+08:00 | Request Beam Quality test instructions for Ingenuity CT / Ingenuity Core 128 / Ingenuity Core | 4.1.3.XX022 |  |  | Subsystem |  |  | Change Proposal | hubenschmidt_j | 0 | 3 |  | Support-Physics | NO | Closed |  | Ingenuity 64
CTUDT00598558 | Refresh | Closed | rodan_d | June 5, 2018 9:04:34 PM GMT+08:00 | Image Quality and Dose Testing Instructions for V 4.1.6 O-level is missing CT Linearity test instructions. | 4.1.3.XX022 |  | The workaround is to send a Service Engineer to the site each month and execute Automatic Performance, which contains the test. However, that test is not accessible to CT User or to O-level Service. | Subsystem | History Review |  | Does Not Apply | yao_j | 0 | 5 | Documentation - Service | Support-Service | NO | Closed |  | Ovation
CTUDT00598558 | Refresh | Closed | rodan_d | June 5, 2018 9:04:34 PM GMT+08:00 | Image Quality and Dose Testing Instructions for V 4.1.6 O-level is missing CT Linearity test instructions. | 4.1.3.XX022 |  | The workaround is to send a Service Engineer to the site each month and execute Automatic Performance, which contains the test. However, that test is not accessible to CT User or to O-level Service. | Subsystem | History Review |  | Does Not Apply | yao_j | 0 | 5 | Documentation - Service | Support-Service | NO | Closed |  | Brilliance 64
CTUDT00598558 | Refresh | Closed | rodan_d | June 5, 2018 9:04:34 PM GMT+08:00 | Image Quality and Dose Testing Instructions for V 4.1.6 O-level is missing CT Linearity test instructions. | 4.1.3.XX022 |  | The workaround is to send a Service Engineer to the site each month and execute Automatic Performance, which contains the test. However, that test is not accessible to CT User or to O-level Service. | Subsystem | History Review |  | Does Not Apply | yao_j | 0 | 5 | Documentation - Service | Support-Service | NO | Closed |  | Ingenuity 64
CTUDT00598557 | Refresh | Closed | rodan_d | June 5, 2018 9:01:46 PM GMT+08:00 | Val: 4.1.7 - the system did not save the correct parameters in ECM for user Exam Card HIGH RES INS/EXP after upgrade/Install. | V4.1.7-WIP-5.1 |  | User can verify and edit the parameters used for a scan prior to acquisition of the scan. | Subsystem |  | NA | Does Not Apply | goswami_s | 0 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00598557 | Refresh | Closed | rodan_d | June 5, 2018 9:01:46 PM GMT+08:00 | Val: 4.1.7 - the system did not save the correct parameters in ECM for user Exam Card HIGH RES INS/EXP after upgrade/Install. | V4.1.7-WIP-5.1 |  | User can verify and edit the parameters used for a scan prior to acquisition of the scan. | Subsystem |  | NA | Does Not Apply | goswami_s | 0 |  |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00598557 | Refresh | Closed | rodan_d | June 5, 2018 9:01:46 PM GMT+08:00 | Val: 4.1.7 - the system did not save the correct parameters in ECM for user Exam Card HIGH RES INS/EXP after upgrade/Install. | V4.1.7-WIP-5.1 |  | User can verify and edit the parameters used for a scan prior to acquisition of the scan. | Subsystem |  | NA | Does Not Apply | goswami_s | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00598555 | Refresh | Closed | rodan_d | June 5, 2018 8:56:58 PM GMT+08:00 | Val: 4.1.7 - The system Duplicated 2 results without being prompted to do so | V4.1.7-WIP-5.1 |  | User can delete the duplicate series manually. | Subsystem |  |  | Change Proposal | sanjeeva_a | 0 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00598554 | Refresh | Closed | rodan_d | June 5, 2018 8:46:42 PM GMT+08:00 | Val: 4.1.7 - The cardiac images created with Slab Viewer will not display ECG Wave and only images in 2D are supported. | V4.1.7 | WIP11.1 | create the coronary result using the surview to get original images. This allows you the full capability in Cardiac Viewer including the ECG Wave display | Subsystem | Cardiac | NA | Fixed | veerabhadraiah | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP14SegReq | Ovation |  | WIP11.1 | February 25, 2019 4:31:14 PM GMT+08:00
CTUDT00598554 | Refresh | Closed | rodan_d | June 5, 2018 8:46:42 PM GMT+08:00 | Val: 4.1.7 - The cardiac images created with Slab Viewer will not display ECG Wave and only images in 2D are supported. | V4.1.7 | WIP11.1 | create the coronary result using the surview to get original images. This allows you the full capability in Cardiac Viewer including the ECG Wave display | Subsystem | Cardiac | NA | Fixed | veerabhadraiah | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP14SegReq | Brilliance 64 |  | WIP11.1 | February 25, 2019 4:31:14 PM GMT+08:00
CTUDT00598554 | Refresh | Closed | rodan_d | June 5, 2018 8:46:42 PM GMT+08:00 | Val: 4.1.7 - The cardiac images created with Slab Viewer will not display ECG Wave and only images in 2D are supported. | V4.1.7 | WIP11.1 | create the coronary result using the surview to get original images. This allows you the full capability in Cardiac Viewer including the ECG Wave display | Subsystem | Cardiac | NA | Fixed | veerabhadraiah | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP14SegReq | Brilliance Big Bore |  | WIP11.1 | February 25, 2019 4:31:14 PM GMT+08:00
CTUDT00598554 | Refresh | Closed | rodan_d | June 5, 2018 8:46:42 PM GMT+08:00 | Val: 4.1.7 - The cardiac images created with Slab Viewer will not display ECG Wave and only images in 2D are supported. | V4.1.7 | WIP11.1 | create the coronary result using the surview to get original images. This allows you the full capability in Cardiac Viewer including the ECG Wave display | Subsystem | Cardiac | NA | Fixed | veerabhadraiah | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP14SegReq | Ingenuity 64 |  | WIP11.1 | February 25, 2019 4:31:14 PM GMT+08:00
CTUDT00598554 | Refresh | Closed | rodan_d | June 5, 2018 8:46:42 PM GMT+08:00 | Val: 4.1.7 - The cardiac images created with Slab Viewer will not display ECG Wave and only images in 2D are supported. | V4.1.7 | WIP11.1 | create the coronary result using the surview to get original images. This allows you the full capability in Cardiac Viewer including the ECG Wave display | Subsystem | Cardiac | NA | Fixed | veerabhadraiah | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP14SegReq | IQon1 |  | WIP11.1 | February 25, 2019 4:31:14 PM GMT+08:00
CTUDT00598553 | Refresh | Closed | rodan_d | June 5, 2018 8:33:29 PM GMT+08:00 | VAL WIP209 BB42: Scan aborted with System Error | 4.2.0.17xxx-WIP209 |  |  | Subsystem |  |  | Change Proposal | sanjeeva_a | 0 | 2 |  | Console4 | YES | Closed |  | Ovation | none
CTUDT00598553 | Refresh | Closed | rodan_d | June 5, 2018 8:33:29 PM GMT+08:00 | VAL WIP209 BB42: Scan aborted with System Error | 4.2.0.17xxx-WIP209 |  |  | Subsystem |  |  | Change Proposal | sanjeeva_a | 0 | 2 |  | Console4 | YES | Closed |  | Brilliance 64 | none
CTUDT00598553 | Refresh | Closed | rodan_d | June 5, 2018 8:33:29 PM GMT+08:00 | VAL WIP209 BB42: Scan aborted with System Error | 4.2.0.17xxx-WIP209 |  |  | Subsystem |  |  | Change Proposal | sanjeeva_a | 0 | 2 |  | Console4 | YES | Closed |  | Brilliance Big Bore | none
CTUDT00598553 | Refresh | Closed | rodan_d | June 5, 2018 8:33:29 PM GMT+08:00 | VAL WIP209 BB42: Scan aborted with System Error | 4.2.0.17xxx-WIP209 |  |  | Subsystem |  |  | Change Proposal | sanjeeva_a | 0 | 2 |  | Console4 | YES | Closed |  | Ingenuity 64 | none
CTUDT00598553 | Refresh | Closed | rodan_d | June 5, 2018 8:33:29 PM GMT+08:00 | VAL WIP209 BB42: Scan aborted with System Error | 4.2.0.17xxx-WIP209 |  |  | Subsystem |  |  | Change Proposal | sanjeeva_a | 0 | 2 |  | Console4 | YES | Closed |  | IQon1 | none
CTUDT00598552 | Refresh | Closed | rodan_d | June 5, 2018 8:27:53 PM GMT+08:00 | VAL WIP209 BB42: Quality Assurance dialogue box on system shows "N/A" for time recommended next | 4.2.0.17xxx-WIP209 |  |  | Subsystem |  |  | Change Proposal | oren_d | 0 | 3 |  | Support-System | NO | Closed |  | Brilliance Big Bore | none
CTUDT00598551 | Refresh | Closed | rodan_d | June 5, 2018 7:29:51 PM GMT+08:00 | CCT - Save icon in CCT Vol display scene | 4.1.6 | WIP17 | NA | Subsystem |  | NA | Fixed | vishnu_devi | 0 | 2 | Console Generic | Console4 | NO | Closed | WIP16SubCom | Ovation |  | WIP17 | May 8, 2019 9:47:35 PM GMT+08:00
CTUDT00598551 | Refresh | Closed | rodan_d | June 5, 2018 7:29:51 PM GMT+08:00 | CCT - Save icon in CCT Vol display scene | 4.1.6 | WIP17 | NA | Subsystem |  | NA | Fixed | vishnu_devi | 0 | 2 | Console Generic | Console4 | NO | Closed | WIP16SubCom | Brilliance 64 |  | WIP17 | May 8, 2019 9:47:35 PM GMT+08:00
CTUDT00598551 | Refresh | Closed | rodan_d | June 5, 2018 7:29:51 PM GMT+08:00 | CCT - Save icon in CCT Vol display scene | 4.1.6 | WIP17 | NA | Subsystem |  | NA | Fixed | vishnu_devi | 0 | 2 | Console Generic | Console4 | NO | Closed | WIP16SubCom | Ingenuity 64 |  | WIP17 | May 8, 2019 9:47:35 PM GMT+08:00
CTUDT00598551 | Refresh | Closed | rodan_d | June 5, 2018 7:29:51 PM GMT+08:00 | CCT - Save icon in CCT Vol display scene | 4.1.6 | WIP17 | NA | Subsystem |  | NA | Fixed | vishnu_devi | 0 | 2 | Console Generic | Console4 | NO | Closed | WIP16SubCom | IQon1 |  | WIP17 | May 8, 2019 9:47:35 PM GMT+08:00
CTUDT00598550 | Refresh | Closed | rodan_d | June 5, 2018 7:24:08 PM GMT+08:00 | FoK: Aarhus: GantryWCFSelfHostedServices has stopped working message during CT user login after restart | Eagle_WIP113.2FoK | WIP13 | Logout and login again to the HOST console. | System |  | NA | Fixed | karkour | 0 | 3 | Gantry Generic | GCX | NO | Closed | SubCom | IQon1 | 9407068 9401541 9405147 | WIP13 | May 7, 2019 11:48:51 PM GMT+08:00
CTUDT00598549 | Refresh | Closed | rodan_d | June 5, 2018 7:22:29 PM GMT+08:00 | a message asking if bugrep needed to be generated popped up right after the helical scan finished. | 4.1.1 | WIP13 | The user should delete the Surview step of Chest-Pelvis (series 3) before performing the Surview (series 1) | System |  | NA | Fixed | Snehashis_Dey | 0 | 3 | Console Software | Console4 | YES | Closed | MergeFrom4.x | Ovation | PR4111818 9330167 | WIP13 | February 4, 2019 4:07:35 PM GMT+08:00
CTUDT00598549 | Refresh | Closed | rodan_d | June 5, 2018 7:22:29 PM GMT+08:00 | a message asking if bugrep needed to be generated popped up right after the helical scan finished. | 4.1.1 | WIP13 | The user should delete the Surview step of Chest-Pelvis (series 3) before performing the Surview (series 1) | System |  | NA | Fixed | Snehashis_Dey | 0 | 3 | Console Software | Console4 | YES | Closed | MergeFrom4.x | Brilliance 64 | PR4111818 9330167 | WIP13 | February 4, 2019 4:07:35 PM GMT+08:00
CTUDT00598549 | Refresh | Closed | rodan_d | June 5, 2018 7:22:29 PM GMT+08:00 | a message asking if bugrep needed to be generated popped up right after the helical scan finished. | 4.1.1 | WIP13 | The user should delete the Surview step of Chest-Pelvis (series 3) before performing the Surview (series 1) | System |  | NA | Fixed | Snehashis_Dey | 0 | 3 | Console Software | Console4 | YES | Closed | MergeFrom4.x | Brilliance Big Bore | PR4111818 9330167 | WIP13 | February 4, 2019 4:07:35 PM GMT+08:00
CTUDT00598549 | Refresh | Closed | rodan_d | June 5, 2018 7:22:29 PM GMT+08:00 | a message asking if bugrep needed to be generated popped up right after the helical scan finished. | 4.1.1 | WIP13 | The user should delete the Surview step of Chest-Pelvis (series 3) before performing the Surview (series 1) | System |  | NA | Fixed | Snehashis_Dey | 0 | 3 | Console Software | Console4 | YES | Closed | MergeFrom4.x | Ingenuity 64 | PR4111818 9330167 | WIP13 | February 4, 2019 4:07:35 PM GMT+08:00
CTUDT00598549 | Refresh | Closed | rodan_d | June 5, 2018 7:22:29 PM GMT+08:00 | a message asking if bugrep needed to be generated popped up right after the helical scan finished. | 4.1.1 | WIP13 | The user should delete the Surview step of Chest-Pelvis (series 3) before performing the Surview (series 1) | System |  | NA | Fixed | Snehashis_Dey | 0 | 3 | Console Software | Console4 | YES | Closed | MergeFrom4.x | IQon1 | PR4111818 9330167 | WIP13 | February 4, 2019 4:07:35 PM GMT+08:00
CTUDT00598548 | Refresh | Closed | rodan_d | June 5, 2018 7:20:44 PM GMT+08:00 | Cardiac scan failed due to S_MPL_SURR_SERIAL_OUT_OF_SYNC | _OVATION_3.2.4 | WIP16 |  | Subsystem |  | NA | Fixed | karkour | 0 | 2 | Gantry Software | GCX | YES | Closed | WIP16SubCom | Ovation | PR5505937 PR6265660 | WIP16 | June 25, 2019 2:51:52 AM GMT+08:00
CTUDT00598548 | Refresh | Closed | rodan_d | June 5, 2018 7:20:44 PM GMT+08:00 | Cardiac scan failed due to S_MPL_SURR_SERIAL_OUT_OF_SYNC | _OVATION_3.2.4 | WIP16 |  | Subsystem |  | NA | Fixed | karkour | 0 | 2 | Gantry Software | GCX | YES | Closed | WIP16SubCom | IQon1 | PR5505937 PR6265660 | WIP16 | June 25, 2019 2:51:52 AM GMT+08:00
CTUDT00598547 | Refresh | Closed | rodan_d | June 5, 2018 7:19:12 PM GMT+08:00 | Cardiac\User Messages: The message suggesting to reconsider using Cardiac DoseRight when arrhythmia was detected doesn't always appear | 4.1.0.XX250-WIP211.3 | WIP13 | NA. This works as specified. | Subsystem |  | NA | Fixed | ravindra_a | 0 | 3 | Console Generic | Console4 | NO | Closed | WIP13SubCom | Ovation |  | WIP13 | January 23, 2019 6:53:40 PM GMT+08:00
CTUDT00598547 | Refresh | Closed | rodan_d | June 5, 2018 7:19:12 PM GMT+08:00 | Cardiac\User Messages: The message suggesting to reconsider using Cardiac DoseRight when arrhythmia was detected doesn't always appear | 4.1.0.XX250-WIP211.3 | WIP13 | NA. This works as specified. | Subsystem |  | NA | Fixed | ravindra_a | 0 | 3 | Console Generic | Console4 | NO | Closed | WIP13SubCom | Brilliance 64 |  | WIP13 | January 23, 2019 6:53:40 PM GMT+08:00
CTUDT00598547 | Refresh | Closed | rodan_d | June 5, 2018 7:19:12 PM GMT+08:00 | Cardiac\User Messages: The message suggesting to reconsider using Cardiac DoseRight when arrhythmia was detected doesn't always appear | 4.1.0.XX250-WIP211.3 | WIP13 | NA. This works as specified. | Subsystem |  | NA | Fixed | ravindra_a | 0 | 3 | Console Generic | Console4 | NO | Closed | WIP13SubCom | Brilliance Big Bore |  | WIP13 | January 23, 2019 6:53:40 PM GMT+08:00
CTUDT00598547 | Refresh | Closed | rodan_d | June 5, 2018 7:19:12 PM GMT+08:00 | Cardiac\User Messages: The message suggesting to reconsider using Cardiac DoseRight when arrhythmia was detected doesn't always appear | 4.1.0.XX250-WIP211.3 | WIP13 | NA. This works as specified. | Subsystem |  | NA | Fixed | ravindra_a | 0 | 3 | Console Generic | Console4 | NO | Closed | WIP13SubCom | Ingenuity 64 |  | WIP13 | January 23, 2019 6:53:40 PM GMT+08:00
CTUDT00598547 | Refresh | Closed | rodan_d | June 5, 2018 7:19:12 PM GMT+08:00 | Cardiac\User Messages: The message suggesting to reconsider using Cardiac DoseRight when arrhythmia was detected doesn't always appear | 4.1.0.XX250-WIP211.3 | WIP13 | NA. This works as specified. | Subsystem |  | NA | Fixed | ravindra_a | 0 | 3 | Console Generic | Console4 | NO | Closed | WIP13SubCom | IQon1 |  | WIP13 | January 23, 2019 6:53:40 PM GMT+08:00
CTUDT00598546 | Refresh | Closed | rodan_d | June 5, 2018 7:15:31 PM GMT+08:00 | ECG data was not saved with the cardiac scan series. | V4.1.6 |  |  | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 3 |  | Console4 | YES | Closed |  | Ovation | 9442043
CTUDT00598546 | Refresh | Closed | rodan_d | June 5, 2018 7:15:31 PM GMT+08:00 | ECG data was not saved with the cardiac scan series. | V4.1.6 |  |  | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 3 |  | Console4 | YES | Closed |  | Brilliance 64 | 9442043
CTUDT00598546 | Refresh | Closed | rodan_d | June 5, 2018 7:15:31 PM GMT+08:00 | ECG data was not saved with the cardiac scan series. | V4.1.6 |  |  | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 3 |  | Console4 | YES | Closed |  | Brilliance Big Bore | 9442043
CTUDT00598546 | Refresh | Closed | rodan_d | June 5, 2018 7:15:31 PM GMT+08:00 | ECG data was not saved with the cardiac scan series. | V4.1.6 |  |  | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 3 |  | Console4 | YES | Closed |  | Ingenuity 64 | 9442043
CTUDT00598546 | Refresh | Closed | rodan_d | June 5, 2018 7:15:31 PM GMT+08:00 | ECG data was not saved with the cardiac scan series. | V4.1.6 |  |  | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 3 |  | Console4 | YES | Closed |  | IQon1 | 9442043
CTUDT00598545 | Refresh | Closed | rodan_d | June 5, 2018 7:11:12 PM GMT+08:00 | GCX.SSRS.iCT.20284 and GCX.SSRS.iCT.20285 | V4.1.7_WIP-6 |  |  | Subsystem |  |  | Does Not Apply | zhou_r | 0 | 3 |  | Couch | NO | Closed |  | Ovation
CTUDT00598545 | Refresh | Closed | rodan_d | June 5, 2018 7:11:12 PM GMT+08:00 | GCX.SSRS.iCT.20284 and GCX.SSRS.iCT.20285 | V4.1.7_WIP-6 |  |  | Subsystem |  |  | Does Not Apply | zhou_r | 0 | 3 |  | Couch | NO | Closed |  | Brilliance 64
CTUDT00598545 | Refresh | Closed | rodan_d | June 5, 2018 7:11:12 PM GMT+08:00 | GCX.SSRS.iCT.20284 and GCX.SSRS.iCT.20285 | V4.1.7_WIP-6 |  |  | Subsystem |  |  | Does Not Apply | zhou_r | 0 | 3 |  | Couch | NO | Closed |  | Brilliance Big Bore
CTUDT00598545 | Refresh | Closed | rodan_d | June 5, 2018 7:11:12 PM GMT+08:00 | GCX.SSRS.iCT.20284 and GCX.SSRS.iCT.20285 | V4.1.7_WIP-6 |  |  | Subsystem |  |  | Does Not Apply | zhou_r | 0 | 3 |  | Couch | NO | Closed |  | Ingenuity 64
CTUDT00598544 | Refresh | Closed | rodan_d | June 5, 2018 6:48:18 PM GMT+08:00 | Incomplete Requirements for Stuck Button Check | _Current Production HW_ |  |  | Subsystem |  | GCX.SSRS.iCT.13955 | Change Proposal | cumpson | 0 | 2 |  | GCX | YES | Closed |  | Ovation
CTUDT00598544 | Refresh | Closed | rodan_d | June 5, 2018 6:48:18 PM GMT+08:00 | Incomplete Requirements for Stuck Button Check | _Current Production HW_ |  |  | Subsystem |  | GCX.SSRS.iCT.13955 | Change Proposal | cumpson | 0 | 2 |  | GCX | YES | Closed |  | Brilliance 64
CTUDT00598544 | Refresh | Closed | rodan_d | June 5, 2018 6:48:18 PM GMT+08:00 | Incomplete Requirements for Stuck Button Check | _Current Production HW_ |  |  | Subsystem |  | GCX.SSRS.iCT.13955 | Change Proposal | cumpson | 0 | 2 |  | GCX | YES | Closed |  | Brilliance Big Bore
CTUDT00598544 | Refresh | Closed | rodan_d | June 5, 2018 6:48:18 PM GMT+08:00 | Incomplete Requirements for Stuck Button Check | _Current Production HW_ |  |  | Subsystem |  | GCX.SSRS.iCT.13955 | Change Proposal | cumpson | 0 | 2 |  | GCX | YES | Closed |  | Ingenuity 64
CTUDT00598544 | Refresh | Closed | rodan_d | June 5, 2018 6:48:18 PM GMT+08:00 | Incomplete Requirements for Stuck Button Check | _Current Production HW_ |  |  | Subsystem |  | GCX.SSRS.iCT.13955 | Change Proposal | cumpson | 0 | 2 |  | GCX | YES | Closed |  | IQon1
CTUDT00598543 | Refresh | Closed | rodan_d | June 5, 2018 6:44:43 PM GMT+08:00 | Lishui Central Hospital system lou out process takes a long time | 4.1.6 |  | Restart the computer. | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 | Console USP Applications | Console4 | NO | Closed |  | Ovation | CFR 9391326 CFR 9434928
CTUDT00598542 | Refresh | Closed | rodan_d | June 5, 2018 6:42:53 PM GMT+08:00 | QueueManager : Deleting\stopping recon processes makes the QueueManager to crash. | 4.1.0.XX225-WIP208.1 |  | Re-launch the Recon Queue Manager. Before pressing delete button, user should select a recon from the list | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 | Console Generic | Console4 | NO | Closed |  | Ovation
CTUDT00598542 | Refresh | Closed | rodan_d | June 5, 2018 6:42:53 PM GMT+08:00 | QueueManager : Deleting\stopping recon processes makes the QueueManager to crash. | 4.1.0.XX225-WIP208.1 |  | Re-launch the Recon Queue Manager. Before pressing delete button, user should select a recon from the list | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 | Console Generic | Console4 | NO | Closed |  | Brilliance 64
CTUDT00598542 | Refresh | Closed | rodan_d | June 5, 2018 6:42:53 PM GMT+08:00 | QueueManager : Deleting\stopping recon processes makes the QueueManager to crash. | 4.1.0.XX225-WIP208.1 |  | Re-launch the Recon Queue Manager. Before pressing delete button, user should select a recon from the list | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 | Console Generic | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00598541 | Refresh | Closed | rodan_d | June 5, 2018 6:40:02 PM GMT+08:00 | Report - Send findings from CS to report using Mass Score protocol , in the impression section 2nd paragraph the comments are displayed for all the scores - Resubmittal of CTUDT00538934 based on new information. | 4.1.6_RfVD; 4BB 4.2; IQon 4.7.0 |  |  | Subsystem |  | NA | Duplicate | redrouthu_r | 0 | 2 |  | iCAP | YES | Closed |  | Ovation | Multiple - to be provided.
CTUDT00598541 | Refresh | Closed | rodan_d | June 5, 2018 6:40:02 PM GMT+08:00 | Report - Send findings from CS to report using Mass Score protocol , in the impression section 2nd paragraph the comments are displayed for all the scores - Resubmittal of CTUDT00538934 based on new information. | 4.1.6_RfVD; 4BB 4.2; IQon 4.7.0 |  |  | Subsystem |  | NA | Duplicate | redrouthu_r | 0 | 2 |  | iCAP | YES | Closed |  | Brilliance 64 | Multiple - to be provided.
CTUDT00598541 | Refresh | Closed | rodan_d | June 5, 2018 6:40:02 PM GMT+08:00 | Report - Send findings from CS to report using Mass Score protocol , in the impression section 2nd paragraph the comments are displayed for all the scores - Resubmittal of CTUDT00538934 based on new information. | 4.1.6_RfVD; 4BB 4.2; IQon 4.7.0 |  |  | Subsystem |  | NA | Duplicate | redrouthu_r | 0 | 2 |  | iCAP | YES | Closed |  | Ingenuity 64 | Multiple - to be provided.
CTUDT00598540 | Refresh | Closed | rodan_d | June 5, 2018 6:36:02 PM GMT+08:00 | security : open ports to for access by any computer: TCP ports, 754, 787, 803, 808, 809, 811, 815 | uCT4.0.1 Release |  | There is no workaround for this issue. | Subsystem |  |  | Does Not Apply | sanjeeva_a | 0 | 3 | Security | Support-Security | YES | Closed |  | Ovation
CTUDT00598540 | Refresh | Closed | rodan_d | June 5, 2018 6:36:02 PM GMT+08:00 | security : open ports to for access by any computer: TCP ports, 754, 787, 803, 808, 809, 811, 815 | uCT4.0.1 Release |  | There is no workaround for this issue. | Subsystem |  |  | Does Not Apply | sanjeeva_a | 0 | 3 | Security | Support-Security | YES | Closed |  | Brilliance 64
CTUDT00598540 | Refresh | Closed | rodan_d | June 5, 2018 6:36:02 PM GMT+08:00 | security : open ports to for access by any computer: TCP ports, 754, 787, 803, 808, 809, 811, 815 | uCT4.0.1 Release |  | There is no workaround for this issue. | Subsystem |  |  | Does Not Apply | sanjeeva_a | 0 | 3 | Security | Support-Security | YES | Closed |  | Brilliance Big Bore
CTUDT00598540 | Refresh | Closed | rodan_d | June 5, 2018 6:36:02 PM GMT+08:00 | security : open ports to for access by any computer: TCP ports, 754, 787, 803, 808, 809, 811, 815 | uCT4.0.1 Release |  | There is no workaround for this issue. | Subsystem |  |  | Does Not Apply | sanjeeva_a | 0 | 3 | Security | Support-Security | YES | Closed |  | Ingenuity 64
CTUDT00598539 | Refresh | Closed | rodan_d | June 5, 2018 6:32:38 PM GMT+08:00 | Service Notes on logout does not Close on clicking on Cross button | 4.1.0.XX224-WIP208 |  | The service notes can be closed by pressing button in the dialog. The close button (X) is inoperative as the service user is required to make an entry in the log. | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 3 | Service Tools | Service Tools | NO | Closed |  | Ovation
CTUDT00598538 | Refresh | Closed | rodan_d | June 5, 2018 6:26:41 PM GMT+08:00 | Update to IFU needed for autolocation. | 4.1.0.XX224-WIP208 | WIP16 |  | Subsystem |  | N/A | Fixed | rajalakshmi_j | 0 | 2 | User Documentation | Support-System | YES | Closed | WIP12SegReq | Ovation | CFR9388301 | WIP16 | May 29, 2019 2:03:11 PM GMT+08:00
CTUDT00598538 | Refresh | Closed | rodan_d | June 5, 2018 6:26:41 PM GMT+08:00 | Update to IFU needed for autolocation. | 4.1.0.XX224-WIP208 | WIP16 |  | Subsystem |  | N/A | Fixed | rajalakshmi_j | 0 | 2 | User Documentation | Support-System | YES | Closed | WIP12SegReq | Brilliance 64 | CFR9388301 | WIP16 | May 29, 2019 2:03:11 PM GMT+08:00
CTUDT00598538 | Refresh | Closed | rodan_d | June 5, 2018 6:26:41 PM GMT+08:00 | Update to IFU needed for autolocation. | 4.1.0.XX224-WIP208 | WIP16 |  | Subsystem |  | N/A | Fixed | rajalakshmi_j | 0 | 2 | User Documentation | Support-System | YES | Closed | WIP12SegReq | Brilliance Big Bore | CFR9388301 | WIP16 | May 29, 2019 2:03:11 PM GMT+08:00
CTUDT00598538 | Refresh | Closed | rodan_d | June 5, 2018 6:26:41 PM GMT+08:00 | Update to IFU needed for autolocation. | 4.1.0.XX224-WIP208 | WIP16 |  | Subsystem |  | N/A | Fixed | rajalakshmi_j | 0 | 2 | User Documentation | Support-System | YES | Closed | WIP12SegReq | Ingenuity 64 | CFR9388301 | WIP16 | May 29, 2019 2:03:11 PM GMT+08:00
CTUDT00598538 | Refresh | Closed | rodan_d | June 5, 2018 6:26:41 PM GMT+08:00 | Update to IFU needed for autolocation. | 4.1.0.XX224-WIP208 | WIP16 |  | Subsystem |  | N/A | Fixed | rajalakshmi_j | 0 | 2 | User Documentation | Support-System | YES | Closed | WIP12SegReq | IQon1 | CFR9388301 | WIP16 | May 29, 2019 2:03:11 PM GMT+08:00
CTUDT00598537 | Refresh | Closed | rodan_d | June 5, 2018 6:24:13 PM GMT+08:00 | User Manual Enhancement Request to include reference to updated 60601-1: 2005 + A1:2012 | v4.1.7.xx019 (akaWIP-5) | WIP16 |  | Subsystem |  | N/A | Fixed | rajalakshmi_j | 0 | 4 | User Documentation | Support-System | NO | Closed | WIP14SegReq | Ovation |  | WIP16 | May 29, 2019 2:02:48 PM GMT+08:00
CTUDT00598537 | Refresh | Closed | rodan_d | June 5, 2018 6:24:13 PM GMT+08:00 | User Manual Enhancement Request to include reference to updated 60601-1: 2005 + A1:2012 | v4.1.7.xx019 (akaWIP-5) | WIP16 |  | Subsystem |  | N/A | Fixed | rajalakshmi_j | 0 | 4 | User Documentation | Support-System | NO | Closed | WIP14SegReq | Ingenuity 64 |  | WIP16 | May 29, 2019 2:02:48 PM GMT+08:00
CTUDT00598536 | Refresh | Closed | rodan_d | June 5, 2018 6:21:54 PM GMT+08:00 | v4.1.6 - Bugrep error displayed on selecting Flip direction for MPR result loaded to slab viewer. | 4.1.6 WIP 105.2 |  | The user is able to load the images into the CT Viewer in Slab mode to reconstruct the same result. the IFU in section "CT and Cardiac Viewer- Slab Mode" describes this process. | Subsystem |  |  | Does Not Apply | sanjeeva_a | 0 | 3 |  | Console4 | NO | Closed |  | Ovation
CTUDT00598536 | Refresh | Closed | rodan_d | June 5, 2018 6:21:54 PM GMT+08:00 | v4.1.6 - Bugrep error displayed on selecting Flip direction for MPR result loaded to slab viewer. | 4.1.6 WIP 105.2 |  | The user is able to load the images into the CT Viewer in Slab mode to reconstruct the same result. the IFU in section "CT and Cardiac Viewer- Slab Mode" describes this process. | Subsystem |  |  | Does Not Apply | sanjeeva_a | 0 | 3 |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00598536 | Refresh | Closed | rodan_d | June 5, 2018 6:21:54 PM GMT+08:00 | v4.1.6 - Bugrep error displayed on selecting Flip direction for MPR result loaded to slab viewer. | 4.1.6 WIP 105.2 |  | The user is able to load the images into the CT Viewer in Slab mode to reconstruct the same result. the IFU in section "CT and Cardiac Viewer- Slab Mode" describes this process. | Subsystem |  |  | Does Not Apply | sanjeeva_a | 0 | 3 |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00598536 | Refresh | Closed | rodan_d | June 5, 2018 6:21:54 PM GMT+08:00 | v4.1.6 - Bugrep error displayed on selecting Flip direction for MPR result loaded to slab viewer. | 4.1.6 WIP 105.2 |  | The user is able to load the images into the CT Viewer in Slab mode to reconstruct the same result. the IFU in section "CT and Cardiac Viewer- Slab Mode" describes this process. | Subsystem |  |  | Does Not Apply | sanjeeva_a | 0 | 3 |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00598536 | Refresh | Closed | rodan_d | June 5, 2018 6:21:54 PM GMT+08:00 | v4.1.6 - Bugrep error displayed on selecting Flip direction for MPR result loaded to slab viewer. | 4.1.6 WIP 105.2 |  | The user is able to load the images into the CT Viewer in Slab mode to reconstruct the same result. the IFU in section "CT and Cardiac Viewer- Slab Mode" describes this process. | Subsystem |  |  | Does Not Apply | sanjeeva_a | 0 | 3 |  | Console4 | NO | Closed |  | IQon1
CTUDT00598525 | Refresh | Closed | rodan_d | June 4, 2018 8:45:30 PM GMT+08:00 | Found by IOCC: Annotations (Text /Arrow / Angle) and Measurements (Lines / ROI / Points) are not visible on ISPACS | WIP 507 |  |  | Subsystem |  | NA | Change Proposal | redrouthu_r | 0 |  | USP | iCAP | NO | Closed |  | IQon1
CTUDT00598524 | Refresh | Closed | rodan_d | June 4, 2018 8:44:33 PM GMT+08:00 | Found by IOCC:There is no option to set the event in which Dose object is generated. | WIP 507 |  |  | Subsystem |  | NA | Change Proposal | kumar_v | 0 | 4 |  | Console4 | NO | Closed |  | IQon1
CTUDT00598523 | Refresh | Closed | rodan_d | June 4, 2018 8:42:19 PM GMT+08:00 | Reconstruction progress of SBI is not in the correct order in the RQM | WIP509.1 |  | NA | Subsystem |  | Console-SRS.15150 | Change Proposal | sanjeeva_a | 0 | 4 |  | Console4 | NO | Closed |  | IQon1
CTUDT00598522 | Refresh | Closed | rodan_d | June 4, 2018 8:41:04 PM GMT+08:00 | IQON HB: In Preference, DDC gets enabled even on unchecking and saving it. | WIP509 |  | 1. Check and uncheck any other preference in “Reconstruction” Preference category. 2. Then change Detector Displacement Correction preference. 3. Save it. | Subsystem |  | N/A | No Plan to Fix | prabhu_m | 0 | 3 |  | Console4 | NO | Closed | WIP14SubCom | IQon1
CTUDT00598521 | Refresh | Closed | rodan_d | June 4, 2018 8:39:46 PM GMT+08:00 | Gantry error message appears when pressing on Enable and movement button on CT Box | WIP510 |  |  | Subsystem |  |  | Change Proposal | cumpson | 0 |  |  | GCX | NO | Closed |  | IQon1
CTUDT00598520 | Refresh | Closed | rodan_d | June 4, 2018 8:38:53 PM GMT+08:00 | Result direction for save SBI is displayed incorrectly when only Spectral and MPR results are planned in clinical scan | WIP508.1 |  | to change SBI direction manually | Subsystem |  |  | Change Proposal | sanjeeva_a | 0 | 3 |  | Console4 | NO | Closed |  | IQon1
CTUDT00598519 | Refresh | Closed | rodan_d | June 4, 2018 8:35:48 PM GMT+08:00 | Improve WES 2009 GPC security by disabling unused services. | WIP506.1 |  |  | Subsystem |  |  | Does Not Apply | cumpson | 0 | 3 |  | GCX | NO | Closed |  | IQon1
CTUDT00598518 | Refresh | Closed | rodan_d | June 4, 2018 8:34:57 PM GMT+08:00 | RFC: MBC algorithm requires additional dicom tag information for investigations and integration | WIP505 |  |  | Subsystem |  |  | Change Proposal | juster_d | 0 | 4 |  | Support_Spectral | NO | Closed |  | IQon1
CTUDT00598517 | Refresh | Closed | rodan_d | June 4, 2018 8:32:03 PM GMT+08:00 | IQ problems near edge of FOV in spectral images | IQon-preRfVD |  | N/A | Subsystem |  | N/A | No Plan to Fix | richards_b | 0 | 4 | CIRS | CIRS | NO | Closed | SubNC | IQon1
CTUDT00598515 | Refresh | Closed | rodan_d | June 4, 2018 8:30:49 PM GMT+08:00 | IQON HB - During perfusion EC result recon which has two seperate conv and SBI recons, Only one recon is shown at a time in RQM. | WIP505 |  |  | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 |  | Console4 | NO | Closed |  | IQon1
CTUDT00598514 | Refresh | Closed | rodan_d | June 4, 2018 8:29:59 PM GMT+08:00 | For CCT and for Perfusion exam cards on editing the result in view scene with Save SBI option checked and not doing any changes in recon parameters, new SBI's are getting created under PD. | WIP505 |  |  | Subsystem |  | Console-SRS-13741 | No Plan to Fix | sanjeeva_a | 0 | 3 |  | Console4 | NO | Closed |  | IQon1
CTUDT00598513 | Refresh | Closed | rodan_d | June 4, 2018 8:29:04 PM GMT+08:00 | v4.1.3 FOK – Scanner never gave the Continue with Manual Scan during this study | WIP509.1 | WIP14 | Press the manual or auto scan button on CT box. | System | Verified1 | None | Fixed | patra_r | 0 | 3 | Console Software | Console4 | YES | Closed | WIP13SegReq | Ovation | PR4214074 PR4393676 PR4499398 | WIP14 | February 19, 2019 3:13:57 PM GMT+08:00
CTUDT00598513 | Refresh | Closed | rodan_d | June 4, 2018 8:29:04 PM GMT+08:00 | v4.1.3 FOK – Scanner never gave the Continue with Manual Scan during this study | WIP509.1 | WIP14 | Press the manual or auto scan button on CT box. | System | Verified1 | None | Fixed | patra_r | 0 | 3 | Console Software | Console4 | YES | Closed | WIP13SegReq | Brilliance 64 | PR4214074 PR4393676 PR4499398 | WIP14 | February 19, 2019 3:13:57 PM GMT+08:00
CTUDT00598513 | Refresh | Closed | rodan_d | June 4, 2018 8:29:04 PM GMT+08:00 | v4.1.3 FOK – Scanner never gave the Continue with Manual Scan during this study | WIP509.1 | WIP14 | Press the manual or auto scan button on CT box. | System | Verified1 | None | Fixed | patra_r | 0 | 3 | Console Software | Console4 | YES | Closed | WIP13SegReq | Ingenuity 64 | PR4214074 PR4393676 PR4499398 | WIP14 | February 19, 2019 3:13:57 PM GMT+08:00
CTUDT00598513 | Refresh | Closed | rodan_d | June 4, 2018 8:29:04 PM GMT+08:00 | v4.1.3 FOK – Scanner never gave the Continue with Manual Scan during this study | WIP509.1 | WIP14 | Press the manual or auto scan button on CT box. | System | Verified1 | None | Fixed | patra_r | 0 | 3 | Console Software | Console4 | YES | Closed | WIP13SegReq | IQon1 | PR4214074 PR4393676 PR4499398 | WIP14 | February 19, 2019 3:13:57 PM GMT+08:00
CTUDT00598512 | Refresh | Closed | rodan_d | June 4, 2018 8:27:59 PM GMT+08:00 | IQON HB:System crashed while editing the axial result | WIP510 |  | NA | Subsystem |  | NA | No Plan to Fix | sanjeeva_a | 0 | 4 |  | Console4 | NO | Closed |  | IQon1
CTUDT00598511 | Refresh | Closed | rodan_d | June 4, 2018 8:26:59 PM GMT+08:00 | Excessive memory consumption by the spectral viewer. | IQon-RfVD |  |  | Subsystem |  | N/A | No Plan to Fix | juster_d | 0 | 4 |  | Support_Spectral | NO | Closed | SegReq_WIP8.1 | IQon1
CTUDT00598510 | Refresh | Closed | rodan_d | June 4, 2018 8:26:12 PM GMT+08:00 | High resolution spectral results: difference compared to standard resolution | WIP501 and v4.7.2 |  | N/A | Subsystem |  | N/A | No Plan to Fix | richards_b | 0 | 4 |  | Support-Physics | NO | Closed | SubNC | IQon1
CTUDT00598509 | Refresh | Closed | rodan_d | June 4, 2018 8:25:19 PM GMT+08:00 | MPR spectral results aren't loading to sCTV with SBI | WIP501 | WIP6 | Load the MPR result without the SBI, or load the SBI and create the results which are needed from the SBI. | Subsystem |  | NA | Fixed | aviw | 0 | 4 | Spectral Viewer | Support_Spectral | NO | Closed | SubCom | IQon1 |  | WIP6 | April 28, 2019 11:46:19 PM GMT+08:00
CTUDT00598508 | Refresh | Closed | rodan_d | June 4, 2018 8:24:32 PM GMT+08:00 | IQ problems in spectral scans with small FOV | WIP501 and v4.7.2 |  | The associated reference to the workaround is found in the product labelling in IFU 459800917203 Rev D, Section 3 Preparing for an Exam/Spectral Results | Subsystem |  | N/A | No Plan to Fix | richards_b | 0 | 3 | CIRS | CIRS | NO | Closed | SubNC | IQon1 | PR7364790
CTUDT00598507 | Refresh | Closed | rodan_d | June 4, 2018 8:23:21 PM GMT+08:00 | Not all images are reconstructed for Spectral Sagittal result of an extended scan | IQon-RfVD |  | User can reconstruct the images through offline recon | Subsystem |  | Console-SRS.10724 | No Plan to Fix | sanjeeva_a | 0 | 3 | Console Software | Console4 | NO | Closed |  | IQon1
CTUDT00598506 | Refresh | Closed | rodan_d | June 4, 2018 8:20:49 PM GMT+08:00 | IV: IQon - Last image of coronal results appeared dark / not as expected with identical window level of series / result | IQon-preRfVD |  | For visualization - change windowing of the last image. If measurements are needed - perform them on the images other than the last image in the sagittal/coronal orientation, or perform measurements on images in other orientations (not sagittal / coronal). | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 3 | Console Generic | Console4 | YES | Closed |  | Ovation
CTUDT00598506 | Refresh | Closed | rodan_d | June 4, 2018 8:20:49 PM GMT+08:00 | IV: IQon - Last image of coronal results appeared dark / not as expected with identical window level of series / result | IQon-preRfVD |  | For visualization - change windowing of the last image. If measurements are needed - perform them on the images other than the last image in the sagittal/coronal orientation, or perform measurements on images in other orientations (not sagittal / coronal). | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 3 | Console Generic | Console4 | YES | Closed |  | Brilliance 64
CTUDT00598506 | Refresh | Closed | rodan_d | June 4, 2018 8:20:49 PM GMT+08:00 | IV: IQon - Last image of coronal results appeared dark / not as expected with identical window level of series / result | IQon-preRfVD |  | For visualization - change windowing of the last image. If measurements are needed - perform them on the images other than the last image in the sagittal/coronal orientation, or perform measurements on images in other orientations (not sagittal / coronal). | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 3 | Console Generic | Console4 | YES | Closed |  | Brilliance Big Bore
CTUDT00598506 | Refresh | Closed | rodan_d | June 4, 2018 8:20:49 PM GMT+08:00 | IV: IQon - Last image of coronal results appeared dark / not as expected with identical window level of series / result | IQon-preRfVD |  | For visualization - change windowing of the last image. If measurements are needed - perform them on the images other than the last image in the sagittal/coronal orientation, or perform measurements on images in other orientations (not sagittal / coronal). | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 3 | Console Generic | Console4 | YES | Closed |  | Ingenuity 64
CTUDT00598506 | Refresh | Closed | rodan_d | June 4, 2018 8:20:49 PM GMT+08:00 | IV: IQon - Last image of coronal results appeared dark / not as expected with identical window level of series / result | IQon-preRfVD |  | For visualization - change windowing of the last image. If measurements are needed - perform them on the images other than the last image in the sagittal/coronal orientation, or perform measurements on images in other orientations (not sagittal / coronal). | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 3 | Console Generic | Console4 | YES | Closed |  | IQon1
CTUDT00598505 | Refresh | Closed | rodan_d | June 4, 2018 8:19:55 PM GMT+08:00 | Artifact near the skull as a result of the multi-resolution stage in structure propagation | IQon-preRfVD |  | The instructions in the customer release notes will include recommendation to use slice thickness of 1mm in head regions | Subsystem | CRN | NA | No Plan to Fix | richards_b | 0 | 4 | CIRS | CIRS | NO | Closed | SubNC | IQon1
CTUDT00598504 | Refresh | Closed | rodan_d | June 4, 2018 8:18:54 PM GMT+08:00 | RfVD: SNUBH: FoK: Antony: Get Ready For surview scan failed with a Gantry error | IQon-RfVD | WIP19 |  | Subsystem | Blocker19 | N/A | Fixed | kline_w | 0 | 2 | Gantry Generic | GCX | YES | Closed | WIP19SegReq | IQon1 | 9401240 9433266 9450165 Minamino – oneEMS 0111592909 Nakakawachi - oneEMS 0111752932 | WIP19 | August 14, 2019 9:09:30 PM GMT+08:00
CTUDT00598503 | Refresh | Closed | rodan_d | June 4, 2018 8:18:00 PM GMT+08:00 | RfVD: Minamino: Could not perform the Short Tube Conditioning | IQon-RfVD |  | Restart Host. This workaround is a valid workaround in the context of the service tasks that are performed. | Subsystem |  | N/A | No Plan to Fix | cumpson | 0 | 4 |  | GCX | NO | Closed |  | Ovation | CFR 9424785
CTUDT00598503 | Refresh | Closed | rodan_d | June 4, 2018 8:18:00 PM GMT+08:00 | RfVD: Minamino: Could not perform the Short Tube Conditioning | IQon-RfVD |  | Restart Host. This workaround is a valid workaround in the context of the service tasks that are performed. | Subsystem |  | N/A | No Plan to Fix | cumpson | 0 | 4 |  | GCX | NO | Closed |  | IQon1 | CFR 9424785
CTUDT00598502 | Refresh | Closed | rodan_d | June 4, 2018 8:17:13 PM GMT+08:00 | RfVD: Lahey: Extra series created on PACS (9 images) | IQon-RfVD |  |  | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 | Console Software | Console4 | NO | Closed |  | Ovation | CFR 9423439
CTUDT00598502 | Refresh | Closed | rodan_d | June 4, 2018 8:17:13 PM GMT+08:00 | RfVD: Lahey: Extra series created on PACS (9 images) | IQon-RfVD |  |  | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 | Console Software | Console4 | NO | Closed |  | Brilliance 64 | CFR 9423439
CTUDT00598502 | Refresh | Closed | rodan_d | June 4, 2018 8:17:13 PM GMT+08:00 | RfVD: Lahey: Extra series created on PACS (9 images) | IQon-RfVD |  |  | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 | Console Software | Console4 | NO | Closed |  | Brilliance Big Bore | CFR 9423439
CTUDT00598502 | Refresh | Closed | rodan_d | June 4, 2018 8:17:13 PM GMT+08:00 | RfVD: Lahey: Extra series created on PACS (9 images) | IQon-RfVD |  |  | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 | Console Software | Console4 | NO | Closed |  | Ingenuity 64 | CFR 9423439
CTUDT00598502 | Refresh | Closed | rodan_d | June 4, 2018 8:17:13 PM GMT+08:00 | RfVD: Lahey: Extra series created on PACS (9 images) | IQon-RfVD |  |  | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 | Console Software | Console4 | NO | Closed |  | IQon1 | CFR 9423439
CTUDT00598501 | Refresh | Closed | rodan_d | June 4, 2018 8:16:21 PM GMT+08:00 | sCTV: Colors are removed from Fusion results after pressing on "Reset all" button | WIP502-4.7.5.43502 | WIP6 | After pressing again on reset all and changing between results in the overlay the colors return. | Subsystem |  | NA | Fixed | aviw | 0 | 4 | Spectral Viewer | Support_Spectral | NO | Closed | SubCom | IQon1 |  | WIP6 | April 29, 2019 3:54:07 PM GMT+08:00
CTUDT00598500 | Refresh | Closed | rodan_d | June 4, 2018 8:15:30 PM GMT+08:00 | RfVD: Lahey - Data Corrupt Message; View2 application crash | IQon-RfVD |  | Need to re-launch the View2 to do offline recon again. | Subsystem | CRN | NA | No Plan to Fix | veerabhadraiah | 0 | 4 | USP | iCAP | NO | Closed |  | Ovation | PR6256680 CFR9423440
CTUDT00598500 | Refresh | Closed | rodan_d | June 4, 2018 8:15:30 PM GMT+08:00 | RfVD: Lahey - Data Corrupt Message; View2 application crash | IQon-RfVD |  | Need to re-launch the View2 to do offline recon again. | Subsystem | CRN | NA | No Plan to Fix | veerabhadraiah | 0 | 4 | USP | iCAP | NO | Closed |  | Brilliance 64 | PR6256680 CFR9423440
CTUDT00598500 | Refresh | Closed | rodan_d | June 4, 2018 8:15:30 PM GMT+08:00 | RfVD: Lahey - Data Corrupt Message; View2 application crash | IQon-RfVD |  | Need to re-launch the View2 to do offline recon again. | Subsystem | CRN | NA | No Plan to Fix | veerabhadraiah | 0 | 4 | USP | iCAP | NO | Closed |  | Brilliance Big Bore | PR6256680 CFR9423440
CTUDT00598500 | Refresh | Closed | rodan_d | June 4, 2018 8:15:30 PM GMT+08:00 | RfVD: Lahey - Data Corrupt Message; View2 application crash | IQon-RfVD |  | Need to re-launch the View2 to do offline recon again. | Subsystem | CRN | NA | No Plan to Fix | veerabhadraiah | 0 | 4 | USP | iCAP | NO | Closed |  | Ingenuity 64 | PR6256680 CFR9423440
CTUDT00598500 | Refresh | Closed | rodan_d | June 4, 2018 8:15:30 PM GMT+08:00 | RfVD: Lahey - Data Corrupt Message; View2 application crash | IQon-RfVD |  | Need to re-launch the View2 to do offline recon again. | Subsystem | CRN | NA | No Plan to Fix | veerabhadraiah | 0 | 4 | USP | iCAP | NO | Closed |  | IQon1 | PR6256680 CFR9423440
CTUDT00598499 | Refresh | Closed | rodan_d | June 4, 2018 8:14:33 PM GMT+08:00 | sCTV: Wrong measurement units in suggested name for batch | WIP503 |  |  | Subsystem | CRN | N/A | No Plan to Fix | juster_d | 0 | 4 | Spectral Viewer | Support_Spectral | NO | Closed | SubNC | IQon1
CTUDT00598498 | Refresh | Closed | rodan_d | June 4, 2018 8:13:43 PM GMT+08:00 | Exam Card Manager crash when changing the scan type from Calcium Scoring to Step&shoot | WIP503 |  | first change scan type to any other but 2d, then to step&Shot. Or start from an axial 3D. | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 3 | Console Software | Console4 | NO | Closed |  | IQon1
CTUDT00598497 | Refresh | Closed | rodan_d | June 4, 2018 8:12:23 PM GMT+08:00 | Iqon HB : Unable to navigate to different results from plan box by clicking on result numbers | WIP503 |  |  | Subsystem |  | Console-SRS.2371 | No Plan to Fix | sanjeeva_a | 0 | 3 |  | Console4 | NO | Closed |  | IQon1
CTUDT00598496 | Refresh | Closed | rodan_d | June 4, 2018 8:11:05 PM GMT+08:00 | RfVD: Anam: End Exam button was grayed out after scan error | IQon-RfVD |  |  | Subsystem |  | NA | No Plan to Fix | Prakash_Chand_Yadav | 0 | 3 | Console Software | Console4 | YES | Closed |  | Ovation | CFM#9432176
CTUDT00598496 | Refresh | Closed | rodan_d | June 4, 2018 8:11:05 PM GMT+08:00 | RfVD: Anam: End Exam button was grayed out after scan error | IQon-RfVD |  |  | Subsystem |  | NA | No Plan to Fix | Prakash_Chand_Yadav | 0 | 3 | Console Software | Console4 | YES | Closed |  | Brilliance 64 | CFM#9432176
CTUDT00598496 | Refresh | Closed | rodan_d | June 4, 2018 8:11:05 PM GMT+08:00 | RfVD: Anam: End Exam button was grayed out after scan error | IQon-RfVD |  |  | Subsystem |  | NA | No Plan to Fix | Prakash_Chand_Yadav | 0 | 3 | Console Software | Console4 | YES | Closed |  | Brilliance Big Bore | CFM#9432176
CTUDT00598496 | Refresh | Closed | rodan_d | June 4, 2018 8:11:05 PM GMT+08:00 | RfVD: Anam: End Exam button was grayed out after scan error | IQon-RfVD |  |  | Subsystem |  | NA | No Plan to Fix | Prakash_Chand_Yadav | 0 | 3 | Console Software | Console4 | YES | Closed |  | Ingenuity 64 | CFM#9432176
CTUDT00598496 | Refresh | Closed | rodan_d | June 4, 2018 8:11:05 PM GMT+08:00 | RfVD: Anam: End Exam button was grayed out after scan error | IQon-RfVD |  |  | Subsystem |  | NA | No Plan to Fix | Prakash_Chand_Yadav | 0 | 3 | Console Software | Console4 | YES | Closed |  | IQon1 | CFM#9432176
CTUDT00598495 | Refresh | Closed | rodan_d | June 4, 2018 8:09:36 PM GMT+08:00 | RfVD: Minamino: Calcium scoring scan not started due to S_MASTER_SEQUENCER_STATE_TIMEOUT_ERR. | IQon-RfVD |  | User rebooted the Host, and the scan succeded | Subsystem |  | N/A | No Plan to Fix | cumpson | 0 | 2 | Gantry Software | GCX | YES | Closed |  | Ovation | 9431158 9442230
CTUDT00598495 | Refresh | Closed | rodan_d | June 4, 2018 8:09:36 PM GMT+08:00 | RfVD: Minamino: Calcium scoring scan not started due to S_MASTER_SEQUENCER_STATE_TIMEOUT_ERR. | IQon-RfVD |  | User rebooted the Host, and the scan succeded | Subsystem |  | N/A | No Plan to Fix | cumpson | 0 | 2 | Gantry Software | GCX | YES | Closed |  | IQon1 | 9431158 9442230
CTUDT00598494 | Refresh | Closed | rodan_d | June 4, 2018 8:08:03 PM GMT+08:00 | IQon-Priority results was not reconstructed based the priority selected | WIP 504 |  |  | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 | Console Software | Console4 | NO | Closed |  | IQon1
CTUDT00598493 | Refresh | Closed | rodan_d | June 4, 2018 8:07:10 PM GMT+08:00 | sCTV: When changing datatype in non fused viewport the color mapping is changed | WIP504 | WIP6 | Select a specific fusion preset after clicking on the Fusion presets group. | Subsystem |  | NA | Fixed | aviw | 0 | 3 | Spectral Viewer | Support_Spectral | NO | Closed | SubCom | IQon1 |  | WIP6 | April 29, 2019 4:04:40 PM GMT+08:00
CTUDT00598492 | Refresh | Closed | rodan_d | June 4, 2018 8:06:19 PM GMT+08:00 | sCTV: Saving Z-eff from 2D as original doesn't work when block measurement unchecked | WIP503 | WIP6 | Delete the effectiveZ result, switch to slab, create the result again,switch back to 2D, and select the result in the series list. Notice that now the result is not colored. It can be saved as original | Subsystem |  | NA | Fixed | aviw | 0 | 3 | Spectral Viewer | Support_Spectral | NO | Closed | SubCom | IQon1 |  | WIP6 | April 29, 2019 4:21:16 PM GMT+08:00
CTUDT00598491 | Refresh | Closed | rodan_d | June 4, 2018 8:05:35 PM GMT+08:00 | IQon HB: Duplicate ECG series gets created for split study when result is edited | WIP503 |  |  | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 | Console Software | Console4 | NO | Closed |  | IQon1
CTUDT00598490 | Refresh | Closed | rodan_d | June 4, 2018 8:04:45 PM GMT+08:00 | RfVD: Ulster: Gantry: waveform stays visible on Touch Panels if disconnecting the PIM before the Leads | IQon-RfVD |  | Cycle system power | Subsystem |  | N/A | No Plan to Fix | cumpson | 0 | 3 | Gantry Software | GCX | YES | Closed |  | Ovation | 9432526
CTUDT00598490 | Refresh | Closed | rodan_d | June 4, 2018 8:04:45 PM GMT+08:00 | RfVD: Ulster: Gantry: waveform stays visible on Touch Panels if disconnecting the PIM before the Leads | IQon-RfVD |  | Cycle system power | Subsystem |  | N/A | No Plan to Fix | cumpson | 0 | 3 | Gantry Software | GCX | YES | Closed |  | IQon1 | 9432526
CTUDT00598489 | Refresh | Closed | rodan_d | June 4, 2018 8:03:39 PM GMT+08:00 | sCTV: KeV value is not displayed in Series List when changing in slab from SBI | WIP504.1 | WIP6 |  | Subsystem |  | NA | Fixed | aviw | 0 | 4 | Spectral Viewer | Support_Spectral | NO | Closed | SubCom | IQon1 |  | WIP6 | April 29, 2019 5:29:43 PM GMT+08:00
CTUDT00598488 | Refresh | Closed | rodan_d | June 4, 2018 8:02:40 PM GMT+08:00 | sCTV: Learn more about doesn't work in fusion overlay result | WIP504.1 | WIP6 | use the drop-down menu from the underlay, or from a non-fusion viewport. | Subsystem |  | NA | Fixed | aviw | 0 | 3 | Spectral Viewer | Support_Spectral | NO | Closed | SubCom | IQon1 |  | WIP6 | April 29, 2019 5:37:54 PM GMT+08:00
CTUDT00598487 | Refresh | Closed | rodan_d | June 4, 2018 8:01:46 PM GMT+08:00 | Spectral result is not covered by Conventional when using Review Before Save | WIP504.1 |  | Assure exact coverage of spectral results by conventional ones during review before save | Subsystem |  | NA | No Plan to Fix | goswami_s | 0 | 3 | Console Software | Console4 | YES | Closed |  | IQon1
CTUDT00598486 | Refresh | Closed | rodan_d | June 4, 2018 8:00:23 PM GMT+08:00 | Pressing pause in the second scan of a timed sequence (perfusion) causing crash | WIP504.2 |  |  | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 3 |  | Console4 | YES | Closed |  | IQon1
CTUDT00598485 | Refresh | Closed | rodan_d | June 4, 2018 7:59:26 PM GMT+08:00 | sCTV: Wrong windowing for Z-Effective appears when the windowing preset is Default | WIP504.2 | WIP20 | right click on the image, in the menu choose Windowing, zeff | Subsystem |  | N/A | Fixed | aviw | 0 | 3 | ApplicationTools | Support_Spectral | NO | Closed | SubCom | IQon1 |  | WIP20 | May 7, 2019 11:46:44 PM GMT+08:00
CTUDT00598484 | Refresh | Closed | rodan_d | June 4, 2018 7:58:34 PM GMT+08:00 | Cannot change recon filter in offline recon of auto-performance scans | WIP505 |  |  | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 |  | Console4 | NO | Closed |  | IQon1
CTUDT00598483 | Refresh | Closed | rodan_d | June 4, 2018 7:57:28 PM GMT+08:00 | sCTV: Manipulation on images in Slab presets are not saved when changing orientation | WIP505 |  |  | Subsystem |  | N/A | No Plan to Fix | juster_d | 0 | 4 |  | Support_Spectral | NO | Closed |  | IQon1
CTUDT00598482 | Refresh | Closed | rodan_d | June 4, 2018 7:56:12 PM GMT+08:00 | System Error appears during performance 100 patient test | WIP505 |  | Retry test | Subsystem |  | N/A | No Plan to Fix | cumpson | 0 | 3 |  | GCX | YES | Closed |  | IQon1
CTUDT00598481 | Refresh | Closed | rodan_d | June 4, 2018 7:55:18 PM GMT+08:00 | System is crashing when trying to export ROI parameters of non-HU results from Exam Application | v4.7.5.43508 |  | to review produced images in the appropriate offline viewer | Subsystem |  | N/A | No Plan to Fix | veerabhadraiah | 0 | 4 |  | iCAP | NO | Closed |  | IQon1
CTUDT00598480 | Refresh | Closed | rodan_d | June 4, 2018 7:53:57 PM GMT+08:00 | Export Dose Log: "export" button throws an exception when there are no valid drives available in the browse Dialog box, | WIP505 |  |  | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 |  | Console4 | NO | Closed |  | IQon1
CTUDT00598479 | Refresh | Closed | rodan_d | June 4, 2018 7:52:35 PM GMT+08:00 | CAT viewer crashed during the import file. CT-NM.SRS.Service.1309:The log viewer shall be able to import the log file with or without markers. | WIP505 |  | Restart Service Tools and try to import again with fweer files | Subsystem |  |  | Does Not Apply | sanjeeva_a | 0 | 3 |  | Service Tools | NO | Closed |  | IQon1
CTUDT00598478 | Refresh | Closed | rodan_d | June 4, 2018 7:51:16 PM GMT+08:00 | sCTV: Datatype name in the text annotation is duplicated when saving the result | WIP505 |  |  | Subsystem |  | N/A | No Plan to Fix | sundeep_s | 0 | 4 |  | Support_Spectral | NO | Closed | WIP9_SegReq | IQon1
CTUDT00598477 | Refresh | Closed | rodan_d | June 4, 2018 7:49:57 PM GMT+08:00 | Localization: error message appears when loading SBI to SCTV after changing UI language | WIP506 |  |  | Subsystem |  | NA | No Plan to Fix | aviw | 0 | 4 | Spectral Viewer | Support_Spectral | NO | Closed |  | IQon1 |  |  | May 7, 2019 8:32:08 PM GMT+08:00
CTUDT00598476 | Refresh | Closed | rodan_d | June 4, 2018 7:48:10 PM GMT+08:00 | Gantry cannot initialize and Not all images created messages appear for 2 Abdomen scans with thickness 1 and 549 images (DMS Buffer) | v4.7.5.43511 |  | To press "Go" button again | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 |  | Console4 | NO | Closed |  | IQon1
CTUDT00598475 | Refresh | Closed | rodan_d | June 4, 2018 7:35:16 PM GMT+08:00 | CTV: Save Results as function doesn't work (Option enabled when RECONSTRUCTION_DIAMETER value greater than 250. | WIP506.1 and v4.7.2 | WIP14 |  | Subsystem | Verified1 | NA | Fixed | redrouthu_r | 0 | 2 | EBW | iCAP | NO | Closed | WIP14SegReq | Ovation |  | WIP14 | March 14, 2019 5:35:45 PM GMT+08:00
CTUDT00598475 | Refresh | Closed | rodan_d | June 4, 2018 7:35:16 PM GMT+08:00 | CTV: Save Results as function doesn't work (Option enabled when RECONSTRUCTION_DIAMETER value greater than 250. | WIP506.1 and v4.7.2 | WIP14 |  | Subsystem | Verified1 | NA | Fixed | redrouthu_r | 0 | 2 | EBW | iCAP | NO | Closed | WIP14SegReq | Brilliance 64 |  | WIP14 | March 14, 2019 5:35:45 PM GMT+08:00
CTUDT00598475 | Refresh | Closed | rodan_d | June 4, 2018 7:35:16 PM GMT+08:00 | CTV: Save Results as function doesn't work (Option enabled when RECONSTRUCTION_DIAMETER value greater than 250. | WIP506.1 and v4.7.2 | WIP14 |  | Subsystem | Verified1 | NA | Fixed | redrouthu_r | 0 | 2 | EBW | iCAP | NO | Closed | WIP14SegReq | Brilliance Big Bore |  | WIP14 | March 14, 2019 5:35:45 PM GMT+08:00
CTUDT00598475 | Refresh | Closed | rodan_d | June 4, 2018 7:35:16 PM GMT+08:00 | CTV: Save Results as function doesn't work (Option enabled when RECONSTRUCTION_DIAMETER value greater than 250. | WIP506.1 and v4.7.2 | WIP14 |  | Subsystem | Verified1 | NA | Fixed | redrouthu_r | 0 | 2 | EBW | iCAP | NO | Closed | WIP14SegReq | Ingenuity 64 |  | WIP14 | March 14, 2019 5:35:45 PM GMT+08:00
CTUDT00598475 | Refresh | Closed | rodan_d | June 4, 2018 7:35:16 PM GMT+08:00 | CTV: Save Results as function doesn't work (Option enabled when RECONSTRUCTION_DIAMETER value greater than 250. | WIP506.1 and v4.7.2 | WIP14 |  | Subsystem | Verified1 | NA | Fixed | redrouthu_r | 0 | 2 | EBW | iCAP | NO | Closed | WIP14SegReq | IQon1 |  | WIP14 | March 14, 2019 5:35:45 PM GMT+08:00
CTUDT00598474 | Refresh | Closed | rodan_d | June 4, 2018 7:33:46 PM GMT+08:00 | "Not all images arrived" message during performance Trauma 150 patient test | WIP507 |  | The user shall perform an offline reconstruction of the acquired data to recover and generate all images obtained. | Subsystem |  | N/A | Works As Specified / Not A Defect | tamari_o | 0 | 4 |  | CIRS | YES | Closed | SubNC | IQon1
CTUDT00598473 | Refresh | Closed | rodan_d | June 4, 2018 7:32:48 PM GMT+08:00 | RfVD: Ansan:Scan initialization error when CPU is busy | IQon-RfVD | WIP16 | to press Go again | System | Verified1 | NA | Fixed | karkour | 0 | 3 | Gantry Software | GCX | NO | Closed | WIP16SubCom | Ovation | TBD | WIP16 | June 25, 2019 2:48:21 AM GMT+08:00
CTUDT00598473 | Refresh | Closed | rodan_d | June 4, 2018 7:32:48 PM GMT+08:00 | RfVD: Ansan:Scan initialization error when CPU is busy | IQon-RfVD | WIP16 | to press Go again | System | Verified1 | NA | Fixed | karkour | 0 | 3 | Gantry Software | GCX | NO | Closed | WIP16SubCom | Brilliance 64 | TBD | WIP16 | June 25, 2019 2:48:21 AM GMT+08:00
CTUDT00598473 | Refresh | Closed | rodan_d | June 4, 2018 7:32:48 PM GMT+08:00 | RfVD: Ansan:Scan initialization error when CPU is busy | IQon-RfVD | WIP16 | to press Go again | System | Verified1 | NA | Fixed | karkour | 0 | 3 | Gantry Software | GCX | NO | Closed | WIP16SubCom | Brilliance Big Bore | TBD | WIP16 | June 25, 2019 2:48:21 AM GMT+08:00
CTUDT00598473 | Refresh | Closed | rodan_d | June 4, 2018 7:32:48 PM GMT+08:00 | RfVD: Ansan:Scan initialization error when CPU is busy | IQon-RfVD | WIP16 | to press Go again | System | Verified1 | NA | Fixed | karkour | 0 | 3 | Gantry Software | GCX | NO | Closed | WIP16SubCom | Ingenuity 64 | TBD | WIP16 | June 25, 2019 2:48:21 AM GMT+08:00
CTUDT00598473 | Refresh | Closed | rodan_d | June 4, 2018 7:32:48 PM GMT+08:00 | RfVD: Ansan:Scan initialization error when CPU is busy | IQon-RfVD | WIP16 | to press Go again | System | Verified1 | NA | Fixed | karkour | 0 | 3 | Gantry Software | GCX | NO | Closed | WIP16SubCom | IQon1 | TBD | WIP16 | June 25, 2019 2:48:21 AM GMT+08:00
CTUDT00598472 | Refresh | Closed | rodan_d | June 4, 2018 7:31:18 PM GMT+08:00 | Val: "Not all images were created” message appeared during reconstruction | WIP507 |  | All images were provided per request. | Subsystem |  | NA | Duplicate | tamari_o | 1 | 4 |  | CIRS | NO | Closed | SubNC | IQon1
CTUDT00598471 | Refresh | Closed | rodan_d | June 4, 2018 7:29:38 PM GMT+08:00 | Val: Scan failed with S_MASTER_SEQUENCER_NO_XRAY message | WIP507 |  | N/A | Subsystem |  | N/A | No Plan to Fix | cumpson | 0 | 3 |  | GCX | YES | Closed |  | Ovation
CTUDT00598471 | Refresh | Closed | rodan_d | June 4, 2018 7:29:38 PM GMT+08:00 | Val: Scan failed with S_MASTER_SEQUENCER_NO_XRAY message | WIP507 |  | N/A | Subsystem |  | N/A | No Plan to Fix | cumpson | 0 | 3 |  | GCX | YES | Closed |  | IQon1
CTUDT00598470 | Refresh | Closed | rodan_d | June 4, 2018 7:28:32 PM GMT+08:00 | Val: Serviceability: Multiple errors during System Status tool test execution. | WIP507 |  | Run Visual diagnostics | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 3 |  | Service Tools | NO | Closed |  | IQon1
CTUDT00598469 | Refresh | Closed | rodan_d | June 4, 2018 7:21:42 PM GMT+08:00 | [IQON HB] The number is missing in the reference surview when the result length goes beyond surview. | WIP508 |  |  | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 |  | Console4 | NO | Closed |  | IQon1
CTUDT00598468 | Refresh | Closed | rodan_d | June 4, 2018 7:20:39 PM GMT+08:00 | sCTV: Text is cut in burnt label when saving MonoE Equivalent with maximum zoom | WIP508 | WIP6 |  | Subsystem |  | N/A | Fixed | aviw | 0 | 4 | Spectral Viewer | Support_Spectral | NO | Closed | SubCom | IQon1 |  | WIP6 | April 29, 2019 7:20:14 PM GMT+08:00
CTUDT00598467 | Refresh | Closed | rodan_d | June 4, 2018 7:19:39 PM GMT+08:00 | When clinical scan is extended more than 2 times then some results(Conventional, Spectral and Save SBI's) are getting stuck in RQM | WIP508.1 |  | 1) Use offline recon to get the images for the extended part. 2) Delete the SBI series from PD and do offline recon to get the correct set of images for the extended part. | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 |  | Console4 | YES | Closed |  | IQon1
CTUDT00598466 | Refresh | Closed | rodan_d | June 4, 2018 7:18:50 PM GMT+08:00 | IQON HB: The Cardiac functionality type of the result changes automatically on pressing "Start Final Recon" | WIP508.1 |  | Execute an additional offline recon request | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 3 |  | Console4 | YES | Closed |  | IQon1
CTUDT00598465 | Refresh | Closed | rodan_d | June 4, 2018 7:17:45 PM GMT+08:00 | CIRS Fails to Get Ready for Pitch less than 0.15 and rotation time 1.5sec | WIP506.1 |  | If the user attempted to scan with a pitch less than 0.15 and a rotation time of 1.5 sec the system would not allow them to proceed with the scan. The user would need to press Go again or increase the pitch or reduce the rotation time in order to proceed. | Subsystem |  | N/A | Works As Specified / Not A Defect | tamari_o | 1 | 3 |  | CIRS | NO | Closed | SubNC | IQon1
CTUDT00598464 | Refresh | Closed | rodan_d | June 4, 2018 7:16:53 PM GMT+08:00 | IQONHB:When Stationary scan is extended twice, more number of images were created for spectral and SBI series of the second time extended scan | WIP509.1 |  | No | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 3 |  | Console4 | NO | Closed |  | IQon1
CTUDT00598463 | Refresh | Closed | rodan_d | June 4, 2018 7:15:58 PM GMT+08:00 | In spectral viewer, when viewing images of spectral results of type Mono Energetic xx keV the annotation overlaps the Burn text hence making it unclear to read. | WIP508.1 |  | use preference to disable the generation of the spectral label | Subsystem |  | N/A | No Plan to Fix | aviw | 0 | 4 |  | Support_Spectral | NO | Closed |  | IQon1
CTUDT00598462 | Refresh | Closed | rodan_d | June 4, 2018 7:15:06 PM GMT+08:00 | When changing the phase in one of the results the message that it was edited appears for all results | WIP509.1 |  |  | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 |  | Console4 | NO | Closed |  | IQon1
CTUDT00598461 | Refresh | Closed | rodan_d | June 4, 2018 7:14:07 PM GMT+08:00 | Copying partial raw data to local folder a raw section that doesn't exist doesn't show any error | WIP509.1 |  | Select the correct range of the raw data. | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 3 |  | Console4 | NO | Closed |  | IQon1
CTUDT00598460 | Refresh | Closed | rodan_d | June 4, 2018 7:13:15 PM GMT+08:00 | After pressing the logout button a pop up that there are reconstructions in progress appears when there aren't any reconstructions in RQM. | WIP509.1 | WIP15 |  | Subsystem |  | NA | Fixed | v_sumanth | 0 | 4 | Console Software | Console4 | NO | Closed | WIP14SegReq | Ovation |  | WIP15 | April 2, 2019 7:13:36 PM GMT+08:00
CTUDT00598460 | Refresh | Closed | rodan_d | June 4, 2018 7:13:15 PM GMT+08:00 | After pressing the logout button a pop up that there are reconstructions in progress appears when there aren't any reconstructions in RQM. | WIP509.1 | WIP15 |  | Subsystem |  | NA | Fixed | v_sumanth | 0 | 4 | Console Software | Console4 | NO | Closed | WIP14SegReq | Brilliance 64 |  | WIP15 | April 2, 2019 7:13:36 PM GMT+08:00
CTUDT00598460 | Refresh | Closed | rodan_d | June 4, 2018 7:13:15 PM GMT+08:00 | After pressing the logout button a pop up that there are reconstructions in progress appears when there aren't any reconstructions in RQM. | WIP509.1 | WIP15 |  | Subsystem |  | NA | Fixed | v_sumanth | 0 | 4 | Console Software | Console4 | NO | Closed | WIP14SegReq | Brilliance Big Bore |  | WIP15 | April 2, 2019 7:13:36 PM GMT+08:00
CTUDT00598460 | Refresh | Closed | rodan_d | June 4, 2018 7:13:15 PM GMT+08:00 | After pressing the logout button a pop up that there are reconstructions in progress appears when there aren't any reconstructions in RQM. | WIP509.1 | WIP15 |  | Subsystem |  | NA | Fixed | v_sumanth | 0 | 4 | Console Software | Console4 | NO | Closed | WIP14SegReq | Ingenuity 64 |  | WIP15 | April 2, 2019 7:13:36 PM GMT+08:00
CTUDT00598460 | Refresh | Closed | rodan_d | June 4, 2018 7:13:15 PM GMT+08:00 | After pressing the logout button a pop up that there are reconstructions in progress appears when there aren't any reconstructions in RQM. | WIP509.1 | WIP15 |  | Subsystem |  | NA | Fixed | v_sumanth | 0 | 4 | Console Software | Console4 | NO | Closed | WIP14SegReq | IQon1 |  | WIP15 | April 2, 2019 7:13:36 PM GMT+08:00
CTUDT00598459 | Refresh | Closed | rodan_d | June 4, 2018 7:12:22 PM GMT+08:00 | Pressing cancel on injector parameters window when injector is armed showing a pop up that parameters changed and injector cannot be updated | WIP509.1 |  | NA | Subsystem |  | Console-SRS.973 | No Plan to Fix | sanjeeva_a | 0 | 4 |  | Console4 | NO | Closed |  | IQon1
CTUDT00598458 | Refresh | Closed | rodan_d | June 4, 2018 7:11:24 PM GMT+08:00 | Val: "Gantry cannot initialize" error appeared multiple times | WIP509.1 |  | user would need to utilize the system within specifications of the accelrate / decelerate duty cycle. | Subsystem |  | N/A | No Plan to Fix | cumpson | 0 | 3 |  | GCX | YES | Closed |  | IQon1
CTUDT00598457 | Refresh | Closed | rodan_d | June 4, 2018 7:10:00 PM GMT+08:00 | Val: Dose notification appeared in exam summary/dose log after lowering DRI to below the notification value | WIP509.1 |  | Look into "Dose Information" on the "Exam Information" page | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 2 |  | Console4 | NO | Closed |  | IQon1
CTUDT00598456 | Refresh | Closed | rodan_d | June 4, 2018 7:08:55 PM GMT+08:00 | Val: Missing Exam Summary Pages | WIP509.1 |  | NA | Subsystem |  | NA | No Plan to Fix | v_sumanth | 0 | 4 |  | Console4 | NO | Closed | BB_SubNC | Ovation
CTUDT00598456 | Refresh | Closed | rodan_d | June 4, 2018 7:08:55 PM GMT+08:00 | Val: Missing Exam Summary Pages | WIP509.1 |  | NA | Subsystem |  | NA | No Plan to Fix | v_sumanth | 0 | 4 |  | Console4 | NO | Closed | BB_SubNC | Brilliance 64
CTUDT00598456 | Refresh | Closed | rodan_d | June 4, 2018 7:08:55 PM GMT+08:00 | Val: Missing Exam Summary Pages | WIP509.1 |  | NA | Subsystem |  | NA | No Plan to Fix | v_sumanth | 0 | 4 |  | Console4 | NO | Closed | BB_SubNC | Brilliance Big Bore
CTUDT00598456 | Refresh | Closed | rodan_d | June 4, 2018 7:08:55 PM GMT+08:00 | Val: Missing Exam Summary Pages | WIP509.1 |  | NA | Subsystem |  | NA | No Plan to Fix | v_sumanth | 0 | 4 |  | Console4 | NO | Closed | BB_SubNC | Ingenuity 64
CTUDT00598456 | Refresh | Closed | rodan_d | June 4, 2018 7:08:55 PM GMT+08:00 | Val: Missing Exam Summary Pages | WIP509.1 |  | NA | Subsystem |  | NA | No Plan to Fix | v_sumanth | 0 | 4 |  | Console4 | NO | Closed | BB_SubNC | IQon1
CTUDT00598455 | Refresh | Closed | rodan_d | June 4, 2018 7:07:45 PM GMT+08:00 | ROI's marked incorrectly in Compare mode of the sCTV | WIP509.1 | WIP6 | click on “reset all” then click the orientation button without scrolling the images. | Subsystem |  | NA | Fixed | aviw | 0 | 3 | Spectral Viewer | Support_Spectral | NO | Closed | SubCom | IQon1 |  | WIP6 | April 29, 2019 7:29:55 PM GMT+08:00
CTUDT00598454 | Refresh | Closed | rodan_d | June 4, 2018 7:06:20 PM GMT+08:00 | IOCC: CT system proposes following SOP Classes which are not claimed in DCS. | WIP507 | DCS for IQon Refresh |  | Subsystem | Verified1 | NA | Fixed | mogilevsky_i | 0 | 4 | Design Documentation | iCAP | NO | Closed | WIP10SegReq | Brilliance 64 |  | DCS for IQon Refresh | April 29, 2019 1:28:12 PM GMT+08:00
CTUDT00598454 | Refresh | Closed | rodan_d | June 4, 2018 7:06:20 PM GMT+08:00 | IOCC: CT system proposes following SOP Classes which are not claimed in DCS. | WIP507 | DCS for IQon Refresh |  | Subsystem | Verified1 | NA | Fixed | mogilevsky_i | 0 | 4 | Design Documentation | iCAP | NO | Closed | WIP10SegReq | Brilliance Big Bore |  | DCS for IQon Refresh | April 29, 2019 1:28:12 PM GMT+08:00
CTUDT00598454 | Refresh | Closed | rodan_d | June 4, 2018 7:06:20 PM GMT+08:00 | IOCC: CT system proposes following SOP Classes which are not claimed in DCS. | WIP507 | DCS for IQon Refresh |  | Subsystem | Verified1 | NA | Fixed | mogilevsky_i | 0 | 4 | Design Documentation | iCAP | NO | Closed | WIP10SegReq | Ingenuity 64 |  | DCS for IQon Refresh | April 29, 2019 1:28:12 PM GMT+08:00
CTUDT00598453 | Refresh | Closed | rodan_d | June 4, 2018 7:04:01 PM GMT+08:00 | IOCC: CT system accepts following SOP Classes which are not claimed in DCS. | WIP507 | DCS for IQon Refresh |  | Subsystem | Verified1 | NA | Fixed | mogilevsky_i | 0 | 4 | Design Documentation | iCAP | NO | Closed | WIP10SegReq | Brilliance 64 |  | DCS for IQon Refresh | April 29, 2019 1:25:51 PM GMT+08:00
CTUDT00598453 | Refresh | Closed | rodan_d | June 4, 2018 7:04:01 PM GMT+08:00 | IOCC: CT system accepts following SOP Classes which are not claimed in DCS. | WIP507 | DCS for IQon Refresh |  | Subsystem | Verified1 | NA | Fixed | mogilevsky_i | 0 | 4 | Design Documentation | iCAP | NO | Closed | WIP10SegReq | Brilliance Big Bore |  | DCS for IQon Refresh | April 29, 2019 1:25:51 PM GMT+08:00
CTUDT00598453 | Refresh | Closed | rodan_d | June 4, 2018 7:04:01 PM GMT+08:00 | IOCC: CT system accepts following SOP Classes which are not claimed in DCS. | WIP507 | DCS for IQon Refresh |  | Subsystem | Verified1 | NA | Fixed | mogilevsky_i | 0 | 4 | Design Documentation | iCAP | NO | Closed | WIP10SegReq | Ingenuity 64 |  | DCS for IQon Refresh | April 29, 2019 1:25:51 PM GMT+08:00
CTUDT00598452 | Refresh | Closed | rodan_d | June 4, 2018 7:02:49 PM GMT+08:00 | Found by IOCC:Even when storage commitment for Dose reports have failed, they can be deleted from the system. | WIP 507 |  | Images should not be deleted before it is confirmed that the images have been trasnferred to the desired location. | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 |  | Console4 | YES | Closed |  | IQon1
CTUDT00598451 | Refresh | Closed | rodan_d | June 4, 2018 7:00:57 PM GMT+08:00 | IQon HB:Bugrep Error window comes up and system error is thrown, when MPR result is deleted and click on GO for the scan. | WIP509.3 |  | workaroud from the investigation: Reduce the number of images and do the scan. the workaround for the issue as agreed in DRB: user needs to restart ""exam application", recover the scans and recons, and use the previous surview for the scan that crashed. | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 |  | Console4 | YES | Closed |  | IQon1
CTUDT00598450 | Refresh | Closed | rodan_d | June 4, 2018 6:56:37 PM GMT+08:00 | IQon HB:Issue found on scan when huge results are planned and on reverting the results are not generated as expected | WIP509.3 |  | User can do offline recon for results as raw data is captured. | Subsystem |  | NA | No Plan to Fix | sowmya_k | 0 | 4 |  | Console4 | YES | Closed | SubNC | IQon1
CTUDT00598449 | Refresh | Closed | rodan_d | June 4, 2018 6:55:43 PM GMT+08:00 | Not all images created pop up message after performing 3-5 patients causing 20 results to erase from RQM | WIP509.3 |  | Perform offline recon | Subsystem |  | NA | Works As Specified / Not A Defect | tamari_o | 1 | 3 |  | CIRS | YES | Closed | SubNC | IQon1
CTUDT00598448 | Refresh | Closed | rodan_d | June 4, 2018 6:54:40 PM GMT+08:00 | Localization of Calcium suppression Index and Electron Density in Preferences -windowing presets are missing. | WIP509.3 | WIP12 | Also, results names can be discerned from the English name, and there are other places in the UI that inform the user about the result | Subsystem |  | NA | Fixed | redrouthu_r | 0 | 2 | EBW | iCAP | NO | Closed | WIP12SegReq | IQon1 |  | WIP12 | February 14, 2019 1:32:54 PM GMT+08:00
CTUDT00598447 | Refresh | Closed | rodan_d | June 4, 2018 6:53:31 PM GMT+08:00 | IQOn HB: Exam summary is not created with all pages. | WIP510 | WIP13 | NA | Subsystem |  | NA | Fixed | datta_i | 0 | 2 | Console Software | Console4 | NO | Closed | WIP13SubCom | IQon1 |  | WIP13 | January 21, 2019 1:20:11 AM GMT+08:00
CTUDT00598446 | Refresh | Closed | rodan_d | June 4, 2018 6:52:40 PM GMT+08:00 | IQON HB: result engine crash | WIP509.4 |  | none | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 |  | Console4 | YES | Closed |  | IQon1
CTUDT00598445 | Refresh | Closed | rodan_d | June 4, 2018 6:51:41 PM GMT+08:00 | Not all images created pop up appears after changing windowing and starting recon in view2 of review before save result | WIP510 |  | User can perform offline reconstruction | Subsystem | SP1 | NA | No Plan to Fix | bollineni_d | 0 | 4 |  | Console4 | NO | Closed | SubNC | IQon1
CTUDT00598444 | Refresh | Closed | rodan_d | June 4, 2018 6:50:45 PM GMT+08:00 | IQon HB:ReviewBeforeSave button is disabled even when result is still not saved. Also system crashes on editing the RBS result and clicking start final recon | WIP510 |  | the offline recon | Subsystem |  | None | No Plan to Fix | sanjeeva_a | 0 | 4 |  | Console4 | NO | Closed |  | IQon1
CTUDT00598443 | Refresh | Closed | rodan_d | June 4, 2018 6:49:48 PM GMT+08:00 | Val: Incorrect series description in local directory for saved CCT image | WIP510 |  |  | Subsystem |  |  | Duplicate | navale_s | 0 | 4 |  | Console4 | NO | Closed |  | IQon1
CTUDT00598442 | Refresh | Closed | rodan_d | June 4, 2018 6:48:49 PM GMT+08:00 | Preview images arrive with long delay to the host and causing start final recon button to be enabled after 30sec to 1min | WIP510 | WIP6 | N/A | Subsystem |  | N/A | Fixed | rakic_k | 0 | 4 | CIRS_SW_Data_Management | CIRS | NO | Closed | WIP6SubCom | IQon1 |  | WIP6 | September 18, 2018 8:35:47 PM GMT+08:00
CTUDT00598441 | Refresh | Closed | rodan_d | June 4, 2018 6:47:51 PM GMT+08:00 | IQon HB: Microsoft .NET framework unhandled exception occurs when tried to logout during reconstruction in progress | WIP 509.4 |  | Exit View2 and then click on LogOut. | Subsystem |  | Console.SRS.2940 | No Plan to Fix | sagar_v | 0 | 4 |  | Console4 | NO | Closed | WIP16SubCom | Ovation
CTUDT00598441 | Refresh | Closed | rodan_d | June 4, 2018 6:47:51 PM GMT+08:00 | IQon HB: Microsoft .NET framework unhandled exception occurs when tried to logout during reconstruction in progress | WIP 509.4 |  | Exit View2 and then click on LogOut. | Subsystem |  | Console.SRS.2940 | No Plan to Fix | sagar_v | 0 | 4 |  | Console4 | NO | Closed | WIP16SubCom | Brilliance 64
CTUDT00598441 | Refresh | Closed | rodan_d | June 4, 2018 6:47:51 PM GMT+08:00 | IQon HB: Microsoft .NET framework unhandled exception occurs when tried to logout during reconstruction in progress | WIP 509.4 |  | Exit View2 and then click on LogOut. | Subsystem |  | Console.SRS.2940 | No Plan to Fix | sagar_v | 0 | 4 |  | Console4 | NO | Closed | WIP16SubCom | Brilliance Big Bore
CTUDT00598441 | Refresh | Closed | rodan_d | June 4, 2018 6:47:51 PM GMT+08:00 | IQon HB: Microsoft .NET framework unhandled exception occurs when tried to logout during reconstruction in progress | WIP 509.4 |  | Exit View2 and then click on LogOut. | Subsystem |  | Console.SRS.2940 | No Plan to Fix | sagar_v | 0 | 4 |  | Console4 | NO | Closed | WIP16SubCom | Ingenuity 64
CTUDT00598441 | Refresh | Closed | rodan_d | June 4, 2018 6:47:51 PM GMT+08:00 | IQon HB: Microsoft .NET framework unhandled exception occurs when tried to logout during reconstruction in progress | WIP 509.4 |  | Exit View2 and then click on LogOut. | Subsystem |  | Console.SRS.2940 | No Plan to Fix | sagar_v | 0 | 4 |  | Console4 | NO | Closed | WIP16SubCom | IQon1
CTUDT00598440 | Refresh | Closed | rodan_d | June 4, 2018 6:46:41 PM GMT+08:00 | Val: Serviceability: During DMS tests performance Tdmc_Tbuf Test failed since Estop was opened | WIP510 |  |  | Subsystem |  | N/A | Works As Specified / Not A Defect | gehtin_e | 0 | 4 |  | DMS | NO | Closed |  | IQon1
CTUDT00598439 | Refresh | Closed | rodan_d | June 4, 2018 6:45:51 PM GMT+08:00 | Val: Serviceability: partial raw data was reconstructed with wider range then required | WIP510 |  |  | Subsystem |  | Console-SRS.15152 | No Plan to Fix | sanjeeva_a | 0 | 4 |  | Console4 | NO | Closed |  | IQon1
CTUDT00598438 | Refresh | Closed | rodan_d | June 4, 2018 6:44:52 PM GMT+08:00 | Performance Test: User input required after test already started (before Case 5 starts) | WIP510 |  | Use phantom without bubble. Formal verification should be done with phantom without bubble. | Subsystem |  | n/a | No Plan to Fix | frider_o | 0 | 4 |  | DMS | NO | Closed |  | IQon1
CTUDT00598437 | Refresh | Closed | rodan_d | June 4, 2018 6:43:56 PM GMT+08:00 | Val: Gantry Initialization Error | WIP510 |  | The user would need to accept the error message and re-try by pressing Go again. | Subsystem | Gantry | NA | Duplicate | richards_b | 0 | 4 |  | CIRS | YES | Closed | SubNC | IQon1
CTUDT00598436 | Refresh | Closed | rodan_d | June 4, 2018 6:42:59 PM GMT+08:00 | Multiple series images cannot be viewed when images are Auto launched in Spectral CT Viewer using Slab viewer tool | WIP507 | WIP6 | Use 2D scene. When loading from patient directory, load both at once, not one after the other. | Subsystem |  | NA | Fixed | aviw | 0 | 4 | Spectral Viewer | Support_Spectral | NO | Closed | SubCom | IQon1 |  | WIP6 | April 29, 2019 7:48:44 PM GMT+08:00
CTUDT00598435 | Refresh | Closed | rodan_d | June 4, 2018 6:42:01 PM GMT+08:00 | CT Scan stays in Initializing state for a while and will never ask to press manual scan in Exam application | 4.1.6 |  | The operator has the option to restart the exam and use previous surview for planning. | Subsystem |  | N/A | No Plan to Fix | cumpson | 0 | 4 |  | GCX | NO | Closed |  | Ovation
CTUDT00598435 | Refresh | Closed | rodan_d | June 4, 2018 6:42:01 PM GMT+08:00 | CT Scan stays in Initializing state for a while and will never ask to press manual scan in Exam application | 4.1.6 |  | The operator has the option to restart the exam and use previous surview for planning. | Subsystem |  | N/A | No Plan to Fix | cumpson | 0 | 4 |  | GCX | NO | Closed |  | Brilliance 64
CTUDT00598435 | Refresh | Closed | rodan_d | June 4, 2018 6:42:01 PM GMT+08:00 | CT Scan stays in Initializing state for a while and will never ask to press manual scan in Exam application | 4.1.6 |  | The operator has the option to restart the exam and use previous surview for planning. | Subsystem |  | N/A | No Plan to Fix | cumpson | 0 | 4 |  | GCX | NO | Closed |  | Brilliance Big Bore
CTUDT00598435 | Refresh | Closed | rodan_d | June 4, 2018 6:42:01 PM GMT+08:00 | CT Scan stays in Initializing state for a while and will never ask to press manual scan in Exam application | 4.1.6 |  | The operator has the option to restart the exam and use previous surview for planning. | Subsystem |  | N/A | No Plan to Fix | cumpson | 0 | 4 |  | GCX | NO | Closed |  | Ingenuity 64
CTUDT00598435 | Refresh | Closed | rodan_d | June 4, 2018 6:42:01 PM GMT+08:00 | CT Scan stays in Initializing state for a while and will never ask to press manual scan in Exam application | 4.1.6 |  | The operator has the option to restart the exam and use previous surview for planning. | Subsystem |  | N/A | No Plan to Fix | cumpson | 0 | 4 |  | GCX | NO | Closed |  | IQon1
CTUDT00598434 | Refresh | Closed | rodan_d | June 4, 2018 6:41:04 PM GMT+08:00 | Validation Test Case Development: IQon HB - Unable to load SBI (warning icon on Spectral Viewer) | SW v.4.7.5.43505 | WIP17 | Use subselection | Subsystem |  | NA | Fixed | s_sa | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP17SubCom | IQon1 |  | WIP17 | May 21, 2019 12:53:57 PM GMT+08:00
CTUDT00598433 | Refresh | Closed | rodan_d | June 4, 2018 6:39:58 PM GMT+08:00 | Validation Test Case Development: iQon HB - 3D Calcium Score plan mis-match to FOV | v4.7.5.43508 |  | Drag the plan box using the mouse and set the FOV to the desired level | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 3 |  | Console4 | NO | Closed |  | Ovation
CTUDT00598433 | Refresh | Closed | rodan_d | June 4, 2018 6:39:58 PM GMT+08:00 | Validation Test Case Development: iQon HB - 3D Calcium Score plan mis-match to FOV | v4.7.5.43508 |  | Drag the plan box using the mouse and set the FOV to the desired level | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 3 |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00598433 | Refresh | Closed | rodan_d | June 4, 2018 6:39:58 PM GMT+08:00 | Validation Test Case Development: iQon HB - 3D Calcium Score plan mis-match to FOV | v4.7.5.43508 |  | Drag the plan box using the mouse and set the FOV to the desired level | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 3 |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00598433 | Refresh | Closed | rodan_d | June 4, 2018 6:39:58 PM GMT+08:00 | Validation Test Case Development: iQon HB - 3D Calcium Score plan mis-match to FOV | v4.7.5.43508 |  | Drag the plan box using the mouse and set the FOV to the desired level | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 3 |  | Console4 | NO | Closed |  | IQon1
CTUDT00598432 | Refresh | Closed | rodan_d | June 4, 2018 6:38:21 PM GMT+08:00 | Tick mark and status message after completion of extended scan recon in Edit Result is not displayed | WIP507 |  | check PD and/or scan viewer for images arriving | Subsystem |  | Console-SRS.2540 | Works As Specified / Not A Defect | babu_s | 1 | 3 |  | Console4 | NO | Closed | WIP14SubCom | IQon1
CTUDT00598430 | Refresh | Closed | rodan_d | June 4, 2018 6:37:17 PM GMT+08:00 | Val: IQon HB selecting 'yes' to a system prompted BugRep collection prompts another BugReport | WIP510 | WIP17 |  | Subsystem |  | Console-SRS.2374 | Fixed | rath_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP16SubCom | Ovation |  | WIP17 | May 6, 2019 6:11:01 PM GMT+08:00
CTUDT00598430 | Refresh | Closed | rodan_d | June 4, 2018 6:37:17 PM GMT+08:00 | Val: IQon HB selecting 'yes' to a system prompted BugRep collection prompts another BugReport | WIP510 | WIP17 |  | Subsystem |  | Console-SRS.2374 | Fixed | rath_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP16SubCom | Brilliance 64 |  | WIP17 | May 6, 2019 6:11:01 PM GMT+08:00
CTUDT00598430 | Refresh | Closed | rodan_d | June 4, 2018 6:37:17 PM GMT+08:00 | Val: IQon HB selecting 'yes' to a system prompted BugRep collection prompts another BugReport | WIP510 | WIP17 |  | Subsystem |  | Console-SRS.2374 | Fixed | rath_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP16SubCom | Ingenuity 64 |  | WIP17 | May 6, 2019 6:11:01 PM GMT+08:00
CTUDT00598430 | Refresh | Closed | rodan_d | June 4, 2018 6:37:17 PM GMT+08:00 | Val: IQon HB selecting 'yes' to a system prompted BugRep collection prompts another BugReport | WIP510 | WIP17 |  | Subsystem |  | Console-SRS.2374 | Fixed | rath_s | 0 | 2 | Console Software | Console4 | NO | Closed | WIP16SubCom | IQon1 |  | WIP17 | May 6, 2019 6:11:01 PM GMT+08:00
CTUDT00598429 | Refresh | Closed | rodan_d | June 4, 2018 6:36:07 PM GMT+08:00 | IQON HB: sCTV: Crash when selecting VNC, then MonoE | WIP507 |  | Re-launch the sCTV and use other order of actions. | Subsystem |  | N/A | No Plan to Fix | aviw | 0 | 4 |  | Support_Spectral | NO | Closed |  | IQon1
CTUDT00598428 | Refresh | Closed | rodan_d | June 4, 2018 6:34:53 PM GMT+08:00 | Image Quality: Spectral results: issues with last images in the reconstructed volume | WIP506 |  | 1. Perform an offline recon with a different slice increment, such that the number of slices is even. 2. The level of artifact is reduced significantly if the result is reconstructed with a level of overlap between the slices which is larger than 1.8 (e.g. increment=0.55mm and thickness=1.0mm as in the attached images). | Subsystem |  | N/A | No Plan to Fix | richards_b | 0 | 3 |  | Support-Physics | YES | Closed | SubNC | IQon1
CTUDT00598427 | Refresh | Closed | rodan_d | June 4, 2018 6:33:49 PM GMT+08:00 | Adobe Reader and Flash player updates | V4.1.7-WIP-4 |  |  | Subsystem |  |  | Does Not Apply | sanjeeva_a | 0 | 3 |  | Console4 | YES | Closed |  | Ovation
CTUDT00598427 | Refresh | Closed | rodan_d | June 4, 2018 6:33:49 PM GMT+08:00 | Adobe Reader and Flash player updates | V4.1.7-WIP-4 |  |  | Subsystem |  |  | Does Not Apply | sanjeeva_a | 0 | 3 |  | Console4 | YES | Closed |  | Brilliance 64
CTUDT00598427 | Refresh | Closed | rodan_d | June 4, 2018 6:33:49 PM GMT+08:00 | Adobe Reader and Flash player updates | V4.1.7-WIP-4 |  |  | Subsystem |  |  | Does Not Apply | sanjeeva_a | 0 | 3 |  | Console4 | YES | Closed |  | Brilliance Big Bore
CTUDT00598427 | Refresh | Closed | rodan_d | June 4, 2018 6:33:49 PM GMT+08:00 | Adobe Reader and Flash player updates | V4.1.7-WIP-4 |  |  | Subsystem |  |  | Does Not Apply | sanjeeva_a | 0 | 3 |  | Console4 | YES | Closed |  | Ingenuity 64
CTUDT00598427 | Refresh | Closed | rodan_d | June 4, 2018 6:33:49 PM GMT+08:00 | Adobe Reader and Flash player updates | V4.1.7-WIP-4 |  |  | Subsystem |  |  | Does Not Apply | sanjeeva_a | 0 | 3 |  | Console4 | YES | Closed |  | IQon1
CTUDT00598425 | Refresh | Closed | rodan_d | June 4, 2018 6:30:07 PM GMT+08:00 | Val: IQon HB workflow crashes when planning a 500mm FOV coronal result on surview | WIP510 |  | The FOV can be adjusted on the axial result. | Subsystem |  |  | Works As Specified / Not A Defect | sanjeeva_a | 1 | 4 |  | Console4 | YES | Closed |  | IQon1
CTUDT00598423 | Refresh | Closed | rodan_d | June 4, 2018 6:27:23 PM GMT+08:00 | CTDI Probe Calibration date shown when starting CTDI test application is incorrect | WIP510 | WIP14 | There is an option to manually type the calibration date of the probe. It shall be documented in manufacturing documentation | Subsystem |  | N/A | Fixed | daibes_s | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP11.1SegReq | Ovation |  | WIP14 | February 11, 2019 5:44:17 PM GMT+08:00
CTUDT00598423 | Refresh | Closed | rodan_d | June 4, 2018 6:27:23 PM GMT+08:00 | CTDI Probe Calibration date shown when starting CTDI test application is incorrect | WIP510 | WIP14 | There is an option to manually type the calibration date of the probe. It shall be documented in manufacturing documentation | Subsystem |  | N/A | Fixed | daibes_s | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP11.1SegReq | Brilliance Big Bore |  | WIP14 | February 11, 2019 5:44:17 PM GMT+08:00
CTUDT00598423 | Refresh | Closed | rodan_d | June 4, 2018 6:27:23 PM GMT+08:00 | CTDI Probe Calibration date shown when starting CTDI test application is incorrect | WIP510 | WIP14 | There is an option to manually type the calibration date of the probe. It shall be documented in manufacturing documentation | Subsystem |  | N/A | Fixed | daibes_s | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP11.1SegReq | Ingenuity 64 |  | WIP14 | February 11, 2019 5:44:17 PM GMT+08:00
CTUDT00598423 | Refresh | Closed | rodan_d | June 4, 2018 6:27:23 PM GMT+08:00 | CTDI Probe Calibration date shown when starting CTDI test application is incorrect | WIP510 | WIP14 | There is an option to manually type the calibration date of the probe. It shall be documented in manufacturing documentation | Subsystem |  | N/A | Fixed | daibes_s | 0 | 3 | DMS Calibrations | DMS | NO | Closed | WIP11.1SegReq | IQon1 |  | WIP14 | February 11, 2019 5:44:17 PM GMT+08:00
CTUDT00598422 | Refresh | Closed | rodan_d | June 4, 2018 6:26:12 PM GMT+08:00 | When a result is reconstructed, which was a Conventional with SBI and Review before save option checked, it generates two SBI series after recon | WIP503 |  | The operator has to delete the first among the duplicate SBI series manually. | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 3 | Console Software | Console4 | NO | Closed |  | IQon1
CTUDT00598420 | Refresh | Closed | rodan_d | June 4, 2018 6:25:13 PM GMT+08:00 | For helical, Axial EC's, Perfusion EC's spectral results reconstruction tick mark symbol never appears against the spectral result in the series list in view2 and in view scene. | WIP505 |  | NA | Subsystem |  | Console-SRS.2288 | No Plan to Fix | sanjeeva_a | 0 | 3 |  | Console4 | NO | Closed |  | IQon1
CTUDT00598419 | Refresh | Closed | rodan_d | June 4, 2018 6:24:08 PM GMT+08:00 | When adding multiple results from an SBI in sCTV 2D scene, the Scroll bar is disabled. | WIP501 | WIP20 |  | Subsystem |  | N/A | Fixed | aviw | 0 | 2 | Spectral Viewer | Support_Spectral | NO | Closed | SubCom | IQon1 |  | WIP20 | May 8, 2019 3:01:42 PM GMT+08:00
CTUDT00598418 | Refresh | Closed | rodan_d | June 4, 2018 6:22:53 PM GMT+08:00 | CDR folder is not cleared when pressing on "Clear" button | WIP501 and v4.7.2 | WIP12 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 4 | USP | iCAP | NO | Closed | WIP12SegReq | Ovation |  | WIP12 | December 14, 2018 5:16:32 PM GMT+08:00
CTUDT00598418 | Refresh | Closed | rodan_d | June 4, 2018 6:22:53 PM GMT+08:00 | CDR folder is not cleared when pressing on "Clear" button | WIP501 and v4.7.2 | WIP12 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 4 | USP | iCAP | NO | Closed | WIP12SegReq | Brilliance 64 |  | WIP12 | December 14, 2018 5:16:32 PM GMT+08:00
CTUDT00598418 | Refresh | Closed | rodan_d | June 4, 2018 6:22:53 PM GMT+08:00 | CDR folder is not cleared when pressing on "Clear" button | WIP501 and v4.7.2 | WIP12 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 4 | USP | iCAP | NO | Closed | WIP12SegReq | Brilliance Big Bore |  | WIP12 | December 14, 2018 5:16:32 PM GMT+08:00
CTUDT00598418 | Refresh | Closed | rodan_d | June 4, 2018 6:22:53 PM GMT+08:00 | CDR folder is not cleared when pressing on "Clear" button | WIP501 and v4.7.2 | WIP12 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 4 | USP | iCAP | NO | Closed | WIP12SegReq | Ingenuity 64 |  | WIP12 | December 14, 2018 5:16:32 PM GMT+08:00
CTUDT00598418 | Refresh | Closed | rodan_d | June 4, 2018 6:22:53 PM GMT+08:00 | CDR folder is not cleared when pressing on "Clear" button | WIP501 and v4.7.2 | WIP12 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 4 | USP | iCAP | NO | Closed | WIP12SegReq | IQon1 |  | WIP12 | December 14, 2018 5:16:32 PM GMT+08:00
CTUDT00598417 | Refresh | Closed | rodan_d | June 4, 2018 6:20:41 PM GMT+08:00 | CTV/sCTV: Creation of Spectral mix from batch of different monoE levels | WIP503 |  | use the two batches instead of generating and using the spectral mix series. | Subsystem |  | N/A | No Plan to Fix | juster_d | 0 | 4 | Spectral Viewer | Support_Spectral | NO | Closed |  | IQon1
CTUDT00598416 | Refresh | Closed | rodan_d | June 4, 2018 6:19:14 PM GMT+08:00 | Iqon HB - Few recon jobs are not getting removed from RQM after deleting | WIP504 |  | User can logout and login again to remove the deleted recons from the rqm. | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 | Console Software | Console4 | NO | Closed |  | IQon1
CTUDT00598415 | Refresh | Closed | rodan_d | June 4, 2018 6:16:25 PM GMT+08:00 | Number of images after 1st extend is not as per planned for an oblique result | WIP504 |  | NA | Subsystem |  | N/A | Works As Specified / Not A Defect | navale_s | 0 | 4 |  | Console4 | NO | Closed | WIP16SubCom | IQon1
CTUDT00598412 | Refresh | Closed | rodan_d | June 4, 2018 6:14:55 PM GMT+08:00 | sCTV: Loading spectral coronal result the sagittal and axial viewports are corrupted | WIP504.1 |  | switch to 2D to see the original results. | Subsystem |  | N/A | Duplicate | juster_d | 0 | 4 | Spectral Viewer | Support_Spectral | NO | Closed |  | IQon1
CTUDT00598410 | Refresh | Closed | rodan_d | June 4, 2018 6:08:02 PM GMT+08:00 | sCTV: Loading spectral coronal result the sagittal and axial viewports are corrupted | WIP504.1 |  | switch to 2D to see the original results. | Subsystem |  | N/A | No Plan to Fix | juster_d | 0 | 4 | Spectral Viewer | Support_Spectral | NO | Closed |  | IQon1
CTUDT00598404 | Refresh | Closed | rodan_d | June 4, 2018 4:43:15 PM GMT+08:00 | The message "The Data-Path results are within limits, corrective action is not recommended" doesn't appear.CT-NM.SRS.Service.225 | WIP505 |  | Service may need to review other logs such as can trace, logger, or CIRS logs for investigation into datapath problems. | Subsystem |  | NA | No Plan to Fix | v_su | 0 | 4 |  | Service Tools | NO | Closed |  | IQon1
CTUDT00598403 | Refresh | Closed | rodan_d | June 4, 2018 4:39:07 PM GMT+08:00 | Re-use of wrong conventional images, extracted from the TempDevice folder. | WIP506.1 |  | Logout and Login at least once after the study has ended before opening the View2.or performing offline recon from raw data and not images | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 |  | Console4 | YES | Closed |  | IQon1
CTUDT00598402 | Refresh | Closed | rodan_d | June 4, 2018 4:37:34 PM GMT+08:00 | Val: Restore Utility presents errors | WIP507 |  | Complete the installation despite the errors (installation will be successful) | Subsystem | Service | N/A | No Plan to Fix | karkour | 0 | 3 |  | GCX | NO | Closed |  | IQon1
CTUDT00598401 | Refresh | Closed | rodan_d | June 4, 2018 4:34:32 PM GMT+08:00 | Val: Incorrect procedure description - Missing Letters | WIP507 |  | NA | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 |  | Console4 | NO | Closed |  | IQon1
CTUDT00598400 | Refresh | Closed | rodan_d | June 4, 2018 4:33:11 PM GMT+08:00 | Val: "Gantry communication is not available. Gantry might be powered off" message popped up during HCOR Basic Infant calibration | WIP507 |  | Close the error message and continue. If the problem persists, the QA application can be restarted. | Subsystem |  | N/A | No Plan to Fix | cumpson | 0 | 4 |  | GCX | NO | Closed |  | Ovation
CTUDT00598400 | Refresh | Closed | rodan_d | June 4, 2018 4:33:11 PM GMT+08:00 | Val: "Gantry communication is not available. Gantry might be powered off" message popped up during HCOR Basic Infant calibration | WIP507 |  | Close the error message and continue. If the problem persists, the QA application can be restarted. | Subsystem |  | N/A | No Plan to Fix | cumpson | 0 | 4 |  | GCX | NO | Closed |  | Brilliance 64
CTUDT00598400 | Refresh | Closed | rodan_d | June 4, 2018 4:33:11 PM GMT+08:00 | Val: "Gantry communication is not available. Gantry might be powered off" message popped up during HCOR Basic Infant calibration | WIP507 |  | Close the error message and continue. If the problem persists, the QA application can be restarted. | Subsystem |  | N/A | No Plan to Fix | cumpson | 0 | 4 |  | GCX | NO | Closed |  | Brilliance Big Bore
CTUDT00598400 | Refresh | Closed | rodan_d | June 4, 2018 4:33:11 PM GMT+08:00 | Val: "Gantry communication is not available. Gantry might be powered off" message popped up during HCOR Basic Infant calibration | WIP507 |  | Close the error message and continue. If the problem persists, the QA application can be restarted. | Subsystem |  | N/A | No Plan to Fix | cumpson | 0 | 4 |  | GCX | NO | Closed |  | Ingenuity 64
CTUDT00598400 | Refresh | Closed | rodan_d | June 4, 2018 4:33:11 PM GMT+08:00 | Val: "Gantry communication is not available. Gantry might be powered off" message popped up during HCOR Basic Infant calibration | WIP507 |  | Close the error message and continue. If the problem persists, the QA application can be restarted. | Subsystem |  | N/A | No Plan to Fix | cumpson | 0 | 4 |  | GCX | NO | Closed |  | IQon1
CTUDT00598399 | Refresh | Closed | rodan_d | June 4, 2018 4:31:42 PM GMT+08:00 | IQON HB: SBI series does not get deleted upon pressing pause and unchecking save sbi or deleting parent conventional+savesbi result | WIP508.1 |  | the workarond is to remove the redundant recon from the RQM. | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 |  | Console4 | NO | Closed |  | IQon1
CTUDT00598398 | Refresh | Closed | rodan_d | June 4, 2018 4:29:11 PM GMT+08:00 | IQONHB: Gantry error and System error occured while performing performing multiple scans ~20scans with more number of images and transferring multiple studies from local to different removable devices like PACS, EBW, USB, Remote devices | WIP508.1 |  |  | Subsystem |  | Console-SRS.10391 | No Plan to Fix | sanjeeva_a | 0 | 3 |  | Console4 | YES | Closed |  | IQon1
CTUDT00598397 | Refresh | Closed | rodan_d | June 4, 2018 4:27:28 PM GMT+08:00 | Val: "ErrorExamCardManager" Application Error While Creating User Exam Card | WIP509.1 | WIP17 |  | Subsystem |  | Console.SRS.850 | Fixed | chokkalingam_r | 0 | 3 | Console Software | Console4 | NO | Closed | WIP16SegReq | Ovation |  | WIP17 | May 16, 2019 11:49:00 AM GMT+08:00
CTUDT00598397 | Refresh | Closed | rodan_d | June 4, 2018 4:27:28 PM GMT+08:00 | Val: "ErrorExamCardManager" Application Error While Creating User Exam Card | WIP509.1 | WIP17 |  | Subsystem |  | Console.SRS.850 | Fixed | chokkalingam_r | 0 | 3 | Console Software | Console4 | NO | Closed | WIP16SegReq | Brilliance 64 |  | WIP17 | May 16, 2019 11:49:00 AM GMT+08:00
CTUDT00598397 | Refresh | Closed | rodan_d | June 4, 2018 4:27:28 PM GMT+08:00 | Val: "ErrorExamCardManager" Application Error While Creating User Exam Card | WIP509.1 | WIP17 |  | Subsystem |  | Console.SRS.850 | Fixed | chokkalingam_r | 0 | 3 | Console Software | Console4 | NO | Closed | WIP16SegReq | Ingenuity 64 |  | WIP17 | May 16, 2019 11:49:00 AM GMT+08:00
CTUDT00598397 | Refresh | Closed | rodan_d | June 4, 2018 4:27:28 PM GMT+08:00 | Val: "ErrorExamCardManager" Application Error While Creating User Exam Card | WIP509.1 | WIP17 |  | Subsystem |  | Console.SRS.850 | Fixed | chokkalingam_r | 0 | 3 | Console Software | Console4 | NO | Closed | WIP16SegReq | Veyron |  | WIP17 | May 16, 2019 11:49:00 AM GMT+08:00
CTUDT00598397 | Refresh | Closed | rodan_d | June 4, 2018 4:27:28 PM GMT+08:00 | Val: "ErrorExamCardManager" Application Error While Creating User Exam Card | WIP509.1 | WIP17 |  | Subsystem |  | Console.SRS.850 | Fixed | chokkalingam_r | 0 | 3 | Console Software | Console4 | NO | Closed | WIP16SegReq | IQon1 |  | WIP17 | May 16, 2019 11:49:00 AM GMT+08:00
CTUDT00598396 | Refresh | Closed | rodan_d | June 4, 2018 4:25:01 PM GMT+08:00 | Val: Queue Manager - Multiple Results Deleted when One Result was Selected | WIP509.1 |  | Perform offline reconstruction | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 3 |  | Console4 | YES | Closed |  | IQon1
CTUDT00598395 | Refresh | Closed | rodan_d | June 4, 2018 4:23:40 PM GMT+08:00 | Few Conventional results with Save SBI wont be reconstructed when merge series is selected in results. | WIP 509.2 |  | do not optimize identical results which have reversed direction. | Subsystem |  | NA | No Plan to Fix | bollineni_d | 0 | 4 |  | Console4 | YES | Closed | SubNC | IQon1
CTUDT00598394 | Refresh | Closed | rodan_d | June 4, 2018 4:22:17 PM GMT+08:00 | Val: Service: "Delete Temporary Data..." Message Stuck After Host Reboot | WIP509.1 |  | Hard reboot | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 3 |  | Console4 | NO | Closed |  | IQon1
CTUDT00598393 | Refresh | Closed | rodan_d | June 4, 2018 4:18:10 PM GMT+08:00 | Val: Duplicated Results with Same Series Number | WIP509.1 |  |  | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 |  | Console4 | NO | Closed |  | IQon1
CTUDT00598389 | Refresh | Closed | manjunath_c | June 1, 2018 4:03:59 PM GMT+08:00 | During edited result reconstruction "System error" is coming | WIP6.1 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | babu_s | 0 |  |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00598387 | Refresh | Closed | manjunath_c | June 1, 2018 4:00:44 PM GMT+08:00 | System error is coming during reconstruction of a study with a combination of IMR and any other recon mode | WIP6.1 |  |  | Subsystem |  |  | Works As Specified / Not A Defect | jose_s | 0 |  |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00598385 | Refresh | Closed | manjunath_c | June 1, 2018 3:37:07 PM GMT+08:00 | After doing 500+ studies bugrep is coming | WIP6.1 |  |  | Subsystem | Bugrep - High Load | NA | Works As Specified / Not A Defect | tamari_o | 1 |  |  | iCAP | NO | Closed | WIP9SegReq | Ovation
CTUDT00598385 | Refresh | Closed | manjunath_c | June 1, 2018 3:37:07 PM GMT+08:00 | After doing 500+ studies bugrep is coming | WIP6.1 |  |  | Subsystem | Bugrep - High Load | NA | Works As Specified / Not A Defect | tamari_o | 1 |  |  | iCAP | NO | Closed | WIP9SegReq | IQon1
CTUDT00598375 | Refresh | Closed | finestine_m | May 31, 2018 9:31:13 PM GMT+08:00 | Refresh - Bad Detectors are not marked when pressing on the BadDet checkbox | WIP7 |  |  | Subsystem |  | n/a | Works As Specified / Not A Defect | gehtin_e | 0 |  |  | DMS | NO | Closed |  | Ingenuity 64
CTUDT00598372 | Refresh | Closed | attali_s | May 31, 2018 9:08:12 PM GMT+08:00 | Refresh - Wrong X Label in Mean Tile TempCo Plot in Stability Test | WIP7 | WIP9.1 |  | Subsystem |  | n/a | Fixed | yosefk | 0 | 2 | DMS Software | DMS | NO | Closed | WIP9_SegReq | Ovation |  | WIP9.1 | September 13, 2018 1:21:06 PM GMT+08:00
CTUDT00598372 | Refresh | Closed | attali_s | May 31, 2018 9:08:12 PM GMT+08:00 | Refresh - Wrong X Label in Mean Tile TempCo Plot in Stability Test | WIP7 | WIP9.1 |  | Subsystem |  | n/a | Fixed | yosefk | 0 | 2 | DMS Software | DMS | NO | Closed | WIP9_SegReq | Brilliance 64 |  | WIP9.1 | September 13, 2018 1:21:06 PM GMT+08:00
CTUDT00598372 | Refresh | Closed | attali_s | May 31, 2018 9:08:12 PM GMT+08:00 | Refresh - Wrong X Label in Mean Tile TempCo Plot in Stability Test | WIP7 | WIP9.1 |  | Subsystem |  | n/a | Fixed | yosefk | 0 | 2 | DMS Software | DMS | NO | Closed | WIP9_SegReq | Ingenuity 64 |  | WIP9.1 | September 13, 2018 1:21:06 PM GMT+08:00
CTUDT00598372 | Refresh | Closed | attali_s | May 31, 2018 9:08:12 PM GMT+08:00 | Refresh - Wrong X Label in Mean Tile TempCo Plot in Stability Test | WIP7 | WIP9.1 |  | Subsystem |  | n/a | Fixed | yosefk | 0 | 2 | DMS Software | DMS | NO | Closed | WIP9_SegReq | IQon1 |  | WIP9.1 | September 13, 2018 1:21:06 PM GMT+08:00
CTUDT00598366 | Refresh | Closed | fitusi_r | May 31, 2018 5:12:31 PM GMT+08:00 | Refresh, IQon -There are no shortcuts for "Service_Tools" and "Backup\Restore" on the Host desktop. | WIP 7.1 | WIP11 | 1. Navigate to “C:\Pms\System” find “LaunchPSC.vbs” file and drag the file to the desktop using the right mouse button and select the option “Create shortcut”. 2. Repeat the same for the following files: LaunchUpgradeUtility.vbs LaunchBnRTools.vbs | Subsystem | Backup / Restore | NA | Fixed | goswami_s | 0 | 4 | Console Software | Service Tools | NO | Closed | WIP11SegReq | Ovation |  | WIP11 | November 15, 2018 5:35:58 PM GMT+08:00
CTUDT00598366 | Refresh | Closed | fitusi_r | May 31, 2018 5:12:31 PM GMT+08:00 | Refresh, IQon -There are no shortcuts for "Service_Tools" and "Backup\Restore" on the Host desktop. | WIP 7.1 | WIP11 | 1. Navigate to “C:\Pms\System” find “LaunchPSC.vbs” file and drag the file to the desktop using the right mouse button and select the option “Create shortcut”. 2. Repeat the same for the following files: LaunchUpgradeUtility.vbs LaunchBnRTools.vbs | Subsystem | Backup / Restore | NA | Fixed | goswami_s | 0 | 4 | Console Software | Service Tools | NO | Closed | WIP11SegReq | Brilliance 64 |  | WIP11 | November 15, 2018 5:35:58 PM GMT+08:00
CTUDT00598366 | Refresh | Closed | fitusi_r | May 31, 2018 5:12:31 PM GMT+08:00 | Refresh, IQon -There are no shortcuts for "Service_Tools" and "Backup\Restore" on the Host desktop. | WIP 7.1 | WIP11 | 1. Navigate to “C:\Pms\System” find “LaunchPSC.vbs” file and drag the file to the desktop using the right mouse button and select the option “Create shortcut”. 2. Repeat the same for the following files: LaunchUpgradeUtility.vbs LaunchBnRTools.vbs | Subsystem | Backup / Restore | NA | Fixed | goswami_s | 0 | 4 | Console Software | Service Tools | NO | Closed | WIP11SegReq | Brilliance Big Bore |  | WIP11 | November 15, 2018 5:35:58 PM GMT+08:00
CTUDT00598366 | Refresh | Closed | fitusi_r | May 31, 2018 5:12:31 PM GMT+08:00 | Refresh, IQon -There are no shortcuts for "Service_Tools" and "Backup\Restore" on the Host desktop. | WIP 7.1 | WIP11 | 1. Navigate to “C:\Pms\System” find “LaunchPSC.vbs” file and drag the file to the desktop using the right mouse button and select the option “Create shortcut”. 2. Repeat the same for the following files: LaunchUpgradeUtility.vbs LaunchBnRTools.vbs | Subsystem | Backup / Restore | NA | Fixed | goswami_s | 0 | 4 | Console Software | Service Tools | NO | Closed | WIP11SegReq | Ingenuity 64 |  | WIP11 | November 15, 2018 5:35:58 PM GMT+08:00
CTUDT00598366 | Refresh | Closed | fitusi_r | May 31, 2018 5:12:31 PM GMT+08:00 | Refresh, IQon -There are no shortcuts for "Service_Tools" and "Backup\Restore" on the Host desktop. | WIP 7.1 | WIP11 | 1. Navigate to “C:\Pms\System” find “LaunchPSC.vbs” file and drag the file to the desktop using the right mouse button and select the option “Create shortcut”. 2. Repeat the same for the following files: LaunchUpgradeUtility.vbs LaunchBnRTools.vbs | Subsystem | Backup / Restore | NA | Fixed | goswami_s | 0 | 4 | Console Software | Service Tools | NO | Closed | WIP11SegReq | IQon1 |  | WIP11 | November 15, 2018 5:35:58 PM GMT+08:00
CTUDT00598358 | Refresh | Closed | bagdady_b | May 31, 2018 12:47:34 PM GMT+08:00 | Refresh- GPC installation process blocked by the Medrad Injector | WIP7.1 |  |  | Subsystem | GPC Installation |  | Does Not Apply | cumpson | 0 |  |  | GCX | NO | Closed |  | IQon1
CTUDT00598354 | Refresh | Closed | rodan_d | May 30, 2018 11:52:55 PM GMT+08:00 | V4.1 FOK - Prospective gated chest exam triggered but did not scan | IQon-preRfVD |  | Perform Retrospective scans; and/or reboot gantry | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 3 | Gantry Generic | GCX | YES | Closed |  | IQon1 | CFR#9084543, PR#2543348
CTUDT00598353 | Refresh | Closed | rodan_d | May 30, 2018 11:50:43 PM GMT+08:00 | RfVD: Sapporo: Kagosima: Surview scan failed due to S_MPL_SURR_SERIAL_OUT_OF_SYNC (Reliability : FM5: Failed Scan: "A problem has occurred, please retry.." message appeared while attempting to perform surview) | IQon-preRfVD |  | No known workaround to prevent this. If this occurs, a rescan may be needed. | Subsystem |  |  | Duplicate | karkour | 0 | 2 | Gantry Generic | GCX | YES | Closed |  | IQon1 | 9440158 9424703
CTUDT00598352 | Refresh | Closed | rodan_d | May 30, 2018 11:48:45 PM GMT+08:00 | ACR Test completed result indicates failure although all test steps passed | WIP510.1 | WIP18 | in case of failure in one of the steps within the test , the entire test should be repeated. .Workaround should be documented in the manufacturing working instruction | Subsystem |  | N/A | Fixed | elviral | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP11.1SegReq | Ovation |  | WIP18 | June 4, 2019 2:56:16 PM GMT+08:00
CTUDT00598352 | Refresh | Closed | rodan_d | May 30, 2018 11:48:45 PM GMT+08:00 | ACR Test completed result indicates failure although all test steps passed | WIP510.1 | WIP18 | in case of failure in one of the steps within the test , the entire test should be repeated. .Workaround should be documented in the manufacturing working instruction | Subsystem |  | N/A | Fixed | elviral | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP11.1SegReq | Brilliance Big Bore |  | WIP18 | June 4, 2019 2:56:16 PM GMT+08:00
CTUDT00598352 | Refresh | Closed | rodan_d | May 30, 2018 11:48:45 PM GMT+08:00 | ACR Test completed result indicates failure although all test steps passed | WIP510.1 | WIP18 | in case of failure in one of the steps within the test , the entire test should be repeated. .Workaround should be documented in the manufacturing working instruction | Subsystem |  | N/A | Fixed | elviral | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP11.1SegReq | Ingenuity 64 |  | WIP18 | June 4, 2019 2:56:16 PM GMT+08:00
CTUDT00598352 | Refresh | Closed | rodan_d | May 30, 2018 11:48:45 PM GMT+08:00 | ACR Test completed result indicates failure although all test steps passed | WIP510.1 | WIP18 | in case of failure in one of the steps within the test , the entire test should be repeated. .Workaround should be documented in the manufacturing working instruction | Subsystem |  | N/A | Fixed | elviral | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP11.1SegReq | IQon1 |  | WIP18 | June 4, 2019 2:56:16 PM GMT+08:00
CTUDT00598350 | Refresh | Closed | fitusi_r | May 30, 2018 11:13:24 PM GMT+08:00 | Refresh, IQon- CIRS configuration issue. | WIP 7.1 | WIP8 | A trained user may go directly into the registry and change the “NumberofNodes” manually instead of running this configuration file on the desktop. | Subsystem | Verified1 | N/A | Fixed | rakic_k | 0 | 3 | CIRS_SW_Installation | CIRS | NO | Closed | WIP8SubCom | Ovation |  | WIP8 | September 18, 2018 8:46:15 PM GMT+08:00
CTUDT00598350 | Refresh | Closed | fitusi_r | May 30, 2018 11:13:24 PM GMT+08:00 | Refresh, IQon- CIRS configuration issue. | WIP 7.1 | WIP8 | A trained user may go directly into the registry and change the “NumberofNodes” manually instead of running this configuration file on the desktop. | Subsystem | Verified1 | N/A | Fixed | rakic_k | 0 | 3 | CIRS_SW_Installation | CIRS | NO | Closed | WIP8SubCom | Brilliance 64 |  | WIP8 | September 18, 2018 8:46:15 PM GMT+08:00
CTUDT00598350 | Refresh | Closed | fitusi_r | May 30, 2018 11:13:24 PM GMT+08:00 | Refresh, IQon- CIRS configuration issue. | WIP 7.1 | WIP8 | A trained user may go directly into the registry and change the “NumberofNodes” manually instead of running this configuration file on the desktop. | Subsystem | Verified1 | N/A | Fixed | rakic_k | 0 | 3 | CIRS_SW_Installation | CIRS | NO | Closed | WIP8SubCom | Brilliance Big Bore |  | WIP8 | September 18, 2018 8:46:15 PM GMT+08:00
CTUDT00598350 | Refresh | Closed | fitusi_r | May 30, 2018 11:13:24 PM GMT+08:00 | Refresh, IQon- CIRS configuration issue. | WIP 7.1 | WIP8 | A trained user may go directly into the registry and change the “NumberofNodes” manually instead of running this configuration file on the desktop. | Subsystem | Verified1 | N/A | Fixed | rakic_k | 0 | 3 | CIRS_SW_Installation | CIRS | NO | Closed | WIP8SubCom | Ingenuity 64 |  | WIP8 | September 18, 2018 8:46:15 PM GMT+08:00
CTUDT00598350 | Refresh | Closed | fitusi_r | May 30, 2018 11:13:24 PM GMT+08:00 | Refresh, IQon- CIRS configuration issue. | WIP 7.1 | WIP8 | A trained user may go directly into the registry and change the “NumberofNodes” manually instead of running this configuration file on the desktop. | Subsystem | Verified1 | N/A | Fixed | rakic_k | 0 | 3 | CIRS_SW_Installation | CIRS | NO | Closed | WIP8SubCom | IQon1 |  | WIP8 | September 18, 2018 8:46:15 PM GMT+08:00
CTUDT00598349 | Refresh | Closed | rodan_d | May 30, 2018 11:11:46 PM GMT+08:00 | IQon 504.2: Wrong Warning message displayed for HR above 75bpm during acquisition | WIP504.2 |  |  | Subsystem | History Review |  | Works As Specified / Not A Defect | sanjeeva_a | 0 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00598348 | Refresh | Closed | rodan_d | May 30, 2018 11:07:17 PM GMT+08:00 | RfVD: Sapporo: Noise in cardiac scan due to low tube output in one of the cycles | RFVD V4.7.2 | WIP12 | The user is able to switch off cardiac DV. | Subsystem |  | N/A | Fixed | gwynn_m | 0 | 3 | DMS Electrical | DMS | YES | Closed | WIP11.1SegReq | Ovation | PR6850427 | WIP12 | February 7, 2019 8:22:13 PM GMT+08:00
CTUDT00598348 | Refresh | Closed | rodan_d | May 30, 2018 11:07:17 PM GMT+08:00 | RfVD: Sapporo: Noise in cardiac scan due to low tube output in one of the cycles | RFVD V4.7.2 | WIP12 | The user is able to switch off cardiac DV. | Subsystem |  | N/A | Fixed | gwynn_m | 0 | 3 | DMS Electrical | DMS | YES | Closed | WIP11.1SegReq | IQon1 | PR6850427 | WIP12 | February 7, 2019 8:22:13 PM GMT+08:00
CTUDT00598347 | Refresh | Closed | rodan_d | May 30, 2018 11:05:37 PM GMT+08:00 | Spectral Parent and Child can have different value for "Single Cycle Reconstruction" and Recon optimisation does not consider "Single Cycle Reconstruction" while optimizing | IQon-RfVD | WIP17 | User should assure both spectral and conventional result are with the same "single-cycle" / "Multi-cycle" parameters. | Subsystem |  | NA | Fixed | tejaswini_nidamanuri | 0 | 3 | Console Software | Console4 | YES | Closed | WIP15SubCom | Ovation |  | WIP17 | May 10, 2019 4:39:37 PM GMT+08:00
CTUDT00598347 | Refresh | Closed | rodan_d | May 30, 2018 11:05:37 PM GMT+08:00 | Spectral Parent and Child can have different value for "Single Cycle Reconstruction" and Recon optimisation does not consider "Single Cycle Reconstruction" while optimizing | IQon-RfVD | WIP17 | User should assure both spectral and conventional result are with the same "single-cycle" / "Multi-cycle" parameters. | Subsystem |  | NA | Fixed | tejaswini_nidamanuri | 0 | 3 | Console Software | Console4 | YES | Closed | WIP15SubCom | Brilliance 64 |  | WIP17 | May 10, 2019 4:39:37 PM GMT+08:00
CTUDT00598347 | Refresh | Closed | rodan_d | May 30, 2018 11:05:37 PM GMT+08:00 | Spectral Parent and Child can have different value for "Single Cycle Reconstruction" and Recon optimisation does not consider "Single Cycle Reconstruction" while optimizing | IQon-RfVD | WIP17 | User should assure both spectral and conventional result are with the same "single-cycle" / "Multi-cycle" parameters. | Subsystem |  | NA | Fixed | tejaswini_nidamanuri | 0 | 3 | Console Software | Console4 | YES | Closed | WIP15SubCom | Ingenuity 64 |  | WIP17 | May 10, 2019 4:39:37 PM GMT+08:00
CTUDT00598347 | Refresh | Closed | rodan_d | May 30, 2018 11:05:37 PM GMT+08:00 | Spectral Parent and Child can have different value for "Single Cycle Reconstruction" and Recon optimisation does not consider "Single Cycle Reconstruction" while optimizing | IQon-RfVD | WIP17 | User should assure both spectral and conventional result are with the same "single-cycle" / "Multi-cycle" parameters. | Subsystem |  | NA | Fixed | tejaswini_nidamanuri | 0 | 3 | Console Software | Console4 | YES | Closed | WIP15SubCom | IQon1 |  | WIP17 | May 10, 2019 4:39:37 PM GMT+08:00
CTUDT00598346 | Refresh | Closed | rodan_d | May 30, 2018 11:03:14 PM GMT+08:00 | 4.1.6 upgrade converted a parameter to an out of range value in a CCT Volume user exam card | 4.1.6 |  |  | Subsystem |  |  | Duplicate | jotawar_d | 0 | 3 |  | Console4 | YES | Closed |  | Ovation | PR6113361
CTUDT00598345 | Refresh | Closed | rodan_d | May 30, 2018 11:00:12 PM GMT+08:00 | BB:Extend behavior is different before and after recon for oblique results | 4.2.0.17xxx-WIP208 | WIP13 |  | System |  | N/A | Fixed | Snehashis_Dey | 0 | 2 | Console Software | Console4 | NO | Closed | MergeFrom4.x | Ovation |  | WIP13 | February 4, 2019 4:11:38 PM GMT+08:00
CTUDT00598344 | Refresh | Closed | rodan_d | May 30, 2018 10:57:14 PM GMT+08:00 | BB42 FOK: Requesting Physician name not saved in image. | V4.1.6 | WIP13 |  | System |  | Console-SRS.2083 | Fixed | Snehashis_Dey | 0 | 3 | Console Generic | Console4 | NO | Closed | MergeFrom4.x | Ovation | CFR9410300 PR5929661 CFR9447844 PR7147837 | WIP13 | February 4, 2019 4:12:00 PM GMT+08:00
CTUDT00598344 | Refresh | Closed | rodan_d | May 30, 2018 10:57:14 PM GMT+08:00 | BB42 FOK: Requesting Physician name not saved in image. | V4.1.6 | WIP13 |  | System |  | Console-SRS.2083 | Fixed | Snehashis_Dey | 0 | 3 | Console Generic | Console4 | NO | Closed | MergeFrom4.x | Brilliance 64 | CFR9410300 PR5929661 CFR9447844 PR7147837 | WIP13 | February 4, 2019 4:12:00 PM GMT+08:00
CTUDT00598344 | Refresh | Closed | rodan_d | May 30, 2018 10:57:14 PM GMT+08:00 | BB42 FOK: Requesting Physician name not saved in image. | V4.1.6 | WIP13 |  | System |  | Console-SRS.2083 | Fixed | Snehashis_Dey | 0 | 3 | Console Generic | Console4 | NO | Closed | MergeFrom4.x | Ingenuity 64 | CFR9410300 PR5929661 CFR9447844 PR7147837 | WIP13 | February 4, 2019 4:12:00 PM GMT+08:00
CTUDT00598344 | Refresh | Closed | rodan_d | May 30, 2018 10:57:14 PM GMT+08:00 | BB42 FOK: Requesting Physician name not saved in image. | V4.1.6 | WIP13 |  | System |  | Console-SRS.2083 | Fixed | Snehashis_Dey | 0 | 3 | Console Generic | Console4 | NO | Closed | MergeFrom4.x | IQon1 | CFR9410300 PR5929661 CFR9447844 PR7147837 | WIP13 | February 4, 2019 4:12:00 PM GMT+08:00
CTUDT00598342 | Refresh | Closed | rodan_d | May 30, 2018 10:54:40 PM GMT+08:00 | BB42 FOK: when sorting IQ checks by report date, the report does not sort | V4.x | WIP14 | No WO | Subsystem |  | N/A | Fixed | yosefk | 0 | 4 | DMS Calibrations | DMS | NO | Closed | WIP13SegReq | Ovation | CFR9427903 PR 6410290 PR6877525 | WIP14 | January 24, 2019 5:45:13 PM GMT+08:00
CTUDT00598340 | Refresh | Closed | rodan_d | May 30, 2018 10:52:15 PM GMT+08:00 | Calcuim Scoring Scans do not handle arrhythmias when requested | 4.1.6.19022 WIP105.2 | WIP13 |  | Subsystem |  | NA | Fixed | sowmya_k | 0 | 4 | Console Software | Console4 | NO | Closed | WIP13SubCom | Ovation |  | WIP13 | January 3, 2019 7:34:01 PM GMT+08:00
CTUDT00598340 | Refresh | Closed | rodan_d | May 30, 2018 10:52:15 PM GMT+08:00 | Calcuim Scoring Scans do not handle arrhythmias when requested | 4.1.6.19022 WIP105.2 | WIP13 |  | Subsystem |  | NA | Fixed | sowmya_k | 0 | 4 | Console Software | Console4 | NO | Closed | WIP13SubCom | Brilliance 64 |  | WIP13 | January 3, 2019 7:34:01 PM GMT+08:00
CTUDT00598340 | Refresh | Closed | rodan_d | May 30, 2018 10:52:15 PM GMT+08:00 | Calcuim Scoring Scans do not handle arrhythmias when requested | 4.1.6.19022 WIP105.2 | WIP13 |  | Subsystem |  | NA | Fixed | sowmya_k | 0 | 4 | Console Software | Console4 | NO | Closed | WIP13SubCom | Ingenuity 64 |  | WIP13 | January 3, 2019 7:34:01 PM GMT+08:00
CTUDT00598339 | Refresh | Closed | rodan_d | May 30, 2018 10:47:39 PM GMT+08:00 | CCT Images are saved as "Orginal", Not "Derived". | 4.1.0.XX224-WIP208 |  |  | Subsystem |  |  | Does Not Apply | oren_d | 0 | 2 | System | Support-System | YES | Closed |  | Ovation
CTUDT00598338 | Refresh | Closed | rodan_d | May 30, 2018 10:44:22 PM GMT+08:00 | Clinical scan fails to start after tracker scan is skipped on iCT | V4.1.7 | WIP16 |  | Subsystem |  | NA | Fixed | karkour | 0 | 3 | Gantry Software | GCX | YES | Closed | WIP16SubCom | Ovation | CFR9374077 PR5131996 PR5901560 PR6006768 | WIP16 | June 12, 2019 3:16:56 AM GMT+08:00
CTUDT00598337 | Refresh | Closed | rodan_d | May 30, 2018 10:40:09 PM GMT+08:00 | Continued intermittent PIM connection issues | iCT3 Legacy |  | Power cycling the system clears the error. | Subsystem |  | N/A | No Plan to Fix | cumpson | 0 | 3 | Gantry Generic | GCX | YES | Closed |  | Ovation | 8948888/PR2215340
CTUDT00598336 | Refresh | Closed | rodan_d | May 30, 2018 10:38:30 PM GMT+08:00 | CTDI reporting should be compliant with IEC 60601-2-44, Edition 3.1 | _Current Production HW_ | WIP16 | n/a | Subsystem |  | N/A | Fixed | rajalakshmi_j | 0 | 2 | User Documentation | Support-Physics | NO | Closed | IFU | Ovation |  | WIP16 | May 29, 2019 2:02:28 PM GMT+08:00
CTUDT00598335 | Refresh | Closed | rodan_d | May 30, 2018 10:35:41 PM GMT+08:00 | Delay of manual start display during bolus tracker scan | 4.1.6xxx30 |  | Restart Console/GPC | Subsystem |  |  | Duplicate | cumpson | 0 | 3 |  | GCX | YES | Closed |  | Ovation
CTUDT00598334 | Refresh | Closed | rodan_d | May 30, 2018 10:28:20 PM GMT+08:00 | Dicom Transfer failure from CT system to PACS | 4.1.1 | WIP13 |  | Subsystem |  | NA | Fixed | Snehashis_Dey | 0 | 2 | Console Software | Console4 | NO | Closed | MergeFrom4.x | Brilliance 64 | 9342124 | WIP13 | February 4, 2019 4:12:31 PM GMT+08:00
CTUDT00598334 | Refresh | Closed | rodan_d | May 30, 2018 10:28:20 PM GMT+08:00 | Dicom Transfer failure from CT system to PACS | 4.1.1 | WIP13 |  | Subsystem |  | NA | Fixed | Snehashis_Dey | 0 | 2 | Console Software | Console4 | NO | Closed | MergeFrom4.x | Ingenuity 64 | 9342124 | WIP13 | February 4, 2019 4:12:31 PM GMT+08:00
CTUDT00598334 | Refresh | Closed | rodan_d | May 30, 2018 10:28:20 PM GMT+08:00 | Dicom Transfer failure from CT system to PACS | 4.1.1 | WIP13 |  | Subsystem |  | NA | Fixed | Snehashis_Dey | 0 | 2 | Console Software | Console4 | NO | Closed | MergeFrom4.x | IQon1 | 9342124 | WIP13 | February 4, 2019 4:12:31 PM GMT+08:00
CTUDT00598333 | Refresh | Closed | rodan_d | May 30, 2018 10:23:53 PM GMT+08:00 | EBW: Patient details are stored in the bugrep | 4.1.x | WIP15 |  | System | Verified1 | NA | Fixed | redrouthu_r | 0 | 2 | USP | iCAP | NO | Closed | WIP14SegReq | Ovation |  | WIP15 | June 13, 2019 12:57:50 PM GMT+08:00
CTUDT00598333 | Refresh | Closed | rodan_d | May 30, 2018 10:23:53 PM GMT+08:00 | EBW: Patient details are stored in the bugrep | 4.1.x | WIP15 |  | System | Verified1 | NA | Fixed | redrouthu_r | 0 | 2 | USP | iCAP | NO | Closed | WIP14SegReq | Brilliance 64 |  | WIP15 | June 13, 2019 12:57:50 PM GMT+08:00
CTUDT00598333 | Refresh | Closed | rodan_d | May 30, 2018 10:23:53 PM GMT+08:00 | EBW: Patient details are stored in the bugrep | 4.1.x | WIP15 |  | System | Verified1 | NA | Fixed | redrouthu_r | 0 | 2 | USP | iCAP | NO | Closed | WIP14SegReq | Ingenuity 64 |  | WIP15 | June 13, 2019 12:57:50 PM GMT+08:00
CTUDT00598332 | Refresh | Closed | rodan_d | May 30, 2018 10:14:15 PM GMT+08:00 | Exam Cards arranged in custom order do not hold position after log out / log in or restart | v4.1.6 |  |  | Subsystem |  |  | Does Not Apply | sanjeeva_a | 0 | 4 |  | Console4 | NO | Closed |  | Ovation
CTUDT00598331 | Refresh | Closed | rodan_d | May 30, 2018 10:09:22 PM GMT+08:00 | Host software doesn't have recent Microsoft security updates | V4.1.6 | WIP15 |  | Subsystem |  | NA | Fixed | dasgupta_s | 0 | 2 | Installation | Console4 | YES | Closed | WIP15SegReq | Ovation |  | WIP15 | April 10, 2019 8:16:07 PM GMT+08:00
CTUDT00598331 | Refresh | Closed | rodan_d | May 30, 2018 10:09:22 PM GMT+08:00 | Host software doesn't have recent Microsoft security updates | V4.1.6 | WIP15 |  | Subsystem |  | NA | Fixed | dasgupta_s | 0 | 2 | Installation | Console4 | YES | Closed | WIP15SegReq | Brilliance 64 |  | WIP15 | April 10, 2019 8:16:07 PM GMT+08:00
CTUDT00598331 | Refresh | Closed | rodan_d | May 30, 2018 10:09:22 PM GMT+08:00 | Host software doesn't have recent Microsoft security updates | V4.1.6 | WIP15 |  | Subsystem |  | NA | Fixed | dasgupta_s | 0 | 2 | Installation | Console4 | YES | Closed | WIP15SegReq | Brilliance Big Bore |  | WIP15 | April 10, 2019 8:16:07 PM GMT+08:00
CTUDT00598331 | Refresh | Closed | rodan_d | May 30, 2018 10:09:22 PM GMT+08:00 | Host software doesn't have recent Microsoft security updates | V4.1.6 | WIP15 |  | Subsystem |  | NA | Fixed | dasgupta_s | 0 | 2 | Installation | Console4 | YES | Closed | WIP15SegReq | Ingenuity 64 |  | WIP15 | April 10, 2019 8:16:07 PM GMT+08:00
CTUDT00598331 | Refresh | Closed | rodan_d | May 30, 2018 10:09:22 PM GMT+08:00 | Host software doesn't have recent Microsoft security updates | V4.1.6 | WIP15 |  | Subsystem |  | NA | Fixed | dasgupta_s | 0 | 2 | Installation | Console4 | YES | Closed | WIP15SegReq | IQon1 |  | WIP15 | April 10, 2019 8:16:07 PM GMT+08:00
CTUDT00598330 | Refresh | Closed | rodan_d | May 30, 2018 10:04:04 PM GMT+08:00 | Iqon HB – Functional results are not getting deleted when deleted from “Edit Phases” dialog. | V4.1.7 |  | In order to delete, user needs to have at least one coronary result. The user can right click on each result and delete without using the "Edit Phases" dialog. | Subsystem |  |  | Does Not Apply | sanjeeva_a | 0 | 4 |  | Console4 | NO | Closed |  | Ovation
CTUDT00598327 | Refresh | Closed | rodan_d | May 30, 2018 9:57:40 PM GMT+08:00 | IQon: Patient details are saved into the bugrep when trying to launch the View2 for the active patient | Eagle_WIP207.1 |  |  | Subsystem |  |  | Does Not Apply | sanjeeva_a | 0 | 2 | Console Generic | Console4 | NO | Closed |  | Ovation
CTUDT00598326 | Refresh | Closed | rodan_d | May 30, 2018 9:55:28 PM GMT+08:00 | IV: 4.1.6 iCT SP - Could not review series 601 until series 607 was reconstructed | 4.1.6.xx023 WIP 106 |  | The IFU in the "Edit Results" section: After a scan is complete and reconstruction has finished, you can create additional, unplanned reconstructions of a specific area. In addition, if you selected the Review Before Save parameter prior to scanning, you can also make changes to the way images are reconstructed. Edit Results has similar functionality to the Plan Window. | Subsystem |  | NA | No Plan to Fix | veerabhadraiah | 0 | 3 |  | iCAP | NO | Closed |  | Ovation
CTUDT00598325 | Refresh | Closed | rodan_d | May 30, 2018 9:53:20 PM GMT+08:00 | IV: Solo - Cycle 2 – Critical Error Upon Logging into PET mode | 4.1.0.XX224-WIP208 | WIP16 | Log out and back in to the user if the errors in this defect are encountered. | Subsystem |  | NA | Fixed | karkour | 0 | 2 | Gantry Software | GCX | NO | Closed | WIP16SubCom | Ovation |  | WIP16 | June 12, 2019 3:16:12 AM GMT+08:00
CTUDT00598324 | Refresh | Closed | rodan_d | May 30, 2018 9:49:25 PM GMT+08:00 | IV: Tracker : wrong GUI progress shown on the scan ruler | 4.1.0.XX246-WIP211.2 |  | There is no workaround for this issue. | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 | Console Generic | Console4 | NO | Closed |  | Ovation
CTUDT00598323 | Refresh | Closed | rodan_d | May 30, 2018 9:46:48 PM GMT+08:00 | IV:Reference Executed Surview did not appear in Exam Summary as expected. | 4.1.3.XX011-WIP309.1 |  | NA | Subsystem |  | Console-SRS.14341 | Works As Specified / Not A Defect | sagar_v | 0 | 3 | Console Generic | Console4 | NO | Closed | WIP16SubCom | Ovation
CTUDT00598323 | Refresh | Closed | rodan_d | May 30, 2018 9:46:48 PM GMT+08:00 | IV:Reference Executed Surview did not appear in Exam Summary as expected. | 4.1.3.XX011-WIP309.1 |  | NA | Subsystem |  | Console-SRS.14341 | Works As Specified / Not A Defect | sagar_v | 0 | 3 | Console Generic | Console4 | NO | Closed | WIP16SubCom | Brilliance 64
CTUDT00598323 | Refresh | Closed | rodan_d | May 30, 2018 9:46:48 PM GMT+08:00 | IV:Reference Executed Surview did not appear in Exam Summary as expected. | 4.1.3.XX011-WIP309.1 |  | NA | Subsystem |  | Console-SRS.14341 | Works As Specified / Not A Defect | sagar_v | 0 | 3 | Console Generic | Console4 | NO | Closed | WIP16SubCom | Brilliance Big Bore
CTUDT00598323 | Refresh | Closed | rodan_d | May 30, 2018 9:46:48 PM GMT+08:00 | IV:Reference Executed Surview did not appear in Exam Summary as expected. | 4.1.3.XX011-WIP309.1 |  | NA | Subsystem |  | Console-SRS.14341 | Works As Specified / Not A Defect | sagar_v | 0 | 3 | Console Generic | Console4 | NO | Closed | WIP16SubCom | Ingenuity 64
CTUDT00598323 | Refresh | Closed | rodan_d | May 30, 2018 9:46:48 PM GMT+08:00 | IV:Reference Executed Surview did not appear in Exam Summary as expected. | 4.1.3.XX011-WIP309.1 |  | NA | Subsystem |  | Console-SRS.14341 | Works As Specified / Not A Defect | sagar_v | 0 | 3 | Console Generic | Console4 | NO | Closed | WIP16SubCom | IQon1
CTUDT00598322 | Refresh | Closed | rodan_d | May 30, 2018 9:34:28 PM GMT+08:00 | Mitigation for CTUDT00159794 - Scan failed to initialize due to S_ROTOR_START_MOTION_TIMEOUT_ERROR | _OVATION_2.5.0.19225 POSTRELEASE | WIP13 |  | Subsystem |  | NA | Fixed | karkour | 0 | 3 | Gantry | GCX | YES | Closed | WIP14SegReq | Ovation | PR5824548 | WIP13 | May 7, 2019 11:50:37 PM GMT+08:00
CTUDT00598322 | Refresh | Closed | rodan_d | May 30, 2018 9:34:28 PM GMT+08:00 | Mitigation for CTUDT00159794 - Scan failed to initialize due to S_ROTOR_START_MOTION_TIMEOUT_ERROR | _OVATION_2.5.0.19225 POSTRELEASE | WIP13 |  | Subsystem |  | NA | Fixed | karkour | 0 | 3 | Gantry | GCX | YES | Closed | WIP14SegReq | IQon1 | PR5824548 | WIP13 | May 7, 2019 11:50:37 PM GMT+08:00
CTUDT00598321 | Refresh | Closed | nuriel_i | May 30, 2018 9:30:45 PM GMT+08:00 | Refresh- Temperature Notification black sections appear on the screen | WIP6.1 | WIP11 |  | Subsystem |  | NA | Fixed | franco_a | 0 | 4 | DMS Software | DMS | NO | Closed | WIP10SegReq | Ovation |  | WIP11 | September 2, 2018 8:08:01 PM GMT+08:00
CTUDT00598321 | Refresh | Closed | nuriel_i | May 30, 2018 9:30:45 PM GMT+08:00 | Refresh- Temperature Notification black sections appear on the screen | WIP6.1 | WIP11 |  | Subsystem |  | NA | Fixed | franco_a | 0 | 4 | DMS Software | DMS | NO | Closed | WIP10SegReq | Ingenuity 64 |  | WIP11 | September 2, 2018 8:08:01 PM GMT+08:00
CTUDT00598321 | Refresh | Closed | nuriel_i | May 30, 2018 9:30:45 PM GMT+08:00 | Refresh- Temperature Notification black sections appear on the screen | WIP6.1 | WIP11 |  | Subsystem |  | NA | Fixed | franco_a | 0 | 4 | DMS Software | DMS | NO | Closed | WIP10SegReq | IQon1 |  | WIP11 | September 2, 2018 8:08:01 PM GMT+08:00
CTUDT00598316 | Refresh | Closed | rodan_d | May 30, 2018 9:11:18 PM GMT+08:00 | Ring artifacts using HDS with iCT V3.0.1 | _OVATION_3.0.1 RfVD _OVATION_3.0.0. |  | Following alternative solutions can be used: 1. Disable HDS. 2. If HDS is enabled, for collimations greater than or equal to 40 mm use of a filter that cuts off less than 13 lp/cm (for example YC) or a larger reconstruction field of view. 3. If HDS is enabled, use of collimations less than 40 mm. | Subsystem |  |  | Does Not Apply | soni_n | 0 | 4 |  | Support-System | YES | Closed |  | Ovation | PR5785862
CTUDT00598315 | Refresh | Closed | rodan_d | May 30, 2018 9:09:44 PM GMT+08:00 | System Status shows failed TDMC_5v_Motor_Drive voltage | V4.1.3 FOK | WIP13 |  | Subsystem |  | N/A | Fixed | ivanf | 0 | 2 | DMS Software | DMS | NO | Closed | WIP12SegReq | Ovation |  | WIP13 | January 17, 2019 4:23:32 PM GMT+08:00
CTUDT00598314 | Refresh | Closed | rodan_d | May 30, 2018 9:07:49 PM GMT+08:00 | The patient size is not displayed in the surview image by the German patient interface | 4.1.3 |  | In some languages, on-screen phrases may be truncated due to space limitations. Patient details may be viewed by selecting the information page icon. | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 |  | Console4 | NO | Closed |  | Ovation | PR5246069
CTUDT00598308 | Refresh | Closed | rodan_d | May 30, 2018 8:56:30 PM GMT+08:00 | V4.1.3 & V4.1.6 When adjusting the countdown setting or when opening injection or voice balloon, the Exam Application crashed | V4.1.6 | WIP12 | The user can restart the application, reselect the patient and use previous surview to recover. Using previous surview is discussed in the IFU. | Subsystem |  | Console-SRS.974, Console-SRS.2451 | Fixed | datta_i | 0 | 4 | Console Software | Console4 | NO | Closed | WIP12SubCom | Ovation | CFR# 9418026 / PR# 6127996 PR5413542 | WIP12 | December 11, 2018 6:53:24 PM GMT+08:00
CTUDT00598308 | Refresh | Closed | rodan_d | May 30, 2018 8:56:30 PM GMT+08:00 | V4.1.3 & V4.1.6 When adjusting the countdown setting or when opening injection or voice balloon, the Exam Application crashed | V4.1.6 | WIP12 | The user can restart the application, reselect the patient and use previous surview to recover. Using previous surview is discussed in the IFU. | Subsystem |  | Console-SRS.974, Console-SRS.2451 | Fixed | datta_i | 0 | 4 | Console Software | Console4 | NO | Closed | WIP12SubCom | Brilliance 64 | CFR# 9418026 / PR# 6127996 PR5413542 | WIP12 | December 11, 2018 6:53:24 PM GMT+08:00
CTUDT00598308 | Refresh | Closed | rodan_d | May 30, 2018 8:56:30 PM GMT+08:00 | V4.1.3 & V4.1.6 When adjusting the countdown setting or when opening injection or voice balloon, the Exam Application crashed | V4.1.6 | WIP12 | The user can restart the application, reselect the patient and use previous surview to recover. Using previous surview is discussed in the IFU. | Subsystem |  | Console-SRS.974, Console-SRS.2451 | Fixed | datta_i | 0 | 4 | Console Software | Console4 | NO | Closed | WIP12SubCom | Ingenuity 64 | CFR# 9418026 / PR# 6127996 PR5413542 | WIP12 | December 11, 2018 6:53:24 PM GMT+08:00
CTUDT00598308 | Refresh | Closed | rodan_d | May 30, 2018 8:56:30 PM GMT+08:00 | V4.1.3 & V4.1.6 When adjusting the countdown setting or when opening injection or voice balloon, the Exam Application crashed | V4.1.6 | WIP12 | The user can restart the application, reselect the patient and use previous surview to recover. Using previous surview is discussed in the IFU. | Subsystem |  | Console-SRS.974, Console-SRS.2451 | Fixed | datta_i | 0 | 4 | Console Software | Console4 | NO | Closed | WIP12SubCom | IQon1 | CFR# 9418026 / PR# 6127996 PR5413542 | WIP12 | December 11, 2018 6:53:24 PM GMT+08:00
CTUDT00598307 | Refresh | Closed | fitusi_r | May 30, 2018 8:55:25 PM GMT+08:00 | Refresh, IQon- When entering the “WorkFlow”, A " Film Viewer " error appears. | WIP 7.1 | WIP9 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP9SubCom | Ovation |  | WIP9 | August 3, 2018 11:59:10 AM GMT+08:00
CTUDT00598307 | Refresh | Closed | fitusi_r | May 30, 2018 8:55:25 PM GMT+08:00 | Refresh, IQon- When entering the “WorkFlow”, A " Film Viewer " error appears. | WIP 7.1 | WIP9 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP9SubCom | Brilliance 64 |  | WIP9 | August 3, 2018 11:59:10 AM GMT+08:00
CTUDT00598307 | Refresh | Closed | fitusi_r | May 30, 2018 8:55:25 PM GMT+08:00 | Refresh, IQon- When entering the “WorkFlow”, A " Film Viewer " error appears. | WIP 7.1 | WIP9 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP9SubCom | Brilliance Big Bore |  | WIP9 | August 3, 2018 11:59:10 AM GMT+08:00
CTUDT00598307 | Refresh | Closed | fitusi_r | May 30, 2018 8:55:25 PM GMT+08:00 | Refresh, IQon- When entering the “WorkFlow”, A " Film Viewer " error appears. | WIP 7.1 | WIP9 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP9SubCom | Ingenuity 64 |  | WIP9 | August 3, 2018 11:59:10 AM GMT+08:00
CTUDT00598307 | Refresh | Closed | fitusi_r | May 30, 2018 8:55:25 PM GMT+08:00 | Refresh, IQon- When entering the “WorkFlow”, A " Film Viewer " error appears. | WIP 7.1 | WIP9 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP9SubCom | IQon1 |  | WIP9 | August 3, 2018 11:59:10 AM GMT+08:00
CTUDT00598305 | Refresh | Closed | rodan_d | May 30, 2018 8:33:29 PM GMT+08:00 | v4.1.6 FOK-extending scan on brain displays differently on exam summary | V4.1.6 |  |  | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 2 |  | Console4 | NO | Closed |  | Ovation | PR6096044
CTUDT00598305 | Refresh | Closed | rodan_d | May 30, 2018 8:33:29 PM GMT+08:00 | v4.1.6 FOK-extending scan on brain displays differently on exam summary | V4.1.6 |  |  | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 2 |  | Console4 | NO | Closed |  | PET - Ingenuity_TF64/128 | PR6096044
CTUDT00598304 | Refresh | Closed | beck_d | May 30, 2018 8:16:48 PM GMT+08:00 | Refresh - Images with ROIs cannot be saved to the Archive | WIP 7 | WIP9 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP9SubCom | Ovation |  | WIP9 | August 3, 2018 11:59:41 AM GMT+08:00
CTUDT00598304 | Refresh | Closed | beck_d | May 30, 2018 8:16:48 PM GMT+08:00 | Refresh - Images with ROIs cannot be saved to the Archive | WIP 7 | WIP9 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP9SubCom | Brilliance 64 |  | WIP9 | August 3, 2018 11:59:41 AM GMT+08:00
CTUDT00598304 | Refresh | Closed | beck_d | May 30, 2018 8:16:48 PM GMT+08:00 | Refresh - Images with ROIs cannot be saved to the Archive | WIP 7 | WIP9 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP9SubCom | Brilliance Big Bore |  | WIP9 | August 3, 2018 11:59:41 AM GMT+08:00
CTUDT00598304 | Refresh | Closed | beck_d | May 30, 2018 8:16:48 PM GMT+08:00 | Refresh - Images with ROIs cannot be saved to the Archive | WIP 7 | WIP9 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP9SubCom | Ingenuity 64 |  | WIP9 | August 3, 2018 11:59:41 AM GMT+08:00
CTUDT00598304 | Refresh | Closed | beck_d | May 30, 2018 8:16:48 PM GMT+08:00 | Refresh - Images with ROIs cannot be saved to the Archive | WIP 7 | WIP9 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP9SubCom | IQon1 |  | WIP9 | August 3, 2018 11:59:41 AM GMT+08:00
CTUDT00598302 | Refresh | Closed | fitusi_r | May 30, 2018 7:58:36 PM GMT+08:00 | Refresh, IQon-Start menu (Taskbar) on the bottom of the application | WIP 7.1 | WIP8.1 |  | Subsystem |  | None | Fixed | mohit_p | 0 | 3 | Console Software | Support-Security | NO | Closed | WIP9SegReq | Ovation |  | WIP8.1 | July 18, 2018 4:12:00 PM GMT+08:00
CTUDT00598302 | Refresh | Closed | fitusi_r | May 30, 2018 7:58:36 PM GMT+08:00 | Refresh, IQon-Start menu (Taskbar) on the bottom of the application | WIP 7.1 | WIP8.1 |  | Subsystem |  | None | Fixed | mohit_p | 0 | 3 | Console Software | Support-Security | NO | Closed | WIP9SegReq | Brilliance 64 |  | WIP8.1 | July 18, 2018 4:12:00 PM GMT+08:00
CTUDT00598302 | Refresh | Closed | fitusi_r | May 30, 2018 7:58:36 PM GMT+08:00 | Refresh, IQon-Start menu (Taskbar) on the bottom of the application | WIP 7.1 | WIP8.1 |  | Subsystem |  | None | Fixed | mohit_p | 0 | 3 | Console Software | Support-Security | NO | Closed | WIP9SegReq | Ingenuity 64 |  | WIP8.1 | July 18, 2018 4:12:00 PM GMT+08:00
CTUDT00598302 | Refresh | Closed | fitusi_r | May 30, 2018 7:58:36 PM GMT+08:00 | Refresh, IQon-Start menu (Taskbar) on the bottom of the application | WIP 7.1 | WIP8.1 |  | Subsystem |  | None | Fixed | mohit_p | 0 | 3 | Console Software | Support-Security | NO | Closed | WIP9SegReq | IQon1 |  | WIP8.1 | July 18, 2018 4:12:00 PM GMT+08:00
CTUDT00598301 | Refresh | Closed | rodan_d | May 30, 2018 7:54:03 PM GMT+08:00 | v4.x Systems allow the users to import exam cards from other systems like iCT v4.x, BR4.x etc. | Eagle_WIP113.2FoK |  | Create Exam Cards using ECM instead Protocol Convertor | Subsystem |  |  | Does Not Apply | sanjeeva_a | 0 | 3 | Console Generic | Console4 | YES | Closed |  | Ovation | None
CTUDT00598301 | Refresh | Closed | rodan_d | May 30, 2018 7:54:03 PM GMT+08:00 | v4.x Systems allow the users to import exam cards from other systems like iCT v4.x, BR4.x etc. | Eagle_WIP113.2FoK |  | Create Exam Cards using ECM instead Protocol Convertor | Subsystem |  |  | Does Not Apply | sanjeeva_a | 0 | 3 | Console Generic | Console4 | YES | Closed |  | PET - Ingenuity_TF64/128 | None
CTUDT00598300 | Refresh | Closed | rodan_d | May 30, 2018 7:51:00 PM GMT+08:00 | Wedge Order Check, collimator diagnostic tool, failed | V4.1.6 | WIP15 WIP14 | The FSE can perform a visual inspection of the wedge order | Subsystem |  | N/A | Fixed | morton_t | 0 | 3 | Gantry Software | Support-Physics | NO | Closed | WIP15SubCom | Ovation | PR6174664 | WIP15 WIP14 | April 9, 2019 8:40:38 PM GMT+08:00
CTUDT00598300 | Refresh | Closed | rodan_d | May 30, 2018 7:51:00 PM GMT+08:00 | Wedge Order Check, collimator diagnostic tool, failed | V4.1.6 | WIP15 WIP14 | The FSE can perform a visual inspection of the wedge order | Subsystem |  | N/A | Fixed | morton_t | 0 | 3 | Gantry Software | Support-Physics | NO | Closed | WIP15SubCom | IQon1 | PR6174664 | WIP15 WIP14 | April 9, 2019 8:40:38 PM GMT+08:00
CTUDT00598297 | Refresh | Closed | fitusi_r | May 30, 2018 6:54:44 PM GMT+08:00 | Refresh,IQon- firewall issue. | WIP 7.1 | WIP8.1 | None | Subsystem | Firewall | None | Duplicate | kapadia_v | 0 | 2 | Console Software | Support-Security | NO | Closed | WIP9SegReq | Ovation |  | WIP8.1 | July 18, 2018 4:11:43 PM GMT+08:00
CTUDT00598297 | Refresh | Closed | fitusi_r | May 30, 2018 6:54:44 PM GMT+08:00 | Refresh,IQon- firewall issue. | WIP 7.1 | WIP8.1 | None | Subsystem | Firewall | None | Duplicate | kapadia_v | 0 | 2 | Console Software | Support-Security | NO | Closed | WIP9SegReq | Brilliance 64 |  | WIP8.1 | July 18, 2018 4:11:43 PM GMT+08:00
CTUDT00598297 | Refresh | Closed | fitusi_r | May 30, 2018 6:54:44 PM GMT+08:00 | Refresh,IQon- firewall issue. | WIP 7.1 | WIP8.1 | None | Subsystem | Firewall | None | Duplicate | kapadia_v | 0 | 2 | Console Software | Support-Security | NO | Closed | WIP9SegReq | Ingenuity 64 |  | WIP8.1 | July 18, 2018 4:11:43 PM GMT+08:00
CTUDT00598297 | Refresh | Closed | fitusi_r | May 30, 2018 6:54:44 PM GMT+08:00 | Refresh,IQon- firewall issue. | WIP 7.1 | WIP8.1 | None | Subsystem | Firewall | None | Duplicate | kapadia_v | 0 | 2 | Console Software | Support-Security | NO | Closed | WIP9SegReq | IQon1 |  | WIP8.1 | July 18, 2018 4:11:43 PM GMT+08:00
CTUDT00598289 | Refresh | Closed | bagdady_b | May 30, 2018 6:00:08 PM GMT+08:00 | Refresh- Patient directory crashes and Pms Remoting service stops | WIP7.1 | WIP9 |  | Subsystem | Patient Directory | NA | Fixed | veerabhadraiah | 0 | 4 | Console USP_EBW | iCAP | NO | Closed | WIP9SubCom | Ovation |  | WIP9 | December 14, 2018 5:15:29 PM GMT+08:00
CTUDT00598289 | Refresh | Closed | bagdady_b | May 30, 2018 6:00:08 PM GMT+08:00 | Refresh- Patient directory crashes and Pms Remoting service stops | WIP7.1 | WIP9 |  | Subsystem | Patient Directory | NA | Fixed | veerabhadraiah | 0 | 4 | Console USP_EBW | iCAP | NO | Closed | WIP9SubCom | Brilliance 6-16 |  | WIP9 | December 14, 2018 5:15:29 PM GMT+08:00
CTUDT00598289 | Refresh | Closed | bagdady_b | May 30, 2018 6:00:08 PM GMT+08:00 | Refresh- Patient directory crashes and Pms Remoting service stops | WIP7.1 | WIP9 |  | Subsystem | Patient Directory | NA | Fixed | veerabhadraiah | 0 | 4 | Console USP_EBW | iCAP | NO | Closed | WIP9SubCom | Brilliance Big Bore |  | WIP9 | December 14, 2018 5:15:29 PM GMT+08:00
CTUDT00598289 | Refresh | Closed | bagdady_b | May 30, 2018 6:00:08 PM GMT+08:00 | Refresh- Patient directory crashes and Pms Remoting service stops | WIP7.1 | WIP9 |  | Subsystem | Patient Directory | NA | Fixed | veerabhadraiah | 0 | 4 | Console USP_EBW | iCAP | NO | Closed | WIP9SubCom | Ingenuity 64 |  | WIP9 | December 14, 2018 5:15:29 PM GMT+08:00
CTUDT00598289 | Refresh | Closed | bagdady_b | May 30, 2018 6:00:08 PM GMT+08:00 | Refresh- Patient directory crashes and Pms Remoting service stops | WIP7.1 | WIP9 |  | Subsystem | Patient Directory | NA | Fixed | veerabhadraiah | 0 | 4 | Console USP_EBW | iCAP | NO | Closed | WIP9SubCom | IQon1 |  | WIP9 | December 14, 2018 5:15:29 PM GMT+08:00
CTUDT00598287 | Refresh | Closed | bagdady_b | May 30, 2018 4:44:54 PM GMT+08:00 | Refresh-sCTV: Unable to load old SBI or an SBI from DEMO folder | WIP7.1 |  |  | Subsystem | sCTV - SBI loading | N/A | Works As Specified / Not A Defect | tamari_o | 1 |  |  | Support_Spectral | YES | Closed |  | IQon1
CTUDT00598285 | Refresh | Closed | bagdady_b | May 30, 2018 4:33:04 PM GMT+08:00 | Refresh- Sleep mode causing disconnection of Host and Gantry | WIP7.1 | WIP10 | Manually configure Console to never sleep on power plugged-in. | Subsystem | Verified1 | None | Fixed | dasgupta_s | 0 | 3 | Console Software | DI | NO | Closed | WIP9SegReq | Brilliance 64 |  | WIP10 | January 9, 2020 6:53:23 PM GMT+08:00
CTUDT00598285 | Refresh | Closed | bagdady_b | May 30, 2018 4:33:04 PM GMT+08:00 | Refresh- Sleep mode causing disconnection of Host and Gantry | WIP7.1 | WIP10 | Manually configure Console to never sleep on power plugged-in. | Subsystem | Verified1 | None | Fixed | dasgupta_s | 0 | 3 | Console Software | DI | NO | Closed | WIP9SegReq | Ingenuity 64 |  | WIP10 | January 9, 2020 6:53:23 PM GMT+08:00
CTUDT00598283 | Refresh | Closed | bagdady_b | May 30, 2018 4:31:32 PM GMT+08:00 | Refresh- An error message "Error during scan" appears when performing AIR calibration | WIP7.1 |  |  | Subsystem | Quality Assurance - Air Calibrattion | N/A | Works As Specified / Not A Defect | tamari_o | 1 |  |  | DMS | NO | Closed |  | IQon1
CTUDT00598281 | Refresh | Closed | bagdady_b | May 30, 2018 4:23:04 PM GMT+08:00 | Refresh- Unable to proceed scanning due to Tube is too hot error | WIP7.1 | CAT Version 4.34.5.0 |  | Subsystem | Tube Heat defect | N/A | Fixed | krisher_i | 0 | 2 | Test Design | Support-System | NO | Closed | WIP13SegReq | IQon1 |  | CAT Version 4.34.5.0 | January 8, 2019 10:32:16 PM GMT+08:00
CTUDT00598267 | Refresh | Closed | bagdady_b | May 29, 2018 10:12:44 PM GMT+08:00 | Refresh- Anonymize changes the study date to incorrect date | WIP7 |  |  | Subsystem |  | NA | Works As Specified / Not A Defect | veerabhadraiah | 0 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00598265 | Refresh | Closed | bagdady_b | May 29, 2018 9:38:47 PM GMT+08:00 | Refresh- sCTV: Remove all Graphics button does not remove Spectral ROI | WIP7 | WIP17 | remove the spectral ROI from its context menu | Subsystem | sCTV - remove all graphics | N/A | Fixed | aviw | 0 | 4 | SDViewer | Support_Spectral | NO | Closed | WIP9_SegReq | IQon1 |  | WIP17 | May 20, 2019 10:32:42 PM GMT+08:00
CTUDT00598263 | Refresh | Closed | bagdady_b | May 29, 2018 9:34:54 PM GMT+08:00 | Refresh- sCTV: Add to running application for SBI doesn't function properly | WIP7 | WIP13 | open the series list, click on the conventional series which match the SBI. Now the data type menu will show all spectral type and you create any spectral result. | Subsystem | sCTV - Add to running application | N/A | Fixed | juster_d | 0 | 3 | Console USP_EBW | Support_Spectral | NO | Closed | SegReq_WIP9 | IQon1 |  | WIP13 | February 20, 2019 6:35:05 PM GMT+08:00
CTUDT00598261 | Refresh | Closed | bagdady_b | May 29, 2018 9:31:18 PM GMT+08:00 | Refresh- sCTV\CTV: Patients age is not calculated from DICOM tags automatically | WIP7 | WIP9.1 |  | Subsystem |  | NA | Fixed | sundeep_s | 0 | 3 | Console USP_EBW | iCAP | YES | Closed | WIP10SegReq | Ovation |  | WIP9.1 | August 28, 2018 6:35:57 PM GMT+08:00
CTUDT00598261 | Refresh | Closed | bagdady_b | May 29, 2018 9:31:18 PM GMT+08:00 | Refresh- sCTV\CTV: Patients age is not calculated from DICOM tags automatically | WIP7 | WIP9.1 |  | Subsystem |  | NA | Fixed | sundeep_s | 0 | 3 | Console USP_EBW | iCAP | YES | Closed | WIP10SegReq | Brilliance 64 |  | WIP9.1 | August 28, 2018 6:35:57 PM GMT+08:00
CTUDT00598261 | Refresh | Closed | bagdady_b | May 29, 2018 9:31:18 PM GMT+08:00 | Refresh- sCTV\CTV: Patients age is not calculated from DICOM tags automatically | WIP7 | WIP9.1 |  | Subsystem |  | NA | Fixed | sundeep_s | 0 | 3 | Console USP_EBW | iCAP | YES | Closed | WIP10SegReq | Brilliance Big Bore |  | WIP9.1 | August 28, 2018 6:35:57 PM GMT+08:00
CTUDT00598261 | Refresh | Closed | bagdady_b | May 29, 2018 9:31:18 PM GMT+08:00 | Refresh- sCTV\CTV: Patients age is not calculated from DICOM tags automatically | WIP7 | WIP9.1 |  | Subsystem |  | NA | Fixed | sundeep_s | 0 | 3 | Console USP_EBW | iCAP | YES | Closed | WIP10SegReq | Ingenuity 64 |  | WIP9.1 | August 28, 2018 6:35:57 PM GMT+08:00
CTUDT00598261 | Refresh | Closed | bagdady_b | May 29, 2018 9:31:18 PM GMT+08:00 | Refresh- sCTV\CTV: Patients age is not calculated from DICOM tags automatically | WIP7 | WIP9.1 |  | Subsystem |  | NA | Fixed | sundeep_s | 0 | 3 | Console USP_EBW | iCAP | YES | Closed | WIP10SegReq | IQon1 |  | WIP9.1 | August 28, 2018 6:35:57 PM GMT+08:00
CTUDT00598259 | Refresh | Closed | bagdady_b | May 29, 2018 9:24:45 PM GMT+08:00 | Refresh- Study date is changeable by using "Change Patient Details" | WIP7 | WIP10 |  | Subsystem | Patient Scene - study date | NA | Fixed | veerabhadraiah | 0 | 3 | Console USP_EBW | iCAP | YES | Closed | WIP10SegReq | Ovation |  | WIP10 | November 27, 2018 7:45:18 PM GMT+08:00
CTUDT00598259 | Refresh | Closed | bagdady_b | May 29, 2018 9:24:45 PM GMT+08:00 | Refresh- Study date is changeable by using "Change Patient Details" | WIP7 | WIP10 |  | Subsystem | Patient Scene - study date | NA | Fixed | veerabhadraiah | 0 | 3 | Console USP_EBW | iCAP | YES | Closed | WIP10SegReq | Brilliance 64 |  | WIP10 | November 27, 2018 7:45:18 PM GMT+08:00
CTUDT00598259 | Refresh | Closed | bagdady_b | May 29, 2018 9:24:45 PM GMT+08:00 | Refresh- Study date is changeable by using "Change Patient Details" | WIP7 | WIP10 |  | Subsystem | Patient Scene - study date | NA | Fixed | veerabhadraiah | 0 | 3 | Console USP_EBW | iCAP | YES | Closed | WIP10SegReq | Brilliance Big Bore |  | WIP10 | November 27, 2018 7:45:18 PM GMT+08:00
CTUDT00598259 | Refresh | Closed | bagdady_b | May 29, 2018 9:24:45 PM GMT+08:00 | Refresh- Study date is changeable by using "Change Patient Details" | WIP7 | WIP10 |  | Subsystem | Patient Scene - study date | NA | Fixed | veerabhadraiah | 0 | 3 | Console USP_EBW | iCAP | YES | Closed | WIP10SegReq | Ingenuity 64 |  | WIP10 | November 27, 2018 7:45:18 PM GMT+08:00
CTUDT00598259 | Refresh | Closed | bagdady_b | May 29, 2018 9:24:45 PM GMT+08:00 | Refresh- Study date is changeable by using "Change Patient Details" | WIP7 | WIP10 |  | Subsystem | Patient Scene - study date | NA | Fixed | veerabhadraiah | 0 | 3 | Console USP_EBW | iCAP | YES | Closed | WIP10SegReq | IQon1 |  | WIP10 | November 27, 2018 7:45:18 PM GMT+08:00
CTUDT00598257 | Refresh | Closed | rodan_d | May 28, 2018 8:52:15 PM GMT+08:00 | Calibrations: RFC: DMS calibrations shall not use EBW paths for its calibrations tables location | Eagle_WIP108 |  |  | Subsystem |  | N/A | Works As Specified / Not A Defect | gehtin_e | 0 |  | DMS_Calibrations | DMS | NO | Closed |  | IQon1
CTUDT00598256 | Refresh | Closed | rodan_d | May 28, 2018 8:50:51 PM GMT+08:00 | Queue Manager: CCT Single/Continuous scan resets the active recon job and resumes it with delay. | Eagle_WIP107 |  |  | Subsystem |  |  | Does Not Apply | sanjeeva_a | 0 | 4 | Console Generic | Console4 | NO | Closed |  | IQon1
CTUDT00598255 | Refresh | Closed | rodan_d | May 28, 2018 8:49:32 PM GMT+08:00 | calcium scoring scan- the phase % is not displayed in the series list. | Eagle_WIP109 |  | phase is visible in Show All and in parameters tab | Subsystem |  |  | Change Proposal | sanjeeva_a | 0 | 4 | Console Generic | Console4 | NO | Closed |  | IQon1
CTUDT00598254 | Refresh | Closed | rodan_d | May 28, 2018 8:48:18 PM GMT+08:00 | IV: The message that appears when trying to lock the system during Short Tube Conditioning (STC) should state "Quality Assurance tests are being performed please wait." | Eagle_WIP111 |  |  | Subsystem |  |  | Change Proposal | cumpson | 0 |  | DMS_Calibrations | GCX | NO | Closed |  | IQon1
CTUDT00598253 | Refresh | Closed | rodan_d | May 28, 2018 8:46:56 PM GMT+08:00 | IV: Ring artifacts in 2D head/Brain scans | Eagle_WIP112.1 |  | N/A | Subsystem |  | N/A | Change Proposal | soni_n | 0 | 4 | Physics | Support-System | NO | Closed | SubNC | IQon1
CTUDT00598252 | Refresh | Closed | rodan_d | May 28, 2018 8:45:11 PM GMT+08:00 | O-level can't change CIRS S3 IP address after reinstall | Eagle_WIP113.2 |  | Workaround is documented in Service manual for olevel installation (459801189421) section 2.2.3 - Setting the CIRS S3 Server IP Address. | Subsystem |  |  | Change Proposal | rubin_y | 0 | 3 | Documentation | Support-Service | NO | Closed |  | IQon1
CTUDT00598251 | Refresh | Closed | rodan_d | May 28, 2018 8:43:36 PM GMT+08:00 | spectral CT viewer cannot create spectral results from a 4 slices SBI | Eagle_WIP113.2 | WIP6 |  | Subsystem |  | NA | Fixed | aviw | 0 | 2 | sCTV | Support_Spectral | NO | Closed | SubCom | IQon1 |  | WIP6 | April 29, 2019 8:20:37 PM GMT+08:00
CTUDT00598249 | Refresh | Closed | rodan_d | May 28, 2018 8:41:50 PM GMT+08:00 | Iqon : The log message for displaying the incoming CCT fluoro images on the viewer is not logged in the log viewer | Eagle_WIP202 | DHF318051 Rev 05 Test Protocol |  | Subsystem |  | Console SSRS: 11572 | Fixed | babu_s | 0 | 4 | Test Design | Console4 | NO | Closed | WIP16SubCom | Ovation |  | DHF318051 Rev 05 Test Protocol | May 9, 2019 9:16:12 PM GMT+08:00
CTUDT00598249 | Refresh | Closed | rodan_d | May 28, 2018 8:41:50 PM GMT+08:00 | Iqon : The log message for displaying the incoming CCT fluoro images on the viewer is not logged in the log viewer | Eagle_WIP202 | DHF318051 Rev 05 Test Protocol |  | Subsystem |  | Console SSRS: 11572 | Fixed | babu_s | 0 | 4 | Test Design | Console4 | NO | Closed | WIP16SubCom | Brilliance 64 |  | DHF318051 Rev 05 Test Protocol | May 9, 2019 9:16:12 PM GMT+08:00
CTUDT00598249 | Refresh | Closed | rodan_d | May 28, 2018 8:41:50 PM GMT+08:00 | Iqon : The log message for displaying the incoming CCT fluoro images on the viewer is not logged in the log viewer | Eagle_WIP202 | DHF318051 Rev 05 Test Protocol |  | Subsystem |  | Console SSRS: 11572 | Fixed | babu_s | 0 | 4 | Test Design | Console4 | NO | Closed | WIP16SubCom | Brilliance Big Bore |  | DHF318051 Rev 05 Test Protocol | May 9, 2019 9:16:12 PM GMT+08:00
CTUDT00598249 | Refresh | Closed | rodan_d | May 28, 2018 8:41:50 PM GMT+08:00 | Iqon : The log message for displaying the incoming CCT fluoro images on the viewer is not logged in the log viewer | Eagle_WIP202 | DHF318051 Rev 05 Test Protocol |  | Subsystem |  | Console SSRS: 11572 | Fixed | babu_s | 0 | 4 | Test Design | Console4 | NO | Closed | WIP16SubCom | Ingenuity 64 |  | DHF318051 Rev 05 Test Protocol | May 9, 2019 9:16:12 PM GMT+08:00
CTUDT00598249 | Refresh | Closed | rodan_d | May 28, 2018 8:41:50 PM GMT+08:00 | Iqon : The log message for displaying the incoming CCT fluoro images on the viewer is not logged in the log viewer | Eagle_WIP202 | DHF318051 Rev 05 Test Protocol |  | Subsystem |  | Console SSRS: 11572 | Fixed | babu_s | 0 | 4 | Test Design | Console4 | NO | Closed | WIP16SubCom | IQon1 |  | DHF318051 Rev 05 Test Protocol | May 9, 2019 9:16:12 PM GMT+08:00
CTUDT00598248 | Refresh | Closed | rodan_d | May 28, 2018 8:40:11 PM GMT+08:00 | RFC:Upgrade McAfee Application control to version 7.0.1 | Eagle_WIP206 |  |  | Subsystem |  |  | Does Not Apply | sanjeeva_a | 0 |  | Security | Console4 | NO | Closed |  | IQon1
CTUDT00598247 | Refresh | Closed | rodan_d | May 28, 2018 8:38:24 PM GMT+08:00 | RFC: CeaserFTP should be replaced with new FTP product | Eagle_WIP206 | WIP17 BB_WIP7 |  | Subsystem |  | N/A | Fixed | dasgupta_s | 0 | 2 | Installation | Console4 | NO | Closed | SubCom | Ovation |  | WIP17 BB_WIP7 | May 21, 2019 7:12:38 PM GMT+08:00
CTUDT00598247 | Refresh | Closed | rodan_d | May 28, 2018 8:38:24 PM GMT+08:00 | RFC: CeaserFTP should be replaced with new FTP product | Eagle_WIP206 | WIP17 BB_WIP7 |  | Subsystem |  | N/A | Fixed | dasgupta_s | 0 | 2 | Installation | Console4 | NO | Closed | SubCom | Brilliance 64 |  | WIP17 BB_WIP7 | May 21, 2019 7:12:38 PM GMT+08:00
CTUDT00598247 | Refresh | Closed | rodan_d | May 28, 2018 8:38:24 PM GMT+08:00 | RFC: CeaserFTP should be replaced with new FTP product | Eagle_WIP206 | WIP17 BB_WIP7 |  | Subsystem |  | N/A | Fixed | dasgupta_s | 0 | 2 | Installation | Console4 | NO | Closed | SubCom | Brilliance Big Bore |  | WIP17 BB_WIP7 | May 21, 2019 7:12:38 PM GMT+08:00
CTUDT00598247 | Refresh | Closed | rodan_d | May 28, 2018 8:38:24 PM GMT+08:00 | RFC: CeaserFTP should be replaced with new FTP product | Eagle_WIP206 | WIP17 BB_WIP7 |  | Subsystem |  | N/A | Fixed | dasgupta_s | 0 | 2 | Installation | Console4 | NO | Closed | SubCom | Ingenuity 64 |  | WIP17 BB_WIP7 | May 21, 2019 7:12:38 PM GMT+08:00
CTUDT00598247 | Refresh | Closed | rodan_d | May 28, 2018 8:38:24 PM GMT+08:00 | RFC: CeaserFTP should be replaced with new FTP product | Eagle_WIP206 | WIP17 BB_WIP7 |  | Subsystem |  | N/A | Fixed | dasgupta_s | 0 | 2 | Installation | Console4 | NO | Closed | SubCom | IQon1 |  | WIP17 BB_WIP7 | May 21, 2019 7:12:38 PM GMT+08:00
CTUDT00598246 | Refresh | Closed | rodan_d | May 28, 2018 8:35:49 PM GMT+08:00 | RFC: Requirements: Add a tolerance to the preview Zoom/Pan operations | Eagle_WIP206.1 |  |  | Subsystem |  |  | Change Proposal | oren_d | 0 | 4 | Requirements | Support-System | NO | Closed |  | IQon1
CTUDT00598245 | Refresh | Closed | rodan_d | May 28, 2018 8:34:03 PM GMT+08:00 | RFC: SRS Requirements for O-Level | Eagle_WIP206.1 |  |  | Subsystem |  |  | Change Proposal | oren_d | 0 |  |  | Support-System | NO | Closed |  | IQon1
CTUDT00598244 | Refresh | Closed | rodan_d | May 28, 2018 8:32:00 PM GMT+08:00 | Message regarding failure of silent spectral mix calibration | Eagle_WIP207.1 |  |  | Subsystem |  | N/A | Does Not Apply | braunstein_d | 0 |  | DMS_Calibrations | DMS | NO | Closed |  | IQon1
CTUDT00598243 | Refresh | Closed | rodan_d | May 28, 2018 8:30:13 PM GMT+08:00 | IV: IQon Serviceability Cycle 2 - Remote Connection Severed when Logging In/Out | Eagle_WIP207.1 |  |  | Subsystem |  |  | Change Proposal | sanjeeva_a | 0 |  | Service Tools | Service Tools | NO | Closed |  | IQon1
CTUDT00598242 | Refresh | Closed | rodan_d | May 28, 2018 8:24:28 PM GMT+08:00 | IQ clinical evaluation- in CTA neck UB vs. B demonstrated noticeable difference in noise level, between LD to VD version. | Eagle_WIP206.1 |  | 1. In Low MonoE iamges - high noise can be expected (as describd in the IFU) 2. For neck studies (no brain) - use the exam type of neck and use B filter; for Brain+Neck - use UB filter. 3. Use higher spectral level reconstruction in case of UB filters with Head+Neck. | Subsystem |  | N/A | Change Proposal | soni_n | 0 |  | Image Quality | Support-System | NO | Closed | SubNC | IQon1
CTUDT00598241 | Refresh | Closed | rodan_d | May 28, 2018 8:18:22 PM GMT+08:00 | IV: IQon - Exam summary fails to include injection summary if injection is paused & study is ended on system GUI prior to ending on the injector | Eagle_WIP112.1 |  |  | Subsystem |  |  | Change Proposal | sanjeeva_a | 0 | 3 | Console Generic | Console4 | NO | Closed |  | IQon1
CTUDT00598240 | Refresh | Closed | rodan_d | May 28, 2018 8:15:53 PM GMT+08:00 | IV: IQon – IFU: Ch. 6 – Icon Images for Directory Manager Functions Missing | Eagle_WIP112 | WIP16 |  | Subsystem |  | NA | Fixed | rajalakshmi_j | 0 | 3 | User Documentation | Support-System | NO | Closed | WIP10SubCom | IQon1 |  | WIP16 | May 29, 2019 1:29:33 PM GMT+08:00
CTUDT00598239 | Refresh | Closed | rodan_d | May 28, 2018 8:11:09 PM GMT+08:00 | Locators and tracker scans may cause background recons to become temporarily suspended | Eagle_WIP206 |  |  | Subsystem |  |  | Does Not Apply | sanjeeva_a | 0 |  | Console Generic | Console4 | NO | Closed |  | IQon1
CTUDT00598238 | Refresh | Closed | rodan_d | May 28, 2018 8:06:22 PM GMT+08:00 | IV: IQon Spectral CT IFU Chapter 7 Subsection CT and Cardiac Viewer- Slab Mode | Eagle_WIP207.1 | WIP16 |  | Subsystem |  | N/A | Fixed | rajalakshmi_j | 0 | 3 | User Documentation | Support-System | NO | Closed | WIP10SubCom | Ovation |  | WIP16 | May 29, 2019 2:01:51 PM GMT+08:00
CTUDT00598238 | Refresh | Closed | rodan_d | May 28, 2018 8:06:22 PM GMT+08:00 | IV: IQon Spectral CT IFU Chapter 7 Subsection CT and Cardiac Viewer- Slab Mode | Eagle_WIP207.1 | WIP16 |  | Subsystem |  | N/A | Fixed | rajalakshmi_j | 0 | 3 | User Documentation | Support-System | NO | Closed | WIP10SubCom | IQon1 |  | WIP16 | May 29, 2019 2:01:51 PM GMT+08:00
CTUDT00598234 | Refresh | Closed | pB_s | May 28, 2018 7:55:44 PM GMT+08:00 | The Console does not modify the TILT value properly to the closest valid value, when the couch height changes. | WIP6.1 | WIP10 |  | Subsystem | Tilt | N/A | Fixed | ramesh_t | 0 | 2 | Test Design | Console4 | NO | Closed | WIP9SegReq | Ingenuity 64 |  | WIP10 | November 29, 2018 4:44:07 PM GMT+08:00
CTUDT00598232 | Refresh | Closed | rodan_d | May 28, 2018 4:43:11 PM GMT+08:00 | iCT 4.1.6: Deleting a ReviewBeforeSave result in series list does not update in Patient Directory But gets updated in series list. | WIP 102 |  |  | Subsystem |  |  | Change Proposal | sanjeeva_a | 0 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00598231 | Refresh | Closed | rodan_d | May 28, 2018 4:41:52 PM GMT+08:00 | 2.5mm Recon thickness is not available with Body Perfusion scantype in iCT 4.1.3 | 4.1.3 |  |  | Subsystem | History Review |  | Change Proposal | sanjeeva_a | 0 |  |  | Console4 | NO | Closed |  | Ovation | PR4805065
CTUDT00598230 | Refresh | Closed | rodan_d | May 28, 2018 4:39:27 PM GMT+08:00 | RfVD: Ansan: Artifact for iQon system in Korea | IQon-RfVD | CIRS-5.1.0.0209 WIP30 | N/A | Subsystem | SP1 | N/A | Fixed | kumpf_c | 0 | 2 | CIRS_SW_ImgNodeSvr | Support-Physics | NO | Closed | SP1 | IQon1 | CFR#9427137 | CIRS-5.1.0.0209 WIP30 | June 4, 2020 2:41:37 PM GMT+08:00
CTUDT00598229 | Refresh | Closed | rodan_d | May 28, 2018 4:37:23 PM GMT+08:00 | Volume display CCT | 4.2.0.17xxx-WIP207 | WIP17 |  | Subsystem |  | Console-SRS.10371 | Fixed | bandil_j | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SubCom | Ovation |  | WIP17 | May 17, 2019 7:31:44 PM GMT+08:00
CTUDT00598229 | Refresh | Closed | rodan_d | May 28, 2018 4:37:23 PM GMT+08:00 | Volume display CCT | 4.2.0.17xxx-WIP207 | WIP17 |  | Subsystem |  | Console-SRS.10371 | Fixed | bandil_j | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SubCom | Brilliance 64 |  | WIP17 | May 17, 2019 7:31:44 PM GMT+08:00
CTUDT00598229 | Refresh | Closed | rodan_d | May 28, 2018 4:37:23 PM GMT+08:00 | Volume display CCT | 4.2.0.17xxx-WIP207 | WIP17 |  | Subsystem |  | Console-SRS.10371 | Fixed | bandil_j | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SubCom | Ingenuity 64 |  | WIP17 | May 17, 2019 7:31:44 PM GMT+08:00
CTUDT00598229 | Refresh | Closed | rodan_d | May 28, 2018 4:37:23 PM GMT+08:00 | Volume display CCT | 4.2.0.17xxx-WIP207 | WIP17 |  | Subsystem |  | Console-SRS.10371 | Fixed | bandil_j | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SubCom | IQon1 |  | WIP17 | May 17, 2019 7:31:44 PM GMT+08:00
CTUDT00598228 | Refresh | Closed | rodan_d | May 28, 2018 4:35:56 PM GMT+08:00 | IQon HB: Need to add a support for ISP6.5.3 | WIP507 |  |  | Subsystem |  |  | Change Proposal | mikia | 0 |  |  | Support_Spectral | NO | Closed |  | IQon1
CTUDT00598227 | Refresh | Closed | rodan_d | May 28, 2018 4:34:26 PM GMT+08:00 | Wrong classification for bone and uric acid on 6 core machine | WIP507 |  | Not needed becaus et will never happen on the machines suppiled as consoles for the scanner. | Subsystem |  | N/A | No Plan to Fix | aviw | 0 | 4 |  | Support_Spectral | NO | Closed |  | IQon1
CTUDT00598226 | Refresh | Closed | rodan_d | May 28, 2018 4:33:04 PM GMT+08:00 | TRG section 6 User information/Essential Performance for Interventional Imaging, Content is not related to interventional imaging | WIP507 |  |  | Subsystem |  |  | Change Proposal | rajalakshmi_j | 0 | 2 | User Documentation | Support-System | NO | Closed |  | IQon1
CTUDT00598225 | Refresh | Closed | rodan_d | May 28, 2018 4:31:02 PM GMT+08:00 | Provide instructions how to calibrate a single Gantry Touch Panel instead of two Touch Panels | IQon-RfVD |  | Replace failed display/touch panel. Impacts customer availability. | Subsystem |  |  | Change Proposal | cumpson | 0 | 3 |  | GCX | NO | Closed |  | IQon1
CTUDT00598224 | Refresh | Closed | rodan_d | May 28, 2018 4:13:27 PM GMT+08:00 | IQon HB1: Update the requirement CT-NM.SRS.Service.354 - Bugrep collection timeout | WIP509.2 |  |  | Subsystem |  |  | Change Proposal | oren_d | 0 |  |  | Support-System | NO | Closed |  | IQon1
CTUDT00598220 | Refresh | Closed | rodan_d | May 28, 2018 4:00:27 PM GMT+08:00 | Found by IOCC : “Anonymize" application is not compliant. | WIP507 | WIP22.2 |  | Subsystem | Blocker | NA | Fixed | srushti_k | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | SubCom_WIP8 | Ovation |  | WIP22.2 | November 29, 2019 3:22:35 PM GMT+08:00
CTUDT00598220 | Refresh | Closed | rodan_d | May 28, 2018 4:00:27 PM GMT+08:00 | Found by IOCC : “Anonymize" application is not compliant. | WIP507 | WIP22.2 |  | Subsystem | Blocker | NA | Fixed | srushti_k | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | SubCom_WIP8 | Brilliance 64 |  | WIP22.2 | November 29, 2019 3:22:35 PM GMT+08:00
CTUDT00598220 | Refresh | Closed | rodan_d | May 28, 2018 4:00:27 PM GMT+08:00 | Found by IOCC : “Anonymize" application is not compliant. | WIP507 | WIP22.2 |  | Subsystem | Blocker | NA | Fixed | srushti_k | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | SubCom_WIP8 | Brilliance Big Bore |  | WIP22.2 | November 29, 2019 3:22:35 PM GMT+08:00
CTUDT00598220 | Refresh | Closed | rodan_d | May 28, 2018 4:00:27 PM GMT+08:00 | Found by IOCC : “Anonymize" application is not compliant. | WIP507 | WIP22.2 |  | Subsystem | Blocker | NA | Fixed | srushti_k | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | SubCom_WIP8 | Ingenuity 64 |  | WIP22.2 | November 29, 2019 3:22:35 PM GMT+08:00
CTUDT00598220 | Refresh | Closed | rodan_d | May 28, 2018 4:00:27 PM GMT+08:00 | Found by IOCC : “Anonymize" application is not compliant. | WIP507 | WIP22.2 |  | Subsystem | Blocker | NA | Fixed | srushti_k | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | SubCom_WIP8 | IQon1 |  | WIP22.2 | November 29, 2019 3:22:35 PM GMT+08:00
CTUDT00598217 | Refresh | Closed | rodan_d | May 28, 2018 3:50:36 PM GMT+08:00 | Reference Jog Brain Perfusion Exam Card needs to be updated. | 4.1.3 and 4.1.6 |  |  | Subsystem |  |  | Change Proposal | sanjeeva_a | 0 |  |  | Support-System | NO | Closed |  | Ovation | PR5842445
CTUDT00598217 | Refresh | Closed | rodan_d | May 28, 2018 3:50:36 PM GMT+08:00 | Reference Jog Brain Perfusion Exam Card needs to be updated. | 4.1.3 and 4.1.6 |  |  | Subsystem |  |  | Change Proposal | sanjeeva_a | 0 |  |  | Support-System | NO | Closed |  | IQon1 | PR5842445
CTUDT00598212 | Refresh | Closed | sengupta_a | May 28, 2018 3:12:59 PM GMT+08:00 | Console.SW.Host displays incorrect value in Console tab in Analysis tab in Service tools | WIP6.1 | WIP 11 |  | Subsystem | System Configuration of Console | N/A | Fixed | jindal_k | 0 | 2 | Test Design | Console4 | NO | Closed | WIP9.1SegReq | Ingenuity 64 |  | WIP 11 | November 29, 2018 7:12:32 PM GMT+08:00
CTUDT00598206 | Refresh | Closed | lin_s | May 27, 2018 10:46:25 PM GMT+08:00 | IOCC: Required template information is missing in created DOSE SR report. | 4.2.x |  | Fill in label if the comment is needed. | Subsystem | Medium_3 | N\A | Duplicate | kaushik_u | 0 | 2 | Console Generic | Console4 | NO | Closed | WIP13SegReq | Ovation
CTUDT00598205 | Refresh | Closed | lin_s | May 27, 2018 10:44:23 PM GMT+08:00 | IQon: Patient details are saved into the bugrep when trying to launch the View2 for the active patient | 4.2.x |  |  | Subsystem |  |  | Duplicate | sanjeeva_a | 0 | 2 | Console Generic | Console4 | NO | Closed |  | Brilliance 64
CTUDT00598205 | Refresh | Closed | lin_s | May 27, 2018 10:44:23 PM GMT+08:00 | IQon: Patient details are saved into the bugrep when trying to launch the View2 for the active patient | 4.2.x |  |  | Subsystem |  |  | Duplicate | sanjeeva_a | 0 | 2 | Console Generic | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00598205 | Refresh | Closed | lin_s | May 27, 2018 10:44:23 PM GMT+08:00 | IQon: Patient details are saved into the bugrep when trying to launch the View2 for the active patient | 4.2.x |  |  | Subsystem |  |  | Duplicate | sanjeeva_a | 0 | 2 | Console Generic | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00598204 | Refresh | Closed | lin_s | May 27, 2018 10:42:46 PM GMT+08:00 | IV: 4.1.6 - "No Bellows connection to patient" in waveform window when Varian RPM device is connected and functional | 4.2.x | WIP17 |  | Subsystem |  | N/A | Fixed | sowmya_k | 0 | 3 | Console Software | Console4 | NO | Closed | WIP17SubCom | Brilliance 64 |  | WIP17 | May 30, 2019 8:57:11 PM GMT+08:00
CTUDT00598204 | Refresh | Closed | lin_s | May 27, 2018 10:42:46 PM GMT+08:00 | IV: 4.1.6 - "No Bellows connection to patient" in waveform window when Varian RPM device is connected and functional | 4.2.x | WIP17 |  | Subsystem |  | N/A | Fixed | sowmya_k | 0 | 3 | Console Software | Console4 | NO | Closed | WIP17SubCom | Ingenuity 64 |  | WIP17 | May 30, 2019 8:57:11 PM GMT+08:00
CTUDT00598202 | Refresh | Closed | lin_s | May 27, 2018 10:40:58 PM GMT+08:00 | IV: Check Box for Continue Previous Exam in Patient Entry | 4.2.x |  | There is no workaround for this issue. | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 | Console Generic | Console4 | NO | Closed |  | Brilliance 64
CTUDT00598202 | Refresh | Closed | lin_s | May 27, 2018 10:40:58 PM GMT+08:00 | IV: Check Box for Continue Previous Exam in Patient Entry | 4.2.x |  | There is no workaround for this issue. | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 | Console Generic | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00598201 | Refresh | Closed | lin_s | May 27, 2018 10:36:22 PM GMT+08:00 | License key .bat file for new iCT product installation will fail to run with pre-installed 4.1.6 software. | 4.2.x |  |  | Subsystem |  |  | Does Not Apply | sanjeeva_a | 0 | 2 |  | Service Tools | NO | Closed |  | Ovation
CTUDT00598201 | Refresh | Closed | lin_s | May 27, 2018 10:36:22 PM GMT+08:00 | License key .bat file for new iCT product installation will fail to run with pre-installed 4.1.6 software. | 4.2.x |  |  | Subsystem |  |  | Does Not Apply | sanjeeva_a | 0 | 2 |  | Service Tools | NO | Closed |  | Brilliance Big Bore
CTUDT00598200 | Refresh | Closed | lin_s | May 27, 2018 10:34:15 PM GMT+08:00 | Localization:Step & Shoot: High HR warning message in Dutch is cut. | 4.2.x |  | There is no workaround for this issue and none is required as the popup is readable even though the last word is cutoff in the vertical direction. | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 | Console Generic | Console4 | NO | Closed |  | Ovation
CTUDT00598200 | Refresh | Closed | lin_s | May 27, 2018 10:34:15 PM GMT+08:00 | Localization:Step & Shoot: High HR warning message in Dutch is cut. | 4.2.x |  | There is no workaround for this issue and none is required as the popup is readable even though the last word is cutoff in the vertical direction. | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 | Console Generic | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00598199 | Refresh | Closed | kadish_n | May 27, 2018 10:32:59 PM GMT+08:00 | No permissions to retrieve Tube-Heat from GLC | WIP6.1 |  |  | Subsystem | Tube Heat Tool |  | Works As Specified / Not A Defect | gehtin_e | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00598196 | Refresh | Closed | lin_s | May 27, 2018 10:31:06 PM GMT+08:00 | Not all DICOM Structured Reports are Accepted at Destination Node due to Missing TextValue (0040,a160) | 4.2.x | BB_WIP3 |  | Subsystem |  | N/A | Fixed | v_sumanth | 0 | 2 | Console Software | Console4 | NO | Closed | WIP12SubCom | Brilliance Big Bore | CFR9420595/PR6189708 | BB_WIP3 | December 6, 2018 5:06:00 PM GMT+08:00
CTUDT00598195 | Refresh | Closed | lin_s | May 27, 2018 10:29:25 PM GMT+08:00 | Out of range thickness read from a user CCT volume exam card treated as valid | 4.2.x | WIP12 |  | System | High_1 | NA | Fixed | kumar_v | 0 | 2 | Console Software | Console4 | YES | Closed | WIP12SubCom | Ovation | PR6113361 | WIP12 | December 6, 2018 1:34:57 PM GMT+08:00
CTUDT00598195 | Refresh | Closed | lin_s | May 27, 2018 10:29:25 PM GMT+08:00 | Out of range thickness read from a user CCT volume exam card treated as valid | 4.2.x | WIP12 |  | System | High_1 | NA | Fixed | kumar_v | 0 | 2 | Console Software | Console4 | YES | Closed | WIP12SubCom | Brilliance 64 | PR6113361 | WIP12 | December 6, 2018 1:34:57 PM GMT+08:00
CTUDT00598195 | Refresh | Closed | lin_s | May 27, 2018 10:29:25 PM GMT+08:00 | Out of range thickness read from a user CCT volume exam card treated as valid | 4.2.x | WIP12 |  | System | High_1 | NA | Fixed | kumar_v | 0 | 2 | Console Software | Console4 | YES | Closed | WIP12SubCom | Ingenuity 64 | PR6113361 | WIP12 | December 6, 2018 1:34:57 PM GMT+08:00
CTUDT00598195 | Refresh | Closed | lin_s | May 27, 2018 10:29:25 PM GMT+08:00 | Out of range thickness read from a user CCT volume exam card treated as valid | 4.2.x | WIP12 |  | System | High_1 | NA | Fixed | kumar_v | 0 | 2 | Console Software | Console4 | YES | Closed | WIP12SubCom | IQon1 | PR6113361 | WIP12 | December 6, 2018 1:34:57 PM GMT+08:00
CTUDT00598194 | Refresh | Closed | lin_s | May 27, 2018 10:27:29 PM GMT+08:00 | Pre-IV: iCT - Patient directory contains duplicate coronal results | 4.2.x |  | The user can delete the duplicate series. the delete operation is defined in the IFU in several places, one of them is in section "Patient List Functions". | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 3 |  | Console4 | NO | Closed |  | Ovation
CTUDT00598194 | Refresh | Closed | lin_s | May 27, 2018 10:27:29 PM GMT+08:00 | Pre-IV: iCT - Patient directory contains duplicate coronal results | 4.2.x |  | The user can delete the duplicate series. the delete operation is defined in the IFU in several places, one of them is in section "Patient List Functions". | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 3 |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00598193 | Refresh | Closed | lin_s | May 27, 2018 10:25:48 PM GMT+08:00 | Real Time Data Tool Filament Current always reads 65535 | 4.2.x | WIP17 BB_WIP7 | No possible workaround. | System |  | NA | Fixed | v_su | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP17SubCom | Brilliance 64 |  | WIP17 BB_WIP7 | May 9, 2019 3:44:13 PM GMT+08:00
CTUDT00598193 | Refresh | Closed | lin_s | May 27, 2018 10:25:48 PM GMT+08:00 | Real Time Data Tool Filament Current always reads 65535 | 4.2.x | WIP17 BB_WIP7 | No possible workaround. | System |  | NA | Fixed | v_su | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP17SubCom | Brilliance Big Bore |  | WIP17 BB_WIP7 | May 9, 2019 3:44:13 PM GMT+08:00
CTUDT00598193 | Refresh | Closed | lin_s | May 27, 2018 10:25:48 PM GMT+08:00 | Real Time Data Tool Filament Current always reads 65535 | 4.2.x | WIP17 BB_WIP7 | No possible workaround. | System |  | NA | Fixed | v_su | 0 | 2 | Service Tools_Software | Service Tools | NO | Closed | WIP17SubCom | Ingenuity 64 |  | WIP17 BB_WIP7 | May 9, 2019 3:44:13 PM GMT+08:00
CTUDT00598192 | Refresh | Closed | lin_s | May 27, 2018 10:23:09 PM GMT+08:00 | Surview length not set when tab is pressed to move to next parameter field | 4.2.x |  | Yes, Press Enter in the length Box instead of Tab. | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 3 |  | Console4 | YES | Closed |  | Ovation
CTUDT00598192 | Refresh | Closed | lin_s | May 27, 2018 10:23:09 PM GMT+08:00 | Surview length not set when tab is pressed to move to next parameter field | 4.2.x |  | Yes, Press Enter in the length Box instead of Tab. | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 3 |  | Console4 | YES | Closed |  | Brilliance Big Bore
CTUDT00598191 | Refresh | Closed | lin_s | May 27, 2018 10:20:31 PM GMT+08:00 | Surview present in Exam summary series is not of the current patient. | 4.2.x | WIP12 |  | System | High_1 | NA | Fixed | datta_i | 0 | 2 | Console-Console4 | Console4 | NO | Closed | WIP12SubCom | Ovation | PR7252944 ESC# 0109806936 ESC# 0109673916 | WIP12 | December 4, 2018 2:17:54 PM GMT+08:00
CTUDT00598191 | Refresh | Closed | lin_s | May 27, 2018 10:20:31 PM GMT+08:00 | Surview present in Exam summary series is not of the current patient. | 4.2.x | WIP12 |  | System | High_1 | NA | Fixed | datta_i | 0 | 2 | Console-Console4 | Console4 | NO | Closed | WIP12SubCom | Brilliance 64 | PR7252944 ESC# 0109806936 ESC# 0109673916 | WIP12 | December 4, 2018 2:17:54 PM GMT+08:00
CTUDT00598191 | Refresh | Closed | lin_s | May 27, 2018 10:20:31 PM GMT+08:00 | Surview present in Exam summary series is not of the current patient. | 4.2.x | WIP12 |  | System | High_1 | NA | Fixed | datta_i | 0 | 2 | Console-Console4 | Console4 | NO | Closed | WIP12SubCom | Ingenuity 64 | PR7252944 ESC# 0109806936 ESC# 0109673916 | WIP12 | December 4, 2018 2:17:54 PM GMT+08:00
CTUDT00598191 | Refresh | Closed | lin_s | May 27, 2018 10:20:31 PM GMT+08:00 | Surview present in Exam summary series is not of the current patient. | 4.2.x | WIP12 |  | System | High_1 | NA | Fixed | datta_i | 0 | 2 | Console-Console4 | Console4 | NO | Closed | WIP12SubCom | IQon1 | PR7252944 ESC# 0109806936 ESC# 0109673916 | WIP12 | December 4, 2018 2:17:54 PM GMT+08:00
CTUDT00598190 | Refresh | Closed | lin_s | May 27, 2018 10:16:21 PM GMT+08:00 | Untag requirement CT-NM.SRS.Service.1128 | 4.2.x |  |  | Subsystem |  |  | Does Not Apply | oren_d | 0 | 3 |  | Support-System | NO | Closed |  | Brilliance 64
CTUDT00598190 | Refresh | Closed | lin_s | May 27, 2018 10:16:21 PM GMT+08:00 | Untag requirement CT-NM.SRS.Service.1128 | 4.2.x |  |  | Subsystem |  |  | Does Not Apply | oren_d | 0 | 3 |  | Support-System | NO | Closed |  | Brilliance Big Bore
CTUDT00598190 | Refresh | Closed | lin_s | May 27, 2018 10:16:21 PM GMT+08:00 | Untag requirement CT-NM.SRS.Service.1128 | 4.2.x |  |  | Subsystem |  |  | Does Not Apply | oren_d | 0 | 3 |  | Support-System | NO | Closed |  | Ingenuity 64
CTUDT00598189 | Refresh | Closed | lin_s | May 27, 2018 10:14:32 PM GMT+08:00 | User cannot change the priority of “Body Perfusion Helical” offline-recon | 4.2.x | WIP16 | NA | Subsystem |  | Console-SRS.9971 | Fixed | sagar_v | 0 | 2 | Console Generic | Console4 | NO | Closed | WIP16SubCom | Ovation |  | WIP16 | April 15, 2019 8:23:32 PM GMT+08:00
CTUDT00598189 | Refresh | Closed | lin_s | May 27, 2018 10:14:32 PM GMT+08:00 | User cannot change the priority of “Body Perfusion Helical” offline-recon | 4.2.x | WIP16 | NA | Subsystem |  | Console-SRS.9971 | Fixed | sagar_v | 0 | 2 | Console Generic | Console4 | NO | Closed | WIP16SubCom | Brilliance 64 |  | WIP16 | April 15, 2019 8:23:32 PM GMT+08:00
CTUDT00598189 | Refresh | Closed | lin_s | May 27, 2018 10:14:32 PM GMT+08:00 | User cannot change the priority of “Body Perfusion Helical” offline-recon | 4.2.x | WIP16 | NA | Subsystem |  | Console-SRS.9971 | Fixed | sagar_v | 0 | 2 | Console Generic | Console4 | NO | Closed | WIP16SubCom | Brilliance Big Bore |  | WIP16 | April 15, 2019 8:23:32 PM GMT+08:00
CTUDT00598189 | Refresh | Closed | lin_s | May 27, 2018 10:14:32 PM GMT+08:00 | User cannot change the priority of “Body Perfusion Helical” offline-recon | 4.2.x | WIP16 | NA | Subsystem |  | Console-SRS.9971 | Fixed | sagar_v | 0 | 2 | Console Generic | Console4 | NO | Closed | WIP16SubCom | Ingenuity 64 |  | WIP16 | April 15, 2019 8:23:32 PM GMT+08:00
CTUDT00598189 | Refresh | Closed | lin_s | May 27, 2018 10:14:32 PM GMT+08:00 | User cannot change the priority of “Body Perfusion Helical” offline-recon | 4.2.x | WIP16 | NA | Subsystem |  | Console-SRS.9971 | Fixed | sagar_v | 0 | 2 | Console Generic | Console4 | NO | Closed | WIP16SubCom | IQon1 |  | WIP16 | April 15, 2019 8:23:32 PM GMT+08:00
CTUDT00598188 | Refresh | Closed | lin_s | May 27, 2018 10:12:45 PM GMT+08:00 | V4.1 :Patient Page: First Name and Middle Name field is not getting empty/blank when clicked on New tab in service mode. | 4.2.x |  | Clear the name fields manually. | Subsystem |  | None | No Plan to Fix | sanjeeva_a | 0 | 4 | Console Generic | Console4 | NO | Closed |  | Ovation
CTUDT00598188 | Refresh | Closed | lin_s | May 27, 2018 10:12:45 PM GMT+08:00 | V4.1 :Patient Page: First Name and Middle Name field is not getting empty/blank when clicked on New tab in service mode. | 4.2.x |  | Clear the name fields manually. | Subsystem |  | None | No Plan to Fix | sanjeeva_a | 0 | 4 | Console Generic | Console4 | NO | Closed |  | Brilliance 64
CTUDT00598188 | Refresh | Closed | lin_s | May 27, 2018 10:12:45 PM GMT+08:00 | V4.1 :Patient Page: First Name and Middle Name field is not getting empty/blank when clicked on New tab in service mode. | 4.2.x |  | Clear the name fields manually. | Subsystem |  | None | No Plan to Fix | sanjeeva_a | 0 | 4 | Console Generic | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00598187 | Refresh | Closed | lin_s | May 27, 2018 10:10:14 PM GMT+08:00 | V4.1 ECM: Newly created exam cards are saved with same name after export and import of same ec. | 4.2.x |  |  | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 | Console Generic | Console4 | NO | Closed |  | Ovation
CTUDT00598187 | Refresh | Closed | lin_s | May 27, 2018 10:10:14 PM GMT+08:00 | V4.1 ECM: Newly created exam cards are saved with same name after export and import of same ec. | 4.2.x |  |  | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 | Console Generic | Console4 | NO | Closed |  | Brilliance 64
CTUDT00598187 | Refresh | Closed | lin_s | May 27, 2018 10:10:14 PM GMT+08:00 | V4.1 ECM: Newly created exam cards are saved with same name after export and import of same ec. | 4.2.x |  |  | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 | Console Generic | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00598186 | Refresh | Closed | lin_s | May 27, 2018 10:08:01 PM GMT+08:00 | v4.1.4 - Examcard selection not activated (grayed out) after the patient info imported from RIS. | 4.2.x |  | The workaround would be to enter the information manually. | Subsystem | History Review |  | Works As Specified / Not A Defect | babu_s | 1 | 3 |  | Console4 | NO | Closed |  | Brilliance 64 | 9442398
CTUDT00598186 | Refresh | Closed | lin_s | May 27, 2018 10:08:01 PM GMT+08:00 | v4.1.4 - Examcard selection not activated (grayed out) after the patient info imported from RIS. | 4.2.x |  | The workaround would be to enter the information manually. | Subsystem | History Review |  | Works As Specified / Not A Defect | babu_s | 1 | 3 |  | Console4 | NO | Closed |  | Brilliance Big Bore | 9442398
CTUDT00598186 | Refresh | Closed | lin_s | May 27, 2018 10:08:01 PM GMT+08:00 | v4.1.4 - Examcard selection not activated (grayed out) after the patient info imported from RIS. | 4.2.x |  | The workaround would be to enter the information manually. | Subsystem | History Review |  | Works As Specified / Not A Defect | babu_s | 1 | 3 |  | Console4 | NO | Closed |  | Ingenuity 64 | 9442398
CTUDT00598185 | Refresh | Closed | lin_s | May 27, 2018 10:06:02 PM GMT+08:00 | v4.1.4 FOK - Surview for Split Study matched to wrong images | 4.2.x |  | The operator could manually send the applicable surview to the PACS to be with the related results. Information regarding copying images to remote devices is provided in the iCT Instructions for Use Manual 459800604532 Rev C, Chapter 6 Using the Directory. | Subsystem |  |  | Works As Specified / Not A Defect | sanjeeva_a | 0 | 3 |  | Console4 | NO | Closed |  | Ovation | PR4967633
CTUDT00598185 | Refresh | Closed | lin_s | May 27, 2018 10:06:02 PM GMT+08:00 | v4.1.4 FOK - Surview for Split Study matched to wrong images | 4.2.x |  | The operator could manually send the applicable surview to the PACS to be with the related results. Information regarding copying images to remote devices is provided in the iCT Instructions for Use Manual 459800604532 Rev C, Chapter 6 Using the Directory. | Subsystem |  |  | Works As Specified / Not A Defect | sanjeeva_a | 0 | 3 |  | Console4 | NO | Closed |  | Brilliance Big Bore | PR4967633
CTUDT00598184 | Refresh | Closed | lin_s | May 27, 2018 10:03:51 PM GMT+08:00 | v4.1.6 -- Series duplication with "Review Before Save" | 4.2.x |  | Delete the Duplicate series from the PD, by selecting the series and do right click and select the delete option. | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 |  | Console4 | NO | Closed |  | Ovation | PR 5691537 PR 5709867
CTUDT00598184 | Refresh | Closed | lin_s | May 27, 2018 10:03:51 PM GMT+08:00 | v4.1.6 -- Series duplication with "Review Before Save" | 4.2.x |  | Delete the Duplicate series from the PD, by selecting the series and do right click and select the delete option. | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 |  | Console4 | NO | Closed |  | Brilliance 64 | PR 5691537 PR 5709867
CTUDT00598184 | Refresh | Closed | lin_s | May 27, 2018 10:03:51 PM GMT+08:00 | v4.1.6 -- Series duplication with "Review Before Save" | 4.2.x |  | Delete the Duplicate series from the PD, by selecting the series and do right click and select the delete option. | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 |  | Console4 | NO | Closed |  | Brilliance Big Bore | PR 5691537 PR 5709867
CTUDT00598184 | Refresh | Closed | lin_s | May 27, 2018 10:03:51 PM GMT+08:00 | v4.1.6 -- Series duplication with "Review Before Save" | 4.2.x |  | Delete the Duplicate series from the PD, by selecting the series and do right click and select the delete option. | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 |  | Console4 | NO | Closed |  | Ingenuity 64 | PR 5691537 PR 5709867
CTUDT00598183 | Refresh | Closed | lin_s | May 27, 2018 10:01:45 PM GMT+08:00 | V4.1.6 : Number of cycles or scans keep reducing while toggling between thickness in Axial 3D tilted scans | 4.2.x | WIP11.1 |  | Subsystem |  | N/A | Fixed | sowmya_k | 0 | 2 | Console Software | Console4 | YES | Closed | WIP11.1SegReq | Brilliance 64 |  | WIP11.1 | November 21, 2018 6:39:03 PM GMT+08:00
CTUDT00598183 | Refresh | Closed | lin_s | May 27, 2018 10:01:45 PM GMT+08:00 | V4.1.6 : Number of cycles or scans keep reducing while toggling between thickness in Axial 3D tilted scans | 4.2.x | WIP11.1 |  | Subsystem |  | N/A | Fixed | sowmya_k | 0 | 2 | Console Software | Console4 | YES | Closed | WIP11.1SegReq | Ingenuity 64 |  | WIP11.1 | November 21, 2018 6:39:03 PM GMT+08:00
CTUDT00598182 | Refresh | Closed | lin_s | May 27, 2018 9:58:39 PM GMT+08:00 | V4.1.6 : Scan increment can be changed to value of more than 100mm | 4.2.x |  |  | Subsystem |  |  | Works As Specified / Not A Defect | sanjeeva_a | 0 | 3 |  | Console4 | NO | Closed |  | Ovation
CTUDT00598182 | Refresh | Closed | lin_s | May 27, 2018 9:58:39 PM GMT+08:00 | V4.1.6 : Scan increment can be changed to value of more than 100mm | 4.2.x |  |  | Subsystem |  |  | Works As Specified / Not A Defect | sanjeeva_a | 0 | 3 |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00598182 | Refresh | Closed | lin_s | May 27, 2018 9:58:39 PM GMT+08:00 | V4.1.6 : Scan increment can be changed to value of more than 100mm | 4.2.x |  |  | Subsystem |  |  | Works As Specified / Not A Defect | sanjeeva_a | 0 | 3 |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00598182 | Refresh | Closed | lin_s | May 27, 2018 9:58:39 PM GMT+08:00 | V4.1.6 : Scan increment can be changed to value of more than 100mm | 4.2.x |  |  | Subsystem |  |  | Works As Specified / Not A Defect | sanjeeva_a | 0 | 3 |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00598181 | Refresh | Closed | lin_s | May 27, 2018 9:55:50 PM GMT+08:00 | v4.1.6 exe file access permission error | 4.2.x |  |  | Subsystem |  |  | Does Not Apply | prabhu_m | 0 | 3 |  | Support-Security | NO | Closed |  | Brilliance 64
CTUDT00598181 | Refresh | Closed | lin_s | May 27, 2018 9:55:50 PM GMT+08:00 | v4.1.6 exe file access permission error | 4.2.x |  |  | Subsystem |  |  | Does Not Apply | prabhu_m | 0 | 3 |  | Support-Security | NO | Closed |  | Brilliance Big Bore
CTUDT00598181 | Refresh | Closed | lin_s | May 27, 2018 9:55:50 PM GMT+08:00 | v4.1.6 exe file access permission error | 4.2.x |  |  | Subsystem |  |  | Does Not Apply | prabhu_m | 0 | 3 |  | Support-Security | NO | Closed |  | Ingenuity 64
CTUDT00598180 | Refresh | Closed | lin_s | May 27, 2018 9:52:22 PM GMT+08:00 | V4.1.6:Duplicate Thin slice series is created when merge series is enabled for 2 acquisitions. | 4.2.x |  | The user can merge the duplicate series using the CT Viewer. The IFU section called Merge Series describes how to do this: Defined groups of results can be saved into the same DICOM series. Select the appropriate group (1 - 5) for your result. All results in the same group are saved into a single series. The label field is used as a descriptor (the label is taken from the first result in the group). | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 3 |  | Console4 | NO | Closed |  | Ovation
CTUDT00598180 | Refresh | Closed | lin_s | May 27, 2018 9:52:22 PM GMT+08:00 | V4.1.6:Duplicate Thin slice series is created when merge series is enabled for 2 acquisitions. | 4.2.x |  | The user can merge the duplicate series using the CT Viewer. The IFU section called Merge Series describes how to do this: Defined groups of results can be saved into the same DICOM series. Select the appropriate group (1 - 5) for your result. All results in the same group are saved into a single series. The label field is used as a descriptor (the label is taken from the first result in the group). | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 3 |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00598180 | Refresh | Closed | lin_s | May 27, 2018 9:52:22 PM GMT+08:00 | V4.1.6:Duplicate Thin slice series is created when merge series is enabled for 2 acquisitions. | 4.2.x |  | The user can merge the duplicate series using the CT Viewer. The IFU section called Merge Series describes how to do this: Defined groups of results can be saved into the same DICOM series. Select the appropriate group (1 - 5) for your result. All results in the same group are saved into a single series. The label field is used as a descriptor (the label is taken from the first result in the group). | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 3 |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00598180 | Refresh | Closed | lin_s | May 27, 2018 9:52:22 PM GMT+08:00 | V4.1.6:Duplicate Thin slice series is created when merge series is enabled for 2 acquisitions. | 4.2.x |  | The user can merge the duplicate series using the CT Viewer. The IFU section called Merge Series describes how to do this: Defined groups of results can be saved into the same DICOM series. Select the appropriate group (1 - 5) for your result. All results in the same group are saved into a single series. The label field is used as a descriptor (the label is taken from the first result in the group). | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 3 |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00598180 | Refresh | Closed | lin_s | May 27, 2018 9:52:22 PM GMT+08:00 | V4.1.6:Duplicate Thin slice series is created when merge series is enabled for 2 acquisitions. | 4.2.x |  | The user can merge the duplicate series using the CT Viewer. The IFU section called Merge Series describes how to do this: Defined groups of results can be saved into the same DICOM series. Select the appropriate group (1 - 5) for your result. All results in the same group are saved into a single series. The label field is used as a descriptor (the label is taken from the first result in the group). | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 3 |  | Console4 | NO | Closed |  | IQon1
CTUDT00598179 | Refresh | Closed | lin_s | May 27, 2018 9:47:16 PM GMT+08:00 | V4.1: Able to Add "Breathing Light" to 2nd scan of sequential dual energy. | 4.2.x |  | In the scan ruler, select the breathing lights in the second step of the sequential dual energy study and change the language to None. If the user needs auto voice or breathing lights for first step of the sequential dual energy study, he can select the step and enable it. | Subsystem |  | N/A | No Plan to Fix | prabhu_m | 0 | 3 | Console Generic | Console4 | NO | Closed |  | Ovation
CTUDT00598179 | Refresh | Closed | lin_s | May 27, 2018 9:47:16 PM GMT+08:00 | V4.1: Able to Add "Breathing Light" to 2nd scan of sequential dual energy. | 4.2.x |  | In the scan ruler, select the breathing lights in the second step of the sequential dual energy study and change the language to None. If the user needs auto voice or breathing lights for first step of the sequential dual energy study, he can select the step and enable it. | Subsystem |  | N/A | No Plan to Fix | prabhu_m | 0 | 3 | Console Generic | Console4 | NO | Closed |  | Brilliance 64
CTUDT00598179 | Refresh | Closed | lin_s | May 27, 2018 9:47:16 PM GMT+08:00 | V4.1: Able to Add "Breathing Light" to 2nd scan of sequential dual energy. | 4.2.x |  | In the scan ruler, select the breathing lights in the second step of the sequential dual energy study and change the language to None. If the user needs auto voice or breathing lights for first step of the sequential dual energy study, he can select the step and enable it. | Subsystem |  | N/A | No Plan to Fix | prabhu_m | 0 | 3 | Console Generic | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00598179 | Refresh | Closed | lin_s | May 27, 2018 9:47:16 PM GMT+08:00 | V4.1: Able to Add "Breathing Light" to 2nd scan of sequential dual energy. | 4.2.x |  | In the scan ruler, select the breathing lights in the second step of the sequential dual energy study and change the language to None. If the user needs auto voice or breathing lights for first step of the sequential dual energy study, he can select the step and enable it. | Subsystem |  | N/A | No Plan to Fix | prabhu_m | 0 | 3 | Console Generic | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00598179 | Refresh | Closed | lin_s | May 27, 2018 9:47:16 PM GMT+08:00 | V4.1: Able to Add "Breathing Light" to 2nd scan of sequential dual energy. | 4.2.x |  | In the scan ruler, select the breathing lights in the second step of the sequential dual energy study and change the language to None. If the user needs auto voice or breathing lights for first step of the sequential dual energy study, he can select the step and enable it. | Subsystem |  | N/A | No Plan to Fix | prabhu_m | 0 | 3 | Console Generic | Console4 | NO | Closed |  | IQon1
CTUDT00598178 | Refresh | Closed | lin_s | May 27, 2018 9:34:05 PM GMT+08:00 | Veyron:First and last image bed positions were flipped in slab viewer when the result direction is set to Head to Feet. | 4.2.x |  | The user may plan in slab viewer regardless of the direction of the images. | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 1 | 3 |  | Console4 | NO | Closed |  | Ovation
CTUDT00598178 | Refresh | Closed | lin_s | May 27, 2018 9:34:05 PM GMT+08:00 | Veyron:First and last image bed positions were flipped in slab viewer when the result direction is set to Head to Feet. | 4.2.x |  | The user may plan in slab viewer regardless of the direction of the images. | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 1 | 3 |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00598178 | Refresh | Closed | lin_s | May 27, 2018 9:34:05 PM GMT+08:00 | Veyron:First and last image bed positions were flipped in slab viewer when the result direction is set to Head to Feet. | 4.2.x |  | The user may plan in slab viewer regardless of the direction of the images. | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 1 | 3 |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00598178 | Refresh | Closed | lin_s | May 27, 2018 9:34:05 PM GMT+08:00 | Veyron:First and last image bed positions were flipped in slab viewer when the result direction is set to Head to Feet. | 4.2.x |  | The user may plan in slab viewer regardless of the direction of the images. | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 1 | 3 |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00598178 | Refresh | Closed | lin_s | May 27, 2018 9:34:05 PM GMT+08:00 | Veyron:First and last image bed positions were flipped in slab viewer when the result direction is set to Head to Feet. | 4.2.x |  | The user may plan in slab viewer regardless of the direction of the images. | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 1 | 3 |  | Console4 | NO | Closed |  | IQon1
CTUDT00598176 | Refresh | Closed | rodan_d | May 27, 2018 8:43:35 PM GMT+08:00 | (Under Service user only) Clicking "NEW" button under patient demographics panel does not clear "First Name" | 4.1.6 | WIP14 | manually clear the patient name and replace with the desired information. | Subsystem |  | Console-SRS.15761 | Fixed | Snehashis_Dey | 0 | 4 | Console Software | Console4 | NO | Closed | MergeFrom4.x | Brilliance 64 |  | WIP14 | February 26, 2019 2:58:22 PM GMT+08:00
CTUDT00598176 | Refresh | Closed | rodan_d | May 27, 2018 8:43:35 PM GMT+08:00 | (Under Service user only) Clicking "NEW" button under patient demographics panel does not clear "First Name" | 4.1.6 | WIP14 | manually clear the patient name and replace with the desired information. | Subsystem |  | Console-SRS.15761 | Fixed | Snehashis_Dey | 0 | 4 | Console Software | Console4 | NO | Closed | MergeFrom4.x | Brilliance Big Bore |  | WIP14 | February 26, 2019 2:58:22 PM GMT+08:00
CTUDT00598176 | Refresh | Closed | rodan_d | May 27, 2018 8:43:35 PM GMT+08:00 | (Under Service user only) Clicking "NEW" button under patient demographics panel does not clear "First Name" | 4.1.6 | WIP14 | manually clear the patient name and replace with the desired information. | Subsystem |  | Console-SRS.15761 | Fixed | Snehashis_Dey | 0 | 4 | Console Software | Console4 | NO | Closed | MergeFrom4.x | Ingenuity 64 |  | WIP14 | February 26, 2019 2:58:22 PM GMT+08:00
CTUDT00598176 | Refresh | Closed | rodan_d | May 27, 2018 8:43:35 PM GMT+08:00 | (Under Service user only) Clicking "NEW" button under patient demographics panel does not clear "First Name" | 4.1.6 | WIP14 | manually clear the patient name and replace with the desired information. | Subsystem |  | Console-SRS.15761 | Fixed | Snehashis_Dey | 0 | 4 | Console Software | Console4 | NO | Closed | MergeFrom4.x | IQon1 |  | WIP14 | February 26, 2019 2:58:22 PM GMT+08:00
CTUDT00598172 | Refresh | Closed | rodan_d | May 27, 2018 8:19:16 PM GMT+08:00 | Veyron specific software, PETPCCULoggerService, is included in non-Veyron software | Veyron Verification WIP128 |  |  | Subsystem |  |  | Change Proposal | cumpson | 0 | 3 | Gantry Software | GCX | NO | Closed |  | IQon1
CTUDT00598171 | Refresh | Closed | rodan_d | May 27, 2018 8:15:38 PM GMT+08:00 | CIRS accepts recon requests that are out of the datasets range | WIP505 |  | Redo the offline reconstruction with the correct start and end locations. | Subsystem |  | NA | No Plan to Fix | richards_b | 0 | 4 |  | Support-System | NO | Closed | SubNC | Ovation
CTUDT00598171 | Refresh | Closed | rodan_d | May 27, 2018 8:15:38 PM GMT+08:00 | CIRS accepts recon requests that are out of the datasets range | WIP505 |  | Redo the offline reconstruction with the correct start and end locations. | Subsystem |  | NA | No Plan to Fix | richards_b | 0 | 4 |  | Support-System | NO | Closed | SubNC | Brilliance 64
CTUDT00598171 | Refresh | Closed | rodan_d | May 27, 2018 8:15:38 PM GMT+08:00 | CIRS accepts recon requests that are out of the datasets range | WIP505 |  | Redo the offline reconstruction with the correct start and end locations. | Subsystem |  | NA | No Plan to Fix | richards_b | 0 | 4 |  | Support-System | NO | Closed | SubNC | Ingenuity 64
CTUDT00598171 | Refresh | Closed | rodan_d | May 27, 2018 8:15:38 PM GMT+08:00 | CIRS accepts recon requests that are out of the datasets range | WIP505 |  | Redo the offline reconstruction with the correct start and end locations. | Subsystem |  | NA | No Plan to Fix | richards_b | 0 | 4 |  | Support-System | NO | Closed | SubNC | Ingenuity BigBore
CTUDT00598171 | Refresh | Closed | rodan_d | May 27, 2018 8:15:38 PM GMT+08:00 | CIRS accepts recon requests that are out of the datasets range | WIP505 |  | Redo the offline reconstruction with the correct start and end locations. | Subsystem |  | NA | No Plan to Fix | richards_b | 0 | 4 |  | Support-System | NO | Closed | SubNC | IQon1
CTUDT00598169 | Refresh | Closed | rodan_d | May 27, 2018 7:56:38 PM GMT+08:00 | Infant scans off of the patient support using the infant cradle could be too noisy due to inaccurate water equivalent diameter estimation. | 4.1.3.XX022 |  | 1. Perform scan on the couch. 2. If performing scan off the couch turn off dose management and plan manually. | Subsystem |  | N/A | No Plan to Fix | richards_b | 0 | 3 |  | Support-Physics | YES | Closed | IFU | Ovation
CTUDT00598169 | Refresh | Closed | rodan_d | May 27, 2018 7:56:38 PM GMT+08:00 | Infant scans off of the patient support using the infant cradle could be too noisy due to inaccurate water equivalent diameter estimation. | 4.1.3.XX022 |  | 1. Perform scan on the couch. 2. If performing scan off the couch turn off dose management and plan manually. | Subsystem |  | N/A | No Plan to Fix | richards_b | 0 | 3 |  | Support-Physics | YES | Closed | IFU | Brilliance 64
CTUDT00598169 | Refresh | Closed | rodan_d | May 27, 2018 7:56:38 PM GMT+08:00 | Infant scans off of the patient support using the infant cradle could be too noisy due to inaccurate water equivalent diameter estimation. | 4.1.3.XX022 |  | 1. Perform scan on the couch. 2. If performing scan off the couch turn off dose management and plan manually. | Subsystem |  | N/A | No Plan to Fix | richards_b | 0 | 3 |  | Support-Physics | YES | Closed | IFU | Brilliance Big Bore
CTUDT00598169 | Refresh | Closed | rodan_d | May 27, 2018 7:56:38 PM GMT+08:00 | Infant scans off of the patient support using the infant cradle could be too noisy due to inaccurate water equivalent diameter estimation. | 4.1.3.XX022 |  | 1. Perform scan on the couch. 2. If performing scan off the couch turn off dose management and plan manually. | Subsystem |  | N/A | No Plan to Fix | richards_b | 0 | 3 |  | Support-Physics | YES | Closed | IFU | Ingenuity 64
CTUDT00598169 | Refresh | Closed | rodan_d | May 27, 2018 7:56:38 PM GMT+08:00 | Infant scans off of the patient support using the infant cradle could be too noisy due to inaccurate water equivalent diameter estimation. | 4.1.3.XX022 |  | 1. Perform scan on the couch. 2. If performing scan off the couch turn off dose management and plan manually. | Subsystem |  | N/A | No Plan to Fix | richards_b | 0 | 3 |  | Support-Physics | YES | Closed | IFU | IQon1
CTUDT00598159 | Refresh | Closed | oren_d | May 27, 2018 3:11:03 PM GMT+08:00 | Val: 4.1.7 - 'Matlab_Runtime.bat' file not available on system | 4.1.7 |  |  | Subsystem |  |  | Change Proposal | rubin_y | 0 |  |  | Support-Service | NO | Closed |  | Brilliance 64
CTUDT00598159 | Refresh | Closed | oren_d | May 27, 2018 3:11:03 PM GMT+08:00 | Val: 4.1.7 - 'Matlab_Runtime.bat' file not available on system | 4.1.7 |  |  | Subsystem |  |  | Change Proposal | rubin_y | 0 |  |  | Support-Service | NO | Closed |  | Ingenuity 64
CTUDT00598154 | Refresh | Closed | oren_d | May 27, 2018 1:23:05 PM GMT+08:00 | Clinical User DEMO folder image files do not represent current IQ for Philips Models for iCT & iCT-SP | V4.1.6 | WIP16 |  | Subsystem | USP | NA | Fixed | sirigiri_s | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP15SegReq | Ovation |  | WIP16 | April 23, 2019 8:33:31 PM GMT+08:00
CTUDT00598154 | Refresh | Closed | oren_d | May 27, 2018 1:23:05 PM GMT+08:00 | Clinical User DEMO folder image files do not represent current IQ for Philips Models for iCT & iCT-SP | V4.1.6 | WIP16 |  | Subsystem | USP | NA | Fixed | sirigiri_s | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP15SegReq | Brilliance 64 |  | WIP16 | April 23, 2019 8:33:31 PM GMT+08:00
CTUDT00598154 | Refresh | Closed | oren_d | May 27, 2018 1:23:05 PM GMT+08:00 | Clinical User DEMO folder image files do not represent current IQ for Philips Models for iCT & iCT-SP | V4.1.6 | WIP16 |  | Subsystem | USP | NA | Fixed | sirigiri_s | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP15SegReq | Brilliance Big Bore |  | WIP16 | April 23, 2019 8:33:31 PM GMT+08:00
CTUDT00598154 | Refresh | Closed | oren_d | May 27, 2018 1:23:05 PM GMT+08:00 | Clinical User DEMO folder image files do not represent current IQ for Philips Models for iCT & iCT-SP | V4.1.6 | WIP16 |  | Subsystem | USP | NA | Fixed | sirigiri_s | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | WIP15SegReq | Ingenuity 64 |  | WIP16 | April 23, 2019 8:33:31 PM GMT+08:00
CTUDT00598130 | Refresh | Closed | fitusi_r | May 24, 2018 10:40:29 PM GMT+08:00 | Refresh,IQon- After performing a procedure in " Quality Assurance " and pressing "EXIT", the "QA" is grayed out and cannot be accessed | WIP 7 |  |  | Subsystem |  | NA | Duplicate | kumar_v | 0 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00598128 | Refresh | Closed | fitusi_r | May 24, 2018 9:44:56 PM GMT+08:00 | Refresh, IQon:When the "End Exam and continue working " button is pressed, the View2 does not appear on the right screen | WIP 7 | WIP11.1 | N/A | Subsystem | End Exam Continue Working | Console-SRS.9747 | Fixed | veerabhadraiah | 0 | 3 | Console USP_EBW | iCAP | NO | Closed | SeqReq_WIP8.1 | IQon1 |  | WIP11.1 | February 25, 2019 4:31:36 PM GMT+08:00
CTUDT00598122 | Refresh | Closed | fitusi_r | May 24, 2018 8:50:38 PM GMT+08:00 | Refresh, IQon- CT user permissions issue. | WIP 7 | WIP8.1 |  | Subsystem |  | N/A | Fixed | mohit_p | 0 | 2 | Console Software | Console4 | YES | Closed | SegReq_WIP9 | Ovation |  | WIP8.1 | July 17, 2018 5:25:20 PM GMT+08:00
CTUDT00598122 | Refresh | Closed | fitusi_r | May 24, 2018 8:50:38 PM GMT+08:00 | Refresh, IQon- CT user permissions issue. | WIP 7 | WIP8.1 |  | Subsystem |  | N/A | Fixed | mohit_p | 0 | 2 | Console Software | Console4 | YES | Closed | SegReq_WIP9 | Brilliance 64 |  | WIP8.1 | July 17, 2018 5:25:20 PM GMT+08:00
CTUDT00598122 | Refresh | Closed | fitusi_r | May 24, 2018 8:50:38 PM GMT+08:00 | Refresh, IQon- CT user permissions issue. | WIP 7 | WIP8.1 |  | Subsystem |  | N/A | Fixed | mohit_p | 0 | 2 | Console Software | Console4 | YES | Closed | SegReq_WIP9 | Ingenuity 64 |  | WIP8.1 | July 17, 2018 5:25:20 PM GMT+08:00
CTUDT00598122 | Refresh | Closed | fitusi_r | May 24, 2018 8:50:38 PM GMT+08:00 | Refresh, IQon- CT user permissions issue. | WIP 7 | WIP8.1 |  | Subsystem |  | N/A | Fixed | mohit_p | 0 | 2 | Console Software | Console4 | YES | Closed | SegReq_WIP9 | IQon1 |  | WIP8.1 | July 17, 2018 5:25:20 PM GMT+08:00
CTUDT00598120 | Refresh | Closed | fitusi_r | May 24, 2018 8:15:15 PM GMT+08:00 | Refresh, IQon-when planning a clinical scan with trigger=timed the button “GO” is grayed out, | WIP 7 |  |  | Subsystem | Timed Scan | N/A | Works As Specified / Not A Defect | tamari_o | 1 | 2 |  | Support-System | NO | Closed | WIP10SegReq | IQon1
CTUDT00598099 | Refresh | Closed | fitusi_r | May 23, 2018 4:08:37 PM GMT+08:00 | Refresh, ICT- CIRS installation of V4.8, 5.1.0.0100 Failed. | 4.1.7 |  |  | Subsystem |  | N/A | Invalid/Uncontrolled Configuration | richards_b | 0 |  |  | CIRS | NO | Closed |  | IQon1
CTUDT00598097 | Refresh | Closed | fitusi_r | May 23, 2018 3:40:36 PM GMT+08:00 | Refresh, IQon- Offline Recon issue. | WIP6.1 |  |  | Subsystem |  | N/A | No Plan to Fix | kumar_v | 0 | 4 |  | Console4 | NO | Closed | WIP9SegReq | IQon1
CTUDT00598095 | Refresh | Closed | fitusi_r | May 23, 2018 3:28:36 PM GMT+08:00 | Refresh,IQon- Can't collect a BugRep. | WIP6.1 |  |  | Subsystem | Bugreport |  | Works As Specified / Not A Defect | babu_s | 0 |  |  | Service Tools | NO | Closed | WIP9SegReq | IQon1
CTUDT00598057 | Refresh | Closed | oren_d | May 22, 2018 2:55:10 AM GMT+08:00 | Worklist/Filtering, when applying any filter that will match only one patient the displayed entry is not marked as selected. | 4.1.7 |  | Manually select the patient record | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 | Console Generic | Console4 | NO | Closed |  | Ovation
CTUDT00598057 | Refresh | Closed | oren_d | May 22, 2018 2:55:10 AM GMT+08:00 | Worklist/Filtering, when applying any filter that will match only one patient the displayed entry is not marked as selected. | 4.1.7 |  | Manually select the patient record | Subsystem |  | N/A | No Plan to Fix | sanjeeva_a | 0 | 4 | Console Generic | Console4 | NO | Closed |  | Brilliance 64
CTUDT00598048 | Refresh | Closed | fitusi_r | May 22, 2018 12:27:04 AM GMT+08:00 | Refresh, IQon-CIRS issue, Rack configuration by the registry in the CIRS servers. | WIP6.1 |  |  | Subsystem |  |  | Duplicate | richards_b | 0 |  |  | CIRS | NO | Closed |  | IQon1
CTUDT00598046 | Refresh | Closed | fitusi_r | May 22, 2018 12:17:54 AM GMT+08:00 | Refresh, IQon-The "CAN CARD" cannot be identified using the "device manager" (HP Host). | WIP6.1 |  |  | Subsystem |  | N/A | System Repaired | oren_d | 0 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00598044 | Refresh | Closed | fitusi_r | May 22, 2018 12:01:52 AM GMT+08:00 | Refresh,IQon-Noises from the Audio card on board (HP Host) | WIP6.1 | WIP 21 |  | Subsystem | Audio driver; SP1 | N/A | Fixed | tiwari_g | 0 | 3 | Console Software | Console4 | NO | Closed | WIP9.1SegReq | Brilliance 64 |  | WIP 21 | October 11, 2019 10:34:19 PM GMT+08:00
CTUDT00598044 | Refresh | Closed | fitusi_r | May 22, 2018 12:01:52 AM GMT+08:00 | Refresh,IQon-Noises from the Audio card on board (HP Host) | WIP6.1 | WIP 21 |  | Subsystem | Audio driver; SP1 | N/A | Fixed | tiwari_g | 0 | 3 | Console Software | Console4 | NO | Closed | WIP9.1SegReq | Brilliance Big Bore |  | WIP 21 | October 11, 2019 10:34:19 PM GMT+08:00
CTUDT00598044 | Refresh | Closed | fitusi_r | May 22, 2018 12:01:52 AM GMT+08:00 | Refresh,IQon-Noises from the Audio card on board (HP Host) | WIP6.1 | WIP 21 |  | Subsystem | Audio driver; SP1 | N/A | Fixed | tiwari_g | 0 | 3 | Console Software | Console4 | NO | Closed | WIP9.1SegReq | Ingenuity 64 |  | WIP 21 | October 11, 2019 10:34:19 PM GMT+08:00
CTUDT00598044 | Refresh | Closed | fitusi_r | May 22, 2018 12:01:52 AM GMT+08:00 | Refresh,IQon-Noises from the Audio card on board (HP Host) | WIP6.1 | WIP 21 |  | Subsystem | Audio driver; SP1 | N/A | Fixed | tiwari_g | 0 | 3 | Console Software | Console4 | NO | Closed | WIP9.1SegReq | IQon1 |  | WIP 21 | October 11, 2019 10:34:19 PM GMT+08:00
CTUDT00598032 | Refresh | Closed | sengupta_a | May 21, 2018 8:09:44 PM GMT+08:00 | Min mAs and mA s fields are not highlighted blue when Doseright option is checked and surview is performed. | WIP6.1 | WIP15 | NO as it's a UI issue | Subsystem | Scan Parameters | N/A | Fixed | sagar_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SegReq | Ovation |  | WIP15 | March 18, 2019 10:05:54 PM GMT+08:00
CTUDT00598032 | Refresh | Closed | sengupta_a | May 21, 2018 8:09:44 PM GMT+08:00 | Min mAs and mA s fields are not highlighted blue when Doseright option is checked and surview is performed. | WIP6.1 | WIP15 | NO as it's a UI issue | Subsystem | Scan Parameters | N/A | Fixed | sagar_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SegReq | Brilliance 64 |  | WIP15 | March 18, 2019 10:05:54 PM GMT+08:00
CTUDT00598032 | Refresh | Closed | sengupta_a | May 21, 2018 8:09:44 PM GMT+08:00 | Min mAs and mA s fields are not highlighted blue when Doseright option is checked and surview is performed. | WIP6.1 | WIP15 | NO as it's a UI issue | Subsystem | Scan Parameters | N/A | Fixed | sagar_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SegReq | Ingenuity 64 |  | WIP15 | March 18, 2019 10:05:54 PM GMT+08:00
CTUDT00598032 | Refresh | Closed | sengupta_a | May 21, 2018 8:09:44 PM GMT+08:00 | Min mAs and mA s fields are not highlighted blue when Doseright option is checked and surview is performed. | WIP6.1 | WIP15 | NO as it's a UI issue | Subsystem | Scan Parameters | N/A | Fixed | sagar_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP14SegReq | IQon1 |  | WIP15 | March 18, 2019 10:05:54 PM GMT+08:00
CTUDT00598029 | Refresh | Closed | singh_p | May 21, 2018 3:17:31 PM GMT+08:00 | Refresh [iCT] : During the timed scan , delete option displayed in context menu of series list for results. | WIP6.1 | WIP9.1 | No | Subsystem |  | N/A | Fixed | bhawna_g | 0 | 4 | Console Software | Console4 | NO | Closed | WIP10SegReq | Ovation |  | WIP9.1 | July 31, 2018 6:12:11 PM GMT+08:00
CTUDT00598029 | Refresh | Closed | singh_p | May 21, 2018 3:17:31 PM GMT+08:00 | Refresh [iCT] : During the timed scan , delete option displayed in context menu of series list for results. | WIP6.1 | WIP9.1 | No | Subsystem |  | N/A | Fixed | bhawna_g | 0 | 4 | Console Software | Console4 | NO | Closed | WIP10SegReq | Brilliance 64 |  | WIP9.1 | July 31, 2018 6:12:11 PM GMT+08:00
CTUDT00598029 | Refresh | Closed | singh_p | May 21, 2018 3:17:31 PM GMT+08:00 | Refresh [iCT] : During the timed scan , delete option displayed in context menu of series list for results. | WIP6.1 | WIP9.1 | No | Subsystem |  | N/A | Fixed | bhawna_g | 0 | 4 | Console Software | Console4 | NO | Closed | WIP10SegReq | Ingenuity 64 |  | WIP9.1 | July 31, 2018 6:12:11 PM GMT+08:00
CTUDT00598029 | Refresh | Closed | singh_p | May 21, 2018 3:17:31 PM GMT+08:00 | Refresh [iCT] : During the timed scan , delete option displayed in context menu of series list for results. | WIP6.1 | WIP9.1 | No | Subsystem |  | N/A | Fixed | bhawna_g | 0 | 4 | Console Software | Console4 | NO | Closed | WIP10SegReq | IQon1 |  | WIP9.1 | July 31, 2018 6:12:11 PM GMT+08:00
CTUDT00598017 | Refresh | Closed | zhang_l | May 17, 2018 11:17:36 PM GMT+08:00 | Patient ,Plan etc. buttons region were obscured after turn on/off 'Automatically hide the taskbar in desktop mode'. | WIP6.1 | N/A |  | Subsystem |  | N/A | Fixed | dasgupta_s | 0 | 3 | Console Software | Support-Security | NO | Closed | SegReq_WIP9 | Ovation |  | N/A | July 26, 2019 3:06:00 PM GMT+08:00
CTUDT00598017 | Refresh | Closed | zhang_l | May 17, 2018 11:17:36 PM GMT+08:00 | Patient ,Plan etc. buttons region were obscured after turn on/off 'Automatically hide the taskbar in desktop mode'. | WIP6.1 | N/A |  | Subsystem |  | N/A | Fixed | dasgupta_s | 0 | 3 | Console Software | Support-Security | NO | Closed | SegReq_WIP9 | Brilliance 64 |  | N/A | July 26, 2019 3:06:00 PM GMT+08:00
CTUDT00598017 | Refresh | Closed | zhang_l | May 17, 2018 11:17:36 PM GMT+08:00 | Patient ,Plan etc. buttons region were obscured after turn on/off 'Automatically hide the taskbar in desktop mode'. | WIP6.1 | N/A |  | Subsystem |  | N/A | Fixed | dasgupta_s | 0 | 3 | Console Software | Support-Security | NO | Closed | SegReq_WIP9 | Ingenuity 64 |  | N/A | July 26, 2019 3:06:00 PM GMT+08:00
CTUDT00598017 | Refresh | Closed | zhang_l | May 17, 2018 11:17:36 PM GMT+08:00 | Patient ,Plan etc. buttons region were obscured after turn on/off 'Automatically hide the taskbar in desktop mode'. | WIP6.1 | N/A |  | Subsystem |  | N/A | Fixed | dasgupta_s | 0 | 3 | Console Software | Support-Security | NO | Closed | SegReq_WIP9 | IQon1 |  | N/A | July 26, 2019 3:06:00 PM GMT+08:00
CTUDT00598012 | Refresh | Closed | zhang_t | May 17, 2018 8:44:34 PM GMT+08:00 | ECG device/Bellow Device not connection displayed in Console UI while the device was connecting. | WIP6.1 | WIP9 | NA | Subsystem |  | NA | Fixed | rathore_cs | 0 | 2 | Console Software | Console4 | NO | Closed | WIP9SegReq | Brilliance 64 |  | WIP9 | August 20, 2018 5:32:15 PM GMT+08:00
CTUDT00598012 | Refresh | Closed | zhang_t | May 17, 2018 8:44:34 PM GMT+08:00 | ECG device/Bellow Device not connection displayed in Console UI while the device was connecting. | WIP6.1 | WIP9 | NA | Subsystem |  | NA | Fixed | rathore_cs | 0 | 2 | Console Software | Console4 | NO | Closed | WIP9SegReq | Ingenuity 64 |  | WIP9 | August 20, 2018 5:32:15 PM GMT+08:00
CTUDT00598008 | Refresh | Closed | manjunath_c | May 17, 2018 7:04:27 PM GMT+08:00 | [Refresh] iQon-HB : Deviation in the default parameter values displayed in ECM compared to what is mentioned in the "DHF282684 Rev03 IQon Spectral CT Reference Exam Cards" Specification | WIP6.1 | WIP13 |  | Subsystem | Reference Ecs | Console.SRS-10648 | Fixed | nanjunda_s | 0 | 2 | Test Design | Console4 | NO | Closed | WIP11SegReq | IQon1 |  | WIP13 | February 12, 2019 6:25:51 PM GMT+08:00
CTUDT00598006 | Refresh | Closed | hiremath_b | May 17, 2018 6:52:17 PM GMT+08:00 | Refresh] Ingenuity/iCT : Phase bar gets deleted when trying to move phase bar within peak to peak range in the 1st cycle in offline cardiac viewer, where as it works fine for last cycle. | WIP6.1 | WIP10 |  | Subsystem |  | None | Fixed | datta_i | 0 | 2 | Console Software | Console4 | NO | Closed | WIP10SegReq | Ovation |  | WIP10 | September 18, 2018 7:50:06 PM GMT+08:00
CTUDT00598006 | Refresh | Closed | hiremath_b | May 17, 2018 6:52:17 PM GMT+08:00 | Refresh] Ingenuity/iCT : Phase bar gets deleted when trying to move phase bar within peak to peak range in the 1st cycle in offline cardiac viewer, where as it works fine for last cycle. | WIP6.1 | WIP10 |  | Subsystem |  | None | Fixed | datta_i | 0 | 2 | Console Software | Console4 | NO | Closed | WIP10SegReq | Brilliance 64 |  | WIP10 | September 18, 2018 7:50:06 PM GMT+08:00
CTUDT00598006 | Refresh | Closed | hiremath_b | May 17, 2018 6:52:17 PM GMT+08:00 | Refresh] Ingenuity/iCT : Phase bar gets deleted when trying to move phase bar within peak to peak range in the 1st cycle in offline cardiac viewer, where as it works fine for last cycle. | WIP6.1 | WIP10 |  | Subsystem |  | None | Fixed | datta_i | 0 | 2 | Console Software | Console4 | NO | Closed | WIP10SegReq | Ingenuity 64 |  | WIP10 | September 18, 2018 7:50:06 PM GMT+08:00
CTUDT00598006 | Refresh | Closed | hiremath_b | May 17, 2018 6:52:17 PM GMT+08:00 | Refresh] Ingenuity/iCT : Phase bar gets deleted when trying to move phase bar within peak to peak range in the 1st cycle in offline cardiac viewer, where as it works fine for last cycle. | WIP6.1 | WIP10 |  | Subsystem |  | None | Fixed | datta_i | 0 | 2 | Console Software | Console4 | NO | Closed | WIP10SegReq | IQon1 |  | WIP10 | September 18, 2018 7:50:06 PM GMT+08:00
CTUDT00598003 | Refresh | Closed | basavaraju_p | May 17, 2018 6:33:51 PM GMT+08:00 | [Refresh] Ingenuity 64 : Deviation in the parameter values displayed in Exam Card Manager compared to what is mentioned in the Scan Type Specification - DHF251214 Rev 06 | WIP6.1 | WIP16 | The user can check or uncheck the Dose Right check box. | Subsystem | Scan Types | Console.SRS 15254 | Fixed | bhawna_g | 0 | 3 | Console Software | Console4 | NO | Closed | SubNC | Ovation |  | WIP16 | April 5, 2019 5:12:13 PM GMT+08:00
CTUDT00598003 | Refresh | Closed | basavaraju_p | May 17, 2018 6:33:51 PM GMT+08:00 | [Refresh] Ingenuity 64 : Deviation in the parameter values displayed in Exam Card Manager compared to what is mentioned in the Scan Type Specification - DHF251214 Rev 06 | WIP6.1 | WIP16 | The user can check or uncheck the Dose Right check box. | Subsystem | Scan Types | Console.SRS 15254 | Fixed | bhawna_g | 0 | 3 | Console Software | Console4 | NO | Closed | SubNC | Brilliance 64 |  | WIP16 | April 5, 2019 5:12:13 PM GMT+08:00
CTUDT00598003 | Refresh | Closed | basavaraju_p | May 17, 2018 6:33:51 PM GMT+08:00 | [Refresh] Ingenuity 64 : Deviation in the parameter values displayed in Exam Card Manager compared to what is mentioned in the Scan Type Specification - DHF251214 Rev 06 | WIP6.1 | WIP16 | The user can check or uncheck the Dose Right check box. | Subsystem | Scan Types | Console.SRS 15254 | Fixed | bhawna_g | 0 | 3 | Console Software | Console4 | NO | Closed | SubNC | Ingenuity 64 |  | WIP16 | April 5, 2019 5:12:13 PM GMT+08:00
CTUDT00598003 | Refresh | Closed | basavaraju_p | May 17, 2018 6:33:51 PM GMT+08:00 | [Refresh] Ingenuity 64 : Deviation in the parameter values displayed in Exam Card Manager compared to what is mentioned in the Scan Type Specification - DHF251214 Rev 06 | WIP6.1 | WIP16 | The user can check or uncheck the Dose Right check box. | Subsystem | Scan Types | Console.SRS 15254 | Fixed | bhawna_g | 0 | 3 | Console Software | Console4 | NO | Closed | SubNC | IQon1 |  | WIP16 | April 5, 2019 5:12:13 PM GMT+08:00
CTUDT00597999 | Refresh | Closed | basavaraju_p | May 17, 2018 5:26:07 PM GMT+08:00 | Deviation in the parameter values displayed in Examcardmanager compared to what is mentioned in the Scan Type Specifiction - DHF251214 Rev 06 | WIP6.1 | WIP12 |  | Subsystem | Scan Types | N/A | Fixed | nanjunda_s | 0 | 2 | Test Design | Console4 | NO | Closed | WIP11SegReq | Brilliance 64 |  | WIP12 | February 12, 2019 5:55:32 PM GMT+08:00
CTUDT00597999 | Refresh | Closed | basavaraju_p | May 17, 2018 5:26:07 PM GMT+08:00 | Deviation in the parameter values displayed in Examcardmanager compared to what is mentioned in the Scan Type Specifiction - DHF251214 Rev 06 | WIP6.1 | WIP12 |  | Subsystem | Scan Types | N/A | Fixed | nanjunda_s | 0 | 2 | Test Design | Console4 | NO | Closed | WIP11SegReq | Ingenuity 64 |  | WIP12 | February 12, 2019 5:55:32 PM GMT+08:00
CTUDT00597982 | Refresh | Closed | naik_p | May 16, 2018 8:57:59 PM GMT+08:00 | ELSCINTCT_ADAPTIVE_FILTER TAG displays wrong value in image parameters and mdus | WIP6.1 |  |  | Subsystem |  | Console-SRS.11342,Console-SRS.14074 | Works As Specified / Not A Defect | kumar_v | 0 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00597979 | Refresh | Closed | sengupta_a | May 16, 2018 7:38:55 PM GMT+08:00 | Idose level dropdown shows values upto 5 only i.e 1,2,3,4,5 | WIP6.1 |  | NA | Subsystem |  | N/A | Works As Specified / Not A Defect | rathore_cs | 0 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00597979 | Refresh | Closed | sengupta_a | May 16, 2018 7:38:55 PM GMT+08:00 | Idose level dropdown shows values upto 5 only i.e 1,2,3,4,5 | WIP6.1 |  | NA | Subsystem |  | N/A | Works As Specified / Not A Defect | rathore_cs | 0 |  |  | Console4 | NO | Closed |  | Brilliance 64
CTUDT00597979 | Refresh | Closed | sengupta_a | May 16, 2018 7:38:55 PM GMT+08:00 | Idose level dropdown shows values upto 5 only i.e 1,2,3,4,5 | WIP6.1 |  | NA | Subsystem |  | N/A | Works As Specified / Not A Defect | rathore_cs | 0 |  |  | Console4 | NO | Closed |  | Brilliance Big Bore
CTUDT00597979 | Refresh | Closed | sengupta_a | May 16, 2018 7:38:55 PM GMT+08:00 | Idose level dropdown shows values upto 5 only i.e 1,2,3,4,5 | WIP6.1 |  | NA | Subsystem |  | N/A | Works As Specified / Not A Defect | rathore_cs | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00597979 | Refresh | Closed | sengupta_a | May 16, 2018 7:38:55 PM GMT+08:00 | Idose level dropdown shows values upto 5 only i.e 1,2,3,4,5 | WIP6.1 |  | NA | Subsystem |  | N/A | Works As Specified / Not A Defect | rathore_cs | 0 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00597977 | Refresh | Closed | basavaraju_p | May 16, 2018 5:56:26 PM GMT+08:00 | In Chest Helical for the Saggittal Result the Actual reconstructed images are not equal to planned images | WIP6.1 |  | NA | Subsystem |  | Console-SRS.15364 | Works As Specified / Not A Defect | rathore_cs | 0 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00597971 | Refresh | Closed | basavaraju_p | May 16, 2018 3:45:30 PM GMT+08:00 | Go button is not enabled for the added scan with the combination of helical and axial | WIP6.1 | WIP8.1 |  | Subsystem |  | Console.SRS 11245 | Fixed | shruthi_u | 0 | 3 | Console Software | Console4 | YES | Closed | SegReq_WIP8.1 | Ovation |  | WIP8.1 | July 12, 2018 8:18:21 PM GMT+08:00
CTUDT00597971 | Refresh | Closed | basavaraju_p | May 16, 2018 3:45:30 PM GMT+08:00 | Go button is not enabled for the added scan with the combination of helical and axial | WIP6.1 | WIP8.1 |  | Subsystem |  | Console.SRS 11245 | Fixed | shruthi_u | 0 | 3 | Console Software | Console4 | YES | Closed | SegReq_WIP8.1 | Brilliance 64 |  | WIP8.1 | July 12, 2018 8:18:21 PM GMT+08:00
CTUDT00597971 | Refresh | Closed | basavaraju_p | May 16, 2018 3:45:30 PM GMT+08:00 | Go button is not enabled for the added scan with the combination of helical and axial | WIP6.1 | WIP8.1 |  | Subsystem |  | Console.SRS 11245 | Fixed | shruthi_u | 0 | 3 | Console Software | Console4 | YES | Closed | SegReq_WIP8.1 | Ingenuity 64 |  | WIP8.1 | July 12, 2018 8:18:21 PM GMT+08:00
CTUDT00597971 | Refresh | Closed | basavaraju_p | May 16, 2018 3:45:30 PM GMT+08:00 | Go button is not enabled for the added scan with the combination of helical and axial | WIP6.1 | WIP8.1 |  | Subsystem |  | Console.SRS 11245 | Fixed | shruthi_u | 0 | 3 | Console Software | Console4 | YES | Closed | SegReq_WIP8.1 | IQon1 |  | WIP8.1 | July 12, 2018 8:18:21 PM GMT+08:00
CTUDT00597958 | Refresh | Closed | naik_p | May 15, 2018 5:12:35 PM GMT+08:00 | Protect field contains value as 'Factory' in VoiceSets.cfg and Phrases.cfg files. | WIP6.1 |  |  | Subsystem | Voice Manager |  | Works As Specified / Not A Defect | sachan_a | 0 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00597956 | Refresh | Closed | basavaraju_p | May 15, 2018 2:59:59 PM GMT+08:00 | Ok button is not visible in Exam card manager | WIP6.1 | WIP9 |  | Subsystem |  | N/A | Fixed | kumar_v | 0 | 3 | Console Software | Console4 | NO | Closed | SegReq_WIP9 | Ovation |  | WIP9 | July 23, 2018 6:16:27 PM GMT+08:00
CTUDT00597956 | Refresh | Closed | basavaraju_p | May 15, 2018 2:59:59 PM GMT+08:00 | Ok button is not visible in Exam card manager | WIP6.1 | WIP9 |  | Subsystem |  | N/A | Fixed | kumar_v | 0 | 3 | Console Software | Console4 | NO | Closed | SegReq_WIP9 | Brilliance 64 |  | WIP9 | July 23, 2018 6:16:27 PM GMT+08:00
CTUDT00597956 | Refresh | Closed | basavaraju_p | May 15, 2018 2:59:59 PM GMT+08:00 | Ok button is not visible in Exam card manager | WIP6.1 | WIP9 |  | Subsystem |  | N/A | Fixed | kumar_v | 0 | 3 | Console Software | Console4 | NO | Closed | SegReq_WIP9 | Ingenuity 64 |  | WIP9 | July 23, 2018 6:16:27 PM GMT+08:00
CTUDT00597956 | Refresh | Closed | basavaraju_p | May 15, 2018 2:59:59 PM GMT+08:00 | Ok button is not visible in Exam card manager | WIP6.1 | WIP9 |  | Subsystem |  | N/A | Fixed | kumar_v | 0 | 3 | Console Software | Console4 | NO | Closed | SegReq_WIP9 | IQon1 |  | WIP9 | July 23, 2018 6:16:27 PM GMT+08:00
CTUDT00597952 | Refresh | Closed | nanjunda_s | May 15, 2018 2:12:49 PM GMT+08:00 | iDose level and spectral level are not displayed as planned in images of result viewers and mdu's. | WIP6.1 | WIP 11 | NA | Subsystem | iDose levels | N/A | Fixed | jindal_k | 0 | 2 | Test Design | Console4 | NO | Closed | WIP10SegReq | Ovation |  | WIP 11 | November 29, 2018 7:11:29 PM GMT+08:00
CTUDT00597952 | Refresh | Closed | nanjunda_s | May 15, 2018 2:12:49 PM GMT+08:00 | iDose level and spectral level are not displayed as planned in images of result viewers and mdu's. | WIP6.1 | WIP 11 | NA | Subsystem | iDose levels | N/A | Fixed | jindal_k | 0 | 2 | Test Design | Console4 | NO | Closed | WIP10SegReq | Brilliance 64 |  | WIP 11 | November 29, 2018 7:11:29 PM GMT+08:00
CTUDT00597952 | Refresh | Closed | nanjunda_s | May 15, 2018 2:12:49 PM GMT+08:00 | iDose level and spectral level are not displayed as planned in images of result viewers and mdu's. | WIP6.1 | WIP 11 | NA | Subsystem | iDose levels | N/A | Fixed | jindal_k | 0 | 2 | Test Design | Console4 | NO | Closed | WIP10SegReq | Brilliance Big Bore |  | WIP 11 | November 29, 2018 7:11:29 PM GMT+08:00
CTUDT00597952 | Refresh | Closed | nanjunda_s | May 15, 2018 2:12:49 PM GMT+08:00 | iDose level and spectral level are not displayed as planned in images of result viewers and mdu's. | WIP6.1 | WIP 11 | NA | Subsystem | iDose levels | N/A | Fixed | jindal_k | 0 | 2 | Test Design | Console4 | NO | Closed | WIP10SegReq | Ingenuity 64 |  | WIP 11 | November 29, 2018 7:11:29 PM GMT+08:00
CTUDT00597952 | Refresh | Closed | nanjunda_s | May 15, 2018 2:12:49 PM GMT+08:00 | iDose level and spectral level are not displayed as planned in images of result viewers and mdu's. | WIP6.1 | WIP 11 | NA | Subsystem | iDose levels | N/A | Fixed | jindal_k | 0 | 2 | Test Design | Console4 | NO | Closed | WIP10SegReq | IQon1 |  | WIP 11 | November 29, 2018 7:11:29 PM GMT+08:00
CTUDT00597950 | Refresh | Closed | naik_p | May 15, 2018 2:01:32 PM GMT+08:00 | Logs for Recieved offFocalTableName is not received and also zMeCombineWeightsTabName tag is not present in scan Mdu. | WIP6.1 |  |  | Subsystem |  | Console-SRS.15446 | Works As Specified / Not A Defect | kumar_v | 0 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00597948 | Refresh | Closed | naik_p | May 15, 2018 1:50:26 PM GMT+08:00 | Overview tab displays orientation columns even when the parameter is unselected. | WIP6.1 | WIP9 |  | Subsystem |  | NA | Fixed | bhawna_g | 0 | 4 | Console Software | Console4 | NO | Closed | WIP11SegReq | Ovation |  | WIP9 | July 24, 2018 5:07:38 PM GMT+08:00
CTUDT00597948 | Refresh | Closed | naik_p | May 15, 2018 1:50:26 PM GMT+08:00 | Overview tab displays orientation columns even when the parameter is unselected. | WIP6.1 | WIP9 |  | Subsystem |  | NA | Fixed | bhawna_g | 0 | 4 | Console Software | Console4 | NO | Closed | WIP11SegReq | Brilliance 64 |  | WIP9 | July 24, 2018 5:07:38 PM GMT+08:00
CTUDT00597948 | Refresh | Closed | naik_p | May 15, 2018 1:50:26 PM GMT+08:00 | Overview tab displays orientation columns even when the parameter is unselected. | WIP6.1 | WIP9 |  | Subsystem |  | NA | Fixed | bhawna_g | 0 | 4 | Console Software | Console4 | NO | Closed | WIP11SegReq | Ingenuity 64 |  | WIP9 | July 24, 2018 5:07:38 PM GMT+08:00
CTUDT00597948 | Refresh | Closed | naik_p | May 15, 2018 1:50:26 PM GMT+08:00 | Overview tab displays orientation columns even when the parameter is unselected. | WIP6.1 | WIP9 |  | Subsystem |  | NA | Fixed | bhawna_g | 0 | 4 | Console Software | Console4 | NO | Closed | WIP11SegReq | IQon1 |  | WIP9 | July 24, 2018 5:07:38 PM GMT+08:00
CTUDT00597921 | Refresh | Closed | vinoliya_j | May 10, 2018 9:06:43 PM GMT+08:00 | When there is a MPR + thin slices, and a axial result with minimum thickness and overlapped increment no thin slices are generated for MPR | WIP6.1 | WIP11.1 |  | Subsystem | Thin Slices | N/A | Fixed | jindal_k | 0 | 2 | Test Design | Console4 | NO | Closed | WIP9SegReq | Ingenuity 64 |  | WIP11.1 | November 29, 2018 7:09:06 PM GMT+08:00
CTUDT00597916 | Refresh | Closed | naik_p | May 10, 2018 8:13:04 PM GMT+08:00 | Recon Priority change when new recon is initiated gives incomplete logs or no logs after even after multiple trials | WIP6.1 | WIP16 | NA | Subsystem | Reconstruction Priority | Console-SRS.15659 | Fixed | vishnu_devi | 0 | 2 | Console Software | Console4 | NO | Closed | WIP16SubCom | IQon1 |  | WIP16 | April 10, 2019 3:29:07 PM GMT+08:00
CTUDT00597912 | Refresh | Closed | sengupta_a | May 10, 2018 7:15:43 PM GMT+08:00 | While changing to another scan type with a head/body helical exam and then changing back to head/body scan type, the idose level values does not refresh . | WIP6.1 |  |  | Subsystem |  | N/A | Works As Specified / Not A Defect | harish_n | 0 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00597909 | Refresh | Closed | pB_s | May 10, 2018 6:23:33 PM GMT+08:00 | System error after 'Start Final Recon' button is pressed in all of the Head group EC's | WIP6.1 |  |  | Subsystem | Brain Axial 2D | N/A | Works As Specified / Not A Defect | kumar_v | 0 | 3 |  | Console4 | YES | Closed | SegReq_WIP9 | Ingenuity 64
CTUDT00597905 | Refresh | Closed | pB_s | May 10, 2018 5:33:44 PM GMT+08:00 | On Changing the interface language from preferences to anything other than English, the application is not localized to the selected language | WIP6.1 | WIP12 | we have fixed the Localization compilation issue in EBW4.slp file. | Subsystem | UI Localization | NA | Fixed | sirigiri_s | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP9SegReq | Ovation |  | WIP12 | December 14, 2018 1:49:09 PM GMT+08:00
CTUDT00597905 | Refresh | Closed | pB_s | May 10, 2018 5:33:44 PM GMT+08:00 | On Changing the interface language from preferences to anything other than English, the application is not localized to the selected language | WIP6.1 | WIP12 | we have fixed the Localization compilation issue in EBW4.slp file. | Subsystem | UI Localization | NA | Fixed | sirigiri_s | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP9SegReq | Brilliance 64 |  | WIP12 | December 14, 2018 1:49:09 PM GMT+08:00
CTUDT00597905 | Refresh | Closed | pB_s | May 10, 2018 5:33:44 PM GMT+08:00 | On Changing the interface language from preferences to anything other than English, the application is not localized to the selected language | WIP6.1 | WIP12 | we have fixed the Localization compilation issue in EBW4.slp file. | Subsystem | UI Localization | NA | Fixed | sirigiri_s | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP9SegReq | Brilliance Big Bore |  | WIP12 | December 14, 2018 1:49:09 PM GMT+08:00
CTUDT00597905 | Refresh | Closed | pB_s | May 10, 2018 5:33:44 PM GMT+08:00 | On Changing the interface language from preferences to anything other than English, the application is not localized to the selected language | WIP6.1 | WIP12 | we have fixed the Localization compilation issue in EBW4.slp file. | Subsystem | UI Localization | NA | Fixed | sirigiri_s | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP9SegReq | Ingenuity 64 |  | WIP12 | December 14, 2018 1:49:09 PM GMT+08:00
CTUDT00597905 | Refresh | Closed | pB_s | May 10, 2018 5:33:44 PM GMT+08:00 | On Changing the interface language from preferences to anything other than English, the application is not localized to the selected language | WIP6.1 | WIP12 | we have fixed the Localization compilation issue in EBW4.slp file. | Subsystem | UI Localization | NA | Fixed | sirigiri_s | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP9SegReq | IQon1 |  | WIP12 | December 14, 2018 1:49:09 PM GMT+08:00
CTUDT00597903 | Refresh | Closed | sengupta_a | May 10, 2018 5:10:27 PM GMT+08:00 | Transfer tab in RQM still transfers the high priority results first to the remote device even when priority transfer checkbox is checked. | WIP6.1 |  |  | Subsystem |  | N/A | Works As Specified / Not A Defect | kumar_v | 0 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00597901 | Refresh | Closed | hiremath_b | May 10, 2018 2:39:35 PM GMT+08:00 | Refresh: Able to edit phases for Amplitude binning type in Pulmo Helical scans in online edit and in View2. | WIP6.1 |  |  | Subsystem | Pulmo Helical scans | N/A | Works As Specified / Not A Defect | hiremath_b | 0 | 2 |  | Console4 | NO | Closed | WIP10SegReq | Ingenuity 64
CTUDT00597898 | Refresh | Closed | hiremath_b | May 10, 2018 2:20:55 PM GMT+08:00 | Refresh : Quality Assurance button does not get enabled even after exiting from it. Needs logout to enable it. | WIP6.1 | WIP8 |  | Subsystem |  | NA | Fixed | kumar_v | 0 | 4 | Console Software | Console4 | NO | Closed | SegReq_WIP8.1 | Ovation |  | WIP8 | July 4, 2018 7:15:12 PM GMT+08:00
CTUDT00597898 | Refresh | Closed | hiremath_b | May 10, 2018 2:20:55 PM GMT+08:00 | Refresh : Quality Assurance button does not get enabled even after exiting from it. Needs logout to enable it. | WIP6.1 | WIP8 |  | Subsystem |  | NA | Fixed | kumar_v | 0 | 4 | Console Software | Console4 | NO | Closed | SegReq_WIP8.1 | Ingenuity 64 |  | WIP8 | July 4, 2018 7:15:12 PM GMT+08:00
CTUDT00597898 | Refresh | Closed | hiremath_b | May 10, 2018 2:20:55 PM GMT+08:00 | Refresh : Quality Assurance button does not get enabled even after exiting from it. Needs logout to enable it. | WIP6.1 | WIP8 |  | Subsystem |  | NA | Fixed | kumar_v | 0 | 4 | Console Software | Console4 | NO | Closed | SegReq_WIP8.1 | IQon1 |  | WIP8 | July 4, 2018 7:15:12 PM GMT+08:00
CTUDT00597891 | Refresh | Closed | sengupta_a | May 9, 2018 9:27:02 PM GMT+08:00 | Context menu on cardiac results shows incorrect naming of options for both coronary and functional results. | WIP6.1 | WIP9 |  | Subsystem |  | N/A | Fixed | kumar_v | 0 | 4 | Console Software | Console4 | NO | Closed | WIP9SegReq | Ovation |  | WIP9 | July 23, 2018 6:14:23 PM GMT+08:00
CTUDT00597891 | Refresh | Closed | sengupta_a | May 9, 2018 9:27:02 PM GMT+08:00 | Context menu on cardiac results shows incorrect naming of options for both coronary and functional results. | WIP6.1 | WIP9 |  | Subsystem |  | N/A | Fixed | kumar_v | 0 | 4 | Console Software | Console4 | NO | Closed | WIP9SegReq | Ingenuity 64 |  | WIP9 | July 23, 2018 6:14:23 PM GMT+08:00
CTUDT00597880 | Refresh | Closed | nanjunda_s | May 9, 2018 6:59:39 PM GMT+08:00 | Critical error occurred when First time login to Olevel. | WIP6.1 | WIP9 | NA | Subsystem |  | NA | Fixed | sachan_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP9SegReq | Ovation |  | WIP9 | August 6, 2018 7:57:30 PM GMT+08:00
CTUDT00597880 | Refresh | Closed | nanjunda_s | May 9, 2018 6:59:39 PM GMT+08:00 | Critical error occurred when First time login to Olevel. | WIP6.1 | WIP9 | NA | Subsystem |  | NA | Fixed | sachan_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP9SegReq | Brilliance 64 |  | WIP9 | August 6, 2018 7:57:30 PM GMT+08:00
CTUDT00597880 | Refresh | Closed | nanjunda_s | May 9, 2018 6:59:39 PM GMT+08:00 | Critical error occurred when First time login to Olevel. | WIP6.1 | WIP9 | NA | Subsystem |  | NA | Fixed | sachan_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP9SegReq | Ingenuity 64 |  | WIP9 | August 6, 2018 7:57:30 PM GMT+08:00
CTUDT00597880 | Refresh | Closed | nanjunda_s | May 9, 2018 6:59:39 PM GMT+08:00 | Critical error occurred when First time login to Olevel. | WIP6.1 | WIP9 | NA | Subsystem |  | NA | Fixed | sachan_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP9SegReq | IQon1 |  | WIP9 | August 6, 2018 7:57:30 PM GMT+08:00
CTUDT00597878 | Refresh | Closed | shekhar_p | May 9, 2018 6:49:32 PM GMT+08:00 | Refresh: Performance logs are not available for injector pause operation | WIP6.1 |  | User can check for the appearance of below logs to calculate the performance "injectionInProgress = false"appears in the log In the logs look for the first appearance of : "Ending operation: Update presentation with local data changes" that comes right after the appearance of the string from the previous step. | Subsystem |  | Console-SRS.10133 | Change Proposal | bn_so | 0 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00597878 | Refresh | Closed | shekhar_p | May 9, 2018 6:49:32 PM GMT+08:00 | Refresh: Performance logs are not available for injector pause operation | WIP6.1 |  | User can check for the appearance of below logs to calculate the performance "injectionInProgress = false"appears in the log In the logs look for the first appearance of : "Ending operation: Update presentation with local data changes" that comes right after the appearance of the string from the previous step. | Subsystem |  | Console-SRS.10133 | Change Proposal | bn_so | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00597875 | Refresh | Closed | vinoliya_j | May 9, 2018 6:13:22 PM GMT+08:00 | During Resp Single phase, a wrong message is displayed and the scan is not allowed. | WIP6.1 | WIP12 | User should press 'Start Leaning' and wait till the learning is complete to initiate the scan. | Subsystem | Resp4D CT | Console-SRS.1507 | Fixed | sowmya_k | 0 | 2 | Couch Software | Console4 | NO | Closed | WIP12SubCom | Brilliance 64 |  | WIP12 | December 3, 2018 5:43:40 PM GMT+08:00
CTUDT00597875 | Refresh | Closed | vinoliya_j | May 9, 2018 6:13:22 PM GMT+08:00 | During Resp Single phase, a wrong message is displayed and the scan is not allowed. | WIP6.1 | WIP12 | User should press 'Start Leaning' and wait till the learning is complete to initiate the scan. | Subsystem | Resp4D CT | Console-SRS.1507 | Fixed | sowmya_k | 0 | 2 | Couch Software | Console4 | NO | Closed | WIP12SubCom | Ingenuity 64 |  | WIP12 | December 3, 2018 5:43:40 PM GMT+08:00
CTUDT00597853 | Refresh | Closed | shekhar_p | May 8, 2018 5:43:50 PM GMT+08:00 | Refresh: System crash-Bugrep error | WIP6.1 | WIP7 | The user is able to close the injector dialog before clicking on tools icon | Subsystem | SegReq_WIP8 | NA | Fixed | kumar_l | 0 | 3 | Console Software | Console4 | YES | Closed | WIP8_SubCom | Ovation |  | WIP7 | July 12, 2018 7:47:33 PM GMT+08:00
CTUDT00597853 | Refresh | Closed | shekhar_p | May 8, 2018 5:43:50 PM GMT+08:00 | Refresh: System crash-Bugrep error | WIP6.1 | WIP7 | The user is able to close the injector dialog before clicking on tools icon | Subsystem | SegReq_WIP8 | NA | Fixed | kumar_l | 0 | 3 | Console Software | Console4 | YES | Closed | WIP8_SubCom | Ingenuity 64 |  | WIP7 | July 12, 2018 7:47:33 PM GMT+08:00
CTUDT00597849 | Refresh | Closed | singh_p | May 8, 2018 5:13:12 PM GMT+08:00 | Refersh [iCT] : Duplicate value of Collimation displayed for Bariatric Spine scan type. | WIP6.1 | WIP8 | NA | Subsystem |  | Console-SRS.15961 | Fixed | sachan_a | 0 | 4 | Console Software | Console4 | NO | Closed | SegReq_WIP8.1 | Ovation |  | WIP8 | July 12, 2018 1:37:23 PM GMT+08:00
CTUDT00597844 | Refresh | Closed | elavarasan_s | May 8, 2018 3:15:13 PM GMT+08:00 | Refresh [Ingenuity] - Bosnian, Bulgarian, Croatian, Estonian, Indonesian, Macedonian, Serbian, Slovak, Ukrainian languages are missing in voice manager and Auto voice option. | WIP6.1 | WIP9.1 | NA | Subsystem |  | Console-SRS.16184 | Fixed | sachan_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP10SegReq | Ovation |  | WIP9.1 | August 17, 2018 1:55:54 PM GMT+08:00
CTUDT00597844 | Refresh | Closed | elavarasan_s | May 8, 2018 3:15:13 PM GMT+08:00 | Refresh [Ingenuity] - Bosnian, Bulgarian, Croatian, Estonian, Indonesian, Macedonian, Serbian, Slovak, Ukrainian languages are missing in voice manager and Auto voice option. | WIP6.1 | WIP9.1 | NA | Subsystem |  | Console-SRS.16184 | Fixed | sachan_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP10SegReq | Brilliance 64 |  | WIP9.1 | August 17, 2018 1:55:54 PM GMT+08:00
CTUDT00597844 | Refresh | Closed | elavarasan_s | May 8, 2018 3:15:13 PM GMT+08:00 | Refresh [Ingenuity] - Bosnian, Bulgarian, Croatian, Estonian, Indonesian, Macedonian, Serbian, Slovak, Ukrainian languages are missing in voice manager and Auto voice option. | WIP6.1 | WIP9.1 | NA | Subsystem |  | Console-SRS.16184 | Fixed | sachan_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP10SegReq | Ingenuity 64 |  | WIP9.1 | August 17, 2018 1:55:54 PM GMT+08:00
CTUDT00597842 | Refresh | Closed | elavarasan_s | May 8, 2018 2:19:42 PM GMT+08:00 | Refresh [Ingenuity] - Actual Pitch is displayed less than Min pitch for IAC/Temporal Bones scan type. | WIP6.1 | WIP8 | User can manually change the pitch if required. | Subsystem |  | Console-SRS.10640 | Fixed | sachan_a | 0 | 3 | Console Software | Console4 | YES | Closed | SegReq_WIP8.1 | Ingenuity 64 |  | WIP8 | July 12, 2018 7:55:11 PM GMT+08:00
CTUDT00597834 | Refresh | Closed | vinoliya_j | May 7, 2018 10:41:59 PM GMT+08:00 | Console.SRS-12888 needs to be updated | WIP6.1 | Host SSRS DHF303211 Rev 04 |  | Subsystem |  | Console-SRS.12888 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | DHF303211 Rev04 | Brilliance 64 |  | Host SSRS DHF303211 Rev 04 | September 18, 2018 6:41:02 PM GMT+08:00
CTUDT00597834 | Refresh | Closed | vinoliya_j | May 7, 2018 10:41:59 PM GMT+08:00 | Console.SRS-12888 needs to be updated | WIP6.1 | Host SSRS DHF303211 Rev 04 |  | Subsystem |  | Console-SRS.12888 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | DHF303211 Rev04 | Ingenuity 64 |  | Host SSRS DHF303211 Rev 04 | September 18, 2018 6:41:02 PM GMT+08:00
CTUDT00597830 | Refresh | Closed | shekhar_p | May 7, 2018 9:41:38 PM GMT+08:00 | Refresh: As per the requirement Change parameter ID expected latency shows incorrect value | WIP6.1 | DHF318051 Rev 14 HOST Verification Test Protocol | None | Subsystem | Test Design | N/A | Fixed | ramesh_t | 0 | 4 | Test Design | Console4 | NO | Closed | WIP13SubCom | Ovation |  | DHF318051 Rev 14 HOST Verification Test Protocol | May 2, 2019 8:44:33 PM GMT+08:00
CTUDT00597830 | Refresh | Closed | shekhar_p | May 7, 2018 9:41:38 PM GMT+08:00 | Refresh: As per the requirement Change parameter ID expected latency shows incorrect value | WIP6.1 | DHF318051 Rev 14 HOST Verification Test Protocol | None | Subsystem | Test Design | N/A | Fixed | ramesh_t | 0 | 4 | Test Design | Console4 | NO | Closed | WIP13SubCom | Brilliance 64 |  | DHF318051 Rev 14 HOST Verification Test Protocol | May 2, 2019 8:44:33 PM GMT+08:00
CTUDT00597830 | Refresh | Closed | shekhar_p | May 7, 2018 9:41:38 PM GMT+08:00 | Refresh: As per the requirement Change parameter ID expected latency shows incorrect value | WIP6.1 | DHF318051 Rev 14 HOST Verification Test Protocol | None | Subsystem | Test Design | N/A | Fixed | ramesh_t | 0 | 4 | Test Design | Console4 | NO | Closed | WIP13SubCom | Brilliance Big Bore |  | DHF318051 Rev 14 HOST Verification Test Protocol | May 2, 2019 8:44:33 PM GMT+08:00
CTUDT00597830 | Refresh | Closed | shekhar_p | May 7, 2018 9:41:38 PM GMT+08:00 | Refresh: As per the requirement Change parameter ID expected latency shows incorrect value | WIP6.1 | DHF318051 Rev 14 HOST Verification Test Protocol | None | Subsystem | Test Design | N/A | Fixed | ramesh_t | 0 | 4 | Test Design | Console4 | NO | Closed | WIP13SubCom | Ingenuity 64 |  | DHF318051 Rev 14 HOST Verification Test Protocol | May 2, 2019 8:44:33 PM GMT+08:00
CTUDT00597830 | Refresh | Closed | shekhar_p | May 7, 2018 9:41:38 PM GMT+08:00 | Refresh: As per the requirement Change parameter ID expected latency shows incorrect value | WIP6.1 | DHF318051 Rev 14 HOST Verification Test Protocol | None | Subsystem | Test Design | N/A | Fixed | ramesh_t | 0 | 4 | Test Design | Console4 | NO | Closed | WIP13SubCom | IQon1 |  | DHF318051 Rev 14 HOST Verification Test Protocol | May 2, 2019 8:44:33 PM GMT+08:00
CTUDT00597823 | Refresh | Closed | pB_s | May 7, 2018 5:31:01 PM GMT+08:00 | Console.SRS.12904: Refresh : Upon changing scan time to the maximum limit field does not highlight and does not display a warning message for CTA run off scan type | WIP6.1 |  | NA | Subsystem | Documentation | Console-SRS.12904 | Forwarded to Different Database | bollineni_d | 0 | 4 |  | Support-Simulator | NO | Closed | SegReq_WIP10 | Brilliance 64
CTUDT00597823 | Refresh | Closed | pB_s | May 7, 2018 5:31:01 PM GMT+08:00 | Console.SRS.12904: Refresh : Upon changing scan time to the maximum limit field does not highlight and does not display a warning message for CTA run off scan type | WIP6.1 |  | NA | Subsystem | Documentation | Console-SRS.12904 | Forwarded to Different Database | bollineni_d | 0 | 4 |  | Support-Simulator | NO | Closed | SegReq_WIP10 | Ingenuity 64
CTUDT00597809 | Refresh | Closed | shekhar_p | May 7, 2018 1:59:25 PM GMT+08:00 | As per requirement Use previous surview ID expected latency shows incorrect value | WIP6.1 | WIP11.1 |  | Subsystem |  | Console-SRS.10103 | Fixed | ravindra_a | 0 | 4 | Console Software | Console4 | NO | Closed | WIP11SegReq | Ovation |  | WIP11.1 | November 12, 2018 6:17:30 PM GMT+08:00
CTUDT00597809 | Refresh | Closed | shekhar_p | May 7, 2018 1:59:25 PM GMT+08:00 | As per requirement Use previous surview ID expected latency shows incorrect value | WIP6.1 | WIP11.1 |  | Subsystem |  | Console-SRS.10103 | Fixed | ravindra_a | 0 | 4 | Console Software | Console4 | NO | Closed | WIP11SegReq | Brilliance 64 |  | WIP11.1 | November 12, 2018 6:17:30 PM GMT+08:00
CTUDT00597809 | Refresh | Closed | shekhar_p | May 7, 2018 1:59:25 PM GMT+08:00 | As per requirement Use previous surview ID expected latency shows incorrect value | WIP6.1 | WIP11.1 |  | Subsystem |  | Console-SRS.10103 | Fixed | ravindra_a | 0 | 4 | Console Software | Console4 | NO | Closed | WIP11SegReq | Brilliance Big Bore |  | WIP11.1 | November 12, 2018 6:17:30 PM GMT+08:00
CTUDT00597809 | Refresh | Closed | shekhar_p | May 7, 2018 1:59:25 PM GMT+08:00 | As per requirement Use previous surview ID expected latency shows incorrect value | WIP6.1 | WIP11.1 |  | Subsystem |  | Console-SRS.10103 | Fixed | ravindra_a | 0 | 4 | Console Software | Console4 | NO | Closed | WIP11SegReq | Ingenuity 64 |  | WIP11.1 | November 12, 2018 6:17:30 PM GMT+08:00
CTUDT00597809 | Refresh | Closed | shekhar_p | May 7, 2018 1:59:25 PM GMT+08:00 | As per requirement Use previous surview ID expected latency shows incorrect value | WIP6.1 | WIP11.1 |  | Subsystem |  | Console-SRS.10103 | Fixed | ravindra_a | 0 | 4 | Console Software | Console4 | NO | Closed | WIP11SegReq | IQon1 |  | WIP11.1 | November 12, 2018 6:17:30 PM GMT+08:00
CTUDT00597790 | Refresh | Closed | zhang_t | May 5, 2018 12:25:42 AM GMT+08:00 | "System is still initializing. Please wait..” window popup and displayed around 1minute | WIP6.1 | WIP9 | NA | Subsystem |  | NA | Fixed | ravindra_a | 0 | 2 | Console Software | Console4 | NO | Closed | SegReq_WIP9 | Brilliance 64 |  | WIP9 | September 21, 2018 4:49:11 PM GMT+08:00
CTUDT00597790 | Refresh | Closed | zhang_t | May 5, 2018 12:25:42 AM GMT+08:00 | "System is still initializing. Please wait..” window popup and displayed around 1minute | WIP6.1 | WIP9 | NA | Subsystem |  | NA | Fixed | ravindra_a | 0 | 2 | Console Software | Console4 | NO | Closed | SegReq_WIP9 | Ingenuity 64 |  | WIP9 | September 21, 2018 4:49:11 PM GMT+08:00
CTUDT00597786 | Refresh | Closed | zhang_t | May 4, 2018 10:56:37 PM GMT+08:00 | Installation version mismatched | WIP6.1 | WIP9.1 | N/A | Subsystem | Installation Validation Tool; Verified1 | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_Installation | CIRS | NO | Closed | WIP9.1SubCom | Ovation |  | WIP9.1 | February 19, 2019 2:21:13 AM GMT+08:00
CTUDT00597786 | Refresh | Closed | zhang_t | May 4, 2018 10:56:37 PM GMT+08:00 | Installation version mismatched | WIP6.1 | WIP9.1 | N/A | Subsystem | Installation Validation Tool; Verified1 | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_Installation | CIRS | NO | Closed | WIP9.1SubCom | Brilliance 64 |  | WIP9.1 | February 19, 2019 2:21:13 AM GMT+08:00
CTUDT00597786 | Refresh | Closed | zhang_t | May 4, 2018 10:56:37 PM GMT+08:00 | Installation version mismatched | WIP6.1 | WIP9.1 | N/A | Subsystem | Installation Validation Tool; Verified1 | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_Installation | CIRS | NO | Closed | WIP9.1SubCom | Brilliance Big Bore |  | WIP9.1 | February 19, 2019 2:21:13 AM GMT+08:00
CTUDT00597786 | Refresh | Closed | zhang_t | May 4, 2018 10:56:37 PM GMT+08:00 | Installation version mismatched | WIP6.1 | WIP9.1 | N/A | Subsystem | Installation Validation Tool; Verified1 | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_Installation | CIRS | NO | Closed | WIP9.1SubCom | Ingenuity 64 |  | WIP9.1 | February 19, 2019 2:21:13 AM GMT+08:00
CTUDT00597786 | Refresh | Closed | zhang_t | May 4, 2018 10:56:37 PM GMT+08:00 | Installation version mismatched | WIP6.1 | WIP9.1 | N/A | Subsystem | Installation Validation Tool; Verified1 | N/A | Fixed | richards_b | 0 | 2 | CIRS_SW_Installation | CIRS | NO | Closed | WIP9.1SubCom | IQon1 |  | WIP9.1 | February 19, 2019 2:21:13 AM GMT+08:00
CTUDT00597780 | Refresh | Closed | zhang_t | May 4, 2018 9:58:28 PM GMT+08:00 | When start final recon, system error occurred. | WIP6.1 |  |  | Subsystem | Brain Helical | N/A | Works As Specified / Not A Defect | babu_s | 0 | 3 |  | Console4 | YES | Closed | SegReq_WIP9 | Brilliance 64
CTUDT00597780 | Refresh | Closed | zhang_t | May 4, 2018 9:58:28 PM GMT+08:00 | When start final recon, system error occurred. | WIP6.1 |  |  | Subsystem | Brain Helical | N/A | Works As Specified / Not A Defect | babu_s | 0 | 3 |  | Console4 | YES | Closed | SegReq_WIP9 | Ingenuity 64
CTUDT00597778 | Refresh | Closed | shekhar_p | May 4, 2018 9:19:54 PM GMT+08:00 | Refresh:Auditlogvivewr.exe stopped working | WIP6.1 | WIP9 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP9SubCom | Ovation |  | WIP9 | August 3, 2018 11:57:50 AM GMT+08:00
CTUDT00597778 | Refresh | Closed | shekhar_p | May 4, 2018 9:19:54 PM GMT+08:00 | Refresh:Auditlogvivewr.exe stopped working | WIP6.1 | WIP9 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP9SubCom | Brilliance 64 |  | WIP9 | August 3, 2018 11:57:50 AM GMT+08:00
CTUDT00597778 | Refresh | Closed | shekhar_p | May 4, 2018 9:19:54 PM GMT+08:00 | Refresh:Auditlogvivewr.exe stopped working | WIP6.1 | WIP9 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP9SubCom | Brilliance Big Bore |  | WIP9 | August 3, 2018 11:57:50 AM GMT+08:00
CTUDT00597778 | Refresh | Closed | shekhar_p | May 4, 2018 9:19:54 PM GMT+08:00 | Refresh:Auditlogvivewr.exe stopped working | WIP6.1 | WIP9 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP9SubCom | Ingenuity 64 |  | WIP9 | August 3, 2018 11:57:50 AM GMT+08:00
CTUDT00597778 | Refresh | Closed | shekhar_p | May 4, 2018 9:19:54 PM GMT+08:00 | Refresh:Auditlogvivewr.exe stopped working | WIP6.1 | WIP9 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP9SubCom | IQon1 |  | WIP9 | August 3, 2018 11:57:50 AM GMT+08:00
CTUDT00597775 | Refresh | Closed | babu_s | May 4, 2018 8:23:21 PM GMT+08:00 | [Refresh] : For Axial scans having more than one cycle the CTDIvol value is displayed with units wrongly arranged. | WIP6.1 | WIP9 |  | Subsystem |  | N/A | Fixed | kumar_v | 0 | 4 | Console Software | Console4 | NO | Closed | SegReq_WIP8.1 | Ovation |  | WIP9 | July 9, 2018 7:13:04 PM GMT+08:00
CTUDT00597775 | Refresh | Closed | babu_s | May 4, 2018 8:23:21 PM GMT+08:00 | [Refresh] : For Axial scans having more than one cycle the CTDIvol value is displayed with units wrongly arranged. | WIP6.1 | WIP9 |  | Subsystem |  | N/A | Fixed | kumar_v | 0 | 4 | Console Software | Console4 | NO | Closed | SegReq_WIP8.1 | Brilliance 64 |  | WIP9 | July 9, 2018 7:13:04 PM GMT+08:00
CTUDT00597775 | Refresh | Closed | babu_s | May 4, 2018 8:23:21 PM GMT+08:00 | [Refresh] : For Axial scans having more than one cycle the CTDIvol value is displayed with units wrongly arranged. | WIP6.1 | WIP9 |  | Subsystem |  | N/A | Fixed | kumar_v | 0 | 4 | Console Software | Console4 | NO | Closed | SegReq_WIP8.1 | Ingenuity 64 |  | WIP9 | July 9, 2018 7:13:04 PM GMT+08:00
CTUDT00597775 | Refresh | Closed | babu_s | May 4, 2018 8:23:21 PM GMT+08:00 | [Refresh] : For Axial scans having more than one cycle the CTDIvol value is displayed with units wrongly arranged. | WIP6.1 | WIP9 |  | Subsystem |  | N/A | Fixed | kumar_v | 0 | 4 | Console Software | Console4 | NO | Closed | SegReq_WIP8.1 | IQon1 |  | WIP9 | July 9, 2018 7:13:04 PM GMT+08:00
CTUDT00597769 | Refresh | Closed | nanjunda_s | May 4, 2018 6:13:23 PM GMT+08:00 | VIPR sub stage tag is available for Locator, Tracker, IMR , High E and Low E reconstructions | WIP6.1 | WIP8 | NA | Subsystem | Seg_Req_WIP8 | Console SRS.13863 | Fixed | sachan_a | 0 | 2 | Console Software | Console4 | YES | Closed | WIP8_SubCom | Ovation |  | WIP8 | June 14, 2018 2:03:14 PM GMT+08:00
CTUDT00597769 | Refresh | Closed | nanjunda_s | May 4, 2018 6:13:23 PM GMT+08:00 | VIPR sub stage tag is available for Locator, Tracker, IMR , High E and Low E reconstructions | WIP6.1 | WIP8 | NA | Subsystem | Seg_Req_WIP8 | Console SRS.13863 | Fixed | sachan_a | 0 | 2 | Console Software | Console4 | YES | Closed | WIP8_SubCom | Brilliance 64 |  | WIP8 | June 14, 2018 2:03:14 PM GMT+08:00
CTUDT00597769 | Refresh | Closed | nanjunda_s | May 4, 2018 6:13:23 PM GMT+08:00 | VIPR sub stage tag is available for Locator, Tracker, IMR , High E and Low E reconstructions | WIP6.1 | WIP8 | NA | Subsystem | Seg_Req_WIP8 | Console SRS.13863 | Fixed | sachan_a | 0 | 2 | Console Software | Console4 | YES | Closed | WIP8_SubCom | Ingenuity 64 |  | WIP8 | June 14, 2018 2:03:14 PM GMT+08:00
CTUDT00597769 | Refresh | Closed | nanjunda_s | May 4, 2018 6:13:23 PM GMT+08:00 | VIPR sub stage tag is available for Locator, Tracker, IMR , High E and Low E reconstructions | WIP6.1 | WIP8 | NA | Subsystem | Seg_Req_WIP8 | Console SRS.13863 | Fixed | sachan_a | 0 | 2 | Console Software | Console4 | YES | Closed | WIP8_SubCom | IQon1 |  | WIP8 | June 14, 2018 2:03:14 PM GMT+08:00
CTUDT00597767 | Refresh | Closed | vinoliya_j | May 4, 2018 6:04:30 PM GMT+08:00 | Incorrect series description for the non gated results in Resp 4D CT scan | WIP6.1 | WIP7 |  | Subsystem |  | NA | Fixed | kumar_l | 0 | 2 | Console Software | Console4 | YES | Closed | WIP8_SubCom | Brilliance 64 |  | WIP7 | July 17, 2018 2:27:31 PM GMT+08:00
CTUDT00597767 | Refresh | Closed | vinoliya_j | May 4, 2018 6:04:30 PM GMT+08:00 | Incorrect series description for the non gated results in Resp 4D CT scan | WIP6.1 | WIP7 |  | Subsystem |  | NA | Fixed | kumar_l | 0 | 2 | Console Software | Console4 | YES | Closed | WIP8_SubCom | Ingenuity 64 |  | WIP7 | July 17, 2018 2:27:31 PM GMT+08:00
CTUDT00597763 | Refresh | Closed | blessily | May 4, 2018 3:22:47 PM GMT+08:00 | If all Respiratory steps are removed from the series list and no Pulmo signal is detected and replaced by an Abdomen scan, still the Console displays the Resp Online Viewer. | WIP 6.1 |  | User can manually minimise the Resp Viewer using the minimise button | Subsystem | Resp4D CT | N/A | No Plan to Fix | bandil_j | 0 | 4 |  | Console4 | NO | Closed | WIP13SubCom | Brilliance 64
CTUDT00597763 | Refresh | Closed | blessily | May 4, 2018 3:22:47 PM GMT+08:00 | If all Respiratory steps are removed from the series list and no Pulmo signal is detected and replaced by an Abdomen scan, still the Console displays the Resp Online Viewer. | WIP 6.1 |  | User can manually minimise the Resp Viewer using the minimise button | Subsystem | Resp4D CT | N/A | No Plan to Fix | bandil_j | 0 | 4 |  | Console4 | NO | Closed | WIP13SubCom | Ingenuity 64
CTUDT00597761 | Refresh | Closed | singh_p | May 4, 2018 2:24:01 PM GMT+08:00 | Refresh [iCT] : IMR recon mode is enabled for Ultra High Resolution. | WIP6.1 | WIP8 | The user the can choose not to select IMR with UH resolution. | Subsystem |  | Console-SRS.7682 | Fixed | shruthi_u | 0 | 3 | Console Software | Console4 | NO | Closed | SegReq_WIP8 | Ovation |  | WIP8 | June 14, 2018 8:45:14 PM GMT+08:00
CTUDT00597759 | Refresh | Closed | singh_p | May 4, 2018 1:53:28 PM GMT+08:00 | Refresh [iCT] : 'Detailed' resolution is not available for Head & Body scan types. | WIP6.1 |  | N/A | Subsystem |  | Console-SRS.15255 | Works As Specified / Not A Defect | bn_so | 0 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00597745 | Refresh | Closed | sengupta_a | May 3, 2018 9:08:30 PM GMT+08:00 | [Refresh] ICT: Idose option is available in recon mode when switching to CCT single/continuous from CCT fluoro in Console. | WIP6.1 | WIP8 | N/A | Subsystem |  | Console-SRS.1128 | Fixed | bn_so | 0 | 2 | Console Software | Console4 | YES | Closed | WIP8.1_SubCom | Ovation |  | WIP8 | June 14, 2018 1:45:36 PM GMT+08:00
CTUDT00597741 | Refresh | Closed | elavarasan_s | May 3, 2018 6:50:35 PM GMT+08:00 | Refresh [Ingenuity] - ELSCINTCT_FOCAL_SPOT_RESOLUTION tag value should be ultrafast for surview scans. But it is displaying as Standard. | WIP6.1 | DHF303211 Rev 03 DHF316705 Rev 03 |  | Subsystem |  | Console-SRS.10711, Console-SRS.10725, Console-SRS.11353, Console-SRS.11588, Console-SRS.13164 | Fixed | mistry_h | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | DHF303211 Rev03 | Ovation |  | DHF303211 Rev 03 DHF316705 Rev 03 | September 17, 2018 4:18:37 PM GMT+08:00
CTUDT00597741 | Refresh | Closed | elavarasan_s | May 3, 2018 6:50:35 PM GMT+08:00 | Refresh [Ingenuity] - ELSCINTCT_FOCAL_SPOT_RESOLUTION tag value should be ultrafast for surview scans. But it is displaying as Standard. | WIP6.1 | DHF303211 Rev 03 DHF316705 Rev 03 |  | Subsystem |  | Console-SRS.10711, Console-SRS.10725, Console-SRS.11353, Console-SRS.11588, Console-SRS.13164 | Fixed | mistry_h | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | DHF303211 Rev03 | Brilliance 64 |  | DHF303211 Rev 03 DHF316705 Rev 03 | September 17, 2018 4:18:37 PM GMT+08:00
CTUDT00597741 | Refresh | Closed | elavarasan_s | May 3, 2018 6:50:35 PM GMT+08:00 | Refresh [Ingenuity] - ELSCINTCT_FOCAL_SPOT_RESOLUTION tag value should be ultrafast for surview scans. But it is displaying as Standard. | WIP6.1 | DHF303211 Rev 03 DHF316705 Rev 03 |  | Subsystem |  | Console-SRS.10711, Console-SRS.10725, Console-SRS.11353, Console-SRS.11588, Console-SRS.13164 | Fixed | mistry_h | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | DHF303211 Rev03 | Brilliance Big Bore |  | DHF303211 Rev 03 DHF316705 Rev 03 | September 17, 2018 4:18:37 PM GMT+08:00
CTUDT00597741 | Refresh | Closed | elavarasan_s | May 3, 2018 6:50:35 PM GMT+08:00 | Refresh [Ingenuity] - ELSCINTCT_FOCAL_SPOT_RESOLUTION tag value should be ultrafast for surview scans. But it is displaying as Standard. | WIP6.1 | DHF303211 Rev 03 DHF316705 Rev 03 |  | Subsystem |  | Console-SRS.10711, Console-SRS.10725, Console-SRS.11353, Console-SRS.11588, Console-SRS.13164 | Fixed | mistry_h | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | DHF303211 Rev03 | Ingenuity 64 |  | DHF303211 Rev 03 DHF316705 Rev 03 | September 17, 2018 4:18:37 PM GMT+08:00
CTUDT00597741 | Refresh | Closed | elavarasan_s | May 3, 2018 6:50:35 PM GMT+08:00 | Refresh [Ingenuity] - ELSCINTCT_FOCAL_SPOT_RESOLUTION tag value should be ultrafast for surview scans. But it is displaying as Standard. | WIP6.1 | DHF303211 Rev 03 DHF316705 Rev 03 |  | Subsystem |  | Console-SRS.10711, Console-SRS.10725, Console-SRS.11353, Console-SRS.11588, Console-SRS.13164 | Fixed | mistry_h | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | DHF303211 Rev03 | IQon1 |  | DHF303211 Rev 03 DHF316705 Rev 03 | September 17, 2018 4:18:37 PM GMT+08:00
CTUDT00597739 | Refresh | Closed | shekhar_p | May 3, 2018 6:12:12 PM GMT+08:00 | Refresh :Gemini Couch is displaying in the options and 12 NC is not displaying for couch type | WIP6.1 | Refresh_WIP15 |  | Subsystem | Couch Type | N/A | Fixed | bilal_m | 0 | 4 | All modules marked for CT plus | DI | NO | Closed | WIP10SegReq | Ovation |  | Refresh_WIP15 | April 16, 2019 7:52:57 PM GMT+08:00
CTUDT00597739 | Refresh | Closed | shekhar_p | May 3, 2018 6:12:12 PM GMT+08:00 | Refresh :Gemini Couch is displaying in the options and 12 NC is not displaying for couch type | WIP6.1 | Refresh_WIP15 |  | Subsystem | Couch Type | N/A | Fixed | bilal_m | 0 | 4 | All modules marked for CT plus | DI | NO | Closed | WIP10SegReq | Brilliance 64 |  | Refresh_WIP15 | April 16, 2019 7:52:57 PM GMT+08:00
CTUDT00597739 | Refresh | Closed | shekhar_p | May 3, 2018 6:12:12 PM GMT+08:00 | Refresh :Gemini Couch is displaying in the options and 12 NC is not displaying for couch type | WIP6.1 | Refresh_WIP15 |  | Subsystem | Couch Type | N/A | Fixed | bilal_m | 0 | 4 | All modules marked for CT plus | DI | NO | Closed | WIP10SegReq | Ingenuity 64 |  | Refresh_WIP15 | April 16, 2019 7:52:57 PM GMT+08:00
CTUDT00597739 | Refresh | Closed | shekhar_p | May 3, 2018 6:12:12 PM GMT+08:00 | Refresh :Gemini Couch is displaying in the options and 12 NC is not displaying for couch type | WIP6.1 | Refresh_WIP15 |  | Subsystem | Couch Type | N/A | Fixed | bilal_m | 0 | 4 | All modules marked for CT plus | DI | NO | Closed | WIP10SegReq | IQon1 |  | Refresh_WIP15 | April 16, 2019 7:52:57 PM GMT+08:00
CTUDT00597734 | Refresh | Closed | singh_p | May 3, 2018 2:22:20 PM GMT+08:00 | Refresh [iCT] : Enhancement field is accepting the value beyond the range for non -U filters. | WIP6.1 |  | NA | Subsystem | Enhancement field | Console-SRS.5163 | Works As Specified / Not A Defect | bandil_j | 0 | 2 |  | Console4 | YES | Closed | SegReq_WIP8 | Ovation
CTUDT00597734 | Refresh | Closed | singh_p | May 3, 2018 2:22:20 PM GMT+08:00 | Refresh [iCT] : Enhancement field is accepting the value beyond the range for non -U filters. | WIP6.1 |  | NA | Subsystem | Enhancement field | Console-SRS.5163 | Works As Specified / Not A Defect | bandil_j | 0 | 2 |  | Console4 | YES | Closed | SegReq_WIP8 | Brilliance 64
CTUDT00597734 | Refresh | Closed | singh_p | May 3, 2018 2:22:20 PM GMT+08:00 | Refresh [iCT] : Enhancement field is accepting the value beyond the range for non -U filters. | WIP6.1 |  | NA | Subsystem | Enhancement field | Console-SRS.5163 | Works As Specified / Not A Defect | bandil_j | 0 | 2 |  | Console4 | YES | Closed | SegReq_WIP8 | Ingenuity 64
CTUDT00597734 | Refresh | Closed | singh_p | May 3, 2018 2:22:20 PM GMT+08:00 | Refresh [iCT] : Enhancement field is accepting the value beyond the range for non -U filters. | WIP6.1 |  | NA | Subsystem | Enhancement field | Console-SRS.5163 | Works As Specified / Not A Defect | bandil_j | 0 | 2 |  | Console4 | YES | Closed | SegReq_WIP8 | IQon1
CTUDT00597731 | Refresh | Closed | naik_p | May 3, 2018 2:01:33 PM GMT+08:00 | v4.1.10 : delay between two timed axial scans (end of the first scan to start of the 2nd) is NOT 1000 seconds.(REQ 5146) | WIP6.1 | WIP9 |  | Subsystem |  | NA | Fixed | kumar_v | 0 | 2 | Console Software | Console4 | YES | Closed | WIP10SegReq | Ovation |  | WIP9 | August 9, 2018 1:52:38 PM GMT+08:00
CTUDT00597731 | Refresh | Closed | naik_p | May 3, 2018 2:01:33 PM GMT+08:00 | v4.1.10 : delay between two timed axial scans (end of the first scan to start of the 2nd) is NOT 1000 seconds.(REQ 5146) | WIP6.1 | WIP9 |  | Subsystem |  | NA | Fixed | kumar_v | 0 | 2 | Console Software | Console4 | YES | Closed | WIP10SegReq | IQon1 |  | WIP9 | August 9, 2018 1:52:38 PM GMT+08:00
CTUDT00597693 | Refresh | Closed | elavarasan_s | April 30, 2018 8:11:10 PM GMT+08:00 | Refresh [Ingenuity] - Not able to enable philips_ops, philips_rnd and dc_user. | WIP6.1 | WIP10 |  | Subsystem |  | NA | Fixed | kumar_v | 0 | 2 | Console Software | Console4 | NO | Closed | WIP9SegReq | Ingenuity 64 |  | WIP10 | September 19, 2018 6:16:54 PM GMT+08:00
CTUDT00597691 | Refresh | Closed | elavarasan_s | April 30, 2018 7:22:16 PM GMT+08:00 | Refresh [Ingenuity] - 0.75 value is not displayed in the Rotation time drop down menu for cardiac gated EC. | WIP6.1 | Host SSRS DHF303211 Rev 04 |  | Subsystem |  | Console-SRS.12599 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | DHF303211 Rev04 | Ingenuity 64 |  | Host SSRS DHF303211 Rev 04 | September 18, 2018 6:41:54 PM GMT+08:00
CTUDT00597681 | Refresh | Closed | singh_p | April 28, 2018 1:30:33 AM GMT+08:00 | iBrain filters are not available for "Head" scan types. | WIP6.1 | WIP10 | NA | Subsystem |  | Console-SRS.16885 | Fixed | sachan_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP9_SegReq | Ovation |  | WIP10 | September 17, 2018 5:56:32 PM GMT+08:00
CTUDT00597681 | Refresh | Closed | singh_p | April 28, 2018 1:30:33 AM GMT+08:00 | iBrain filters are not available for "Head" scan types. | WIP6.1 | WIP10 | NA | Subsystem |  | Console-SRS.16885 | Fixed | sachan_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP9_SegReq | Brilliance 64 |  | WIP10 | September 17, 2018 5:56:32 PM GMT+08:00
CTUDT00597681 | Refresh | Closed | singh_p | April 28, 2018 1:30:33 AM GMT+08:00 | iBrain filters are not available for "Head" scan types. | WIP6.1 | WIP10 | NA | Subsystem |  | Console-SRS.16885 | Fixed | sachan_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP9_SegReq | Brilliance Big Bore |  | WIP10 | September 17, 2018 5:56:32 PM GMT+08:00
CTUDT00597681 | Refresh | Closed | singh_p | April 28, 2018 1:30:33 AM GMT+08:00 | iBrain filters are not available for "Head" scan types. | WIP6.1 | WIP10 | NA | Subsystem |  | Console-SRS.16885 | Fixed | sachan_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP9_SegReq | Ingenuity 64 |  | WIP10 | September 17, 2018 5:56:32 PM GMT+08:00
CTUDT00597681 | Refresh | Closed | singh_p | April 28, 2018 1:30:33 AM GMT+08:00 | iBrain filters are not available for "Head" scan types. | WIP6.1 | WIP10 | NA | Subsystem |  | Console-SRS.16885 | Fixed | sachan_a | 0 | 2 | Console Software | Console4 | NO | Closed | WIP9_SegReq | IQon1 |  | WIP10 | September 17, 2018 5:56:32 PM GMT+08:00
CTUDT00597677 | Refresh | Closed | pB_s | April 27, 2018 7:22:17 PM GMT+08:00 | Pitch and rotation time values do not change upon changing of heart rate | WIP6.1 | DHF303211 Rev 03 DHF316705 Rev 03 |  | Subsystem |  | Console SRS.1586, Console-SRS.13377, Console-SRS.3590 | Fixed | mistry_h | 0 | 2 | Requirements Documentation | Console4 | NO | Closed | DHF303211 Rev04 | Ingenuity 64 |  | DHF303211 Rev 03 DHF316705 Rev 03 | September 17, 2018 4:18:20 PM GMT+08:00
CTUDT00597675 | Refresh | Closed | nanjunda_s | April 27, 2018 6:34:25 PM GMT+08:00 | Language name displayed in Voice manager is incorrect for "Chinese-standard mandarin","Slovak" and also for "Portuguese" | WIP6.1 | WIP11 | NA | Subsystem |  | Console-SRS.14521 | Fixed | ravindra_a | 0 | 4 | Console Software | Console4 | NO | Closed | SegReq_WIP9 | Brilliance 64 |  | WIP11 | October 29, 2018 2:57:50 PM GMT+08:00
CTUDT00597675 | Refresh | Closed | nanjunda_s | April 27, 2018 6:34:25 PM GMT+08:00 | Language name displayed in Voice manager is incorrect for "Chinese-standard mandarin","Slovak" and also for "Portuguese" | WIP6.1 | WIP11 | NA | Subsystem |  | Console-SRS.14521 | Fixed | ravindra_a | 0 | 4 | Console Software | Console4 | NO | Closed | SegReq_WIP9 | Ingenuity 64 |  | WIP11 | October 29, 2018 2:57:50 PM GMT+08:00
CTUDT00597675 | Refresh | Closed | nanjunda_s | April 27, 2018 6:34:25 PM GMT+08:00 | Language name displayed in Voice manager is incorrect for "Chinese-standard mandarin","Slovak" and also for "Portuguese" | WIP6.1 | WIP11 | NA | Subsystem |  | Console-SRS.14521 | Fixed | ravindra_a | 0 | 4 | Console Software | Console4 | NO | Closed | SegReq_WIP9 | IQon1 |  | WIP11 | October 29, 2018 2:57:50 PM GMT+08:00
CTUDT00597673 | Refresh | Closed | nanjunda_s | April 27, 2018 6:21:15 PM GMT+08:00 | v4.7.7:Refresh:Electronic Density and Calcium Supression Index pages are blank in Exam Summary with pink background. | WIP6.1 | WIP8 | NA | Subsystem | SegReq_WIP8 | Console-SRS.11711 | Fixed | sachan_a | 0 | 2 | Console Software | Console4 | YES | Closed | WIP8_SubCom | IQon1 |  | WIP8 | June 14, 2018 1:57:43 PM GMT+08:00
CTUDT00597671 | Refresh | Closed | nanjunda_s | April 27, 2018 6:16:54 PM GMT+08:00 | Burn spectral label in Spectral CT viewer is displayed incorrectly for all spectral types, where as it is working as expected in Result viewer. | WIP6.1 | WIP10 |  | Subsystem |  | N/A | Fixed | nanjunda_s | 0 | 2 | Test Design | Support_Spectral | NO | Closed | WIP11SegReq | IQon1 |  | WIP10 | October 22, 2018 2:34:01 PM GMT+08:00
CTUDT00597669 | Refresh | Closed | nanjunda_s | April 27, 2018 6:08:46 PM GMT+08:00 | v4.7.7:Refresh:For HIGH and LOW spectral result types Spatial Weight is 0 in recon mdus instead of 250 | WIP6.1 |  |  | Subsystem |  | Console-SRS.8295 | Works As Specified / Not A Defect | shruthi_u | 0 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00597667 | Refresh | Closed | nanjunda_s | April 27, 2018 6:03:38 PM GMT+08:00 | v4.1.10:Ing-> Head Dual Energy and Body Dual Energy scan types are not available. | WIP6.1 |  |  | Subsystem |  | N/A | Works As Specified / Not A Defect | harish_n | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00597664 | Refresh | Closed | pB_s | April 27, 2018 5:56:37 PM GMT+08:00 | [Refresh] Ingenuity : The Console is not supporting TILT functionality without surview | WIP6.1 | WIP10 | NA | Subsystem |  | Console.SRS-2390 | Fixed | sachan_a | 0 | 2 | Console Software | Console4 | YES | Closed | SegReq_WIP8.1 | Ingenuity 64 |  | WIP10 | September 17, 2018 7:54:51 PM GMT+08:00
CTUDT00597663 | Refresh | Closed | nanjunda_s | April 27, 2018 5:55:49 PM GMT+08:00 | v4.1.10: Ing->default options are not selected in Plan Geometry Limitations section in preferences(REQ 15924) | WIP6.1 | WIP8 | User can manually select the required options under Plan Geometry limitations section | Subsystem |  | Console-SRS.15924 | Fixed | bn_so | 0 | 3 | Console Software | Console4 | NO | Closed | WIP8.1_SubCom | Ingenuity 64 |  | WIP8 | June 28, 2018 12:50:10 PM GMT+08:00
CTUDT00597659 | Refresh | Closed | nanjunda_s | April 27, 2018 5:47:55 PM GMT+08:00 | v4.1.10:Refresh: wip6.1 During installation firewall block error was observed and screens were flickering. | WIP6.1 | WIP10 | N/A | Subsystem |  | N/A | Fixed | mohit_p | 0 | 4 | Console Software | Console4 | NO | Closed | WIP11SegReq | Ovation |  | WIP10 | October 16, 2018 5:26:40 PM GMT+08:00
CTUDT00597659 | Refresh | Closed | nanjunda_s | April 27, 2018 5:47:55 PM GMT+08:00 | v4.1.10:Refresh: wip6.1 During installation firewall block error was observed and screens were flickering. | WIP6.1 | WIP10 | N/A | Subsystem |  | N/A | Fixed | mohit_p | 0 | 4 | Console Software | Console4 | NO | Closed | WIP11SegReq | Brilliance 64 |  | WIP10 | October 16, 2018 5:26:40 PM GMT+08:00
CTUDT00597659 | Refresh | Closed | nanjunda_s | April 27, 2018 5:47:55 PM GMT+08:00 | v4.1.10:Refresh: wip6.1 During installation firewall block error was observed and screens were flickering. | WIP6.1 | WIP10 | N/A | Subsystem |  | N/A | Fixed | mohit_p | 0 | 4 | Console Software | Console4 | NO | Closed | WIP11SegReq | Brilliance Big Bore |  | WIP10 | October 16, 2018 5:26:40 PM GMT+08:00
CTUDT00597659 | Refresh | Closed | nanjunda_s | April 27, 2018 5:47:55 PM GMT+08:00 | v4.1.10:Refresh: wip6.1 During installation firewall block error was observed and screens were flickering. | WIP6.1 | WIP10 | N/A | Subsystem |  | N/A | Fixed | mohit_p | 0 | 4 | Console Software | Console4 | NO | Closed | WIP11SegReq | Ingenuity 64 |  | WIP10 | October 16, 2018 5:26:40 PM GMT+08:00
CTUDT00597659 | Refresh | Closed | nanjunda_s | April 27, 2018 5:47:55 PM GMT+08:00 | v4.1.10:Refresh: wip6.1 During installation firewall block error was observed and screens were flickering. | WIP6.1 | WIP10 | N/A | Subsystem |  | N/A | Fixed | mohit_p | 0 | 4 | Console Software | Console4 | NO | Closed | WIP11SegReq | IQon1 |  | WIP10 | October 16, 2018 5:26:40 PM GMT+08:00
CTUDT00597657 | Refresh | Closed | nanjunda_s | April 27, 2018 5:43:49 PM GMT+08:00 | v4.1.10:Refresh:When logged into CT user 2nd time after installation, observed “system error”,exam application crash and bugrep errors | WIP6.1 |  |  | Subsystem | User Login | NA | Works As Specified / Not A Defect | babu_s | 0 |  |  | iCAP | NO | Closed | WIP9SegReq | Ovation
CTUDT00597655 | Refresh | Closed | nanjunda_s | April 27, 2018 5:39:49 PM GMT+08:00 | v4.1.10:Refresh: Unable to edit phase for functional result in offline ECG viewer | WIP6.1 |  |  | Subsystem |  | N/A | Works As Specified / Not A Defect | harish_n | 0 |  |  | Console4 | NO | Closed |  | Ovation
CTUDT00597655 | Refresh | Closed | nanjunda_s | April 27, 2018 5:39:49 PM GMT+08:00 | v4.1.10:Refresh: Unable to edit phase for functional result in offline ECG viewer | WIP6.1 |  |  | Subsystem |  | N/A | Works As Specified / Not A Defect | harish_n | 0 |  |  | Console4 | NO | Closed |  | Ingenuity 64
CTUDT00597655 | Refresh | Closed | nanjunda_s | April 27, 2018 5:39:49 PM GMT+08:00 | v4.1.10:Refresh: Unable to edit phase for functional result in offline ECG viewer | WIP6.1 |  |  | Subsystem |  | N/A | Works As Specified / Not A Defect | harish_n | 0 |  |  | Console4 | NO | Closed |  | IQon1
CTUDT00597653 | Refresh | Closed | nanjunda_s | April 27, 2018 5:33:53 PM GMT+08:00 | v4.1.10:Refresh: Unable to load all results to “cardiac viewer” from auto launch. | WIP6.1 |  | The user can manually launch the required results to Cardiac Viewer. This is a known workaround. | Subsystem | Auto Launch - Cardiac Viewer | NA | No Plan to Fix | redrouthu_r | 0 | 3 |  | iCAP | NO | Closed |  | Ovation
CTUDT00597653 | Refresh | Closed | nanjunda_s | April 27, 2018 5:33:53 PM GMT+08:00 | v4.1.10:Refresh: Unable to load all results to “cardiac viewer” from auto launch. | WIP6.1 |  | The user can manually launch the required results to Cardiac Viewer. This is a known workaround. | Subsystem | Auto Launch - Cardiac Viewer | NA | No Plan to Fix | redrouthu_r | 0 | 3 |  | iCAP | NO | Closed |  | Ingenuity 64
CTUDT00597653 | Refresh | Closed | nanjunda_s | April 27, 2018 5:33:53 PM GMT+08:00 | v4.1.10:Refresh: Unable to load all results to “cardiac viewer” from auto launch. | WIP6.1 |  | The user can manually launch the required results to Cardiac Viewer. This is a known workaround. | Subsystem | Auto Launch - Cardiac Viewer | NA | No Plan to Fix | redrouthu_r | 0 | 3 |  | iCAP | NO | Closed |  | IQon1
CTUDT00597651 | Refresh | Closed | nanjunda_s | April 27, 2018 5:30:40 PM GMT+08:00 | v4.8- After drawing annotations on the images and trying to save it Bug rep error shows up. Happens in both quick review and CT viewer | WIP6.1 | WIP9 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP9SegReq | Ovation |  | WIP9 | August 3, 2018 11:56:58 AM GMT+08:00
CTUDT00597651 | Refresh | Closed | nanjunda_s | April 27, 2018 5:30:40 PM GMT+08:00 | v4.8- After drawing annotations on the images and trying to save it Bug rep error shows up. Happens in both quick review and CT viewer | WIP6.1 | WIP9 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP9SegReq | Brilliance 64 |  | WIP9 | August 3, 2018 11:56:58 AM GMT+08:00
CTUDT00597651 | Refresh | Closed | nanjunda_s | April 27, 2018 5:30:40 PM GMT+08:00 | v4.8- After drawing annotations on the images and trying to save it Bug rep error shows up. Happens in both quick review and CT viewer | WIP6.1 | WIP9 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP9SegReq | Brilliance Big Bore |  | WIP9 | August 3, 2018 11:56:58 AM GMT+08:00
CTUDT00597651 | Refresh | Closed | nanjunda_s | April 27, 2018 5:30:40 PM GMT+08:00 | v4.8- After drawing annotations on the images and trying to save it Bug rep error shows up. Happens in both quick review and CT viewer | WIP6.1 | WIP9 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP9SegReq | Ingenuity 64 |  | WIP9 | August 3, 2018 11:56:58 AM GMT+08:00
CTUDT00597651 | Refresh | Closed | nanjunda_s | April 27, 2018 5:30:40 PM GMT+08:00 | v4.8- After drawing annotations on the images and trying to save it Bug rep error shows up. Happens in both quick review and CT viewer | WIP6.1 | WIP9 |  | Subsystem |  | NA | Fixed | veerabhadraiah | 0 | 2 | Console USP_EBW | iCAP | NO | Closed | WIP9SegReq | IQon1 |  | WIP9 | August 3, 2018 11:56:58 AM GMT+08:00
CTUDT00597649 | Refresh | Closed | nanjunda_s | April 27, 2018 5:25:42 PM GMT+08:00 | v4.1.10 : Dose Log viewer displays incorrect software version | WIP6.1 | WIP8 | No | Subsystem |  | Console-SRS.15409 | Fixed | kr_m | 0 | 4 | Console Software | Console4 | NO | Closed | WIP8.1_SubCom | Ovation |  | WIP8 | July 25, 2018 6:30:08 PM GMT+08:00
CTUDT00597649 | Refresh | Closed | nanjunda_s | April 27, 2018 5:25:42 PM GMT+08:00 | v4.1.10 : Dose Log viewer displays incorrect software version | WIP6.1 | WIP8 | No | Subsystem |  | Console-SRS.15409 | Fixed | kr_m | 0 | 4 | Console Software | Console4 | NO | Closed | WIP8.1_SubCom | Ingenuity 64 |  | WIP8 | July 25, 2018 6:30:08 PM GMT+08:00
CTUDT00597649 | Refresh | Closed | nanjunda_s | April 27, 2018 5:25:42 PM GMT+08:00 | v4.1.10 : Dose Log viewer displays incorrect software version | WIP6.1 | WIP8 | No | Subsystem |  | Console-SRS.15409 | Fixed | kr_m | 0 | 4 | Console Software | Console4 | NO | Closed | WIP8.1_SubCom | IQon1 |  | WIP8 | July 25, 2018 6:30:08 PM GMT+08:00
CTUDT00597647 | Refresh | Closed | nanjunda_s | April 27, 2018 5:09:43 PM GMT+08:00 | v4.1.10 : Able to load extended results to slab viewer when result is loaded from PD.(req# 10436) | WIP6.1 | Host SSRS DHF303211 Rev04 | NA | Subsystem |  | Console-SRS.10436 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | DHF303211 Rev04 | Ovation |  | Host SSRS DHF303211 Rev04 | September 18, 2018 6:41:36 PM GMT+08:00
CTUDT00597647 | Refresh | Closed | nanjunda_s | April 27, 2018 5:09:43 PM GMT+08:00 | v4.1.10 : Able to load extended results to slab viewer when result is loaded from PD.(req# 10436) | WIP6.1 | Host SSRS DHF303211 Rev04 | NA | Subsystem |  | Console-SRS.10436 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | DHF303211 Rev04 | Brilliance 64 |  | Host SSRS DHF303211 Rev04 | September 18, 2018 6:41:36 PM GMT+08:00
CTUDT00597647 | Refresh | Closed | nanjunda_s | April 27, 2018 5:09:43 PM GMT+08:00 | v4.1.10 : Able to load extended results to slab viewer when result is loaded from PD.(req# 10436) | WIP6.1 | Host SSRS DHF303211 Rev04 | NA | Subsystem |  | Console-SRS.10436 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | DHF303211 Rev04 | Ingenuity 64 |  | Host SSRS DHF303211 Rev04 | September 18, 2018 6:41:36 PM GMT+08:00
CTUDT00597647 | Refresh | Closed | nanjunda_s | April 27, 2018 5:09:43 PM GMT+08:00 | v4.1.10 : Able to load extended results to slab viewer when result is loaded from PD.(req# 10436) | WIP6.1 | Host SSRS DHF303211 Rev04 | NA | Subsystem |  | Console-SRS.10436 | Fixed | bandil_j | 0 | 2 | Host SSRS | Console4 | NO | Closed | DHF303211 Rev04 | IQon1 |  | Host SSRS DHF303211 Rev04 | September 18, 2018 6:41:36 PM GMT+08:00
CTUDT00597601 | Refresh | Closed | bagdady_b | April 23, 2018 9:42:49 PM GMT+08:00 | sCTV: When saving Original series from 2D from SBI the burnt annotations don't appear for the user | WIP6 | WIP18 |  | Subsystem | sCTV - save operation | N/A | Fixed | aviw | 0 | 2 | Console USP_EBW | Support_Spectral | YES | Closed | SegReq_WIP8.1 | IQon1 |  | WIP18 | June 20, 2019 4:21:40 PM GMT+08:00

## Remote Connection issue
INDX | Product | Customer | SN | Country | PN | Date Report | Remarks | Type | Note | BugRep | CTUDT | USP team discussion | One EMS | Action Owner | Actions | Related to CTUDT00630958/CTUDT00631602 | Function owner | Complain status | Complain statusVlookup
1899-12-31 | BBRT | PAVILLON DU BOIS | 1954-10-07 | France | 307569 | 2020-09-30 | Big Bore:Transfert stuck to PACS need to restart Host See attached email to get bugreport | PACS | PACS Transfer | https://fileshare.ehv.campus.philips.com/private/20201029-RE5A3BC4971434E3A8F9C9100961991EE | CTUDT00630958-TBD | Initial investigation looked similar to the defect. But asked for the below query and did not get response . So unable to conclude "We checked the screenshots shared again and we can see that in both the images the patient details for which the PACS is stuck seems to be same (from the blurred images.). Can you please try deleting the first 3 studies for the GEPACS and check the transfer issue is still there or not." | CFM | Sankari | Refer PR430689 | No | BB | #N/A | #N/A
1900-01-01 | BBRT | PAVILLON DU BOIS | 1954-10-07 | France | PR430689 | 2020-11-24 | less than before a installation of a patch but it is not acceptable for our customers. | PACS |  | https://teams.microsoft.com/_#/files/General?threadId=19%3A76cc42a0c8fa4bd7877f7de12141f84c%40thread.tacv2&ctx=channel&context=PR%2523430689&rootfolder=%252Fsites%252FBigBoreRTRefreshFOKCoreTeam%252FShared%2520Documents%252FGeneral%252FFOK%2520CFM(CFR)%2520complaints%2520detail%2520information%252FBig%2520Bore%252FPR%2523430689 | CTUDT00632114 | The issue is because of Broken connection and is not related to CTUDT00630958 |  | Dongping | Check if there is exsting issue or not, if not, open a new one. | TBC - broken connect | BB | #N/A | #N/A
1900-01-02 | CT5000 | St Joseph Westgate Medical Center | 2009-07-07 | United States | 315544 | 2020-10-02 | IT network set up | PACS | PACS config | no need, seeting issue. Not a defect | Not a defect | NA | 116234503 | NA | NA | NA | Service | Closed - No Further Investigation | Closed - No Further Investigation
1900-01-03 | CT5000 | Gesundheit Nord gGmbh | 2009-07-08 | Germany | 267626 | 2020-09-07 | Problems when sending images to the PACS | PACS | No Bugrep | TBD | CTUDT00632127 |  | 116078260 | Brook | Check the issue still happen on site or not, and collect information. | TBD | Engineering | Closed - Non-Complaint | Closed - Non-Complaint
1900-01-04 | CT5000 | The Regional Medical Center | 2009-07-11 | US | 2419-07-05 | 2020-07-15 | Work list populating issue | Worklist | Worklist | https://share.philips.com/sites/STS020181029010003/Shared%20Documents/Forms/AllItems.aspx?sortField=Modified&isAscending=false&viewid=972f714e%2D8b0a%2D4d20%2Da2de%2Dd93340813f00&id=%2Fsites%2FSTS020181029010003%2FShared%20Documents%2FBugReports%2FCTUDT00631283 | CTUDT00631283 | Worklist issue. Belongs to Console | 115744348 | Brook | This AR has closed Please check with Site if issue reproduced or not. | No | Engineering | Closed - Non-Complaint | Closed - Non-Complaint
1900-01-05 | CT5000 | Seton Medical Center Harker He | 2009-07-12 | USA | 339755 | 2020-10-15 | Can't archive to Pacs | PACS | No Bugrep | TBD | CTUDT00632127 | NA | 116303710 | Brook | Check the issue still happen on site or not, and collect information. | TBD | Engineering | Closed - No Further Investigation | Closed - No Further Investigation
1900-01-06 | CT5000 | Manati Medical Center | 2009-07-13 | Puerto Rico | 359739 | 2020-10-23 | Problems with transmission of images to the PACS | PACS | No Bugrep | TBD | CTUDT00632127 |  | 116356503 | Brook | Check the issue still happen on site or not, and collect information. | TBD | Engineering | Closed - No Further Investigation | Closed - No Further Investigation
1900-01-07 | CT5000 | Manati Medical Center | 2009-07-13 | Puerto Rico | 398673 | 2020-11-10 | Does not send images to PACS | PACS | No Bugrep |  | CTUDT00632127 |  | 116458505 | Brook | Check the issue still happen on site or not, and collect information. | TBD | Engineering | Closed - No Further Investigation | Closed - No Further Investigation
1900-01-08 | CT5000 | Manati Medical Center | 2009-07-13 | Puerto Rico | 423095 | 2020-11-19 | PROBLEM WITH IMAGE TRANSFER | PACS | No Bugrep |  | CTUDT00632127 |  | 116519533 | Brook | Check the issue still happen on site or not, and collect information. | TBD | Engineering | Evaluation | Closed - No Further Investigation
1900-01-09 | CT5000 | Manati Medical Center | 2009-07-13 | Puerto Rico | 423104 | 2020-11-19 | Image transfer problems | PACS | No Bugrep |  | CTUDT00632127 |  | 116519563 | Brook | Check the issue still happen on site or not, and collect information. | TBD | Engineering | Evaluation | Closed - No Further Investigation
1900-01-10 | CT5000 | T.C Saglik Bakanligi Bagcilar | 2009-07-16 | Turkey | 267191 | 2020-09-07 | During pacsa patient transfer, the queue freezes. | PACS | PACS Transfer | https://docs.philips.com/:f:/g/personal/brook_shi_philips_com/Etsd-YhLjDdOhrruPVQwk6UBmh8-ayDX3DTKfk8csX83sQ?e=1CyDRM | CTUDT00631602-TBD |  | 116076599 | Sankari | Need to do investigation and confirm link to existing AR or not. If Not, will open a new one. | TBD | Engineering | Closed - No Further Investigation | Closed - No Further Investigation
1900-01-11 | CT5000 | T.C Saglik Bakanligi Bagcilar -> Bağcılar State Hospital | 2009-07-16 | Turkey | 308737 | 2020-10-01 | It freezes and the patient list does not appear. | Worklist | It freezes and the patient list does not appear. | https://docs.philips.com/:f:/g/personal/brook_shi_philips_com/Etsd-YhLjDdOhrruPVQwk6UBmh8-ayDX3DTKfk8csX83sQ?e=1CyDRM | CTUDT00631283 |  | 116224696 | Shubhdeep | Need to check if worklist issue, open AR or not. | No | Engineering | Closed - No Further Investigation | Closed - No Further Investigation
1900-01-12 | CT5000 | T.C Saglik Bakanligi Bagcilar | 2009-07-16 | Turkey | 226030 | 2020-08-10 | Image transfer problem | PACS | PACS Transfer | https://docs.philips.com/:f:/g/personal/brook_shi_philips_com/Etsd-YhLjDdOhrruPVQwk6UBmh8-ayDX3DTKfk8csX83sQ?e=1CyDRM | Not a defect | NA | 115937930 | NA | NA | NA | Engineering | Closed - No Further Investigation | Closed - No Further Investigation
1900-01-13 | CT5000 | T.C Saglik Bakanligi Bagcilar | 2009-07-16 | Turkey | 234535 | 2020-08-17 | Unable to connect to worklist | Worklist | Worklist | https://docs.philips.com/:f:/g/personal/brook_shi_philips_com/Etsd-YhLjDdOhrruPVQwk6UBmh8-ayDX3DTKfk8csX83sQ?e=1CyDRM | Not a defect | NA | 115971072 | NA | NA | NA | Others | Closed - No Further Investigation | Closed - No Further Investigation
1900-01-14 | CT5000 | Bağcılar State Hospital | 2009-07-16 | Turkey | 279732 | 2020-09-10 | PACS transfer stuck issue more than once per day | PACS | PACS Transfer | https://docs.philips.com/:f:/g/personal/brook_shi_philips_com/Etsd-YhLjDdOhrruPVQwk6UBmh8-ayDX3DTKfk8csX83sQ?e=1CyDRM | Not a defect | Imaged are blcoked due to filters applied. This is in IFU. | CFM |  |  | No | Engineering | Closed - Done | Closed - Done
1900-01-15 | CT5000 | SCP GROUPE D IMAGERIE MEDICALE | 2009-07-25 | France | 307590 | 2020-09-07 | PACS transfer stuck issue once per day See attached email to get bugreport | PACS | PACS Transfer | https://fileshare.ehv.campus.philips.com/private/20201029-W2A0229652C4B4BE891BCA8C5B49E3E11 https://fileshare.ehv.campus.philips.com/private/20201029-W1EC23EBCC57E4436BD70EC32F8B52DD1 https://fileshare.ehv.campus.philips.com/private/20201030-W340A1E21594646AEB947F571E9A67942 | CTUDT00632114 | The issue is because of Broken connection and is not related to CTUDT00630958 | CFM | Tammy | Check if there is exsting issue or not, if not, open a new one. | TBC - broken connect | Engineering | Investigation | Closed - Done
1900-01-16 | CT5000 | SCP GROUPE D IMAGERIE MEDICALE | 2009-07-25 | France | 309138 | 2020-09-30 | Problem with image transfer | PACS | PACS Transfer | \\cnhszhphc1ws300\Input\Silva\China FOK Bugreport\TW PR 307590 309138 Transfer stuck once a day to PACS and need to restart Host | CTUDT00632114 | The issue is because of Broken connection and is not related to CTUDT00630958 | 116221699 | Tammy | Check if there is exsting issue or not, if not, open a new one. | TBC - broken connect | Engineering | Closed - No Further Investigation | Closed - No Further Investigation
1900-01-17 | CT5000 | SCP GROUPE D IMAGERIE MEDICALE | 2009-07-25 | France | 372032 | 2020-10-29 | Changing the DICOM parameters. | PACS | PACS Transfer Stuck at Dec-03/15 | https://www.clouddrive.philips.com.cn/file/vxhr0ismieyebmcv83c57da6yer3igdf# password:123 https://share.philips.com/sites/Hotfix2ReleaseComplaints/Shared%20Documents/General/CT5000%20Coudekerque%20France%2040019/DEID_2020_12_15_10_16_05_MESTIRI__SUIVI_VOLUME_DIRECTORY.zip | CTUDT00631602-TBD |  | 116385298 | Sankari | Need to do investigation and confirm link to existing AR or not. If Not, will open a new one. | TBD | Service | Closed - Non-Complaint | Closed - Non-Complaint
1900-01-18 | CT5000 | SCP GROUPE D IMAGERIE MEDICALE | 2009-07-25 | France | 422951 | 2020-11-20 | Changing the DICOM parameters. | PACS | PACS Transfer Stuck at Dec-03/15 | https://www.clouddrive.philips.com.cn/file/vxhr0ismieyebmcv83c57da6yer3igdf# password:123 https://share.philips.com/sites/Hotfix2ReleaseComplaints/Shared%20Documents/General/CT5000%20Coudekerque%20France%2040019/DEID_2020_12_15_10_16_05_MESTIRI__SUIVI_VOLUME_DIRECTORY.zip | CTUDT00631602-TBD |  | 116524671 | Sankari | Need to do investigation and confirm link to existing AR or not. If Not, will open a new one. | TBD | Service | Closed - Non-Complaint | Closed - Non-Complaint
1900-01-19 | CT5000 | SCP GROUPE D IMAGERIE MEDICALE | 2009-07-25 | France | 307590 | 2020-09-30 | Transfer stuck once a day to PACS and need to restart Host | Worklist | worklist | \\cnhszhphc1ws300\Input\Silva\China FOK Bugreport\TW PR 307590 309138 Transfer stuck once a day to PACS and need to restart Host | CTUDT00631283 | Worklist issue. Belongs to Console | CFM | Shubhdeep | confirm | No | Engineering | Investigation | Closed - Done
1900-01-20 | CT5000 | S.C.S. HOSP. GRAL. DE FUERTEVE | 2009-07-26 | Spain | 332249 | 2020-10-13 | it is determined that the slow data transfer is due to a remote system (PACS) | PACS | PACS Transfer | The CT5000 PACS issue in Spain has been resolved. The problem was in the PACS, not our system. \\cnhszhphc1ws300\Input\Silva\China FOK Bugreport\S.C.S. HOSP. GRAL. DE FUERTEVE | Not a defect |  | 116287625 | Sankari | Need to do investigation and confirm link to existing AR or not. If Not, will open a new one. | TBD | Engineering | Closed - No Further Investigation | Closed - No Further Investigation
1900-01-21 | CT5000 | S.C.S. HOSP. GRAL. DE FUERTEVE | 2009-07-26 | Spain | 348670 | 2020-10-19 | Communication failure. It does not send to PACS. It seems to be a problem with PACS. | PACS | PACS Transfer | The CT5000 PACS issue in Spain has been resolved. The problem was in the PACS, not our system. \\cnhszhphc1ws300\Input\Silva\China FOK Bugreport\S.C.S. HOSP. GRAL. DE FUERTEVE | Not a defect |  | 116317508 | Sankari | Need to do investigation and confirm link to existing AR or not. If Not, will open a new one. | TBD | Engineering | Closed - No Further Investigation | Closed - No Further Investigation
1900-01-22 | CT5000 | S.C.S. HOSP. GRAL. DE FUERTEVE | 2009-07-26 | Spain | 354900 | 2020-10-22 | Problems with slowness when sending to PACS from the TAC. | PACS | PACS Transfer | The CT5000 PACS issue in Spain has been resolved. The problem was in the PACS, not our system. \\cnhszhphc1ws300\Input\Silva\China FOK Bugreport\S.C.S. HOSP. GRAL. DE FUERTEVE | Not a defect |  | 116343656 | Sankari | Need to do investigation and confirm link to existing AR or not. If Not, will open a new one. | TBD | Engineering | Closed - Duplicate | Closed - Duplicate
1900-01-23 | CT5000 | S.C.S. HOSP. GRAL. DE FUERTEVE | 2009-07-26 | Spain | 392810 | 2020-11-06 | Network transmission issues. | PACS | PACS Transfer Stuck at Nov-23/Dec-03 | https://www.clouddrive.philips.com.cn/file/hsqqzpkko331q5b2icrjp0xvymbmu28z# password:123 | CTUDT00631602-TBD |  | 116429840 | Sankari | Need to do investigation and confirm link to existing AR or not. If Not, will open a new one. | TBD | Engineering | Closed - No Further Investigation | Closed - No Further Investigation
1900-01-24 | CT5000 | S.C.S. HOSP. GRAL. DE FUERTEVE | 2009-07-26 | Spain | 439686 | 2020-11-27 | shipments to Pacs are intermittently blocked | PACS | PACS Transfer Stuck at Nov-23/Dec-03 | https://www.clouddrive.philips.com.cn/file/hsqqzpkko331q5b2icrjp0xvymbmu28z# password:123 | CTUDT00631602-TBD |  | 116561211 | Sankari | Need to do investigation and confirm link to existing AR or not. If Not, will open a new one. | TBD | Engineering | Investigation - Pending Translation | Closed - Done
1900-01-25 | CT5000 | S.C.S. HOSP. GRAL. DE FUERTEVE | 2009-07-26 | Spain | 466844 | 2020-12-09 | Dicom images export problem | PACS | PACS Transfer Stuck at Nov-23/Dec-03 | https://www.clouddrive.philips.com.cn/file/hsqqzpkko331q5b2icrjp0xvymbmu28z# password:123 | CTUDT00631602-TBD |  | CFM | Sankari | Need to do investigation and confirm link to existing AR or not. If Not, will open a new one. | TBD | Engineering |  | Closed - Done
1900-01-26 | CT5000 | McLaren Macomb | 2009-07-27 | United States | 360185 | 2020-10-24 | Having issues sending some exams to PACS | PACS | No Bugrep | TBD | CTUDT00632127 |  | 116357364 | Brook | Check the issue still happen on site or not, and collect information. | TBD | Engineering | Closed - No Further Investigation | Closed - No Further Investigation
1900-01-27 | CT5000 | NMG Zizhiqu No.4 Hosp | 2009-07-29 | China | 424524 | 2020-11-20 | 02N/Ingenuity CT 728326/unstable image transfer | Printer | No Bugrep |  | Not a defect |  | 116523860 | N/A | Can not Print due to configuration. This is not a PACS transfer issue 1 per week, use GE PACS, 30-40 / day | No | Engineering | Closed - No Further Investigation | Closed - No Further Investigation
1900-01-28 | CT5000 | Central Vermont Medical Center | 2009-07-30 | United States | 371889 | 2020-10-29 | 1398CDFD2 Problems images going to PACS | PACS | No Bugrep |  | CTUDT00632127 |  | 116386841 | Brook | Check the issue still happen on site or not, and collect information. | TBD | Engineering | Evaluation | Closed - No Further Investigation
1900-01-29 | CT5000 | Central Vermont Medical Center | 2009-07-30 | United States | 378368 | 2020-11-02 | Issue sending to PACS | PACS | No Bugrep |  | CTUDT00632127 |  | 116400792 | Brook | Check the issue still happen on site or not, and collect information. | TBD | Engineering | Closed - No Further Investigation | Closed - No Further Investigation
1900-01-30 | CT5000 | Odessa Regional Medical Center | 2009-08-04 | United States | 456455 | 2020-12-05 | 1399CDFD1 Cannot Transfer To PACS & Injector Not Triggering | PACS | No Bugrep |  | CTUDT00632127 |  | 116608163 | Brook | Check the issue still happen on site or not, and collect information. | TBD | Engineering | Evaluation | Closed - Done
1900-01-31 | CT5000 | Centro Hospitalar de Vila Nova | 2009-08-05 | Portugal | 470009 | 2020-12-10 | stop sending images to PACS forcing you to restart the equipment | PACS | No Bugrep |  | CTUDT00632127 |  | 116636550 | Brook | Check the issue still happen on site or not, and collect information. | TBD | Engineering |  | Closed - No Further Investigation
1900-02-01 | CT5000 | HEN Puyang People's Hosp | 2009-08-06 | China | 465836 | 2020-12-09 | The CT host cannot transmit image to PACS (restart in every 5 patients) | PACS | PACS Transfer Stuck at Dec-09/15/16 | https://www.clouddrive.philips.com.cn/file/lopzyqrnfqzby73a5v6wjho282qpwy21# password:123 | CTUDT00631602 | Dose SR not support by LianYing machine, request FSE uncheck Dose SR tranfser to LianYing machine and continue mornitor if still stuch per 5 studies | CFM | Sankari | Additional two bugreport collected, and SpoolerSnapshot.xml shared. Need investigate Brook: please request FSE remove Dose SR transfer to Lian ying device | Yes | Engineering |  | Closed - Duplicate
1900-02-02 | IQon | Lille | 2064-04-12 | France | PR397498 | 2020-11-10 | Slow transfer to ISP | ISP |  | https://share.philips.com/:f:/r/sites/SendingImagestoPACSissues/Shared%20Documents/General?csf=1&web=1&e=1JhLoV | CTUDT00631631 | [25-Nov] - From the available logs we are not seeing any errors which might cause the delay in transfer to ISP. The configured settings done for ISP are also correct. Probably the delay is caused because of the ISP device FSE confirmed , that the slowness is seen on ISP and not on transfer part as the PACS configured for the transfer worked well. There was also upgrade done for ISP on this site after this ISP slowness is observed. |  |  |  |  | iQon | #N/A | #N/A
1900-02-03 | IQon | JS Wuxi No.4 People's Hospital (910018638) | 2064-04-14 | China | PR405360 | 2020-12-11 | image transfer to pacs problem | PACS, ISP |  | https://www.clouddrive.philips.com.cn/file/wun2cy0oxxgvl37l9jnxuhb1hv50ovjw# 密码:123 | Not a defect | [14-Dec] From the logs we see that all the transfers has been initiated to PACS and ISP and the transfers which has been initiated are completed successfully. From the config file we see that the Store Association timeout for PACS has been set to 50. We requested the site to increase the Association Timeout to 120 sec as it is the default value for all Dicom services. |  |  |  |  | iQon | #N/A | #N/A
1900-02-04 | IQon | JS Wuxi No.4 People's Hospital (910018638) | 2064-04-14 | China | N/A | 16 Dec 2020 | image transfer to ISP problem | ISP |  | https://www.clouddrive.philips.com.cn/file/hlnfuv9ggs8pahkrdaeh1l9qpjvxao8z# password 123 | CTUDT00630958 | [22-Dec] the AR timer was 5 and AssociationTimeoutInSeconds was set to 120 according to engineering recommendations The problem time was on 16 Dec 2020 between 15:30 – 17:30 and was resolved by reset spooler Since we see that all the transfers happens to both ISP and PACS, the user can select one of the device (possibly PACS) for store , and later initiate transfer from PACS to ISP as a workaround to reduce the load on transfer and observe whether it reduces the frequency. |  |  |  |  | iQon | #N/A | #N/A
1900-02-05 | IQon | Dirksland/Stichting | 2064-04-15 | Netherlands | PR371303 | 2020-11-10 | Sending to PACS failure | PACS |  | https://share.philips.com/:f:/r/sites/SendingImagestoPACSissues/Shared%20Documents/General?csf=1&web=1&e=1JhLoV | CTUDT00630958 | [25-Nov] - The respective completed log is not there. After the transfer to PACS has resumed after the shutdown/restart of system. Hence this might be a Transfer stuck issue with CTUDT00630958 (fixed SP1) |  |  |  |  | iQon | #N/A | #N/A
1900-02-06 | IQon | Lyon | 2064-04-27 | Franch | PR464118 | 2020-12-09 | image transfer issues | ISP PACS |  | https://teams.microsoft.com/_#/files/General?threadId=19%3A728fd2ceb3984215b4e275f21155cf6b%40thread.tacv2&ctx=channel&context=Lyon%2520IQon%252060019%2520France%2520PR%2523464118&rootfolder=%252Fsites%252FHotfix2ReleaseComplaints%252FShared%2520Documents%252FGeneral%252FLyon%2520IQon%252060019%2520France%2520PR%2523464118 | Not a defect | [10-Dec] - 2 devices configured for Auto store (1 ISP and 1 PACS) Investigation raised that there are no transfer failures/ stuck. All recons and transfers were completed in a reasonable amount of time. Requested from FSE on next similar event to recreate bugrep, screenshots, capture of recon queue and transfer queue. |  |  |  |  | iQon | #N/A | #N/A
1900-02-07 | IQon | JS Province People Hospital (264125) | 2064-04-28 | China | PR465809 | 2020-12-09 | auto store failed |  |  |  | CTUDT00629903 | Workaround: Restart the Host |  |  |  | No | iQon | #N/A | #N/A
1900-02-08 | iCT | Nor Med | 2138-03-16 | Israel | PR271959 | 2020-09-10 | Image transfer to PACS problem. The selected studies in status “waiting in queue” the only workaround reset spooler or host restart | PACS |  | P:\Raw_Database2\Mor\87004 IMAGE TRANSFER TO PACS FAILURE | CTUDT00630958 | Resolved for WIP33 .Fix not complete and needed additional fix from CTUDT00631602 |  |  |  | Yes | iQon | #N/A | #N/A
1900-02-09 | iCT | Nor Med | 2138-03-16 | Israel | PR335321 | 2020-09-10 | Image transfer to PACS problem. The selected studies in status “waiting in queue” the only workaround reset spooler or host restart | PACS |  | P:\Raw_Database2\Mor\87004 IMAGE TRANSFER TO PACS FAILURE | CTUDT00630958 | Resolved for WIP33 .Fix not complete and needed additional fix from CTUDT00631602 |  |  |  | Yes | iQon | #N/A | #N/A
1900-02-10 | BBRT | ONCOLOGIE 78 | 200013 | France | PR429308 | 2020-11-25 | Transfer problem. | PACS | No Bugrep |  | CTUDT00632127 |  |  | Jacky | Check the issue still happen on site or not, and collect information. | TBD | iQon | #N/A | #N/A
1900-02-11 | CT5000 | MURWILLUMBAH DISTRICT HOSPITAL | 341000 | Australia | 357261 | 2020-10-22 | Doctors are reporting that not all the images are being sent at times. TROUBLESHOOTING: 22 Oct 2020 Discussed the network infrastructure which is a mixture of private and public systems. The issue is with regards to studies transferring to Helius PACS (which is a Private managed system). The other PACS node is a GE PACS and is a publicly managed system. Discussed the option to reset the spooler as a possible option but discussed that the network infrastructure the most likely cause for the incomplete transfers. An Engineer is scheduled to attend site on the 23rd Oct and will proceed to reset the spooler as had issues at the time with the olevel password. | PACS | No Bugrep | TBD | CTUDT00632127 |  | 116342442 | Brook | Check the issue still happen on site or not, and collect information. | TBD | Engineering | Closed - No Further Investigation | Closed - No Further Investigation
1900-02-12 | CT5000 | Dalian Branch of Shengjing Hospital of China Medical University | 345001 | China | 354667 | 2020-10-20 | -> PACS transfer stuck issue from 4-5 times per day to once per day | PACS | PACS Transfer | https://www.clouddrive.philips.com.cn/file/e36j96w4k5diiuqerldmrd402m7x89y3# Password: 123 | CTUDT00630958 | Resolved for WIP33 .Fix not complete and needed additional fix from CTUDT00631602 | CFM |  |  | Yes. This is covered in another excel too | Engineering | Closed - Done | Closed - Done
1900-02-13 | CT5000 | Dalian Branch of Shengjing Hospital of China Medical University | 345001 | China | 356232 | 2020-10-21 | Frequent crashes, restart can restore it. Research and development says it may be a defect; transmitting images frequently get stuck. Resetting/restarting the spooler doesn't solve it. After reboot, it is back to normal. PRS TR status: connected Repair Action: Suggest onsite service procedures: It is recommended to collect the bug report and submit escalation to L3. | PACS |  | No | CTUDT00632127 |  |  | Brook |  |  | Engineering | Closed - No Further Investigation | Closed - No Further Investigation
1900-02-14 | CT5000 | Jinggangshan Hospital, the First Affiliated Hospital of Nanchang University, Jiangxi | 345002 | China | 329176 | 2020-10-10 | connect PACS | PACS | PACS config | no need, seeting issue. Not a defect | Not a defect | NA | 116277286 |  |  | NA | Service | Closed - Non-Complaint | Closed - Non-Complaint
1900-02-15 | CT5000 | Jinggangshan Hospital, the First Affiliated Hospital of Nanchang University, Jiangxi | 345002 | China | 348693 | 2020-10-19 | Need to add network node, report error when open LAN config. PRS TR status: connected Repair Action: Suggest onsite service procedures: Need backend to modify the network card setting, remove ipv6. | PACS | PACS Config | no need, seeting issue. Not a defect | Not a defect | NA | 116317254 |  |  | NA | Engineering | Closed - No Further Investigation | Closed - No Further Investigation
1900-02-16 | CT5000 | General Hospital of Jizhong Energy Jingxing Mining Group Co. LTD | 345003 | China | 295724 | 2020-09-23 | 02N/Ingenuity Core128/Elite China/The network transmission does not match the hospital system | PACS | PACS config | no need, seeting issue. Not a defect | Not a defect | NA | 116175398 |  |  | NA | Others | Closed - No Further Investigation | Closed - No Further Investigation
1900-02-17 | CT5000 | SD Laiyang People Hospital | 345005 | China | 279232 | 2020-09-14 | 02N/Core128/Can not print Original Problem Reported by Customer Can not print Problem Description system can not film Troubleshooting Action start spooler Repair Action start spooler | PACS | PACS Transfer Stuck at Nov-14/Dec-06 | https://www.clouddrive.philips.com.cn/file/rneppykohyn6i1uegg40p2uav6oysftq# 密码:123 | CTUDT00632127 |  | 116108800 | Sankari | https://www.clouddrive.philips.com.cn/file/rneppykohyn6i1uegg40p2uav6oysftq# 密码:123 | TBD | Engineering | Closed - No Further Investigation | Closed - No Further Investigation
1900-02-18 | CT5000 | SD Laiyang People Hospital | 345005 | China | 319756 | 2020-10-06 | 02N/Ingenuity Core128/Elite China/worklist fault | Worklist | No Bugrep |  | CTUDT00631283-TBD |  | 116243434 | Shubhdeep | https://www.clouddrive.philips.com.cn/file/rneppykohyn6i1uegg40p2uav6oysftq# 密码:123 | NO | Engineering | Closed - No Further Investigation | Closed - No Further Investigation
1900-02-19 | CT5000 | SD Laiyang People Hospital | 345005 | China | 341665 | 2020-10-15 | image can't transfer | PACS | PACS Transfer Stuck at Nov-14/Dec-06 | https://www.clouddrive.philips.com.cn/file/rneppykohyn6i1uegg40p2uav6oysftq# 密码:123 | CTUDT00632127 |  | 116300206 | Sankari | https://www.clouddrive.philips.com.cn/file/rneppykohyn6i1uegg40p2uav6oysftq# 密码:123 | TBD | Engineering | Closed - No Further Investigation | Closed - No Further Investigation
1900-02-20 | CT5000 | SD Laiyang People Hospital | 345005 | China | 460334 | 2020-11-09 | Image can't transfer to PACS | PACS | PACS Transfer Stuck at Nov-14/Dec-06 | https://www.clouddrive.philips.com.cn/file/rneppykohyn6i1uegg40p2uav6oysftq# 密码:123 | CTUDT00632127 |  | CFM | Sankari | https://www.clouddrive.philips.com.cn/file/rneppykohyn6i1uegg40p2uav6oysftq# 密码:123 | TBD | Engineering |  | Closed - Done
1900-02-21 | CT5000 | XIFENG CHINESE MEDICINE HOSPIT | 345014 | China | 427597 | 2020-11-23 | 02G/Ingenuity Core128/Elite China/Images cannot be transmitted | PACS | No Bugrep |  | CTUDT00632127 |  | 116531139 | Brook | ke.hu@philips.com 18585068636, PRS connected FSE plan to onsite in next days, and collect bugreport | TBD | Engineering | Closed - No Further Investigation | Closed - No Further Investigation
1900-02-22 | CT5000 | SD Anqiu T.C.M Hosp. | 345017 | China | 437631 | 2020-11-26 | 02N/Ingenuity Core128/Elite China/can not transfer image | PACS | PACS Transfer Stuck at Nov-14/Dec-06 | https://www.clouddrive.philips.com.cn/file/wy4uku66h0kpsorozctww912qezcadsz# password:123 | CTUDT00631602-TBD |  | 116554777 | Sankari | Need to do investigation and confirm link to existing AR or not. If Not, will open a new one. | TBD | Engineering | Closed - No Further Investigation | Closed - No Further Investigation
1900-02-23 | CT5000 | SD Anqiu T.C.M Hosp. | 345017 | China | 457075 | 2020-12-05 | 02N/Core128/Can not transfer image | PACS | PACS Transfer Stuck at Nov-14/Dec-06 | https://www.clouddrive.philips.com.cn/file/wy4uku66h0kpsorozctww912qezcadsz# password:123 | CTUDT00631602-TBD |  | 116607368 | Sankari | Need to do investigation and confirm link to existing AR or not. If Not, will open a new one. | TBD | Engineering | Closed - No Further Investigation | Closed - No Further Investigation
1900-02-24 | CT5000 | SX Datong Yanggao People Hosp. | 345025 | China | 462639 | 2020-12-08 | 02N/Ingenuity Core128/Elite China/Cannot transmit data | PACS | No Bugrep |  | CTUDT00632127 |  | 116620331 | Brook | Liu Kun 18636689229, PRS not work corrected, FSE will check schedule to onsite collect bugreport | TBD | Engineering | Closed - No Further Investigation | Closed - No Further Investigation
1900-02-25 | CT5000 | SX Datong Yanggao People Hosp. | 345025 | China | 462732 | 2020-12-08 | 02N/Ingenuity Core128/Elite China/Image transmission failure | PACS | No Bugrep |  | CTUDT00632127 |  | 116620490 | Brook | Liu Kun 18636689229, PRS not work corrected, FSE will check schedule to onsite collect bugreport | TBD | Engineering | Closed - No Further Investigation | Closed - Duplicate
1900-02-26 | CT5000 | SX Datong Yanggao People Hosp. | 345025 | China | 473938 | 2020-12-14 | The CT host cannot transmit image to PACS. | PACS | No Bugrep |  | CTUDT00632127 |  | CFM | Brook | Liu Kun 18636689229, PRS not work corrected, FSE will check schedule to onsite collect bugreport | TBD | Engineering |  | Closed - Done
1900-02-27 | BBRT | La Tour Reseau de soins SA | 760023 | Switzerland | PR419954 | 2020-11-18 | Image surview not pass transfer to PACS. | PACS | No Bugrep |  | CTUDT00632127 |  |  | Jacky | Check the issue still happen on site or not, and collect information. | TBD | BB |  | #N/A
1900-02-28 | BBRT | NHS Lothian | 768000 | France | PR422357 | 2020-11-19 | The simulator is disconnected from RT archive quite frequently | PACS | No Bugrep |  | CTUDT00632127 |  |  | Jacky | Check the issue still happen on site or not, and collect information. | TBD | BB |  | #N/A
1900-03-01 | IQon | Chaing Mai Univ | 860134 | Thailand | PR217800 | 2020-08-03 | Transfer icon didn't showed on some series that already send to PACS | PACS |  | \\ingbtcpic5dtn20\ImageRecon_Thailand Issue\Thailand Site Other Issues\11thBugrep \\ingbtcpic5dtn20\ImageRecon_Thailand Issue\Thailand Site Other Issues\TH iQon of CMU | CTUDT00630689 | Defect fixed in Hotfix2 |  |  |  | No | iQon |  | #N/A
1900-03-02 | IQon | Chaing Mai Univ | 860134 | Thailand | PR217257 | 2020-08-03 | Some series that using "review before save" didn't auto send to PACS after reconstruction finished and show duplicate series (one of the original and nother one is edited ) | Exam Manager + PACS |  | \\ingbtcpic5dtn20\ImageRecon_Thailand Issue\Thailand Site Other Issues\EXAM_MANAGER | CTUDT00630640 | Defect fixed in Hotfix2 |  |  |  | No | iQon |  | #N/A
1900-03-03 | IQon | Chaing Mai Univ | 860134 | Thailand | PR217257 | 2020-08-04 | ExamManager has stopped working error ( Auto recon image didn't send to PACS and create duplicate series). | Exam Manager + PACS | No Bugrep |  | CTUDT00630640 | Defect fixed in Hotfix2 |  |  |  | No | iQon |  | #N/A
1900-03-04 | IQon | Chaing Mai Univ | 860134 | Thailand | PR217257 | 2020-08-05 | ExamManager has stopped working.(Auto recon image didn't send to PACS and create duplicate series). | Exam Manager + PACS | No Bugrep |  | CTUDT00630640 | Defect fixed in Hotfix2 |  |  |  | No | iQon |  | #N/A
1900-03-05 | IQon | Chaing Mai Univ | 860134 | Thailand | PR217257 | 2020-08-10 | ExamManager has stopped working.(Auto recon image didn't send to PACS and create duplicate series). | Exam Manager + PACS |  | \\ilqhfaatc1msclr\Storage\Raw_Database2\TamirZ | CTUDT00630640 | Defect fixed in Hotfix2 |  |  |  | No | iQon |  | #N/A
1900-03-06 | IQon | Chaing Mai Univ | 860134 | Thailand | PR217800 | 2020-08-17 | Image transfer icon didnt show up even the images were send to PACS | PACS |  | \\ingbtcpic5dtn20\ImageRecon_Thailand Issue\Thailand Site Other Issues\TH iQon of CMU | CTUDT00630689 | Defect fixed in Hotfix2 |  |  |  | No | iQon |  | #N/A
1900-03-07 | IQon | Chaing Mai Univ | 860134 | Thailand | PR217800 | 2020-08-19 | Image transfer icon was not show. But all Patient image was sent on PACS | PACS |  | \\ingbtcpic5dtn20\ImageRecon_Thailand Issue\Thailand Site Other Issues\11thBugrep | CTUDT00630689 | Defect fixed in Hotfix2 |  |  |  | No | iQon |  | #N/A
1900-03-08 | IQon | Chaing Mai Univ | 860134 | Thailand | PR217800 | 2020-08-20 | Image transfer icon was not show. But all Patient image was sent on PACS | PACS | No Bugrep |  | CTUDT00630689 | Defect fixed in Hotfix2 |  |  |  | No | iQon |  | #N/A
1900-03-09 | IQon | Chaing Mai Univ | 860134 | Thailand | PR217800 | 2020-08-21 | Image transfer icon was not show. But all Patient image was sent on PACS | PACS | No Bugrep |  | CTUDT00630689 | Defect fixed in Hotfix2 |  |  |  | No | iQon |  | #N/A
1900-03-10 | IQon | Chaing Mai Univ | 860134 | Thailand | N/A | 2020-08-21 | - Encounter SBI series not sent to ISP and PACS |  | No Bugrep | Not reproduciable. | Not a defect |  |  |  |  | NA | iQon |  | #N/A

## EngineeringTrack
Type | Complaint2AR
Defect Track | 103
TBC | 6
TBD | 40
Not a defect | 21
total | 170
ID | Number
CTUDT00632127 | 20
CTUDT00630138/630160 | 13
CTUDT00631602 | 7
TBC-TubAcr CTUDT00630085 | 7
CTUDT00631367/CTUDT00628691 | 5
CTUDT00631156 | 4
CTUDT00631283 | 3
CTUDT00631634 | 3
CTUDT00631669 | 3
CTUDT00631671 | 3
TBC | 3
CTUDT00632114 | 2
CTUDT00632116 | 2
CTUDT00630136 ECR-074469 | 2
CTUDT00630969 | 2
CTUDT00631143 | 2
CTUDT00631165 | 2
CTUDT00632110 | 2
CTUDT00631545 | 2
TBC-Lexan Ring ECR-075975 | 2
CTUDT00632121 | 1
CTUDT00632125 | 1
CTUDT00629920 | 1
CTUDT00630134 ECR-075587 | 1
CTUDT00630864 | 1
CTUDT00630958 | 1
CTUDT00630971 | 1
CTUDT00630991 | 1
CTUDT00631039 | 1
CTUDT00631055 | 1
CTUDT00631058 | 1
CTUDT00631134 | 1
CTUDT00631140 | 1
CTUDT00631142 | 1
CTUDT00631636 | 1
CTUDT00631638 | 1
CTUDT00631717 | 1
TBC_Disk-Bugrep | 1
TBC-Cannot Recon | 1
TBC-CCT Performance | 1

## Service+OPS
Leve1 | PR# | Level2 | Root causes | Countermeasures | Owner | Timing | Count of Complaint (PR#) No. | Percentage | Progress | Note
Service |  | PACS/Printer Setup | Customer request to add new device Image transfer has issue between system and remote devices | Temporary: onsite to configure/modify PACS settings Permanent: share remote skills to enable remote configure possibility | Brook Shi/L3 | Dec 2020 | 1900-01-06 | 0.80049 | 50%
Service |  | Service Parts Order | Parts failure on customer site Order service parts from SPS for DEFOA to fix issue | Temporary: parts order related case -> non a complaint Permanent: not create separately parts order related case | Brook Shi/L3 | Dec 2020 | 1900-01-03 | 0.45743 | 50%
Service |  | PRS Setup | Refresh difference on remoteservice user account activation PRS firewall rule defect | Temporary: on site to fix Permanent: align with L3 to ask FSE to implement workaround and complete setup during installation | Brook Shi/L3 | Dec 2020 | 1900-01-03 | 0.45743 | 75%
APP |  | Q&A/Password mgr/Refresh difference | Refresh difference on security design | More training and education to Application and customer, application to be updated | Kevin/Caroline | Dec 2020 | 1900-01-13 | 1.601
OPS | 247210/342373 | ECG Leads | ECG Leadset type is wrongly sent to customer | Temporary: order part from SPS or locally Permanent: supplier action? | David/Tomer | Dec 2020 | 1900-01-01 | 0.22871 | ?
OPS | 2316-06-23 | De-coupling | Couch FW is not matched with system software, due to couch was de-coupling from system integration | Temporary: Redo couch programing Permanent: ECR-076249 to cut in decoupling process | David/Tomer | Dec 2020 | 1899-12-31 | 0.11436 | ?
OPS | 2316-06-26 | Power Brush_DOA | Power brush is crooked | Temporary: order part from SPS or locally Permanent: | David/Tomer | Dec 2020 | 1899-12-31 | 0.11436 | ?
OPS | 2316-06-22 | DVI Adaptor | DVI Adaptor is missing | Temporary: order part from SPS or locally Permanent: | David/Tomer | Dec 2020 | 1899-12-31 | 0.11436 | ?
OPS | 349706 | Washer size | Washer size is not match with anchor bolt | Temporary: order part from SPS or locally Permanent: | David/Tomer | Dec 2020 | 1899-12-31 | 0.11436 | ?
OPS | 2316-06-27 | Color mismatch | Console table color mismatch with console | Temporary: order part from SPS or locally Permanent: new supplier cut in | David/Tomer | Dec 2020 | 1899-12-31 | 0.11436 | ?
OPS | 2316-07-23 | GCP_DOA | Gantry control panel tilt LED is not light up | Temporary: order part from SPS or locally Permanent: | David/Tomer | Dec 2020 | 1899-12-31 | 0.11436 | ?
OPS | DEFOA | DEFOA CD-29578 | Front cover bulge | Temporary: go through DEFOA process to replace part Permanent: ECR-XXX | Shimrit |  | - | - | ?
OPS | DEFOA | DEFOA CD-30329 | FLT1 ( RED LED ) on the R2D board | Temporary: go through DEFOA process to replace part Permanent: ECR-XXX | Shimrit |  | - | - | ?
OPS | DEFOA | DEFOA CD-30791 | Plinth couch base cover mismatch | Temporary: go through DEFOA process to replace part Permanent: | Shimrit |  | - | - | ?
OPS | DEFOA | DEFOA CD-30934 | X-ray system cannot comply. (RHOST) | Temporary: go through DEFOA process to replace part Permanent: | Shimrit |  | - | - | ?
OPS | DEFOA | DEFOA CD-30513 | Joystick Intv Couch is not delivered with upgrade - missing | Temporary: go through DEFOA process to replace part Permanent: | Shimrit |  | - | - | ?
OPS | 240242&DEFOA | DEFOA CD-30584 | part order for Bruce Schuler | Not applicable, FSE not required | Shimrit |  |  |  | N/A
OPS | DEFOA | DEFOA CD-31191 | Eye phantom holder | Temporary: go through DEFOA process to replace part Permanent: ECR-XXX cut in date, 10-Nov-2020 @Tomer to check | Shimrit |  | - | - | Complete
OPS | DEFOA | DEFOA CD-31889 | Partial Ring artifact (DMS Module) | Temporary: go through DEFOA process to replace part Permanent: | Shimrit |  | - | - | DMS known design issue, no further invstigation
OPS | DEFOA | DEFOA CD-32123 | Eye phantom holder | Temporary: go through DEFOA process to replace part Permanent: ECR-XXX cut in date, 10-Nov-2020 @Tomer to check | Shimrit |  | - | - | Complete
OPS | DEFOA | DEFOA CD-32401 | Big noise in CT Box during gantry rotating (Rotor belt) | Temporary: go through DEFOA process to replace part Permanent: | Shimrit |  | - | - | ?
OPS | DEFOA | DEFOA CD-32457 | BD Calibraiton failed (DMS Module) | Temporary: go through DEFOA process to replace part Permanent: | Shimrit |  | - | - | DMS known design issue, no further invstigation
OPS | 440917&DEFOA | DEFOA CD-32475 | BD Calibraiton failed (DMS Module) | Temporary: go through DEFOA process to replace part Permanent: | Shimrit |  | - | - | DMS known design issue, no further invstigation
OPS | DEFOA | DEFOA CD-31911 | 459801386961 - CABLE, IBOX2 P7 TO USB EXTENDER DC POWER | Temporary: go through DEFOA process to replace part Permanent: Agile - SAP MBP mismatch was corrected from 1-Nov-2020. | Shimrit |  | - | - | Complete
OPS | DEFOA | DEFOA CD-33107 | 453567135962 - EXTENDED PHANTOM PERFORMANCE KIT | Temporary: go through DEFOA process to replace part Permanent: | Shimrit |  | - | - | ?
OPS | 496927 | IMR Cable | Ingenuity CT - Whilst installing an IMR Ingenuity 5000 we noticed the cables for the IMR rack were missing, DEFOA number to be raised | Temporary: order part from SPS or locally Permanent: | David/Tomer

## IB_Status
| 2020-04-17 | [28 empty] | Customer Temperature
| DO NOT USE CUSTOM SORT TO CREATE VIEWS. USE THE FILTER SELECT ONLY CLEAR ALL FILTERS BEFORE ADDING SITES OR INSERTING NEW ROWS. | [28 empty] | Wk23 | Wk24 | Wk25 | Wk26 | Wk27 | Wk28 | WK29 | WK30 | WK31 | WK32 | WK33 | WK34 | WK35 | WK36 | WK37 | WK38 | WK39 | WK40 | WK41 | WK42 | WK43 | WK44 | WK45 | WK46 | WK47 | WK48 | WK49 | WK50 | WK51 | WK52
| Region | Site Name | Country / State |  |  |  | S/N | System New / Upgrade | CT Site Shipping | Version Type | Site ready for Instal? | System arrival date | All Material Arrived (Y/N)? | Start Install | Install Date | Installation time [hours] | Remark | Related FRU P/N | FRU Description | Qty | FRU Cost [$ USD] | During/After Install | Service Workorder | Installation Complete Date | system connect to RSN? | APP Start? | Warrenty | System Operational Date? | IATD/ PA Record Completed? | Wk23 | Wk24 | Wk25 | Wk26 | Wk27 | Wk28 | WK29 | WK30 | WK31 | WK32 | WK33 | WK34 | WK35 | WK36 | WK37 | WK38 | WK39 | WK40 | WK41 | WK42 | WK43 | WK44 | WK45 | WK46 | WK47 | WK48 | WK49 | WK50 | WK51 | WK52 | Total scan Patients | Wk 15 Apr 11 | Wk 16 Apr 18 | Wk 17 Apr 25 | Wk 18 May 02 | Wk 19 May 09 | Wk 20 May 16 | Wk 25 June 20 | Wk 26 June 27 | Wk 27 July 4 | Wk 28 July 11 | Wk 29 July 18 | Wk 30 July 25 | Wk 31 Aug 1 | Wk 32 Aug 8 | Wk 33 Aug 15 | Wk 34 Aug 22 | Wk 35 Aug 29 | Wk 36 Sep.5 | Wk 37 Sep.12 | Wk 38 Sep.19 | Wk 39 Sep.26 | Wk 40 Oct.3 | Wk 41 Oct.10 | Wk 42 Oct.18 | Wk 43 Oct.25 | Wk 44 Nov.1 | Wk 45 Nov.8 | Wk 46 Nov.15 | Wk 47 Nov.22 | Wk 48 Nov.29 | Wk 49 Dec. 6 | Wk 50 Dec.13 | Wk 51 Dec.20 | Wk 52 Dec.27 | Wk 1 Jan.03 | Wk 2 Jan.09 | Wk 3 Jan.16 | Wk 4 Jan.23 | Wk 5 Jan.30 | Wk 6 Feb.06 | Wk 7 Feb.13 | Wk 8 Feb.20 | Wk 9 Feb.27 | Wk 10 Mar.06 | Wk 11 Mar.13 | Wk 12 Mar.20 | Wk 13 Mar.27 | Wk 14 Apr.3 | Wk 15 Apr.10 | Wk 16 Apr.17 | Wk 17 Apr.24 | Wk 18 May.01 | Wk 19 May.08 | Wk 20 May.15 | Wk 21 May.22 | Wk 22 May.29 | Wk 23 Jun.05 | Wk 24 Jun.12 | Wk 25 Jun.19 | Wk 26 Jun.26 | Wk 27 Jul.03 | Wk 28 Jul.10 | Wk 29 Jul.17 | Wk 30 Jul.24
1899-12-31 | NA | The Regional Medical Center | USA | 4398300% | 728323 | TRUE | 40005 | New | Done | 4.1.10 | Yes | 6/1/2020 | Yes | 4398400% | 4398300% | 52 | Electrical schedule had delayed system power on 4 hours. (48h) | 459801415161 | HARMONIZED PHANTOM HEAD SECTION | 1 | 556 | During |  |  | Yes | Y | Monthly Contract to 2024 | 2020-06-16 | Yes | Customer Temperature Wk25 Wk25: 1899-12-31; Customer Temperature WK33 WK33: 1899-12-31; Customer Temperature WK35 WK35: 1900-01-02
1900-01-01 | NA | Willis Knighton Health System | USA | 4399900% | 728323 | TRUE | 40004 | New | Done | 4.1.10 | Yes | 6/15/2020 | Yes | 4399900% | 4399900% | 54.5 | Electrical schedule had delayed system power on 3 hours/Missing items for CCT option/Contractor PM too much people on site 2.5 hours (49h) | [7 empty] | Yes | Y | 26.06.2020 | 2020-07-27 | Yes | Customer Temperature WK39 WK39: 7/8
1900-01-02 | NA | Clay County Medical Center | USA | 4400000% | 728323 | TRUE | 40003 | New | Done | 4.1.10 | Yes | 6/15/2020 | Yes | 4400000% | 4400000% | 65.5 | Front cover DEFOA 8 hours, addition time for connection the PDU with at least 4 hours. (53.5h) | [5 empty] | 50863203 |  | Yes | Y | 29.06.2020 | 2020-06-30 | Yes | Customer Temperature Wk26 Wk26: 1900-01-01; Customer Temperature Wk27 Wk27: 1900-01-01; Customer Temperature WK31 WK31: 1900-01-01; Customer Temperature WK32 WK32: 1900-01-01; Customer Temperature WK39 WK39: 1900-01-01
1900-01-03 | EMEA | T.C Saglik Bakanligi Bagcilar | Turkey | 4401400% | 728323 | TRUE | 40010 | New | Done | 4.1.11 | Yes | 2020-06-30 | Yes | 4401400% | 4401400% | 24 | TBC Equipment # 85287945 | [6 empty] | 20/21 not work | Yes | Y |  | 2020-07-20 | TBD | Customer Temperature WK43 WK43: 1900-01-13
1900-01-04 | NA | St Joseph Westgate Medical Center | USA | 4402500% | 728321 | TRUE | 40001 | New | Done | 4.1.10 | Yes | 7/8/2020 | Yes | 4402500% | 4402500% |  | not finished, not full time FSE in installtion | 453566200321 | PHANTOM WATER BOTTLE | 6 | 3.92 | During |  |  | No | Y | 10.08.2020 |  | 4404300% | bar: Customer Temperature WK29 WK29 to Customer Temperature WK46 WK46
1900-01-05 | NA | Seton Medical Center Harker He | USA | 4403200% | 728323 | TRUE | 40006 | New | Done | 4.1.10 | Yes | 6/15/2020 | Yes | 4403200% | 4403200% |  | TBC if start or not | [7 empty] | No | Y | 31.08.2020 |  |  | Customer Temperature WK39 WK39: 1900-01-08; Customer Temperature WK40 WK40: 1900-01-10
1900-01-06 | EMEA | Gesundheit Nord gGmbh | Germany | 4403900% | 728326 | TRUE | 40002 | New | [5 empty] | 4403900% | 4403900% | bar: Customer Temperature WK31 WK31 to Customer Temperature WK46 WK46
1900-01-07 | NA | METRO MAYAGUEZ INC | Puerto Rico | 4403900% | 728321 | TRUE | 40009 | New | Done | 4.1.10 |  |  |  | 4403900% | 4403900% | [11 empty] | 31.08.2020 |  |  | Customer Temperature WK36 WK36: 1900-01-03; Customer Temperature WK42 WK42: 1900-01-12; Customer Temperature WK46 WK46: 1900-01-18
1900-01-08 | LA | CMD Clinica Integrales Rancagu | Chile | 4405300% | 728323 | TRUE | 40008 | New | Done | 4.1.10 |  |  |  | 4404100% | 4405300% | [11 empty] | 28.09.2020 | bar: Customer Temperature WK31 WK31 to Customer Temperature WK46 WK46
1900-01-09 | NA | Mount Grant General Hospital | USA | 4406000% | 728321 | TRUE | 40012 | New | Done | 4.1.10 |  |  |  | 4406000% | 4406000% | [10 empty] | Y | 30.08.2020 |  |  | Customer Temperature WK50 WK50: 1900-01-19
1900-01-10 | EMEA | SCP GROUPE D IMAGERIE MEDICALE | France | 4406800% | 728321 | TRUE | 40019 | New | [6 empty] | 4406800% | [10 empty] | Y | 07.09.2020 |  |  | Customer Temperature WK37 WK37: 1900-01-05
1900-01-11 | EMEA | S.C.S. HOSP. GRAL. DE FUERTEVE | Spain | 4406800% | 728323 | TRUE | 40020 | New | Done | 4.1.10 |  |  |  | 4406000% | 4406800% | [14 empty] | Customer Temperature WK40 WK40: 1900-01-09; Customer Temperature WK42 WK42: 1900-01-11
1900-01-12 | APAC | Jinggangshan Hospital, the First Affiliated Hospital of Nanchang University, Jiangxi | China | 4406900% | 7283-24 | TRUE | 345002 | New | [6 empty] | 4406900% | [11 empty] | 28.08.2020 | bar: Customer Temperature WK35 WK35 to Customer Temperature WK46 WK46
1900-01-13 | APAC | Huaying Mountain Guangneng General Hospital | China | 4407000% | 7283-24 | TRUE | 345000 | New | [6 empty] | 4407000% | [10 empty] | Y | 02.09.2020 |  |  | Customer Temperature WK37 WK37: 1900-01-04
1900-01-14 | APAC | Dalian Branch of Shengjing Hospital of China Medical University | China | 4407100% | 7283-24 | TRUE | 345001 | New | [6 empty] | 4407100% | [11 empty] | 14.09.2020 |  |  | Customer Temperature WK43 WK43: PACS
1900-01-15 | APAC | General Hospital of Jizhong Energy Jingxing Mining Group Co. LTD | China | 4407500% | 7283-24 | TRUE | 345003 | New | [6 empty] | 4407500% | [11 empty] | 14.09.2020 | bar: Customer Temperature WK36 WK36 to Customer Temperature WK46 WK46
1900-01-16 | APAC | Laiyang hospital | China | 4408000% | 7283-24 | TRUE | 345005 | New | [6 empty] | 4408000% | [10 empty] | Y | 10.09.2020 | bar: Customer Temperature WK36 WK36 to Customer Temperature WK46 WK46
1900-01-17 | APAC | Wuhan Ren 'ai Hospital, Hubei Province | China | 4408100% | 7283-24 | TRUE | 345004 | New | [6 empty] | 4408100% | [11 empty] | 16.09.2020 | bar: Customer Temperature WK37 WK37 to Customer Temperature WK46 WK46
1900-01-18 | NA | Manati Medical Center | Puerto Rico | 4408100% | 728323 | TRUE | 40007 | New | [6 empty] | 4408100% | bar: Customer Temperature WK37 WK37 to Customer Temperature WK46 WK46
1900-01-19 | EMEA | Basaksehir Sehir Hastanesi | Turkey | 4408300% | 728323 | TRUE | 337000 | New | Done | 4.1.10 |  |  |  | No data yet | 4408300% | [14 empty] | Customer Temperature WK43 WK43: 1900-01-14; Customer Temperature WK45 WK45: 1900-01-17
1900-01-20 | APAC | JL Taonan T.C.M. Hosp. | China | 4409100% | 7283-24 | TRUE | 345006 | New | [6 empty] | 4409100% | [11 empty] | 18.09.2020 | bar: Customer Temperature WK38 WK38 to Customer Temperature WK46 WK46
1900-01-21 | NA | HOSPITAL METROPOLITANO SAN GER | Puerto Rico | 4409500% | 728323 | TRUE | 40017 | New | [6 empty] | 4409500% | bar: Customer Temperature WK39 WK39 to Customer Temperature WK46 WK46
1900-01-22 | APAC | Senhai Hospital of Harbin | China | 4409600% | 7283-24 | TRUE | 345007 | New | [6 empty] | 4409600% | [11 empty] | 4409900% | bar: Customer Temperature WK39 WK39 to Customer Temperature WK46 WK46
| EMEA | Centro Hospitalar de Vila Nova | Portugal | [11 empty] | 4409700% | bar: Customer Temperature WK39 WK39 to Customer Temperature WK46 WK46
1900-01-23 | APAC | Henan Fengquan District Dakuai | China | 4409900% | 7283-24 | TRUE | 345009 | New | [6 empty] | 4409900% | bar: Customer Temperature WK39 WK39 to Customer Temperature WK46 WK46
1900-01-24 | APAC | MURWILLUMBAH DISTRICT HOSPITAL | Australia | 4410400% | 728323 | TRUE | 341000 | New | [6 empty] | 4410400% | [14 empty] | Customer Temperature WK43 WK43: 16/17
1900-01-25 | NA | HOSPITAL DR SUSONI INC | Puerto Rico | 4410900% | 728321 | TRUE | 40013 | New | [6 empty] | 4410900% | bar: Customer Temperature WK41 WK41 to Customer Temperature WK46 WK46
1900-01-26 | NA | McLaren Macomb | USA | 4410900% | 728323 | TRUE | 40021 | New | [6 empty] | 4410900% | bar: Customer Temperature WK41 WK41 to Customer Temperature WK46 WK46
1900-01-27 | APAC | NMG Zizhiqu No.4 Hosp | China | 4411800% | 728323 | TRUE | 40023 | New | [6 empty] | 4411800% | bar: Customer Temperature WK42 WK42 to Customer Temperature WK46 WK46
1900-01-28 | NA | Central Vermont Medical Center | USA | 4412000% | 728323 | TRUE | 40024 | New | [6 empty] | 4412000% | bar: Customer Temperature WK42 WK42 to Customer Temperature WK46 WK46
1900-01-29 | APAC | SD Chest Hosp. | China | 4412600% | 728323 | TRUE | 40025 | New | [6 empty] | 4412600% | bar: Customer Temperature WK43 WK43 to Customer Temperature WK46 WK46
1900-01-30 | APAC | Hebei Laonianbing Hospital | China | 4412800% | 7283-24 | 345015 | 345015 | New | [6 empty] | 4412800% | bar: Customer Temperature WK43 WK43 to Customer Temperature WK46 WK46
1900-01-31 | APAC | XIFENG CHINESE MEDICINE HOSPIT | China | 4413100% | 728323 | 345014 | 345014 | New | [6 empty] | 4413100% | bar: Customer Temperature WK44 WK44 to Customer Temperature WK46 WK46
1900-02-01 | APAC | SD Anqiu T.C.M Hosp. | China | 4413100% | 7283-24 | 345017 | 345017 | New | [6 empty] | 4413100% | bar: Customer Temperature WK44 WK44 to Customer Temperature WK46 WK46
1900-02-02 | APAC | Fujian Medical Device Co. LTD. | China | 4413300% | 728323 | 345010 | 345010 | New | [6 empty] | 4413300% | bar: Customer Temperature WK44 WK44 to Customer Temperature WK46 WK46
1900-02-03 | APAC | Guanzhou Panyu District Third | China | 4413500% | 7283-24 | 345013 | 345013 | New | [6 empty] | 4413500% | bar: Customer Temperature WK44 WK44 to Customer Temperature WK46 WK46
1900-02-04 | APAC | HLJ Huachuan People Hosp. | China | 4413600% | 7283-24 | 345018 | 345018 | New | [6 empty] | 4413600% | bar: Customer Temperature WK44 WK44 to Customer Temperature WK46 WK46
1900-02-05 | EMEA | Ministerstwo Zdrowia | PL | 4413700% |  | 341001 | 341001 | [7 empty] | 4413700% | bar: Customer Temperature WK45 WK45 to Customer Temperature WK46 WK46
1900-02-06 | NA | VA Medical Center Biloxi | USA | 4413800% | 7283-24 | 2009-07-24 | 40018 | [7 empty] | 4413800% | bar: Customer Temperature WK45 WK45 to Customer Temperature WK46 WK46
1900-02-07 | APAC | HEB Caofeidian Dist. Hosp. | China | 4414000% | 7283-24 | 345016 | 345016 | [7 empty] | 4414000% | bar: Customer Temperature WK45 WK45 to Customer Temperature WK46 WK46
1900-02-08 | APAC | Fujian Medical Device Co. LTD. | China | 4414700% | 7283-24 | 345011 | 345011 | [7 empty] | 4414700% | bar: Customer Temperature WK46 WK46
1900-02-09 | APAC | Fujian Medical Device Co. LTD. | China | 4414700% | 7283-24 | 345012 | 345012 | [7 empty] | 4414700% | bar: Customer Temperature WK46 WK46
1900-02-10 | APAC | Suzhou Guangci Tumor Hospital | China |  | 7283-24 | 4414700% | [8 empty] | 4414900% | bar: Customer Temperature WK46 WK46
1900-02-11 | EMEA | STEPPING HILL HOSPITAL | United Kingdom |  | 7283-24 | [9 empty] | 4415100%
1900-02-12 | EMEA | A.O. SPEDALI CIVILI DI BRESCIA | Italy |  | 7283-24 | [9 empty] | 4415100%
1900-02-13 | APAC | SX Datong Yanggao People Hosp. | China |  | 7283-24
1900-02-14 | APAC | HEN Puyang People's Hosp | China |  | 7283-24 | [9 empty] | 1899-12-31 | [14 empty] | Customer Temperature WK50 WK50: PACS
1900-02-15 | [14 empty] | 1900-01-01
1900-01-02
1900-01-03
1900-01-04
1900-01-04
1899-12-31
1900-01-01
1900-01-02
[23 empty] | 1900-01-03 | Linkage broken on site | [5 empty] | Customer Temperature WK30 WK30: 1900-01-07; Customer Temperature WK31 WK31: New license file not installed on system
[23 empty] | 1900-01-04 | Couch travel range issue | [5 empty] | Customer Temperature WK30 WK30: 1900-01-08; Customer Temperature WK31 WK31: Table locked.
|  | Clay County Medical Center | [20 empty] | 1900-01-05 | Couch power supply failure | [10 empty] | Site Name | Country / State | ING | System | Installation Complete | Version Type | Key Market Contact Person | WK23 | WK24 | Wk25 | WK26 | WK27 | WK28 | WK29 | WK30 | WK31 | WK32 | WK33 | WK34 | WK35 | WK36 | [5 empty] | NA | The Regional Medical Center | ORANGEBURG | SC | USA | 1014823 | 4514513299 | 6/1/2020 | 5/10/2020 | Pick up 5/14 | 728323 | Ingenuity CT 728326 | 2009-07-11 | HF | 2020-06-01 | 2020-06-16 | Completed | HF
[23 empty] | 1900-01-06 | Customer forget password | [12 empty] | S/N | New / Upgrade | [Date] | [21 empty] | NA | Willis Knighton Health System | SHREVEPORT | LA | USA | 1072902 | 4514878790 | 2020-06-15 | 5/4/2020 | Pick up 5/10 | 728323 | Ingenuity CT 728326 | 2009-07-10 | HF | 2020-06-17 | 2020-07-27 | Completed | #N/A |  |  | 2009-07-11 | New
[78 empty] | #N/A |  |  | 2009-07-10 | New
[35 empty] | The Regional Medical Center | USA | 2009-07-11 | New | 2020-06-02 | 4.1.10 | Nick Stephenson | [19 empty] | NA | Clay County Medical Center | CLAY CENTER | KS | USA | 1010629 | 4514488405 | 2020-06-15 | 5/4/2020 | Pick up 5/10 | 728323 | Ingenuity CT 728326 | 2009-07-09 | HF | 2020-06-18 | 2020-06-30 | Completed | #N/A |  |  | 2009-07-09 | New
[35 empty] | Willis Knighton Health System | USA | 2009-07-10 | New | 2020-06-17 | 4.1.10 | Cristopher Charlton | [19 empty] | EMEA | T.C Saglik Bakanligi Bagcilar -> Bağcılar State Hospital | ISTANBUL | TR | Turkey | 1114350 | 4515160371 | 2020-06-30 | 2020-06-14 | Pick up 6/21 | 728323 | Ing Elite System (R) | 2009-07-16 | HF | 2020-07-02 | 2020-07-20 | Completed | HF |  |  | 2009-07-16 | New | bar: Wk25 to WK36
[35 empty] | Clay County Medical Center | USA | 2009-07-09 | New | 2020-06-18 | 4.1.10 | Craig Denny | [19 empty] | NA | St Joseph Westgate Medical Center | Glendale | AZ | USA | 1024871 | 4514581808 | 7/8/2020 | 5/20/2020 | Pick up 5/20 | 728321 | CT Ingenuity Core | 2009-07-07 | HF | 2020-07-13 | 2020-10-05 | -> APP | #N/A |  |  | 2009-07-07 | New | bar: Wk25 to WK36
[25 empty] | Column Labels | [9 empty] | T.C Saglik Bakanligi Bagcilar | Turkey | 2009-07-16 | New | 2020-07-02 | 4.1.10 | [20 empty] | NA | Seton Medical Center Harker He | Temple | TX | USA | 1113748 | 4515157906 | 2020-06-09 | 5/25/2020 | Pick up 5/28 | 728323 | Ingenuity CT 728326 | 2009-07-12 | HF | 2020-07-20 | 2020-09-01 | APP | HF |  |  | 2009-07-12 | New | bar: WK27 to WK36
[24 empty] | Row Labels | 1900-01-23 | 1900-02-20 | 1900-02-22 12:00 | 1900-03-05 12:00 | Grand Total | [5 empty] | St Joseph Westgate Medical Center | USA | 2009-07-07 | New | 2020-07-13 | 4.1.10 | Kirt Nitsche | [19 empty] | NA | METRO MAYAGUEZ INC | Mayaguez | PR | Puerto Rico | 1102238 | 4515077102 | 2020-07-20 | 2020-06-14 | Pick up 6/15 | 728321 | Ing Core System (R) | 2009-07-15 | HF | 2020-07-27 | 2020-08-18 | APP | HF |  |  | 2009-07-15 | New
[24 empty] | 40003 | [10 empty] | Seton Medical Center Harker He | USA | 2009-07-12 | New | 2020-07-20 | 4.1.10 | Sean McGrath | [19 empty] | EMEA | Gesundheit Nord gGmbh | Bremen | DE | Germany | 1105795 | 4515107307 | 2020-07-13 | 2020-06-17 | Pick up 6/21 | 728326 | Ing Elite IMR System (R) | 2009-07-08 | HF | 2020-07-27 | 2020-08-17 | APP | #N/A
[24 empty] | 40004 | [10 empty] | METRO | USA | 2009-07-15 | New | 2020-07-27 | 4.1.10 | Virgilio Valencia | [19 empty] | LA | CMD Clinica Integrales Rancagu | Santiago | CL | Chile | 1082161 | 4514933061 | 2020-07-25 | 2020-06-01 | Pick up 6/2 | 728323 | Ing Elite System (R) | 2009-07-14 | HF | 2020-08-10 | 2020-08-31 | APP | HF |  |  | 2009-07-14 | New
[24 empty] | 40005 | [10 empty] | MAYAGUEZ INC | [25 empty] | NA | Mount Grant General Hospital | HAWTHORNE | NV | USA | 1095204 | 4515019084 | 2020-08-03 | 2020-06-15 | Pick up 6/22 | 728321 | Ing Core System (R) | 2009-07-18 | HF | 2020-08-17 | 2020-08-25 | APP Done | HF |  |  | 2009-07-08 | New
[24 empty] | 40010 | [10 empty] | CMD Clinica Integrales Rancagu | CHile | 2009-07-14 | New | 2020-07-20 | 4.1.10 | Diego Valencia | [19 empty] | EMEA | S.C.S. HOSP. GRAL. DE FUERTEVE | PUERTO DEL ROSARIO | ES | Spain | 1139379 | 4515319114 | 2020-09-03 | 2020-08-05 | Pick up 8/9 | 728323 | Ing Elite System (R) | 2009-07-26 | HF | 2020-08-25 | 2020-09-21 | -> APP | HF
[24 empty] | Grand Total | [10 empty] | Gesundheit Nord gGmbh |  | 2009-07-08 | New | 2020-08-07 | 4.1.10 | N/A | [19 empty] | EMEA | SCP GROUPE D IMAGERIE MEDICALE | Dunkerque | FR | France | 1117152 | 4515179983 | 2020-08-28 | 2020-07-29 | Pick up 7/30 | 728321 | Ing Core System (R) | 2009-07-25 | HF | 2020-08-25 | 2020-08-31 | APP | HF
[61 empty] | APAC | Jinggangshan Hospital, the First Affiliated Hospital of Nanchang University, Jiangxi | Jinggangshan |  | China | 1161327 |  | 2020-08-24 | 2020-08-21 | Shipped | 7283-24 | Ingenuity Core128/Elite China | 345002 | SZ | 2020-08-26 | 2020-08-31 | APP | SZ
[61 empty] | APAC | Huaying Mountain Guangneng General Hospital | Sichuan |  | China | 1038972 |  | 2020-08-25 | 2020-08-21 | Shipped | 7283-24 | Ingenuity Core128/Elite China | 345000 | SZ | 2020-08-27 | 2020-09-04 | -> APP | HF
[61 empty] | APAC | General Hospital of Jizhong Energy Jingxing Mining Group Co. LTD | Jizhong |  | China | 1095408 |  | 2020-08-31 | 2020-08-28 | Shipped | 7283-24 | Ingenuity Core128/Elite China | 345003 | SZ | 2020-09-01 | No date yet | Installating | SZ
[61 empty] | APAC | Wuhan Ren 'ai Hospital, Hubei Province | Wuhan |  | China | 1158838 |  | 2020-09-02 | 2020-08-28 | Shipped | 7283-24 | Ingenuity Core128/Elite China | 345004 | SZ | 2020-09-02 | No date yet | Installating | SZ
[61 empty] | APAC | Dalian Branch of Shengjing Hospital of China Medical University | Dalian |  | China | 1141569 |  | 2020-08-28 | 2020-08-21 | Shipped | 7283-24 | Ingenuity Core128/Elite China | 345001 | SZ | 2020-08-28 | 2020-08-31 | APP | HF
[43 empty] | 1899-12-31 | Couch Movement issue | FCO72800735 | Done
[43 empty] | 1900-01-01 | Front Cover DEFOA | ECR-075587 Design improvement | Done
[43 empty] | 1900-01-02 | Gantry Shuts off | ECR-075587 Cover design improvement (E-127321) | Done
[43 empty] | 1900-01-03 | Linkage broken on site | Under investigation, CTUDT00631671 | WIP
|  | S/N | Installation time [hours] |  |  |  | Target | Version Type | Site ready for Instal? |  | All Material Arrived (Y/N)? | Start Install | Installation time [hours] | [28 empty] | WK24: 1900-01-04; Wk25: Couch travel range too short; WK26: ECR-076784 Couch MTS Update; WK27: Done
|  | 2009-07-11 | 52 |  |  |  | 51 | 4.1.10 | Yes |  | Yes | 4398400% | 52 | [28 empty] | WK24: 1900-01-05; Wk25: Couch power supply failure; WK26: False alarm with these cases. Bad contact in front panel. No parts were replaced.; WK27: N/A
|  | 2009-07-10 | 54.5 |  |  |  | 51 | 4.1.10 | Yes |  | Yes | 4399900% | 54.5 | [28 empty] | WK24: 1900-01-06; Wk25: Customer forget password; WK26: Reset password, Clinical/service training/IFU already included instruction; WK27: N/A
|  | 2009-07-09 | 65.5 |  |  |  | 51 | 4.1.10 | Yes |  | No | 4400000% | 65.5 | [28 empty] | WK24: 1900-01-07; Wk25: License file not installed on system; WK26: Field not download the correct license; WK27: N/A
|  | 2009-07-16 | TBC |  |  |  | 51 | 4.1.11 | Yes |  | Yes | 4401400% | TBC | [28 empty] | WK24: 1900-01-08; Wk25: Table locked.; WK26: The cover was damaged by the customer; WK27: N/A
|  | 2009-07-07 | TBC |  |  |  | 51 | 4.1.10 | Yes |  | Yes | 4402500% | TBC | [28 empty] | WK24: 1900-01-09; Wk25: Shows ring-shaped artifact; WK26: Ring artifacts due a defective modules. To be replaced; WK27: WIP
|  | 2009-07-12 | TBC |  |  |  | 51 | 4.1.10 | Yes |  | Yes | 4403200% | TBC | [28 empty] | WK24: 1900-01-10; Wk25: Slow PACS transfer; WK26: Under investigation, remote system (PACS) malfunction; WK27: WIP
[43 empty] | 1900-01-11 | System won't power up | "453566496941,RELAY,24VDC PLUGIN W/COVER20 Amp fuse blown to system" | WIP
[43 empty] | 1900-01-12 | Not shooting | sound is heard when the bed moves back and forth | WIP
[43 empty] | 1900-01-13 | Table error | instant mechanical impingement due to the weight of the patient | WIP
Note on Q3: Shi, Brook: Include Install+Option+Variance=51
Note on AG4: Shi, Brook: Couch motion issue
Note on AH6: Shi, Brook: Front Cover DEFOA
Note on D96: Shi, Brook: Include Install+Option+Variance=51
Note on N96: Shi, Brook: Include Install+Option+Variance=51

## System Usage
Site Information |  |  |  |  | System Informaiton |  |  | System Performance
Region | Site name | City | Region (in USA) | Country | Remote | BB Type [Oncology / Radiology] | S/N | Head/Neck #Patients | Chest #Patients | Abdomen/Pelvis #Patients | Extremities | Pulmo 4DCT #Patients | Site profile | Specific questions | Total #Patients
NA | The Regional Medical Center | USA |  |  |  |  | 2009-07-11
NA | Willis Knighton Health System | USA |  |  |  |  | 2009-07-10
NA | Clay County Medical Center | USA |  |  |  |  | 2009-07-09
EMEA | T.C Saglik Bakanligi Bagcilar -> Bağcılar State Hospital | Turkey |  |  |  |  | 2009-07-16
NA | St Joseph Westgate Medical Center | USA |  |  |  |  | 2009-07-07
NA | Seton Medical Center Harker He | USA |  |  |  |  | 2009-07-12
EMEA | Gesundheit Nord gGmbh | Germany |  |  |  |  | 2009-07-08
NA | METRO MAYAGUEZ INC | Puerto Rico |  |  |  |  | 2009-07-15
LA | CMD Clinica Integrales Rancagu | Chile |  |  |  |  | 2009-07-14
NA | Mount Grant General Hospital | USA |  |  |  |  | 2009-07-18
EMEA | SCP GROUPE D IMAGERIE MEDICALE | France |  |  |  |  | 2009-07-25
EMEA | S.C.S. HOSP. GRAL. DE FUERTEVE | Spain |  |  |  |  | 2009-07-26
APAC | Jinggangshan Hospital, the First Affiliated Hospital of Nanchang University, Jiangxi | China |  |  |  |  | 345002
APAC | Huaying Mountain Guangneng General Hospital | China |  |  |  |  | 345000
APAC | Dalian Branch of Shengjing Hospital of China Medical University | China |  |  |  |  | 345001
APAC | General Hospital of Jizhong Energy Jingxing Mining Group Co. LTD | China |  |  |  |  | 345003
APAC | SD Laiyang People Hospital | China |  |  |  |  | 345005
NA | Manati Medical Center | Puerto Rico |  |  |  |  | 2009-07-13
APAC | Wuhan Ren 'ai Hospital, Hubei Province | China |  |  |  |  | 345004
EMEA | HITACHI HEALTHCARE TURKEY | Turkey |  |  |  |  | 337000
APAC | JL Taonan T.C.M. Hosp. | China |  |  |  |  | 345006
NA | HOSPITAL METROPOLITANO SAN GER | Puerto Rico |  |  |  |  | 2009-07-23
APAC | HLJ Harbin Senhai Hospital | China |  |  |  |  | 345007
APAC | Henan Fengquan District Dakuai | China |  |  |  |  | 345009
APAC | MURWILLUMBAH DISTRICT HOSPITAL | Australia |  |  |  |  | 341000
EMEA | Uniwersytecki Szpital Kliniczny | PL |  |  |  |  | 341001
EMEA | Centrum Opieki Medycznej | PL |  |  |  |  | 341002
NA | HOSPITAL DR SUSONI INC | Puerto Rico |  |  |  |  | 2009-07-19
NA | McLaren Macomb | USA |  |  |  |  | 2009-07-27
NA | Leesburg Regional Medical Cent | USA |  |  |  |  | 2009-07-28
APAC | NMG Zizhiqu No.4 Hosp | China |  |  |  |  | 2009-07-29
NA | Central Vermont Medical Center | USA |  |  |  |  | 2009-07-30
APAC | Kinan Hospital | JP |  |  |  |  | 340000
APAC | SD Chest Hosp. | China |  |  |  |  | 2009-07-31
EMEA | Centro Hospitalar de Vila Nova | Portugal |  |  |  |  | 2009-08-05
APAC | Hebei Laonianbing Hospital | China |  |  |  |  | 345015
APAC | XIFENG CHINESE MEDICINE HOSPIT | China |  |  |  |  | 345014
APAC | SD Anqiu T.C.M Hosp. | China |  |  |  |  | 345017
APAC | Fujian Medical Device Co. LTD. | China |  |  |  |  | 345010
APAC | GD Dongguan Dalingshan Hosp. | China |  |  |  |  | 345013
APAC | HLJ Huachuan People Hosp. | China |  |  |  |  | 345018
NA | VA Medical Center Biloxi | USA |  |  |  |  | 2009-07-24
APAC | HEB Caofeidian Dist. Hosp. | China |  |  |  |  | 345016
APAC | Fujian Medical Device Co. LTD. | China |  |  |  |  | 345011
APAC | Fujian Medical Device Co. LTD. | China |  |  | Yes |  | 345012
APAC | Suzhou Guangci Tumor Hospital | China |  |  | Yes |  | 345019
EMEA | STEPPING HILL HOSPITAL | UK |  |  | Yes |  | 2009-07-22
EMEA | A.O. SPEDALI CIVILI DI BRESCIA | Italy |  |  |  |  | 2009-08-01
NA | Odessa Regional Medical Center | USA |  |  |  |  | 2009-08-04
LA | DIAGNOSIS Radiologia Corporation | Dominican Republic |  |  |  |  | 2009-08-18
APAC | SX Datong Yanggao People Hosp. | China |  |  |  |  | 345025
APAC | HEN Puyang People's Hosp | China |  |  |  |  | 2009-08-06
EMEA | Szpital Tymczasowy w Płocku | PL |  |  |  |  | 341007
APAC | QH Haibei Zhou No. 2 | China |  |  |  |  | 345029
EMEA | Kongresowe Szpital Tymczasowy | PL |  |  |  |  | 341008
APAC | JS Nanjing U.A. Zhongda Hosp. | China |  |  |  |  | 345021
EMEA | Avesta lasarett | Sweden |  |  |  |  | 2009-08-15
APAC | Yangfangdian hospital | China |  |  |  |  | 2009-08-10
APAC | Nikko Memorial Hospital | Japan |  |  |  |  | 341011
APAC | JS Yangzhou No.1 People Hosp. Xiqu Hosp. | China |  |  |  |  | 345022
APAC | Zhenjiang M.C.A. Hosp. | China |  |  |  |  | 345023
APAC | Saiseikai Gosu Hospital | JP |  |  |  |  | 343000
APAC | JS People Hosp. Hexi Branch | China |  |  |  |  | 345024
APAC | GX Tengxian People's Hospital | China |  |  |  |  | 2009-08-11
NA | Lackey Memorial Hospital | USA |  |  |  |  | 2009-07-21
APAC | BJ C. & W. Medicine Hosp. | China |  |  |  |  | 2009-08-07
APAC | JS Xuzhou M.C.A. No.3 Hosp. | China |  |  |  |  | 2009-08-14
APAC | Wuxi Jiangyuan Hosp. | China |  |  |  |  | 345020
APAC | NMG Alashan Youqi People Hosp | China |  |  |  |  | 345031
APAC | JL Songyuan Central Hosptial | China |  |  |  |  | 345008
NA | Unity Health Systems | United States |  |  |  |  | 2009-08-12
APAC | Guiyang Fuyou Hosp | China |  |  |  |  | 2009-08-08
APAC | SD Jinan No.4 Peo. Hosp. | China |  |  |  |  | 2009-08-13
APAC | NMG Zhengxiang Baiqi Hosp. | China |  |  |  |  | 345033
APAC | JS Nanjing M.U.A Yifu Hosp. | China |  |  |  |  | 345028
LA | HOSPITAL BARROS LUCO TRUDEAU | Chile |  |  |  |  | 341005
NA | Washington Radiology Associate/Manati Medical Center | USA |  |  |  |  | 2009-07-17
EMEA | MAZ ZARAGOZA M.A.T.E.P.S.S.11 | Spain |  |  |  |  | 2009-08-17
APAC | SC Yibin No.2 T.C.M Hosp. | China |  |  |  |  | 345035
LA | Hogar Clinica San Juan de Dios | Arequipa |  |  |  |  | 2009-08-09
APAC | NMG Xinganmeng Hos. | China |  |  |  |  | 345032
APAC | Tongling Boai Hospital Co LTD | China |  |  |  |  | 345034
APAC | JS No.2 Chinese Medicine Hosp. | China |  |  |  |  | 345027
APAC | JX Province Children Hosp. | China |  |  |  |  | 2009-08-16
[8 empty] | Total Patient Scan - Oncology: | [6 empty] | 0
[8 empty] | Total Patient Scan - Radiology: | [6 empty] | 0
Note on N2: Shi, Brook: * Emergency department (inc. abdomen, head) * Cardio-vascular imaging (‘head&neck CTA’, ‘Runoff’) * Cardiac site (‘step & shoot’, ‘retrospective’) * Interventional radiology (‘CCT’, ‘Biopsy’) * Radiation oncology (‘4DCT’ ‘Pulmo gated’) * Diagnostic oncology (‘liver’, ‘lung cancer’) * High throughput – more than 10 pts/hour, more than 80 pts/12 hours * Geography distribution
Note on O2: Shi, Brook: * ‘Brain routine’ , ‘abdomen routine’ * (‘head&neck CTA’, ‘Runoff’) * (‘step & shoot’, ‘retrospective’) * Utilization of IMR, Spin-Spin * Improvement from previous version: throughput or perceived recon speed/delays

## Certified Apps Specialists (hidden)
Course Code | First Name | Last Name | Country | Date Passed
AW8078 IQon Train the Trainer Exam Oct 2015 | FRANCA | BROUWER, DE | NLD | 2016-01-19
AW8078 IQon Train the Trainer Exam Oct 2015 | EMILY | BURKE | UNITED STATES | 2015-10-22
AW8078 IQon Train the Trainer Exam Oct 2015 | CATHERINE | CAMPBELL | CAN | 2015-10-26
AW8078 IQon Train the Trainer Exam Oct 2015 | DYAN | CURRY | USA | 2015-10-22
AW8078 IQon Train the Trainer Exam Oct 2015 | STEPHEN | DIDONATO | USA | 2016-03-03
AW8078 IQon Train the Trainer Exam Oct 2015 | ANTHONY | DISALVO | USA | 2015-10-22
AW8078 IQon Train the Trainer Exam Oct 2015 | yoo | dongeui |  | 2015-10-22
AW8078 IQon Train the Trainer Exam Oct 2015 | TAISUKE | FUJIOKA | JPN | 2016-01-14
AW8078 IQon Train the Trainer Exam Oct 2015 | PATRICK | GEORGE | BEL | 2015-10-22
AW8078 IQon Train the Trainer Exam Oct 2015 | DIANE | GRESSLEY | USA | 2015-10-26
AW8078 IQon Train the Trainer Exam Oct 2015 | WILLIAM | HALL | USA | 2015-10-23
AW8078 IQon Train the Trainer Exam Oct 2015 | MICHELE | HIRCHAK | USA | 2016-03-03
AW8078 IQon Train the Trainer Exam Oct 2015 | MARY JAYNE | HOGAN | USA | 2016-03-03
AW8078 IQon Train the Trainer Exam Oct 2015 | Allison | Lecky |  | 2015-10-22
AW8078 IQon Train the Trainer Exam Oct 2015 | SHARON | MOEN | USA | 2016-03-03
AW8078 IQon Train the Trainer Exam Oct 2015 | RON | O'CONNOR | USA | 2016-01-14
AW8078 IQon Train the Trainer Exam Oct 2015 | GESA | RADTKE | CHE | 2016-01-14
AW8078 IQon Train the Trainer Exam Oct 2015 | HANNE | SCHOU | DNK | 2015-10-22
AW8078 IQon Train the Trainer Exam Oct 2015 | ELIZABETH | SHERRARD | USA | 2016-03-03
AW8078 IQon Train the Trainer Exam Oct 2015 | TRACI | SPENCER-MEDLOCK | USA | 2015-11-23
AW8078 IQon Train the Trainer Exam Oct 2015 | SOEREN | THOMSEN | DNK | 2015-10-22
AW8078 IQon Train the Trainer Exam Oct 2015 | CHRISTOPHER | TRITLE | USA | 2015-10-22
AW8078 IQon Train the Trainer Exam Oct 2015 | cathy | wei |  | 2016-03-03
AW8078 IQon Train the Trainer Exam Oct 2015 | JULIA | WELLONS | USA | 2015-10-22
Total Passed: 24

## Service Training-SpDS (hidden)
Service Training - SpDS | Workstation (IX)
Server – Tower
Server – Rack
Details for training-path and course description can be found at InCenter:
http://incenter.medical.philips.com/Default.aspx?tabid=4828

## SpDS PO & SO (hidden)
Customer | Creation date | ISP SpDS Sales Order | Customer PO Number | HW Configuration (Server/WS) | RDD date | CDD date | FSD date | Final user country | Final user name | Finale user city | FSE the Install the ISP SpDS
2365-12-01 | 4219300% | 512830 | 4508879941 | Server - Rack | 4229700% | 4236900% | 4236600% | BE | CLINIQUES UNIVERSITAIRES SAINT | Bruxelles
980230 | 4219800% | 515281 | 4508925637 | Server - Tower | 4233100% | 4233100% | 4232500% | CH | Clinique de Genolier | Genolier VD
455053 | 4219900% | 515575 | 4508930316 | Server - Tower | 4227900% | 4227900% | 4226900% | KR | Seoul National University Hosp | Seoul
455053 | 4220000% | 515851 | 4508934709 | Empty – need to be corrected in the order | 4231300% | 4231300% | 4230300% | KR | KOREA UNIV.HOSP.ANSAN | Ansan
980230 | 4222200% | 519398 | 4509006158 | Empty – need to be corrected in the order | 4228300% | 4228300% | 4227700% | CH | Clinique de Genolier | Genolier
570285 | 4223500% | 521470 | 4509044166 | Server - Tower | 4230500% | 4230500% | 4229300% | KR | Kyungsang National Univ. Hosp. | Jinju
300278 | 4224400% | 522995 | 4509076792 | Server - Rack | 4234500% | 4234500% | 4234000% | DK | Region Midtjylland | Aarhus N

## Install Audits (hidden)
Region | System type: Ingenuity | Site | Rigging to the room | Template | Equipment positioning | Unpacking system components | Level Gantry | Gantry and Table cabling | CRC internal cabling | System cable interconect | System power up | Align & level couch | Verify couch, tilt and laser cals | Install license key | Set up network | Peform air calibrations | Install covers | Acceptance test | Cleaning | Physics checks | Regional variance 20% | Corrective maintenance | Total | Total plus CM
|  | Average Install | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!