## refresh2
itd | Head_line | CEshi1
N/A | ceshi1as | ceshi1df () *
# | ceshi1zx | ceshi1vf
3 | ceshi1 as | Pending during There is a problem with my current Python script

## refresh_Field_IssueSE
id | FOK_Action_Item_ID | Status_1 | Country | Site_Name | S_N | Date_Reported | wk | Field_Issues_Heading | Field_Issue_Description__Field_reported | Customer_Impact_Effect | Bugreport | CA | Target_to_Close__in_5_Workday | PA | Target_to_Define_action_plan__in_10_Workday | Picture | During__After_Installation | Lebel2 | Level3_ | Root_Cause | Countermeasures | Action__PA | Owner | Timing | Viper__One_EMS_No | Case_Status | Complaint__PR_No | tru | status_2 | PR_Status | Complaint_Status | Defect__CTUDT | Application_Specialist__Clinical_Science_Comments | Customer_Feedback_ | Highlight | Assign_Owner | Customer_Facing_Issues | Funct_Owner_1 | Investigation_Results | Correction_Plan | Action | Status | Failure_mode | Categoly | SSS | Status_comment | Start_Clinical_Date | Between_failure | Fix_Priority_ | Mitigation_Workaround | IB_Exist__New_Found | Related_FRU_P_N | FRU_Description | Details__Comments | Level1 | Level2 | Level3 | Level2__3 | System_Type_0 | silva | Check_Level3_Type | Type1_level2 | type2_silva | MonthRolling | level1_check | level2_check | level3_check | level2_3_check | System_Type_0_check | Month_Rollingcheck | ifequal | silva_check__compared_with_BI | Check_Level3_Type__BJ | Type1_level2__BK | type2_silva__BL | Funct_Owner | PR_status_Check | IB_Existing_Check | CTUDTcheck | SSS_Check | Check_ | End_
# | FOK Action Item ID # | Status | Country | Site Name | S/N | Date Reported | wk | Field Issues Heading | Field Issue Description (Field reported) | Customer Impact/Effect | Bugreport | CA | Target to Close (in 5 Workday) | PA | Target to Define action plan (in 10 Workday) | Picture | During / After Installation | Lebel2 | Level3 | Root Cause | Countermeasures | Action (PA) | Owner | Timing | Viper/One EMS No. | Case Status | Complaint (PR#) No. | tru | status | PR Status | Complaint Status | Defect (CTUDT#) | Application Specialist / Clinical Science Comments | Customer Feedback | Highlight | Assign Owner | Customer_Facing_Issues | Funct. Owner | Investigation Results | Correction Plan | Action | Status | Failure mode | Categoly | SSS | Status comment | Start Clinical Date | Between failure | Fix Priority | Mitigation/ Workaround | IB Exist/New Found | Related FRU P/N | FRU Description | Details / Comments | Level1 | Level2 | Level3 | Level2+3 | System-Type-0 | silva | Check-Level3-Type | Type1-level2 | type2-silva | 3MonthRolling | level1-check | level2-check | level3-check | level2+3-check | System-Type-0-check | 3MonthRollingcheck | ifequal | silva-check (compared with BI) | Check-Level3-Type (BJ) | Type1-level2 (BK) | type2-silva (BL) | Funct. Owner | PR status Check | IB Existing Check | CTUDTcheck | SSS Check | Check | End
1 | 43993 | TRUE | US | The Regional Medical Center | 40005 | 43993 | 24 | Cable Harness been squeezed | Cable Harness Rear Cover Close Up been pressed by screw hole, see picture | No |  | Adjust and fixing cable | HFA investigation | Drawing update and supplier quality control ECR-075587 submitted to fix drawing update | ECR-075587 | - | During |  |  |  | Short term | Long term |  |  | CFM |  | 152152 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | Valid | CTUDT00630134 |  |  | N | Wu ZhiMing | Installation Covers | Engineering | Cable Harness Rear Cover Close Up been pressed by screw hole The cable harness was assemblied in supplier side | Supplier Quality Issue | need PE/SQE to check and adjust the cable tie position away from the screw hole | Closed | Manufacturing Issue | System Hardware | Supplier Quality | Ongoing, ZhiMing is aligning with supplier to add additional cable clip for above the screw hole 6/16 update: drawing will update to add more cable clip |  |  | Low | N/A, Cable work with function, continue monitoring |  | N/A | N/A | System issue | system | Gantry-HW | gantry cover | Gantry-HW-gantry cover | Gantry-HW-gantry cover | CTUDT00630134 | gantry cover |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
2 | 43993 | TRUE | US | The Regional Medical Center | 40005 | 43993 | 24 | E-Stop can not close | During initial power on and estop closure, no EStop Prompt appeared. System configuration check found a deviation with Couch software. Screen capture acquired | No |  | Re-programing couch software | Done | Integrate couch during MFG testing Workshop was hold between OPS and service to define action Timeline: 23-Jul update the status (David) | 3-Aug-2020: David and Tomer |  | During | Console-SW | Console-SW-Workflow | UI was cut in from BB on Rail | local team to use IVC control or CT BOX control. | CTUDT00630141 in SP1 scope |  | 31-Jan-21 | CFM |  | 152116 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | Valid | ECR-076249 |  |  | N | David | E-Stop | Operation | During initial power on and estop closure, no EStop Prompt appeared. System configuration check found a deviation with Couch software. | Decoupling or Coupling | Need check with David and make sure the couch has been tested in factory with correct sofftware version, de-coupling?? o OWNER: Elena ACTION: confirm what version of software was used to test this Couch. Provide any evidence/details you can for this item such as the test DHR record, any screenshots have for this test process, etc. – Couch SN is 905005, was tested together with system 40005 with SW version 4.1.10.25065 | Closed | Manufacturing Issue | System Hardware | De-coupling | Fixed by FSE |  |  | Low | E-Stop closed after Re-programing couch, system works |  | N/A | N/A | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | 0 | No | y | y | y | y | y | y | y | y | y | y | #REF!
3 | 43993 | TRUE | US | The Regional Medical Center | 40005 | 43993 | 24 | DVI Adaptor is missing | 2 Display ports to DVI adaptors are needed to attach both monitors, however only 1 was shipped with system. Missing 2nd Adaptor. We borrowed a 2nd one to attach and utilize dual displays. DVI - Display port adaptor | No |  | Obtained Locally | Done | Add check if adaptor exist in package before ship out Timeline: SZ: Already include the check piont for the adapter before teal down the system HF: Tomer, 23-Jul update the status 0728 update ECO in process to add to outgoing inspection; refer to ECO-128239 | 3-Aug-2020: David and Tomer |  | During | [7 empty] | CFM |  | 152115 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | Valid | E-128239 |  |  | N | David | Missing Item | Operation | Two adapters in console sublevel, it should be used for Mfg testing. The adaptor is necessary to connect the DVI cable from the console harness to the Display Port on the back of CRC Rack. |  | Shelvin offline confirm with David for WI and detail consider to add check to make sure adapter in packing box 6/12 suggest to add a note/remark in WI not to remove the adapator 6/17 update align with HF to check | Closed | Missing on Arrival | Missing Item | DVI Adaptor | Fixed by FRU |  |  | Middle | obtain a 2nd one locally to attach and utilize dual displays |  | N/A | N/A | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
4 | 43993 | TRUE | US | The Regional Medical Center | 40005 | 43993 | 24 | Power brush with pressure | Concerned about the adjustment of one of the power brush blocks. I observed the trailing tips of one of the brushes in the 120v Tracks as being higher than the isolation dividers. This indicates too much pressure on brushes which can be a result of improper shimming. Photo Acquired. | No |  | Redo brush alignment | Done | Add check for brush trailing tips in right postion in MFG Alignment spec and steps in current WI Need engineering input/spec how to measure the gap and define the position, any CTQ related the brush alignment. Timeline: David, Brook 23-Jul update the status | 11-Aug-2020: David and Tomer 18-Aug-2020: David |  | During | [7 empty] | CFM |  | 152119 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | Valid | Not A Defect |  |  | N | David | Hardware Issue | Operation | From the picture showed, the power bruch was too close to slip ring, cause the The wire warped |  | Partial Gantry from Suzhou, David and Elena confrim the SN and shiming for this system,check if correct and avoid issue happen Monitoring next site instllation. Take photos for new brush 6/18 update plan was for them to address on scheduled PM 6/29 update Pending on new Power brush arrive and replace | Closed | Manufacturing Issue | System Hardware | Power Brush_DOA | Fixed by FSE (re-adjustment) |  |  | High | Alignment is correct, reminded him to order another brush block and to monitor the existing one. suggest replacing it if the brush remains raised |  | N/A | N/A | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
5 | 43993 | TRUE | US | The Regional Medical Center | 40005 | 43993 | 24 | Console table Color mismatch with CRC Cabinet | Color mismatch between the CRC Cabinet and Console table is very noticeable and not very attractive. Photo Captured but doesn’t demonstrate the contrast as significant as actually seen in the room. | Aesthetic |  | No plan to replace due customer use their own table | Done | Cut in new Console table 459801054601, CBOM will update to fix the issue Timeline: Tomer, check if check point add or supplier qulity check for the console table ensure the table is white, 23-Jul update the status  Tomer input: HFA has two suppliers for the table with two different 12NC’s and varying levels of whiteness. One version from supplier Hontech is within the whiteness spec required. For BB and ING shipments from HFA, Logistics will only ship the table from supplier Hontech. Roni working out details with Logistics team; for shipping only from Hontech, and how to deal with the inventory. | 11-Aug-2020: David and Tomer Tomer align with Logistic team to check stocks |  | During | [7 empty] | CFM |  | 152120 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | Valid | Not A Defect |  |  | Y | David | User Experience | Operation | Computer table is mushroom color, CRC Cabinet is whiter color. Need check C-BOM/PCM/Supply chain to cut in whiter table for Refresh version, consider to ship a new table for this customer | 1. Correct C-BOM to cut in new white table for forward production 2. consider to ship a new table for this customer | How many old table stock in Haifa/Suzhou factory? Haifa 12NC for old table 459801054601 (new color, 57 stock new color in Suzhou) or 455012005721 (old color, 50 open PO for old color table. Haifa 17, Suzhou 0) Xie Hong inform Haifa stop shiping old color for Refresh orders (even for 4.1.7 with G5.1), raise ECR to trigger C-BOM/PCM update 6/11 Update Mor.Igor is confirming with Sammia supplier for the old table. Jiaan, Xunliang, is confirming Hontec for new table color. SQE（XIa Bin） in Suzhou/Eleva is HFA?: if need add spot check during IQC? 6/29 update L3 asked FSE raise DEFOA to highlight this issue, waiting feedback | Closed | Supplier Quality issues | Option Hardware | Color mismatch | Customer cancel |  |  | High | No workround, consider to ship a new table for this customer if color mismatch | Refresh | N/A | N/A | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
6 | 43993 | TRUE | US | The Regional Medical Center | 40005 | 43993 | 24 | Phantom knob bent | Adjusting knob on phantom holder is bent. Need to find out if we can order only this adjusting knob or must order entire phantom holder assy. A follow up to the original claim of bent phantom adjusting knob. We found further damage to the phantom. The clear plastic around the perimeter of the phantom has a break. So, it appears that the phantom suffered damage due to unusual force. There are photos including a GIF demonstrating the damage in the Day 5 folder. | No |  | Raise DEFOA for the damaged part | Done | Adjust knob to the end position to aviod damage during transpotation. Change the knob from one to two for spare use Change the knob to FRU ECR submitted | ECR-074469 |  | During | [7 empty] | CFM |  | 152140 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | Valid | CTUDT00630136 |  |  | N | Ada | Installation Other | Engineering | Knob is not FRU part for field replacement. phantom suffered damage due to unusual force. Additional issue: handle of phantom cracked |  | Set the adjusting knob as FRU. Change the screw qty from 1 to 2 in Body phantom BOM 6/11 update keep monitor next installation 6/29 update asked FSE raise DEFOA to highlight this issue, waiting feedback | Closed | Defects on Arrival | System Hardware | Supplier Quality | Conitnue Monitoring |  |  | Middle | New Phantom orderred from SPS and arrived on site |  | 459801415161 | HARMONIZED PHANTOM HEAD SECTION | System issue | system | Couch-HW | Phantom | Couch-HW-Phantom | Phantom | CTUDT00630136 | Phantom |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
7 | 43993 | TRUE | US | The Regional Medical Center | 40005 | 43993 | 24 | HCOR calibration failed | When performing HCOR calibrations we were unsuccessful in aligning the phantom via "Auto Centering". We reverted to manual alignment to complete the HCOR calibrations. | Extra time for calibration |  | Manual alignment to complete the HCOR calibrations. | Done | No action planned, keep monitoring Long term: Innovation idea on Auto Centering function | N/A |  | During | DMS_SW | [6 empty] | CFM |  | 152141 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | Valid | Not A Defect |  |  | N | Silva | Calibration/Performance | Engineering | When performing HCOR calibrations we were unsuccessful in aligning the phantom via "Auto Centering" angel limitation |  | Investigate on system and update the result, Pending for Field feedback with detail information 6/11 update keep monitor next installation | Closed | System Performance Issue | Calibration/Image Quality | Auto Centering | Fixed by FSE (re-adjustment) |  |  | Middle | Manual alignment to complete the HCOR calibrations. Auto centering successfully once 3 phantom attached |  |  |  | System issue | system | DMS-SW | Phantom Detect | DMS-SW-Phantom Detect | DMS-SW-Phantom Detect | Not A Defect | Phantom Detect |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
8 | 43999 | TRUE | US | The Regional Medical Center | 40005 | 43999 | 25 | No infant phantom and auto centering failed | We are missing the Infant phantom and unable to perform cals related to it. Noticed that Infant cals were not a separate item in the Calibration menu and also noticed the Help instructions available during HCOR calibrations show illustrations of the Harmonized phantom to include the head, body AND infant phantom. So, this raises the question: IS this why auto centering failed? Does the infant phantom need to be installed in order for auto centering to succeed? (Update: Spoke with Brook and he checked the PCM to find that the infant phantom is an option. So, we must investigate the auto centering issue. Maybe we need to more precisely set up the phantom holder.) | Extra time for calibration |  | Uncheck the Infant phantom in preference setting | Done | Infant phantom always been used and tested in OPS, service calibration manual inlucde step to the uncheck step according to customer order. | 459801587311 Rev d |  | During | [7 empty] | CFM |  | 170518 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | Valid | CTUDT00629920 |  |  | N | Silva | Calibration/Performance | Engineering | 1. H-phantom Infant phantom now only used for infant performance test to confirm system IQ performance for pediatric patient. All the scanner’s calibration can be covered with current H-phantom head and body. No need infant phantom for calibration. 2. Any official release document relate to this change: a. DHF337053 is review TRR for removing infant phantom from HCor calibration marked several test covered information. b. DHF179202 physics spec and H-phantom install picture (some errors are included in this document, David is drafting the update version. CTUDT00629755 and CTUDT00629920 are relate to this issue) |  | 6/11 update Proposal: Need update service manual or ops wi to add the uncheck step according to customer order. From RD perspective, customer no need prepare infant phantom for scanner usage. Silva: confirm with Amitai for investigation result and defect handling Brook: check with manual will be impacted. 6/12 update Siva: drive related defect handling and input for service manual or ops wi update Keep monitor next installation | Closed | System Performance Issue | Calibration/Image Quality | Infant Phantom | Conitnue Monitoring |  |  | High | Uncheck the Infant option in Preference setting need more picture and image of testing, grid chart/DICOM picture, as well as message popup from system Step1, select the calibration from Patient Directory, Right click and select View DICOM Info Step2. check the path at the bottom of DICOM Info Screen. Step3, Open above path and copy all files. Question: Whether system can not perform infant scan without Infant phantom? no impact on infant scan |  |  |  | System issue | system | DMS-SW | Help Page | DMS-SW-Help Page | DMS-SW-Help Page | CTUDT00629920 | Help Page |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
9 | 43993 | TRUE | US | The Regional Medical Center | 40005 | 43993 | 24 | All phantom labels are saying System Phantom | All phantom labels are saying System Phantom | Extra time for calibration |  | Add label for phantom detail (ship to Chuck 7/7) | Done | Design improvement to add detail section of phantom on the label ECR submitted | ECR-074469 |  | During | [7 empty] | CFM |  | 152148 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | Valid | CTUDT00630136 |  |  | N | Ada | Applications/Usability | Engineering | Phantom label without detail infromation e.g. infor from agile 459800443051 Philips Materials • PHANTOM ASSEMBLY-HEAD |  | Add phatom detail after P/N _HEAD SECTION for H-Phantom section Conbine action with item 7 in ECR scope 6/11 update ECR under team reivew to add detail section of phantom in the label 6/29 update, will send the new lables to Chuck/Carlos for the site correction | Closed | Engineering Issue | System Hardware | Phantom Label | Conitnue Monitoring |  |  | Low | Training to customer |  |  |  | System issue | system | Couch-HW | Phantom | Couch-HW-Phantom | Phantom | CTUDT00630136 | Phantom |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
10 | 43993 | TRUE | US | The Regional Medical Center | 40005 | 43993 | 24 | Gantry panel LED (tilt) mismatch | Found at end of day: We noticed an instance of the rear gantry panels not reporting identical status. A photo was provided. Notice that Left rear panel tilt LEDs are BOTH lit. But right rear panel only has 1 tilt enable LED lit. This requires further investigation. | can not show real tilt status in rear left control panel. |  | Order part from SPS 459801485021 on shelf, ready to order | Done | MFG test check for LED/out going check status for LED from supplier checkpoint available in process; confirming with the Test team the actual process and DMR #. Will provide once available; target 28-JUL once in DMR367625 step. 420-430, and another time in DMR367526 step 110, no led functionality, only controls | 11-Aug-2020: David and Tomer ECO ongoing. Target close 17-Aug |  | During | [7 empty] | CFM |  | 152146 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | Valid | CTUDT00630866 |  |  | N | David | Hardware Issue | Operation | Awaiting FSE confirm and feedback to do investigation, if still not work, consider to order new panel 459801485021 |  | William/David/Elena, Check if we have control panel LED Status check in factory? @Chuck, Please check what is the status of system. Always not able to lit up? Cabling check 6/11 update After replace consider to return to BIU for investigation 6/29 update ensure FSE follow DEFOA to get correct part 459801485021 to replace | Closed | Defects on Arrival | System Hardware | GCP_DOA | under field confirmation |  |  | Middle | Chuck ask FSE Thursday to verify it... I am not certain that the fill range of tilt was tested in this case |  | 459801485021 | CONTROL PANEL ASSEMBLY-REAR RIGHT | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
11 | 43993 | TRUE | US | The Regional Medical Center | 40005 | 43993 | 24 | Cosmetic plugs not fit | The cosmetic plugs that cover the holes on the tilting portion of the gantry do not fit. See the photos in Day 5 folder. Other plugs located on inside of Left & Right gantry leg fit properly. Also the couch & Footswitch plugs fit properly. | Aesthetic |  | Find other plugs in installation kit can fit | Done | No action planned, keep monitoring |  |  | During | [7 empty] | CFM |  | 152142 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | Valid | Not A Defect |  |  | N | William | Installation Other | Service | The screw hole in the stator box with thread and may not fit with 0.75 plug. Suggest to use 0.625 plug for this position |  | William, what is the intend position for the plugs? This position is not able to insert the plugs. @Chuck, Please share an overview picture of the cosmetic plugs. According which requirement to insert the plugs? Investigate on test bay to defince plugs 6/11 update William: check spec of the shipping bold and see which plug is fit For Field: try to insert the 0.625 plugs on this position 6/29 update check with FSE to see if we have other plugs in installation kit can fit | Closed | Installation Process Issue | Others | Misunderstand | Fixed by FSE (re-install) |  |  | Low | The screw hole in the stator box with thread and may not fit with 0.75 plug. Suggest to use 0.625 plug for this position |  |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
12 | 43999 | TRUE | US | The Regional Medical Center | 40005 | 43999 | 25 | couch not moving | Today applications was training at the site and they had intermittently during the day several problem with the couch not moving when pressing the buttons. I attach a video and and the bugreport link below They are trying now removing the rear footswitch , may be a source for the problem | P2 |  | Use normal couch motion button but not unload button | Done | Design change to fix issue FCO72800735 | FCO72800735 |  | After | [7 empty] | 115619737 |  | 159883 | TRUE | Closed - Done | Closed - Done | Valid | CTUDT00630138/630160 |  |  | N | Robin | Lockup or Crash | Engineering | First error , 2020-06-16 14:35 :08 the horizontal position is out of Limit After the first error : the table always can’t reach the target position One time the AMC red LED shows incorrect status after couch not moving |  | Solution: It’s best to recalibration the horizontal position. 6/18 update Order relay and AMC for this site to debug and fix issue 6/29 update BIU working on solution | Closed | Engineering Issue | System Hardware | FCO72800735 | FSE did the horizontal position recalibration, but not work, system was down, National service support suggest to order |  |  | High | suggest “manual” patient unload until issue can be resolved |  |  |  | System issue | System | Couch-SW | Horizontal Move Block | Couch-SW-Horizontal Move Block | Couch-SW-Horizontal Move Block | CTUDT00630138/630160 | Horizontal Move Block | Couch-SW | Couch-SW-Horizontal Move Block | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
13 | #N/A | #N/A | US | Clay County Medical Center | 40003 | 44000 | 25 | Front cover bulge | The right side of the cover has a bulge in it about 12 inches up from the bottom. It is worth noting that the bulge starts right after a supporting L bracket on the inside stops. The bulge is pretty large. We tried lowering the right side of the cover to shift the cover down and to the right which fixed the bulge rubbing, but then the collimator rubbed on the inside. We tried adjusting the left side up but it was already at It is maximum. We then tried shifting the cover left and right. We eventually got the collimator to not rub and the bulge to clear but a few hours later the bulge got caught while tilting, and luckily we stopped it before it broke the cover. We tried sanding it down but if we sand it down to where it will clear safely the material will be too thin. So we ordered a new cover. We are trying to bend in the bulge but since this is fiberglass we don_t think It is going to work. | Tilt not work |  | Raise DEFOA to replace part, part sent 10-Jul 24-July arrive on hospital | Done | Supplier Quality check (10) added for cover width and flatness before ship to factory Engineering: ECR-075587 Tomer input: confirming with the Test team what the actual process is in testing for tilt and verifying gaps to see if this item is covered in normal process. Need to check if also have any outgoing inspection process to cover this gap. Will provide specific DMR if available; target 28-JUL | 3-Aug-2020: David and Tomer DMR367625 step. 320 |  | During | [7 empty] | CD-29578 |  | DEFOA CD-29578 | #N/A | #N/A | DEFOA | Valid | ECR-075587 |  |  | Y | Andy/William | Installation Covers | Operation | Cover flatness too big and not fit to gap between stator box and gantry. | Haifa and Suzhou OPS is aligning with supplier how to ship the Good part to for this DEFOA | update the spec to control the flatness of Front and Rear cover DEFOA submit #. CD-29554 awaiting for shipment | Closed | Supplier Quality issues | System Hardware | DEFOA | waiting shipment |  |  | High |  |  |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
14 | 43999 | TRUE | US | The Regional Medical Center | 40005 | 43999 | 25 | problem with not being able to advance the couch into the gantry with the load/unload foot pedals or the gantry control panel | 1 • After returning from lunch, we attempted to perform a mock scan since we did not think we could scan patients because the PACS transfer was not working. Upon setting up the scan, we found that we could not move the table in or out of the gantry using the foot pedals, gantry controls or the controls on the scan control box. When we pressed the load/unload foot pedals, the table would only go up or down, but it would not advance into the gantry. Pressing the grey or white gantry in/out controls also would not advance the couch. I had them turn off the gantry key switch. We restarted the gantry with the key switch and this resolved the problem. Approx. 1:05pm 2 • The customer attempted to get a patient for a PE scan on the table. We again, had the problem with not being able to advance the couch into the gantry with the load/unload foot pedals or the gantry control panel. This time turning the gantry on & off with the key did not resolve the issue. We shut the system all the way down to the wall (both computers and Teal). It still did not resolve the issue. The patient was taken off the table and moved to their other scanner. I again shut the entire system down to the wall, and after the 2nd time shutting it down, the in and out couch controls started working. This was around 3pm timeframe. • A bugreport was made and I contacted 3rd level support. They were having difficulties logging in or had a slow connection. I downloaded the bug report to send but with the slow hospital wi-fi connection at this facility, it took over 1 hour to send the file. The file was sent to Carols Molinari and Sam Pierce. 3 • They attempted to get another PE scan. We again had the gantry issue with driving the table in/out of the gantry. It would not advance with the load/unload foot pedals, the gantry in/out controls or the scan control box. After multiple attempts at shutting the system completely down, we still could not get it to work again. The patient had to be taken off the table and was moved to their other scanner again. I shut the system all the way to the wall 3 different times, and still could not get it to work. A service call was placed, Case #0115619737 • One of the technologists suggested I disconnect the round foot switch we had installed earlier in the day to see if that seemed to be causing the problem. We did not seem to have any issues prior to having the switch plugged in, but we also were not doing patient scans. After disconnecting the round foot switch, the in/out controls started working again. I do not know if this is coincidence or if this has resolved the issue, only time will tell. | Couch Fast move/load/unload function doesnot work properly |  | Use normal couch motion button but not unload button | Done | Duplicate with #12 | FCO72800735 |  | After | [7 empty] | CFM |  | 175760 | TRUE | Closed - Done | Closed - Done | Duplicate (12) | CTUDT00630138/630160 | APP Training |  | N | Robin | Lockup or Crash | Engineering | 1st happened: 1:05PM, restarted the gantry with the key switch and this resolved the problem 2nd happened: 3PM restart system 3rd happened: After multiple attempts at shutting the system completely down, we still could not get it to work again. This did not change after AMC replacement. So far we replaced encoders, AMC, CCB. Horizontal calibration. Limit switches checking. Main thing still happening couch slams into the rear end stop while driving at high speed with UNLOAD or FAST OUT It is software defect in the Ingenuity 5000 |  | Software fix, detail see defect | Closed | Engineering Issue | System Hardware | FCO72800735 | bugreport for latest Orangeburg site https://www.clouddrive.philips.com.cn/file/hvbre3d0gk8c0vijfoo5iaimj3igdmyj# 密码:c02lz0k9 |  |  | High | suggest “manual” patient unload until issue can be resolved |  |  |  | System issue | System | Couch-SW | Horizontal Move Block | Couch-SW-Horizontal Move Block | Couch-SW-Horizontal Move Block | CTUDT00630138/630160 | Horizontal Move Block | Couch-SW | Couch-SW-Horizontal Move Block | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
15 | 44014 | TRUE | US | Clay County Medical Center | 40003 | 44014 | 27 | Couch locked | 6. Bariatric table at site and system locked up on large patient. System needed to be powered off to wall to recover. (Bug report pending) Couch will not move horizontally. | P4 |  | Use normal couch motion button but not unload button, FCO will be deployed to fix site issue |  | Duplicate with #14 | FCO72800735 |  | After | [7 empty] | 115700330 |  | 177781 | TRUE | Closed - Done | Closed - Done | Valid | CTUDT00630138/630160 |  |  | N | Robin | Lockup or Crash | Engineering | not quite understand. waiting more detail of this issue from application. As well as Bugrep |  |  | Closed | Engineering Issue | System Hardware | FCO72800735 |  |  |  | High |  |  |  |  | System issue | System | Couch-SW | Horizontal Move Block | Couch-SW-Horizontal Move Block | Couch-SW-Horizontal Move Block | CTUDT00630138/630160 | Horizontal Move Block | Couch-SW | Couch-SW-Horizontal Move Block | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
16 | 44022 | TRUE | US | The Regional Medical Center | 40005 | 44022 | 28 | Work list populating issue | Q7: Service engineers description of problem worklist is not populating Q8: What troubleshooting was performed? What were the results? checked the worklist and found that it was working properly at this time Q9: What actions were taken to repair the equipment? checked the IT parameters and monitored the system operation | P3 |  | Retry Bugreport received |  | No action plan, will keep observation this case and check the connections. | N/A |  | After | [7 empty] | 115744348 | Closed | 189748 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | Valid | CTUDT00631283 |  |  | N | Silva | PACS | Engineering |  |  |  | Closed | Site Readiness Issues | System/LAN/DICOM Config | PACS | PACS | [7 empty] | System issue | system | Console-SW | Remote Connection-Worklist | Console-SW-Remote Connection-Worklist | Remote Connection | CTUDT00631283 | Remote Connection |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
17 | 44015 | TRUE | US | Clay County Medical Center | 40003 | 44015 | 27 | couch error. will not allow to reset | Buttons work to go up and down, but not in and out. Patient moved to the Brilliance 16 scanner to be completed. | P2 |  | Cool down motor 30mins |  | FCO72800735 | FCO72800735 |  | After | [7 empty] | 115711918 | New | 180048 | TRUE | Closed - Done | Closed - Done | Duplicate (15) | CTUDT00630138/630160 |  | The customer temperature is fine. I explained that it was my fault for driving the table all the way out with the grey controls and emphasized the importance of not using the grey switch | Y | Robin | Lockup or Crash | Engineering |  |  |  | Closed | Engineering Issue | System Hardware | FCO72800735 | [8 empty] | System issue | System | Couch-SW | Horizontal Move Block | Couch-SW-Horizontal Move Block | Couch-SW-Horizontal Move Block | CTUDT00630138/630160 | Horizontal Move Block | Couch-SW | Couch-SW-Horizontal Move Block | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
18 | #N/A | #N/A | US | METRO MAYAGUEZ INC | 40009 | 44056 | 33 | FLT1 ( RED LED ) on the R2D board | • “When the gantry (CT Ingenuity 5000) was power up for the first time, I notice the FLT1 ( RED LED ) on the R2D board On. This light is on when there is something wrong with the R2D board”. • “Although the gantry did initialized and rotate fine, the FLT Red LED did not switch off”. • “When the new R2D was installed the FLT Red LED was OFF and the rest of the Green LED were counting in binary as the gantry rotate”. | No impact |  | gantry did initialized and rotate fine, no impact on system function |  | improvement change to create new 12NC 459800467343 to distinguish from the 459800467342. ECR-075566 | @Silva and team, please share if you have any comments on the RED LED summary and field recommendation, if no issue this will deliver to field this Thursday |  | During | [7 empty] | 115760046 |  | DEFOA CD-30329 | #N/A | #N/A | DEFOA | Valid | CTUDT00630796 ECR-075566 |  |  |  | Silva/Ya Ping | Hardware Issue | Operation |  |  |  | Closed | Defects on Arrival | HW | DEFOA | [5 empty] | IB Exist |  |  | System issue |  |  |  |  | Gantry-HW-R2D |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | y
19 | 44034 | TRUE | Germany | Gesundheit Nord gGmbh | 40002 | 44034 | 30 | Extra material is required for service | We need the following material for the service on the system. Please open extra WO since the material for service is needed, not for the installation. Kit, EXP / BRILLIANCE FUSE KIT 453567045792 1x Kit, Lubrication Kit, Patient Support 453567023141 1x Kit, SGL Gantry Jumper By-Pass Service Kit 453567001111 1x | P5 |  | Not a defect, FSE order some spare part for this site in advance |  | No acion needed |  |  | During | [7 empty] | 115800635 | Closed | 201842 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | Valid | Not A Defect |  |  | N | Brook | Works as Designed | Service |  |  |  | Closed | Installation Process Issue |  | Service Parts Order | N/A (Not a Defect) | [7 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
20 | 44029 | TRUE | US | Willis Knighton Health System | 40004 | 44029 | 29 | IV pole mount | IV pole mount "Q7: Service engineers description of problem" : "Confirmed, intermittent", "Q8: What troubleshooting was performed? What were the results? " : "HCDiag performansce", "Q9: What actions were taken to repair the equipment?" : "Installed manual IV handle mount", "Q10: Attach or document any test and inspection data used to perform the repair to specified requirements." : "na", "Q11: List all calibrated Tools and Test equipment (if applicable) used during servicing. Include make, model, and serial numbers" : "Fluke 287 SN33510063 Calibrated 10/10/2018 Due 11/20/2020", "internal A: Did you speak with the RSE/NSS about Diagnostic Information?" : "Yes", "internal B: Did the RSE/NSS and FSE agree to a plan of action?" : "Yes" | P5 |  | need confirm if FSE follow service manual to install the IV Pole (Peter?) |  | 459800074701 Rev F ING Patient Support Panning Handle Joystick manual 459801429121(IV Pole tooling maintained as FRU) | Done |  | During | [7 empty] | 115780086 | Closed | 196896 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | Valid | Not A Defect |  |  | N | Brook | Installation Other | Service |  |  |  | Closed | Installation Process Issue |  | Service workmanship | Fixed by FSE (re-install) | [7 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
21 | 44035 | TRUE | Turkey | T.C Saglik Bakanligi Bagcilar | 40010 | 44035 | 30 | Failure in first installation | Installed new Ingenuity5000 system as a FOK in Istanbul and after first day of operation rotation issue appeared and solved by replacing R2D 459800467341 | Yes |  | Under investigation bugreport ready for analysis, |  | Raise from CFM Under investigation |  |  | After | [7 empty] | CFM |  | 202722 | TRUE | Closed - Done | Closed - Done | Valid | CTUDT00630864 |  |  |  | Silva | Failure to Plan | Engineering |  |  |  | Closed | Hardware Failure |  | R2D | [8 empty] | System issue | System | Gantry-HW | R2D | Gantry-HW-R2D | Gantry-HW-R2D | CTUDT00630864 | R2D | Gantry-HW | Gantry-HW-R2D | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
22 | 44041 | TRUE | Turkey | T.C Saglik Bakanligi Bagcilar | 40010 | 44041 | 31 | xray error | xray hatası | P3 |  | Issue no longer occurring, no work performed by Philips |  | N/A |  |  | After | Gantry_HW | Tube_arc | [5 empty] | 115835618 | Closed | 210461 | TRUE | Closed - Cancelled | Closed - Cancelled | Valid | Not A Defect |  |  |  | Silva | Works as Designed | Engineering |  |  |  | Closed | System Performance Issue |  | Single case | Single case | [7 empty] | System issue | system | IC-HW | Tub Arc | IC-HW-Tub Arc | Gantry-HW-Tub Arc | Not A Defect | Tub Arc |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
23 | 44053 | TRUE | Turkey | T.C Saglik Bakanligi Bagcilar | 40010 | 44053 | 33 | Image transfer problem | Görüntü aktarım sorunu Issue no longer occurring, no work performed by Philips Solution for the Customer : Customer Declined Service, Support, Or Repair | P3 |  | Site Readiness Issues Bugreport received |  | N/A |  | 0723 0807 0810 | After | Remote Device | [6 empty] | 115937930 | Closed | 226030 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | Valid | Not a Defect |  |  |  | Silva | Works as Designed | Engineering |  |  |  | Closed | Site Readiness Issues |  | PACS | PACS | [7 empty] | System issue | System | Console-SW | Remote Connection-Not a Defect | Console-SW-Remote Connection-Not a Defect | Console-SW-Remote Connection | Not a Defect | Remote Connection | Console-SW | Console-SW-Remote Connection-Not a Defect | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
24 | 44060 | TRUE | Turkey | T.C Saglik Bakanligi Bagcilar | 40010 | 44060 | 34 | Unable to connect to worklist | Workliste bağlanmıyor Worklist server IP can be pinged. However, the remote node could not be examined because it was not possible to login with OLEVEL. Site visit required. It was determined that the problem was caused by the worklist provider. The problem was solved after meeting with the relevant people and restarting the worklist server infrastructure system. | P3 |  | Need bugreport, timestamp of the issue |  | N/A |  | https://docs.philips.com/:f:/g/personal/brook_shi_philips_com/EqKesoK_Dd9Jnz6skcO9kDgBsvBScyHCtUeIgYlFlUFY-A?e=kcDJ5a | After | [7 empty] | 115971072 | Closed | 234535 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | Valid | Not a Defect |  |  |  | Silva | Works as Designed | Others |  |  |  | Closed | Site Readiness Issues |  | PACS Vendor | PACS | [7 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
25 | 43986 | TRUE | US | Clay County Medical Center | 40003 | 43986 | 23 | Parts ID - CT dose probe | Parts ID - CT dose probe | P5 |  | OneEMS internal remarks Closing case, FSE Don using app to place parts ID on installation work order. |  | done |  |  | During | [7 empty] | 115561317 | Closed | 143519 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | Valid | Not a Defect |  |  |  | Silva | Works as Designed | Service |  |  |  | Closed | Installation Process Issue |  | Service Parts Order | Field work order | [7 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
26 | 44063 | TRUE | Turkey | T.C Saglik Bakanligi Bagcilar | 40010 | 44063 | 34 | Masadan ses geliyor | Masadan ses geliyor | P3 |  | The system was delivered actively by replacing the horizontal movement coupling, which was mechanically broken, with a new one. 453567099401,COUPLING/SPIDER |  | N/A |  |  | After | [7 empty] | 115990681 | Closed | 239416 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | Valid | CTUDT00632179 |  |  |  | Silva | Patient Handling | Engineering |  |  |  | Closed |  |  | Motor coupling | Fixed by FRU | [5 empty] | 453567099401,COUPLING/SPIDER |  | System issue | System | Couch-HW | Motor Coupling | Couch-HW-Motor Coupling | Couch-HW-Motor Coupling | CTUDT00632179 | Motor Coupling | Couch-HW | Couch-HW-Motor Coupling | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
27 | 44058 | TRUE | US | The Regional Medical Center | 40005 | 44058 | 33 | Gantry shut down | Gantry shut down | P5 |  | Couch issue 1.00000,459800440403,COMMON GANTRY MOTION PROCESSOR ASSY ROHS |  | Couch issue FCO72800735 Deployed |  |  | After | [7 empty] | 115966993 | Incomplete | 239676 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | Duplicate (12) | CTUDT00630138/630160 |  |  |  | Silva | Failure to scan | Engineering |  |  |  | Closed | Engineering Issue |  | FCO72800735 | [6 empty] | 459800440403,COMMON GANTRY MOTION PROCESSOR ASSY ROHS |  | System issue | System | Couch-SW | Horizontal Move Block | Couch-SW-Horizontal Move Block | Couch-SW-Horizontal Move Block | CTUDT00630138/630160 | Horizontal Move Block | Couch-SW | Couch-SW-Horizontal Move Block | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
28 | 44064 | TRUE | US | The Regional Medical Center | 40005 | 44064 | 34 | Needs to know the correct temperature in order to use the scanner | Needs to know the correct temperature in order to use the scanner | P4 |  | Temperature Requirements The ideal stable room temperature during operation is 22 degrees Celsius (72 degrees Fahrenheit). • The room-temperature range of operation should be stabilized to within 18 to 24 degrees Celsius (64 to 75 degrees Fahrenheit). This temperature range must be maintained at all times. • Temperature variations within the range (18-24 degrees Celsius; 64-75 degrees Fahrenheit) must not exceed 3 degrees Celsius (5 degrees Fahrenheit) during operation. • Temperature gradient must never exceed 3 degrees Celsius (5 degrees Fahrenheit) per hour. • The range of relative humidity in the room must be 35% to 70% (non-condensing). • Storage and transportation temperature should be -20 to -50 degrees Celsius; -4 to -122 degrees Fahrenheit) . The range of relative humidity must be 20% to 85% (non-condensing). |  | 459801063524 IFU Includes temp infor |  |  | After | [7 empty] | 115995104 | Closed | 240262 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | Valid | Not a defect |  |  |  | Silva | Applications/Usability | Customer |  |  |  | Closed | Positive Customer Feedback |  | Q&A | N/A (Not a Defect) | [7 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
29 | 44068 | TRUE | US | Mount Grant General Hospital | 40012 | 44068 | 35 | AAMI leads not shipping to NA | North American (Red, White, Black) leads did not accompany ECG monitor. European leads (Green Yellow Red) arrived with the equipment. Customer is not familiar with color identification when placing leads on patient. Other details: Missing hardware Equipment ID: 83439741 | Yes |  | Check SO which ECG ordered From PCM description, there are two selectable ECG, need check which one was in the sales order? Is it match which country? Dongping & Silva&XieHong: Check SRS impact and background of change in 2020/Jan? what is 12NC for North American (Red, White, Black) leads can ship to customer? |  | Under investigation with order desk and ECG contact |  |  | After | [7 empty] | CFM Gary Nearing |  | 247210 | TRUE | Closed - Done | Closed - Done | Valid | Not A Defect |  | Yes |  | David/Xie Hong | Applications/Usability | Operation |  |  |  | Closed |  |  | ECG Leads | Open | [7 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
30 | 44068 | TRUE | US | Mount Grant General Hospital | 40012 | 44068 | 35 | Cracked Compensator message on new install | While preforming a routine air calibration the following message occurred 4 min into the 18 min calibration: An artifact, which may be related to a cracked compensator, has been detected. Finish the calibration, run “IQ Check” test and check if the artifacts are observed on images. In any case, please inform service that a cracked compensator warning was displayed. The air calibration was restarted and failed again at the 4 minute point. The calibration was started for the third time and completed the without error. An automated QA was competed with the harmonized phantom before and after the air calibration error. The pre calibration images appeared to be normal, there were no artifacts recognized, there was nothing in scan apature. A brain and abdomen scan was preformed the day before with normal image quality. Don Parr NSS and Bruce Schuler, Primary FSE were notified by phone. I bug report was completed. Please contact Don Parr NSS for further information. Other details: Missing hardware Equipment ID: 83439741 | Yes |  | Beam and DMS check is integrated with Air Calibration for CT user in Refresh version, Factory MTS need update to include Beam and DMS check test |  | Under investigation Awaiting further field feedback check the distance of phantom all installed and then perfrom Air calibration to see if phantom in the beam, from end of phsicis layer to the beam. |  |  | After | [7 empty] | CFM Gary Nearing |  | 247224 | TRUE | Closed - Done | Closed - Done | Valid | CTUDT00630969 |  | Yes |  | Silva | Calibration/Performance | Engineering |  |  |  | Closed |  |  | Defective DMS module | Defective DMS module | 44068 | 0 | [5 empty] | System issue | System | DMS-HW | Compensator | DMS-HW-Compensator | DMS-HW-Compensator | CTUDT00630969 | Compensator | DMS-HW | DMS-HW-Compensator | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
31 | 44068 | TRUE | US | Willis Knighton Health System | 40004 | 44068 | 35 | scannable range line on pad doesn’t match table line | FOK scanner. Question if the correct couch pad was received. | Yes |  | blue is correct color for Refresh Ingenuity Mattress. It’s developed based on old mushroom Mattress, just change the color and keep same for all dimensions including scan line position. Need FSE on site check 1. What is the distance between the front end of Carbon Top and Marker on the Carbon Top? 2. What it the actual dimension of scan line position? 3. When the mattress was put on the Carbon Top and align with the front end of Carbon top, what is distance between the scan line on the mattress and marker on the Carbon top? |  | Under investigation this is not new issue check exist PR and Silva will take a look on this issue, but I think we need create defect to trace. Samuel will have a discussion with Robin for detail. |  |  | After | [7 empty] | CFM |  | 249304 | TRUE | Closed - Done | Closed - Done | Valid | CTUDT00631638 |  |  |  | Ada | User Experience | Engineering |  |  |  | Closed |  |  | Scan Line | [8 empty] | System issue | System | Couch-HW | Mattress scan line | Couch-HW-Mattress scan line | Couch-HW-Mattress scan line | CTUDT00631638 | Mattress scan line | Couch-HW | Couch-HW-Mattress scan line | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
32 | 44069 | TRUE | US | Willis Knighton Health System | 40004 | 44069 | 35 | Gantry not ready | Gantry not ready message after scanner sat idle for a few | Yes |  | Under investigation, bugreport https://docs.philips.com/:f:/g/personal/brook_shi_philips_com/EhWQLJDoQYFKglYEjCQFBk4BKNpiGXJvRK59SCeJIoShgQ?e=SJL0um |  | Under investigation @Silva&Tammy, bug analysis |  |  | After | [7 empty] | CFM |  | 261069 | TRUE | Closed - Done | Closed - Done | Valid | CTUDT00631039 |  |  |  | Silva | Failure to scan | Engineering |  |  |  | Closed |  |  | Gantry not ready | [8 empty] | System issue | System | Couch-SW | Horizontal Move Block | Couch-SW-Horizontal Move Block | Couch-SW-Horizontal Move Block | CTUDT00631039 | Horizontal Move Block | Couch-SW | Couch-SW-Horizontal Move Block | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
33 | 44068 | TRUE | US | Willis Knighton Health System | 40004 | 44068 | 35 | no phantom detected | During IQ check, scanner stated no phantom was detected. However, the phantom was clearly on the scout. The IQ check procedure was exited and reinitiated. IQ check ran ok after | Yes |  | Under investigation, bugreport https://docs.philips.com/:f:/g/personal/brook_shi_philips_com/EhWQLJDoQYFKglYEjCQFBk4BKNpiGXJvRK59SCeJIoShgQ?e=SJL0um |  | Under investigation @Silva&Tammy, bug analysis |  |  | After | [7 empty] | CFM |  | 249299 | TRUE | Closed - Done | Closed - Done | Valid | CTUDT00630971 |  |  |  | Silva | Calibration/Performance | Engineering |  |  |  | Closed |  |  | Phantom not detected | [8 empty] | System issue | System | DMS-SW | Phantom Detect | DMS-SW-Phantom Detect | DMS-SW-Phantom Detect | CTUDT00630971 | Phantom Detect | DMS-SW | DMS-SW-Phantom Detect | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
34 | 44070 | TRUE | Puerto Rico | METRO MAYAGUEZ INC | 40009 | 44070 | 35 | Pedal of CT Fluoroscopy is not working | Pedal of CT Fluoroscopy is not working | P3 |  | Field handling, reconnect |  |  |  |  | After | Gantry_HW | Pedal of CT Fluoroscopy | HW?SW? not clear |  |  |  |  | 116030772 | Closed | 251271 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | Valid | Not A Defect |  |  |  | NA | Works as Designed | Engineering |  |  |  | Closed |  |  | Pedal Connector | Fixed by FSE (re-connect) | [7 empty] | System issue | System | Gantry-HW | CCT pedal | Gantry-HW-CCT pedal | Gantry-HW-CCT pedal | Not A Defect | CCT pedal | Gantry-HW | Gantry-HW-CCT pedal | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
35 | 44070 | TRUE | US | The Regional Medical Center | 40005 | 44070 | 35 | GANTRY CONTROL PANEL ARE NOT WORKING | GANTRY CONTROL PANEL ARE NOT WORKING closed this case as Fixed 1.00000,459801484981,CONTROL PANEL ASSEMBLY-FRONT RIGHT 1.00000,459801485021,CONTROL PANEL ASSEMBLY-REAR RIGHT 1.00000,459801484991,CONTROL PANEL ASSEMBLY-FRONT LEFT 1.00000,459800325772,BRUSH, PWR BLOCK PCB ASSEMBLY | P1 |  | Field handling, parts orderred, plan to replace the panels and implement FCO72800735 Sep-1 link to #xxx |  | FCO72800735 |  |  | After | [7 empty] | 116020708 | Closed | 249503 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | Duplicate (12) | CTUDT00630138/630160 |  |  |  | Robin | Lockup or Crash | Engineering |  |  |  | Closed |  |  | FCO72800735 | [8 empty] | System issue | System | Couch-SW | Horizontal Move Block | Couch-SW-Horizontal Move Block | Couch-SW-Horizontal Move Block | CTUDT00630138/630160 | Horizontal Move Block | Couch-SW | Couch-SW-Horizontal Move Block | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
36 | 44070 | TRUE | US | Mount Grant General Hospital | 40012 | 44070 | 35 | Table noise | Table base is creaking loudly when 12” from the completely down position and 12” away in the upward movement. Video with sound included in this report. Equipment ID : 83439741 | Yes |  | repositioned the spring and used black grease on the top and bottom. No more noise. |  | Defect raise to track design improvement, WI update? |  |  | After | [7 empty] | CFM |  | 250901 | TRUE | Closed - Done | Closed - Done | Valid | CTUDT00631058 |  |  |  | Robin | Results Not as Expected | Engineering |  |  |  | Closed |  |  | Couch Spring | [8 empty] | System issue | System | Couch-HW | Vertical spring | Couch-HW-Vertical spring | Couch-HW-Vertical spring | CTUDT00631058 | Vertical spring | Couch-HW | Couch-HW-Vertical spring | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
37 | 44068 | TRUE | Turkey | T.C Saglik Bakanligi Bagcilar | 40010 | 44068 | 35 | E-stop opens | E-stop açıyor | P3 |  | FCO735 implemented, no further action |  | FCO735 |  |  | After | [7 empty] | 116011193 | In progress | 245011 | TRUE | Closed - Done | Closed - Done | Valid | CTUDT00630138/630160 |  |  |  | Robin | E-Stop | Engineering |  |  |  | Closed |  |  | FCO72800735 | [8 empty] | System issue | System | Couch-SW | Horizontal Move Block | Couch-SW-Horizontal Move Block | Couch-SW-Horizontal Move Block | CTUDT00630138/630160 | Horizontal Move Block | Couch-SW | Couch-SW-Horizontal Move Block | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
38 | 44068 | TRUE | Germany | Gesundheit Nord gGmbh | 40002 | 44068 | 35 | hor. Table movement is temporarily not possible | über Applikation Brachtendorf/Bösemann hor. Tischbewegung ist zeitweise nicht mögl. vertikale Bewegung funktioniert, Tisch läßt sich nicht ohne Betätigung der TapeSwitches manuell bewegen "via application Brachtendorf / Bösemann hor. Table movement is temporarily not possible. vertical movement works, table cannot be moved manually without actuating the tape switches " | P4 |  | FCO735 implemented, no further action |  | FCO735 |  |  | After | [7 empty] | 116011421 | In progress | 245035 | TRUE | Closed - Done | Closed - Done | Valid | CTUDT00630138/630160 |  |  |  | Robin | Lockup or Crash | Engineering |  |  |  | Closed |  |  | FCO72800735 | [8 empty] | System issue | System | Couch-SW | Horizontal Move Block | Couch-SW-Horizontal Move Block | Couch-SW-Horizontal Move Block | CTUDT00630138/630160 | Horizontal Move Block | Couch-SW | Couch-SW-Horizontal Move Block | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
39 | 44063 | TRUE | US | Mount Grant General Hospital | 40012 | 44063 | 34 | part order for Bruce Schuler | part order for Bruce Schuler 1.00000,459801234171,LCC CRC RACK USB 3.0 EXT CBL 1.00000,455300027201,"CBL,USB A/M TO B/M 0.5M | P5 |  | DEFOA CD-0000030584 Per FSE parts are no longer needed on site.. |  | Operation investigation |  |  | During | [7 empty] | 115988967 | Incomplete | 240242 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | Duplicate (77) | Not A Defect |  |  |  | David | Missing Item | Operation |  |  |  | WIP |  | Missing Item | Not Applicable | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
40 | #N/A | #N/A | China | General Hospital of Jizhong Energy Jingxing Mining Group Co. LTD | 345003 | 44077 | 36 | Plinth couch base cover mismatch | The part 459801540481 shall be shipped for Ingenuity couch base cover, but 459800832761 (which is for ICT, IQON couch) received on site. | Yes |  | New part shipped to Site. |  | OPS investigate the rootcause and define prevent action (from supplier and Suzhou IQC) |  |  | During | [7 empty] | 116010968 |  | DEFOA CD-30791 | #N/A | #N/A | DEFOA | Valid | Not A Defect |  |  |  | David/Lynn/Erika? | Missing Item | Operation |  |  |  | Closed |  |  | DEFOA | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
41 | 44083 | TRUE | China | Huaying Mountain Guangneng General Hospital | 345000 | 44083 | 37 | Cannot move couch horizontally | System has implemented Couch patch in Suzhou factory(4.1.10.25202, couch FW:5.3.39.1500). Couch was not able to move horizontally system could not scan. check the errorlog, error "S_HMC_AXIS_Wrong_DEST_REACHED " was been found. | P3 |  | Investigation and Bugreport analysis (travel range indicate as 1903) |  | WI update to make sure the travel range is correct, ECR will be raised to define action |  |  | After | [7 empty] | 116080577 | In Process | 269690 | TRUE | Closed - Done | Closed - Done | Valid | CTUDT00631156 |  |  |  | Changxian/Jitao | Lockup or Crash | Engineering |  |  |  | Closed |  |  | Horizontal movement | [8 empty] | System issue | System | Couch-SW | Horizontal Move Block | Couch-SW-Horizontal Move Block | Couch-SW-Horizontal Move Block | CTUDT00631156 | Horizontal Move Block | Couch-SW | Couch-SW-Horizontal Move Block | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
42 | 44087 | TRUE | China | Huaying Mountain Guangneng General Hospital | 345000 | 44087 | 37 | LAN Config tool cannot open | System configured ISP and PRS remote service (one fixed IP for hospital, one fixed IP for 4G wireless route), then LAN config tool open with error | Yes |  | This seems to be a corner case when we have multiple IPs configured the Lanconfig is not able to load more than 2 IPs. Hence throwing the error. Also only when two IPs are configured the Lanconfig is taking the second IP as IPv6. Since before Refresh there was no option for IPv6 and IPv4 to be displayed separately and there was only one field to display IP, this was not getting displayed. This needs to be fixed to handle. Please raise a defect. |  |  |  |  | After | [7 empty] | CFM |  | 277226 | TRUE | Closed - Done | Closed - Done | Valid | CTUDT00631055 |  |  |  | Silva/Tammy | LanConfig/DICOM | Engineering |  |  |  | Closed |  |  | LAN Config Tool | [8 empty] | System issue | System | Console-SW | Lan Config | Console-SW-Lan Config | Console-SW-Lan Config | CTUDT00631055 | Lan Config | Console-SW | Console-SW-Lan Config | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
43 | 44087 | TRUE | China | Huaying Mountain Guangneng General Hospital | 345000 | 44087 | 37 | Message popup when Ingenuity image add into ISP film print | Message popup every time when Ingenuity image add into ISP film print, images from Ingenuity 4.1.7 does not have below message popup. The Film application is a generic application that is used to prepare results for printing and for miscellaneous activities not affiliated with IntelliSpace Portal, Spectral images are non-specific annotations and other types of identification that enable proper interpretation are retained and can be seen in other viewers. Images added in film application successfully after click the OK the message. | Yes |  | From the issue shared we understood there was images sent to ISP from Refresh Ing and Ing IB version. For the first one , when images are sent ISP print , the popup mentioned came repeatedly for each image. Where as for the second no popup came. Based on discussion with ISP, the popup from ISP film may come depending on 1. How the images are sent to ISP film. Is it from ISP review application to ISP film (or) from ISP PD to ISP Film. Please note directly it cannot be send to ISP film from another system. So confirm how the images are sent to ISP film. 2. Type of images sent. From shared images in bugrep, we see that the images are MPR\derived. Was it the same images(derived images) used from 4.1.7 as well. If not, please explain the type of images sent to ISP , and was it the same set of images sent from 4.1.7 as well. |  | Defect raised |  |  | After | [7 empty] | CFM |  | 277227 | TRUE | Closed - Done | Closed - Done | Valid | CTUDT00631142 |  |  |  | Silva/Tammy | User Experience | Engineering |  |  |  | Closed |  |  | ISP | [8 empty] | System issue | System | Console-SW | Spectral Error Message | Console-SW-Spectral Error Message | Console-SW-Spectral Error Message | CTUDT00631142 | Spectral Error Message | Console-SW | Console-SW-Spectral Error Message | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
44 | 44077 | TRUE | US | Mount Grant General Hospital | 40012 | 44077 | 36 | Non standard installation | Adobe reader version higher than execept version, the system is configured with PRS There is Scheduled Adobe Acrobat Update Task in every day which may trigger automatically version update and cause version difference | No |  | Need disable the auto upgrade? Any impact on function? |  | defect raised, handling by DRB. |  |  | After | [7 empty] | CFM |  | 268516 | TRUE | Closed - Done | Closed - Done | Valid | CTUDT00630991 |  |  |  | Silva/Tammy | Results Not as Expected | Engineering |  |  |  | Closed |  |  | Non standard installation | [8 empty] | System issue | System | Console-SW | InstallationVersion | Console-SW-InstallationVersion | Console-SW-InstallationVersion | CTUDT00630991 | InstallationVersion | Console-SW | Console-SW-InstallationVersion | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
45 | 44075 | TRUE | US | The Regional Medical Center | 40005 | 44075 | 36 | Customer feedback to be improved | While onsite implement an FCO, the Customer provided feedback for things she would like to see improved. The first feedback was the design of the Patient Support. The Customer feels it is very difficult to move a gurney into position to transfer to the Patient Support Tabletop due to the Tabletop Handle protruding outward. The second feedback is the Customer would like to see the Surview display as it is being acquired in the event they feel they have enough coverage and prevent unnecessary exposure. The third feedback is the Customer would like to have a means to move the Tabletop into the exact required position during a biopsy or interventional procedure without having to move the tabletop manually. | Yes |  | review the customer feedback and evaluate the clinical benefit 1. in next generation Couch 2. check console spec 3. IVC feature |  | Defect TO BE raised |  |  | After | [7 empty] | CFM |  | 266893 | TRUE | Closed - Done | Closed - Done | Valid | CTUDT00631134 |  |  |  | Kevin | Positive Feedback | Engineering |  |  |  | Closed | Positive Customer Feedback |  | Improvement | [8 empty] | System issue | System | Couch-HW | RFC | Couch-HW-RFC | Couch-HW-RFC | CTUDT00631134 | RFC | Couch-HW | Couch-HW-RFC | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
46 | 44083 | TRUE | Puerto Rico | METRO MAYAGUEZ INC | 40009 | 44083 | 37 | Enable Remote Access not present in Home Menu | Enable Remote Access not present in Home Menu |  |  | Sridhar, Sandeep, investigation Even retore the full backup of the system, the remote access settting can not be recovered. (Lifang checked with R&D system BR64) |  |  |  |  | After | [7 empty] | CFM |  | 275190 | TRUE | Closed - Done | Closed - Done | Valid | CTUDT00631143 |  |  |  | Silva | System Config. Issues | Engineering |  |  |  | Closed |  |  | DoD Config Tool | [8 empty] | System issue | System | Console-SW | Service Tool | Console-SW-Service Tool | Console-SW-Service Tool | CTUDT00631143 | Service Tool | Console-SW | Console-SW-Service Tool | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
47 | 44083 | TRUE | US | Willis Knighton Health System | 40004 | 44083 | 37 | Rotation issue in Slab viewer | While using the slab viewer tool in Edit Results or Recon mode, it is not possible to rotate the image until it has been swiveled first. In v.4.1.7 (Brilliance 64, Ingenuity or iCT) and v.4.7.5 (IQon) it was possible to rotate the image without having to perform a swivel motion first. This is a loss of functionality with the new Refresh software version. |  |  | Defect to be raised |  |  |  |  | After | [7 empty] | CFM |  | 278078 | TRUE | Closed - Done | Closed - Done | Valid | CTUDT00631136 |  |  |  | Kevin | Work Flow | Application |  |  |  | Closed | Engineering Issue |  | Slab Viewer | [8 empty] | System issue | system | Console-SW | SlabView | Console-SW-SlabView | Console-SW-SlabView |  | 0 |  |  | No | 0 | 0 | 0 | 0 | 0 | No | y | y | y | y | y | y | y | y | y | y | y
48 | 44083 | TRUE | US | Willis Knighton Health System | 40004 | 44083 | 37 | EStop on Couch Zero | When the customer pressed the zero couch button on the right side of the gantry, the estop opened. Then all red Estop buttons were checked and the Estop Key was then used to close the Estop. Gantry came back on and scanning continued. |  |  | Need bugreport, timestamp of the issue |  |  |  |  | After | [7 empty] | CFM |  | 278069 | TRUE | Closed - Done | Closed - Done | Valid | CTUDT00631165 |  |  |  | Robin | Failure to scan | Engineering |  |  |  | Closed |  |  | E-Stop | [8 empty] | System issue | System | Couch-HW | Horizontal motor | Couch-HW-Horizontal motor | Couch-HW-Horizontal motor | CTUDT00631165 | Horizontal motor | Couch-HW | Couch-HW-Horizontal motor | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
49 | 44083 | TRUE | US | The Regional Medical Center | 40005 | 44083 | 37 | Recon - Slab Tool-Image Rotation not working properly | This is one of the new v.4.1.10 CT 5000 (Ingenuity scanner). While using the slab viewer tool in Edit Results or Recon mode, it is not possible to rotate the image until it has been swiveled first. In v.4.1.7 (Brilliance 64, Ingenuity or iCT) and v.4.7.5 (IQon) it was possible to rotate the image without having to perform a swivel motion first. This is a loss of functionality with the new Refresh software version. Like this, it is now an extra mouse click for the Operator. If image is not out of alignment, the Operator has to swivel the image a little before they can change the rotation of the image. This slows the Operator’s workflow down. Workflow steps: The workflow is to take an image series to Recon (or use Edit Results). Click on the Slab Viewer icon at the bottom of the page. Click on the swivel icon (or use the Rt. & Lt. mouse shortcut), position the mouse on the any of the outer portions of the viewport and rotate the image. In v.4.1.7, the image would rotate without performing any other functions. Now with v. 4.1.10, it is not possible to rotate the image unless you perform the swivel motion first. I have a 2 videos demonstrating this workflow. One from a v4.1.10 system and one from a v.4.1.7 system. |  |  | Defect to be raised Attached is the video I made of how the rotate function works in v4.1.7. It does not require you to swivel the image prior to being able to rotate the image. v.4.1.10 is requiring the swivel function to be use first prior to being able to rotate. |  |  |  |  | After | [7 empty] | CFM |  | 274937 | TRUE | Closed - Done | Closed - Done | Valid | CTUDT00631136 |  |  |  | Kevin | Work Flow | Application |  |  |  | Closed | Positive Customer Feedback |  | Slab Viewer | [8 empty] | System issue | system | Console-SW | SlabView | Console-SW-SlabView | Console-SW-SlabView |  | 0 |  |  | No | 0 | 0 | 0 | 0 | 0 | No | y | y | y | y | y | y | y | y | y | y | y
50 | 44084 | TRUE | France | SCP GROUPE D IMAGERIE MEDICALE | 40019 | 44084 | 37 | NO MORE COUCH HORIZONTAL MOUVMENT | Couch Linkage broken |  |  | to be checked if Refresh couch have cut in new linkage, and if this is reported from 4.1.7 system or 4.1.10 4535-670-23325 STD PATIENT SUPPORT WHT, IN SN: 905010 |  | check if the new linkage cut into production, how many shipped system with old linkage. ESPF Check and update if needed |  |  | After | [7 empty] | CFM |  | 272532 | TRUE | Closed - Done | Closed - Done | Valid | CTUDT00631671 |  |  |  | Robin | Failure to scan | Engineering |  |  |  | Closed |  |  | Linkage | [8 empty] | System issue | System | Couch-HW | Horizontal linkage | Couch-HW-Horizontal linkage | Couch-HW-Horizontal linkage | CTUDT00631671 | Horizontal linkage | Couch-HW | Couch-HW-Horizontal linkage | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
51 | 44064 | TRUE | US | The Regional Medical Center | 40005 | 44064 | 34 | Patient Support: Table will not advance into the gantry. |  | P3 | [6 empty] | After | [7 empty] | 115999394 |  | 262755 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | Duplicate (12) | CTUDT00630138/630160 |  |  |  | Robin | Failure to scan | Engineering |  |  |  | Closed |  |  | FCO72800735 | [8 empty] | System issue | System | Couch-SW | Horizontal Move Block | Couch-SW-Horizontal Move Block | Couch-SW-Horizontal Move Block | CTUDT00630138/630160 | Horizontal Move Block | Couch-SW | Couch-SW-Horizontal Move Block | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
52 | 44067 | TRUE | Germany | Gesundheit Nord gGmbh | 40002 | 44067 | 35 | There is a message about the remote control. We can not hide this message. | There is a message about the remote control. We can not hide this message. We only can put it in another position. It´s very disturbing for the customer. When we put this field into another position on the right monitor, the system lost this information and after a restart the field is again on the left monitor. The position of this field is not so good, because the field is in front of the tool buuton, where the customer makes the short TC. |  |  | Defect raised |  |  |  |  | After | [7 empty] | CFM |  | 243621 | TRUE | Closed - Done | Closed - Done | Valid | CTUDT00631140 |  |  |  | Silva | Applications/Usability | Engineering |  |  |  | Closed |  |  | Remote Service Tool | [8 empty] | System issue | System | Console-SW | Service Tool | Console-SW-Service Tool | Console-SW-Service Tool | CTUDT00631140 | Service Tool | Console-SW | Console-SW-Service Tool | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
53 | 44074 | TRUE | Germany | Gesundheit Nord gGmbh | 40002 | 44074 | 36 | Noises in the gantry | 1. Problem Table does not move horizontally Troubleshooting: Analysis of gantry logs: S_HMC_HOR_MOTOR_RELAY_OPEN Relay has already been exchanged after power cycle function ok Creation of bug reports and forwarding for escalation to TSI G.Brücker 2. Failure Noises from gantry. please 1x 459801413791 IBOX2 USB Extender kit 1x 455300027201 USB cable order to the customer The USB Extender for the IBox has been replaced. In the tests carried out, the system corresponds to the manufacturer_s specifications and is ready for clinical use. | P4 |  | USB deactivation and activation Sound of the IBox connection can be heard at times Field handling, USB connection from the gantry to the CT host has been exchanged. In the tests carried out, the system corresponds to the manufacturer_s specifications and is ready for clinical use. USB connection from gantry to CT host breaks down temporarily. 1.00000,459801413791,IBOX2 USB EXTENDER Kit 1.00000,455300027201,CBL,USB A/M TO B/M 0.5M |  | TBD |  |  | After | [7 empty] | 116041656 | Incomplete | 254677 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | Valid | Not A Defect |  |  |  | William | Failure to scan | Engineering |  |  |  | Closed |  |  | USB Extender | [8 empty] | System issue | System | Gantry-HW | IBox | Gantry-HW-IBox | Gantry-HW-IBox | Not A Defect | IBox | Gantry-HW | Gantry-HW-IBox | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
54 | 44074 | TRUE | Puerto Rico | METRO MAYAGUEZ INC | 40009 | 44074 | 36 | Pedal of CT Fluoroscopy is not working | Pedal of CT Fluoroscopy is not working | P3 |  | Found the connection of the pedal on the GPU not making good connection. Connector was reseted. |  |  |  |  | After | Gantry_HW | Pedal of CT Fluoroscopy | HW?SW? not clear | FSE to handle | Keep monitoring | William Yang |  | 116047112 | Closed | 256431 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | Duplicate (34) | Not A Defect |  |  |  | NA | Works as Designed | Engineering |  |  |  | Closed |  |  | Pedal Connector | [8 empty] | System issue | System | Gantry-HW | CCT pedal | Gantry-HW-CCT pedal | Gantry-HW-CCT pedal | Not A Defect | CCT pedal | Gantry-HW | Gantry-HW-CCT pedal | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
55 | 44074 | TRUE | China | Jinggangshan Hospital, the First Affiliated Hospital of Nanchang University, Jiangxi | 345002 | 44074 | 36 | Application/Ingenuity Core128/Elite China | IT Password forget | P5 |  | Field handling |  |  |  |  | After | [7 empty] | 116043077 | Closed | 255185 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | Valid | Not A Defect |  |  |  | Kevin | Applications/Usability | Application |  |  |  | Closed |  |  | User Account | User Password | [7 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
56 | 44075 | TRUE | Turkey | T.C Saglik Bakanligi Bagcilar | 40010 | 44075 | 36 | The device activates e-stop. | Upon the user_s complaint about e-stop, system error records were examined by connecting to the system remotely. In the examination, it was observed that Ghost rebooted himself and the problem was caused by this. A site visit is required for detailed investigation. The system is operational. Gantry laser cables must be examined by making a site visit. Cable damage can cause Ghost reboot. If the cable is intact, the source of the fault is likely to be Ghost. | P3 |  | Field handling |  | need check if this is cable non-fixed issue, Repeat case as PG |  |  | After | [7 empty] | 116051295 | Closed | 257297 | TRUE | Closed - Done | Closed - Done | Valid | CTUDT00631634 |  |  |  | William | E-Stop | Engineering |  |  |  | Closed |  |  | Laser Cable | [8 empty] | System issue | System | Gantry-HW | gantry laser | Gantry-HW-gantry laser | Gantry-HW-GantryLaser | CTUDT00631634 | gantry laser | Gantry-HW | Gantry-HW-gantry laser | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
57 | 44075 | TRUE | China | Jinggangshan Hospital, the First Affiliated Hospital of Nanchang University, Jiangxi | 345002 | 44075 | 36 | 02G/Ingenuity Core128/Elite China/Account number and password | 已联系xie sheng，谢经理告知设备在南昌大学第一附属医院井冈山中医院，告知客户昨天已来电过，此问题不属于维修工程师处理范围，谢经理告知已联系应用培训老师，过两天会到场跟进培训 | P3 |  | 已联系xie sheng，谢经理告知设备在南昌大学第一附属医院井冈山中医院，告知客户昨天已来电过，此问题不属于维修工程师处理范围，谢经理告知已联系应用培训老师，过两天会到场跟进培训 |  |  |  |  | After | [7 empty] | 116049951 | Closed | 257233 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | Duplicate (55) | Not A Defect |  |  |  | Kevin | Applications/Usability | Application |  |  |  | Closed |  |  | User Account | User Password | [7 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
58 | 44076 | TRUE | US | Clay County Medical Center | 40003 | 44076 | 36 | Gantry can not get message | Install Power Monitor, Filament Calibration, Long Tube conditioning and troubleshoot HV. Notes has been added with title Internal Remarks Body: Planner Initials: TH Actions Taken: I pulled the case from the board and reviewed the follow up activity note. FSE Jeffrey Stauffer requested to have a follow up visit for 09/15 Description by Engineer Description: Customer Function/Role: CT Tech Device Use: General CT exams Expected and Unexpected Results: Expected is uninterupted ct exams. System is giving intermittent High Voltage Faults. User Impact: Delayed scans Patient Impact: Delayed Scans Current Software Version: 4.1.10 Downloaded and reviewed bug report and recommended that Jeff run the TC Boost procedure to condition tube3 | P3 |  | Field handling |  | Get bugreport from field for investigation |  |  | After | [7 empty] | 116060736 |  | 261212 | TRUE | Closed - Done | Closed - Done | Valid | CTUDT00630085 |  |  |  | Silva | Failure to Plan | Engineering |  |  |  | Closed |  |  | Tube arc | Tube Arc | [7 empty] | System issue | System | IC-HW | Tub Arc | IC-HW-Tub Arc | IC-HW-Tub Arc | CTUDT00630085 | Tub Arc | IC-HW | IC-HW-Tub Arc | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
59 | 44076 | TRUE | Puerto Rico | METRO MAYAGUEZ INC | 40009 | 44076 | 36 | Gantry Shuts off | Problem Description by Engineer Description: Customer Function/Role: Technology Radiology Device Use: Interventional radiology Expected and Unexpected Results: Gantry shutdown when laser is switch on. User Impact: Study will faile. Patient Impact: No patient intact Current Software Version: 4.1.10 New Work Order Activity has been created with type: Diagnostic performed by Engineer Description: Findings: - Gantry fails when the laser button is pressed. - Found the +5v laser cable touching the metal edge of the PDMC frame. This cusses the GPU to reset. When this happen the gantry shut down also. | P2 |  | Field handling |  | under investigation with Gantry team to check the cable connection, SQ work with supplier to improve the cabling of the |  | ECR-075587 Ing cover design change | After | [7 empty] | 116061078 | Closed | 261207 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | Valid | CTUDT00631634 |  |  |  | Silva | Investigating | Engineering |  |  |  | Closed |  |  | Supplier Quality | [8 empty] | System issue | System | Gantry-HW | gantry laser | Gantry-HW-gantry laser | Gantry-HW-GantryLaser | CTUDT00631634 | gantry laser | Gantry-HW | Gantry-HW-gantry laser | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
60 | 44077 | TRUE | US | Clay County Medical Center | 40003 | 44077 | 36 | Couch motion errors | changed Escalation Accepted from true to false. | P3 | [6 empty] | After | [7 empty] | 116066094 |  | 263001 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | Valid | CTUDT00630138/630160 |  |  |  | Robin | Lockup or Crash | Engineering |  |  |  | Closed |  |  | Invalid complaint | Invalid complaint | [7 empty] | System issue | System | Couch-SW | Horizontal Move Block | Couch-SW-Horizontal Move Block | Couch-SW-Horizontal Move Block | CTUDT00630138/630160 | Horizontal Move Block | Couch-SW | Couch-SW-Horizontal Move Block | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
61 | 44077 | TRUE | US | Mount Grant General Hospital | 40012 | 44077 | 36 | Q/O: Warranty | • Swap M27 with M9. Still getting cracked compensator error during air cal after changing module 27 Parts Consumed Description: 1.00000,459800911271,TDMS V2 M.Det Class A FRU | P5 |  | Defective module replaced. |  |  |  |  | After | [7 empty] | 116061562 |  | 261462 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | Duplicate (30) | CTUDT00630969 |  |  |  | Silva | DMS | Engineering |  |  |  | Closed |  |  | Defective DMS module | Defective DMS module | 44068 | #NUM! |  |  |  | `459800911271 |  | System issue | System | DMS-HW | Compensator | DMS-HW-Compensator | DMS-HW-Compensator | CTUDT00630969 | Compensator | DMS-HW | DMS-HW-Compensator | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
62 | 44078 | TRUE | Puerto Rico | METRO MAYAGUEZ INC | 40009 | 44078 | 36 | There is an image artifact | INTERMITTENT PROBLEM FSE said this service was already done, but the ticket is created for documenting and follow up. | P4 |  | Field handling, check the Detector calibration result after replaced The bad detectors were placed on the bad detector list. This compensate for the faulty detector. A new appointment will be schedule to replaced the detectors. |  | Get scan protocol and image from service if possible (Peter, Brook) |  |  | After | [7 empty] | 116073264 | Closed | 264727 | TRUE | Closed - Done | Closed - Done | Valid | CTUDT00632110 |  |  |  | Silva | Investigating | Engineering |  |  |  | Closed |  |  | Defective DMS module | Defective DMS module | [7 empty] | System issue | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | DMS Module | DMS-HW | DMS-HW-DMS Module | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
63 | 44078 | TRUE | China | Huaying Mountain Guangneng General Hospital | 345000 | 44078 | 36 | Couch can not move | 床不能水平运动 设备本身技术问题： 是。 Troubleshooting Action: 远程诊断与测试结果： 很难判断问题 PRS TR状态： 不通，请工程师到场检查。 Repair Action: 建议现场维修步骤： 建议工程师到场检查log，电源，Ghost。。。 | P2 |  | Travel range is too short to cause the couch movement issue. Adjust the trave range from 1903 to 1906 |  | ECR-076784 Couch MTS Update to control the range during manufucturing |  |  | After | [7 empty] | 116069567 |  | 264536 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | Duplicate (41) | CTUDT00631156 |  |  |  | Robin | Failure to scan | Engineering |  |  |  | Closed |  |  | Horizontal movement | [8 empty] | System issue | System | Couch-SW | Horizontal Move Block | Couch-SW-Horizontal Move Block | Couch-SW-Horizontal Move Block | CTUDT00631156 | Horizontal Move Block | Couch-SW | Couch-SW-Horizontal Move Block | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
64 | 44080 | TRUE | China | Huaying Mountain Guangneng General Hospital | 345000 | 44080 | 36 | 02N/Ingenuity Core128/Elite China/Scanning couch fault | Scanning couch fault | P3 |  | Need investigate and define action in couch production line to avoid same issue reported |  |  |  |  | After | [7 empty] | 116074943 |  | 265579 | TRUE | Closed - Done | Closed - Done | Duplicate (41) | CTUDT00631156 |  |  |  | Robin | Failure to scan | Engineering |  |  |  | Closed |  |  | Horizontal movement | [8 empty] | System issue | System | Couch-SW | Horizontal Move Block | Couch-SW-Horizontal Move Block | Couch-SW-Horizontal Move Block | CTUDT00631156 | Horizontal Move Block | Couch-SW | Couch-SW-Horizontal Move Block | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
65 | 44082 | TRUE | China | Huaying Mountain Guangneng General Hospital | 345000 | 44082 | 37 | 02N/Ingenuity Core128/Elite China/Couch fault | 偶发床不能水平运动 客户问题详细描述或故障现象： 重新开机正常 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： S_HMC_AXIS_WRONG_DEST_REACHED 远程诊断与测试结果： 新装机，怀疑床水平校准问题 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议工程师到场做床水平校准，检测丝杠水平运动是否顺畅。 | P3 |  | Need investigate and define action in couch production line to avoid same issue reported |  |  |  |  | After | [7 empty] | 116080577 |  | 266795 | TRUE | Closed - Duplicate | Closed - Duplicate | Duplicate (41) | CTUDT00631156 |  |  |  | Robin | Failure to scan | Engineering |  |  |  | Closed |  |  | Horizontal movement | [8 empty] | System issue | system | Couch-SW | Horizontal Move Block | Couch-SW-Horizontal Move Block | Couch-SW-Horizontal Move Block | CTUDT00631156 | Horizontal Move Block |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
66 | 44081 | TRUE | Germany | Gesundheit Nord gGmbh | 40002 | 44081 | 37 | Problems when sending images to the PACS | 1. Customer was instructed in the short scan preparation. Basic operating instructions. 2. The data connection to the Pacs is sometimes very slow or stops. After restarting the CT Host, all data was transferred. Failure could not be simulated. Medical technology contacts the Pacs manufacturer to discuss possible restrictions. 3. Joystick upgrade still needs to be completed. An appointment will be made. In the tests carried out, the system corresponds to the manufacturer_s specifications and is ready for clinical use. | P3 |  | Need bugreport, timestamp of the issue, waiting FSE Halama, Thomas <thomas.halama@philips.com> feedback |  |  |  |  | After | [7 empty] | 116078260 | Fixed | 267626 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | Valid | CTUDT00632127 |  |  |  | NA | PACS | Engineering |  |  |  | Closed |  |  | PACS | PACS | [7 empty] | System issue | system | Console-SW | Remote Connection | Console-SW-Remote Connection | Remote Connection | CTUDT00632127 | Remote Connection |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
67 | 44081 | TRUE | Turkey | T.C Saglik Bakanligi Bagcilar | 40010 | 44081 | 37 | Software freezes. | During pacsa patient transfer, the queue freezes. | P3 |  | Need bugreport, timestamp of the issue Pacs_a transfer sırasında queue manager da donmalar yaşanıyor. Günde 5-6 kez windows donuyor |  | Need bugreport, timestamp of the issue |  |  | After | [7 empty] | 116076599 |  | 267191 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | Valid | CTUDT00631602 |  |  |  | Silva | Lockup or Crash | Engineering |  |  |  | Closed |  |  | PACS | PACS | [7 empty] | System issue | System | Console-SW | Remote Connection-1602TBD | Console-SW-Remote Connection-1602TBD | Console-SW-Remote Connection | CTUDT00631602 | Remote Connection | Console-SW | Console-SW-Remote Connection-1602TBD | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
68 | 44084 | TRUE | US | Willis Knighton Health System | 40004 | 44084 | 37 | Open Estop Intermittently | duplicate with #48 ESTOP issue when reset couch to zero | P4 |  | Field handling |  |  |  |  | After | [7 empty] | 116099753 |  | 273427 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | Duplicate (48) | CTUDT00631165 |  |  |  | Robin | Failure to scan | Engineering |  |  |  | Closed |  |  | E-Stop | [8 empty] | System issue | System | Couch-HW | Horizontal motor | Couch-HW-Horizontal motor | Couch-HW-Horizontal motor | CTUDT00631165 | Horizontal motor | Couch-HW | Couch-HW-Horizontal motor | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
69 | 44084 | TRUE | us | The Regional Medical Center | 40005 | 44084 | 37 | Noise coming from gantry and CT box | Noise coming from gantry and CT box Noise coming from gantry and CT box. Recommend replacing CRC- Power Distribution Module 459801032622. | P4 |  | Field handling Diagnostic performed by Engineer Description: system reboot Debrief Status has been added as The system fully meets specification and returned to use |  | TBD |  |  | After | Console-HW | [6 empty] | 116099229 | Closed | 273273 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | Valid | Not A Defect |  |  |  | William | Noise | Engineering |  |  |  | Closed |  |  | Gantry Noise | Noise | [7 empty] | System issue | System | Console-HW | CRC Rack | Console-HW-CRC Rack | Console-HW-CRC Rack | Not A Defect | CRC Rack | Console-HW | Console-HW-CRC Rack | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
70 | #N/A | #N/A | China | Wuhan Ren _ai Hospital, Hubei Province | 345004 | 44089 | 38 | X-ray system cannot comply. | Why? (E.g. identified and confirmed root causes) After diagnosis and investigation, according to the Logger error S_GENERATOR_MONITOR_MA_BALANCE_ERROR, the RHOST SG164 ASSEMBLY with CPM (RoHS) was replaced, and the preheating and bed positioning film scanning were performed again, and both returned to normal, and subsequent debugging can also pass. How? (E.g. to replicate symptom, Which steps taken?) 1. Check the logger file, prompt S_GENERATOR_MONITOR_MA_BALANCE_ERROR 2. Check whether the bulb and the four high voltage cathode and anode plugs are abnormal, and the results are normal. 3. Clean up the four cathode and anode plugs and reapply the silicone grease, and perform preheating and scanning of the bed positioning film again, but an error is reported. 4. Check the RHOST SG164 ASSEMBLY with CPM (RoHS) and replace it with a new one. Perform pre-heating and scanning of the bed positioning film again to return to normal. | Yes |  | request the field part to be sent back for failure analysis. |  | OPS investigate the rootcause and define prevent action (from supplier and Suzhou IQC) |  |  | During | [7 empty] | 116068133 |  | DEFOA CD-30934 | #N/A | #N/A | DEFOA | Valid | Not A Defect |  |  |  | David | Failure to scan | Operation |  |  |  | WIP |  |  | DEFOA | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
71 | 44081 | TRUE | France | SCP GROUPE D IMAGERIE MEDICALE | 40019 | 44081 | 37 | Linkage broken | Linkage broken on site | P2 |  | Replace the linkage |  | check if the new linkage cut into production, how many shipped system with old linkage. ESPF Check and update if needed |  |  | After | [7 empty] | 116076690 | Closed | 265518 | TRUE | Closed - Duplicate | Closed - Duplicate | Duplicate (50) | CTUDT00631671 |  |  |  | Robin | Failure to scan | Engineering |  |  |  | Closed |  |  | Linkage | [8 empty] | System issue | system | Couch-HW | Horizontal linkage | Couch-HW-Horizontal linkage | Horizontal linkage | CTUDT00631671 | Horizontal linkage |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
72 | 44089 | TRUE | Turkey | T.C Saglik Bakanligi Bagcilar | 40010 | 44089 | 38 | Dicom files are being waited on queue after some time. | dicom files are being waited on queue after some time. It seems there is a dicom send issue on the system. Aftr reboot and/or reset spooler via service it starts to send dicom files to external nodes again. Need to solve tis issue asap because system is being used heavily and customer does not want to do restart every in an hours time. | Yes |  | Need bugreport, timestamp of the issue |  | TBD |  |  | After | [7 empty] | CFM |  | 279732 | TRUE | Closed - Done | Closed - Done | Duplicate (23) | Not a Defect |  |  |  | Silva | PACS | Engineering |  |  |  | Closed |  |  | PACS | PACS | [7 empty] | System issue | System | Console-SW | Remote Connection-Not a Defect | Console-SW-Remote Connection-Not a Defect | Console-SW-Remote Connection | Not a Defect | Remote Connection | Console-SW | Console-SW-Remote Connection-Not a Defect | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
73 | 44075 | TRUE | us | The Regional Medical Center | 40005 | 44075 | 36 | Cross hair | Customer wanted a long cross hair on the axials. Explained that does not exist on 4.1.7 | P5 |  | Does 4.1.10 has a long cross hair on the axials |  | TBD |  |  | After | [7 empty] | 116053427 |  | 256112 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | Valid | CTUDT00631636 |  |  |  | Silva | Applications/Usability | Engineering |  |  |  | Closed |  |  | RFC | [8 empty] | System issue | system | Console-SW | RFC | Console-SW-RFC | Console-SW-RFC | CTUDT00631636 | RFC |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
74 | 44067 | TRUE | US | Willis Knighton Health System | 40004 | 44067 | 35 | Missing Images - Horz Couch Errors | Operator error. Did not move couch forward ULV shows HMC and VMC collision errors - and other errors. | P3 |  | Field handling Operator error. Did not move couch forward |  | TBD |  |  | After | [7 empty] | 116006767 |  | 243770 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | Valid | Not a Defect |  |  |  | Robin | Operator error | Customer |  |  |  | Closed |  | N/A | Operator error | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
75 | 44088 | TRUE | Germany | Gesundheit Nord gGmbh | 40002 | 44088 | 38 | Connect hand switch | Connect hand switch The hand switch at the end of the table still has to be electrically connected about Mr. Halama / PH Appointment 1 man / 3 hours Appointment with Ms. Wagner for the 28.09. 2 p.m. agreed. | P3 |  | Field handling |  |  |  |  | After | [7 empty] | 116109645 |  | 279252 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | Valid | Not a Defect |  |  |  | TBD | Applications/Usability | Service |  |  |  | Closed |  |  | Cabling | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
76 | #N/A | #N/A | Germany | Gesundheit Nord gGmbh | 40002 | 44060 | 34 | Joystick Intv Couch is not delivered with upgrade - missing | 4598 013 19771 - Kit Joystick Intv Couch is not delivered with upgrade - missing Item blocked at SPS site : Block: SPS BPO Planning: Block applicable for: Temporary Expiration date: 9-Jan-2021 Remove D-chain <(>&<)> SLI to SSD on the moment the part has stock. What? There is an order protection block in place. This is a new material which is not yet ready for ordering. Why? We need to complete setup and sourcing of this material before the block can be released. How? Market Parts Orderdesk: Contact ASAP to request block lift. SPS ASAP Instruction: This part has a Dchain 51. If there’s stock available ASAP is allowed to raise a service request to the MDM team to unblock the Dchain temporarily. | P5 |  | DEFOA for 4598 013 19771 - Kit Joystick Intv Couch Commerial DEFOA CD-30513 per order 1105795 bom, 4598-015-86581 KIT- iCT/ING INTERV CNTL, EXT WHT Z2N is included, which include the 4598 013 19771 - Kit Joystick Intv Couch and other parts, did the interventional control kit shipped to customer? |  | Operation investigation |  |  | During | [7 empty] | 115745150 |  | DEFOA CD-30513 | #N/A | #N/A | DEFOA | Valid | Not a defect |  |  |  | David | Missing Item | Operation |  |  |  | WIP |  |  | DEFOA | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
77 | #N/A | #N/A | US | Mount Grant General Hospital | 40012 | 44063 | 34 | part order for Bruce Schuler | Don Parr and I are working on this install. I have tried to order the parts through the DFOA app but the app is asking for an 8 digit SWO# that is no longer valid due to ServiceMax. I need two parts that did not arrive with the system. Both of them are for the ECG interface to the CT being installed. These parts are needed tomorrow to continue the installation on schedule. The install is scheduled to be completed by Friday as Application training is starting on Monday. The parts are as follows: USB extender cable 455300027201 CRC USB Cable 459801234171 | P5 |  | Field handling, DEFOA CD-0000030584 Check with Opertaion for the missing items which is not only case? |  | Operation investigation |  |  | During | [7 empty] | 115949349 |  | DEFOA CD-30584 | #N/A | #N/A | DEFOA | Valid | Not a defect |  |  |  | David | Missing Item | Operation |  |  |  | WIP |  |  | DEFOA | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
78 | 44092 | TRUE | China | Dalian Branch of Shengjing Hospital of China Medical University | 345001 | 44092 | 38 | 02N/Ingenuity Core128/Elite China/A computer crash | 02N/Ingenuity Core128/Elite China/A computer crash 服务内容：计算机死机，长时间未关机，重启计算机。 服务结果：维修后，设备恢复正常工作。 | P3 |  | Need bugreport, timestamp of the issue |  |  |  |  | After | [7 empty] | 116138079 |  | 287934 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | Valid | CTUDT00631367 |  |  |  | Silva | Lockup or Crash | Engineering |  |  |  | Closed |  |  | Yellow Screen | [8 empty] | System issue | System | Console-SW | WF-Yellow Screen | Console-SW-WF-Yellow Screen | Console-SW-Yellow-Screen | CTUDT00631367 | Yellow Screen | Console-SW | Console-SW-WF-Yellow Screen | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
79 | 44092 | TRUE | United States | McLaren Macomb | 40021 | 44092 | 38 | AT - 143465 | New CT floor leveling $5100.00 FSE- Tony Mona called in to order Part# 453567972921 ( qty 4) Ship to Site Attn: Jim Dord/ED NEW CTC Installations For Standard Delivery. I submitted the part order for FSE. | P5 |  | pre installation_ Self leveling floor ready for installation |  |  |  |  | During | [7 empty] | 116141136 |  | 288529 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | Valid | Not a Defect |  |  |  | Brook | Installation Other | Service |  |  |  | Closed | Shipped Incomplete |  | Service Parts Order | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
80 | 44088 | TRUE | China | SD Laiyang People Hospital | 345005 | 44088 | 38 | 02N/Core128/Can not print | 02N/Core128/Can not print Original Problem Reported by Customer Can not print Problem Description system can not film Troubleshooting Action start spooler Repair Action start spooler | P5 |  | Need bugreport, timestamp of the issue |  |  |  |  | After | [7 empty] | 116108800 | Closed | 279232 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | Valid | CTUDT00632127 |  |  |  | Silva | LanConfig/DICOM | Engineering |  |  |  | Closed |  |  | PACS | PACS | [7 empty] | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
81 | 44088 | TRUE | Spain | S.C.S. HOSP. GRAL. DE FUERTEVE | 40020 | 44088 | 38 | Clinical applications. | Aplicaciones clínicas. | P5 |  | Application Q&A |  |  |  |  | After | [7 empty] | 116115136 | Closed | 279980 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | Valid | Not a Defect |  |  |  | TBD | Applications/Usability | Application |  |  |  | Closed | Positive Customer Feedback |  | Q&A | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
82 | #N/A | #N/A | China | Senhai Hospital of Harbin | 345007 | 44097 | 39 | Eye phantom holder | Missing handle for legacy phantom | P5 |  | DEFOA for the holder |  | ECR to correct the BOM |  |  | During | [7 empty] | 116119982 |  | DEFOA CD-31191 | #N/A | #N/A | DEFOA |  | CTUDT00631640 |  |  |  | TBD | Missing Item | Operation |  |  |  | Closed |  |  | DEFOA | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
83 | 44096 | TRUE | China | Dalian Branch of Shengjing Hospital of China Medical University | 345001 | 44096 | 39 | yellow screen (buttons not working | occurred about once per 3 days, host has to be restarted ipv6 unchecked for the second IP for PRS router | P4 |  | under investigation, defect to be raised |  |  |  | https://www.clouddrive.philips.com.cn/file/i7ju2w01gvwogpvi026iij0jsozonr6d# 密码:123 | After | [7 empty] | 116151345 |  | 293042 | TRUE | Closed - Duplicate | Closed - Duplicate | Duplicate (84) | CTUDT00631367 |  |  |  | Silva | Lockup or Crash | Engineering |  |  |  | Closed |  |  | Yellow Screen | Yellow screen | [7 empty] | System issue | system | Console-SW | WF-Yellow Screen | Console-SW-WF-Yellow Screen | Yellow Screen | CTUDT00631367 | Yellow Screen |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
84 | 44098 | TRUE | China | Dalian Branch of Shengjing Hospital of China Medical University | 345001 | 44098 | 39 | yellow screen (buttons not working | occurred about once per 3 days, host has to be restarted | P4 |  | under investigation, defect to be raised |  |  |  |  | After | [7 empty] | CFM |  | 295524 | TRUE | Closed - Done | Closed - Done | Valid | CTUDT00631367 |  |  |  | Silva | Lockup or Crash | Engineering |  |  |  | Closed |  |  | Yellow Screen | Yellow screen | [7 empty] | System issue | System | Console-SW | WF-Yellow Screen | Console-SW-WF-Yellow Screen | Console-SW-Yellow-Screen | CTUDT00631367 | Yellow Screen | Console-SW | Console-SW-WF-Yellow Screen | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
85 | 44094 | TRUE | China | Wuhan Ren _ai Hospital, Hubei Province | 345004 | 44094 | 38 | 02N/Ingenuity Core128/Elite China/Unable to log in | BU Q&R Feedback: No Reason: Corrective Maintenance 无法登陆软件，重启后恢复。 服务结果：维修后，设备恢复正常工作。 Problem Reported by Customer : Unable to log in Next Step Action: Xu Junbiao 15572725806 call back Safety Questions:Case: 0116144388 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Qi Xiaoyu Last updated by: Qi Xiaoyu | P3 |  | more information needed, do we have bugreport 徐均标 15572725806 call back |  |  |  |  | After | [7 empty] | 116144388 | Closed | 289233 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD_Bugreport |  |  |  | Silva | Lockup or Crash | Engineering |  |  |  | WIP |  |  | Reboot | [8 empty] | System issue | System | Console-SW | Login | Console-SW-Login | Console-SW-Login | TBD_Bugreport | Login | Console-SW | Console-SW-Login | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
86 | 44095 | TRUE | USA | Willis Knighton Health System | 40004 | 44095 | 39 | 0306CDFM2 laser turns off the gantry and creates an estop | Work Order #: WO-02062198 Atz, Michael The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Customer Function/Role: BIO MED Clint system rebooted to the wall and the FSE Spence Williams is aware please dispatch to him for on-site service 020/09/21 15:44:26.629 S_MP_MPSHELL_SERIAL_LIFE_SYNC_ERROR 2020/09/21 15:36:24.434 S_ROTOR_MCO ---- ERROR_STRING S_MC_DEVICE_PS_FAILURE --- 2020/09/21 15:28:04.559__01 S_MPL_VX_SEND_ERROR Explanation Fatal error causes a controller reset. vxworks msgq send error. The tasks queue is filled up. Indicates the task is not getting enough running time to pick up its MPL messages in real time. Maybe due to excessive CAN traffic. Action Try a gantry reboot. Explanation Couch power supply failure. Device Use: Diagnostic Expected and Unexpected Results: scan and view User Impact: unable to scan Patient Impact: wait delay Current Software Version: 4.1.10 Diagnostic performed by Engineer : Customer Function/Role: BIO MED Clint system rebooted to the wall and the FSE Spence Williams is aware please dispatch to him for on-site service 020/09/21 15:44:26.629 S_MP_MPSHELL_SERIAL_LIFE_SYNC_ERROR 2020/09/21 15:36:24.434 S_ROTOR_MCO ---- ERROR_STRING S_MC_DEVICE_PS_FAILURE --- 2020/09/21 15:28:04.559__01 S_MPL_VX_SEND_ERROR Explanation Fatal error causes a controller reset. vxworks msgq send error. The tasks queue is filled up. Indicates the task is not getting enough running time to pick up its MPL messages in real time. Maybe due to excessive CAN traffic. Action Try a gantry reboot. Explanation Couch power supply failure. Device Use: Diagnostic Expected and Unexpected Results: scan and view User Impact: unable to scan Patient Impact: wait delay Current Software Version: 4.1.10 Follow Up Required : Customer Function/Role: BIO MED Clint system rebooted to the wall and the FSE Spence Williams is aware please dispatch to him for on-site service 020/09/21 15:44:26.629 S_MP_MPSHELL_SERIAL_LIFE_SYNC_ERROR 2020/09/21 15:36:24.434 S_ROTOR_MCO ---- ERROR_STRING S_MC_DEVICE_PS_FAILURE --- 2020/09/21 15:28:04.559__01 S_MPL_VX_SEND_ERROR Explanation Fatal error causes a controller reset. vxworks msgq send error. The tasks queue is filled up. Indicates the task is not getting enough running time to pick up its MPL messages in real time. Maybe due to excessive CAN traffic. Action Try a gantry reboot. Explanation Couch power supply failure. Device Use: Diagnostic Expected and Unexpected Results: scan and view User Impact: unable to scan Patient Impact: wait delay Current Software Version: 4.1.10 Check all connections in the table and power supply -- since the flouro option was recently added ----------------------------------------------------------------------------- Work Order #: WO-02062285 22-Sep-2020 Williams Spence The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: BIO MED Clint system rebooted to the wall and the FSE Spence Williams is aware please dispatch to him for on-site service 020/09/21 15:44:26.629 S_MP_MPSHELL_SERIAL_LIFE_SYNC_ERROR 2020/09/21 15:36:24.434 S_ROTOR_MCO ---- ERROR_STRING S_MC_DEVICE_PS_FAILURE --- 2020/09/21 15:28:04.559__01 S_MPL_VX_SEND_ERROR Explanation Fatal error causes a controller reset. vxworks msgq send error. The tasks queue is filled up. Indicates the task is not getting enough running time to pick up its MPL messages in real time. Maybe due to excessive CAN traffic. Action Try a gantry reboot. Explanation Couch power supply failure. Device Use: Diagnostic Expected and Unexpected Results: scan and view User Impact: unable to scan Patient Impact: wait delay Current Software Version: 4.1.10 Follow Up Required : need couch motor/ clutch assy 453566494823 ship to UPS Hold 2000819 10:30am Resolution : Order couch clutch / motor assy 453566494823 UPS hold 2000819 for 10:30am 9/22/2020 ----------------------------------------------------------------------------- | P2 |  | False alarm with these cases. Bad contact in front panel. No parts were replaced. |  |  |  |  | After | [7 empty] | 116150622 |  | 290534 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00631634 |  |  |  | Robin | Failure to scan | Engineering |  |  |  | Closed |  |  | False alarm | [8 empty] | System issue | System | Gantry-HW | gantry laser | Gantry-HW-gantry laser | Gantry-HW-GantryLaser | CTUDT00631634 | gantry laser | Gantry-HW | Gantry-HW-gantry laser | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
87 | 44096 | TRUE | China | Wuhan Ren _ai Hospital, Hubei Province | 345004 | 44096 | 39 | 02N/Ingenuity Core128/Elite China/Tube warm up failed | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer:Tube warm-up failed Safety Questions:Case: 0116153281 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingli Last updated by: Lu Mingli 无法预热，重启主机重建机及机架后恢复。 服务结果：维修后，设备恢复正常工作。 | P3 |  | more information needed, do we have bugreport 徐均标 15572725806 call back | [12 empty] | 116153281 | Closed | 293646 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD_Bugreport |  |  |  | Silva | Failure to scan | Engineering |  |  |  | WIP |  |  | Reboot | [8 empty] | System issue | System | Gantry-SW | Other | Gantry-SW-Other | Gantry-SW-Other | TBD_Bugreport | Other | Gantry-SW | Gantry-SW-Other | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
88 | 44096 | TRUE | USA | Willis Knighton Health System | 40004 | 44096 | 39 | continue troubleshooting and install couch motor | Work Order #: WO-02066876 Williams Spence Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: | P1 |  | False alarm with these cases. Bad contact in front panel. No parts were replaced. | [12 empty] | 116150994 |  | 294287 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not a Defect |  |  |  | Robin | Failure to scan | Engineering |  |  |  | Closed |  |  | False alarm | [8 empty] | System issue | System | Gantry-HW | Control Panel | Gantry-HW-Control Panel | Gantry-HW-Control Panel | Not a Defect | Control Panel | Gantry-HW | Gantry-HW-Control Panel | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
89 | 44097 | TRUE | USA | Seton Medical Center Harker He | 40006 | 44097 | 39 | having issues logging in and the gantry is not turning on | Work Order #: WO-02082414 Boyd Craig The system does not meet specification and is out of use Follow Up Required : FSE on site to establish log on to system and show customer. Diagnostic performed by Engineer : At log in tech enters ct and message incorrect password or user name is displayed. Unable to log in as Emergency as well. Work Order #: WO-02082964 24-Sep-2020 Beato Tom The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: Unable to log on to system User Impact: Patient Impact: Current Software Version: Resolution : Entered the initial user password for system. Attempted to changed the password to blank for user, but appears windows 10 does not allow that option for the CT system. Configured new user password for tech. Tested system and downloaded exam cards for user. ----------------------------------------------------------------------------FSE got a call from Physicist asking where the SMPTE pattern was located. Apparently the DEMO folder was removed in version 4.1.10 software. Had to research this to verify but on inspection using remote connection, no DEMO folder was found. Verified this with L3. Provided documentation and SMPTE pattern to be installed. | P2 |  | Customer forget password. System working after reset the password. Another 2 sites from China report similar cases about the password forget. |  | Follow up with Caroline to define actions to improve the user instruction and training material and highlight the key changes between win7 and win10. -- Kevin Access control in IFU 459801656011 Clinical Training materal https://www.learningconnection.philips.com/en/ct-user-access-controls Service Training Course | [10 empty] | 116184131 |  | 294608 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not a Defect |  |  |  | Kevin | User Experience | Application |  |  |  | Closed |  |  | User Account | User Password | [7 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
90 | 44096 | TRUE | China | SD Laiyang People Hospital | 345005 | 44096 | 39 | 02N/Ingenuity Core128/Elite China/can not warm up | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer:can not warm up Safety Questions:Case: 0116151056 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Chen YuanHui Last updated by: Chen YuanHui 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer_s problems: Scanning failed. Detailed description of Customer_s problems or malfunctions: Scanning failed. It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: N/A Remote troubleshoot and test result: FTP and TELNET are faulty, unable to download PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Onsite check error log SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support External Remarks: 22.09.2020 07:24:28 UTC Xingbo Yang (86038250) service details: Reset e-stop button. service result: device returns to normal working condition after service Internal Remarks: 22.09.2020 07:24:31 UTC Xingbo Yang (86038250) "RDI Accuracy:yes" $Q7: (Confirmed reported problem & failed to finish system initialization) $Q8: ( review error logger and check e-stop loop path & find emergency stop button pressed) $Q9: (Reset e-stop button) $Q10: (Test and inspection data refer to: PA & Device returned to full functionality) $Q11: Calibrated tools name(NA) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date: 9/ 10 / 2020 ^Current Tube usage: scan second_2039 ----------------------------------------------------------------------------- | P3 |  | ( review error logger and check e-stop loop path & fund emergency stop button pressed) $Q9: ( reset e-stop button ) | [12 empty] | 116151056 |  | 294998 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not a Defect |  |  |  | Silva | Failure to scan | Customer |  |  |  | Closed |  |  | Operator error | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
91 | 44097 | TRUE | China | JL Taonan T.C.M. Hosp. | 345006 | 44097 | 39 | 02N/Ingenuity Core128/Elite China/High Voltage sparking | BU Q&R Feedback: No SAP Customer Symptom Code: BOOT Customer Symptom Code: Boot up/Power-up Problem Reason: Corrective Maintenance Engineer Failure Code: Configuration Error Repair Activity Code: Configuration Change FSE Notes : This PO is a Reference Number only Problem Reported by Customer : High voltage sparking Safety Questions : Case: 0116175144 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Yuan Mingze Last updated by: Yuan Mingze N/A T2 Activities : Problem Description: tube arcs Troubleshooting Action: tube arcs Repair Action: ask customer to perform tube condition SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | P3 |  | tube arcs, ask customer to perform tube condition | [12 empty] | 116175144 |  | 295028 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00630085 |  |  |  | Silva | Calibration/Performance | Engineering |  |  |  | Closed |  |  | Tube arc | Tube Arc | [7 empty] | System issue | System | IC-HW | Tub Arc | IC-HW-Tub Arc | IC-HW-Tub Arc | CTUDT00630085 | Tub Arc | IC-HW | IC-HW-Tub Arc | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
92 | 44097 | TRUE | China | General Hospital of Jizhong Energy Jingxing Mining Group Co. LTD | 345003 | 44097 | 39 | 02N/Ingenuity Core128/Elite China/The network transmission does not match the hospital system | T2 Activities : Problem Description: 客户问题概述： 医生报修，有些扫描参数照片，PACS不能接受 客户问题详细描述或故障现象： 具体问题，医生描述不清 设备本身技术问题： 是。 Troubleshooting Action: 远程诊断与测试结果： 需要工程师到场，联系PACS公司人员共同解决问题 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议教客户使用飞利浦微信报修 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 25.09.2020 05:56:05 UTC Litu Yang (86038906)服务内容： [传输问题，修改配置恢复。] 服务结果： [设备正常工作。] Internal Remarks : 25.09.2020 06:02:43 UTC Litu Yang (86038906) $Q7: (Confirmed the reported problem. transfer issue) $Q8: (checked the dicom parameter，port was wrong，need to modify) $Q9: (readd the dicom port ) $Q10: (Test and inspection data refer to: PA. Device returned to full functionality.) $Q11: Calibrated Tools name (NA） For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date: 09/13/2020 ^Current Tube usage: scan second_9727 ----------------------------------------------------------------------------- | P5 |  | try to get more infor | [12 empty] | 116175398 |  | 295724 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not a Defect |  |  |  | Silva | PACS | Others |  |  |  | Closed |  |  | PACS Setup | PACS | [7 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
93 | 44098 | TRUE | Turkey | Basaksehir Sehir Hastanesi HITACHI HEALTHCARE TURKEY | 337000 | Console-SW | 39 | Date and time information match | Work Order #: WO-02085350 24-Sep-2020 Tunay Uysal The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A | P3 |  | need more information | [12 empty] | 116187066 | Closed | 297084 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not a Defect |  |  |  | TBD | User Experience | Service |  |  |  | Closed |  |  | Installation setting | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
94 | 44114 | TRUE | China | Jinggangshan Hospital, the First Affiliated Hospital of Nanchang University, Jiangxi | 345002 | 44114 | 41 | 02N/Ingenuity Core128/Elite China/connect PACS | BU Q&R Feedback: No SAP Customer Symptom Code: CENV Customer Symptom Code: Customer Environmental Concern Reason: Corrective Maintenance Engineer Failure Code: Environmental Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer : connect PACS Safety Questions : Case: 0116277286 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Lili Last updated by: Lu Lili 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A FSE Notes : This PO is a Reference Number only T2 Activities : Problem Description: $Q7: (Confirmed reported problem. & The device is ok,just need to add PACS.) Troubleshooting Action: $Q8: (The device is ok,just need to add PACS.No error,no troubleshooting. & The system is ok,just need to add PACS.) Repair Action: $Q9: (Guide the customer Add PACS.) Test/Inspection: $Q10: (Test and Inspection data: Let Hosp to functions test passed. & Device status: The Device had returned to fully functions.) Test Equipment: $Q11: Calibrated Tools:N/A ----------------------------------------------------------------------------- | P5 | [14 empty] | 116277286 |  | 329176 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | Opened - Pending Translation | Not a Defect |  |  |  | Silva | PACS | Service |  |  |  | Closed |  |  | PACS Setup | PACS Setup | [7 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
95 | 44113 | TRUE | China | General Hospital of Jizhong Energy Jingxing Mining Group Co. LTD | 345003 | 44113 | 41 | fast-moving patient table power off intermittently | \\cnhszhphc1ws300\Input\Silva\China FOK Bugreport\TW PR 324680 305910 fast-moving patient table power off intermittently |  | \\cnhszhphc1ws300\Input\Silva\China FOK Bugreport\TW PR 324680 305910 fast-moving patient table power off intermittently | Horizatal, couch movement calibration | [14 empty] | 324680 | TRUE | Closed - Done | Closed - Done | Opened | CTUDT00631669 |  |  |  | Changxian | E-Stop | Engineering | Open |  |  | Closed |  |  | Horizontal movement | [8 empty] | System issue | System | Couch-SW | Horizontal Move Block | Couch-SW-Horizontal Move Block | Couch-SW-Horizontal Move Block | CTUDT00631669 | Horizontal Move Block | Couch-SW | Couch-SW-Horizontal Move Block | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
96 | 44110 | TRUE | Spain | S.C.S. HOSP. GRAL. DE FUERTEVE | 40020 | 44110 | 41 | Aplicaciones remoto | Work Order #: WO-02145998 07-Oct-2020 Mario Vacas Jerez The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Regulatory Questions : N/A Diagnostic performed by Engineer : Se da soporte remoto a la espera de poder realizar la formación presencial. Durante el mismo se crean los protocolos de adquisición, se imparte una formación básica que permite la realización de estudios y se resuelven las posibles dudas surgidas. Resolution : Se da soporte remoto a la espera de poder realizar la formación presencial. Durante el mismo se crean los protocolos de adquisición, se imparte una formación básica que permite la realización de estudios y se resuelven las posibles dudas surgidas. | P5 |  | Application support | [12 empty] | 116245985 |  | 319684 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | Opened - Pending Translation | Not a Defect |  |  |  | Kevin | Applications/Usability | Application |  |  |  | Closed |  | Application | Q&A | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
97 | 44110 | TRUE | China | SD Laiyang People Hospital | 345005 | 44110 | 41 | 02N/Ingenuity Core128/Elite China/worklist fault | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer : worklist fault Next Step Action : hu haiyang call back Safety Questions : Case: 0116243434 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingli Last updated by: Lu Mingli 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action : FSE 胡海洋-18801767774 来电告知此单先关闭不派工，客户已经下班了，明天他会继续指导客户，有需要的会再致电800 ----------------------------------------------------------------------------- | P3 |  | worklist fault | [12 empty] | 116243434 |  | 319756 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | Opened - Pending Translation | CTUDT00631283 |  |  |  | Silva | PACS | Engineering | Open |  |  | Closed |  |  | PACS | PACS | [7 empty] | System issue | System | Console-SW | Remote Connection-Worklist | Console-SW-Remote Connection-Worklist | Console-SW-Remote Connection | CTUDT00631283 | Remote Connection | Console-SW | Console-SW-Remote Connection-Worklist | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
98 | 44110 | TRUE | Germany | Gesundheit Nord gGmbh | 40002 | 44110 | 41 | Beautification work on the table | Work Order #: WO-02144559 Halama Thomas Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | P5 |  | On October 30th, Mr. Haller will be on duty and you will be asked to take the constancy test on this day | [12 empty] | 116245421 |  | 319878 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | Opened - Pending Translation | Not A Defect |  |  |  | TBD | Investigating | Others |  |  |  | WIP |  | Investigating | Painting | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
99 | 44106 | TRUE | United States | St Joseph Westgate Medical Center | 40001 | 44106 | 40 | Physicist testing | Work Order #: WO-02133491 03-Oct-2020 Leon Peter Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A | P5 |  | SMPTE from monitor calibration | [12 empty] | 116234472 |  | 315543 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | Opened | TBC |  |  |  | TBD | User Experience | Others |  |  |  | Closed |  |  | Refresh Diff | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
100 | 44106 | TRUE | United States | St Joseph Westgate Medical Center | 40001 | 44106 | 40 | IT network set up | Work Order #: WO-02133502 Leon Peter Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | P5 |  | PACS setup | [12 empty] | 116234503 |  | 315544 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | Evaluation | Not a Defect |  |  |  | Silva | PACS | Service | Open |  |  | Closed |  |  | PACS Setup | PACS Setup | [7 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
101 | 44105 | TRUE | Turkey | T.C Saglik Bakanligi Bagcilar -> Bağcılar State Hospital | 40010 | 44105 | 40 | It freezes and the patient list does not appear. | Work Order #: WO-02121972 Serkan Yavaşça Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: The user reported some SW errors. Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | P3 |  | Need bugreport, timestamp of the issue PACS tansfer issue | [12 empty] | 116224696 |  | 308737 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | Opened | CTUDT00631283 |  |  |  | Silva | PACS | Engineering | Open |  |  | Closed |  |  | Worklist | PACS | [7 empty] | System issue | System | Console-SW | Remote Connection-Worklist | Console-SW-Remote Connection-Worklist | Console-SW-Remote Connection | CTUDT00631283 | Remote Connection | Console-SW | Console-SW-Remote Connection-Worklist | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
102 | 44105 | TRUE | Spain | S.C.S. HOSP. GRAL. DE FUERTEVE | 40020 | 44105 | 40 | Table locked. | Work Order #: WO-02122797 01-Oct-2020 José Carlos Pulido Acosta Problem Description by Engineer: Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Regulatory Questions : N/A Diagnostic performed by Engineer: N/A ----------------------------------------------------------------------------- | P2 | \\cnhszhphc1ws300\Input\Silva\China FOK Bugreport\TW PR 311441 | Parts Consumed Description: 1.00000,459801485251,RING-LEXAN GN5 Parts Consumed Description: 1.00000,459800623445,FRU, AMC MOTOR CONTROL MODULE I do a short resume for you: José Carlos went to do the FCO, and he implemented it, after this started the problems, then he asked to the customer to see if they had problems before his arrived, then they told him a strange story about that someone hit the central cover of the scanner and damaged the cover, and after check, reinstall software, changes parts and others finally the problem was solved by the changes of the AMC. |  | Changxian to investigate | [10 empty] | 116225664 | In Process | 311441 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | Opened | Not a Defect |  |  |  | Changxian | Oprator error | Customer |  |  |  | Closed |  |  | Operator error | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
103 | 44105 | TRUE | United States | Seton Medical Center Harker He | 40006 | 44105 | 40 | Phantom not seen message during IQ check | Phantom not seen message during IQ check. This occurred twice during the week. Once on Oct 1, 2020 and once on Oct 2, 2020 two bug reports have been created. Will send once tracking number is created |  | \\cnhszhphc1ws300\Input\Silva\China FOK Bugreport\TW PR 315698 | Tammy support to do investigation |  |  |  |  | After | [9 empty] | 315698 | TRUE | Closed - Done | Closed - Done | Investigation | CTUDT00631545 |  |  |  | Tammy/Silva | Calibration/Performance | Engineering |  |  |  | Closed |  |  | Phantom not detected | [8 empty] | System issue | System | DMS-SW | Phantom Detect | DMS-SW-Phantom Detect | DMS-SW-Phantom Detect | CTUDT00631545 | Phantom Detect | DMS-SW | DMS-SW-Phantom Detect | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
104 | 44104 | TRUE | France | SCP GROUPE D IMAGERIE MEDICALE | 40019 | 44104 | 40 | Transfer stuck once a day to PACS and need to restart Host |  |  | \\cnhszhphc1ws300\Input\Silva\China FOK Bugreport\TW PR 307590 309138 Transfer stuck once a day to PACS and need to restart Host | SW Team investigation | [14 empty] | 307590 | TRUE | Closed - Done | Closed - Done | Investigation | CTUDT00632114 |  |  |  | Tammy/Silva | PACS | Engineering |  |  |  | Closed |  |  | PACS | PACS | [7 empty] | System issue | System | Console-SW | Remote Connection-Broken | Console-SW-Remote Connection-Broken | Console-SW-Remote Connection | CTUDT00632114 | Remote Connection | Console-SW | Console-SW-Remote Connection-Broken | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
105 | 44104 | TRUE | France | SCP GROUPE D IMAGERIE MEDICALE | 40019 | 44104 | 40 | Problem with image transfer | Work Order #: WO-02118605 Aurelien ZAJAC The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: user Device Use: yes Expected and Unexpected Results: n/a User Impact: no Patient Impact: no Current Software Version: Regulatory Questions : N/A Diagnostic performed by Engineer : Le client signale des problèmes de transfert. Effacement des fichiers batch en remote et redémarrage du host. Resolution : Le scanner est de nouveau pleinement opérationnel. | P3 | \\cnhszhphc1ws300\Input\Silva\China FOK Bugreport\TW PR 307590 309138 Transfer stuck once a day to PACS and need to restart Host | SW Team investigation | [12 empty] | 116221699 |  | 309138 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632114 |  |  |  | Tammy/Silva | PACS | Engineering |  |  |  | Closed |  |  | PACS | PACS | [7 empty] | System issue | System | Console-SW | Remote Connection-Broken | Console-SW-Remote Connection-Broken | Console-SW-Remote Connection | CTUDT00632114 | Remote Connection | Console-SW | Console-SW-Remote Connection-Broken | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
106 | 44104 | TRUE | Puerto Rico | METRO MAYAGUEZ INC | 40009 | 44104 | 40 | Missing remote connection Icon | Work Order #: WO-02118637 Zayas, Amilcar Work Order #: WO-02118644 Hernandez, Alvin Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A Integration : A related Work Order for this Case and Installed Product is pending integration with SAP. This Work Order will not be sent to SAP until the pending integration on the related Work Order is complete | P5 |  | if DOD config tool run, need reinstall software to fix the connection |  | Defect linked | [10 empty] | 116221725 | Closed | 309970 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | Opened | CTUDT00631143 |  |  |  | Silva | Remote Service Tool | Engineering |  |  |  | Closed |  | Remote Service Tool | DoD Config Tool | [8 empty] | System issue | System | Console-SW | Service tool | Console-SW-Service tool | Console-SW-Service tool | CTUDT00631143 | Service tool | Console-SW | Console-SW-Service tool | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
107 | 44103 | TRUE | United States | Seton Medical Center Harker He | 40006 | 44103 | 40 | 0711CDFK1 needs iDose enabled | Work Order #: WO-02110693 01-Oct-2020 Beato Tom The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : Checked license file and determined idose was not listed as a permanent entitlement. Downloaded the proper license file and installed on system. Verified idose installed properly. ----------------------------------------------------------------------------- Work Order #: WO-02110701 Carpenter, Kevin The system partly meets specification and is returned to use with restrictions Follow Up Required : Check that SWare Option Key includes iDose feature and run a test iDose scan while onsite. Refer customer to CT Clinical Apps if needed. Diagnostic performed by Engineer : iDose not working ~ SWare Option Key may be needed, or CT Clin Apps. SWVer = 4.1.10.25202 Problem Description by Engineer : Customer Function/Role: Rebecca Device Use: Diag Expected and Unexpected Results: iDose not working ~ SWare Option Key may be needed, or CT Clin Apps. User Impact: None Patient Impact: Cant scan Pts with iDose feature Current Software Version: 4.1.10.25202 ----------------------------------------------------------------------------- | P2 |  | Field not download the correct license | [12 empty] | 116213363 | Closed | 304928 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | Evaluation | Not A Defect |  |  |  | N/A | License | Service |  |  |  | Closed |  | Not a defect | Installation license | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
108 | 44103 | TRUE | China | General Hospital of Jizhong Energy Jingxing Mining Group Co. LTD | 345003 | 44103 | 40 | WeChat:02N/Ingenuity Core128/Elite China/The rack frequently loses power | ( Gantry power down，fast moving table occured，maybe accidental，need reboot gantry) 客户问题概述： 床运动时频繁掉电 客户问题详细描述或故障现象： 床运动时频繁掉电 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 床运动时频繁掉电 远程诊断与测试结果： 无法下载日志，FSE怀疑4.1.10软件Bug PRS TR状态： 不通，请工程师到场检查。 Repair Action: 建议现场维修步骤： 检查修复PRS，联系TSE查询bug详情，如必要Escalation。 建议备件... | P4 | \\cnhszhphc1ws300\Input\Silva\China FOK Bugreport\TW PR 324680 305910 fast-moving patient table power off intermittently | under defect investigation | [12 empty] | 116209040 | Closed | 305910 | TRUE | Closed - Done | Closed - Done | Opened | CTUDT00631669 |  |  |  | Changxian | E-Stop | Engineering |  |  |  | Closed |  |  | Horizontal movement | [8 empty] | System issue | System | Couch-SW | Horizontal Move Block | Couch-SW-Horizontal Move Block | Couch-SW-Horizontal Move Block | CTUDT00631669 | Horizontal Move Block | Couch-SW | Couch-SW-Horizontal Move Block | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
109 | 44103 | TRUE | France | SCP GROUPE D IMAGERIE MEDICALE | 40019 | 44103 | 40 | Noise during rotation. | Work Order #: WO-02109691 30-Sep-2020 Jean-luc CODRON The system fully meets specification and returned to use Regulatory Questions : N/A Diagnostic performed by Engineer : réglage courroie et tendeur rotor Problem Description by Engineer : bruit en rotation | P3 |  | rotating noise, belt and rotor tensioner adjustment, check if have more detail from field about the issue and action to fix the issue. | [12 empty] | 116212670 | Closed | 307252 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | Closed - No Further Investigation | ECR-075975 |  |  |  | William | Noise | Engineering |  |  |  | Closed |  | Noise | Gantry Noise | [8 empty] | System issue | System | Gantry-HW | Lexan Ring | Gantry-HW-Lexan Ring | Gantry-HW-Lexan Ring | ECR-075975 | Lexan Ring | Gantry-HW | Gantry-HW-Lexan Ring | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
110 | 44103 | TRUE | United States | Seton Medical Center Harker He | 40006 | 44103 | 40 | table which can trap fluids and dirt. | Couch pad arrived with extreme wrinkles in the Velcro. This causes gaps between the velcro and the table which can trap fluids and dirt. When the customer tried to straighten out the wrinkles after removing the clear film from the adhesive, they attempted to place this back on so the velcro could be separated. Then the clear film could not be removed a 2nd time. Service will need to replace the Velcro. Two pictures are attached. |  |  | Ada work with SQ to do investigation | [14 empty] | 315368 | TRUE | Closed - Done | Closed - Done | Investigation | CTUDT00632175 |  |  |  | Ada | User Experience | Engineering |  |  |  | Closed |  |  | Couch Pad | [8 empty] | System issue | System | Couch-HW | Mattress | Couch-HW-Mattress | Couch-HW-Mattress | CTUDT00632175 | Mattress | Couch-HW | Couch-HW-Mattress | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
111 | 44102 | TRUE | Turkey | T.C Saglik Bakanligi Bagcilar -> Bağcılar State Hospital | 40010 | 44102 | 40 | They are asking the number of tube examinations | Work Order #: WO-02099825 Soner Saglam The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Tube kullanım değerleri istendi. Patient Impact: Current Software Version: Resolution : İstenilen değerler gönderildi. | P3 |  | Closed - Non-Complaint | [12 empty] | 116202093 |  | 300117 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | Closed - Non-Complaint | Not A Defect |  |  |  | N/A | Applications/Usability | Application |  |  |  | Closed |  |  | Q&A | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
112 | 44119 | TRUE | USA | Seton Medical Center Harker He | 40006 | 44119 | 42 | Trying to get images to PAC_s, will not move images. Only icon that is an option is to "remove images" | Work Order #: WO-02218745 Miller, John M The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role:Technician/Sonographer/Therapist Device Use: Diagnostic Expected and Unexpected Results: can not archive to Pacs1 User Impact: None Patient Impact: None Current Software Version: Diagnostic performed by Engineer : Customer states they images are not transferring to Pacs1. Tested network with Ping to the Pacs1 and the test passed. Instructed customer to restart the Host PC. After restart the images are now transferring to Pacs1. The customer preferred to callback if needed rather than a follow up call. ----------------------------------------------------------------------------- | P3 |  | collect bugreport | [12 empty] | 116303710 | Incomplete | 339755 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering |  |  |  | Closed |  |  | PACS | PACS | [7 empty] | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
113 | 44117 | TRUE | USA | McLaren Macomb | 40021 | 44117 | 42 | IEC cardiac leads |  |  |  | Send SO to Xiehong HFA investigate | [14 empty] | 342373 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not A Defect | [5 empty] | Operation |  |  |  | WIP |  |  | ECG Leads | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
114 | 44123 | TRUE | Turkey | HITACHI HEALTHCARE TURKEY | 337000 | 44123 | 43 | Hastalar databasede gözükmüyor Patients are not visible in database | Work Order #: WO-02231941 Problem Description by Engineer: Database problem Resolution: Although the patient database is empty in the device, the hard disk seems full. Hard disk control and repair procedures were performed in the system, and patients who did not appear after this procedure came back. It has been observed that these patients are deleted from the system together with the user and free up memory. The device is active. | P3 |  | collect bugreport, does system restart at least once per week? | [12 empty] | 116318132 | Fixed | 346419 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632261 | [5 empty] | Engineering |  |  |  | Closed |  |  | Disk full | [8 empty] | System issue | System | Console-SW | DiskFull | Console-SW-DiskFull | Console-SW-DiskFull | CTUDT00632261 | DiskFull | Console-SW | Console-SW-DiskFull | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
115 | 44120 | TRUE | Turkey | HITACHI HEALTHCARE TURKEY | 337000 | 44120 | 42 | Çekim sırasında hata veriyor | Work Order #: WO-02222895 16-Oct-2020 Can Özgenlik The system fully meets specification and returned to use Problem Description by Engineer : Error while shooting Resolution: In the examination made on the device, there were errors related to slipring. The slipring was cleaned and oiled. Trial shots were made with the Phantom. No problem at the moment, the device is active | P3 |  | No bugreport & picture, keep monitoring | [5 empty] | Gantry_HW | Slipring | Contamination on slipring | Clearing and oiling | Suggest use UV for disfication | William Yang |  | 116308161 | Closed | 342077 | TRUE | Closed - Done | Closed - Done |  | NotEngineering-SlipringClean | [5 empty] | Engineering |  |  |  | WIP |  |  | Slipring | [8 empty] | System issue | System | Gantry-HW | SlipRingClean | Gantry-HW-SlipRingClean | Gantry-HW-SlipRingClean | NotEngineering-SlipringClean | SlipRingClean | Gantry-HW | Gantry-HW-SlipRingClean | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
116 | 44118 | TRUE | Turkey | T.C Saglik Bakanligi Bagcilar -> Bağcılar State Hospital | 40010 | 44118 | 42 | X-ray error | Work Order #: WO-02209529 New Work Order Activity has been created with type: Resolution Description: XRay tube and HV tanks connectors and plugs were checked for Xray errors.HV silicones were mistaken.Long TC was made. The user must definitely do Short Tube conditioning every morning and when he / she is not using the device for more than 2 hours. If this is not done, there may be a problem with the Xray tube and a user error will occur. ----------------------------------------------------------------------------- | P3 |  | Tube Arc | [12 empty] | 116294975 | In progress | 338688 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00630085 | [5 empty] | Engineering |  |  |  | Closed |  |  | Tube Arc | Tube Arc | [7 empty] | System issue | System | IC-HW | Tub Arc | IC-HW-Tub Arc | IC-HW-Tub Arc | CTUDT00630085 | Tub Arc | IC-HW | IC-HW-Tub Arc | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
117 | 44120 | TRUE | Spain | S.C.S. HOSP. GRAL. DE FUERTEVE | 40020 | 44120 | 42 | Tiene un usuario bloqueado y no permite iniciar sessión TAC - Planta 1 ES6305749270 | New Work Order Activity has been created with type: Diagnostic performed by Engineer Description: I speak with Juan Pedro, the ct user has been blocked, they are working with the emergency user, I connect to the computer and change the password. | P3 |  | Password | [12 empty] | 116309560 | Closed | 342752 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not a defect | [5 empty] | Application |  |  |  | Closed |  |  | User Account | User Password | [7 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
118 | 44120 | TRUE | Spain | S.C.S. HOSP. GRAL. DE FUERTEVE | 40020 | 44120 | 42 | Artefacto de línea en el surview frontal. Se adjunta imagen. | New Work Order Activity has been created with type: Diagnostic performed by Engineer Description: We are informed of a slight artifact in the surview images. We connected remotely, seeing that the current images do not present an artifact that, most likely, was resolved by calibrating air. | P3 |  | The case is closed, since no artifacts are seen in the images displayed, surely corrected by the air calibrations. | [12 empty] | 116309987 | Closed | 344954 | TRUE | Closed - Done | Closed - Done |  | Not A Defect | [5 empty] | Engineering |  |  |  | Closed |  |  | Artifact | [8 empty] | System issue | System | DMS-SW | Artifact | DMS-SW-Artifact | DMS-SW-Artifact | Not A Defect | Artifact | DMS-SW | DMS-SW-Artifact | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
119 | 44119 | TRUE | Spain | S.C.S. HOSP. GRAL. DE FUERTEVE | 40020 | 44119 | 42 | Spanish keyboard installation | Parts Consumed Description: 1.00000,459801078761,Keyboard, USB, Spanish | P5 |  | Spanish keyboard installation | [12 empty] | 116304027 | Closed | 342885 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not A Defect | [5 empty] | Service |  |  |  | Closed |  |  | Keyboard | [6 empty] | 1.00000,459801078761,Keyboard, USB, Spanish |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
120 | 44119 | TRUE | Spain | S.C.S. HOSP. GRAL. DE FUERTEVE | 40020 | 44119 | 42 | Table covers make noise when going up. | Work Order #: WO-02219116 José Carlos Pulido Acosta Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A Regulatory Questions : N/A | P5 |  | No problem found | [5 empty] | Couch_HW | [6 empty] | 116304054 | In progress | 342890 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not A Defect | [5 empty] | Engineering |  |  |  | Closed |  |  | Table Noise | [8 empty] | System issue | System | Couch-HW | Cover | Couch-HW-Cover | Couch-HW-Cover | Not A Defect | Cover | Couch-HW | Couch-HW-Cover | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
121 | 44117 | TRUE | Spain | S.C.S. HOSP. GRAL. DE FUERTEVE | 40020 | 44117 | 42 | Al realizar un cráneo sale la mitad y va muy lento. | New Work Order Activity has been created with type: Diagnostic performed by Engineer Description: After various tests and checks, it is determined that the slow data transfer is due to a remote system (PACS) malfunction and not due to a malfunction. team_s functioning. I speak with Francisco Gordillo. The team is very slow making transfers to the pacs and, furthermore, some skull study has not been completed | P2 |  | Per field feedback, The CT5000 PACS issue in Spain has been resolved. The problem was in the PACS, not our system. | [12 empty] | 116287625 | Closed | 332249 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not a Defect | [5 empty] | Engineering |  |  |  | Closed |  |  | PACS | PACS | [7 empty] | System issue | System | Console-SW | Remote Connection-Not a Defect | Console-SW-Remote Connection-Not a Defect | Console-SW-Remote Connection | Not a Defect | Remote Connection | Console-SW | Console-SW-Remote Connection-Not a Defect | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
122 | 44119 | TRUE | Puerto Rico | METRO MAYAGUEZ INC | 40009 | 44119 | 42 | Shows ring-shaped artifact | Detector module 21 and 31 was ordered. Both modules were received and installed. The following calibration were done. 1. Bad Detector. 2. Phantom calibration. 3. Air Calibration. 4. Constancy Test. 5. IQ Check. 6. Beam and DMS Check | P2 |  | Defective module replaced. | [12 empty] | 116302238 | Closed | 339255 | TRUE | Closed - Done | Closed - Done |  | CTUDT00631717 | [5 empty] | Engineering |  |  |  | Closed |  |  | Defective DMS module | Defective DMS module | 44061 | #NUM! |  |  |  | 2.00000,459800911271,TDMS V2 M.Det Class A |  | System issue | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00631717 | DMS Module | DMS-HW | DMS-HW-DMS Module | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
123 | 44118 | TRUE | Puerto Rico | Manati Medical Center | 40007 | 44118 | 42 | SETTING PRS | PRS Setup | P5 | [14 empty] | 116298672 | New | 340418 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not a Defect |  |  |  |  | PRS Setup | Service |  |  |  | Closed |  |  | PRS Setup | PRS Setup | [7 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
124 | 44122 | TRUE | China | Dalian Branch of Shengjing Hospital of China Medical University | 345001 | 44122 | 42 | wechat:02N/Ingenuity Core128/Elite China/an error | 扫描时床报错，连轴器断裂，更换连轴器。 服务结果：维修后，设备恢复正常工作。check error log,SMC_DEVICE_HOR_ENCODER_NOT_SYNC,Visual inspection,&coupling broken) | P3 |  | Check with FSE, 453567099401 COUPLING 1/4"-6MM FLEXIBLE replaced to fix issue. 轴没有对中导致连轴器损坏 | [5 empty] | Couch_HW | [6 empty] | 116314537 | Closed | 345758 | TRUE | Closed - Done | Closed - Done |  | CTUDT00632179 |  |  |  | Ada Wang | Failure to scan | Engineering |  |  |  | Closed |  |  | Motor coupling | [6 empty] | 453567099401 COUPLING 1/4"-6MM FLEXIBLE |  | System issue | System | Couch-HW | Motor Coupling | Couch-HW-Motor Coupling | Couch-HW-Motor Coupling | CTUDT00632179 | Motor Coupling | Couch-HW | Couch-HW-Motor Coupling | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
125 | 44119 | TRUE | China | SD Laiyang People Hospital | 345005 | 44119 | 42 | 02N/Ingenuity Core128/Elite China/image can not transfer | 图像传输失败。 客户问题详细描述或故障现象： 不能传图到星河 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 星河回传可以用,两台设备到pacs也可以i用 远程诊断与测试结果： 网络问题 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 检查网络,配置remote service, olevel, reset spooler (86038250)服务内容：服务账户解锁。 服务结果：维修后，设备恢复正常工作。 ----------------------------------------------------------------------------- | P5 |  | collect bugreport & picture | [12 empty] | 116300206 | Closed | 341665 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering |  |  |  | Closed |  |  | PACS | PACS | [7 empty] | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
126 | 44118 | TRUE | China | Huaying Mountain Guangneng General Hospital | 345000 | 44118 | 42 | 02N/Ingenuity Core128/Elite China/Image artifacts | Ring artifact on image 扫描出现环形伪影，需要工程师到场维修 服务内容: 更换探测器模块后，设备恢复正常。 | P3 |  | Defective module replaced. | [12 empty] | 116295321 | Closed | 341883 | TRUE | Closed - Done | Closed - Done |  | CTUDT00632110 | [5 empty] | Engineering |  |  |  | Closed |  |  | Defective DMS module | Defective DMS module | 44078 | #NUM! |  |  |  | `459800911281 |  | System issue | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | DMS Module | DMS-HW | DMS-HW-DMS Module | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
127 | 44111 | TRUE | China | Huaying Mountain Guangneng General Hospital | 345000 | 44111 | 41 | 02N/Ingenuity Core128/Elite China/Unable to log in to the system | 用户名密码错误，远程指导恢复登录。 服务结果: 维修后，设备恢复正常工作。 | P3 |  | Follow up with Caroline | [12 empty] | 116248769 | Closed | 330802 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not a defect | [5 empty] | Application |  |  |  | Closed |  |  | User Account | User Password | [7 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
128 | 44119 | TRUE | Australia | MURWILLUMBAH DISTRICT HOSPITAL | 341000 | 44119 | 42 | CT Application Specialist whilst on site reported that a message appeared reporting the maximum number of hard arcs were reached during the short tube conditioning | Work Order #: WO-02221101 Steve Teixeira The system fully meets specification and returned to use Problem Description by Engineer : ________________________________________________________________ DATE THE ISSUE OCCURRED: 16 Oct 2020 ________________________________________________________________ CUSTOMER REPORTED PROBLEM OR EVENT DESCRIPTION: 16 Oct 2020 CT Application Specialist whilst on site reported that a message appeared reporting the maximum number of hard arcs were reached during the short tube conditioning. Specialist reported that it occurred after a Generator test. After the issue occurred the Users did full a shutdown and power up and another short tube conditioning which worked. No arcs were seen in the logs after the power cycle. ________________________________________________________________ CURRENT SYSTEM STATUS: In clinical use. ________________________________________________________________ Diagnostic performed by Engineer : ________________________________________________________________ CUSTOMER REPORTED PROBLEM OR EVENT DESCRIPTION: 16 Oct 2020 CT Application Specialist whilst on site reported that a message appeared reporting the maximum number of hard arcs were reached during the short tube conditioning. Specialist reported that it occurred after a Generator test. After the issue occurred the Users did full a shutdown and power up and another short tube conditioning which worked. No arcs were seen in the logs after the power cycle. ________________________________________________________________ SOFTWARE REVISION OF THE SYSTEM: Ver. 4.1.10 ________________________________________________________________ TROUBLESHOOTING: Based on symptoms recommended the Filament calibration be completed followed by a long tube conditioning. CT Specialist to carry out the recommendations. Will follow up with the site over the next few days. ________________________________________________________________ ----------------------------------------------------------------------------- | P4 |  | Tube Arc | [5 empty] | Gantry_HW | Tube_arc | [5 empty] | 116305481 | In progress | 340568 | TRUE | Closed - Done | Closed - Done |  | CTUDT00630085 | [5 empty] | Engineering |  |  |  | Closed |  |  | Tube Arc | Tube Arc | [7 empty] | System issue | System | IC-HW | Tub Arc | IC-HW-Tub Arc | IC-HW-Tub Arc | CTUDT00630085 | Tub Arc | IC-HW | IC-HW-Tub Arc | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
129 | 44117 | TRUE | Australia | MURWILLUMBAH DISTRICT HOSPITAL | 341000 | 44117 | 42 | Horizontal Table Motion intermittently stalls | Work Order #: WO-02202163 Ian Fishwick The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: CT Radiographer Device Use: Patient positioning Expected and Unexpected Results: After CT couch moved to full out position the CT couch Horizontal motion function is lost, requiring CT Power cycle with wait time in off position. User Impact: Delay in getting next patient positioned. Patient Impact: No patient impact reported. Current Software Version: 4.1.10 Diagnostic performed by Engineer : _______________________________________________________________ PROBLEM DESCRIPTION After CT couch moved to full out position the CT couch Horizontal motion function is lost, requiring CT Power cycle with wait time in off position. ________________________________________________________________ DATE THE ISSUE OCCURRED: 14 Oct 2020 ________________________________________________________________ TROUBLESHOOTING: System SW Version is: 4.1.10 Source of this issue is due to CT System Defect resolved in CT upgrade FCO72800735, Couch Software Patch Upgrade For INGENUITY, BIG BORE AND BR64 This upgrade will need to be scheduled and implemented by Field team. ________________________________________________________________ RECOMMENDATION: Philips CT Engineer will need to implement CT upgrade FCO72800735, Couch Software Patch Upgrade For INGENUITY, BIG BORE AND BR64, this upgrade is released, I have informed CT National Support and Local CT Engineering teams of this requirement. The workaround to implement prior to resolution by the upgrade is to not allow the CT couch to move to the fully out position and if it does to move it back in a few cm. ________________________________________________________________ CUSTOMER UPDATE: Site aware of upgrade requirement and ________________________________________________________________ CURRENT SYSTEM STATUS: Operational. ________________________________________________________________ ----------------------------------------------------------------------------- | P5 |  | FCO72800735 | [12 empty] | 116286751 | In progress | 334516 | TRUE | Closed - Duplicate | Closed - Duplicate |  | CTUDT00630138/630160 |  |  |  | Robin | Lockup or Crash | Engineering | Closed |  |  | Closed |  |  | FCO72800735 | [8 empty] | System issue | system | Couch-SW | Horizontal Move Block | Couch-SW-Horizontal Move Block | Couch-SW-Horizontal Move Block | CTUDT00630138/630160 | Horizontal Move Block |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
130 | 44125 | TRUE | United States | Clay County Medical Center | 40003 | 44125 | 43 | Intermittent Xray Errors | Work Order #: WO-02255625 21-Oct-2020 Stauffer Jeffrey Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: CT exams Expected and Unexpected Results: Expected: ct exams without any interuptions. Unexpected: Intermittent xray errors during exams. User Impact: Delayed Scans Patient Impact: Delayed scans Current Software Version: 4.1 Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | P4 |  | Insufficient Privileges | [5 empty] | Gantry_HW | Tube_arc | [5 empty] | 116338399 | In progress | 353584 | TRUE | Closed - Done | Closed - Done |  | CTUDT00630085 | [5 empty] | Engineering |  |  |  | Closed |  |  | Tube Arc | [8 empty] | System issue | System | IC-HW | Tub Arc | IC-HW-Tub Arc | IC-HW-Tub Arc | CTUDT00630085 | Tub Arc | IC-HW | IC-HW-Tub Arc | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
131 | 44126 | TRUE | United States | Seton Medical Center Harker He | 40006 | 44126 | 43 | Needs application support | Work Order #: WO-02264109 Curry, Dyan The system fully meets specification and returned to use Problem description by Clinical Support Specialist : Customer Function: Technologist Device Use: in use Expected and Unexpected Results: protocol help Patient Impact: none Current Software version : 4.1.7 Diagnostic performed by Engineer : helped customer create EC for t/l spine wo/w. Scanning in 2 series with mprs assigned to each accession number. ----------------------------------------------------------------------------- | P5 | [14 empty] | 116349388 | Closed | 356806 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not A Defect | [5 empty] | Application |  |  |  | Closed |  |  | Appointment | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
132 | 44127 | TRUE | United States | Seton Medical Center Harker He | 40006 | 44127 | 43 | Customer requesting support- questions about the exam split study and connection issues | Problem description by Clinical Support Specialist : Customer Function/Role: Technologist Device Use: In use Expected and Unexpected Results: Exams Split User Impact: none Patient Impact: none Current Software Version: Ing64 Diagnostic performed by Engineer : Darla, needed help on reviewing her T-L Spine W/WO Exam Card to ensure the duel exam was splitting accession numbers associated to the images. A review of the entire EC was also accessed to ensure Cor/Sag were build accordingly. Wrist/Hand EC reviewed. Recon reviewed. No further questions at this time. -----------------------------------------------------------------------------Customer scanned a T-L spine and all the results are hanging with the L-spine in PACS. Customer very frantic. Walked her through figuring out what was what. She had multiple recons of everything. We were able to recon bone and soft tissue windows on both sets, w/o and with contrast. Fixed the accession number so that the t-spine would transfer correctly to PACS. This did work. She would like a call back tomorrow to go over the protocol, to check and make sure the PROCEDURE was set up correctly. MBH ----------------------------------------------------------------------------- | P5 | [14 empty] | 116356138 | Incomplete | 359597 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not a defect | [5 empty] | Application |  |  |  | WIP |  |  | Exam Card | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
133 | 44127 | TRUE | Puerto Rico | Manati Medical Center | 40007 | 44127 | 43 | Problems with transmission of images to the PACS | The customer reported problems sending images to the PACS. When was the first occurrence? Within... : 23-10-2020 16:0:0 How often is issue occurring? : Continuously Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: No further information is available as per the customer Patient Impact/Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer | P3 | [14 empty] | 116356503 | In progress | 359739 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering |  |  |  | Closed |  |  | PACS | PACS | [7 empty] | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
134 | 44127 | TRUE | Chile | CMD Clinica Integrales Rancagu | 40008 | 44127 | 43 | CL1675 -DAMAGE TO THE GANTRY | Problem Description by Engineer: Break in gantry cover. Follow Up Required: Replacement of broken cap. Diagnostic performed by Engineer: Replacement of broken part is required. ----------------------------------------------------------------------------- | P3 |  | Front cover replacement in gantry is performed. | [12 empty] | 116356507 | Closed | 359854 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not a defect | [5 empty] | Customer |  |  |  | Closed |  |  | Cover | [6 empty] | 1.00000,459801484881,ASSEMBLY-FRONT COVER |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
135 | 44125 | TRUE | Turkey | T.C Saglik Bakanligi Bagcilar -> Bağcılar State Hospital | 40010 | 44125 | 43 | Masa yatay harekette ses çıkartıyor | Sound is heard during horizontal movement of the table Resolution: The horizontal movement mechanism of the table has been checked. The problem of friction of the table side cover has been fixed. The device is active. | P3 |  | Table horizontal movement mechanism has been checked. The problem of rubbing the table side cover has been fixed. The device is active. | [5 empty] | Couch_HW | [6 empty] | 116338856 | In progress | 354694 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not a defect | [5 empty] | Engineering |  |  |  | Closed |  |  | Table Noise | [8 empty] | System issue | System | Couch-HW | Cover | Couch-HW-Cover | Couch-HW-Cover | Not a defect | Cover | Couch-HW | Couch-HW-Cover | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
136 | 44127 | TRUE | Turkey | T.C Saglik Bakanligi Bagcilar -> Bağcılar State Hospital | 40010 | 44127 | 43 | table error | sound is heard when the bed moves back and forth The pin on the bearing back and forth shaft is fixed in place | P2 |  | FSE mechanically pushed the bolt back to its original location as below. FSE told me that there is not any issue now but is linkage assy the name for this assy | [12 empty] | 116354795 | Closed | 360975 | TRUE | Closed - Done | Closed - Done |  | CTUDT00631671 | [5 empty] | Engineering |  |  |  | Closed |  |  | Table Noise | [8 empty] | System issue | System | Couch-HW | Horizontal linkage | Couch-HW-Horizontal linkage | Couch-HW-Horizontal linkage | CTUDT00631671 | Horizontal linkage | Couch-HW | Couch-HW-Horizontal linkage | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
137 | 44125 | TRUE | United States | VA Medical Center Biloxi | 40018 | 44125 | 43 | Onsite support | device is a complete harddown and Service need to be dispatch for onsite. This is a New Install So the Product is under warranty . Need Service ASAP | P5 |  | Need FCO done for couch issue | [5 empty] | Couch_HW | [6 empty] | 116341851 | Incomplete | 356905 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | CTUDT00630138/630160 | [5 empty] | Engineering |  |  |  | Closed |  |  | FCO72800735 | [8 empty] | System issue | system | Couch-SW | Horizontal Move Block | Couch-SW-Horizontal Move Block | Couch-SW-Horizontal Move Block | CTUDT00630138/630160 | Horizontal Move Block |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
138 | 44123 | TRUE | Spain | S.C.S. HOSP. GRAL. DE FUERTEVE | 40020 | 44123 | 43 | Communication failure. It does not send to PACS. It seems to be a problem with PACS. | Numerous tests are carried out together with hospital informatics. In principle, while the shipments were failing, the only relevant thing found was that in the PACS log the association requests did not appear on the part of the TAC, but instead, the PACS requested something called pooling association that apparently is common among the PACS own servers and stations (PACS managers are informed). On the other hand, a series of modifications are made to the TAC configuration and an obsolete node that shared IP with the TAC is eliminated from the PACS but none of these changes should solve the problem, however, the problem stops occurring and the shipments are now done correctly. The next day it is verified that the equipment continues to send correctly. intermittent problems to send to pacs, the sending queue is restarted. The problem reappears. To the portal the shipments work well. They are doing a migration of the pacs. | P3 | [14 empty] | 116317508 | In progress | 348670 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not a Defect | [5 empty] | Engineering |  |  |  | Closed |  |  | PACS | PACS | [7 empty] | System issue | System | Console-SW | Remote Connection-Not a Defect | Console-SW-Remote Connection-Not a Defect | Console-SW-Remote Connection | Not a Defect | Remote Connection | Console-SW | Console-SW-Remote Connection-Not a Defect | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
139 | 44126 | TRUE | Spain | S.C.S. HOSP. GRAL. DE FUERTEVE | 40020 | 44126 | 43 | Problems with slowness when sending to PACS from the TAC. | transmission problems to pacs. | P4 | [14 empty] | 116343656 | Closed | 354900 | TRUE | Closed - Duplicate | Closed - Duplicate |  | Not a Defect | [5 empty] | Engineering |  |  |  | Closed |  |  | PACS | PACS | [7 empty] | System issue | system | Console-SW | Remote Connection-Not a Defect | Console-SW-Remote Connection-Not a Defect | Remote Connection | Not a Defect | Remote Connection |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
140 | 44125 | TRUE | United States | McLaren Macomb | 40021 | 44125 | 43 | when getting partial images and wants to know if theirs a way to change the viewing on the partials | Problem description by Clinical Support Specialist : Customer Function: Technologist Device Use: in use Expected and Unexpected Results: protocol questions Patient Impact: none Current Software version : 4.1.7 Diagnostic performed by Engineer : Brain evolving images are too light. Is there a setting to change this? Reached out to Chris Beason for assistance. Chris does not know of a setting to change evolving images window setting. Called customer and asked his to take a picture with patient info removed and email it to me . Will f/u with customer tomorrow morning. Diagnostic performed by Engineer : 586-493-3024 Mike Emailing Dyan images of heads that are light in the evolving stage. Requesting callback at this number when she returns. MBH Diagnostic performed by Engineer : received images from Customer . Evolving images appear to be sufficient . They are not final reconned images. Mike agrees nd says its ok to close case. | P4 |  | Kevin offline check | [12 empty] | 116340913 | Incomplete | 353682 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not a defect | [5 empty] | Application |  |  |  | WIP |  |  | Q&A | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
141 | 44128 | TRUE | United States | McLaren Macomb | 40021 | 44128 | 43 | Issues with sending images to PACS | Diagnostic performed by Engineer : Spokw with Roc. Intial problemwas being unable to log into the system. The site has a unique password assigned to the CT user. The system was locked out. Rob was able to login after a timeout period. The most likely cause was a password spelling issue. Rob states that the site is having issues with certain protocols not sending to PACS, Was unable to have Rob log into olevel. This has either been changed or is a different software release. I am dispatching the case for Monday service as it will best be handled with a FSE and the PACS admin working live on the networking. Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: Having issues sending some exams to PACS User Impact: None Patient Impact: None Current Software Version: Follow Up Required : Onsite FSE needed to work out networking issues to determine if it is actually network or protocol related. | P4 | [14 empty] | 116357364 |  | 360185 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering |  |  |  | Closed |  |  | PACS | PACS | [7 empty] | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
142 | 44123 | TRUE | United States | Central Vermont Medical Center | 40024 | 44123 | 43 | does not seal with front cover properly. | Gantry Lexan Ring does not seal with front cover properly. This is a potential infection control issue for the customer, patients, and the FSE. The gantry is already Covid-19 compromised upon delivery. Lexan ring on new covers appears similar to the original Ingenuity lexan ring and does not facilitate a tight seal between ring and cover. There is a 3mm gap between the surfaces allowing a point of entry for contrast media and other fluids into the gantry at the bottom of the lexan ring. All other surfaces of the ring make a tight seal. |  |  | Lexan ring size is bigger than gantry hole E-128610 ECR-075975 Ingenuity Refresh Lexan Ring Optimized | [14 empty] | 349662 | TRUE | Closed - Done | Closed - Done |  | ECR-075975 | [5 empty] | Engineering |  |  |  | Closed |  |  | Lexan Ring/Contrast | [8 empty] | System issue | System | Gantry-HW | Lexan Ring | Gantry-HW-Lexan Ring | Gantry-HW-Lexan Ring | ECR-075975 | Lexan Ring | Gantry-HW | Gantry-HW-Lexan Ring | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
143 | 44123 | TRUE | United States | Central Vermont Medical Center | 40024 | 44123 | 43 | installation hardware kit | The installation hardware kit (12NC 459801527551) contains a quantity of 21 washers (12NC 453567240831). In the kit we recieved for the installation, none of the 21 washers had a correct inner diameter to be used in conjunction with gantry and couch anchor screw (453567132381 and 453566494191). See attached photos |  |  | 12NC 453567240831 washer size too big | [14 empty] | 349706 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | TBC_wahser | [5 empty] | Operation |  |  |  | WIP |  |  | Washer size | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
144 | 44123 | TRUE | United States | Central Vermont Medical Center | 40024 | 44123 | 43 | Computer fails to boot prior to applications turnover | HP Z4 Computer fails to boot prior to applications turnover System was fully installed, all calibrations and performance testing was completed. Dual monitors were both functioning. After creating Physicist user account, when the system rebooted, it failed to boot again. - The z4 computer was brought to minimum config but the CRC Rack USB cables for KB & Mouse were used to keep KB & Mouse functionality while in minimum config state. The PC would still not boot. There were no diagnostic codes, the power LED on the PC was steady white, and the BIOS logo would not appear on the screen or would hang indefinitely at the HP Sure Start logo. When the power LED was pressed again, the PC shut immediately off, it was not booted without video. A replacement Z4 computer was installed and the same problem persisted. Troubleshooting revealed the keyboard USB cable inside the CRC rack (12NC 459801540281) the Keyboard USB Extension cable (12NC 459801032501) was preventing bootup of the Z4 PC. - PC at minimum boot config and all USB connections except KB USB cable resulted in no boot - Keyboard disconnected at end table end of Console harness (455012202371)resulted in no boot - Keyboard USB extension cable in console table harness (455012202371) was unplugged at back of CRC rack resulted in no boot - Keyboard cable (459801032501) inside CRC rack disconnected at back of host PC resulted in successful boot - The original host computer was installed and the keyboard USB cable ((459801032501) was disconnected at the back of the host PC. The keyboard was plugged into the barcode reader USB port on the table harness cable(455012202371) and the system was fully functional. reinserting the Keyboard USB would halt boot again. |  |  | Console can not boot | [14 empty] | 349769 | TRUE | Closed - Done | Closed - Done |  | Not a defect | [5 empty] | Engineering |  |  |  | Closed |  |  | Host PC | [8 empty] | System issue | System | Console-HW | Keyboard USB cable | Console-HW-Keyboard USB cable | Console-HW-Keyboard USB cable | Not a defect | Keyboard USB cable | Console-HW | Console-HW-Keyboard USB cable | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
145 | 44125 | TRUE | United States | Seton Medical Center Harker He | 40006 | 44125 | 43 | XR-29 Certificate | Work Order #: WO-02253427 Hirchak, Michele The system fully meets specification and returned to use Problem description by Clinical Support Specialist : Customer Function/Role:: Tech Device Use: In Use Expected and Unexpected Results: XR-29 certificate request. User Impact: none Patient Impact: none Current Software Version: 4.1.7 Diagnostic performed by Engineer : Emailed certificate to customer and FSE. MBH ----------------------------------------------------------------------------- | P5 |  | Emailed certificate to customer and FSE | [12 empty] | 116339755 | Closed | 352752 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not A Defect | [5 empty] | Customer |  |  |  | Closed |  |  | Request | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
146 | 44124 | TRUE | United States | Central Vermont Medical Center | 40024 | 44124 | 43 | System isn_t receiving ECG data from monitor | workstation is shutting down by itself | P3 |  | workstation is shutting down by itself | [12 empty] | 116329251 | Closed | 353794 | TRUE | Closed - Done | Closed - Done |  | Not A Defect | [5 empty] | Customer |  |  |  | Closed |  |  | Mis-use | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
147 | 44126 | TRUE | United States | Central Vermont Medical Center | 40024 | 44126 | 43 | Head artifact issues | The system partly meets specification and is returned to use with restrictions When was the first occurrence? Within... : 22-10-2020 Expected and Unexpected results: Head Artifacts Diagnostic performed by Engineer : Called and left voicemail for the customer to call back to determine next step. Internal Comments : Error # and/or description of error : Head Artifacts Completed or Recommended Repair Action "Internal" : Suggest checking mylar and detector faces for debris. Suggest reconstructing image without Image Balance and no calibration. Helps determine if calibration is causing issue. Suggest using ring ID, bad detectors and physics analysis to determine where the bad detectors reside. Suggest possible phantom calibration. Also visually check the Compensator for cracks. Refer to the image quality troubleshooting guide for further assistance. Follow Up Required : What are the next Troubleshooting steps : Suggest checking mylar and detector faces for debris. Suggest reconstructing image without Image Balance and no calibration. Helps determine if calibration is causing issue. Suggest using ring ID, bad detectors and physics analysis to determine where the bad detectors reside. Suggest possible phantom calibration. Also visually check the Compensator for cracks. Refer to the image quality troubleshooting guide for further assistance. | P4 |  | Called customer to schedule w/o and was advised no onstie work order is needed and onsite can be cancelled | [12 empty] | 116349692 | Closed | 356828 | TRUE | Closed - Done | Closed - Done |  | Not a defect | [5 empty] | Engineering |  |  |  | Closed |  |  | Defective DMS module | [8 empty] | System issue | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | Not a defect | DMS Module | DMS-HW | DMS-HW-DMS Module | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
148 | 44123 | TRUE | Turkey | HITACHI HEALTHCARE TURKEY | 337000 | 44123 | 43 | Remote connection set up | By configuring the device for remote connection, the system was connected to the Philips server. The remote connection is active. | P3 |  | Remote connection set up | [12 empty] | 116318140 | Closed | 348086 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not A Defect | [5 empty] | Service |  |  |  | Closed |  |  | PRS Setup | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
149 | 44123 | TRUE | Turkey | HITACHI HEALTHCARE TURKEY | 337000 | 44123 | 43 | Error during acquisition. | In the examination made on the device, errors related to slipring were observed. Slipring and brushes are cleared. Air filters have been cleaned. It was thought that the source of the errors occurring in the slipring in a short time might be the cleaning material used to clean the device. It was discussed with users that not to use bleach and similar sub+E179:I179stances, and that cleaning can be done with the substances in the procedure determined by Philips. The device was delivered actively by shooting trials with the Phantom. | P3 |  | COVID-19 cleaning？ | [12 empty] | 116318133 | Closed | 348089 | TRUE | Closed - Done | Closed - Done |  | TBD_Slipring | [5 empty] | Customer |  |  |  | Closed |  |  | Cleaning Material | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
150 | 44126 | TRUE | Turkey | HITACHI HEALTHCARE TURKEY | 337000 | 44126 | 43 | Not shooting | When the device_s errors were examined, it was seen that it gave an error about the table. It is possible that there was an instant mechanical impingement due to the weight of the patient. The shooting of the same patient was repeated, and it was seen that the shooting was done without any problems. | P2 |  | What error was poped up? Collect bugreport and timestamp The shooting of the same patient was repeated, and it was seen that the shooting was done without any problems. | [12 empty] | 116346805 | Closed | 355976 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not a Defect | [5 empty] | Engineering |  |  |  | Closed |  |  | Table error | Missing information | [7 empty] | System issue | System | Couch-HW | Horizontal drive assembly | Couch-HW-Horizontal drive assembly | Couch-HW-Horizontal drive assembly | Not a Defect | Horizontal drive assembly | Couch-HW | Couch-HW-Horizontal drive assembly | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
151 | 44124 | TRUE | Australia | MURWILLUMBAH DISTRICT HOSPITAL | 341000 | 44124 | 43 | Random red artefacts | When performing surview scans on a head scan, it was noted that there were random red artefacts. These did not appear on the scan itself. These artefacts I have not seen in previous software releases. No patient impact. |  |  | DVI Adaptor failure with poor quality, multi site have report this adaptor failure during installation. | [14 empty] | 347941 | TRUE | Closed - Done | Closed - Done |  | CTUDT00632173 | [5 empty] | Engineering |  |  |  | Closed |  |  | DVI Adaptor | [8 empty] | System issue | System | Console-HW | Adaptor | Console-HW-Adaptor | Console-HW-Adaptor | CTUDT00632173 | Adaptor | Console-HW | Console-HW-Adaptor | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
152 | 44124 | TRUE | Australia | MURWILLUMBAH DISTRICT HOSPITAL | 341000 | 44124 | 43 | Table Stop | When using gantry controls for horizontal slide of table, it hits the table limits very hard with a loud noise both moving the table in and out. |  |  | FCO72800735 | [14 empty] | 347942 | TRUE | Closed - Done | Closed - Done |  | CTUDT00630138/630160 | [5 empty] | Engineering |  |  |  | Closed |  |  | FCO72800735 | [8 empty] | System issue | System | Couch-SW | Horizontal Move Block | Couch-SW-Horizontal Move Block | Couch-SW-Horizontal Move Block | CTUDT00630138/630160 | Horizontal Move Block | Couch-SW | Couch-SW-Horizontal Move Block | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
153 | 44124 | TRUE | Australia | MURWILLUMBAH DISTRICT HOSPITAL | 341000 | 44124 | 43 | Update Licence keys | Work Order #: WO-02243530 21-Oct-2020 Ruben MARQUEZ The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Radiographer Device Use: in use Expected and Unexpected Results: Load new License User Impact: NA Patient Impact: NA Current Software Version: 4.1.10 Diagnostic performed by Engineer : Checked and uploaded the new license. | P5 |  | uploaded the new license. | [12 empty] | 116329299 | Closed | 351089 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not A Defect | [5 empty] | Service |  |  |  | Closed |  |  | Installation license | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
154 | 44125 | TRUE | Australia | MURWILLUMBAH DISTRICT HOSPITAL | 341000 | 44125 | 43 | "Gantry cannot initialize.." message | Work Order #: WO-02244196 Steve Teixeira The system does not meet specification and is out of use DATE THE ISSUE OCCURRED: 21 Oct 2020. 21 Oct 2020 Customer called to report that "Gantry cannot initialize.." message appeared attempting a scan. Engineer had just left site after updating the license key. A short tube conditioning was successfully completed but attempting a surview scan on the next patient produced a "Gantry cannot initialize.." message. A patient was scanned previously today without issue. CURRENT SYSTEM STATUS: System down. System would not carry out exam protocols. Couch error. REPLACE 453566496941,RELAY,24VDC PLUGIN W/COVER | P2 |  | What is the couch type? Why upload license key? Supplier Quality check of 453566496941,RELAY,24VDC PLUGIN W/COVER. | [5 empty] | Couch_HW | [6 empty] | 116330215 | Closed | 351011 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not a Defect | [5 empty] | Engineering |  |  |  | Closed |  |  | Table error | HW Replace/Keep obs | [5 empty] | 453566496941,RELAY,24VDC PLUGIN W/COVER |  | System issue | System | Couch-HW | Motor relay | Couch-HW-Motor relay | Couch-HW-Motor relay | Not a Defect | Motor relay | Couch-HW | Couch-HW-Motor relay | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
155 | 44126 | TRUE | Australia | MURWILLUMBAH DISTRICT HOSPITAL | 341000 | 44126 | 43 | System won_t power up | System not starting up, Gantry will not start but computers OK. Check over system. Found 20 Amp fuse blown to system. Replace . Carry out functional testing. Return to service. 23 Oct 2020 Account Manager called on behalf of the customer to report that customer was unable to power up the system this morning. Spoke to the customer and informed that a long tube conditioning was completed yesterday and the gantry turned off (approx 6 PM) as usual. This morning on arrival a clicking noise was heard coming from the computer rack and the CT Box display was flashing. Engineer is on site and investigating the issue. | P2 |  | Check if can get field information | [5 empty] | Gantry_HW | 20A Fuse | Not clear | Replace Fuse | Keep monitoring | William Yang |  | 116350661 | Closed | 357074 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not a Defect | [5 empty] | Engineering |  |  |  | Closed |  |  | Fuse blown | HW Replace/Keep obs | [7 empty] | System issue | System | Gantry-HW | Other | Gantry-HW-Other | Gantry-HW-Other | Not a Defect | TBD-NoEnoughInfo | Gantry-HW | Gantry-HW-Other | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
156 | 44126 | TRUE | Australia | MURWILLUMBAH DISTRICT HOSPITAL | 341000 | 44126 | 43 | Network issues | Doctors are reporting that not all the images are being sent at times. TROUBLESHOOTING: 22 Oct 2020 Discussed the network infrastructure which is a mixture of private and public systems. The issue is with regards to studies transferring to Helius PACS (which is a Private managed system). The other PACS node is a GE PACS and is a publicly managed system. Discussed the option to reset the spooler as a possible option but discussed that the network infrastructure the most likely cause for the incomplete transfers. An Engineer is scheduled to attend site on the 23rd Oct and will proceed to reset the spooler as had issues at the time with the olevel password. | P3 | [14 empty] | 116342442 | In progress | 357261 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering |  |  |  | Closed |  |  | PACS | PACS | [7 empty] | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
157 | 44128 | TRUE | Australia | MURWILLUMBAH DISTRICT HOSPITAL | 341000 | 44128 | 43 | order new Fuses | 45357045792 x1 fuse Kit | P3 | [14 empty] | 116357401 | In progress | 361737 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not A Defect | [5 empty] | Service |  |  |  | Closed |  |  | Service Parts Order | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
158 | 44122 | TRUE | China | Dalian Branch of Shengjing Hospital of China Medical University | 345001 | 44122 | 42 | 02N/Ingenuity Core128/Elite China/cannot scan | 服务内容：扫描初始化报错，发射接收盒未启动，重启机架。 服务结果：维修后，设备恢复正常工作 | P3 |  | CAN Trace/bugreport | [5 empty] | Gantry_HW | TX/RX | Not clear/no error found | Reboot system | Keep monitoring | William Yang |  | 116315056 | Closed | 346875 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD_OBS | [5 empty] | Engineering |  |  |  | WIP |  |  | TX/RX | [8 empty] | System issue | System | Gantry-HW | Other | Gantry-HW-Other | Gantry-HW-Other | TBD_OBS | TBD-NoEnoughInfo | Gantry-HW | Gantry-HW-Other | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
159 | 44125 | TRUE | China | Dalian Branch of Shengjing Hospital of China Medical University | 345001 | 44125 | 43 | 02N/Ingenuity Core128/Elite China/Frequent crashes | Frequent crashes, restart can restore it. Research and development says it may be a defect; transmitting images frequently get stuck. Resetting/restarting the spooler doesn_t solve it. After reboot, it is back to normal. PRS TR status: connected Repair Action: Suggest onsite service procedures: It is recommended to collect the bug report and submit escalation to L3. | P4 | [14 empty] | 116330094 | In progress | 356232 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering |  |  |  | Closed |  |  | PACS | [8 empty] | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
160 | 44126 | TRUE | China | Dalian Branch of Shengjing Hospital of China Medical University | 345001 | 44126 | 43 | host transfer image hang up <CT> | PACS Issue | [17 empty] | 354667 | TRUE | Closed - Done | Closed - Done |  | CTUDT00630958 | [5 empty] | Engineering |  |  |  | Closed |  |  | PACS | PACS | [7 empty] | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00630958 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
161 | 44123 | TRUE | China | Jinggangshan Hospital, the First Affiliated Hospital of Nanchang University, Jiangxi | 345002 | 44123 | 43 | 02N/Ingenuity Core128/Elite China/connect PACS | Need to add network node, report error when open LAN config. PRS TR status: connected Repair Action: Suggest onsite service procedures: Need backend to modify the network card setting, remove ipv6. | P5 | [14 empty] | 116317254 | In progress | 348693 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not a Defect | [5 empty] | Engineering |  |  |  | Closed |  |  | Lan config | PACS | [7 empty] | System issue | System | Console-SW | Lan config | Console-SW-Lan config | Console-SW-Lan config | Not a Defect | Lan config | Console-SW | Console-SW-Lan config | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
162 | 44123 | TRUE | China | General Hospital of Jizhong Energy Jingxing Mining Group Co. LTD | 345003 | 44123 | 43 | 02N/Ingenuity Core128/Elite China/gantry power off | 客户问题概述： 扫描失败。 客户问题详细描述或故障现象： 床水平运动故障 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： ESCALATION CASE， 研发建议更换AMC和MOTOR 服务内容： [掉电，更换水平电机和驱动。] 服务结果： [设备正常工作。] | P3 |  | Defect raised | [12 empty] | 116315837 | Closed | 348881 | TRUE | Closed - Done | Closed - Done |  | CTUDT00631669 | [5 empty] | Engineering |  |  |  | Closed |  |  | Horizontal movement | [8 empty] | System issue | System | Couch-SW | Horizontal Move Block | Couch-SW-Horizontal Move Block | Couch-SW-Horizontal Move Block | CTUDT00631669 | Horizontal Move Block | Couch-SW | Couch-SW-Horizontal Move Block | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
163 | 44124 | TRUE | China | SD Laiyang People Hospital | 345005 | 44124 | 43 | 02N/Ingenuity Core128/Elite China/Add printer | 服务内容：添加网络节点。 服务结果：维修后，设备恢复正常工作。 | P5 | [14 empty] | 116323497 | Closed | 351633 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not A Defect | [5 empty] | Service |  |  |  | Closed |  |  | Printer config | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
164 | #N/A | #N/A | China | SD Xiongke hospital | 40025 | 44127 | 43 | Partial Ring artifact | All calibration test are passed, but there is a partial ring artifact on patient scan |  |  | Defective module replaced. | [14 empty] | DEFOA CD-31889 | #N/A | #N/A | DEFOA |  | TBC |  |  |  |  | DOA | Operation |  |  |  | WIP |  |  | DMS Module | Defective DMS module | [7 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
165 | 44132 | TRUE | United States | McLaren Macomb | 40021 | 44132 | 44 | Application question: results are not match. | Diagnostic performed by Engineer : The customer indicated that some of the scans were not acquiring fully and missing images. After some investigation it appeared the technologist scanning did not open the parameter enough to acquire the wanted data (this could not be proven). The manager was aided to check the scan types as they would separate the result boxes when scanning, this was not the issue. It was then suggested on more involved exam cards the technologist follow the steps for acquisitional work flow on the system. The manager will monitor the situation and run a bug report should they encounter more missing data during acquisition. No further questions at this time. | P5 |  | No action needed, keep monitoring | [12 empty] | 116378659 | Fixed | 368541 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not a Defect | [5 empty] | Application |  |  |  | Closed |  |  | Mis-use | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
166 | 44132 | TRUE | United States | The Regional Medical Center | 40005 | 44132 | 44 | usb device disconnecting and reconnecting to host computer | Aliyah said that it sounds like a usb device disconnecting and reconnecting to host computer. She keeps hearing the windows chime sound intermittently. She heard it at 12:10 pm, 12:20 pm and 12:24 pm. She made a bug report but ports are blocked on prs so I could not log in. Review bug report host event logs at times 12:10 pm, 12:20 pm and 12:24 pm. and look for errors. The usb extenders for ibox may be disconnecting or another usb device. Check ibox. | P3 |  | checked all the computer connections and IBOX connections, The IBOX USB cable was redirected to make sure it maintains a proper connection because the cable was a little loose | [12 empty] | 116380407 | Closed | 369333 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not a Defect | [5 empty] | Service |  |  |  | Closed |  |  | USB Cabling | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
167 | 44132 | TRUE | Puerto Rico | Manati Medical Center | 40007 | 44132 | 44 | low quality image, there is noise | Diagnostic performed by Engineer Description: validated that the physician made a quality test that didnt pass, explained that the system has the own daily qc check but if additional test is required an onsite visit will be scheduled | P3 |  | validated that the physician made a quality test that didnt pass, explained that the system has the own daily qc check but if additional test is required an onsite visit will be scheduled. | [5 empty] | DMS_SW | [6 empty] | 116380980 | In Process | 369643 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not a defect | [5 empty] | Engineering |  |  |  | Closed |  |  | IQ | [8 empty] | System issue | System | DMS-SW | Other | DMS-SW-Other | DMS-SW-Other | Not a defect | Other | DMS-SW | DMS-SW-Other | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
168 | 44132 | TRUE | Turkey | T.C Saglik Bakanligi Bagcilar -> Bağcılar State Hospital | 40010 | 44132 | 44 | IP address change | Resolution : RSN aktif olmadığı için telefon desteği verilerek IP güncellemesi yaptırıldı. Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : Kullanıcının isteği üzerine IP değişikliği yapılacak. | P3 |  | IP address change | [12 empty] | 116376027 | Closed | 370633 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not A Defect | [5 empty] | Service |  |  |  | Closed |  |  | IP Change | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
169 | 44132 | TRUE | Turkey | HITACHI HEALTHCARE TURKEY | 337000 | 44132 | 44 | Remote connection settings will be completed | In order for the device to provide remote connection access, the product code was written correctly in the system files on the device and re-registered in the Philips PRS server. The device is active. | P3 |  | No action needed | [12 empty] | 116376588 | Closed | 370867 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not A Defect | [5 empty] | Service |  |  |  | Closed |  |  | PRS Setup | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
170 | 44133 | TRUE | United States | Central Vermont Medical Center | 40024 | 44133 | 44 | 1398CDFD2 Problems images going to PACS | Diagnostic performed by Engineer Description: System doesn_t appear to be in CT System Config to know SW Version. FSE Russ Morgan is aware and has pulled logs for analysis. Paul (Clin Apps) just states needed case created. My understanding issue is with PACS node, some images start transferring then stop. After host restart they start transferring again. Reset of spooler has already been performed according to Clin Apps. | P4 |  | PACS transfer issue | [12 empty] | 116386841 | In Process | 371889 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering |  |  |  | Closed |  |  | PACS | [8 empty] | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
171 | 44133 | TRUE | France | SCP GROUPE D IMAGERIE MEDICALE | 40019 | 44133 | 44 | Changing the DICOM parameters. | Diagnostic performed by Engineer : modif config DICOM PACS | P3 |  | config DICOM PACS | [12 empty] | 116385298 | Closed | 372032 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not A Defect | [5 empty] | Service |  |  |  | Closed |  |  | PACS Setup | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
172 | 44134 | TRUE | Turkey | T.C Saglik Bakanligi Bagcilar -> Bağcılar State Hospital | 40010 | 44134 | 44 | IP Ayarlarının yapılması | Outpatient CT remote connection settings were corrected, Emergency CT Protek pacs were introduced. Problem Description by Engineer: PROTEK PACS introduction and fixing remote connection settings | P3 |  | config DICOM PACS | [12 empty] | 116394291 | Closed | 374529 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not A Defect | [5 empty] | Service |  |  |  | Closed |  |  | PACS Setup | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
173 | 44134 | TRUE | Puerto Rico | Manati Medical Center | 40007 | 44134 | 44 | Dose info is not passing to the PACS. The images pass but not the dose image used. | Problem Description by Engineer: Dose report does not reach the PACS. Resolution: Verified system configuration and found missing selections in the dose management. Contacted Pacs administrator and we ran some tests. Now the dose report is send to the Pacs | P3 |  | Dose report not send to PACS | [12 empty] | 116396474 | Closed | 375636 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not A Defect | [5 empty] | Service |  |  |  | Closed |  |  | PACS Setup | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
174 | 44134 | TRUE | China | JL Taonan T.C.M. Hosp. | 345006 | 44134 | 44 | 02N/Ingenuity Core128/Elite China/The tube can not warm up and light | 预热中断，间歇性 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： PRS通但是无法 下日志 M2M不通 远程诊断与测试结果： 怀疑新装机打火 PRS TR状态： 不通，请工程师到场检查。 Repair Action: 建议现场维修步骤： 现场检查日志，如果打火的话检查高压电缆，完成长训，灯丝校正 建议备件12NC及名称： N/A | P5 |  | Tube Arc, fixed by filement and tube conditioning | [5 empty] | Gantry_HW | Tube_arc | [5 empty] | 116391370 | Closed | 376038 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00630085 | [5 empty] | Engineering |  |  |  | Closed |  |  | Tube arc | [8 empty] | System issue | System | IC-HW | Tub Arc | IC-HW-Tub Arc | IC-HW-Tub Arc | CTUDT00630085 | Tub Arc | IC-HW | IC-HW-Tub Arc | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
175 | 44134 | TRUE | Spain | S.C.S. HOSP. GRAL. DE FUERTEVE | 40020 | 44134 | 44 | No envia las imagenes del TAC al PACS. | New Work Order Activity has been created with type: Problem Description by Engineer Description: Sending images from Tac to Pacs do not work. Towards Portal and other stations, it does work. At first, the tac technician tells me that the transfers are working well, but when I speak with Juan Pedro, he confirms that they are still unable to transfer; they have to send the images from the Portal to reach the Pacs. Follow Up Required Description: It is necessary to check if the configuration of the Pacs node, in the Tac, is compatible with the reception of images in that one. Practically, I am told, this transfer has never worked well since this new Tac was installed. | P3 |  | PACS vendor issue | [12 empty] | 116393445 | In progress | 376747 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not a defect | [5 empty] | Others |  |  |  | Closed |  |  | PACS Vendor | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
176 | 44137 | TRUE | United States | Central Vermont Medical Center | 40024 | 44137 | 45 | Issue sending to PACS | Diagnostic performed by Engineer Description: Images would not transfer to pacs.. Walked customer through rebooting host and customer was able to transfer images | P3 |  | PACS transfer issue | [12 empty] | 116400792 | Closed | 378368 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering |  |  |  | Closed |  |  | PACS | [8 empty] | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
177 | 44137 | TRUE | Turkey | HITACHI HEALTHCARE TURKEY | 337000 | 44137 | 45 | Device down | Resolution: The remote connection could not be made to the device and the LOG file could not be viewed. According to the photos sent by the user, there may be slipring cleaning or Aplane error. field service required. Diagnostic performed by Engineer: The device that gives Gantry initializing error does not start shooting. Problem Description by Engineer: Customer Function / Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Follow Up Required: The remote connection could not be made to the device and the LOG file could not be examined. According to the photos sent by the user, there may be slipring cleaning or Aplane error. field service required. | P2 |  | The device that gives Gantry initializing error does not start shooting. The device started taking images after complete restart. Device controls were made and no problems were found. | [12 empty] | 116401343 | Closed | 378612 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD_Slipring | [5 empty] | Customer |  |  |  | Closed |  |  | Cleaning Material | Keep obs | [7 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
178 | #N/A | #N/A | China | Fujian Medical Device Co. LTD. | 345010 | 44144 | 46 | Eye phantom holder | 459801551931 -> 459800207531 HANDLE ASSEMBLY FOR BROAD EYE PHANTOM |  |  | Send correct part to field | [14 empty] | DEFOA CD-32123 | #N/A | #N/A | DEFOA |  | TBC |  |  |  |  | Missing Item | Operation |  |  |  | WIP |  |  | DEFOA | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
179 | 44144 | TRUE | Germany | Gesundheit Nord gGmbh | 40002 | 44144 | 46 | The head examination cannot be reconstructed | MPR cannot be created from axial images Diagnostic performed by Engineer: Troubleshooting: Escalation via FSE Follow Up Required: On-site troubleshooting by FSE necessary FSE is supposed to reconstruct errors. Pick up bug report and images and contact TSI | P3 |  | Bugreport https://share.philips.com/sites/BigBoreRTRefreshFOKCoreTeam/Shared%20Documents/General/Bugreps/Bremen/S00370.zip | [12 empty] | 116437334 | Closed | 394655 | TRUE | Closed - Done | Closed - Done |  | CTUDT00632183 | [5 empty] | Engineering |  |  |  | Closed |  |  | Recon error | [8 empty] | System issue | System | Console-SW | CannotRecon | Console-SW-CannotRecon | Console-SW-CannotRecon | CTUDT00632183 | MPR Recon | Console-SW | Console-SW-CannotRecon | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
180 | 44141 | TRUE | Spain | S.C.S. HOSP. GRAL. DE FUERTEVE | 40020 | 44141 | 45 | Network transmission issues. | They cannot send studies over the network, we do a test and the jam occurs by sending to syngovia, an attempt is made to delete the studies and send again but the problem persists. I connect to the computer and restart the delivery queue. After the restart the studies are transferred successfully. It is kept under observation. | P4 |  | Remote Device After the restart the studies are transferred successfully. | [12 empty] | 116429840 | Closed | 392810 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00631602 | [5 empty] | Engineering |  |  |  | Closed |  |  | PACS | [8 empty] | System issue | System | Console-SW | Remote Connection-1602TBD | Console-SW-Remote Connection-1602TBD | Console-SW-Remote Connection | CTUDT00631602 | Remote Connection | Console-SW | Console-SW-Remote Connection-1602TBD | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
181 | 44141 | TRUE | United States | Central Vermont Medical Center | 40024 | 44141 | 45 | Follow up to Case 0116349692 | Head Artifacts Completed or Recommended Repair Action "Internal" : Suggest checking mylar and detector faces for debris. Suggest reconstructing image without Image Balance and no calibration. Helps determine if calibration is causing issue. Suggest using ring ID, bad detectors and physics analysis to determine where the bad detectors reside. Suggest possible phantom calibration. Also visually check the Compensator for cracks. Refer to the image quality troubleshooting guide for further assistance. | P4 |  | Poor paient positioning, contrast… | [12 empty] | 116427605 | In Progress | 392069 | TRUE | Closed - Duplicate | Closed - Duplicate |  | Not A Defect | [5 empty] | Application |  |  |  | Closed |  |  | Mis-use | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
182 | 44140 | TRUE | United States | St Joseph Westgate Medical Center | 40001 | 44140 | 45 | Q/O: Assist application support with configuration changes | Contact Peter Leion, fse Method of Contact: assignment Action: rescheduled Comments: per fse fse customer requested for reschedule 11/9@1pm | P5 |  | Appointment, follow up | [12 empty] | 116425372 | In Progress | 391098 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not A Defect | [5 empty] | Application |  |  |  | Closed |  |  | Appointment | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
183 | 44140 | TRUE | Germany | Gesundheit Nord gGmbh | 40002 | 44140 | 45 | Beautification work on the table | Beautification work on the table see also CASE 0116245421 about Mr. Halama please plan for about 1 hour at the next opportunity 1x 459800932031 PAINT, TOUCH-UP 90273 ALPINE WHITE 1x 459800932021 PAINT, AEROSOL 90273 ALPINE WHITE in the MT | P5 |  | 90273- Carbon top 90007-White cover 11609-Gray_couch Check color paint 12NC with Ada | [12 empty] | 116423362 | New | 390152 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not A Defect | [5 empty] | Others |  |  |  | Closed |  |  | Painting | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
184 | 44139 | TRUE | United States | Seton Medical Center Harker He | 40006 | 44139 | 45 | the table scan button is not working under the table | Diagnostic performed by Engineer : Inspected system. Table does not move downward using CT box controls, but moves using gantry controls. Could not find error in logs. Will order new CT box Follow Up Required : Replace CT box | P3 |  | 453566200321 phantom water 459801705022,CT BOX ASSEMBLY_BEST | [12 empty] | 116418866 | In Progress | 387967 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632185 | [5 empty] | Engineering |  |  |  | Closed |  |  | CT Box | [8 empty] | System issue | System | Gantry-HW | CT Box | Gantry-HW-CT Box | Gantry-HW-CT Box | CTUDT00632185 | CT Box | Gantry-HW | Gantry-HW-CT Box | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
185 | 44139 | TRUE | Puerto Rico | METRO MAYAGUEZ INC | 40009 | 44139 | 45 | Fluoro CT pedal exposes but does not shows images sometimes and rebuilds are very slow | There is an intermittent error while doing fluoro. Onsite visit is required for further troubleshooting. | P3 |  | waiting Bugreport from Carlos | [12 empty] | 116418473 | In Progress | 387771 | TRUE | Closed - Done | Closed - Done |  | TBD-CCT Performance_Bugrep | [5 empty] | Engineering |  |  |  | WIP |  |  | CCT | [8 empty] | System issue | System | Console-SW | CCT Performance | Console-SW-CCT Performance | Console-SW-CCT Performance | TBD-CCT Performance_Bugrep | CCT Performance | Console-SW | Console-SW-CCT Performance | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
186 | 44138 | TRUE | United States | McLaren Macomb | 40021 | 44138 | 45 | 1004CDFJ2 Phantom not detected during IQ | Diagnostic performed by Engineer : Customer states the IQ is failing with can not find phantom. Instructed customer on setting of phantom for IQ. The customer would like local service to check to see if there is an issue. Follow Up Required : Customer states the IQ is failing with can not find phantom. Instructed customer on setting of phantom for IQ. The customer would like local service to check to see if there is an issue. ----------------------------------------------------------------------------- Resolution : Checked system and advised customer on QC phantom positioning. power cycle system ----------------------------------------------------------------------------- | P5 |  | Repeat,Bugreport? | [12 empty] | 116411229 | Incomplete | 381788 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00631545/CTUDT00630971 |  |  |  |  | Work Flow | Engineering |  |  |  | Closed |  |  | Phantom detection | [8 empty] | System issue | System | DMS-SW | Phantom Detect | DMS-SW-Phantom Detect | DMS-SW-Phantom Detect | CTUDT00631545/CTUDT00630971 | Phantom Detect | DMS-SW | DMS-SW-Phantom Detect | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
187 | 44137 | TRUE | China | JL Taonan T.C.M. Hosp. | 345006 | 44137 | 45 | wechat: 02N/Ingenuity Core128/Elite China/login password | answer customer question | P3 |  | answer customer question about password Kevin Add in FOK APP tips | [12 empty] | 116401368 | Closed | 380809 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not A Defect |  |  |  | Kevin Zhao | User Experience | Application |  |  |  | Closed |  |  | User Account | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
188 | 44151 | TRUE | Turkey | HITACHI HEALTHCARE TURKEY | 337000 | 44151 | 47 | table breakdown | Problem Description by Engineer Description: Table does not shoot due to movement problem. Torque wrench may be required when making a site visit and checking table mechanical settings. The tool order has been entered. A malfunction causing mechanical jam in the table has been detected. Repaired. The system was tested with the phantom in the clinic and delivered ready to use. | P3 |  | Horizontal calibration of the table was done. PM activity performed by Can Özgenlik was accompanied by OJT. The system was tested with the phantom in the clinic and delivered ready to use. | [12 empty] | 116487509 | In progress | 410363 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not a Defect |  |  |  |  | Failure to scan | Engineering |  |  |  | Closed |  |  | Table error | [8 empty] | System issue | System | Couch-HW | Mechanism Jam | Couch-HW-Mechanism Jam | Couch-HW-Mechanism Jam | Not a Defect | Mechanism Jam | Couch-HW | Couch-HW-Mechanism Jam | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
189 | 44151 | TRUE | China | SD Anqiu T.C.M Hosp. | 345017 | 44151 | 47 | The device occasionally gets stuck after a patient scans and needs to be restarted before it can be used again. As a result, Customers are upset and not satisfied. | Yellow screen The device occasionally gets stuck after a patient scans and needs to be restarted before it can be used again. As a result, Customers are upset and not satisfied. | Yes |  | The device occasionally gets stuck after a patient scans and needs to be restarted before it can be used again. Silva Add this issue in yellow screen CTUDT | [14 empty] | 409880 | TRUE | Closed - Done | Closed - Done |  | CTUDT00631367 |  |  |  | Silva Lin | Host | Engineering |  |  |  | Closed |  |  | Yellow Screen | [8 empty] | System issue | System | Console-SW | WF-Yellow Screen | Console-SW-WF-Yellow Screen | Console-SW-Yellow-Screen | CTUDT00631367 | Yellow Screen | Console-SW | Console-SW-WF-Yellow Screen | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
190 | 44148 | TRUE | United States | Seton Medical Center Harker He | 40006 | 44148 | 46 | needs support building a protocol | helped customer create new shoulder protocol for specific post processing. Add overlapping axial scans in bone and soft tissue with 250 mm FOV. Saved protocol to user side | P5 |  | helped customer create new shoulder protocol for specific post processing. Add overlapping axial scans in bone and soft tissue with 250 mm FOV. Saved protocol to user side | [12 empty] | 116482244 | Closed | 407724 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not a Defect |  |  |  | Kevin Zhao | User Experience | Application |  |  |  | Closed |  |  | Q&A | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
191 | 44147 | TRUE | Puerto Rico | Manati Medical Center | 40007 | 44147 | 46 | PRS conection verification | PRS conection verification | P5 |  | Did changes to this system PRS configuration. Now I can acces it via telnet and ftp. | [12 empty] | 116473418 | Closed | 404816 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not a Defect |  |  |  | Brook Shi | PRS Setup | Service |  |  |  | Closed |  |  | PRS Setup | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
192 | 44146 | TRUE | China | SD Anqiu T.C.M Hosp. | 345017 | 44146 | 46 | 02N/Ingenuity Core128/Elite China/Gantry can not power up | 早上使用主机死机，客户重启主机并断开ESTOP后重新加电失败。 设备本身技术问题： 否。 Troubleshooting Action: 关键故障代码或现象： 检查日志发ESTOP闭合后maindriver 没有闭合。 远程诊断与测试结果： 指导客户大关机架后恢复正常。 PRS TR状态： 通。 Repair Action: 建议现场维修步骤： 1.指导客户每天重启主机 。如果确实经常死机，收集BUGreport反馈。 | P4 |  | Yellow Screen 死机的数量 现象 时间点 bugreport | [12 empty] | 116462445 | Closed | 402481 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00631367 |  |  |  | Silva Lin | Host | Engineering |  |  |  | Closed |  |  | Yellow Screen | [8 empty] | System issue | System | Console-SW | WF-Yellow Screen | Console-SW-WF-Yellow Screen | Console-SW-Yellow-Screen | CTUDT00631367 | Yellow Screen | Console-SW | Console-SW-WF-Yellow Screen | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
193 | 44146 | TRUE | Australia | MURWILLUMBAH DISTRICT HOSPITAL | 341000 | 44146 | 46 | Red speckle dots on console display | Work Order #: WO-02373083 David BROWN Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A | P3 |  | Replace DP to DVI adaptor. | [12 empty] | 116462922 | Closed | 402457 | TRUE | Closed - Done | Closed - Done |  | CTUDT00632173 |  |  |  | Shelvin Yuan | Image Quality | Engineering |  |  |  | Closed |  |  | DVI Adaptor | [8 empty] | System issue | System | Console-HW | Adaptor | Console-HW-Adaptor | Console-HW-Adaptor | CTUDT00632173 | Adaptor | Console-HW | Console-HW-Adaptor | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
194 | 44145 | TRUE | Puerto Rico | Manati Medical Center | 40007 | 44145 | 46 | Does not send images to PACS | Resolution : During my visit the system did not fail. Anyway I Verified the Pacs configuration and reconfigured it as a remote workstation. Together with Pacs administrator ran several tests and system did not fail. | P3 |  | PACS Transfer issue Brook Ask field to get bugreport Lilian to update in Remote Device CTUDT with bugreport | [12 empty] | 116458505 | Closed | 398673 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 |  |  |  | Silva Lin | Remote Device | Engineering |  |  |  | Closed |  |  | PACS | [8 empty] | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
195 | 44144 | TRUE | Puerto Rico | METRO MAYAGUEZ INC | 40009 | 44144 | 46 | The Stretchermoves the gantry is turN Off | Diagnostic performed by Engineer : The errors were verified in the view log and it was found that the ABS encoder is not synchronizing with the relative encoder. Problem Description by Engineer : When the Table Top is moved, the table stops and the gantry turns off. Resolution : An inspection was made to the couch and it was found that the motor coupling is defective. The motor coupling is replaced, then the table top is calibrated horizontally. Horizontal movement tests were made with patient and the table top responded without any problem. | P2 |  | The motor coupling is replaced, then the table top is calibrated horizontally. Horizontal movement tests were made with patient and the table top responded without any problem. Confirm if can get the field part return to Suzhou | [12 empty] | 116440184 | Closed | 395546 | TRUE | Closed - Done | Closed - Done |  | CTUDT00632179 |  |  |  | Ada Wang | Failure to Plan | Engineering |  |  |  | Closed |  |  | Motor coupling | [6 empty] | 453567099401 | COUPLING/SPIDER | System issue | System | Couch-HW | Motor Coupling | Couch-HW-Motor Coupling | Couch-HW-Motor Coupling | CTUDT00632179 | Motor Coupling | Couch-HW | Couch-HW-Motor Coupling | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
196 | 44144 | TRUE | Turkey | T.C Saglik Bakanligi Bagcilar -> Bağcılar State Hospital | 40010 | 44144 | 46 | Gantry is not ready | Problem Description by Engineer: The device interrupts the capture. Resolution: S_CTRL_LINK_GHOST_GOT_CRC_FRAMING_ERROR error was seen. The maintenance that will be done in the recent past will be pulled forward and slipring cleaning will be done. | P3 |  | The maintenance that will be done in the recent past will be pulled forward and slipring cleaning will be done. Share bugreport if possible add COVID-19 Cleaning APP tips | [12 empty] | 116438089 | Closed | 395026 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | NotEngineering-SlipringClean |  |  |  | William Yang | Failure to Plan | Engineering |  |  |  | WIP |  |  | Slipring | [8 empty] | System issue | System | Gantry-HW | SlipRingClean | Gantry-HW-SlipRingClean | Gantry-HW-SlipRingClean | NotEngineering-SlipringClean | SlipRingClean | Gantry-HW | Gantry-HW-SlipRingClean | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
197 | 44140 | TRUE | United States | Seton Medical Center Harker He | 40006 | 44140 | 45 | The new Ingenuity 5000 table pad now has Velcro the entire length of the pad. This makes it impossible to use the arm rest. The arm rest are designed to slide between the table and the pad when needed. The Velcro prevents the arm rest from sliding between the table and pad now. See pics attached A work around is to slide the arm rest between the patient and the pad. However, this does NOT create a stable arm rest and will make starting IVs very difficult. | The new Ingenuity 5000 table pad now has Velcro the entire length of the pad. This makes it impossible to use the arm rest. The arm rest are designed to slide between the table and the pad when needed. The Velcro prevents the arm rest from sliding between the table and pad now. See pics attached A work around is to slide the arm rest between the patient and the pad. However, this does NOT create a stable arm rest and will make starting IVs very difficult. | Yes |  | A work around is to slide the arm rest between the patient and the pad. However, this does NOT create a stable arm rest and will make starting IVs very difficult. | [14 empty] | 393742 | TRUE | Closed - Done | Closed - Done |  | CTUDT00632125 |  |  |  | Ada Wang | User Experience | Engineering |  |  |  | Closed |  |  | RFC_Pad Velcro | [8 empty] | System issue | System | Couch-HW | Arm Rest | Couch-HW-Arm Rest | Couch-HW-Arm Rest | CTUDT00632125 | Arm Rest | Couch-HW | Couch-HW-Arm Rest | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
198 | 44138 | TRUE | China | Hebei Laonianbing Hospital | 345015 | 44138 | 45 | Wechat：02G/Ingenuity Core128/Elite China/Open edit permissions | Open exam edit permissions | P3 |  | CT user no permission to edit exam cards after config in IT User, fixed issue Add to FOK App tips | [12 empty] | 116407844 | Closed | 383073 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not A Defect |  |  |  | Kevin Zhao | User Experience | Application |  |  |  | WIP |  |  | User Account | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
199 | #N/A | #N/A | China | Fujian Medical Device Co. LTD. | 345011 | 44151 | 47 | Big noise in CT Box during gantry rotating | Belt is too long and not adjust well, Belt replaced to fix issue |  |  | under investigation | [14 empty] | DEFOA CD-32401 | #N/A | #N/A | DEFOA |  | Not a Defect |  |  |  | Frank/David |  | Operation |  |  |  | WIP |  |  | DEFOA | [8 empty] | System issue |  |  |  |  | Gantry-HW-Belt |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
200 | #N/A | #N/A | China | Fujian Medical Device Co. LTD. | 345012 | 44158 | 48 | BD Calibraiton failed | BD Calibraiton failed with Module 34, entire module 34 with bad detector Detector module 34 all channel bad detector (calibration history and air scan images see below cloud drive)--- DEFOA CD-32457 was raised Sep 21: 16:23- 18:52, perform 5 times BD calibration, passed in 5th time, what was done in factory to make it passed? Nov 20: Failed with Module 34 bias and Nosie test. Module SN: 37767; Return Material Label: R402160918 |  |  | under investigation https://www.clouddrive.philips.com.cn/file/t8sku5om88f6uu5jajgzdakd4sfv1yp6# Password:123 | [14 empty] | DEFOA CD-32457 | #N/A | #N/A | DEFOA |  | Not a Defect |  |  |  | David |  | Operation |  |  |  | WIP |  |  | DMS Module DOA | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
201 | #N/A | #N/A | China | SX Datong Yanggao People Hosp. | 345025 | 44158 | 48 | BD Calibraiton failed | Module 13 (Channel 1&2) and 26 (Channel 1) BD calibration failed |  |  | under investigation | [14 empty] | DEFOA CD-32475 | #N/A | #N/A | DEFOA |  | Not A Defect |  |  |  | David |  | Operation |  |  |  | WIP |  |  | DMS Module DOA | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
202 | 44144 | TRUE | Poland | Ministerstwo Zdrowia | 341001 | 44144 | 46 | Repair | Preparing the camera for work. Software configuration. | P4 | [14 empty] | 116438212 | Closed | 397291 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not a defect | [5 empty] | Service |  |  |  | Closed |  |  | Printer config | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
203 | 44151 | TRUE | Germany | Gesundheit Nord gGmbh | 40002 | 44151 | 47 | Problems with reconstruction | MPR Recon with axial CCT head recordings are sometimes not shown completely. The supervisor service was started on CIRS Server 2. Service switched off and set to manual. In the tests carried out, the system complies with the manufacturer_s specifications and is ready for clinical use. | P3 |  | Console team review with operation and supplier for the CIRS SW installation instructions | [12 empty] | 116488003 | Closed | 413115 | TRUE | Closed - Done | Closed - Done |  | CTUDT00632116 | [5 empty] | Engineering |  |  |  | Closed |  |  | CIRS Supervisor | [8 empty] | System issue | System | Console-SW | MPR Recon | Console-SW-MPR Recon | Console-SW-MPR Recon | CTUDT00632116 | MPR Recon | Console-SW | Console-SW-MPR Recon | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
204 | 44151 | TRUE | Turkey | T.C Saglik Bakanligi Bagcilar -> Bağcılar State Hospital | 40010 | 44151 | 47 | Detector failure | No defective detector was detected in the device during the examination. When the device covers are opened, it has been observed that contrast material leaks on the TDMS cover, causing artifacts. Air Cal. process has been done. A trial shoot was taken and the system was delivered in working condition. | P3 |  | it has been observed that contrast material leaks on the TDMS cover, causing artifacts. | [12 empty] | 116487629 | Closed | 410488 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD_Contrast intrusion | [5 empty] | Customer |  |  |  | WIP |  |  | Lexan Ring/Contrast | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
205 | 44151 | TRUE | France | SCP GROUPE D IMAGERIE MEDICALE | 40019 | 44151 | 47 | Ring artifact in image due to detector module failure | The client reports a large ring on exams. Remote image recovery and analysis. Module 19 or 25 incriminated. Wall mounted A / M but the problem persists. Module to be replaced. Diagnostic performed by Engineer: modules 18 and 02 swapped | P3 |  | It is a temporary swap, the module will be replaced so we can live with it. | [12 empty] | 116488275 | Closed | 410639 | TRUE | Closed - Done | Closed - Done |  | CTUDT00632110 | [5 empty] | Engineering |  |  |  | Closed |  |  | Defective DMS module | [8 empty] | System issue | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | DMS Module | DMS-HW | DMS-HW-DMS Module | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
206 | 44152 | TRUE | Germany | Gesundheit Nord gGmbh | 40002 | 44152 | 47 | The customer reports with the MIP in the lung window, that the displayed layers are displayed alternately soft and hard | At MIP, the customer reports in the lung window that the displayed layers are alternately displayed soft and hard. Patient pictures from Nov 12 are on the CDR folder and can be viewed from the application. Please pass this on to the application hotline. Software version of the Ingenuity is 4.1.10! Unclear, get pic if possible | P4 |  | info Unclear, need Check with FSE, get pic/bugreport if possible | [12 empty] | 116503157 | New | 417020 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632116 | [5 empty] | Engineering |  |  |  | Closed |  |  | MPR | [8 empty] | System issue | System | CIRS-SW | MPR Recon | CIRS-SW-MPR Recon | CIRS-SW-MPR Recon | CTUDT00632116 | MPR Recon | CIRS-SW | CIRS-SW-MPR Recon | No | y | Console-SW | y | Console-SW-MPR Recon | Console-SW-MPR Recon | No | y | y | y | Console-SW | Console-SW-MPR Recon | y | y | y | y | y | y
207 | 44152 | TRUE | Spain | S.C.S. HOSP. GRAL. DE FUERTEVE | 40020 | 44152 | 47 | Change of configuration of the PACS of Atrys. | Change of configuration of the PACS of Atrys. | P5 | [14 empty] | 116493720 | Closed | 415727 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not a defect | [5 empty] | Service | [6 empty] | PACS config | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
208 | 44152 | TRUE | Australia | MURWILLUMBAH DISTRICT HOSPITAL | 341000 | 44152 | 47 | Display issues | 18 Nov 2020 Customer called to report display issues on the LHS monitor, including the display being vertically offset by a few millimeters and intermittently producing weird looking images. 18 Nov 2020 Swapping the video cables the erratic display moved to the other monitor. Requested the customer disconnect the suspected signal cable connector. Left the system as a one monitor configuration. Informed the customer that the Engineer would arrange time to attend site to investigate the issue further. Ordered the suspected parts. | P3 |  | Check with FSE to inform for monitor setting, not happen on previous verisons | [12 empty] | 116506031 | In Progress | 414927 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632173 | [5 empty] | Engineering |  |  |  | WIP |  |  | Monitor | [8 empty] | System issue | System | Console-HW | Adaptor | Console-HW-Adaptor | Console-HW-Adaptor | CTUDT00632173 | Adaptor | Console-HW | Console-HW-Adaptor | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
209 | 44152 | TRUE | China | XIFENG CHINESE MEDICINE HOSPIT | 345014 | 44152 | 47 | 02G/Ingenuity Core128/Elite China/can not scan the image | recon failed, CIRS reinstalled, reference detector damaged, part replaced pacs transfer issue (`4 times) fixed by spooler reset, FSE will inform customer to collect bugreport when face issue again | P3 |  | CIRS Software reinstalled to fix recon issue, reference detector replace to fix ACQ error | [12 empty] | 116493732 | Closed | 415747 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD_OBS | [5 empty] | Engineering |  |  |  | WIP |  |  | Scan fail | [8 empty] | System issue | System | Gantry-HW | Reference-Detector | Gantry-HW-Reference-Detector | Gantry-HW-Reference-Detector | TBD_OBS | Reference-Detector | Gantry-HW | Gantry-HW-Reference-Detector | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
210 | 44153 | TRUE | Germany | Gesundheit Nord gGmbh | 40002 | 44153 | 47 | Laser enlinement does not work correctly | inner laser is not positioned correctly Adjustment of inner laser and also outer laser | P3 |  | Dose laser assembly screw looses, design improve | [12 empty] | 116510760 | Closed | 416689 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | ECR-077066 | [5 empty] | Engineering |  |  |  | WIP |  |  | Laser alignment | [8 empty] | System issue | System | Gantry-HW | Laser | Gantry-HW-Laser | Gantry-HW-GantryLaser | ECR-077066 | gantry laser | Gantry-HW | Gantry-HW-Laser | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
211 | 44153 | TRUE | China | NMG Zizhiqu No.4 Hosp | 40023 | 44153 | 47 | 02N/Ingenuity CT 728326/Combination lock problem | 客户问题概述： 主机密码锁定 客户问题详细描述或故障现象： 多次登录后密码锁定 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 密码锁定无法登录 远程诊断与测试结果： PRS 不通，无法进行远程重置，客户IT密码说登录不了 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 到场重置密码 | P3 |  | Add APP Tip | [12 empty] | 116506309 | Closed | 418312 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not a defect | [5 empty] | Application |  |  |  | Closed |  |  | User Account | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
212 | 44154 | TRUE | Puerto Rico | Manati Medical Center | 40007 | 44154 | 47 | PROBLEM WITH IMAGE TRANSFER | PROBLEM WITH IMAGE TRANSFER | P3 |  | check if have bugreport to support investigation | [12 empty] | 116519533 | In Progress | 423095 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering |  |  |  | WIP |  |  | PACS | [8 empty] | System issue | System | Console-SW | remote Connection | Console-SW-remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
213 | 44154 | TRUE | Puerto Rico | Manati Medical Center | 40007 | 44154 | 47 | Image transfer problems | There is a problem when sending studies to the PACS | P3 |  | check if have bugreport to support investigation | [12 empty] | 116519563 | In Progress | 423104 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering |  |  |  | WIP |  |  | PACS | [8 empty] | System issue | System | Console-SW | remote Connection | Console-SW-remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
214 | 44155 | TRUE | France | SCP GROUPE D IMAGERIE MEDICALE | 40019 | 44155 | 47 | Changing the DICOM parameters. | modification of DICOM parameters | P3 | [14 empty] | 116524671 | Incomplete | 422951 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not a defect | [5 empty] | Service | [6 empty] | PACS Config | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
215 | 44159 | TRUE | Puerto Rico | METRO MAYAGUEZ INC | 40009 | 44159 | 48 | Ingenuity 5000 CT - Customer dissatisfaction | The costumer is not happy with the performance of Its new Ingenuity CT 5000. He is concerned with this system. Sense day one, this CT started to have the following problems. 1 PRS Communication problem. The applications personal was not able to connect remotely to the CT with the customer to complete the applications. The remote share Desktop operation did not work. 2 The Patient support top, did not move at all, if the patient support top was moved all the way to one of the two end sides. Technician had to wait like 5 to 10 minutes before the before the table top coul be move with the control buttons again. 3 Circular artifact appears on the studies images, caused by two defected detector modules. 4. Fluro image fail to appear on the display monitor when the fluro pedal was pressed for the first time. After that the fluro work ok. The Dr. complained about having a needle inside a patient and not been able to see it the first time he presses the fluro pedal. The pedal had to be pressed a second time in order to see the image. 5. Gantry shut down when laser was power on and the CCT fluro protocol was used. 6. Fluro pedal connector was connected on the wrong connector on the Ghost board. 7. Technician complain about having to do a lot of steps to get the Angio reconstruction completed. 8. Technician complain about the reconstruction time. They find it to be slower than expected. 9. Radiologist complain about the time it takes to pull out remote study images from the Ingenuity CT to their workstation for analysis. They say that the studies that are pull from the Newsoft CT are received faster and completed than the ones from the Ingenuity CT. They had to wait for the transfer to be completed and some time not all the studies are received. Most of this problem had been solved by now. Some of them happened with the patient on the system. I had include an email from the customer with the actual disatisfaction from this system. |  |  | Hot site, Export current complaint list of this site 1 PRS Communication problem. The applications personal was not able to connect remotely to the CT with the customer to complete the applications. The remote share Desktop operation did not work. 2 The Patient support top, did not move at all, if the patient support top was moved all the way to one of the two end sides. Technician had to wait like 5 to 10 minutes before the before the table top coul be move with the control buttons again. 3 Circular artifact appears on the studies images, caused by two defected detector modules. 4. Fluro image fail to appear on the display monitor when the fluro pedal was pressed for the first time. After that the fluro work ok. The Dr. complained about having a needle inside a patient and not been able to see it the first time he presses the fluro pedal. The pedal had to be pressed a second time in order to see the image. 5. Gantry shut down when laser was power on and the CCT fluro protocol was used. 6. Fluro pedal connector was connected on the wrong connector on the Ghost board. It was connected on the Tilt Override from factory 7. Technician complain about having to do a lot of steps to get the Angio reconstruction completed. Today they Will have additional on site applications. The initial applicatons were remote over Skype without the possibility of LOTS 8. Technician complain about the reconstruction time. They find it to be slower than expected. Not reported any more. If problem reapears after application sets up protocols they’ll créate a bugreport 9. Radiologist complain about the time it takes to pull out remote study images from the Ingenuity CT to their workstation for analysis. They say that the studies that are pull from the Newsoft CT are received faster and completed than the ones from the Ingenuity CT. They had to wait for the transfer to be completed and some time not all the studies are received. This seems to be a PACS problema and not CT. We’ll keep monitoring, no data available | [14 empty] | 426445 | TRUE | Closed - Done | Closed - Done |  | TBD_OBS | [5 empty] | Engineering |  |  |  | WIP |  |  | Dissatisfaction | [8 empty] | System issue | System | Console-SW | Performance | Console-SW-TBD-OBS | Console-SW-TBD-OBS | TBD_OBS | TBD | Console-SW | Console-SW-TBD-OBS | No | y | y | y | Console-SW-TBD | Console-SW-TBD | No | y | y | y | y | Console-SW-TBD | y | y | y | y | y | y
216 | 44155 | TRUE | China | NMG Zizhiqu No.4 Hosp | 40023 | 44155 | 47 | 02N/Ingenuity CT 728326/unstable image transfer | T2 Activities : Problem Description: image can not send to PACS smooth Repair Action: instructor customer do reset the spooler on host Test/Inspection: n/a Test Equipment: n/a SAP Transfer : Case Transferred to SA | P3 |  | Remote Device | [12 empty] | 116523860 | Closed | 424524 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not a defect | [5 empty] | Engineering |  |  |  | WIP |  |  | PACS | [8 empty] | System issue | System | Console-SW | Remote Connection-Not a Defect | Console-SW-Remote Connection-Not a Defect | Console-SW-Remote Connection | Not a defect | Remote Connection | Console-SW | Console-SW-Remote Connection-Not a Defect | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
217 | 44155 | TRUE | China | SD Laiyang People Hospital | 345005 | 44155 | 47 | 02N/Ingenuity Core128/Elite China/Chip failure and boot up | T2 Activities : Problem Description: 客户问题概述： 控制面板按键失灵。 客户问题详细描述或故障现象： 面板按键不灵 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： logger无报错。 远程诊断与测试结果： 重启机架无效。 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 检查左面板，必要时更换。 建议备件12NC及名称： PPD 459801484981 | P3 |  | `459801484981 GCP button check if can return to Suzhou | [12 empty] | 116523854 | Closed | 435982 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632181 | [5 empty] | Engineering |  |  |  | WIP |  |  | GCP | [8 empty] | System issue | System | Gantry-HW | Gantry Panel | Gantry-HW-Gantry Panel | Gantry-HW-Gantry Panel | CTUDT00632181 | Control Panel | Gantry-HW | Gantry-HW-Gantry Panel | No | y | y | y | y | y | No | y | y | Gantry Panel | y | y | y | y | y | y | y | y
218 | 44158 | TRUE | Turkey | T.C Saglik Bakanligi Bagcilar -> Bağcılar State Hospital | 40010 | 44158 | 48 | Gantry failure | Problem Description by Engineer: The system has failed to shoot. Diagnostic performed by Engineer: In the intervention performed; The port number 22 of the system local IP has been opened to enable logging with remote access. After RSE extracts and interprets the logs, the malfunction will be intervened again. Follow Up Required: After RSE pulls and interprets the logs, the fault will be intervened again. No detail information | P3 |  | After Reboot, the system was monitored and no problem was observed. | [12 empty] | 116532485 | In Progress | 426224 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD_OBS | [5 empty] | Others |  |  |  | WIP |  |  | Unclear | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
219 | 44158 | TRUE | Turkey | HITACHI HEALTHCARE TURKEY | 337000 | 44158 | 48 | Pacs transfer problem | Because the IP settings of Pacs to which the shots are transmitted on the device have changed, new IP settings have been defined on the device. The system was tested with the phantom in the clinic and delivered ready to use. | P3 |  | Customer | [12 empty] | 116531704 | In Progress | 427933 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not a defect | [5 empty] | Customer |  |  |  | Closed |  |  | PACS Config | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
220 | 44158 | TRUE | Turkey | HITACHI HEALTHCARE TURKEY | 337000 | 44158 | 48 | Microphone malfunction | The microphone button was found to be broken. The relevant part has been ordered. Once the piece arrives, the service will be arranged again. | P3 |  | Replace part check if CT-BOX can return to Suzhou 459801705022,CT BOX ASSEMBLY_BEST | [12 empty] | 116531726 | In Progress | 427944 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632185 | [5 empty] | Engineering |  |  |  | WIP |  |  | CT Box | [8 empty] | System issue | System | Gantry-HW | CT Box | Gantry-HW-CT Box | Gantry-HW-CT Box | CTUDT00632185 | CT Box | Gantry-HW | Gantry-HW-CT Box | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
221 | 44158 | TRUE | China | XIFENG CHINESE MEDICINE HOSPIT | 345014 | 44158 | 48 | 02G/Ingenuity Core128/Elite China/Images cannot be transmitted | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer : Images cannot be transmitted Safety Questions : Case: 0116531139 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Zixuan Last updated by: Wang Zixuan | P3 |  | Remote Device | [12 empty] | 116531139 | Closed | 427597 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering |  |  |  | WIP |  |  | PACS | [8 empty] | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
222 | 44159 | TRUE | Germany | Gesundheit Nord gGmbh | 40002 | 44159 | 48 | CT-controlled puncture. Patient was lying in prone position - images upside down | Request for rotated images After analytical evaluation with Mr. Kraschewski together, one can say that the pictures are technically all in order and are only displayed unfamiliar to the user. Since there will be another application on site in December, I asked the customer to have the dart positions explained again or to modify the settings! | P4 |  | Meet design, can rotate image 180 degree Check if have related instruction in IFU. Yes 459801063524_D IFU Page234-235 | [12 empty] | 116544595 | Closed | 432032 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not a defect | [5 empty] | Application |  |  |  | Closed |  |  | User unfamiliar | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
223 | 44159 | TRUE | Australia | MURWILLUMBAH DISTRICT HOSPITAL | 341000 | 44159 | 48 | Unable to change the window levels | Work Order #: WO-02472956 Steve Teixeira The system fully meets specification and returned to use Diagnostic performed by Engineer : ________________________________________________________________ CUSTOMER REPORTED PROBLEM, QUERY OR EVENT DESCRIPTION: 18 Nov 2020 Customer called to report that unable to change the window settings for a C Spine Recon, as part of a Neck Angio study. Clicking on [Show All] the detailed pop up window does not appear to allow the customer to change the window settings. ________________________________________________________________ APPLICATION SUPPORT: CT Application Specialist contacted the customer and assisted with the following workflow suggestion: 1. Load the soft recon images into Quick Review utility. 2. Apply Image enhancement 3. Adjust the window levels to the customer_s liking. 4. Save the batch. ________________________________________________________________ CUSTOMER REPORTED PROBLEM, QUERY OR EVENT DESCRIPTION: 18 Nov 2020 Customer called to report that unable to change the window settings for a C Spine Recon, as part of a Neck Angio study. Clicking on [Show All] the detailed pop up window does not appear to allow the customer to change the window settings. ________________________________________________________________ CASE CLOSED: As Application support provided as requested, no further action will be taken on this case. | P3 |  | 1. Load the soft recon images into Quick Review utility. 2. Apply Image enhancement 3. Adjust the window levels to the customer_s liking. 4. Save the batch. | [12 empty] | 116541072 | Closed | 427594 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not a defect | [5 empty] | Application |  |  |  | Closed |  |  | Application support | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
224 | 44161 | TRUE | Portugal | Centro Hospitalar de Vila Nova | 40030 | 44161 | 48 | Extra training on the CT Ingenuity 5000 equipment. | Follow-up training. Optimization of protocols and support for exams. | P5 |  | App training | [12 empty] | 116553140 | In Progress | 436584 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not a defect | [5 empty] | Application | [6 empty] | App training | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
225 | 44161 | TRUE | China | SD Anqiu T.C.M Hosp. | 345017 | 44161 | 48 | 02N/Ingenuity Core128/Elite China/can not transfer image | BU Q&R Feedback: No SAP Customer Symptom Code: ACQP Customer Symptom Code: Imaging/Patient Recording/Acquisition Problem Reason: Corrective Maintenance Engineer Failure Code: Configuration Error Repair Activity Code: Configuration Change Problem Reported by Customer : can not transfer image Safety Questions : Case: 0116554777 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Guomei Li Last updated by: Guomei Li 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: can not transfer image Troubleshooting Action: image transfer stuck Repair Action: reset spooler in prs Test/Inspection: image transfer normal_ SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | P3 |  | Remote Device | [12 empty] | 116554777 | Closed | 437631 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00631602 | [5 empty] | Engineering |  |  |  | WIP |  |  | PACS | [8 empty] | System issue | System | Console-SW | Remote Connection-1602TBD | Console-SW-Remote Connection-1602TBD | Console-SW-Remote Connection | CTUDT00631602 | Remote Connection | Console-SW | Console-SW-Remote Connection-1602TBD | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
226 | 44162 | TRUE | Germany | Gesundheit Nord gGmbh | 40002 | 44162 | 48 | Eightfold key arrows released | The buttons are made of Silicone Rubber, suggest not use Raucodrape® Surgical Incise Drape to protect it if not allowed. | P3 |  | William investigation Share with KM the acceptable material for cover the button | [12 empty] | 116560849 | In Progress | 438049 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632181 | [5 empty] | Engineering |  |  |  | WIP |  |  | GCP | [8 empty] | System issue | System | Gantry-HW | Gantry Panel | Gantry-HW-Gantry Panel | Gantry-HW-Gantry Panel | CTUDT00632181 | Control Panel | Gantry-HW | Gantry-HW-Gantry Panel | No | y | y | y | y | y | No | y | y | Gantry Panel | y | y | y | y | y | y | y | y
227 | 44162 | TRUE | Spain | S.C.S. HOSP. GRAL. DE FUERTEVE | 40020 | 44162 | 48 | A study with contrast hangs. | Work Order #: WO-02492808 Carlos Saumell Lasala Regulatory Questions : N/A Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: | P3 |  | Regardless of the previously existing problem, due to which shipments to Pacs are intermittently blocked (it is already attended to from another case), no new incident is observed. | [12 empty] | 116561211 | In Progress | 439686 | TRUE | Closed - Done | Closed - Done |  | CTUDT00631602 | [5 empty] | Engineering |  |  |  | WIP |  |  | PACS | [8 empty] | System issue | System | Console-SW | Remote Connection-1602TBD | Console-SW-Remote Connection-1602TBD | Console-SW-Remote Connection | CTUDT00631602 | Remote Connection | Console-SW | Console-SW-Remote Connection-1602TBD | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
228 | 44164 | TRUE | China | SX Datong Yanggao People Hosp. | 345025 | 44164 | 48 | 02G/Ingenuity Core128/Elite China/Image has artifacts | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer : Image has artifacts Safety Questions : Case: 0116564579 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Yuan Mingze Last updated by: Yuan Mingze 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action : IA#52688481 加OP给Site 1 Liu Kun Next Step Action : 装机经理曹永峰已批复，OP-30已加，派给FSE Liu Kun ----------------------------------------------------------------------------- | P3 |  | DEFOA CD-32475 during installation | [12 empty] | 116564579 | Closed | 440917 | TRUE | Closed - Done | Closed - Done |  | Not A Defect | [5 empty] | Operation |  |  |  | Closed |  |  | Defective DMS module | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
229 | 44166 | TRUE | Germany | Gesundheit Nord gGmbh | 40002 | 44166 | 49 | MPR and MIP cannot be created from axial layers | MPR and MIP cannot be created from axial layers |  |  | plan to reload sw take a backup of exam cards. Check the exam card from bugreport and test on suzhou system | [12 empty] | CFM |  | 444629 | TRUE | Closed - Done | Closed - Done |  | CTUDT00632116 | [5 empty] | Engineering |  |  |  | WIP |  |  | MPR | [8 empty] | System issue | System | CIRS-SW | MPR Recon | CIRS-SW-MPR Recon | CIRS-SW-MPR Recon | CTUDT00632116 | MPR Recon | CIRS-SW | CIRS-SW-MPR Recon | No | y | Console-SW | y | Console-SW-MPR Recon | Console-SW-MPR Recon | No | y | y | y | Console-SW | Console-SW-MPR Recon | y | y | y | y | y | y
230 | 44165 | TRUE | Germany | Gesundheit Nord gGmbh | 40002 | 44165 | 49 | Support requested. | Work Order #: WO-02499888 Michael Gahlen Diagnostic performed by Engineer : N/A Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: | P3 |  | Axial slices do not create MPRs Images and bug report sent to help desk | [12 empty] | 116569288 | In Progress | 444106 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632116 | [5 empty] | Engineering |  |  |  | WIP |  |  | MPR | [8 empty] | System issue | System | CIRS-SW | MPR Recon | CIRS-SW-MPR Recon | CIRS-SW-MPR Recon | CTUDT00632116 | MPR Recon | CIRS-SW | CIRS-SW-MPR Recon | No | y | Console-SW | y | Console-SW-MPR Recon | Console-SW-MPR Recon | No | y | y | y | Console-SW | Console-SW-MPR Recon | y | y | y | y | y | y
231 | 44166 | TRUE | Puerto Rico | METRO MAYAGUEZ INC | 40009 | 44166 | 49 | Break is stuck to table | Resolution : The scissor mechanisms to be greased. The patient support was tested by moving it up and down. The noise did not happened again. The system was left fully operational. Problem Description by Engineer : Problem description by engineer : noise in vertical move and when gantry rotor is rotating Diagnostic performed by Engineer : The noise that was reported by the customer is coming from the patient support scissors bearing. They are lacking from grease. The scissor mechanisms need to be greased. | P4 |  | Table noise due to lacking from grease | [12 empty] | 116578364 | Incomplete | 445485 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not a Defect | [5 empty] | Engineering | Engineering | Engineering | Engineering | Closed | Engineering | Engineering | Table noise | [8 empty] | System issue | System | Couch-HW | Table Noise | Couch-HW-Table Noise | Couch-HW-Table Noise | Not a Defect | Table Noise | Couch-HW | Couch-HW-Table Noise | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
232 | 44165 | TRUE | Puerto Rico | METRO MAYAGUEZ INC | 40009 | 44165 | 49 | Has a noise in the gantry | Resolution : The USB sound was coming from faulty Ibox connection on the left gantry leg. Problem Description by Engineer : Problem description by engineer : The gantry is generating a noise compatible with a usb device is getting disconnected. Onsite visit is required to check the I Box usb cables. Resolution : The Ibox was repositioned the USB connection was fix. The s system was tested and left it fully operational. | P3 |  | IBox USB cable connection | [12 empty] | 116569212 | Closed | 441582 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not a Defect | [5 empty] | Service |  |  |  | Closed |  |  | USB Cabling | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
233 | 44169 | TRUE | China | SD Chest Hosp. | 40025 | 44169 | 49 | 02N/Ingenuity CT 728326/gantry fault | Next Step Action : PPD 453567027195 1&455014003271 1已发24H至医院。FSE黄兆坤-86-18654536122 Next Step Action : 018654536122 FSE Huang Zhaokun 来电告知此单转派金晓 Internal Remarks : 05.12.2020 12:34:03 UTC Xiao Jin (86038251) "RDI Accuracy:Yes" $Q7: (Confirmed reported problem, gantry fault, E-stop open) $Q8: (Logger check, found error:__CT BOX UNINITIATED RESET. CT box check, SDB board check, CRC power supply check, found the SDB board and the CT box problem. ) $Q9: ( Change the SDB board"Part NO: 455014003271" and the CT box, Port NO.453567027195, system callibration, the system is ok. ) $Q10: (Test and inspection data refer to PA. Power on and scan test, Passed. the system is ok. Device returned to full functionality) $11: Calibrated Tools name(Multimeter Fluke 287) S/N (14110098)Next Calibration date (08/23/2021) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube usage: scan second_ External Remarks : 05.12.2020 12:34:00 UTC Xiao Jin (86038251)服务内容：机架掉电，CT box报错，检查CT box， 检查SDB, 检查主机CRC电源，更换CT BOX 和SDB板。机器正常工作。 服务结果：维修后设备恢复正常运行 ----------------------------------------------------------------------------- | P3 |  | Power cable connection issue | [12 empty] | 116598971 | In Process | 457012 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not a Defect | [5 empty] | Service |  |  |  | Closed |  |  | Cable loose | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
234 | 44170 | TRUE | United States | Odessa Regional Medical Center | 40029 | 44170 | 49 | 1399CDFD1 Cannot Transfer To PACS & Injector Not Triggering | CT Tech reports studies are not auto-transferring to PACS (states auto-transfer was working before) but Tech can manually transfer to PACS. Also, Tech reports Injector is not triggering. Tech also reports "the sequence has stopped due to error" when attempting short TC at 0806 on 12/5/20. Follow Up Required Description: Error # and/or description of error : Cannot Transfer To PACS & Injector Not Triggering Completed or Recommended Repair Action "Internal" : Inspect auto-send functionality in DICOM/LAN Config and in preferences and in the exam card configuration. Inspect injector triggering functionality, cabling/connections. Repair/replace as necessary. | P3 |  | Known issue Transfer To PACS struk issue | [12 empty] | 116608163 | In Progress | 456455 | TRUE | Closed - Done | Closed - Done |  | CTUDT00632127 | [5 empty] | Engineering |  |  |  | WIP |  |  | PACS | [8 empty] | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
235 | 44169 | TRUE | United States | Odessa Regional Medical Center | 40029 | 44169 | 49 | 0711CDFM6Unable to turn off DICOM structured dose report. | Work Order #: WO-02535540 Mark Finnigan Problem Description by Engineer : Problem description by engineer : Clinical apps is onsite for first applications session and is having an issue with DICOM structured reporting. When was the first occurrence? Within... : 04-Dec-2020 How often is issue occurring? : Unknown Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: No further information is available as per the customer Patient Impact/Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer Current Software Version: 4.1.10 Diagnostic performed by Engineer : Spoke with Philips clinical apps Rhonda Swain. She is onsite and having a problem with DICOM structered reporting setup. Internal Comments : Error # and/or description of error : Structured reporting is not configuring properly and results in either no structured report being generated or the structured report going to unwanted destinations (PACS) Completed or Recommended Repair Action "Internal" : Structured reporting is not configuring properly and results in either no structured report being generated or the structured report going to unwanted destinations (PACS) Had Michelle Hirchak from applications support join the call and Michelle figured out a work around where we go into LAN config and block the DSR on the PACS node. Dispatching the case to the primary FSE and requesting he open a CFR on the case. Michelle Hirchak will do the write up for the FSE to submit. Follow Up Required : What are the next Troubleshooting steps : File a CFR as requested. Error # and/or description of error : Structured reporting is not configuring properly and results in either no structured report being generated or the structured report going to unwanted destinations (PACS) Completed or Recommended Repair Action "Internal" : Structured reporting is not configuring properly and results in either no structured report being generated or the structured report going to unwanted destinations (PACS) Had Michelle Hirchak from applications support join the call and Michelle figured out a work around where we go into LAN config and block the DSR on the PACS node. Dispatching the case to the primary FSE and requesting he open a CFR on the case. Michelle Hirchak will do the write up for the FSE to submit. ----------------------------------------------------------------------------- Work Order #: WO-02535857 McCain, Jimmie Problem Description by Engineer : Problem description by engineer : Clinical apps is onsite for first applications session and is having an issue with DICOM structured reporting. When was the first occurrence? Within... : 04-Dec-2020 How often is issue occurring? : Unknown Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: No further information is available as per the customer Patient Impact/Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer Current Software Version: 4.1.10 Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | P4 |  | Is possible to go into LAN config and block the DSR on the PACS node? For Dose SR transition issue, two suggestions base on whether the customer create SR in Local: 1. If the customer would not create SR per Exam in Local, Suggest the user Unclick “Create Dose Structured Report” selection under Dose Management UI in Preference, in this case the SR would not be created in Load or PACS. 2. If the customer would like to create Exam SR in Local but not transfer SR to PACS, Suggest the user Enable “Create Dose Structured Report” selection and Not choose PACS in “Select Device” window, And Transfer the data using Auto Storage (add PACS) during exam plan in this case the SR would not be sent to PACS. BTW, if the user send(copy to) the exam data including SR into PACS under Directory UI, the SR would be sent to PACS. | [12 empty] | 116602409 | In Progress | 454928 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not a defect | [5 empty] | Application |  |  |  | Closed |  |  | DSR Report | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
236 | 44165 | TRUE | Portugal | Centro Hospitalar de Vila Nova | 40030 | 44165 | 49 | Gantry tilt button repair | Gantry angulation button repair | P3 |  | GCP replaced | [12 empty] | 116568593 | New | 443925 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632181 | [5 empty] | Engineering |  |  |  | WIP |  |  | GCP | [8 empty] | System issue | System | Gantry-HW | Gantry Panel | Gantry-HW-Gantry Panel | Gantry-HW-Gantry Panel | CTUDT00632181 | Control Panel | Gantry-HW | Gantry-HW-Gantry Panel | No | y | y | y | y | y | No | y | y | Gantry Panel | y | y | y | y | y | y | y | y
237 | 44167 | TRUE | China | Huaying Mountain Guangneng General Hospital | 345000 | 44167 | 49 | 02N/Ingenuity Core128/Elite China/scan fault | 客户问题概述： 图像伪影。 客户问题详细描述或故障现象： 环形伪影 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 定位相竖条形伪影，螺旋为环形伪影 远程诊断与测试结果： 探测器模块故障 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 更换探测器模块，完成校准 建议备件12NC及名称： 459800911281module SAP Transfer : Dispatched for Onsite Support Next Step Action : PPD 四川华蓥市华蓥山广能集团总医院 24H FSE：李全飞-86-13699457740 探测器模块 TDMS V2 M.Det Class B FRU 459800911281 1 已处理 Next Step Action : li quanfei External Remarks : 04.12.2020 11:36:54 UTC Quanfei Li (86001471)故障现象: 扫描出现环形伪影，需要工程师到场维修 服务内容: 更换探测器模块后，设备恢复正常。 Internal Remarks : 04.12.2020 11:36:57 UTC Quanfei Li (86001471) $Q7: (Confirmed the reported problem,Image artifacts . ) $Q8: (Test the device.confirmed DMS-MODULE with problem.) $Q9: (replace DMS-MODULE（12NC：459800911281） and do calibration.) $Q10: (Test and inspection data refer to:PA. the device returned to full service.) $Q11: (Calibrated Tools name (N/A) ----------------------------------------------------------------------------- | P3 |  | Ring Artifact, replaced TDMS V2 M.Det Class B FRU 459800911281 收集SN， 额外信息根据DMS team反馈 | [12 empty] | 116581235 | Closed | 449621 | TRUE | Closed - Done | Closed - Done |  | CTUDT00632110 | [5 empty] | Engineering |  |  |  | WIP |  |  | Defective DMS module | [8 empty] | System issue | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | DMS Module | DMS-HW | DMS-HW-DMS Module | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
238 | 44167 | TRUE | China | Dalian Branch of Shengjing Hospital of China Medical University | 345001 | 44167 | 49 | 02N/Ingenuity Core128/Elite China/couch fault | 服务内容：CT BOX降床按键不灵敏，CT BOX问题，更换CT BOX。 服务结果：维修后，设备恢复正常工作。 | P3 |  | CT Box 459801704602 - CT BOX ASSEMBLY_BEST | [12 empty] | 116582303 | Closed | 450087 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632185 | [5 empty] | Engineering |  |  |  | WIP |  |  | CT Box | [8 empty] | System issue | System | Gantry-HW | CT Box | Gantry-HW-CT Box | Gantry-HW-CT Box | CTUDT00632185 | CT Box | Gantry-HW | Gantry-HW-CT Box | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
239 | 44170 | TRUE | China | SD Laiyang People Hospital | 345005 | 44170 | 49 | 02N/Ingenuity Core128/Elite China/Cannot access account | 服务内容：指导客户添加节点文件。 服务结果：维修后，设备恢复正常工作。 Internal Remarks : 07.12.2020 00:05:44 UTC Xingbo Yang (86038250) "RDI Accuracy:yes" $Q7: (Confirmed reported problem & host cannot visit the pacs server) $Q8: ( review lan config & need to add a remote host ) $Q9: ( instruct customer for net node confuguration ) $Q10: (Test and inspection data refer to: PA & Device returned to full functionality) $Q11: Calibrated tools name(NA) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date: 9 / 22 / 2020 ^Current Tube usage: scan second_29909 | P3 |  | Olevel account locked by too many failed try, suggest to reduce password complexity | [12 empty] | 116607000 | New | 455237 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not a defect | [5 empty] | Application |  |  |  | Closed |  |  | User account | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
240 | 44168 | TRUE | China | Hebei Laonianbing Hospital | 345015 | 44168 | 49 | 02N/Ingenuity Core128/Elite China/The test value is not up to standard and needs to be calibrated | Problem Description:third-party testing equipment is not up to standard Troubleshooting Action: new installation site, CT no problem Repair Action: site1 engineer feedback that, next time will be on site, adjusted the equipment Test/Inspection: Functional Check PASSED. The device returned to full functionality. Test Equipment: NA | P3 |  | Customer request FSE to support 3rd party test, CT no problem | [12 empty] | 116586764 | Closed | 452179 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not a defect | [5 empty] | Service | [6 empty] | Support 3rd party test | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
241 | 44170 | TRUE | China | SD Anqiu T.C.M Hosp. | 345017 | 44170 | 49 | 02N/Core128/Can not transfer image | (86038250)服务内容：restart spooler。 服务结果：维修后，设备恢复正常工作。 Next Step Action : 杨工已调查内容如下： 12/6到场了解到客户反映的问题主要如下： 1 扫描程序偶发死机现象，需要重新login恢复 2 客户界面偶发黄屏，需要重启host恢复 3 偶发本地图像无法传输到pacs等远程主机 TSE明确以上问题属于CT5000的BUG，escalation 已经建立，研发部门表示明年一季度做出升级。给到客户的临时的应急方案是每天强制重启主机1-2次最大限度避免死机问题，遇到传输问题时手动进行缓冲池清空。 就以上问题和应急方案的使用跟孙老师进行了充分说明，对方表示理解并保持观察 | P3 |  | Known issue, suggest to check and use recommended remote devices Setting by following service tips | [12 empty] | 116607368 | New | 457075 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00631602 | [5 empty] | Engineering |  |  |  | WIP |  |  | PACS | [8 empty] | System issue | System | Console-SW | Remote Connection-1602TBD | Console-SW-Remote Connection-1602TBD | Console-SW-Remote Connection | CTUDT00631602 | Remote Connection | Console-SW | Console-SW-Remote Connection-1602TBD | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
242 | 44175 | TRUE | China | SX Datong Yanggao People Hosp. | 345025 | 44175 | 50 | 02N/Ingenuity Core128/Elite China/Artifact | Troubleshooting Action: instruct customer to do air calibration Repair Action: instruct customer to do air calibration Test/Inspection: Functional Check PASSED. Device returned to full functionality. | P3 |  | Air Calibration | [12 empty] | 116634678 | Closed | 469491 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not A Defect | [5 empty] | Engineering |  |  |  | Closed |  |  | Artifact | [8 empty] | System issue | System | DMS-SW | Air Calibration | DMS-SW-Air Calibration | DMS-SW-Air Calibration | Not A Defect | Air Calibration | DMS-SW | DMS-SW-Air Calibration | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
243 | 44173 | TRUE | China | SX Datong Yanggao People Hosp. | 345025 | 44173 | 50 | 02N/Ingenuity Core128/Elite China/Cannot transmit data | BU Q&R Feedback: No SAP Customer Symptom Code: BOOT Customer Symptom Code: Boot up/Power-up Problem Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: Reboot/Restart Problem Reported by Customer:cannot transmit data Safety Questions:Case: 0116620331 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Chen YuanHui Last updated by: Chen YuanHui 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Field Service Engineer Notes:This PO is a Reference Number only T2 Activities: Problem Description: Cannot transmit data to pacs Troubleshooting Action: Cannot transmit data to pacs Repair Action: guide the customer reset spooler, solve the problem Test/Inspection: Functional Check PASSED. The device returned to full functionality. Test Equipment: NA ----------------------------------------------------------------------------- | P3 |  | Peter please help get bugreport | [12 empty] | 116620331 | Closed | 462639 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering |  |  |  | Closed |  |  | PACS | [8 empty] | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
244 | 44173 | TRUE | China | SX Datong Yanggao People Hosp. | 345025 | 44173 | 50 | 02N/Ingenuity Core128/Elite China/Image transmission failure | BU Q&R Feedback: No SAP Customer Symptom Code: IONW Customer Symptom Code: Interoperability/Network Issue Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: Instructed Customer Problem Reported by Customer:Image transmission failure Safety Questions:Case: 0116620490 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Zhang Jing Last updated by: Zhang Jing 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer_s problems: CT host image transmission failure. Detailed description of Customer_s problems or malfunctions: CT host image transmission failure. It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: Cannot transmit Remote troubleshoot and test result: Check image transfer, try reset Spooler PRS TR status: All passed through. Repair Action: reset spooler, then restore Test/Inspection: Use Normal Test Equipment: N/A ----------------------------------------------------------------------------- | P3 |  | PACS duplicate with 462639 | [12 empty] | 116620490 | Closed | 462732 | TRUE | Closed - Duplicate | Closed - Duplicate |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS | [8 empty] | System issue | system | Console-SW | Remote Connection | Console-SW-Remote Connection | Remote Connection | CTUDT00632127 | Remote Connection |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
245 | 44174 | TRUE | China | SD Chest Hosp. | 40025 | 44174 | 50 | 02N/Ingenuity CT 728326/power off | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : power off T2 Activities : Problem Description: 客户问题概述： 间歇性CT操作盒报错 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： CT BOX UNINITIATED RESET 远程诊断与测试结果： 已经更换CT操作盒，怀疑CRC电源问题 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议工程师检查更换电源 建议备件12NC及名称： 电源：459801032622，CT操作盒：453567027196 SAP Transfer : Case Transferred to SAP Next Step Action : PPD 453567027196，需要等国外解锁后才能订货，与TSE核实是否有其他替代号待回复 External Remarks : 11.12.2020 03:21:11 UTC SUP JCOUSER (86038405)服务内容：Gantry输入电压不稳，装机时teal接线错误，重新链接teal输出电源线。 服务结果：维修后，设备恢复正常。 Internal Remarks : 11.12.2020 03:21:14 UTC Zhaokun Huang (86038405) "RDI Accuracy: NO" $Q7: ( Confirmed reported problem & Estop open intermediate. ) $Q8: (Troubleshooting find Gantry input power unstable & Teal PDU output power cables connect incorrect ) $Q9: ( Reconnect Teal output power cables) $Q10: (Test and inspection data refer to: PA . Device returned to full functionality.) $Q11: Calibrated Tools name ( Multimeter Fluke 287 ) S/N (23350170) Next Calibration date( 04/16/2021 ) ^Current Tube installed date: N/A ^Current Tube usage: N/A Next Step Action : 018654536122 FSE Huang Zhaokun call in and dispatch him. SAP Transfer : Dispatched for Onsite Support Next Step Action : PPD CTBOX ASSEMBLY RAFI/TILTED (RoHS) 453567027195 1 CRC- Power Distribution Module 459801032622 1 已发24H到医院，FSE:黄兆坤-86-18654536122 Safety Questions : Case: 0116627925 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Guomei Li Last updated by: Guomei Li 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action : 已回访 ----------------------------------------------------------------------------- | P3 |  | PDU Power Cable connection | [12 empty] | 116627925 | Closed | 466578 | TRUE | Closed - Done | Closed - Done |  | Not A Defect | [5 empty] | Service |  |  |  | Closed |  |  | Cable loose | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
246 | 44174 | TRUE | United States | Odessa Regional Medical Center | 40029 | 44174 | 50 | 0708CFFD2 One Monitor Mode | Work Order #: WO-02564643 Maier, Donnie The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: No patient Involved Expected and Unexpected Results: Stuck in One Monitor Mode User Impact: No Patient Impact: No Current Software Version: 4.1 Diagnostic performed by Engineer : It appears they were rearranging monitors and system went into one monitor mode. I talked cu thru a complete power cycle of the system (Host/CIRS/Gantry) and now system is back to functioning properly with both monitors. Resolution : I talked cu thru a complete power cycle of the system (Host/CIRS/Gantry) and now system is functioning properly with both monitors and no longer in one monitor mode. | P3 |  | restart system | [12 empty] | 116626928 | Closed | 462291 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not A Defect | [5 empty] | Engineering |  |  |  | Closed |  |  | Monitor/adaptor | [8 empty] | System issue | System | Console-HW | Monitor | Console-HW-Monitor | Console-HW-Monitor | Not A Defect | Monitor | Console-HW | Console-HW-Monitor | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
247 | 44174 | TRUE | United States | Mount Grant General Hospital | 40012 | 44174 | 50 | 1102CDFD1 Cannot Energize CT/UPS Post-Emergency Stop | Work Order #: WO-02569122 Holeman, David The system does not meet specification and is out of use Problem Description by Engineer : Problem description by engineer : CT Tech reports he pressed Emergency OFF button and cannot re-energize CT System. When was the first occurrence? Within... : 09-Dec-2020 08:00:00 How often is issue occurring? : First occurrence Malfunction area : Others Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: Tech Cannot Energize CT/UPS Post-Emergency Stop Button Patient Impact/Outcome: Potential Delayed Scans User Impact: Cannot Energize CT/UPS Post-Emergency Stop Current Software Version: 4.1.7 Diagnostic performed by Engineer : Inquired status of wall breaker and Teal PDU breaker. Tech reports wall breaker is OFF and Teal PDU breaker is tripped. Advised Tech to reset the Teal breaker and turn ON (Teal input power green light = OFF). Advised Tech to reset wall breaker and turn ON (still no Teal light). Inquired if start/stop switches are present on the console wall = yes, but taped over - does not believe they are being used. Tech also states the CT has a Staco First Line UPS and is unsure how to start it/energize it - no quick guide or instructions present. Tech declines further remote support and is requesting FSE to call him and step him through the UPS and CT startup/energize procedure. RSE called FSE and left voicemail in attempt to expedite resolution. Internal Comments : Error # and/or description of error : Cannot Energize CT/UPS Post-Emergency Stop Button Completed or Recommended Repair Action "Internal" : Provide phone support to guide Tech on Staco First-Line UPS startup/energizing and CT system energizing. Provide quick guide of UPS and CT power-on for future reference. Follow Up Required : Error # and/or description of error : Cannot Energize CT/UPS Post-Emergency Stop Button Completed or Recommended Repair Action "Internal" : Tech declined further remote support and requested FSE to call site and provide phone support to guide Tech on Staco First-Line UPS startup/energizing and CT system energizing. Provide quick guide of UPS and CT power-on for future reference. ----------------------------------------------------------------------------- | P2 |  | Customer not familiar with how to Energize CT/UPS Post-Emergency Stop Button | [12 empty] | 116631485 | Closed | 464628 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not A Defect | [5 empty] | Customer |  |  |  | Closed |  |  | Mis-use | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
248 | 44176 | TRUE | United States | Odessa Regional Medical Center | 40029 | 44176 | 50 | 1299CDFD1 Maximum Hard Arcs & Short TC Frequency Inquiry | Work Order #: WO-02578475 McCain, Jimmie Problem Description by Engineer : Problem description by engineer : CT Tech reports "Maximum Hard Arcs Reached" message while scanning around 0730 on 12/10/20). When was the first occurrence? Within... : 10-Dec-2020 07:30:00 How often is issue occurring? : First occurrence Malfunction area : Tube&Gen Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: Scan Interrupt Due To Arcs Patient Impact/Outcome: None User Impact: Scan Interrupt Due To Arcs Current Software Version: 4.1.10 Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | P4 |  | Peter please help get bugreport | [12 empty] | 116640360 | In Progress | 468419 | TRUE | Closed - Done | Closed - Done |  | CTUDT00630085 | [5 empty] | Engineering |  |  |  | Closed |  |  | Tube Arc | [8 empty] | System issue | System | IC-HW | Tub Arc | IC-HW-Tub Arc | IC-HW-Tub Arc | CTUDT00630085 | Tub Arc | IC-HW | IC-HW-Tub Arc | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
249 | 44174 | TRUE | France | SCP GROUPE D IMAGERIE MEDICALE | 40019 | 44174 | 50 | Artifact. ACQ module is defective. | Artifact Faulty ACQ module Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Regulatory Questions : N/A Diagnostic performed by Engineer : N/A | P3 |  | 459800911271,TDMS V2 M.Det Class A FRU | [12 empty] | 116631946 | In Progress | 468008 | TRUE | Closed - Done | Closed - Done |  | TBD-ACQ Module | [5 empty] | Engineering |  |  |  | WIP |  |  | Defective DMS module | [8 empty] | System issue | System | Console-HW | ACQ-Module | Console-HW-ACQ-Module | Console-HW-ACQ-Module | TBD-ACQ Module | ACQ-Module | Console-HW | Console-HW-ACQ-Module | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
250 | 44172 | TRUE | Chile | CMD Clinica Integrales Rancagu | 40008 | 44172 | 50 | CL1675 - WEIRD NOISE WITHIN THE PATIENT STRETCHER. | Vertical table movement is lubricated, movement arms and vertical auger are verified. Suggest add lubricated in every 6 Month | P3 |  | Dec.3 SCE approved to cut in improved grease FRU 453566132481 on stock | [12 empty] | 116618168 | In Progress | 458940 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not a Defect | [5 empty] | Service |  |  |  | Closed |  |  | lubricate | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
251 | 44176 | TRUE | Japan | Kinan Hospital | 340000 | 44176 | 50 | CZ:CT アプリケーションサポートパッケージ | BU Q&R Feedback: No Reason: Application Support Problem Reported by Customer : CZ:CT アプリケーションサポートパッケージ Safety Questions : Case: 0116641941 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: ZHAO Cindy Last updated by: ZHAO Cindy 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A OneEMS internal remarks : 大谷DB SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 14.12.2020 01:50:50 UTC NATSUKI OOTANI (81100565)【ご依頼内容】 操作に関する質問と改善要望の聞き取り 【作業内容詳細】 心臓検査に関して説明とテストスキャンにて動作を確認していただきました。 改善要望の聞き取りを行いました。 【最終確認】 装置が正常に動作することを確認いたしました。 Internal Remarks : 14.12.2020 01:50:53 UTC NATSUKI OOTANI (81100565) 1.新型コロナウイルス感染患者もしくは感染の疑いがある患者に対して治療検疫 が行われていますか？ はい 2.感染患者のいる治療室および病床に入る必要がありますか？(作業場並びに移 動経路含む) いいえ ----------------------------------------------------------------------------- | P5 |  | application support ----- | [12 empty] | 116641941 | New | 473862 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not a Defect | [5 empty] | Application | [6 empty] | application support | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
252 | 44169 | TRUE | United States | Odessa Regional Medical Center | 40029 | 44169 | 49 | Exam Summary does not show the dose for the scan unless she has the SR report turned on. | Exam Summary does not show the dose for the scan unless she has the SR report turned on. Can not uncheck PACS as send to Destination if PACS is listed as a Default Destination. Exam Summary does not show the dose for the scan unless the SR report is turned on. Can not uncheck PACS as a send to Destination if PACS is listed as a Default Destination. |  |  | Checked on 4.1.10(see below), Dose info displayed in Exam summary after SR unchecked. Need screenshot or Bugrep from site for this issue. 2. Cannot deselect PACs as send to destination for Structured Dose Report if Pacs is selected as a default send to the destination. -----It WAS. 3. SR will not go to PACS -----It is an issue related to transition, need BugRep from site | [12 empty] | CFM |  | 467713 | TRUE | Closed - Done | Closed - Done |  | CTUDT00632121 | [5 empty] | Engineering |  |  |  | Closed |  |  | DSR Report | [8 empty] | System issue | System | Console-SW | Dose SR | Console-SW-Dose SR | Console-SW-Dose SR | CTUDT00632121 | Dose SR | Console-SW | Console-SW-Dose SR | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
253 | 44174 | TRUE | Spain | S.C.S. HOSP. GRAL. DE FUERTEVE | 40020 | 44174 | 50 | Dicom images export problem | System is experiencing intermittent Dicom images export problems. System exports images correctly to Philips Node(Intellispace portal) but it shows intermittent problems when exporting images to GE PACS or Siemens Work Station Syngovia. For some reason, queue to those nodes hangs up and system is not able to export images to them until Host computer is rebooted. It happens sometimes once a week, sometimes more and sometimes less frequently. After talking with other colleague, at least another system experienced same issue in Spain. |  |  | Peter please help get bugreport | [12 empty] | CFM |  | 466844 | TRUE | Closed - Done | Closed - Done |  | CTUDT00631602 | [5 empty] | Engineering |  |  |  | Closed |  |  | PACS | [8 empty] | System issue | System | Console-SW | Remote Connection-1602TBD | Console-SW-Remote Connection-1602TBD | Console-SW-Remote Connection | CTUDT00631602 | Remote Connection | Console-SW | Console-SW-Remote Connection-1602TBD | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
254 | 44144 | TRUE | China | SD Laiyang People Hospital | 345005 | 44144 | 46 | Image can not transfer to PACS | 患者因肺结节于11月7日入院，行胸部CT检查。11月8日打印胶片时，发现图像不能传输到打印界面，无法给患者打印胶片。重启电脑后仍不能解决，后清除打印内存，顺利完成打印 |  |  | Peter please help get bugreport | [12 empty] | CFM |  | 460334 | TRUE | Closed - Done | Closed - Done |  | CTUDT00632127 | [5 empty] | Engineering |  |  |  | Closed |  |  | PACS | [8 empty] | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
255 | 44175 | TRUE | Japan | Kinan Hospital | 340000 | 44175 | 50 | JS: Cannot connect via the RSN. | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : JS: Cannot connect via the RSN. | P5 |  | PRS Setup | [12 empty] | 116634723 | New | 469218 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not a defect | [5 empty] | Service |  |  |  | Closed |  |  | PRS Setup | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
256 | 44175 | TRUE | Portugal | Centro Hospitalar de Vila Nova | 40030 | 44175 | 50 | stop sending images to PACS forcing you to restart the equipment | The new TC has presented several problems The new TC of su has presented the following problems: - a "xiadeira" noise when the gantry is angled; - stop sending images to PACS forcing you to restart the equipment, resending is not enough; - "tube hot" information on the use of "abd / pelv" protocols with 120 kV and 140 kv as shown attached images. | P3 |  | Peter please help get bugreport | [12 empty] | 116636550 | New | 470009 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering |  |  |  | Closed |  |  | PACS | [8 empty] | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
257 | 44173 | TRUE | China | Guanzhou Panyu District Third | 345013 | 44173 | 50 | PRS Connection/02N/Core128 | 李宗贤-13751567086 开PRS 连接单 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 08.12.2020 10:15:08 UTC Zongxian Li (86001504)服务内容：安装远程诊断。 服务结果：工作正常。 | P5 |  | PRS Setup | [12 empty] | 116620747 | Closed | 459727 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not a defect | [5 empty] | Service | [6 empty] | PRS Setup | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
258 | 44173 | TRUE | United States | Central Vermont Medical Center | 40024 | 44173 | 50 | Q/O:Squeak in the gantry- | Ship 453567003451 or 453566492411 to do belt adjustment | P5 |  | belt adjustment | [12 empty] | 116619280 | In Progress | 462176 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | ECR-077012 | [5 empty] | Engineering |  |  |  | Closed |  |  | Belt | [8 empty] | System issue | System | Gantry-HW | Belt | Gantry-HW-Belt | Gantry-HW-Belt | ECR-077012 | Belt | Gantry-HW | Gantry-HW-Belt | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
259 | 44175 | TRUE | Germany | Gesundheit Nord gGmbh | 40002 | 44175 | 50 | Reconstruction images cannot be calculated | Reconstruction images cannot be calculated | P3 |  | Peter please help get bugreport | [12 empty] | 116637057 | Closed | 467235 | TRUE | Closed - Duplicate | Closed - Duplicate |  | CTUDT00632116 | [5 empty] | Engineering | [6 empty] | Recon | [8 empty] | System issue | System | Console-SW | TBD | Console-SW-TBD | Console-SW-TBD | CTUDT00632116 | TBD |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
260 | 44179 | TRUE | France | SCP GROUPE D IMAGERIE MEDICALE | 40019 | 44179 | 51 | Rings on the images. | Aritifact, analyse bad detector; update table; calibration | P3 |  | analyse bad detector; update table; calibration | [12 empty] | 116655928 | In Progress | 474298 | TRUE | Closed - Done | Closed - Done |  | CTUDT00632110 | [5 empty] | Engineering |  |  |  | Closed |  |  | Artifact | [8 empty] | System issue | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | DMS Module | DMS-HW | DMS-HW-DMS Module | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
261 | 44175 | TRUE | China | HEN Puyang People_s Hosp | 40031 | 44175 | 50 | The CT host cannot transmit image to PACS | CT__Product Complaint__Ingenuity CT The host cannot_transmit_image_to_PACS__HEN Puyang People_s Hosp The CT host cannot transmit image to PACS. In the Queue manager view the transmit state is always in a waiting in Queue state and not have transmit speed. Frequency of issue:always. Action:Reconfigure the PACS DICOM nodes ,Reset the spooler,the transmit issue still exist. |  |  | Peter please help get bugreport | [12 empty] | CFM |  | 465836 | TRUE | Closed - Duplicate | Closed - Duplicate |  | CTUDT00631602 | [5 empty] | Engineering |  |  |  | Closed |  |  | PACS | [8 empty] | System issue | system | Console-SW | Remote Connection-1602 | Console-SW-Remote Connection-1602 | Remote Connection | CTUDT00631602 | Remote Connection |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
262 | 44179 | TRUE | China | SX Datong Yanggao People Hosp. | 345025 | 44179 | 51 | The CT host cannot transmit image to PACS. | The CT host cannot transmit image to PACS. In the Queue manager view the transmit state is always in a waiting in Queue state and not have transmit speed. Frequency of issue:always. Action:Reconfigure the PACS DICOM nodes ,Reset the spooler,the transmit issue still exist. |  |  | Peter please help get bugreport | [12 empty] | CFM |  | 473938 | TRUE | Closed - Done | Closed - Done |  | CTUDT00632127 | [5 empty] | Engineering |  |  |  | Closed |  |  | PACS | [8 empty] | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
263 | 44174 | TRUE | China | SD Chest Hosp. | 40025 | 44174 | 50 | WeChat:02N/Ingenuity CT 728326/Add IP address | T2 Activities : Problem Description: add pacs Troubleshooting Action: add pacs Repair Action: remote add pacs , solve the problem Test/Inspection: Functional Check PASSED. Device returned to full functionality. SAP Transfer : Case Transferred to SAP | P3 |  | PACS Setup remotely | [12 empty] | 116628065 | Closed | 466092 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not a defect | [5 empty] | Service | [6 empty] | PACS Setup | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
264 | #N/A | #N/A | Portugal | Centro Hospitalar de Vila Nova | 40030 | 44097 | 39 | CD-0000031911 Centro Hospitalar de Vila Nova Rua Conceicao Fernandes Vila Nova De Gaia 13 4434-502 Portugal | 459801386961 - CABLE, IBOX2 P7 TO USB EXTENDER DC POWER | P5 | [14 empty] | 116182544 | Closed | DEFOA CD-31911 | #N/A | #N/A | DEFOA |  | Not a defect | [5 empty] | Operation | [6 empty] | DEFOA | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
265 | #N/A | #N/A | China | JL Songyuan Central Hosptial | 345008 | 44184 | 51 | Blue phantom cracked | CD-0000033107 |  |  | resend Phantom to field | [14 empty] | DEFOA CD-33107 | #N/A | #N/A | DEFOA |  | Not a defect | [5 empty] | Operation | [6 empty] | DEFOA | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
266 | 44182 | TRUE | United States | The Regional Medical Center | 40005 | 44182 | 51 | 0698CDFM2 Beeping from the CT box | Atz, Michael The system partly meets specification and is returned to use with restrictions Beeping from the CT box customer states that this has happened before Diagnostic performed by Engineer : Beeping from the CT box customer states that this has happened before advise to replace the CRC power supply and the CT box post troubleshooting Internal Comments : Error # and/or description of error : Beeping from the CT box customer states that this has happened before advise to replace the CRC power supply and the CT box post troubleshooting please check part numbers and order the ones that pertain to your system CT Box ingenuity CRC- Power Distribution Module 459801032622 -- CTBOX ASSEMBLY RAFI/TILTED (RoHS) 453567027195 Follow Up Required : Error # and/or description of error : Beeping from the CT box customer states that this has happened before Beeping from the CT box customer states that this has happened before advise to replace the CRC power supply and the CT box post troubleshooting please check part numbers and order the ones that pertain to your system CT Box ingenuity CRC- Power Distribution Module 459801032622 -- CTBOX ASSEMBLY RAFI/TILTED (RoHS) 453567027195 | P2 |  | When the ICT noise reduction board cut in to CT5000? 459800367081 Can replace the noise reduction board? Is it possible to return to Suzhou? | [12 empty] | 116697683 | In Progress | 483939 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632185 | [5 empty] | Engineering |  |  |  | WIP |  |  | CT Box | [8 empty] | System issue | System | Gantry-HW | CT Box | Gantry-HW-CT Box | Gantry-HW-CT Box | CTUDT00632185 | CT Box | Gantry-HW | Gantry-HW-CT Box | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
267 | 44184 | TRUE | United States | Lackey Memorial Hospital | 40015 | 44184 | 51 | Low dose chest screening | Work Order #: WO-02649555 Mujica, David The system fully meets specification and returned to use Problem description by Clinical Support Specialist : Customer Function/Role: Technologist Device Use: In use Expected and Unexpected Results: Low Dose Chest questions User Impact: none Patient Impact: none Current Software Version: 4.1 Diagnostic performed by Engineer : Rich and Michae were aided in questions on a low dose CT chest screening. Michae was also sent information on low dose chest screenings. Applications will be on site to aide in construction of this Exam Card next week. No further support needed at this time. | P5 |  | Applications will be on site to aide in construction of this Exam Card next week. No further support needed at this time. | [12 empty] | 116709413 | Closed | 487699 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not a defect | [5 empty] | Application | [6 empty] | Application support | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
268 | 44184 | TRUE | United States | Lackey Memorial Hospital | 40015 | 44184 | 51 | how to tilt gantry |  | P4 |  | not clear | [12 empty] | 116708798 | New | 489238 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD_OBS | [5 empty] | Application |  |  |  | WIP |  |  | Not clear | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
269 | 44182 | TRUE | United States | McLaren Macomb | 40021 | 44182 | 51 | NEMO JIM CORDS | Parts Consumed : 1.00000,459800484501,CONTROL PANEL ASSEMBLY-FRONT LH | P5 |  | 459800484501,CONTROL PANEL ASSEMBLY-FRONT LH | [12 empty] | 116698434 | Closed | 487099 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | CTUDT00632181 | [5 empty] | Engineering | [6 empty] | GCP | [8 empty] | System issue | System | Gantry-HW | Gantry Panel | Gantry-HW-Gantry Panel | Gantry-HW-Gantry Panel | CTUDT00632181 | Control Panel |  |  | No | y | y | y | y | y | No | y | y | Gantry Panel | y | y | y | y | y | y | y | y
270 | 44182 | TRUE | United States | Odessa Regional Medical Center | 40029 | 44182 | 51 | Noise Heard when Gantry Rotates | Work Order #: WO-02636487 Paine Matthew The system does not meet specification and is out of use Problem Description by Engineer : Customer Function/Role: Technologist Device Use: Diagnostic Expected and Unexpected Results: Noise Heard when Gantry Rotates User Impact: Unknown Patient Impact: Unknown Current Software Version:N/A Diagnostic performed by Engineer : Customer states that when Gantry Rotates a Noise Can be Heard Internal Comments : Contacted Primary FSE Jimmie McCain. He was going to reschedule his PM to take the call. Follow Up Required : Customer Needs Service ASAP. Noise can be heard from Gantry when it is Rotating. | P4 |  | adjustment of lexan ring (betweem collimator) Confirmed noise from gantry. Filed area of cover that was rubbing on the a-plane collimator, closed covers and confirmed proper clearance. | [12 empty] | 116697736 | Closed | 483967 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not a defect | [5 empty] | Engineering |  |  |  | Closed |  |  | Gantry Noise | [8 empty] | System issue | System | Gantry-HW | Noise | Gantry-HW-Noise | Gantry-HW-Noise | Not a defect | Noise | Gantry-HW | Gantry-HW-Noise | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
271 | 44181 | TRUE | Portugal | Centro Hospitalar de Vila Nova | 40030 | 44181 | 51 | O equipamento falha intermitentemente quande se faz o bolus traking manual. | The equipment fails intermittently when manual bolus traking is performed. Problem Description by Engineer: Customer Function / Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer: The reported problem seems to be related to the communication with the injector. The description manifests itself intermittently. Since the existing injector is not connected to the Tac, the SAS option was disabled. If the problem does not appear, it is considered solved. | P3 |  | The reported problem seems to be related to the communication with the injector. | [12 empty] | 116690347 | In Progress | 484112 | TRUE | Closed - Done | Closed - Done |  | Not a defect | [5 empty] | Engineering |  |  |  | Closed |  |  | Cable loose | [8 empty] | System issue | System | Gantry-HW | Cable loose | Gantry-HW-Cable loose | Gantry-HW-Cable loose | Not a defect | Cable loose | Gantry-HW | Gantry-HW-Cable loose | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
272 | 44181 | TRUE | China | HEN Puyang People_s Hosp | 40031 | 44181 | 51 | A DMS module can not work correctly within a month since the new system installation. |  |  |  | Need get SN of module, BD calibration result, file location xxxxx | [14 empty] | 479232 | TRUE | Closed - Done | Closed - Done |  | CTUDT00632110 | [5 empty] | Engineering |  |  |  | Closed |  |  | Defective DMS module | [8 empty] | System issue | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | DMS Module | DMS-HW | DMS-HW-DMS Module | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
273 | 44180 | TRUE | China | HEN Puyang People_s Hosp | 40031 | 44180 | 51 | Ring artifacts and error reporting, replace detector module | 客户问题概述： 图像伪影。 客户问题详细描述或故障现象： 胸部扫描外环见一带状环形伪影 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 带状环形伪影，指导客户做空气校正无效 远程诊断与测试结果： 室内温度或探测器不稳定 PRS TR状态： 不通，请工程师到场检查。 Repair Action: 建议现场维修步骤： 1，到场检查室内温度2，做空气校正3，更换探测器 Test/Inspection: n/a Test Equipment: n/a External Remarks : 16.12.2020 13:10:41 UTC SUP JCOUSER (86038950)服务内容：客户报修设备有伪影，到场检查发现环形伪影，检查发现3号探测器模块故障。 服务结果：更换3号探测器模块，设备恢复正常 | P3 |  | module was damaged within one month after handover to customer, and around 50 patients/day • System installed at Nov-27. (BD calibration files are attached • The DMS module 3 was failed BD calibration at Dec-15 because of customer report ring artifacts • After DMS module replaced the new BD calibration was performed and passed at Dec-16 The defective module SN is 017880 and returned to repair center. Milstein, Dmitry <dmitry.milstein@philips.com> I see that the whole tile is disconnected from module. This is not typical failure mode of single bad detector. We will follow up with repair center to check what the root cause might be | [12 empty] | 116662071 | In Process | 479060 | TRUE | Closed - Done | Closed - Done |  | CTUDT00632110 | [5 empty] | Engineering |  |  |  | Closed |  |  | Defective DMS module | [8 empty] | System issue | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | DMS Module | DMS-HW | DMS-HW-DMS Module | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
274 | 44181 | TRUE | Japan | Kinan Hospital | 340000 | 44181 | 51 | 03DS: KKI: Mr. Kadosaka: Simple scan stopped partway through >> S_GENERATOR_MONITOR_KV_PERFORMANCE_ERROR > T [TRUNCATED] | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 03DS:KKI:ｶﾄﾞｻｶ様:単純撮影中にスキャンストップ発生＞＞S_GENERATOR_MONITOR_KV_PERFORMANCE_ERROR＞継続して検査可 Safety Questions : Case: 0116686543 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: KAZUMASA KIMURA Last updated by: KAZUMASA KIMURA T2 Activities : Problem Description: Problem description by engineer : 対象装置:CT 故障個所:03High Voltage Energy 最初に発生したのはいつか?(~以内):1 Hour 問題発生の頻度は?:First Occurrence エンドユーザｰのワークフローのどの時点において問題が発生しているのか?:検査・測定中(During Acqustion/Scan) 最近どのようなイベントが発生したか?:N.A. 顧客は問題について具体的なシステムイベントレポートを作成しているか?:いいえ。お客様にログファイル、レポートは作成して頂けませんでした。 Troubleshooting Action: What steps are taken during the system diagnosis : 対応後の結果:改善された 故障個所:03High Voltage Energy 診断中にどのようなステップを実施したか?: 【コール内容】 単純撮影中にスキャンストップ発生 一昨日:2回、昨日:2回、本日:3回発生したようです。 【対応】 継続して検査可 【ログ確認結果】 以下のエラーが発生しております。 S_GENERATOR_MONITOR_KV_PERFORMANCE_ERROR 高圧系のエラーが頻発して発生しているようです。 Error # and/or description of error : N/A Repair Action: Completed or Recommended Repair Action "External" : サービス担当の訪問作業。 OneEMS internal remarks : Problem description by engineer : 対象装置:CT 故障個所:03High Voltage Energy 最初に発生したのはいつか?(~以内):1 Hour 問題発生の頻度は?:First Occurrence エンドユーザｰのワークフローのどの時点において問題が発生しているのか?:検査・測定中(During Acqustion/Scan) 最近どのようなイベントが発生したか?:N.A. 顧客は問題について具体的なシステムイベントレポートを作成しているか?:いいえ。お客様にログファイル、レポートは作成して頂けませんでした。 What steps are taken during the system diagnosis : 対応後の結果:改善された 故障個所:03High Voltage Energy 診断中にどのようなステップを実施したか?: 【コール内容】 単純撮影中にスキャンストップ発生 一昨日:2回、昨日:2回、本日:3回発生したようです。 【対応】 継続して検査可 【ログ確認結果】 以下のエラーが発生しております。 S_GENERATOR_MONITOR_KV_PERFORMANCE_ERROR 高圧系のエラーが頻発して発生しているようです。 Error # and/or description of error : N/A Completed or Recommended Repair Action "External" : サービス担当の訪問作業。 Completed or Recommended Repair Action "Internal" : --Troubleshooting Action-- 対応後の結果:改善された 故障個所:03High Voltage Energy 診断中にどのようなステップを実施したか?: 【コール内容】 単純撮影中にスキャンストップ発生 一昨日:2回、昨日:2回、本日:3回発生したようです。 【対応】 継続して検査可 【ログ確認結果】 以下のエラーが発生しております。 S_GENERATOR_MONITOR_KV_PERFORMANCE_ERROR 高圧系のエラーが頻発して発生しているようです。 --Repair Action-- Recommended 12Nc: PR1 459800638522:RHOST SG164 ASSEMBLY with CPM (RoHS) QTY:1 SLOC:大阪(U615) PR1 453567560022:80 KW System Control Board QTY:1 SLOC:大阪(U615) PR1 453567560012:Spellman 80 KW RoHS APM Assy# QTY:1 SLOC:大阪(U615) PR1 453567560002:Spellman 80 KW RoHS CPM Assy QTY:1 SLOC:BackOrder --Subject-- 03DS:KKI:ｶﾄﾞｻｶ様:単純撮影中にスキャンストップ発生＞＞S_GENERATOR_MONITOR_KV_PERFORMANCE_ERROR＞継続して検査可 --Problem Description-- 対象装置:CT 故障個所:03High Voltage Energy 最初に発生したのはいつか?(~以内):1 Hour 問題発生の頻度は?:First Occurrence エンドユーザｰのワークフローのどの時点において問題が発生しているのか?:検査・測定中(During Acqustion/Scan) 最近どのようなイベントが発生したか?:N.A. 顧客は問題について具体的なシステムイベントレポートを作成しているか?:いいえ。お客様にログファイル、レポートは作成して頂けませんでした。 ----------------------------------------------------------------------------- | P4 |  | ask FSE to collect bugreport S_GENERATOR_MONITOR_KV_PERFORMANCE_ERROR | [12 empty] | 116686543 | In Process | 479926 | TRUE | Closed - Done | Closed - Done |  | TBD_KV_Perf_Bugrep | [5 empty] | Engineering |  |  |  | WIP |  |  | HV Generator | [8 empty] | System issue | System | IC-HW | APM/CPM | IC-HW-APM/CPM | IC-HW-APM/CPM | TBD_KV_Perf_Bugrep | APM/CPM | IC-HW | IC-HW-APM/CPM | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
275 | 44179 | TRUE | Japan | Kinan Hospital | 340000 | 44179 | 51 | JS: CT Application support package. | BU Q&R Feedback: No Reason: Application Support Problem Reported by Customer : JS: CT Application support package. OneEMS internal remarks : Otani DB Safety Questions:Case: 0116653279 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: SHA Joy Last updated by: SHA Joy | P5 |  | Application support | [12 empty] | 116653279 | New | 475980 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not a defect | [5 empty] | Application | [6 empty] | Application support | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
276 | 44181 | TRUE | Poland | Uniwersytecki Szpital Kliniczny | 341001 | 44181 | 51 | Installation completed + tests. | Work Order #: WO-02627360 18-Dec-2020 Gerard Dąbrówka The system fully meets specification and returned to use Problem Description by Engineer : asysta podczas testów odbiorczych Resolution : Testy wykonane przez Synektik z wynikiem pozytywnym | P4 |  | Service support acceptance test | [12 empty] | 116688919 | Incomplete | 483628 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not a defect | [5 empty] | Service | [6 empty] | 3rd party test | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
277 | 44180 | TRUE | China | SD Laiyang People Hospital | 345005 | 44180 | 51 | 02N/Ingenuity Core128/Elite China/There is no linkage between the equipment and the protective door | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Instructed Customer SAP Transfer:Case Transferred to SAP Problem Reported by Customer: There is no linkage between the equipment and the protective door Repair Action: Suggest onsite service procedures: According to the hospital_s request, cooperate with the hospital to connect the gate control switch Test/Inspection: n/a Test Equipment: n/a SAP Transfer:Dispatched for On site Support | P5 |  | Door switch connection | [12 empty] | 116663866 | In Process | 479262 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not a defect | [5 empty] | Service | [6 empty] | Door switch | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
278 | 44179 | TRUE | China | HLJ Harbin Senhai Hospital | 345007 | 44179 | 51 | 02N/Core128/Connect PACS | 张先生 18600090656 自称是飞利浦的售前销售，告知医院的5台设备都要连接PACS，但是时间都还没有和PACS厂家协商好， 需要工程师先和院办的罗主任13224606999 先电话沟通确认情况 Next Step Action : FSE：靳军13796075213 告知今天可以开SWO给他 Next Step Action : NTD：与FSE 核实是否需要今天就开单处理 FSE Notes : This PO is a Reference Number only T2 Activities : Problem Description: 客户问题概述： 图像传输失败。 客户问题详细描述或故障现象： 主机电脑需要和pacs进行连接。 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 主机电脑需要和pacs进行连接。 远程诊断与测试结果： 已和工程师确认，明天连接PACS。 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议明天RSE先与客户进行连接，如果需要到场地更改设置，必要时到场地更改。 建议备件12NC及名称： N/A更改主机IP，添加pacs，打印机，worklist 服务结果:设备恢复正常 | P5 |  | NAC, It is only for DICOM device connection | [12 empty] | 116655856 | Closed | 476747 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not a defect | [5 empty] | Service | [6 empty] | PACS config | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
279 | 44179 | TRUE | China | Hebei Laonianbing Hospital | 345015 | 44179 | 51 | 02N/Ingenuity Core128/Elite China/Insufficient uniformity | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: Insufficient uniformity Safety Questions:Case: 0116656167 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Qi Xiaoyu Last updated by: Qi Xiaoyu 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer_s problems: Other types of error. Detailed description of Customer_s problems or malfunctions: The hospital asked to go to the hospital to assist status monitoring It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: The hospital asked to go to the hospital to assist status monitoring Remote troubleshoot and test result: The hospital asked to go to the hospital to assist status monitoring PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: The engineer is going to the hospital to assist the inspection SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support External Remarks: 15.12.2020 02:14:52 UTC Jianping Sun (86039023) service details: [cooperate with metrology testing.] Service results: The device returned to normal working conditions.] Internal Remarks: 18.12.2020 04:27:57 UTC Jianping Sun (86039023) just standby, no CM action related. ----------------------------------------------------------------------------- | P3 |  | 配合计量检测 | [12 empty] | 116656167 | In Process | 479107 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not a defect | [5 empty] | Service | [6 empty] | 3rd party test | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
280 | 44182 | TRUE | China | SD Anqiu T.C.M Hosp. | 345017 | 44182 | 51 | 02N/Ingenuity Core128/Elite China/power off | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : power off Safety Questions : Case: 0116695036 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liangyun Geng Last updated by: Liangyun Geng 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 间歇性机架掉电 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： S_COUCHMGR_HOR_MOTION_STOP_ERROR 远程诊断与测试结果： 怀疑床引起机架掉电 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议工程师做床水平运动校准，检查水平运动阻力，编码器，耦合器。。 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 20.12.2020 18:42:49 UTC Xingbo Yang (86038250)服务内容：调节水平运动阻力。 服务结果：维修后，设备恢复正常工作。 Internal Remarks : 20.12.2020 18:42:52 UTC Xingbo Yang (86038250) "RDI Accuracy:yes" $Q7: (Confirmed reported problem & e-stop open ) $Q8: ( review error logger "S_COUCHMGR_HOR_MOTION_STOP_ERROR "& high horizontal resistance) $Q9: ( clean leadscrew drive rod) $Q10: (Test and inspection data refer to: PA & Device returned to full functionality) $Q11: Calibrated tools name(NA) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date: 09 / 24 / 2020 ^Current Tube usage: scan second_42049 ----------------------------------------------------------------------------- | P3 |  | ask FSE to collect bugreport "S_COUCHMGR_HOR_MOTION_STOP_ERROR" | [12 empty] | 116695036 | In Process | 485771 | TRUE | Closed - Done | Closed - Done |  | CTUDT00632189 | [5 empty] | Engineering |  |  |  | Closed |  |  | Horizontal movement | [8 empty] | System issue | System | Couch-HW | Estop open | Couch-HW-Estop open | Couch-SW-Horizontal Move Block | CTUDT00632189 | Estop open | Couch-HW | Couch-HW-Estop open | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
281 | 44190 | TRUE | China | SD Anqiu T.C.M Hosp. | 345017 | 44190 | 52 | 02N/Ingenuity Core128/Elite China/Can not power on gantry | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : Can not power on gantry Safety Questions : Case: 0116739559 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Yuan Mingze Last updated by: Yuan Mingze 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support T2 Activities : Problem Description: 客户问题概述： ESTOP断开。 客户问题详细描述或故障现象： 机架打角度过程中掉电 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： S_TILT_C_NO_MOVE 远程诊断与测试结果： 间接性打角度过程中机架掉电，结合现象及logger，考虑Tilt POT故障 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 现场进一步检查故障原因，1 检查Tilt POT是否转动顺畅；2 检查gantry motion power board给tilt电机供电及接线；3 检查GMP 至gantry motion power board tilt enable信号；4 检查tilt电机及各限位开关；5 必要时更换tilt pot并观察 建议备件12NC及名称： "459800441121 DUAL POT ASSEMBLY" 客户需求备注： Please Dispatch Next Step Action : PPD DUAL POT ASSEMBLY 459800441121 1 已安排24H到医院，FSE：杨兴波-86-13583235930 External Remarks : 26.12.2020 15:25:03 UTC Xingbo Yang (86038250)服务内容：执行倾斜校正。 服务结果：维修后，设备恢复正常工作。 Internal Remarks : 26.12.2020 15:25:06 UTC Xingbo Yang (86038250) "RDI Accuracy:yes" $Q7: (Confirmed reported problem & e-stop open when gantry tilt movement) $Q8: ( review error logger "S_TILT_C_NO_MOVE "& doublt that It is a matter of relative tilt) $Q9: ( perform tilt calibration ) $Q10: (Test and inspection data refer to: PA & Device returned to full functionality) $Q11: Calibrated tools name(NA) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date: 10/ 23 / 2020 ^Current Tube usage: scan second_53945 ----------------------------------------------------------------------------- | P3 |  | PPD DUAL POT ASSEMBLY 459800441121 ECR is preparing to control the assembly process in factory. | [12 empty] | 116739559 | Closed | 501607 | TRUE | Closed - Done | Closed - Done |  | CTUDT00632190 | [5 empty] | Engineering |  |  |  | Closed |  |  | Dual Pot | [8 empty] | System issue | System | Gantry-HW | Estop | Gantry-HW-Estop | Gantry-HW-Estop | CTUDT00632190 | Estop | Gantry-HW | Gantry-HW-Estop | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
282 | 44190 | TRUE | China | SD Chest Hosp. | 40025 | 44190 | 52 | wechat:02N/Ingenuity CT 728326/The host is not connected to the workstation | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : The host is not connected to the workstation Safety Questions : Case: 0116739097 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Qi Xiaoyu Last updated by: Qi Xiaoyu 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： CT操作台不能传输图像到工作站 设备本身技术问题： 是。 Troubleshooting Action: 远程诊断与测试结果： 不能远程reset spooler，联系工程师黄兆坤，到场检查网络 PRS TR状态： 不通，请工程师到场检查。 Repair Action: 建议现场维修步骤： 建议工程师到场检查网络，reset spooler。。。 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 26.12.2020 13:53:04 UTC SUP JCOUSER (86038405)服务内容：CT到ISP传图失败，远程协助reset spooler。 服务结果：远程指导后，机器恢复正常。 Internal Remarks : 27.12.2020 11:45:54 UTC Zhaokun Huang (86038405) "RDI Accuracy: Yes." $Q7: ( Confirmed reported problem & Image from host transfer to ISP failed ) $Q8: ( Support customer reset spooler & It‘s Bug about CT5000.) $Q9: ( Reset spooler.) $Q10: (Test and inspection data refer to: No PA Required . Device returned to full functionality.) $Q11: Calibrated Tools name ( N/A） ^Current Tube installed date: N/A ^Current Tube usage: N/A ----------------------------------------------------------------------------- | P3 |  | 1st report PACS issue, reset spooler remotely by TSE, 100 patients per day, already guide customer how to reset by themselves. No bugreport | [12 empty] | 116739097 | Closed | 501165 | TRUE | Closed - Done | Closed - Done |  | CTUDT00632127 | [5 empty] | Engineering |  |  |  | Closed |  |  | PACS | [8 empty] | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
283 | 44190 | TRUE | China | SD Chest Hosp. | 40025 | 44190 | 52 | Complaint/Ingenuity CT 728326/TSE wengjinming | BU Q&R Feedback: No Reason: Customer Information Safety Questions : Case: 0116739535 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Xie Lili Last updated by: Xie Lili 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Problem Reported by Customer : Complaint TSE wengjinming Next Step Action : "12/25:TFM通知RSE Mgr刘国平跟进 翁工反馈之前在和FSE沟通故障内容" Next Step Action : 12/25:TFM通知RSE Mgr刘国平跟进 翁工反馈之前在和FSE沟通故障内容 12/28：CZH回访客户确认，对我司后续处理表示满意，同意关闭投诉 ----------------------------------------------------------------------------- | P3 |  | E-Stop open during unload patient, CAN Trace/logger will be shared by FSE, Huang zhaokun. Under Changxian investigation 1: check the AMC driver cable connection; 2: do several time unload button function to try to reproduce; 3: if still happened change the AMC driver. | [12 empty] | 116739535 | Closed | 501147 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | CTUDT00632187 | [5 empty] | Engineering | [6 empty] | AMC/CCB | [8 empty] | System issue | system | Couch-SW | Estop | Couch-SW-Estop | Couch-SW-Estop | CTUDT00632187 | Estop |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | 0 | y
284 | 44190 | TRUE | China | General Hospital of Jizhong Energy Jingxing Mining Group Co. LTD | 345003 | 44190 | 52 | Wechat：02N/Ingenuity Core128/Elite China/cannot boot up reconstruction | BU Q&R Feedback: No SAP Customer Symptom Code: HANG Customer Symptom Code: Hang/Crash Reason: Corrective Maintenance Engineer Failure Code: Wear-out/Breakage/Mechanical Failure Repair Activity Code: Reboot/Restart Safety Questions : Case: 0116738968 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Tao Fengming Last updated by: Tao Fengming 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Problem Reported by Customer : cannot boot up reconstruction T2 Activities : Problem Description: gantry can not expose. Troubleshooting Action: Instruct customer to restart gantry&rack. Repair Action: Instruct customer to restart gantry&rack. Test/Inspection: customer confirm can scan patient and image ok.functional Check PASSED. Device returned to full functionality. Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | P3 |  | restart gantry&rack | [12 empty] | 116738968 | Closed | 500937 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD_OBS | [5 empty] | Engineering |  |  |  | WIP |  |  | Reboot | [8 empty] | System issue | System | Gantry-HW | Restart | Gantry-HW-Restart | Gantry-HW-Restart | TBD_OBS | Restart | Gantry-HW | Gantry-HW-Restart | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
285 | 44190 | TRUE | China | SD Chest Hosp. | 40025 | 44190 | 52 | No timely response |  | NA |  | Not a Product Complaint | [14 empty] | 500175 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not a defect | [5 empty] | Others | [15 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
286 | 44188 | TRUE | China | Zhenjiang M.C.A. Hosp. | 345023 | 44188 | 52 | Wechat:02G/Ingenuity Core128/Elite China/Door machine linkage control fault | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer : Door machine linkage control fault Safety Questions : Case: 0116730913 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Bai Jingwen Last updated by: Bai Jingwen 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action : 与MM核实石松松 18868112815跟进，石工电话未接通，已短信告知确认是否需要加op Next Step Action : 在IA #52717502上加operation，派工FSE Shi Songsong Next Step Action : OP已加，0030 ----------------------------------------------------------------------------- | P3 |  | Door machine linkage control fault | [12 empty] | 116730913 | Closed | 499680 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not a defect | [5 empty] | Service |  |  |  | Closed |  |  | Door switch | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
287 | 44188 | TRUE | China | GD Dongguan Dalingshan Hosp. | 345013 | 44188 | 52 | 02N/Ingenuity Core128/Elite China/Change the parameters | BU Q&R Feedback: No SAP Customer Symptom Code: CENV Customer Symptom Code: Customer Environmental Concern Reason: Corrective Maintenance Engineer Failure Code: Environmental Cause Repair Activity Code: Configuration Change Problem Reported by Customer : Change the parameters Safety Questions:Case: 0116730461 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Qian Wang Last updated by: Qian Wang 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities: Troubleshooting Action: instruct customer to change the parameter Repair Action: instruct customer to change the parameter Test/Inspection: Functional Check PASSED. The device returned to full functionality. Test Equipment: N/A SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | P3 |  | Unclear what parameter have changed | [12 empty] | 116730461 | Closed | 498824 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not a defect | [5 empty] | Customer | [15 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
288 | 44186 | TRUE | China | SD Laiyang People Hospital | 345005 | 44186 | 52 | 02N/Ingenuity Core128/Elite China/report error | 客户问题概述： 扫描失败。 客户问题详细描述或故障现象： 客户反馈机架扫描时报错，需重新扫描。 服务内容：校正床板位置。 服务结果：维修后，设备恢复正常工作。 | P3 |  | 扫描失败。 校正床板位置。 | [12 empty] | 116715853 | Closed | 495369 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-Calibration | [5 empty] | Engineering |  |  |  | WIP |  |  | Couch calibration | [8 empty] | System issue | System | Gantry-HW | Calibration | Gantry-HW-Calibration | Gantry-HW-Calibration | TBD-Calibration | Calibration | Gantry-HW | Gantry-HW-Calibration | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
289 | 44188 | TRUE | Japan | Saiseikai Gosu Hospital | 343000 | 44188 | 52 | It could not provide the MPR by using the divided images. | This issue was occurred when we were using Ingenuity CT Version4.1.10. In derectory, we chosed thin slice brain image obtained by non-helical scan and click on CT Viewer. Then we found that the images were divided into three series. We couldn_t make MPR images entirely by the divided images. On the series list which was located left part of the screen, we could select Merge by right mouse click. By choosing Merge, the image covered whole brain and we could create MPR images entirely. But the thin slice image should be displayed entirely without choosing Merge. It means that if we choose thin slice image obtained by non-helical scan and open CT Viewer, the entire brain image should be displayed on the viewport. This issue didn_t occur everytime we conducted non-helical scan. The image in which problem occurred was obtained by non-helical scan with tilting. This problem took place on Dec.21 and we came to know about it on the same day. |  |  | Need bugreport to do analysis @Peter | [14 empty] | 494951 | TRUE | Closed - Done | Closed - Done |  | CTUDT00632116 | [5 empty] | Engineering |  |  |  | WIP |  |  | MPR | [8 empty] | System issue | System | CIRS-SW | MPR Recon | CIRS-SW-MPR Recon | CIRS-SW-MPR Recon | CTUDT00632116 | MPR Recon | CIRS-SW | CIRS-SW-MPR Recon | No | y | Console-SW | y | Console-SW-MPR Recon | Console-SW-MPR Recon | No | y | y | y | Console-SW | Console-SW-MPR Recon | y | y | y | y | y | y
290 | 44187 | TRUE | Puerto Rico | METRO MAYAGUEZ INC | 40009 | 44187 | 52 | 0116720221 ANILLOS EN LA IMAGEN | Work Order #: WO-02660189 21-Dec-2020 Hernandez, Alvin The system does not meet specification and is out of use Problem Description by Engineer : Problem Reported: Customer reported Ring Artifact during patients study images. Device Use: Device used for regular anatomy diagnostics. Expected and Unexpected Results: Patient studies can not be made User Impact: System inoperative. Current Software Version:4.1.10 Follow Up Required : Detector Module 22 & 23 are defetive Resolution : The contract solution was cleaned from in side the gantry parts. Still there is a center rotating ring that was not possible to fix. For this artifact 2 new detector modules were ordered. They will be replace as soon as I receive them. Diagnostic performed by Engineer : Images were evaluated and show a rotating center ring along multiple extended ring from the center to outside of the image. After carefully examination found contrast fluid in side the gantry. All over the inside of the protective cover, on top the detectors and in the x-ray port. This was the cause of the ring artifacts on the images. ----------------------------------------------------------------------------- Work Order #: WO-02660993 22-Dec-2020 Hernandez, Alvin The system fully meets specification and returned to use Diagnostic performed by Engineer : Found dry contrast on the center of the ultra high grid at the DMS. This was causing a round rotation artifact motion on the High resolution images. Resolution : The contras was removed from the high resolution grid. Air calibration was done. The system was tested and left fully operational. Problem Description by Engineer : Customer Reported Problem: Images have ring artifact in thin slices studies. ----------------------------------------------------------------------------- | P2 |  | Found dry contrast on the center of the ultra high grid at the DMS. This was causing a round rotation artifact motion on the High resolution images, For this artifact 2 new detector modules were ordered. Adjust lexan ring, cleaning the ring | [12 empty] | 116720221 | Closed | 491601 | TRUE | Closed - Done | Closed - Done |  | TBD_Contrast intrusion | [5 empty] | Customer |  |  |  | Closed |  |  | Lexan Ring/Contrast | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
291 | 44188 | TRUE | Spain | S.C.S. HOSP. GRAL. DE FUERTEVE | 40020 | 44188 | 52 | Problems when sending to PACS. | Various tests and checks are performed on the communications between the computer and the remote PACS system. It is observed that the equipment cancels the sending of images to the PACS, presenting an error message, because it does not get a response from the PACS during the predefined time interval (30 seconds). Time interval is modified and extended to 120 seconds. DICOM communication test is performed to PACS with error result for Dicom STORE service. Alternative communication tests are performed from an external node to the TAC, and it is observed that the response time of the PACS is also extremely long for the storage of images from this other node (ISP). Communication between the TAC and the rest of the Dicom nodes (ISP and Syngovía) is found to be very fast and fluid. It is recommended to check the correct operation of the remote PACS system. | P3 |  | PACS transfer, repeat bugreport: https://www.clouddrive.philips.com.cn/file/5bo82p1cok2fzh1k6a915v3tnpferlno# 密码:123 | [12 empty] | 116732434 | Closed | 497317 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering |  |  |  | Closed |  |  | PACS | [8 empty] | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
292 | 44158 | TRUE | United Kingdom | STEPPING HILL HOSPITAL | 40016 | 44158 | 48 | Ingenuity CT - noticed the brightness of the positioning lasers were dimmer |  | RFC |  | Laser brightness | [14 empty] | 496938 | TRUE | Closed - Done | Closed - Done |  | RFC | [5 empty] | Engineering |  |  |  | Closed |  |  | Laser alignment | [8 empty] | System issue | System | Gantry-HW | RFC | Gantry-HW-RFC | Gantry-HW-RFC | RFC | RFC | Gantry-HW | Gantry-HW-RFC | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
293 | 44158 | TRUE | United Kingdom | STEPPING HILL HOSPITAL | 40016 | 44158 | 48 | Ingenuity CT - noticed the new style of water filled system phantoms were difficult to use. |  | RFC |  | H-phantom improvement | [14 empty] | 496937 | TRUE | Closed - Done | Closed - Done |  | RFC | [5 empty] | Engineering |  |  |  | Closed |  |  | H-Phantom imp | [8 empty] | System issue | System | Couch-HW | RFC | Couch-HW-RFC | Couch-HW-RFC | RFC | RFC | Couch-HW | Couch-HW-RFC | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
294 | 44158 | TRUE | United Kingdom | STEPPING HILL HOSPITAL | 40016 | 44158 | 48 | Ingenuity CT - Whilst installing an IMR Ingenuity 5000 we noticed the cables for the IMR rack were missing |  | DEFOA |  | Cable DEFOA, need FSE provide DEFOA # | [14 empty] | 496927 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not a defect | [5 empty] | Operation |  |  |  | Closed |  |  | Part missing | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
295 | 44105 | TRUE | United States | Seton Medical Center Harker He | 40006 | 44105 | 40 | Phantom not detect during IQ check, Change Request(Defect) CTUDT00631545 was submitted for this issue-Second occurrence. |  |  |  | CTUDT00631545 Phantom not detect during IQ check | [14 empty] | 502039 | TRUE | Closed - Done | Closed - Done |  | CTUDT00631545 | [5 empty] | Engineering |  |  |  | Closed |  |  | H-Phantom detect | [8 empty] | System issue | System | DMS-SW | Phantom Detect | DMS-SW-Phantom Detect | DMS-SW-Phantom Detect | CTUDT00631545 | Phantom Detect | DMS-SW | DMS-SW-Phantom Detect | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
296 | 44189 | TRUE | United States | Lackey Memorial Hospital | 40015 | 44189 | 52 | WILL NOT SEND THE PACS/NETWORKING | Work Order #: WO-02676523 24-Dec-2020 Denison Chuck The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: X-ray Dept. Head Device Use: Philips C.T. scanner Expected and Unexpected Results: Images export to packs. User Impact: Sending images Patient Impact: None: Burn disk as backup. Current Software Version:4.1.10 Resolution : Worked with customer on issue of sending to PACS across the network. ----------------------------------------------------------------------------- | P2 |  | PACS Transfer, FCO72800737 to be implemented https://www.clouddrive.philips.com.cn/file/blii80od6ouakcpe1iqrtdgycn9g74k1# 密码:123 | [12 empty] | 116737273 | Closed | 499505 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering |  |  |  | Closed |  |  | PACS | [8 empty] | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
297 | 44188 | TRUE | United States | Lackey Memorial Hospital | 40015 | 44188 | 52 | Network connection | Work Order #: WO-02674055 23-Dec-2020 Denison Chuck The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: X-ray Dept. Head. Device Use: Philips C.T. scanner. Expected and Unexpected Results: Sending images to V-RAD after hours site. User Impact: Sending images across the network. Patient Impact: None. Was routed another way out. Current Software Version:4.1.10 Resolution : Worked on the config. set up in the Host P.C. . It all checked o.k. . Had to change the data rate a which the data was been sent over the network of the hospital. All is good . waited a bit to make sure all was o.k. ----------------------------------------------------------------------------- | P4 |  | 1st report PACS issue, Sending images to V-RAD after hours site. https://www.clouddrive.philips.com.cn/file/blii80od6ouakcpe1iqrtdgycn9g74k1# 密码:123 | [12 empty] | 116733916 | Incomplete | 497453 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering |  |  |  | Closed |  |  | PACS | [8 empty] | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
298 | 44187 | TRUE | United States | VA Medical Center Biloxi | 40018 | 44187 | 52 | Customer requesting Support-Cannot get the Estop closed | Work Order #: WO-02668012 23-Dec-2020 Mclain Jon The system fully meets specification and returned to use Resolution : Replaced Vertical transducer. Calibrated vertical and made backup. Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: ----------------------------------------------------------------------------- | P2 |  | Replaced Vertical transducer. Calibrated vertical and made backup. Ada follow up | [12 empty] | 116727610 | Closed | 497270 | TRUE | Closed - Done | Closed - Done |  | TBD_Vertical Transducer | [5 empty] | Engineering |  |  |  | WIP |  |  | Vertical Transducer | [8 empty] | System issue | System | Couch-HW | VerticalTransducer | Couch-HW-VerticalTransducer | Couch-HW-VerticalTransducer | TBD_Vertical Transducer | VerticalTransducer | Couch-HW | Couch-HW-VerticalTransducer | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
299 | 44187 | TRUE | United States | McLaren Macomb | 40021 | 44187 | 52 | 4.1.7 brain perfusion | Work Order #: WO-02666063 Curry, Dyan The system partly meets specification and is returned to use with restrictions Problem description by Clinical Support Specialist : Customer Function: technologist Device Use: in use Expected and Unexpected Results: post processing brain perfusion questions User Impact: none Patient Impact: none Current Software version: 4.1.7 Diagnostic performed by Engineer : tech is busy with patients. Called back, gone for the day. F/U tomorrow Diagnostic performed by Engineer : joined GTM with Evelyn. Evelyn went over BP package , advising customer how to minimize motion on images and place ROI_s. Also how to use tools for correct perfusion maps. ----------------------------------------------------------------------------- | P5 |  | Is it 4.1.7? advising customer how to minimize motion on images and place ROI_s. Also how to use tools for correct perfusion maps | [12 empty] | 116726303 | Closed | 494331 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Q&A | [5 empty] | Application | [15 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
300 | 44186 | TRUE | United States | Central Vermont Medical Center | 40024 | 44186 | 52 | System restricted, issue is system isn_t sending studies to PACS | Work Order #: WO-02658852 Daigle Jim Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | P3 |  | PACS transfer, repeat | [12 empty] | 116719079 | In Progress | 493887 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering |  |  |  | Closed |  |  | PACS | [8 empty] | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
301 | 44180 | TRUE | United States | Odessa Regional Medical Center | 40029 | 44180 | 51 | The table pad now has velcro almost the entire length of the pad. This makes it almost impossible to place the arm rest |  | RFC |  | Arm rest can not attached | [14 empty] | 491545 | TRUE | Closed - Done | Closed - Done |  | CTUDT00632125 | [5 empty] | Engineering |  |  |  | Closed |  |  | Arm rest | [8 empty] | System issue | System | Couch-HW | Arm Rest | Couch-HW-Arm Rest | Couch-HW-Arm Rest | CTUDT00632125 | Arm Rest | Couch-HW | Couch-HW-Arm Rest | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
302 | 44193 | TRUE | United States | The Regional Medical Center | 40005 | 44193 | 53 | 0798CDFM3 usb disconnect sound | USB disconnect beep FSE requesting WO to order parts Patient Impact/Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer Current Software Version: Not provided Parts Consumed : 1.00000,459801413791,IBOX2 USB EXTENDER Kit 1.00000,455300027201,CBL,USB A/M TO B/M 0.5M | P3 |  | Suggest to use Bizlink USB Cable | [12 empty] | 116747022 | Closed | 503067 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-USBExtended | [5 empty] | Engineering |  |  |  | WIP |  |  | USB Extender | [8 empty] | System issue | System | Gantry-HW | USBExtended | Gantry-HW-USBExtended | Gantry-HW-USBExtended | TBD-USBExtended | USBExtended | Gantry-HW | Gantry-HW-USBExtended | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
303 | 44195 | TRUE | United States | Lackey Memorial Hospital | 40015 | 44195 | 53 | 1202CDOD1 Maximum Number Of Hard Arcs Errors | "Time of Event" "12/30/2020 8:29:00 AM" "Time of Collection" "12/30/2020 8:29:37 AM" "Problem Name" "MAX NUMBER OF HARD ARCS" Submitter "SHEA BOLES" CT Tech reports Max Number of hard arcs errors on 12/30 at 0830 and on 12/21 during onsite Clinical Applications training. Last gantry power-cycle = 8 days earlier so error has occurred again since gantry power-cycle. Logger indicates: 2020/12/30 08:28 S_GEN_FAULT_CATHODE_TUBE_ARC 5:5 CPM Arc 2020/12/30 08:28 S_GEN_FAULT_ANODE_TUBE_ARC 5:0 APM Arc 2020/12/30 08:28 S_GEN_FAULT_ANALOG_FAULT 2:5 Analog Fault 2020/12/30 08:28 S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT NORMAL Bug Report additional info: Description: High voltage arc *** HV Error Info: *** GXray EXP=1772 MA err -26.0* 2:5 Analog Fault 5:0 APM Arc 5:5 CPM Arc Follow Up Required : Inspect HV connections, candlesticks, silicone grease, etc. Perform long TC. Perform filament calibrations if necessary. Inspect Power Brush Blocks / slip ring for poor connections. If necessary, perform no-load test. ----------------------------------------------------------------------------- | P5 |  | bugreport is downloading for investigation https://www.clouddrive.philips.com.cn/file/blii80od6ouakcpe1iqrtdgycn9g74k1# PASSWORD:123 | [12 empty] | 116758107 | In Progress | 507320 | TRUE | Closed - Done | Closed - Done |  | CTUDT00632177 | [5 empty] | Engineering |  |  |  | Closed |  |  | Tube Arc | [8 empty] | System issue | System | IC-HW | Hard Arcs | IC-HW-Hard Arcs | IC-HW-Tub Arc | CTUDT00632177 | Hard Arcs | IC-HW | IC-HW-Hard Arcs | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
304 | 44194 | TRUE | United States | Lackey Memorial Hospital | 40015 | 44194 | 53 | Clinical support questions about building protocol for scanner. | Work Order #: WO-02691059 Mujica, David The system fully meets specification and returned to use Problem description by Clinical Support Specialist : Customer Function/Role: Technologist Device Use: in use Expected and Unexpected Results: Peds Exam Cards User Impact: none Patient Impact: none Current Software Version: 4.1 Diagnostic performed by Engineer : Michea was aided to find the smpte pattern test for the physicist, but unfortunately the Demo folder was not created on this new system and the smpte could not be found on the system. An FSE has already been dispatched and Michea will coordinate a time for him to install this test. Michea also had questions on Pediatric Head Exam Cards. Slice thickness for reconstruction was covered as well as ACR guidelines for this exam card. No further questions at this time. ----------------------------------------------------------------------------- | P3 |  | Clinical support Q&A | [12 empty] | 116753405 | Incomplete | 507380 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not a defect | [5 empty] | Application | [15 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
305 | 44196 | TRUE | United Kingdom | STEPPING HILL HOSPITAL | 40016 | 44196 | 53 | Date and time wrong on scanner | Work Order #: WO-02702006 31-Dec-2020 Martin Greenbank The system fully meets specification and returned to use Regulatory Questions : N/A Diagnostic performed by Engineer : see below Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : CM - Successful Corrective maintenance Visit with reported problem - Confirmed and corrected reported problem. corrected date and time on cirs s1 and tested for normal operation. Device is now within manufacturer’s specification. | P3 |  | Date and time setting already in Installation manual | [12 empty] | 116764880 | Closed | 509084 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not a defect | [5 empty] | Service |  |  |  | Closed |  |  | Time sync | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
306 | 44197 | TRUE | China | SD Chest Hosp. | 40025 | 44197 | 1 | wechat:02N/Ingenuity CT 728326/Blue screen | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : Blue screen Next Step Action : RSE hanwennian 回电客户 Safety Questions : Case: 0116768184 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Guomei Li Last updated by: Guomei Li | P3 |  | Blue screen after yellow screen bugreport ready https://www.clouddrive.philips.com.cn/file/dzaopb6yzuotz05hnptuj5ttub8gm0db# 密码:123 | [12 empty] | 116768184 | New | 510004 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00628691 | [5 empty] | Engineering |  |  |  | Closed |  |  | Yellow Screen | [8 empty] | System issue | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00628691 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
307 | 44193 | TRUE | China | SD Chest Hosp. | 40025 | 44193 | 53 | 02N/Ingenuity CT 728326/gantry power off | 客户问题概述： 机架掉电 客户问题详细描述或故障现象： 机架掉电 设备本身技术问题： 是。 Troubleshooting Action: 客户问题概述： 机架掉电 客户问题详细描述或故障现象： 机架掉电 设备本身技术问题： 是。 Repair Action: 建议现场维修步骤： 1.检查水平机械阻力 2.检查AMC供电，线缆等 建议备件12NC及名称： 无 客户需求备注： 无SAP Transfer : Dispatched for Onsite SupportExternal Remarks : 29.12.2020 05:02:15 UTC SUP JCOUSER (86038405)服务内容：检查床水平运动阻力，检查AMC信号线与电源线。 服务结果：检修后，设备恢复正常使用。Internal Remarks : 31.12.2020 15:58:27 UTC Zhaokun Huang (86038405) "RDI Accuracy: YES." $Q7: ( Confirmed reported problem & Estop open when unload couch ) $Q8: ( Check error log , report "COUCH_HOR_MOTION_STOP_ERROR", Check Cantrace find AMC driver haven’t give the answer to couch & AMC Driver intermediate failed.) $Q9: ( Reconnect AMC driver signal and power cables.) $Q10: ( Test and inspection data refer to: PA . Device returned to full functionality.) $Q11: Calibrated Tools name (Multimeter Fluke 287 ) S/N (23350170 ) Next Calibration date ( 04/16/2021 ) ^Current Tube installed date: N/A ^Current Tube usage: N/A | P3 |  | E-Stop open CAN TRACE, LOGGER ready https://www.clouddrive.philips.com.cn/file/dzaopb6yzuotz05hnptuj5ttub8gm0db# 密码:123 | [12 empty] | 116742199 | Closed | 503538 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632187 | [5 empty] | Engineering |  |  |  | Closed |  |  | Horizontal movement | [8 empty] | System issue | System | Couch-SW | Estop | Couch-HW-Estop | Couch-SW-Horizontal Move Block | CTUDT00632187 | Estop | Couch-SW | Couch-HW-Estop | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
308 | 44195 | TRUE | United States | Odessa Regional Medical Center | 40029 | 44195 | 53 | CT was hung up during a interventional procedure. | Work Order #: WO-02695631 Mark Finnigan The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: Used for chest tube placement. User Impact: System froze during the scan. Patient Impact: procedure was completed after the CT was restarted. Current Software Version: Diagnostic performed by Engineer : Spoke with James Ortiz. The scanner froze up while the site was performing an interventional exam. Unable to get apps support at that moment as there were all tied up with customer calls. The customer eventually restarted the CT and when back up completed the interventional procedure. Requested that the customer create a bugreport . Answered yest to the safety question and performed the debrief. Dispatch case and call ing the FSE so that they can perform the next step actions and get the NSS involved. Follow Up Required : Spoke with James Ortiz. The scanner froze up while the site was performing an interventional exam. Unable to get apps support at that moment as there were all tied up with customer calls. The customer eventually restarted the CT and when back up completed the interventional procedure. Requested that the customer create a bugreport . Answered yest to the safety question and performed the debrief. Dispatch case and call ing the FSE so that they can perform the next step actions and get the NSS involved. ----------------------------------------------------------------------------- Work Order #: WO-02695707 McCain, Jimmie Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: Used for chest tube placement. User Impact: System froze during the scan. Patient Impact: procedure was completed after the CT was restarted. Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | P4 |  | The scanner froze up while the site was performing an interventional exam. The customer eventually restarted the CT and when back up completed the interventional procedure. Requested that the customer create a bugreport | [12 empty] | 116758524 | In Progress | 507446 | TRUE | Closed - Done | Closed - Done |  | TBD_Bugreport | [5 empty] | Engineering |  |  |  | WIP |  |  | Interventional | [8 empty] | System issue | System | Console-SW | SystemStuck | Console-SW-SystemStuck | Console-SW-SystemStuck | TBD_Bugreport | SystemStuck | Console-SW | Console-SW-SystemStuck | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
309 | 44200 | TRUE | Japan | Kinan Hospital | 340000 | 44200 | 2 | An error occurs and Air Calibration does not run. >> SU000C0A00024: Error during calibration | BU Q&R Feedback: NoReason: Corrective MaintenanceProblem Reported by Customer : 03CS:KK:FSE大山さん:8時ごろE-STOP OPENした。>>アークが発生していました。再CLOSEにて復旧。----------------------------------------------------------------------------- | P4 |  | E-STOP opened around 8 o_clock. >> An arc was occurring. Restored by reCLOSE. -> The case was wrongly registered and canceled. This problem was reported from the other site. | [12 empty] | 116770459 | In Process | 511535 | TRUE | Closed - Duplicate | Closed - Duplicate |  | NAC | [5 empty] | Engineering |  |  |  | WIP |  |  | Invalid | [8 empty] | System issue | System | Gantry-HW | Estop | Gantry-HW-Estop | Gantry-HW-Estop | NAC | Estop |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
310 | 44193 | TRUE | Japan | Kinan Hospital | 340000 | 44193 | 53 | There is a strange noise when the sleeper moves up and down. | There is a strange noise when the sleeper moves up and down. [Phenomenon confirmation] I was able to confirm. [Details of work] There was an abnormal noise on one side of the gas spring inside the sleeper. The abnormal noise has disappeared by adjusting the gas spring mounting position. There is no problem with the operation of the device. | P5 |  | The abnormal noise has disappeared by adjusting the gas spring mounting position. | [12 empty] | 116742200 | New | 503455 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-GasPring | [5 empty] | Engineering |  |  |  | Closed |  |  | Couch Spring | [8 empty] | System issue | System | Gantry-HW | Gas Pring | Gantry-HW-Gas Pring | Gantry-HW-Gas Pring | TBD-GasPring | Gas Pring | Gantry-HW | Gantry-HW-Gas Pring | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
311 | 44198 | TRUE | Japan | Nikko Memorial Hospital | 341011 | 44198 | 1 | Protocol change work | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 01DS:RS:FSE経由:プロトコール変更作業 FSE Notes : This PO is a Reference Number only T2 Activities : Problem Description: 対象装置:CT 故障個所:01Cosole 最初に発生したのはいつか?(~以内):1 Hour 問題発生の頻度は?:Unknow エンドユーザｰのワークフローのどの時点において問題が発生しているのか?:その他(Other) 最近どのようなイベントが発生したか?:N.A. 顧客は問題について具体的なシステムイベントレポートを作成しているか?:N.A. Troubleshooting Action: 対応後の結果:改善された 故障個所:01Cosole 診断中にどのようなステップを実施したか?: FSE対応 Repair Action: サービス担当の訪問作業。 OneEMS internal remarks : --Troubleshooting Action-- 対応後の結果:改善された 故障個所:01Cosole 診断中にどのようなステップを実施したか?: FSE対応 --Repair Action-- --Subject-- 01DS:RS:FSE経由:プロトコール変更作業 --Problem Description-- 対象装置:CT 故障個所:01Cosole 最初に発生したのはいつか?(~以内):1 Hour 問題発生の頻度は?:Unknow エンドユーザｰのワークフローのどの時点において問題が発生しているのか?:その他(Other) 最近どのようなイベントが発生したか?:N.A. 顧客は問題について具体的なシステムイベントレポートを作成しているか?:N.A. Safety Questions : Case: 0116768898 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: RUI SATAKE Last updated by: RUI SATAKE 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Dispatched for Onsite Support SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | P4 |  | Problem Reported by Customer: 01DS: RS: Via FSE: Protocol change work | [12 empty] | 116768898 | In Process | 511493 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not a defect | [5 empty] | Application |  |  |  | Closed |  |  | Protocol | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
312 | 44196 | TRUE | China | SD Laiyang People Hospital | 345005 | 44196 | 53 | 02N/Ingenuity Core128/Elite China/Star Station switched to Nebula Station | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : Star Station switched to Nebula Station Safety Questions : Case: 0116761157 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Chen YuanHui Last updated by: Chen YuanHui 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 需要把CT连接到星云工作站 客户问题详细描述或故障现象： 需要把CT连接到星云工作站 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： NA 远程诊断与测试结果： 客户认为星河工作站不能满足后处理要求，希望将CT图像传输到星云工作站处理。星云工作站是磁共振配备的工作站，已经出保。联系工程师反馈，不确定星云配置的选项，且告知客户联系过销售。 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 到场尝试连接工作站节点，看是否有CT图像后处理功能，若有商务因素，及时联系相关人员 SAP Transfer : Dispatched for Onsite Support SAP Transfer : Case Transferred to SAP External Remarks : 03.01.2021 15:25:36 UTC Xingbo Yang (86038250)服务内容：添加远程节点。 服务结果：维修后，设备恢复正常工作。 Internal Remarks : 03.01.2021 15:25:39 UTC Xingbo Yang (86038250) "RDI Accuracy:yes" $Q7: (Confirmed reported problem & need to configured remote workstaiton) $Q8: (check the local network & need to configured remote workstation ) $Q9: ( add remote node to local configuration ) $Q10: (Test and inspection data refer to: PA & Device returned to full functionality) $Q11: Calibrated tools name(NA) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date: 9 / 25 /2020 ^Current Tube usage: scan second_67284 ----------------------------------------------------------------------------- | P3 |  | Change PACS Configuration | [12 empty] | 116761157 | In Process | 510450 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not a defect | [5 empty] | Customer | [15 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
313 | 44194 | TRUE | China | HLJ Huachuan People Hosp. | 345018 | 44194 | 53 | 02N/Ingenuity Core128/Elite China/cannot boot up | Next Step Action: Customer report for repair, workstation fail to boot up. Unable to query workstation info in the system. Contacted site1, jin jun informed that he has not been to the site, doesn_t know the details. Contact ASP huicong to confirm site info Next Step Action: Huicong replied, need Sales to follow up. Contact Presales M.M. Wang Fengjuan Heilongjiang region manager 13903651586 Told us not to worry, she will contact sales to follow-up | P5 |  | Workstation fail to boot up. Unable to query workstation info in the system. Peter Check if CT complaint or ISP related and how to diff. | [12 empty] | 116749597 | Closed | 506384 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD_Workstation | [5 empty] | Engineering |  |  |  | WIP |  |  | Workstation | [8 empty] | System issue | System | Console-SW | DeviceConfig | Console-SW-DeviceConfig | Console-SW-DeviceConfig | TBD_Workstation | DeviceConfig | Console-SW | Console-SW-DeviceConfig | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
314 | 44195 | TRUE | China | Zhenjiang M.C.A. Hosp. | 345023 | 44195 | 53 | wechat:02G/Ingenuity Core128/Elite China/The detector artifact is serious | Next Step Action: contacted worker Shi, for now, no need to add operation to the machine order, previous one can still be used.----------------------------------------------------------------------------- | P3 |  | Installation with BD, BD detected during clinical training, calibration data from D:/pms/data/physics/history/Bad Det is available in Cloud drive, https://www.clouddrive.philips.com.cn/file/2qnbhckeg7aw21n1xcar36heuog33t2k# 密码:123 Share it to DMS team to investigate | [12 empty] | 116755596 | Closed | 508484 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632110 | [5 empty] | Engineering |  |  |  | Closed |  |  | Defective DMS module | [8 empty] | System issue | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | DMS Module | DMS-HW | DMS-HW-DMS Module | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
315 | 44202 | TRUE | Chile | CMD Clinica Integrales Rancagu | 40008 | 44202 | 2 | CL1675 - PROBLEM RELATED TO THE LOWERING BUTTON WITHIN THE CT BOX (DETECTED DURING PREVENTIVE MAINTENANCE). | Work Order #: WO-02729561 Osorio, Juan The system fully meets specification and returned to use. Problem Description by Engineer: Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer: Lowering button in CT BOX is experiencing issues when being presses, as it activates intermittently. Resolution: Replacement of the CT Box. Follow-up is generated. Follow Up Required: Clinica Redsalud Rancagua (integral clinic). Libertador Bernardo O_Higgins 634, Rancagua D. R., Medical Devices, +56 9 6156xxxx | P5 |  | Can field share some picture/info about the button failure, if possible to return the CT Box to Suzhou for investigation? | [12 empty] | 116786665 | In Progress | 518961 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632185 | [5 empty] | Engineering |  |  |  | Closed |  |  | CT Box | [8 empty] | System issue | System | Gantry-HW | CT Box | Gantry-HW-CT Box | Gantry-HW-CT Box | CTUDT00632185 | CT Box | Gantry-HW | Gantry-HW-CT Box | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
316 | 44200 | TRUE | United States | Lackey Memorial Hospital | 40015 | 44200 | 2 | needs to removes object off image | Work Order #: WO-02713636 Mujica, David The system fully meets specification and returned to use Problem description by Clinical Support Specialist : Customer Function/Role: Technologist Device Use: In use Expected and Unexpected Results: 3D post processing User Impact: none Patient Impact: none Current Software Version: 4.1 Diagnostic performed by Engineer : Michea was aided to acquire 3D rotations of the circle of wilis in CT viewer for batching. Adjusting threshold values to trigger the CTA scan sooner was also discussed for CTA head/neck. No further instruction is needed at this time. ----------------------------------------------------------------------------- Work Order #: WO-02715705 Curry, Dyan The system fully meets specification and returned to use Problem Description by Engineer : Customer Function: technologist Device Use: in use Expected and Unexpected Results: how to remove bone for CTA User Impact: none Patient Impact: none Current Software version: 4.1.7 Diagnostic performed by Engineer : used REACTS to assist. Helped customer load images in CT Viewer and use clipping tools to remove table and bones. showed customer how mouse middle mouse and shower button to remove scatter. Brain images look sub par for contract. Loaded neck images and bone remove was done with no issues. There is a lot of venous contamination. Check EC card , changed threshold from 150 to 120. Customer was using a "difference " method, advised her to just let the scanner start as EC is. Customer asks for a follow up call tomorrow Diagnostic performed by Engineer : follow up with customer. OK to close. ----------------------------------------------------------------------------- | P5 |  | it just answered some questions and trained the user to operate in Viewer and adjust the threshold for CTA scan. | [12 empty] | 116777000 | Closed | 512718 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not a defect | [5 empty] | Application | [6 empty] | NAC_Q&A | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
317 | 44202 | TRUE | France | SCP GROUPE D IMAGERIE MEDICALE | 40019 | 44202 | 2 | Network tests | Work Order #: WO-02725808 06-Jan-2021 Jean-luc CODRON The system fully meets specification and returned to use Regulatory Questions : N/A Problem Description by Engineer : tests reseau transfert PACS Diagnostic performed by Engineer : tests reseau transfert PACS | P3 |  | is it related to PACS transfer issue? Any detail information could share? | [12 empty] | 116789009 | Closed | 518260 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering |  |  |  | Closed |  |  | PACS | [8 empty] | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
318 | 44201 | TRUE | France | SCP GROUPE D IMAGERIE MEDICALE | 40019 | 44201 | 2 | Free space for images are limited due to saturation of log taking place on D drive | D:\40019 | CFM | https://www.clouddrive.philips.com.cn/file/uiilqgfqtj8gp9nyfff378n1w0ripy2h# 密码:123 DEID_2020_12_15_10_16_05_MESTIRI__SUIVI_VOLUME_DIRECTORY | https://www.clouddrive.philips.com.cn/file/uiilqgfqtj8gp9nyfff378n1w0ripy2h# 密码:123 DEID_2020_12_15_10_16_05_MESTIRI__SUIVI_VOLUME_DIRECTORY | [12 empty] | CFM |  | 514848 | TRUE | Closed - Done | Closed - Done |  | CTUDT00632261 | [5 empty] | Engineering |  |  |  | Closed |  |  | TBC_Disk full | [8 empty] | System issue | System | Console-SW | DiskFull | Console-SW-DiskFull | Console-SW-DiskFull | CTUDT00632261 | DiskFull | Console-SW | Console-SW-DiskFull | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
319 | 44203 | TRUE | Spain | S.C.S. HOSP. GRAL. DE FUERTEVE | 40020 | 44203 | 2 | Transfers to Pacs stucked. | Work Order #: WO-02734348 Carlos Saumell Lasala Regulatory Questions : N/A Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A | P4 | https://www.clouddrive.philips.com.cn/file/4z92zhovr6yv6xiiki797tk50vssfafh# 密码:123 DEID_2021_01_06_10_02_18_PHILIPS_COMUNICACION.zip | https://www.clouddrive.philips.com.cn/file/4z92zhovr6yv6xiiki797tk50vssfafh# 密码:123 DEID_2021_01_06_10_02_18_PHILIPS_COMUNICACION.zip | [12 empty] | 116795725 | In Progress | 523523 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering |  |  |  | Closed |  |  | PACS | [8 empty] | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
320 | 44201 | TRUE | China | NMG Zizhiqu No.4 Hosp | 40023 | 44201 | 2 | 02N/Ingenuity CT 728326/Images of failure | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : Images of failure Safety Questions : Case: 0116780692 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Zixuan Last updated by: Wang Zixuan 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 图像质量问题 客户问题详细描述或故障现象： 图像质量问题 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： NA 远程诊断与测试结果： 客户反映图像质量太差，严重影响诊断。客户意见很大，急需处理 PRS TR状态： 无PRS。 Repair Action: 建议现场维修步骤： 现场做图像质量测试，做图像校准；联系应用老师，确认扫描协议设置。必要时上escalation，上CFM. 客户需求备注： 客户很急，工程师尽快回电客户 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Internal Remarks : 11.01.2021 05:49:28 UTC Feng Song (86038267) RDI：yes $Q7: (Confirmed reported problem , Images of failure ) $Q8:(Image quality low ,DRI level too low) Q9 ( Adjust scan protocal and perform hcor and air calibration ,,then issue was solved ) $Q10: (Test and inspection data refer to PA. Device returned to full functionality..) $Q11:Calibrated tools refer to PA External Remarks : 11.01.2021 05:49:25 UTC Feng Song (86038267)客户联系人：影像科贾主任13848146018 服务内容:客户提出图像质量差，调整扫描条件并做相关校正后，问题排除。 服务结果:服务完成后，设备恢复正常 其它重要事项：无 ----------------------------------------------------------------------------- | P3 |  | HCOR/Performance test passed, acceptance test fail (head), FSE is ordering new phantom to check if It is phantom issue. Export the calibration history of failed item as well as DICOM images. Receive images https://www.clouddrive.philips.com.cn/file/f7mr107a8k5ksdxcm07j94zq84zrr6iu# 密码:123 | [12 empty] | 116780692 | In Process | 517348 | TRUE | Closed - Done | Closed - Done |  | TBD_IQ | [5 empty] | Engineering |  |  |  | WIP |  |  | TBD_IQ | [8 empty] | System issue | System | Couch-HW | H-Phantom | Couch-HW-H-Phantom | Couch-HW-H-Phantom | TBD_IQ | H-Phantom | Couch-HW | Couch-HW-H-Phantom | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
321 | 44204 | TRUE | United States | Central Vermont Medical Center | 40024 | 44204 | 2 | UPS replacement + PACS transfer issue | Work Order #: WO-02740150 Cooke Steven Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A Worked with Staco rep to start up the ups. | P5 |  | the problems is that they delivered a 50HZ full system UPS. It was replaced by the correct one during the weekend. I’m checking if this is part of the system configuration or this is something they purchased separately from the original PO to check if it is a shipping / delivery or configuration problem. https://teams.microsoft.com/_#/files/%E5%B8%B8%E8%A7%84?threadId=19%3A76cc42a0c8fa4bd7877f7de12141f84c%40thread.tacv2&ctx=channel&context=CVMC-CT5000-SN40024-4.1.10.25208-ESC0116719079-PR%2520BD&rootfolder=%252Fsites%252FBigBoreRTRefreshFOKCoreTeam%252FShared%2520Documents%252FGeneral%252FIngenuity%2520Bugreps%252FCVMC-CT5000-SN40024-4.1.10.25208-ESC0116719079-PR%2520BD | [12 empty] | 116803273 | In Progress | 526205 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | TBD_UPS | [5 empty] | Engineering |  |  |  | Closed |  |  | TBD_UPS | [8 empty] | System issue | System | Gantry-HW | UPS | Gantry-HW-UPS | Gantry-HW-UPS | TBD_UPS | UPS |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
322 | 44201 | TRUE | United States | Odessa Regional Medical Center | 40029 | 44201 | 2 | Ring Artifact | Work Order #: WO-02719515 McCain, Jimmie Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A | P3 |  | Contrast intrusion (present on the face of the detector and the Mylar protective cover. | [12 empty] | 116782625 | In Progress | 518106 | TRUE | Closed - Done | Closed - Done |  | ECR-075975 | [5 empty] | Engineering |  |  |  | Closed |  |  | Lexan Ring/Contrast | [8 empty] | System issue | System | Gantry-HW | Lexan Ring | Gantry-HW-Lexan Ring | Gantry-HW-Lexan Ring | ECR-075975 | Lexan Ring | Gantry-HW | Gantry-HW-Lexan Ring | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
323 | 44170 | TRUE | United States | Odessa Regional Medical Center | 40029 | 44170 | 49 | 1399CDFD1 Cannot Transfer To PACS & Injector Not Triggering | source case 116608163/456455, separate to two PR | CFM |  | source case 116608163/456455, separate to two PR Studies not auto-sending to PACS and injector not triggering. Verified injector settings for SAS and auto-send settings in Preferences/Exam Cards set. Scanner performing properly per the settings in Preferences and Exam Card. Enhance application training | [12 empty] | Split |  | 512353 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not a defect | [5 empty] | Engineering |  |  |  | Closed |  |  | Separate | [8 empty] | System issue | System | Gantry-HW | Injector | Gantry-HW-Injector | Gantry-HW-Injector | Not a defect | Injector | Gantry-HW | Gantry-HW-Injector | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
324 | 44200 | TRUE | China | Yangfangdian hospital | 40035 | 44200 | 2 | 02N/Ingenuity CT 728326/Locating lamp fault | BU Q&R Feedback: No Reason: Corrective Maintenance Next Step Action: The customer says he expects the field service engineer on site in the afternoon. Problem Reported by Customer: Locating lamp fault Safety Questions:Case: 0116771012 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Mengen Xie Last updated by: Mengen Xie 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Case Transferred to SAP T2 Activities:Problem Description: Summary of customer_s problems: Other types of error. Detailed description of Customer_s problems or malfunctions: Laser positioning light is not accurate It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: Laser positioning light is not accurate Remote troubleshoot and test result: Need to adjust laser lamp. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Need to adjust laser lamp. SAP Transfer:Dispatched for On site Support External Remarks : 05.01.2021 03:48:01 UTC Xing Zheng (86039460) Service Details: The laser lamp_s positioning is inaccurate Service Result: The equipment is restored to its full functionality. Internal Remarks: 05.01.2021 03:48:04 UTC Xing Zheng (86039460) $Q7：Confirmed reported problem:Locating lamp fault $ Q8 Performed visual inspection, the locating laser fault $Q9: adjust the laser locating. $Q10：Test and inspection data refer to: PA. Device returned to full functionality. $Q11:N/A Internal Remarks : 06.01.2021 15:46:36 UTC HongHong Han (86039436)$Q7：Confirmed reported problem:Locating lamp fault $ Q8 Performed visual inspection, the locating laser fault $Q9: adjust the laser locating. $Q10：Test and inspection data refer to: PA. The device returned to full functionality. $Q11:N/A External Remarks : 06.01.2021 15:46:33 UTC HongHong Han (86039436)) Service Details: The laser lamp_s positioning is inaccurate. Service Result: The equipment is restored to its full functionality. ----------------------------------------------------------------------------- | P3 |  | Link to Andy_s ECR about laser position change ECR-077066 | [12 empty] | 116771012 | In Process | 516689 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | ECR-077066 | [5 empty] | Engineering |  |  |  | Closed |  |  | Laser alignment | [8 empty] | System issue | System | Gantry-HW | Laser | Gantry-HW-Laser | Gantry-HW-GantryLaser | ECR-077066 | gantry laser | Gantry-HW | Gantry-HW-Laser | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
325 | 44200 | TRUE | China | SD Jinan No.4 Peo. Hosp. | 40038 | 44200 | 2 | PRS connection/02G/Ingenuity CT 728326 | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer:PRS Connection Safety Questions:Case: 0116771228 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Qian Wang Last updated by: Qian Wang 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action:FSE Chen Tao 18560023293 call in and disc him Email IB to add operation to IA#: 52849499, dispatch FSE Chen Tao Next Step Action: OP added, 0020 ----------------------------------------------------------------------------- | P5 |  | PRS connection | [12 empty] | 116771228 | Closed | 513964 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not a defect | [5 empty] | Service | [6 empty] | PRS Connection | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
326 | 44201 | TRUE | Turkey | HITACHI HEALTHCARE TURKEY | 337000 | 44201 | 2 | Uzak bağlantı için sistemin PRS’e kaydı. | Registration of the system to PRS for remote connection. | P3 |  | PRS connection | [12 empty] | 116781138 | Closed | 517320 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not a defect | [5 empty] | Service | [6 empty] | PRS Connection | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
327 | 44200 | TRUE | Japan | Kinan Hospital | 340000 | 44200 | 2 | 03CS:KK:FSE Oyama: An E-stop open occurred at around 8:00.>>An arc had occurred.Recovery achieved upon doing [TRUNCATED] | BU Q&R Feedback: No SAP Customer Symptom Code: HANG Customer Symptom Code: Hang/Crash Reason: Corrective Maintenance Engineer Failure Code: Electrical/Electronic Failure Repair Activity Code: Reboot/Restart OneEMS internal remarks: --Troubleshooting Action-- Result of response: Problem solved. Failure location: 03High Voltage Energy What steps did you perform to diagnose the problem? An arc occurred.Recovery achieved upon doing a CLOSE again. I had Oyama contact the customer, and the customer will monitor the situation. From the logger, 2021/01/04 07:59:37.166__02 [--o] ----------------- Series 4 Ready ----------------- 2021/01/04 07:59:38.567__01 [-oo] - - - - - - Shot 0 Starts - - - - - - 2021/01/04 07:59:41.673__01 [-oo] - - - - - - Shot 1 Starts - - - - - - 2021/01/04 07:59:44.671__01 [-oo] - - - - - - Shot 2 Starts - - - - - - 2021/01/04 07:59:47.718__01 [-oo] - - - - - - Shot 3 Starts - - - - - - 2021/01/04 07:59:50.720 [-oo] - - - - - - Shot 4 Starts - - - - - - 2021/01/04 07:59:53.719__01 [-oo] - - - - - - Shot 5 Starts - - - - - - 2021/01/04 07:59:56.765__01 [-oo] - - - - - - Shot 6 Starts - - - - - - 2021/01/04 07:59:59.812 [ooo] ----------------- Series 4 Complete ----------------- 2021/01/04 08:00:10.054__06 [---] ------ Series 5 Load Structure ----- 2021/01/04 08:00:17.291__02 [--o] ----------------- Series 5 Ready ----------------- 2021/01/04 08:00:18.366 [-oo] - - - - - - Shot 0 Starts - - - - - - 2021/01/04 08:00:20.191__06 <*HV*> szInfo: SoftArcReg 0x211* 2021/01/04 08:00:20.242 <*HV*> Type: NORMAL ErrNum: S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT szInfo: P1:0/P2: 2021/01/04 08:00:20.263 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_ANALOG_FAULT szInfo:2:5 Analog Fault P1:2/P2:32 2021/01/04 08:00:20.265 [CANCEL] CStateShot:Series 5 canceled, due to GANTRY_CANCEL_SOURCE_NORMAL_ERROR, in state CStateShot 2021/01/04 08:00:20.272__02 [***] - - - - - - Study Exit - - - - - - 2021/01/04 08:00:20.273 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_ANODE_TUBE_ARC 2021/01/04 08:00:20.288 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_CATHODE_TUBE_ARC 2021/01/04 08:00:20.292 <WARN> Type: WARN ErrNum: S_GEN_MGR_SYNC_OFF_OUT_OF_SEQUENCE P1:0/P2:0 2021/01/04 08:00:22.309__01 <*ERROR*> Type: NORMAL ErrNum: S_MASTER_SEQUENCER_EXIT_STATE_ERR P1:4096/P2:8369 2021/01/04 08:00:37.308 <*ERROR*> Type: NORMAL ErrNum: S_TIMER_WATCHDOG_CAUSED_LAST_RESET P1:4096/P2:0 2021/01/04 08:00:38.288 <WARN> Type: WARN ErrNum: S_DTH_USER_ALREADY_REGISTERED P1:0/P2:0 2021/01/04 08:00:38.292 <WARN> Type: WARN ErrNum: S_GAFSTARTUPTASK_DTH_BAD_PARAM P1:1/P2:-1 2021/01/04 08:00:38.296 <WARN> Type: WARN ErrNum: S_GAFSTARTUPTASK_DTH_BAD_PARAM P1:1/P2:-1 2021/01/04 08:00:38.299 <WARN> Type: WARN ErrNum: S_GAFSTARTUPTASK_DTH_BAD_PARAM P1:1/P2:-1 2021/01/04 08:00:38.302 <WARN> Type: WARN ErrNum: S_GAFSTARTUPTASK_DTH_BAD_PARAM P1:1/P2:-1 2021/01/04 08:00:38.336__01 <WARN> Type: WARN ErrNum: S_ROTOR_DTH_BAD_PARAM P1:-1/P2:0 2021/01/04 08:00:38.341__03 <WARN> Type: WARN ErrNum: S_DTH_USER_ALREADY_REGISTERED P1:0/P2:0 2021/01/04 08:00:38.432 <*ERROR*> Type: NORMAL ErrNum: S_MP_MPSHELL_SERIAL_LIFE_SYNC_ERROR szInfo:MP Latest Fatal Error Data. P1:3.670017E+07/P2:3997697 2021/01/04 08:00:38.948__01 <WARN> Type: WARN ErrNum: S_RMC_DTH_BAD_PARAM P1:-1/P2:0 2021/01/04 08:00:38.950 <WARN> Type: WARN ErrNum: S_RMC_MP_DTH_BAD_PARAM P1:-1/P2:0 2021/01/04 08:02:28.347 <WARN> Type: WARN ErrNum: S_MPS_WAITING_FOR_MAIN_DRIVE_TO_CLOSE P1:0/P2:0 szInfo:MP: Waiting for MDC to close.. 2021/01/04 08:02:29.433 <WARN> Type: WARN ErrNum: S_MPS_WAITING_FOR_MAIN_DRIVE_TO_CLOSE P1:0/P2:0 szInfo:MP: Waiting for MDC to close.. 2021/01/04 08:02:30.038 [ESTOP CLOSE] Type:EVENT_MINOR Msg:GantryEstopClosedBkst in CStateConfigured --Repair Action-- --Subject-- 03CS:KK:FSE Oyama: An E-stop open occurred at around 8:00.>>An arc had occurred.Recovery achieved upon doing a CLOSE again. --Problem Description-- Equipment affected: CT Failure location: 03High Voltage Energy When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence When, in the end user_s work flow, is the problem occurring?: During Examination ・Measurement preparation (Setting up Protocol) What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem? No.The customer has not created a log file or report. --4/8Elements-- Q1. Was the reported problem confirmed? Confirmed. Q2. Were any other problems confirmed? None. Q3. What did you do to resolve the problem?: An arc occurred.Recovery achieved upon doing a CLOSE again. I had Oyama contact the customer, and the customer will monitor the situation. From the logger, 2021/01/04 07:59:37.166__02 [--o] ----------------- Series 4 Ready ----------------- 2021/01/04 07:59:38.567__01 [-oo] - - - - - - Shot 0 Starts - - - - - - 2021/01/04 07:59:41.673__01 [-oo] - - - - - - Shot 1 Starts - - - - - - 2021/01/04 07:59:44.671__01 [-oo] - - - - - - Shot 2 Starts - - - - - - 2021/01/04 07:59:47.718__01 [-oo] - - - - - - Shot 3 Starts - - - - - - 2021/01/04 07:59:50.720 [-oo] - - - - - - Shot 4 Starts - - - - - - 2021/01/04 07:59:53.719__01 [-oo] - - - - - - Shot 5 Starts - - - - - - 2021/01/04 07:59:56.765__01 [-oo] - - - - - - Shot 6 Starts - - - - - - 2021/01/04 07:59:59.812 [ooo] ----------------- Series 4 Complete ----------------- 2021/01/04 08:00:10.054__06 [---] ------ Series 5 Load Structure ----- 2021/01/04 08:00:17.291__02 [--o] ----------------- Series 5 Ready ----------------- 2021/01/04 08:00:18.366 [-oo] - - - - - - Shot 0 Starts - - - - - - 2021/01/04 08:00:20.191__06 <*HV*> szInfo: SoftArcReg 0x211* 2021/01/04 08:00:20.242 <*HV*> Type: NORMAL ErrNum: S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT szInfo: P1:0/P2: 2021/01/04 08:00:20.263 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_ANALOG_FAULT szInfo:2:5 Analog Fault P1:2/P2:32 2021/01/04 08:00:20.265 [CANCEL] CStateShot:Series 5 canceled, due to GANTRY_CANCEL_SOURCE_NORMAL_ERROR, in state CStateShot 2021/01/04 08:00:20.272__02 [***] - - - - - - Study Exit - - - - - - 2021/01/04 08:00:20.273 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_ANODE_TUBE_ARC 2021/01/04 08:00:20.288 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_CATHODE_TUBE_ARC 2021/01/04 08:00:20.292 <WARN> Type: WARN ErrNum: S_GEN_MGR_SYNC_OFF_OUT_OF_SEQUENCE P1:0/P2:0 2021/01/04 08:00:22.309__01 <*ERROR*> Type: NORMAL ErrNum: S_MASTER_SEQUENCER_EXIT_STATE_ERR P1:4096/P2:8369 2021/01/04 08:00:37.308 <*ERROR*> Type: NORMAL ErrNum: S_TIMER_WATCHDOG_CAUSED_LAST_RESET P1:4096/P2:0 2021/01/04 08:00:38.288 <WARN> Type: WARN ErrNum: S_DTH_USER_ALREADY_REGISTERED P1:0/P2:0 2021/01/04 08:00:38.292 <WARN> Type: WARN ErrNum: S_GAFSTARTUPTASK_DTH_BAD_PARAM P1:1/P2:-1 2021/01/04 08:00:38.296 <WARN> Type: WARN ErrNum: S_GAFSTARTUPTASK_DTH_BAD_PARAM P1:1/P2:-1 2021/01/04 08:00:38.299 <WARN> Type: WARN ErrNum: S_GAFSTARTUPTASK_DTH_BAD_PARAM P1:1/P2:-1 2021/01/04 08:00:38.302 <WARN> Type: WARN ErrNum: S_GAFSTARTUPTASK_DTH_BAD_PARAM P1:1/P2:-1 2021/01/04 08:00:38.336__01 <WARN> Type: WARN ErrNum: S_ROTOR_DTH_BAD_PARAM P1:-1/P2:0 2021/01/04 08:00:38.341__03 <WARN> Type: WARN ErrNum: S_DTH_USER_ALREADY_REGISTERED P1:0/P2:0 2021/01/04 08:00:38.432 <*ERROR*> Type: NORMAL ErrNum: S_MP_MPSHELL_SERIAL_LIFE_SYNC_ERROR szInfo:MP Latest Fatal Error Data. P1:3.670017E+07/P2:3997697 2021/01/04 08:00:38.948__01 <WARN> Type: WARN ErrNum: S_RMC_DTH_BAD_PARAM P1:-1/P2:0 2021/01/04 08:00:38.950 <WARN> Type: WARN ErrNum: S_RMC_MP_DTH_BAD_PARAM P1:-1/P2:0 2021/01/04 08:02:28.347 <WARN> Type: WARN ErrNum: S_MPS_WAITING_FOR_MAIN_DRIVE_TO_CLOSE P1:0/P2:0 szInfo:MP: Waiting for MDC to close.. 2021/01/04 08:02:29.433 <WARN> Type: WARN ErrNum: S_MPS_WAITING_FOR_MAIN_DRIVE_TO_CLOSE P1:0/P2:0 szInfo:MP: Waiting for MDC to close.. 2021/01/04 08:02:30.038 [ESTOP CLOSE] Type:EVENT_MINOR Msg:GantryEstopClosedBkst in CStateConfigured Q4: Is the equipment operating normally following the response? Good. Q5. How did the examination finish? None. Q6. Were there any mistaken findings (misdiagnosis, mistaken treatment) due to the problem? None. Q7. Was there any unintended unnecessary exposure? None. Q8. Were there any artifacts? None. T2 Activities:Problem Description: subject equipment: CT Failure location: 03High Voltage Energy When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence When, in the end user_s work flow, is the problem occurring?: During Examination ・Measurement preparation (Setting up Protocol) What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem? No.The customer has not created a log file or report. Troubleshooting Action: Result after troubleshooting: Improved Failure location: 03High Voltage Energy What steps did you perform to diagnose the problem? An arc occurred.Recovery achieved upon doing a CLOSE again. I had Oyama contact the customer, and the customer will monitor the situation. From the logger, 2021/01/04 07:59:37.166__02 [--o] ----------------- Series 4 Ready ----------------- 2021/01/04 07:59:38.567__01 [-oo] - - - - - - Shot 0 Starts - - - - - - 2021/01/04 07:59:41.673__01 [-oo] - - - - - - Shot 1 Starts - - - - - - 2021/01/04 07:59:44.671__01 [-oo] - - - - - - Shot 2 Starts - - - - - - 2021/01/04 07:59:47.718__01 [-oo] - - - - - - Shot 3 Starts - - - - - - 2021/01/04 07:59:50.720 [-oo] - - - - - - Shot 4 Starts - - - - - - 2021/01/04 07:59:53.719__01 [-oo] - - - - - - Shot 5 Starts - - - - - - 2021/01/04 07:59:56.765__01 [-oo] - - - - - - Shot 6 Starts - - - - - - 2021/01/04 07:59:59.812 [ooo] ----------------- Series 4 Complete ----------------- 2021/01/04 08:00:10.054__06 [---] ------ Series 5 Load Structure ----- 2021/01/04 08:00:17.291__02 [--o] ----------------- Series 5 Ready ----------------- 2021/01/04 08:00:18.366 [-oo] - - - - - - Shot 0 Starts - - - - - - 2021/01/04 08:00:20.191__06 <*HV*> szInfo: SoftArcReg 0x211* 2021/01/04 08:00:20.242 <*HV*> Type: NORMAL ErrNum: S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT szInfo: P1:0/P2: 2021/01/04 08:00:20.263 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_ANALOG_FAULT szInfo:2:5 Analog Fault P1:2/P2:32 2021/01/04 08:00:20.265 [CANCEL] CStateShot:Series 5 canceled, due to GANTRY_CANCEL_SOURCE_NORMAL_ERROR, in state CStateShot 2021/01/04 08:00:20.272__02 [***] - - - - - - Study Exit - - - - - - 2021/01/04 08:00:20.273 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_ANODE_TUBE_ARC 2021/01/04 08:00:20.288 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_CATHODE_TUBE_ARC 2021/01/04 08:00:20.292 <WARN> Type: WARN ErrNum: S_GEN_MGR_SYNC_OFF_OUT_OF_SEQUENCE P1:0/P2:0 2021/01/04 08:00:22.309__01 <*ERROR*> Type: NORMAL ErrNum: S_MASTER_SEQUENCER_EXIT_STATE_ERR P1:4096/P2:8369 2021/01/04 08:00:37.308 <*ERROR*> Type: NORMAL ErrNum: S_TIMER_WATCHDOG_CAUSED_LAST_RESET P1:4096/P2:0 2021/01/04 08:00:38.288 <WARN> Type: WARN ErrNum: S_DTH_USER_ALREADY_REGISTERED P1:0/P2:0 2021/01/04 08:00:38.292 <WARN> Type: WARN ErrNum: S_GAFSTARTUPTASK_DTH_BAD_PARAM P1:1/P2:-1 2021/01/04 08:00:38.296 <WARN> Type: WARN ErrNum: S_GAFSTARTUPTASK_DTH_BAD_PARAM P1:1/P2:-1 2021/01/04 08:00:38.299 <WARN> Type: WARN ErrNum: S_GAFSTARTUPTASK_DTH_BAD_PARAM P1:1/P2:-1 2021/01/04 08:00:38.302 <WARN> Type: WARN ErrNum: S_GAFSTARTUPTASK_DTH_BAD_PARAM P1:1/P2:-1 2021/01/04 08:00:38.336__01 <WARN> Type: WARN ErrNum: S_ROTOR_DTH_BAD_PARAM P1:-1/P2:0 2021/01/04 08:00:38.341__03 <WARN> Type: WARN ErrNum: S_DTH_USER_ALREADY_REGISTERED P1:0/P2:0 2021/01/04 08:00:38.432 <*ERROR*> Type: NORMAL ErrNum: S_MP_MPSHELL_SERIAL_LIFE_SYNC_ERROR szInfo:MP Latest Fatal Error Data. P1:3.670017E+07/P2:3997697 2021/01/04 08:00:38.948__01 <WARN> Type: WARN ErrNum: S_RMC_DTH_BAD_PARAM P1:-1/P2:0 2021/01/04 08:00:38.950 <WARN> Type: WARN ErrNum: S_RMC_MP_DTH_BAD_PARAM P1:-1/P2:0 2021/01/04 08:02:28.347 <WARN> Type: WARN ErrNum: S_MPS_WAITING_FOR_MAIN_DRIVE_TO_CLOSE P1:0/P2:0 szInfo:MP: Waiting for MDC to close.. 2021/01/04 08:02:29.433 <WARN> Type: WARN ErrNum: S_MPS_WAITING_FOR_MAIN_DRIVE_TO_CLOSE P1:0/P2:0 szInfo:MP: Waiting for MDC to close.. 2021/01/04 08:02:30.038 [ESTOP CLOSE] Type:EVENT_MINOR Msg:GantryEstopClosedBkst in CStateConfigured Repair Action: [Confirmation of problem] Confirmed. Work details: An arc occurred.Recovery achieved upon doing a CLOSE again. I had Oyama contact the customer, and the customer will monitor the situation. From the logger, 2021/01/04 07:59:37.166__02 [--o] ----------------- Series 4 Ready ----------------- 2021/01/04 07:59:38.567__01 [-oo] - - - - - - Shot 0 Starts - - - - - - 2021/01/04 07:59:41.673__01 [-oo] - - - - - - Shot 1 Starts - - - - - - 2021/01/04 07:59:44.671__01 [-oo] - - - - - - Shot 2 Starts - - - - - - 2021/01/04 07:59:47.718__01 [-oo] - - - - - - Shot 3 Starts - - - - - - 2021/01/04 07:59:50.720 [-oo] - - - - - - Shot 4 Starts - - - - - - 2021/01/04 07:59:53.719__01 [-oo] - - - - - - Shot 5 Starts - - - - - - 2021/01/04 07:59:56.765__01 [-oo] - - - - - - Shot 6 Starts - - - - - - 2021/01/04 07:59:59.812 [ooo] ----------------- Series 4 Complete ----------------- 2021/01/04 08:00:10.054__06 [---] ------ Series 5 Load Structure ----- 2021/01/04 08:00:17.291__02 [--o] ----------------- Series 5 Ready ----------------- 2021/01/04 08:00:18.366 [-oo] - - - - - - Shot 0 Starts - - - - - - 2021/01/04 08:00:20.191__06 <*HV*> szInfo: SoftArcReg 0x211* 2021/01/04 08:00:20.242 <*HV*> Type: NORMAL ErrNum: S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT szInfo: P1:0/P2: 2021/01/04 08:00:20.263 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_ANALOG_FAULT szInfo:2:5 Analog Fault P1:2/P2:32 2021/01/04 08:00:20.265 [CANCEL] CStateShot:Series 5 canceled, due to GANTRY_CANCEL_SOURCE_NORMAL_ERROR, in state CStateShot 2021/01/04 08:00:20.272__02 [***] - - - - - - Study Exit - - - - - - 2021/01/04 08:00:20.273 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_ANODE_TUBE_ARC 2021/01/04 08:00:20.288 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_CATHODE_TUBE_ARC 2021/01/04 08:00:20.292 <WARN> Type: WARN ErrNum: S_GEN_MGR_SYNC_OFF_OUT_OF_SEQUENCE P1:0/P2:0 2021/01/04 08:00:22.309__01 <*ERROR*> Type: NORMAL ErrNum: S_MASTER_SEQUENCER_EXIT_STATE_ERR P1:4096/P2:8369 2021/01/04 08:00:37.308 <*ERROR*> Type: NORMAL ErrNum: S_TIMER_WATCHDOG_CAUSED_LAST_RESET P1:4096/P2:0 2021/01/04 08:00:38.288 <WARN> Type: WARN ErrNum: S_DTH_USER_ALREADY_REGISTERED P1:0/P2:0 2021/01/04 08:00:38.292 <WARN> Type: WARN ErrNum: S_GAFSTARTUPTASK_DTH_BAD_PARAM P1:1/P2:-1 2021/01/04 08:00:38.296 <WARN> Type: WARN ErrNum: S_GAFSTARTUPTASK_DTH_BAD_PARAM P1:1/P2:-1 2021/01/04 08:00:38.299 <WARN> Type: WARN ErrNum: S_GAFSTARTUPTASK_DTH_BAD_PARAM P1:1/P2:-1 2021/01/04 08:00:38.302 <WARN> Type: WARN ErrNum: S_GAFSTARTUPTASK_DTH_BAD_PARAM P1:1/P2:-1 2021/01/04 08:00:38.336__01 <WARN> Type: WARN ErrNum: S_ROTOR_DTH_BAD_PARAM P1:-1/P2:0 2021/01/04 08:00:38.341__03 <WARN> Type: WARN ErrNum: S_DTH_USER_ALREADY_REGISTERED P1:0/P2:0 2021/01/04 08:00:38.432 <*ERROR*> Type: NORMAL ErrNum: S_MP_MPSHELL_SERIAL_LIFE_SYNC_ERROR szInfo:MP Latest Fatal Error Data. P1:3.670017E+07/P2:3997697 2021/01/04 08:00:38.948__01 <WARN> Type: WARN ErrNum: S_RMC_DTH_BAD_PARAM P1:-1/P2:0 2021/01/04 08:00:38.950 <WARN> Type: WARN ErrNum: S_RMC_MP_DTH_BAD_PARAM P1:-1/P2:0 2021/01/04 08:02:28.347 <WARN> Type: WARN ErrNum: S_MPS_WAITING_FOR_MAIN_DRIVE_TO_CLOSE P1:0/P2:0 szInfo:MP: Waiting for MDC to close.. 2021/01/04 08:02:29.433 <WARN> Type: WARN ErrNum: S_MPS_WAITING_FOR_MAIN_DRIVE_TO_CLOSE P1:0/P2:0 szInfo:MP: Waiting for MDC to close.. 2021/01/04 08:02:30.038 [ESTOP CLOSE] Type:EVENT_MINOR Msg:GantryEstopClosedBkst in CStateConfigured 【Confirmation content】 Good. Problem Reported by Customer: 03CS: KK: FSE Oyama: E-stop opened around 8:00.>>An arc had occurred.Recovery achieved upon doing a CLOSE again. Safety Questions:Case: 0116771864 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: KENICHI KOJIMA Last updated by: KENICHI KOJIMA 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | P4 |  | > Arc errors (ARC_FAULT, ANODE_TUBE_ARC, CATHOSE_TUBE_ARC) were found. No bugrep was created. No action was taken. | [12 empty] | 116771864 | Fixed | 514000 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00630085 | [5 empty] | Engineering |  |  |  | Closed |  |  | TBD_Escalate to L3 | [8 empty] | System issue | System | IC-HW | Tub Arc | IC-HW-Tub Arc | IC-HW-Tub Arc | CTUDT00630085 | Tub Arc | IC-HW | IC-HW-Tub Arc | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
328 | 44202 | TRUE | Japan | Nikko Memorial Hospital | 341011 | 44202 | 2 | 01CS:KK:アプリ中川さん:昨日の8時40分ごろの検査でDose infoが作られなかった。>>撮影時にエラーは特になし。リコンのログが更新されておりリコン側のエラーは確認できず。 | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 01CS:KK:アプリ中川さん:昨日の8時40分ごろの検査でDose infoが作られなかった。>>撮影時にエラーは特になし。リコンのログが更新されておりリコン側のエラーは確認できず。 ----------------------------------------------------------------------------- | P4 |  | No particular error during shooting. The recon log has been updated and no error can be confirmed on the recon side. Philips application specialist reported that the dose info was not generated after the scans. RCE checked the logs but no error was found. No action taken. Kevin follow up to check | [12 empty] | 116786894 | In Process | 516467 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD_unclear | [5 empty] | Application |  |  |  | WIP |  |  | TBD_unclear | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
329 | 44202 | TRUE | China | Dalian Branch of Shengjing Hospital of China Medical University | 345001 | 44202 | 2 | 02N/Ingenuity Core128/Elite China/Maintenance record | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer: Maintenance record Safety Questions:Case: 0116786931 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Mengen Xie Last updated by: Mengen Xie 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action: Customer need all maintenance records of 2020 of all large devices from Dalian Campus of China Medical University Affiliated Shengjing Hospital, and Liaoning Province Dalian Economic and Technological Development Zone Hospital. Mailed SWO group and Billing Group to provide ----------------------------------------------------------------------------- | P5 |  | Closed - Non-Complaint | [12 empty] | 116786931 | Closed | 519400 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not a defect | [5 empty] | Application | [6 empty] | NAC_Customer report | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
330 | 44205 | TRUE | China | SD Laiyang People Hospital | 345005 | 44205 | 2 | 02N/Ingenuity Core128/Elite China/Image quality failure | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : Image quality failure Safety Questions : Case: 0116807939 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Zhang Jing Last updated by: Zhang Jing 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 图像质量问题 客户问题详细描述或故障现象： 图像质量问题 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： NA 远程诊断与测试结果： 客户反馈冠脉扫描，血管模糊，达不到他们临床要求，之前有应用老师调试过，有改善但是还是达不到要求。希望派资深的人员前来处理 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 做图像性能测试，必要时上escalation和CFM SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 10.01.2021 14:52:18 UTC Xingbo Yang (86038250)服务内容：执行图像质量校正。 服务结果：维修后，设备恢复正常工作。 Internal Remarks : 10.01.2021 14:52:20 UTC Xingbo Yang (86038250) "RDI Accuracy:yes" $Q7: (Confirmed reported problem & unsatisfactory image quality) $Q8: (quick IQ check & need to do air cal ) $Q9: ( perform air calibration ) $Q10: (Test and inspection data refer to: PA & Device returned to full functionality) $Q11: Calibrated tools name(NA) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date: 10/ 15 /2020 ^Current Tube usage: scan second_70294 ----------------------------------------------------------------------------- | P3 |  | Any Bugreport could share? Train customer to performed calibration regually, Check how frequently customer perform calibration | [12 empty] | 116807939 | In Process | 527013 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | NotEngineering-Service | [5 empty] | Engineering |  |  |  | WIP |  |  | quick IQ check & need to do air cal | [8 empty] | System issue | System | NotEngineering | ShouldbeCustomer | NotEngineering-ShouldbeCustomer | NotEngineering-ShouldbeCustomer | NotEngineering-Service | NotEngineering-ShouldbeCustomer | NotEngineering | NotEngineering-ShouldbeCustomer | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
331 | 44204 | TRUE | China | Henan Fengquan District Dakuai | 345009 | 44204 | 2 | Application/Ingenuity Core128/Elite China | BU Q&R Feedback: No Reason: Customer Information FSE Notes : This PO is a Reference Number only Next Step Action : 应用老师 陈险峰 13971190179 跟进 Next Step Action : 回访陈老师，已解决，关单。 Problem Reported by Customer : Application Safety Questions : Case: 0116802196 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Yuan Mingze Last updated by: Yuan Mingze 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | P5 |  | Closed - Non-Complaint | [12 empty] | 116802196 | Closed | 526059 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not a defect | [5 empty] | Application | [6 empty] | NAC_Q&A | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
332 | 44202 | TRUE | China | Henan Fengquan District Dakuai | 345009 | 44202 | 2 | 02N/Ingenuity Core128/Elite China/Transmission parameters | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: Transmission parameters Field Service Engineer Notes:This PO is a Reference Number only Safety Questions:Case: 0116787460 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Qian Wang Last updated by: Qian Wang 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer_s problems: Add Cloud workstation Detailed description of Customer_s problems or malfunctions: Add Cloud workstation It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: NA Remote troubleshoot and test result: Guided to enter Olevel interface and found software error. Restarted the computer and could not enter LAN setting. Consider software issue PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Check software, restore backup, and if necessary, re-install software.add workstation node Memo of Customer request: Service requester would like an engineer to be on the field in the afternoon. Need the engineer to call back service requester SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support External Remarks: 07.01.2021 03:39:58 UTC SUP JCOUSER (86038950) Service details: Customer request service, need to check and add DICOM node, remotely instruct customer over the phone. Service result: Remotely instruct customer over the phone, check and add DICOM node. Other important issues: [None] Internal Remarks : 07.01.2021 03:40:00 UTC Yongping Luo (86038950) $Q7: (Confirm the problem &The Customer did not know How to check and add DICOM node ) $Q8: (The Customer did not know How to check and add DICOM node, remotely guidance the customer check and add DICOM Node) $Q9: (remotely guidance the customer check and add DICOM Node, the device return to full function) $Q10: ( Device returned to service) $Q11: Calibrated Tools N/A ----------------------------------------------------------------------------- | P3 |  | could not enter LAN setting | [12 empty] | 116787460 | In Process | 516685 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | CTUDT00631055 | [5 empty] | Engineering |  |  |  | Closed |  |  | LAN Config Tool | [8 empty] | System issue | System | Console-SW | Lan Config | Console-SW-Lan Config | Console-SW-Lan Config | CTUDT00631055 | Lan Config |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
333 | 44201 | TRUE | China | GD Dongguan Dalingshan Hosp. | 345013 | 44201 | 2 | 02N/Core128/Need to move monitor and computer. | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: Need to move monitor and computer. Safety Questions:Case: 0116779403 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Pan Chuchu Last updated by: Pan Chuchu 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action NTD: Customer want engineer to be onsite The equipment has no malfunction. The customer just wanted to ask the engineer to go to the field to help move the post-processed monitor and put it along with the computer on the other side T2 Activities:Problem Description: Summary of customer_s problems: The equipment has no malfunction. The customer just wanted to ask the engineer to go to the field to help move the post-processed monitor and put it along with the computer on the other side Detailed description of Customer_s problems or malfunctions: The equipment has no malfunction. The customer just wanted to ask the engineer to go to the field to help move the post-processed monitor and put it along with the computer on the other side It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: NA Remote troubleshoot and test result: Contacted the doctor who requested service, who said the engineer from the Equipment Dept had been contacted to move the workstation. The hospital engineer is unwilling to do it, asked the doctor to request service, and asking the engineer from the manufacturer to go to the hospital to handle it.The customer is not in a hurry. Ask the engineer to call back and make an appointment about the time to go on the field. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Contact the customer and make an appointment to move the workstation. SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Next Step Action: Li Zongxian Internal Remarks:05.01.2021 07:56:00 UTC Zongxian Li (86001504) $Q1: () $Q2: () $Q3: () $Q4: () RDI Accuracy:Yes $Q7: (Confirmed the reported problem, ISP need to move to another position) $Q8: (ISP need to move to another position) $Q9: (Moved ISP to another position) $Q10: (Test and inspection data refer to: N/A, no PA required. Device returned to full functionality.) $Q11: $Calibrated Tools name (N/A) S/N (N/A) Next Calibration date (N/A) External Remarks: 05.01.2021 07:55:57 UTC Zongxian Li (86001504) service details: move post processing workstation. Service_s result: device operating normally ----------------------------------------------------------------------------- | P5 |  | 44201 | [12 empty] | 116779403 | In Process | 517114 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not a defect | [5 empty] | Customer | [6 empty] | NAC | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
334 | 44201 | TRUE | China | HLJ Huachuan People Hosp. | 345018 | 44201 | 2 | 02N/Ingenuity Core128/Elite China/The screen is yellow. | BU Q&R Feedback: No Reason: Corrective Maintenance Safety Questions:Case: 0116779816 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Guomei Li Last updated by: Guomei Li 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Problem Reported by Customer: The screen is yellow. Next Step Action: ss Mgr yuan call in T2 Activities:Problem Description: Summary of customer_s problems: The machine crashes intermittently Detailed description of Customer_s problems or malfunctions: The machine crashes intermittently It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: N/A Remote troubleshoot and test result: N/A PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Suggest to replace the hard disc and reinstall the software system. We recommend spare parts 12NC and names: HP TurboDrive 512GB TLC Z4 SSD w/adapter: 459801652541. Memo of Customer request: None SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Next Step Action: PPD 459801652541 *1, out of stock, ordered, FSE Jin jun-86-13796075213 ----------------------------------------------------------------------------- | P3 |  | could you please share Bugreport? FSE create bug report and escalation , he will do it when he on site. | [12 empty] | 116779816 | In Process | 516921 | TRUE | Closed - Done | Closed - Done |  | CTUDT00632520 | [5 empty] | Engineering |  |  |  | Closed |  |  | Yellow screen | [8 empty] | System issue | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00631367 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | CTUDT00631367 | y | y
335 | 44204 | TRUE | China | JS Nanjing M.U.A Yifu Hosp. | 345028 | 44204 | 2 | 02G/Ingenuity Core128/Elite China/Add worklist | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer : Add worklist Safety Questions:Case: 0116800271 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Zixuan Last updated by: Wang Zixuan 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | P3 |  | Closed - Non-Complaint | [12 empty] | 116800271 | Closed | 525394 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not a defect | [5 empty] | Customer | [6 empty] | NAC_Customer request | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
336 | 44203 | TRUE | China | JS Nanjing M.U.A Yifu Hosp. | 345028 | 44203 | 2 | 02G/Ingenuity Core128/Elite China/Add worklist | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer : Add worklist Next Step Action: MM Zhang Yi confirm Site 1 FSE Yao Luping Phone FSE Yao, can resolve remotely, no need to add OP to IA, phone if need to add later. Safety Questions:Case: 0116794954 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Yuan Mingze Last updated by: Yuan Mingze 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | P5 |  | Closed - Non-Complaint | [12 empty] | 116794954 | Closed | 523223 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not a defect | [5 empty] | Customer | [6 empty] | NAC_Customer request | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
337 | 44180 | TRUE | Sweden | Avesta lasarett | 40040 | 44180 | 51 | Helical brain scan, the preview images care displayed with wrong window. | A new installation of CT 5000 with software 4.1.10. in Avesta Sweden. When they scan helical brain, the preview images care displayed with wrong window. In the corner of the images it says C = 30 and W = 70. But the preview images are not displayed as these parameters. They are too bright. see attached images. When they click “Start recon” and get the original images, they are displayed correctly. This issue is only seen on helical brain, but every time they do a helical brain scan. We tried to build a new protocol from factory set, but we saw the same issue in the preview. The system is running normal and is in clinical use. KMR Note: The become Aware date was set to 15-Dec-2020 and the reason for delay: I got Covid19 infected the same day, and forwarded this to PM, FSE and my colleague. No one have done a CF on this. 13-Jan-2021 |  |  | https://share.philips.com/sites/BigBoreRTRefreshFOKCoreTeam/Shared%20Documents/General/Ingenuity%20Bugreps/DEID_2021_01_12_12_26_07_Magnus_CT_Brain.zip | [14 empty] | 533097 | TRUE | Closed - Done | Closed - Done |  | CTUDT00632584 | [5 empty] | Engineering |  |  |  | WIP |  |  | Preview image | [8 empty] | System issue | System | Console-SW | PreviewImage-WW-WC | Console-SW-PreviewImage-WW-WC | Console-SW-PreviewImage-WW-WC | CTUDT00632584 | WWWC | Console-SW | Console-SW-PreviewImage-WW-WC | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
338 | 44160 | TRUE | United States | Central Vermont Medical Center | 40024 | 44160 | 48 | System hangs intermittently sending to PACS | Images are in transfer queue as waiting but never attempt to transfer. Host PC Reboot required to restore communication Customer undergoing network attack; All facilities that had Philips products installed experienced difficulty transferring to PACS. It wasn_t until attack was resolved that the problem persisted and we were aware of a service related issue. Clinical education normally makes these CFR_s when they occur during handover training, and that did not occur. creating it now. 5-Nov-2020 |  |  | PACS transfer issue, bugreport downloading https://www.clouddrive.philips.com.cn/file/rygbo40ulmom9a92px7jv6cbh620lapu# 密码:123 | [14 empty] | 539797 | TRUE | Closed - Done | Closed - Done |  | CTUDT00632127 | [5 empty] | Engineering |  |  |  | Closed |  |  | PACS | [8 empty] | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
339 | 44175 | TRUE | Portugal | Centro Hospitalar de Vila Nova | 40030 | 44175 | 50 | The new CT has had several problems | PR 533642 will address: "tube hot" information on the use of "abd/pelv" protocols with 120 kV and 140kv |  |  | PR 470009 will address: "a squeaking noise when tilting the gantry" Refer to 0116636550 need bugreport | [14 empty] | 533642 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD_Tube hot_bugreport | [5 empty] | Engineering |  |  |  | WIP |  |  | Tube hot | [8 empty] | System issue | System | Gantry-SW | TubHot | Gantry-SW-TubHot | Gantry-SW-TubHot | TBD_Tube hot_bugreport | TubHot | Gantry-SW | Gantry-SW-TubHot | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
340 | 44175 | TRUE | Portugal | Centro Hospitalar de Vila Nova | 40030 | 44175 | 50 | The new CT has had several problems | PR 533603 will address: "stopping the transfer of images to PACS forcing to restart the equipment" |  |  | need bugreport | [14 empty] | 533603 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering |  |  |  | Closed |  |  | PACS | [8 empty] | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
341 | 44195 | TRUE | United States |  | 40015 | 44195 | 53 | 1202CDOD1 Maximum Number Of Hard Arcs Errors | add PR in CTUDT00632177 |  |  | Duplicate with another EMS case 116758107/ PR#507320 | [14 empty] | 533506 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632177 | [5 empty] | Engineering |  |  |  | Closed |  |  | Tube Arc | [8 empty] | System issue | System | IC-HW | Hard Arcs | IC-HW-Hard Arcs | IC-HW-Tub Arc | CTUDT00632177 | Hard Arcs | IC-HW | IC-HW-Hard Arcs | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
342 | 44207 | TRUE | United States | Leesburg Regional Medical Cent | 40022 | 44207 | 3 | Q/O Need new flat screws for flat plated | Work Order #: WO-02759292 13-Jan-2021 Neuman Terry The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: CT Scanner Expected and Unexpected Results: can not use flat pannel User Impact: no flat pallet use Patient Impact: none Current Software Version:4.1.10 Diagnostic performed by Engineer : visual inspection Parts Consumed : 1.00000,453566268391,THUMB SCREW(AQU-TABLE) 1.00000,453566268391,THUMB SCREW(AQU-TABLE) | P3 |  | 1.00000,453566268391,THUMB SCREW(AQU-TABLE) 1.00000,453566268391,THUMB SCREW(AQU-TABLE) | [12 empty] | 116821137 | Closed | 532108 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | NotEngineering-PartReconnection | [5 empty] | Others |  |  |  | WIP |  |  | THUMB SCREW | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
343 | 44208 | TRUE | Portugal | Centro Hospitalar de Vila Nova | 40030 | 44208 | 3 | Presence of ring artifact in CT images. | Work Order #: WO-02761943 12-Jan-2021 Jaime Granja The system fully meets specification and returned to use Regulatory Questions : N/A Problem Description by Engineer: Customer Function/Role: Administrative. Device Use: Diagnostics Expected and Unexpected Results: Ring-shaped artifacts after eventual contrast agent entry into gantry. User Impact: Possible misinterpretation of images. Patient Impact: Possible need for repetition. Current Software Version: 4.1.10.25202 Resolution: The presence of crystallized contrast agent was detected, deposited on the detector_s protection window. Cleaned. Air calibration and tests with head and body phantoms. Monitoring of the first exams after repair. Awaiting for evolution to rule out any instability of any of the modules caused by the contrast deposit still in a liquid state. As a precaution and considering the safety and hygiene measures imposed by the COVID-19 coronavirus and to avoid contagion risks to our customers and to Philips staff this work order is not signed by the customer (technician Amanda Silva), who was informed about it and agreed. | P3 |  | contrast entry into gantry, check if have gap in lexan ring. The presence of crystallized contrast agent was detected, deposited on the detector_s protection window. | [12 empty] | 116824046 | Closed | 531504 | TRUE | Closed - Done | Closed - Done |  | ECR-075975 | [5 empty] | Engineering |  |  |  | Closed |  |  | Lexan Ring/Contrast | [8 empty] | System issue | System | Gantry-HW | Lexan Ring | Gantry-HW-Lexan Ring | Gantry-HW-Lexan Ring | ECR-075975 | Lexan Ring | Gantry-HW | Gantry-HW-Lexan Ring | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
344 | 44208 | TRUE | United States |  | 40037 | 44208 | 3 | XR-29 certificate | Work Order #: WO-02765571 Hirchak, Michele The system fully meets specification and returned to use Problem description by Clinical Support Specialist : Customer Function/Role:: Physicist Device Use: In Use Expected and Unexpected Results:XR-29 certificate User Impact: None Patient Impact: None 4.1.71.7Current Software Version:4.1.7 Diagnostic performed by Engineer : EMailed certificate to customer, Chris Beason/APPS, and Joseph Healey, FSE. MBH ----------------------------------------------------------------------------- | P5 |  | NAC | [12 empty] | 116827718 | Closed | 531392 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | NAC | [5 empty] | Customer | [15 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
345 | 44208 | TRUE | Japan | Saiseikai Gosu Hospital | 343000 | 44208 | 3 | SC:Change to Blue Allow volume setting | BU Q&R Feedback: No Reason: Application Support Problem Reported by Customer: SC:Change to Blue Allow volume setting Safety Questions:Case: 0116822052 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: CAI Susie Last updated by: CAI Susie 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A OneEMS internal remarks: Kawashige DB SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Support ----------------------------------------------------------------------------- | P5 |  | is it Blue Go button volume setting? Peter check if this volume setting is a valid complaint | [12 empty] | 116822052 | New | 532528 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | NAC? | [5 empty] | Customer |  |  |  | WIP |  |  | Volume seeting | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
346 | 44209 | TRUE | China | NMG Zizhiqu No.4 Hosp | 40023 | 44209 | 3 | 02N/Ingenuity CT 728326/Need water film. | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: Need water film. Safety Questions:Case: 0116830095 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Yuan Mingze Last updated by: Yuan Mingze 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action: FSE Su Ruijie call in and dispatch him. SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | P5 |  | HCOR/Performance test passed, acceptance test fail (head), FSE is ordering new phantom to check if It is phantom issue. Export the calibration history of failed item as well as DICOM images. Share How to export images to FSE Add follow up case to 116780692 https://www.clouddrive.philips.com.cn/file/f7mr107a8k5ksdxcm07j94zq84zrr6iu# 密码:123 Adjust outer laser, and recalibration, IQ issue fixed | [12 empty] | 116830095 | In Process | 535531 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | TBD_Outer laser adjustment | [5 empty] | Engineering |  |  |  | WIP |  |  | Laser alignment | [8 empty] | System issue | System | Gantry-HW | LaserAlignment | Gantry-HW-LaserAlignment | Gantry-HW-laser | TBD_Outer laser adjustment | gantry laser |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
347 | 44209 | TRUE | Peru | Hogar Clinica San Juan de Dios | 40034 | 44209 | 3 | 40034-A VISIT IS REQUIRED TO ENABLE REMOTE CONNECTION | 40034-A VISIT IS REQUIRED TO ENABLE REMOTE CONNECTION | P5 |  | A VISIT IS REQUIRED TO ENABLE REMOTE CONNECTION, NAC? | [12 empty] | 116836828 | New | 538084 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not a defect | [5 empty] | Customer |  |  |  | Closed |  |  | PRS Setup | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
348 | 44209 | TRUE | United States | Unity Health Systems | 40037 | 44209 | 3 | Need a ip address added to the CT Scanner | Need a ip address added to the CT Scanner |  |  | NAC | [12 empty] | 116835958 | Closed | 537855 | TRUE | Closed - Duplicate | Closed - Duplicate |  | NAC | [5 empty] | Customer |  |  |  | WIP |  |  | NAC | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
349 | 44209 | TRUE | United States | Unity Health Systems | 40037 | 44209 | 3 | 0308CDFD1 Mechanical Screeching Noise During Couch Vertical Movements | Work Order #: WO-02772770 Holeman, David The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Problem description by engineer : Mechanical Screeching Noise During Couch Vertical Movements When was the first occurrence? Within... : 13-Jan-2021 08:00:00 How often is issue occurring? : Continuously Malfunction area : Table Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Not provided Expected and Unexpected results: Mechanical Screeching Noise During Couch Vertical Movements Patient Impact/Outcome: None User Impact: Noise During Couch Vertical Movement Current Software Version: Not provided Diagnostic performed by Engineer : Tech reports Clinical Applications training is onsite and the patient support started making a mechanical "screeching" noise today during vertical (up/down) movements. Internal Comments : Error # and/or description of error : Mechanical Screeching Noise During Couch Vertical Movements Completed or Recommended Repair Action "Internal" : Inspect couch vertical operation and inspect the following for binding or noise: covers and mechanics, ball screw, slides, thrust bearing, Helper spring and scissor arm for mechanical binding or noise. Clean, relubricate and adjust the vertical movement mechanics: ballscrew, slide rails, bearings, scissor arm and Helper spring. If necessary, replace any failed components. Follow Up Required : Error # and/or description of error : Mechanical Screeching Noise During Couch Vertical Movements Completed or Recommended Repair Action "Internal" : Inspect couch vertical operation and inspect the following for binding or noise: covers and mechanics, ball screw, slides, thrust bearing, Helper spring and scissor arm for mechanical binding or noise. Clean, relubricate and adjust the vertical movement mechanics: ballscrew, slide rails, bearings, scissor arm and Helper spring. If necessary, replace any failed components. ----------------------------------------------------------------------------- Work Order #: WO-02772828 13-Jan-2021 Shultz Bob The system fully meets specification and returned to use Problem Description by Engineer : Problem description by engineer : Mechanical Screeching Noise During Couch Vertical Movements When was the first occurrence? Within... : 13-Jan-2021 08:00:00 How often is issue occurring? : Continuously Malfunction area : Table Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Not provided Expected and Unexpected results: Mechanical Screeching Noise During Couch Vertical Movements Patient Impact/Outcome: None User Impact: Noise During Couch Vertical Movement Current Software Version: Not provided Resolution : Noise when table is moved vertically ----------------------------------------------------------------------------- | P3 |  | Couch noise The couch rods needing grease within 30 days of install seems to be a somewhat common theme. Ada support to check with OPS how grease (old or new?) add and control. Couch SN: 925035 | [12 empty] | 116835700 | Incomplete | 534528 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | E-125059 | [5 empty] | Engineering |  |  |  | WIP |  |  | Couch Noise | [8 empty] | System issue | System | Couch-HW | Lubricant | Couch-HW-Lubricant | Couch-HW-Lubricant | E-125059 | Lubricant | Couch-HW | Couch-HW-Lubricant | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
350 | 44210 | TRUE | United Kingdom | STEPPING HILL HOSPITAL | 40016 | 44210 | 3 | There is no horizontal movement on the table. | Work Order #: WO-02780300 Graeme Baxter The system partly meets specification and is returned to use with restrictions Regulatory Questions : N/A Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Follow Up Required : issue with the couch Horizontal movement , when can`t access the system remotely , needs an FSE to site to download logs and investigate Diagnostic performed by Engineer : issue with the couch Horizontal movement , when can`t access the system remotely , needs an FSE to site to download logs and investigate | P2 |  | FSE forget to do CPM programing, After program the GMP board, issue fixed | [12 empty] | 116842856 | In Progress | 537815 | TRUE | Closed - Done | Closed - Done |  | CTUDT00630138/630160 | [5 empty] | Engineering |  |  |  | Closed |  |  | FCO72800735 | [8 empty] | System issue | System | Couch-SW | Horizontal Move Block | Couch-SW-Horizontal Move Block | Couch-SW-Horizontal Move Block | CTUDT00630138/630160 | Horizontal Move Block | Couch-SW | Couch-SW-Horizontal Move Block | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
351 | 44210 | TRUE | Puerto Rico | HOSPITAL METROPOLITANO SAN GER | 40017 | 44210 | 3 | System locked | Work Order #: WO-02781535 Valenzuela Velasquez, Eduardo Ariel The system does not meet specification and is out of use Problem Description by Engineer : The user account is locked. Follow Up Required : Reset the user account. Enable also the IT user account: the user manages the account there. Diagnostic performed by Engineer : The user account is locked and the remote connection is down. It is needen to reset the user account. | P2 |  | defect to be raised_Password Reset The password reset lock time should be updated. | [12 empty] | 116843995 | In Progress | 538217 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632236 | [5 empty] | Engineering |  |  |  | Closed |  |  | Password reset | [8 empty] | System issue | System | Console-SW | User Account | Console-SW-User Account | Console-SW-User Account | CTUDT00632236 | User Account | Console-SW | Console-SW-User Account | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
352 | 44210 | TRUE | Chile | HOSPITAL BARROS LUCO TRUDEAU | 341005 | 44210 | 3 | CL1688 THE FIRST ACQUISITION IS DELIVERED WITH NOISE IN THE IMAGE | Work Order #: WO-02778522 Valenzuela Velasquez, Eduardo Ariel The system fully meets specification and returned to use Image quality problems. Air calibration is requested. Phantom tests are performed and images are smooth. Then patients are made without presenting problems. | P2 |  | Air calibration is requested. Phantom tests are performed and images are smooth. Then patients are made without presenting problems. Image quality issue, does issue fixed by air calibration? Check how frequently customer perform calibration and can customer calibrate regualy. If possible export the calibration failed DICOM images for investigation follow the PPT. | [12 empty] | 116841497 | Closed | 537889 | TRUE | Closed - Done | Closed - Done |  | Not a defect | [5 empty] | Customer |  |  |  | WIP |  |  | Air calibration | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
353 | 44210 | TRUE | Japan | Saiseikai Gosu Hospital | 343000 | 44210 | 3 | 07CS:KK:Mr. Moriwaki: He wanted to contact FSE Kawashige.>> I asked Kawashige to call him back. | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer : 07CS:KK:Mr. Moriwaki: He wanted to contact FSE Kawashige.>> I asked Kawashige to call him back. ----------------------------------------------------------------------------- | P5 |  | NAC? | [12 empty] | 116839252 | Closed | 535954 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | NAC? | [5 empty] | Customer | [15 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
354 | 44211 | TRUE | United States |  | 40015 | 44211 | 3 | XR 29 compliance certificate | Work Order #: WO-02794391 Curry, Dyan The system fully meets specification and returned to use Problem description by Clinical Support Specialist : Customer Function: technologist Device Use: in use Expected and Unexpected Results: XR 29 certificates User Impact: none Patient Impact: none Current Software version: 4.1.7 Diagnostic performed by Engineer : created XR29 certificate , uploaded certificate to TEAMS, emailed to customer | P5 |  | NAC | [12 empty] | 116855945 | In Progress | 541177 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | NAC | [5 empty] | Customer | [15 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
355 | 44211 | TRUE | China | GX Tengxian People_s Hospital | 40036 | 44211 | 3 | 02n/Ingenuity CT 728326/Machine preheating failure & key failure | BU Q&R Feedback: No SAP Customer Symptom Code: BOOT Customer Symptom Code: Boot up/Power-up Problem Reason: Corrective Maintenance Engineer Failure Code: Electrical/Electronic Failure Repair Activity Code: Reboot/Restart Problem Reported by Customer : Machine preheating failure & key failure Safety Questions : Case: 0116846910 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Lili Last updated by: Lu Lili 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A FSE Notes : This PO is a Reference Number only T2 Activities : Problem Description: Machine preheating failure & key failure Troubleshooting Action: check log, mdc not closing Repair Action: restart gantry Test/Inspection: estop close normal ----------------------------------------------------------------------------- | P3 |  | Machine preheating failure & key failure check log, mdc not closing restart gantry need time of issue happened CAN Trace downloaded: https://www.clouddrive.philips.com.cn/file/4iffbmq2mu5hwdx1a5py1t3v62h09cel# 密码:123 @Jerry | [12 empty] | 116846910 | Fixed | 542180 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD_Restart | [5 empty] | Engineering |  |  |  | WIP |  |  | Reboot | [8 empty] | System issue | System | Gantry-HW | BootUpFailure | Gantry-HW-BootUpFailure | Gantry-HW-BootUpFailure | TBD_Restart | BootUpFailure | Gantry-HW | Gantry-HW-BootUpFailure | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
356 | 44211 | TRUE | China | SD Anqiu T.C.M Hosp. | 345017 | 44211 | 3 | 02N/Ingenuity Core128/Elite China/Data out of sync | BU Q&R Feedback: No Reason: Corrective Maintenance Safety Questions : Case: 0116845537 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Qian Wang Last updated by: Qian Wang 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Problem Reported by Customer : Data out of sync SAP Transfer : Case Transferred to SAP T2 Activities : Problem Description: ECG signal has noise and HR too low Troubleshooting Action: Check ECG signal wave, base is impacted by 50 Hz electric noise, check earth label and fixed. HR issue has no repeat, Nemo ECG cable and monitor Repair Action: check earth label and fixed, nemo ECG cable and monitor Test/Inspection: N/A Test Equipment: N/A ----------------------------------------------------------------------------- | P3 |  | nemo ECG cable and monitor | [12 empty] | 116845537 | Fixed | 541590 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | NAC? Peter confirm | [5 empty] | Customer |  |  |  | WIP |  |  | ECG | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
357 | 44209 | TRUE | China | SD Chest Hosp. | 40025 | 44209 | 3 | Estop opened and log report | E-Stop during couch movement |  |  | Philips aware date是客户向飞利浦员工投诉的日期，请确认12/28/2020是客户报修的日期，还是客户投诉的日期？若是投诉日期，因已超过5个工作日提交CFM表, 请详述为何迟交CFM表的原因在CFM表的reason for delay. | [14 empty] | 532347 | TRUE | Closed - Duplicate | Closed - Duplicate |  | CTUDT00632187 | [5 empty] | Engineering |  |  |  | Closed |  |  | Horizontal movement | [8 empty] | System issue | System | Couch-SW | Estop | Couch-SW-Estop | Couch-SW-Estop | CTUDT00632187 | Estop |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
358 |  |  | United States |  | 310115 | 44177 | 50 | New Upgrade - Password Not Working | Work Order #: WO-02587689 Carpenter, Kevin The system does not meet specification and is out of use Resolution : Cust called FSE directly as he had left his cell# onsite b/c the SWare upgrade was completed afterhours on Friday when the customer was gone. He helped them understand how to use the Password that he had left onsite. Diagnostic performed by Engineer : Attempted older standard passwords - none worked CT RSE also unable to connect via Telnet SSH Problem Description by Engineer : Customer Function/Role: CTTech Mandy Device Use: Diag Expected and Unexpected Results: New Upgrade - Password Not Working User Impact: None Patient Impact: Cant Scan Current Software Version: IDK - Recently upgraded & no PRS ----------------------------------------------------------------------------- | [15 empty] | 116649070 | Closed | 472464 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not a Defect | [5 empty] | Customer | [15 empty] | Upgrade issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
359 |  |  | United States | NORTHWEST RADIOLOGY/MRI | 300057 | 44181 | 51 | 0701CDFK1 CT Clin Apps is OnSite - Unable to record breathing instructions. | Work Order #: WO-02615589 Powell, Jamie The system partly meets specification and is returned to use with restrictions Diagnostic performed by Engineer : No diags Problem Description by Engineer : Customer Function/Role: CTTech Bridget Device Use: Diag Expected and Unexpected Results: Can record new breathign instructions User Impact: None Patient Impact: None Current Software Version: Unknown Follow Up Required : CT Clin Apps is OnSite - Unable to record breathing instructions. | P5 |  | I know exactly what is going on. When I did the Win 10 software upgrade, it had me disconnect the line in and out. After the load, I plugged them back in and the CT box screamed. Had to leave line in unplugged. I will have to plug back in and see if I still have the issue. May have just been feedback. Cable connection issue | [12 empty] | 116676660 | Closed | 478616 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not a Defect | [5 empty] | Service |  |  |  | Closed |  |  | Line in out cable | [5 empty] | Refresh |  |  | Upgrade issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
360 |  |  | United States | Upson Regional Medical Center | 310115 | 44182 | 51 | 0711CDFC2 After upgrade protocols changed | Work Order #: WO-02631005 Boyd Craig The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: After upgrade protocols changed User Impact: None Patient Impact: None Current Software Version: Diagnostic performed by Engineer : After upgrade protocols changed. Unable to access protocols to make changes. PRS access not available for Telnet SSH2. Tech tried CT as password but did not work. Follow Up Required : FSE on site to change password so tech_s have access. Confirm settings for remote access. ----------------------------------------------------------------------------- Work Order #: WO-02631181 18-Dec-2020 Lee John The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: After upgrade protocols changed User Impact: None Patient Impact: None Current Software Version: Resolution : Instructed customer ----------------------------------------------------------------------------- | P3 |  | FSE on site to change password so tech_s have access, application training need enhance | [12 empty] | 116692043 | Closed | 481697 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not a Defect | [5 empty] | Customer | NAC |  |  | Closed |  |  | User account | Upgrade |  |  |  |  | Refresh |  |  | Upgrade issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
361 |  |  | United States | NORTHWEST RADIOLOGY/MRI | 300057 | 44203 | 2 | CIRS is making a squeaking noise. | Work Order #: WO-02733817 Mark Finnigan The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: CIRS is making a noise. It is still functioning. User Impact: None Patient Impact: None Current Software Version: Diagnostic performed by Engineer : Spoke with Bridgett Williams. She states that the CIRS started making a squealing noise today as if a fan were going bad. Open CIRS find the source of the problem. Follow Up Required : Spoke with Bridgett Williams. She states that the CIRS started making a squealing noise today as if a fan were going bad. Open CIRS find the source of the problem. ----------------------------------------------------------------------------- Work Order #: WO-02733877 07-Jan-2021 Williams John The system partly meets specification and is returned to use with restrictions Diagnostic performed by Engineer : Found nsource of noise. Also ACQ errors Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: CIRS is making a noise. It is still functioning. User Impact: None Patient Impact: None Current Software Version: Follow Up Required : Replace fan and Optic tx/rx | P4 |  | Site is a FOK Ingenuity running 4.1.10 software. Issues with resends . Provided approval to order new ACQ as the TX/RX were already replaced. Also found connection issue on RCOM. Set up remote connection PRS, RemoteService, Firewall, and COPSSH and was able to connect to the system. Collected BugReport to upload for L3 evaluation. Software may need to be reloaded due to a boot up issue. 1.00000,453567538334,MODULE, ROTOR ELECTRONICS TX 1.00000,453567538322,Module, Optical RX Stator Moog Velocity 1.00000,459800992271,Z840 Processor/Memory Fans 1.00000,459800992281,Z840 Rear System Fans Diamond selection system, those issue are not related to Refresh scope. | [12 empty] | 116796880 | In Progress | 521038 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not a defect (Diamond selection) | [5 empty] | Others |  |  |  | Closed |  |  | Diamond select | [5 empty] | Refresh |  |  | Upgrade issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
362 |  |  | United States | Lakeland Regional Medical Cent | 320474 | 44204 | 2 | The customer contacted me because they were unable to copy exam cards (ECs) from their Ingenuity CT scanner (v4.1.7) to |  |  |  | Exam lost after Upgrade from 4.1.7 to 4.1.10, how did FSE do exam cards buckup and restore? Did FSE try to follow upgrade manual and see if issue still happen? | [14 empty] | 535037 | #N/A | #N/A | Closed - Done |  | CTUDT00632298 | [5 empty] | Engineering |  |  |  | Closed |  |  | Exam Card | Upgrade |  |  |  |  | Refresh |  |  | Upgrade issue | System | Console-SW | ExamCard | Console-SW-ExamCard | Console-SW-ExamCard | CTUDT00632298 | EC Upgrade | Console-SW | Console-SW-ExamCard | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
363 |  |  | United States | Lakeland Regional Medical Cent | 320474 | 44209 | 3 | Q/O: CFM needs to be filled out concerning Windows 10 upgrade | Work Order #: WO-02772540 13-Jan-2021 Watson Adam The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : Grabbed bug report from system, and created a CFM. | P5 |  | is it duplicate with above CFM | [12 empty] | 116835566 | In Progress | 535117 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | TBD | [5 empty] | TBD |  |  |  | Upgrade | [11 empty] | Upgrade issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
364 |  |  | United States | Lakeland Regional Medical Cent | 320474 | 44211 | 3 | Q/O: Customer cannot log into CT user account | Work Order #: WO-02792863 15-Jan-2021 Watson Adam The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : Customer reached out to engineer stating the tech cannot log in the CT user account. Engineer remotely worked with the CT tech to log into the system. There was no problem with the system. The tech was able to log in successfully after engineer walked them through the Win 10 log in process. | P5 |  | Password management changes in Win10 | [12 empty] | 116854675 | In Progress | 540761 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not a Defect | [5 empty] | Customer |  |  |  | Closed |  |  | User account | Upgrade |  |  |  |  | Refresh |  |  | Upgrade issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
365 |  |  | United States | St Joseph Westgate Medical Center | 40001 | 44218 | 4 | Gantry Cant Comply | The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Biomed Device Use: Diag Expected and Unexpected Results: Gantry Cant Comply Errors User Impact: No Patient Impact: No Current Software Version: 4 Diagnostic performed by Engineer : All ports closed. Biomed (Kevin) 509-934-9165 states it has happened 2 or 3 times today and would like a FSE Dispatched. Follow Up Required : Analyze logs locally. All ports appear to be closed in PRS. Kevin (Biomed) 509-934-9165 states it has happened 2 or 3 times today and would like FSE dispatched. | 3 |  | All ports appear to be closed in PRS. Kevin (Biomed) 509-934-9165 states it has happened 2 or 3 times today and would like FSE dispatched. Need bugreport | [12 empty] | 116901062 | In Progress | 557028 | TRUE | Closed - Done | Closed - Done |  | TBD_Bugreport | [5 empty] | Engineering | Need get more detail about the issue and what action |  |  | WIP |  |  | Gantry can not comply | [8 empty] | System issue | System | TBD-Bugreport | TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | TBD_Bugreport | Upg | TBD-Bugreport | TBD-Bugreport-TBD_Bugreport | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
366 |  |  | United States | St Joseph Westgate Medical Center | 40001 | 44217 | 4 | PM Inquiry | The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Biomed Device Use: No Patient Involved Expected and Unexpected Results: Need PM Plan/Coop Helpdesk Info User Impact: No Patient Impact: No Current Software Version: Unknown Diagnostic performed by Engineer : Updated Mr. Freedman_s email address in Service Max. Provided case# created and emailed Coop Helpdesk/PM Planner Contact information. Resolution : Updated Mr. Freedman_s email address in Service Max. Provided case# created and emailed Coop Helpdesk/PM Planner Contact information. | 5 |  | PM Planning | [12 empty] | 116893140 | Closed | 553728 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not a defect | [5 empty] | Customer | NAC | [14 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
367 |  |  | Puerto Rico | Manati Medical Center | 40007 | 44217 | 4 | NOISE IN THE GANTRY WHILE BROKEN | Work Order #: WO-02828430 Caldelas Aldana, Carlos Manuel The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: user Device Use: diagnostic Expected and Unexpected Results: noise abnormal User Impact: noise when using it Patient Impact: n/a Current Software Version:unknown Diagnostic performed by Engineer : ther is a noise when gantry is rotating like friction or something hitting insider covers. no errors or other issues reported. Follow Up Required : check rotor elements for something loose or cover alignment | 5 |  | Next step is FSE onsite to check rotor elements for something loose or cover alignment | [12 empty] | 116892005 | In Progress | 554327 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD_Noise | [5 empty] | Engineering | TBD |  |  | WIP |  |  | Gantry Noise | [8 empty] | System issue | System | Gantry-HW | TBD_Noise | Gantry-HW-TBD_Noise | Gantry-HW-TBD_Noise | TBD_Noise | TBD_Noise | Gantry-HW | Gantry-HW-TBD_Noise | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
368 |  |  | United States | Lackey Memorial Hospital | 40015 | 44217 | 4 | Sending Images to V-Rad After Hrs Having Issues | Denison Chuck The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: X-ray Manager Device Use: Philips C.T. 5000 Expected and Unexpected Results: None. Waiting on patch. User Impact: None Patient Impact: None Current Software Version: 4.1.10 Resolution : Worked creating an CFS for the customer and working with my NSS on the problem of sending images over to V-Rad services. | 5 |  | PACS Transfer issue | [12 empty] | 116892035 | Closed | 556091 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering | Collect screenshot or failed transfer and bugreport for engineering investigation |  |  | Closed |  |  | PACS | [8 empty] | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
369 |  |  | United States | Lackey Memorial Hospital | 40015 | 44215 | 4 | QO: Sending Images to V-Rad After Hrs Having Issues | Denison Chuck The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: X-ray manager Device Use: C.T. Expected and Unexpected Results: None User Impact: None Patient Impact: None Current Software Version:4.1.10 Resolution : Worked on the C.T. sending to V-rad. Change some settings around. Monitor system. Follow Up Required : Need to do Customer Feed Back and Escalation on the case. | 3 |  | PACS Transfer issue | [12 empty] | 116878221 | Closed | 550815 | TRUE | Closed - Done | Closed - Done |  | CTUDT00632127 | [5 empty] | Engineering | Collect screenshot or failed transfer and bugreport for engineering investigation |  |  | Closed |  |  | PACS | [8 empty] | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
370 |  |  | United States | Lackey Memorial Hospital | 40015 | 44215 | 4 | Not sending images as should to V-Rad after hour services | Duplicate 550815 | CFM |  | Duplicate 550815 | [14 empty] | 554151 | TRUE | Closed - Duplicate | Closed - Duplicate |  | CTUDT00632127 | [5 empty] | Engineering | Duplicate 550815 |  |  | Closed |  |  | PACS | [8 empty] | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Remote Connection | CTUDT00632127 | Remote Connection |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
371 |  |  | United States | Lackey Memorial Hospital | 40015 | 44195 | 53 | CT laser problem, The problem occurred during PM (preventive maintenance) The case was closed as No further investigation by SRRT, addressed by PR 547541; | CT laser problem, The problem occurred during PM (preventive maintenance) The case was closed as No further investigation by SRRT, addressed by PR 547541; Resolution Description: adjusted lasers, found table had shfted realigned and retorqued bolts, replaced bad bolt, too long | 5 |  | Resolution Description: adjusted lasers, re-alien lasers in front cover. | [12 empty] | Source case116758107 |  | 547541 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD_Outer laser adjustment | [5 empty] | Engineering | Need get more detail about which laser (inner or outer?) |  |  | WIP |  |  | Laser alignment | [8 empty] | System issue | System | Gantry-HW | LaserAlignment | Gantry-HW-LaserAlignment | Gantry-HW-GantryLaser | TBD_Outer laser adjustment | gantry laser | Gantry-HW | Gantry-HW-LaserAlignment | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
372 |  |  | United States | Lackey Memorial Hospital | 40015 | 44195 | 53 | Network switcher problem, The problem occurred during PM (preventive maintenance) The case was closed as No further investigation by SRRT, addressed by PR 547542; | Network switcher problem, The problem occurred during PM (preventive maintenance) The case was closed as No further investigation by SRRT, addressed by PR 547542; | 5 |  | Changed out the Network switcher. / Switched the CRC door around for better excess | [12 empty] | Source case116758107 |  | 547542 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | NotEngineering-Connection | [5 empty] | Others | Need get more detail about the switch issue |  |  | WIP |  |  | Network switch | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
373 |  |  | United States | Lackey Memorial Hospital | 40015 | 44195 | 53 | Closed - Non-Complaint |  | 5 | [14 empty] | Source case116758107 |  | 547543 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not a defect | [5 empty] | Customer | Closed - Non-Complaint | [14 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
374 |  |  | United States | Lackey Memorial Hospital | 40015 | 44195 | 53 | Table shifted. The problem occurred during PM (preventive maintenance) The case was closed as No further investigation by SRRT, addressed by PR 552777 | Table shifted. The problem occurred during PM (preventive maintenance) The case was closed as No further investigation by SRRT, addressed by PR 552777 | 5 |  | found table had shfted realigned and retorqued bolts, replaced bad bolt, too long , Had to move couch base over and re-torque bolts | [12 empty] | Source case116758107 |  | 552777 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not a defect | [5 empty] | Service | Need get more detail about the table shift issue |  |  | Closed |  |  | Table fixed | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
375 |  |  | Puerto Rico | HOSPITAL METROPOLITANO SAN GER | 40017 | 44218 | 4 | SOFTWARE REFILL IS REQUIRED TO COMPLETE THE FCO72800737 | SOFTWARE REFILL IS REQUIRED TO COMPLETE THE FCO72800737 | 5 |  | FCO planning and create ticket | [12 empty] | 116899958 | In Progress | 558670 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not a defect | [5 empty] | Customer | NAC | [5 empty] | HF2 | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
376 |  |  | United States | Leesburg Regional Medical Cent | 40022 | 44215 | 4 | 0798CDOD1 Studies Arriving In Random Order To Destination Nodes | The system fully meets specification and returned to use Problem Description by Engineer : Problem description by engineer : Studies Arriving In Random Order To Multiple Destination Nodes When was the first occurrence? Within... : 05-Jan-2021 08:00:00 How often is issue occurring? : Continuously Malfunction area : Console Information to support the complaint handling process Customer Function/Role: IT/PACS Admin Device Use: Diagnostic Expected and Unexpected results: Studies Arriving In Random Order To Multiple Destination Nodes Patient Impact/Outcome: None User Impact: Studies Arriving In Random Order To Multiple Destination Nodes Current Software Version: 4.1.10.25202 Diagnostic performed by Engineer : Michelle, CT Tech states ongoing issue since installation: Studies Arriving In Random Order To Destination Nodes. Systeme installed about 2 weeks ago and this has been an issue since installation. FSE and Clinical Applications aware of issue and recommended PACS side configure for proper sorting. Tech states when studies (in correct order on the CT side) are sent to PACS or Terra RECON, the study series and sometimes images are arriving in a jumbled, out-of-order, random order. Advised Tech that on the Philips CT side, the LAN configuration is for IP address, AE title, port and some other settings, but no settings or filters can be configured that "sort" the studies and the order the studies arrive at a DICOM destination node. Parm, PACS Rep is requesting onsite support to engage with FSE and determine what needs to be configured in order for the incoming studies to be sorted properly. System software version: v4.1.10.25202 Internal Comments : Error # and/or description of error : Studies Arriving In Random Order To Multiple Destination Nodes Completed or Recommended Repair Action "Internal" : Work with PACS Rep and Terra RECON Rep to identify and assist with Sorting/Filtering configuration settings needed to properly sort incoming studies. If issue identified on Philips CT side software version: v4.1.10.25202, submit CFM. Consult with NSS as necessary. Follow Up Required : Error # and/or description of error : Studies Arriving In Random Order To Multiple Destination Nodes Completed or Recommended Repair Action "Internal" : Work with PACS Rep and Terra RECON Rep to identify and assist with Sorting/Filtering configuration settings needed to properly sort incoming studies. If issue identified on Philips CT side software version: v4.1.10.25202, submit CFM. Consult with NSS as necessary. ----------------------------------------------------------------------------- Work Order #: WO-02809902 21-Jan-2021 Neuman Terry Problem Description by Engineer : Problem description by engineer : Studies Arriving In Random Order To Multiple Destination Nodes When was the first occurrence? Within... : 05-Jan-2021 08:00:00 How often is issue occurring? : Continuously Malfunction area : Console Information to support the complaint handling process Customer Function/Role: IT/PACS Admin Device Use: Diagnostic Expected and Unexpected results: Studies Arriving In Random Order To Multiple Destination Nodes Patient Impact/Outcome: None User Impact: Studies Arriving In Random Order To Multiple Destination Nodes Current Software Version: 4.1.10.25202 Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 3 |  | PACS Rep is requesting onsite support to engage with FSE and determine what needs to be configured in order for the incoming studies to be sorted properly. | [12 empty] | 116875783 | In Progress | 546672 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD_PACS config | [5 empty] | Others | Is it PACS configuration issue? Or CT config issue? |  |  | Closed |  |  | PACS | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
377 |  |  | China | NMG Zizhiqu No.4 Hosp | 40023 | 44218 | 4 | 02N/Ingenuity CT 728326/stand by | Next Step Action : FSE：苏瑞杰-13304719335 call in and disp him SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : Su Ruijie External Remarks : 25.01.2021 06:00:41 UTC Ruijie Su (86039043)服务结果： [观察设备调整图像。设备恢复正常。] 其它重要事项：无 Internal Remarks : 25.01.2021 06:00:44 UTC Ruijie Su (86039043) Q7: Confirmed reported problem,stand by $Q8: On sit service stand by . need adjust image $Q9: adjust image max MAS form 300 to 350. . and filter form YA to B . . . system is ok! $Q10:Test and inspection data refer to: PA. Device returned to full functionality. $Q11: Calibrated Toolsname ( Multimeter Fluke 287C) S/N (31060028) Next Calibrationdate (29/10/2021) 2( Torque Wrench Kit ) S/N E201601998 Next Calibration date09/02 /2021 For Mainland China CV / XR, CT engineers only.Please fill in: ^ Current Tube installed date 10/01/2017 ^ Current Tube usage: scan second 10376 | 5 |  | Support For 疾控中心检测, issue was fixed by re-calibrated test, the stand by case is follow up for 116896529 | [12 empty] | 116896236 | New | 557637 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | TBD_Outer laser adjustment | [5 empty] | Engineering | NAC？ FSE ask CCC to update |  |  | WIP |  |  | Laser alignment | [8 empty] | System issue | System | Gantry-HW | LaserAlignment | Gantry-HW-LaserAlignment | Gantry-HW-laser | TBD_Outer laser adjustment | gantry laser |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
378 |  |  | China | NMG Zizhiqu No.4 Hosp | 40023 | 44218 | 4 | 02N/Ingenuity CT 728326/Image quality failure | External Remarks : 25.01.2021 03:20:06 UTC Ruijie Su (86039043)服务结果： [调整激光灯和校正。设备恢复正常。] 其它重要事项：无 Internal Remarks : 25.01.2021 03:20:09 UTC Ruijie Su (86039043) Q7: Confirmed reported problem,Image quality failure $Q8: On sit service Image quality failure . need adjust laster light and air calibration. $Q9: adjust laster light and air calibration. . . system is ok! $Q10:Test and inspection data refer to: PA. Device returned to full functionality. $Q11: Calibrated Toolsname ( Multimeter Fluke 287C) S/N (31060028) Next Calibrationdate (29/10/2021) 2( Torque Wrench Kit ) S/N E201601998 Next Calibration date09/02 /2021 For Mainland China CV / XR, CT engineers only.Please fill in: ^ Current Tube installed date 10/01/2017 ^ Current Tube usage: scan second 40376 T2 Activities : Problem Description: 客户问题概述： 其他类型故障。 客户问题详细描述或故障现象： 质控部门检测图像质量不达标需要工程师到现场支持 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 图像质量参数不达标 远程诊断与测试结果： 需要到场做图像质量校正 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 1，到场做quick IQ检查图像质量2，做水模校正3，和质控部门一起做图像质量检查 Test/Inspection: N/A Test Equipment: N/A SAP Transfer : Dispatched for Onsite Support ----------------------------------------------------------------------------- | 5 |  | Phantom not been replaced, due to the holder was tilted. Support For 疾控中心检测, issue was fixed by re-calibrated test, Outer laser not positioning very well | [12 empty] | 116896529 | In Process | 557821 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD_Outer laser adjustment | [5 empty] | Engineering | Outer laser |  |  | WIP |  |  | Laser alignment | [8 empty] | System issue | System | Gantry-HW | LaserAlignment | Gantry-HW-LaserAlignment | Gantry-HW-GantryLaser | TBD_Outer laser adjustment | gantry laser | Gantry-HW | Gantry-HW-LaserAlignment | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
379 |  |  | China | NMG Zizhiqu No.4 Hosp | 40023 | 44215 | 4 | 02N/Ingenuity CT 728326/Add workstation | BU Q&R Feedback: No SAP Customer Symptom Code: CENV Customer Symptom Code: Customer Environmental Concern Reason: Corrective Maintenance Engineer Failure Code: Connection/Cable/Interconnect Repair Activity Code: Calibration/Adjustment Safety Questions : Case: 0116872713 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Yu Zhao Last updated by: Yu Zhao 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Problem Reported by Customer : Add workstation T2 Activities : Repair Action: 同0116862789 ， 工程师黄勇到场支持 Test/Inspection: N/A Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 | [14 empty] | 116872713 | Closed | 548234 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not a defect | [5 empty] | Customer | Closed - Non-Complaint | [14 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
380 |  |  | China | NMG Zizhiqu No.4 Hosp | 40023 | 44214 | 4 | 02N/Ingenuity CT 728326/Add parameters | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: Add parameters Safety Questions:Case: 0116862789 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingli Last updated by: Lu Mingli 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Dispatched for On site Support T2 Activities:Problem Description: Summary of customer_s problems: Other types of error. Detailed description of Customer_s problems or malfunctions: Added new nodes It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: Added new nodes Remote troubleshoot and test result: Unable to add remotely, need to go to the site PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Added new nodes SAP Transfer:Case Transferred to SAP External Remarks: 19.01.2021 11:55:19 UTC Yong Huang (86038368) service details: [add workstation] Service result: [The device returns to normal working condition after maintenance.] Other important issues: [None] Internal Remarks: 19.01.2021 11:55:22 UTC Yong Huang (86038368) RDI Accuracy:Yes/No $Q7: (Confirm reported problem, Add parameters ) $Q8: (Performed visual inspection, need to add workstation ) $Q9: (Add workstation AE, IP and port ) $Q10: (Test and inspection data refer to: PA. Device returned to full functionality.) $Q11:Calibrated Tools(N/A). ----------------------------------------------------------------------------- | 5 | [14 empty] | 116862789 | Closed | 548948 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not a defect | [5 empty] | Customer | Closed - Non-Complaint | [14 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
381 |  |  | Portugal | Centro Hospitalar de Vila Nova | 40030 | 44217 | 4 | Imagem artefactual no TC do SU da PHILIPS Ingenuity | The system fully meets specification and returned to use Problem Description by Engineer: Customer Function / Role: Radiology Technician Device Use: Diagnostics Expected and Unexpected Results: Intermittent artifact in the form of a central point. User Impact: Risk of misinterpretation Patient Impact: Risk of misinterpretation Current Software Version: 4.1.10.25202 Regulatory Questions: N / A Resolution: Cleaning of traces of crystallized contrast in the detector. Air calibration of the equipment. Monitoring of the first exams after intervention. Preventively, taking into account the safety and hygiene measures decreed by the appearance of the coronavirus (COVID-19), and to avoid possible risks of contagion to our customers and Philips personnel, this worksheet is not signed by the customer (Téc. Fátima Dias), to which it was informed and shows its conformity. | 3 |  | Contrast intrusion, fixed by Cleaning of traces of crystallized contrast in the detector. 1. Reminder for cleaning 2. Rem for laxan ring and gap adjustment | [12 empty] | 116890789 | Closed | 553451 | TRUE | Closed - Done | Closed - Done |  | ECR-075975 | [5 empty] | Engineering | Any piture of lexan ring gap? Need to double check and adjust the lexan ring if necessary |  |  | Closed |  |  | Lexan Ring/Contrast | [8 empty] | System issue | System | Gantry-HW | Lexan Ring | Gantry-HW-Lexan Ring | Gantry-HW-Lexan Ring | ECR-075975 | Lexan Ring | Gantry-HW | Gantry-HW-Lexan Ring | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
382 |  |  | China | HEN Puyang People_s Hosp | 40031 | 44217 | 4 | 02N/Ingenuity CT 728326/couch fault | T2 Activities : Problem Description: 客户问题概述： ESTOP断开。 客户问题详细描述或故障现象： 扫描掉电 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： couch relay open 远程诊断与测试结果： 建议更换couch motor relay PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 派工 建议备件12NC及名称： relay SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : PPD 453566496941 1 罗永平-86-15226131133 24H 已出货 Internal Remarks : 23.01.2021 13:20:32 UTC Yongping Luo (86038950) $Q7:Confirmed report problem，Gantry can not scan. $Q8:Perform visiual check on site and find the HOR_MOTOR_RELAY_OPEN and conform the system cannot scan nornaly. $Q9:repalce the Relay,Device return to full function. $Q10: test and inspection data refer to PA. Device returned to full functionality. $Q11:Calibrated Tools refer to PA. External Remarks : 23.01.2021 13:20:30 UTC SUP JCOUSER (86038950)服务内容：设备无法扫描，床运动正常，检查发现设备偶尔报水平运动电机继电器故障。 服务结果：更换水平电机继电器，设备恢复正常工作。 其它重要事项：无 。 | 3 |  | Order 453566496941 上午出现两次，在即将扫描进床过程中报错 HOR_MOTOR_RELAY_OPEN | [12 empty] | 116888763 | In Process | 554942 | TRUE | Closed - Done | Closed - Done |  | NotEngineering-PartReplacement | [5 empty] | Engineering | Couch Relay, (might also check cable connection) return to Suzhou for investigation (waiting FSE to check if failure part still available) Chang xian Work with Robin to sync LCM improvement direction |  |  | WIP |  |  | Relay | [8 empty] | System issue | System | Couch-HW | Relay | Couch-HW-Relay | Couch-HW-Relay | NotEngineering-PartReplacement | Relay | Couch-HW | Couch-HW-Relay | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
383 |  |  | China | HEN Puyang People_s Hosp | 40031 | 44217 | 4 | 02N/Ingenuity CT 728326/couch has noise | Next Step Action : luo yongping called in ,called back ,disp him SAP Transfer : Case Transferred to SAP T2 Activities : Problem Description: 客户问题概述： 床上下运动有噪声 设备本身技术问题： 是。 Troubleshooting Action: 远程诊断与测试结果： luo yongping called in，怀疑滑轨缺润滑油 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 需要工程师到场检查添加润滑油 建议备件12NC及名称： 润滑油：453566132481 SAP Transfer : Dispatched for Onsite Support Next Step Action : PPD LUBRICANT 453566132481 1，危险品只能安排72H出货，FSE罗永平-86-15226131133 External Remarks : 23.01.2021 14:15:50 UTC SUP JCOUSER (86038950)服务内容：扫描床升降有噪声，检查发现升降滑动杆缺油。 服务结果：升降滑动杆加油后，设备恢复正常工作。 其它重要事项：无 Internal Remarks : 23.01.2021 14:15:53 UTC Yongping Luo (86038950) $Q7:Confirmed report problem，The Couch have noise when vertical move. $Q8:Perform visiual check on site and find the Vertical move sliding bar lack of lubricating oil. $Q9:Add lubricating oil ,Device return to full function. $Q10: test and inspection data refer to PA. Device returned to full functionality. $Q11:Calibrated Tools refer to PA. ----------------------------------------------------------------------------- | 3 |  | 升降滑动杆加油后，设备恢复正常工作。 LUBRICANT 453566132481 在低位开始运动瞬间有异响， Ada support to check with OPS how grease (old or new?) add and control. Couch SN: 905018 initialization时间大概23秒左右，是否正常? Silva check Linkage to defect and requirement 10099, investigate the bugrep is necessary https://www.clouddrive.philips.com.cn/file/rgb5hu0ebxhckfdh5ggllb5z3s0dxl1l# 密码:123 | [12 empty] | 116890436 | In Process | 555424 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | E-125059 | [5 empty] | Engineering | Check couch SN and share to Robin to check if new lubricant cut in or not STD Couch SN: 905018. |  |  | WIP |  |  | Couch Noise | [8 empty] | System issue | System | Couch-HW | Lubricant | Couch-HW-Lubricant | Couch-HW-Lubricant | E-125059 | Lubricant | Couch-HW | Couch-HW-Lubricant | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
384 |  |  | China | SD Jinan No.4 Peo. Hosp. | 40038 | 44214 | 4 | 02N/Ingenuity CT 728326/Software yellow screen | T2 Activities : Problem Description: 客户问题概述： 扫描床水平运动故障。 客户问题详细描述或故障现象： 系统黄屏之后左边显示器不亮了 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 左边显示器不亮了 远程诊断与测试结果： 显示器故障或线缆故障 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 1，检查显示器信号线缆2，更换显示器 建议备件12NC及名称： 1，显示器（459801140821） Test/Inspection: N/A Test Equipment: N/A Next Step Action : PPD 459801140821*1 金晓-86-13370562316 已安排出货 External Remarks : 19.01.2021 14:27:39 UTC Tao Chen (86038373)故障现象: 单屏显示 服务内容: 插拔视频线和供电线 服务结果: 维修后,设备恢复正常. Internal Remarks : 19.01.2021 14:27:42 UTC Tao Chen (86038373) “RDI Accuracy: YES" $Q7: (Confirmed reported problem, single display of CT host) $Q8: ( Check display in display settings, found only one display) $Q9: ( Fasten video cable and power cable for LCD ) $Q10: ( Test and inspection data refer to: PA. Device returned to full functionality) $Q11: Calibrated Tools name ( N/A) ^Current Tube installed date: N/A ^Current Tube usage: N/A ----------------------------------------------------------------------------- | 3 |  | 459801140821 - 19" LCD display monitor EA193Mi AND DVI adaptor | [12 empty] | 116861304 | Closed | 545509 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00631367 | [5 empty] | Engineering | Download CAN trace to ChangXian for Couch investigation: 20200118 CAN Trace already cleaned, share the bugreport for yellow screen Check if FSE does DVI adaptor has issue too? Brook |  |  | Closed |  |  | DVI Adaptor | [8 empty] | System issue | System | Console-SW | WF-Yellow Screen | Console-SW-WF-Yellow Screen | Console-SW-Yellow-Screen | CTUDT00631367 | Yellow Screen | Console-SW | Console-SW-WF-Yellow Screen | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
385 |  |  | Australia | MURWILLUMBAH DISTRICT HOSPITAL | 341000 | 44212 | 3 | Had a big storm the previous night and now system comes on as normal but will not scan. Have 2 errors - CIRS and... | The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: CT Radiographer Device Use: System Restart after storm activity. Expected and Unexpected Results: System not able to initialise after local Storm Activity. User Impact: System is down Patient Impact: None reported Current Software Version: 4.1.10 LOG REVIEW: Indicates that there was errors relating to DMS communication and power, plus Gantry and couch power. Details of errors seen. 2021/01/16 13:03:22.119__001 S_DMC_SHOT_COMM_COLLAPSE 2021/01/16 13:03:24.087 S_HMC_COUCH_24V_IS_OFF ERROR_NORMAL 2021/01/16 13:03:24.317 S_MP_CTRL_POS_ERROR_MAX_MIN_VOLTS 2021/01/16 13:03:24.644 S_RMC_ERROR_MCO ERROR_NORMAL PROBLEM DESCRIPTION Site experienced a large storm and when they attempted to initialise the system the following day the system would not scan with Ganrtry and CIRS error notifications to the user. DATE THE ISSUE OCCURRED: 16 Jan 2020, 12:20 am System SW Version is: 4.1.10 Called site early 18 Jan 2020 and discussed with Elizabeth, review of system logs does indicated that the system did scan later that day from 2:59 PM, Elizabeth indicated that on call Radiographer John was able to resolve by full system power cycle. ________________________________________________________________ LOG REVIEW: Indicates that there was errors relating to DMS communication and power, plus Gantry Rotation and couch power. ________________________________________________________________ RECOMMENDATION: Issue occurred after storm activity, and was resolved by system reset, no action required if no recurrence under normal conditions. ________________________________________________________________ CUSTOMER UPDATE: I have discussed with site contact Elizabeth. ________________________________________________________________ CURRENT SYSTEM STATUS: System is operational. ________________________________________________________________ ----------------------------------------------------------------------------- | 2 |  | Issue occurred after storm activity, and was resolved by system reset, no action required if no recurrence under normal conditions. | [12 empty] | 116857185 | Closed | 543404 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not a defect | [5 empty] | Others | Unexpect external condition |  |  | Closed |  |  | Storm | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
386 |  |  | Japan | Nikko Memorial Hospital | 341011 | 44221 | 5 | 01CS:RS:ｱﾝﾄﾞｳ様:リングアーチファクト発生>>Module 12/13or31/30>> | OneEMS internal remarks Description: --Troubleshooting Action-- Post-response results: Not improved Failure location: 04Gantry Exclude HV What steps did you take during the diagnosis ?: Confirmed ring artifacts on Module 12/13 or 31/30. FSE compatible. --Repair Action-- Recommended 12Nc: PR1 459800911271: TDMS V2 M.Det Class A FRU QTY: 1 SLOC: Sapporo (U612) --Subject-- 04DS: RS: Ando: Ring artifact occurred >> Module 12/13 or 31/30 >> Air calibration not resolved --Problem Description-- Target device: CT Failure location: 04Gantry Exclude HV When did it first occur? (Within ~): 1 Hour How often do problems occur ?: Continuously At what point in the end-user workflow is the problem occurring ?: During Reconstruction / Post Processing What happened recently ?: N.A. Does the customer create a specific system event report for the problem ?: N.A. | 3 |  | The detector module to be returned is s/n39367. Only the data showing the problem are the snapshots attached here. https://www.clouddrive.philips.com.cn/file/3t9k1o2j1kq9kpb64d3tzgko0zcpuk5w# 密码:123 | [12 empty] | 116904565 | In Process | 559486 | TRUE | Closed - Done | Closed - Done |  | CTUDT00632110 | [5 empty] | Engineering | DMS Module Defective Check SN of return module and BD calibration data Need waiting FSE handling |  |  | Closed |  |  | Defective DMS module | [8 empty] | System issue | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | DMS Module | DMS-HW | DMS-HW-DMS Module | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
387 |  |  | Japan | Tomakomai City Hospital | 341014 | 44221 | 5 | TM: Technician in charge: Streak Artifact occurred. Occurs from the upper left to the lower right in the Helical photography of the head. There is no abnormality when shooting with the head on the Externder | Streak Artifact occurs. Occurs from the upper left to the lower right in the Helical photography of the head. No abnormality when shooting with the head on the Externder | 4 |  | artifact, The streak artifact was from the edge of mattress. | [12 empty] | 116906176 | In Process | 559680 | TRUE | Closed - Done | Closed - Done |  | Not a defect | [5 empty] | Customer | DMS Module Defective? Need waiting FSE handling |  |  | Closed |  |  | Mis-use | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
388 |  |  | Japan | Saiseikai Gosu Hospital | 343000 | 44218 | 4 | HY: There is a strange noise from the gantry | Problem Reported by Customer : HY:ガントリーから異音する OneEMS internal remarks : 河重DB SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support | 5 |  | The noise is from around the belt tensioner, emitted during high speed rotation over 0.5sec. FSE is planning to adjust the tensioner position. | [12 empty] | 116896552 | New | 557783 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | ECR-077012 | [5 empty] | Engineering | is it from IBOX Boot up? Or scratch during roatating? |  |  | WIP |  |  | Belt tension | [8 empty] | System issue | System | Gantry-HW | Belt | Gantry-HW-Belt | Gantry-HW-Belt | ECR-077012 | Belt | Gantry-HW | Gantry-HW-Belt | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
389 |  |  | China | Dalian Branch of Shengjing Hospital of China Medical University | 345001 | 44214 | 4 | White ring Image artifacts due to DMS Detector Module 20 fault, replaced the DMS Detector Module. | T2 Activities : Problem Description: 客户问题概述： 图像伪影。 客户问题详细描述或故障现象： 图像环形伪影。 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： log无报错。 远程诊断与测试结果： 空气校准无效。已和工程师确认故障。 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议用物理分析排查模块，必要时更换。 建议备件12NC及名称： PPD 459800594971 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : PPD 459800594971无货，已与TSE 确认可触替代， 探测器模块 TDMS V2 M.Det Class B FRU 459800911281*1 已发货，24H 场地，FSE 于志龙-86-18640955881 External Remarks : 19.01.2021 14:12:12 UTC SUP JCOUSER (86038273)服务内容：图像中心白色环形伪影，20号探测器无输出，更换20号探测器。 服务结果：维修后，设备恢复正常工作。 Internal Remarks : 19.01.2021 14:12:15 UTC Zhilong Yu (86038273) RDI Accuracy:Yes $Q7: (Confirmed reported problem &white ring artifact on image center) $Q8: (Visual inspection,run bad decetor&no 20 moudle no oupute) $Q9: (replace no 20 decetor moudle) $Q10: (Test and Inspection data refer to pa & Device returned to full functionality) $Q11:Calibrated Tools name (NA) ----------------------------------------------------------------------------- | 3 |  | 图像中心白色环形伪影，20号探测器无输出，更换20号探测器。 | [12 empty] | 116860719 | Closed | 545399 | TRUE | Closed - Done | Closed - Done |  | CTUDT00632110 | [5 empty] | Engineering | DMS Module Defective Check SN of return module and BD calibration data Check if order correct class A module? https://www.clouddrive.philips.com.cn/file/ol7vg3rlshbe2yr8r1s4wucjzvfi3jb8# 密码:123 |  |  | Closed |  |  | Defective DMS module | [8 empty] | System issue | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | DMS Module | DMS-HW | DMS-HW-DMS Module | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
390 |  |  | China | JL Taonan T.C.M. Hosp. | 345006 | 44219 | 4 | 02N/Ingenuity Core128/Elite China/Print the fault | SAP Transfer : Dispatched for Onsite Support Internal Remarks : 23.01.2021 09:31:30 UTC Tian Hua (86038446) RDI Accuracy ：No $Q7: (Confirmed reported problem&Cannot print) $Q8: (Review logger and check the Lan config，The Host computer IP lost ） $Q9: (Reset the host IP and test） $Q10: (Test and inspection data refer to: PA. Device returned to full functionality) $Q11 Calibrated Tools name (FLUKE287C) S/N (27420003) Next Calibration date (03/31/2021) T2 Activities : Problem Description: 客户问题概述： 不能打印图像 客户问题详细描述或故障现象： 不能打印图像 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： N/A 远程诊断与测试结果： N/A PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 网卡设置在系统中丢失，现场设置网卡IP地址. 建议备件12NC及名称： N/A 客户需求备注： 无 External Remarks : 23.01.2021 09:31:27 UTC Tian Hua (86038446)服务内容： 胶片打印间断失败，主机ip不识别，重新设置识别。 服务结果： 维修后，设备恢复正常工作。 其它重要事项： 无。 | 3 |  | Observation : From logs we have seen that printing problem from host was because of ConnectionClosedException of DicomPrinter . So connection was gracefully closed and end print failed because of the same. Hence printing issue was not because of FilmView errors as mentioned in below mail chain(screen shots) . And also from the logs we can see that , there was 4 restarts during the time of issue , and of all the 4 times , there was not one printing which was succeeded. Hence printing issue was not resolved by CT restart which again is related to the printer not connecting issue. Based on above , this is clearly remote printer issue and is not related to CT. Please check this with the respective device vendor. Logs : 2021.01.22 09:28:10:632476 Event Filming Send from Film View to printer started, 112 images FilmViewApp SEND_FROM_FILM_TO_PRINTER_START 2021.01.22 09:28:11:680314 Error DicomPrinter DicomPrinter : EndPrint failed with exception =Philips.PmsCT.Platform.StorageDevices.ConnectionClosedException: Connection Gracefully Closed 2021.01.22 09:28:11:680314 Error PrintSpooler PrintSpooler : Print operation failed. Status changed to Failed. Exception = Philips.PmsCT.Platform.StorageDevices.ConnectionClosedException: Connection Gracefully Closed | [12 empty] | 116902097 | In Process | 559220 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not a defect | [5 empty] | Others | Collect bugreport for engineering investigation https://www.clouddrive.philips.com.cn/file/kvvupeh6f2ets1nhs2ggbheza6i64t7d# 密码:123 |  |  | Closed |  |  | Printing | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
391 |  |  | China | XIFENG CHINESE MEDICINE HOSPIT | 345014 | 44217 | 4 | PRS connection/02N/Ingenuity Core128/Elite China | BU Q&R Feedback: No Reason: Customer Relations Problem Reported by Customer : PRS connection Safety Questions : Case: 0116892457 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Qian Wang Last updated by: Qian Wang 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action : Hu ke SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support ----------------------------------------------------------------------------- | 5 |  | PRS connection | [12 empty] | 116892457 | New | 556383 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not a defect | [5 empty] | Customer | NAC， confirmed | [5 empty] | PRS Setup | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
392 |  |  | China | XIFENG CHINESE MEDICINE HOSPIT | 345014 | 44217 | 4 | 02N/Ingenuity Core128/Elite China/report error | Problem Reported by Customer : report error Next Step Action : 018100859923 谢老师二度催单，TSE电话忙线中，已告知客户 T2 Activities : Problem Description: 客户问题概述： 定位扫描之后，进行螺旋扫描间歇性报错 客户问题详细描述或故障现象： Gantry cann initialize 设备本身技术问题： 是。 Troubleshooting Action: 远程诊断与测试结果： 怀疑机架旋转，准直器。。。问题 PRS TR状态： 不通，请工程师到场检查。 Repair Action: 建议现场维修步骤： 需要工程师到场检查log，判断问题 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : Hu Ke External Remarks : 21.01.2021 14:56:41 UTC SUP JCOUSER (86029494)服务内容：确认报修故障，间歇性扫描报错。现场检查，发现滑环污渍，医院使用84消毒液消毒过机房。清洁滑环后设备恢复。 服务结果：维修后设备恢复正常。 Internal Remarks : 22.01.2021 13:29:42 UTC Ke Hu (86029494) "RDI Accuracy: NA" $Q7: (Confirmed reported problem,report error when scanning) $Q8: (Troubleshooting, onsite check logger, it reported CRC frame error.So I check the slip ring and clean the slip ring.Then device return to functionality.) $Q9: (Repair steps, clean the slip ring.) $Q10: (Test and Inspection data:Refer to PA.Device status: Device return to functionality) $Q11: Calibrated Tools name : na. For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date: na ^Current Tube usage: na ----------------------------------------------------------------------------- | 3 |  | Cl containment | [12 empty] | 116888865 | Closed | 557554 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not a defect | [5 empty] | Customer | Cleaning |  |  | Closed |  |  | Cleaning Material | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
393 |  |  | China | HEB Caofeidian Dist. Hosp. | 345016 | 44218 | 4 | 02N/Ingenuity Core128/Elite China/Image transfer problem | Next Step Action : SS 管建斌 018603299918来电报修 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support T2 Activities : Problem Description: 客户问题概述： 图像传输失败。 客户问题详细描述或故障现象： 主机传图到PACS会停止，不能完全传输 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 传图一半停止，有时能传完，有时需要重启主机才能继续传输 远程诊断与测试结果： 怀疑是软件问题 PRS TR状态： 无PRS。 Repair Action: 建议现场维修步骤： 1，检查软件报错2，做bugreport 提交esclation 3，写CFM反馈 Test/Inspection: N/A Test Equipment: N/A | 5 |  | PACS Transfer issue, similar with Henan Puyang site, ISP 1 time/day, other PACS 5/6 times/day | [12 empty] | 116895358 | In Process | 557731 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00631602 | [5 empty] | Engineering | Collect screenshot or failed transfer and bugreport for engineering investigation, Bugreport grabbed, PRS was rejected to connect by customer, waiting FSE on site to upload https://www.clouddrive.philips.com.cn/file/x53kir40f7kiepxe8bbqewogu1bzbggp# password:123 |  |  | Closed |  |  | PACS | [8 empty] | System issue | System | Console-SW | Remote Connection-1602 | Console-SW-Remote Connection-1602 | Console-SW-Remote Connection | CTUDT00631602 | Remote Connection | Console-SW | Console-SW-Remote Connection-1602 | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
394 |  |  | China | HEB Caofeidian Dist. Hosp. | 345016 | 44214 | 4 | 02n/Ingenuity Core128/Elite China/The display does not display images | T2 Activities:Problem Description: Summary of customer_s problems: failure related to computer software. Detailed description of Customer_s problems or malfunctions: Dual/split screen changed to single screen It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: Dual/split screen changed to single screen Remote troubleshoot and test result: Need to set up display on the back end PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: It feels like the monitor display is malfunctioning. There is no remote access to the back end to check the display quantity Modify display mode in the back end Test/Inspection: N/A Test Equipment: N/A Next Step Action:Hu Yue External Remarks:19.01.2021 13:26:21 UTC Yue Hu (86038448) Service details (required): Display problem. Change graphic card setting Service result (required): after the maintenance, the device returns to normal working condition Internal Remarks:19.01.2021 13:26:24 UTC Yue Hu (86038448) $Q7:(Confirmed reported problem: The display does not display images $Q8: (onsite check, re-configuration display setting $Q9: (re-configure display setting, system return to service $Q10: (Test and inspection data refer to: PA. Device returned to full service) $Q11: Calibrated Tools name (FLUKE287) S/N (20060015) next calibration data (05/17/2021) ----------------------------------------------------------------------------- | 3 |  | DVI Adaptor damaged | [12 empty] | 116862671 | In Process | 546669 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632173 | [5 empty] | Engineering | DVI adapator, FSE is going send back to Suzhou for investigation |  |  | Closed |  |  | DVI Adaptor | [8 empty] | System issue | System | Console-HW | Adaptor | Console-HW-Adaptor | Console-HW-Adaptor | CTUDT00632173 | Adaptor | Console-HW | Console-HW-Adaptor | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
395 |  |  | China | JS Yangzhou No.1 People Hosp. Xiqu Hosp. | 345022 | 44216 | 4 | 01/Ingenuity Core128/Elite China/need password | Next Step Action : 已联系qi liangdong，齐工告知需要找装机工程师询问。 已联系鸿乐嘉 ----------------------------------------------------------------------------- | 3 |  | Password forget | [12 empty] | 116880827 | New | 548380 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not a defect | [5 empty] | Customer | NAC， confirmed | [14 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
396 |  |  | China | Zhenjiang M.C.A. Hosp. | 345023 | 44219 | 4 | wechat：05R/Ingenuity Core128/Elite China/Operation manual | Next Step Action : 需要操作手册(电子版) 徐昊天 18761995321跟进回电 ----------------------------------------------------------------------------- | 3 |  | Operation manual is not found | [12 empty] | 116902741 | Closed | 559283 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not a defect | [5 empty] | Customer | NAC， confirmed | [14 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
397 |  |  | United States | NORTHWEST RADIOLOGY/MRI | 300057 | 44216 | 4 | 0711CFFM2 CTA Gated protocol is slow but not other protocol | Work Order #: WO-02823813 Atz, Michael The system fully meets specification and returned to use Problem Description by Engineer : Problem description by engineer : CTA Gated protocol is slow but not other protocol Information to support the complaint handling process Customer Function/Role: Not provided Device Use: Not provided Expected and Unexpected results: Not provided Patient Impact/Outcome: Not provided User Impact: Not provided Current Software Version: 4.1.7.25026 Resolution : Protocol rebuild Internal Comments : Error # and/or description of error : none Completed or Recommended Repair Action "Internal" : I spoke to Dyan from APPS she suggests that they rebuild the protocol from scratch and Bridget Williams is going to do this with the onsite APPS person Diagnostic performed by Engineer : I spoke to Dyan from APPS she suggests that they rebuild the protocol from scratch and Bridget Williams is going to do this with the onsite APPS person ----------------------------------------------------------------------------- | 3 |  | I spoke to Dyan from APPS she suggests that they rebuild the protocol from scratch and Bridget Williams is going to do this with the onsite APPS person | [12 empty] | 116887001 | Closed | 551005 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not a defect | [5 empty] | Application | is it a protocol issue or upgade issue? |  |  | Closed |  |  | Protocol | Upgrade |  |  |  |  | Refresh |  |  | Upgrade issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
398 |  |  | United States | Upson Regional Medical Center | 310115 | 44214 | 4 | 0198CDOM2 recon box error restart the recon box John Lee is aware | Work Order #: WO-02799908 Atz, Michael The system does not meet specification and is out of use Follow Up Required : Error # and/or description of error : Red dot Completed or Recommended Repair Action "Internal" : Ingenuity recon box error restart the recon box red dot is present check model and S/N and troubleshoot the CIRS before ordering parts -- Selfcool NGBP,5A,1G,PCIe x8,No5V,(SN5BP) 459800056953 -- GPU, Nvidia Tesla C2075 459800192241 -- CIRS Ver.Server G4-VS501,2.40GHz,500 MB 453567533572 -- CIRS Vertical Server, G4-VS501, 2.4GHz 453567533573 -- Savvio 15K.3 146GB SAS Drive FRU Kit 459801632781 -- Nvidia K4200 GPU Card 459800916081 Internal Comments : Error # and/or description of error : Red dot Completed or Recommended Repair Action "Internal" : Ingenuity recon box error restart the recon box red dot is present check model and S/N and troubleshoot the CIRS before ordering parts -- Selfcool NGBP,5A,1G,PCIe x8,No5V,(SN5BP) 459800056953 -- GPU, Nvidia Tesla C2075 459800192241 -- CIRS Ver.Server G4-VS501,2.40GHz,500 MB 453567533572 -- CIRS Vertical Server, G4-VS501, 2.4GHz 453567533573 -- Savvio 15K.3 146GB SAS Drive FRU Kit 459801632781 -- Nvidia K4200 GPU Card 459800916081 Problem Description by Engineer : Problem description by engineer : recon box error restart the recon box John Lee is aware When was the first occurrence? Within... : 17-Jan-2021 How often is issue occurring? : Continuously Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: No further information is available as per the customer Patient Impact/Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer Current Software Version: Not provided Diagnostic performed by Engineer : The CT tech has rebooted several times ----------------------------------------------------------------------------- Work Order #: WO-02799992 18-Jan-2021 Ramzy Duraid The system fully meets specification and returned to use Problem Description by Engineer : Problem description by engineer : recon box error restart the recon box John Lee is aware When was the first occurrence? Within... : 17-Jan-2021 How often is issue occurring? : Continuously Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: No further information is available as per the customer Patient Impact/Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer Current Software Version: Not provided Resolution : Reloaded CIRS software and re-seated the GPU BD. System is fully functional. ----------------------------------------------------------------------------- | 2 |  | Reloaded CIRS software and re-seated the GPU BD. System is fully functional | [12 empty] | 116861545 | Closed | 543578 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not a defect | [5 empty] | Service | does the issue means GPU not been installed in the beginning? |  |  | WIP |  |  | Upgrade | Upgrade |  |  |  |  | Refresh |  |  | Upgrade issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
399 |  |  | United States | Lakeland Regional Medical Cent | 320474 | 44196 | 53 | Gantry is making a clicking noise when rotating. | Work Order #: WO-02703435 Mark Finnigan The system fully meets specification and returned to use Follow Up Required : Spoke with Nicole Neuman. She states that the gantry is making a clicking noise during rotation. Determine if this is a cover rub issue. If not it may be the mechanical brake on the motor or the belt idler that is creating the noise. Diagnostic performed by Engineer : Spoke with Nicole Neuman. She states that the gantry is making a clicking noise during rotation. Determine if this is a cover rub issue. If not it may be the mechanical brake on the motor or the belt idler that is creating the noise. Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: Clicking noise when rotating. User Impact: None Patient Impact: None Current Software Version: ----------------------------------------------------------------------------- Work Order #: WO-02703476 31-Dec-2020 Sher Mark The system does not meet specification and is out of use Follow Up Required : Next Flight Out BELT, 5-RIB POLY-VEE, 198" 453566492401 Q 1 ATT: Philips CT WCP Lakeland Regional Medical Center 1324 Lakeland Hills Blvd Lakeland, FL. 33805-4543 on site service Aubrey Brown FSE cell 8137270443 Diagnostic performed by Engineer : troubleshot and found clicking noise comming from rotor belt having bad spots Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: Clicking noise when rotating. User Impact: None Patient Impact: None Current Software Version: ----------------------------------------------------------------------------- Work Order #: WO-02704705 02-Jan-2021 Garcia Marrero, Jorge The system fully meets specification and returned to use Diagnostic performed by Engineer : Installation and adjustment of rotational belt. System is fully oprrational. Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: Clicking noise when rotating. User Impact: None Patient Impact: None Current Software Version: Covid19 PP1 Parts Consumed : 1.00000,453566492401,POLY-VEE BELT ----------------------------------------------------------------------------- | 3 |  | Installation and adjustment of rotational belt. 1.00000,453566492401,POLY-VEE BELT | [12 empty] | 116766276 | Closed | 509227 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | ECR-077012 | [5 empty] | Engineering | Belt noise |  |  | WIP |  |  | Belt Noise | Upgrade |  |  |  |  | Non-Refresh |  |  | Upgrade issue | System | Gantry-HW | Belt | Gantry-HW-Belt | Gantry-HW-Belt | ECR-077012 | Belt | Gantry-HW | Gantry-HW-Belt | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | y
400 |  |  | Spain | Policlínica Nuestra Señora del Rosario | 336190 | 44124 | 43 | Failure in the TAC couch which does not prevent working, but it rubs and it does not work correctly. | "Work Order #: WO-02240337 21-Oct-2020 Miguel Ribot Martinez The system partly meets specification and is returned to use with restrictions Diagnostic performed by Engineer: The tabletop rubs against a side cover when moving longitudinally. When the tabletop moves longitudinally at its fast speed, it does not automatically slow down when reaching the extremes and collides with the mechanical stop. Problem Description by Engineer: Customer Function / Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Regulatory Questions: N / A Resolution: The left side cover of the table is repositioned well and the board no longer rubs it when moving. The encoders and their wiring are checked and are not damaged. The operability of the "" limit switches "" of the longitudinal movement is verified. The calibration of the longitudinal movement of the board is carried out. After these actions, the problem of the impact of the board at the ends is not solved. Follow Up Required: Change table top and implement software patch. Resolution: The problem that the stretcher goes against the internal and external stops when it is made to move with the fast speed, is a software bug that is solved by implementing fco 72800735 that goes up to version 4.1.10.25202. It is necessary to continue the intervention. Our Service Center will contact you to plan it as soon as possible. -------------------------------------------------- --------------------------- Work Order #: WO-02260923 02-Nov-2020 Miguel Ribot Martinez The system fully meets specification and returned to use Resolution: The encoder of the horizontal movement of the table is reprogrammed, after which it no longer hits with the mechanical stop. The table stops smoothly by software when reaching the ends with the highest speed of movement. The broken upper left side cover is replaced with a new one. The equipment is configured so that remote access to it works. The equipment remains working properly. Problem Description by Engineer: Customer Function / Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Regulatory Questions: N / A Automatic IP Update: No Automatic IP Updates were processed for this Part as the Traceability Flag on Product 453567504262 is blank Parts Consumed: 1.00000,453567504262, Cover, Subframe LH Pat Supp -------------------------------------------------- --------------------------- | 5 |  | 1.00000,453567504262, Cover, Subframe LH Pat Supp FCO72800735 implemented and reprograming | [12 empty] | 116326389 | Closed | 352817 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00630138/630160 | [5 empty] | Engineering | FCO72800735 |  |  | Closed |  |  | FCO72800735 | Upgrade |  |  |  |  | Refresh |  |  | Upgrade issue | System | Couch-SW | Horizontal Move Block | Couch-SW-Horizontal Move Block | Couch-SW-Horizontal Move Block | CTUDT00630138/630160 | Horizontal Move Block | Couch-SW | Couch-SW-Horizontal Move Block | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
401 |  |  | United States | Mount Grant General Hospital | 40012 | 44225 | 5 | CT will not scan. | Work Order #: WO-02883500 Mark Finnigan The system does not meet specification and is out of use Problem Description by Engineer : Problem description by engineer : Spoke with CT tech Duane Henry. Duane states that the system would not perform a scan and gave a invalid table position error. Power cycling the system did not help. When was the first occurrence? Within... : 28-Jan-2021 How often is issue occurring? : First occurrence Malfunction area : Table Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: No further information is available as per the customer Patient Impact/Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer Current Software Version: 4.1.10 Diagnostic performed by Engineer : Used AEL to locate the error messages. Internal Comments : Error # and/or description of error : S_VMC_INVALID_CALIBRATION. Issue with the vertical calibration and/or the NVRAM settings on the CCB. Completed or Recommended Repair Action "Internal" : Found the invalid table calibration error which opened Estop. Then there are repeating NORMAL error messages stating that a vertical calibration is necessary. Follow Up Required : Error # and/or description of error : S_VMC_INVALID_CALIBRATION. Issue with the vertical calibration and/or the NVRAM settings on the CCB. Completed or Recommended Repair Action "Internal" : Found the invalid table calibration error which opened Estop. Then there are repeating NORMAL error messages stating that a vertical calibration is necessary. ----------------------------------------------------------------------------- Work Order #: WO-02883557 Schuler Bruce Problem Description by Engineer : Problem description by engineer : Spoke with CT tech Duane Henry. Duane states that the system would not perform a scan and gave a invalid table position error. Power cycling the system did not help. When was the first occurrence? Within... : 28-Jan-2021 How often is issue occurring? : First occurrence Malfunction area : Table Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: No further information is available as per the customer Patient Impact/Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer Current Software Version: 4.1.10 Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 2 |  | Found the invalid table calibration error which opened Estop. Then there are repeating NORMAL error messages stating that a vertical calibration is necessary. Need more information from field | [12 empty] | 116944433 | In Progress | 570574 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-NoEnoughInfo | [5 empty] | Engineering | Does this system implement FCO72800735 Couch FCO? If yes, please help collect the CAN trace when issue happened, share the file as well as timestamp for engineering investigation. |  |  | WIP |  |  | Couch Calibration | [8 empty] | System issue | System | Couch-HW | TBD-NoEnoughInfo | Couch-HW-TBD-NoEnoughInfo | Couch-HW-TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo | Couch-HW | Couch-HW-TBD-NoEnoughInfo | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
402 |  |  | Spain | S.C.S. HOSP. GRAL. DE FUERTEVE | 40020 | 44222 | 5 | The abdomen - neck study is stopped. | Work Order #: WO-02862448 Carlos Saumell Lasala The system partly meets specification and is returned to use with restrictions Diagnostic Performed by Engineer: I check the error logs, at the time the incident occurred today, but no significant error message regarding the couch position is observed. Problem Description by Engineer: I talk to María del Mar Espinosa. In the studies with three sequences: thorax-abdomen-neck, the equipment performs the first two correctly, but when it is preparing for the third one, it places the couch in a different position from the one programmed in the surview. But it does not show any error. It always happens with the same protocol. It has failed other times, in the same way. Regulatory Questions : N/A Follow Up Required: We should go, simulate a study identical to the one from the protocol that is failing, and see if there is something predefined in the study affecting the final couch positioning. ----------------------------------------------------------------------------- | 3 |  | Need more information about how issue happened, is it planned sequence or share the failed exam card . Yiming & Kevin tried to following, the studies with three sequences: thorax-abdomen-neck, and issue is not occurred on test system. Need share the failed exam card for further investigation. | [12 empty] | 116925580 | In Progress | 562643 | TRUE | Closed - Done | Closed - Done |  | TBD-NoEnoughInfo | [5 empty] | Engineering | Share the exam cards and CAN trace file for engineering investigation. |  |  | WIP |  |  | Neck study | [8 empty] | System issue | System | Console-SW | TBD-NoEnoughInfo | Console-SW-TBD-NoEnoughInfo | Console-SW-TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo | Console-SW | Console-SW-TBD-NoEnoughInfo | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
403 |  |  | United States | Odessa Regional Medical Center | 40029 | 44222 | 5 | Repair/Replace Table Top | Work Order #: WO-02865654 McCain, Jimmie Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 3 |  | Order 459801429121 PULL OUT TOOL FOR IV POLE "Please reschedule for 2-13-2021 at 1pm, 4 hours on site, I have confirmed with the customer." | [12 empty] | 116927861 | In Progress | 566710 | TRUE | Closed - Done | Closed - Done |  | NotEngineering-PartOrder | [5 empty] | Engineering | Not clear, 459801429121 PULL OUT TOOL FOR IV POLE |  |  | WIP |  |  | IV Pole | [8 empty] | System issue | System | Gantry-HW | NotEngineering-PartOrder | Gantry-HW-NotEngineering-PartOrder | Gantry-HW-NotEngineering-PartOrder | NotEngineering-PartOrder | NotEngineering-PartOrder | Gantry-HW | Gantry-HW-NotEngineering-PartOrder | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
404 |  |  | Portugal | Centro Hospitalar de Vila Nova | 40030 | 44223 | 5 | Image with artifacts on CT equipment of Urgencies Department. | Work Order #: WO-02871779 Jaime Granja Regulatory Questions : N/A Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 3 |  | Artefactual image on the SU CT equipment 1. Reminder Customer to do CT cleaning in time 2. Reminder Service to do Lexan ring and gap adjustment | [12 empty] | 116933408 |  | 569273 | TRUE | Closed - Done | Closed - Done |  | ECR-075975 | [5 empty] | Engineering | Not clear 1.21 report contrast introsion. |  |  | WIP |  |  | Lexan Ring/Contrast | [8 empty] | System issue | System | Gantry-HW | Lexan Ring | Gantry-HW-Lexan Ring | Gantry-HW-Lexan Ring | ECR-075975 | Lexan Ring | Gantry-HW | Gantry-HW-Lexan Ring | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
405 |  |  | United States | Unity Health Systems | 40037 | 44224 | 5 | table is squeaking | Work Order #: WO-02878567 28-Jan-2021 Taylor, Durrance The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : Issue went away on its own and has not come back. ----------------------------------------------------------------------------- | 4 |  | Issue went away on its own and has not come back. Spring position need to be adjusted | [12 empty] | 116940220 | In Progress | 570126 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00631058 | [5 empty] | Engineering | Not clear |  |  | WIP |  |  | Couch Noise | [8 empty] | System issue | System | Couch-HW | Vertical spring | Couch-HW-Vertical spring | Couch-HW-Vertical spring | CTUDT00631058 | Vertical spring | Couch-HW | Couch-HW-Vertical spring | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
406 |  |  | Turkey | HITACHI HEALTHCARE TURKEY | 337000 | 44223 | 5 | Air calibration fails | Work Order #: WO-02870359 Soner Saglam The system partly meets specification and is returned to use with restrictions Follow Up Required : Air cal fails. Compensator analysis was completed successfully. Site visit is required. Problem Description by Engineer : Ring artefact in the system. Resolution : Air cal fails. Compensator analysis was completed successfully. Site visit is required. ----------------------------------------------------------------------------- | 3 |  | There is ring artifact in the system. Service was given for the problem of not performing air calibration reported in the device. The protection cover of the detectors, inside and outside of the gantry cover were cleaned and air calibration was tried again. It was seen that it was completed without any problems. Test shot was made with the Phantom, no problem was observed, the device is active. | [12 empty] | 116932731 | In Progress | 565863 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | ECR-075975 | [5 empty] | Engineering | CT cleaning? contrast introsion? |  |  | WIP |  |  | Lexan Ring/Contrast | [8 empty] | System issue | System | Gantry-HW | Lexan Ring | Gantry-HW-Lexan Ring | Gantry-HW-Lexan Ring | ECR-075975 | Lexan Ring | Gantry-HW | Gantry-HW-Lexan Ring | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
407 |  |  | Chile | HOSPITAL BARROS LUCO TRUDEAU | 341005 | 44224 | 5 | Potential Security Event: CL1688 Printer configuration | Work Order #: WO-02879693 Valenzuela Velázquez, Eduardo Ariel The system fully meets specification and returned to use. Problem Description by Engineer: Configure printer. Resolution: We configured the printer and performed a printing test; passed | 3 |  | Printer configure | [12 empty] | 116941126 | In Progress | 570034 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not a defect | [5 empty] | Customer | Not a complaint |  |  | Closed |  |  | Printer config | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
408 |  |  | Chile | HOSPITAL BARROS LUCO TRUDEAU | 341005 | 44221 | 5 | CL1688- SOME RECONSTRUCTION APPLICATIONS WERE DELETED. IT CANNOT EXPORT OR EXPORT. | Work Order #: WO-02852170 Valenzuela Velázquez, Eduardo Ariel The system partly meets specification and is returned to use with restrictions Problem Description by Engineer: Problems with CT Viewer application. Diagnostic performed by Engineer: CT viewer does not show all the layouts. Applications sent the file to the client to load the layout but they are not loaded in user mode. Follow Up Required: Load layouts from service mode. ----------------------------------------------------------------------------- | 3 |  | Load layouts from service mode. @Kevin, could you please check normal opertion steps for the layout? | [12 empty] | 116915843 | In Progress | 560832 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-NoEnoughInfo | [5 empty] | Engineering | Not clear what is problem and field action done |  |  | WIP |  |  | Layout | [8 empty] | System issue | System | Console-SW | TBD-NoEnoughInfo | Console-SW-TBD-NoEnoughInfo | Console-SW-TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo | Console-SW | Console-SW-TBD-NoEnoughInfo | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
409 |  |  | China | Jinggangshan Hospital, the First Affiliated Hospital of Nanchang University, Jiangxi | 345002 | 44224 | 5 | 02N/Ingenuity Core128/Elite China/Out of hard disk memory | BU Q&R Feedback: No SAP Customer Symptom Code: DIDM Customer Symptom Code: Patient or Image Data Management/Post Processing Data Reason: Corrective Maintenance Engineer Failure Code: Environmental Cause Repair Activity Code: Clean / Lubricate Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/Out of hard disk memory Safety Questions : Case: 0116938865 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Cheng Zhenhua Last updated by: Cheng Zhenhua 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Troubleshooting Action: instruct customer to delete the image Repair Action: instruct customer to delete the image Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: N/A ----------------------------------------------------------------------------- | 3 |  | 在OneEMS里面看到0116938865报硬盘空间不够，在线RSE指导客户删除图像后解决。 | [12 empty] | 116938865 | Fixed | 571453 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632261 | [5 empty] | Engineering | Out of momery |  |  | WIP |  |  | Disk full | [8 empty] | System issue | System | Console-SW | DiskFull | Console-SW-DiskFull | Console-SW-DiskFull | CTUDT00632261 | DiskFull | Console-SW | Console-SW-DiskFull | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
410 |  |  | China | SD Laiyang People Hospital | 345005 | 44223 | 5 | Wechat:02n/Ingenuity Core128/Elite China/The patient number information in the upper left corner cannot be deleted | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : The patient number information in the upper left corner cannot be deleted Safety Questions : Case: 0116930039 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Yu Zhao Last updated by: Yu Zhao 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action : TSE:使用问题，安排培训老师给客户，回电话. Next Step Action : Application oncall 值班邓明明,18929200533 电话不接，已短信告知。 T2 Activities : Problem Description: 客户问题概述： 有些病人图像不能删除 客户问题详细描述或故障现象： 报错：单个病人超出700幅 设备本身技术问题： 是。 Troubleshooting Action: 远程诊断与测试结果： 联系工程师马俊，怀疑软件问题 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议工程师到场检查，后台删除病人图像，重新安装软件。。 SAP Transfer : Dispatched for Onsite Support SAP Transfer : Case Transferred to SAP Next Step Action : Wang Rui External Remarks : 27.01.2021 16:08:11 UTC Rui Wang (86090282)服务内容： 后台删除病人图像，清理系统内存 服务结果： 设备恢复正常交付医院使用 Internal Remarks : 28.01.2021 13:02:48 UTC Rui Wang (86090282) "RDI Accuracy:Yes" $Q7: (Confirmed reported problem &print has problem） $Q8: (reboot gantry，check the set，the set has problem) $Q9: ( reset the system Settings) $Q10: (Test and inspection data refer to: PA & Device returned to full functionality) $Q11: Calibrated Tools name ( N/A) ^Current Tube installed date: N/A ^Current Tube usage: N/A ----------------------------------------------------------------------------- | 3 |  | 建议工程师到场检查，后台删除病人图像，重新安装软件。。 此场地FSE反馈医院技师误操作导致打印胶片较多，软件提示超过700幅，FSE到场删除多余打印任务恢复正常。 Tammy, Check if SW have popup message for custmoer to decide to print or not? And see if there is any improvement opportunity. | [12 empty] | 116930039 | In Process | 570653 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Too many print job | [5 empty] | Customer | Check if dump files are too many in D disk? |  |  | WIP |  |  | Mis-use | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
411 |  |  | China | XIFENG CHINESE MEDICINE HOSPIT | 345014 | 44223 | 5 | RMSE:02N/Ingenuity Core128/Elite China/Image cannot be transmitted | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : Image cannot be transmitted Safety Questions : Case: 0116930399 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Zixuan Last updated by: Wang Zixuan 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Internal Remarks : 28.01.2021 14:02:49 UTC Ke Hu (86029494) “RDI Accuracy: NA” $Q7: (Confirmed reported problem.images transfer has failed.) $Q8: (Troubleshooting,teach customer to run spooler reset.) $Q9: (Repair steps,spooler reset. ) $Q10: (Test and Inspection data:na.Device status: Device return to functionality.) $Q11: (Calibrated Tools name:NA.) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date: NA ^Current Tube usage: NA. External Remarks : 28.01.2021 14:02:46 UTC SUP JCOUSER (86029494)服务内容：客户要求为工作站添加网络节点，远程指导客户添加网络节点。 服务结果：远程指导客户添加网络节点。 External Remarks : 28.01.2021 14:04:25 UTC SUP JCOUSER (86029494)服务内容：传图失败，远程指导客户运行spooler reset后传图恢复正常。 服务结果：远程指导后设备恢复正常。 ----------------------------------------------------------------------------- | 3 |  | 传图失败，远程指导客户运行spooler reset后传图恢复正常。 share bugreport for analysis | [12 empty] | 116930399 | New | 567807 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering | PACS |  |  | Closed |  |  | PACS | [8 empty] | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
412 |  |  | China | HEB Caofeidian Dist. Hosp. | 345016 | 44223 | 5 | Escalation for SWO#53007548 | Next Step Action : FSE 徐国梁 13869524913 早上来电开此单，已告知线上非常忙碌，现处理完已短信通知徐工 ----------------------------------------------------------------------------- | 3 |  | 客户问题概述： 图像传输失败。 客户问题详细描述或故障现象： 主机传图到PACS会停止，不能完全传输 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 传图一半停止，有时能传完，有时需要重启主机才能继续传输 远程诊断与测试结果： 怀疑是软件问题 PRS TR状态： 无PRS。 Repair Action: 建议现场维修步骤： 1，检查软件报错2，做bugreport 提交esclation 3，写CFM反馈 | [12 empty] | 116930481 | New | 564770 | TRUE | Closed - Duplicate | Closed - Duplicate |  | CTUDT00632127 | [5 empty] | Engineering | Not clear |  |  | WIP |  |  | PACS Transfer | [8 empty] | System issue | system | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
413 |  |  | China | HEB Caofeidian Dist. Hosp. | 345016 | 44223 | 5 | 02N/Ingenuity Core128/Elite China/Stand By | BU Q&R Feedback: No Reason: Customer Relations Problem Reported by Customer : Stand By Next Step Action : FSE 胡越 call in and disp him. Safety Questions : Case: 0116929722 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Yu Zhao Last updated by: Yu Zhao 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Internal Remarks : 29.01.2021 03:16:14 UTC Yue Hu (86038448) $Q7:(Confirmed reported problem，stand by $Q8: （onsitecheck，download bugreport $Q9: (download bugreport，system return to service $Q10: (Test and inspection data refer to: PA. Device returned to full service) $Q11: Calibrated Tools name (FLUKE287 ) S/N (20060015) next calibration data (05/17/2021) External Remarks : 29.01.2021 03:16:12 UTC Yue Hu (86038448)服务内容（必填）： 传图问题，抓bugreport 服务结果（必填）： 维修后，设备恢复正常工作 ----------------------------------------------------------------------------- | 5 |  | 传图问题，抓bugreport 服务结果（必填）： 维修后，设备恢复正常工作 https://www.clouddrive.philips.com.cn/file/x53kir40f7kiepxe8bbqewogu1bzbggp# 密码:123 | [12 empty] | 116929722 | New | 567512 | TRUE | Closed - Done | Closed - Done |  | CTUDT00632127 | [5 empty] | Engineering | PACS |  |  | Closed |  |  | PACS | [8 empty] | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
414 |  |  | China | HEB Caofeidian Dist. Hosp. | 345016 | 44222 | 5 | Image transmission problem |  | CFM |  | Peter/Brook Check with field what happened, is it Duplicate with above? | [14 empty] | 562872 | TRUE | Closed - Duplicate | Closed - Duplicate |  | Duplicate | [5 empty] | Engineering | Duplicate with above, NAC? |  |  | WIP |  |  | PACS | [8 empty] | System issue | system | NA | Duplicate | NA-Duplicate | NA-Duplicate | Duplicate | NA-Duplicate | NA | NA-Duplicate | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
415 |  |  | China | Suzhou Guangci Tumor Hospital | 345019 | 44222 | 5 | Application/Ingenuity Core128 | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer : Application Safety Questions:Case: 0116923259 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Pan Chuchu Last updated by: Pan Chuchu 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Field Service Engineer Notes:This PO is a Reference Number only Next Step Action: IA#52655790 was installed by FSE Guang Maopeng 18170605112, customer request first-time training Contacted Guan and said this hospital seems to have done for the 1st time. He will contact the customer to follow up with his agenda ----------------------------------------------------------------------------- | 5 |  | customer request first-time training | [12 empty] | 116923259 | Closed | 564486 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not a defect | [5 empty] | Application | Not a complaint |  |  | Closed |  |  | Application support | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
416 |  |  | China | NMG Alashan Youqi People Hosp | 345031 | 44221 | 5 | 02n/Ingenuity Core128/Elite China/There was no response when the key was turned | BU Q&R Feedback: No SAP Customer Symptom Code: HANG Customer Symptom Code: Hang/Crash Reason: Corrective Maintenance Engineer Failure Code: Wear-out/Breakage/Mechanical Failure Repair Activity Code: Calibration/Adjustment Problem Reported by Customer : There was no response when the key was turned FSE Notes : This PO is a Reference Number only Safety Questions : Case: 0116906694 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Lili Last updated by: Lu Lili 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: There was no response when the key was turned Troubleshooting Action: The emergency stop switch is pressed and held through PRS. Repair Action: Know that the customer restores the emergency stop switch. Test/Inspection: customer confirm can scan patient and image ok.functional Check PASSED. Device returned to full functionality. Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | The emergency stop switch is pressed and held through PRS. Repair Action: Know that the customer restores the emergency stop switch. | [12 empty] | 116906694 | Fixed | 562259 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Not a defect | [5 empty] | Customer | Not a complaint, emergency stop switch is pressed |  |  | Closed |  |  | Mis-use | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
417 |  |  | China | NMG Zhengxiang Baiqi Hosp. | 345033 | 44221 | 5 | 02N/Ingenuity Core128/Elite China/Add a node | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: Instructed Customer Safety Questions : Case: 0116906941 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Yu Zhao Last updated by: Yu Zhao 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A FSE Notes : This PO is a Reference Number only Problem Reported by Customer : Add a node T2 Activities : Problem Description: add new workstation Troubleshooting Action: add new workstation Repair Action: guide customer to add new workstation node Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: NA SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | add new workstation | [12 empty] | 116906941 | Fixed | 562377 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | Not a defect | [5 empty] | Customer | Not a complaint | [14 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
418 |  |  | China | SC Yibin No.2 T.C.M Hosp | 345035 | 44217 | 4 | Spellman 80 KW Cathode Power Module Assembly DEFOA | When short Tube Conditioning, it cannot be exposed. |  |  | Bad spare parts SN： 110228455-B03132 | [12 empty] | DEFOA |  | DEFOA CD-33610 | #N/A | #N/A | DEFOA |  | DEFOA | [5 empty] | Operation | [15 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
419 |  |  | China | SC Yibin No.2 T.C.M Hosp | 345035 | 44217 | 4 | 2 DMS Module DEFOA | Scanning the Phantom found artifacts, and doing bad detectors found problems with detectors 29 and 31. |  |  | Bad spare parts SN：039338 040171 | [12 empty] | DEFOA |  | DEFOA CD-33611 | #N/A | #N/A | DEFOA |  | DEFOA | [5 empty] | Operation | [6 empty] | DMS Module DOA | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
420 |  |  | Chile | CMD Clinica Integrales Rancagu | 40008 | 44246 |  | CL1675-RUIDO EXTRAÑO EN LA CAMILLA DEL PACIENTE | Work Order #: WO-03013870 Valenzuela Velasquez, Eduardo Ariel Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- STRANGE NOISE ON PATIENT TURN Noise on patient table. Check / lubricate patient table movement. | 3 |  | Lubrication | [12 empty] | 117065354 | In Progress | 631296 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | E-125059 | [5 empty] | Engineering | [6 empty] | Table Noise | [8 empty] | System issue | System | Couch-HW | Lubricant | Couch-HW-Lubricant | Couch-HW-Lubricant | E-125059 | Lubricant | Couch-HW | Couch-HW-Lubricant | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
421 |  |  | Chile | HOSPITAL BARROS LUCO TRUDEAU | 341005 | 44243 |  | CL1688 STRANGE NOISE IN THE TABLE | Work Order #: WO-02990203 Valenzuela Velázquez, Eduardo Ariel The system partly meets specification and is returned to use with restrictions Problem Description by Engineer: Noise when lowering the table. Diagnostic performed by Engineer: Noise when lowering the table to a minimum height. Follow Up Required: Vertical spring check-up. Check the interior of the bed due to noise. ----------------------------------------------------------------------------- | 3 |  | Vertical spring check-up. Check the interior of the bed due to noise. @Ada, check the progress of Couch spring ECR | [12 empty] | 117044142 | In Progress | 616782 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00631058 | [5 empty] | Engineering | [6 empty] | Spring position | [8 empty] | System issue | System | Couch-HW | Vertical spring | Couch-HW-Vertical spring | Couch-HW-Vertical spring | CTUDT00631058 | Vertical spring | Couch-HW | Couch-HW-Vertical spring | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
422 |  |  | China | SD Jinan No.4 Peo. Hosp. | 40038 | 44246 |  | 02N/Ingenuity CT 728326/tube is too hot | T2 Activities : Problem Description: 客户问题概述： 计算机软件相关故障。 客户问题详细描述或故障现象： 连续扫描几个病人后会报球管过热 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 软件提示球管过热，需要等待一会儿才能继续扫描 远程诊断与测试结果： 需要检查软件设置和计量设置 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 1，到场检查扫描协议的dose index等 Test/Inspection: N/A Test Equipment: N/A Next Step Action : 已回访 External Remarks : 19.02.2021 13:58:49 UTC Tao Chen (86038373)故障现象: 球管报过热 服务内容: 设置mAs最大最小值 服务结果: 维修后,设备恢复正常. Internal Remarks : 19.02.2021 13:58:52 UTC Tao Chen (86038373) “RDI Accuracy: YES" $Q7: (Confirmed reported problem, tube was too hot intermittently) $Q8: ( Check mA value, found it was too high) $Q9: ( Decreased DoseRight index value and set range of min and max mAs ) $Q10: ( Test and inspection data refer to: PA. Device returned to full functionality) $Q11: Calibrated Tools name ( N/A) ^Current Tube installed date: N/A ^Current Tube usage: N/A ----------------------------------------------------------------------------- | 3 |  | Need check what is the error popup? Ask to provide bugreport and detail scan setting before issue and after correction. Need Jerry and Silva tammy check the scan matrix on 4.1.x, and see if have similar issue on 4.1.10 Related CTUDT: MinmA change to 20 by CTUDT00595661 | [12 empty] | 117059334 | Closed | 630708 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-ScanParameter | [5 empty] | Engineering | [6 empty] | Scan parameter | [8 empty] | System issue | System | TBD-Bugreport | TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | TBD-ScanParameter | TBD_Bugreport | TBD-Bugreport | TBD-Bugreport-TBD_Bugreport | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
423 |  |  | China | JX Yugan Chudong hospital | 337500 | 44245 |  | 02G/Ingenuity Core128/Elite China/consulting | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer : consulting Safety Questions : Case: 0117052597 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Mengen Xie Last updated by: Mengen Xie 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action : 致电MM谢晟013970873443，告知机器还没安装完，他先回电客户 Next Step Action : 致电MM谢晟013970873443，告知没有故障，可以关单了 ----------------------------------------------------------------------------- | 3 |  | Peter follow up to check, this is NAC | [12 empty] | 117052597 | Closed | 625362 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Customer | [6 empty] | NAC | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
424 |  |  | China | XIFENG CHINESE MEDICINE HOSPIT | 345014 | 44246 |  | 02N/Ingenuity Core128/Elite China/gantry fault | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Clean / Lubricate Problem Reported by Customer : gantry fault Safety Questions : Case: 0117059346 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Mengen Xie Last updated by: Mengen Xie 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 计算机软件相关故障。 客户问题详细描述或故障现象： 开始扫描时报错，不能进行扫描 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： S_CTRL_LINK_GHOST_GOT_CRC_FRAMING_ERROR GHOST端有重发 远程诊断与测试结果： 考虑信号刷和滑环问题 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 1，检查信号刷和滑环，并作清洁2，检查配重校准 Test/Inspection: N/A Test Equipment: N/A T2 Activities : Repair Action: replace the signal brush (459800325761) SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : PPD Brush, Signal Block Assembly 459800325761 1 已安排24H到医院，FSE：胡克-86-18585068636 External Remarks : 20.02.2021 15:01:26 UTC SUP JCOUSER (86029494)服务内容：确认报修故障，扫描报错，现场检查发现滑环过脏，清洁滑环和信号刷后设备恢复正常。 服务结果：维修后设备恢复正常。 Internal Remarks : 21.02.2021 16:46:47 UTC Guanglin Tang (86039448) Training on the job ----------------------------------------------------------------------------- | 3 |  | Resend error due to hostiptal cleaning with 84 disinfectant, system works well after slipring and brush cleaning | [12 empty] | 117059346 | In Process | 629441 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | NotEngineering-SlipringClean | [5 empty] | Engineering | [6 empty] | Slip ring cleaning | [8 empty] | System issue | System | Gantry-HW | SlipRingClean | Gantry-HW-SlipRingClean | Gantry-HW-SlipRingClean | NotEngineering-SlipringClean | SlipRingClean | Gantry-HW | Gantry-HW-SlipRingClean | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
425 |  |  | China | HEB Caofeidian Dist. Hosp. | 345016 | 44235 |  | 02N/Ingenuity Core128/Elite China/screen fault | BU Q&R Feedback: No Reason: Corrective Maintenance T2 Activities : Problem Description: 客户问题概述： 显示器故障 客户问题详细描述或故障现象： 一个显示器好的一个显示器不亮没信号输入 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 显示器没信号输入 远程诊断与测试结果： 显示器没信号输入 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 检查显示器的信号线，检查显示器 Problem Reported by Customer : screen fault Safety Questions : Case: 0116997231 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingli Last updated by: Lu Mingli 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 09.02.2021 13:19:16 UTC Yue Hu (86038448)服务内容（必填）： 显示器无显示，重启后恢复 服务结果（必填）： 维修后，设备恢复正常工作 Internal Remarks : 09.02.2021 13:19:18 UTC Yue Hu (86038448) $Q7: (Confirmed reported problem,screen fault ) $Q8: （screen fault reboot system and host computer) $Q9: ( instruct customer,system return to service ) $Q10: (N/A) $Q11: Calibrated Tools name (N/A) S/N (N/A) Next Calibration date (N/A) ----------------------------------------------------------------------------- | 3 |  | Field replaced DVI Adaptor and will try to return for analysis need shelvin to support investigation | [12 empty] | 116997231 | Closed | 594513 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632173 | [5 empty] | Engineering | [6 empty] | DVI Adaptor | [8 empty] | System issue | System | Console-HW | Adaptor | Console-HW-Adaptor | Console-HW-Adaptor | CTUDT00632173 | Adaptor | Console-HW | Console-HW-Adaptor | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
426 |  |  | China | Wuxi Jiangyuan Hosp. | 345020 | 44245 | 8 | 02N/Ingenuity Core128/cannot do short tube condition | Problem Description: cannot do short tube condition Troubleshooting Action: check logger via PRS, Relay open Repair Action: need reboot system Test/Inspection: after reboot ，scan normal.Functional Check PASSED. Device returned to full functionality. Test Equipment: N/A | 2 |  | Need check how fequently customer restart system Rely improvement? Ask to provide can trace for analysis Need check with Robin | [12 empty] | 117054185 | Closed | 621882 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD_Couch Relay | [5 empty] | Engineering | [6 empty] | Relay | [8 empty] | System issue | System | Couch-HW | Relay | Couch-HW-Relay | Couch-HW-Relay | TBD-Couch Relay | Relay | Couch-HW | Couch-HW-Relay | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
427 |  |  | China | NMG Alashan Youqi People Hosp | 345031 | 44239 |  | 02n/Ingenuity Core128/Elite China/Unable to boot | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer: Unable to boot Safety Questions:Case: 0117028134 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Yu Zhao Last updated by: Yu Zhao 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action: Feng Yongsheng call back customer first. Have texted FSE Feng of customer_s phone number, if need to dispatch, call back 800. Service Request 3: Inner Mongolia Alashan Right Banner People_s Hospital. can not start the machine. NN Newly installed machine 64CT Dr. Zhang 13948060336 The customer didn_t provide the equipment information. Next Step Action: Feng Yongsheng PH 18647210453, call back to him again at noon Next Step Action: call back FSE Feng, customer forgot password, solved, the order can be closed. ----------------------------------------------------------------------------- | 3 |  | Password forgot | [12 empty] | 117028134 | Closed | 611136 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Application | [6 empty] | Password | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
428 |  |  | United States | McLaren Macomb | 40021 | 44238 |  | 0711CFN84 Where to change Remote Support password | Work Order #: WO-02969407 Carpenter, Kevin The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Biomed Device Use: Diag Expected and Unexpected Results: Biomed wants to change password for _ Remoteservice _ , per his HCF IT Dept request. User Impact: None Patient Impact: None Current Software Version: Win10 4.1.10 Diagnostic performed by Engineer : Biomed wants to change password for _ Remoteservice _ , per his HCF IT Dept request. Referred biomed to InCenter service document : Ingenuity_Series_and_Brilliance_64_Software_Installation_Manual_for_v4.1.10 459801587361 Rev H See page 204 in Section 5 Cautioned biomed of possible remote support access / support complications if password was nonstandard. Resolution : Referred biomed to InCenter service document : Ingenuity_Series_and_Brilliance_64_Software_Installation_Manual_for_v4.1.10 459801587361 Rev H See page 204 in Section 5 Cautioned biomed of possible remote support access / support complications if password was nonstandard. ----------------------------------------------------------------------------- | 3 |  | Closed - Non-Complaint | [12 empty] | 117022588 | Closed | 607122 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Service | [6 empty] | Closed - Non-Complaint | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
429 |  |  | China | SD Tai_an T.C.M. Hosp. No.2 Ho | 345037 | 44244 |  | 02N/Ingenuity Core128/Elite China/password lock | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: Password lock Safety Questions:Case: 0117046324 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Zixuan Last updated by: Wang Zixuan 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support External Remarks:17.02.2021 14:22:50 UTC Tao Chen (86038373) Error description: The login password is locked Service content: Reset password and set up network Service outcome: Device returns to normal working condition after servicing. Internal Remarks : 18.02.2021 01:45:28 UTC Tao Chen (86038373) “RDI Accuracy: YES" $Q7: (Confirmed reported problem, password was locked) $Q8: ( Do remote desktop connection, found it was failed) $Q9: ( Reset password with IST card) $Q10: (Test and inspection data refer to: PA. Device returned to full functionality) $Q11: Calibrated Tools name ( N/A) ^Current Tube installed date: N/A ^Current Tube usage: N/A ----------------------------------------------------------------------------- | 3 |  | The login password is locked | [12 empty] | 117046324 | Closed | 625173 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Application | [6 empty] | Password | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
430 |  |  | China | SD Tai_an T.C.M. Hosp. No.2 Ho | 345037 | 44235 |  | wechat: 02n/Ingenuity Core128/Elite China/Loss of patient information | BU Q&R Feedback: No SAP Customer Symptom Code: CENV Customer Symptom Code: Customer Environmental Concern Reason: Corrective Maintenance Engineer Failure Code: Environmental Cause Repair Activity Code: Configuration Change Problem Reported by Customer: loss of patient information Safety Questions:Case: 0116998644 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Guomei Li Last updated by: Guomei Li 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Field Service Engineer Notes:This PO is a Reference Number only T2 Activities:Problem Description: Summary of customer_s problems: Need to add query and scan It is a technology problem with the equipment itself: Yes. Troubleshooting Action: remote diagnosis and testing result: Password is locked. Need an engineer to go to the field to unlock and set up PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Need an engineer to go to the field to unlock and set up T2 Activities:Troubleshooting Action: instruct customer to change the lan config Repair Action: instruct customer to change the lan config Test/Inspection: Functional Check PASSED. The device returned to full functionality. Test Equipment: N/A SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | Password forgot | [12 empty] | 116998644 | Closed | 595663 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Application | [6 empty] | Password | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
431 |  |  | Germany | Gesundheit Nord gGmbh | 40002 | 44236 |  | Images did not run correctly into the archive. | Work Order #: WO-02947514 Meier, Sebastian The system fully meets specification and returned to use Problem Description by Engineer: Ms. Schichtel +494214082643 The customer describes that after a thorax image in the overview, the change from the soft tissue window to the lung window gives rise to the final reconstruction being calculated in the lung window despite setting it back to the soft tissue window. Diagnostic performed by Engineer: Asked the customer to pay attention that It is actually being set back to the soft tissue window. Ms. Schichte will test that on a dummy. Resolution: Testing on the Dani worked. The customer allowed himself active time when changing between the windowing (lung and soft tissue window). ----------------------------------------------------------------------------- | 3 |  | Kevin support to help understand and try it | [12 empty] | 117008129 | Closed | 599982 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-WW | [5 empty] | Engineering | [6 empty] | lung and soft tissue window | [8 empty] | System issue | System | Console-SW | WW | Console-SW-WW | Console-SW-WW | TBD-WW | WWWC | Console-SW | Console-SW-WW | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
432 |  |  | Italy | A.O. SPEDALI CIVILI DI BRESCIA | 40026 | 44235 |  | Speaker from console to CAT does not work | Work Order #: WO-02939338 SIGNORINI, DIEGO The system partly meets specification and is returned to use with restrictions Diagnostic performed by Engineer: Checked connections of audio cables of the host. Everything OK. Probably the speaker is broken. Spare part order will follow as well as a machine downtime of one hour Follow Up Required: Order speaker and arrange one hour of machine down time Problem Description by Engineer: Customer_s Function/Role: Tech Device Use: Yes Expected and Unexpected Results: na User Impact: No Patient Impact: No Current Software Version: 4.1.7 reported that speaker of the gantry does not work ----------------------------------------------------------------------------- | 3 |  | Closed - Duplicate | [12 empty] | 117000733 | In Progress | 593004 | TRUE | Closed - Duplicate | Closed - Duplicate | [7 empty] | Others | [6 empty] | Closed - Duplicate | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
433 |  |  | Japan | Nikko Memorial Hospital | 341011 | 44243 |  | 01CS:RS:Mr. Nishida: The exam summary was not created. >> It was not created after he restarted the host either. | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 01CS:RS:Mr. Nishida: The dose report was not created. Safety Questions:Case: 0117040103 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: RUI SATAKE Last updated by: RUI SATAKE 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 4 |  | Check what Does report is customer want? Share screenshot and bugreport. Related CTUDT00632121 FOK: Exam Summary does not show the dose for the scan unless has the SR report turned on & Can not uncheck PACS as a send to Destination if PACS is listed as a Default Destination. | [12 empty] | 117040103 | In Process | 613848 | TRUE | Closed - Done | Closed - Done |  | TBD-DoseReport | [5 empty] | Engineering | [6 empty] | Dose report | [8 empty] | System issue | System | Console-SW | Dose SR | Console-SW-Dose SR | Console-SW-Dose SR | TBD-DoseReport | Dose SR | Console-SW | Console-SW-Dose SR | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
434 |  |  | Japan | Nikko Memorial Hospital | 341011 | 44241 |  | 03CM:TM:Mr. Ando: Arcing occurred during Short TC. >> He was able to continue scanning. >> SoftArcReg/S_GEN_ [TRUNCATED] | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 03CM:TM:Mr. Ando: Arcing occurred. Safety Questions:Case: 0117031930 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: TAKEYUKI MIYAKAWA Last updated by: TAKEYUKI MIYAKAWA 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Problem description by engineer : Equipment affected: CT Failure location: 03High Voltage Energy When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence When, in the end user_s work flow, is the problem occurring?: During Examination ・Measurement preparation (Setting up Protocol) What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem?: NA Engineer_s Comment: _============================== $version = "v4.1.10.25202" 10:14 / Call received from Mr. Ando. Call description: An arc occurred. The X-ray tube was replaced the other day, and arcing also occurred on Friday. _============================== Troubleshooting Action: What steps are taken during the system diagnosis: Result of response: Problem solved. Failure location: 03High Voltage Energy What steps did you perform to diagnose the problem? _================================== Results of checking the log: It seems arcing occurred from the CPM in the last stage of Short TC. 2021/02/14_10:01:42.163__06 <*HV*> szInfo: SoftArcReg 0x211* P1:0/P2: 2021/02/14_10:01:42.259 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_ANALOG_FAULT szInfo:2:5 Analog Fault P1:2/P2:32 2021/02/14_10:01:42.262 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_CATHODE_TUBE_ARC szInfo: 5:5 CPM Arc P1:5/P2: 2021/02/14_10:01:42.268__01 [CANCEL] Series 3 cancelled, due to GANTRY_CANCEL_SOURCE_NORMAL_ERROR, in state CStatePreshot 2021/02/14_10:01:42.276__02 <*HV*> Type: NORMAL ErrNum: S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT szInfo: P1:0/P2: 2021/02/14_10:01:44.320__01 <*ERROR*> Type: NORMAL ErrNum: S_MASTER_SEQUENCER_EXIT_STATE_ERR P1:4096/P2:8369 _================================== Error # and/or description of error: NA Repair Action: Completed or Recommended Repair Action "External": [Confirmation of problem] Confirmed. Work details: I checked the log and reported to the customer. 【Confirmation content】 Good. OneEMS internal remarks : Problem description by engineer : Equipment affected: CT Failure location: 03High Voltage Energy When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence When, in the end user_s work flow, is the problem occurring?: During Examination ・Measurement preparation (Setting up Protocol) What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem?: NA Engineer_s Comment: _============================== $version = "v4.1.10.25202" 10:14 / Call received from Mr. Ando. Call description: An arc occurred. The X-ray tube was replaced the other day, and arcing also occurred on Friday. _============================== What steps are taken during the system diagnosis: Result of response: Problem solved. Failure location: 03High Voltage Energy What steps did you perform to diagnose the problem? _================================== Results of checking the log: It seems arcing occurred from the CPM in the last stage of Short TC. 2021/02/14_10:01:42.163__06 <*HV*> szInfo: SoftArcReg 0x211* P1:0/P2: 2021/02/14_10:01:42.259 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_ANALOG_FAULT szInfo:2:5 Analog Fault P1:2/P2:32 2021/02/14_10:01:42.262 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_CATHODE_TUBE_ARC szInfo: 5:5 CPM Arc P1:5/P2: 2021/02/14_10:01:42.268__01 [CANCEL] Series 3 cancelled, due to GANTRY_CANCEL_SOURCE_NORMAL_ERROR, in state CStatePreshot 2021/02/14_10:01:42.276__02 <*HV*> Type: NORMAL ErrNum: S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT szInfo: P1:0/P2: 2021/02/14_10:01:44.320__01 <*ERROR*> Type: NORMAL ErrNum: S_MASTER_SEQUENCER_EXIT_STATE_ERR P1:4096/P2:8369 _================================== Error # and/or description of error: NA Completed or Recommended Repair Action "External" : [Confirmation of problem] Confirmed. Work details: I checked the log and reported to the customer. 【Confirmation content】 Good. Completed or Recommended Repair Action "Internal": --Troubleshooting Action-- Result of response: Problem solved. Failure location: 03High Voltage Energy What steps did you perform to diagnose the problem? _================================== Results of checking the log: It seems arcing occurred from the CPM in the last stage of Short TC. 2021/02/14_10:01:42.163__06 <*HV*> szInfo: SoftArcReg 0x211* P1:0/P2: 2021/02/14_10:01:42.259 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_ANALOG_FAULT szInfo:2:5 Analog Fault P1:2/P2:32 2021/02/14_10:01:42.262 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_CATHODE_TUBE_ARC szInfo: 5:5 CPM Arc P1:5/P2: 2021/02/14_10:01:42.268__01 [CANCEL] Series 3 cancelled, due to GANTRY_CANCEL_SOURCE_NORMAL_ERROR, in state CStatePreshot 2021/02/14_10:01:42.276__02 <*HV*> Type: NORMAL ErrNum: S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT szInfo: P1:0/P2: 2021/02/14_10:01:44.320__01 <*ERROR*> Type: NORMAL ErrNum: S_MASTER_SEQUENCER_EXIT_STATE_ERR P1:4096/P2:8369 _================================== --Repair Action-- --Subject-- 03CM:TM:Mr. Ando: Arcing occurred during Short TC. >> He was able to continue scanning. >> SoftArcReg/S_GEN_FAULT_CATHODE_TUBE_ARC --Problem Description-- Equipment affected: CT Failure location: 03High Voltage Energy When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence When, in the end user_s work flow, is the problem occurring?: During Examination ・Measurement preparation (Setting up Protocol) What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem?: NA Engineer_s Comment: _============================== $version = "v4.1.10.25202" 10:14 / Call received from Mr. Ando. Call description: An arc occurred. The X-ray tube was replaced the other day, and arcing also occurred on Friday. _============================== --4/8Elements-- Q1. Was the reported problem confirmed? Confirmed. Q2. Were any other problems confirmed? None. Q3. What did you do to resolve the problem?: I checked the log and reported to the customer. Q4: Is the equipment operating normally following the response? Good. Q5. How did the examination finish? None. Q6. Were there any mistaken findings (misdiagnosis, mistaken treatment) due to the problem? None. Q7. Was there any unintended unnecessary exposure? None. Q8. Were there any artifacts? None. ----------------------------------------------------------------------------- | 4 |  | Ask field to get Bugreport for engineering investigation | [12 empty] | 117031930 | In Process | 611032 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00630085 | [5 empty] | Engineering | [6 empty] | HV Arc | [8 empty] | System issue | System | IC-HW | Tub Arc | IC-HW-Tub Arc | IC-HW-Tub Arc | CTUDT00630085 | Tub Arc | IC-HW | IC-HW-Tub Arc | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
435 |  |  | Japan | Nikko Memorial Hospital | 341011 | 44239 |  | 03CM: TM: Error during patient scan >> Scanning possible >> SoftArcReg/: S_GEN_FAULT_ANODE_TUBE_ARC/S_GEN_FA [TRUNCATED] | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 03 Arcing occurred in a wrist scan. Safety Questions:Case: 0117025381 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: TAKEYUKI MIYAKAWA Last updated by: TAKEYUKI MIYAKAWA 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Problem description by engineer : Equipment affected: CT Failure location: 03High Voltage Energy When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user_s work flow did the problem occur?: (During Acquisition/Scan) What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem?: NA Engineer_s Comment: _============================== $version = "v4.1.10.25202" 14:47 / Call from Mr. Tsukakashi: Call description: We replaced the tube yesterday, but an error, which may have been an arc, was observed on the wrist. _============================== Troubleshooting Action: What steps are taken during the system diagnosis: Result of response: Problem solved. Failure location: 03High Voltage Energy What steps did you perform to diagnose the problem? _================================== Results of checking the log: The following errors were recorded. 2021/02/12_14:29:58.583__04 <*HV*> szInfo: SoftArcReg 0x211* P1:0/P2: 2021/02/12_14:29:58.638 <*HV*> Type: NORMAL ErrNum: S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT szInfo: P1:0/P2: 2021/02/12_14:29:58.678 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_ANALOG_FAULT szInfo:2:5 Analog Fault P1:2/P2:32 2021/02/12_14:29:58.680 [CANCEL] CStateShot:Series 72 cancelled, due to GANTRY_CANCEL_SOURCE_NORMAL_ERROR, in state CStateShot 2021/02/12_14:29:58.684 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_ANODE_TUBE_ARC szInfo: 5:0 APM Arc P1:5/P2: 2021/02/12_14:29:58.685 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_CATHODE_TUBE_ARC szInfo: 5:5 CPM Arc P1:5/P2: 2021/02/12_14:30:00.738 <*ERROR*> Type: NORMAL ErrNum: S_MASTER_SEQUENCER_EXIT_STATE_ERR P1:4096/P2:8369 _================================== Error # and/or description of error: NA Repair Action: Completed or Recommended Repair Action "External": [Confirmation of problem] Confirmed. Work details: I checked the log and reported to the customer. 【Confirmation content】 Good. OneEMS internal remarks : Problem description by engineer : Equipment affected: CT Failure location: 03High Voltage Energy When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user_s work flow did the problem occur?: (During Acquisition/Scan) What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem?: NA Engineer_s Comment: _============================== $version = "v4.1.10.25202" 14:47 / Call from Mr. Tsukakashi: Call description: We replaced the tube yesterday, but an error, which may have been an arc, was observed on the wrist. _============================== What steps are taken during the system diagnosis: Result of response: Problem solved. Failure location: 03High Voltage Energy What steps did you perform to diagnose the problem? _================================== Results of checking the log: The following errors were recorded. 2021/02/12_14:29:58.583__04 <*HV*> szInfo: SoftArcReg 0x211* P1:0/P2: 2021/02/12_14:29:58.638 <*HV*> Type: NORMAL ErrNum: S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT szInfo: P1:0/P2: 2021/02/12_14:29:58.678 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_ANALOG_FAULT szInfo:2:5 Analog Fault P1:2/P2:32 2021/02/12_14:29:58.680 [CANCEL] CStateShot:Series 72 cancelled, due to GANTRY_CANCEL_SOURCE_NORMAL_ERROR, in state CStateShot 2021/02/12_14:29:58.684 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_ANODE_TUBE_ARC szInfo: 5:0 APM Arc P1:5/P2: 2021/02/12_14:29:58.685 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_CATHODE_TUBE_ARC szInfo: 5:5 CPM Arc P1:5/P2: 2021/02/12_14:30:00.738 <*ERROR*> Type: NORMAL ErrNum: S_MASTER_SEQUENCER_EXIT_STATE_ERR P1:4096/P2:8369 _================================== Error # and/or description of error: NA Completed or Recommended Repair Action "External" : [Confirmation of problem] Confirmed. Work details: I checked the log and reported to the customer. 【Confirmation content】 Good. Completed or Recommended Repair Action "Internal": --Troubleshooting Action-- Result of response: Problem solved. Failure location: 03High Voltage Energy What steps did you perform to diagnose the problem? _================================== Results of checking the log: The following errors were recorded. 2021/02/12_14:29:58.583__04 <*HV*> szInfo: SoftArcReg 0x211* P1:0/P2: 2021/02/12_14:29:58.638 <*HV*> Type: NORMAL ErrNum: S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT szInfo: P1:0/P2: 2021/02/12_14:29:58.678 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_ANALOG_FAULT szInfo:2:5 Analog Fault P1:2/P2:32 2021/02/12_14:29:58.680 [CANCEL] CStateShot:Series 72 cancelled, due to GANTRY_CANCEL_SOURCE_NORMAL_ERROR, in state CStateShot 2021/02/12_14:29:58.684 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_ANODE_TUBE_ARC szInfo: 5:0 APM Arc P1:5/P2: 2021/02/12_14:29:58.685 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_CATHODE_TUBE_ARC szInfo: 5:5 CPM Arc P1:5/P2: 2021/02/12_14:30:00.738 <*ERROR*> Type: NORMAL ErrNum: S_MASTER_SEQUENCER_EXIT_STATE_ERR P1:4096/P2:8369 _================================== --Repair Action-- --Subject-- 03CTM: Error during patient scan >> Scanning possible >> SoftArcReg/: S_GEN_FAULT_ANODE_TUBE_ARC/S_GEN_FAULT_CATHODE_TUBE_ARC --Problem Description-- Equipment affected: CT Failure location: 03High Voltage Energy When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user_s work flow did the problem occur?: (During Acquisition/Scan) What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem?: NA Engineer_s Comment: _============================== $version = "v4.1.10.25202" 14:47 / Call from Mr. Tsukakashi: Call description: We replaced the tube yesterday, but an error, which may have been an arc, was observed on the wrist. _============================== --4/8Elements-- Q1. Was the reported problem confirmed? Confirmed. Q2. Were any other problems confirmed? None. Q3. What did you do to resolve the problem?: I checked the log and reported to the customer. Q4: Is the equipment operating normally following the response? Good. Q5. How did the examination finish? None. Q6. Were there any mistaken findings (misdiagnosis, mistaken treatment) due to the problem? None. Q7. Was there any unintended unnecessary exposure? None. Q8. Were there any artifacts? None. ----------------------------------------------------------------------------- | 4 |  | Ask field to get Bugreport for engineering investigation | [12 empty] | 117025381 | In Process | 606198 | TRUE | Closed - Done | Closed - Done |  | CTUDT00630085 | [5 empty] | Engineering | [6 empty] | HV Arc | [8 empty] | System issue | System | IC-HW | Tub Arc | IC-HW-Tub Arc | IC-HW-Tub Arc | CTUDT00630085 | Tub Arc | IC-HW | IC-HW-Tub Arc | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
436 |  |  | Japan | Nikko Memorial Hospital | 341011 | 44236 |  | Since after installation, HV arcs continues to occur. Install date:Dec-28-2020 | ・Since after installation, HV arcs continues to occur. Install date:Dec-28-2020 ・Stable operation ・Scan stopped. It was occurred noisy Image. ・Normal operation. Scanning. ・Yes. Noise image appeared. continue scan. ・10times/1month ・No impact to the patient. ・cleaned HV connectors. Filament calibration. several times Long Tube Conditioning. Replaced HV cables. 2021/02/02_13:59:44.037__01 [-oo] - - - - - - Shot 0 Starts - - - - - - 2021/02/02_13:59:49.381__04 <*HV*> szInfo: SoftArcReg 0x211* 2021/02/02_13:59:49.433 <*HV*> Type: NORMAL ErrNum: S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT szInfo: P1:0/P2: 2021/02/02_13:59:49.440 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_ANALOG_FAULT szInfo:2:5 Analog Fault P1:2/P2:32 2021/02/02_13:59:49.443 [CANCEL] CStateShot:Series 66 cancelled, due to GANTRY_CANCEL_SOURCE_NORMAL_ERROR, in state CStateShot 2021/02/02_13:59:49.445__02 [***] - - - - - - Study Exit - - - - - - 2021/02/02_13:59:49.473 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_ANODE_TUBE_ARC 2021/02/02_13:59:49.476 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_CATHODE_TUBE_ARC 2021/02/02_13:59:49.479 <WARN> Type: WARN ErrNum: S_GEN_MGR_SYNC_OFF_OUT_OF_SEQUENCE P1:0/P2:0 2021/02/02_13:59:51.537 <*ERROR*> Type: NORMAL ErrNum: S_MASTER_SEQUENCER_EXIT_STATE_ERR P1:4096/P2:8369 2021/02/09_10:57:54.227__04 <*HV*> szInfo: SoftArcReg 0x211* 2021/02/09_10:57:54.274 <*HV*> Type: NORMAL ErrNum: S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT szInfo: P1:0/P2: 2021/02/09_10:57:54.280 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_ANALOG_FAULT szInfo:2:5 Analog Fault P1:2/P2:32 2021/02/09_10:57:54.284 [CANCEL] CStateShot:Series 49 cancelled, due to GANTRY_CANCEL_SOURCE_NORMAL_ERROR, in state CStateShot 2021/02/09_10:57:54.287__01 [***] - - - - - - Study Exit - - - - - - 2021/02/09_10:57:54.288 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_ANODE_TUBE_ARC 2021/02/09_10:57:54.317 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_CATHODE_TUBE_ARC 2021/02/09_10:57:56.338__01 <*ERROR*> Type: NORMAL ErrNum: S_MASTER_SEQUENCER_EXIT_STATE_ERR P1:4096/P2:8369 2021/02/09_10:58:10.088 [---] ------ Series 50 Load Structure ----- 2021/02/09_10:58:20.998__02 [--o] ----------------- Series 50 Ready ----------------- I sent BUGREP to Haifa using SDT. | CFM |  | Peter follow up to get bugreport from submitter | [14 empty] | 598500 | TRUE | Closed - Done | Closed - Done |  | CTUDT00630085 | [5 empty] | Engineering | [6 empty] | HV Arc | [8 empty] | System issue | System | IC-HW | Tub Arc | IC-HW-Tub Arc | IC-HW-Tub Arc | CTUDT00630085 | Tub Arc | IC-HW | IC-HW-Tub Arc | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
437 |  |  | Japan | Nikko Memorial Hospital | 341011 | 44236 |  | 03DM:RS:Mr. Ando: A streak artifact appeared on the final two or three slices. >> S_RHOST_HVG_HSS_GENERATOR_ [TRUNCATED] | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 03DS:RS:ｱﾝﾄﾞｳ様:最後の2,3Sliceでストリークアーチファクトが発生>> Safety Questions : Case: 0117005010 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: RUI SATAKE Last updated by: RUI SATAKE 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 対象装置:CT 故障個所:03High Voltage Energy 最初に発生したのはいつか?(~以内):1 Hour 問題発生の頻度は?:Multiple times as week エンドユーザｰのワークフローのどの時点において問題が発生しているのか?:検査・測定中(During Acqustion/Scan) 最近どのようなイベントが発生したか?:N.A. 顧客は問題について具体的なシステムイベントレポートを作成しているか?:N.A. Troubleshooting Action: 対応後の結果:改善された 故障個所:03High Voltage Energy 診断中にどのようなステップを実施したか?: その後の撮影ではストリーク発生せず。 該当の時間ログからエックス線管球のアークの発生を確認。 FSE対応。 Repair Action: サービス担当の訪問作業。 OneEMS internal remarks : --Troubleshooting Action-- 対応後の結果:改善された 故障個所:03High Voltage Energy 診断中にどのようなステップを実施したか?: その後の撮影ではストリーク発生せず。 該当の時間ログからエックス線管球のアークの発生を確認。 FSE対応。 --Repair Action-- Recommended 12Nc: PR1 989000003302:HV CABLE PHILIPS MRC CT-O4S 0,95m QTY:1 SLOC:札幌(U612) PR1 989000005011:HV-ANODECABLE PHILIPS MRC880 CT 1,84m QTY:1 SLOC:札幌(U612) --Subject-- 03DM:RS:ｱﾝﾄﾞｳ様:最後の2,3Sliceでストリークアーチファクトが発生>>S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT/SoftArcReg 0x211* /ANODE/CATHODE --Problem Description-- 対象装置:CT 故障個所:03High Voltage Energy 最初に発生したのはいつか?(~以内):1 Hour 問題発生の頻度は?:Multiple times as week エンドユーザｰのワークフローのどの時点において問題が発生しているのか?:検査・測定中(During Acqustion/Scan) 最近どのようなイベントが発生したか?:N.A. 顧客は問題について具体的なシステムイベントレポートを作成しているか?:N.A. SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 11.02.2021 13:14:44 UTC HODAKA HASHIMOTO (81100553)[ご依頼内容] 最後の2,3Sliceでストリークアーチファクトが発生 [現象確認] 確認出来ませんでした。 [作業内容詳細] 2021年2月3日の現象を思われる画像を確認。 2021年2月9日 高圧系周辺に何らかの問題があると判断し発生器からエックス線管にかけての高圧ケーブルの交換を実施。 また、エージング作業複数回実施。 照射時間：23817秒 アーク数：Anode 9回 Cathode 18回 同時 8回 2021年2月10日 昨日の作業立ち合い中、同現象再発。 2021年2月11日 ここ数日の現象がエックス線管によるものと判断しエックス線管交換を実施。 交換前の最終データ 照射時間：25335秒 アーク数：Anode 10回 Cathode 19回 同時 9回 交換後、装置に問題ないことを確認。テストスキャンにてイメージクオリティにも問題ないことを確認しております。 交換後の新しいエックス線管データ 照射時間：合計時間 27113秒（新エックス線管：1778秒） アーク数：Anode 13回 Cathode 23回 （新エックス線管 Anode3回 Cathode 4回 同時2回） [他の故障/処置] 無し [装置の状態] 装置使用可 [作業の状況] 作業完了 11-Feb-2021 81100553 FSE Handover Tool1.3.4.32612.DB13 20200715 Internal Remarks : 11.02.2021 13:14:47 UTC HODAKA HASHIMOTO (81100553) [交換パーツ] 交換パーツ有り 12Nc:989000003302 名称:HV CABLE PHILIPS MRC CT-O4S 0,95m QTY:1 12Nc:989000005011 名称:HV-ANODECABLE PHILIPS MRC880 CT 1,84m QTY:1 12Nc:459800089031 名称:Tool, XRT Strap Pad Assy QTY:1 12Nc:453566495091 名称:1XRT SERVICE TOOL ASSY EXP QTY:1 12Nc:989000086352 名称:PHILIPS MRC 880 Replacement QTY:1 [ミスオーダーパーツレポート] ミス無し [様子見情報] 様子見無し [TS Requests] 返却パーツ無し [CT追加情報] Q1その検査はどの様に終了しましたか?: Q1 Noの為対象外 Q2 問題により間違った所見（誤診、誤治療）の有無: 無し Q3 アーチファクトの有無: 無し Q4 意図しない不要な曝射の有無: 無し [修理済証] 修理済証の貼り付け有り [その他情報共有事項] １、新型コロナウィルス感染患者もしくは感染の疑いがある患者に対して治療/ 検疫が行われていますか？ はい ２、感染患者のいる治療室及び病床に入る必要がありますか？（作業場並びに移 いいえ 機器担当者と報告完了方法を合意できましたか？ YES 当日通常報告完了 11-Feb-2021 81100553 FSE Handover Tool1.3.4.32612.DB13 20200715 ----------------------------------------------------------------------------- | 3 |  | Please share the image with artifact 12Nc:989000003302 名称:HV CABLE PHILIPS MRC CT-O4S 0,95m QTY:1 12Nc:989000005011 名称:HV-ANODECABLE PHILIPS MRC880 CT 1,84m QTY:1 12Nc:459800089031 名称:Tool, XRT Strap Pad Assy QTY:1 12Nc:453566495091 名称:1XRT SERVICE TOOL ASSY EXP QTY:1 12Nc:989000086352 名称:PHILIPS MRC 880 Replacement QTY:1 | [12 empty] | 117005010 | In Process | 594753 | TRUE | Closed - Done | Closed - Done |  | TBD-Artifact | [5 empty] | Engineering | [6 empty] | Streak artifact | [8 empty] | System issue | System | DMS-SW | Artifact | DMS-SW-Artifact | DMS-SW-Artifact | TBD-Artifact | Artifact | DMS-SW | DMS-SW-Artifact | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
438 |  |  | Japan | KYOTO MEDICAL CENTER | 341013 | 44245 |  | 01CS:KKI: SAS synchronization was interrupted and the Tracker scan could not be started | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 01CS:KKI: SAS synchronization was interrupted and the Tracker scan could not be started ----------------------------------------------------------------------------- | 4 |  | Peter support to get bugreport from submitter Please share the bugreport which mentioned in Service Notes. Meanwhile please help check below elements: 1. what injector type and connection was used for this site? 2. Please measure TB252-1 & 2 AC Voltage before estop closed and after estop closed (~60mv). 3. What is power environment at the site | [12 empty] | 117052141 | In Process | 620984 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-SAS | [5 empty] | Engineering | [6 empty] | SAS | [8 empty] | System issue | System | Gantry-HW | Injector | Gantry-HW-Injector | Gantry-HW-Injector | TBD-SAS | Injector | Gantry-HW | Gantry-HW-Injector | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
439 |  |  | Japan | KYOTO MEDICAL CENTER | 341013 | 44242 |  | 04DS: TM: via FSE: The gantry makes a rattling noise when it rotates >> FSE response | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 04DS:TM:FSE経由:Gantry回転時にカラカラという異音>> SAP Transfer : Dispatched for Onsite Support SAP Transfer : Case Transferred to SAP Safety Questions : Case: 0117032973 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: TAKEYUKI MIYAKAWA Last updated by: TAKEYUKI MIYAKAWA 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: Problem description by engineer : 対象装置:CT 故障個所:04Gantry Exclude HV 最初に発生したのはいつか?(~以内):1 Hour 問題発生の頻度は?:First Occurrence エンドユーザｰのワークフローのどの時点において問題が発生しているのか?:その他(Other) 最近どのようなイベントが発生したか?:N.A. 顧客は問題について具体的なシステムイベントレポートを作成しているか?:N.A. Engineer_s Comment: _============================== コール内容 Gantry回転時にカラカラという異音 _============================== Troubleshooting Action: What steps are taken during the system diagnosis : 対応後の結果:改善されなかった 故障個所:04Gantry Exclude HV 診断中にどのようなステップを実施したか?: _================================== 対応 FSE対応とします _================================== Error # and/or description of error : NA Repair Action: Completed or Recommended Repair Action "External" : サービス担当の訪問作業。 OneEMS internal remarks : Problem description by engineer : 対象装置:CT 故障個所:04Gantry Exclude HV 最初に発生したのはいつか?(~以内):1 Hour 問題発生の頻度は?:First Occurrence エンドユーザｰのワークフローのどの時点において問題が発生しているのか?:その他(Other) 最近どのようなイベントが発生したか?:N.A. 顧客は問題について具体的なシステムイベントレポートを作成しているか?:N.A. Engineer_s Comment: _============================== コール内容 Gantry回転時にカラカラという異音 _============================== What steps are taken during the system diagnosis : 対応後の結果:改善されなかった 故障個所:04Gantry Exclude HV 診断中にどのようなステップを実施したか?: _================================== 対応 FSE対応とします _================================== Error # and/or description of error : NA Completed or Recommended Repair Action "External" : サービス担当の訪問作業。 Completed or Recommended Repair Action "Internal" : --Troubleshooting Action-- 対応後の結果:改善されなかった 故障個所:04Gantry Exclude HV 診断中にどのようなステップを実施したか?: _================================== 対応 FSE対応とします _================================== --Repair Action-- --Subject-- 04DS:TM:FSE経由:Gantry回転時にカラカラという異音>>FSE対応 --Problem Description-- 対象装置:CT 故障個所:04Gantry Exclude HV 最初に発生したのはいつか?(~以内):1 Hour 問題発生の頻度は?:First Occurrence エンドユーザｰのワークフローのどの時点において問題が発生しているのか?:その他(Other) 最近どのようなイベントが発生したか?:N.A. 顧客は問題について具体的なシステムイベントレポートを作成しているか?:N.A. Engineer_s Comment: _============================== コール内容 Gantry回転時にカラカラという異音 _============================== Internal Remarks : 22.02.2021 00:50:10 UTC YUKO YAMAMOTO (81100269) [交換パーツ] 交換パーツ有り 12Nc:453567487011 名称:MOTOR, RPM AC 2 HP, 230 VAC, 3PH QTY:1 12Nc:453566492401 名称:POLY-VEE BELT QTY:1 12Nc:459801265631 名称:VERT DRIVE REPLACEMENT TOOL KIT QTY:1 [ミスオーダーパーツレポート] ミス無し [様子見情報] 様子見無し [TS Requests] 返却パーツ無し [CT追加情報] Q1その検査はどの様に終了しましたか?: Q1 Noの為対象外 Q2 問題により間違った所見（誤診、誤治療）の有無: 無し Q3 アーチファクトの有無: 無し Q4 意図しない不要な曝射の有無: 無し [修理済証] 修理済証の貼り付け有り [その他情報共有事項] １、新型コロナウィルス感染患者もしくは感染の疑いがある患者に対して治療/ 検疫が行われていますか？ はい ２、感染患者のいる治療室及び病床に入る必要がありますか？（作業場並びに移 動経路含む） いいえ 機器担当者と報告完了方法を合意できましたか？ YES 翌日立会い 19-Feb-2021 81100269 FSE Handover Tool1.3.4.32612.DB13 20200715 External Remarks : 22.02.2021 00:50:07 UTC YUKO YAMAMOTO (81100269)[ご依頼内容] ガントリー回転時に異音がする [現象確認] 確認出来ました。 [作業内容詳細] 異音の対応として、ガントリー回転用モーターとドライブベルトの交換を実施しました。 [他の故障/処置] 無し [装置の状態] 装置使用可 [作業の状況] 作業完了 [その他] 同行作業者Ope:30 /種類:CMAIL/作業者:TE005山中康司 /作業日 2021年2月19日 16:00-20:30/作業時間4.5Hr ----------------------------------------------------------------------------- | 5 |  | check with field, suggest to adjust the belt force before parts replacement. does one or all below parts cause noise issue? 12Nc: 453567487011 Name: MOTOR, RPM AC 2 HP, 230 VAC, 3PH QTY: 1 12Nc: 453566492401 Name: POLY-VEE BELT QTY: 1 12Nc: 459801265631 Name: VERT DRIVE REPLACEMENT TOOL KIT QTY: 1 | [12 empty] | 117032973 | In Process | 611475 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-GantryNoise | [5 empty] | Engineering | [6 empty] | Gantry noise | [8 empty] | System issue | System | Gantry-HW | Noise | Gantry-HW-Noise | Gantry-HW-Noise | TBD-GantryNoise | Noise | Gantry-HW | Gantry-HW-Noise | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
440 |  |  | Japan | Tomakomai City Hospital | 341014 | 44244 |  | 01CS:KKI: Image transfer to Server 2 is not possible >> Restarting the CT host resolved the problem | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 01CS:KKI: Image transfer to Server 2 is not possible >> Restarting the CT host resolved the problem ----------------------------------------------------------------------------- | 4 |  | PACS Transfer issue, suggest to implement FCO72800740 | [12 empty] | 117045719 | In Process | 617214 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS | [8 empty] | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
441 |  |  | Japan | Tomakomai City Hospital | 341014 | 44239 |  | Didn_t become "Ready" during scan and would not proceed >> | 01CS: TM: Mr. Masuda: I had two plans, but the initialization did not start immediately after the first scan, and it took a long time to become Ready >> No error code. It may not be due to trouble --Problem Description-- Target device: CT Failure location: 01Cosole When did it first occur? (Within ~): 1 Hour How often do problems occur ?: First Occurrence At what point in the end-user workflow is the problem occurring ?: During inspection / measurement (During Acqustion / Scan) What happened recently ?: N.A. Does the customer create a specific system event report for the problem ?: N.A. Engineer_s Comment: _============================== $ version = "v4.1.10.25202" | 4 |  | took a long time to get ready because the initialization didn_t start immediately after the first scan. No error code suggest to implement HF2&3, and see if have improvement on the time | [12 empty] | 117024772 | In Process | 605955 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-Initilization | [5 empty] | Engineering | [6 empty] | Initialization time | [8 empty] | System issue | System | Console-SW | Performance | Console-SW-Performance | Console-SW-Performance | TBD-Initilization | Performance | Console-SW | Console-SW-Performance | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
442 |  |  | Japan | Saiseikai Gosu Hospital | 343000 | 44249 |  | 01CS:KKI:昨日、CIRS S1の電源が落ちていた＞＞電源を投入し、その後正常動作 | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 01CS:KKI:昨日、CIRS S1の電源が落ちていた＞＞電源を投入し、その後正常動作 ----------------------------------------------------------------------------- CIRS S1 was turned off >> Turn on the power and then operate normally | 4 |  | CIRS power off | [12 empty] | 117069439 | In Process | 631911 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | CIRS HW | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
443 |  |  | Japan | Saiseikai Gosu Hospital | 343000 | 44245 |  | 01CS:KK:Mr. Moriwaki: Image transfer to the ISP stops frequently.>>Recovered by restarting the host. | Does the customer create a specific system event report for the problem ?: No. The customer was unable to create a log file or report. Troubleshooting Action: Result after response: Improved Failure location: 01Cosole What steps did you take during the diagnosis ?: Recovered by restarting the Host. Since it occurs frequently, we decided to investigate. Corresponding to Mr. Kawashige. Repair Action: Service representative visit work. OneEMS internal remarks: --Troubleshooting Action-- Post-response results: Improved Failure location: 01Cosole What steps did you take during the diagnosis ?: Recovered by restarting the Host. Since it occurs frequently, we decided to investigate. | 4 |  | PACS Transfer issue, suggest to implement FCO72800740 | [12 empty] | 117052051 | In Process | 620946 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS | [8 empty] | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
444 |  |  | Peru | Hogar Clinica San Juan de Dios | 40034 | 44243 |  | 40034 - ESTABLISH REMOTE CONNECTION | 40034 - ESTABLISH REMOTE CONNECTION | 5 |  | Remote connection | [12 empty] | 117043481 | New | 619523 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Service | [6 empty] | PRS Setup | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
445 |  |  | Spain | S.C.S. HOSP. GRAL. DE FUERTEVE | 40020 | 44237 |  | It does not send to Portal. | Work Order #: WO-02954236 16-Feb-2021 Gabriel Limiñana Perez The system fully meets specification and returned to use. Problem Description by Engineer: The equipment does not send to portal. Regulatory Questions : N/A Resolution: After connecting to the equipment, we observe that the studies are not being sent to Portal since the previous day in the afternoon. Although the sending from TAC to the Portal does not work, when connecting to the Portal it was possible to take the studies from the TAC. This action is performed with the pending studies, as well as subsequent sending from portal to EBW upon request by a doctor. Once the user stopped using the TAC for a while, we proceeded to restart the host after which the equipment was repaired. We confirm that in the following studies are automatically sent without any problem. There is already a case escalated to the factory as well as a customer feedback opened for the repetition of sending problems to different nodes that are solved by restarting the host. ----------------------------------------------------------------------------- | 3 |  | PACS Transfer issue, suggest to implement FCO72800740 EBW complaint | [12 empty] | 117014316 | In Progress | 612260 | #N/A | #N/A | Evaluation |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS | [5 empty] | Non-Refresh |  |  | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | y
446 |  |  | Sweden | Avesta lasarett | 40040 | 44244 |  | The ceiling-mounted monitor is not starting up. | Work Order #: WO-02994141 20-Feb-2021 TJALLDIN HANS GUNNAR Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 3 |  | 459801140821 - 19" LCD display monitor EA193Mi | [12 empty] | 117047474 | In Progress | 622079 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632173 | [5 empty] | Engineering | [6 empty] | Monitor | [8 empty] | System issue | System | Console-HW | Monitor | Console-HW-Monitor | Console-HW-Monitor | CTUDT00632173 | Monitor | Console-HW | Console-HW-Monitor | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
447 |  |  | Japan | KYOTO MEDICAL CENTER | 341013 | 44235 |  | JY: CT Application support package. | BU Q&R Feedback: No Reason: Application Support Problem Reported by Customer : JY: CT Application support package. Safety Questions:Case: 0117004061 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: YAN Joanna Last updated by: YAN Joanna 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A OneEMS internal remarks : Otani DB SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Support External Remarks: 17.02.2021 07:55:44 UTC NATSUKI OOTANI (81100565) [Request details] Check problem and provide data on scattered radiation. Work details: I retrieved the data for reporting purposes. I gave the customer scattered radiation data documentation. Internal Remarks : 17.02.2021 07:55:47 UTC NATSUKI OOTANI (81100565) 1. Are patients with COVID-19 infection or patients who are suspected of having COVID-19 infection being treated or quarantined ? Yes 2. Is it necessary to enter the treatment room or the bed of infected patients?(Including workplace and transfer routes) No ----------------------------------------------------------------------------- | 5 |  | Closed - Non-Complaint | [12 empty] | 117004061 | New | 598220 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Application | [6 empty] | Closed - Non-Complaint | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
448 |  |  | United States | The Regional Medical Center | 40005 | 44236 |  | 0899CDX84 Periodic USB plugin / plugout Sound coming from CT box & gantry | Work Order #: WO-02946239 Ali, Adam The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Problem description by engineer : Periodic USB plug-in / plug-out Sound coming from CT box & gantry When was the first occurrence? Within... : 03-Feb-2021 How often is issue occurring? : Intermittent Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: Not provided Patient Impact/Outcome: Not provided User Impact: n/a Current Software Version: Not provided Diagnostic performed by Engineer : Periodic USB plug-in / plug-out Sound coming from CT box & gantry. Repeat issue to 0116976739 Internal Comments : Error # and/or description of error : Periodic USB plug-in / plug-out Sound coming from CT box & gantry. Completed or Recommended Repair Action "Internal" : Continue to further troubleshoot as per previous case 0116976739 Replace USB cables on IBOX , GPC, CRC, HOST Follow Up Required : What are the next Troubleshooting steps : Continue to further troubleshoot as per previous case 0116976739 Replace USB cables on IBOX , GPC, CRC, HOST Error # and/or description of error : Periodic USB plug-in / plug-out Sound coming from CT box & gantry. Completed or Recommended Repair Action "Internal" : Continue to further troubleshoot as per previous case 0116976739 Replace USB cables on IBOX , GPC, CRC, HOST ----------------------------------------------------------------------------- Work Order #: WO-02946353 18-Feb-2021 Jensen George The system fully meets specification and returned to use Problem Description by Engineer : Problem description by engineer : Periodic USB plug-in / plug-out Sound coming from CT box & gantry When was the first occurrence? Within... : 03-Feb-2021 How often is issue occurring? : Intermittent Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: Not provided Patient Impact/Outcome: Not provided User Impact: n/a Current Software Version: Not provided Diagnostic performed by Engineer : Problem intermittent and could not be reproduced. ----------------------------------------------------------------------------- | 4 |  | Suggest field to order the improved USB cable to replace the USB extender kit. 459801731761 USB EXTENDER CABLE FROM IBOX TO CONSOLE | [12 empty] | 117007079 | Closed | 596054 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-USBExtended | [5 empty] | Engineering | [6 empty] | USB connection | [8 empty] | System issue | System | Gantry-HW | USBExtended | Gantry-HW-USBExtended | Gantry-HW-USBExtended | TBD-USBExtended | USBExtended | Gantry-HW | Gantry-HW-USBExtended | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
449 |  |  | United States | Lackey Memorial Hospital | 40015 | 44245 | 8 | bottom portion of the scan is not diagnostic images | Work Order #: WO-03003680 Maier, Donnie The system does not meet specification and is out of use Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: Diag Expected and Unexpected Results: Couch Free Floating User Impact: No Patient Impact: No Current Software Version: 4.1.10 Diagnostic performed by Engineer : PFSE Chuck Denison is already aware and enroute. No access to system remotely (New System). Initially cu states they were able to scan (didn_t automatically start injection) but last images were undiagnostic. Throughout entire image. Cu did get error message unclear of what error was. I talked cu thru exercising all 5 estop buttons, footswitches (no rear footswitch), etc. We power cycled system completely (Host/CIRS/Gantry) and still have same symptom (Click, but no thunk). Keyswitch prompt goes away but couch is still free floating. I had cu create another bug report after complete power cycle. I left a VM for PFSE Denison as to our actions but further onsite troubleshooting is necessary. Follow Up Required : PFSE Chuck Denison is already aware and enroute. No access to system remotely (New System). Initially cu states they were able to scan (didn_t automatically start injection) but last images were undiagnostic. Throughout entire image. Cu did get error message unclear of what error was. I talked cu thru exercising all 5 estop buttons, footswitches (no rear footswitch), etc. We power cycled system completely (Host/CIRS/Gantry) and still have same symptom (Click, but no thunk). Keyswitch prompt goes away but couch is still free floating. I had cu create another bug report after complete power cycle. I left a VM for PFSE Denison as to our actions but further onsite troubleshooting is necessary. ----------------------------------------------------------------------------- Work Order #: WO-03004657 18-Feb-2021 Denison Chuck The system does not meet specification and is out of use Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: Diag Expected and Unexpected Results: Couch Free Floating User Impact: No Patient Impact: No Current Software Version: 4.1.10 Diagnostic performed by Engineer : Trouble shoot Problem with C.T. scanner not coming to ready and will not scan. Follow Up Required : Replace the G-host PCB. and Check C.T. performance. Tools & Equipment : null,2023-02-28 00:00:00 ----------------------------------------------------------------------------- Work Order #: WO-03006643 19-Feb-2021 Donald Vic The system does not meet specification and is out of use Problem Description by Engineer : Customer Function/Role: Radiology Tech Device Use: clinical use Expected and Unexpected Results: while scanning a patient the CT images had artifacts. The customer tried restarting the Host computer to resolve the issue. Once the Host was back online and ESTOP closed the table was in free float and the rotor would not find home. User Impact: The CT was not usable Patient Impact: No patient impacted Current Software Version: 4.1.10 Diagnostic performed by Engineer : The results of the troubleshooting lead to U44 being defective on the GHOST. The missing signal needed from U44 was naught MKDON. This has been a recurring issue at this site. The issue was initially showing up on a Brilliance 16 CT and now it is presenting on a Ingenuity CT 5000 which was installed in DEC 2020. The issue has historically presented after power fluctuation at the site. The CT is protected by a STACO 25kVA UPS. | 2 |  | Need bugreport after power cycle check what injector type and connection was used for this system? Check slipring as well to see if the voltage are in correct status. @Yaping, support investigation | [12 empty] | 117056364 | In Progress | 623561 | TRUE | Closed - Done | Closed - Done |  | TBD-Ghost | [5 empty] | Engineering | [6 empty] | Ghost | [8 empty] | System issue | System | Gantry-HW | Ghost | Gantry-HW-Ghost | Gantry-HW-Ghost | TBD-Ghost | Ghost | Gantry-HW | Gantry-HW-Ghost | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
450 |  |  | United States | Lackey Memorial Hospital | 40015 | 44238 |  | Making an ongoing noise. | Work Order #: WO-02965131 12-Feb-2021 Denison Chuck The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Dept. Head Device Use: Philips C.T. Expected and Unexpected Results: None C.T. working as should. User Impact: None Patient Impact: None Current Software Version:4.1.8 Resolution : Checked the Host Computer. Nothing problems in the Device manager. Nothing in the Logs showed any problems. ----------------------------------------------------------------------------- Work Order #: WO-02971660 12-Feb-2021 Denison Chuck The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Dept. Head Device Use: Philips C.T. scanner Expected and Unexpected Results: Working as should User Impact: None Patient Impact: None Current Software Version:4.1.8 Diagnostic performed by Engineer : Worked with customer to make sure Philips C.T. scanner was working as it should and no more noise from the Host CPU. ----------------------------------------------------------------------------- | 5 |  | Worked with customer to make sure Philips C.T. scanner was working as it should and no more noise from the Host CPU. | [12 empty] | 117023304 | Closed | 607520 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-HostNoise | [5 empty] | Engineering | [6 empty] | Host noise | [8 empty] | System issue | System | Console-HW | Noise | Console-HW-Noise | Console-HW-Noise | TBD-HostNoise | Noise | Console-HW | Console-HW-Noise | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
451 |  |  | United States | Lackey Memorial Hospital | 40015 | 44237 |  | Issues With Patient Monitor | Work Order #: WO-02958364 11-Feb-2021 Turley Rich The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: CT Technologist Device Use: Normal Expected and Unexpected Results: ECG Monitor arrived User Impact: none Patient Impact: none Current Software Version: Resolution : assembled the ECG monitor, tested , turned over to customer with it charging | 5 |  | ECG connection | [12 empty] | 117017705 | In Progress | 602379 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | ECG | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
452 |  |  | United States | McLaren Macomb | 40021 | 44247 |  | 0699CDN84 Injections not working correctly | Error # and/or description of error : Injections not working correctly Completed or Recommended Repair Action "Internal" : Insure all preferences for SAS are turned, tech thinks they are. Verify connections from injector to scanner. Confirm signal from injector to scanner. Note to FSE: please keep PRS / remote access passwords updated in PRS for this site. I was unable to login remotely. In-House Bio-Med resolved issue | 3 |  | need get bugreport from submitter Meanwhile please help check below elements: 1. what injector type and connection was used for this site? 2. Please measure TB252-1 & 2 AC Voltage before estop closed and after estop closed (~60mv). 3. What is power environment at the site | [12 empty] | 117068197 | In Progress | 631498 | TRUE | Closed - Done | Closed - Done |  | TBD-SAS | [5 empty] | Engineering | [6 empty] | SAS | [8 empty] | System issue | System | Gantry-HW | Injector | Gantry-HW-Injector | Gantry-HW-Injector | TBD-SAS | Injector | Gantry-HW | Gantry-HW-Injector | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
453 |  |  | United States | Leesburg Regional Medical Cent | 40022 | 44236 |  | FSE locked out of CT system | Work Order #: WO-02948637 Mujica, David The system partly meets specification and is returned to use with restrictions Problem description by Clinical Support Specialist : Customer Function/Role: FSE Device Use: Partially in use Expected and Unexpected Results: System locked due to incorrect password User Impact: none Patient Impact: none Current Software Version: 4.1 Diagnostic performed by Engineer : FSE Don called in because he locked the customer out of the CT system because of too many failed attempts. Shut down and restart attempted, but did not work. Also, using the correct password was attempted but the system indicates too many failed attempts. At this time CT clinical support cannot remedy this issue and FSE Don was told to call his NSS for support. No further clinical support needed at this time. | 4 |  | system locked due to incorrect password | [12 empty] | 117008985 | Closed | 597214 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Application | [6 empty] | Password | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
454 |  |  | United States | Leesburg Regional Medical Cent | 40022 | 44236 |  | Failed to send images to PACS; Cleared sending queue cache files to resolve. | Work Order #: WO-02946726 Carpenter, Kevin The system partly meets specification and is returned to use with restrictions Diagnostic performed by Engineer : Forwarded repeat case to Field and advise FSE in follow-up activity. Problem Description by Engineer : Customer Function/Role: CTTech Device Use: Diag Expected and Unexpected Results: Ongoing problem sending images to PACS. FSE please consult NSS as cust reports this as Repeat Problem. CT must be rebooted to fix - Other CT does not have this problem, nor do other modalities. User Impact: None Patient Impact: None Current Software Version: 4.1.10 Follow Up Required : Ing SWVer 4.1.10 ~ Images fail to PACS unless CTHost is rebooted. Ongoing problem sending images to PACS. FSE please consult NSS as cust reports this as Repeat Problem. CT must be rebooted to fix - Other CT does not have this problem, nor do other modalities. CT RSE notes that if CTHost NIC is is config_ed as _ AUTO _ in WinOS for Speed / Duplex then that can cause the reported problem. In those cases, FSE asked HCF IT Dept for network Speed/Duplex, then change CTHost settings in WinOS from AUTO to the manual settings given by HCF. Discuss with NSS as this may help. Also run TRACERT against PACS IP Address and check CTHost stats for Dropped / Packet Collisions. ----------------------------------------------------------------------------- Work Order #: WO-02946784 09-Feb-2021 Stoutt Chad The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: CTTech Device Use: Diag Expected and Unexpected Results: Ongoing problem sending images to PACS. FSE please consult NSS as cust reports this as Repeat Problem. CT must be rebooted to fix - Other CT does not have this problem, nor do other modalities. User Impact: None Patient Impact: None Current Software Version: 4.1.10 Resolution : Cleared sending queue cache files after discussing problem with specialist. ----------------------------------------------------------------------------- | 3 |  | PACS Transfer issue, suggest to implement FCO72800740 | [12 empty] | 117007377 | In Progress | 596260 | TRUE | Closed - Done | Closed - Done |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS | [8 empty] | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
455 |  |  | China | SX Datong Yanggao People Hosp. | 345025 | 44239 |  | 02N/Core128/Consultation | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer: consultation Safety Questions:Case: 0117024765 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Pan Chuchu Last updated by: Pan Chuchu 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action: Informed Chinese New Year shift Engineer Liu Guoping 18066898085 follow up with customer first on feedback, SMS has been sent Next Step Action: Teacher Zhang 17860615325 feedback, logon account was locked when turn on the machine in the morning. Need to consult how to handle ----------------------------------------------------------------------------- | 5 |  | Closed - Non-Complaint | [12 empty] | 117024765 | Closed | 609232 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Application | [6 empty] | Closed - Non-Complaint | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
456 |  |  | United States | Unity Health Systems | 40037 | 44246 |  | Q/O:Power Monitor Install | Expected and Unexpected Results: Installed Power Monitor. Will Remove on 2/26/21. | 5 |  | what is power monitor? For customer power enrionment? | [12 empty] | 117062670 | In Progress | 630622 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Others | [6 empty] | Power monitor | [8 empty] | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
457 |  |  | United States | Unity Health Systems | 40037 | 44246 |  | 0999CDN83 Artifact and Noise | Work Order #: WO-03015101 Morrow, Dan The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Problem description by engineer : Line artifact and Noise during gantry rotation When was the first occurrence? Within... : 19-Feb-2021 How often is issue occurring? : First occurrence Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Artifact and Noise Patient Impact/Outcome: none User Impact: Artifact Current Software Version: Not provided Diagnostic performed by Engineer : PRS erroring out. Advised customer PO required for tonight and weekend service. Customer to call back with PO. Internal Comments : Error # and/or description of error : No PRS Completed or Recommended Repair Action "Internal" : PRS erroring out. Advised customer PO required for tonight and weekend service. Customer to call back with PO. Follow Up Required : What are the next Troubleshooting steps : Review logs and Artifact. Error # and/or description of error : No PRS Completed or Recommended Repair Action "Internal" : PRS erroring out. Advised customer PO required for tonight and weekend service. Customer to call back with PO. ----------------------------------------------------------------------------- | 3 |  | Collect BD calibration data and images for investigation. Does the module replacement fix the issue？ | [12 empty] | 117066155 | In Progress | 629198 | TRUE | Closed - Done | Closed - Done |  | CTUDT00632110 | [5 empty] | Engineering | [6 empty] | Artifact | [8 empty] | System issue | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | DMS Module | DMS-HW | DMS-HW-DMS Module | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
458 |  |  | United States | Unity Health Systems | 40037 | 44244 | 8 | 0304CDY73 Gantry cannot initialize error. States FSE is already on the way. Tried to reboot the system and received... | Work Order #: WO-02996491 Lockwood Thomas Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- Work Order #: WO-02996504 Sheffield, Scott The system fully meets specification and returned to use Problem Description by Engineer : Problem description by engineer : Gantry cannot initialize error. Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No further information is available as per the customer Expected and Unexpected results: Gantry cannot initialize error. Patient Impact/Outcome: delay User Impact: delay Current Software Version: Not provided Diagnostic performed by Engineer : I downloaded logs and found couch errors below. Internal Comments : Error # and/or description of error : 2021/02/17 10:21:09.539 S_HMC_HOR_MOTOR_RELAY_OPEN ERROR_NORMAL Hor Motor Relay open 2021/02/17 10:21:09.539 S_HMC_HOR_MOTOR_RELAY_OPEN ERROR_NORMAL Hor Motor Relay open Completed or Recommended Repair Action "Internal" : The logs show that it has been 17 days since last power system power cycle. I had Judy power cycle the system. After power cycle I had her do a test scan of the table top using chest protocol and head protocol. The system scanned fine and is working normally at this time. I called FSE Tom and discussed log findings with him. Tom is going to site and will follow up with customer on this issue. System is working at this time. AEL Recommendations: Check / Replace the relay (K103). 2. Found cold solder joints inside of the relay (K103). 3. Found couch relay k103 that had cold solder joints on the wires for the contacts. Replaced k103 recommended part# 453566496941 - RELAY,24VDC PLUGIN W/COVER ----------------------------------------------------------------------------- | 2 |  | CT Gantry had not been powered down for 17 days. Instructed technologist and will instruct on procedure Rely improvement? Need check with Robin | [12 empty] | 117049348 | Closed | 619682 | TRUE | Closed - Done | Closed - Done |  | NotEngineering-PartReplacement | [5 empty] | Engineering | [6 empty] | Relay | [8 empty] | System issue | System | Couch-HW | Relay | Couch-HW-Relay | Couch-HW-Relay | NotEngineering-PartReplacement | Relay | Couch-HW | Couch-HW-Relay | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
459 |  |  | United States | Wellstar Kennestone Hospital | 68114 | 44237 |  | 0501CDN * Constancy Failing | Work Order #: WO-02955695 Carpenter, Kevin The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Customer Function/Role: HCF Dispatcher Roland on behalf of Beth Black Device Use: Diag Expected and Unexpected Results: Constancy Fails User Impact: None Patient Impact: None Current Software Version: IDK Follow Up Required : Beth Black says Constancy is Failing. Diagnostic performed by Engineer : None - call placed by HCF dispatcher ~ FSE requested ----------------------------------------------------------------------------- Work Order #: WO-02955727 11-Feb-2021 Bourne Barrett The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: HCF Dispatcher Roland on behalf of Beth Black Device Use: Diag Expected and Unexpected Results: Constancy Fails User Impact: None Patient Impact: None Current Software Version: IDK Diagnostic performed by Engineer : Fixed air bubble in phantom, leveled phantom holder and ran acceptance test, then constancy passed. ----------------------------------------------------------------------------- | 3 |  | Non-Refresh | [12 empty] | 117015523 | Closed | 600965 | #N/A | #N/A | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Non-Refresh | [5 empty] | Non-Refresh |  |  | Upgrade issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | y
460 |  |  | United States | Moore County Hospital District | 300033 | 44243 |  | 0302CDY70 Foot Board - lock doesn_t work/ Part ID | Work Order #: WO-02989901 Mark Finnigan The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Tech Device Use: No patient involved. Expected and Unexpected Results: Customer states he needs a new footrest extension. User Impact: None Patient Impact: Current Software Version: Follow Up Required : Spoke with Zane Pratt. Zane states that he needs a new footrest extension because the one that shipped with the system does not lock into the slot properly. His FSE asked that he place this call. The part number for the footrest is 453567134162 Diagnostic performed by Engineer : Spoke with Zane Pratt. Zane states that he needs a new footrest extension because the one that shipped with the system does not lock into the slot properly. His FSE asked that he place this call. The part number for the footrest is 453567134162 ----------------------------------------------------------------------------- Work Order #: WO-02989920 17-Feb-2021 Young Jerry The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Tech Device Use: No patient involved. Expected and Unexpected Results: Customer states he needs a new footrest extension. User Impact: None Patient Impact: Current Software Version: Resolution : Ordered replacement footboard. Follow Up Required : Order footboard. ----------------------------------------------------------------------------- Work Order #: WO-02995032 19-Feb-2021 Young Jerry The system fully meets specification and returned to use Resolution : Confirmed part delivery Problem Description by Engineer : Customer Function/Role: Tech Device Use: No patient involved. Expected and Unexpected Results: Customer states he needs a new footrest extension. User Impact: None Patient Impact: Current Software Version: Parts Consumed : 1.00000,453567134162,ASSEMBLY, FOOTREST EXTENSION ----------------------------------------------------------------------------- | 5 |  | Non-Refresh | [12 empty] | 117043830 | Closed | 616084 | #N/A | #N/A | #N/A | [7 empty] | Others | [6 empty] | Non-Refresh | [5 empty] | Non-Refresh |  |  | Upgrade issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | #N/A | y |  |  | y
461 |  |  | United States | Select Specialty Hosp-Tallahas | 300037 | 44238 |  | The patient maximum position scan range line on the couch pad is not properly aligned with the mark for the actual maxim | The patient maximum position scan range line on the couch pad is not properly aligned with the mark for the actual maximum scan range on the couch itself. Feedback type: Technical Issue Other Details: n/a Submitter Email: Melissa.lyter@philips.com CFM Received Date: 2/11/2021 | CFM |  | Silva follow with Jeff to conclude the next step action, whether need design change on pad scran line. | [14 empty] | 612763 | TRUE | Closed - Done | Closed - Done |  | CTUDT00631638 | [5 empty] | Engineering | [6 empty] | Scan Line | [5 empty] | Refresh |  |  | Upgrade issue | System | Couch-HW | Mattress scan line | Couch-HW-Mattress scan line | Couch-HW-Mattress scan line | CTUDT00631638 | Mattress scan line | Couch-HW | Couch-HW-Mattress scan line | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
462 |  |  | United States | Select Specialty Hosp-Tallahas | 300037 | 44238 |  | They have concerns about patient movement during initialization. | Customer believes the scanner initialization prior to performing a surview, and in between the dual surviews takes too long. They have concerns about patient movement during initialization. May of their patients at this facility are long-term care patients who are ventilated or cannot care for themselves. These types of patients have a hard time holding still and the long initialization times ~20 seems like a very long time when you have someone who has a difficult time holding still. They are comparing the initialization time to their GE scanner the customer previously had. Customer feels as though this is a flaw in the system. Feedback type: Enhancement Request Other Details: n/a Submitter Email: Melissa.lyter@philips.com CFM Received Date: 2/11/2021 | CFM |  | Initialization time Meet current design spec. need check if software update to HF2 which has improvement on the performance? What is the initialization time on 4.1.7, if possible share bugreport for engineering investigation | [14 empty] | 612741 | #N/A | #N/A | Closed - Done |  | TBD-Initilization | [5 empty] | Engineering | [6 empty] | Initialization time | [5 empty] | Refresh |  |  | Upgrade issue | System | Console-SW | Performance | Console-SW-Performance | Console-SW-Performance | TBD-Initilization | Performance | Console-SW | Console-SW-Performance | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
463 |  |  | United States | Select Specialty Hosp-Tallahas | 300037 | 44238 |  | The customer is requesting that we perform real-time surviews. | The customer is requesting that we perform real-time surviews. They have another vendor_s scanner (GE) that acquires real-time surview. They feel this is an important feature that our system is lacking. Feedback type: Enhancement Request Other Details: n/a Submitter Email: Melissa.lyter@philips.com CFM Received Date: 2/11/2021 | CFM |  | RFC_Real-time surview | [14 empty] | 612717 | #N/A | #N/A | Closed - Done |  | TBD-RFC | [5 empty] | Engineering | [6 empty] | RFC_Real-time surview | [5 empty] | Refresh |  |  | Upgrade issue | System | Console-SW | RealTimeSurview | Console-SW-RealTimeSurview | Console-SW-RealTimeSurview | RFC | RealTimeSurview | Console-SW | Console-SW-RealTimeSurview | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
464 |  |  | United States | Christ Hospital | 310034 | 44239 |  | nemo/ jeff | Work Order #: WO-02971894 Parts Consumed : 2.00000,453566503471,LIMIT SWITCH HORIZ IN ----------------------------------------------------------------------------- | 3 |  | Non-Refresh | [12 empty] | 117028902 | Closed | 610599 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Others | [6 empty] | Non-Refresh | [5 empty] | Non-Refresh |  |  | Upgrade issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | y
465 |  |  | United States | SSH-Pensacola | 310051 | 44235 |  | customer requesting support-needs to know what type of contrast and dose to use | Work Order #: WO-02942414 Mujica, David The system fully meets specification and returned to use Problem description by Clinical Support Specialist : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : Doncile was aided to put contrast usage on images and descriptor field using text boxes and saving to the scout image. Doncile had no further questions or need for instruction at this time. ----------------------------------------------------------------------------- | 5 |  | Non-Refresh | [12 empty] | 117003094 | Closed | 593825 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Others | [6 empty] | Non-Refresh | [5 empty] | Non-Refresh |  |  | Upgrade issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | y
466 |  |  | United States | St Jude Medical Center | 310165 | 44240 |  | 1201CDN84 Xray System Cannot Comply | Work Order #: WO-02974385 Holeman, David The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Problem description by engineer : Xray System Cannot Comply When was the first occurrence? Within... : 12-Feb-2021 17:00:00 How often is issue occurring? : First occurrence Malfunction area : Tube&Gen Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected results: Xray System Cannot Comply Patient Impact/Outcome: Delayed Scan User Impact: Contrast Scan Interrupted Current Software Version: 3.6.6.25204 Diagnostic performed by Engineer : Biomed (not at CT site) states that CT Tech reports "Xray System Cannot Comply" errors around 1715 on 2/12/21. Tech was able to complete the scan. Inquired if any subsequent scans = unknown. Techs are very busy and unavailable for remote service actions. Last gantry power-cycle = 23 days earlier. Inspected last filament calibrations: 04/09/2020 10:37 AM <DIR> Filament 12/29/2020 11:00 AM <DIR> Filament2 Robert Spidell, Biomed Phone (714) 309-4742 Email robert.spidell@ii-techknow.com authorized after-hours service and provided PO# JD-3643428 Internal Comments : Error # and/or description of error : 2021/02/12 17:16 S_GEN_FAULT_CATHODE_TUBE_ARC WARN 5:5 CPM Arc 2021/02/12 17:16 S_GEN_FAULT_ANODE_TUBE_ARC WARN 5:0 APM Arc 2021/02/12 17:16 S_GEN_FAULT_ANALOG_FAULT WARN 2:5 Analog Fault 2021/02/12 17:16 S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT NORMAL Completed or Recommended Repair Action "Internal" : Inspect, clean, re-grease HV connectors. If persistent arcing, perform no-load test with XRT disconnected. If no arcs, perform filament 1 and filament 2 calibrations. Perform long tube conditioning. Follow Up Required : Error # and/or description of error : 2021/02/12 17:16 S_GEN_FAULT_CATHODE_TUBE_ARC WARN 5:5 CPM Arc 2021/02/12 17:16 S_GEN_FAULT_ANODE_TUBE_ARC WARN 5:0 APM Arc 2021/02/12 17:16 S_GEN_FAULT_ANALOG_FAULT WARN 2:5 Analog Fault 2021/02/12 17:16 S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT NORMAL Completed or Recommended Repair Action "Internal" : Inspect, clean, re-grease HV connectors. If persistent arcing, perform no-load test with XRT disconnected. If no arcs, perform filament 1 and filament 2 calibrations. Perform long tube conditioning. ----------------------------------------------------------------------------- Work Order #: WO-02974429 13-Feb-2021 Buzzetta Salvatore The system fully meets specification and returned to use Problem Description by Engineer : Problem description by engineer : Xray System Cannot Comply When was the first occurrence? Within... : 12-Feb-2021 17:00:00 How often is issue occurring? : First occurrence Malfunction area : Tube&Gen Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected results: Xray System Cannot Comply Patient Impact/Outcome: Delayed Scan User Impact: Contrast Scan Interrupted Current Software Version: 3.6.6.25204 Resolution : long tube conditioning needed ----------------------------------------------------------------------------- | 3 |  | Non-Refresh | [12 empty] | 117030480 | Incomplete | 609306 | #N/A | #N/A | #N/A | [7 empty] | Others | [6 empty] | Non-Refresh | [5 empty] | Non-Refresh |  |  | Upgrade issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | #N/A | y |  |  | y
467 |  |  | United States | Lakeland Regional Medical Cent | 320474 | 44239 |  | MIP images are not retaining the correct increment after Service restores exam cards (ECs) or when the customer imports | MIP images are not retaining the correct increment after Service restores exam cards (ECs) or when the customer imports ECs with the exam card manager (ECM). After an ECM import or a system restore the EC increment will follow the system configuration options of contiguous or overlap instead of following what was previously set. The customer wants all results in the EC to remain as they had originally set. In this case the customer’s protocol is for the coronal and sagittal MIP images to be 10mm x 2mm. The axial MIP images should be 6mm x 2mm. After their iPatient S (v4.1.10) their coronal and sagittal MIPs were changed to 10mm x 5mm. The axial MIPs were changed to 6mm x 3mm. This is frustrating for the individuals responsible for maintaining exam cards. If the site is a busy site and they have a lot of ECs with MIP reconstructions this can be very time consuming. Feedback type: Enhancement request Other Details: other Submitter Email: melissa.lyter@philips.com CFM Received Date: 2/15/2021 | CFM |  | Send out to Silva Sridhar review | [14 empty] | 615486 | TRUE | Closed - Done | Closed - Done |  | CTUDT00632298 | [5 empty] | Engineering | [6 empty] | Exam Card | [5 empty] | Refresh |  |  | Upgrade issue | System | Console-SW | ExamCard | Console-SW-ExamCard | Console-SW-ExamCard | CTUDT00632298 | EC Upgrade | Console-SW | Console-SW-ExamCard | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
468 |  |  | United States | Lakeland Regional Medical Cent | 320474 | 44239 |  | when the customer_s exam cards were reinstalled the exam were listed in alphabetical order instead of retaining the cust | After iPatient S upgrade (v4.1.7  4.1.10), when the customer_s exam cards were reinstalled the exam were listed in alphabetical order instead of retaining the customer’s preferential order. They had to go in and reorder all the exam cards which depending on how many exam cards the customer has can take too much time. Especially at a busy facility where they cannot take the system down for an extended period in addition to the time it was already down for the upgrade/repair. The customer believes that after an upgrade or install of exam cards from a Service backup that the Exam Cards should be re-installed in their original order. Feedback type: Enhancement request Other Details: other Submitter Email: melissa.lyter@philips.com CFM Received Date: 2/15/2021 | CFM |  | Exam Card order CTUDT00632298 | [14 empty] | 615485 | TRUE | Closed - Done | Closed - Done |  | CTUDT00632455 | [5 empty] | Engineering | [6 empty] | Exam Card | [5 empty] | Refresh |  |  | Upgrade issue | System | Console-SW | ExamCard | Console-SW-ExamCard | Console-SW-ExamCard | CTUDT00632298 | EC Upgrade | Console-SW | Console-SW-ExamCard | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
469 |  |  | United States | Lakeland Regional Medical Cent | 320474 | 44238 |  | Q/O: Customer needs new Pads. | Work Order #: WO-02963288 17-Feb-2021 Watson Adam The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : Replaced worn out pads Parts Consumed : 6.00000,453567559721,Pad, Brilliance Foot Extension 1.00000,459800042241,Cushion-Head/ Armrest ----------------------------------------------------------------------------- | 5 |  | Non-Refresh | [12 empty] | 117021848 | Closed | 608055 | #N/A | #N/A | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Non-Refresh | [5 empty] | Non-Refresh |  |  | Upgrade issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | y
470 |  |  | United States | Lakeland Regional Medical Cent | 320474 | 44250 | 9 | 0711CFXXX Cannot Bootup Host Post Power-Cycle | Work Order #: WO-03029694 Holeman, David The system fully meets specification and returned to use Diagnostic performed by Engineer : CT Tech reports she performed a routine Host shutdown and upon reboot a Windows blue screen with "change user" is displayed and is not booting up into CT application. Inquired if any other messages or errors displayed = no. While on phone with Tech, the Philips CT application did launch and Tech is now operating normally. Host = HP z440 v4.1.10.25208 Problem Description by Engineer : Problem description by engineer : Cannot Bootup Host Post Power-Cycle When was the first occurrence? Within... : 23-Feb-2021 08:00:00 How often is issue occurring? : First occurrence Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: Cannot Bootup Host Post Power-Cycle Patient Impact/Outcome: None User Impact: Cannot Bootup Host Post Power-Cycle Current Software Version: 4.1.7.25026 Internal Comments : Error # and/or description of error : Cannot Bootup Host Post Power-Cycle Resolution : After Host shutdown/restart, Bootup and CT application launch are working successfully. ----------------------------------------------------------------------------- | 2 |  | Did FSE/CUSTOMER save bugreprot at that moment? Any force shutdown on this system? Tammy share the log path to support download from PRS | [12 empty] | 117080455 | Closed | 634733 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-PowerOff | [5 empty] | Engineering | [6 empty] | Host Power Off | [5 empty] | Refresh |  |  | Upgrade issue | System | Gantry-HW | PowerOff | Gantry-HW-PowerOff | Gantry-HW-PowerOff | TBD-PowerOff | PowerOff | Gantry-HW | Gantry-HW-PowerOff | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
471 |  |  | Chile | CMD Clinica Integrales Rancagu | 40008 | 44251 |  | CL1675-NO SE PUEDE INICIAR EL GANTRY. | "Work Order #: WO-03041283 Valenzuela Velasquez, Eduardo Ariel The system does not meet specification and is out of use Problem Description by Engineer: Gantry error, does not allow studies. Diagnostic performed by Engineer: Communication error occurs in gantry (control-link). Review of slipring and brushes is needed. Follow Up Required: Cleaning and lubrication of slipring. Cleaning brushes. Review with the client how they disinfect the room for Covid-19. -------------------------------------------------- --------------------------- Work Order #: WO-03044188 25-Feb-2021 Osorio, Juan The system fully meets specification and returned to use Problem Description by Engineer: Gantry error, does not allow studies. Resolution: Slipring (signal tracks) and brush cleaning, which show obvious traces of dirt. Bump tests run smoothly. Diagnostic performed by Engineer: Diagnostic performed remotely indicates that the equipment has errors in its data transfer line (CRC Errors) Test & Verification: Test and verifications OK -------------------------------------------------- --------------------------- | 3 |  | Slipring (signal tracks) and brush cleaning, which show obvious traces of dirt. | [12 empty] | 117090360 | Closed | 644627 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | NotEngineering-SlipringClean | [5 empty] | Engineering | [6 empty] | Slip ring cleaning | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | SlipRingClean | Gantry-HW-SlipRingClean | Gantry-HW-SlipRingClean | NotEngineering-SlipringClean | SlipRingClean | Gantry-HW | Gantry-HW-SlipRingClean | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
472 |  |  | China | JS Nanjing M.U.A Yifu Hosp. | 345028 | 44253 |  | 02N/Ingenuity Core128/Add node | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Connection/Cable/Interconnect Repair Activity Code: Instructed Customer Problem Reported by Customer : Need add network node for deepwise Safety Questions : Case: 0117099040 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Cao Last updated by: Liu Cao 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: Customer need to add node Troubleshooting Action: Cannot open LAN config, due to IPV6 enable. Instruct dis-plug network cable and add node Repair Action: Instruct customer use LAN config Test/Inspection: N/A Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 5 |  | Defect_LAN Config implement FCO740 to fix this issue | [12 empty] | 117099040 | Closed | 651786 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00631055 | [5 empty] | Engineering | [6 empty] | LAN Config | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Lan Config | Console-SW-Lan Config | Console-SW-Lan Config | CTUDT00631055 | Lan Config | Console-SW | Console-SW-Lan Config | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
473 |  |  | China | Suzhou Guangci Tumor Hospital | 345019 | 44252 |  | 02N/Ingenuity Core128/The scan is interrupted | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : The scan is interrupted Safety Questions : Case: 0117092408 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingsong Last updated by: Lu Mingsong 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 扫描中断 客户问题详细描述或故障现象： 扫描软件运行卡顿 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： N/A 远程诊断与测试结果： N/A PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 检查扫描软件系统，必要时重装软件系统. 建议备件12NC及名称： N/A 客户需求备注： 无 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 26.02.2021 04:50:38 UTC Haotian Xu (86090280) 服务内容： 扫描软件运行卡顿 ，修改参数，清理主机内存 服务结果： 维修后，设备工作正常。 Internal Remarks : 26.02.2021 11:40:32 UTC Haotian Xu (86090280) $Q7: (confirmed reported problem，The problem is the software stuck and the software was performed slowly） $Q8: (check the software and the Protocol parameters and found the parameters need to be changed) $Q9: （change the Protocol parameters and delete the junk files in the computer ) $Q10: Test and inspection data refer to PA. Device returned to full functionality. $Q11: Calibrated Tools refer to PA ----------------------------------------------------------------------------- | 3 |  | 扫描软件运行卡顿 ，修改参数，清理主机内存 Did customer restart daily? Which parameter cause slowly? Any bugreprot saved? | [12 empty] | 117092408 | Closed | 652031 | TRUE | Closed - Done | Closed - Done |  | TBD-SW-Slow | [5 empty] | Engineering | [6 empty] | Software slow | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Performance | Console-SW-Performance | Console-SW-Performance | TBD-SW-Slow | Performance | Console-SW | Console-SW-Performance | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
474 |  |  | China | SD Chest Hosp. | 40025 | 44251 |  | WeChat:02N/Ingenuity CT 728326/Image quality is poor | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : Image quality is poor Safety Questions : Case: 0117084845 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Bai Jingwen Last updated by: Bai Jingwen 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 图像质量不好，需要调节 客户问题详细描述或故障现象： 胸部薄层扫描，图像不清晰 设备本身技术问题： 是。 Troubleshooting Action: 远程诊断与测试结果： 联系工程师黄兆坤，准备和培训老师引起到场检查 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议工程师到场检查球管位置，焦点，膜体校准。。。 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support ----------------------------------------------------------------------------- | 5 |  | 建议工程师到场检查球管位置，焦点，膜体校准。。。 | [12 empty] | 117084845 | In Process | 643555 | TRUE | Closed - Done | Closed - Done |  | TBD-SW-IQ | [5 empty] | Application | [6 empty] | Poor IQ | [5 empty] | Refresh |  |  | System issue | System | DMS-SW | IQ | DMS-SW-IQ | DMS-SW-IQ | TBD-SW-IQ | IQ | DMS-SW | DMS-SW-IQ | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
475 |  |  | China | HEB Caofeidian Dist. Hosp. | 345016 | 44249 |  | 02N/Ingenuity Core128/Elite China/software error | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Install Update Problem Reported by Customer:Software error Safety Questions:Case: 0117070390 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Chen YuanHui Last updated by: Chen YuanHui 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action: FSE Hu Yue 013703597273 call in and dispatch him T2 Activities:Problem Description: Summary of customer_s problems: FCO is going to go to the field to open a ticket. Detailed description of Customer_s problems or malfunctions: FCO is going to go to the field to open a ticket. It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: NA Remote troubleshoot and test result: FCO is on site to open ticket and record PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Perform FCO, generate order and record Test/Inspection: N/A Test Equipment: N/A SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support External Remarks:23.02.2021 02:42:53 UTC Yue Hu (86038448) Service details (required): Image transfer issue. Software patching Service result (required): after the maintenance, the device returns to normal working condition Internal Remarks:23.02.2021 02:43:11 UTC Yue Hu (86038448) $Q7: (Confirmed reported problem: software error. $Q8: (software error, need install update $Q9: (install update, system return to service $Q10: (Test and inspection data refer to: PA. Device returned to full service) $Q11: Calibrated Tools name (FLUKE287) S/N (20060015) next calibration data (05/17/2021) ----------------------------------------------------------------------------- | 3 |  | FCO740 | [12 empty] | 117070390 | In Process | 639186 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
476 |  |  | Japan | KYOTO MEDICAL CENTER | 341013 | 44256 |  | CIRSの通信エラーでスキャン進まず。Goのクリック３回目で撮影が走った>> | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : CIRSの通信エラーでスキャン進まず。Goのクリック３回目で撮影が走った>> ----------------------------------------------------------------------------- The scan does not proceed due to a CIRS communication error. Shooting ran with the third click of Go >> | 4 |  | An error occurred when pressing Go (transition to scan screen at the third time) [Phenomenon confirmation] I was able to confirm. [Details of work] Since there was a difference between the planned current value and the measured value of the tube current, an error occurred at that time. As a response, we adjusted the filament. | [12 empty] | 117109700 | In Process | 654411 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Scan | [5 empty] | Refresh |  |  | System issue | System | CIRS-SW | TBD-NoEnoughInfo | CIRS-SW-TBD-NoEnoughInfo | CIRS-SW-TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo | CIRS-SW | CIRS-SW-TBD-NoEnoughInfo | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
477 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 341006 | 44250 |  | Problem with sending images to PACS. | Work Order #: WO-03029009 Soner Saglam The system fully meets specification and returned to use Problem Description by Engineer : Problem with sending images to PACS. Follow Up Required : Complete FCO72800740 with site visit. Resolution : Complete FCO72800740 with site visit. ----------------------------------------------------------------------------- | 3 |  | FCO740 | [12 empty] | 117079847 | In Progress | 633972 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
478 |  |  | United States | McLaren Macomb | 40021 | 44255 | 9 | 0903CDY72 Ring Artifacts | Work Order #: WO-03061629 Boyd Craig The system does not meet specification and is out of use Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: Ring Artifacts User Impact: None Patient Impact: None Current Software Version: Diagnostic performed by Engineer : FSE, Tony Mona and Steven Mauri, on site and said a Detector module is needed to repair system. FSE requesting case be opened. Follow Up Required : FSE on site to evaluate artifact issue. Please page to FSE: Steven Mauri. Manager of site Mike Adams is approving service this weekend. ----------------------------------------------------------------------------- Moved detector module 24 to 40 and calibrated system, performed bad detectors, phantom, air, constancy, and IQ, system is with in spec. | 2 |  | Detector module swapped, but suggest to swap within Class A group which allow BD detector rule (slots 9-10 or 33-34), Does issue solved by module swap or replaced? If replace, please share the serial number of defective module. | [12 empty] | 117107731 | Fixed | 653857 | TRUE | Closed - Done | Closed - Done |  | CTUDT00632110 | [5 empty] | Engineering | [6 empty] | Artifact | [5 empty] | Refresh |  |  | System issue | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | DMS Module | DMS-HW | DMS-HW-DMS Module | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
479 |  |  | United States | Unity Health Systems | 40037 | 44251 | 9 | Gantry can not comply error due to tube arcing; Ran filament calibration and long tube conditioning to correct. | Work Order #: WO-03037907 24-Feb-2021 Taylor, Durrance The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Device Use: System Not in Use when Call was placed. Down. Expected and Unexpected Results: Tube arcing. Expected of new tubes. More conditioning was perfromed User Impact: Patient Impact: Current Software Version: 4.1.10 Diagnostic performed by Engineer : Performed Filament cals and then ran Long Tube Conditioning. Zero arcs. Ran Filament cals after tube conditioning and then air cals with zero arcs. Then I did a scan with the same exam card parameters that caused an an arc in the past and had no issues. Ran Acceptance and IQ check and everything passed successfully. New tube needed more conditioning | 2 |  | Performed Filament cals and then ran Long Tube Conditioning. Zero arcs. Keep monitoring, get bugreport for analysis if possible | [12 empty] | 117087824 | In Progress | 642468 | TRUE | Closed - Done | Closed - Done |  | CTUDT00630085 | [5 empty] | Engineering | [6 empty] | Tube Arc | [5 empty] | Refresh |  |  | System issue | System | IC-HW | Tub Arc | IC-HW-Tub Arc | IC-HW-Tub Arc | CTUDT00630085 | Tub Arc | IC-HW | IC-HW-Tub Arc | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
480 |  |  | Chile | CMD Clinica Integrales Rancagu | 40008 | 44263 |  | Revisar PRS y deshabilitar auto-delete | Work Order #: WO-03107947 Caldelas Aldana, Carlos Manuel The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: usuarfio Device Use: fuera de uso Expected and Unexpected Results: prs no funciona correctamente y no se puede deshabilitar opcion de autoborrado User Impact: problemas con la base de datos Patient Impact: n/a Current Software Version:4.1.10 Diagnostic performed by Engineer : derivado de otro soporte, se confirmó tienen problemas con la opcion de autodelete y es necesario acudir a desactivarla desde modo servicio. Follow Up Required : Revisar porque no se puede conectar en remote console, validar firewal exceptions, hacer pruebas funcionales para lograr hacer remote desktop conectio via RSN/M2M. Deshabilitar auto-delete en directorio para que se encarguen ellso de mantener manualmente base de datos con al menos 15% libre ----------------------------------------------------------------------------- | 5 |  | PRS Connection | [12 empty] | 117151500 | In Progress | 678517 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Service | [6 empty] | PRS Setup | [5 empty] | Refresh |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
481 |  |  | Chile | CMD Clinica Integrales Rancagu | 40008 | 44260 |  | CL1675 FALLA EN FUNCIONAMIENTO DEL EQUIPO AUN CON EL 70% DE MEMORIA | "Work Order #: WO-03099317 Caldelas Aldana, Carlos Manuel The system fully meets specification and returned to use Problem Description by Engineer: Customer Function / Role: user Device Use: diagnosis Expected and Unexpected Results: Database problem User Impact: deleting studies Patient Impact: n / a Current Software Version: 4.1.10 Diagnostic performed by Engineer: validated had problems with the derived database that had little space and the active auto-delete option, so it deleted some studies and was still there without showing others. Remote inspection was made confirming damage to the dicom spooler, it was successfully reset via remote, restoring the dispatch queue and local directory functionality. Resolution: it is recommended to always leave at least 15% free of local disk Follow Up Required: Check PRS and disable auto-delete | 4 |  | check if implemnent HF3, if yes, share bugreport and pic for investigate | [12 empty] | 117142748 | In Progress | 676435 | TRUE | Closed - Done | Closed - Done |  | TBD-Bugrep | [5 empty] | Engineering | [6 empty] | Database problem | [5 empty] | Refresh |  |  | System issue | System | TBD-Bugreport | TBD_Bugreport | TBD-Translation | TBD-Translation | TBD-Bugrep | TBD-Translation | TBD-Bugreport | TBD-Translation | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
482 |  |  | China | SD Chest Hosp. | 40025 | 44263 |  | 02N/Ingenuity CT 728326/application stuck | 主机蓝屏 客户问题详细描述或故障现象： 主机蓝屏 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 主机蓝屏，重启后报系统错误 远程诊断与测试结果： 主机蓝屏，装软件 PRS TR状态：全通 Repair Action: 建议现场维修步骤： 1，做磁盘整理 2，清理系统缓存 3，重新安装系统 4，收集bugreport 提交esclation 和CFM Test/Inspection: N/A | 3 |  | 执行了HF3 FCO, patient data missing 客户就强制关机了，起来进入黄屏 工程师说用ist进到维修模式启动软件，再logout，用客户账号登录就好了 | [12 empty] | 117149926 | In Process | 678034 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00631367 | [5 empty] | Engineering | [6 empty] | New Yellow + Blue screen | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
483 |  |  | China | Huaying Mountain Guangneng General Hospital | 345000 | 44261 |  | 02N/Ingenuity Core128/Elite China/gantry power off | N/A T2 Activities : Problem Description: 客户问题概述： 其他类型故障。 客户问题详细描述或故障现象： 机架掉电 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： S_COUCHMGR_HOR_MOTION_STOP_ERROR S_MANMOT_HORIZONTAL_STOP_ERROR 远程诊断与测试结果： 机架掉电，说前几天有relayopen的故障，床的水平运动故障 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 检查床的水平运动部件，根据床的诊断手册作相应troubleshooting 建议备件12NC及名称： Relay, 24VDC plugin with cover 453566496941//BRILLIANCE COUCH CNTL PCB ASSY 459800477572 Next Step Action : PPD 459800477572 1 & 453566496941 1 已安排手工单24H到医院，FSE李全飞-86-13699457740 Next Step Action : 李全飞 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support ----------------------------------------------------------------------------- | 3 |  | 客户问题概述： 其他类型故障。 客户问题详细描述或故障现象： 机架掉电 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： S_COUCHMGR_HOR_MOTION_STOP_ERROR S_MANMOT_HORIZONTAL_STOP_ERROR 远程诊断与测试结果： 机架掉电，说前几天有relayopen的故障，床的水平运动故障 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 检查床的水平运动部件，根据床的诊断手册作相应troubleshooting 建议备件12NC及名称： Relay, 24VDC plugin with cover 453566496941//BRILLIANCE COUCH CNTL PCB ASSY 459800477572 | [12 empty] | 117145330 | In Process | 675205 | TRUE | Closed - Done | Closed - Done |  | TBD-Translation | [5 empty] | Engineering | [6 empty] | CCB | [5 empty] | Refresh |  |  | System issue | System | TBD-Bugreport | TBD_Bugreport | TBD-Translation | TBD-Translation | TBD-Translation | TBD-Translation | TBD-Bugreport | TBD-Translation | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
484 |  |  | China | HLJ Harbin Senhai Hospital | 345007 | 44260 | 10 | 02N/Ingenuity Core128/Elite China/artifacts | Next Step Action : 客户还需要到场应用培训，已电话+邮件联系杨俊 13817647801 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support T2 Activities : Problem Description: 客户问题概述： 图像伪影 客户问题详细描述或故障现象： 图像伪影 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： NA 远程诊断与测试结果： 环形伪影，联系工程师可能探测器通道有问题 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 1.分析原始数据和dicom确认问题 2. 重启机架看模块是否恢复 3.屏蔽通道或者更换模块等 建议备件12NC及名称： 459800911281 TDMS V2 M.Det Class B FRU T2 Activities : Problem Description: 客户问题概述： 图像伪影 客户问题详细描述或故障现象： 图像伪影 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： NA 远程诊断与测试结果： 环形伪影已经出现一段时间了，联系工程师可能探测器通道有问题 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 1.分析原始数据和dicom确认问题 2. 重启机架看模块是否恢复 3.屏蔽通道或者更换模块 4.做图像质量校准等 建议备件12NC及名称： 459800911281 TDMS V2 M.Det Class B FRU Next Step Action : PPD 459800911281*1 倪刚-86-18686766464 已安排24H到医院 External Remarks : 06.03.2021 02:49:03 UTC Gang Ni (86038229)"SUAT. Ready for Clinical Use local time" : "2021-03-06 10:30" 服务内容：屏蔽坏通道，空气校正 服务结果：设备恢复 Internal Remarks : 06.03.2021 10:21:49 UTC Gang Ni (86038229) $Q7: (Confirmed reported problem & artifacts) $Q8: (on site review image; check rawdata; bad detector & found bad detector) $Q9: (add bad detector; air calibration) $Q10: (Test and inspection data refer to: PA. & Device returned to full functionality) $Q11: Calibrated Tools name：none ----------------------------------------------------------------------------- | 2 |  | DMS Module not replace, module 13 BD calibrated | [12 empty] | 117137960 | Closed | 674710 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632110 | [5 empty] | Engineering | [6 empty] | Artifact | [5 empty] | Refresh |  |  | System issue | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | DMS Module | DMS-HW | DMS-HW-DMS Module | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
485 |  |  | China | GX Tengxian People_s Hospital | 40036 | 44259 |  | 02N/Ingenuity CT 728326/Image fuzzy | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer : Image fuzzy Safety Questions : Case: 0117131137 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Bai Jingwen Last updated by: Bai Jingwen 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: PACS client get fuzzy image Troubleshooting Action: PACS client get fuzzy image Repair Action: Customer may consult PACS engineer Test/Inspection: N/A Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | PACS vendor issue | [12 empty] | 117131137 | Closed | 670997 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | PACS Vendor | [5 empty] | Refresh |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
486 |  |  | China | JS No.2 Chinese Medicine Hosp. | 345027 | 44259 |  | Wechat:02n/Ingenuity Core128/Elite China/Scan the interrupt | T2 Activities : Problem Description: 客户问题概述： 扫描中断 客户问题详细描述或故障现象： 扫描中断 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： X-ray system cannot comply。 远程诊断与测试结果： N/A PRS TR状态： 不通，请工程师到场检查。 Repair Action: 建议现场维修步骤： 检查高压系统. 建议备件12NC及名称： N/A 客户需求备注： 无 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Internal Remarks : 08.03.2021 11:27:09 UTC Liangdong Qi (86036024) RDI Accuracy: Yes $Q7: (Confirmed reported problem, scan can not be performed.) $Q8: (Analysis error report, found Heating Unit was over high when scan can not be performed.) $Q9: (Changed scan protocol, redused exposure dose.) $Q10: (Test and Inspection data refer to PA & Device return to fully function) $Q11: Calibrated Tools (NA). External Remarks : 08.03.2021 11:27:06 UTC Liangdong Qi (86036024)服务内容：调整扫描协议. 服务结果： 维修后，设备恢复正常工作. ----------------------------------------------------------------------------- | 3 |  | Heating Unit was over high when scan can not be performed.Changed scan protocol, redused exposure dose. | [12 empty] | 117130861 | Closed | 670845 | TRUE | Closed - Done | Closed - Done |  | TBD_Tube hot_bugreport | [5 empty] | Engineering | [6 empty] | Tube heat | [5 empty] | Refresh |  |  | System issue | System | Gantry-SW | TubHot | Gantry-SW-TubHot | Gantry-SW-TubHot | TBD_Tube hot_bugreport | TubHot | Gantry-SW | Gantry-SW-TubHot | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
487 |  |  | China | SD Chest Hosp. | 40025 | 44258 |  | 02N/Ingenuity CT 728326/optical driver problem | 光驱故障 客户问题详细描述或故障现象： 光驱故障 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： NA 远程诊断与测试结果： 工程师反馈主机光驱弹不出来，需要两个光驱 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 检查光驱和供电；更换光驱 建议备件12NC及名称： 459801631381 HPE HP 9.5mm Slim SuperMulti DVD Writer | 3 |  | 光驱更换 | [12 empty] | 117125211 | In Process | 667327 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-DVDDrive | [5 empty] | Engineering | [6 empty] | DVD Drive | [5 empty] | Refresh |  |  | System issue | System | Console-HW | DVDDrive | Console-HW-DVDDrive | Console-HW-DVDDrive | TBD-DVDDrive | DVDDrive | Console-HW | Console-HW-DVDDrive | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
488 |  |  | China | JL Songyuan Central Hosptial | 345008 | 44258 |  | 02N/Ingenuity Core128/Elite China/system hung up | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : system hung up SAP Transfer : Case Transferred to SAP T2 Activities : Problem Description: 客户问题概述： ESTOP断开。 客户问题详细描述或故障现象： 间歇性estop open 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： "S_MC_DEVICE_HOR_ENCODERS_NOT_SYNCED " 远程诊断与测试结果： NONE PRS TR状态： 通。 Repair Action: 建议现场维修步骤： 1.检查编码器连接线。2.更换编码器，运行水平校正 建议备件12NC及名称： 453566500911 ENCODER CT ABSOLUTE Next Step Action : PPD备件ENCODER CT ABSOLUTE 453566500911*1已安排24H到场地,FSE：迟晓东-86-18943690010 External Remarks : 04.03.2021 11:50:01 UTC Tian Hua (86038446)服务内容： 更换床水平绝对编码器，校准，打火，处理高压接头，校准灯丝，长训管。 服务结果： 维修后，设备恢复正常工作。 其它重要事项： 无。 Internal Remarks : 04.03.2021 11:50:04 UTC Tian Hua (86038446) RDI Accuracy ：No $Q7: (Confirmed reported problem&Scan failed) $Q8: (Review logger and check the Couch and gantry，The Couch ABSOLUTE ENCODER failed ，HV ARC） $Q9: (Replaced the ABSOLUTE ENCODER and Clean HV cable ，calibration） $Q10: (Test and inspection data refer to: PA. Device returned to full functionality) $Q11 Calibrated Tools name (FLUKE287C) S/N (27420003) Next Calibration date (03/31/2021) Safety Questions : Case: 0117124149 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Qian Wang Last updated by: Qian Wang 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Dispatched for Onsite Support ----------------------------------------------------------------------------- | 3 |  | 更换床水平绝对编码器，校准，打火，处理高压接头，校准灯丝，长训管。 | [12 empty] | 117124149 | Closed | 666409 | TRUE | Closed - Done | Closed - Done |  | TBD-Translation | [5 empty] | Engineering | [6 empty] | Hori. Encoder | [5 empty] | Refresh |  |  | System issue | System | TBD-Bugreport | TBD_Bugreport | TBD-Translation | TBD-Translation | TBD-Translation | TBD-Translation | TBD-Bugreport | TBD-Translation | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
489 |  |  | China | JL Songyuan Central Hosptial | 345008 | 44256 |  | 02N/Ingenuity Core128/Elite China/Cannot transfer image | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: cannot transfer image Next Step Action: 13756701686 call in, duplicate maintenance call Safety Questions:Case: 0117108425 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Yu Zhao Last updated by: Yu Zhao 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Field Service Engineer Notes:This PO is a Reference Number only SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support T2 Activities:Problem Description: Summary of customer_s problems: host Image transfer failure Detailed description of Customer_s problems or malfunctions: host Image transfer failure It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: NA Remote troubleshoot and test result: Customer feedback, host image transfer to PACS occasionally failed, newly installed machine, customer hope to resolve onsite. PRS TR status: It is not through, ask engineer for maintenance on site. Repair Action: Suggest onsite service procedures: Reset spooler, check hospital network, check node setup. Suggest to have customer make an appointment with PACS engineer to go to the field to troubleshoot External Remarks: 02.03.2021 02:03:37 UTC Tian Hua (86038446) Service details: Transmission to PACS failed, reboot, restore. After doing FCO72800740, the problem persisted. Service outcome: device returns to normal working condition after service Other significant matters: None. Internal Remarks:02.03.2021 02:03:39 UTC Tian Hua (86038446) RDI Accuracy ：No $Q7: (Confirmed reported problem: Cannot transfer image to PACS) $Q8: (Review logger and check the HOST computer, The HOST computer transfer image to PACS failed) $Q9: (Done the fco72800740 and Escalation) $Q10: (Test and inspection data refer to: PA. Device returned to full functionality) $Q11 Calibrated Tools name (FLUKE287C) S/N (27420003) Next Calibration date (03/31/2021) ----------------------------------------------------------------------------- | 5 |  | PACS New, CTUDT00632520 | [12 empty] | 117108425 | In Process | 661517 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS New | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
490 |  |  | China | JX Yugan Chudong hospital | 337500 | 44256 |  | Recognizable image ring artifacts, perfrom URH align, UHR DFS calibration and Air calibration. | Problem Reported by Customer:image has artifacts Field Service Engineer Notes:This PO is a Reference Number only SAP Transfer:Case Transferred to SAP T2 Activities:Problem Description: Summary of customer_s problems: Image artifact. Detailed description of Customer_s problems or malfunctions: There are artifacts to some of the later images of the abdomen scan sequence It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: N/A Remote troubleshoot and test result: N/A PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Image with artifacts has MAs display too low. Perform raw data analysis of images with artifacts. Check image with artifacts HV auto adjustment and Aplane position parameters during exposure. We recommend spare parts 12NC and names: N/A Memo of Customer request: None SAP Transfer:Dispatched for On site Support External Remarks: 01.03.2021 15:31:18 UTC SUP JCOUSER (86038347) Service details: CT image has small thin annular artifacts. Perform ultra high calibration; CT image has faint ring artifacts, perform air calibration; CT images have thick, grainy artifacts, positioning images are incomplete which cause mA adjustment improper. Suggest customer to perform proper positioning scan and planned scan. service result: device returns to normal working condition after service Internal Remarks: 03.03.2021 01:55:49 UTC Xuehui Yan (86038347) RDI Accuracy:Yes $Q7: Confirmed reported problem Image has artifacts. $Q8: Check the UHR alignment, found 0.02mm misalignment; Check raw data, found some detector output normal; Check the ma and surview, helical scan‘s length. found the length of surview was less than helical, so the ma of segment that not surviewed was very low. $Q9: Perform URH align & UHR DFS calibration & air calibration & tell customer to do complete surview before plan scan. $Q10: Test and inspection data refer to: PA. Device returned to full functionality. $Q11: Calibrated Tools name (N/A) S/N (N/A) Next Calibration date (N/A). ^Current Tube installed date: MM/DD/ YYYY. ^Current Tube usage: scan second_N/A. External Remarks : 03.03.2021 08:23:28 UTC Jiawei Li (86091272)UDI: (01)00884838059863(21)337500 01.03.2021 15:31:18 UTC SUP JCOUSER (86038347) Service Details: The CT image has small, thin ring artifacts. We performed the ultra high calibration; the CT image has faint ring artifacts. We performed an air calibration; the CT images have thick, grainy artifacts and the positioning images are incomplete, which causes the mA adjustment not be properly modulated. It is recommended to have the customer perform the proper positioning scan and the planned scan. service result: device returns to normal working condition after service Internal Remarks : 03.03.2021 08:23:34 UTC Jiawei Li (86091272)$Q7: Confirmed reported problem Image has artifacts. $Q8: Check the UHR alignment，found 0.02mm misalignment；Check raw data，found some detector output normal；Check the ma and surview、helical scan‘s length，found the length of surview was less than helical，so the ma of segment that not surviewed was very low. $Q9: Perfrom URH align & UHR DFS calibration & air calibration & tell customer to do complete surview before plan scan. $Q10: Test and inspection data refer to: PA. The device returned to full functionality. $Q11: Calibrated Tools name (N/A) S/N (N/A) Next Calibration date (N/A). ^Current Tube installed date: MM/DD/ YYYY. ^Current Tube usage: scan second_N/A. Internal Remarks: 03.03.2021 13:30:55 UTC Jiawei Li (86091272)RDI Accuracy:Yes $Q7: Confirmed reported problem Image has artifacts. $Q8: Check the UHR alignment，found 0.02mm misalignment；Check raw data，found some detector output normal；Check the ma and surview、helical scan‘s length，found the length of surview was less than helical，so the ma of segment that not surviewed was very low. $Q9: Perfrom URH align & UHR DFS calibration & air calibration & tell customer to do complete surview before plan scan. $Q10: Test and inspection data refer to: PA. The device returned to full functionality. $Q11: Calibrated Tools name (N/A) S/N (N/A) Next Calibration date (N/A). ^Current Tube installed date: MM/DD/ YYYY. ^Current Tube usage: scan second_N/A. 03.03.2021 08:23:34 UTC Jiawei Li (86091272)$Q7: Confirmed reported problem Image has artifacts. $Q8: Check the UHR alignment，found 0.02mm misalignment；Check raw data，found some detector output normal；Check the ma and surview、helical scan‘s length，found the length of surview was less than helical，so the ma of segment that not surviewed was very low. $Q9: Perfrom URH align & UHR DFS calibration & air calibration & tell customer to do complete surview before plan scan. $Q10: Test and inspection data refer to: PA. The device returned to full functionality. $Q11: Calibrated Tools name (N/A) S/N (N/A) Next Calibration date (N/A). ^Current Tube installed date: MM/DD/ YYYY. ^Current Tube usage: scan second_N/A. External Remarks : 03.03.2021 13:30:52 UTC Jiawei Li (86091272)UDI: (01)00884838059863(21)337500 01.03.2021 15:31:18 UTC SUP JCOUSER (86038347) Service Details: The CT image has small, thin ring artifacts. We performed the ultra high calibration; the CT image has faint ring artifacts. We performed an air calibration; the CT images have thick, grainy artifacts and the positioning images are incomplete, which causes the mA adjustment not be properly modulated. It is recommended to have the customer perform the proper positioning scan and the planned scan. service result: device returns to normal working condition after service 03.03.2021 08:23:28 UTC Jiawei Li (86091272)UDI: ( 01) 00884838059863 (21) 337500 01.03.2021 15:31:18 UTC SUP JCOUSER (86038347) Service Details: The CT image has small, thin ring artifacts. We performed the ultra high calibration; the CT image has faint ring artifacts. We performed an air calibration; the CT images have thick, grainy artifacts and the positioning images are incomplete, which causes the mA adjustment not be properly modulated. It is recommended to have the customer perform the proper positioning scan and the planned scan. service result: device returns to normal working condition after service ----------------------------------------------------------------------------- | 3 |  | Check the UHR alignment, found 0.02mm misalignment; Check raw data, found some detector output normal; Check the ma and surview, helical scan‘s length. found the length of surview was less than helical, so the ma of segment that not surviewed was very low. $Q9: Perform URH align & UHR DFS calibration & air calibration & tell customer to do complete surview before plan scan. | [12 empty] | 117108316 | Closed | 659607 | TRUE | Closed - Done | Closed - Done |  | TBD-UHR | [5 empty] | Engineering | [6 empty] | UHR | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | UHR | Gantry-HW-UHR | Gantry-HW-UHR | TBD-UHR | UHR | Gantry-HW | Gantry-HW-UHR | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
491 |  |  | China | JL Songyuan Central Hosptial | 345008 | 44256 |  | 02N/Ingenuity Core128/Elite China/Unable to start | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: Unable to start Safety Questions:Case: 0117108432 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Yu Zhao Last updated by: Yu Zhao 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Field Service Engineer Notes:This PO is a Reference Number only T2 Activities:Problem Description: Summary of customer_s problems: The machine occasionally loses power. Detailed description of Customer_s problems or malfunctions: The machine occasionally loses power. It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: NA Remote troubleshoot and test result: The customer said the machine frequently loses power. Shut down can restore. Newly installed machine. The customer needs an engineer to go to the field as soon as possible.PRS FTP and TELNET are not working. Need an engineer to go to the site to check the settings according to the manual PRS TR status: It is not through, ask engineer for maintenance on site. Repair Action: Suggest onsite service procedures: Troubleshooting on the site based on error log Memo of Customer request: Customer is anxious, need engineer to contact customer ASAP SAP Transfer:Dispatched for On site Support SAP Transfer:Case Transferred to SAP External Remarks: 02.03.2021 04:38:33 UTC Tian Hua (86038446) Service details: Scan interruption, bed encoder synchronizing failure. Calibrate horizontally. Signal transmission interruption. Clean signal brush. Service outcome: device returns to normal working condition after service Other significant matters: None. Internal Remarks:02.03.2021 04:38:35 UTC Tian Hua (86038446) RDI Accuracy ：No $Q7: (Confirmed reported problem&Scan failed) $Q8: (Review logger and check the Couch, The gantry signal failed) $Q9: (Calibration the couch and clean the gantry) $Q10: (Test and inspection data refer to: PA. Device returned to full functionality) $Q11 Calibrated Tools name (FLUKE287C) S/N (27420003) Next Calibration date (03/31/2021) ----------------------------------------------------------------------------- | 3 |  | Scan interruption, bed encoder synchronizing failure. Calibrate horizontally. Signal transmission interruption. Clean signal brush. Service outcome: device returns to normal working condition after service | [12 empty] | 117108432 | Closed | 657059 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-Encoder | [5 empty] | Engineering | [6 empty] | Hori. Encoder | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | Encoder | Couch-HW-Encoder | Couch-HW-Encoder | TBD-Encoder | Encoder | Couch-HW | Couch-HW-Encoder | No | y | Gantry-HW | y | Gantry-HW-Encoder | Gantry-HW-Encoder | No | y | y | y | Gantry-HW | Gantry-HW-Encoder | y | y | y | y | y | y
492 |  |  | China | SD Jinan No.4 Peo. Hosp. | 40038 | 44247 |  | Application/Ingenuity CT 728326 | BU Q&R Feedback: No Reason: Customer Information Field Service Engineer Notes:This PO is a Reference Number only Problem Reported by Customer : Application Next Step Action: Jun Yang 13817647801 yang.yj.jun@philips.com follow Safety Questions:Case: 0117066975 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Yu Zhao Last updated by: Yu Zhao 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 5 |  | Closed - Non-Complaint | [12 empty] | 117066975 | Closed | 655410 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Others | [12 empty] | Refresh |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
493 |  |  | Japan | Higashi-Ohmi General Medical C | 341015 | 44260 |  | Message for Otani.The scheduled time has been changed from 16:00 to 16:45. | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer: Message for Otani.The scheduled time has been changed from 16:00 to 16:45. ----------------------------------------------------------------------------- | 5 |  | NAC, schedule change | [12 empty] | 117139690 | In Process | 671444 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Others | [12 empty] | Refresh |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
494 |  |  | Japan | Higashi-Ohmi General Medical C | 341015 | 44258 |  | CT:APP:NO:ﾋｻｷ様:質問 | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Application Support Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer : Q:①胸腹撮影での腹部のみのMPRの作り方 ②造影のタイミングをDICOMタグに載せられるか ③歯科でのO-MAR ④MPRのシリーズNo A:①Sub-selectionの操作をお伝えしました ②ラベルに入力していただければ残ります。自動的にはタグには入りません。 ③O-MARは整形を対象としており、2500HUかつ10ピクセル以上のものが金属として認識されますので、歯科は難しい場合もあります。FOVによって変わりますので検討ください ④スキャンの順番に沿ったシリーズNoがつくようになっており、コンソールでのシリーズNoの変更はできません。 Safety Questions : Case: 0117125311 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Ootani Natsuki Last updated by: Ootani Natsuki 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 5 |  | Closed - Non-Complaint | [12 empty] | 117125311 | Fixed | 667000 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Others | [12 empty] | Refresh |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
495 |  |  | Japan | Higashi-Ohmi General Medical C | 341015 | 44257 |  | Streaking and beam hardening artifacts from the mattress appear on the brain image(attached the images). | Streaking and beam hardening artifacts from the mattress appear on the brain image(attached the images). Customer perform head exam for emergency patient without using the head holder and they set patient feet-fist because it is difficult to set the patient with ventilator head-first. This issue has already been reported from other facilities. For example, PR No.101766 and 373156, Escalation Case No. 0115066709 and 0115066718. As the customer_s request to Philips, they told that they want us to prepare the mattress that dose not produce artifacts. | CFM |  | Artifacts | [12 empty] | CFM |  | 657763 | TRUE | Closed - Done | Closed - Done |  | CTUDT00632793 | [5 empty] | Engineering | [6 empty] | Artifacts | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | Matress | Couch-HW-Matress | Couch-HW-Matress | CTUDT00632793 | Matress | Couch-HW | Couch-HW-Matress | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
496 |  |  | Japan | Nikko Memorial Hospital | 341011 | 44257 |  | 04DS:KKI: A streak artifact occurred. | You can see something like a streak artifact in the trunk. [Phenomenon confirmation] I was able to confirm. Streak-like artifacts were seen in 2 patients. [Details of work] A streak artifact appeared at the part where the 17th and 26th detector modules overlap. I checked the RAW data, but couldn_t determine which side the artifact came from. We replaced the 17th detector with a new one and swapped the 26th and 34th detectors. (17⇐NEW 26⇔34) After performing various calibrations, we conducted an image quality test and confirmed that there were no problems. We performed a test scan and confirmed that there were no problems. 12Nc: 459800911271 Name: TDMS V2 M.Det Class A FRU QTY: 1 | 4 |  | DMS Module, check defective SN 459800911271:TDMS V2 M.Det Class A FRU QTY:1 | [12 empty] | 117116746 | In Process | 657180 | TRUE | Closed - Done | Closed - Done |  | CTUDT00632110 | [5 empty] | Engineering | [6 empty] | DMS Module | [5 empty] | Refresh |  |  | System issue | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | DMS Module | DMS-HW | DMS-HW-DMS Module | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
497 |  |  | Peru | Hogar Clinica San Juan de Dios | 40034 | 44259 |  | CIRCULAR ARTIFACT IN THE CENTER OF THE IMAGE | CIRCULAR ARTIFACT IN THE CENTER OF THE IMAGE | 3 |  | Artifacts | [12 empty] | 117136613 | New | 674066 | TRUE | Closed - Duplicate | Closed - Duplicate |  | TBD-Translation | [5 empty] | Engineering | [6 empty] | Artifacts | [5 empty] | Refresh |  |  | System issue | System | TBD | Translation | TBD-Translation | TBD-Translation | TBD-Translation | TBD-Translation | TBD | TBD-Translation | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
498 |  |  | Spain | S.C.S. HOSP. GRAL. DE FUERTEVE | 40020 | 44263 |  | Does not rebuild images | "Work Order #: WO-03108247 Carlos Saumell Lasala The system does not meet specification and is out of use They cannot do rebuilds; the equipment gives a communication error with the rebuilder. Regulatory Questions: N / A Follow Up Required: We must go on to check the startup and operation of the CIRS computer. Our engineer, Carlos Delgado, knows the problem. Problem Description by Engineer: They cannot perform rebuilds; the equipment gives a communication error with the rebuilder. Diagnostic performed by Engineer: I can not get the remote connection to the rebuilder; the CIRS computer is not turned off, but the operating system should not load properly. | 3 |  | Recon | [12 empty] | 117151764 | In Progress | 678931 | TRUE | Closed - Done | Closed - Done |  | TBD-Translation | [5 empty] | Engineering | [6 empty] | Recon | [5 empty] | Refresh |  |  | System issue | System | TBD-Bugreport | TBD_Bugreport | TBD-Translation | TBD-Translation | TBD-Translation | TBD-Translation | TBD-Bugreport | TBD-Translation | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
499 |  |  | Spain | S.C.S. HOSP. GRAL. DE FUERTEVE | 40020 | 44260 |  | One of the studies carried out yesterday afternoon does not appear (patient with contrast) | One of the acquired patients did not appear on the Local list. When I spoke with María, the Tac technician, she had already managed to have the missing images again. She has rebuilt them again from the raw data. | 3 |  | It is confirmed with the client that the remote partner contacted the user and the patient images were retrieved. The user indicates that the acquired images do not appear in the list. | [12 empty] | 117140249 | In Progress | 675539 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-Translation | [5 empty] | Engineering | [6 empty] | Image missing | [5 empty] | Refresh |  |  | System issue | System | TBD-Bugreport | TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | TBD-Translation | TBD_Bugreport | TBD-Bugreport | TBD-Bugreport-TBD_Bugreport | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
500 |  |  | Spain | MAZ ZARAGOZA M.A.T.E.P.S.S.11 | 40042 | 44256 |  | The table locks. | Work Order #: WO-03065279 Mar-04-2021 Diego Soria The system fully meets specification and returned to use. Diagnostic performed by Engineer: The system works properly. We are told the table was found to be unlocked twice and it wasn_t working when using buttons. They had to restart the system completely to make it work. Checked errors and found that the main contactor is not closed and the system is waiting for it to close. Need to go back to check the error in situ and detect whether it repeats. Problem Description by Engineer: Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Regulatory Questions : N/A Follow Up Required: Check if the reason for the notice is repeated. ----------------------------------------------------------------------------- | 3 |  | We are told the table was found to be unlocked twice and it wasn_t working when using buttons. They had to restart the system completely to make it work. After speaking with the service staff, the reason for the notice is not reproduced, he has not had a problem with the table again. They are working without problems. The equipment is ready for use. | [12 empty] | 117111562 | In Progress | 658898 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-Button | [5 empty] | Engineering | [6 empty] | Button | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | UnlockButton | Couch-HW-UnlockButton | Couch-HW-UnlockButton | TBD-Button | UnlockButton | Couch-HW | Couch-HW-UnlockButton | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
501 |  |  | Sweden | Ludvika Lasarett | 40049 | 44259 |  | CT box blacked out. Attached e-mail: | Fuse F1 on SDB blown. Changed fuse as well as CT box. 1.00000,459801705022,CT BOX ASSEMBLY_BEST 1.00000,455014003271,SDB Board (only the board itself) | 3 |  | CT Box, share more detail 1.00000,459801705022,CT BOX ASSEMBLY_BEST 1.00000,455014003271,SDB Board (only the board itself) | [12 empty] | 117134507 | In Progress | 672877 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632185 | [5 empty] | Engineering | [6 empty] | CT-Box | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | CT Box | Gantry-HW-CT Box | Gantry-HW-CT Box | CTUDT00632185 | CT Box | Gantry-HW | Gantry-HW-CT Box | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
502 |  |  | Sweden | Avesta lasarett | 40040 | 44258 |  | Gantry error, customer rebooted system | Work Order #: WO-03082295 Knudsen Martin The system fully meets specification and returned to use Problem Description by Engineer : Gantry error Diagnostic performed by Engineer : Found Overcurrent on APM and CPM. Customer did a reboot and system working again. If error reoccurs, we need to start with Filament calibration. ----------------------------------------------------------------------------- | 3 |  | Found Overcurrent on APM and CPM. Customer did a reboot and system working again. If error reoccurs, we need to start with Filament calibration. | [12 empty] | 117127033 | In Progress | 667781 | TRUE | Closed - Done | Closed - Done |  | TBD-APM/CPM | [5 empty] | Engineering | [6 empty] | APM/CPM | [5 empty] | Refresh |  |  | System issue | System | IC-HW | APM/CPM | IC-HW-APM/CPM | IC-HW-APM/CPM | TBD-APM/CPM | APM/CPM | IC-HW | IC-HW-APM/CPM | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
503 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 341006 | 44263 |  | Gantry sounds | Work Order #: WO-03106101 Soner Saglam The system partly meets specification and is returned to use with restrictions Resolution: Gantry sound comes from right back, belt should be checked. Problem Description by Engineer: Gantry sounds inside. Follow Up Required: Gantry right behind the sound, belt should be checked. | 3 |  | vThe gantry of the CT Device was checked, it was detected that there was a noise from the belt part while turning, The belt order was made, the device was delivered in working condition. | [12 empty] | 117149473 | In Progress | 678317 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | ECR-077012 | [5 empty] | Engineering | [6 empty] | Belt | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Belt | Gantry-HW-Belt | Gantry-HW-Belt | ECR-077012 | Belt | Gantry-HW | Gantry-HW-Belt | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
504 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 341006 | 44263 |  | Pacs transfer problem | Work Order #: WO-03106102 Soner Saglam The system fully meets specification and returned to use Problem Description by Engineer: There is a problem sending images to PACS. Diagnostic performed by Engineer: 0117119587 followed by. | 3 |  | FCO application number 7280040 has been done. The device is in working condition. March 10, 2021 at 8:09 AM | [12 empty] | 117149476 | In Progress | 678116 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
505 |  |  | Turkey | T.C Saglik Bakanligi Bagcilar -> Bağcılar State Hospital | 40010 | 44258 |  | Table noise during horizontal movement | Work Order #: WO-03080565 03-Mar-2021 Serkan Yavaşça The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : The malfunctioning Horizontal Linkage was ordered. Follow Up Required : Horizontal driver replacement. Back cover switch replacement. | 3 |  | Linkage 453566147251 - SWITCH FRNT INTERLCK 459801539352 - FRU, LINKAGE ASSY- STD COUCH | [12 empty] | 117125453 | In Progress | 663184 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00631671 | [5 empty] | Engineering | [6 empty] | Linkage | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | Horizontal linkage | Couch-HW-Horizontal linkage | Couch-HW-Horizontal linkage | CTUDT00631671 | Horizontal linkage | Couch-HW | Couch-HW-Horizontal linkage | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
506 |  |  | Turkey | HITACHI HEALTHCARE TURKEY | 337000 | 44256 |  | Ring artifact in bone scans. | Work Order #: WO-03065769 03-Mar-2021 Tunay Uysal The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Artefact in bone imaging. Resolution : Opaque material was found in high resolution part and DMS. These parts were cleaned. Air calibration, bad detector, IQ Check were completed. No artifacts were found. The system was delivered in functional condition after testing with phantom. ----------------------------------------------------------------------------- | 3 |  | Artifacts, contrast intrusion High resolution part and DMS opaque substance were detected in the examination. Relevant parts have been cleaned. Air calibration, bad detector, IQ Check was done. No artifact was seen. The system was tested with the phantom in the clinic and delivered ready to use. | [12 empty] | 117111940 | In Progress | 659143 | TRUE | Closed - Done | Closed - Done |  | ECR-075975 | [5 empty] | Engineering | [6 empty] | Lexan Ring/Contrast | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Lexan Ring | Gantry-HW-Lexan Ring | Gantry-HW-Lexan Ring | ECR-075975 | Lexan Ring | Gantry-HW | Gantry-HW-Lexan Ring | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
507 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 341006 | 44256 |  | Scans are aborted. | Work Order #: WO-03065344 Emre Gül The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : Error displayed in some scans. Bugrep was analyzed. CIRS raid disk and rotation frame related errors were found. On site diagnosis is required. Device is active. Follow Up Required : Diagnose CIRS raid disk and rotation frame related errors on site. ----------------------------------------------------------------------------- Work Order #: WO-03065397 02-Mar-2021 Suat Firat The system fully meets specification and returned to use Problem Description by Engineer : Scans are aborted. Resolution : Error log was analyzed. Recon computer was reset. The system was delivered in working condition after testing. ----------------------------------------------------------------------------- | 3 |  | Diagnose CIRS raid disk and rotation frame related errors on site. The error log check of the device was done, the recon computer was reset, the test shot was made and the device was delivered in working condition. | [12 empty] | 117110402 | Incomplete | 655192 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-ServiceTool | [5 empty] | Engineering | [6 empty] | CIRS | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Service tool | Console-SW-Service tool | Console-SW-Service tool | TBD-ServiceTool | Service tool | Console-SW | Console-SW-Service tool | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
508 |  |  | United States | Doctors Hospital LLC | 40014 | 44259 | 10 | Gantry spontaneously turns off, Gantry is also tilting again | Gantry spontaneously turns off, Gantry is also tilting again Created case and Customer claims Warranty (recent installation) | 2 |  | 0000868840 is Created on 2021-03-11 16:57:22 and Action Due Date is 2021-03-11 00:00:00 With Action Owner Role Quotation And Hub GBS-NA action description is below please cancel case and quote, duplicate | [12 empty] | 117136960 | New | 674239 | TRUE | Closed - Done | Closed - Done |  | TBD-GMP | [5 empty] | Engineering | [6 empty] | GMP | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | GMP | Gantry-HW-GMP | Gantry-HW-GMP | TBD-GMP | GMP | Gantry-HW | Gantry-HW-GMP | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
509 |  |  | United States | Odessa Regional Medical Center | 40029 | 44169 |  | Cannot deselect PACS as send to destination for Structured Dose Report; It is WAS(Works as specified). | Split from PR#467713 | CFM |  | Dose SR | [12 empty] | CFM |  | 667155 | TRUE | Closed - Done | Closed - Done | [7 empty] | Customer | [6 empty] | Dose SR | [5 empty] | Refresh |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
510 |  |  | United States | Unity Health Systems | 40037 | 44246 |  | Gantry noise when rotation, not replicated | Split from PR#629198 | CFM |  | Split from PR#629198 | [12 empty] | CFM |  | 666211 | TRUE | Closed - Done | Closed - Done |  | TBD-GantryNoise | [5 empty] | Engineering | [6 empty] | Gantry noise | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Noise | Gantry-HW-Noise | Gantry-HW-Noise | TBD-GantryNoise | Noise | Gantry-HW | Gantry-HW-Noise | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
511 |  |  | United States | Doctors Hospital LLC | 40014 | 44257 |  | Gantry keeps shutting off | Work Order #: WO-03076168 05-Mar-2021 Lynn Matthew The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : Checked error logs, found gantry tilt errors. Checked gantry tilt mechanism, found faulty potentiometers which track the gantry tilt motion. Need to replace the parts. Advised customer not to tilt gantry until the parts are installed. System is otherwise operational. Follow Up Required : Need to install new tilt potentiometers ----------------------------------------------------------------------------- | 4 |  | Tilt potentiometer | [12 empty] | 117121214 | In Progress | 664477 | TRUE | Closed - Done | Closed - Done |  | TBD-GMP | [5 empty] | Engineering | [6 empty] | GMP | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | GMP | Gantry-HW-GMP | Gantry-HW-GMP | TBD-GMP | GMP | Gantry-HW | Gantry-HW-GMP | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
512 |  |  | United States | Unity Health Systems | 40037 | 44258 |  | Patient can barely hear instructions; Changed the Volume from 24 to 75 | Work Order #: WO-03081979 03-Mar-2021 Taylor, Durrance The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Device Use: in use Expected and Unexpected Results: User Impact: volume for gantry speaker to low Patient Impact: patient can barely hear instructions Current Software Version: 4.1.10 Resolution : Powered down system and removed covers. Inspected gantry audio board and cable and speaker. Reseated cables to speaker and audio board. Then ran gantry audio test. Volume is still to low. Ran gantry speaker test while adjusting the audio settings and having the customer listen to the speaker. Volume was set at 24. I changed it to 75. Much louder than before. ----------------------------------------------------------------------------- | 5 |  | adjusting the audio settings | [12 empty] | 117126780 | Closed | 664104 | TRUE | Closed - Done | Closed - Done | [7 empty] | Customer | [6 empty] | Audio setting | [5 empty] | Refresh |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
513 |  |  | United States | McLaren Macomb | 40021 | 44255 | 9 | Ring Artifact | Link to 0117107731 Diagnostic performed by Engineer Description: Moved detector module 24 to 40 and calibrated system, performed bad detectors, phantom, air, constancy, and IQ, system is with in spec | 2 |  | Cancel or Duplicate | [12 empty] | 117107698 | Canceled | 655237 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Duplicate | [5 empty] | Engineering | [6 empty] | Artifacts | [5 empty] | Refresh |  |  | System issue | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | Duplicate | DMS Module | DMS-HW | DMS-HW-DMS Module | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
514 |  |  | United States | Wellstar Kennestone Hospital | 68114 | 44262 |  | 0711CDN84 Blue screen error with windows system dump and system error | Work Order #: WO-03104770 Powell, Jamie The system does not meet specification and is out of use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : Blue screens and shutdown.. Customer also having multiple application errors Looked in host event logs but no errors found Check hard drive for errors Reload software Follow Up Required : Blue screens and shutdown.. Customer also having multiple application errors Looked in host event logs but no errors found Check hard drive for errors Reload software ----------------------------------------------------------------------------- Work Order #: WO-03104776 08-Mar-2021 Bourne Barrett The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : Replaced Host hard drive and loaded software. ----------------------------------------------------------------------------- | 2 |  | System still not upgrade to 4.1.10 | [12 empty] | 117147281 | In Progress | 677496 | #N/A | #N/A | #N/A | [7 empty] | Others | [12 empty] | Non-Refresh |  |  | Upgrade issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | #N/A | y |  |  | y
515 |  |  | United States | Christ Hospital | 310034 | 44257 |  | Table Does Not Always Move | Work Order #: WO-03078255 Paine Matthew The system does not meet specification and is out of use Diagnostic performed by Engineer : Customer states that they have been having intermittent table issues were it stops moving with the buttons and foot pedal, but can be moved by the tape switch controls. Customer gave permission to analyze logs via PRS. Problem Description by Engineer : Customer Function/Role: Bio-Med Device Use: Diagnostic Expected and Unexpected Results: Table Does Not Always Move User Impact: Unknown Patient Impact: Unknown Current Software Version:N/A Follow Up Required : Recommended Follow up Part for FSE: CONTROL PANEL ASSEMBLY-FRONT LH 459800484501 Error Associated with this :S_COMMAND_BUTTON_STUCK_ERROR-Stuck Left HOut. ----------------------------------------------------------------------------- | 4 |  | System still not upgrade to 4.1.10 | [12 empty] | 117122876 | In Progress | 660926 | #N/A | #N/A | #N/A | [7 empty] | Others | [12 empty] | Non-Refresh |  |  | Upgrade issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | #N/A | y |  |  | y
516 |  |  | United States | Doctors Hospital LLC | 40014 | 44258 |  | Scanned a T and L spine decube | • Feedback Type: Technical Issue • Other Details: other • Submitter Email: Lori.Vierneisel@Philips.com • CFM Form Received Date (Outlook email received date): 03/03/2021 - Scanned a T and L spine decube o Noticed that for the Axial Bone Series (2.1)on the T spine the images were correctly labeled as if the patient had been scanned supine. They were also labeled as Derived o On the L Sp that was not the case on the Axial Bone series were labeled Original and we had to redo the images in CT Viewer to flip them so it looked like the pt had been scanned supine Work Around - Prior to scanning the patient if you tilted series 2.1 even 1 degree the series was then labeled Derived. Upon reconstruction, the images were as if the patient was scanned supine. Saves the technologists from loading the images into the CT Viewer to manipulate and rotate images after scanning into desired planes for the radiologist. | WANG, Peter |  | defect raised CTUDT00632618 | [14 empty] | 673825 | TRUE | Closed - Done | Closed - Done |  | CTUDT00632618 | [5 empty] | Engineering | [6 empty] | Spine | [5 empty] | Refresh |  |  | System issue | System | Console-SW | PoorIQ | Console-SW-PoorIQ | Console-SW-PoorIQ | CTUDT00632618 | PoorIQ | Console-SW | Console-SW-PoorIQ | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
517 |  |  | United States | Doctors Hospital LLC | 40014 | 44260 |  | Customer would like the DOB format changed to | Enhancement request: For Veradius Unity and Zenition product with available optional wireless footswitch. It would be ideal to have a charge cable built into the MVS where the footswitch docks. The pigtail charge could come out of the right side of the front cover, powered from the MVS and connect to the wireless footswitch that is hanging in It is dock. This would eliminate the need for an additional charger and pocket for the charger. It would minimize chances of loosing the charger and leaving the footswitch behind when prepping for a case, assuming the footswitch is charged away from the MVS. | WANG, Peter |  | Raise defect for RFC | [14 empty] | 687623 | TRUE | Closed - Done | Closed - Done | [7 empty] | Customer | [6 empty] | RFC | [5 empty] | Refresh |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
518 |  |  | China | Hebei Laonianbing Hospital | 345015 | 44262 | 10 | 02N/Ingenuity Core128/Elite China/The boot key won_t close | BU Q&R Feedback: No Reason: Corrective Maintenance External Remarks : 13.03.2021 07:02:50 UTC Jianping Sun (86039023)"SUAT. Ready for Clinical Use local time" : "2021-03-09 11:00"服务内容：estop不能闭合，到场查看estop回路，重新安装信号刷。 服务结果：维修后，设备恢复正常工作。 Internal Remarks : 13.03.2021 07:03:00 UTC Jianping Sun (86039023) RDI Accuracy：yes. $Q7:(Confirmed the reported problem. estop cannot closed) $Q8:(Check the eventlogger found estop loop) $Q9:(Analysis the error log VIA ATP，reinstall signal brush) $Q10:(Test and Inspection data refer to PA.Device returened to full futionnallity.) $Q11:(Calibrated tools name（FLUKE287）S/N（23350065）Next calibration date:(06/04/2021)) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date:N/A ^Current Tube usage:N/A Problem Reported by Customer : The boot key won_t close Safety Questions : Case: 0117146624 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Qi Xiaoyu Last updated by: Qi Xiaoyu 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 有ESTOP闭合提示框，拧钥匙后有继电器吸合声，但随即断开。 设备本身技术问题： 是。 Troubleshooting Action: 远程诊断与测试结果： PRS不通，很难判断问题 PRS TR状态： 不通，请工程师到场检查。 Repair Action: 建议现场维修步骤： 建议工程师到场检查log，教客户使用飞利浦微信报修 Next Step Action : Sun Jianping SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support ----------------------------------------------------------------------------- | 2 |  | estop不能闭合，到场查看estop回路，重新安装信号刷。 服务结果：维修后，设备恢复正常工作。 | [12 empty] | 117146624 | Closed | 680509 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-Translation | [5 empty] | Engineering | [6 empty] | Signal Brush | [5 empty] | Refresh |  |  | System issue | System | NotEngineering | ShouldbeOthers | TBD-Translation | TBD-Translation | TBD-Translation | TBD-Translation | NotEngineering | TBD-Translation | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
519 |  |  | China | XIFENG CHINESE MEDICINE HOSPIT | 345014 | 44263 |  | 02N/Ingenuity Core128/Elite China/error report | Problem Description: when do a ECG scan it reported an error Troubleshooting Action: ACQ error, ERROR - Invalid DFS setting only have one time, contact the engineer he said to have an observation. Repair Action: reboot the gantry the system return full function Test/Inspection: Functional Check PASSED. Device returned to full functionality. It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: "ACQ ERROR, RESEND =0 ERROR - Invalid DFS setting BufferSel: 3 ReadingOffset: 0xC7D8200 Source: AcqWorkItem(227) ProcessID: 5064 ThreadID: 4424" Remote troubleshoot and test result: The coronary artery scan is interrupted. Routine scanning is normal. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Check rotation related parts. We recommend spare parts 12NC and names: R2D Assembly 459800467342 T2 Activities : Problem Description: when do a ECG scan it reported an error Troubleshooting Action: ACQ error, ERROR - Invalid DFS setting only have one time, contact the engineer he said to have an observation. Repair Action: reboot the gantry the system return full function Test/Inspection: Functional Check PASSED. The device returned to full functionality. Test Equipment: NA SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | reboot the gantry the system return full function, collect bugreport if issue still occur | [12 empty] | 117149523 | Fixed | 681789 | TRUE | Closed - Done | Closed - Done |  | CTUDT00630408 | [5 empty] | Engineering | [6 empty] | ACQ error | [5 empty] | Refresh |  |  | System issue | System | CIRS-SW | ACQ_NOT_OK | CIRS-SW-ACQ_NOT_OK | CIRS-SW-ACQ_NOT_OK | ACQ_NOT_OK | ACQ_NOT_OK | CIRS-SW | CIRS-SW-ACQ_NOT_OK | No | y | TBD-Bugreport | TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | No | y | TBD_Bugreport | TBD_Bugreport | TBD-Bugreport | TBD-Bugreport-TBD_Bugreport | y | y | y | TBD_Bugreport | y | y
520 |  |  | Japan | Nikko Memorial Hospital | 341011 | 44263 |  | 04DM:KK: Ring artifact occurred.>> The artifact appeared around M18D273 or M25D400. | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 04DM:KK: A ring artifact occurred.>> The artifact appeared around M18D273 or M25D400. T2 Activities:Problem Description: subject equipment: CT Problem location: 04 Gantry Exclude HV When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user_s work flow did the problem occur?: (During Acquisition/Scan) What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem? No.The customer has not created a log file or report. Troubleshooting Action: Result of response: Problem not solved. Problem location: 04 Gantry Exclude HV What steps did you perform to diagnose the problem? Ring artifact occurrence. The artifact appeared around M18D273 or M25D400. I asked Hashimoto to respond. Repair Action: A service engineer visited the site. OneEMS internal remarks: --Troubleshooting Action-- Result of response: Problem not solved. Problem location: 04 Gantry Exclude HV What steps did you perform to diagnose the problem? Ring artifact occurrence. The artifact appeared around M18D273 or M25D400. I asked Hashimoto to respond. --Repair Action-- Recommended 12Nc: --None-- 459800911281:TDMS V2 M.Det Class B FRU QTY:1 SLOC:--None-- --None-- 459800911271:TDMS V2 M.Det Class A FRU QTY:1 SLOC:--None-- --Subject-- 04DM:KK:Mr. Sasaki: A ring artifact occurred.>> The artifact appeared around M18D273 or M25D400. --Problem Description-- Equipment affected: CT Problem location: 04 Gantry Exclude HV When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user_s work flow did the problem occur?: (During Acquisition/Scan) What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem? No.The customer has not created a log file or report. Safety Questions:Case: 0117147425 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: KENICHI KOJIMA Last updated by: KENICHI KOJIMA 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Support External Remarks:10.03.2021 22:20:16 UTC HODAKA HASHIMOTO (81100553)[Details of request] Ring artifact occurred again. Problem confirmation: I confirmed it. Work details: The ring artifact that occurred on March 2 occurred again. I confirmed the problem by looking at the images.The artifact appeared in the same area as the one that appeared last time. While I was working to identify the cause, I discovered that Module 26 inside the detector was experiencing a signal sensitivity problem. I replaced Module 26. I performed various calibrations and tests to confirm that the problem had been resolved. Other problems/Measures taken: None. Equipment status The equipment can be used. Work status Work finished. 11-Mar-2021 81100553 FSE Handover Tool1.3.4.32612.DB13 20200715 Internal Remarks:10.03.2021 22:20:18 UTC HODAKA HASHIMOTO (81100553) Replacement parts: Replacement parts: As below 12Nc:459800911281 Description: TDMS V2 M.Det Class B FRU QTY:1 [Report on mistakes in parts order] No mistakes. Situation monitoring information: Not monitoring the situation TS Request: No parts to return. Additional CT information: Q1. How did the examination finish? Q1. No (not applicable) Q2. Were there any mistaken findings (misdiagnosis, mistaken treatment) due to the problem? None. Q3. Were there any artifacts? None. Q4. Was there any unintended unnecessary exposure? None. Repair sticker: I affixed a Repaired sticker. [Other information to share] 1. Was treatment/quarantine carried out for a patient who has or is suspected of having new coronavirus? Yes 2. Is it necessary to go into treatment rooms/beds where there are infected patients?(Including workplace and routes of movement) No Did you agree on how to complete the report with the person in charge of devices? Yes. I completed normal reporting that day. 11-Mar-2021 81100553 FSE Handover Tool1.3.4.32612.DB13 20200715 ----------------------------------------------------------------------------- | 3 |  | 12Nc:459800911281 Description: TDMS V2 M.Det Class B FRU QTY:1 collect SN info | [12 empty] | 117147425 | In Process | 680541 | TRUE | Closed - Done | Closed - Done |  | CTUDT00632110 | [5 empty] | Engineering | [6 empty] | DMS Module | [5 empty] | Refresh |  |  | System issue | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | DMS Module | DMS-HW | DMS-HW-DMS Module | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
521 |  |  | Chile | HOSPITAL BARROS LUCO TRUDEAU | 341005 | 44263 |  | CL1688 - AFTER A STUDY, THE GANTRY TURNED OFF. | Work Order #: WO-03109395 González, Carlos The system fully meets specification and returned to use. Problem Description by Engineer : Problem description by engineer : GANTRY SHUTS OFF. When was the first occurrence? Within... : 08-Mar-2021 11:00:00 How often is issue occurring? : First occurrence Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected results: No further information is available as per the customer Patient Impact/Outcome: No further information is available as per the customer. User Impact: No further information is available as per the customer Current Software Version: 4.1.10 Internal Comments : Error # and/or description of error : N/A Diagnostic performed by Engineer: The error log was checked and it did not show anything definitive as a cause for the shutdown. The case was kept open from Monday until today Thursday and the error did not occur again. I close the case. ----------------------------------------------------------------------------- | 3 |  | The error log was checked and it did not show anything definitive as a cause for the shutdown. The case was kept open from Monday until today Thursday and the error did not occur again. I close the case. | [12 empty] | 117152680 | Incomplete | 683436 | TRUE | Closed - Done | Closed - Done | [7 empty] | Others | [6 empty] | Not clear | [5 empty] | Refresh |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
522 |  |  | China | HLJ Huachuan People Hosp. | 345018 | 44264 |  | 02N/Core128/Customer Feedback. | Next Step Action: emailed to inform ASP YiYuanji to follow up with customer feedback and explain related Next Step Action: ASP YiYuanji: After engineer_s maintenance, the machine is back to normal. On-Duty Bian 15946595552. Have notified the engineer to contact the hospital teacher again, thanks. Next Step Action: PM SWO# 52648294 Ingenuity Core128, Heilongjiang Province Huachuan County People_s Hospital, dispatch ASP YiYuanji to service on site. Teacher Si 18845473456 called and said the equipment could not be turned on today after service was done last night. Customer is complaining, saying the device can not be opened after the service, which is worse than no service done.. Also ask ASP YiYuanji FSE to coordinate with the customer to communicate. If necessary to appease the customer. ----------------------------------------------------------------------------- | 3 |  | equipment could not be turned on today after service was done | [12 empty] | 117155894 | Closed | 684971 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Not clear | [5 empty] | Refresh |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
523 |  |  | China | Suzhou Guangci Tumor Hospital | 345019 | 44264 |  | 02N/Ingenuity Core128/Estop open | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer:Estop open Safety Questions:Case: 0117157310 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Cao Last updated by: Liu Cao T2 Activities:Problem Description: Summary of customer_s problems: Fail to close estop after power off. Detailed description of Customer_s problems or malfunctions: Fail to close estop after power off. It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: S_COUCHMGR_HOR_MOTION_STOP_ERROR S_HMC_VELOCITY_ENCODER_FAILURE S_MPS_WAITING_FOR_MAIN_DRIVE_TO_CLOSE Check Estop button if it release normal Remote troubleshoot and test result: Need to go onsite to confirm the power off cause and the resistance of the bed PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: 1. power off reason, confirm if bed horizontal movement has resistance 2. Check estop switch, gull wing, ACCU 24V indicator, SK confirm the cause of power on failure Test Equipment: N/A SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support External Remarks: 10.03.2021 06:38:50 UTC Haotian Xu (86090280) Service description: Power on failure. Check Estop switch, then power on is normal Service outcome: After the service, the equipment goes back to normal. Internal Remarks: 10.03.2021 14:54:11 UTC Haotian Xu (86090280) $Q7: (Confirm reported problem. Gantry estop open and can not power up) $Q8: (check the couch and E-stop button) $Q9: (restart the gantry) $Q10: (Test and inspection data refer to PA. Device returned to full functionality.) $Q11: Calibrated tools refer to PA. ----------------------------------------------------------------------------- | 3 |  | 服务内容： 上电失败，检查Estop开关后能正常上电 服务结果： 服务后，设备恢复正常使用。 No bugreport collected Forward Changxian to investigate | [12 empty] | 117157310 | Closed | 681008 | TRUE | Closed - Done | Closed - Done |  | TBD_Bugreport | [5 empty] | Engineering | [6 empty] | E-Stop | [5 empty] | Refresh |  |  | System issue | System | TBD-Bugreport | TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | TBD_Bugreport | TBD_Bugreport | TBD-Bugreport | TBD-Bugreport-TBD_Bugreport | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
524 |  |  | Japan | Nikko Memorial Hospital | 341011 | 44264 |  | JY: Power supply replacement. | BU Q&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP Problem Reported by Customer: JY: Power supply replacement OneEMS internal remarks: Hashimoto DB SAP Transfer: Dispatched for Onsite Support External Remarks:10.03.2021 22:25:06 UTC HODAKA HASHIMOTO (81100553)[Details of request] The power supply was replaced. Problem confirmation: I couldn_t confirm the problem. Work details: This occurred on March 2nd and 8th.As a countermeasure for ring artifacts, I replaced the power supply unit for the detector. I confirmed that the power output and input values are within the specifications after the replacements. Other problems/Measures taken: None. Equipment status The equipment can be used. Work status Work finished. 11-Mar-2021 81100553 FSE Handover Tool1.3.4.32612.DB13 20200715 Internal Remarks:10.03.2021 22:25:09 UTC HODAKA HASHIMOTO (81100553) Replacement parts: Replacement parts: As below 12Nc:459800359381 Description:POWER SUPPLY, TRIPLE OUTPUT QTY:2 [Report on mistakes in parts order] No mistakes. Situation monitoring information: Not monitoring the situation TS Request: No parts to return. Additional CT information: Q1. How did the examination finish? Q1. No (not applicable) Q2. Were there any mistaken findings (misdiagnosis, mistaken treatment) due to the problem? None. Q3. Were there any artifacts? None. Q4. Was there any unintended unnecessary exposure? None. Repair sticker: I affixed a Repaired sticker. [Other information to share] 1. Was treatment/quarantine carried out for a patient who has or is suspected of having new coronavirus? Yes 2. Is it necessary to go into treatment rooms/beds where there are infected patients?(Including workplace and routes of movement) No Did you agree on how to complete the report with the person in charge of devices? YES Contacted the customer regarding the status the next day. 11-Mar-2021 81100553 FSE Handover Tool1.3.4.32612.DB13 20200715 ----------------------------------------------------------------------------- | 5 |  | 12Nc:459800359381 Description:POWER SUPPLY, TRIPLE OUTPUT QTY:2 | [12 empty] | 117156092 | New | 684998 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632110 | [5 empty] | Engineering | [6 empty] | DMS? | [5 empty] | Refresh |  |  | System issue | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | DMS Module | DMS-HW | DMS-HW-DMS Module | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
525 |  |  | United States | Doctors Hospital LLC | 40014 | 44264 | 11 | Unable to scan, ran potentiometer calibrations | Work Order #: WO-03118329 09-Mar-2021 Testerman Bennett The system fully meets specification and returned to use Diagnostic performed by Engineer : review logs, run potientiometer cals retestest tilt circuit Problem Description by Engineer : Customer Function/Role: unable to scan Device Use: in use Expected and Unexpected Results: review error logs User Impact: un able to scan Patient Impact: Current Software Version: | 2 |  | review logs, run potientiometer cals retestest tilt circuit Ask FSE to check if this is related to belt tension too tight and share CAN Trace if have | [12 empty] | 117161330 | Closed | 684678 | TRUE | Closed - Done | Closed - Done |  | TBD_Bugreport | [5 empty] | Engineering | [6 empty] | E-Stop | [5 empty] | Refresh |  |  | System issue | System | TBD-Bugreport | TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | TBD_Bugreport | TBD_Bugreport | TBD-Bugreport | TBD-Bugreport-TBD_Bugreport | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
526 |  |  | China | HLJ Huachuan People Hosp. | 345018 | 44264 | 11 | 02N/Core128/Can not boot up | BU Q&R Feedback: No Reason: Corrective Maintenance Safety Questions : Case: 0117155891 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Pan Chuchu Last updated by: Pan Chuchu 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Problem Reported by Customer : Can not boot up T2 Activities : Problem Description: 客户问题概述： 扫描失败。 客户问题详细描述或故障现象： 扫描失败。 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 扫描失败。 远程诊断与测试结果： 无法远程登录，怀疑打火问题 PRS TR状态： 通，但是无法下载logger Repair Action: 建议现场维修步骤： 1 查看logger 确认问题 2 按照手册设置remoteservice 登录密码并检查端口 SAP Transfer : Dispatched for Onsite Support SAP Transfer : Case Transferred to SAP External Remarks : 09.03.2021 13:03:14 UTC Gang Ni (86038229)"SUAT. Ready for Clinical Use local time" : "2021-03-09 19:00" 服务内容：清理搞压电缆，灯丝校正，长训 服务结果：设备恢复 Internal Remarks : 10.03.2021 11:00:32 UTC Gang Ni (86038229) $Q7: (Confirmed reported problem & Can not boot up) $Q8: (on site review errorlog; noload test; check HV cable & need to clean HV cable) $Q9: (clean HV cable; filament calibration Long TC) $Q10: (Test and inspection data refer to: PA. & Device returned to full functionality) $Q11: Calibrated Tools name：none ----------------------------------------------------------------------------- | 2 |  | 服务内容：清理搞压电缆，灯丝校正，长训 服务结果：设备恢复 | [12 empty] | 117155891 | Closed | 689354 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-Translation | [5 empty] | Engineering | [6 empty] | TC Boost | [5 empty] | Refresh |  |  | System issue | System | NotEngineering | ShouldbeService | TBD-Translation | TBD-Translation | TBD-Translation | TBD-Translation | NotEngineering | TBD-Translation | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
527 |  |  | China | Tongling Boai Hospital Co LTD | 345034 | 44264 |  | 02N/Ingenuity Core128/add PACS DICOM node | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Calibration/Adjustment External Remarks : 11.03.2021 08:23:36 UTC Kai Zhou (86036469)服务内容：软件重新添加pacs节点。 服务结果：服务后设备使用正常。 Internal Remarks : 12.03.2021 16:45:43 UTC Kai Zhou (86036469) RDI Accuracy:Yes $Q7: (Confirmed reported problem,need transfer picture to pacs） $Q8: (check the software，found the ip was set wrong） $Q9: (reset the lan config，trouble disppeared.) $Q10:(Test and inspection data refer to PA . Device returned to full functionality） $:Q11(calibration tools:NA） Problem Reported by Customer : the service account have no password . SAP Transfer : Dispatched for Onsite Support T2 Activities : Problem Description: 客户问题概述： 计算机软件相关故障。 客户问题详细描述或故障现象： 服务器没有管理员密码不能添加节点 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 没有管理员账号和密码 远程诊断与测试结果： 需要到场更改管理员账户密码 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 1，检查服务器设置，更改密码，2检查远程服务设置，启用远程桌面 | 3 |  | 服务内容：软件重新添加pacs节点。 服务结果：服务后设备使用正常。 | [12 empty] | 117157434 | In Process | 685968 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Service | [6 empty] | PACS Config | [5 empty] | Refresh |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
528 |  |  | Japan | Nikko Memorial Hospital | 341011 | 44264 |  | 03DM: RS: An error occurred during shooting and interrupted >> Continue shooting as it is >> S_GEN_FAULT_ANODE_TUBE_ARC | The phenomenon was confirmed on March 9. We will visit you for regular inspection on the next 10 days. We checked the HV cable and tested the reproducibility of the phenomenon, but there were no problems. Today, the generator (anode side) is replaced based on past repair results. Confirm that there is no problem with the device. We have confirmed that there is no problem in various Cal and tests. In addition, a dehumidifier was installed today. | 3 |  | Recommended 12Nc: PR1 453567560012: Spellman 80 KW RoHS APM Assy # QTY: 1 SLOC: Sapporo (U612) PR1 453567560002: Spellman 80 KW RoHS CPM Assy QTY: 1 SLOC: Fukuoka (U617) PR1 459800089031: Tool, XRT Strap Pad Assy QTY: 1 SLOC: Sapporo (U612) PR1 453566495101: GENERATOR SERVICE TOOL QTY: 1 SLOC: Sapporo (U612) | [12 empty] | 117156347 | In Process | 685485 | TRUE | Closed - Done | Closed - Done |  | TBD-APM/CPM | [5 empty] | Engineering | [6 empty] | APM/CPM | [5 empty] | Refresh |  |  | System issue | System | IC-HW | APM/CPM | IC-HW-APM/CPM | IC-HW-APM/CPM | TBD-APM/CPM | APM/CPM | IC-HW | IC-HW-APM/CPM | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
529 |  |  | Puerto Rico | HOSPITAL DR SUSONI INC | 40013 | 44264 | 11 | The post processor is slow, they are in the middle of applications | Work Order #: WO-03117299 Caldelas Aldana, Carlos Manuel The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Problem description by engineer : slow application How often is issue occurring? : First occurrence Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: application is slow Patient Impact/Outcome: n/a User Impact: application is slow Current Software Version: Not provided Diagnostic performed by Engineer : several attempts to investigate furhter about the problem with customer without answer, remote conection available but without collaboration cant solve or advance in diagnostics. FSE is already confirmed is arriving onsite. Remote support not concluded. Internal Comments : Error # and/or description of error : application is slow Completed or Recommended Repair Action "Internal" : n/a ----------------------------------------------------------------------------- Work Order #: WO-03118340 09-Mar-2021 Ramos Rivera, Pablo The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Reconstructions are very slow. Gantry opens estop Diagnostic performed by Engineer : Found that belt in the tilt post was excessively tight. I realese it just a bit. Verified the health of raw data disks in the cirs amd they are ok. Found no problem in cirs. I recomend, if problem continue, software should be reloaded. In my last test it worked properly. ----------------------------------------------------------------------------- | 2 |  | Reconstructions are very slow. Gantry opens estop Diagnostic performed by Engineer : Found that belt in the tilt post was excessively tight. I realese it just a bit. Verified the health of raw data disks in the cirs amd they are ok. Found no problem in cirs. I recomend, if problem continue, software should be reloaded. In my last test it worked properly. Reloaded software to the host computer. Tested it and now is working properly. | [12 empty] | 117160602 | Closed | 684086 | TRUE | Closed - Done | Closed - Done |  | ECR-077012 | [5 empty] | Engineering | [6 empty] | Tilt Belt improvement | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Belt | Gantry-HW-Belt | Gantry-HW-Belt | ECR-077012 | Belt | Gantry-HW | Gantry-HW-Belt | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
530 |  |  | Sweden | Avesta lasarett | 40040 | 44264 |  | Right after bolus tracking, the system stops. Logs shows arcs in both anode and cathode. Long tube cond. sug [TRUNCATED] | Work Order #: WO-03113799 09-Mar-2021 JOHANSSON FREDRIK The system fully meets specification and returned to use Problem Description by Engineer : Immediately after bonus tracking and before scan, the system stops. Logs shows arcing on both anode and cathode side. Resolution : Performed long tube conditioning and filament calibration pass 2. ----------------------------------------------------------------------------- | 3 |  | Performed long tube conditioning and filament calibration pass 2. | [12 empty] | 117157452 | Closed | 681428 | TRUE | Closed - Done | Closed - Done |  | CTUDT00630085 | [5 empty] | Engineering | [6 empty] | Arc | [5 empty] | Refresh |  |  | System issue | System | IC-HW | Tub Arc | IC-HW-Tub Arc | IC-HW-Tub Arc | CTUDT00630085 | Tub Arc | IC-HW | IC-HW-Tub Arc | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
531 |  |  | China | Fujian Medical Device Co. LTD. | 345011 | 44263 |  | 02N/Ingenuity Core128/Elite China/need filter | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: need filter Safety Questions:Case: 0117149389 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created by: Wennian Han Last updated by: Wennian Han 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Field Service Engineer Notes:This PO is a Reference Number only SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | 5 | [14 empty] | 117149389 | Closed | 681415 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Others | [12 empty] | Refresh |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
532 |  |  | China | Fujian Medical Device Co. LTD. | 345010 | 44263 |  | 02N/Ingenuity Core128/Elite China/need filter | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : need filter Safety Questions : Case: 0117149189 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wennian Han Last updated by: Wennian Han 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A FSE Notes : This PO is a Reference Number only SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 5 | [14 empty] | 117149189 | Closed | 681357 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Others | [12 empty] | Refresh |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
533 |  |  | China | Fujian Medical Device Co. LTD. | 345012 | 44263 |  | 02N/Ingenuity Core128/Elite China/need filter | BU Q&R Feedback: No Reason: Corrective Maintenance Safety Questions:Case: 0117149281 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created by: Wennian Han Last updated by: Wennian Han 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Problem Reported by Customer: need filter SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | 3 | [14 empty] | 117149281 | Closed | 681363 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Others | [12 empty] | Refresh |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
534 |  |  | China | Dalian Branch of Shengjing Hospital of China Medical University | 345001 | 44265 | 10 | Host yellow screen after FCO72800740 implemented | Host yellow screen after FCO72800740 implemented | WANG, Peter |  | New Yellow Screen + PACS Transfer issue, bugreport provided for eng analysis, check if duplicate with above 694348 | [14 empty] | 685008 | TRUE | Closed - Done | Closed - Done |  | CTUDT00631367 | [5 empty] | Engineering | [6 empty] | New Yellow Screen | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
535 |  |  | Portugal | Centro Hospitalar de Vila Nova | 40030 | 44265 |  | Studies locking frequently with indication of heating of the tube, even when no studies have been performed | Blocking exams often with an indication of tube heating, even when no exams have been done yet | 3 |  | check the exam card and scan parameters which cause heating of the tube? Jerry follow up to investigate. | [12 empty] | 117165930 | New | 690772 | TRUE | Closed - Done | Closed - Done |  | TBD-APM/CPM | [5 empty] | Engineering | [6 empty] | CPM | [5 empty] | Refresh |  |  | System issue | System | IC-HW | APM/CPM | IC-HW-APM/CPM | IC-HW-APM/CPM | TBD-APM/CPM | APM/CPM | IC-HW | IC-HW-APM/CPM | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
536 |  |  | United States | Lackey Memorial Hospital | 40015 | 44265 | 11 | 1203CDN84 X-ray system cannot comply | Work Order #: WO-03123459 Ali, Adam The system does not meet specification and is out of use Problem Description by Engineer : Problem description by engineer : X-ray system cannot comply When was the first occurrence? Within... : 10-Mar-2021 How often is issue occurring? : Continuously Malfunction area : Tube&Gen Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Exam does not start Patient Impact/Outcome: Cannot be scanned & diagnosed User Impact: Cannot start an exam Current Software Version: 4.1.7 Diagnostic performed by Engineer : System will not start exam, it gives breathing instructions and then gives error _X-ray system cannot comply_. Internal Comments : Error # and/or description of error : 2021/03/10 04:42:53.339 S_GEN_FAULT_GRID_FBK_IMBALANCE ERROR_WARNING Shot Logger.mdb Unchecked Description: Grid Feedback >450V difference from commanded value. Grid voltages are in imbalance 2021/03/10 04:42:53.341 S_RHOST_HVG_HSS_GENERATOR_NON_ARC_FAULT ERROR_NORMAL Shot Logger.mdb Checked 2021/03/10 04:42:53.453__001 HOST_USR_POPUP USR_POPUP Shot Usplog.log Unchecked Showing message to user : X-ray system cannot comply. Please retry. Completed or Recommended Repair Action "Internal" : Check the feedback fiber cable. Check RHOST connectors. Check Cathode Power Module. Follow Up Required : What are the next Troubleshooting steps : Check the feedback fiber cable. Check RHOST connectors. Check Cathode Power Module. Error # and/or description of error : 2021/03/10 04:42:53.339 S_GEN_FAULT_GRID_FBK_IMBALANCE ERROR_WARNING Shot Logger.mdb Unchecked Description: Grid Feedback >450V difference from commanded value. Grid voltages are in imbalance 2021/03/10 04:42:53.341 S_RHOST_HVG_HSS_GENERATOR_NON_ARC_FAULT ERROR_NORMAL Shot Logger.mdb Checked 2021/03/10 04:42:53.453__001 HOST_USR_POPUP USR_POPUP Shot Usplog.log Unchecked Showing message to user : X-ray system cannot comply. Please retry. Completed or Recommended Repair Action "Internal" : Check the feedback fiber cable. Check RHOST connectors. Check Cathode Power Module. ----------------------------------------------------------------------------- Work Order #: WO-03123629 11-Mar-2021 Denison Chuck The system does not meet specification and is out of use Problem Description by Engineer : Problem description by engineer : X-ray system cannot comply When was the first occurrence? Within... : 10-Mar-2021 How often is issue occurring? : Continuously Malfunction area : Tube&Gen Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Exam does not start Patient Impact/Outcome: Cannot be scanned & diagnosed User Impact: Cannot start an exam Current Software Version: 4.1.7 Follow Up Required : Trouble shot C.T. / Grid error in the Gen. Will not scan. replace the CPM Module. Diagnostic performed by Engineer : Trouble shot C.T. / Grid error in the Gen. Will not scan. ----------------------------------------------------------------------------- Work Order #: WO-03126198 10-Mar-2021 Donald Vic The system does not meet specification and is out of use Problem Description by Engineer : Customer Function/Role: Radiology tech Device Use: Clinical use Expected and Unexpected Results: xray cannot comply message User Impact: system is down Patient Impact: no impact Current Software Version: 4.1.10 Diagnostic performed by Engineer : Order KPM, RHOST, system control bd. Follow Up Required : Replace kpm, rhost, system control bd. Tools & Equipment : MULTIMETER,2023-01-31 00:00:00 ----------------------------------------------------------------------------- Work Order #: WO-03127239 11-Mar-2021 Donald Vic The system fully meets specification and returned to use Parts Consumed : 1.00000,453567560002,Spellman 80 KW RoHS CPM Assy 1.00000,453567017751,RHOST SG164 ASSY. Tools & Equipment : TORQUE WRENCH,2022-02-28 00:00:00 Diagnostic performed by Engineer : I was advised to use the XRAY troubleshooting guide to diagnosis the problem. The guide directed me to measure the voltages on the filament feedback circuit on TP32 and TP28 on the system control board. If the voltages did not equal each other then to replace the KPM. the voltage on TP31 was 1V and voltage on TP 28 was .0043V. I was told to order the KPM, System control board, and RHOST. I was told to start with the system control board when troubleshooting. We installed the system control board and issue was still present. So we installed the KPM and RHOST to get the customer back up ASAP. Problem Description by Engineer : Customer Function/Role: Radiology tech Device Use: Clinical use Expected and Unexpected Results: xray cannot comply message User Impact: system is down Patient Impact: no impact Current Software Version: 4.1.10 ----------------------------------------------------------------------------- Work Order #: WO-03129356 11-Mar-2021 Denison Chuck The system fully meets specification and returned to use Tools & Equipment : TORQUE WRENCH,2022-02-28 00:00:00 Problem Description by Engineer : Problem description by engineer : X-ray system cannot comply When was the first occurrence? Within... : 10-Mar-2021 How often is issue occurring? : Continuously Malfunction area : Tube&Gen Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Exam does not start Patient Impact/Outcome: Cannot be scanned & diagnosed User Impact: Cannot start an exam Current Software Version: 4.1.7 Diagnostic performed by Engineer : Assisted Victor on the install of Cathode Module and Calibrate system. ----------------------------------------------------------------------------- | 2 |  | Diagnostic performed by Engineer Description: I was advised to use the XRAY troubleshooting guide to diagnosis the problem. The guide directed me to measure the voltages on the filament feedback circuit on TP32 and TP28 on the system control board. If the voltages did not equal each other then to replace the KPM. the voltage on TP31 was 1V and voltage on TP 28 was .0043V. I was told to order the KPM, System control board, and RHOST. I was told to start with the system control board when troubleshooting. We installed the system control board and issue was still present. So we installed the KPM and RHOST to get the customer back up ASAP. Parts Consumed Description: 1.00000,453567560002,Spellman 80 KW RoHS CPM Assy 1.00000,453567017751,RHOST SG164 ASSY Ask FSE to provide SN for eng investigate with supplier and repair report | [12 empty] | 117166369 | Closed | 686765 | TRUE | Closed - Done | Closed - Done |  | TBD-TubGenerate | [5 empty] | Engineering | [6 empty] | CPM/Rhost | [5 empty] | Refresh |  |  | System issue | System | IC-HW | APM/CPM | IC-HW-APM/CPM | IC-HW-APM/CPM | TBD-TubGenerate | APM/CPM | IC-HW | IC-HW-APM/CPM | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
537 |  |  | China | NMG Xinganmeng Hos. | 345032 | 44265 |  | 02N/Ingenuity Core128/Elite China/report error | BU Q&R Feedback: No Reason: Corrective Maintenance External Remarks : 15.03.2021 07:02:16 UTC Guanglu Zhu (86034061)服务内容: 设备放射防护验收检测中高分辨跟CT值参数未在范围，调整检测参数，恢复正常 服务结果: 服务完成后,设备恢复正常工作 Problem Reported by Customer : report error Safety Questions : Case: 0117165675 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Zhang Jing Last updated by: Zhang Jing 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : Zhu Guanglu T2 Activities : Problem Description: 客户问题概述： 其他类型故障。 客户问题详细描述或故障现象： 电脑主界面显示已启用远程控制，无法使用电脑 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 软件问题 远程诊断与测试结果： 查看log，未发现异常，但是主机有远程推送通知，主机host重启后也无法使用，软件问题 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 检查软件设置 ----------------------------------------------------------------------------- | 3 |  | 服务内容: 设备放射防护验收检测中高分辨跟CT值参数未在范围，调整检测参数，恢复正常 服务结果: 服务完成后,设备恢复正常工作 | [12 empty] | 117165675 | In Process | 690795 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Radiation Test | [5 empty] | Refresh |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
538 |  |  | Japan | Tomakomai City Hospital | 341014 | 44265 |  | 07CS: TM: Login password has been locked >> Explained how to reset password. Log in with IT and Reset Password | My login password has been locked >> Explained how to reset the password. Log in with IT and Reset Password | 4 |  | Log in with IT and Reset Password | [12 empty] | 117170263 | Fixed | 689111 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Password | [5 empty] | Refresh |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
539 |  |  | Sweden | Avesta lasarett | 40040 | 44265 |  | Remote FTP unreachable, need to check network configuration. |  | 3 |  | Remote connection | [12 empty] | 117167271 | New | 691612 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Service | [6 empty] | PRS Setup | [5 empty] | Refresh |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
540 |  |  | China | General Hospital of Jizhong Energy Jingxing Mining Group Co. LTD | 345003 | 44266 |  | Wechat：02N/Ingenuity Core128/Elite China/image cannot transfer ISP | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : image cannot transfer ISP Safety Questions : Case: 0117173873 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingli Last updated by: Lu Mingli 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： CT机不能传输图像到场isp工作站， 客户问题详细描述或故障现象： reset spooler，网络正常，重新启动主机，工作站，还是不行 设备本身技术问题： 是。 Troubleshooting Action: 远程诊断与测试结果： 怀疑软件问题 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议工程师到场检查重新安装软件，网络设置 Next Step Action : Yang Litu SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 14.03.2021 07:39:56 UTC Litu Yang (86038906)服务内容： [无法传图ISP，恢复备份。] 服务结果： [设备正常工作。] Internal Remarks : 14.03.2021 09:16:34 UTC Litu Yang (86038906) $Q7: (Confirmed the reported problem. transfer issue) $Q8: (image can not transfer to ISP，reset spooler does not work，need on site) $Q9: (Re-add the ISP，no effect，but restore the previous back，it‘s solved ) $Q10: (Test and inspection data refer to: PA. Device returned to full functionality.) $Q11: Calibrated Tools name (NA) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date: 10/19/2020 ^Current Tube usage: scan second_153727 ----------------------------------------------------------------------------- | 3 |  | 服务内容： [无法传图ISP，恢复备份。] 服务结果： [设备正常工作。] $Q8: (image can not transfer to ISP，reset spooler does not work，need on site) $Q9: (Re-add the ISP，no effect，but restore the previous back，it‘s solved ) | [12 empty] | 117173873 | Closed | 696028 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS Transfer | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
541 |  |  | China | Dalian Branch of Shengjing Hospital of China Medical University | 345001 | 44266 |  | Escalation/SWO：52557242 | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : Escalation Safety Questions : Case: 0117171709 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Mengen Xie Last updated by: Mengen Xie 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action : FSE 徐国梁013869524913来电开escalation单 ----------------------------------------------------------------------------- | 3 |  | New Yellow Screen + PACS Transfer issue, bugreport provided for eng analysis | [12 empty] | 117171709 | New | 694348 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | CTUDT00632641 | [5 empty] | Engineering | [6 empty] | New Yellow Screen | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Yellow Screen | CTUDT00632520 | Yellow Screen |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
542 |  |  | China | SD Jinan No.4 Peo. Hosp. | 40038 | 44266 | 11 | 02N/Ingenuity CT 728326/need reset password | BU Q&R Feedback: No SAP Customer Symptom Code: DIDM Customer Symptom Code: Patient or Image Data Management/Post Processing Data Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: Instructed Customer Problem Reported by Customer : need reset password T2 Activities : Problem Description: need reset password Troubleshooting Action: need reset password Repair Action: guide customer reset password Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: N/A SAP Transfer : Case Transferred to SAP Safety Questions : Case: 0117173133 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Guoping Last updated by: Liu Guoping 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 2 |  | guide customer reset password, enhance application training for upgrade site | [12 empty] | 117173133 | Closed | 690798 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Customer | [6 empty] | Password | [5 empty] | Refresh |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
543 |  |  | China | GX Tengxian People_s Hospital | 40036 | 44266 |  | 02N/Ingenuity CT 728326/couch has noise | BU Q&R Feedback: No Reason: Corrective Maintenance SAP Transfer : Case Transferred to SAP Problem Reported by Customer : couch has noise Safety Questions : Case: 0117170658 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Xie Lili Last updated by: Xie Lili N/A T2 Activities : Problem Description: 客户问题概述： 设备异响。 客户问题详细描述或故障现象： 升级床异响 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： "升级床升降异响 " 远程诊断与测试结果： 需要润滑 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 润滑床滑轨 建议备件12NC及名称： " " SAP Transfer : Dispatched for Onsite Support ----------------------------------------------------------------------------- | 3 |  | 建议现场维修步骤： 润滑床滑轨 | [12 empty] | 117170658 | In Process | 694052 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-Translation | [5 empty] | Engineering | [6 empty] | Couch Noise | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | Lubricant | Couch-HW-Lubricant | Couch-HW-Lubricant | Couch-HW-Lubricant | Lubricant | Couch-HW | Couch-HW-Lubricant | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
544 |  |  | China | JS No.2 Chinese Medicine Hosp. | 345027 | 44266 |  | wechat:02N/Ingenuity Core128/scan break off | BU Q&R Feedback: No Reason: Corrective Maintenance Internal Remarks : 12.03.2021 13:27:25 UTC Liangdong Qi (86036024) RDI Accuracy: Yes $Q7: (Confirmed reported problem, Xray system cannot comply please retry.) $Q8: (Analysis error report, found error " S_GEN_Fault_anode tube Arc" when scan failed, CPM Arc, APM arc.) $Q9: (Performed long tube condioning.) $Q10: (Test and Inspection data refer to PA & Device return to fully function) $Q11: Calibrated Tools (NA). External Remarks : 12.03.2021 13:27:22 UTC Liangdong Qi (86036024)服务内容：长训球管. 服务结果： 维修后，设备恢复正常工作. Problem Reported by Customer : scan break off N/A SAP Transfer : Dispatched for Onsite Support T2 Activities : Problem Description: 客户问题概述： 扫描失败。 客户问题详细描述或故障现象： 扫描失败。 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： xray system cannot comply please retry 远程诊断与测试结果： 怀疑高压问题，远程PRSping的通，但是无法连接，端口都是封闭的 PRS TR状态： 不通，请工程师到场检查。 | 3 |  | 服务内容：长训球管. 服务结果： 维修后，设备恢复正常工作. | [12 empty] | 117170786 | Closed | 693870 | TRUE | Closed - Done | Closed - Done |  | TBD-Translation | [5 empty] | Engineering | [6 empty] | TC Boost | [5 empty] | Refresh |  |  | System issue | System | IC-HW | Tube Arc | IC-HW-Tube Arc | IC-HW-Tube Arc | TBD-Translation | Tube Arc | IC-HW | IC-HW-Tube Arc | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
545 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 341006 | 44266 |  | Cardiac Shooting Problem | Action Request: 0000878967 has an overdue Action Due Date 2021-03-17 00:00:00, related to the Escalated case 0117173796 with the Customer Temperature as Red with Action title as Cardiag Image artifact | 3 |  | Customer Temperature as Red with Action title as Cardiag Image artifact | [12 empty] | 117173796 | In Progress | 695731 | TRUE | Closed - No Further Investigation | Closed - Done |  | TBD-CardiacShootIssue | [5 empty] | Engineering | [6 empty] | Not clear | [5 empty] | Refresh |  |  | System issue | System | Gantry-SW | IBOX2 | Gantry-SW-IBOX2 | Gantry-SW-IBOX2 | TBD-CardiacShootIssue | IBOX2 | Gantry-SW | Gantry-SW-IBOX2 | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
546 |  |  | Turkey | T.C Saglik Bakanligi Bagcilar -> Bağcılar State Hospital | 40010 | 44266 |  | Stuck in SW | Lags in PACS send and freezes in system SW. | 3 |  | New Yellow Screen + PACS Transfer issue, collect bugreport for eng analysis | [12 empty] | 117173543 | Closed | 695570 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00631367 | [5 empty] | Engineering | [6 empty] | New Yellow Screen | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
547 |  |  | Japan | Higashi-Ohmi General Medical C | 341015 | 44267 |  | JY: CT Application support package. | BU Q&R Feedback: No Reason: Application Support Problem Reported by Customer : JY: CT Application support package. SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Support Safety Questions:Case: 0117177678 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: YAN Joanna Last updated by: YAN Joanna 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A OneEMS internal remarks : Otani DB ----------------------------------------------------------------------------- | 5 |  | Application support Peter support to check if this is NAC | [12 empty] | 117177678 | New | 698234 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Application | [6 empty] | Application support | [5 empty] | Refresh |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
548 |  |  | China | Lindian County Hospital | 345038 | 44267 |  | 02G/Ingenuity Core128/Elite China/gantry fault | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer : gantry fault Safety Questions : Case: 0117178840 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Xie Lili Last updated by: Xie Lili Next Step Action : 孙玉财 Yucai Sun 13889873962，电话FSE nigang 18686766464 未接，mail 短信 Next Step Action : IA # 53166932 上加operation，派工 nigang工程师 Next Step Action : 问题已经解决了，如果operation没加，就不用再加了 | 5 |  | 问题已经解决了，如果operation没加，就不用再加了 | [12 empty] | 117178840 | Closed | 698918 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Not clear | [5 empty] | Refresh |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
549 |  |  | China | SX Datong Yanggao People Hosp. | 345025 | 44267 | 11 | 02N/Ingenuity Core128/Elite China/crash | BU Q&R Feedback: No SAP Customer Symptom Code: BOOT Customer Symptom Code: Boot up/Power-up Problem Reason: Corrective Maintenance Engineer Failure Code: Incomplete System Repair Activity Code: Reboot/Restart Problem Reported by Customer : crash Safety Questions : Case: 0117177522 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Yu Zhao Last updated by: Yu Zhao 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: the system can not boot up Troubleshooting Action: the system can not boot up PRS could not log in SHH2 and SFTP Repair Action: reboot two times the hole system ,the device return full function Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: NA SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 2 |  | hf2 reboot two times the hole system ,the device return full function, check if yellow screen? | [12 empty] | 117177522 | Closed | 698406 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-Translation | [5 empty] | Engineering | [6 empty] | SW Crash | [5 empty] | Refresh |  |  | System issue | System | TBD-Bugreport | TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | TBD-Translation | TBD_Bugreport | TBD-Bugreport | TBD-Bugreport-TBD_Bugreport | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
550 |  |  | Japan | Tomakomai City Hospital | 341014 | 44267 |  | JS:ベッド水平時異音がする | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : JS:ベッド水平時異音がする OneEMS internal remarks : 菅原Tel SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Internal Remarks : 13.03.2021 08:28:46 UTC MOTOTSUGU SUGAWARA (81100170) There is a strange noise when the bed is horizontally operated. [Phenomenon confirmation] I was able to confirm. [Details of work] There seemed to be a strange noise coming from the clutch for the tape switch in the horizontal movement part. Parts need to be replaced, so we will ask you to replace them at a later date. Other than abnormal noise, it does not affect the operation. We confirmed the operation and test scan to confirm that there were no problems. [Other failures / treatments] Yes The treatment top plate does not work well. → The alignment of the bed mounting part has been adjusted. [Device status] Equipment can be used [Work status] done | 5 |  | The treatment top plate does not work well. → The alignment of the bed mounting part has been adjusted. | [12 empty] | 117179052 | New | 700567 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-Service | [5 empty] | Engineering | [6 empty] | Couch Noise | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | Vertical spring | Couch-HW-Vertical spring | Couch-HW-Vertical spring | TBD-Service | Vertical spring | Couch-HW | Couch-HW-Vertical spring | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
551 |  |  | Peru | Hogar Clinica San Juan de Dios | 40034 | 44267 |  | 86543406 ARTEFACTO CIRCULAR EN EL CENTRO DE LA IMAGEN | air calibration reboot | 3 |  | air calibration reboot | [12 empty] | 117183199 | New | 700262 | TRUE | Closed - Done | Closed - Done |  | TBD-Translation | [5 empty] | Engineering | [6 empty] | Not clear | [5 empty] | Refresh |  |  | System issue | System | NotEngineering | ShouldbeCustomer | NotEngineering-ShouldbeCustomer | NotEngineering-ShouldbeCustomer | TBD-Translation | NotEngineering-ShouldbeCustomer | NotEngineering | NotEngineering-ShouldbeCustomer | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
552 |  |  | Puerto Rico | HOSPITAL DR SUSONI INC | 40013 | 44267 |  | Device time is 4 hours late | Work Order #: WO-03141725 Caldelas Aldana, Carlos Manuel Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 3 |  | Date and Time adjust | [12 empty] | 117182322 | In Progress | 699997 | TRUE | Closed - Done | Closed - Done | [7 empty] | Service | [6 empty] | Date and Time | [5 empty] | Refresh |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
553 |  |  | United States | St Joseph Westgate Medical Center | 40001 | 44267 |  | Trying to send to remote destination. | Work Order #: WO-03140982 Morrow, Dan The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Biomed Device Use: No Patient involved Expected and Unexpected Results: Adjust destination User Impact: none Patient Impact: none Current Software Version: 4.1.10 Diagnostic performed by Engineer : CT not free right now. Customer will call back. Customer called back and cannot log into olevel to adjust the lan config. Tried Philips1 and #Welcomehost123 passwords. Dispatching FSE to reset password and adjust LAN config. Follow Up Required : Cannot log into olevel to adjust the lan config. Tried Philips1 and #Welcomehost123 passwords. Dispatching FSE to reset password and adjust LAN config. ----------------------------------------------------------------------------- Work Order #: WO-03141606 15-Mar-2021 Leon Peter Problem Description by Engineer : Customer Function/Role: Biomed Device Use: No Patient involved Expected and Unexpected Results: Adjust destination User Impact: none Patient Impact: none Current Software Version: 4.1.10 Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 5 |  | reset password and adjust LAN config. | [12 empty] | 117181765 | In Progress | 696981 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Password | [5 empty] | Refresh |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
554 |  |  | China | JX Yugan Chudong hospital | 337500 | 44268 |  | wechat：02N/Ingenuity Core128/Elite China/system hung up | BU Q&R Feedback: No SAP Customer Symptom Code: HANG Customer Symptom Code: Hang/Crash Reason: Corrective Maintenance Engineer Failure Code: Wear-out/Breakage/Mechanical Failure Repair Activity Code: Reboot/Restart Problem Reported by Customer : system hung up Safety Questions : Case: 0117183430 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Qian Wang Last updated by: Qian Wang T2 Activities : Problem Description: The host computer crashes and the inmage transfer often fails. Troubleshooting Action: restart computer and reset spooler. Repair Action: guide customer to restart computer and reset spooler. Test/Inspection: customer confirm can scan patient and image ok.functional Check PASSED. Device returned to full functionality. Test Equipment: N/A SAP Transfer : Case Transferred to SAP | 3 |  | New Yellow Screen + PACS Transfer issue, collect bugreport for eng analysis | [12 empty] | 117183430 | Fixed | 698147 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00631367 | [5 empty] | Engineering | [6 empty] | New Yellow Screen | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
555 |  |  | Japan | Higashi-Ohmi General Medical C | 341015 | 44270 |  | This is enhancement request about Dose modulation in the range without Surview. | This is enhancement request about Dose modulation in the range without Surview. |  |  | Raise defect | [14 empty] | 701480 | TRUE | Investigation | Closed - Done |  | TBD-RFC | [5 empty] | Engineering | [6 empty] | RFC | [5 empty] | Refresh |  |  | System issue | System | Console-SW | RFC | Console-SW-RFC | Console-SW-RFC | RFC | RFC | Console-SW | Console-SW-RFC | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
556 |  |  | United States | St Jude Medical Center | 310165 | 44268 |  | 0609CDN84 Thumping Noise in Gantry | Work Order #: WO-03144902 Carpenter, Kevin The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Biomed Device Use: Diag Expected and Unexpected Results: Gantry Nosie User Impact: None Patient Impact: None Current Software Version: IDK Diagnostic performed by Engineer : No Diags Resolution : Cust called to cancel service for Monday ----------------------------------------------------------------------------- | 3 |  | Non-Refresh | [12 empty] | 117184394 |  | 700840 | #N/A | #N/A | Closed - No Further Investigation | [20 empty] | Non-Refresh |  |  | Upgrade issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | y
557 |  |  | United States | Lakeland Regional Medical Cent | 320474 | 44265 |  | Q/O: images are hanging up in the queue | Work Order #: WO-03127574 12-Mar-2021 Watson Adam The system fully meets specification and returned to use Resolution : After every scan, system would attempt to send to MPPS_Cerner the dose structured report automatically. The structured report is failing to send to MPPS_Cerner but sends normally to IDose. IP address, host, and port address of MPPS_Cerner check out correctly. Hospital IT is investigating why MPPS Cerner is not accepting the image transfer. CT is operating normally. Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: ----------------------------------------------------------------------------- | 3 |  | Dose structure report | [12 empty] | 117169819 |  | 693271 | #N/A | #N/A | Closed - No Further Investigation |  | TBD-DoseReport | [5 empty] | Engineering | [6 empty] | Dose structure report | [5 empty] | Refresh |  |  | Upgrade issue | System | Console-SW | Dose SR | Console-SW-Dose SR | Console-SW-Dose SR | TBD-DoseReport | Dose SR | Console-SW | Console-SW-Dose SR | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
558 |  |  | United Kingdom | FORTH VALLEY PC NHS TRUST | 68183 | 44266 |  | newly installed and gantry cannot comply, did a long tube recon but the system will not respond and is OOA | Work Order #: WO-03130290 Graeme Baxter The system does not meet specification and is out of use Diagnostic performed by Engineer : gantry can`t comply stopping them scan , i will download the logs and investigate From the logs i can see S_MASTER_SEQUENCER_TILT_POS_ERR we need to investigate the tilt circuitry Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Follow Up Required : gantry can`t comply stopping them scan , i will download the logs and investigate From the logs i can see S_MASTER_SEQUENCER_TILT_POS_ERR we need to investigate the tilt circuitry | 2 |  | CD-0000034558 | [12 empty] | 117172639 |  | 690595 | #N/A | #N/A | #N/A | [20 empty] | Non-Refresh |  |  | Upgrade issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | #N/A | y |  |  | y
559 |  |  | United States | Wellstar Kennestone Hospital | 68117 | 44265 |  | Gating not working | Work Order #: WO-03131478 13-Mar-2021 Lee John The system fully meets specification and returned to use Resolution : Reloaded ibox driver Problem Description by Engineer : Problem description by engineer : Customer states the Gating in not working. When was the first occurrence? Within... : 10-Mar-2021 How often is issue occurring? : Continuously Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: 3rd Party Device Use: Diagnostic Expected and Unexpected results: Gating not working Patient Impact/Outcome: can not have scan. User Impact: can not use Gating. Current Software Version: 4.1.7 ----------------------------------------------------------------------------- Work Order #: WO-03131743 10-Mar-2021 Parts Consumed : 1.00000,459800166651,I-BOX-2 BASE BOARD ASSY 1.00000,459801731761,USB EXTENDER CABLE FROM IBOX TO CONSOLE 2.00000,459801731761,USB EXTENDER CABLE FROM IBOX TO CONSOLE ----------------------------------------------------------------------------- Work Order #: WO-03125224 Miller, John M The system does not meet specification and is out of use Problem Description by Engineer : Problem description by engineer : Customer states the Gating in not working. When was the first occurrence? Within... : 10-Mar-2021 How often is issue occurring? : Continuously Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: 3rd Party Device Use: Diagnostic Expected and Unexpected results: Gating not working Patient Impact/Outcome: can not have scan. User Impact: can not use Gating. Current Software Version: 4.1.7 Follow Up Required : Error # and/or description of error : Gating not working Completed or Recommended Repair Action "Internal" : Suggest checking the cabling from the EKG monitor to the gantry. Verify the connections at the Usb hub and test the Ibox. 459801413791 New EXTenders USB, released and available!!! Material Number 459801731761 - USB EXTENDER CABLE FROM IBOX TO CONSOLE Material Number 459800166652 - I-Box-2 Base board assy Internal Comments : Error # and/or description of error : Gating not working Completed or Recommended Repair Action "Internal" : Customer opened call for onsite local FSE_s Suggest checking the cabling from the EKG monitor to the gantry. Verify the connections at the Usb hub and test the Ibox. 459801413791 New EXTenders USB, released and available!!! Material Number 459801731761 - USB EXTENDER CABLE FROM IBOX TO CONSOLE Material Number 459800166652 - I-Box-2 Base board assy Diagnostic performed by Engineer : Collect information from customer. ----------------------------------------------------------------------------- Work Order #: WO-03125326 10-Mar-2021 Lee John The system partly meets specification and is returned to use with restrictions Follow Up Required : Replace ibox Problem Description by Engineer : Problem description by engineer : Customer states the Gating in not working. When was the first occurrence? Within... : 10-Mar-2021 How often is issue occurring? : Continuously Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: 3rd Party Device Use: Diagnostic Expected and Unexpected results: Gating not working Patient Impact/Outcome: can not have scan. User Impact: can not use Gating. Current Software Version: 4.1.7 Diagnostic performed by Engineer : Troubleshot ibox ----------------------------------------------------------------------------- | 2 |  | Non-Refresh | [12 empty] | 117167986 |  | 687642 | #N/A | #N/A | #N/A | [20 empty] | Non-Refresh |  |  | Upgrade issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | #N/A | y |  |  | y
560 |  |  | United States | Lakeland Regional Medical Cent | 320474 | 44265 | 11 | 0711CFXXX Cant Login to System | Work Order #: WO-03120963 Maier, Donnie The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: No Patient Involved Expected and Unexpected Results: Cant Login User Impact: No Patient Impact: No Current Software Version: 4.1.10 Diagnostic performed by Engineer : Emergency login would not allow CT Tech to login either. I had her go to Change Password and change to ctpwc1. Originally password was ctpwc. Cu updated their password sheet as well. Cu is able to login and scan patients normally now. Michelle will call back if she has any further issues. Resolution : I has cu select change password as she is no longer able to login with original password or emergency login. Cu updated their password on their password sheet as well. | 2 |  | Password management, enhance application training for upgrade site, highlight the difference between 4.1.7 (formal release) to customer. | [12 empty] | 117163284 |  | 685078 | #N/A | #N/A | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Password | [5 empty] | Refresh |  |  | Upgrade issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
561 |  |  | Netherlands | Rijnstate Ziekenhuis Arnhem | 320103 | 44264 |  | Mattress of table is worn | Work Order #: WO-03119612 EDDY ELFRING The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer: Quote created to order mattress. ----------------------------------------------------------------------------- | 5 |  | Non-Refresh | [12 empty] | 117162315 |  | 684399 | #N/A | #N/A | Closed - No Further Investigation | [20 empty] | Non-Refresh |  |  | Upgrade issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | y
562 |  |  | United States | NORTHWEST RADIOLOGY/MRI | 300057 | 44264 |  | 0304CDY72 Intermittent gantry shutoff | Problem Description by Engineer : Problem description by engineer : Intermittent Estop open. Happened once this morning and once last week. Reclose Estops to clear issue. Bug report made. Follow Up Required : What are the next Troubleshooting steps : ACS / AMC Watchdog Time Out. The ACS /AMC has stopped responding to the Couch Control Board. Communication between CCB and ACS / AMC as watched by the CCB Couch Carbon Top Register bit R/Motion Controller Live is 1. This can also be observed by LED 30 on the CCB which should blink every two seconds when the communication between CCB and ACS / AMC is up. Recommendation: 1. Check CCB LED 30, it should flash every 30 Sec. 2. Monitor the ACS 7 segment displays for the following sequence: a) As soon as power is applied to the gantry (blue button) the ACS should read 8 (this is a power on self test of the ACS) for about 8 seconds. b) This should proceed to a flashing 1 (standard table), 3 (bariatric table) or 4 (extended table) while Estop is open. c) Once Estop is Closed the prior flashing number (1,3 or 4) should proceed to a sold number (1,3 or 4). d) If the ACS stays on 8 or never shows any following numbers, it did not pass It is power on self test. Check It is power feeds and connections per the below documents before replacing. 3. Read the Couch theory to understand the switch ON procedure. 4. Follow the couch troubleshooting document 5. Replace ACS / AMC or CCB. 6.. Check / replace ACS / AMC to CCB cable 459800623445 - FRU, AMC MOTOR CONTROL MODULE Error # and/or description of error : S_HMC_ACS_WATCHDOG_FAIL S_HMC_CTRL_A_MOTOR_SHORT_CIRCUIT_OR_OVERTEMP Completed or Recommended Repair Action "Internal" : Reviewed logs. Issue with horizontal control watchdog. Dispatching FSE to check horizontal control Diagnostic performed by Engineer : Reviewed logs. Issue with horizontal control watchdog. Dispatching FSE to check horizontal control. Internal Comments : Error # and/or description of error : S_HMC_ACS_WATCHDOG_FAIL S_HMC_CTRL_A_MOTOR_SHORT_CIRCUIT_OR_OVERTEMP Completed or Recommended Repair Action "Internal" : Reviewed logs. Issue with horizontal control watchdog. Dispatching FSE to check horizontal control ----------------------------------------------------------------------------- Work Order #: WO-03115832 10-Mar-2021 Williams John The system fully meets specification and returned to use Problem Description by Engineer : Problem description by engineer : Intermittent Estop open. Happened once this morning and once last week. Reclose Estops to clear issue. Bug report made. When was the first occurrence? Within... : 01-Mar-2021 How often is issue occurring? : Multiple times per week Malfunction area : Table Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: Estop opened Patient Impact/Outcome: none User Impact: Reclose Estops Current Software Version: 4.1.7.25026 Resolution : Reseat all couch cables on AMC and CCC ----------------------------------------------------------------------------- | 4 |  | Reseat all couch cables on AMC and CCC | [12 empty] | 117159422 |  | 682490 | #N/A | #N/A | Closed - Done |  | TBD-AMC | [5 empty] | Engineering | [6 empty] | AMC/CCB | [5 empty] | Non-Refresh |  |  | Upgrade issue | System | Couch-HW | AMC | Couch-HW-AMC | Couch-HW-AMC | TBD-AMC | AMC | Couch-HW | Couch-HW-AMC | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | y
563 |  |  | Chile | HOSPITAL BARROS LUCO TRUDEAU | 341005 | 44274 |  | CL1688- PROBLEMA CON EL SISTEMA DE ENVIÓ DE IMÁGENES | "Work Order #: WO-03180947 Valenzuela Velasquez, Eduardo Ariel The system partly meets specification and is returned to use with restrictions Internal Comments: Error # and / or description of error: Some studies do not reach the PACS in full. Completed or Recommended Repair Action "" Internal "": - Review the client_s data network that does not present interruptions. - Review communication with the PACS. - Carry out tests for sending images to the PACS. Diagnostic performed by Engineer: Intermittent Failure - Some exams do not fully reach the PACS. Communication tests between the equipment and the PACS are performed without problems. On-site review is required to perform shipping tests and test the data network. Problem Description by Engineer: Problem description by engineer: There are problems in sending images to the PACS. Information to support the complaint handling process Customer Function / Role: Technician / Sonographer Device Use: Diagnostic Expected and Unexpected results: No further information is available as per the customer Patient Impact / Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer Current Software Version: 4.1.10 Follow Up Required: Error # and / or description of error: Some studies do not reach the PACS in full. Completed or Recommended Repair Action "" Internal "": - Review the client_s data network that does not present interruptions. - Review communication with the PACS. - Carry out tests for sending images to the PACS. | 4 |  | Need check if HF3 FCO72800740 implemented or not | [12 empty] | 117218929 | In Progress | 723096 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS New? | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
564 |  |  | China | QH Haibei Zhou No. 2 | 345029 | 44274 |  | PRS Connection/02N/Ingenuity Core128/Elite China | BU Q&R Feedback: No Reason: Customer Relations Problem Reported by Customer : PRS Connection Next Step Action : FSE 许永华 call in and disp him. Safety Questions : Case: 0117215360 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Yu Zhao Last updated by: Yu Zhao 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support ----------------------------------------------------------------------------- | 5 |  | Not a compliant | [12 empty] | 117215360 | New | 721872 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Others | [6 empty] | Not a compliant | [5 empty] | Refresh |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
565 |  |  | China | JS No.2 Chinese Medicine Hosp. | 345027 | 44274 |  | Wechat:02N/Ingenuity Core128/Elite China/Preheating alarm | BU Q&R Feedback: No Reason: Corrective Maintenance T2 Activities : Problem Description: 客户问题概述： 预热失败。 客户问题详细描述或故障现象： 做预热时扫描失败 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： max arc reached 远程诊断与测试结果： 设备有打火 PRS TR状态： 不通，请工程师到场检查。 Repair Action: 建议现场维修步骤： 根据报错日志查看打火原因，并做相关troubleshooting Problem Reported by Customer : Preheating alarm Safety Questions : Case: 0117214413 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Bai Jingwen Last updated by: Bai Jingwen 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support ----------------------------------------------------------------------------- | 3 |  | LTC performed and keep monitoring | [12 empty] | 117214413 | In Process | 721553 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00630085 | [5 empty] | Engineering | [6 empty] | Tube Arc | [5 empty] | Refresh |  |  | System issue | System | IC-HW | Tub Arc | IC-HW-Tub Arc | IC-HW-Tub Arc | CTUDT00630085 | Tub Arc | IC-HW | IC-HW-Tub Arc | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
566 |  |  | China | GD Dongguan Dalingshan Hosp. | 345013 | 44274 |  | 02N/Core128/no worklist and Cannot transfer image | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Core128/no worklist and Cannot transfer image SAP Transfer : Dispatched for Onsite Support Safety Questions : Case: 0117216532 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Bo Last updated by: Wang Bo 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 其他类型故障。 客户问题详细描述或故障现象： 主机无法传输图像和打印胶片。 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： log无报错。 远程诊断与测试结果： 重启无效。已和客户确认网络和reset spooler，已重启。 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议检查网卡设置，检查pacs及相机参数设置。 建议备件12NC及名称： N/A SAP Transfer : Case Transferred to SAP Next Step Action : HUang shaoxiang Next Step Action : 转zhuguanglu External Remarks : 20.03.2021 05:02:29 UTC Guanglu Zhu (86034061)服务内容: 主机网络故障，经检测确认为ip地址与另外一台设备冲突，更改设备IP地址，完成相应备份 服务结果: 服务完成后,设备恢复正常工作 Internal Remarks : 22.03.2021 12:40:48 UTC Guanglu Zhu (86034061) $ RDI Accuracy: yes $Q7: ( Confirmed reported problem+host can not transfer image and worklist update) $Q8: (analysis network and test found ip address collide with one of the dicom printer ) $Q9: (adjust dicom printer of collision IP address ,and test network function it return to normal ) $Q10: (Test and Inspection data refer to PA & Device status return to fully function) $Q11: Calibrated Tools name (na ) S/N (na ) Next Calibration date (na ) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date: na ^Current Tube exposure:NA ----------------------------------------------------------------------------- | 3 |  | IP Conflict, system works well after reconfiguration | [12 empty] | 117216532 | In Process | 718680 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Service | [6 empty] | IP Config | [5 empty] | Refresh |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
567 |  |  | China | Wuxi Jiangyuan Hosp. | 345020 | 44273 |  | 02N/Ingenuity Core128/Elite China/ups failed | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: ups failed Safety Questions:Case: 0117207840 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Jianqiang Last updated by: Wang Jianqiang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer_s problems: UPS can not start. Detailed description of Customer_s problems or malfunctions: UPS can not start. It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: UPS can not start. Remote troubleshoot and test result: UPS can not start. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Check 亻 power supply, check ups SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support ----------------------------------------------------------------------------- | 3 |  | 服务内容： UPS无法启动，更换UPS主机和电池 服务结果： 维修后，设备工作正常。 Check part # that replaced on field and add in FRU cost list | [12 empty] | 117207840 | In Process | 717716 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD_UPS | [5 empty] | Engineering | [6 empty] | UPS | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | UPS | Gantry-HW-UPS | Gantry-HW-UPS | TBD_UPS | UPS | Gantry-HW | Gantry-HW-UPS | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
568 |  |  | China | Wuxi Jiangyuan Hosp. | 345020 | 44273 |  | 02N/Ingenuity Core128/Elite China/UPS Fault | BU Q&R Feedback: No SAP Customer Symptom Code: BOOT Customer Symptom Code: Boot up/Power-up Problem Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: Reboot/Restart Problem Reported by Customer : UPS Fault Safety Questions : Case: 0117207723 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Mengen Xie Last updated by: Mengen Xie 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A FSE Notes : This PO is a Reference Number only T2 Activities : Problem Description: UPS fault Troubleshooting Action: repeated with case 0117207840 Repair Action: repeated with case 0117207840 Test/Inspection: N/A Test Equipment: N/A ----------------------------------------------------------------------------- | 3 |  | repeated with case 0117207840 | [12 empty] | 117207723 | Fixed | 717665 | TRUE | Closed - Duplicate | Closed - Duplicate | [7 empty] | Others | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
569 |  |  | China | SD Tai_an T.C.M. Hosp. No.2 Ho | 345037 | 44272 |  | 02N/Ingenuity Core128/The head cushion was loss | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : The head cushion was loss SAP Transfer : Case Transferred to SAP Safety Questions : Case: 0117201402 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Weng Jinming Last updated by: Weng Jinming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 5 |  | The head cushion was loss | [12 empty] | 117201402 | Closed | 713743 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Head cushion | [5 empty] | Refresh |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
570 |  |  | China | JS Nanjing U.A. Zhongda Hosp. | 345021 | 44272 |  | Wechat:02N/Ingenuity Core128/Elite China/Time fault | BU Q&R Feedback: No SAP Customer Symptom Code: HANG Customer Symptom Code: Hang/Crash Reason: Corrective Maintenance Engineer Failure Code: Configuration Error Repair Activity Code: Configuration Change Problem Reported by Customer : Time fault Safety Questions : Case: 0117200007 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Bai Jingwen Last updated by: Bai Jingwen 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP T2 Activities : Problem Description: Modify system time Troubleshooting Action: Help customers enter the system and modify the system time Repair Action: Help customers enter the system to modify system time and guide by telephone Test/Inspection: customer confirm can scan patient and image ok.functional Check PASSED. Device returned to full functionality. Test Equipment: N/A ----------------------------------------------------------------------------- | Zhai, Lisa | [14 empty] | 117200007 | Closed | 712979 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [20 empty] | Refresh |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
571 |  |  | China | Dalian Branch of Shengjing Hospital of China Medical University | 345001 | 44270 |  | 02N/Ingenuity Core128/Elite China/Add a node | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: Add a node Safety Questions:Case: 0117186968 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Mengen Xie Last updated by: Mengen Xie 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer_s problems: CT machine needs to add peripherals It is a technology problem with the equipment itself: Yes. Troubleshooting Action: remote diagnosis and testing result: We need an engineer to go on site to assist. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Need an engineer to cooperate on the field. Teach the customer how to use Philips WeChat to report for repair SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support External Remarks: 16.03.2021 04:06:47 UTC SUP JCOUSER (86038273) service details: help the hospital to add nodes. service result: device returns to normal working condition after service Internal Remarks: 16.03.2021 04:08:54 UTC Zhilong Yu (86038273) RDI Accuracy:Yes $Q7: (Confirmed reported problem: isp need add node) $Q8: (isp need add node) $Q9:(add node) $Q10: (NA) ----------------------------------------------------------------------------- | Lu, Xueying | [14 empty] | 117186968 | Closed | 711440 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [20 empty] | Refresh |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
572 |  |  | China | SD Laiyang People Hospital | 345005 | 44270 |  | 02N/Ingenuity Core128/Elite China/Connect a wire | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : Connect a wire Safety Questions:Case: 0117187124 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Guomei Li Last updated by: Guomei Li 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer_s problems: Other types of error. Detailed description of Customer_s problems or malfunctions: The hospital needs to connect the gate lock and add a PACS node. It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: The log has no errors. Remote troubleshoot and test result: Restart but didn_t work.We tried to resolve it remotely. The customer refused and asked to have a field service engineer go to the site to handle it. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: It is recommended to help the customer to connect the gate lock and add a PACS node. We recommend spare parts 12NC and names: N/A SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support External Remarks : 17.03.2021 04:49:48 UTC Rui Wang (86090282) Service Details: Add a PACS node, install the gate switch Service outcome: It was added successfully; delivered to the hospital for use. Internal Remarks : 17.03.2021 04:57:27 UTC Rui Wang (86090282) "RDI Accuracy:Yes" $Q7: (Confirmed reported problem & Mounting door switch） $Q8: (check error，need door switch，Door switches need to be installed) $Q9: ( install the door switch ) $Q10:(Test and inspection data refer to: PA & Device returned to full functionality) $Q11: Calibrated Tools name ( N/A) ----------------------------------------------------------------------------- | 5 |  | Customer request to add a PACS node, install the gate switch | [12 empty] | 117187124 | Fixed | 704602 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Service | [6 empty] | PACS Config+Door switch install | [5 empty] | Refresh |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
573 |  |  | China | QH Haibei Zhou No. 2 | 345029 | 44270 |  | 02N/Core128/Scan abort intermittent | T2 Activities : Problem Description: 客户问题概述： 计算机软件相关故障。 客户问题详细描述或故障现象： 主机软件扫描时报错。 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： A problem has occurred，please retry。无PRS。 远程诊断与测试结果： 重启无效。 PRS TR状态： 无PRS。 Repair Action: 建议现场维修步骤： 建议根据报错，排查软件故障。如果不是机架报错建议收集bugreport，反馈给TSE。 建议备件12NC及名称： N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : xu yonghua Next Step Action : 已短信客户设备信息，以便后续报修 External Remarks : 22.03.2021 03:54:22 UTC Yonghua Xu (86001334)科室：放射科CT室 联系人：郭荣寿 联系电话：15309708576 故障现象：设备不能曝光。 解决方案/维修措施：设备不能曝光。HOST错误日志提示旋转主控板Rhost没有收到曝光信号，检测DMC到旋转主控板Rhost通讯线路正常 。替换旋转主控板Rhost和DMC故障依旧。判断高压系统控制板或者高压阴极模块故障。更换高压系统控制板故障依旧，进而更换高压阴极模块。 完成相应校准，完成备份。 服务后设备正常工作。 Internal Remarks : 22.03.2021 03:54:25 UTC Yonghua Xu (86001334) “RDI Accuracy: No” $Q7:Confirmation reported problem,The device cannot be exposed. $Q8:The device cannot be exposed.The HOST error log prompts S_MASTER_SEQUENCER_SHOT_OUT_OF_SEQUENCE_OR_ERR.P1=64. RHOST did not detect Xray ON.Check that the communication line from DMC to Rhost is normal.The replacement Rhost and DMC still failed.Determine failure of HV system control board or HV cathode power module. $Q9:Replace the HV system board panel failure is still, and then replace the HV cathode power module. $Q10:Test and Inspection data refer to PA , Do a few patient, Observation of a certain period of time. The system is ok. Device returned to full functionality. $Q11:Calibrated tools name refer to PA. ----------------------------------------------------------------------------- | 3 |  | 替换旋转主控板Rhost和DMC故障依旧。判断高压系统控制板或者高压阴极模块故障。更换高压系统控制板故障依旧，进而更换高压阴极模块。 完成相应校准，完成备份。 Check parts replaced on field and update FRU cost list | [12 empty] | 117186873 | In Process | 704386 | TRUE | Closed - Done | Closed - Done |  | TBD-APM/CPM | [5 empty] | Engineering | [6 empty] | CPM | [5 empty] | Refresh |  |  | System issue | System | IC-HW | APM/CPM | IC-HW-APM/CPM | IC-HW-APM/CPM | TBD-APM/CPM | APM/CPM | IC-HW | IC-HW-APM/CPM | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
574 |  |  | China | JX Province Children Hosp. | 40041 | 44270 |  | 02G/Ingenuity CT 728326/Password error | BU Q&R Feedback: No Reason: Customer Information Safety Questions:Case: 0117186120 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Zixuan Last updated by: Wang Zixuan 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Problem Reported by Customer : Password error Next Step Action: 019979813233CALLED in, xie sheng follow ----------------------------------------------------------------------------- | 3 |  | Incorrect password | [12 empty] | 117186120 | Closed | 704026 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Password | [5 empty] | Refresh |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
575 |  |  | Germany | Gesundheit Nord gGmbh | 40002 | 44272 |  | An error message has occurred. | Work Order #: WO-03165609 18-Mar-2021 Arno weeps The system fully meets specification and returned to use Problem Description by Engineer: Customer reports problem, system cannot be reached remotely Resolution: Error analysis carried out. Slip ring and breakage cleaned. Function checked. Tools & Equipment: ELECTRICAL SAFETY TESTER, 2021-05-08 00:00:00 | 3 |  | Error analysis carried out. Slip ring and breakage cleaned. Function checked. | [12 empty] | 117203874 | Closed | 708724 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | NotEngineering-SlipringClean | [5 empty] | Engineering | [6 empty] | Slip ring clean | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | SlipRingClean | Gantry-HW-SlipRingClean | Gantry-HW-SlipRingClean | NotEngineering-SlipringClean | SlipRingClean | Gantry-HW | Gantry-HW-SlipRingClean | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
576 |  |  | Japan | Saiseikai Gosu Hospital | 343000 | 44277 |  | 01CS:KKI:昨日、ShortTCが走らなかった＞＞CTホスト再起動にて復旧＞Scan Failed due to error: Gantry cannot initialize. Please change scan... | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 01CS:KKI:昨日、ShortTCが走らなかった＞＞CTホスト再起動にて復旧＞Scan Failed due to error: Gantry cannot initialize. Please change scan parameters or restart the Gantry. If problem persists, please contact Service. Safety Questions : Case: 0117222824 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: KAZUMASA KIMURA Last updated by: KAZUMASA KIMURA | 4 |  | FSE handling, check if have bugreport available, collect bugreport | [12 empty] | 117222824 | In Process | 726860 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-Translation | [5 empty] | Engineering | [6 empty] | In progress | [5 empty] | Refresh |  |  | System issue | System | TBD-Bugreport | TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | TBD-Translation | TBD_Bugreport | TBD-Bugreport | TBD-Bugreport-TBD_Bugreport | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
577 |  |  | Japan | Nikko Memorial Hospital | 341011 | 44274 |  | I would like to contact FSE regarding the top plate mattress | Problem Reported by Customer : 天板マットレスの件でFSEと連絡希望 Safety Questions : Case: 0117215869 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: RUI SATAKE Last updated by: RUI SATAKE | 5 |  | Q&A? Check if question only or issue reported from customer | [12 empty] | 117215869 | In Process | 722076 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | TBD | [6 empty] | NAC? | [5 empty] | Refresh |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
578 |  |  | Japan | Saiseikai Gosu Hospital | 343000 | 44273 |  | The gantry makes a strange noise. | BU Q&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Dispatched for Onsite Support Problem Reported by Customer: SC: The gantry makes a strange noise Safety Questions:Case: 0117206921 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: CAI Susie Last updated by: CAI Susie 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A OneEMS internal remarks : Kobayashi DB SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | FSE handling, check if have bugreport available | [12 empty] | 117206921 | New | 717207 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD_Bugreport | [5 empty] | Engineering | [6 empty] | In progress | [5 empty] | Refresh |  |  | System issue | System | TBD-Bugreport | TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | TBD_Bugreport | TBD_Bugreport | TBD-Bugreport | TBD-Bugreport-TBD_Bugreport | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
579 |  |  | Japan | Higashi-Ohmi General Medical C | 341015 | 44273 |  | Applications inquiry. >> Nakagawa will respond. | App inquiry >> Correspondence to Mr. Nakagawa Safety Questions: Case: 0117208856 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: TAKEYUKI MIYAKAWA Last updated by: TAKEYUKI MIYAKAWA 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm / alert capability, did it alarm / alert as it should have at the time the issue was discovered? N / A 4. Was this an out of box failure? N / A T2 Activities: Problem Description: Problem description by engineer: Target device: CT Failure location: 07 Others When did it first occur? (Within ~): Unknow How often do problems occur ?: Unknow At what point in the end-user workflow is the problem occurring ?: Other What happened recently ?: N.A. Does the customer create a specific system event report for the problem ?: N.A. Engineer_s Comment: _============================== ① If you send the next Pt before end_exam, two Pt will be displayed in the schedule. Is there a setting to delete the previous Pt information after scanning? ② I took a picture of Dental. Reconstruction was performed because there was a request to create 3D, but the soft tissue condition function could not be selected. (3) Partial transfer is performed using sub-selection. When sub-selection is selected, I think the list is blue, but today some of them are not selected in white, and only one is selected in blue. ④ Too_hot appears when multi-layered trunk photography is performed. As you told me, if you change the shooting conditions and press GO, a message such as "Wait for XX seconds" will appear and you will not be able to change the conditions. Should I wait? (5) Contrast-enhanced photography was performed with BT, but because it leaked, both Injection and Scan were stopped. Do I have to set up BT from the beginning? ⑥ There are two electronic medical records, but information was sent from an electronic medical record that is not dedicated to CT. ⑦ When will the mat be delivered? ⑧ The next web_meeting is closed on Tuesday, so 26 (Friday) may be good. _============================== Troubleshooting Action: What steps are taken during the system diagnosis: Post-response results: Improved Failure location: 07 Others What steps were taken during the diagnosis ?: _================================== Correspondence ① Basically such specifications. We asked FSE Takada-san to take action. (2) The items that can be selected are narrowed down by Scan_Type. By changing to Sinus or Face, you will be able to select soft tissue functions. Please create it by saving it as a different name. ③ I don_t think it will take over the previous operation. I will investigate. ④ You can set the conditions again by pressing the pause button. If you wait too long, please reset it. ⑤ Although it is a small amount, the timing may change because the contrast medium has been applied to the body once. For that reason, the timing is reset once. Sorry to trouble you, but please reset it. ⑥ I think it is the setting on the electronic medical record side. Please contact your manufacturer. ⑦ We asked Mr. FSE Takada to answer. ⑧ Otani will contact you. _================================== Error # and / or description of error: NA Repair Action: Completed or Recommended Repair Action "" External "": [Phenomenon confirmation] Verified [Details of work] Answer the question [Confirmation details] Good OneEMS internal remarks: Problem description by engineer: Target device: CT Failure location: 07 Others When did it first occur? (Within ~): Unknow How often do problems occur ?: Unknow At what point in the end-user workflow is the problem occurring ?: Other What happened recently ?: N.A. Does the customer create a specific system event report for the problem ?: N.A. Engineer_s Comment: _============================== ① If you send the next Pt before end_exam, two Pt will be displayed in the schedule. Is there a setting to delete the previous Pt information after scanning? ② I took a picture of Dental. Reconstruction was performed because there was a request to create 3D, but the soft tissue condition function could not be selected. (3) Partial transfer is performed using sub-selection. When sub-selection is selected, I think the list is blue, but today some of them are not selected in white, and only one is selected in blue. ④ Too_hot appears when multi-layered trunk photography is performed. As you told me, if you change the shooting conditions and press GO, a message such as "Wait for XX seconds" will appear and you will not be able to change the conditions. Should I wait? (5) Contrast-enhanced photography was performed with BT, but because it leaked, both Injection and Scan were stopped. Do I have to set up BT from the beginning? ⑥ There are two electronic medical records, but information was sent from an electronic medical record that is not dedicated to CT. ⑦ When will the mat be delivered? ⑧ The next web_meeting is closed on Tuesday, so 26 (Friday) may be good. _============================== What steps are taken during the system diagnosis: Post-response results: Improved Failure location: 07 Others What steps were taken during the diagnosis ?: _================================== Correspondence ① Basically such specifications. We asked FSE Takada-san to take action. (2) The items that can be selected are narrowed down by Scan_Type. By changing to Sinus or Face, you will be able to select soft tissue functions. Please create it by saving it as a different name. ③ I don_t think it will take over the previous operation. I will investigate. ④ You can set the conditions again by pressing the pause button. If you wait too long, please reset it. ⑤ Although it is a small amount, the timing may change because the contrast medium has been applied to the body once. For that reason, the timing is reset once. Sorry to trouble you, but please reset it. ⑥ I think it is the setting on the electronic medical record side. Please contact your manufacturer. ⑦ We asked Mr. FSE Takada to answer. ⑧ Otani will contact you. _================================== Error # and / or description of error: NA Completed or Recommended Repair Action "" External "": [Phenomenon confirmation] Verified [Details of work] Answer the question [Confirmation details] Good | 5 |  | Q&A | [12 empty] | 117208856 | Fixed | 713740 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [14 empty] | NAC? | [5 empty] | Refresh |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
580 |  |  | Japan | Saiseikai Gosu Hospital | 343000 | 44273 |  | 01CS: KKI: ExamSummary was not created after shooting | T2 Activities: Problem Description: Problem description by engineer: Target device: CT Failure location: 01Cosole When did it first occur? (Within ~): 1 Hour How often do problems occur ?: First Occurrence At what point in the end-user workflow is the problem occurring ?: After inspection / measurement (Post Acquistion / Scan) What happened recently ?: N.A. Does the customer create a specific system event report for the problem ?: No. The customer was unable to create a log file or report. | 4 |  | Exam Summary was not created after scan, wait FSE to handling onsite @Kevin, double check if have spec to enable or disable ES | [12 empty] | 117207141 | In Process | 712960 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-ExamSummary | [5 empty] | Engineering | [6 empty] | Exam Summary | [5 empty] | Refresh |  |  | System issue | System | Console-SW | ExamSummary | Console-SW-ExamSummary | Console-SW-ExamSummary | TBD-ExamSummary | ExamSummary | Console-SW | Console-SW-ExamSummary | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
581 |  |  | Japan | Higashi-Ohmi General Medical C | 341015 | 44272 |  | 01CS: Tm: Hisaki: CT-Host freezes >> | $ version = "" v4.1.10.25208 "" 10:43 Incoming call Call content I can not press the CT-Host button. All icons such as New patient on the left side of the screen are blacked out and cannot be pressed. Switching screens with a wrench mark does not open when trying to open ExamCard Manager Logout / Login is done, but it does not start with a blue screen. _============================== What steps are taken during the system diagnosis: Post-response results: Improved Failure location: 01Cosole What steps were taken during the diagnosis ?: _================================== Correspondence I received Sign Out-> Restart with Ctrl + alt + del. After that, it works normally. Log confirmation result No noticeable error. It is probable that the CT-Host was a temporary malfunction and was resolved by restarting the computer. | 4 |  | HF2 Yellow screen and then blue screen happened after restart, need bugreport for analysis, @Peter @Lilian, Silva, reproduce on test bay | [12 empty] | 117200243 | Fixed | 712300 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632520 | [5 empty] | Engineering | [6 empty] | Yellow Screen | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | CTUDT00631367 | y | y
582 |  |  | Japan | Saiseikai Gosu Hospital | 343000 | 44272 |  | 01CS:RS: The bug report window popped up when he clicked "Start final recon" after a scan. >> The reconstruc [TRUNCATED] | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 01CS:RS: The bug report window popped up when he clicked "Start final recon" after a scan. >> The reconstruction was performed without a problem. >> | 4 |  | Bugreport error message popup, need bugreport for analysis, @Peter | [12 empty] | 117200886 | In Process | 712065 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-BugRepError | [5 empty] | Engineering | [6 empty] | Bugreport error | [5 empty] | Refresh |  |  | System issue | System | Console-SW | BugRepError | Console-SW-BugRepError | Console-SW-BugRepError | TBD-BugRepError | BugRepError | Console-SW | Console-SW-BugRepError | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
583 |  |  | Japan | Tomakomai City Hospital | 341014 | 44272 |  | 03CS:KK: Scan stopped partway through.>> Examinations were possible with a rescan.>> S_GEN_FAULT_ANODE_TUBE_ARC | BU Q&R Feedback: No SAP Customer Symptom Code: HANG Customer Symptom Code: Hang/Crash Reason: Corrective Maintenance Engineer Failure Code: Electrical/Electronic Failure Repair Activity Code: Instructed Customer Problem Reported by Customer: 03CS:KK: Scan stopped partway through.>> Examinations were possible with a rescan.>> S_GEN_FAULT_ANODE_TUBE_ARC T2 Activities:Problem Description: subject equipment: CT Failure location: 03High Voltage Energy When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user_s work flow did the problem occur?: (During Acquisition/Scan) What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem? No.The customer has not created a log file or report. Troubleshooting Action: Result after troubleshooting: Improved Failure location: 03High Voltage Energy What steps did you perform to diagnose the problem? An APM arc had occurred. Scanning possible with rescan. I had FSE Hashimoto contact the customer, and they will monitor the situation. From the logger, ============================================================ 2021/03/17 12:45:04.820 [START STUDY] GSE::CState::StateBkst::STUDY_STARTED was sent to GPC 2021/03/17 12:47:52.352 [---] ------ Series 124 Load Structure ----- 2021/03/17 12:48:08.398__02 [--o] ----------------- Series 124 Ready ----------------- 2021/03/17 12:48:11.543__01 [-oo] - - - - - - Shot 0 Starts - - - - - - 2021/03/17 12:48:17.031 [ooo] ----------------- Series 124 Complete ----------------- 2021/03/17 12:48:28.961__06 [---] ------ Series 125 Load Structure ----- 2021/03/17 12:48:36.808__02 [--o] ----------------- Series 125 Ready ----------------- 2021/03/17 12:48:45.201__01 [-oo] - - - - - - Shot 0 Starts - - - - - - 2021/03/17 12:48:50.586 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_ANALOG_FAULT szInfo:2:5 Analog Fault P1:2/P2:32 2021/03/17 12:48:50.631 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_ANODE_TUBE_ARC 2021/03/17 12:48:50.637__03 [CANCEL] CStateShot:Series 125 cancelled, due to GANTRY_CANCEL_SOURCE_NORMAL_ERROR, in state CStateShot 2021/03/17 12:48:50.642__02 [***] - - - - - - Study Exit - - - - - - 2021/03/17 12:48:50.643 <*HV*> Type: NORMAL ErrNum: S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT szInfo: P1:0/P2: 2021/03/17 12:48:50.654 <WARN> Type: WARN ErrNum: S_GEN_MGR_SYNC_OFF_OUT_OF_SEQUENCE P1:0/P2:0 2021/03/17 12:48:52.697__01 <*ERROR*> Type: NORMAL ErrNum: S_MASTER_SEQUENCER_EXIT_STATE_ERR P1:4096/P2:8369 2021/03/17 12:49:32.077__02 [---] ------ Series 126 Load Structure ----- 2021/03/17 12:49:36.148__02 [--o] ----------------- Series 126 Ready ----------------- 2021/03/17 12:49:43.271__01 [-oo] - - - - - - Shot 0 Starts - - - - - - 2021/03/17 12:49:52.366 [ooo] ----------------- Series 126 Complete ----------------- Repair Action: [Confirmation of problem] Confirmed. Work details: An APM arc had occurred. Scanning possible with rescan. I had FSE Hashimoto contact the customer, and they will monitor the situation. | 4 |  | APM, need to created a log file or report. | [12 empty] | 117200842 | Fixed | 707768 | TRUE | Closed - Done | Closed - Done |  | CTUDT00630085 | [5 empty] | Engineering | [6 empty] | HV Arc | [5 empty] | Refresh |  |  | System issue | System | IC-HW | Tub Arc | IC-HW-Tub Arc | IC-HW-Tub Arc | CTUDT00630085 | Tub Arc | IC-HW | IC-HW-Tub Arc | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
584 |  |  | Japan | KYOTO MEDICAL CENTER | 341013 | 44270 |  | CZ: The monitor sometimes does not display. | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: CZ: The monitor sometimes does not display. Safety Questions:Case: 0117186872 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: ZHAO Cindy Last updated by: ZHAO Cindy 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A OneEMS internal remarks: Yamamoto DB SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Support ----------------------------------------------------------------------------- | 5 |  | Probably because of DVI adaptor, recommand to replace 459801659061 FRU, DP to DVI-D adapter | [12 empty] | 117186872 | New | 704316 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632173 | [5 empty] | Engineering | [6 empty] | DVI adapter | [5 empty] | Refresh |  |  | System issue | System | Console-HW | Adaptor | Console-HW-Adaptor | Console-HW-Adaptor | CTUDT00632173 | Adaptor | Console-HW | Console-HW-Adaptor | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
585 |  |  | Japan | Higashi-Ohmi General Medical C | 341015 | 44270 |  | 04DS: TM: A strange sound, like a light brushing occurs when the gantry is rotating >> Response by FSE | There_s a light brushing sound when the gantry rotates. Problem confirmation: I confirmed it. I confirmed that a strange noise was occurring. Work details: Regarding the problem of the gantry make a strange sound when it rotates, I investigated the cause of the noise and found that the gantry cone cover was in contact with a portion of the gantry_s rotating section, and the sound was caused by this contact. I adjusted the position of the gantry cover, and that solved the problem of interference with the gantry_s rotating component. We apologize for any trouble caused. Other problems/Measures taken: None. Equipment status The equipment can be used. Work status Work finished. 16-Mar-2021 81100262 FSE Handover Tool1.3.4.32612.DB13 20210115 Internal Remarks : 18.03.2021 03:10:29 UTC MASAHIRO MIYAKE (81100262) Replacement parts: No replacement parts. [Report on mistakes in parts order] No mistakes. Situation monitoring information: Not monitoring the situation TS Request: No parts to return. Additional CT information: Q1. How did the examination finish? Q1. No (not applicable) Q2. Were there any mistaken findings (misdiagnosis, mistaken treatment) due to the problem? None. Q3. Were there any artifacts? None. Q4. Was there any unintended unnecessary exposure? None. Repair sticker: I affixed a Repaired sticker. [Other information to share] 1. Was treatment/quarantine carried out for a patient who has or is suspected of having new coronavirus? No 2. Is it necessary to go into treatment rooms/beds where there are infected patients?(Including workplace and routes of movement) No Did you agree on how to complete the report with the person in charge of devices? Yes. I completed normal reporting that day. 16-Mar-2021 81100262 FSE Handover Tool1.3.4.32612.DB13 20210115 ----------------------------------------------------------------------------- | 5 |  | the gantry cone cover was in contact with a portion of the gantry_s rotating section, and the sound was caused by this contact. I adjusted the position of the gantry cover, and that solved the problem of interference with the gantry_s rotating component. | [12 empty] | 117192668 | In Process | 703716 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-GantryNoise | [5 empty] | Engineering | [6 empty] | Gantry Noise | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Noise | Gantry-HW-Noise | Gantry-HW-Noise | TBD-GantryNoise | Noise | Gantry-HW | Gantry-HW-Noise | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
586 |  |  | Poland | Centrum Opieki Medycznej | 341002 | 44270 |  | The hospital starts a worklist. | "Work Order #: WO-03151855 Matthew Ramadan The system fully meets specification and returned to use Problem Description by Engineer: Worklist configuration according to the indications of an IT specialist: IP = 11/10/19.2 AET = WLING341002 PORT = 6000. After configuring according to these parameters, the communication with the worklist looks correct, but the answer to the worklist_s query looks empty. For further analysis with an IT specialist. Diagnostic performed by Engineer: Worklist configuration according to IT indications: IP = 11/10/19.2 AET = WLING341002 PORT = 6000. After configuring according to these parameters, the communication with the worklist looks correct, but the answer to the worklist_s query looks empty. For further analysis with an IT specialist. " | 5 |  | Worklist configuration according to IT indications | [12 empty] | 117190946 | Closed | 702800 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Service | [6 empty] | Worklist config | [5 empty] | Refresh |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
587 |  |  | Puerto Rico | HOSPITAL DR SUSONI INC | 40013 | 44264 |  | Reconstructions are very slow | PR684086 split | [17 empty] | 724089 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-ReconSlow | [5 empty] | Engineering | [6 empty] | Recon Slow | [5 empty] | Refresh |  |  | System issue | System | CIRS-SW | ReconSlow | CIRS-SW-ReconSlow | CIRS-SW-ReconSlow | TBD-ReconSlow | ReconSlow | CIRS-SW | CIRS-SW-ReconSlow | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
588 |  |  | Puerto Rico | HOSPITAL DR SUSONI INC | 40013 | 44272 |  | Request support to program printer in the device | Work Order #: WO-03165610 Caldelas Aldana, Carlos Manuel Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 3 |  | Add printer | [12 empty] | 117204317 | In Progress | 715516 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Service | [6 empty] | Add printer | [5 empty] | Refresh |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
589 |  |  | Puerto Rico | HOSPITAL DR SUSONI INC | 40013 | 44272 | 11 | Gantry is turning off | Work Order #: WO-03166247 Caldelas Aldana, Carlos Manuel The system partly meets specification and is returned to use with restrictions Internal Comments : Error # and/or description of error : S_CTBOX_UNINITIATED_RESET S_LIFESYNCMGRTASK_LIFESYNC_FAIL / P1=2 Completed or Recommended Repair Action "Internal" : N/A Diagnostic performed by Engineer : validated that suddenly the ctbox appears like power on and off and the estop opens. Validated with remote conection that logfiles shows a lot of recurrences of initialization issues in the CTBOX. Aditional, some warnings about variation in vertical position of the patient suppor detected. Problem Description by Engineer : Problem description by engineer : estop opens When was the first occurrence? Within... : 16-Mar-2021 19:00:00 How often is issue occurring? : Intermittent Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: estop opens intermitent Patient Impact/Outcome: aborted study User Impact: gantry shutsdown Current Software Version: 4.1.10 Follow Up Required : What are the next Troubleshooting steps : Check CTBOX conections to SDB and CRC Rack. IF everything ok, replace CTBOX. Clean the complete metal string pot (vertical pot. encoder) in the patient support to avoid the recurence of lectures of position variating during scan. Check FRC and power conections to RHost Controller. During visit, check PRS configuration, FTP is not working RSN nor M2M, maybe SFTP is not enabled in firewall settings (Check remote configuration manual). Error # and/or description of error : S_CTBOX_UNINITIATED_RESET S_LIFESYNCMGRTASK_LIFESYNC_FAIL / P1=2 Completed or Recommended Repair Action "Internal" : N/A ----------------------------------------------------------------------------- | 2 |  | CT-Box intermittent lost power, CT-Box was replaced. 459801705022,CT BOX ASSEMBLY_BEST | [12 empty] | 117204766 | In Progress | 711407 | TRUE | Closed - Done | Closed - Done |  | CTUDT00632185 | [5 empty] | Engineering | [6 empty] | CT-Box | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | CT Box | Gantry-HW-CT Box | Gantry-HW-CT Box | CTUDT00632185 | CT Box | Gantry-HW | Gantry-HW-CT Box | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
590 |  |  | Puerto Rico | HOSPITAL DR SUSONI INC | 40013 | 44270 |  | REQUEST FOR RELOCATION OF THE WORKING TABLE | Work Order #: WO-03152517 15-Mar-2021 Ramos Rivera, Pablo The system fully meets specification and returned to use Problem Description by Engineer : Due to space restriction in the control room customer wants to replace the technologist tanle for smaller. Resolution : Disconnected and moved temporarly the keyboard, mouse, monitors and ct box while hospital personnel replaced the control table. Once they finnishes I reassembled them and tested system. ----------------------------------------------------------------------------- | [15 empty] | 117191250 | Closed | 703229 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [20 empty] | Refresh |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
591 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 341006 | 44274 |  | Bug report will be received in cardiac shooting | "Work Order #: WO-03177886 19-Mar-2021 Turgay Baylan The system partly meets specification and is returned to use with restrictions Problem Description by Engineer: Customer Function / Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution: After the CT device cardiac capture, bugrep was taken and added for escalation. The necessary park orders for Ibox have been entered. Follow Up Required: part order -------------------------------------------------- --------------------------- | 3 |  | The necessary park orders for Ibox | [12 empty] | 117216251 | In Progress | 718957 | TRUE | Closed - Done | Closed - Done |  | TBD-Translation | [5 empty] | Engineering | [6 empty] | Ibox | [5 empty] | Refresh |  |  | System issue | System | NotEngineering | ShouldbeOthers | NotEngineering-ShouldbeOthers | NotEngineering-ShouldbeOthers | TBD_Bugreport | NotEngineering-ShouldbeOthers | NotEngineering | NotEngineering-ShouldbeOthers | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
592 |  |  | Turkey | T.C Saglik Bakanligi Bagcilar -> Bağcılar State Hospital | 40010 | 44274 |  | Freezing SW and stuttering image transfer | Freezing SW and stuttering image transfer "Work Order #: WO-03178247 Soner Saglam The system partly meets specification and is returned to use with restrictions Problem Description by Engineer: SW freezing and not being able to send images to PACS. Resolution: Problems persist despite FCO740 implementation, case escalated. Follow Up Required: Problems persist despite FCO740 application, case escalated. | 3 |  | Yellow screen and slow, check if have further bugreport other then we have | [12 empty] | 117216620 | In Progress | 718902 | TRUE | Investigation | Closed - Done |  | CTUDT00631367 | [5 empty] | Engineering | [6 empty] | New Yellow Screen | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
593 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 341006 | 44273 |  | Ganty laser malfunction. | Work Order #: WO-03171801 Soner Saglam The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Gantry laser malfunction was reported. Resolution : On site diagnosis is required. Follow Up Required : On site diagnosis is required. ----------------------------------------------------------------------------- | 3 |  | FSE handling, check if have bugreport available | [12 empty] | 117210409 | In Progress | 716904 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD_Outer laser adjustment | [5 empty] | Engineering | [6 empty] | Laser | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | LaserAlignment | Gantry-HW-LaserAlignment | Gantry-HW-GantryLaser | TBD_Outer laser adjustment | gantry laser | Gantry-HW | Gantry-HW-LaserAlignment | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
594 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 341006 | 44273 |  | Unable to create recon in high and subsequent imaging sequences. | Work Order #: WO-03171800 Soner Saglam The system fully meets specification and returned to use Problem Description by Engineer : SW recon error. Resolution : SW installation will be tracked with 0117210389. ----------------------------------------------------------------------------- | 3 | [14 empty] | 117210425 | Closed | 716905 | TRUE | Closed - Duplicate | Closed - Duplicate | [20 empty] | Refresh |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
595 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 341006 | 44273 |  | Unable to display images when shooting. | Work Order #: WO-03171799 Soner Saglam The system fully meets specification and returned to use Problem Description by Engineer : SW error. Resolution : Followed up with 0117210389. ----------------------------------------------------------------------------- | 3 | [14 empty] | 117210392 | Closed | 716881 | TRUE | Closed - Duplicate | Closed - Duplicate | [20 empty] | Refresh |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
596 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 341006 | 44273 |  | Recon generation is not as fast as it used to be. | Work Order #: WO-03171798 Soner Saglam The system fully meets specification and returned to use Problem Description by Engineer : SW problem. Resolution : SW installation will be tracked with 0117210389. ----------------------------------------------------------------------------- | 3 |  | Duplicate with 0117210389. | [12 empty] | 117210418 | Closed | 716856 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Duplicate | [5 empty] | Engineering | [6 empty] | Recon Slow | [5 empty] | Refresh |  |  | System issue | System | TBD-Bugreport | TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | Duplicate | TBD_Bugreport | TBD-Bugreport | TBD-Bugreport-TBD_Bugreport | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
597 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 341006 | 44273 |  | Problem with data transfer to PACS. | Work Order #: WO-03171796 Soner Saglam The system fully meets specification and returned to use Problem Description by Engineer : The system freezes when sending images to PACS. Resolution : Followed up with 0117194764. ----------------------------------------------------------------------------- | 3 | [14 empty] | 117210388 | Closed | 716813 | TRUE | Closed - Duplicate | Closed - Duplicate | [20 empty] | Refresh |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
598 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 341006 | 44273 |  | Sistemde kasma ve donma oluyor | Work Order #: WO-03171797 Soner Saglam The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Various problems with SW have been reported. Resolution: The system needs to load SW. Follow Up Required: System SW must be loaded. ----------------------------------------------------------------------------- Work Order #: WO-03172766 21-Mar-2021 Turgay Baylan The system fully meets specification and returned to use CT Scanner host computer and cirs computer software installed. Host patches have been installed. System calibrations were made, shooting and delivered in working condition. ----------------------------------------------------------------------------- | 3 |  | SW reinstall and HF3 FCO implement | [12 empty] | 117210389 | Closed | 715220 | TRUE | Closed - Done | Closed - Done |  | TBD-Service | [5 empty] | Engineering | [6 empty] | Reload | [5 empty] | Refresh |  |  | System issue | System | CIRS-SW | Installation | CIRS-SW-Installation | CIRS-SW-Installation | TBD-Service | Installation | CIRS-SW | CIRS-SW-Installation | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
599 |  |  | Turkey | T.C Saglik Bakanligi Bagcilar -> Bağcılar State Hospital | 40010 | 44270 |  | Table Failure | [16 empty] | 117189630 | Fixed | 706594 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [20 empty] | Refresh |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
600 |  |  | Turkey | T.C Saglik Bakanligi Bagcilar -> Bağcılar State Hospital | 40010 | 44270 |  | Gantry failure | [16 empty] | 117189629 | Fixed | 706589 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [20 empty] | Refresh |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
601 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 341006 | 44271 |  | CT5000 HearthRate on the MP5 and system console are not synchronous | [18 empty] | 705866 | TRUE | Investigation | Closed - Done |  | TBD_HR_MP5 | [5 empty] | Engineering | [6 empty] | MP5 | [5 empty] | Refresh |  |  | System issue | System | TBD-Bugreport | TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | TBD_Bugreport | TBD_Bugreport | TBD-Bugreport | TBD-Bugreport-TBD_Bugreport | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
602 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 341006 | 44271 |  | Noise in the gantry. | Work Order #: WO-03155681 Soner Saglam The system fully meets specification and returned to use Problem Description by Engineer : Noise in the gantry was reported. Resolution : Closed with 0117149473 (The gantry of the CT Device was checked, it was detected that there was a noise from the belt part while turning, The belt order was made, the device was delivered in working condition.). ----------------------------------------------------------------------------- | 3 |  | Belt noise | [12 empty] | 117194761 | Closed | 705851 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | ECR-077012 | [5 empty] | Engineering | [6 empty] | Belt Noise | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Belt | Gantry-HW-Belt | Gantry-HW-Belt | ECR-077012 | Belt | Gantry-HW | Gantry-HW-Belt | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
603 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 341006 | 44271 |  | PACS Transfer Problem | Work Order #: WO-03155682 Soner Saglam The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Image transfer to PACS is slow. Resolution : The problem continues after FCO no 740. Spooler reset and case escalation is advised. Follow Up Required: The problem continues after FCO no 740. Spooler reset and case escalation is advised. ----------------------------------------------------------------------------- Work Order #: WO-03156950 16-Mar-2021 Turgay Baylan The system fully meets specification and returned to use Problem Description by Engineer : Image transfer to PACS is slow. Resolution : Bug report was taken when the image transfer to PACS has failed and restart spooler was completed in system service tools. RSN configuration was completed. Image transfer to PACS is slow, the case was escalated. ----------------------------------------------------------------------------- | 3 |  | PACS Transfer on HF3 | [12 empty] | 117194764 | In Progress | 705827 | TRUE | Evaluation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS New | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
604 |  |  | Turkey | HITACHI HEALTHCARE TURKEY | 337000 | 44270 |  | IP update of the transfer server. | Work Order #: WO-03149356 Soner Saglam The system fully meets specification and returned to use Resolution : Unable to connect to remote desktop. Connection problem. DOSEWATCH should be updated as 176.235.200.20. Problem Description by Engineer : Remote node update. Follow Up Required : Unable to connect remotely. Connection problem. DOSEWATCH should be updated as 176.235.200.20. ----------------------------------------------------------------------------- | 3 |  | Remote node update. | [12 empty] | 117189009 | In Progress | 704780 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Service | [6 empty] | Node config | [5 empty] | Refresh |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
605 |  |  | United States | Leesburg Regional Medical Cent | 40022 | 44277 |  | not able to pull worklists | Work Order #: WO-03185685 Morrow, Dan The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Problem description by engineer : Intermittent worklist not populating. When was the first occurrence? Within... : 21-Mar-2021 How often is issue occurring? : Intermittent Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: Manual input Patient Impact/Outcome: none User Impact: Manual patient input Current Software Version: Not provided Diagnostic performed by Engineer : Rebooted Host. System pingable. Advised customer to contact PACS admin first thing tomorrow to check PACS settings. Customer aware that PACS system has had problems in the past with worklist to CT. Will leave case open 24 hours. Internal Comments : Error # and/or description of error : No error Completed or Recommended Repair Action "Internal" : Rebooted Host. System pingable. Advised customer to contact PACS admin first thing tomorrow to check PACS settings. Customer aware that PACS system has had problems in the past with worklist to CT. Will leave case open 24 hours. Resolution : PACS admin fixed worklist on PACS side. OK to close. ----------------------------------------------------------------------------- | 4 |  | PACS admin fixed worklist on PACS side. OK to close. | [12 empty] | 117222858 | Fixed | 724101 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | PACS Vendor | [5 empty] | Refresh |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
606 |  |  | United States | The Regional Medical Center | 40005 | 44275 | 11 | Gantry | Work Order #: WO-03183375 Maier, Donnie The system does not meet specification and is out of use Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: No Patient Involved Expected and Unexpected Results: Power Event in facility and no power to gantry User Impact: No Patient Impact: No Current Software Version:Unknown Diagnostic performed by Engineer : I talked to CT Tech Tara and had her check to see if she had an on/off button on wall in front of her. She was unable to find one. She states they had a power event in facility. I asked her to look for teal and she was able to locate it (New Shorter Teal). I had her turn breaker off wait 5 minutes and turn it back on but there but blue and green lights are not lit. Cu is going to call her engineering dept. and have them check the incoming power as it appears we do not have any. I let her know I will put it for HTM 3/22AM dispatch for now and if they get it back up and running to call and let us know and we will cancel dispatch. I also let her know if her manager/director requests AH service this weekend, please get Name/Title/Email/Contact Phone Number to speed up response time. New System so it doesn_t appear to be in Radar, CT System Config, and ULV shows they lost power @18:41 hrs and hasn_t come back up obviously. CT Tech states they have another CT Scanner and probably will use that one this weekend. | 2 |  | I talked cu thru a complete power cycle of the system (Host/CIRS/Gantry). Issue ended up being within their facility (incoming power). I checked back and they are back up and running | [12 empty] | 117220689 | Fixed | 721668 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-PowerOff | [5 empty] | Engineering | [6 empty] | Power off | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | PowerOff | Gantry-HW-PowerOff | Gantry-HW-PowerOff | TBD-PowerOff | PowerOff | Gantry-HW | Gantry-HW-PowerOff | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
607 |  |  | United States | Doctors Hospital LLC | 40014 | 44271 | 11 | Artifact | Work Order #: WO-03159606 18-Mar-2021 Testerman Bennett The system fully meets specification and returned to use Follow Up Required : continue monitor site Diagnostic performed by Engineer : baddets,fee, check power Problem Description by Engineer : Customer Function/Role: line on pilot Device Use: in use Expected and Unexpected Results: run diagnostics User Impact: lines on pilot Patient Impact: Current Software Version: ----------------------------------------------------------------------------- Work Order #: WO-03172771 19-Mar-2021 Testerman Bennett The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: line on pilot Device Use: in use Expected and Unexpected Results: run diagnostics User Impact: lines on pilot Patient Impact: Current Software Version: Diagnostic performed by Engineer : reboot of system | 2 |  | 459801127432 - PCIe CAN CARD PCIe-402-1 Follow Up Required Description: continue monitor sit | [12 empty] | 117198388 | Closed | 710241 | TRUE | Closed - Done | Closed - Done |  | TBD-CanCard | [5 empty] | Engineering | [6 empty] | CAN Card | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | CanCard | Gantry-HW-CanCard | Gantry-HW-CanCard | TBD-CanCard | CanCard | Gantry-HW | Gantry-HW-CanCard | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
608 |  |  | United States | Lackey Memorial Hospital | 40015 | 44271 |  | when letting table down it made a crashing sound | Work Order #: WO-03158512 16-Mar-2021 Donald Vic The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Radiology Tech Device Use: Clinical Use Expected and Unexpected Results: The tech entered the control room to run daily cals. The tech found a ESTOP prompt message on the screen. The gantry displays were blank. The key on CT Box was still in the ESTOP closed position. Later that morning while lowering a patient support to off-load a patient the table made a crushing noise. User Impact: No impact Patient Impact: No impact Current Software Version: 4.1.10 Diagnostic performed by Engineer : Inspected the system Resolution : I put the patient support in the service position. I inspected the skirts, the screw drive, and bearings for defects. I also inspected the scissor lift for defects. There were no defects found. I ran the bug report through the AEL program to determine what may have caused ESTOP to open unintentionally. I did not find any causes for the ESTOP to open by itself. I will keep in contact with the customer on the system’s behavior to determine if there any other issues or occurrences of ESTOP dropping out unintentionally. | 5 |  | not find any causes for the ESTOP to open by itself. FSE keep in contact with the customer | [12 empty] | 117197499 | Closed | 708736 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD_Bugreport | [5 empty] | Engineering | [6 empty] | E-Stop | [5 empty] | Refresh |  |  | System issue | System | TBD-Bugreport | TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | TBD_Bugreport | TBD_Bugreport | TBD-Bugreport | TBD-Bugreport-TBD_Bugreport | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
609 |  |  | United States | Central Vermont Medical Center | 40024 | 44265 |  | When Networking images to pacs intermittently the | PACS Transfer issue on HF3, Bugreport sent out to Srusti reviewed Please note that the Store Association timeout value for iSite has been set to 30 seconds in this case, the default association timeout values for Dicom services is 120 seconds. Hence please change the association timeout value to 120. |  |  | PACS Transfer issue on HF3, Bugreport sent out to Srusti reviewed | [14 empty] | 703648 | TRUE | Closed - Done | Closed - Done |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS New | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
610 |  |  | United Kingdom | FORTH VALLEY PC NHS TRUST | 68183 | 44276 |  | recon box error is coming up | Work Order #: WO-03185153 21-Mar-2021 Paul Crane The system fully meets specification and returned to use Resolution : Replaced fibre gantry cirs Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Regulatory Questions : N/A ----------------------------------------------------------------------------- | 3 | [14 empty] | 117222465 |  | 723923 | #N/A | #N/A | #N/A | [20 empty] | Non-Refresh |  |  | Upgrade issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | #N/A | y |  |  | y
611 |  |  | United States | Christ Hospital | 310034 | 44272 |  | NEMO | Work Order #: WO-03168413 ----------------------------------------------------------------------------- | 3 | [14 empty] | 117206555 |  | 716997 | #N/A | #N/A | Evaluation | [20 empty] | Non-Refresh |  |  | Upgrade issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | y
612 |  |  | United States | St Jude Medical Center | 310165 | 44273 |  | 0504CDN81 E-Stop opened while idle | Work Order #: WO-03174983 Weaver David The system does not meet specification and is out of use Problem Description by Engineer : E-Stop opened while idle CALLER FUNCTION/ROLE Biomed DEVICE INTENDED USE Diagnostic WHEN WAS THE FIRST OCCURRENCE? WITHIN . . . Same day HOW OFTEN IS ISSUE OCCURRING? First occurrence WHERE IN THE END-USER WORKFLOW IS THE ISSUE OCCURRING? During acquisition/scan WHAT RECENT EVENTS HAVE OCCURRED? During normal use NEGATIVE OPERATOR IMPACT Can not scan NEGATIVE PATIENT IMPACT/PROCEDURE OUTCOME Can not scan patients HAS CUSTOMER CREATED A SPECIFIC SYSTEM EVENT REPORT OF ISSUE? No Diagnostic performed by Engineer : E-Stop opened while idle. Follow Up Required : ULV showed S_RTDH_HC_TIMEOUT_ERROR. and S_MPLTRANSPORTREC_ERREPORT_TOUCAN_ERROR after second power cycle. Problem could be CAN bus noise because Ghost did not see RHost online. This is a critical need, site down. Please page FSE immediately. ----------------------------------------------------------------------------- | 2 |  | 1.00000,453567374153,POWER SUPPLY 120 VAC DUAL OUTPUT Found no Estop GUI. Diagnosed a defective P3 power supply on gantry rotor. Ordered part as Expedited Shipping (Same Day) from Local Depot. Replaced power supply. Powered up and tested system. No issues encountered. Monitored site during first patient use. No issues. | [12 empty] | 117213060 |  | 716541 | #N/A | #N/A | #N/A | [14 empty] | 120VAC | [5 empty] | Non-Refresh |  |  | Upgrade issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | #N/A | y |  |  | y
613 |  |  | United States | Lehigh Valley Hospital Hazleto | 300167 | 44273 |  | XR-29 certif | Work Order #: WO-03174727 Hirchak, Michele The system fully meets specification and returned to use Problem description by Clinical Support Specialist : Customer Function/Role: TECH Device Use: In Use Expected and Unexpected Results: XR-29 certificate. Pt. Impact:: NONE User Impact: NONE Current Software Version: 4.1.7 Diagnostic performed by Engineer : Emailed certificate. She is all set. MBH ----------------------------------------------------------------------------- | 5 | [14 empty] | 117212824 |  | 716391 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [20 empty] | Non-Refresh |  |  | Upgrade issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | y
614 |  |  | United States | Wellstar Kennestone Hospital | 68114 | 44273 |  | 1302CDY73 Dosewise is not working | Work Order #: WO-03174570 Bourne Barrett Problem Description by Engineer : Problem description by engineer : Customer states that DoseWise is not working. When was the first occurrence? Within... : 18-Mar-2021 How often is issue occurring? : First occurrence Malfunction area : Console Information to support the complaint handling process Customer Function/Role: 3rd Party Device Use: No patient involved Expected and Unexpected results: DoseWise not working Patient Impact/Outcome: Not provided User Impact: can not use DoseWise. Current Software Version: 4.1.7 Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- Work Order #: WO-03174545 Miller, John M The system partly meets specification and is returned to use with restrictions Diagnostic performed by Engineer : Collect information from customer. Problem Description by Engineer : Problem description by engineer : Customer states that DoseWise is not working. When was the first occurrence? Within... : 18-Mar-2021 How often is issue occurring? : First occurrence Malfunction area : Console Information to support the complaint handling process Customer Function/Role: 3rd Party Device Use: No patient involved Expected and Unexpected results: DoseWise not working Patient Impact/Outcome: Not provided User Impact: can not use DoseWise. Current Software Version: 4.1.7 Internal Comments : Error # and/or description of error : DoseWise not working. Completed or Recommended Repair Action "Internal" : Customer did not have anymore information. Suggest checking Network, cabling and Lan Config settings. Follow Up Required : Error # and/or description of error : DoseWise not working. Completed or Recommended Repair Action "Internal" : Customer did not have anymore information. Suggest checking Network, cabling and Lan Config settings. ----------------------------------------------------------------------------- | 3 | [14 empty] | 117212644 |  | 716331 | #N/A | #N/A | #N/A | [20 empty] | Non-Refresh |  |  | Upgrade issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | #N/A | y |  |  | y
615 |  |  | United States | SSH-Pensacola | 310051 | 44272 |  | QO: Replacing IBox2 usb extenders | Work Order #: WO-03166639 18-Mar-2021 Siker Shawn The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : Ordered new USB extenders Follow Up Required : Replace USB extenders ----------------------------------------------------------------------------- Work Order #: WO-03172112 19-Mar-2021 Siker Shawn The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : Received parts ----------------------------------------------------------------------------- | 4 |  | 1.00000,459801731761,USB EXTENDER CABLE FROM IBOX TO CONSOLE 1.00000,455300027201,CBL,USB A/M TO B/M 0.5M Assisted FSE with implementing 4.1.7 to 4.1.10 Upgrade. Also implemented FCO_s 72800735, 72800737 and 72800740. System fully operational | [12 empty] | 117205085 |  | 714961 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | TBD-USBExtended | [5 empty] | Engineering | [6 empty] | USB extender | [5 empty] | Refresh |  |  | Upgrade issue | System | Gantry-HW | USBExtended | Gantry-HW-USBExtended | Gantry-HW-USBExtended | TBD-USBExtended | USBExtended |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
616 |  |  | China | Zhenjiang M.C.A. Hosp. | 345023 | 44282 |  | PRS Connection/02N/Core128 | Problem Reported by Customer : PRS Connection Next Step Action : FSE:马申申 18217339631 来电开此单 SAP Transfer : Case Transferred to SAP External Remarks : 27.03.2021 11:15:11 UTC SUP JCOUSER (86036526)服务内容：RSN不通，重新注册后正常 服务结果：维修后，设备恢复正常工作 Internal Remarks : 27.03.2021 11:15:14 UTC Shenshen Ma (86036526) RDI Accuracy：Yes $Q7: (Confirmed reported problem,the system can not connect PRS ) $Q8: (Check the network connection，found M2M was not registered in the system) $Q9: (Re-register M2M) $Q10: (Test and Inspection refer to PA，device return to fully function) $Q11:$Calibrated Tools：N/A SAP Transfer : Dispatched for Onsite Support | 5 |  | PRS Setup | [12 empty] | 117275571 |  | 747277 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Service | [6 empty] | PRS Setup | [5 empty] | Refresh |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
617 |  |  | Japan | Tomakomai City Hospital | 341014 | 44282 |  | 04DS:RS:E-STOP Close画面がポップアップしない>>全コントローラEAL_REQ/ANSの更新なし | Problem Reported by Customer: 04DM: RS: E-STOP Close screen does not pop up >> SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Support T2 Activities: Problem Description: Target device: CT Failure location: 04Gantry Exclude HV When did it first occur? (Within ~): 1 Hour How often do problems occur ?: Continuously At what point in the end-user workflow is the problem occurring ?: At system boot (Boot Up) What happened recently ?: N.A. Does the customer create a specific system event report for the problem ?: N.A. Troubleshooting Action: Result after response: Not improved Failure location: 04Gantry Exclude HV What steps were taken during the diagnosis ?: Even though it restarted from TEAL, it did not recover and it was judged that the system was down. There is no record of CAN EAL_ANS for all primary controllers in the log. Suspect a malfunction of the gantry main board or poor cable contact. Repair Action: Service representative visit work. OneEMS internal remarks: --Troubleshooting Action-- Post-response results: Not improved Failure location: 04Gantry Exclude HV What steps were taken during the diagnosis ?: Even though it restarted from TEAL, it did not recover and it was judged that the system was down. There is no record of CAN EAL_ANS for all primary controllers in the log. Suspect a malfunction of the gantry main board or poor cable contact. --Repair Action-- Recommended 12Nc: PR1 453567057167: FRC ASSEMBLY QTY: 1 SLOC: Sapporo (U612) PR1 453566492542: PWR SUPPLY, SWITCHING 12V, 140W QTY: 1 SLOC: Sapporo (U612) PR1 459800325761: Brush, Signal Block Assembly QTY: 1 SLOC: Tokyo (U611) PR1 459800325772: BRUSH, PWR BLOCK PCB ASSEMBLY QTY: 2 SLOC: Tokyo (U611) PR1 459800298761: AC CONTROL UNIT PCB QTY: 1 SLOC: Tokyo (U611) PR1 453567110325: BRILLIANCE GHOST ASSEMBLY QTY: 1 SLOC: Tokyo (U611) | 2 |  | 453566492542 PWR SUPPLY, SWITCHING 12V, 140W After replace PWR SUPPLY，SWITCHING 12V system works well --- Closed | [12 empty] | 117275501 |  | 747215 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-Translation | [5 empty] | Engineering | [6 empty] | E-Stop | [5 empty] | Refresh |  |  | System issue | System | NotEngineering | ShouldbeOthers | NotEngineering-ShouldbeOthers | NotEngineering-ShouldbeOthers | NotEngineering-ShouldbeOthers | NotEngineering-ShouldbeOthers | NotEngineering | NotEngineering-ShouldbeOthers | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
618 |  |  | China | HLJ Huachuan People Hosp. | 345018 | 44282 |  | 02N/Core128/Gantry is rotating noise | T2 Activities : Problem Description: 客户问题概述： 机架旋转时有噪声 客户问题详细描述或故障现象： 机架旋转时有噪声 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： N/A 远程诊断与测试结果： N/A PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 检查机架内部旋转部分。 建议备件12NC及名称： N/A 客户需求备注： 无 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : Site1 FSE 靳军13796075213 告知可以派给他，已发派工短信通知到工程师 Internal Remarks : 29.03.2021 00:59:48 UTC Jun Jin (86039442) $Q1: () $Q2: () $Q3: () $Q4: () $Q7:(Confirmed reported problem &noise when rotating ) $Q8: (when rotating，the tension wheel rubs against the ftame，then can‘t found 0 position ) $Q9: (adjust tension wheel and replace gmp) $Q10( Test and inspection data refer to : PA . Device returned to full functionality.) $Q11( Calibrated tools name (N/A） External Remarks : 29.03.2021 00:59:45 UTC Jun Jin (86039442)服务内容： 调整张紧轮，更换gmp 服务结果：设备恢复正常 | 3 |  | 服务内容： 调整张紧轮，更换gmp 服务结果：设备恢复正常 | [12 empty] | 117275416 |  | 747166 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-Translation | [5 empty] | Engineering | [6 empty] | Gantry noise | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | GMP | Gantry-HW-GMP | Gantry-HW-GMP | Gantry-HW-GMP | GMP | Gantry-HW | Gantry-HW-GMP | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
619 |  |  | China | Lindian County Hospital | 345038 | 44281 |  | WeChat:02N/Ingenuity Core128/Elite China/Unable to preheat | T2 Activities : Problem Description: 客户问题概述： 预热失败。 客户问题详细描述或故障现象： 预热报错，扫描病人正常 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 预热报错，扫描病人正常 远程诊断与测试结果： 后台log未见明显报错指向，现场具体根据实际情况观察异常 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 根据现场实际情况做相关troubleshooting SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 26.03.2021 23:58:37 UTC Gang Ni (86038229)服务内容：指导用户重启设备 服务结果：设备恢复 Internal Remarks : 26.03.2021 23:58:49 UTC Gang Ni (86038229) $Q7: (Confirmed reported problem & Unable to preheat) $Q8: (review errorlog; check couch position & couch position error) $Q9: (Reset couch) $Q10: (functional test pass & Device returned to full functionality) $Q11: Calibrated Tools name：none | 3 |  | 服务内容：指导用户重启设备 服务结果：设备恢复 | [12 empty] | 117268403 |  | 747030 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-Translation | [5 empty] | Engineering | [6 empty] | Warm up | [5 empty] | Refresh |  |  | System issue | System | TBD-Bugreport | TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | TBD_Bugreport | TBD-Bugreport | TBD-Bugreport-TBD_Bugreport | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
620 |  |  | China | Yangfangdian hospital | 40035 | 44281 |  | 02N/Ingenuity CT 728326/stand by | BU Q&R Feedback: No Reason: Customer Relations Problem Reported by Customer : stand by Next Step Action : 下周一上午10点需要工程师到场协助 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Safety Questions : Case: 0117270394 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Qian Wang Last updated by: Qian Wang | 5 |  | add in CD-record, Add a network recorder for workstaion and CT host. | [12 empty] | 117270394 |  | 745827 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | TBD | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
621 |  |  | China | Dongguan Donghua Hospital CO LTD | 40059 | 44281 |  | 02G/Ingenuity CT 728326/standby | Next Step Action : 已mailIB，在IA#53232184上加operation，派工FSE Lin Zhitao Next Step Action : OP已加0020 ----------------------------------------------------------------------------- | 5 |  | 派工FSE Lin Zhitao | [12 empty] | 117268904 |  | 745753 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | TBD | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
622 |  |  | China | The Second People_s Hospital | 345039 | 44280 |  | Wechat:02N/Ingenuity Core128/Elite China/UPS fault | BU Q&R Feedback: No SAP Customer Symptom Code: BOOT Customer Symptom Code: Boot up/Power-up Problem Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: Reboot/Restart Problem Reported by Customer : UPS fault Safety Questions : Case: 0117260173 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Bai Jingwen Last updated by: Bai Jingwen T2 Activities : Problem Description: UPS alarm Troubleshooting Action: need power down the UPS Repair Action: guide customer power down the UPS Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: N/A SAP Transfer : Case Transferred to SAP | 3 |  | power down the UPS | [12 empty] | 117260173 |  | 741064 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD_UPS | [5 empty] | Engineering | [6 empty] | UPS | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | UPS | Gantry-HW-UPS | Gantry-HW-UPS | TBD_UPS | UPS | Gantry-HW | Gantry-HW-UPS | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
623 |  |  | China | JS Yangzhou No.1 People Hosp. Xiqu Hosp. | 345022 | 44278 |  | 02N/Ingenuity Core128/Elite China/add new printer | BU Q&R Feedback: No SAP Customer Symptom Code: ACQP Customer Symptom Code: Imaging/Patient Recording/Acquisition Problem Reason: Corrective Maintenance Engineer Failure Code: Configuration Error Repair Activity Code: Configuration Change T2 Activities: Problem Description: need add new printer Troubleshooting Action: need add new printer Repair Action: guide customer add new node Test/Inspection: Functional Check PASSED. The device returned to full functionality. Test Equipment: N/A SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | 5 |  | guide customer add new node | [12 empty] | 117231215 |  | 733103 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | TBD | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
624 |  |  | China | SD Chest Hosp. | 40025 | 44279 |  | 02N/Ingenuity CT 728326 | BU Q&R Feedback: No SAP Customer Symptom Code: HANG Customer Symptom Code: Hang/Crash Reason: Corrective Maintenance Engineer Failure Code: Configuration Error Repair Activity Code: Configuration Change Safety Questions : Case: 0117242328 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Bo Last updated by: Wang Bo Problem Reported by Customer : 02N/Ingenuity CT 728326 T2 Activities : Problem Description: Loss of patient information Troubleshooting Action: Assist the customer to rebuild the rawdata data manually Repair Action: Assist the customer to rebuild the rawdata data manually Test/Inspection: customer confirm can scan patient and image ok.functional Check PASSED. Device returned to full functionality. Test Equipment: N/A SAP Transfer : Case Transferred to SAP | 3 |  | Assist the customer to rebuild the rawdata data manually | [12 empty] | 117242328 |  | 731723 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-Translation | [5 empty] | Engineering | [6 empty] | Rawdate | [5 empty] | Refresh |  |  | System issue | System | TBD-Bugreport | TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | TBD_Bugreport | TBD-Bugreport | TBD-Bugreport-TBD_Bugreport | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
625 |  |  | China | Yangfangdian hospital | 40035 | 44278 |  | 02N/Ingenuity CT 728326/stand by | External Remarks : 25.03.2021 01:53:34 UTC Xing Zheng (86039460)服务内容：配合激光灯验收 服务结果：设备恢复正常使用。 External Remarks : 25.03.2021 02:58:15 UTC Guochao Fan (86039456)服务内容：配合激光灯验收 服务结果：设备恢复正常使用。 Internal Remarks : 25.03.2021 02:58:20 UTC Guochao Fan (86039456) "RDI Accuracy: Yes" $Q7: Confirm reported problems:Built-in laser lamp detection. $Q8: Trouble shooting steps:for Built-in laser lamp detection .the Laser lamps need to be adjusted $Q9: Function returns to normal after the Laser lamps adjusted。 $Q10: Test and Inspection data refer to PA& Device status:Device returned to full functionally $Q11: N/A Internal Remarks : 25.03.2021 15:50:18 UTC Xing Zheng (86039460)"RDI Accuracy: Yes" $Q7: Confirm reported problems:Built-in laser lamp detection. $Q8: Trouble shooting steps:for Built-in laser lamp detection .the Laser lamps need to be adjusted $Q9: Function returns to normal after the Laser lamps adjusted。 $Q10: Test and Inspection data refer to PA& Device status:Device returned to full functionally $Q11: N/A | 5 |  | 服务内容：配合激光灯验收 服务结果：设备恢复正常使用。 | [12 empty] | 117230254 |  | 731690 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint |  | TBD_Outer laser adjustment | [5 empty] | Engineering | [6 empty] | Laser | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | LaserAlignment | Gantry-HW-LaserAlignment | Gantry-HW-laser | TBD_Outer laser adjustment | gantry laser |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
626 |  |  | China | JS Yangzhou No.1 People Hosp. Xiqu Hosp. | 345022 | 44278 |  | 02N/Ingenuity Core128/Elite China/add new printer | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 02N/Ingenuity Core128/Elite China Safety Questions:Case: 0117233822 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Guoping Last updated by: Liu Guoping 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? Yes 4. Was this an out of box failure? Yes ----------------------------------------------------------------------------- | 3 |  | Add new printer | [12 empty] | 117233822 |  | 728282 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | TBD | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
627 |  |  | China | HEB Caofeidian Dist. Hosp. | 345016 | 44278 |  | 02N/Ingenuity Core128/Elite China/change | BU Q&R Feedback: No SAP Customer Symptom Code: HANG Customer Symptom Code: Hang/Crash Reason: Corrective Maintenance Engineer Failure Code: Configuration Error Repair Activity Code: Configuration Change T2 Activities:Problem Description: customer need to change time Troubleshooting Action: Help customers modify time Repair Action: Help customers modify time Test/Inspection: customer confirm can scan patient and image ok.functional Check PASSED. The device returned to full functionality. Test Equipment: N/A SAP Transfer:Case Transferred to SAP | 3 |  | Change date and time | [12 empty] | 117232531 |  | 728055 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | TBD | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
628 |  |  | Puerto Rico | HOSPITAL DR SUSONI INC | 40013 | 44280 |  | The device does not work and has a burning smell | Work Order #: WO-03211010 Valenzuela Velasquez, Eduardo Ariel The system does not meet specification and is out of use Problem Description by Engineer : Problem description by engineer : Gantry error. It doesn´t allow to work. Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: No further information is available as per the customer Patient Impact/Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer Current Software Version: 4.1.10 Diagnostic performed by Engineer : The error is permanent and after a complete system restart, the CT is still showing same message. Internal Comments : Error # and/or description of error : System error, please retry. Completed or Recommended Repair Action "Internal" : The remote access to download files is having problems. Follow Up Required : Error # and/or description of error : System error, please retry. Completed or Recommended Repair Action "Internal" : The remote access to download files is having problems. FSE is traveling to the site. ----------------------------------------------------------------------------- Work Order #: WO-03211508 25-Mar-2021 Rios, William The system fully meets specification and returned to use Problem Description by Engineer : Problem description by engineer : Gantry error. It doesn´t allow to work. Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: No further information is available as per the customer Patient Impact/Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer Current Software Version: 4.1.10 Resolution : Turn ON the reconstructor and perform functional test to the system.. unit is ok and ready for patient use. Also found the the floor where the CT Gantry is located when Gantry start the rotation at maximum speed the floor start vibrating, Notify customer that need to inspect with the Engineer department. Diagnostic performed by Engineer : Customer report that system dose not let scan, when verify system, found that the Reconstructor computer was Off.. also report when gantry is rotating the patient table start vibrating during scan. ----------------------------------------------------------------------------- | 2 |  | Turn ON the reconstructor and perform functional test to the system.. unit is ok and ready for patient use. | [12 empty] | 117263791 |  | 738099 | TRUE | Closed - Done | Closed - Done |  | TBD-GantryError | [5 empty] | Engineering | [6 empty] | Gantry error | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Gantry Error | Gantry-HW-Gantry Error | Gantry-HW-Gantry Error | TBD-GantryError | Gantry Error | Gantry-HW | Gantry-HW-Gantry Error | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
629 |  |  | China | SD Tai_an T.C.M. Hosp. No.2 Ho | 345037 | 44277 |  | 02N/Ingenuity Core128/Elite China/Head cushion | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Replace Part T2 Activities:Problem Description: need head cushion Troubleshooting Action: nemo Head cushion Repair Action: nemo head cushion 459800826031 STANDARD HEADREST CUSHION KIT 1 453567104512 Head Cushion (Formally FCO# 72800264) 1 Test Equipment: N/A SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | 5 |  | 459800826031 STANDARD HEADREST CUSHION KIT 1 453567104512 Head Cushion (Formally FCO# 72800264) | [12 empty] | 117223381 |  | 727595 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-HeadCushion | [5 empty] | Engineering | [6 empty] | Head cushion | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | HeadCushion | Couch-HW-HeadCushion | Couch-HW-HeadCushion | TBD-HeadCushion | HeadCushion | Couch-HW | Couch-HW-HeadCushion | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
630 |  |  | China | GD Dongguan Dalingshan Hosp. | 345013 | 44277 |  | 02N/Ingenuity Core128/Elite China/ the computer forced turn off | Problem Reported by Customer: the computer forced turn off T2 Activities:Problem Description: Summary of customer_s problems: failure related to computer software. Detailed description of Customer_s problems or malfunctions: After entering the patient to settle in place before coming out, the computer is shut down without a reason It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: Computer hardware is malfunctioning Remote troubleshoot and test result: Computer hardware is malfunctioning PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Check software and hardware related failure SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support | 5 |  | Need bugreport for analysis | [12 empty] | 117223479 |  | 727329 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00631367 | [5 empty] | Engineering | [6 empty] | New Yellow Screen | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
631 |  |  | Japan | Saiseikai Gosu Hospital | 343000 | 44280 |  | CT:APP:NO:ﾓﾘﾜｷ様:質問 | BU Q & R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Application Support Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure SAP Transfer: Case Transferred to SAP Problem Reported by Customer: Q: I plan to photograph a patient who weighs more than 140 kg. Should I raise the DRI? A: Even if you raise the DRI, there is a high possibility that you will get caught in the tube limit, so try to keep the output by reducing the pitch or slowing down the rotation time. If there is a lot of noise, it is recommended to cover it with iDose. Safety Questions: Case: 0117262135 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Ootani Natsuki Last updated by: Ootani Natsuki | 5 |  | NAC? | [12 empty] | 117262135 |  | 741978 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Application | [6 empty] | Q&A | [5 empty] | Refresh |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
632 |  |  | Japan | KYOTO MEDICAL CENTER | 341013 | 44280 |  | Auto network transfer stucked after installed FCO72800740 | Duplate with below case 117241556? | WANG, Peter |  | Duplicate? Do we have bugreport? | [14 empty] | 737030 | TRUE | Investigation | Closed - Duplicate |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | New PACS | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
633 |  |  | Japan | KYOTO MEDICAL CENTER | 341013 | 44279 |  | JY:Even after the patch was installed, auto transfer fails. | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: JY:Even after the patch was installed, auto transfer fails. Safety Questions:Case: 0117241556 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: YAN Joanna Last updated by: YAN Joanna N/A OneEMS internal remarks: Yamamoto DB SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Support | 5 |  | CTUDT00632127 | [12 empty] | 117241556 |  | 735974 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | New PACS | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
634 |  |  | Japan | Higashi-Ohmi General Medical C | 341015 | 44280 |  | 01DS:RS:PACSへの画像転送がWaitingのまま進まず>> | BU Q & R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 01DS: RS: Image transfer to PACS does not proceed as Waiting >> T2 Activities: Problem Description: Target device: CT Failure location: 01Cosole When did it first occur? (Within ~): 1 Hour How often do problems occur ?: First Occurrence At what point in the end-user workflow is the problem occurring ?: Network What happened recently ?: N.A. Does the customer create a specific system event report for the problem ?: N.A. Troubleshooting Action: Result after response: Improved Failure location: 01Cosole What steps were taken during the diagnosis ?: Resolved after restarting the host Repair Action: [Phenomenon confirmation] Verified [Details of work] Resolved after restarting the host [Confirmation details] Good Safety Questions: Case: 0117259687 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: RUI SATAKE Last updated by: RUI SATAKE | 4 |  | CTUDT00632127 | [12 empty] | 117259687 |  | 735921 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | New PACS | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
635 |  |  | Poland | Centrum Opieki Medycznej | 341002 | 44281 |  | "UP" key does not work | Work Order #: WO-03218197 Mateusz Ramadan Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A | 4 |  | 1.00000,459800484501,CONTROL PANEL ASSEMBLY-FRONT LH | [12 empty] | 117271694 |  | 746105 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632181 | [5 empty] | Engineering | [6 empty] | GCP | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Gantry Panel | Gantry-HW-Gantry Panel | Gantry-HW-Gantry Panel | CTUDT00632181 | Control Panel | Gantry-HW | Gantry-HW-Gantry Panel | No | y | y | y | y | y | No | y | y | Gantry Panel | y | y | y | y | y | y | y | y
636 |  |  | Poland | Centrum Opieki Medycznej | 341002 | 44279 |  | The camera cannot be turned on. | Problem Description by Engineer: Diagnostics. Disc Problems Found: 1. RHOST - hangs randomly. 2. GMP - Freezes Randomly. Plates to be replaced have been ordered. Diagnostic performed by Engineer: Diagnostics. Disc Problems Found: 1. RHOST - hangs randomly. 2. GMP - Freezes Randomly. Plates to be replaced have been ordered. -------------------------------------------------- --------------------------- Work Order #: WO-03204286 26-Mar-2021 Matthew Ramadan The system fully meets specification and returned to use Problem Description by Engineer: Replacing the RHOST board. GMP plate replacement. The fault has been removed. Diagnostic performed by Engineer: Replacing the RHOST board. GMP plate replacement. The fault has been removed. Tools & Equipment: TORQUE SCREWDRIVER, 2021-06-03 00:00:00 MULTIMETER, 2021-07-15 00:00:00 Parts Consumed: 1.00000,459800440404, Common Gantry Motion Processor Assy 1.00000,459800638522, RHOST SG164 ASSEMBLY with CPM (RoHS) | 4 |  | Parts Consumed: 1.00000,459800440404, Common Gantry Motion Processor Assy 1.00000,459800638522, RHOST SG164 ASSEMBLY with CPM (RoHS) | [12 empty] | 117247653 |  | 742909 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-GMP | [5 empty] | Engineering | [6 empty] | GMP+Rhost | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | GMP | Gantry-HW-GMP | Gantry-HW-GMP | TBD-GMP | GMP | Gantry-HW | Gantry-HW-GMP | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
637 |  |  | Poland | Swietokrzyski Szpital tymczasowy | 341009 | 44279 |  | System configuration. | Work Order #: WO-03203001 24-Mar-2021 Marcin Knyps The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : Camera network configuration, PACS and worklist connection. Checking the operation of data transfer and downloading data from the worklist - correct. Acceptance, constancy, quick iq image quality test - in accordance with the manufacturer_s specification. apparatus prepared for training. | 4 |  | NAC? | [12 empty] | 117251351 |  | 735621 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Service | [6 empty] | Network config | [5 empty] | Refresh |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
638 |  |  | Poland | Uniwersytecki Szpital Kliniczny | 341001 | 44279 |  | System configuration. | Work Order #: WO-03203019 24-Mar-2021 Marcin Knyps The system fully meets specification and returned to use Diagnostic performed by Engineer : N/A Problem Description by Engineer: Configuration of the device was finished. Resolution: Network configuration of the device, PACS and worklist connection. Check of data transfer and data download from the worklist - correct. Acceptance, constancy, quick IQ image quality test - in accordance with the manufacturer_s specification. ￼￼￼ Device is prepared for training ----------------------------------------------------------------------------- | 4 |  | NAC | [12 empty] | 117250660 |  | 735602 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [14 empty] | Network config | [5 empty] | Refresh |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
639 |  |  | Spain | S.C.S. HOSP. GRAL. DE FUERTEVE | 40020 | 44277 |  | Problems with the reconstruction in a patient_s neck. | Work Order #: WO-03188377 Carlos Saumell Lasala The system fully meets specification and returned to use Regulatory Questions : N/A Problem Description by Engineer : I speak with Mar, Tac technician. In one particular patient, composed of several series, neck reconstruction was not to the liking of the doctors. This is a study conducted on January 22. Diagnostic performed by Engineer: I review the images of this neck study and verify that it is acquired with a high resolution bone protocol. And the doctors wanted to see soft parts. Being the month of January, the raw data are no longer found in the CIRS, to try a new reconstruction. Resolution: I indicate to the client that, when wanting to visualize soft parts, they should use another protocol or another reconstruction with the raw data. But now these are no longer available, being from some time ago (January). They should try new rebuilds when the raw data is still available in CIRS. | 3 |  | Doctors wanted to see soft parts on the images of this neck study and verify that it is acquired with a high resolution bone protocol | [12 empty] | 117226317 |  | 728992 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Mis-use | [5 empty] | Refresh |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
640 |  |  | Sweden | Ludvika Lasarett | 40049 | 44278 |  | The pictures do not go over to PACS. E-mail below: | Work Order #: WO-03193889 BORGERSEN EIRIK The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : problem sending to Pacs Diagnostic performed by Engineer: We think this problem was solved with a software patch in FCO72800740. I will ask the service center to get in touch and plan this in. With friendly greetings Eirik ----------------------------------------------------------------------------- | 3 |  | FCO72800740 | [12 empty] | 117231207 |  | 732244 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS Transfer | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
641 |  |  | Sweden | Avesta lasarett | 40040 | 44277 |  | The images do not transfer to PACS. Most restart all the time. | Work Order #: WO-03189544 BORGERSEN EIRIK The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : problem sending to PACS Diagnostic performed by Engineer : Problem sending to PACS. System is currently not on lates SW patch, we belive the errror will be solved in the 740 patch below. This is already planned in Avesta in august, but I will tell service center to find a new date EQ 79152830 SN 40040 Avesta DT2 SE5301 (Ingenuity 4.1.10.25202) 4.1.10.25202 FCO 72800737 FCO 72800740 | 3 |  | FCO72800740 | [12 empty] | 117227216 |  | 729678 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS Transfer | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
642 |  |  | Turkey | T.C Saglik Bakanligi Bagcilar -> Bağcılar State Hospital | 40010 | 44279 |  | There are artifacts in images. | Work Order #: WO-03202494 Soner Saglam The system partly meets specification and is returned to use with restrictions Resolution: The gantry should be opened up and the detector should be checked as it is suspected that there is opaque substance on it. Problem Description by Engineer: There are artifacts in the images. Follow-Up Required: The gantry should be opened up and the detector should be checked as it is suspected that there is opaque substance on it. ----------------------------------------------------------------------------- | 3 |  | In the examination, it was observed that contrast material was leaking into the device. Contrast material was cleaned and air calibrated. The system was delivered in an active condition by trial shooting. | [12 empty] | 117249456 |  | 733734 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | ECR-075975 | [5 empty] | Engineering | [6 empty] | Lexan Ring/Contrast | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Lexan Ring | Gantry-HW-Lexan Ring | Gantry-HW-Lexan Ring | ECR-075975 | Lexan Ring | Gantry-HW | Gantry-HW-Lexan Ring | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
643 |  |  | United States | Seton Medical Center Harker He | 40006 | 44280 |  | Remote access to scanner is not working correctly. | Work Order #: WO-03214634 Mark Finnigan The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Tech Device Use: No patient involved. Expected and Unexpected Results: Customer says remote connection is not working correctly. User Impact: Unknow Patient Impact: None Current Software Version: 4.1.10 Internal Comments : Called back and spoke with Philips onsite apllication suport Beth Sherrard. We tried to establish a remote desktop and could not. An error message returned stating the server is not running. This may very well be a port or firewall issue. The local FSE will need to work closely with the onsite IT to resolve this issue. Follow Up Required : Spoke with Darla Rose. She states that her remote connection to the site is not workingg correctly and that this has been a long standing issue. Clinical applications are onsite and they tried to use remote desktop sharing and it did not work. This is a 4.1.10 system. It is recommended that the FSE connect to PRS from his laptop while at the site and test all of the various applications. If an issue is uncovered he can then work directly with the sote network team to resolve the issue. Called back and spoke with Philips onsite apllication suport Beth Sherrard. We tried to establish a remote desktop and could not. An error message returned stating the server is not running. This may very well be a port or firewall issue. The local FSE will need to work closely with the onsite IT to resolve this issue. Diagnostic performed by Engineer : Spoke with Darla Rose. She states that her remote connection to the site is not workingg correctly and that this has been a long standing issue. Clinical applications are onsite and they tried to use remote desktop sharing and it did not work. This is a 4.1.10 system. It is recommended that the FSE connect to PRS from his laptop while at the site and test all of the various applications. If an issue is uncovered he can then work directly with the sote network team to resolve the issue. ----------------------------------------------------------------------------- | 5 |  | Firewall setting need to enable ports for PRS connection by following service manual 459801829121 | [12 empty] | 117266871 |  | 739953 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-PRS | [5 empty] | Engineering | [6 empty] | Remote connection | [5 empty] | Refresh |  |  | System issue | System | Console-SW | PRSPort | Console-SW-PRSPort | Console-SW-PRSPort | TBD-PRS | PRSPort | Console-SW | Console-SW-PRSPort | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
644 |  |  | United States | Mount Grant General Hospital | 40012 | 44280 |  | 1302CDN83 x246 Sending images stopped halfway | Work Order #: WO-03213523 Maier, Donnie The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: No Patient Involved Expected and Unexpected Results: Images stop sending halfway thru User Impact: No Patient Impact: No Current Software Version: 4.1.10 Diagnostic performed by Engineer : Kevin (CT Tech) states he had to shutdown computers and bring them back to finish sending the images. Once today and maybe 3 times yesterday. I tried to talk him thru resetting spooler via olevel but 4.1.10 system and username and password must be different than usual olevel. I tried pulling logs via our script but got Keyboard interactive error (CopSSH?) Cu has created a bug report. System is fairly new and under warranty. System is not in Radar, CT System Config, etc. Dispatching FSE to analyze BR locally, etc. No PFSE/SFSE/TFSE listed under IP# to collaborate with. Follow Up Required : Kevin (CT Tech) states he had to shutdown computers and bring them back to finish sending the images. Once today and maybe 3 times yesterday. I tried to talk him thru resetting spooler via olevel but 4.1.10 system and username and password must be different than usual olevel. I tried pulling logs via our script but got Keyboard interactive error (CopSSH?) Cu has created a bug report. System is fairly new and under warranty. System is not in Radar, CT System Config, etc. Dispatching FSE to analyze BR locally, etc. Cu also states they have database is only at 20% full. | 3 |  | What is SW version? HF2?HF3? PRS connection need to be checked by follow the service manual 459801829121 share bugreport for analysis | [12 empty] | 117265815 |  | 739373 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS Transfer | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
645 |  |  | United States | Seton Medical Center Harker He | 40006 | 44279 |  | 0699CFXXX GCC during Calcium Scoring | Work Order #: WO-03204712 Maier, Donnie The system fully meets specification and returned to use Diagnostic performed by Engineer : My script collecting logs from this win10 system doesn_t appear to be working. ULV however shows S_MASTER_SEQUENCER_ACQ_NOT_OK Normal Error 10:02:47 which is time Beth (Clin Apps) told me the Gantry Cant Comply Error occurred. Apps state they restarted the scan and everything is working properly again. ULV shows 139.59 days since last gantry power cycle. System doesn_t appear to be in Radar. I called back to Beth (Clin Apps) and Darla CT Tech and confirmed that yes they have not been shutting off gantry. Cu will power cycle system completely when they get a chance and call back if they have any further issues. Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: Diag Expected and Unexpected Results: GCC during Calcium Scoring Exam User Impact: No Patient Impact: No Current Software Version: 4.1.10 Resolution : A complete power cycle of the system was performed. | 4 |  | ULV shows 139.59 days since last gantry power cycle. Need train customer to restart whole system once per week and host restart once per day | [12 empty] | 117255675 |  | 734600 | TRUE | Closed - Done | Closed - Done |  | CTUDT00630408 | [5 empty] | Engineering | [6 empty] | ACQ_Not_ok | [5 empty] | Refresh |  |  | System issue | System | CIRS-SW | ACQ_NOT_OK | CIRS-SW-ACQ_NOT_OK | CIRS-SW-ACQ_NOT_OK | ACQ_NOT_OK | ACQ_NOT_OK | CIRS-SW | CIRS-SW-ACQ_NOT_OK | No | y | Gantry-HW | y | Gantry-HW-ACQ_NOT_OK | Gantry-HW-ACQ_NOT_OK | No | y | TBD-ACQ-Not-OK | y | Gantry-HW | Gantry-HW-ACQ_NOT_OK | y | y | y | TBD-ACQ-Not-OK | y | y
646 |  |  | United States | The Regional Medical Center | 40005 | 44279 |  | CUSTOMER SUPPORT REQUEST | Work Order #: WO-03201982 Powell, Jamie The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : Customer_s system was still logged in as service after PM was done yesterday. Walked customer through logging out of service and was able to log in as CT Diagnostic performed by Engineer : Customer_s system was still logged in as service after PM was done yesterday. Walked customer through logging out of service and was able to log in as CT | 3 |  | Service User login | [12 empty] | 117247373 |  | 733050 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Service | [6 empty] | Service User login | [5 empty] | Refresh |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
647 |  |  | United States | Clay County Medical Center | 40003 | 44278 |  | Xray tube arc error during short tube conditioning | Work Order #: WO-03197920 23-Mar-2021 Stauffer Jeffrey The system fully meets specification and returned to use Diagnostic performed by Engineer : Looked in error logs and saw tube arcs at time of error. Problem Description by Engineer : Customer Function/Role: CT Technician Device Use: General CT Exams Expected and Unexpected Results: Expected no errors during scanning. Unexpected - xray tube arc errors during short tube conditioning. User Impact: Delays in scanning next patient. Patient Impact: Delayed exams. Current Software Version: 4.1.10.25202 Resolution : Performed long tube conditioning, Performed filament calibration pass 2, cleaned off detector and tube face and performed air calibrations. | 4 |  | TC Boost | [12 empty] | 117239760 |  | 730596 | TRUE | Closed - Done | Closed - Done |  | CTUDT00630085 | [5 empty] | Engineering | [6 empty] | Tube Arc | [5 empty] | Refresh |  |  | System issue | System | IC-HW | Tub Arc | IC-HW-Tub Arc | IC-HW-Tub Arc | CTUDT00630085 | Tub Arc | IC-HW | IC-HW-Tub Arc | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
648 |  |  | China | Henan Fengquan District Dakuai | 345009 | 44277 |  | 02N/Ingenuity Core128/Elite China/consult FCO | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer : consult FCO Safety Questions : Case: 0117223459 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Zixuan Last updated by: Wang Zixuan 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: Consult about 72800740 Troubleshooting Action: Explain as software patch Repair Action: Explain as patch and FSE will contact before onsite Test/Inspection: N/A Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | 44277 | [12 empty] | 117223459 |  | 727433 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue |  |  |  |  | 0 |  | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
649 |  |  | United States | St Joseph Westgate Medical Center | 40001 | 44232 |  | CT box down button not working; Replaced the CT box to resolve. | Reviewed the case and indicated that the CT box was also replaced in this record. Checked in the “Defective part information" of the "Parts used" and confirmed that the down button wasn_t working and therefore CT box was replaced. As there were two different issues identified in this record which were 1) Gantry can not comply errors 2) CT box down button not working, therefore, this PR 589001 has been split into two to have: Other PR( * PR 589001 addressing the gantry can not comply errors) * PR 727823 addressing CT box down button not working. Work Order #: WO-02932899 09-Feb-2021 Miracle James The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : Tested all kv signals with NSS. Jtagged and flashed RHOST. Set up remote access on system Follow Up Required : Check system and finalize ct box replacement Parts Consumed : 1.00000,459801705022,CT BOX ASSEMBLY_BEST Tools & Equipment : null,2022-07-31 00:00:00 null,2021-03-31 00:00:00 ----------------------------------------------------------------------------- Work Order #: WO-02963344 16-Feb-2021 Miracle James The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : Finalized ct box replacement and tested movements | XU, Miley |  | CT box down button not working; Replaced the CT box to resolve. | [14 empty] | 727823 | TRUE | Closed - Done | Closed - Done |  | CTUDT00632185 | [5 empty] | Engineering | [6 empty] | CT Box | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | CT Box | Gantry-HW-CT Box | Gantry-HW-CT Box | CTUDT00632185 | CT Box | Gantry-HW | Gantry-HW-CT Box | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
650 |  |  | Netherlands | Rijnstate Ziekenhuis Arnhem | 320103 | 44279 |  | System error. | Work Order #: WO-03201788 ERWIN OVERMAN The system fully meets specification and returned to use Problem Description by Engineer : System error, please restart the system Diagnostic performed by Engineer: bugrep shows that user has not restarted Cirs for 54 days. Instruct user | [17 empty] | 737277 | #N/A | #N/A | Investigation | [14 empty] | CIRS Reboot | [5 empty] | Non-Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | y
651 |  |  | Japan | Tomakomai City Hospital | 341014 | 44279 |  | SC: CT application support package | 1. Treatment / quarantine for patients infected with or suspected of being infected with the new coronavirus Is it done? Yes 2. Do I need to enter the treatment room or bed where the infected patient is? (Work place lineup Including the movement route) No Operation instructions for Ingenuty CT and ISP11n [Work content] Operation explanation based on the manual | 5 |  | Operation explanation based on the manual | [12 empty] | 117242197 | Closed | 748468 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | NAC | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
652 |  |  | Chile | HOSPITAL BARROS LUCO TRUDEAU | 341005 | 44284 | 13 | CL1688 THE SYSTEM TURNED OFF. | Work Order #: WO-03236030 31-Mar-2021 Valenzuela Cariaga, Luciano The system fully meets specification and returned to use Resolution: FRC gantry power source is changed. Team tests ok. Operative team. Problem Description by Engineer: Problem description by engineer: Gantry shuts down Information to support the complaint handling process Customer Function / Role: Technician Device Use: Diagnostic Expected and Unexpected results: No errors Patient Impact / Outcome: Cancellation of exams. User Impact: Cancellation of exams. Current Software Version: 4.1.10 Parts Consumed: 1.00000,453567057167, FRC ASSEMBLY Test & Verification: tests and verifications ok | 2 |  | 1.00000,453567057167, FRC ASSEMBLY After replace FRC system works well--- Closed | [12 empty] | 117283656 | Closed | 752814 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | FRC-Replace | [5 empty] | Engineering | [6 empty] | FRC | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | FRC-Replace | Gantry-HW-FRC-Replace | Gantry-HW-FRC-Replace | FRC-Replace | FRC-Replace | Gantry-HW | Gantry-HW-FRC-Replace | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
653 |  |  | China | GD Dongguan Dalingshan Hosp. | 345013 | 44284 | 13 | 02N/Ingenuity Core128/Elite China/fault | BU Q&R Feedback: No Reason: Corrective Maintenance T2 Activities:Problem Description: Summary of customer_s problems: failure related to computer software. Detailed description of Customer_s problems or malfunctions: Often the patient turns gray and cannot be used, and it will recover after restarting It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: Software issue. Remote troubleshoot and test result: Software problem. After reboot, it can be used again, but it occurred several times. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Checked the software and collected bugreport SAP Transfer:Case Transferred to SAP Safety Questions:Case: 0117277749 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Jianqiang Last updated by: Wang Jianqiang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Problem Reported by Customer: fault SAP Transfer:Dispatched for On site Support ----------------------------------------------------------------------------- | 4 |  | Yellow screen, bugreport sent out already for analysis to detemine if in SP1 scope | [12 empty] | 117277749 | In Process | 750561 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00631367 | [5 empty] | Engineering | [6 empty] | New Yellow screen | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
654 |  |  | Japan | Tomakomai City Hospital | 341014 | 44284 | 13 | JY: Laser marker is dark | BU Q&R Feedback: No Reason: Corrective Maintenance Safety Questions:Case: 0117277875 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: YAN Joanna Last updated by: YAN Joanna 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Problem Reported by Customer: JY: Laser marker is dark SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Support OneEMS internal remarks: Sugahara DB ----------------------------------------------------------------------------- | 5 |  | Laser marker is dark | [12 empty] | 117277875 | New | 750675 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-RFC | [5 empty] | Engineering | [6 empty] | RFC | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | LaserRFC | Gantry-HW-LaserRFC | Gantry-HW-GantryLaser | RFC | gantry laser | Gantry-HW | Gantry-HW-LaserRFC | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
655 |  |  | United States | Holston Medical Group | 40053 | 44284 | 13 | No power to gantry, enabled the blue button on the gantry to power on the system | Work Order #: WO-03228426 Sheffield, Scott The system does not meet specification and is out of use Problem Description by Engineer : Customer Function/Role: tech Device Use: boot up Expected and Unexpected Results: no power to gantry User Impact: delay Patient Impact: none Current Software Version: Diagnostic performed by Engineer : gantry will not turn on. This is a new install. I was working with Jim and I had him look for a start / stop button and he could not fine one. The breaker was on. I had him look for teal but what he described did not sound like the teal. I spoke to FSE Preston and he said that the blue button was off and he will go to site and turn in it on. Follow Up Required : gantry will not turn on. This is a new install. I was working with Jim and I had him look for a start / stop button and he could not fine one. The breaker was on. I had him look for teal but what he described did not sound like the teal. I spoke to FSE Preston and he said that the blue button was off and he will go to site and turn in it on. Jim Rhoten cell 276-608-6467 ----------------------------------------------------------------------------- Work Order #: WO-03228692 30-Mar-2021 McDowell Preston The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: tech Device Use: boot up Expected and Unexpected Results: no power to gantry User Impact: delay Patient Impact: none Current Software Version: Enabled the blue button on the gantry to power on the System.Verified proper turn on of system, and performed Tube conditioning. Instructed customer to wait for one hour and then perform Air Calibration, prior to Application support arriving later today. Resolution : Enabled the blue button on the gantry to power on the System.Verified proper turn on of system, and performed Tube conditioning. Instructed customer to wait for one hour and then perform Air Calibration, prior to Application support arriving later today. ----------------------------------------------------------------------------- | 2 |  | blue button on the gantry not release | [12 empty] | 117281444 | Incomplete | 748848 | TRUE | Closed - Done | Closed - Done | [7 empty] | Service | [6 empty] | Gantry reset | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
656 |  |  | Japan | Nikko Memorial Hospital | 341011 | 44285 | 13 | CZ:Delivery of vacuum cleaner due to failure. | BU Q&R Feedback: No Reason: Corrective Maintenance SAP Transfer: Case Transferred to SAP Safety Questions:Case: 0117285448 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: ZHAO Cindy Last updated by: ZHAO Cindy 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Problem Reported by Customer: CZ:Delivery of vacuum cleaner due to failure. OneEMS internal remarks: Hashimoto DB SAP Transfer: Dispatched for Onsite Support ----------------------------------------------------------------------------- | 5 |  | Delivery of vacuum cleaner due to failure. | [12 empty] | 117285448 | New | 755468 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Others | [6 empty] | Others | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
657 |  |  | Japan | Nikko Memorial Hospital | 341011 | 44285 | 13 | JS:Replacement of the relief mat. | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: JS:Replacement of the relief mat. Safety Questions:Case: 0117285360 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: SHA Joy Last updated by: SHA Joy 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A OneEMS internal remarks: Hashimoto DB SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Support ----------------------------------------------------------------------------- | 5 |  | Replacement of the relief mat. | [12 empty] | 117285360 | New | 755457 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | TBD | [6 empty] | NAC？ | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
658 |  |  | Japan | Tomakomai City Hospital | 341014 | 44285 | 13 | CZ:Replacement of the relief mat. | BU Q&R Feedback: No Reason: Corrective Maintenance OneEMS internal remarks: Hashimoto DB SAP Transfer: Dispatched for Onsite Support Problem Reported by Customer: CZ:Replacement of the relief mat. Safety Questions:Case: 0117285333 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: ZHAO Cindy Last updated by: ZHAO Cindy 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 5 |  | Replacement of the relief mat. | [12 empty] | 117285333 | New | 755445 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | TBD | [6 empty] | NAC？ | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
659 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 341006 | 44285 | 13 | Recognizable image artifact | Work Order #: WO-03234513 Soner Saglam The system partly meets specification and is returned to use with restrictions Problem Description by Engineer: Artefact problem was reported. Resolution : Detector check and artefact diagnosis with site visit. Follow-Up Required: There is no RSN connection. Detector check and artefact diagnosis with site visit. ----------------------------------------------------------------------------- Work Order #: WO-03234837 31-Mar-2021 Turgay Baylan The system fully meets specification and returned to use Problem Description by Engineer: Artefact problem was reported. Resolution : There was no connection between the CT scanner and monitor. HDMI cable was malfunctioning. The problem was solved after cable replacement. Contrast agents in the detector were cleaned. A test scan was completed after air calibration. No artefacts were found and the device was delivered in working condition. ----------------------------------------------------------------------------- | 3 |  | There was no connection between the CT scanner and monitor. HDMI cable was malfunctioning. The problem was solved after cable replacement. Contrast agents in the detector were cleaned. A test scan was completed after air calibration. | [12 empty] | 117287736 | Closed | 752139 | TRUE | Closed - Done | Closed - Done |  | ECR-075975 | [5 empty] | Engineering | [6 empty] | Monitor cable + contrast | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Lexan Ring | Gantry-HW-Lexan Ring | Gantry-HW-Lexan Ring | ECR-075975 | Lexan Ring | Gantry-HW | Gantry-HW-Lexan Ring | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
660 |  |  | United States | Holston Medical Group | 40053 | 44285 | 13 | Replace Phantom | No info yet | 3 |  | Duplicate with 117290194 | [12 empty] | 117291869 | In Progress | 758880 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | TBD | [6 empty] | Duplicate? | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
661 |  |  | United States | Holston Medical Group | 40053 | 44285 | 13 | System Damage During Applications | Work Order #: WO-03236924 Paine Matthew The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Applications Specialist Device Use: Diagnostic Expected and Unexpected Results: Q/C Phantom Fell Off and Cracked Mylar Ring and Damaged Phantom User Impact: Unknown Patient Impact: Unknown Current Software Version:N/A Diagnostic performed by Engineer : Applications Specialist states that during Q/C Phantom Testing, the Q/C Phantom fell off the Holder onto the Mylar Ring and Damaged both the Mylar Ring and Phantom. Follow Up Required : FSE Needs to be scheduled to assess damage to Mylar Ring and Q/C Phantom. Internal Comments : No Primary or Secondary Listed . Applications Specialist stated that Preston McDowell was filling in for Primary FSE. ----------------------------------------------------------------------------- Work Order #: WO-03236991 30-Mar-2021 McDowell Preston Problem Description by Engineer : Customer Function/Role: Applications Specialist Device Use: Diagnostic Expected and Unexpected Results: Q/C Phantom Fell Off and Cracked Mylar Ring and Damaged Phantom User Impact: Unknown Patient Impact: Unknown Current Software Version:N/A Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 4 |  | The Physics phantom slipped and dropped and caused damage to the Phantom and the Lexicon Ring. Ordered both via the DEFOA appl on the Phone. Will wait to arrive to install. Caroline/Peter/Silva: communicate with APP or submitter to know the detail and do risk assement, share the H-phantom related materia for app training Carlos: can we know where FSE order the parts? from SPS or factory? | [12 empty] | 117290194 | In Progress | 753398 | FALSE | Investigation | Closed - Done | [7 empty] | Application | [6 empty] | Mistake | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
662 |  |  | China | Huaying Mountain Guangneng General Hospital | 345000 | 44286 | 13 | 02N/Ingenuity Core128/Elite China/Change login password | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure T2 Activities:Problem Description: Change login password Troubleshooting Action: Instruct customer to Change password Repair Action: Instruct customer to Change password Test/Inspection: Functional Check PASSED. The device returned to full functionality. Test Equipment: N/A SAP Transfer:Case Transferred to SAP | 3 |  | change password | [12 empty] | 117293462 | Closed | 759619 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | TBD | [6 empty] | NAC? | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
663 |  |  | United States | Holston Medical Group | 40053 | 44286 | 13 | 0798CDN82 Worklist not updating & Transfer to PACS slow or fails | Work Order #: WO-03248105 01-Apr-2021 McDowell Preston The system fully meets specification and returned to use Problem Description by Engineer : DESCRIPTION OF PROBLEM (EXPECTED/UNEXPECTED RESULTS) Worklist not updating & Transfer to PACS slow or fails CALLER FUNCTION/ROLE Technician/Sonographer/Therapist DEVICE INTENDED USE Diagnostic WHEN WAS THE FIRST OCCURRENCE? WITHIN . . . Same day HOW OFTEN IS ISSUE OCCURRING? Continuously WHERE IN THE END-USER WORKFLOW IS THE ISSUE OCCURRING? Post acquisition/scan WHAT RECENT EVENTS HAVE OCCURRED? During normal use NEGATIVE OPERATOR IMPACT Worklist and xfer to PACS problems NEGATIVE PATIENT IMPACT/PROCEDURE OUTCOME N/A HAS CUSTOMER CREATED A SPECIFIC SYSTEM EVENT REPORT OF ISSUE? No DEVICE SOFTWARE VERSION 4.1.10 changed some config info for the worklist and the Pacs destinations. Verified proper operation of both at this time. System application support training is occurring this week, and no issues at this time. Resolution : changed some config info for the worklist and the Pacs destinations. Verified proper operation of both at this time. System application support training is occurring this week, and no issues at this time. ----------------------------------------------------------------------------- | 3 |  | changed some config info for the worklist and the Pacs destinations. Verified proper operation of both at this time. System application support training is occurring this week, and no issues at this time. | [12 empty] | 117300811 | Closed | 758597 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Service | [6 empty] | PACS Config | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
664 |  |  | China | NMG Alashan Youqi People Hosp | 345031 | 44287 | 13 | 02N/Ingenuity Core128/Elite China/Add printer | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Connection/Cable/Interconnect Repair Activity Code: Instructed Customer Problem Reported by Customer : 02N/Ingenuity Core128/Elite China Safety Questions : Case: 0117305960 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Cao Last updated by: Liu Cao 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: Add printer Troubleshooting Action: Add printer Repair Action: Add printer Test/Inspection: N/A Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | Add （PACS）DICOM Node | [12 empty] | 117305960 | Closed | 765399 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | NAC | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
665 |  |  | Japan | CIAR BAHIA SA | 341017 | 44287 |  | Request a message to FSE who will visit for repair today. >> Report to FSE Kawazoe. | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer : 今日修理訪問するFSEへ伝言依頼。>>FSE川添さんに報告。 Safety Questions : Case: 0117304164 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: KENICHI KOJIMA Last updated by: KENICHI KOJIMA 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 5 |  | Customer request | [12 empty] | 117304164 | Closed | 764318 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | NAC | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
666 |  |  | Japan | CIAR BAHIA SA | 341017 | 44287 | 13 | 01DS:KKI: Image transfer to PACS stopped at 40%>>Restarting the CT-host solved the problem. | Image transfer to PACS is stopped at 40% >> Improved by restarting CT host No HF3 FCO record in One EMS | 4 |  | Need bugreport | [12 empty] | 117303034 | In Process | 763906 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS transfer | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
667 |  |  | Turkey | T.C Saglik Bakanligi Bagcilar -> Bağcılar State Hospital | 40010 | 44287 | 13 | Scans are aborted intermittently. | Work Order #: WO-03252216 Soner Saglam The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Scan was aborted. Survey related problems were reported. Resolution : The device should be diagnosed on site. SFTP cannot be performed. Check connection with COPSSH. Send bugreps to RSE. Follow Up Required : The device should be diagnosed on site. SFTP cannot be performed. Check connection with COPSSH. Send bugreps to RSE. ----------------------------------------------------------------------------- When the logs in the reported date and time range were tried to be examined, it was observed that the Logger.mdb file was stuck. For this reason, nothing could be detected because logger.mdb and user created bugrep are nonfunctional. It was suggested that the user contact the operation center when faced with the same situation. The system is active. | 3 |  | It was suggested that the user contact the operation center when faced with the same situation. The system is active. | [12 empty] | 117305332 | In Progress | 761566 | TRUE | Closed - Done | Closed - Done |  | TBD_Bugreport | [5 empty] | Engineering | [6 empty] | Scan aborted | [5 empty] | Refresh |  |  | System issue | System | Console-SW | TBD_Bugreport | Console-SW-TBD_Bugreport | Console-SW-TBD_Bugreport | TBD_Bugreport | TBD_Bugreport | Console-SW | Console-SW-TBD_Bugreport | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
668 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 341006 | 44287 | 13 | Table failure | Work Order #: WO-03251516 01-Apr-2021 Turgay Baylan The system fully meets specification and returned to use Resolution : CT scanner was turned off and on from the cable under the table. The system was delivered in working condition after servicing. Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: ----------------------------------------------------------------------------- | 3 |  | It is understood that the CT scanner device has been attached and removed the cable under the table. | [12 empty] | 117304339 | Closed | 760845 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Cable connection | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
669 |  |  | United States | Unity Health Systems | 40037 | 44287 | 13 | April Monthly QA | Checked with FSE Tom Lockwood and he requested to assign with FSE Taylor Durrance. FSE Taylor Durrance requested wo dispatched 04/08/2021 3pm-5pm Next Steps: Called Customer Andrea General and left message for her to cb to provide updated appt date and time" | 5 |  | Customer request update application training | [12 empty] | 117307720 | In Progress | 766041 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | NAC | [6 empty] | NAC? | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
670 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 341006 | 44288 | 13 | It turned off spontaneously during shooting | "Work Order #: WO-03261354 Soner Saglam The system does not meet specification and is out of use Problem Description by Engineer: The device cannot be turned on. Resolution: The system is not receiving power from the UPS. Work should be done with the UPS company. CCC was informed. | 3 |  | The system is not receiving power from the UPS. Work should be done with the UPS company. CCC was informed. | [12 empty] | 117313551 | Closed | 765818 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | UPS | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
671 |  |  | China | GD Dongguan Dalingshan Hosp. | 345013 | 44289 | 13 | WeChat：02N/Ingenuity Core128/Elite China/Disk space is too full | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : Disk space is too full Next Step Action : 18066898085 刘国平 回电 Safety Questions : Case: 0117316200 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Guomei Li Last updated by: Guomei Li 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Dispatched for Onsite Support Next Step Action : 派工 刘振华 18826107797 External Remarks : 04.04.2021 11:35:55 UTC Zhenhua Liu (86034125)故障概述：Host主机D盘被沾满，无法扫描 服务内容：重装系统和软件，恢复系统设置。 服务结果：维修后，机器恢复正常使用。 其他重要事项：无 Internal Remarks : 04.04.2021 11:35:58 UTC Zhenhua Liu (86034125) "RDI Accuracy: Yes" $Q1: () $Q2: () $Q3: () $Q4: () $Q7: (Confirmation & Description of problem) Confirmed reported problem, system could not scan because storage room is fully occupied. $Q8: (Troubleshooting & Results) check and find out Host computer disk D is fully occupied. $Q9: (Repair steps) reinstall the system and software, recover the settings. $Q10: (Test and inspection data refer to PA . Device returned to full functionality ) $Q11: Calibrated Tools name ( refer to PA ) S/N ( ) Next Calibration date ( refer to PA ) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date: N/A ^Current Tube usage: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | 故障概述：Host主机D盘被沾满，无法扫描 服务内容：重装系统和软件，恢复系统设置。 服务结果：维修后，机器恢复正常使用。 | [12 empty] | 117316200 | New | 769265 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632261 | [5 empty] | Engineering | [6 empty] | Disk full | [5 empty] | Refresh |  |  | System issue | System | Console-SW | DiskFull | Console-SW-DiskFull | Console-SW-DiskFull | CTUDT00632261 | DiskFull | Console-SW | Console-SW-DiskFull | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
672 |  |  | Japan | CIAR BAHIA SA | 341017 | 44289 | 13 | 01CS: KK: When I try to shoot three-phase after plain shooting, the mA column turns red and Tube hot is displayed and I cannot scan. >> can not you shoot somehow? So I asked the app for advice. | 01CS: KK: Hoshi: When I try to shoot three-phase after plain shooting, the mA column turns red and Tube hot is displayed and I cannot scan. >> can not you shoot somehow? So I asked the app for advice. --Problem Description-- Target device: CT Failure location: 01Cosole When did it first occur? (Within ~): 1 Hour How often do problems occur ?: First Occurrence At what point in the end-user workflow is the problem occurring ?: During inspection / measurement (During Acqustion / Scan) What happened recently ?: N.A. Does the customer create a specific system event report for the problem ?: No. The customer was unable to create a log file or report. --4 / 8Elements-- Q1 Did you confirm the reported problem? Verified Q2 Did you confirm other problems? None Q3 What did you do to solve the problem? Since the temperature of the Tube has risen, the mA of the plan turns red and scanning cannot be started. I asked him to lower the mA value, but I couldn_t solve it, so I received advice from Mr. Kujirai, an app, and the inspection was possible. Is the device after Q4 compatible? Good Q5 How did the inspection end? No extra irradiation due to stop before X-ray irradiation Q6 Presence or absence of incorrect findings (misdiagnosis, mistreatment) due to problems: None Q7 Presence or absence of unintended unnecessary exposure: None Q8 Presence or absence of artifacts: None | 4 |  | Ask field to provide a video or detail parameters for further analysis | [12 empty] | 117315878 | Fixed | 769094 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD_Tube hot_bugreport | [5 empty] | Engineering | [6 empty] | Tube Overheat | [5 empty] | Refresh |  |  | System issue | System | Gantry-SW | TubHot | Gantry-SW-TubHot | Gantry-SW-TubHot | TBD_Tube hot_bugreport | TubHot | Gantry-SW | Gantry-SW-TubHot | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
673 |  |  | China | HLJ Harbin Senhai Hospital | 345007 | 44290 |  | 02N/Core128/Add （PACS）DICOM Node | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer : Add （PACS）DICOM Node Next Step Action : 已通知值班工程师 曹镠 15969425856先跟进客户判断是否需要派工，短信已发 Safety Questions : Case: 0117317345 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Pan Chuchu Last updated by: Pan Chuchu 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 5 |  | Add （PACS）DICOM Node | [12 empty] | 117317345 | Closed | 769819 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | NAC | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
674 |  |  | China | HLJ Huachuan People Hosp. | 345018 | 44290 | 13 | 02N/Ingenuity Core128/Elite China/couch noise | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : couch noise Safety Questions : Case: 0117317309 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Zixuan Last updated by: Wang Zixuan 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 05.04.2021 09:13:00 UTC Jun Jin (86039442)服务内容：更换联轴器，做水平矫正 服务结果：设备恢复正常 Internal Remarks : 05.04.2021 09:13:03 UTC Jun Jin (86039442) $Q1: () $Q2: () $Q3: () $Q4: () $Q7:(Confirmed reported problem &couch noise ) $Q8: (check noise ，found coupling break) $Q9: (replace coupling and horizontal correction) $Q10( Test and inspection data refer to : PA . Device returned to full functionality.) $Q11( Calibrated tools na (N/A） ----------------------------------------------------------------------------- | 3 |  | 服务内容：更换联轴器，做水平矫正 服务结果：设备恢复正常 | [12 empty] | 117317309 | New | 769788 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632179 | [5 empty] | Engineering | [6 empty] | Coupling | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | Motor Coupling | Couch-HW-Motor Coupling | Couch-HW-Motor Coupling | CTUDT00632179 | Motor Coupling | Couch-HW | Couch-HW-Motor Coupling | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
675 |  |  | Japan | Saiseikai Gosu Hospital | 343000 | 44290 | 13 | 04DS:KK: Ring artifact occurred. | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 04DS:KK: A ring artifact occurred. SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Support T2 Activities:Problem Description: subject equipment: CT Problem location: 04 Gantry Exclude HV When did the problem first occur (within what period of time)?: Same day How often does the problem occur?: Continuously At what point in the end user_s work flow did the problem occur?: (During Acquisition/Scan) What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem? No.The customer has not created a log file or report. Troubleshooting Action: Result of response: Problem not solved. Problem location: 04 Gantry Exclude HV What steps did you perform to diagnose the problem? It seems the problem started with an examination performed in the evening of April 2. I suspect there is a problem with module 17 or 26. I asked Kawashige-san to respond. | 2 |  | We were able to confirm the tendency to form a ring with the replaced module. When Air calibration was performed under the target conditions, the phenomenon disappeared, but it was determined that there was a problem with the detector. [Details of work] I swapped the M16 module with the M9 and installed a new module on the M9. After replacement, make adjustments and check the operation well. [Other failures / treatments] None [Device status] Equipment can be used [Work status] done [Other] Parts used Part number: 459800911271 / Part name: TDMS V2 M.Det Class A FRU / Serial number: 33443 / Quantity: 1 | [12 empty] | 117317157 | In Process | 769655 | TRUE | Investigation | Closed - Done |  | CTUDT00632110 | [5 empty] | Engineering | [6 empty] | Artifact | [5 empty] | Refresh |  |  | System issue | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | DMS Module | DMS-HW | DMS-HW-DMS Module | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
676 |  |  | Japan | CIAR BAHIA SA | 341017 | 44291 |  | The manual data transport was stacked. | Duplicate with 117303034 | WANG, Peter |  | PACS transfer | [14 empty] | 769649 | TRUE | Investigation | Closed - Done |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS Transfer | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
677 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 341006 | 44291 | 14 | The device turns itself off while the table goes out | The system partly meets specification and is returned to use with restrictions Follow Up Required: RSN connection was broken, bugrep could not be retrieved. Site visit required. Problem Description by Engineer: The table sometimes opens an e-stop when going out. Resolution: RSN connection lost, bugrep was not received. Site visit required. | 3 |  | The table of the CT device was horizontally calibrated, the bug report was analyzed. The system was delivered with the patient and in working condition. | [12 empty] | 117319531 | In Progress | 770278 | TRUE | Closed - Done | Closed - Done |  | NotEngineering-HorizontalCalibration | [5 empty] | Engineering | [6 empty] | E-Stop Couch | [5 empty] | Refresh |  |  | System issue | System | Couch-SW | TBD_CouchEstop | Couch-SW-TBD_CouchEstop | Couch-SW-TBD_CouchEstop | TBD_CouchEstop | TBD_CouchEstop | Couch-SW | Couch-SW-TBD_CouchEstop | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
678 |  |  | United Kingdom |  | 336286 | 44285 |  | Resend errors | Work Order #: WO-03235864 01-Apr-2021 Paul Crane The system fully meets specification and returned to use Regulatory Questions : N/A Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : Logs Parts Consumed : 1.00000,453567560151,Kit, Moog Velocity Laser Power ----------------------------------------------------------------------------- Follow-up / New Job Template Hospital Name : Forth Valley Customer Contact Name (not FSE or Philips Employee) :Leeanne Henderson Contact Telephone number:01324567053 WO no: TBD Serial/Pin/Tech ID: GB101683-014-001 (new INGENUITY if number different) System out of action? No System Priority: 3 - System Restricted Customer Temperature: Yellow Required FSE skill level: 3 - Adequate/Experienced Detail of work/Problem Description: Resend errors Number of previous visits on SWO? 1 No. of FSE/s to fit: 1 Hours to fit (excluding travel time): 1 Q1 Was the equipment in ‘Clinical use’ at the time the issue was discovered?: No Q2 Was any patient or user harmed?: No Part Number : 453567560151 Description: Moog Velocity Kit Part Number : 453567538334 Description: MOOG ROTOR TX Part Number : 453567538322 Description: MOOG STATOR RX | 3 |  | 453567560151,Kit, Moog Velocity Laser Power | [12 empty] | 117289134 |  | 757253 | #N/A | #N/A | #N/A | [7 empty] | Non-Refresh | [6 empty] | Non-Refresh | [5 empty] | Non-Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | #N/A | y |  |  | y
679 |  |  | United States | NORTHWEST RADIOLOGY/MRI | 300057 | 44286 | 13 | Is receiving a pop up _crs cannot initialize_ has had to restart everything. Gave a couple more errors and then… | Work Order #: WO-03246848 01-Apr-2021 Brooks Jessica IP & Issue: 54851181; Is receiving a pop up _crs cannot initialize_ has had to restart everything. Gave a couple more errors and then finally started to work. Action taken Called site spoke to Bridget confirmed issue. Stated she ran a bug report due to same thing happened 2 weeks ago. Rebooted several time finally started but did get error stating gantry could not initialize. Called FSE Jessica Brooks accepted wo scheduled for 4.1 @ 9 am explained what the customer stated. Outcome: Dispatched Next Step Action NA Last Updated 3.31.2021 10:04 am cst ----------------------------------------------------------------------------- Assisted FSE with checking disk health on the RAID. All drives show Healthy. Reviewed logs and S1 is complaining about not being able to delete some Raw data files. She will run the disk test. Reviewed the process with her. | 5 |  | crs cannot initialize_ | [12 empty] | 117299903 |  | 762182 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD_NoEnoughInfo | [5 empty] | Engineering | [6 empty] | CIRS | [5 empty] | Refresh |  |  | Upgrade issue | System | CIRS-SW | TBD_NoEnoughInfo | CIRS-SW-TBD_NoEnoughInfo | CIRS-SW-TBD_NoEnoughInfo | TBD_NoEnoughInfo | TBD_NoEnoughInfo | CIRS-SW | CIRS-SW-TBD_NoEnoughInfo | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
680 |  |  | United States |  | 300143 | 44287 |  | fluoro & console monitors quit working | Work Order #: WO-03250390 Watson Adam Approved 3 monitors FSE needed a matching set for the control room and the CCT monitor was failing. Part # 4553-0000-8914 | 3 |  | 19__ COLOR LCD MONITOR Q19-A 4553-0000-8914 qty 3 | [12 empty] | 117302827 |  | 763760 | #N/A | #N/A | #N/A | [7 empty] | Non-Refresh | [6 empty] | Non-Refresh | [5 empty] | Non-Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | #N/A | y |  |  | y
681 |  |  | United States |  | 300057 | 44288 |  | 0998CDY73 Physicist ACR Tests Out Of Spec | Work Order #: WO-03262034 Holeman, David The system fully meets specification and returned to use Problem Description by Engineer : Problem description by engineer : Physicist ACR Tests Out Of Spec When was the first occurrence? Within... : 02-Apr-2021 08:00:00 How often is issue occurring? : First occurrence Malfunction area : Others Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: Physicist ACR Tests Out Of Spec Patient Impact/Outcome: None User Impact: Physicist ACR Tests Out Of Spec Current Software Version: 4.1.7.25026 Diagnostic performed by Engineer : CT Tech states Physicist ACR testing has some tests that are out-of-specification. Inquired what tests failed and what were measurements = Tech does not have the Physicist report. Advised Tech to acquire a copy of the Physicist ACR Report and any discrepancies and make available to the FSE onsite. Secondary FSE is onsite and is aware. Customer is requesting scheduling date/time: Monday, 4/5/21 at 1130. Follow Up Required : Error # and/or description of error : Physicist ACR Tests Out Of Spec Completed or Recommended Repair Action "Internal" : Review Physicist ACR Report and discrepancies. Ensure ACR Accreditation performance tests & measurements are in accordance with guidelines and tests. If necessary, perform any pertinent adjustments/calibrations. Internal Comments : Error # and/or description of error : Physicist ACR Tests Out Of Spec Completed or Recommended Repair Action "Internal" : Review Physicist ACR Report and discrepancies. Ensure ACR Accreditation performance tests & measurements are in accordance with guidelines and tests. If necessary, perform any pertinent adjustments/calibrations. ----------------------------------------------------------------------------- Work Order #: WO-03262081 Problem Description by Engineer : Problem description by engineer : Physicist ACR Tests Out Of Spec When was the first occurrence? Within... : 02-Apr-2021 08:00:00 How often is issue occurring? : First occurrence Malfunction area : Others Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: Physicist ACR Tests Out Of Spec Patient Impact/Outcome: None User Impact: Physicist ACR Tests Out Of Spec Current Software Version: 4.1.7.25026 Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 4 |  | Non-Refresh | [12 empty] | 117314072 |  | 766295 | #N/A | #N/A | Evaluation | [7 empty] | Non-Refresh | [6 empty] | Non-Refresh | [5 empty] | Non-Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | y
682 |  |  | Argentina | Hospital Dr. Bernardo Houssay | 40027 | 44294 |  | CD-0000035172 | ECG Cable Replacement April 05 to 09. Complete installation hours please be informed we have closed the below case as _not applicable_ due to the passed CAD of 03-31-2021. If the CAD is incorrect please have this updated in SAP and inform us the correct CAD so we may re-open the case. Replacement has been arranged via SPS. ASO 0126730921 + LINE ITEM 11 + 12NC 459801731761 - USB EXTENDER CABLE FROM IBOX TO CONSOLE (REPLECED BY 459801413791 - IBOX2 USB EXTENDER Kit) When did you detect the part was missing? During installation Is the missing item part of a larger box? Not What is the suspected reason for this part to be missing? Ordered but not shipped by factory Did you (or the FSE) put a remark on the packing list / shipping papers on the missing items? Not ** TEXT FROM FSE: I am attaching the form for MISSING part. I_m only going to order the cable. There is no evidence. One cable is missing. Get around it by buying a similar one locally. But we need the original *** | [17 empty] | DEFOA CD-35172 |  |  | DEFOA | [7 empty] | Operation | [12 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | n | y | y | y | y | y
683 |  |  | Chile | CMD Clinica Integrales Rancagu | 40008 | 44294 |  | CL1675-CHANGES IN SCHEDULE | Work Order #: WO-03291530 Alec, Preciado The system fully meets specification and returned to use. Internal Comments : Error # and/or description of error : Incorrect time Completed or Recommended Repair Action "Internal" : Support is provided with Daniel Reyes. A space is requested to perform remote connection and reconfiguration. The equipment is restarted and the correct time is confirmed. Diagnostic performed by Engineer: Support is provided with Daniel Reyes. A space is requested to perform remote connection and reconfiguration. The equipment is restarted and the correct time is confirmed. Problem Description by Engineer : Problem description by engineer : The time of the equipment is obsolete. When was the first occurrence? Within... : April 08, 2021 14:00:00 How often is issue occurring? : First occurrence Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Correct time and system working without errors. Patient Impact/Outcome: It does not impact. User Impact: The processes are not recorded with the correct time. Current Software Version: 4.1.10 Resolution: Error # and/or description: Incorrect time Completed or Recommended Repair Action "Internal" : Support is provided with Daniel Reyes. A space is requested to perform remote connection and reconfiguration. The equipment is restarted and the correct time is confirmed. ----------------------------------------------------------------------------- | 3 |  | A space is requested to perform remote connection and reconfiguration. The equipment is restarted and the correct time is confirmed. | [12 empty] | 117342245 | Closed | 783822 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Service | [6 empty] | Remote configuration | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
684 |  |  | China | JX Yugan Chudong hospital | 337500 | 44296 |  | WeChat：02N/Ingenuity Core128/Elite China/Image artifacts | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : Image artifacts T2 Activities : Problem Description: 客户问题概述： 使用内听道扫描序列，医生反应有细条环状伪影1 设备本身技术问题： 是。 Troubleshooting Action: 远程诊断与测试结果： 怀疑球管焦点，准直器校准问题 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议工程师到场检查球管焦点，准直器校准。。 SAP Transfer : Dispatched for Onsite Support SAP Transfer : Case Transferred to SAP Safety Questions : Case: 0117350270 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Zixuan Last updated by: Wang Zixuan | 5 |  | 检查射线通路，发现Aplane下方，球管射线口附近有一金属片，去掉后金属片 system running 2 months with intermitent artifact The foreign object is the chip that used for hock the DFS slit, normal assemble together with tube, total 4 chips with al material 453567268731. AI: Need gantry and manufucturing team investigation | [12 empty] | 117350270 | In Process | 791265 | FALSE | Investigation | Closed - Done |  | CTUDT00632833 | [5 empty] | Engineering | [6 empty] | Flat Spring | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | FlatSpring | Gantry-HW-FlatSpring | Gantry-HW-FlatSpring | CTUDT00632833 | FlatSpring | Gantry-HW | Gantry-HW-FlatSpring | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
685 |  |  | China | SD Tai_an T.C.M. Hosp. No.2 Ho | 345037 | 44294 |  | 02N/Ingenuity Core128/Elite China/Image artifact | T2 Activities : Problem Description: 客户问题概述： 间隙性伪影 客户问题详细描述或故障现象： 间隙性伪影 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： N/A 远程诊断与测试结果： N/A PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 检查DMS的加热模块，检查DMS温度恒定性. 建议备件12NC及名称： N/A 客户需求备注： 无 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : chen tao External Remarks : 11.04.2021 09:48:56 UTC Tao Chen (86038373)故障现象: 图像间歇性伪影 服务内容: 检查射线通路，发现Aplane下方，球管射线口附近有一金属片，去掉后金属片 服务结果: 维修后,设备恢复正常. | 3 |  | 检查射线通路，发现Aplane下方，球管射线口附近有一金属片，去掉后金属片 system running 2 months with intermitent artifact The foreign object is the chip that used for hock the DFS slit, normal assemble together with tube, total 4 chips with al material 453567268731. AI: Need gantry and manufucturing team investigation | [12 empty] | 117337760 | In Process | 785302 | TRUE | Investigation | Closed - Done |  | CTUDT00632833 | [5 empty] | Engineering | [6 empty] | Flat Spring | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | FlatSpring | Gantry-HW-FlatSpring | Gantry-HW-FlatSpring | CTUDT00632833 | FlatSpring | Gantry-HW | Gantry-HW-FlatSpring | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
686 |  |  | China | XJ Yili County Friendship Hospital | 40048 | 44293 |  | 02N/Ingenuity CT 728326/Cooperate with equipment acceptance | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity CT 728326 Safety Questions : Case: 0117331016 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingsong Last updated by: Lu Mingsong 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 配合设备验收 客户问题详细描述或故障现象： 配合设备验收 设备本身技术问题： 否。 Troubleshooting Action: 关键故障代码或现象： N/A 远程诊断与测试结果： N/A PRS TR状态： 无PRS。 Repair Action: 建议现场维修步骤： FSE到场，配合设备验收 建议备件12NC及名称： N/A 客户需求备注： 无 SAP Transfer : Dispatched for Onsite Support SAP Transfer : Case Transferred to SAP Next Step Action : sang jianjun ----------------------------------------------------------------------------- | 5 |  | 服务内容: 系统验收时发现无法扫描，现场检查发现主机无法连接到重建计算机，经检查网线故障，重新连接网线，系统测试正常，系统正常验收 服务结果：维修后，设备恢复正常工作 | [12 empty] | 117331016 | In Process | 780935 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Service | [6 empty] | Cable connection | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
687 |  |  | China |  | 345037 | 44293 |  | 02N/Ingenuity Core128/Elite China/Consult about couch adherence | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Next Step Action : FSE 陈涛18560023293 call in and disc him Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ Safety Questions : Case: 0117330998 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Cao Last updated by: Liu Cao 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: FSE consult about couch cushion adherence Troubleshooting Action: FSE consult about couch cushion adherence Repair Action: FSE consult about couch cushion adherence Test/Inspection: N/A Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | FSE consult about couch cushion adherence | [12 empty] | 117330998 | Fixed | 780926 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Customer | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
688 |  |  | China |  | 345012 | 44293 |  | 02N/Ingenuity Core128/Handover | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer : 02N/Ingenuity Core128 T2 Activities : Problem Description: 客户问题概述： 到场配合验收。 客户问题详细描述或故障现象： 到场配合验收。 设备本身技术问题： 否。 Troubleshooting Action: 关键故障代码或现象： NA 远程诊断与测试结果： 到场配合验收。 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 到场配合验收。 Test/Inspection: N/A Test Equipment: N/A SAP Transfer : Dispatched for Onsite Support SAP Transfer : Case Transferred to SAP External Remarks : 08.04.2021 03:39:40 UTC SUP JCOUSER (86036224)服务内容：车载CT到场配合验收。到场检查并测试车载CT功能，测试通过。 服务结果：维修后，设备恢复正常工作。 Internal Remarks : 08.04.2021 03:39:42 UTC Feng Li (86036224) RDI Accuracy: Yes $Q7: (Confirmed reported problem.) $Q8: (The Car CT on site check and pass test.) $Q9: (The Car CT on site check and pass test.) $Q10: (Test and inspection data refer to: PA.+Device returned to full functionality.) $Q11: Calibrated Tools name (FLUKE 287C Multimeter) S/N (16490128) NextCalibration date (10/20/2021) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date: NA ^Current Tube usage: NA Safety Questions : Case: 0117329536 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Cao Last updated by: Liu Cao 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 5 |  | 车载CT到场配合验收。到场检查并测试车载CT功能，测试通过。 | [12 empty] | 117329536 | Closed | 780040 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Customer | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
689 |  |  | China |  | 345011 | 44293 |  | 02N/Core128/Cooperate with equipment acceptance | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Core128/ Next Step Action : FSE:李枫-18650369579 来电 SAP Transfer : Dispatched for Onsite Support SAP Transfer : Case Transferred to SAP External Remarks : 08.04.2021 04:36:12 UTC SUP JCOUSER (86036224)服务内容：车载CT到场配合验收。到场检查并测试车载CT功能，测试通过。 服务结果：维修后，设备恢复正常工作。 Internal Remarks : 08.04.2021 04:36:19 UTC Feng Li (86036224) RDI Accuracy: Yes $Q7: (Confirmed reported problem.) $Q8: (The Car CT on site check and pass test.) $Q9: (The Car CT on site check and pass test.) $Q10: (Test and inspection data refer to: PA.+Device returned to full functionality.) $Q11: Calibrated Tools name (FLUKE 287C Multimeter) S/N (16490128) NextCalibration date (10/20/2021) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date: NA ^Current Tube usage: NA Safety Questions : Case: 0117329511 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingsong Last updated by: Lu Mingsong 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 配合设备验收 客户问题详细描述或故障现象： 配合设备验收 设备本身技术问题： 否。 Troubleshooting Action: 关键故障代码或现象： N/A 远程诊断与测试结果： N/A PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： FSE去现场配合设备验收 建议备件12NC及名称： N/A 客户需求备注： 无 ----------------------------------------------------------------------------- | 5 |  | 车载CT到场配合验收。到场检查并测试车载CT功能，测试通过。 | [12 empty] | 117329511 | Closed | 780022 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Customer | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
690 |  |  | China | SD Tai_an T.C.M. Hosp. No.2 Ho | 345037 | 44292 |  | 02n/Ingenuity Core128/Elite China/Image artifacts | BU Q&R Feedback: No Reason: Corrective Maintenance Next Step Action : fse chen tao SAP Transfer : Dispatched for Onsite Support Problem Reported by Customer : Image artifacts Safety Questions : Case: 0117325521 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Lili Last updated by: Lu Lili 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP External Remarks : 07.04.2021 13:35:20 UTC Tao Chen (86038373)故障现象: 图像暗 服务内容: 室内空调未开启，开启稳定后做空气校正 服务结果: 维修后,设备恢复正常. Internal Remarks : 10.04.2021 15:29:40 UTC Tao Chen (86038373) “RDI Accuracy: YES" $Q7: (Confirmed reported problem, low CT number on CT images intermittently) $Q8: ( Check plot of reference detector, found it was normal. Check the temperature of image room and air-conditioner, found the temperature was high and air-conditioner was off) $Q9: ( Turn on air-conditioner to decrease image room temperature) $Q10: ( Test and inspection data refer to: PA. Device returned to full functionality) $Q11: Calibrated Tools name ( N/A) ^Current Tube installed date: N/A ^Current Tube usage: N/A ----------------------------------------------------------------------------- | 3 |  | 室内空调未开启，开启稳定后做空气校正 | [12 empty] | 117325521 | New | 777212 | TRUE | Closed - Done | Closed - Done | [7 empty] | Customer | [6 empty] | Artifact | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
691 |  |  | China | HLJ Harbin Senhai Hospital | 345007 | 44292 |  | Recognizable image artifact occurred, added bad detector channel | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02n/Ingenuity Core128/Elite China/ Next Step Action : 045151131578 call in Safety Questions : Case: 0117323867 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingsong Last updated by: Lu Mingsong 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 图像伪影 客户问题详细描述或故障现象： 图像伪影 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： N/A 远程诊断与测试结果： N/A PRS TR状态： 无PRS。 Repair Action: 建议现场维修步骤： 对伪影图像进行原始数据分析，做探测器坏通道检测，必要时更换探测器. 建议备件12NC及名称： TDMS V2 M.Det Class B FRU: 459800911281 客户需求备注： 无 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : PPD 459800911281*1 已安排24H到医院 倪刚-86-18686766464 External Remarks : 07.04.2021 05:15:18 UTC Gang Ni (86038229)服务内容：关闭坏通道 服务结果：设备恢复 Internal Remarks : 07.04.2021 11:52:58 UTC Gang Ni (86038229) $Q7: (Confirmed reported problem & artifact) $Q8: (onsite review rawdata; bad detectot & has a bad channel) $Q9: (add bad detector) $Q10: (Test and inspection data refer to: PA. & Device returned to full functionality) $Q11: Calibrated Tools name：none ----------------------------------------------------------------------------- | 3 |  | 服务内容：关闭坏通道 服务结果：设备恢复 | [12 empty] | 117323867 | Closed | 776099 | TRUE | Closed - Done | Closed - Done |  | CTUDT00632110 | [5 empty] | Engineering | [6 empty] | Artifact | [5 empty] | Refresh |  |  | System issue | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | DMS Module | DMS-HW | DMS-HW-DMS Module | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
692 |  |  | China | SD Anqiu T.C.M Hosp. | 345017 | 44292 |  | 02N/Core128/system stuck | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Calibration/Adjustment Problem Reported by Customer: 02N/Core128/ Next Step Action: Field Service Engineer: Yang Xingbo -13583235930 called in T2 Activities:Problem Description: Summary of customer_s problems: The system is lagging. It has a yellow screen. Detailed description of Customer_s problems or malfunctions: The system is lagging. It has a yellow screen. It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: The system is lagging. It has a yellow screen. Remote troubleshoot and test result: Need to go to the site to confirm. Collect bug report PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Need to go to the site to confirm. Collect bug report Test Equipment: N/A SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Safety Questions:Case: 0117323738 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Cao Last updated by: Liu Cao 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 5 |  | The system is lagging. It has a yellow screen. | [12 empty] | 117323738 | In Process | 776034 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00631367 | [5 empty] | Engineering | [6 empty] | New Yellow Screen | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
693 |  |  | China | Tangshan | 40044 | 44292 |  | 02G/Ingenuity CT 728326/artifacts | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer: artifacts Safety Questions:Case: 0117322621 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Bo Last updated by: Wang Bo 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action: Hu Yue 13703597273 contacted customer, mail IB to add operation to IA# 53081098, dispatch FSE Hu Yue Next Step Action: Op 0020 added ----------------------------------------------------------------------------- | 3 |  | Bad detector calibration and accept bad channel | [12 empty] | 117322621 | Closed | 775532 | TRUE | Closed - Done | Closed - Done |  | CTUDT00632110 | [5 empty] | Engineering | [6 empty] | Artifact | [5 empty] | Refresh |  |  | System issue | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | DMS Module | DMS-HW | DMS-HW-DMS Module | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
694 |  |  | Japan |  | 341017 | 44295 |  | CT:APP:NO:ﾎｼ様:質問 | Problem Reported by Customer: Q: ① The modality display on the PACS side is CT / SR or SR / CT. Why? Can you erase SR? (2) Is it possible to set the Series No of the Exam Summary to 4 digits? A: (1) Since the modality name of RDSR is SR, it cannot be deleted. It was confirmed that the modality notation on the CT console is unified in CT / SR. I asked the person in charge of PACS for confirmation. (2) Currently, it is not possible. I will submit CFM. | 5 |  | Contact with submitter and get detail, check if CFM raised | [12 empty] | 117346325 | Fixed | 789984 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Application | [6 empty] | DSR | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
695 |  |  | Japan |  | 341017 | 44295 |  | HY:CT application support package. | BU Q&R Feedback: No Reason: Application Support SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Support Problem Reported by Customer: HY:CT application support package. OneEMS internal remarks : Otani DB Safety Questions:Case: 0117345520 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Hui Yi Last updated by: Wang Hui Yi | 5 |  | application support package. | [12 empty] | 117345520 | New | 789589 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Application | [6 empty] | NAC? | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
696 |  |  | Japan | Higashi-Ohmi General Medical C | 341015 | 44295 |  | 01CS:KK:End exam後次の患者様登録画面に切り替わったが何も入力できない。Logout/Loginしたが青い画面のまま先に進まない。>>Alt+Ctrl+DelよりLogoutしていただき再起動にて復旧。 | Problem Reported by Customer: 01CS: KK: End exam After the exam, I switched to the next patient registration screen, but I can not enter anything. I logged out / logged in, but the blue screen remains and I can not proceed. >> Log out from Alt + Ctrl + Del and restart to recover. T2 Activities: Problem Description: Target device: CT Failure location: 01Cosole When did it first occur? (Within ~): 1 Hour How often do problems occur ?: First Occurrence At what point in the end-user workflow is the problem occurring ?: Setting up Protocol What happened recently ?: N.A. Does the customer create a specific system event report for the problem ?: No. The customer was unable to create a log file or report. Troubleshooting Action: Result after response: Improved Failure location: 01Cosole What steps were taken during the diagnosis ?: There is no particular error regarding the matter that could not be entered on the new patient registration screen. There was no particular error when logging in after that, but it seems that the app could not operate normally. Recovered by restarting the Host. I was allowed to wait and see. Repair Action: [Phenomenon confirmation] Confirmed | 4 |  | End exam After the exam, I switched to the next patient registration screen, but I can not enter anything. I logged out / logged in, but the blue screen remains and I can not proceed. | [12 empty] | 117343905 | Fixed | 788953 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00631367 | [5 empty] | Engineering | [6 empty] | New Yellow Screen | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
697 |  |  | Japan | Maki Hospital | 341017 | 44293 |  | 05CS:KKI: The X-ray irradiation sound is loud. >> I asked FSE Yamamoto to respond. | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 05DS/KKI / The X-ray irradiation sound is loud. Safety Questions:Case: 0117330273 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: KAZUMASA KIMURA Last updated by: KAZUMASA KIMURA 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: subject equipment: CT Failure location:: 05 Peripheral When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user_s work flow did the problem occur?: During mechanical movements (During geometry/mechanical movements) What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem? No.The customer has not created a log file or report. Troubleshooting Action: Result after troubleshooting: Improved Failure location:: 05 Peripheral What steps did you perform to diagnose the problem? Call description: The X-ray irradiation sound is loud. Response I reported the issue to FSE Yoshiyuki Yamamoto. Repair Action: [Confirmation of problem] Confirmed. Work details: Call description: The X-ray irradiation sound is loud. Response I reported the issue to FSE Yoshiyuki Yamamoto. 【Confirmation content】 Good. OneEMS internal remarks: --Troubleshooting Action-- Result of response: Problem solved. Failure location:: 05 Peripheral What steps did you perform to diagnose the problem? Call description: The X-ray irradiation sound is loud. Response I reported the issue to FSE Yoshiyuki Yamamoto. --Repair Action-- --Subject-- 05CS:KKI:Mr. Hoshi: The X-ray irradiation sound is loud. >> I reported the issue to FSE Yoshiyuki Yamamoto. --Problem Description-- Equipment affected: CT Failure location:: 05 Peripheral When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user_s work flow did the problem occur?: During mechanical movements (During geometry/mechanical movements) What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem? No.The customer has not created a log file or report. --4/8Elements-- Q1. Was the reported problem confirmed? Confirmed. Q2. Were any other problems confirmed? None. Q3. What did you do to resolve the problem?: Call description: The X-ray irradiation sound is loud. Response I reported the issue to FSE Yoshiyuki Yamamoto. Q4: Is the equipment operating normally following the response? Good. Q5. How did the examination finish? None. Q6. Were there any mistaken findings (misdiagnosis, mistaken treatment) due to the problem? None. Q7. Was there any unintended unnecessary exposure? None. Q8. Were there any artifacts? None. ----------------------------------------------------------------------------- | 4 |  | Sound Adjustment | [12 empty] | 117330273 | In Process | 780484 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Service | [6 empty] | Sound Adjustment | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
698 |  |  | Japan | Maki Hospital | 341017 | 44293 |  | 07CS:KKI: The telephone set power is not on.>> I told him that the telephone line would become available from April 13. | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Application Support Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer: 07CS/KKI/ The telephone set power is not on.>> I told him that the telephone line would become available from April 13. Safety Questions:Case: 0117330257 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: KAZUMASA KIMURA Last updated by: KAZUMASA KIMURA 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: subject equipment: CT Location of fault: 07 Other When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence When, in the end user_s work flow, is the problem occurring?: During Examination ・Measurement preparation (Setting up Protocol) What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem? No.The customer has not created a log file or report. Troubleshooting Action: Result after troubleshooting: Improved Location of fault: 07 Other What steps did you perform to diagnose the problem? Call description: The telephone set power is not on. Response I told him that the telephone line would become available from April 13. Repair Action: [Confirmation of problem] Confirmed. Work details: Call description: The telephone set power is not on. Response I told him that the telephone line would become available from April 13. 【Confirmation content】 Good. OneEMS internal remarks: --Troubleshooting Action-- Result of response: Problem solved. Location of fault: 07 Other What steps did you perform to diagnose the problem? Call description: The telephone set power is not on. Response I told him that the telephone line would become available from April 13. --Repair Action-- --Subject-- 07CS:KKI:Mr. Hoshi: The telephone set power is not on.>> I told him that the telephone line would become available from April 13. --Problem Description-- Equipment affected: CT Location of fault: 07 Other When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence When, in the end user_s work flow, is the problem occurring?: During Examination ・Measurement preparation (Setting up Protocol) What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem? No.The customer has not created a log file or report. --4/8Elements-- Q1. Was the reported problem confirmed? Confirmed. Q2. Were any other problems confirmed? None. Q3. What did you do to resolve the problem?: Call description: The telephone set power is not on. Response I told him that the telephone line would become available from April 13. Q4: Is the equipment operating normally following the response? Good. Q5. How did the examination finish? None. Q6. Were there any mistaken findings (misdiagnosis, mistaken treatment) due to the problem? None. Q7. Was there any unintended unnecessary exposure? None. Q8. Were there any artifacts? None. SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 5 |  | The telephone set power is not on.>> I told him that the telephone line would become available from April 13. | [12 empty] | 117330257 | Fixed | 780468 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Telephone | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
699 |  |  | Portugal | Centro Hospitalar de Vila Nova | 40030 | 44293 |  | Aquando da emissão de rx diz: XRay failed, try again" | Work Order #: WO-03289261 08-Apr-2021 Jaime Granja The system fully meets specification and returned to use Regulatory Questions : N/A Problem Description by Engineer : Customer Function/Role: Radiology Technician Device Use: Diagnostics Expected and Unexpected Results: X ray error just before an acquisition. Then it was repeated and went uneventful User Impact: None Patient Impact: None Current Software Version: 4.1.10.25202 Diagnostic performed by Engineer : Coincident with the reported error, an mA error was detected in the system error file. As a preventive measure, a filament calibration was carried out following the recommendation of the service manuals for these cases. Monitoring of some exams that ran without incident. | 4 |  | As a preventive measure, a filament calibration was carried out following the recommendation of the service manuals for these cases. Monitoring of some exams that ran without incident. | [12 empty] | 117333504 | Incomplete | 782625 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | NotEngineering-Calibration | [5 empty] | Others | [6 empty] | Xray failed | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
700 |  |  | Puerto Rico | HOSPITAL DR SUSONI INC | 40013 | 44293 |  | The table when go up, make noise | Work Order #: WO-03282937 Caldelas Aldana, Carlos Manuel The system partly meets specification and is returned to use with restrictions Internal Comments : Error # and/or description of error : noise in patient suport Completed or Recommended Repair Action "Internal" : n/a Diagnostic performed by Engineer : abnormal noise detected in vertical movement in patient support. no errors appear. Problem Description by Engineer : Problem description by engineer : abnormal noise in patient support When was the first occurrence? Within... : 07-Apr-2021 11:00:00 How often is issue occurring? : First occurrence Malfunction area : Table Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: abnormal noise in couch Patient Impact/Outcome: n/a User Impact: noise in patient support Current Software Version: Not provided Follow Up Required : What are the next Troubleshooting steps : check lubrication on patient support on vertical movment related mechanism. Discard covers problems. Error # and/or description of error : noise in patient suport Completed or Recommended Repair Action "Internal" : n/a ----------------------------------------------------------------------------- Work Order #: WO-03283341 07-Apr-2021 Ramos Rivera, Pablo The system fully meets specification and returned to use Problem Description by Engineer : Problem description by engineer : abnormal noise in patient support When was the first occurrence? Within... : 07-Apr-2021 11:00:00 How often is issue occurring? : First occurrence Malfunction area : Table Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: abnormal noise in couch when moving vertically Patient Impact/Outcome: n/a User Impact: noise in patient support Current Software Version: Not provided Resolution : Found that the vertical movement guides needed lubrication. Lubricated the guides and now there is no noise. The exam room temperature is high and humidity is too low. ----------------------------------------------------------------------------- | 3 |  | Found that the vertical movement guides needed lubrication. Lubricated the guides and now there is no noise. The exam room temperature is high and humidity is too low. | [12 empty] | 117334112 | Closed | 778984 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | E-125059 | [5 empty] | Engineering | [6 empty] | Couch Noise | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | Lubricant | Couch-HW-Lubricant | Couch-HW-Lubricant | E-125059 | Lubricant | Couch-HW | Couch-HW-Lubricant | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
701 |  |  | Spain | S.C.S. HOSP. GRAL. DE FUERTEVE | 40020 | 44292 |  | Images are not transferred to the PACS. | Work Order #: WO-03273671 Carlos Saumell Lasala The system fully meets specification and returned to use Resolution: After the transfer restart, shipments to Pacs are resumed, and the client confirms to me after two days that these shipments continue to work well. Diagnostic performed by Engineer: The transfer spooler is restarted remotely, without asking the client to restart the host computer. Problem Description by Engineer: Image transfers from Tac to Pacs have been blocked. Regulatory Questions: N / A ----------------------------------------------------------------------------- | 3 |  | The transfer spooler is restarted remotely, without asking the client to restart the host computer. | [12 empty] | 117325824 | Closed | 777378 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS Transfer | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
702 |  |  | Spain | MAZ ZARAGOZA M.A.T.E.P.S.S.11 | 40042 | 44292 |  | It makes a strange sound | Work Order #: WO-03274440 07-Apr-2021 Diego Soria The system fully meets specification and returned to use Regulatory Questions : N/A Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : In high speed rotation study it makes sharp noise in the gantry I open the gantry and turn the rotor, the noise comes from the rear area on the motor belt. Resolution: I have checked the condition of the belt for the transmission of rotation in the gantry. I only hear the noise when turning at high speeds. It is not a loud sound. I reattach the strap and run tests. I do not hear the noise again, I inform that I will come back to check and inform me if the reason for the notice is reproduced. The system is working properly. Follow Up Required: Check if the reason for the ad is reproduced. | 3 |  | I reattach the strap and run tests. I do not hear the noise again, I inform that I will come back to check and inform me if the reason for the notice is reproduced. The system is working properly. | [12 empty] | 117326472 | Closed | 776799 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | ECR-077012 | [5 empty] | Engineering | [6 empty] | Belt | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Belt | Gantry-HW-Belt | Gantry-HW-Belt | ECR-077012 | Belt | Gantry-HW | Gantry-HW-Belt | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
703 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 341006 | 44295 |  | The screen of the device is frozen | The screen of the device is frozen | 3 |  | The screen of the device is frozen Ask field to collect bugreport | [12 empty] | 117345727 | Closed | 789690 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00631367 | [5 empty] | Engineering | [6 empty] | New Yellow Screen | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
704 |  |  | Turkey | HITACHI HEALTHCARE TURKEY | 337000 | 44295 |  | Customer wants back up | Customer wants back up | 3 |  | A full back up, including new protocols, was purchased and delivered to the customer. The system was tested with the phantom in the clinic and delivered ready to use. | [12 empty] | 117346244 | In Progress | 786249 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Backup | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
705 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 341006 | 44294 |  | Freezes during shooting | Work Order #: WO-03286997 Soner Saglam The system fully meets specification and returned to use Problem Description by Engineer : SW_de donma. Resolution : There is no RSN connection. Cannot log from the system. With the field visit, the connection problem should be resolved and the issue should be examined by taking bugrep. Follow Up Required: No RSN connection. Cannot log from the system. With the field visit, the connection problem should be resolved and the issue should be examined by taking bugrep. ----------------------------------------------------------------------------- | 3 |  | Freezing in SW. | [12 empty] | 117338353 | Closed | 781324 | TRUE | Investigation | Closed - Done |  | CTUDT00631367 | [5 empty] | Engineering | [6 empty] | New Yellow Screen | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
706 |  |  | United States | Unity Health Systems | 40037 | 44297 |  | UPS beeping intermittent | No further issues after resetting UPS. OK to close. | 4 |  | No further issues after resetting UPS. OK to close. | [12 empty] | 117351331 | In Progress | 791159 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | UPS | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
707 |  |  | United States | Leesburg Regional Medical Cent | 40022 | 44295 |  | 0903CDY70 ring artifact | Expected and Unexpected results: ring artifact Patient Impact/Outcome: none User Impact: ring artifact Current Software Version: 4.1.10 Diagnostic performed by Engineer : Angelo said they are getting rings on images. I downloaded vis Secure FTP a few images he gave me and confirmed Module 7 or 36 is showing in the images. I was unable to download any bad detector tables or air tables. System locks up with this message "Unexpected close was received from the remote system for subsystem SFTP. Please contact the server_s administrator for further information. " Customer ran air cals and a test scan and stated he did not see the ring but I am not comfortable not being able to see what is going on so dispatching the case. Internal Comments : Error # and/or description of error : ring artifact Completed or Recommended Repair Action "Internal" : Contact the customer and confirm whether the air calibration resolved the issue. If not use the image quality T/S manual and determine best fix for the module. Might move it our or replace it. Confirm PN listed below prior to ordering. 4598005911281 TDMS V2 Module Class B FRU Follow Up Required : Error # and/or description of error : ring artifact Completed or Recommended Repair Action "Internal" : Contact the customer and confirm whether the air calibration resolved the issue. If not use the image quality T/S manual and determine best fix for the module. Might move it our or replace it. Confirm PN listed below prior to ordering. 4598005911281 TDMS V2 Module Class B FRU ----------------------------------------------------------------------------- Work Order #: WO-03297843 10-Apr-2021 Neuman Terry The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Problem description by engineer : ring artifact When was the first occurrence? Within... : 09-Apr-2021 How often is issue occurring? : First occurrence Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: ring artifact Patient Impact/Outcome: none User Impact: ring artifact Current Software Version: 4.1.10 Follow Up Required : replace detector module Diagnostic performed by Engineer : using FEE test determined module 36 to be bad, moved to module 41 and simptons followed ----------------------------------------------------------------------------- | 3 |  | using FEE test determined module 36 to be bad, moved to module 41 and simptons followed | [12 empty] | 117348642 | In Progress | 787933 | TRUE | Closed - Done | Closed - Done |  | CTUDT00632110 | [5 empty] | Engineering | [6 empty] | Artifact | [5 empty] | Refresh |  |  | System issue | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | DMS Module | DMS-HW | DMS-HW-DMS Module | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
708 |  |  | United States |  | 40053 | 44293 |  | XR-29 needed | Work Order #: WO-03284690 Boyd Craig The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: XR-29 needed User Impact: None Patient Impact: None Current Software Version: Resolution : Request XR-29. Certificate emailed to customer. Diagnostic performed by Engineer : New installation. Customer requesting XR-29 for system. ----------------------------------------------------------------------------- Work Order #: WO-03288357 Hirchak, Michele The system fully meets specification and returned to use Diagnostic performed by Engineer : Emailed certificate to customer and FSE. MBH Problem description by Clinical Support Specialist : Customer Function/Role: TECH Device Use: In Use Expected and Unexpected Results: XR-29 certificate request User Impact: NONE Current Software Version: 4.1 No rescan, no loss of data, no reinjection. ----------------------------------------------------------------------------- | 3 |  | New installation. Customer requesting XR-29 for system. | [12 empty] | 117335435 | Closed | 779727 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Customer | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
709 |  |  | United States | Clay County Medical Center | 40003 | 44292 |  | PRS and Network/HIS issues. | Work Order #: WO-03277360 06-Apr-2021 Stauffer Jeffrey The system fully meets specification and returned to use Diagnostic performed by Engineer : Looked at network settings in LAN Config and saw missing information. Problem Description by Engineer : Customer Function/Role: Radiology Tech Device Use: General ct exams Expected and Unexpected Results: Unexpected cannot access PRS or Worklist. User Impact: None Patient Impact: None Current Software Version: 4.1.10 Resolution : Changed LAN Config settings. ----------------------------------------------------------------------------- | 3 |  | Worklist Configuration | [12 empty] | 117328833 | Closed | 775302 | TRUE | Investigation | Closed - Done | [7 empty] | Service | [6 empty] | Worklist Configuration | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
710 |  |  | United States | Select Specialty Hospitals | 310051 | 44297 |  | 3-System Restricted | Work Order #: WO-03301930 Morrow, Dan The system does not meet specification and is out of use Diagnostic performed by Engineer : Downloaded Bug Report. Horizontal encoders not synced. Dispatching FSE. Problem Description by Engineer : Problem description by engineer : Estops keep opening reference case 0117342543 When was the first occurrence? Within... : 11-Apr-2021 How often is issue occurring? : Multiple times a day Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Unable to use Patient Impact/Outcome: none User Impact: Unable to use Current Software Version: 3.5.5.25007 Internal Comments : Error # and/or description of error : S_MC_DEVICE_HOR_ENCODERS_NOT_SYNCED S_MANMOT_HORIZONTAL_START_ERROR Completed or Recommended Repair Action "Internal" : Downloaded Bug Report. Horizontal encoders not synced. Dispatching FSE. Follow Up Required : What are the next Troubleshooting steps : ERROR_STRING : S_MC_DEVICE_HOR_ENCODERS_NOT_SYNCED Explanation : Horizontal abs and relative encoders do not match. Action : 1. Reset couch and retry, check encoders/baseboard/cable. 2. Check Fuse F3 0.5A in CCB this is 15Vdc protection for Absolute Encoder Experience: 1. Found table MOTOR COUPLING between Absolute encoder and lead screw broken. MOTOR COUPLING replaced. 2. Absolute encoder losing its calibration. Replaced Absolute encoder. 3. Coupling loose between encoders, had to tighten set screws and calibrate horizontal- 4.. Found J1 – RS232/422 connector when touched would trip out the ESTOP and generate this error. Inspection found issue to be due to the ACS pcb where the JI socket is mounted on. Replaced the ACS successfully. 5. E-stop opened at end of max in or out from table movement. Checked connector for calibrating table horizontal movement, found in one case this was still connected. 6. Replaced both encoders to solve this Error # and/or description of error : S_MC_DEVICE_HOR_ENCODERS_NOT_SYNCED S_MANMOT_HORIZONTAL_START_ERROR Completed or Recommended Repair Action "Internal" : Downloaded Bug Report. Horizontal encoders not synced. Dispatching FSE. ----------------------------------------------------------------------------- | 3 |  | Replaced absolute and positive encoders, tested system, ready for use. 1.00000,453567399251,Assy, Absolute Encoder 1.00000,453567399231,Assy, Position Encoder | [12 empty] | 117351913 | In Progress | 791610 | #N/A | #N/A | Closed - Done |  | NotEngineering-PartReplacement | [5 empty] | Engineering | [6 empty] | Couch E-Stop | [5 empty] | Non-Refresh |  |  | Upgrade issue | system | Couch-HW | Encoder | Couch-HW-Encoder | Couch-HW-Encoder | Couch-HW-Encoder | Encoder | Couch-HW | Couch-HW-Encoder | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | y
711 |  |  | United States | Select Specialty Hospitals | 310051 | 44294 |  | E-stop tripping | Work Order #: WO-03291941 Boyd Craig The system does not meet specification and is out of use Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: E-stop tripping User Impact: None Patient Impact: None Current Software Version: Diagnostic performed by Engineer : Power cycle system. Cycle E-stops. No PRS access. E-stop still tripping after CT box key is turned. Follow Up Required : Hold till morning. FSE on site to evaluate E-stop circuit. Check error logs. No remote PRS access. ----------------------------------------------------------------------------- Work Order #: WO-03292104 09-Apr-2021 Siker Shawn The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: E-stop tripping User Impact: None Patient Impact: None Current Software Version: Resolution : Recalibrated horizontal encoders on table and tested system. Performed air cals and iq check, system is operational. ----------------------------------------------------------------------------- | 2 |  | Recalibrated horizontal encoders on table and tested system. Performed air cals and iq check, system is operational. | [12 empty] | 117342543 | Closed | 783902 | TRUE | Closed - Done | Closed - Done |  | NotEngineering_CouchEstop_Redocalibration | [5 empty] | Engineering | [6 empty] | Couch E-Stop | [5 empty] | Refresh |  |  | Upgrade issue | System | NotEngineering | ShouldbeCustomer | NotEngineering-ShouldbeCustomer | NotEngineering-ShouldbeCustomer | NotEngineering-ShouldbeCustomer | NotEngineering-ShouldbeCustomer | NotEngineering | NotEngineering-ShouldbeCustomer | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
712 |  |  | United States | The Christ Hospital-Ft Wright | 300205 | 44291 |  | need password for the exam card editor/ appliction support | need password for the exam card editor/ appliction support Attached is a quote for service(s)/part(s) requested on case #0117320598 for $1,155.00. | 5 |  | Password | [12 empty] | 117320598 | New | 773992 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Application | [6 empty] | Password | [5 empty] | Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
713 |  |  | United States | Lakeland Regional Medical Cent | 320474 | 44292 |  | Q/O System is not sending Dose report to MPPS Cerner | Work Order #: WO-03274839 06-Apr-2021 Watson Adam The system fully meets specification and returned to use Resolution : After a patient is scanned, the system is generating the dose structured report and automatically sending to Dose Wise and MPPS Cerner. The MPPS Cerner port is failing to receive. Checked PWC CT against a known working CT’s IP configuration set up and they are the same. Worked with hospital IT to determine issue and found the system is attempting to send but the problem is on Cerner side. Hospital IT has opened a ticket on the Cerner side to address the issue. Problem Description by Engineer : Customer Function/Role: | 3 |  | MPPS Cerner | [12 empty] | 117326874 | Closed | 773899 | TRUE | Investigation | Investigation | [7 empty] | Others | [6 empty] | MPPS Cerner | [5 empty] | Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
714 |  |  | Argentina | Hospital Dr. Bernardo Houssay | 40027 | 44298 |  | THERE IS NO HORIZONTAL MOVEMENT DURING STUDY | problem Description by Engineer: Horizontal movement failure in study We contact engineer Christian, the equipment does not have a remote connection. Support is requested to extract the error log manually, but by metrics they do not have authorization. Follow Up Required: Follow Up Required What are the next Troubleshooting steps : Error log is checked and the failure is replicated. Check electronics in the patient table, ACS (AMC), perform flash to CCB, horizontal calibration. Mechanical check. Resolution: Firmware was re-installed in the bed controller. Diagnostic performed by Engineer: The equipment_s software status logs are checked. | 3 |  | Firmware was re-installed in the bed controller. Diagnostic performed by Engineer: The equipment_s software status logs are checked. | [12 empty] | 117355700 | Closed | 792858 | TRUE | Closed - Done | Closed - Done |  | NotEngineering_Hwconnection | [5 empty] | Engineering | [6 empty] | Horizontal Movement | [8 empty] | System issue | System | NotEngineering | ShouldbeOthers | NotEngineering-ShouldbeOthers | NotEngineering-ShouldbeOthers | NotEngineering-ShouldbeOthers | NotEngineering-ShouldbeOthers | NotEngineering | NotEngineering-ShouldbeOthers | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
715 |  |  | Chile | HOSPITAL BARROS LUCO TRUDEAU | 341005 | 44301 |  | CL1688-TEST TIME OFFSET | Expected and Unexpected Results: Time is synchronized between host and CIRS. Patient Impact/Outcome: It does not impact. User Impact: It does not impact. Current Software Version: 4.1.10 Diagnostic Performed by Engineer: Remote connection is performed, time is configured in CIRS, working correctly. |  |  | Remote connection is performed, time is configured in CIRS, working correctly. | [12 empty] | 117383785 | Closed | 807917 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | NotEngineering - TimeSync | [5 empty] | Service | [6 empty] | Time Sync | [8 empty] | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
716 |  |  | China | JX Yugan Chudong hospital | 337500 | 44305 |  | Ingenutiy Core CT Flatspring fragment |  |  |  | Design improvement and Manufuture check the assembly process | [14 empty] | 815389 | TRUE | Closed - Duplicate | Closed - Duplicate |  | CTUDT00632833 | [5 empty] | Engineering | [6 empty] | Flat Spring | [8 empty] | System issue | system | Gantry-HW | FlatSpring | Gantry-HW-FlatSpring | Gantry-HW-FlatSpring | CTUDT00632833 | FlatSpring | Gantry-HW | Gantry-HW-FlatSpring | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
717 |  |  | China | Lindian County Hospital | 345038 | 44301 |  | WeChat:02N/Core128/Tube high temperature | Customer Symptom Code: Monitor Alarm Failure Reason: Corrective Maintenance Engineer Failure Code: Incompatible Use Repair Activity Code: Instructed Customer Problem Reported by Customer : Tube high temperature Troubleshooting Action: instruct customer modify the absolute Min mAs to 50 MAX 260 Repair Action: instruct to customer Test/Inspection: Functional Check PASSED. Device returned to full functionality. |  |  | instruct customer modify the absolute Min mAs to 50 MAX 260 确认是否是由于球管真的过热，能否录视频显示详细参数 | [12 empty] | 117376681 | Fixed | 809605 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD_Tube hot_bugreport | [5 empty] | Engineering | [6 empty] | Tube Overheat | [8 empty] | System issue | System | Gantry-SW | TubHot | Gantry-SW-TubHot | Gantry-SW-TubHot | TBD_Tube hot_bugreport | TubHot | Gantry-SW | Gantry-SW-TubHot | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
718 |  |  | China | NMG Alashan Youqi People Hosp | 345031 | 44301 |  | consult image quality | Problem Description: The customer thinks the images are a bit blurry, especially after the head images are enlarged it appears blurry. Troubleshooting Action: Guided the customer to not try to use filter functions to reconstruct images. Modify different scan conditions and see if they have changed Repair Action: The customer agrees to try different reconstruction conditions to reconstruct images Test/Inspection: N/A Test Equipment: N/A SAP Transfer:Case Transferred to SAP |  |  | Guided the customer to not try to use filter functions to reconstruct images. Modify different scan conditions and see if they have changed | [12 empty] | 117376317 | Fixed | 809525 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | NotEngineering - NotcorrectParameter | [5 empty] | Application | [6 empty] | Image Quality | [8 empty] | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
719 |  |  | China | SD Chest Hosp. | 40025 | 44301 |  | Can not start up,blue screen | Problem Reported by Customer : Can not start up,blue screen T2 Activities : Problem Description: 客户问题概述： 计算机软件相关故障。 客户问题详细描述或故障现象： 设备蓝屏，启动不到登录界面 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 设备蓝屏，启动不到登录界面 远程诊断与测试结果： 设备主机启动有问题，所以PRS不通，与工程师沟通后需到场检查 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 检查软件，如有必要进行重装系统 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 15.04.2021 13:11:53 UTC SUP JCOUSER (86038405)服务内容：软件卡死后无法进入，用IST权限进入后配置。 服务结果：维修后，设备恢复正常使用。 External Remarks : 16.04.2021 10:49:48 UTC SUP JCOUSER (86038451)服务内容：软件卡死后无法进入，用IST权限进入后配置。 服务结果：维修后，设备恢复正常使用。 |  |  | 服务内容：软件卡死后无法进入，用IST权限进入后配置。 服务结果：维修后，设备恢复正常使用。 | [12 empty] | 117374910 | In Process | 809156 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00631367 | [5 empty] | Engineering | [6 empty] | Blue screen | [8 empty] | System issue | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
720 |  |  | China | Jinggangshan Hospital, the First Affiliated Hospital of Nanchang University, Jiangxi | 345002 | 44301 |  | Data loss | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer : Data loss Safety Questions : Case: 0117372980 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: Data loss Troubleshooting Action: Remotely instruct the customer to full rebuild data. Repair Action: Remotely instruct the customer to full rebuild data. Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- |  |  | Remotely instruct the customer to full rebuild data. | [12 empty] | 117372980 | Fixed | 808309 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Recon | [8 empty] | System issue | System | TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
721 |  |  | China |  | 345036 | 44300 |  | Can not log in | BU Q&R Feedback: No SAP Customer Symptom Code: MARF Customer Symptom Code: Monitor Alarm Failure Reason: Corrective Maintenance Engineer Failure Code: Incompatible Use Repair Activity Code: Instructed Customer Problem Reported by Customer : Can not log in Safety Questions : Case: 0117368253 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Bai Jingwen Last updated by: Bai Jingwen 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: consulting the olevel_service password Troubleshooting Action: instruct the new system olevel service password to customer Repair Action: instruct to customer Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: NA SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | Lu, Xueying |  | instruct the new system olevel service password to customer | [12 empty] | 117368253 | Fixed | 804981 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [14 empty] | Password | [8 empty] | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
722 |  |  | China | SD Jinan No.4 Peo. Hosp. | 40038 | 44300 |  | clean the database | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Configuration Error Repair Activity Code: Database Recreation / Recovery Problem Reported by Customer: 02N/Ingenuity CT 728326 Next Step Action: Q6 = DICOM image database have a lot of image cannot delete, need clean the database T2 Activities:Problem Description: Some files in the image database cannot be deleted Troubleshooting Action: Use PRS to clean older DICOM data Repair Action: Clean host database | Lu, Xueying |  | Use PRS to clean older DICOM data collect bugreport | [12 empty] | 117368007 | Fixed | 804834 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Disk full | [8 empty] | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
723 |  |  | China | HEN Puyang People_s Hosp | 40031 | 44300 |  | The screen turn yellow | T2 Activities : Problem Description: 客户问题概述： 每周一次黄屏故障。 客户问题详细描述或故障现象： 每周一次黄屏故障。 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 升级后，大概每周一次黄屏故障。 远程诊断与测试结果： 到场收集bugreport，提交反馈 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 到场收集bugreport，提交反馈 Test Equipment: N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 16.04.2021 13:38:26 UTC SUP JCOUSER (86038950)服务内容：扫描软件，界面黄屏，屏幕冻结，鼠标点击无响应。 服务结果：重启主机，重置软件，设备恢复正常，继续观察。 其它重要事项：无 Internal Remarks : 16.04.2021 13:38:28 UTC Yongping Luo (86038950) $Q7:Confirmed report problem，Patient examination procedure interface appears yellow screen, and the interface is frozen . $Q8:Perform visiual check on site ,Patient examination procedure interface appears yellow screen, and the interface is frozen, Restart the host, reset softwarethe APP return to normaly. $Q9:Restart the host, reset software ,Device return to full function. $Q10: test and inspection data refer to PA. Device returned to full functionality. $Q11:Calibrated Tools refer to PA. ----------------------------------------------------------------------------- |  |  | 服务内容：扫描软件，界面黄屏，屏幕冻结，鼠标点击无响应。 服务结果：重启主机，重置软件，设备恢复正常，继续观察。 | [12 empty] | 117367559 | Fixed | 804519 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00631367 | [5 empty] | Engineering | [6 empty] | New Yellow Screen | [8 empty] | System issue | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
724 |  |  | China | SD Chest Hosp. | 40025 | 44300 |  | yellow screen and consulting change password | T2 Activities : Problem Description: 客户问题概述： 其他类型故障。 客户问题详细描述或故障现象： 设备黄屏后重启主机登录提示修改密码，需要工程师到场解决 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 设备黄屏后重启主机登录提示修改密码，需要工程师到场解决 远程诊断与测试结果： 黄屏故障，重启已恢复会正常，现场查看采集故障信息，另密码还有6天过期， PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 修改密码，查看黄屏原因 SAP Transfer : Case Transferred to SAP Next Step Action : HUANG ZHAOKUN External Remarks : 15.04.2021 13:03:25 UTC Zhaokun Huang (86038405)服务内容：远程指导客户修改即将过期的登录密码。 服务结果：设置后，机器恢复正常使用。 Internal Remarks : 16.04.2021 16:30:33 UTC Zhaokun Huang (86038405) “RDI Accuracy: NO" $Q7: (Confirmed reported problem, CT login password soon to expire.) $Q8: ( CT login password need change between 180 days.) $Q9: ( Remote support customer change CT login password.) $Q10: ( Test and inspection data refer to: No PA Required. Device returned to full functionality) $Q11: Calibrated Tools name ( N/A) ^Current Tube installed date: N/A ^Current Tube usage: N/A ----------------------------------------------------------------------------- |  |  | 服务内容：远程指导客户修改即将过期的登录密码。 服务结果：设置后，机器恢复正常使用 | [12 empty] | 117366742 | In Process | 803874 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00631367 | [5 empty] | Engineering | [6 empty] | New Yellow Screen | [8 empty] | System issue | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
725 |  |  | China | JS No.2 Chinese Medicine Hosp. | 345027 | 44300 |  | The detector is selected to have sound | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : The detector is selected to have sound T2 Activities : Problem Description: 客户问题概述： 其他类型故障。 客户问题详细描述或故障现象： 机架旋转起来声音异常，需工程师到场检查 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 机架旋转有噪音 远程诊断与测试结果： 机架不旋转声音正常，旋转起来声音异常 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 检查旋转相关部件，是否和盖子有摩擦，检查皮带以及张紧轮等 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 18.04.2021 14:45:38 UTC Liangdong Qi (86036024)服务内容：调整机架前盖. 服务结果： 维修后，设备恢复正常工作. Internal Remarks : 18.04.2021 14:45:41 UTC Liangdong Qi (86036024) "RDI Accuracy: Yes“ $Q7: (Confirmed reported problem, there was abnormal sound sometimes when gantry rotating.) $Q8: (No error report, chencked the front cover of gantry and found it could have friction with rotor part.) $Q9: (Adjusted the front cover of gantry.) $Q10: (Test and Inspection data refer to PA & Device return to fully function) $Q11: Calibrated Tools (NA). ----------------------------------------------------------------------------- |  |  | 服务内容：调整机架前盖. 服务结果： 维修后，设备恢复正常工作. | [12 empty] | 117366077 | Closed | 803468 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | NotEngineering | [5 empty] | Others | [6 empty] | Gantry Noise | [8 empty] | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
726 |  |  | China | GD Dongguan Dalingshan Hosp. | 345013 | 44299 |  | Change the setting | T2 Activities:Problem Description: Summary of customer_s problems: Other types of error. Detailed description of Customer_s problems or malfunctions: Unable to enter Edit protocol. Not able to edit the scan protocol It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: Unable to enter Edit protocol. Not able to edit the scan protocol Remote troubleshoot and test result: Cannot enter Edit protocol. Unable to edit the scan protocol. Password error, cannot be solved remotely PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Checked the configuration and change the protocol according to the customer_s requirement SAP Transfer:Dispatched for On site Support ----------------------------------------------------------------------------- |  |  | Checked the configuration and change the protocol according to the customer_s requirement | [12 empty] | 117360309 | In Process | 799382 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Application | [6 empty] | User permission | [8 empty] | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
727 |  |  | China |  | 345009 | 44299 |  | need second training | BU Q&R Feedback: No Reason: Customer Information Next Step Action: Jun Yang 13817647801 notified the trainer to follow up Problem Reported by Customer: need second training Safety Questions:Case: 0117359722 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Tang Nan Last updated by: Tang Nan 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action: 013523259896 feedback, second training is going to be carried out on Friday, but need Shenfei Yun engineer to be also on site on Friday Shenfeiyun is a new type of CT workstation. This is a product that Shenzhou Company purchased from Philips and resold to the hospital. for service request, have Shenzhou Company follow up on its own. If you encounter customer request service online for ShenfeiYun workstation, currently resort all to contact: 15901552977 Ren Xiao zhe I am very busy online and will process again in the afternoon Next Step Action: 15901552977 Ren Xiaoze said he would coordinate this matter, sent SMS Next Step Action: Customer called in to cancel training, informed 15901552977 Ren Xiao zhe and YANG JUN 13817647801 ----------------------------------------------------------------------------- |  |  | Customer need second training | [12 empty] | 117359722 | Closed | 799004 | FALSE | NAC? | Closed - Non-Complaint | [14 empty] | Application Training | [8 empty] | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
728 |  |  | China | JS Nanjing M.U.A Yifu Hosp. | 345028 | 44298 |  | the system can not scan | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 05R/Core128/ FSE Notes : This PO is a Reference Number only Safety Questions : Case: 0117354733 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Jianqiang Last updated by: Wang Jianqiang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 扫描失败。 客户问题详细描述或故障现象： 保养时发现，扫描失败，有打火现象 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： "S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT 5:0 APC ARC//S_MASTER_SEQUENCER_NO_XRAY//S_MASTER_SEQUENCER_EXIT_STATE_ERR " 远程诊断与测试结果： 5:0 APM Arc PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： "Recommendation: Review the above *** HV Error Info: *** and follow the X-Ray System Troubleshooting Guide under AEL help (?) / HW Troubleshooting. 1. Put silicone grease on high voltage connector 2. Perform short tube condition 3. Run no load test in order to find fault part (Arcs are indication of a very high probability of a faulty tube, KPM and APM are very rarely the cause of arcs). 4. Check the Power Brush Blocks / slip ring for poor connections 5. Performed Filament Calibrations followed by Long Tube Conditioning." 建议备件12NC及名称： PHILIPS MRC 880 989000086352 //Spellman 80 KW RoHS APM Assy 453567560012 //459800089031 //453566495091 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 15.04.2021 13:58:05 UTC Liuping Yao (86036129)服务内容： 设备不能扫描，球管打火。 灯丝校正和长训球管。 服务结果： 维修后，设备工作正常。 Internal Remarks : 15.04.2021 13:58:08 UTC Liuping Yao (86036129) "RDI Accuracy:Yes" $Q7: (Confirm reported problem. System cannot scan.) $Q8: (Reviewed the error log and found system cannot scan for tube arcing.) $Q9: (Check the HV connectors and nothing abnormal was found. Perform filament calibration and long tube conditioning.) $Q10: (Test and inspection data refer to PA. Device returned to full functionality.) $Q11: Calibrated tools refer to PA. ----------------------------------------------------------------------------- | WANG LIN, LV |  | 服务内容： 设备不能扫描，球管打火。 灯丝校正和长训球管。 服务结果： 维修后，设备工作正常。 | [12 empty] | 117354733 | Closed | 795709 | TRUE | Closed - Done | Closed - Done |  | CTUDT00630085 | [5 empty] | Engineering | [6 empty] | Tube Arc | [8 empty] | System issue | System | IC-HW | Tub Arc | IC-HW-Tub Arc | IC-HW-Tub Arc | CTUDT00630085 | Tub Arc | IC-HW | IC-HW-Tub Arc | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
729 |  |  | China | Zhenjiang M.C.A. Hosp. | 345023 | 44298 |  | PRS Connection | BU Q&R Feedback: No Reason: Customer Relations SAP Transfer:Dispatched for On site Support Problem Reported by Customer:PRS Connection Next Step Action: FSE Ma Shen shen call in and disp him SAP Transfer:Case Transferred to SAP Safety Questions:Case: 0117353237 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Yu Zhao Last updated by: Yu Zhao 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A External Remarks: 12.04.2021 10:42:21 UTC SUP JCOUSER (86036526) Service details: PRS shows M2M is not connected, re register PRS Service outcome: after the repair, the equipment is back to normal. Internal Remarks: 12.04.2021 10:42:24 UTC Shenshen Ma (86036526) RDI Accuracy: Yes $Q7: (Confirmed reported problem, M2M can not find in PRS) $Q8: (Login the system find PSA can not find) $Q9: (Re-register the PRS) $Q10: (Test and Inspection refer to PA. device return to full function) $Q11:Calibrated Tools(N/A). ----------------------------------------------------------------------------- | Zhai, Lisa |  | PRS Connection | [12 empty] | 117353237 | Closed | 794717 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Service | [6 empty] | PRS Connection | [8 empty] | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
730 |  |  | Japan |  | 341017 | 44302 |  | CT:APP:AK:ﾎｼ様:質問 | Problem Reported by Customer: Q: I have scanned two people in the same order. I want to divide it, so please tell me how. A: We went with you from change Patient Detail. FSE Notes: This PO is a Reference Number only OneEMS internal remarks: Q1: Did you confirm the reported problem? :Confirmed Q2: Did you confirm any other problems? ;Nothing Q3: What did you do to solve the problem? 1: If resolved by phone: Explanation Q4: Is the device in good condition after the support? : Good Safety Questions: Case: 0117393778 | Xu, Xiaoting |  | Application question | [12 empty] | 117393778 | Fixed | 812529 | FALSE | NAC? | Closed - Non-Complaint | [14 empty] | Q&A | [8 empty] | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
731 |  |  | Japan | Saiseikai Gosu Hospital | 343000 | 44302 |  | 04CM:KKI:頭部ヘリカル撮影にてリングアーチファクト発生 | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 04DM:KKI:頭部ヘリカル撮影にてリングアーチファクト発生 Safety Questions : Case: 0117393444 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: KAZUMASA KIMURA Last updated by: KAZUMASA KIMURA 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | WANG LIN, LV |  | Ring artifact occurs in head helical photography | [12 empty] | 117393444 | In Process | 812498 | FALSE | Investigation | Closed - Done |  | TBD-NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Ring Artifact | [8 empty] | System issue | System | TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
732 |  |  | Japan |  | 341017 | 44302 |  | CT:APP:AK:ﾄｶﾞｼ様:質問 | Problem Reported by Customer: Q: I have scanned two people in the same order. I want to divide it, so please tell me how. A: We went with you from change Patient Detail. OneEMS internal remarks: Q1: Did you confirm the reported problem? :Confirmed Q2: Did you confirm any other problems? ;Nothing Q3: What did you do to solve the problem? 1: If resolved by phone: Explanation Q4: Is the device in good condition after the support? : Good | Xu, Xiaoting |  | Application question | [12 empty] | 117393646 | Fixed | 812511 | FALSE | Duplicate with 812529 | Closed - Non-Complaint | [14 empty] | Q&A | [8 empty] | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
733 |  |  | Japan | Higashi-Ohmi General Medical C | 341015 | 44300 |  | 01CS: TM: When I tried to shoot, Gantry cannot … | Engineer_s Comment: _============================== $ version = "" v4.1.10.25210 "" 1:56 incoming call Call content When I tried to shoot, the message Gantry cannot Initialize.Change scan parameter was displayed at the initialization stage and the shooting did not run. I asked the patient to do the test with the phantom, but it also did not run with an error. Also, pressing the Zero Set button does not change the value. When the sleeper is operated, the value of the sleeper position changes. I was a well-built patient, can that cause it? _============================== Troubleshooting Action: What steps are taken during the system diagnosis: Post-response results: Improved Failure location: 01Cosole What steps were taken during the diagnosis ?: _================================== Correspondence Check the log. Judged that it is unlikely to depend on the patient. We had you restart including CT-Host, CIRS, TEAL. After restarting, ZeroSet also responds and shooting is possible. | Zhu, Chun |  | We had you restart including CT-Host, CIRS, TEAL. After restarting, ZeroSet also responds and shooting is possible. | [12 empty] | 117371702 | In Process | 807299 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-NotEnoughInformation | [5 empty] | Others | [6 empty] | Gantry cannot Initialize | [8 empty] | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
734 |  |  | Japan | Maki Hospital | 341017 | 44299 |  | JY: Adjusting scan irradiation noise | Problem confirmation: I confirmed it. Work details: There is no way to adjust the beeping (during scans) noise on the application. I adjusted the part of the gantry that makes noise. Other problems/Measures taken: There is one. With regard to the error that occurs when images are sent to the PACS, this can be reduced by changing the application template from PACS to Remote System at external facilities. So I changed the template. Equipment status The equipment can be used. Work status Work finished. | Xu, Xiaoting |  | Since I was able to adjust the Ready sound at the start of scanning on the software, I adjusted it on the software. | [12 empty] | 117360732 | New | 799679 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Volume adjustment | [8 empty] | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
735 |  |  | Japan | Saiseikai Gosu Hospital | 343000 | 44298 |  | 04DM:TM:Vertical lines appeared in one Surview scan.Since then, there has been no abnormality in subsequent [Truncated] | T2 Activities:Problem Description: Problem description by engineer : Equipment affected: CT Problem location: 04 Gantry Exclude HV When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user_s work flow did the problem occur?: (During Acquisition/Scan) What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem?: NA Engineer_s Comment: _============================== Call description: Vertical lines appeared in one Surview scan.Since then, there has been no abnormality in subsequent scans. _============================== External Remarks : 13.04.2021 09:45:15 UTC YUKI KOBAYASHI (81100095)[4 Elements] 1. Was the reported problem confirmed?: Confirmed. 2. Were any other problems confirmed?: 3. What did you do to resolve the problem?: I am canceling this SWO as it duplicates the content of SWO53390892. We responded to this issue on April 12 under that SWO. 4. Is the equipment operating normally following the response? Good. 13-Apr-2021 81100095 FSE Handover Tool1.3.4.32612.DB13 20210115 ----------------------------------------------------------------------------- | LIU, Ankor |  | Vertical lines appeared in one Surview scan.Since then, there has been no abnormality in subsequent scans. Recommended 12Nc: PR1 459800911271: TDMS V2 M.Det Class A FRU QTY: 1 SLOC: Osaka (U615) | [12 empty] | 117352301 | Fixed | 794133 | TRUE | Investigation | Closed - Done |  | CTUDT00632110 | [5 empty] | Engineering | [6 empty] | Vertical line | [8 empty] | System issue | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | DMS Module | DMS-HW | DMS-HW-DMS Module | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
736 |  |  | Spain | MAZ ZARAGOZA M.A.T.E.P.S.S.11 | 40042 | 44298 | 15 | Unit is down | They inform me that at the end of a study, and taking the patient off the table, the gantry is turned off. It takes them a couple of starts to get the system working again. In the error log I see that when removing the patient there is an e-stop due to movement of the table not allowed. Then when starting up again, the main contactor does not close properly and you have to try again. The rest of the morning they work without problems, during the diagnosis, I carry out test studies, the system works correctly. | 2 |  | In the error log I see that when removing a patient there is an e-stop due to not allowed couch movement. Then when starting the equipment again, the main contactor does not close correctly and they have to try again. Check if have bugreport can share engineering investigation, and perform couch calibration if necessary | [12 empty] | 117357337 | Closed | 795837 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | Power Issue | [5 empty] | Engineering | [6 empty] | Couch E-Stop | [8 empty] | System issue | System | Gantry-HW | PowerOff | Gantry-HW-PowerOff | Gantry-HW-PowerOff | TBD-PowerOff | PowerOff | Gantry-HW | Gantry-HW-PowerOff | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
737 |  |  | Sweden | Avesta lasarett | 40040 | 44302 |  | Problems with phantom positioning at IQ Check | Users have problems with monthly IQ check on our two latest CTs, they center the phantom according to the laser but after scan the system asks the user to move the phantom about 6.9mm to the right, if the users do as described they get a new message to move the phantom about 22mm to left, do as the instruction states so at the next scan they get to move the phantom about 12mm to the right and then they get stuck in this loop by moving the phantom back and forth, a user had tried to push through this IQ check but gave up after 15 minutes moving back and forth. They showed me that if you redo the IQ check and at the instructions move 6.9mm to the right, only move about 1.5mm to the right, so it goes on. Thank you if your technicians can take a look at this when you come out and install FCO on DT Ludvika SE5333 and DT2 Avesta SE5301. Follow Up Required : When doing monthly check, user have problems centering phantom. Agreed to have a FSE look at this when doing FCO upgrade on WO-03073973, scheduled 28APR21. |  |  | The guidance of centering phantom is not accurate | [12 empty] | 117397366 | In Progress | 813853 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00630107 | [5 empty] | Engineering | [6 empty] | Phantom positioning | [8 empty] | System issue | System | DMS-SW | Phantom Detect | DMS-SW-Phantom Detect | DMS-SW-Phantom Detect | CTUDT00630107 | Phantom Detect | DMS-SW | DMS-SW-Phantom Detect | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
738 |  |  | Sweden | Ludvika Lasarett | 40049 | 44302 |  | Problems with phantom positioning at IQ Check | Work Order #: WO-03334617 Knudsen Martin The system partly meets specification and is returned to use with restrictions Diagnostic performed by Engineer : When doing monthly check, user have problems centering phantom. Agreed to have a FSE look at this when doing FCO upgrade on WO-033282269, not yet scheduled. Problem Description by Engineer : When doing monthly check, user have problems centering phantom. Follow Up Required : When doing monthly check, user have problems centering phantom. Agreed to have a FSE look at this when doing FCO upgrade on WO-033282269, not yet scheduled. No parts. -MK ----------------------------------------------------------------------------- |  |  | The guidance of centering phantom is not accurate | [12 empty] | 117397371 | In Progress | 813857 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00630107 | [5 empty] | Engineering | [6 empty] | Phantom positioning | [8 empty] | System issue | System | DMS-SW | Phantom Detect | DMS-SW-Phantom Detect | DMS-SW-Phantom Detect | CTUDT00630107 | Phantom Detect | DMS-SW | DMS-SW-Phantom Detect | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
739 |  |  | Turkey | T.C Saglik Bakanligi Bagcilar -> Bağcılar State Hospital | 40010 | 44305 |  | Spill Guard order | Work Order #: WO-03341490 Soner Saglam The system fully meets specification and returned to use Problem Description by Engineer: ordering spill guards for opaque. Resolution: Spill guard ordered. Follow Up Required: Spill guard has been ordered. ----------------------------------------------------------------------------- |  |  | Parts Consumed Description: 1.00000,453566136722,APRON PACKING-SPILL GUARD | [12 empty] | 117403995 | In Progress | 815531 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Others | [6 empty] | SPILL GUARD | [8 empty] | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
740 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 341006 | 44305 |  | Spill Guard siparişi | Work Order #: WO-03341457 Soner Saglam The system fully meets specification and returned to use Problem Description by Engineer : Opak kaçmasını engellemek için spill guard siparişi. Resolution : Spill guard sipariş edildi. Follow Up Required : Spill guard sipariş edildi. |  |  | Parts Consumed Description: 1.00000,453566136722,APRON PACKING-SPILL GUARD | [12 empty] | 117403969 | In Progress | 815517 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Others | [6 empty] | SPILL GUARD | [8 empty] | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
741 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 341006 | 44305 |  | There are artifacts in the images | "Work Order #: WO-03341105 Soner Saglam The system partly meets specification and is returned to use with restrictions Resolution: Aplane analyzed, no crack. Images were examined. There are 2 distant non-homogeneous rind structures. There is a high probability that opaque material got into the detector. Control with site visit is required. Problem Description by Engineer: Images contain ring artifacts. Follow Up Required: Aplane analyzed, no cracks. Images were examined. There are 2 distant non-homogeneous rind structures. There is a high probability that opaque material got into the detector. Control with site visit is required. |  |  | There is a high probability that opaque material got into the detector. | [12 empty] | 117403475 | In Progress | 815406 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | ECR-075975 | [5 empty] | Engineering | [6 empty] | Ring Artifact | [8 empty] | System issue | System | Gantry-HW | Lexan Ring | Gantry-HW-Lexan Ring | Gantry-HW-Lexan Ring | ECR-075975 | Lexan Ring | Gantry-HW | Gantry-HW-Lexan Ring | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
742 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 341006 | 44302 |  | Device freezes. | Work Order #: WO-03332907 Turgay Baylan Diagnostic performed by Engineer : N/A Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: ----------------------------------------------------------------------------- |  |  | The bug report was taken and shared with Remote Service Engineer. The system has been delivered in working condition. | [12 empty] | 117395678 | In Progress | 813185 | FALSE | Investigation | Closed - Done |  | CTUDT00631367 | [5 empty] | Engineering | [6 empty] | New Yellow Screen | [8 empty] | System issue | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
743 |  |  | Turkey | T.C Saglik Bakanligi Bagcilar -> Bağcılar State Hospital | 40010 | 44302 |  | artefact in thin section | "Work Order #: WO-03334803 Soner Saglam The system fully meets specification and returned to use Problem Description by Engineer: Artifact in thin sections. Resolution: UHR filter may have got opaque, site visit required. In the examination, it was observed that contrast material got into the UHR filter of the device. Contrast material was cleaned and air cal was applied. The system has been tested and delivered in working condition. |  |  | opaque material got into the detector. | [12 empty] | 117397528 | In Progress | 810380 | TRUE | Closed - Done | Closed - Done |  | ECR-075975 | [5 empty] | Engineering | [6 empty] | Ring Artifact | [8 empty] | System issue | System | Gantry-HW | Lexan Ring | Gantry-HW-Lexan Ring | Gantry-HW-Lexan Ring | ECR-075975 | Lexan Ring | Gantry-HW | Gantry-HW-Lexan Ring | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
744 |  |  | Turkey | T.C Saglik Bakanligi Bagcilar -> Bağcılar State Hospital | 40010 | 44301 |  | There are artifacts in images. | "Work Order #: WO-03325607 Soner Saglam The system partly meets specification and is returned to use with restrictions Problem Description by Engineer: Images have ring artefact. Resolution: There are 2 ring artefacts in the images that are not symmetrical with the screen. Site visit required. The Compensator analysis program failed to analyze. Follow Up Required: There are 2 ring artefacts that are not symmetrical with the screen. Site visit required. The Compensator analysis program failed to analyze. -------------------------------------------------- --------------------------- Work Order #: WO-03325888 15-Apr-2021 Ugur, Demirbas The system fully meets specification and returned to use Problem Description by Engineer: Images have ring artefact. Resolution: The contrast material was detected on the detector and cleaned. Air cal was done and the system was delivered in working condition. |  |  | opaque material got into the detector. | [12 empty] | 117373824 | Incomplete | 804443 | TRUE | Closed - Done | Closed - Done |  | ECR-075975 | [5 empty] | Engineering | [6 empty] | Lexan Ring/Contrast | [8 empty] | System issue | System | Gantry-HW | Lexan Ring | Gantry-HW-Lexan Ring | Gantry-HW-Lexan Ring | ECR-075975 | Lexan Ring | Gantry-HW | Gantry-HW-Lexan Ring | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
745 |  |  | United States | McLaren Macomb | 40021 | 44301 |  | Tube Heat Issue | Work Order #: WO-03328908 Mujica, David The system fully meets specification and returned to use Problem description by Clinical Support Specialist : Customer Function/Role: Technologist Device Use: In use Expected and Unexpected Results: Tube Heat Issues User Impact: none Patient Impact: none Current Software Version: 4.1 Diagnostic performed by Engineer : Michael was aided to decrease tube heat of their CT PE/Abd/Pel Exam Cards. Cycle times in the bolus tracker were decreased by 33% and a lower DRI were used to decrease tube heat. Michael was thankful for the aide and has successfully implemented these changes to the Exam Cards. Michael has no further questions or need for clinical support at this time. No rescan, no loss of data, no reinjection. |  |  | Cycle times in the bolus tracker were decreased by 33% and a lower DRI were used to decrease tube heat. | [12 empty] | 117381762 | In Progress | 806332 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD_Tube hot_bugreport | [5 empty] | Engineering | [6 empty] | Tube Overheat | [8 empty] | System issue | System | Gantry-SW | TubHot | Gantry-SW-TubHot | Gantry-SW-TubHot | TBD_Tube hot_bugreport | TubHot | Gantry-SW | Gantry-SW-TubHot | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
746 |  |  | United States | Clay County Medical Center | 40003 | 44300 |  | Fails QA scans | Work Order #: WO-03323757 14-Apr-2021 Stauffer Jeffrey The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Radiology Technician Device Use: General CT Exams Expected and Unexpected Results: Expected that the system would pass Daily QA scans. Unexpected - hard to get to pass Daily QA scan. System has you move phantoms multiple times. User Impact: Delayed scans. Patient Impact: None Current Software Version: 4.17.10 Diagnostic performed by Engineer : Looked and phantoms and found the head phantom was tilted down. Found head phantom needed moved to the left and the body phantom needed moving to the right. Resolution : Tilted head phantom was tilted back. Moved head phantom to the left and the body phantom to the right. System then ran and passed the daily QA scan. Tightened adjustment pins on phantoms for repeatable scans. ----------------------------------------------------------------------------- | WANG LIN, LV |  | Tilted head phantom was tilted back. Moved head phantom to the left and the body phantom to the right. System then ran and passed the daily QA scan. Tightened adjustment pins on phantoms for repeatable scans. | [12 empty] | 117372009 | Closed | 803376 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00630107 | [5 empty] | Engineering | [6 empty] | Phantom positioning | [8 empty] | System issue | System | DMS-SW | Phantom Detect | DMS-SW-Phantom Detect | DMS-SW-Phantom Detect | CTUDT00630107 | Phantom Detect | DMS-SW | DMS-SW-Phantom Detect | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
747 |  |  | United States |  | 40015 | 44299 |  | customer has question about acr | Work Order #: WO-03314706 Hirchak, Michele The system fully meets specification and returned to use Problem description by Clinical Support Specialist : Customer Function/Role: TECH Device Use: In Use Expected and Unexpected Results: ACR Current Software Version: 4.1.7 No rescan, no loss of data, no reinjection Diagnostic performed by Engineer : Talked about where to find the mA station. Discussed using the EC manager as a calculator. Also, emailed ACR guide. Keeping open until Friday. MBH |  |  | customer has question about acr | [12 empty] | 117363982 | Closed | 797612 | FALSE | NAC? | Closed - Non-Complaint | [14 empty] | Q&A | [8 empty] | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
748 |  |  | United States | Holston Medical Group | 40053 | 44285 |  | What did not work as they expected? Physics phantom does not securely attach to the head phantom | Physics phantom does not securely attach to the head phantom | XU, Miley |  | Design improvement and add instruction to guide phantom attachment | [14 empty] | 797302 | TRUE | Investigation | Closed - Duplicate |  | NotEngineering-Clinical CTUDT00632803 | [5 empty] | Engineering | [6 empty] | Physics Phantom | [8 empty] | System issue | system | Couch-HW | Phantom | Couch-HW-Phantom | Couch-HW-Phantom | CTUDT00632803 | Phantom | Couch-HW | Couch-HW-Phantom | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
749 |  |  | United States | Select Specialty Hospitals | 310051 | 44302 | 15 | Table cuts off the Gantry | Work Order #: WO-03337215 16-Apr-2021 Siker Shawn The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : Gullwing switch needs replacing Follow Up Required : Replace gullwing switch Thomas Marx will return monday to replace switch | 2 |  | 1.00000,453566147251,SWITCH FRNT INTERLCK 1.00000,459800712461,FRU, AMC upgrade KITS Collect bugreport for eng investigation | [12 empty] | 117399414 |  | 812254 | TRUE | Closed - Done | Closed - Done |  | TBD_CouchEstop | [5 empty] | Engineering | [6 empty] | Couch E-Stop | [5 empty] | Refresh |  |  | Upgrade issue | System | Couch-SW | TBD_CouchEstop | Couch-SW-TBD_CouchEstop | Couch-SW-TBD_CouchEstop | TBD_CouchEstop | TBD_CouchEstop | Couch-SW | Couch-SW-TBD_CouchEstop | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
750 |  |  | United States | NORTHWEST RADIOLOGY/MRI | 300057 | 44302 |  | 4.1.10 Rings on images | Work Order #: WO-03337555 Hirchak, Michele The system fully meets specification and returned to use Problem description by Clinical Support Specialist : Customer Function/Role: Tech Device Use: In Use Expected and Unexpected Results: Rings on high resolution and standard resolution. User Impact: None Patient Impact: None Current Software Version: 4.1.10 No rescan, no loss of data, no reinjection. Diagnostic performed by Engineer : LOTS did not work. Customer has rings on high res hip/knee/ankle protocol. Had her run a recon with higher levels of iDose and without iDose. Had her run another air cal. As well. Conf RSE on the line. He will page the FSE for Monday. MBH ----------------------------------------------------------------------------- Work Order #: WO-03337594 Weaver David The system fully meets specification and returned to use Problem Description by Engineer : DESCRIPTION OF PROBLEM (EXPECTED/UNEXPECTED RESULTS) Rings on images CALLER FUNCTION/ROLE Technician/Sonographer/Therapist DEVICE INTENDED USE Diagnostic WHEN WAS THE FIRST OCCURRENCE? WITHIN . . . <1 Hour HOW OFTEN IS ISSUE OCCURRING? First occurrence WHERE IN THE END-USER WORKFLOW IS THE ISSUE OCCURRING? Pre-acquisition/scan WHAT RECENT EVENTS HAVE OCCURRED? During normal use NEGATIVE OPERATOR IMPACT Can not scan patients NEGATIVE PATIENT IMPACT/PROCEDURE OUTCOME delay in scanning HAS CUSTOMER CREATED A SPECIFIC SYSTEM EVENT REPORT OF ISSUE? No DEVICE SOFTWARE VERSION 4.1.10 Diagnostic performed by Engineer : Rings on images on all scans. Resolution : Rings on images on all scans. ----------------------------------------------------------------------------- | 5 |  | Called to f/u with Bridgette. She said all rings are gone. OK to close | [12 empty] | 117399597 |  | 812141 | TRUE | Closed - Done | Closed - Done |  | NotEngieering-Clinical | [5 empty] | Application | [6 empty] | Ring Artifact | [5 empty] | Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
751 |  |  | United States |  | 310167 | 44300 |  | Accidentally hit E-Stop button | Work Order #: WO-03323734 Mujica, David The system fully meets specification and returned to use Problem description by Clinical Support Specialist : Customer Function/Role: Technologist Device Use: In use Expected and Unexpected Results: E-stop triggered User Impact: none Patient Impact: none Current Software Version: 4.1 Diagnostic performed by Engineer : Kinn accidentally hit the e-stop button on the gantry and was aided to disengage the e-stop button and shut down the CT system to restart the gantry. The shut down was successful and by turning the key on the console, the gantry successfully was initiated. The CT scanner is working properly. Kinn has no further need for clinical support. No rescan, no loss of data, no reinjection. ----------------------------------------------------------------------------- | 4 |  | Accidentally hit E-Stop button | [12 empty] | 117371981 |  | 802950 | #N/A | #N/A | Closed - No Further Investigation | [14 empty] | E-Stop | [5 empty] | Non-Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | y
752 |  |  | United States |  | 310034 | 44299 |  | nemo / dennis | Work Order #: WO-03314367 Parts Consumed : 2.00000,453567560301,Nut, Hex, Nylon 6/6 2.00000,453567560291,Set Screw-Nylon ----------------------------------------------------------------------------- | E, Sravanthi |  | nemo / dennis | [12 empty] | 117363723 |  | 801854 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [14 empty] | nemo | [5 empty] | Non-Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | y
753 |  |  | Netherlands |  | 52101 | 44300 |  | System is slow after restart due to blue screen. | Diagnostic performed by Engineer: system slow after restart bluescreen, application now restarts but crashes after exam card selection Resolution: Acornis restored from the C drive together with a service backup. System tested and released for use. | 2 |  | restored from the C drive together with a service backup. System tested and released for use. Andy, please Communicate with Local team to inform the Exteranl tool is not allowed | [12 empty] | 117369335 |  | 801611 | #N/A | #N/A | Evaluation - Pending Translation | [14 empty] | Blue screen | [5 empty] | Non-Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | y
754 |  |  | United States |  | 310034 | 44298 |  | 0308CDY70 intermittent couch issues | Work Order #: WO-03307520 Mcginnis, Mike The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Problem description by engineer : intermittent couch issues How often is issue occurring? : Intermittent Malfunction area : Table Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected results: intermittent couch issues Patient Impact/Outcome: none User Impact: intermittent couch issues Current Software Version: 3.6.8.25004 Follow Up Required : Error # and/or description of error : 2021/04/12 10:57:17.656 S_MC_DEVICE_COUCH_VERTICAL_HW_FAULT ERROR_WARNING 2021/04/12 10:57:17.750 S_MC_DEVICE_COUCH_VERTICAL_HW_FAULT ERROR_NORMAL 2021/04/12 10:57:18.828 S_COUCHMGR_UNEXPECTED_EVENT_ERROR ERROR_NORMAL 2021/04/12 10:57:34.531 S_TIMER_WATCHDOG_CAUSED_LAST_RESET ERROR_NORMAL 2021/04/12 10:57:36.640 S_MP_MPSHELL_SERIAL_LIFE_SYNC_ERROR ERROR_NORMAL Completed or Recommended Repair Action "Internal" : Check the green bar safety switch and all cables associated with it. 453567535032 KIT-ENCODER BELT; 453567247021 Limit Switch, Safety Bar Brkt Limit switch located at the end of the green vertical support; 453566491851 CBL, Vertical Limits/Brake; 453567484781 Inverter Panel Assembly Is the complete unit; 453566502431 Vertical Endcoder Assy; 459800033291 Limit Switch Assembly 459800033291 Internal Comments : Error # and/or description of error : 2021/04/12 10:57:17.656 S_MC_DEVICE_COUCH_VERTICAL_HW_FAULT ERROR_WARNING 2021/04/12 10:57:17.750 S_MC_DEVICE_COUCH_VERTICAL_HW_FAULT ERROR_NORMAL 2021/04/12 10:57:18.828 S_COUCHMGR_UNEXPECTED_EVENT_ERROR ERROR_NORMAL 2021/04/12 10:57:34.531 S_TIMER_WATCHDOG_CAUSED_LAST_RESET ERROR_NORMAL 2021/04/12 10:57:36.640 S_MP_MPSHELL_SERIAL_LIFE_SYNC_ERROR ERROR_NORMAL Completed or Recommended Repair Action "Internal" : Check the green bar safety switch and all cables associated with it. 453567535032 KIT-ENCODER BELT; 453567247021 Limit Switch, Safety Bar Brkt Limit switch located at the end of the green vertical support; 453566491851 CBL, Vertical Limits/Brake; 453567484781 Inverter Panel Assembly Is the complete unit; 453566502431 Vertical Endcoder Assy; 459800033291 Limit Switch Assembly 459800033291 Diagnostic performed by Engineer : Dennis, biomed stated they have been having issues with the couch motion. He has replaced vertical string pot, inverter and couch board. ----------------------------------------------------------------------------- Work Order #: WO-03307843 Hanlin Christopher +I756:I967 |  |  | 44298 | [12 empty] | 117357843 |  | 793409 | #N/A | #N/A | Evaluation | [14 empty] | Non-Refresh | [5 empty] | Non-Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | y
755 |  |  | Argentina | Hospital Dr. Bernardo Houssay | 40027 | 44305 |  | A DENSE, RING-SHAPED IMAGE IS DISPLAYED IN THE ACQUISITION. | Work Order #: WO-03344466 Alec, Preciado The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Problem description by engineer : Artifact after air calibration. When was the first occurrence? Within... : April 19, 2021 10:00:00 How often is issue occurring? : Multiple times a day Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected Results: Clean image, no artifact Patient Impact/Outcome: Delayed medical care and diagnosis. User Impact: Delayed patient care. Current Software Version: Not provided. Diagnostic Performed by Engineer: We contact user, field service engineer to provide feedback who will go to the site. Internal Comments : Error # and/or description of error : Artifact in the image. Completed or Recommended Repair Action "Internal" : Ring-type artifact is displayed in the image. We suggest them air calibration, bad detector, tables FS0- FS1, Aplane, compensator and external objects check-up. It does not have PRS. Follow Up Required : What are the next Troubleshooting steps : Ring-type artifact is displayed in the image. We suggest them air calibration, bad detector, tables FS0- FS1, Aplane, compensator and external objects check-up. Error # and/or description of error : Artifact in the image. Completed or Recommended Repair Action "Internal" : Ring-type artifact is displayed in the image. We suggest them air calibration, bad detector, tables FS0- FS1, Aplane, compensator and external objects check-up. It does not have PRS. ----------------------------------------------------------------------------- Work Order #: WO-03346062 Meza, Edgardo Arturo Problem Description by Engineer : Problem description by engineer : Artifact after air calibration. When was the first occurrence? Within... : April 19, 2021 10:00:00 How often is issue occurring? : Multiple times a day Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected Results: Clean image, no artifact Patient Impact/Outcome: Delayed medical care and diagnosis. User Impact: Delayed patient care. Current Software Version: Not provided. Diagnostic performed by Engineer: N/A | 3 |  | Detector module No. 2 damaged is identified. | [12 empty] | 117406681 | In Progress | 817035 |  | Investigation | Closed - Done |  | CTUDT00632110 | [5 empty] | Engineering | [6 empty] | DMS Module | [5 empty] | Refresh |  |  | System issue | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | DMS Module | DMS-HW | DMS-HW-DMS Module | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
756 |  |  | China | SD Tai_an T.C.M. Hosp. No.2 Ho | 345037 | 44311 |  | 02n/Ingenuity Core128/Elite China/artifact | Problem Reported by Customer : 02n/Ingenuity Core128/Elite China Next Step Action : FSE 陈涛18560023293 call in and disp him Next Step Action : 已回访 SAP Transfer : Case Transferred to SAP T2 Activities : Problem Description: 客户问题概述： 图像伪影 客户问题详细描述或故障现象： 图像伪影 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： NA 远程诊断与测试结果： 报修图像伪影，之前出现过类似伪影，是由于钨片的Flat Spring断裂掉到准直器造成，这次需要订三个一起换掉。 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 检查Flat Spring，检查原始数据分析是否其他原因造成伪影，做图像质量校准 建议备件12NC及名称： " 453567268731 Flat Spring" SAP Transfer : Dispatched for Onsite Support Next Step Action : PPD 453567268731*3 正在核实替代号，请先派工，金晓-86-13370562316 | 5 |  | Replace flat spring (*3) | [12 empty] | 117438939 | In Process | 839450 |  | Investigation | Closed - Done |  | CTUDT00632833 | [5 empty] | Engineering | [6 empty] | Flat Spring | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | FlatSpring | Gantry-HW-FlatSpring | Gantry-HW-FlatSpring | CTUDT00632833 | FlatSpring | Gantry-HW | Gantry-HW-FlatSpring | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
757 |  |  | China | JS No.2 Chinese Medicine Hosp. | 345027 | 44311 |  | 02n/Ingenuity Core128/Elite China | BU Q&R Feedback: No SAP Customer Symptom Code: BOOT Customer Symptom Code: Boot up/Power-up Problem Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer : 02n/Ingenuity Core128/Elite China Safety Questions : Case: 0117438385 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Cao Last updated by: Liu Cao 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: Call for logger checking Troubleshooting Action: S_tilt_C_no_move error, not more repeat. Check P1/P2 is same at 497. Could be motor or PCB issue. Repair Action: No more error repeat, FSE will follow this issue Test Equipment: N/A ----------------------------------------------------------------------------- | 3 |  | No more error repeat, FSE will follow this issue, FSE check the belt tension force is too high | [12 empty] | 117438385 | Fixed | 838961 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | ECR-077012 | [5 empty] | Engineering | [6 empty] | Tension Force | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Belt | Gantry-HW-Belt | Gantry-HW-Belt | ECR-077012 | Belt | Gantry-HW | Gantry-HW-Belt | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
758 |  |  | China | JL Songyuan Central Hosptial | 345008 | 44310 |  | 02N/Core128/Yellow screen | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : Yellow screen SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Safety Questions : Case: 0117437773 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Pan Chuchu Last updated by: Pan Chuchu 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action : FSE 华天-18946514633申请开此维修单给他后续找TSE申请新的Escalation Internal Remarks : 25.04.2021 01:28:45 UTC Tian Hua (86038446) RDI Accuracy ：No $Q7: (Confirmed reported problem&Computer Yellow screen ) $Q8: (Check the computer ，Software problem and Update Escalation ） $Q9: (Update Escalation） $Q10: (Test and inspection data refer to: PA. Device returned to full functionality) $Q11 Calibrated Tools name (FLUKE287C) S/N (27420003) Next Calibration date (03/17/2022) External Remarks : 25.04.2021 01:28:43 UTC Tian Hua (86038446)服务内容： 软件偶发黄屏，上Escalation。 服务结果： 重启后，设备恢复正常工作。 其它重要事项： 无。 ----------------------------------------------------------------------------- | 5 |  | New Yellow Screen issue on HF3, waiting workaround publish | [12 empty] | 117437773 | Closed | 838623 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00631367 | [5 empty] | Engineering | [6 empty] | New Yellow Screen | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
759 |  |  | China | SD Laiyang People Hospital | 345005 | 44310 |  | 02N/Ingenuity Core128/Elite China/software crash | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ Safety Questions : Case: 0117437575 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Guoping Last updated by: Liu Guoping 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action : Ma jun SAP Transfer : Dispatched for Onsite Support SAP Transfer : Case Transferred to SAP T2 Activities : Problem Description: 客户问题概述： 计算机软件相关故障。 客户问题详细描述或故障现象： 计算机反应速度慢，黄屏 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 计算机反应速度慢，黄屏 远程诊断与测试结果： 怀疑软件问题 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 1检查计算机剩余空间是否合适 2 检查床是否有异物 3 重装软件 建议备件12NC及名称： ----------------------------------------------------------------------------- | 3 |  | New Yellow Screen issue on HF3, waiting workaround publish | [12 empty] | 117437575 | In Process | 838420 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00631367 | [5 empty] | Engineering | [6 empty] | New Yellow Screen | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
760 |  |  | China | JL Songyuan Central Hosptial | 345008 | 44310 |  | escalation/EQ:83704366 | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer:escalation Safety Questions:Case: 0117437412 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Zhang Jing Last updated by: Zhang Jing 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 3 |  | New Yellow Screen | [12 empty] | 117437412 | New | 838302 |  |  | Closed - Non-Complaint | [14 empty] | New Yellow Screen | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
761 |  |  | China | Guangzhou No8 Peoples Hospital | 345032 | 44308 |  | 02N/Ingenuity Core128/ring artifact | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity Core128/ Next Step Action : 02036549013 查老师来电 SAP Transfer : Dispatched for Onsite Support T2 Activities : Problem Description: 客户问题概述： 图像伪影。 客户问题详细描述或故障现象： 图像半环形伪影 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 图像半环形伪影 远程诊断与测试结果： 有坏通道 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 根据伪影手册做相关的伪影故障排除 建议备件12NC及名称： 459800911281 TDMS V2 M.Det Class B FRU Safety Questions : Case: 0117423105 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Jianqiang Last updated by: Wang Jianqiang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP Next Step Action : PPD TDMS V2 M.Det Class B FRU 459800911281 1 已安排24H到医院，FSE朱光录-86-13902259150 Next Step Action : zhu guanglu External Remarks : 23.04.2021 04:04:54 UTC Guanglu Zhu (86034061)服务内容: 高密度环形伪影，12模块探测器有坏通道，更换探测器模块，完成相应校正，恢复正常 服务结果: 服务完成后,设备恢复正常工作 Internal Remarks : 25.04.2021 05:19:51 UTC Guanglu Zhu (86034061) $ RDI Accuracy: yes $Q7: ( Confirmed reported problem+ring artifact in images ) $Q8: (analysis rawdata confirm detector 177 has bad ) $Q9: (replace 12 module and finish recommended callibration ) $Q10: (Test and Inspection data refer to PA & Device status return to fully function) $Q11: Calibrated Tools name (na ) S/N (na ) Next Calibration date (na ) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date: na ^Current Tube exposure:NA ----------------------------------------------------------------------------- | 3 |  | Ring Artifact due to module 12, issue fixed by replacing module | [12 empty] | 117423105 | Closed | 831697 |  | Closed - Done | Closed - Done |  | CTUDT00632110 | [5 empty] | Engineering | [6 empty] | DMS Module | [5 empty] | Refresh |  |  | System issue | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | DMS Module | DMS-HW | DMS-HW-DMS Module | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
762 |  |  | China | SD Chest Hosp. | 40025 | 44307 |  | 02N/Ingenuity CT 728326/MODIFT TITLE | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer : 02N/Ingenuity CT 728326/ Safety Questions : Case: 0117418579 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Cao Last updated by: Liu Cao 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: Need modify image title Troubleshooting Action: Need modify image title Repair Action: Instruct customer Test/Inspection: N/A Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | Instruct customer to modify image title | [12 empty] | 117418579 | Fixed | 828482 |  | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Application | [6 empty] | Image title | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
763 |  |  | China | Guiyang Fuyou Hosp | 40033 | 44306 |  | 02N/Ingenuity CT 728326/check system | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity CT 728326/ SAP Transfer : Case Transferred to SAP T2 Activities : Problem Description: 客户问题概述： 新装机，需要工程师检查场地，教客户开关机，日常保养 设备本身技术问题： 是。 Troubleshooting Action: 远程诊断与测试结果： 新装机，需要工程师检查场地，教客户开关机，日常保养 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 新装机，需要工程师检查场地，教客户开关机，日常保养 客户需求备注： 联系报修人，确认时间 SAP Transfer : Dispatched for Onsite Support Safety Questions : Case: 0117411517 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Weng Jinming Last updated by: Weng Jinming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action : zhou zhongxiang External Remarks : 22.04.2021 16:13:07 UTC SUP JCOUSER (86038997)服务内容：确认报修故障，新装机，需要工程师检查场地，教客户开关机，日常保养 。现场指导客户。 服务结果：服务后设备恢复正常。 ----------------------------------------------------------------------------- | 5 |  | New installed system, intruct customer to power on/off and routine maintainance of system. | [12 empty] | 117411517 | In Process | 824020 |  |  | Closed - Non-Complaint | [14 empty] | NAC? | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
764 |  |  | China | Wuhan Ren _ai Hospital, Hubei Province | 345004 | 44306 |  | praise/Ingenuity Core128/HU FANGTAO | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer: praise Next Step Action: Cheng Zhenhua follow Safety Questions:Case: 0117410771 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Tang Nan Last updated by: Tang Nan 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 3 |  | No issue report from customer | [12 empty] | 117410771 | Closed | 823559 |  |  | Closed - Non-Complaint | [14 empty] | NAC? | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
765 |  |  | China | Wuhan Ren _ai Hospital, Hubei Province | 345004 | 44306 |  | 02N/Ingenuity Core128/modify CT ip | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity Core128 FSE Notes : This PO is a Reference Number only T2 Activities : Problem Description: 客户问题概述： 需要修改host IP 客户问题详细描述或故障现象： 需要修改host IP 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 需要修改host IP 远程诊断与测试结果： 远程指导客户进了LAN CONFIG，网关改不了， PRS TR状态： 不通，请工程师到场检查。 Repair Action: 建议现场维修步骤： 根据客户需要修改IP地址，重新配置PACS 打印机等等相关连接设备 Safety Questions : Case: 0117409738 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Jianqiang Last updated by: Wang Jianqiang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 21.04.2021 15:25:00 UTC SUP JCOUSER (86036493)[服务内容]：医院IP变动，重新添加节点后恢复。 [服务结果]：维修后，设备恢复正常。 Internal Remarks : 21.04.2021 15:25:03 UTC FangTao Hu (86036493) RDI Accuracy：Yes $Q7: (Confirm reported problem &components connection fault ) $Q8: (hospital IT change the IP of PACS and prints leads to connection broken) $Q9: (Add ports and adjust IP adress ,system return to normal ) $Q10: (No PA required ,Device return to full functionality) $Q11: Calibrated Tools name (N/A) S/N (N/A) Next Calibration date (N/A) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date: N/A ^Current Tube usage: scan second_N/A ----------------------------------------------------------------------------- | 3 |  | Change host IP per customer IT assignment | [12 empty] | 117409738 | Closed | 823037 |  | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | IP Config | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
766 |  |  | China | The Second People_s Hospital | 345039 | 44306 |  | wechat:02n/Ingenuity Core128/Elite China/Run slow | BU Q&R Feedback: No SAP Customer Symptom Code: HANG Customer Symptom Code: Hang/Crash Reason: Corrective Maintenance Engineer Failure Code: Wear-out/Breakage/Mechanical Failure Repair Activity Code: Reboot/Restart Problem Reported by Customer:run slow Safety Questions:Case: 0117409659 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Lili Last updated by: Lu Lili 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: host software run slow Troubleshooting Action: reboot computer Repair Action: reboot computer Test/Inspection: customer confirm can scan patient and image ok.functional Check PASSED. The device returned to full functionality. Test Equipment: N/A SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | host software run slow, instruct customer to reboot system, The device returned to full functionality. Intermitent issue, no bugreport saved, similar with YS | [12 empty] | 117409659 | Closed | 822981 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00631367 | [5 empty] | Engineering | [6 empty] | New Yellow Screen | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
767 |  |  | China | HEN Puyang People_s Hosp | 40031 | 44303 |  | 02n/Ingenuity CT 728326/tilt error | Next Step Action : 客户013903937617再次来电，转接韩工 T2 Activities : Problem Description: 客户问题概述： 角度问题 客户问题详细描述或故障现象： 角度问题 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： S TILT C POT1 PO2 KO, S TILT C POT CALIB KO 远程诊断与测试结果： 打角度后掉电，无法恢复，重启无效。 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 检查角度电位器，检查皮带张进度，检查角度电位器到GMP路相关 建议备件12NC及名称： " 459800441121 DUAL POT ASSEMBLY" SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : PPD DUAL POT ASSEMBLY 459800441121 1 已安排24H发货，FSE罗永平-86-15226131133 Next Step Action : Guo Haibin onsite，已发派工短信。 External Remarks : 17.04.2021 10:09:56 UTC Haibin Guo (86038308)"SUAT. Ready for Clinical Use local time" : "2021-04-17 18:30" External Remarks : 17.04.2021 10:22:18 UTC SUP JCOUSER (86038308)故障描述：设备无法开机；垂直编码器不匹配；重新做倾斜校正，调整编码器 服务内容：重新做倾斜校正，调整编码器 服务结果：服务 完成后，设备恢复正常使用 其他事项：无 Internal Remarks : 20.04.2021 05:28:50 UTC Haibin Guo (86038308) RDI Accuracy: YES Q7:Can not confirmed report problem，infact problem was system cannot power up $Q8:perform visiual check on site and found S TILT C POT1 PO2 KO, S TILT C POT CALIB KO and conform the tilt pot need be calibrated or replaced | 2 |  | After adjust the dual pot and belt as well as tilt calibration, system works well, check and update PM accordingly | [12 empty] | 117400375 | Closed | 818618 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | ECR-077012 | [5 empty] | Engineering | [6 empty] | Dual Pot belt | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Belt | Gantry-HW-Belt | Gantry-HW-Belt | ECR-077012 | Belt | Gantry-HW | Gantry-HW-Belt | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
768 |  |  | China | XIFENG CHINESE MEDICINE HOSPIT | 345014 | 44305 |  | 02n/Ingenuity Core128/The Gantry cann_t powerup | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 02n/Ingenuity Core128/Elite China SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support T2 Activities:Problem Description: Summary of customer_s problems: There is an E-stop closing prompt box. After twisting the key, there is no relay suction sound. It is a technology problem with the equipment itself: Yes. Troubleshooting Action: remote diagnosis and testing result: Suspect power supply, doctor_s operation issue PRS TR status: It is not through, ask engineer for maintenance on site. Repair Action: Suggest onsite service procedures: Suggest engineer to check power supply on site, and how the doctor can operate it.。。 Safety Questions:Case: 0117402506 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Weng Jinming Last updated by: Weng Jinming N/A External Remarks : 20.04.2021 04:02:24 UTC Guanglin Tang (86039448)"SUAT. Ready for Clinical Use local time" : "2021-04-19 11:45" External Remarks: 20.04.2021 05:14:11 UTC SUP JCOUSER (86039448) error description: prompted ESTOP close. After turned the key there is no relay suction sound. Error analysis: After turning key, HOST and GHOST failed to connect.Guided the customer to reboot the software and gantry, and the connection succeeded. Failure result: boot up success. Internal Remarks : 20.04.2021 05:14:14 UTC Guanglin Tang (86039448) "RDI Accuracy:YSE" $Q7: (Confirmed reported problem,The gantry will not boot.) $Q8: (Troubleshooting,HOST and GHOST did not connect successfully, The connection was successful after the reboot.) $Q10: (Test and Inspection data:Refer to PA.Device status: Device return to functionality.) $Q11: Calibrated Tools name: na. For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date:NA ^Current Tube usage: NA ----------------------------------------------------------------------------- | 2 |  | Check if have CAN trace of issue, CAN trace received, send to Jerry review | [12 empty] | 117402506 | Closed | 818116 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-NoEnoughInfo | [5 empty] | Engineering | [6 empty] | System Reboot | [5 empty] | Refresh |  |  | System issue | System | Gantry-SW | TBD-NoEnoughInfo | Gantry-SW-TBD-NoEnoughInfo | Gantry-SW-TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo | Gantry-SW | Gantry-SW-TBD-NoEnoughInfo | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
769 |  |  | China | SX Datong Yanggao People Hosp. | 345025 | 44304 |  | 02N/Ingenuity Core128/Elite China/fault | BU Q&R Feedback: No SAP Customer Symptom Code: HANG Customer Symptom Code: Hang/Crash Reason: Corrective Maintenance Engineer Failure Code: Configuration Error Repair Activity Code: Reboot/Restart Problem Reported by Customer: fault Safety Questions:Case: 0117402095 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Mengen Xie Last updated by: Mengen Xie 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: The printer face cannot display the added image Troubleshooting Action: reboot host Repair Action: reboot host, Look again Test/Inspection: customer confirm can scan patient and image ok.functional Check PASSED. The device returned to full functionality. Test Equipment: N/A SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | The printer face cannot display the added image, reboot host, Look again, customer confirm can scan patient and image ok.functional Check PASSED. The device returned to full functionality. Intermitent issue, no bugreport saved | [12 empty] | 117402095 | Fixed | 817735 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-NotEnoughInformation | [5 empty] | Others | [6 empty] | Print | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
770 |  |  | France | SCP GROUPE D IMAGERIE MEDICALE | 40019 | 44312 |  | Ordered part: | Work Order #: WO-03382675 Aurelien ZAJAC The system partly meets specification and is returned to use with restrictions Regulatory Questions : N/A Problem Description by Engineer : Customer Function/Role: user Device Use: yes Expected and Unexpected Results: n/a User Impact: no Patient Impact: no Current Software Version: Diagnostic performed by Engineer: part ordered. Resolution: Spare part requested for replacement. Follow Up Required: Order: CT Box: 459801705022 Qty: 1 Part to be delivered to the scanner console to the attention of JL Codron. ----------------------------------------------------------------------------- Work Order #: WO-03382693 ----------------------------------------------------------------------------- | 4 |  | CT Box request for replacement. Collect picture or video to indicate button issue? | [12 empty] | 117442322 | In Progress | 839666 |  |  | Closed - Non-Complaint | [14 empty] | CT Box | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
771 |  |  | France | SCP GROUPE D IMAGERIE MEDICALE | 40019 | 44308 |  | Button is out of order. | Work Order #: WO-03364005 Aurelien ZAJAC The system fully meets specification and returned to use Regulatory Questions : N/A Problem Description by Engineer : Customer Function/Role: user Device Use: yes Expected and Unexpected Results: n/a User Impact: no Patient Impact: no Current Software Version: 4.1.7 Diagnostic performed by Engineer: The customer reports that the bed lowering button is malfunctioning. The CT box must be replaced. Resolution: Service on site required. Follow Up Required: Order: CT Box: 459800484611 Qty: 1. Part to be delivered to the scanner console to the attention of JL Codron. ----------------------------------------------------------------------------- | 3 |  | CT Box request for replacement. Collect picture or video to indicate button issue? | [12 empty] | 117425041 | In Progress | 828627 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632185 | [5 empty] | Engineering | [6 empty] | CT Box | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | CT Box | Gantry-HW-CT Box | Gantry-HW-CT Box | CTUDT00632185 | CT Box | Gantry-HW | Gantry-HW-CT Box | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
772 |  |  | Japan | Maki Hospital | 341017 | 44309 |  | アーチファクト | In the image of the phantom, the central part is black. In the shooting data of the fruit, you can see a small ring-shaped artifact in the center. Regarding fruit artifacts, the person in charge of the app has reported to the customer that it is conspicuous by reducing the FOV. _================================== --Repair Action-- --Subject-- 04CS: TM: Hoshi-sama: An artifact has occurred. The central part is blackened by the phantom. Fine ring artifacts in the center of the fruit photography data >> Inspection continues. Scheduled for inspection at a later date | 4 |  | Regarding fruit artifacts, the person in charge of the app has reported to the customer that it is conspicuous by reducing the FOV. | [12 empty] | 117431484 | Fixed | 835976 |  | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Mis-use | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
773 |  |  | Japan | Higashi-Ohmi General Medical C | 341015 | 44306 |  | Application inquiry >> Koizumi will respond | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer: Application inquiry >> Koizumi will respond ----------------------------------------------------------------------------- | 5 |  | Application inquiry >> Koizumi will respond | [12 empty] | 117410792 | In Process | 823572 |  | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Application | [6 empty] | Q&A | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
774 |  |  | Japan | Tomakomai City Hospital | 341014 | 44305 |  | JY: ECG TESTER failure, replacement | BU Q&R Feedback: No Reason: Corrective Maintenance Safety Questions:Case: 0117402641 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: YAN Joanna Last updated by: YAN Joanna 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Problem Reported by Customer: JY: ECG tester failure, replacement SAP Transfer: Case Transferred to SAP OneEMS internal remarks: Hashimoto DB SAP Transfer: Dispatched for Onsite Support External Remarks:22.04.2021 05:52:37 UTC HODAKA HASHIMOTO (81100553)[Details of request] ECG tester problem. Problem confirmation: I confirmed it. Work details: I delivered an ECG simulator. Other problems/Measures taken: None. Equipment status The equipment can be used. Work status Work finished. 22-Apr-2021 81100553 FSE Handover Tool1.3.4.32612.DB13 20200715 Internal Remarks:22.04.2021 05:52:40 UTC HODAKA HASHIMOTO (81100553) Replacement parts: Replacement parts: As below 12Nc:453566095882 Name:ECG TESTER QTY:1 [Report on mistakes in parts order] No mistakes. Situation monitoring information: Not monitoring the situation TS Request: No parts to return. Additional CT information: Q1. How did the examination finish? Q1. No (not applicable) Q2. Were there any mistaken findings (misdiagnosis, mistaken treatment) due to the problem? None. Q3. Were there any artifacts? None. Q4. Was there any unintended unnecessary exposure? None. Repair sticker: I affixed a Repaired sticker. [Other information to share] 1. Was treatment/quarantine carried out for a patient who has or is suspected of having new coronavirus? Yes 2. Is it necessary to go into treatment rooms/beds where there are infected patients?(Including workplace and routes of movement) No Did you agree on how to complete the report with the person in charge of devices? YES Contacted the customer regarding the status the next day. 22-Apr-2021 81100553 FSE Handover Tool1.3.4.32612.DB13 20200715 ----------------------------------------------------------------------------- | 5 |  | ECG tester failure, replacement, order 12Nc:453566095882 Name:ECG TESTER QTY:1 Check why need order this part? | [12 empty] | 117402641 | New | 818199 |  | Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | ECG | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
775 |  |  | Japan | Nikko Memorial Hospital | 341011 | 44305 |  | JY: ECG TESTER failure, replacement | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: JY: ECG tester failure, replacement Safety Questions:Case: 0117402620 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: YAN Joanna Last updated by: YAN Joanna 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A OneEMS internal remarks: Hashimoto DB SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Support External Remarks:22.04.2021 05:51:26 UTC HODAKA HASHIMOTO (81100553)[Details of request] ECG tester problem. Problem confirmation: I confirmed it. Work details: I delivered an ECG simulator. Other problems/Measures taken: None. Equipment status The equipment can be used. Work status Work finished. 22-Apr-2021 81100553 FSE Handover Tool1.3.4.32612.DB13 20200715 Internal Remarks:22.04.2021 05:51:29 UTC HODAKA HASHIMOTO (81100553) Replacement parts: Replacement parts: As below 12Nc:453566095882 Name:ECG TESTER QTY:1 [Report on mistakes in parts order] No mistakes. Situation monitoring information: Not monitoring the situation TS Request: No parts to return. Additional CT information: Q1. How did the examination finish? Q1. No (not applicable) Q2. Were there any mistaken findings (misdiagnosis, mistaken treatment) due to the problem? None. Q3. Were there any artifacts? None. Q4. Was there any unintended unnecessary exposure? None. Repair sticker: I affixed a Repaired sticker. [Other information to share] 1. Was treatment/quarantine carried out for a patient who has or is suspected of having new coronavirus? Yes 2. Is it necessary to go into treatment rooms/beds where there are infected patients?(Including workplace and routes of movement) No Did you agree on how to complete the report with the person in charge of devices? YES Contacted the customer regarding the status the next day. 22-Apr-2021 81100553 FSE Handover Tool1.3.4.32612.DB13 20200715 ----------------------------------------------------------------------------- | 5 |  | ECG tester failure, replacement, order 12Nc:453566095882 Name:ECG TESTER QTY:1 | [12 empty] | 117402620 | New | 818184 |  | Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | ECG | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
776 |  |  | Sweden | Ludvika Lasarett | 40049 | 44312 |  | Saturday at 4.00 PM they had issues with initialize gantry. Doesn´t work this morning. Bug rep made. | Work Order #: WO-03381271 Knudsen Martin The system does not meet specification and is out of use Problem Description by Engineer : Gets gantry error when using system Diagnostic performed by Engineer : Logfiles show Heater error and Ref Converter error. Logged on to system and did Diagnostic_s, indicating RCOM or DMC error. Will order RCOM and DMC for replacement. Follow Up Required : Heater issue and Ref conv issue. Alle related to RCOM/DMC. VD shows these errors. Begin with replacing RCOM. Please schedule FSE and order parts. -MK | 2 |  | Gets gantry error when using system, schedule FSE and order parts, collect bugreport for analysis | [12 empty] | 117440693 | In Progress | 839559 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-Bugrep | [5 empty] | Engineering | [6 empty] | RCOM/DMC | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | TBD-Bugrep | Gantry-HW-TBD-Bugrep | Gantry-HW-TBD-Bugrep | TBD-Bugrep | TBD_Bugreport | Gantry-HW | Gantry-HW-TBD-Bugrep | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
777 |  |  | United States | Central Vermont Medical Center | 40024 | 44312 |  | Gating is not working | Work Order #: WO-03383304 Miller, John M The system partly meets specification and is returned to use with restrictions Diagnostic performed by Engineer : Collect information from customer. Problem Description by Engineer : Problem description by engineer : Customer states the Gating is not connecting. When was the first occurrence? Within... : 26-Apr-2021 How often is issue occurring? : Continuously Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Gating not connecting Patient Impact/Outcome: can not have Gating Scan. User Impact: can not use gating Current Software Version: Not provided Internal Comments : Error # and/or description of error : Gating not connecting Completed or Recommended Repair Action "Internal" : Customer states they have tried power cycling and checking all connections. Suggest checking the cabling from the EKG monitor to the gantry. Verify the connections at the Usb hub and test the Ibox. 459801413791 New EXTenders USB, released and available!!! Material Number 459801731761 - USB EXTENDER CABLE FROM IBOX TO CONSOLE Material Number 459800166652 - I-Box-2 Base board assy Material Number 459801787001 - Interconnect cable from GHOST to BELLOW Follow Up Required : Error # and/or description of error : Gating not connecting Completed or Recommended Repair Action "Internal" : Customer states they have tried power cycling and checking all connections. Suggest checking the cabling from the EKG monitor to the gantry. Verify the connections at the Usb hub and test the Ibox. 459801413791 New EXTenders USB, released and available!!! Material Number 459801731761 - USB EXTENDER CABLE FROM IBOX TO CONSOLE Material Number 459800166652 - I-Box-2 Base board assy ----------------------------------------------------------------------------- Work Order #: WO-03383421 Daigle Jim Problem Description by Engineer : Problem description by engineer : Customer states the Gating is not connecting. When was the first occurrence? Within... : 26-Apr-2021 How often is issue occurring? : Continuously Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Gating not connecting Patient Impact/Outcome: can not have Gating Scan. User Impact: can not use gating Current Software Version: Not provided Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 3 |  | Replaced iBox box cable and or30 m cable, Material Number 459801731761 - USB EXTENDER CABLE FROM IBOX TO CONSOLE Material Number 459800166652 - I-Box-2 Base board check which part is finally fix the issue? | [12 empty] | 117442880 | In Progress | 839886 |  | Investigation | Closed - Done |  | NotEngineering-CabelReplace | [5 empty] | Others | [6 empty] | Ibox and USB Cable | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
778 |  |  | United States | Lackey Memorial Hospital | 40015 | 44309 |  | The FSE is on site and the table was making a noise. Was not receiving an error message or anything. | Work Order #: WO-03376194 25-Apr-2021 Denison Chuck The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: X-Ray Supervisor Device Use: Philips C.T. scanner Expected and Unexpected Results: Adjustments to the table top assembly. User Impact: None. system up. Patient Impact: None System up. Current Software Version:4.1.10 Diagnostic performed by Engineer : Took Table top apart. Made repairs to the Cover support and made adjustments to correct the table performance. | 5 |  | Took Table top apart. Made repairs to the Cover support and made adjustments to correct the table performance. | [12 empty] | 117435989 | In Progress | 837830 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | NotEngineering-Assembly | [5 empty] | Others | [6 empty] | Couch Noise | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
779 |  |  | United States | McLaren Macomb | 40021 | 44309 |  | 1002CDN81 constancy phantom was dropped and broken | Work Order #: WO-03374555 Sheffield, Scott The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: biomed Device Use: not is use Expected and Unexpected Results: phantom was dropped and broken User Impact: can not do qa Patient Impact: none Current Software Version: Diagnostic performed by Engineer : Customer dropped phantom and broke it. Customer gave po # 4500017862 for the phantom. Follow Up Required : Customer dropped phantom and broke it. Customer gave po # 4500017862 for the phantom. 453567135965 - EXTENDED PHANTOM PERFORMANCE KIT without the low contrast pins, used in BigBore v4.2 or above, Ingenuity/ Ingenuity Core/Ingenuity Core128,Ingenuity Flex; for others, please order 453567135962; must order 1pcs 459801778031 for phantom replacement; order 453566200321 to replenish wat ----------------------------------------------------------------------------- Work Order #: WO-03375192 Mona Tony Problem Description by Engineer : Customer Function/Role: biomed Device Use: not is use Expected and Unexpected Results: phantom was dropped and broken User Impact: can not do qa Patient Impact: none Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 5 |  | Is there any injure on operator? Is Body section dropped? Any picture or details could share? | [12 empty] | 117434793 | In Progress | 834467 |  | Investigation | Closed - Done |  | CTUDT00632803 | [5 empty] | Engineering | [6 empty] | Phantom drop | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | Phantom | Couch-HW-Phantom | Couch-HW-Phantom | CTUDT00632803 | Phantom | Couch-HW | Couch-HW-Phantom | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | NotEngineering-CTUDT00632803 | y | y
780 |  |  | United States | The Regional Medical Center | 40005 | 44307 |  | Customer feels that the post-threshold delay is too long on PE exams | Enhancement request Customer feels that the post-threshold delay is too long on PE exams. They are bolus tracking at the level of the carina (mid-chest). The couch has to move to the top of the chest to begin the scan. The post-threshold delay is usually around 7 seconds. This does include the breathing instructions. They feel that this time is too long when trying to scan while there is a good contrast bolus in the pulmonary arteries. They commented that on other vendor_s scanners that they have worked on, the table moves much faster into scan position after performing the bolus tracking scan. They would like to see the couch move to the beginning of the scan faster. Even if the breath-hold instructions are removed, the post-threshold delay is still approximately 6 seconds. The expectation is that the post-threshold delay should be no more than 2-3 seconds. It was in clinical use when they encountered the concern. Procedure was continued. It happens every time a PE exam is performed. They feel they get poor contrast enhancement on some PE exams due to this long delay. No actions taken. There is no fix for the issue, the system works as designed. | CFM |  | Customer feels that the post-threshold delay is too long on PE exams, share to Silva and Lilian to review | [14 empty] | 833710 |  | Investigation | Closed - Done | [7 empty] | Application | [6 empty] | RFC_post-threshold delay | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
781 |  |  | United States | The Regional Medical Center | 40005 | 44307 |  | Customer would like the ability to change the result direction during a basic interventional procedure. | Enhancement request Customer would like the ability to change the result direction during a basic interventional procedure. When the patient is on the head head first during an interventional procedure when they use the arrow keys or the wheel to scroll through the images, as they press the up arrow key or roll the mouse wheel upwards, it scroll towards the feet. Arrow down or rolling the wheel down goes towards the head. This is confusing to the operator and physician during the procedure. They would like the ability to reverse the result direction so when using the mouse or wheel in an upward fashion makes the images scroll towards the head, and scrolling down will have the images scroll towards the feet. The Radiologist, Dr. Forte said any other scanner he has worked on has allowed for the ability to reverse the results. It was in clinical use when they encountered the concern. Procedure was continued. It happens every time the patient is positioned head first. Customer states they feel the fact that the images scroll in the opposite direction, they could mistake which way the needle needs to be manipulated (superior vs inferior). No actions taken. No fix for the issue, the system works as designed. | CFM |  | Customer would like the ability to change the result direction during a basic interventional procedure. share to Silva and Lilian to review | [14 empty] | 833690 |  | Investigation | Closed - Done | [7 empty] | Application | [6 empty] | RFC_interventional | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
782 |  |  | United States | Seton Medical Center Harker He | 40006 | 44308 |  | Scan Stopped During a Calcium Score | Work Order #: WO-03369736 Paine Matthew The system partly meets specification and is returned to use with restrictions Follow Up Required : Gantry Acq System can not Comply without Resends. Error S_MASTER_SEQUENCER_ACQ_NOT_OK. CIRS Error- SafetyCheckThread ACQ board reported error. Recommend Running ACQ Diagnostics. Check Health of RAID in CIRS. Run DataPath Diagnostics. Customer states issue only happens during Cardiac Studies. Diagnostic performed by Engineer : Customer states that they recently purchased the Cardiac Package. Customer states that during a Cardiac Calcium Score, they receive a Gantry ACQ can not Comply Message. Customer states a Bug Report was made for Incident and Authorized Remote Access through PRS. Logs indicate Issue with ACQ Cancelling the Scan. Problem Description by Engineer : Customer Function/Role: Technologist Device Use: Diagnostic Expected and Unexpected Results: Gantry ACQ can not Comply During Calcium Score User Impact: Unknown Patient Impact: Unknown Current Software Version:4.1.10 Internal Comments : Contacted Primary FSE Tom Beato. Recommend checking the Raid of CIRS, Run ACQ Diagnostics, Datapath Diagnostics. Try to Duplicate Issue if possible by performing Cardiac Scan_s. No Resends seen. ----------------------------------------------------------------------------- Work Order #: WO-03370392 23-Apr-2021 Beato Tom The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Technologist Device Use: Diagnostic Expected and Unexpected Results: Gantry ACQ can not Comply During Calcium Score User Impact: Unknown Patient Impact: Unknown Current Software Version:4.1.10 Diagnostic performed by Engineer : Checked the view logs and bug report. Verified one ACQ not ok error. Performed DMS diagnostics and data path diagnostics tests. All results passed. Will order loopback test kit to perform Acquisitor diagnostics. Follow Up Required : Perform ACQ diagnostics | 4 |  | Checked the view logs and bug report. Verified one ACQ not ok error. Performed DMS diagnostics and data path diagnostics tests. All results passed. Will order loopback test kit to perform Acquisitor diagnostics. Ask to share bugreport for analysis. To team review May cause rescan on surview!! | [12 empty] | 117430254 |  | 830714 |  | Closed - Done | Closed - Done |  | TBD_Bugreport | [5 empty] | Engineering | [6 empty] | Scan stop | [5 empty] | Refresh |  |  | System issue | System | TBD-Bugreport | TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | TBD_Bugreport | TBD-Bugreport | TBD-Bugreport-TBD_Bugreport | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
783 |  |  | United States | Unity Health Systems | 40037 | 44307 |  | v 4.1.10.2521 Tube heat on everything | Work Order #: WO-03358780 Hirchak, Michele The system partly meets specification and is returned to use with restrictions Problem description by Clinical Support Specialist : Customer Function/Role: Tech Device Use: In Use Expected and Unexpected Results: Tube heat on all scans. User Impact: None Patient Impact: None Current Software Version: 4.1.10 No rescan, no loss of data, no reinjection. Diagnostic performed by Engineer : Customer stated that on most scans they get tube heat messages. It tells them they have to wait up to 30 minutes to be able to scan. Techs have decreased technique and still get the message. I went over protocols with her, they all seem fine. Also, she said when they try to use the select Current Patient option, the screen goes blank and they have to log out and log back in in order to get it to work properly. I have sent this to an RSE to evaluate. This is not a protocol issue. MBH ----------------------------------------------------------------------------- Work Order #: WO-03359133 Mark Finnigan The system does not meet specification and is out of use Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: Issues with tube heat User Impact: Unknown. Patient Impact: Current Software Version: Diagnostic performed by Engineer : The case was started by Michelle Hirchak in clinical applications. The customer is seeing excessive tube heat warnings and also complained about a issue with the software . When the customer selects the "Use Current" for the patient selection, the instruction does not execute and the application goes back to the original step. Brought NSS Rich Hladun into the case. He downloaded bugreports and looked at some parameters. He would like the case dispatched and have a FSE onsite to work with. Dispatching the case. Follow Up Required : The case was started by Michelle Hirchak in clinical applications. The customer is seeing excessive tube heat warnings and also complained about a issue with the software . When the customer selects the "Use Current" for the patient selection, the instruction does not execute and the application goes back to the original step. Brought NSS Rich Hladun into the case. He downloaded bugreports and looked at some parameters. He would like the case dispatched and have a FSE onsite to work with. Dispatching the case. ----------------------------------------------------------------------------- Work Order #: WO-03360517 Taylor, Durrance Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: Issues with tube heat User Impact: Unknown. Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- Work Order #: WO-03367700 22-Apr-2021 Lockwood Thomas The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : System resistricted when awaiting tube heat to countdown. CT Unit is investigating ----------------------------------------------------------------------------- | 3 |  | Tube heat and software stuck issue, yellow scrren? needs bugreport for analysis. | [12 empty] | 117420104 |  | 826206 |  | Evaluation | Closed - No Further Investigation |  | TBD_Tube hot_bugreport | [5 empty] | Engineering | [6 empty] | Tube Overheat | [5 empty] | Refresh |  |  | System issue | System | Gantry-SW | TubHot | Gantry-SW-TubHot | Gantry-SW-TubHot | TBD_Tube hot_bugreport | TubHot | Gantry-SW | Gantry-SW-TubHot | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
784 |  |  | United States | Advent Health Wesley Chapel | 300106 | 44308 |  | 0706CDY73 Attempting to print cd on scanner but saying its locked due to scanning in progress but not scanning... | Work Order #: WO-03367666 Sheffield, Scott The system partly meets specification and is returned to use with restrictions Internal Comments : Error # and/or description of error : CD is locked because scan is in progress Completed or Recommended Repair Action "Internal" : Check CDR configuration. Diagnostic performed by Engineer : I had customer reboot host computer. Still got same problem after reboot. Applications sent me an alternate method to burn a cd and that had the same result. I called FSE he is going to check the cd configuration. Problem Description by Engineer : Problem description by engineer : When trying to burn Cd getting message saying: CD is locked because scan is in progress Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No further information is available as per the customer Expected and Unexpected results: When trying to burn Cd getting message saying; CD is locked because scan is in progress Patient Impact/Outcome: No further information is available as per the customer User Impact: delay in burning a cd Current Software Version: 4.1.7.25030 Follow Up Required : Error # and/or description of error : CD is locked because scan is in progress Completed or Recommended Repair Action "Internal" : Check CDR configuration. ----------------------------------------------------------------------------- Work Order #: WO-03368698 22-Apr-2021 Sher Mark The system fully meets specification and returned to use Problem Description by Engineer : Problem description by engineer : When trying to burn Cd getting message saying: CD is locked because scan is in progress Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No further information is available as per the customer Expected and Unexpected results: When trying to burn Cd getting message saying; CD is locked because scan is in progress Patient Impact/Outcome: No further information is available as per the customer User Impact: delay in burning a cd Current Software Version: 4.1.7.25030 Diagnostic performed by Engineer : changed cd burner drive letter to G from e and tested system.. issue resolved ----------------------------------------------------------------------------- | 3 |  | changed cd burner drive letter to G from e and tested system.. issue resolved | [12 empty] | 117428491 |  | 830419 |  | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | CD burner | [5 empty] | Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
785 |  |  | United States | Advent Health Wesley Chapel | 300106 | 44307 |  | Q/O: images numbers not the image number expected | Work Order #: WO-03359751 22-Apr-2021 Sher Mark The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : king with applications customer corrected converted protocol form 4.1.7 upgrade to windows 10 .. slice thickness on step and shoot were wrong, not converted | 3 |  | Step and Shot, need align with Caroline to know more detail 10mm -> 15mm bring biopy issue and pedal scan no image. | [12 empty] | 117420814 |  | 830143 |  | Investigation | Investigation |  | TBD-CTDUT | [5 empty] | Engineering | [6 empty] | Step and Shot | [5 empty] | Refresh |  |  | Upgrade issue | System | NotEngineering | ShouldbeClinical | NotEngineering-ShouldbeClinical | NotEngineering-ShouldbeClinical | NotEngineering-ShouldbeClinical | NotEngineering-ShouldbeClinical | NotEngineering | NotEngineering-ShouldbeClinical | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
786 |  |  | United States |  | 300106 | 44305 |  | Q/O Need bore mylar replaced and phantom | Work Order #: WO-03342940 20-Apr-2021 Sher Mark The system fully meets specification and returned to use Parts Consumed : 1.00000,459800048742,RING-LEXAN GN5 1.00000,455012004391,CALIBRATION PHANTOM 1.00000,453567122551,SURVIEW GAUGE 1.00000,453567122541,LASER ALIGNMENT GAUGE Diagnostic performed by Engineer : received parts for Pm, checked on system Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: ----------------------------------------------------------------------------- | 3 |  | 1.00000,459800048742,RING-LEXAN GN5 1.00000,455012004391,CALIBRATION PHANTOM 1.00000,453567122551,SURVIEW GAUGE 1.00000,453567122541,LASER ALIGNMENT GAUGE | [12 empty] | 117405409 |  | 819987 |  |  | Closed - Non-Complaint | [14 empty] | Mylar/Phantom | [5 empty] | Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
787 |  |  | United States | Upson Regional Medical Center | 333009 | 44301 |  | System jumping to another result during the planning step on exam card utilizing bolus tracking. This results in breakin |  | CFM |  | Mouse Jump | [14 empty] | 816865 |  | Investigation | Investigation |  | CTUDT00632986 | [5 empty] | Engineering | [6 empty] | Mouse Jump | [5 empty] | Refresh |  |  | Upgrade issue | System | Console-SW | ResultSelectionJump | Console-SW-ResultSelectionJump | Console-SW-ResultSelectionJump | CTUDT00632986 | ResultSelectionJump | Console-SW | Console-SW-ResultSelectionJump | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
788 |  |  | United States | Upson Regional Medical Center | 310115 | 44300 |  | System jumping to another result during the planning step on exam card utilizing bolus tracking. |  | CFM |  | Mouse Jump | [14 empty] | 816831 |  | Investigation | Investigation |  | CTUDT00632986 | [5 empty] | Engineering | [6 empty] | Mouse Jump | [5 empty] | Refresh |  |  | Upgrade issue | System | Console-SW | ResultSelectionJump | Console-SW-ResultSelectionJump | Console-SW-ResultSelectionJump | CTUDT00632986 | ResultSelectionJump | Console-SW | Console-SW-ResultSelectionJump | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
789 |  |  | United States | Upson Regional Medical Center | 333009 | 44301 |  | A system lag between mouse clicks during the scan plan setup with an exam card that has more than about 10 results. This |  | CFM |  | Performace | [14 empty] | 816806 |  | Investigation | Investigation |  | TBD-Performance | [5 empty] | Engineering | [6 empty] | Performace | [5 empty] | Refresh |  |  | Upgrade issue | System | Console-SW | TBD-Performance | Console-SW-TBD-Performance | Console-SW-TBD-Performance | TBD-Performance | TBD-Performance | Console-SW | Console-SW-TBD-Performance | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
790 |  |  | United States | Upson Regional Medical Center | 310115 | 44301 |  | A system lag between mouse clicks during the scan plan setup with an exam card that has more than about 10 results. This |  | CFM |  | Performace | [14 empty] | 816778 |  | Investigation | Investigation |  | CTUDT00632986 | [5 empty] | Engineering | [6 empty] | Performace | [5 empty] | Refresh |  |  | Upgrade issue | System | Console-SW | SerielsJump | Console-SW-SerielsJump | Console-SW-SerielsJump | Duplicate | ResultSelectionJump | Console-SW | Console-SW-SerielsJump | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
791 |  |  | United States |  | 310034 | 44308 |  | NEMO | Work Order #: WO-03369989 ----------------------------------------------------------------------------- | 3 |  | Non-Refresh | [12 empty] | 117430454 |  | 835227 |  | #N/A | Evaluation | [14 empty] | Non-Refresh | [5 empty] | Non-Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | y
792 |  |  | United States |  | 310034 | 44307 |  | nemo | Work Order #: WO-03361780 ----------------------------------------------------------------------------- | 3 |  | Non-Refresh | [12 empty] | 117422240 |  | 831122 |  | #N/A | Evaluation | [14 empty] | Non-Refresh | [5 empty] | Non-Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | y
793 |  |  | United States |  | 310034 | 44307 |  | nemo | Work Order #: WO-03361745 Parts Consumed : 1.00000,459800484511,CONTROL PANEL ASSEMBLY-FRONT RH ----------------------------------------------------------------------------- | 3 |  | Non-Refresh | [12 empty] | 117422216 |  | 831104 |  |  | Closed - Non-Complaint | [14 empty] | Non-Refresh | [5 empty] | Non-Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | y
794 |  |  | United States |  | 310167 | 44307 |  | Failing to send images | Work Order #: WO-03361513 Morrow, Dan The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Problem description by engineer : Failing to send Images. Customer states that IT checked their end and issue with scanner. When was the first occurrence? Within... : 21-Apr-2021 How often is issue occurring? : First occurrence Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Unable to send images Patient Impact/Outcome: none User Impact: Unable to send images Current Software Version: 4.1.7.25028 Diagnostic performed by Engineer : No PRS connection. Customer checked lan cable and is OK. Dispatching FSE. Follow Up Required : What are the next Troubleshooting steps : Check LAN connection to Host. Verify port has connection. Check NIC card. Error # and/or description of error : No PRS. Completed or Recommended Repair Action "Internal" : No PRS connection. Customer checked lan cable and is OK. Dispatching FSE. Internal Comments : Error # and/or description of error : No PRS. Completed or Recommended Repair Action "Internal" : No PRS connection. Customer checked lan cable and is OK. Dispatching FSE. ----------------------------------------------------------------------------- | 3 |  | Non-Refresh | [12 empty] | 117422081 |  | 826822 |  | #N/A | Investigation | [14 empty] | Non-Refresh | [5 empty] | Non-Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | y
795 |  |  | China | ZJ YUEQING PEOPLE HOSPITAL | 40054 | 44321 |  | Application/Ingenuity CT | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer : Application FSE Notes : This PO is a Reference Number only Next Step Action : 邮件ASP鸿乐嘉跟进， 短信协调员 曾彩花(MATC/US 东区/南区) 17875512591 放射科 林拉子 13362770096 申请首次培训，装机单53515511派工ASP鸿乐嘉 | 5 |  | NAC_APP training | [12 empty] | 117494304 |  | 872990 |  | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Application | [6 empty] | NAC_APP training | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
796 |  |  | China | XJ Urumqi No1 People Hosp XJ Urumqi Children Hosp | 40058 | 44315 |  | 02N/Ingenuity CT 728326/stand by | BU Q&R Feedback: No Reason: Customer Relations Problem Reported by Customer: stand by SAP Transfer:Dispatched for On site Support External Remarks: 01.05.2021 04:34:19 UTC Jianjun Sang (86038298) Service details: equipment acceptance. Service outcome: Now the device returned to service fully. Internal Remarks:01.05.2021 04:34:21 UTC Jianjun Sang (86038298) $Q7: (Confirm with reported problem: equipment acceptance.) $Q8: (equipment acceptance, check the equipment.) $Q9: (equipment acceptance.) $Q10: (Test and inspection data refer to: PA. The device returned to full functionality. $Q11: Calibrated Tools name (N/A) S/N (N/A) Next Calibration date (N/A) | 5 |  | NAC_FSE support equipment acceptance test | [12 empty] | 117464487 |  | 855913 |  | #N/A | Closed - non-complaint | [7 empty] | Service | [6 empty] | NAC_FSE support equipment acceptance test | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
797 |  |  | United States | Unity Health Systems | 40037 | 44307 |  | Excessive Tube Heat Warnings and issue with software |  | CFM |  | Excessive Tube Heat Warnings and issue with software | [14 empty] | 843526 |  | Investigation | Closed - Done |  | TBD_Tube hot_bugreport | [5 empty] | Engineering | [6 empty] | Tube Overheat | [5 empty] | Refresh |  |  | System issue | System | Gantry-SW | TubHot | Gantry-SW-TubHot | Gantry-SW-TubHot | TBD_Tube hot_bugreport | TubHot | Gantry-SW | Gantry-SW-TubHot | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
798 |  |  | Turkey | HITACHI HEALTHCARE TURKEY | 337000 | 44313 |  | Compensator analysis. | Work Order #: WO-03391860 Soner Saglam The system fully meets specification and returned to use Problem Description by Engineer : Compensator crack check. Resolution : No cracks were found. | 3 |  | Compensator analysis | [12 empty] | 117450887 |  | 848977 |  | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Service | [6 empty] | Compensator analysis | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
799 |  |  | Chile | CMD Clinica Integrales Rancagu | 40008 | 44313 |  | CL1675-THE GANTRY IS RESTARTING | Work Order #: WO-03394637 Valenzuela Velázquez, Eduardo Ariel The system fully meets specification and returned to use. Problem Description by Engineer : Problem description by engineer : Gantry shuts off Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected Results: No gantry errors. Patient Impact/Outcome: Cancelled examinations. User Impact: Canceled examinations. Current Software Version: 4.1.10 Follow Up Required: Error Number and/or Description: S_MC_DEVICE_HOR_ENCODERS_NOT_SYNCED Completed or Recommended Repair Action "Internal" : Absolute encoder coupler replacement. Internal Comments : Error # and/or description of error : S_MC_DEVICE_HOR_ENCODERS_NOT_SYNCED Completed or Recommended Repair Action "Internal" : Absolute encoder coupler replacement. Diagnostic Performed by Engineer: An error is displayed in the position sensors of the horizontal couch motion. A broken coupling part is found ion the couch. With client_s help, the part is repaired. Equipment is left working but the part should be replaced. | 4 |  | Manuf WI update to control the process, need wait ECO # for WI. | [12 empty] | 117453038 |  | 850654 |  | Closed - Done | Closed - Done |  | CTUDT00632179 | [5 empty] | Engineering | [6 empty] | Coupling | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | Motor Coupling | Couch-HW-Motor Coupling | Couch-HW-Motor Coupling | CTUDT00632179 | Motor Coupling | Couch-HW | Couch-HW-Motor Coupling | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
800 |  |  | Poland | Agencja Rezerw Materiałowych | 341010 | 44320 |  | Physicist training | Work Order #: WO-03433693 Robert Kamola Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A | 4 |  | APP training | [12 empty] | 117488965 |  | 868913 |  | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Application | [6 empty] | APP training | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
801 |  |  | Turkey | T.C Saglik Bakanligi Bagcilar -> Bağcılar State Hospital | 40010 | 44319 |  | CTbox key malfunction. | Work Order #: WO-03427934 Soner Saglam The system fully meets specification and returned to use Problem Description by Engineer : Unable to use the CTbox down key. Resolution : Ctbox replacement. Follow Up Required : Ctbox replacement. | 3 |  | William do investigation | [12 empty] | 117484484 |  | 862117 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632185 | [5 empty] | Engineering | [6 empty] | CT Box | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | CT Box | Gantry-HW-CT Box | Gantry-HW-CT Box | CTUDT00632185 | CT Box | Gantry-HW | Gantry-HW-CT Box | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
802 |  |  | Turkey | T.C Saglik Bakanligi Bagcilar -> Bağcılar State Hospital | 40010 | 44316 |  | Loss of communication errors. | Work Order #: WO-03418930 Soner Saglam The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : e-stop is activated. logger.mdb was analyzed. S_COUCHMGR_HOR_MOTION_STOP_ERROR, S_MANMOT_HORIZONTAL_STOP_ERROR errors were found. Horizantal axis check and calibration are recommended. NOTE: Full gantry recover was advised to the customer. The customer reported that they restarted the gantry without any success. In the examination, it was found that the device gave an error about horizontal movement and opened an estop. The coupling was checked and horizontal encoder calibration was done. It was checked whether there was any congestion in the horizontal motion. 10 trial patients were enrolled. No problems were observed. The system has been delivered in operation. | 3 |  | Check and collect bugreport for anlysis. To ChangXian, if known issue? | [12 empty] | 117477048 |  | 857371 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632187 | [5 empty] | Engineering | [6 empty] | Couch SW issue | [5 empty] | Refresh |  |  | System issue | System | Couch-SW | Estop | Couch-SW-Estop | Couch-SW-Estop | CTUDT00632187 | Estop | Couch-SW | Couch-SW-Estop | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
803 |  |  | Turkey | T.C Saglik Bakanligi Bagcilar -> Bağcılar State Hospital | 40010 | 44313 |  | Compensator analysis. | Work Order #: WO-03391863 Soner Saglam The system fully meets specification and returned to use Problem Description by Engineer : Compensator crack check. Resolution : No cracks were found. ----------------------------------------------------------------------------- | 3 |  | Compensator analysis | [12 empty] | 117450896 |  | 848980 |  | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Service | [6 empty] | Compensator analysis | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
804 |  |  | United States | St Barnabas Hospital | 40028 | 44317 |  | 0711CFXXX Transfer queue failing | Work Order #: WO-03424316 Mcginnis, Mike The system fully meets specification and returned to use Problem Description by Engineer : Problem description by engineer : Transfer queue failing When was the first occurrence? Within... : 01-May-2021 How often is issue occurring? : First occurrence Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Transfer queue failing Patient Impact/Outcome: none User Impact: had to place service call Current Software Version: 4.1.10 Diagnostic performed by Engineer : Iris said they had 5 cases stuck in the transfer Q. I had her delete the transfer of all images and we then went back in and transferred all of the one at a time. I discovered speaking to her that on one of those patients they had an accession number incorrect and had changed it. Everything passed when sending the corrected patient. Internal Comments : Error # and/or description of error : Transfer queue failing Completed or Recommended Repair Action "Internal" : Removed failed Q and sent all exams to the pacs with no issues. Resolution : Removed failed Q and sent all exams to the pacs with no issues. ----------------------------------------------------------------------------- | 3 |  | Transfer queue failing | [12 empty] | 117481437 |  | 860994 |  | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Transfer queue failing | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
805 |  |  | United States | St Barnabas Hospital | 40028 | 44314 |  | Table horizontal not working. | Work Order #: WO-03404474 Mark Finnigan The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: Customer experienced an issue with table horizontal motion. User Impact: Unknown Patient Impact: Current Software Version: Diagnostic performed by Engineer : Philips CT applications support Susan Shick was doing onsite applications support when she encountered an issue where table horizontal motion was not working. She called for support. The emote connection is established but the refresh passwords were not allowing login. When I called the site back Susan had gotten the table working. I recommended that she take the system down to the breaker when she had a chance. I also called the FSE Mohamed Abbas and let him know about the PRS connectivity issue so that he can get that resolved. Internal Comments : No additional calls from Susan Sheck from the site. Left her a VM to call back if she needs additional support | 2 |  | Check and collect bugreport for anlysis. | [12 empty] | 117462864 |  | 850492 |  | Closed - Done | Closed - Done |  | TBD_Bugreport | [5 empty] | Engineering | [6 empty] | Horizontal | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | TBD-bugrep | Couch-HW-TBD-bugrep | Couch-HW-TBD-bugrep | TBD_CouchEstop | TBD_Bugreport | Couch-HW | Couch-HW-TBD-bugrep | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
806 |  |  | China |  | 40050 | 44316 |  | 02G/Ingenuity CT 728326/add printer dicom node | BU Q&R Feedback: No SAP Customer Symptom Code: CENV Customer Symptom Code: Customer Environmental Concern Reason: Corrective Maintenance Engineer Failure Code: Configuration Error Repair Activity Code: Configuration Change Problem Reported by Customer: 02G/Ingenuity CT 728326 Safety Questions:Case: 0117475950 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Gaoge Last updated by: Wang Gaoge 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action: Q6 = add DICOM printer SAP Transfer:Case Transferred to SAP T2 Activities:Problem Description: need to add printer node Troubleshooting Action: Instruct customer to enter admin account and add printer node Repair Action: Instruct customer to enter admin account and add printer node Test/Inspection: N/A Test Equipment: N/A ----------------------------------------------------------------------------- | Zhai, Lisa |  | NAC | [12 empty] | 117475950 |  | 859811 |  |  | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
807 |  |  | United States |  | 40021 | 44315 |  | LS - 065137 | Work Order #: WO-03412685 30-Apr-2021 Mona Tony The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : Replaced phantom Parts Consumed : 1.00000,459801415171,HARMONIZED PHANTOM PHYSICS SECTION ----------------------------------------------------------------------------- | E, Sravanthi |  | NAC | [12 empty] | 117470844 |  | 857952 |  |  | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
808 |  |  | United States | Southwest Mississippi Regional | 40055 | 44314 |  | 1203CDY72 System stops scan during surview | Work Order #: WO-03403330 Weaver David The system partly meets specification and is returned to use with restrictions Follow Up Required : S_GEN_FAULT_CATHODE_TUBE_ARC, S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT Clean high voltage connectors. If this happens with low (or very) high mA and the arcing is in the beginning of scan (Look in szInfo for a short exposure time like GXray EXP=18 (=18 milliseconds)) it could be related to the filament calibration. You might want to run filament calibration followed by long tube reconditioning in such a case. If it happens not in the beginning of scan a long tube reconditioning followed by an hour rest (to degas the tube) might be beneficial. Also look for soft arcs if these are on both anode and cathode side. If it happens all the time run a no load test both with tube connected and tube disconnected. Start with low kV and gradually increase. If there is arcing with tube connected and not when tube is disconnected then the problem is most likely in the tube. Check for burn marks on the HT connectors. Also you may want to listen to arcing sounds coming from one of the power modules (only with no load test!!!). Problem Description by Engineer : DESCRIPTION OF PROBLEM (EXPECTED/UNEXPECTED RESULTS) System stops scan during surview CALLER FUNCTION/ROLE Technician/Sonographer/Therapist DEVICE INTENDED USE Diagnostic WHEN WAS THE FIRST OCCURRENCE? WITHIN . . . <1 Week HOW OFTEN IS ISSUE OCCURRING? Intermittent WHERE IN THE END-USER WORKFLOW IS THE ISSUE OCCURRING? During acquisition/scan WHAT RECENT EVENTS HAVE OCCURRED? During normal use NEGATIVE OPERATOR IMPACT Scan must be restarted NEGATIVE PATIENT IMPACT/PROCEDURE OUTCOME Delay in scanninf HAS CUSTOMER CREATED A SPECIFIC SYSTEM EVENT REPORT OF ISSUE? Yes 4.2 Diagnostic performed by Engineer : System stops scan during surview ----------------------------------------------------------------------------- Work Order #: WO-03403375 30-Apr-2021 Donald Vic The system fully meets specification and returned to use Resolution : Following the recommendations of RSE I ran the long tube conditioning twice and no arcs were present during the long tube conditioning. Diagnostic performed by Engineer : Following the recommendations of RSE I ran the long tube conditioning twice and no arcs were present during the long tube conditioning. Problem Description by Engineer : DESCRIPTION OF PROBLEM (EXPECTED/UNEXPECTED RESULTS) System stops scan during surview CALLER FUNCTION/ROLE Technician/Sonographer/Therapist DEVICE INTENDED USE Diagnostic WHEN WAS THE FIRST OCCURRENCE? WITHIN . . . <1 Week HOW OFTEN IS ISSUE OCCURRING? Intermittent WHERE IN THE END-USER WORKFLOW IS THE ISSUE OCCURRING? During acquisition/scan WHAT RECENT EVENTS HAVE OCCURRED? During normal use NEGATIVE OPERATOR IMPACT Scan must be restarted NEGATIVE PATIENT IMPACT/PROCEDURE OUTCOME Delay in scanninf HAS CUSTOMER CREATED A SPECIFIC SYSTEM EVENT REPORT OF ISSUE? Yes 4.2 ----------------------------------------------------------------------------- | 4 |  | Peter check if have Hazad ? | [12 empty] | 117459746 |  | 849718 |  | Closed - Done | Closed - Done |  | CTUDT00630085 | [5 empty] | Engineering | [6 empty] | Tube Arc | [5 empty] | Refresh |  |  | System issue | System | IC-HW | Tub Arc | IC-HW-Tub Arc | IC-HW-Tub Arc | CTUDT00630085 | Tub Arc | IC-HW | IC-HW-Tub Arc | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
809 |  |  | United States | St Barnabas Hospital | 40028 | 44313 |  | Machine beeping | Work Order #: WO-03391382 28-Apr-2021 Joseph Casale The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : beep sound from ac unit not system rear gantry controls cables in wrong possition corrected | 4 |  | beep sound from ac unit not system rear gantry controls cables in wrong possition corrected | [12 empty] | 117450493 |  | 848652 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | NotEngineering | [5 empty] | Operation | [6 empty] | Gantry Noise | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
810 |  |  | China |  | 40023 | 44314 |  | 02N/Ingenuity CT 728326/change password | BU Q&R Feedback: No SAP Customer Symptom Code: HANG Customer Symptom Code: Hang/Crash Reason: Corrective Maintenance Engineer Failure Code: Configuration Error Repair Activity Code: Configuration Change Problem Reported by Customer: 02N/Ingenuity CT 728326 T2 Activities:Problem Description: Password expired Troubleshooting Action: Help customers reset and modify passwords Repair Action: Help customers reset and modify passwords Test/Inspection: customer confirm can scan patient and image ok.functional Check PASSED. The device returned to full functionality. Test Equipment: N/A Safety Questions:Case: 0117455662 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Bo Last updated by: Wang Bo 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | Zhai, Lisa |  | NAC | [12 empty] | 117455662 |  | 851693 |  |  | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
811 |  |  | United States | Odessa Regional Medical Center | 40029 | 44314 |  | 1102CDN83 Power outage last night. UPS failure. | Work Order #: WO-03400353 Morrow, Dan The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Problem description by engineer : Power outage and UPS did not remain on for console processing. Console turned off with loss of power. When was the first occurrence? Within... : 28-Apr-2021 How often is issue occurring? : First occurrence Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: Console remain on Patient Impact/Outcome: none User Impact: Unable to process images Current Software Version: Not provided Follow Up Required : What are the next Troubleshooting steps : Test UPS. Possible control or battery failure. 459800306382 UPS (Emerson), 230V, 50/60Hz, 3 KVA 459800306392 EXTERNAL BATT PACK, 3KVA UPS, 72VDC Error # and/or description of error : No PRS Completed or Recommended Repair Action "Internal" : Emerson Staco installed. Dispatching FSE to test UPS. Diagnostic performed by Engineer : Emerson Staco installed. Dispatching FSE to test UPS. Internal Comments : Error # and/or description of error : No PRS Completed or Recommended Repair Action "Internal" : Emerson Staco installed. Dispatching FSE to test UPS. ----------------------------------------------------------------------------- | 3 |  | 459800306382 UPS (Emerson), 230V, 50/60Hz, 3 KVA 459800306392 EXTERNAL BATT PACK, 3KVA UPS, 72VDC | [12 empty] | 117459016 |  | 849131 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-NoEnoughInfo | [5 empty] | Engineering | [6 empty] | UPS | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | TBD-NoEnoughInfo | Gantry-HW-TBD-NoEnoughInfo | Gantry-HW-TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo | Gantry-HW | Gantry-HW-TBD-NoEnoughInfo | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
812 |  |  | China | Guangzhou No8 Peoples Hospital | 345032 | 44312 |  | 02N/Ingenuity Core128/Elite China | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 02N/Ingenuity Core128/Elite China SAP Transfer:Dispatched for On site Support Safety Questions:Case: 0117439674 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Guoping Last updated by: Liu Guoping 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Case Transferred to SAP External Remarks: 29.04.2021 05:27:23 UTC Zhenhua Liu (86034125) Failure description: machine noise Service details: check machine movement, wipe slip ring Service result: After service, the machine is back to normal. Other important issues: None Internal Remarks : 29.04.2021 05:27:29 UTC Zhenhua Liu (86034125) "RDI Accuracy:Yes" $Q1: () $Q2: () $Q3: () $Q4: () $Q7: (Confirmation & Description of problem) Confirmed reported problem, gantry make some noise when rotor run in highest speed, but could not repeat the abnormality. $Q8: (Troubleshooting & Results) check the rotor and case, main belt, but could not find anything strange, but find out the slip ring is quite dirty. $Q9: (Repair steps, clean the slip ring.) $Q10: (Test and inspection data refer to PA . Device returned to full functionality ) $Q11: Calibrated Tools name ( refer to PA ) S/N ( ) Next Calibration date ( refer to PA ) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date: N/A ^Current Tube usage: N/A ----------------------------------------------------------------------------- | 3 |  | Check if the brush is installed correctly | [12 empty] | 117439674 |  | 841678 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | NotEngineering-OPS | [5 empty] | Operation | [6 empty] | Brush/Slipring | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
813 |  |  | United States | McLaren Macomb | 40021 | 44319 |  | 0899CDY71 IBox Fails Diags | Work Order #: WO-03428412 Carpenter, Kevin The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Customer Function/Role: Biomed Jim Device Use: Diag Expected and Unexpected Results: iBox failed diags, even after gantry power cycle User Impact: None Patient Impact: can not do gated Pt studies Current Software Version: IDK Diagnostic performed by Engineer : Confirmed that trained biomed had power cycled gantry but that iBox still failed diags. CT RSE colab_ed with pFSE Tony Mona ~ agree to ship iBox Part to his home FunLoc Follow Up Required : Biomed reports iBox failed diags, even after gantry power cycle. pFSE Tony Mona req_s iBox PCB part be shipped to his home / funloc ----------------------------------------------------------------------------- | 3 |  | 459800166651,I-BOX-2 BASE BOARD ASSY | [12 empty] | 117484842 |  | 862222 |  | Investigation | Closed - Done |  | TBD-NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Ibox | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | TBD-NoEnoughInfo | Gantry-HW-TBD-NoEnoughInfo | Gantry-HW-TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo | Gantry-HW | Gantry-HW-TBD-NoEnoughInfo | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
814 |  |  | United States | Lackey Memorial Hospital | 40015 | 44319 |  | Not sending to PACS | Work Order #: WO-03429838 04-May-2021 Denison Chuck The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: X-Ray Supervisor Device Use: Philips C.T. scanner. Expected and Unexpected Results: None was a follow up. User Impact: None C.T. Up. Patient Impact: None C.T. Up. Current Software Version:4.1.10 Diagnostic performed by Engineer : Done follow up on table adjustments. and checked the performance of the FCO that was performed on C.T. scanner. ----------------------------------------------------------------------------- | 5 |  | FCO740? | [12 empty] | 117485900 |  | 866397 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | HF3 | [5 empty] | Engineering | [6 empty] | Couch adjustment | [5 empty] | Refresh |  |  | System issue | System | Console-SW | HF3 | Console-SW-HF3 | Console-SW-HF3 | HF3 | HF3 | Console-SW | Console-SW-HF3 | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
815 |  |  | United States |  | 40015 | 44314 |  | needs help with a protocol | Work Order #: WO-03400997 Hirchak, Michele The system fully meets specification and returned to use Diagnostic performed by Engineer : Michea had questions on CTA Runoff. Threshold, pre and post injection delay, flow rate, saline flush, feet to knee scan and IV placement were discussed and explored on their Ingenuity CT system. Michea will discuss what other parameters could be changed with the radiologist to produce diagnostic imaging when performing CTA Runoffs on larger patients for approval. Michea had no further questions on this issue. No rescan, no loss of data, no reinjection. Problem description by Clinical Support Specialist : Customer Function/Role: Technologist Device Use: In use Expected and Unexpected Results: Runoff Questions User Impact: none Patient Impact: none Current Software Version: 4.1 ----------------------------------------------------------------------------- | Dantha, Rajashekar |  | NAC | [12 empty] | 117459462 |  | 849489 |  |  | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
816 |  |  | Japan | Kinan Hospital | 340000 | 44312 |  | 02DS:TM:(via FSE): The patient couch does not become free. >> FSE response. | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02DS:TM:(via FSE): The patient couch does not become free. >> FSE response. Safety Questions:Case: 0117441020 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: TAKEYUKI MIYAKAWA Last updated by: TAKEYUKI MIYAKAWA 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Support SAP Transfer:OBJ: Case_Activity__c - unable to obtain exclusive access to this record T2 Activities:Problem Description: Problem description by engineer : Equipment affected: CT Failure location: 02Couch When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user_s work flow did the problem occur?: (During Acquisition/Scan) What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem?: NA Engineer_s Comment: _============================== Call description: (via FSE): The patient couch does not become free. _============================== Troubleshooting Action: What steps are taken during the system diagnosis: Result of response: Problem not solved. Failure location: 02Couch What steps did you perform to diagnose the problem? _================================== Response FSE will respond. _================================== Error # and/or description of error: NA Repair Action: Completed or Recommended Repair Action "External": The service engineer for the site visited and worked on the equipment. OneEMS internal remarks : Problem description by engineer : Equipment affected: CT Failure location: 02Couch When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user_s work flow did the problem occur?: (During Acquisition/Scan) What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem?: NA Engineer_s Comment: _============================== Call description: (via FSE): The patient couch does not become free. _============================== What steps are taken during the system diagnosis: Result of response: Problem not solved. Failure location: 02Couch What steps did you perform to diagnose the problem? _================================== Response FSE will respond. _================================== Error # and/or description of error: NA Completed or Recommended Repair Action "External" : The service engineer for the site visited and worked on the equipment. Completed or Recommended Repair Action "Internal": --Troubleshooting Action-- Result of response: Problem not solved. Failure location: 02Couch What steps did you perform to diagnose the problem? _================================== Response FSE will respond. _================================== --Repair Action-- Recommended 12Nc: PR1 453566494823:CLUTCH - HORIZONTAL DRIVE QTY:1 SLOC: Osaka (U615) --Subject-- 02DS:TM:via FSE: The patient couch does not become free. >> FSE response. --Problem Description-- Equipment affected: CT Failure location: 02Couch When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user_s work flow did the problem occur?: (During Acquisition/Scan) What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem?: NA Engineer_s Comment: _============================== Call description: (via FSE): The patient couch does not become free. _============================== SAP Message: Components Added Succesfully to SWO!- SAP Message: LOGI Task added Successfully!- External Remarks:29.04.2021 11:39:09 UTC TETSUO OHYAMA (81100250) [Request details] The bed does not become free. Problem confirmation: I confirmed it. Work details: There was a problem with the clutch in the horizontal direction of the bed. I replaced the clutch part. There is no problem with the operation of the equipment. Other problems/Measures taken: None. Equipment status The equipment can be used. Work status Work finished. 26-Apr-2021 81100250 FSE Handover Tool1.3.4.32612.DB13 20201001 Internal Remarks : 29.04.2021 11:39:11 UTC TETSUO OHYAMA (81100250) Replacement parts: Replacement parts: As below 12Nc:453566494823 Description:CLUTCH - HORIZONTAL DRIVE QTY:1 [Report on mistakes in parts order] No mistakes. Situation monitoring information: Not monitoring the situation TS Request: No parts to return. Additional CT information: Q1. How did the examination finish? Q1. No (not applicable) Q2. Were there any mistaken findings (misdiagnosis, mistaken treatment) due to the problem? None. Q3. Were there any artifacts? None. Q4. Was there any unintended unnecessary exposure? None. Repair sticker: I affixed a Repaired sticker. [Other information to share] 1. Was treatment/quarantine carried out for a patient who has or is suspected of having new coronavirus? Yes 2. Is it necessary to go into treatment rooms/beds where there are infected patients?(Including workplace and routes of movement) No Did you agree on how to complete the report with the person in charge of devices? YES On-site the next day 26-Apr-2021 81100250 FSE Handover Tool1.3.4.32612.DB13 20201001 ----------------------------------------------------------------------------- | 3 |  | 453566494823:CLUTCH - HORIZONTAL DRIVE | [12 empty] | 117441020 |  | 842585 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD_Bugreport | [5 empty] | Engineering | [6 empty] | CLUTCH | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | TBD-bugrep | Couch-HW-TBD-bugrep | Couch-HW-TBD-bugrep | TBD_CouchEstop | TBD_Bugreport | Couch-HW | Couch-HW-TBD-bugrep | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
817 |  |  | China | JS Yangzhou No.1 People Hosp. Xiqu Hosp. | 345022 | 44317 |  | 02N/Ingenuity Core128/Elite China/cannot scan | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: Cannot scan Safety Questions:Case: 0117481105 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingli Last updated by: Lu Mingli 服务内容：删除无用文件，腾空D盘空间. 服务结果： 维修后，设备恢复正常工作. | 3 |  | PACS transfer failed and bring D disk full, after disk clean and implement FCO740, system works well | [12 empty] | 117481105 |  | 861292 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | HF3 | [5 empty] | Engineering | [6 empty] | Disk full | [5 empty] | Refresh |  |  | System issue | System | Console-SW | HF3 | Console-SW-HF3 | Console-SW-HF3 | HF3 | HF3 | Console-SW | Console-SW-HF3 | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
818 |  |  | China | JS No.2 Chinese Medicine Hosp. | 345027 | 44317 |  | 02N/Ingenuity Core128/Elite China | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 02N/Ingenuity Core128/Elite China SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Safety Questions:Case: 0117480558 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Zhang Jing Last updated by: Zhang Jing 服务内容：重启主机和医院网络系统，调整系统参数. 服务结果： 维修后，设备恢复正常工作 | 3 |  | the worklist lost and system freeze occasionally No error report, the network system of hospital could have issue. The system parameter of host server need to be changed. Rebooted the network system of hospital and changed the system parameter of host server. | [12 empty] | 117480558 |  | 861090 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | NotEngineering-Customer | [5 empty] | Customer | [6 empty] | Worklist | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
819 |  |  | China | JS No.2 Chinese Medicine Hosp. | 345027 | 44313 |  | Escalation/EQ#88392971 | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: EQ#88392971 Next Step Action: FSE Qi Liangdong -86-18662702811 call to order Safety Questions:Case: 0117447512 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Cheng Zhenhua Last updated by: Cheng Zhenhua On this case the host was not rebooted for more than 45 hours (almost 2 days) please instruct the user to reboot the host on daily bases Suggest to perform workaround. | 4 |  | Workaround performed | [12 empty] | 117447512 |  | 846706 |  | Evaluation | Closed - No Further Investigation |  | CTUDT00631367 | [5 empty] | Engineering | [6 empty] | New Yellow Screen | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
820 |  |  | China | Tangshan | 40044 | 44313 |  | 02N/Ingenuity CT 728326 | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 02N/Ingenuity CT 728326 Safety Questions:Case: 0117449055 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Guoping Last updated by: Liu Guoping 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 3 |  | Unknown | [12 empty] | 117449055 |  | 847759 |  | Evaluation | Closed - non-complaint | [7 empty] | Others | [6 empty] | Unknown | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
821 |  |  | Japan |  | 341015 | 44313 |  | CT:APP:NO:Mr. Hisaki: Question. | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Application Support Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer : Q: He wants to know how to create function analysis data A: I had him create it from "Edit Phase". Safety Questions:Case: 0117449279 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Ootani Natsuki Last updated by: Ootani Natsuki 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | Zhu, Chun |  | NAC | [12 empty] | 117449279 |  | 847927 |  |  | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
822 |  |  | Japan |  | 341017 | 44313 |  | CT:APP:NO:Mr. Noda: Question | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Application Support Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer: Q: There was a mistake in the patient name so he wants to know how to correct it. A: I had him do it with me from Change Patient Detail. Safety Questions:Case: 0117449259 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Ootani Natsuki Last updated by: Ootani Natsuki 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | Zhu, Chun |  | NAC | [12 empty] | 117449259 |  | 847904 |  |  | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
823 |  |  | China | HLJ Harbin Senhai Hospital | 345007 | 44312 |  | ESTOP Open | Summary of customer_s problems: The customer didn_t transfer. The engineer said it was intermittently power off Detailed description of Customer_s problems or malfunctions: Intermittently power off problem It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: Intermittently power off, no PRS Remote troubleshoot and test result: Go on site to confirm the log. If necessary order spare parts to be replaced $Q7: (Confirmed reported problem: gantry failure.) $Q8: (check log, couch horizontal failure ) $Q9: (do horizontal calibration) $Q10( Test and inspection data refer to : PA . Device returned to full functionality.) $Q11( Calibrated tools name (N/A) External Remarks: 27.04.2021 02:19:40 UTC Jun Jin (86039442) service details: bed horizontal adjustment Service Result: The device returned to normal. External Remarks: 27.04.2021 02:35:21 UTC Yadong Liu (86039473) service details: bed horizontal adjustment Service Result: The device returned to normal. | 3 |  | bed horizontal adjustment | [12 empty] | 117440075 |  | 841942 |  | Investigation | Closed - Done |  | TBD_Bugreport | [5 empty] | Engineering | [6 empty] | Horizontal calibration | [5 empty] | Refresh |  |  | System issue | System | Couch-SW | TBD-bugrep | Couch-SW-TBD-bugrep | Couch-SW-TBD-bugrep | TBD_CouchEstop | TBD_Bugreport | Couch-SW | Couch-SW-TBD-bugrep | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
824 |  |  | China | Huaying Mountain Guangneng General Hospital | 345000 | 44316 |  | Wechat:02N/Ingenuity Core128/Elite China/Abnormal belt noise | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : Abnormal belt noise Safety Questions:Case: 0117474672 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Cheng Zhenhua Last updated by: Cheng Zhenhua 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Case Transferred to SAP T2 Activities:Problem Description: Summary of customer_s problems: The device makes a strange sound. Detailed description of Customer_s problems or malfunctions: There is noise at the back of the gantry, like the belt is bad. It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: There is noise at the back of the gantry, like the belt is bad. Remote troubleshoot and test result: Contacted the field service engineer, the belt is suspected. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Analyze the source of the noise. If necessary replace the belt. We recommend spare parts 12NC and names: BELT, 5-RIB POLY-VEE, 198" 453566492401 SAP Transfer:Dispatched for On site Support Next Step Action: PPD 453566492401 *1, have arranged for 24H delivery, Field Service Engineer Li Quanfei - 86-13699457740 External Remarks:05.05.2021 08:59:28 UTC Quanfei Li (86001471) Service details: Check the equipment and confirm rotation drive belt is making noises. After replacing the belt and adjusting the position, the machine is back to normal. Service outcome: Device returns to normal working condition after service Internal Remarks: 05.05.2021 08:59:34 UTC Quanfei Li (86001471) $Q7: Confirmation the report problem. Abnormal belt noise. $Q:8:check the device.found the noise from the belt. Need to replace it. $Q9:Replace the "BELT, 5-RIB POLY-VEE, 198" (12nc453566492401)"; the device returned to full service. $Q10: (Test and inspection data refer to: PA. Device returned to full functionality.) $Q11: Calibrated Tools:refer to N/A ----------------------------------------------------------------------------- | 3 |  | BELT, 5-RIB POLY-VEE, 198" (12nc453566492401) | [12 empty] | 117474672 |  | 866174 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | ECR-077012 | [5 empty] | Engineering | [6 empty] | Belt | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Belt | Gantry-HW-Belt | Gantry-HW-Belt | ECR-077012 | Belt | Gantry-HW | Gantry-HW-Belt | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
825 |  |  | China | Huaying Mountain Guangneng General Hospital | 345000 | 44314 |  | 02N/Ingenuity Core128/Elite China/Ask for equipment parameters | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ Safety Questions:Case: 0117454467 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Bo Last updated by: Wang Bo 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action: mail ASP Kangmei chengda follow Next Step Action: Deng Shaoyun of our company has replied, please be aware ----------------------------------------------------------------------------- | 3 |  | NAC_Ask for equipment parameters | [12 empty] | 117454467 |  | 851173 |  | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Customer | [6 empty] | NAC_Ask for equipment parameters | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
826 |  |  | Poland | Agencja Rezerw Materiałowych | 341010 | 44312 |  | Starting the device. | Work Order #: WO-03383399 28-Apr-2021 Mr. Robert Kamola The system fully meets specification and returned to use Problem Description by Engineer: Starting the CT scanner pre-installed in the container. Tools & Equipment : MULTIMETER,2021-11-04 00:00:00 ELECTRICAL SAFETY TESTER,2021-09-14 00:00:00 Resolution: Starting the CT scanner pre-installed in the container. Calibrations and system tests were performed. Image quality tests with a positive result. . Assistance during acceptance tests. . Check of the CT application. Backup upload - system tests and calibrations were performed again. . The device was handed over for operation. ----------------------------------------------------------------------------- Work Order #: WO-03383386 27-Apr-2021 Mr. Mariusz Pachla The system fully meets specification and returned to use Diagnostic performed by Engineer: Starting the device after the container has been transported to the target location. Continuation of service by Robert Kamola. Tools & Equipment : MULTIMETER,2021-06-10 00:00:00 Problem Description by Engineer: Starting the device. ----------------------------------------------------------------------------- | 5 |  | NAC_Calibrations and system tests | [12 empty] | 117442996 |  | 843833 |  | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Others | [6 empty] | NAC_Calibrations and system tests | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
827 |  |  | Germany | Gesundheit Nord gGmbh | 40002 | 44313 |  | Error message occurred | Work Order #: WO-03390628 Carsten Becker The system fully meets specification and returned to use Problem Description by Engineer: Mr. Glogner-Umbeer (user) He does not have any info regarding the reported error Follow-up required: Setup of SSH access with the correct credentials Setup of the PSA Call-back in the remote in order to test access Download of the indicated bug reports for offline error analysis Diagnostic performed by Engineer: System can indeed be reached by remote but dial-in via SSH (FTP or Telnet) not possible since the known credentials do not apply Error message: "Password authentication failed. Please verify that the username and password are correct." In addition, remote access via M2M is not set up on the system. ----------------------------------------------------------------------------- | 3 |  | COPSSH setting might need to check | [12 empty] | 117449763 |  | 844076 |  | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Service | [6 empty] | PRS Config | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
828 |  |  | China | GD Dongguan Dalingshan Hosp. | 345013 | 44319 |  | WeChat:02N/Ingenuity Core128/Elite China/gantry is disconnected automatically | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : gantry is disconnected automatically External Remarks: 04.05.2021 09:47:52 UTC Shaoxiang Huang (86034120) Service details: [When the support is tilting, it loses power. Replace the tilt encoder and do the corresponding calibrations.】 Service outcome: The machine is in normal use. Precautions: [None]服务内容： 【支架打角度的时候出现掉电，更换角度倾斜编码器并且做对应的校正。】 | 3 |  | Replace the tilt encoder and do the corresponding calibrations | [12 empty] | 117483552 |  | 864446 |  | Investigation | Closed - Done |  | NotEngineering-partreplacement | [5 empty] | Others | [6 empty] | Tilt | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
829 |  |  | United States | Clay County Medical Center | 40003 | 44321 |  | Hard to run IQ Check | Work Order #: WO-03440191 05-May-2021 Stauffer Jeffrey The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: General CT exams Expected and Unexpected Results: Expected - the use of the IQ Check would be easy and pass. Unexpected - the IQ Check intermittently fails for no apparent reason and it tells to move the phantoms in inconsistent manners. User Impact: Delays in getting IQ Check to pass. Patient Impact: None Current Software Version:4.1.10 Diagnostic performed by Engineer : Attempted to perform IQ Check as customer does. IQ Check would pass intermittently. Sometimes system software would tell to move the phantom one direction then back the other direction. Follow Up Required : Contact NSS to discuss issue and file Customer Feedback form. | 4 |  | Sometimes system software would tell to move the phantom one direction then back the other direction. | [12 empty] | 117494009 |  | 869677 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00630107 | [5 empty] | Engineering | [6 empty] | Phantom positioning | [5 empty] | Refresh |  |  | System issue | System | DMS-SW | Phantom Detect | DMS-SW-Phantom Detect | DMS-SW-Phantom Detect | CTUDT00630107 | Phantom Detect | DMS-SW | DMS-SW-Phantom Detect | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
830 |  |  | Peru | Hogar Clinica San Juan de Dios | 40034 | 44312 |  | CT installation in the city of Huancayo- For PO | CT installation in the city of Huancayo- For PO | 5 |  | NAC_CT Installation? | [12 empty] | 117446425 |  | 846218 |  | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Service | [6 empty] | NAC_CT Installation? | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
831 |  |  | China |  | 40059 | 44312 |  | WeChat:02G/Ingenuity CT 728326/Move the distribution box location | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer: Move the distribution box location Safety Questions:Case: 0117440143 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Cheng Zhenhua Last updated by: Cheng Zhenhua 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action: have requested to add operation to IA#53232184 and dispatch FSE Lin Zhitao Reason: Customer need to cooperate with migration of power supply cabinet Next Step Action: OP added 0030. ----------------------------------------------------------------------------- | Lu, Xueying |  | NAC | [12 empty] | 117440143 |  | 842007 |  |  | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
832 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 341006 | 44313 |  | Compensator analysis. | Work Order #: WO-03391871 Soner Saglam The system fully meets specification and returned to use Problem Description by Engineer : Compensator crack check. Resolution : No cracks were found. ----------------------------------------------------------------------------- | 3 |  | Compensator analysis | [12 empty] | 117450924 |  | 849003 |  | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Service | [6 empty] | Compensator analysis | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
833 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 341006 | 44313 |  | Table movement is blocked. | Work Order #: WO-03390062 Soner Saglam The system fully meets specification and returned to use Problem Description by Engineer : E-stop was activated. Resolution : Likely e-stop activation during cleaning. The system is functional. ----------------------------------------------------------------------------- | 3 |  | Likely e-stop activation during cleaning. The system is functional. | [12 empty] | 117449191 |  | 843956 |  | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | E-Stop button | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
834 |  |  | United States |  | 300033 | 44321 |  | 0799CFXXX BU reports daily CTHost power down restart is overdue and RADAR shows it is generally not done, which may... | Work Order #: WO-03442627 Carpenter, Kevin The system fully meets specification and returned to use Diagnostic performed by Engineer : BU reports daily CTHost power down restart is overdue and RADAR shows it is generally not done, which may cause CT system to not preform optimally. Problem Description by Engineer : Customer Function/Role: BU / RADAR Device Use: Diag Expected and Unexpected Results: Remote ProActive BU Alert User Impact: None Patient Impact: None Current Software Version: IDK Resolution : BU reports daily CTHost power down restart is overdue and RADAR shows it is generally not done, which may cause CT system to not preform optimally. CT RSE will recheck system in a few days ~ if CTHost has still not been restarted then this case will be dispatched to the Field so the local FSE can speak to the customer. ----------------------------------------------------------------------------- | 5 |  | Non-Refresh | [12 empty] | 117496283 |  | 869716 |  | #N/A | Investigation | [14 empty] | Non-Refresh | [5 empty] | Non-Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | y
835 |  |  | United Kingdom |  | 336286 | 44320 |  | Tube warm up failed - error insure system is not in a working state | Work Order #: WO-03433992 Jeffrey Thompson The system fully meets specification and returned to use Regulatory Questions : N/A Problem Description by Engineer : Problem description by engineer : Spoke to Melanie who confirmed the issue, they tried to do a short tube conditioning and the system showed an error ensure system is in a working state. They then did a QA and this was successful and then could complete the short tube conditioning. Bug report has been completed and advised will check logs. ***** No patient re-scans or re-injections ***** When was the first occurrence? Within... : 04-May-2021 07:05:00 How often is issue occurring? : First occurrence Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Radiographer Device Use: No patient involved Expected and Unexpected results: Attempted short tube conditioning and unable to carry this out Patient Impact/Outcome: No patient re-scans or re-injections User Impact: Unable to complete short tube conditioning Current Software Version: 4.1.7.25026 Diagnostic performed by Engineer : Logs checked and can see the horizontal appears to have been left on a positive limit S_HMC_AXIS_AT_POSITIVE_LIMIT (Couch is at horizontal outer limit. (soft limit)) Which may have caused the issue. Always good to move the pallet a few millimetres forward, although on second check of the logs the encoder may have shown it had been moved. This was followed by a S_HMC_AXIS_DEST_COLLISION error which could highlight that also the vertical height of the table may also not have been correct for the short tube conditioning to complete. Hence when the QA carried out it would have moved the table and moved it into specification for the short tube conditioning. Internal Comments : Error # and/or description of error : S_HMC_AXIS_AT_POSITIVE_LIMIT (Couch is at horizontal outer limit. (soft limit)) Which may have caused the issue. Always good to move the pallet a few millimetres forward, although on second check of the logs the encoder may have shown it had been moved. This was followed by a S_HMC_AXIS_DEST_COLLISION error which could highlight that also the horizontal position and possibly the vertical height of the table may also not have been correct for the short tube conditioning to complete. Hence when the QA carried out it would have moved the table and moved it into specification for the short tube conditioning. Unable to check due to cantrace issues. Completed or Recommended Repair Action "Internal" : No further action required. System is in full clinical use with no further issues reported. Resolution : No further problems reported - case closed. ----------------------------------------------------------------------------- | 4 |  | Non-Refresh | [12 empty] | 117489232 |  | 869069 |  | #N/A | Opened | [14 empty] | Non-Refresh | [5 empty] | Non-Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | y
836 |  |  | Spain | Policlínica Nuestra Señora del Rosario | 336190 | 44320 |  | The CT is blocked. Does not recognize the password | Work Order #: WO-03433461 Emilio Olmo Soria The system fully meets specification and returned to use. Diagnostic performed by Engineer: The login of the equipment is blocked; after waiting it unlocks by itself, possibly the passwords were entered several times wrongly and it got blocked for a while. Now it lets you enter and the equipment has started correctly. Problem Description by Engineer: The CT is blocked. Does not recognize the password Regulatory Questions : N/A Resolution: The equipment has been unlocked and they are already working. | 2 |  | Password | [12 empty] | 117488731 |  | 864494 |  | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Application | [6 empty] | Password | [5 empty] | Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
837 |  |  | United States | Select Specialty Hospital | 32261 | 44320 |  | 0699CDN84 Gantry will not expose | Work Order #: WO-03432556 Mark Finnigan The system does not meet specification and is out of use Problem Description by Engineer : Customer Function/Role: Tech Device Use: Unknown Expected and Unexpected Results: Customer reports that they can not scan. User Impact: Unknown Patient Impact: Current Software Version: Diagnostic performed by Engineer : The case came in on the Web Portal so there was no discussion with the customer. Called the provided phone number and it went to voicemail. Tried to log into the site via PRS and the connection will not establish. All ports are blocked so no remote diagnosis is possible. Dispatching the case for first thing in the morning 5/4. Will leave the customer a VM stating this. Follow Up Required : The case came in on the Web Portal so there was no discussion with the customer. Called the provided phone number and it went to voicemail. Tried to log into the site via PRS and the connection will not establish. All ports are blocked so no remote diagnosis is possible. Dispatching the case for first thing in the morning 5/4. Will leave the customer a VM stating this. ----------------------------------------------------------------------------- Found communication error between gantry and CIRS in error logs. Reseated fiber optics on stator transmitter. Tried multiple times to restart, and allow system to shut off completely for over 30 minutes then restarted. Attempted to run quality checks, and system will not allow the gantry to expose. System went down for a period of 1 hour, and then I turned it back on and ran my checks. The checks worked, and then I attempted to scan a patient. The system would not allow me to scan patient. I then tried to run checks again, but system would not allow me to scan. I checked both computers, and all other accessories that operate the machine. All systems are fine, and the tube heat utility says that the tube heat is fine. Every time I try to expose for checks now it will only give me the generalized error that short tube conditioning Failed. | 3 |  | Gantry CIRS communication | [12 empty] | 117487898 |  | 863920 |  | Closed - Done | Closed - Done |  | TBD_Bugreport | [5 empty] | Engineering | [6 empty] | Gantry CIRS communication | [5 empty] | Refresh |  |  | Upgrade issue | System | TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
838 |  |  | United States | Advent Health Wesley Chapel | 300099 | 44319 |  | 1302CFXXX Issue transferring to PACS | Work Order #: WO-03429011 Powell, Jamie The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : Logged in via remote console and took a look at customer_s queue manager which was empty.. Tried transferring to PACS which was successful. Stayed on the line while customer scanned another patient and she was able to transfer to PACS with no issues. Resolution : Logged in via remote console and took a look at customer_s queue manager which was empty.. Tried transferring to PACS which was successful. Stayed on the line while customer scanned another patient and she was able to transfer to PACS with no issues. ----------------------------------------------------------------------------- | 3 |  | PACS Transfer | [12 empty] | 117485312 |  | 862476 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS Transfer | [5 empty] | Refresh |  |  | Upgrade issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
839 |  |  | United States |  | 68116 | 44319 |  | Remote Alert: "CT-PRO-GANTRY-COUCH_Vertical_Motion | Work Order #: WO-03428856 Powell, Jamie The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: M2M Device Use: Diagnostic Expected and Unexpected Results: N/A User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : Remote Alert: "CT-PRO-GANTRY-COUCH_Vertical_Motion Follow Up Required : "CT-PRO-GANTRY-COUCH_Vertical_Motion Dear colleague, We ask your attention for the following Proactive Alert. To prevent repeating cases, keep the case open until the system issue has been solved. PLANNING ADVICE Please contact customer in order to make an appointment. In order to prevent a customer call, please follow-up . PROBLEM DESCRIPTION The Logfiles show problems with the vertical couch movement. This model alerts when too many of following errors are showing: S_VMC_AXIS_UNALLOWED_MOVE S_VMC_AXIS_AT_COLLISION S_MC_DEVICE_VERT_POT_OPEN S_MC_DEVICE_VERT_POT_OPEN_OR_SHORTED Caused by a mismatch between expected and actual vertical position of the couch, or a vertical movement happened without being commanded. The attached graph is showing which were found and which one triggered the alert. Although not alerting on below errors, these are also shown for your reference: S_MC_AXIS_VERT_INVERTER_FAULT S_MC_DEVICE_INVERTER_FAILURE 0x20000DD S_MC_DEVICE_ESTOP_IS_OPEN S_MC_DEVICE_COUCH_VERTICAL_HW_FAULT S_VMC_AXIS_HW_DETECTED_OBSTRUCTION S_VMC_AXIS_STOP_REASON_UNKNOWN S_VMC_AXIS_WRONG_DEST_REACHED CUSTOMER IMPACT Workflow-, scan interrupts or estop opens. OCCURRENCES OVERVIEW See the attached PDF graph (Attachments) in OneEMS / ServiceMax for this case. RECOMMENDED ACTION Verify the vertical movement over the whole range. Verify limitswitches, cabling and clean string pot by carefully extending it over its complete range. Parts involved are the vertical string pot, vertical drive inverter and CCB. When the graph also shows: S_MC_AXIS_VERT_INVERTER_FAULT S_MC_DEVICE_INVERTER_FAILURE S_MC_DEVICE_COUCH_VERTICAL_HW_FAULT, Table moves in small steps.. Additionally verify: - O-Ring attached to the vertical Encoder. - the switch of the safety green bar. - the cable to the switch of the green safety green bar. - and adjust the vertical separation limit switch - CK1 ESTIMATED TECHNICAL REPAIR TIME Depending on findings. ADDITIONAL INFORMATION For more guidance, see Section 2.2 Vertical Motion Troubleshooting, of the Brilliance/Ingenuity, Patient Support Theory and Troubleshooting Guide, 459800414331. If you have any questions or remarks, please let us know. Kind Regards, Remote.Monitoring.Team.CT@philips.com IS CS – Remote Services" ----------------------------------------------------------------------------- | 5 |  | Non-Refresh | [12 empty] | 117485192 |  | 862341 |  | #N/A | Evaluation | [14 empty] | Non-Refresh | [5 empty] | Non-Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | y
840 |  |  | United States |  | 68114 | 44318 |  | 0704CDN84 Host not booting up. | Work Order #: WO-03425284 Miller, John M The system does not meet specification and is out of use Problem Description by Engineer : Host won_t boot. Customer Function/Role:Technician/Sonographer/Therapist Device Use: Diagnostic Expected and Unexpected Results:Host won_t boot. User Impact: None Patient Impact: None Current Software Version: 4.1.7 Follow Up Required : Customer states the Host won_t boot all the way up. Hangs in a Blue Screen. Suggest running the Dell or HP Diagnostic (Dell press F12 during bootup, HP press Esc then F2) to determine the health of the PC. Also check the health of the local drive. Also suggest re-seating memory modules. Suggest replacing hard drive and reload software on new drive. Diagnostic performed by Engineer : Host won_t boot. Internal Comments : Customer states the Host won_t boot all the way up. Hangs in a Blue Screen. Suggest running the Dell or HP Diagnostic (Dell press F12 during bootup, HP press Esc then F2) to determine the health of the PC. Also check the health of the local drive. Also suggest re-seating memory modules. Suggest replacing hard drive and reload software on new drive. ----------------------------------------------------------------------------- Work Order #: WO-03425289 Scheyer William Problem Description by Engineer : Host won_t boot. Customer Function/Role:Technician/Sonographer/Therapist Device Use: Diagnostic Expected and Unexpected Results:Host won_t boot. User Impact: None Patient Impact: None Current Software Version: 4.1.7 Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 2 |  | Non-Refresh | [12 empty] | 117482235 |  | 861389 |  | #N/A | Evaluation | [14 empty] | Non-Refresh | [5 empty] | Non-Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | y
841 |  |  | United States | Lakeland Regional Medical Cent | 320474 | 44318 |  | 1101CFXXX can not get gantry on | Work Order #: WO-03424664 Mcginnis, Mike The system fully meets specification and returned to use Internal Comments : Error # and/or description of error : can not get gantry on Completed or Recommended Repair Action "Internal" : Jeff said the tech called him in. She told him the screen froze so they restarted the system and now they can not get the gantry on. I questioned him and the host was on but the fans in the gantry were not. I had him check for a start/stop switch in the control room. He found it and hit start and the gantry can not on. Closed ESTOP ok so we ran a test scan and he said everything looks good at this point. Resolution : Jeff said the tech called him in. She told him the screen froze so they restarted the system and now they can not get the gantry on. I questioned him and the host was on but the fans in the gantry were not. I had him check for a start/stop switch in the control room. He found it and hit start and the gantry can not on. Closed ESTOP ok so we ran a test scan and he said everything looks good at this point. Diagnostic performed by Engineer : Jeff said the tech called him in. She told him the screen froze so they restarted the system and now they can not get the gantry on. I questioned him and the host was on but the fans in the gantry were not. I had him check for a start/stop switch in the control room. He found it and hit start and the gantry can not on. Closed ESTOP ok so we ran a test scan and he said everything looks good at this point. Problem Description by Engineer : Problem description by engineer : can not get gantry on When was the first occurrence? Within... : 01-May-2021 How often is issue occurring? : First occurrence Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: can not get gantry on Patient Impact/Outcome: none User Impact: system was down Current Software Version: 4.1.7.25026 ----------------------------------------------------------------------------- | 2 |  | E-Stop button | [12 empty] | 117481691 |  | 861092 |  | Closed - Done | Closed - Done | [7 empty] | Customer | [6 empty] | E-Stop button | [5 empty] | Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
842 |  |  | United States |  | 300167 | 44316 |  | 0599CFXXX E-stop not closing | Work Order #: WO-03422666 Boyd Craig The system does not meet specification and is out of use Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: E-stop not closing User Impact: None Patient Impact: None Current Software Version: Diagnostic performed by Engineer : E-stop not closing when turning key on CT Box. Power cycle system. Excise E-stops. Boot system. OK. Resolution : Power cycle system. Excise E-stops. Boot system. OK. | 2 |  | Non-Refresh | [12 empty] | 117480067 |  | 859309 |  | #N/A | Investigation | [14 empty] | Non-Refresh | [5 empty] | Non-Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | y
843 |  |  | United States | Upson Regional Medical Center | 310115 | 44316 |  | 0612CDY72 E-Stop Opened during contrast injection | Work Order #: WO-03420543 Weaver David The system does not meet specification and is out of use Problem Description by Engineer : DESCRIPTION OF PROBLEM (EXPECTED/UNEXPECTED RESULTS) E-Stop Opened during contrast injection. CALLER FUNCTION/ROLE Technician/Sonographer/Therapist DEVICE INTENDED USE Diagnostic WHEN WAS THE FIRST OCCURRENCE? WITHIN . . . Same day HOW OFTEN IS ISSUE OCCURRING? First occurrence WHERE IN THE END-USER WORKFLOW IS THE ISSUE OCCURRING? During acquisition/scan WHAT RECENT EVENTS HAVE OCCURRED? During normal use NEGATIVE OPERATOR IMPACT Delay in scanning NEGATIVE PATIENT IMPACT/PROCEDURE OUTCOME Can not be scanned HAS CUSTOMER CREATED A SPECIFIC SYSTEM EVENT REPORT OF ISSUE? No 4.1.7 Diagnostic performed by Engineer : E-Stop Opened during contrast injection. Follow Up Required : Saw S_MP_CTRL_NO_MORE_AT_SPEED in logger.mdb. Speed is off more than 2%. Parameter 1 is wanted speed in RPM/100 (IT IS AT 8000). Parameter 2 is actual speed in RPM/100. (IT IS AT 0) If P2=0 then rotation lost its speed and position. This may be caused by Ctrl link errors. Check the counters in the EAL Report. The Rotor did not arrive to the desired speed. Recommendation: 1. Perform Spindle block commission. 2. Check the belt tension. 3. Perform Rotor balancing test. 4. Replace the Spindle block. 5. Replace GMP. ----------------------------------------------------------------------------- Seeing a lot of GPU configuration errors in the logs, and the FSE stated that there were GPU issues a few months back, but were resolved with reseating the board. As there is currently no stock available, requested the FSE reseat the board again to see if the issue goes away temporarily. Work Order #: WO-03420710 04-May-2021 Quach Duy The system fully meets specification and returned to use Problem Description by Engineer : DESCRIPTION OF PROBLEM (EXPECTED/UNEXPECTED RESULTS) E-Stop Opened during contrast injection. CALLER FUNCTION/ROLE Technician/Sonographer/Therapist DEVICE INTENDED USE Diagnostic WHEN WAS THE FIRST OCCURRENCE? WITHIN . . . Same day HOW OFTEN IS ISSUE OCCURRING? First occurrence WHERE IN THE END-USER WORKFLOW IS THE ISSUE OCCURRING? During acquisition/scan WHAT RECENT EVENTS HAVE OCCURRED? During normal use NEGATIVE OPERATOR IMPACT Delay in scanning NEGATIVE PATIENT IMPACT/PROCEDURE OUTCOME Can not be scanned HAS CUSTOMER CREATED A SPECIFIC SYSTEM EVENT REPORT OF ISSUE? No 4.1.7 Diagnostic performed by Engineer : intermittent estop open. view logger. checked gantry belt. checked tension on belt. logger state framing err. Tools & Equipment : MULTIMETER,2021-07-31 00:00:00 Resolution : clean power/signal brushed and slip rings ----------------------------------------------------------------------------- | 3 |  | E-Stop_contrast injection | [12 empty] | 117478456 |  | 858161 |  | Closed - Done | Closed - Done |  | ECR-075975 | [5 empty] | Engineering | [6 empty] | E-Stop_contrast injection | [5 empty] | Refresh |  |  | Upgrade issue | System | Gantry-HW | Lexan Ring | Gantry-HW-Lexan Ring | Gantry-HW-Lexan Ring | ECR-075975 | Lexan Ring | Gantry-HW | Gantry-HW-Lexan Ring | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
844 |  |  | United States |  | 310034 | 44315 |  | NEMO | Work Order #: WO-03412540 Parts Consumed : 4.00000,453567027551,CORNER, PEDESTAL COVERS ----------------------------------------------------------------------------- | 3 |  | Non-Refresh | [12 empty] | 117470756 |  | 857895 |  |  | Closed - Non-Complaint | [14 empty] | Non-Refresh | [5 empty] | Non-Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | y
845 |  |  | United States |  | 310165 | 44314 |  | Part Order Per NSS | Work Order #: WO-03404261 Foltz John Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 5 |  | Non-Refresh | [12 empty] | 117462685 |  | 854751 |  | #N/A | Evaluation | [14 empty] | Non-Refresh | [5 empty] | Non-Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | y
846 |  |  | United States |  | 310165 | 44314 |  | 0308CDN84 Remote Alert: "CT-PRO-GANTRY-S_MC_DEVICE_COUCH_VERTICAL_HW_FAULT | Work Order #: WO-03400813 Holeman, David The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: M2M/RADAR Device Use: No Patient Involved Expected and Unexpected Results: Remote Alert: "CT-PRO-GANTRY-S_MC_DEVICE_COUCH_VERTICAL_HW_FAULT User Impact: None Patient Impact: None Current Software Version: 4.1.7 Diagnostic performed by Engineer : Remote Alert: "CT-PRO-GANTRY-S_MC_DEVICE_COUCH_VERTICAL_HW_FAULT Follow Up Required : PROBLEM DESCRIPTION This model alerts when too many of the following error shows in the logs: S_MC_DEVICE_COUCH_VERTICAL_HW_FAULT Couch does not move up or down smoothly but operator needs to press the button several times to lower or raise the couch. This usually happens when (un)loading the patient and is often the combination of a mechanical obstruction and a (heavy) load. If the couch only moves in small steps of 0.5 or 1mm then the O-ring to the vertical encoder is broken. The attached graph is showing the details of the occurrences and the alert reason. CUSTOMER IMPACT When vertical position is eventually reached, customer can continue ok with scanning OCCURRENCES OVERVIEW See the attached PDF graph (Attachments) in OneEMS / ServiceMax for this case. RECOMMENDED ACTION First check the couch vertical motion to operate smoothly over its complete range. Preferably this should be done with a load simulating a patient in the fully out position. This allows the couch to go completely down. In case the problem occurs all the time: - Check / replace the O-Ring attached to the vertical encoder. - Check / replace the switch of the safety green bar. - For Extended couch only: Check/ replace the two vertical safety switches in series with the green safety bar switch. - Check / replace the cable to the switch of the green safety green bar. - Check CK1 at the Inverter PCB. - Replace the Inverter PCB ----------------------------------------------------------------------------- | 5 |  | Non-Refresh | [12 empty] | 117458748 |  | 849370 |  | #N/A | Evaluation | [14 empty] | Non-Refresh | [5 empty] | Non-Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | y
847 |  |  | United States | Advent Health Wesley Chapel | 300099 | 44313 |  | 4.1.10 CTA head and neck images freze and boulus tracking isssues |  | 5 |  | CTA head and neck images freze and boulus tracking isssues | [12 empty] | 117451229 |  | 849298 |  | Closed - Done | Closed - Done |  | TBD-bugrep | [5 empty] | Engineering | [6 empty] | CTA head and neck images freze and boulus tracking isssues | [5 empty] | Refresh |  |  | Upgrade issue | System | Console-SW | TBD-bugrep | Console-SW-TBD-bugrep | Console-SW-TBD-bugrep | TBD-bugrep | TBD_Bugreport | Console-SW | Console-SW-TBD-bugrep | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
848 |  |  | Netherlands | Rijnstate Ziekenhuis Arnhem | 52101 | 44314 |  | Table makes sudden movement in certain position | Work Order #: WO-03400251 NIELS MELICK VAN The system fully meets specification and returned to use Follow Up Required: check table movement Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer: Check table movement ----------------------------------------------------------------------------- Work Order #: WO-03400291 30-Apr-2021 EDDY ELFRING The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : Checked firmware couch controller and calibrated the horizontal table movement. Made a service backup. ----------------------------------------------------------------------------- | 5 |  | Checked firmware couch controller and calibrated the horizontal table movement. Made a service backup. | [12 empty] | 117458894 |  | 848976 |  | Investigation | Investigation |  | TBD-bugrep | [5 empty] | Engineering | [6 empty] | Table sudden movement in certain position | [5 empty] | Refresh |  |  | Upgrade issue | System | Couch-SW | TBD-bugrep | Couch-SW-TBD-bugrep | Couch-SW-TBD-bugrep | TBD_CouchEstop | TBD_Bugreport | Couch-SW | Couch-SW-TBD-bugrep | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
849 |  |  | United States | Advent Health Wesley Chapel | 300099 | 44314 |  | 0711CDN82 both host monitors go off during scans | Work Order #: WO-03399245 Sheffield, Scott The system partly meets specification and is returned to use with restrictions Internal Comments : Error # and/or description of error : host system event logs Type : Error Date : 4/28/2021 Time : 7:43:04 AM Event : 1001 Source : BugCheck Category : None User : N/A Computer : 128-CT-ED Description: The computer has rebooted from a bugcheck. The bugcheck was: 0x00000119 (0x0000000000000002, 0xffffffffc000000d, 0xffff8281e8e5c2a0, 0xffffb40820b60820). A dump was saved in: C:\Windows\MEMORY.DMP. Report Id: 05acd35c-5605-46ff-8885-89d49869366a. Type : Error Date : 4/28/2021 Time : 8:10:32 AM Event : 10001 Source : DCOM Category : None User : \S-1-5-21-3853370449-718592429-2989568474-1007 Computer : 128-CT-ED Description: Unable to start a DCOM Server: microsoft.windows.immersivecontrolpanel as Unavailable/Unavailable. The error: "5" Happened while starting this command: "C:\Windows\ImmersiveControlPanel\SystemSettings.exe" -ServerName:microsoft.windows.immersivecontrolpanel Type : Error Date : 4/28/2021 Time : 8:10:24 AM Event : 10001 Source : DCOM Category : None User : \S-1-5-21-3853370449-718592429-2989568474-1007 Computer : 128-CT-ED Description: Unable to start a DCOM Server: CortanaUI as Unavailable/Unavailable. The error: "5" Happened while starting this command: "C:\windows\SystemApps\Microsoft.Windows.Cortana_cw5n1h2txyewy\SearchUI.exe" -ServerName:CortanaUI.AppXa50dqqa5gqv4a428c9y1jjw7m3btvepj.mca Type : Error Date : 4/28/2021 Time : 7:43:03 AM Event : 6008 Source : EventLog Category : None User : N/A Computer : 128-CT-ED Description: The previous system shutdown at 7:26:55 AM on ?4/?28/?2021 was unexpected. Type : Error Date : 4/28/2021 Time : 7:43:34 AM Event : 7034 Source : Service Control Manager Category : None User : N/A Computer : 128-CT-ED Description: The GantryServiceHost service terminated unexpectedly. It has done this 3 time(s). host application event logs Type : Error Date : 4/28/2021 Time : 8:10:24 AM Event : 5973 Source : Microsoft-Windows-Immersive-Shell Category : Activation of app %1 failed with error: %2 See the Microsoft-Windows-TWinUI/Oper User : \S-1-5-21-3853370449-718592429-2989568474-1007 Computer : 128-CT-ED Description: Activation of app Microsoft.Windows.Cortana_cw5n1h2txyewy!CortanaUI failed with error: -2147024891 See the Microsoft-Windows-TWinUI/Operational log for additional information. Type : Error Date : 4/28/2021 Time : 8:10:08 AM Event : 1000 Source : Application Error Category : Application Crashing Events User : N/A Computer : 128-CT-ED Description: Faulting application name: Philips.PmsCT.Service.AutomaticDownloadApplication.exe, version: 1.0.0.0, time stamp: 0x5deb7d46 Faulting module name: KERNELBASE.dll, version: 6.2.14393.3085, time stamp: 0x5d1d7bde Exception code: 0xe0434f4d Fault offset: 0x000dc1e2 Faulting process id: 0x(null) Faulting application start time: 0x(null) Faulting application path: (null) Faulting module path: (null) Report Id: (null) Faulting package full name: (null) Faulting package-relative application ID: (null) Type : Error Date : 4/28/2021 Time : 7:43:42 AM Event : 5973 Source : Microsoft-Windows-Immersive-Shell Category : Activation of app %1 failed with error: %2 See the Microsoft-Windows-TWinUI/Oper User : \S-1-5-21-3853370449-718592429-2989568474-1007 Computer : 128-CT-ED Description: Activation of app Microsoft.Windows.Cortana_cw5n1h2txyewy!CortanaUI failed with error: -2147024891 See the Microsoft-Windows-TWinUI/Operational log for additional information. Completed or Recommended Repair Action "Internal" : It appears that the host software is crashing. Consult with NSS about needing to reload host computer software. Diagnostic performed by Engineer : I had Russ create a bug report. The problem happened at 7:30 am today but has been happening since Sunday he said. I downloaded logs and found alot of errors in host system and application event logs. It appears that the host software is crashing. Below are some errors from the host event logs. Problem Description by Engineer : Problem description by engineer : both host monitors go off during scans. This only happens during scans. Both host monitors display go black and the monitor power buttons turn amber color. He said that the host computer has lights on when this happens he has to turn off host computer and turn it back on. The monitors work after rebooting and he is able to finish the scan. It happened today at 7:30 am. I had him create a bug report. FSE has been working on this issue but no previous work order for this issue. Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: host computer monitors go black during scans. Patient Impact/Outcome: delay User Impact: delay Current Software Version: 4.1.7.25030 Follow Up Required : Error # and/or description of error : host system event logs Type : Error Date : 4/28/2021 Time : 7:43:04 AM Event : 1001 Source : BugCheck Category : None User : N/A Computer : 128-CT-ED Description: The computer has rebooted from a bugcheck. The bugcheck was: 0x00000119 (0x0000000000000002, 0xffffffffc000000d, 0xffff8281e8e5c2a0, 0xffffb40820b60820). A dump was saved in: C:\Windows\MEMORY.DMP. Report Id: 05acd35c-5605-46ff-8885-89d49869366a. Type : Error Date : 4/28/2021 Time : 8:10:32 AM Event : 10001 Source : DCOM Category : None User : \S-1-5-21-3853370449-718592429-2989568474-1007 Computer : 128-CT-ED Description: Unable to start a DCOM Server: microsoft.windows.immersivecontrolpanel as Unavailable/Unavailable. The error: "5" Happened while starting this command: "C:\Windows\ImmersiveControlPanel\SystemSettings.exe" -ServerName:microsoft.windows.immersivecontrolpanel Type : Error Date : 4/28/2021 Time : 8:10:24 AM Event : 10001 Source : DCOM Category : None User : \S-1-5-21-3853370449-718592429-2989568474-1007 Computer : 128-CT-ED Description: Unable to start a DCOM Server: CortanaUI as Unavailable/Unavailable. The error: "5" Happened while starting this command: "C:\windows\SystemApps\Microsoft.Windows.Cortana_cw5n1h2txyewy\SearchUI.exe" -ServerName:CortanaUI.AppXa50dqqa5gqv4a428c9y1jjw7m3btvepj.mca Type : Error Date : 4/28/2021 Time : 7:43:03 AM Event : 6008 Source : EventLog Category : None User : N/A Computer : 128-CT-ED Description: The previous system shutdown at 7:26:55 AM on ?4/?28/?2021 was unexpected. Type : Error Date : 4/28/2021 Time : 7:43:34 AM Event : 7034 Source : Service Control Manager Category : None User : N/A Computer : 128-CT-ED Description: The GantryServiceHost service terminated unexpectedly. It has done this 3 time(s). host application event logs Type : Error Date : 4/28/2021 Time : 8:10:24 AM Event : 5973 Source : Microsoft-Windows-Immersive-Shell Category : Activation of app %1 failed with error: %2 See the Microsoft-Windows-TWinUI/Oper User : \S-1-5-21-3853370449-718592429-2989568474-1007 Computer : 128-CT-ED Description: Activation of app Microsoft.Windows.Cortana_cw5n1h2txyewy!CortanaUI failed with error: -2147024891 See the Microsoft-Windows-TWinUI/Operational log for additional information. Type : Error Date : 4/28/2021 Time : 8:10:08 AM Event : 1000 Source : Application Error Category : Application Crashing Events User : N/A Computer : 128-CT-ED Description: Faulting application name: Philips.PmsCT.Service.AutomaticDownloadApplication.exe, version: 1.0.0.0, time stamp: 0x5deb7d46 Faulting module name: KERNELBASE.dll, version: 6.2.14393.3085, time stamp: 0x5d1d7bde Exception code: 0xe0434f4d Fault offset: 0x000dc1e2 Faulting process id: 0x(null) Faulting application start time: 0x(null) Faulting application path: (null) Faulting module path: (null) Report Id: (null) Faulting package full name: (null) Faulting package-relative application ID: (null) Type : Error Date : 4/28/2021 Time : 7:43:42 AM Event : 5973 Source : Microsoft-Windows-Immersive-Shell Category : Activation of app %1 failed with error: %2 See the Microsoft-Windows-TWinUI/Oper User : \S-1-5-21-3853370449-718592429-2989568474-1007 Computer : 128-CT-ED Description: Activation of app Microsoft.Windows.Cortana_cw5n1h2txyewy!CortanaUI failed with error: -2147024891 See the Microsoft-Windows-TWinUI/Operational log for additional information. Completed or Recommended Repair Action "Internal" : It appears that the host software is crashing. Consult with NSS about needing to reload host computer software. ----------------------------------------------------------------------------- Work Order #: WO-03399885 30-Apr-2021 Garcia Marrero, Jorge The system fully meets specification and returned to use Problem Description by Engineer : Problem description by engineer : both host monitors go off during scans. This only happens during scans. Both host monitors display go black and the monitor power buttons turn amber color. He said that the host computer has lights on when this happens he has to turn off host computer and turn it back on. The monitors work after rebooting and he is able to finish the scan. It happened today at 7:30 am. I had him create a bug report. FSE has been working on this issue but no previous work order for this issue. Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: host computer monitors go black during scans. Patient Impact/Outcome: delay User Impact: delay Current Software Version: 4.1.7.25030 Diagnostic performed by Engineer : Re-loading of operating system and applications in host computer. Replacement of ordered parts. System is fully operational. ----------------------------------------------------------------------------- | 3 |  | 1.00000,459801633871,16GB DDR4-2666 (2x8GB) RegRAM 1.00000,459801127432,PCIe CAN CARD PCIe-402-1 1.00000,459801536801,NVIDIA Quadro P400 2GB (3)mDP GFX 1.00000,459800962351,LCC CRC RACK SDB BD ASSY Need check bugreport | [12 empty] | 117458073 |  | 848690 |  | Closed - Done | Closed - Done |  | TBD-bugrep | [5 empty] | Engineering | [6 empty] | Black screen | [5 empty] | Refresh |  |  | Upgrade issue | System | Console-HW | BlackScreen | Console-HW-BlackScreen | Console-HW-BlackScreen | BlackScreen | BlackScreen | Console-HW | Console-HW-BlackScreen | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
850 |  |  | United States |  | 68116 | 44313 |  | 0999CDY72 Ring artifacts | Work Order #: WO-03394635 Boyd Craig The system partly meets specification and is returned to use with restrictions Follow Up Required : FSE on site to evaluate artifact. Recommendations: Clean gantry window. Check for anything in the field of view. Perform air calibrations. Perform Test scans. Analyze images. Run bad detector test. Diagnostic performed by Engineer : PRS accessed. No errors showing. Customer is third party and not on site. Requesting case be paged out. Issue started a few days ago. Run Compensator Analyzer. No indication of crack in the compensator was found. Problem Description by Engineer : Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected Results: Ring artifacts User Impact: None Patient Impact: None Current Software Version: ----------------------------------------------------------------------------- Work Order #: WO-03394698 28-Apr-2021 Blackford Cable The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected Results: Ring artifacts User Impact: None Patient Impact: None Current Software Version: Diagnostic performed by Engineer : ring artifacts in iac protical ----------------------------------------------------------------------------- | 3 |  | Non-Refresh | [12 empty] | 117453042 |  | 846067 |  | #N/A | Investigation | [14 empty] | Non-Refresh | [5 empty] | Non-Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | y
851 |  |  | United States | Wellstar Kennestone Hospital | 68117 | 44313 |  | 0101CFXXX Failed during bolus scan | Work Order #: WO-03394484 Mcginnis, Mike The system fully meets specification and returned to use Problem Description by Engineer : Problem description by engineer : Failed during bolus scan How often is issue occurring? : First occurrence Malfunction area : Others Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Failed during bolus scan Patient Impact/Outcome: patient was rescanned, customer stated no one was harmed User Impact: had to restart the scan Current Software Version: 4.1.7.25037 Diagnostic performed by Engineer : Spoke to Melissa, Philips Clinical Applications. She was on site doing follow up training. They had an issue earlier today where a bolus scan stopped. They restarted the host and cirs and scanned the patient ok. Said the system has been working ok since then. The wanted me to review the bug report and let the NSS Sam Pierce know of the incident as he will be on site tomorrow morning to do a Win10 upgrade on the system. I downloaded the bug report and log files via our scripts. Neither of them contained the cirs event logs. See error information found below. Notified NSS Sam Pierce of the findings Internal Comments : Error # and/or description of error : 2021/04/27 14:58:47.730__001 S_MASTER_SEQUENCER_ACQ_NOT_OK ERROR_NORMAL Completed or Recommended Repair Action "Internal" : Their was not enough information to determine the cause of the error. No CIRS logs in bug report or script logs. Notified NSS Sam Pierce of the findings ----------------------------------------------------------------------------- | 3 |  | Need bugreport for analysis | [12 empty] | 117452801 |  | 845982 |  | Closed - Done | Closed - Done |  | TBD-Bugrep | [5 empty] | Engineering | [6 empty] | bolus scan | [5 empty] | Refresh |  |  | Upgrade issue | System | Gantry-SW | TBD-Bugrep | Gantry-SW-TBD-Bugrep | Gantry-SW-TBD-Bugrep | TBD-Bugrep | TBD-Bugrep | Gantry-SW | Gantry-SW-TBD-Bugrep | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
852 |  |  | United States |  | 300106 | 44312 |  | LS - 061663 AND LS - 061661 | Work Order #: WO-03386674 Parts Consumed : 1.00000,459800041131,Cushion-Extended Patient Support 1.00000,459800126661,Cushion-47.24 Inch Mattress Pad ----------------------------------------------------------------------------- | 5 |  | Non-Refresh | [12 empty] | 117445782 |  | 845664 |  |  | Closed - Non-Complaint | [14 empty] | Non-Refresh | [5 empty] | Non-Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | y
853 |  |  | United States | Lakeland Regional Medical Cent | 320474 | 44313 |  | Q/O: Cannot connect to system remotely | Work Order #: WO-03392258 27-Apr-2021 Watson Adam The system fully meets specification and returned to use Resolution : Attempted to remotely log into system but found a connection could not be established. Went on site, reset remoteservice password, and resolved the remote connection issue. Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: ----------------------------------------------------------------------------- | 5 |  | PRS Config | [12 empty] | 117451257 |  | 845148 |  | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Service | [6 empty] | PRS Config | [5 empty] | Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
854 |  |  | Argentina | Hospital Dr. Bernardo Houssay | 40027 | 44324 |  | THE INJECTOR PUMP GETS ERROR 311 AND IS LOCKED | It is validated with Ing Bianco, data is obtained from the failing injector, a visit from the injector supplier is required for review. MEDRAD STELLANT SN: 60764296 207916 | 4 |  | It is validated with Ing Bianco, data is obtained from the failing injector, a visit from the injector supplier is required for review. | [12 empty] | 117518232 |  | 885667 |  | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Injector |  |  | Completed | [5 empty] | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
855 |  |  | China | JS No.2 Chinese Medicine Hosp. | 345027 | 44317 |  | 02N/Ingenuity Core128/Elite China system freeze |  | CFM |  | Split from PR 885945 | [14 empty] | 885945 |  | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Split |  |  | Completed | [5 empty] | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
856 |  |  | China | SD Jinan No.4 Peo. Hosp. | 40038 | 44325 |  | 02N/Ingenuity CT 728326/yellow screen | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity CT 728326/ T2 Activities : Problem Description: 客户问题概述： 计算机软件相关故障。 客户问题详细描述或故障现象： 工程师来电开单。 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 已和工程师确认故障。 远程诊断与测试结果： 重启无效。 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议做workaruond升级。 建议备件12NC及名称： N/A Safety Questions : Case: 0117518647 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Bo Last updated by: Wang Bo 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 10.05.2021 01:10:40 UTC Tao Chen (86038373)故障现象: CT host主机黄屏 服务内容: 运行workaround方案 服务结果: 维修后,设备恢复正常. Internal Remarks : 10.05.2021 01:17:36 UTC Tao Chen (86038373) “RDI Accuracy: YES" $Q7: (Confirmed reported problem, CT host yellow screen) $Q8: ( Software bug of CT 5000 host) $Q9: ( Do workaround for CT host) $Q10: ( Test and inspection data refer to: PA. Device returned to full functionality) $Q11: Calibrated Tools name ( N/A) ^Current Tube installed date: N/A ^Current Tube usage: N/A ----------------------------------------------------------------------------- | 3 |  | Follow up case 0116861304 to implement workaround proactively | [12 empty] | 117518647 |  | 885874 |  | Closed - Duplicate | Closed - Duplicate | [14 empty] | Yellow Screen Workaround | [8 empty] | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
857 |  |  | China | SD Chest Hosp. | 40025 | 44324 |  | 02N/Ingenuity CT 728326/ yellow screen | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity CT 728326 Next Step Action : Huang Zhaokun call in SAP Transfer : Case Transferred to SAP External Remarks : 08.05.2021 11:12:10 UTC SUP JCOUSER (86038405)服务内容：执行CT5000黄屏workround。 服务结果：维修后，设备正常工作。 Internal Remarks : 10.05.2021 11:45:43 UTC Zhaokun Huang (86038405) "RDI Accuracy: NO." $Q7: ( Confirmed reported problem & When End exam , Yellow Screen.) $Q8: ( Execute Yellow Screen Workaround from BIU & Win10 BUG.) $Q9: ( Execute Yellow Screen Workaround from BIU.) $Q10: ( Test and inspection data refer to: PA . Device returned to full functionality.) $Q11: Calibrated Tools name (N/A) ^Current Tube installed date: N/A ^Current Tube usage: N/A Safety Questions : Case: 0117517816 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Jianqiang Last updated by: Wang Jianqiang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 其他类型故障。 客户问题详细描述或故障现象： host主机黄屏 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： host主机黄屏 远程诊断与测试结果： host主机黄屏 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 根据技术支持建议修改文件 SAP Transfer : Dispatched for Onsite Support ----------------------------------------------------------------------------- | 3 |  | Follow up case 0117366742 to implement workaround proactively | [12 empty] | 117517816 |  | 885388 |  | Closed - Duplicate | Closed - Duplicate | [14 empty] | Yellow Screen Workaround | [8 empty] | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
858 |  |  | China | NMG Xinganmeng Hos. | 345030 | 44324 |  | 02N/Ingenuity Core128/Elite China/consultation | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: Instructed Customer Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ FSE Notes : This PO is a Reference Number only Safety Questions : Case: 0117517783 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wennian Han Last updated by: Wennian Han 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: consultation Troubleshooting Action: customer want to know the max KV and MA value Repair Action: answer customer Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: NA SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | NAC? | [12 empty] | 117517783 |  | 885377 |  | Closed - Non-Complaint | Closed - Non-Complaint | [14 empty] | Consultation | [8 empty] | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
859 |  |  | China | HEN Puyang People_s Hosp | 40031 | 44324 |  | 02N/Ingenuity CT 728326/estop open | BU Q&R Feedback: No Reason: Corrective Maintenance SAP Transfer : Case Transferred to SAP Problem Reported by Customer : 02N/Ingenuity CT 728326/ T2 Activities : Problem Description: 客户问题概述： ESTOP断开。 客户问题详细描述或故障现象： 机架无法启动。 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： lifesync P1=8 ,wait_for_acs_init 远程诊断与测试结果： 重启无效。 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议检查AMC启动状态，检查AMC到CCC通讯I/O连线，必要时更换AMC排除故障。 建议备件12NC及名称： PPD 459800712461 External Remarks : 10.05.2021 12:54:28 UTC Yongping Luo (86038950)服务内容： 水平移动床，设备急停，弹出钥匙框。 服务结果：重新校准水平位置，设备恢复正常。 其它重要事项：无 Internal Remarks : 10.05.2021 12:54:30 UTC Yongping Luo (86038950) $Q7:Confirmed report problem，ESTOP Open. $Q8:Perform visiual check on site ,when Move the table the Estop open, recalibration the couch horizontal position,system return full function. $Q9:recalibration the couch horizontal position ,Device return to full function. $Q10: test and inspection data refer to PA. Device returned to full functionality. $Q11:Calibrated Tools refer to PA. Safety Questions : Case: 0117517194 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Bo Last updated by: Wang Bo 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Dispatched for Onsite Support Next Step Action : PPD 459800712461*1 24H FSE:罗永平-86-15226131133 已安排出货 ----------------------------------------------------------------------------- | 3 |  | 459800712461 FRU, AMC upgrade KITS: check if replace the AMC on field? Parts was good returned @Zhou, Robin@ZHANG, Chang Xian please support investigation and check if this is known issue? CAN trace and logger saved: https://teams.microsoft.com/_#/files/%E5%B8%B8%E8%A7%84?threadId=19%3A76cc42a0c8fa4bd7877f7de12141f84c%40thread.tacv2&ctx=channel&context=Couch%2520E-Stop_8_May_40031&rootfolder=%252Fsites%252FBigBoreRTRefreshFOKCoreTeam%252FShared%2520Documents%252FGeneral%252FIngenuity%2520Bugreps%252FChina%2520KM%252FCouch%2520E-Stop_8_May_40031 | [12 empty] | 117517194 |  | 885116 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD_CouchEstop | [5 empty] | Engineering | [6 empty] | Horizontal calibration |  |  | Completed | [5 empty] | System issue | System | Couch-SW | TBD_CouchEstop | Couch-SW-TBD_CouchEstop | Couch-SW-TBD_CouchEstop | TBD_CouchEstop | TBD_CouchEstop | Couch-SW | Couch-SW-TBD_CouchEstop | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
860 |  |  | China | SD Laiyang People Hospital | 345005 | 44324 |  | 02N/Ingenuity Core128/Elite China | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity Core128/Elite China SAP Transfer : Dispatched for Onsite Support SAP Transfer : Case Transferred to SAP T2 Activities : Problem Description: 客户问题概述： 图像伪影。 客户问题详细描述或故障现象： 图像每隔几层会有弧形伪影 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 图像弧形伪影 远程诊断与测试结果： 怀疑有坏通道 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 检查屏蔽坏通道，如有必要进行更换模块 建议备件12NC及名称： 459800911281 V2 探测器模块 Safety Questions : Case: 0117517173 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Jianqiang Last updated by: Wang Jianqiang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 3 |  | 459800911281 TDMS V2 M.Det Class B FRU 1 24H已出货 Check if replace DMS on field? 图像有伪影，检查发现第4块module有坏通道，屏蔽无效，更换module，第40块有坏通道屏蔽恢复。 Further step: engineering inputs for service manual update to deliver the IQ troubleshooting knowledge on Refresh version based on DMS team investigation on the failure parts return for analysis. | [12 empty] | 117517173 |  | 885101 |  | Closed - Done | Closed - Done |  | CTUDT00632110 | [5 empty] | Engineering | [6 empty] | DMS Module |  |  | Completed | [5 empty] | System issue | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | DMS Module | DMS-HW | DMS-HW-DMS Module | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
861 |  |  | China | SHX XIANYANG CENTRAL HOSP. | 40050 | 44323 |  | 02G/Ingenuity CT 728326/transfer failed | BU Q&R Feedback: No SAP Customer Symptom Code: HANG Customer Symptom Code: Hang/Crash Reason: Corrective Maintenance Engineer Failure Code: Configuration Error Repair Activity Code: Reboot/Restart T2 Activities: Problem Description: transfer failed Troubleshooting Action: reboot gantry and software Repair Action: reboot gantry and software Test/Inspection: customer confirm can scan patient and image ok.functional Check PASSED. The device returned to full functionality. Test Equipment: N/A Problem Reported by Customer: 02G/Ingenuity CT 728326 Safety Questions:Case: 0117511285 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Bo Last updated by: Wang Bo 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | 44323 | [12 empty] | 117511285 |  | 883130 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | Image Transfer |  |  | Completed | [5 empty] | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
862 |  |  | China | JX Province Children Hosp. | 40041 | 44323 |  | application/02N/Ingenuity CT 728326 | BU Q&R Feedback: No Reason: Customer Information Safety Questions:Case: 0117511124 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Qian Wang Last updated by: Qian Wang 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Problem Reported by Customer: 02N/Ingenuity CT 728326 Next Step Action: Chen Xianfeng 13971190179 ----------------------------------------------------------------------------- | 5 |  | NAC | [12 empty] | 117511124 |  | 883040 |  | Closed - Non-Complaint | Closed - Non-Complaint | [14 empty] | Application | [8 empty] | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
863 |  |  | China | HEN Puyang People_s Hosp | 40031 | 44323 |  | 02n/Ingenuity CT 728326/couch fault | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02n/Ingenuity CT 728326/ SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : PPD FRU, CBL, HANDLE TAPE SWITCH 459800623252 2 已安排24H发货，FSE罗永平-86-15226131133 External Remarks : 09.05.2021 13:50:50 UTC Yongping Luo (86038950)服务内容： 检查床水平按键不移动，只能手动移动，查看报错S HMC TAPE SWITCH PRESSED 。 服务结果：更换TAPE Switch后，设备恢复正常。 其它重要事项：无 Internal Remarks : 09.05.2021 13:50:53 UTC Yongping Luo (86038950) $Q7:Confirmed report problem，the couch cannot movement by press the button. $Q8:Perform visiual check on site ,find the Error code S HMC TAPE SWITCH PRESSED , replace the Tape switch,Device return to full function . $Q9:replace the Tape switch ,Device return to full function. $Q10: test and inspection data refer to PA. Device returned to full functionality. $Q11:Calibrated Tools refer to PA. Safety Questions : Case: 0117510466 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wennian Han Last updated by: Wennian Han 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 床故障 客户问题详细描述或故障现象： 床故障 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： S HMC TAPE SWITCH PRESSED 远程诊断与测试结果： 床水平按键不动，水平灯是亮的，有压带开关错 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 检查压带开关相关 建议备件12NC及名称： " 459800623252 FRU, CBL, HANDLE TAPE SWITCH" ----------------------------------------------------------------------------- | 3 |  | 459800623252 FRU, CBL, HANDLE TAPE SWITCH | [12 empty] | 117510466 |  | 882684 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | Tape Switch |  |  | Completed | [5 empty] | System issue | System | Couch-HW | TabSwitch | Couch-HW-TabSwitch | Couch-HW-TabSwitch | CTUDT00632127 | TabSwitch | Couch-HW | Couch-HW-TabSwitch | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
864 |  |  | China | SHX Ankang T.C.M.Hosp. | 345041 | 44322 |  | 02N/Ingenuity Core128/Elite China/Image quality issue | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ Safety Questions:Case: 0117506321 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingsong Last updated by: Lu Mingsong 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer_s problems: Image quality problem Detailed description of Customer_s problems or malfunctions: Image quality problem It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: N/A Remote troubleshoot and test result: N/A PRS TR status: It is not through, ask engineer for maintenance on site. Repair Action: Suggest onsite service procedures: Carry out air calibration and module calibration. We recommend spare parts 12NC and names: N/A Memo of Customer request: The customer asked FSE to be on the field. SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support ----------------------------------------------------------------------------- | 3 |  | 服务内容： 图像质量差，重新做空气矫正，H值校正，phantom校正，调整扫描参数。 服务结果：设备恢复正常。 | [12 empty] | 117506321 |  | 879228 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | ShouldbeCustomer | [5 empty] | Engineering | [6 empty] | IQ |  |  | Completed | [5 empty] | System issue | System | NotEngineering | ShouldbeCustomer | NotEngineering-ShouldbeCustomer | NotEngineering-ShouldbeCustomer | ShouldbeCustomer | NotEngineering-ShouldbeCustomer | NotEngineering | NotEngineering-ShouldbeCustomer | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
865 |  |  | China | JX Yugan Chudong hospital | 337500 | 44322 |  | wechat:02n/Ingenuity Core128/Elite China/Bed drop button discontinuous failure | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: Bed drop button discontinuous failure Safety Questions:Case: 0117503991 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Lili Last updated by: Lu Lili 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer_s problems: CTBOX key of downward failure Detailed description of Customer_s problems or malfunctions: CTBOX key of downward failure It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: NA Remote troubleshoot and test result: The customer said ctbox downward key was pressed, but did not respond. The downward key on gantry panel can work normally. The machine is working normally right now.Consider CTBOX itself.12NC on the CTBOX label sent by the customer cannot be found in espf, nor can be found in the system by parts group. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Check CT BOX, check if operation is correct etc.CFM is recommended for related problems. We recommend spare parts 12NC and names: " 459801705022 CT BOX" SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Next Step Action: PPD 459801705022 1 out of stock and has been ordered. FSE xie sheng -86-13970873443 ----------------------------------------------------------------------------- | 5 |  | 459801705022 CT BOX Wait failure part return to Suzhou for analysis | [12 empty] | 117503991 |  | 877637 |  | Investigation | Closed - Done |  | CTUDT00632185 | [5 empty] | Engineering | [6 empty] | CT Box |  |  | Completed | [5 empty] | System issue | System | Gantry-HW | CT Box | Gantry-HW-CT Box | Gantry-HW-CT Box | CTUDT00632185 | CT Box | Gantry-HW | Gantry-HW-CT Box | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
866 |  |  | Japan | Maki Hospital | 341017 | 44326 |  | This is an enhancement request for Quick viewer or CT viewer on console. |  | CFM | [16 empty] | 885927 |  | Evaluation | Closed - non-complaint | [7 empty] | Customer | [6 empty] | RFC |  |  | Completed | [5 empty] | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
867 |  |  | Japan | Maki Hospital | 341017 | 44326 |  | This is an enhancement request for CT viewer on console. |  | CFM | [16 empty] | 885918 |  | Evaluation | Closed - non-complaint | [7 empty] | Customer | [6 empty] | RFC |  |  | Completed | [5 empty] | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
868 |  |  | Japan | Maki Hospital | 341017 | 44324 |  | 01CS: KK: When I create a MIP in a series that shot the head Axial, Slice is divided and recognized. | BU Q&R Feedback: No Reason: Application Support Problem Reported by Customer : 01CS:KK:When a MIP is created in a series of photographs of the head Axial, Slice is divided and recognized. ----------------------------------------------------------------------------- | 4 |  | Without sufficient information | [12 empty] | 117517635 |  | 885285 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632610 | [5 empty] | Engineering | [6 empty] | MIP |  |  | Completed | [5 empty] | System issue | System | Console-SW | MIP | Console-SW-MIP | Console-SW-MIP | CTUDT00632610 | EC Upgrade |  |  | No | y | y | y | y | y | No | y | 0 | y | y | y | y | y | y | 0 | y | y
869 |  |  | Poland | Agencja Rezerw Materiałowych | 341010 | 44323 |  | DICOM configuration. | Work Order #: WO-03460167 07-May-2021 Mr. Robert Kamola The system fully meets specification and returned to use Problem Description by Engineer: DICOM configuration. Resolution: DICOM configuration of the new CT scanner for RIS/PACS and ISP descriptive stations. Testing network connections. Device is operational. ----------------------------------------------------------------------------- | 4 |  | NAC | [12 empty] | 117513096 |  | 882301 |  | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Service | [6 empty] | System Configuration |  |  | Completed | [5 empty] | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
870 |  |  | United States | The Regional Medical Center | 40005 | 44326 |  | Need to power on the CT system | Work Order #: WO-03467911 N, Manjunatha The system fully meets specification and returned to use Diagnostic performed by Engineer : Tara came is to find the gantry is switched off. She has tried to power cycle at the mains breaker. Walked though on how to create bugreport. Logs shows issue on the gantry tilt. Problem Description by Engineer : Problem description by engineer : Cannot close Estop / Gantry is off When was the first occurrence? Within... : 10-May-2021 How often is issue occurring? : Continuously Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: Not provided Patient Impact/Outcome: None User Impact: Cannot start up gantry Current Software Version: 4.1.7 Resolution : The call was transferred into wrong queue as we support ISP product hence informed customer that will create case and will inform CT technician to call back for providing further support on this regards, Follow Up Required : What are the next Troubleshooting steps : Read Pot1 and Pot2 in service mode. If one is more than 1000, check the connector of that Pot. Otherwise, recalibrate Pot1 and Pot2 at 0 degrees with a value of 512 +/- 10. Error # and/or description of error : 2021/05/10 01:00:33.721 S_TILT_C_POT_CALIB_KO ERROR_FATAL EStop Initializing Logger.mdb Checked The Calibration between POT1 and POT2 is too big, or the POT connector is OFF Completed or Recommended Repair Action "Internal" : Read Pot1 and Pot2 in service mode. If one is more than 1000, check the connector of that Pot. Otherwise, recalibrate Pot1 and Pot2 at 0 degrees with a value of 512 +/- 10. Internal Comments : Error # and/or description of error : 2021/05/10 01:00:33.721 S_TILT_C_POT_CALIB_KO ERROR_FATAL EStop Initializing Logger.mdb Checked The Calibration between POT1 and POT2 is too big, or the POT connector is OFF Completed or Recommended Repair Action "Internal" : Read Pot1 and Pot2 in service mode. If one is more than 1000, check the connector of that Pot. Otherwise, recalibrate Pot1 and Pot2 at 0 degrees with a value of 512 +/- 10. ----------------------------------------------------------------------------- Work Order #: WO-03468683 10-May-2021 Holt Douglas Problem Description by Engineer : Problem description by engineer : Cannot close Estop / Gantry is off When was the first occurrence? Within... : 10-May-2021 How often is issue occurring? : Continuously Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: Not provided Patient Impact/Outcome: None User Impact: Cannot start up gantry Current Software Version: 4.1.7 Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- Work Order #: WO-03468666 Ali, Adam The system does not meet specification and is out of use Problem Description by Engineer : Problem description by engineer : Cannot close Estop / Gantry is off When was the first occurrence? Within... : 10-May-2021 How often is issue occurring? : Continuously Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: Not provided Patient Impact/Outcome: None User Impact: Cannot start up gantry Current Software Version: 4.1.7 Follow Up Required : What are the next Troubleshooting steps : Read Pot1 and Pot2 in service mode. If one is more than 1000, check the connector of that Pot. Otherwise, recalibrate Pot1 and Pot2 at 0 degrees with a value of 512 +/- 10. Error # and/or description of error : 2021/05/10 01:00:33.721 S_TILT_C_POT_CALIB_KO ERROR_FATAL EStop Initializing Logger.mdb Checked The Calibration between POT1 and POT2 is too big, or the POT connector is OFF Completed or Recommended Repair Action "Internal" : Read Pot1 and Pot2 in service mode. If one is more than 1000, check the connector of that Pot. Otherwise, recalibrate Pot1 and Pot2 at 0 degrees with a value of 512 +/- 10. Internal Comments : Error # and/or description of error : 2021/05/10 01:00:33.721 S_TILT_C_POT_CALIB_KO ERROR_FATAL EStop Initializing Logger.mdb Checked The Calibration between POT1 and POT2 is too big, or the POT connector is OFF Completed or Recommended Repair Action "Internal" : Read Pot1 and Pot2 in service mode. If one is more than 1000, check the connector of that Pot. Otherwise, recalibrate Pot1 and Pot2 at 0 degrees with a value of 512 +/- 10. Diagnostic performed by Engineer : Tara came is to find the gantry is switched off. She has tried to power cycle at the mains breaker. Walked though on how to create bugreport. Logs shows issue on the gantry tilt. ----------------------------------------------------------------------------- | 3 |  | Reseated connections in the e stop circuit | [12 empty] | 117520499 |  | 885976 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | ECR-077012 | [5 empty] | Engineering | [6 empty] | Dual pot |  |  | Completed | [5 empty] | System issue | System | Gantry-HW | Belt | Gantry-HW-Belt | Gantry-HW-Belt | ECR-077012 | Belt | Gantry-HW | Gantry-HW-Belt | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
871 |  |  | United States | St Barnabas Hospital | 40028 | 44325 |  | 0711CFXXX Software question | Work Order #: WO-03466761 Boyd Craig The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: Software question User Impact: None Patient Impact: None Current Software Version: Resolution : Instructed tech on manual reconstructions. Images good. Diagnostic performed by Engineer : Tech needed to make some manual reconstructions. ----------------------------------------------------------------------------- | 3 |  | NAC | [12 empty] | 117519067 |  | 885610 |  | Closed - Non-Complaint | Closed - Non-Complaint | [14 empty] | Guidance | [8 empty] | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
872 |  |  | United States | St Barnabas Hospital | 40028 | 44323 |  | recon thin slices | Work Order #: WO-03462463 Curry, Dyan The system fully meets specification and returned to use Problem description by Clinical Support Specialist : Customer Function/ Role: technologist Device Use: in use name Expected and Unexpected Results: recon thins User Impact: none Patient Impact: none Current Software version: 4.1.7 Diagnostic performed by Engineer : helped customer recon thin images in IRS and save to local. ----------------------------------------------------------------------------- | 4 |  | helped customer recon thin images in IRS and save to local. NAC | [12 empty] | 117514938 |  | 880838 |  | Closed - Non-Complaint | Closed - Non-Complaint | [14 empty] | Guidance | [8 empty] | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
873 |  |  | United States | Unity Health Systems | 40037 | 44322 |  | 0711CFXXX system intermittently freezing when using perfusion exam card | Work Order #: WO-03456966 Swain Wesley The system fully meets specification and returned to use Diagnostic performed by Engineer : Download and analyze bugreport. Consulted Dyan Curry from ClinApps Problem Description by Engineer : Customer Function/Role: CT User Device Use: Rapid Perfusion Exam Card Expected and Unexpected Results: System freezing up intermittently User Impact: Delay Patient Impact: Delay Current Software Version: 4.1.10 Resolution : Dyan Curry from ClinApps informed that this is a known issue with rapid protocol. Dyan provided workaround as follows. Scan Surview, Plan 2nd perfusion study, Plan 1st perfusion study. Scan normally. Angela will integrate this into their workflow. | 3 |  | Dyan provided workaround as follows. Scan Surview, Plan 2nd perfusion study, Plan 1st perfusion study. Scan normally. Angela will integrate this into their workflow. Is it a known issue with defect tracking? | [12 empty] | 117509271 |  | 876664 |  | Investigation | Closed - Done |  | TBD_Bugreport | [5 empty] | Engineering | [6 empty] | Freezing |  |  | Completed | [5 empty] | System issue | System | Console-SW | TBD-bugrep | Console-SW-TBD-bugrep | Console-SW-TBD-bugrep | TBD_Bugreport | TBD_Bugreport |  |  | No | y | y | y | y | y | No | y | 0 | y | y | y | y | y | y | 0 | y | y
874 |  |  | United States | Leesburg Regional Medical Cent | 40022 | 44322 |  | 0308CDY72 Noise during table vertical motion | Work Order #: WO-03454978 Swain Wesley The system partly meets specification and is returned to use with restrictions Follow Up Required : inspect covers for rubbing. inspect / lubricate vertical guide rails. inspect lubricate ball screw and thrust bearings. Unable to provide part numbers for replacements as PRS will not allow telnet session and was unable to determine couch type. Diagnostic performed by Engineer : rubbing noise when operating table vertical drive Problem Description by Engineer : Customer Function/Role: CT User Device Use: table vertical motion Expected and Unexpected Results: noisy vertical drive User Impact: none Patient Impact: none Current Software Version: n/a ----------------------------------------------------------------------------- | 3 |  | cleaned of old grease from verticle scissors tubes and relubricated tested and no more noise, followed up next day still good | [12 empty] | 117507740 |  | 875081 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00631545/CTUDT00630971 | [5 empty] | Engineering | [6 empty] | Couch Noise |  |  | Completed | [5 empty] | System issue | System | Couch-HW | Table Noise | Couch-HW-Table Noise | Couch-HW-Table Noise | CTUDT00631545/CTUDT00630971 | Table Noise | Couch-HW | Couch-HW-Table Noise | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
875 |  |  | United States | Clay County Medical Center | 40003 | 44320 |  | The customer has had multiple issues running the IQ check. It will tell them to move the phantom multiple times | Phantom detection issue | CFM |  | Bugreport received https://www.clouddrive.philips.com.cn/file/pkokusrr69689w09ew0y63k5dq7h7acj# Password:123 | [14 empty] | 873986 |  | Investigation | Closed - Done |  | CTUDT00631545/CTUDT00630971 | [5 empty] | Engineering | [6 empty] | Phantom detection |  |  | Completed | [5 empty] | System issue | System | DMS-SW | Phantom Detect | DMS-SW-Phantom Detect | DMS-SW-Phantom Detect | CTUDT00631545/CTUDT00630971 | Phantom Detect | DMS-SW | DMS-SW-Phantom Detect | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
876 |  |  | United States | Bibb Medical Center | 310175 | 44323 |  | Q/O: Image Artifact | Work Order #: WO-03462944 10-May-2021 Foster Bobby The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : checked compensator, cleaned lexan ring and detector cover | 5 |  | Resolution Description: checked compensator, cleaned lexan ring and detector cover | [12 empty] | 117515516 |  | 884589 |  | Closed - Done | Closed - Done |  | ECR-075975 | [5 empty] | Engineering | [6 empty] | Artifact |  |  | Complete |  |  | Refresh |  |  | Upgrade issue | System | Gantry-HW | Lexan Ring | Gantry-HW-Lexan Ring | Gantry-HW-Lexan Ring | ECR-075975 | Lexan Ring | Gantry-HW | Gantry-HW-Lexan Ring | No | y | y | y | y | y | No | y | 0 | y | y | y | y | y | y | 0 | y | y
877 |  |  | United States | Advent Health Wesley Chapel | 300099 | 44324 | 18 | 0705CDN84 customer locked out of system, multiple issues | Work Order #: WO-03465132 08-May-2021 Stoutt Chad The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Problem description by engineer : Customer stated they have had nothing but problems since the upgrade. She said there are " a Slew" of bug reports on the system. The keyboard has been intermittent and now they are completely locked out of the system. When was the first occurrence? Within... : 07-May-2021 How often is issue occurring? : Multiple times a day Malfunction area : Others Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: customer locked out of system, multiple issues Patient Impact/Outcome: moved to another scanner User Impact: system down Current Software Version: 4.1.7.25030 Diagnostic performed by Engineer : Onsite 5/8/2021 to diagnose problem. System was in use upon arrival. Customer indicated that a few buttons on keyboard were not working. Keyboard was replaced with part ordered by PFSE. Customer also indicated they are experiencing intermittent blank/black screens (both) after a surview or scan. Customer then has to force reboot on Host PC and perform offline recons. Suspected Host PC video card to be the cause of the issue. Follow Up Required : Replace host PC video card ----------------------------------------------------------------------------- Work Order #: WO-03465120 Mcginnis, Mike The system does not meet specification and is out of use Problem Description by Engineer : Problem description by engineer : Customer stated they have had nothing but problems since the upgrade. She said there are " a Slew" of bug reports on the system. The keyboard has been intermittent and now they are completely locked out of the system. When was the first occurrence? Within... : 07-May-2021 How often is issue occurring? : Multiple times a day Malfunction area : Others Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: customer locked out of system, multiple issues Patient Impact/Outcome: moved to another scanner User Impact: system down Current Software Version: 4.1.7.25030 Diagnostic performed by Engineer : Spoke to Diana. She stated they have had nothing but problems since the upgrade. She said there are " a Slew" of bug reports on the system. The keyboard has been intermittent and now they are completely locked out of the system. I am getting errors when I attempt to remote to the system to review logs. Internal Comments : Error # and/or description of error : customer locked out of system, multiple issues Completed or Recommended Repair Action "Internal" : Reset the password, according to the customer you may have to install a different keyboard. She said one was supposed to be there today but they haven_t seen it so one is on the way. Once you are back into the system review the log files and troubleshoot per your findings. Could not connect to the system for review of the logs remotely Follow Up Required : Error # and/or description of error : customer locked out of system, multiple issues Completed or Recommended Repair Action "Internal" : Reset the password, according to the customer you may have to install a different keyboard. She said one was supposed to be there today but they haven_t seen it so one is on the way. Once you are back into the system review the log files and troubleshoot per your findings. Could not connect to the system for review of the logs remotely ----------------------------------------------------------------------------- Work Order #: WO-03466034 10-May-2021 Sher Mark The system fully meets specification and returned to use Problem Description by Engineer : Problem description by engineer : Customer stated they have had nothing but problems since the upgrade. She said there are " a Slew" of bug reports on the system. The keyboard has been intermittent and now they are completely locked out of the system. When was the first occurrence? Within... : 07-May-2021 How often is issue occurring? : Multiple times a day Malfunction area : Others Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: customer locked out of system, multiple issues Patient Impact/Outcome: moved to another scanner User Impact: system down Current Software Version: 4.1.7.25030 Follow Up Required : outstanding NFO part to be returned.. I have in possession Diagnostic performed by Engineer : bypassed computer cabinet back interface and ran directly to host for the monitor video to correct video errors ----------------------------------------------------------------------------- Work Order #: WO-03471813 08-May-2021 Watson Adam Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 2 |  | Part#:459801536801 Description: NVIDIA Quadro P400 2GB (3)mDP GFX Quantity: 1 | [12 empty] | 117517087 |  | 882762 |  | Closed - Non-Complaint | Closed - Non-Complaint | [14 empty] | Black screen | [5 empty] | Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
878 |  |  | United States | Advent Health Wesley Chapel | 300099 | 44322 |  | Surview Scan Plan -Exam Aborts | Work Order #: WO-03456576 Carpenter, Kevin The system partly meets specification and is returned to use with restrictions Follow Up Required : Philips Onsite Clin Apps ( Mellissa ) relays on going issues with CTHost : Since Wni10 upgrade : During Plan Scan on Surview - Exam Aborts 12:08pm Monitors go blank ( sometimes just one ) 12:58pm Keyboard - control key / num lock not working - then letter "B" She shook keyboard and letter _ B _ began working again. Cust says service has already worked on this and CTHost may have already been replaced once. FSE while bad keyboard likely would not cause monitors to go blank ~ a bad keyboard has been know to cause bizarre problems. Please consult NSS, but in mean time please do replace keyboard. CT RSE is ordering one to site. Diagnostic performed by Engineer : Reviewed logs - ordered keyboard to site ( per onsite Philips Clin Apps ) ~ Advised FSE to consult with NSS as this is a repeat problem. Diagnostic performed by Engineer : Reviewed CTHost EVT logs from very large bugrep ( took a long time to d/l ) ~ Forwarded same to FSE assigned to onsite, pFSE and regional NSS. Problem Description by Engineer : Customer Function/Role: Philips Onsite Clin Apps ( Mellissa ) Device Use: Diag Expected and Unexpected Results: CTHost Lockups and odd behavior User Impact: None Patient Impact: None Current Software Version: IDK ----------------------------------------------------------------------------- Work Order #: WO-03456629 10-May-2021 Sher Mark The system fully meets specification and returned to use Diagnostic performed by Engineer : worked with Adam on differend wo to correct vlicker issues Problem Description by Engineer : Customer Function/Role: Philips Onsite Clin Apps ( Mellissa ) Device Use: Diag Expected and Unexpected Results: CTHost Lockups and odd behavior User Impact: None Patient Impact: None Current Software Version: IDK ----------------------------------------------------------------------------- | 3 |  | Parts Consumed Description: 1.00000,459800918431,Cherry USB keyboard English Silver | [12 empty] | 117508913 |  | 876293 |  | Investigation | Investigation |  | NotEngineering-partreplace | [5 empty] | Engineering | [6 empty] | Keyboard |  |  | Complete |  |  | Refresh |  |  | Upgrade issue | System | Console-HW | Keyboard | Console-HW-Keyboard | Console-HW-Keyboard | NotEngineering-partreplace | Keyboard | Console-HW | Console-HW-Keyboard | No | y | y | y | y | y | No | y | 0 | y | y | y | y | y | y | 0 | y | y
879 |  |  | Argentina | Hospital Dr. Bernardo Houssay | 40027 | 44326 |  | A CIRCLE IS DISPLAYED IN THE IMAGE THAT PREVENTS THE STUDIES FROM BEING CLEARLY SEEN. | Work Order #: WO-03476495 Alec, Preciado The system partly meets specification and is returned to use with restrictions Follow Up Required : What are the next Troubleshooting steps : Check the artifact, error log, checked raw data in the system, identify damaged detector and add it to the bad detector list, run corresponding calibrations to exclude it. Error # and/or description of error : ring artifact in images Completed or Recommended Repair Action "Internal" : It is confirmed with field service engineer Edgardo that the system showed a damaged detector last time in a module different from the one that was replaced, a visit on site is required to run tests on modules. Diagnostic performed by Engineer: We contact the user and the field service engineer, we identify that there is a detector in a module that generates an artifact, a visit to the site is required to check it. Problem Description by Engineer : Problem description by engineer : A CIRCLE IS DISPLAYED IN THE IMAGE THAT PREVENTS THE STUDIES FROM BEING CLEARLY SEEN. When was the first occurrence? Within... : 11-May-2021 09:00:00 How often is issue occurring? : Multiple times a day Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected results: System without artifacts Patient Impact/Outcome: Delay in attention User Impact: Delay in work load. Current Software Version: 4.1.7 Internal Comments : Error # and/or description of error : ring artifact in images Completed or Recommended Repair Action "Internal" : It is confirmed with field service engineer Edgardo that the system showed a damaged detector last time in a module different from the one that was replaced, a visit on site is required to run tests on modules. ----------------------------------------------------------------------------- Work Order #: WO-03478687 11-May-2021 Meza, Edgardo Arturo The system fully meets specification and returned to use. Problem Description by Engineer: A RING IS DISPLAYED IN THE IMAGE PREVENTING THE STUDIES FROM BEING CLEARLY SEEN. Resolution: Air calibration is performed and it is tested with a patient. Artifact is removed. Diagnostic Performed by Engineer: Calibration table is corrupt. ----------------------------------------------------------------------------- | 4 |  | Calibration table is corrupt | [12 empty] | 117523859 |  | 891772 |  | Closed - Done | Closed - Done |  | NotEngineering | [5 empty] | Engineering | [6 empty] | Artifact |  |  | Completed |  |  | Refresh |  |  | System issue | System | NotEngineering | ShouldbeCustomer | NotEngineering-ShouldbeCustomer | NotEngineering-ShouldbeCustomer | NotEngineering | NotEngineering-ShouldbeCustomer | NotEngineering | NotEngineering-ShouldbeCustomer | No | y | y | y | y | y | No | y | 0 | y | y | y | y | y | y | 0 | y | y
880 |  |  | Chile | HOSPITAL BARROS LUCO TRUDEAU | 341005 | 44326 |  | CL1688 - STUDIO SETTING PROBLEMS, TUBE TURNING PROBLEMS, DOES NOT START STUDIES AND HAS NOISE IN THE GANTRY | Work Order #: WO-03479792 Alec, Preciado The system fully meets specification and returned to use Diagnostic performed by Engineer : Se contacta a usuario turno de la tarde, indica que ha trabajado sin problema por 4 días, revisión de logs, se contacta a encargada del área, se identifica ruido en banda de rotor y rotor. Indican que la temperatura del área es de 20°C, aún así perciben problemas de desempeño de tubo. Problem Description by Engineer : Problem description by engineer : CL1688 - PROBLEMAS DE SETEO DE ESTUDIOS, PROBLEMAS CON GIRO DE TUBO, NO INICIA LOS ESTUDIOS Y TIENE RUIDO EN EL GANTRY When was the first occurrence? Within... : 11-May-2021 15:00:00 How often is issue occurring? : First occurrence Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Not provided Expected and Unexpected results: No further information is available as per the customer Patient Impact/Outcome: no impacta User Impact: no impacta Current Software Version: 4.1.10 Internal Comments : Error # and/or description of error : WARNINGS Optical slip ring signal error. S_DMC_SLIP_RING S_COUCHMG_MOTION_START_ERROR Completed or Recommended Repair Action "Internal" : Se identifican Warnings, no se identifican errores normales o fatales que indicaran que el sistema presenta problemas. Follow Up Required : What are the next Troubleshooting steps : Se requiere limpieza del punto de encuentro entre rotor y banda de motor, presenta ruido por suciedad o un desajuste, buscar por debris en la parte inferior de la banda. Revisar el Gap del tensor, de acuerdo a manual de reemplazo, así como revisión por desgaste o rotura en banda. Se requiere corroborar las condiciones de temperatura de la sala y realizar pruebas de desempeño al tubo, así como revisión de protocolos con usuario para identificar si el tubo presenta mensaje de calentamiento de tubo. Realizar calibración horizontal, se identifican algunos warnings relacionados con los limites y el avance de la carbon top. Error # and/or description of error : WARNINGS Optical slip ring signal error. S_DMC_SLIP_RING S_COUCHMG_MOTION_START_ERROR Completed or Recommended Repair Action "Internal" : Se identifican Warnings, no se identifican errores normales o fatales que indicaran que el sistema presenta problemas. ----------------------------------------------------------------------------- Work Order #: WO-03484019 Osorio, Juan Problem Description by Engineer : Problem description by engineer : CL1688 - PROBLEMAS DE SETEO DE ESTUDIOS, PROBLEMAS CON GIRO DE TUBO, NO INICIA LOS ESTUDIOS Y TIENE RUIDO EN EL GANTRY When was the first occurrence? Within... : 11-May-2021 15:00:00 How often is issue occurring? : First occurrence Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Not provided Expected and Unexpected results: No further information is available as per the customer Patient Impact/Outcome: no impacta User Impact: no impacta Current Software Version: 4.1.10 Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 3 |  | Cleaning of the meeting point between rotor and motor belt is required, it presents noise due to dirt or a mismatch, look for debris in the lower part of the belt. Check the Gap of the tensioner, according to the replacement manual, as well as check for wear or breakage in the belt. It is required to corroborate the temperature conditions of the room and perform performance tests on the tube, as well as a review of protocols with the user to identify if the tube presents a tube heating message. Perform horizontal calibration, some warnings related to the limits and the advance of the carbon top are identified. | [12 empty] | 117523046 |  | 891053 |  | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Multi issue |  |  | Completed |  |  | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
881 |  |  | China | ZJ YUEQING PEOPLE HOSPITAL | 40054 | 44331 |  | 02N/Ingenuity CT 728326 | BU Q&R Feedback: No SAP Customer Symptom Code: CENV Customer Symptom Code: Customer Environmental Concern Reason: Corrective Maintenance Engineer Failure Code: Environmental Cause Repair Activity Code: Reboot/Restart Problem Reported by Customer : 02N/Ingenuity CT 728326 Safety Questions : Case: 0117548137 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Weng Jinming Last updated by: Weng Jinming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Troubleshooting Action: instruct customer to restart the system Repair Action: instruct customer to restart the system Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | instruct customer to restart the system | [12 empty] | 117548137 |  | 907990 |  | Closed - Non-Complaint | Closed - Non-Complaint | [14 empty] | E-Stop | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
882 |  |  | China | Lindian County Hospital | 345038 | 44331 |  | 02N/Ingenuity Core128/Elite China/ECG signal Loss | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ Safety Questions : Case: 0117548523 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Weng Jinming Last updated by: Weng Jinming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action : PPD 459800166651*1 459801731761*1备件已系统下单，安排24H到场，FSE：倪刚-86-18686766464（替代号已确认） External Remarks : 17.05.2021 00:29:48 UTC Gang Ni (86038229)服务内容：连接ibox电缆 服务结果：设备恢复 T2 Activities : Problem Description: 客户问题概述： CT机操作台没有心电信号 客户问题详细描述或故障现象： 心电监护仪正常 设备本身技术问题： 是。 Troubleshooting Action: 远程诊断与测试结果： 怀疑心电电缆，I-box，驱动软件。。。问题 PRS TR状态： 全通。 Repair Action: 远程诊断与测试结果： 怀疑心电电缆，I-box，驱动软件。。。问题 PRS TR状态： 全通。 T2 Activities : Problem Description: 客户问题概述： CT机操作台没有心电信号 客户问题详细描述或故障现象： 心电监护仪正常 设备本身技术问题： 是。 Troubleshooting Action: 远程诊断与测试结果： 怀疑心电电缆，I-box，驱动软件。。。问题 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议工程师到场检查心电电缆，驱动软件，I-box，教客户使用飞利浦微信报修 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support ----------------------------------------------------------------------------- | 3 |  | Reconnect ibox cable | [12 empty] | 117548523 |  | 907958 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Ibox |  |  | Completed |  |  | Refresh |  |  | System issue | System | Gantry-HW | TBD-NoEnoughInfo | Gantry-HW-TBD-NoEnoughInfo | Gantry-HW-TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo | Gantry-HW | Gantry-HW-TBD-NoEnoughInfo | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
883 |  |  | China | XJ Yili County Friendship Hospital | 40048 | 44331 |  | 02N/Ingenuity CT 728326/The couch move stuck | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity CT 728326 Safety Questions : Case: 0117548506 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Weng Jinming Last updated by: Weng Jinming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support T2 Activities : Problem Description: 客户问题概述： 新装机，床水平运动卡顿 设备本身技术问题： 是。 Troubleshooting Action: 远程诊断与测试结果： 怀疑外壳，丝杠。。问题 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议工程师到场检查外壳，丝杠，教客户使用飞利浦微信报修 ----------------------------------------------------------------------------- | 3 |  | 床水平运动卡顿, 建议工程师到场检查外壳，丝杠 | [12 empty] | 117548506 |  | 907940 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | AssemblyAdjust | [5 empty] | Engineering | [6 empty] | Vertical Transducer |  |  | Completed |  |  | Refresh |  |  | System issue | System | Couch-HW | Table move stuck | Couch-HW-Table move stuck | Couch-HW-Table move stuck | AssemblyAdjust | Table move stuck | Couch-HW | Couch-HW-Table move stuck | No | y | y | y | y | y | No | y | 0 | y | y | y | y | y | y | 0 | y | y
884 |  |  | China | Tongling Boai Hospital Co LTD | 345034 | 44330 |  | 02N/Ingenuity Core128/TILE LOSE POWER | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Replace Part Problem Reported by Customer : 02N/Ingenuity Core128/ Safety Questions : Case: 0117543360 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Cao Last updated by: Liu Cao 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP T2 Activities : Problem Description: 客户问题概述： 机架打角度掉电，不能归零。 客户问题详细描述或故障现象： 机架打角度掉电，不能归零。 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： S_TILT_C_NO_MOVE S_ROTOR_MCO 远程诊断与测试结果： 检查pot示数，需要时进行更换； pot正常时检查倾斜电机液压杆两侧是否平衡，右侧电源板指示，接线等 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 1.检查pot示数，需要时进行更换； 2.pot正常时检查倾斜电机液压杆两侧是否平衡，右侧电源板指示，接线等 PPD 24H: 459800441121 DUAL POT ASSEMBLY 1 Test Equipment: N/A SAP Transfer : Dispatched for Onsite Support Next Step Action : PPD备件DUAL POT ASSEMBLY 459800441121 1已安排24H到场地，FSE：周凯-86-15056051758 ----------------------------------------------------------------------------- | 3 |  | 服务内容：打倾角后机架不能上电，mco旋转电机过温，处理温控接头。 服务结果：服务后设备恢复正常。 | [12 empty] | 117543360 |  | 906008 |  | Closed - Done | Closed - Done |  | NotEngineering | [5 empty] | Engineering | [6 empty] | Tilt |  |  | Completed |  |  | Refresh |  |  | System issue | System | NotEngineering | ShouldbeService | NotEngineering-ShouldbeService | NotEngineering-ShouldbeService | NotEngineering | NotEngineering-ShouldbeService | NotEngineering | NotEngineering-ShouldbeService | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
885 |  |  | China | SD Tai_an T.C.M. Hosp. No.2 Ho | 345037 | 44329 |  | 02N/Ingenuity Core128/Elite China/Replace Flat Spring | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 05R/Ingenuity Core128/Elite China Safety Questions : Case: 0117538657 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingsong Last updated by: Lu Mingsong 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | replace flat spring | [12 empty] | 117538657 |  | 902531 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632833 | [5 empty] | Engineering | [6 empty] | Flat Spring |  |  | Completed |  |  | Refresh |  |  | System issue | System | Gantry-HW | FlatSpring | Gantry-HW-FlatSpring | Gantry-HW-FlatSpring | CTUDT00632833 | FlatSpring | Gantry-HW | Gantry-HW-FlatSpring | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
886 |  |  | China | HLJ Harbin Senhai Hospital | 345007 | 44329 | 19 | 02N/Ingenuity Core128/Elite China/warm up error | T2 Activities : Problem Description: 客户问题概述： 预热失败。 客户问题详细描述或故障现象： 预热失败。 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 不能预热 Repair Action: 建议现场维修步骤： 根据后台报错做相关troubleshooting Safety Questions : Case: 0117538024 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 14.05.2021 13:56:57 UTC Haijun Du (86038210)"SUAT. Ready for Clinical Use local time" : "2021-05-14 18:00" 服务内容：更换阴极高压。 服务结果：维修后，设备恢复正常工作。 Internal Remarks : 14.05.2021 13:57:26 UTC Haijun Du (86038210) $Q7: (Confirmed reported problem &can not scan) $Q8: (check the CT log，S_RHOST_HVG_HSS_GENERATOR_NON_ARC_FAULT 1:1 Grid error) $Q9: (ok after replaced CPM) | 2 |  | $Q7: (Confirmed reported problem &can not scan) $Q8: (check the CT log，S_RHOST_HVG_HSS_GENERATOR_NON_ARC_FAULT 1:1 Grid error) $Q9: (ok after replaced CPM) 服务内容：更换阴极高压。 服务结果：维修后，设备恢复正常工作。 | [12 empty] | 117538024 |  | 902103 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-APM/CPM | [5 empty] | Engineering | [6 empty] | CPM |  |  | Completed |  |  | Refresh |  |  | System issue | System | IC-HW | APM/CPM | IC-HW-APM/CPM | IC-HW-APM/CPM | TBD-APM/CPM | APM/CPM | IC-HW | IC-HW-APM/CPM | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
887 |  |  | China | Wuhan Ren _ai Hospital, Hubei Province | 345004 | 44329 |  | 02N/Ingenuity Core128/Elite China/system can not log in | BU Q&R Feedback: No SAP Customer Symptom Code: BOOT Customer Symptom Code: Boot up/Power-up Problem Reason: Corrective Maintenance Engineer Failure Code: Incompatible Use Repair Activity Code: Instructed Customer Safety Questions : Case: 0117537938 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Gaoge Last updated by: Wang Gaoge 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Problem Reported by Customer : 02N/Ingenuity Core128/Elite China T2 Activities : Problem Description: 主机启动后登录CT账户自动注销，不能登录 Troubleshooting Action: 指导客户开机后等待三五分钟，让后台进程完全启动后再次登录，此时登录正常。 Repair Action: 指导客户等待几分钟让后台进程完成启动后再登录 Test/Inspection: N/A Test Equipment: N/A Next Step Action : Q6 = system can not log in and the system log out when input the user name and password SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | 指导客户开机后等待三五分钟，让后台进程完全启动后再次登录，此时登录正常。 | [12 empty] | 117537938 |  | 902014 |  | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Login |  |  | Completed |  |  | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
888 |  |  | China | ZJ YUEQING PEOPLE HOSPITAL | 40054 | 44328 |  | 02N/Ingenuity CT 728326change cover | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity CT 728326 Safety Questions : Case: 0117533931 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Tao Fengming Last updated by: Tao Fengming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 其他类型故障。 客户问题详细描述或故障现象： 更换外壳 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 更换外壳 远程诊断与测试结果： 需要更换外壳 Repair Action: 建议现场维修步骤： 更换外壳 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : huang yi Internal Remarks : 13.05.2021 03:05:19 UTC Yi Huang (86036249) RDI Accuracy:Yes $Q7: (Confirmed reported problem&system can not boot up) $Q8: (check errorlog found lifesync 8, no keyprompt, found CCC LED change from 22 to 00) $Q9: (replace CCC) $Q10: (Test and inspection data refer to: PA.Device returned to full functionality.) $Q11: Calibrated tools name(N/A) S/N (N/A) Next calibration date (N/A) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date:N/A ^Current Tube usage: second:N/A External Remarks : 13.05.2021 03:05:17 UTC SUP JCOUSER (86036249)服务内容：机架不能上电，lifesync8，床控制板状态异常，更换控制板。 服务结果：维修后，设备恢复正常工作。 其它重要事项：无 ----------------------------------------------------------------------------- | 3 |  | 服务内容：机架不能上电，lifesync8，床控制板状态异常，更换控制板。 服务结果：维修后，设备恢复正常工作。 | [12 empty] | 117533931 |  | 899292 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD_CouchEstop | [5 empty] | Engineering | [6 empty] | CCB |  |  | Completed |  |  | Refresh |  |  | System issue | System | Couch-SW | TBD-bugrep | Couch-SW-TBD-bugrep | Couch-SW-TBD-bugrep | TBD_CouchEstop | TBD_Bugreport | Couch-SW | Couch-SW-TBD-bugrep | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
889 |  |  | China | SHX XIANYANG CENTRAL HOSP. | 40050 | 44328 |  | 02G/Ingenuity CT 728326/transfer image failed | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer : 02G/Ingenuity CT 728326 T2 Activities : Problem Description: 客户问题概述： 图像传输故障 客户问题详细描述或故障现象： 图像传输故障 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： NA 远程诊断与测试结果： 工程师已经指导客户检查网络正常，节点信息这次，服务模式提示密码不对无法reset spooler PRS TR状态： 不通，请工程师到场检查。 Repair Action: 建议现场维修步骤： 需要到场检查图像传输相关，用工程师权限设置olevel账户密码 Safety Questions : Case: 0117533853 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wennian Han Last updated by: Wennian Han 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action : 设备暂未入保，在IA# 53341371上加operation，派工fse 刘军，以邮件IA Mgr批复 Next Step Action : Op-20已加 ----------------------------------------------------------------------------- | 3 |  | 需要到场检查图像传输相关，用工程师权限设置olevel账户密码, reset spooler | [12 empty] | 117533853 |  | 899231 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | Image Transfer |  |  | Completed |  |  | Refresh |  |  | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
890 |  |  | China | ZJ YUEQING PEOPLE HOSPITAL | 40054 | 44328 |  | 02N/Ingenuity CT 728326/replace shell | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: Instructed Customer Problem Reported by Customer : replace shell Safety Questions : Case: 0117532354 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Tao Fengming Last updated by: Tao Fengming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: need check option Troubleshooting Action: need customer check option Repair Action: guide customer check option Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | NAC | [12 empty] | 117532354 |  | 898301 |  |  | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
891 |  |  | China | SD Laiyang People Hospital | 345005 | 44328 |  | 02N/Ingenuity Core128/add node | BU Q&R Feedback: No Reason: Corrective Maintenance SAP Transfer : Dispatched for Onsite Support Problem Reported by Customer : 02N/Ingenuity Core128/ Safety Questions : Case: 0117532244 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingsong Last updated by: Lu Mingsong 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 添加节点 客户问题详细描述或故障现象： 添加节点 设备本身技术问题： 否。 Troubleshooting Action: 关键故障代码或现象： N/A 远程诊断与测试结果： olevel_service账号不能登录，无法使用lanconfig PRS TR状态： 不通，请工程师到场检查。 Repair Action: 建议现场维修步骤： 现场添加节点. 建议备件12NC及名称： N/A 客户需求备注： 无 SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 5 |  | olevel_service账号不能登录，无法使用lanconfig PRS TR状态： 不通，请工程师到场检查。 | [12 empty] | 117532244 |  | 898208 |  | Closed - Non-Complaint | Closed - Non-Complaint | [14 empty] | LAN Config | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
892 |  |  | China | ZJ YUEQING PEOPLE HOSPITAL | 40054 | 44327 | 19 | 02G/Ingenuity CT 728326/can not power on | T2 Activities:Problem Description: Summary of customer_s problems: Unable to turn the machine on. Detailed description of Customer_s problems or malfunctions: Unable to turn the machine on. It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: NA Remote troubleshoot and test result: The trainer was training on site to demonstrate the normal switch on and off. When turned on, he found that the machine could not be powered on. host has no key box. The customer had contacted the installation engineer. Under instruction, he had reboot several times but could not recover PRS TR status: Only M2M is not connected, ask FSE check onsite. Repair Action: Suggest onsite service procedures: Check customer operation to see if It is correct, check CAN communication related, such as main control board and power supply Memo of Customer request: The customer said they should have prepared for machine launch ceremony tomorrow. The situation is very urgent now and need an engineer to go to the site as soon as possible Next Step Action: Device is yet under warranty. Contacted MM. manager Yuan inform the site has yet been allocated to a FSE. Ask installation engineer to follow up first. from System query found IA#53515511. Dispatched ASP honglejia, mailed honglejia to follow up first ----------------------------------------------------------------------------- | 2 |  | ok after replace CCB | [12 empty] | 117529382 |  | 895802 |  | Evaluation | Closed - No Further Investigation |  | TBD_CouchEstop | [5 empty] | Engineering | [6 empty] | CCB |  |  | Completed |  |  | Refresh |  |  | System issue | System | Couch-SW | TBD-bugrep | Couch-SW-TBD-bugrep | Couch-SW-TBD-bugrep | TBD_CouchEstop | TBD_Bugreport | Couch-SW | Couch-SW-TBD-bugrep | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
893 |  |  | China | SD Laiyang People Hospital | 345005 | 44327 |  | 02N/Ingenuity Core128/add the workstation | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer: 01/Access CT Safety Questions:Case: 0117527641 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Weng Jinming Last updated by: Weng Jinming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Need add workstation Troubleshooting Action: FSE_admin cannot login with password Philips! 1/Philips!/Philips1 Repair Action: Customer will try to confirm password and call back Test/Inspection: N/A Test Equipment: N/A SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | NAC | [12 empty] | 117527641 |  | 894459 |  |  | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
894 |  |  | China | SD Anqiu T.C.M Hosp. | 345017 | 44327 |  | 02N/Core128/Rebuild the image failure | BU Q&R Feedback: No Reason: Corrective Maintenance Next Step Action : FSE Yang xingbo 013583235930 call in T2 Activities:Problem Description: Summary of customer_s problems: Occasionally there are no images after scanning. Detailed description of Customer_s problems or malfunctions: After scanning, there is raw data, but no reconstructed images. It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: N/A Remote troubleshoot and test result: N/A PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: The field service engineer asked to open escalation on a service order. We recommend spare parts 12NC and names: N/A Memo of Customer request: None Problem Reported by Customer: 02N/Core128/ Safety Questions:Case: 0117526548 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingsong Last updated by: Lu Mingsong 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support ----------------------------------------------------------------------------- | 5 |  | 扫描结束后有原始数据，没有重建出图像 工程师要求开单上ESCALATION | [12 empty] | 117526548 |  | 893672 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-NoEnoughInfo-TBD-NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Recon image |  |  | Completed |  |  | Refresh |  |  | System issue | System | TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
895 |  |  | China | NMG Zhengxiang Baiqi Hosp. | 345033 | 44327 |  | 02N/Ingenuity Core128/Elite China/gantry can not power on | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure SAP Transfer : Case Transferred to SAP Problem Reported by Customer : 02N/Ingenuity Core128/Elite China Safety Questions : Case: 0117525846 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingsong Last updated by: Lu Mingsong 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: gantry can not power on Troubleshooting Action: Remotely instruct the customer to restart the system. Repair Action: Remotely instruct the customer to restart the system. Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: N/A ----------------------------------------------------------------------------- | 3 |  | Remotely instruct the customer to restart the system. Test/Inspection: Functional Check PASSED. Device returned to full functionality. | [12 empty] | 117525846 |  | 893210 |  | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Reboot |  |  | Completed |  |  | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
896 |  |  | China | Guiyang Fuyou Hosp | 40033 | 44326 |  | 02N/Ingenuity CT 728326/Consult factory date | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer : Consult factory date Next Step Action : mail to IB 确认 Safety Questions : Case: 0117519675 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Qi Xiaoyu Last updated by: Qi Xiaoyu 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 3 |  | Consult factory date | [12 empty] | 117519675 |  | 888636 |  | Closed - Non-Complaint | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
897 |  |  | China | JL Songyuan Central Hosptial | 345008 | 44326 |  | Application/Ingenuity Core128 | BU Q&R Feedback: No Reason: Customer Information Field Service Engineer Notes:This PO is a Reference Number only Problem Reported by Customer : Application Safety Questions:Case: 0117519599 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Pan Chuchu Last updated by: Pan Chuchu 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action: Installation Order IA#52827232 dispatched ASP Huicong, Jilin Province Songyuan City Central Hospital_s Ingenuity Core128, the CT Room, Teacher Xia informed that training has not been done yet, now they need to apply for clinical training for the first time. Please ask Huicong to assist the customer with applying. CT Room, Teacher Xia 18904381395 ----------------------------------------------------------------------------- | 5 |  | apply for clinical training for the first time | [12 empty] | 117519599 |  | 888541 |  | Closed - Non-Complaint | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
898 |  |  | Japan | Maki Hospital | 341017 | 44331 |  | 07CS:TM:Quality assureranceのメニューに入るとWarningメッセージが表示される>>Tiltが0.5度かかっていた為でした | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Application Support Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer : Quality assureranceのメニューに入るとWarningメッセージが表示される>>Tiltが0.5度かかっていた為でした T2 Activities : Problem Description: Problem description by engineer : 対象装置:CT 故障個所:07その他 最初に発生したのはいつか?(~以内):1 Hour 問題発生の頻度は?:First Occurrence エンドユーザｰのワークフローのどの時点において問題が発生しているのか?:検査・測定準備中(Setting up Protocol) 最近どのようなイベントが発生したか?:N.A. 顧客は問題について具体的なシステムイベントレポートを作成しているか?:N.A. Engineer_s Comment: _============================== コール内容 Quality assureranceのメニューに入るとWarningメッセージが表示される _============================== Troubleshooting Action: What steps are taken during the system diagnosis : 対応後の結果:改善された 故障個所:07その他 診断中にどのようなステップを実施したか?: _================================== 対応 メッセージを読んでいただいたところTiltを0に戻してくださいというものでした。 Tiltが0.5度かかっていた為でしたので０にしていただきました _================================== Error # and/or description of error : NA Repair Action: Completed or Recommended Repair Action "External" : 【現象確認】 確認済み 【作業内容詳細】 質問内容を回答 【確認内容】 良好 OneEMS internal remarks : Problem description by engineer : 対象装置:CT 故障個所:07その他 最初に発生したのはいつか?(~以内):1 Hour 問題発生の頻度は?:First Occurrence エンドユーザｰのワークフローのどの時点において問題が発生しているのか?:検査・測定準備中(Setting up Protocol) 最近どのようなイベントが発生したか?:N.A. 顧客は問題について具体的なシステムイベントレポートを作成しているか?:N.A. Engineer_s Comment: _============================== コール内容 Quality assureranceのメニューに入るとWarningメッセージが表示される _============================== What steps are taken during the system diagnosis : 対応後の結果:改善された 故障個所:07その他 診断中にどのようなステップを実施したか?: _================================== 対応 メッセージを読んでいただいたところTiltを0に戻してくださいというものでした。 Tiltが0.5度かかっていた為でしたので０にしていただきました _================================== Error # and/or description of error : NA Completed or Recommended Repair Action "External" : 【現象確認】 確認済み 【作業内容詳細】 質問内容を回答 【確認内容】 良好 Completed or Recommended Repair Action "Internal" : --Troubleshooting Action-- 対応後の結果:改善された 故障個所:07その他 診断中にどのようなステップを実施したか?: _================================== 対応 メッセージを読んでいただいたところTiltを0に戻してくださいというものでした。 Tiltが0.5度かかっていた為でしたので０にしていただきました _================================== --Repair Action-- --Subject-- 07CS:TM:ﾎｼ様:Quality assureranceのメニューに入るとWarningメッセージが表示される>>Tiltが0.5度かかっていた為でした --Problem Description-- 対象装置:CT 故障個所:07その他 最初に発生したのはいつか?(~以内):1 Hour 問題発生の頻度は?:First Occurrence エンドユーザｰのワークフローのどの時点において問題が発生しているのか?:検査・測定準備中(Setting up Protocol) 最近どのようなイベントが発生したか?:N.A. 顧客は問題について具体的なシステムイベントレポートを作成しているか?:N.A. Engineer_s Comment: _============================== コール内容 Quality assureranceのメニューに入るとWarningメッセージが表示される _============================== --4/8Elements-- Q1 報告された問題は確認できましたか? 確認済み Q2 他の問題は確認できましたか? 無し Q3 問題を解決するために何をおこないましたか? 質問内容を回答 Q4 対応後の装置は順調ですか? 良好 Q5 その検査はどの様に終了しましたか? 無し Q6 問題により間違った所見（誤診、誤治療）の有無: 無し Q7 意図しない不要な曝射の有無: 無し Q8 アーチファクトの有無: 無し Safety Questions : Case: 0117548087 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: TAKEYUKI MIYAKAWA Last updated by: TAKEYUKI MIYAKAWA 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 5 |  | It was because Tilt took 0.5 degrees | [12 empty] | 117548087 |  | 907991 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | NotEngineering | [5 empty] | Engineering | [6 empty] | Tilt |  |  | Completed |  |  | Refresh |  |  | System issue | System | NotEngineering | ShouldbeCustomer | NotEngineering-ShouldbeCustomer | NotEngineering-ShouldbeCustomer | NotEngineering | NotEngineering-ShouldbeCustomer | NotEngineering | NotEngineering-ShouldbeCustomer | No | y | y | y | y | y | No | y | 0 | y | y | y | y | y | y | 0 | y | y
899 |  |  | Japan | Maki Hospital | 341017 | 44330 |  | 07CS:TM:操作室からの声が伝わりづらい。良い方法は無いか>>担当FSEに内容を報告。またアプリ担当者より現状で考えられる対応方法を回答 | BU Q&R Feedback: No Reason: Application Support Problem Reported by Customer : 操作室からの声が伝わりづらい。マイク部分にフィルムを貼っているかの確認。良い方法が無いかの確認 Safety Questions : Case: 0117544211 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: TAKEYUKI MIYAKAWA Last updated by: TAKEYUKI MIYAKAWA 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: Problem description by engineer : 対象装置:CT 故障個所:07その他 最初に発生したのはいつか?(~以内):1 Hour 問題発生の頻度は?:Unknow エンドユーザｰのワークフローのどの時点において問題が発生しているのか?:その他(Other) 最近どのようなイベントが発生したか?:N.A. 顧客は問題について具体的なシステムイベントレポートを作成しているか?:N.A. Engineer_s Comment: _============================== コール内容 操作室からの声が伝わりづらい。良い方法は無いか _============================== Troubleshooting Action: What steps are taken during the system diagnosis : 対応後の結果:改善された 故障個所:07その他 診断中にどのようなステップを実施したか?: _================================== 対応 担当FSEに内容を報告。 またアプリ担当者より現状で考えられる対応方法を回答 （マイクから少し距離をおいて話していただく、音量を調整いただく等） _================================== Error # and/or description of error : NA Repair Action: Completed or Recommended Repair Action "External" : 【現象確認】 確認済み 【作業内容詳細】 質問内容を回答 【確認内容】 良好 OneEMS internal remarks : Problem description by engineer : 対象装置:CT 故障個所:07その他 最初に発生したのはいつか?(~以内):1 Hour 問題発生の頻度は?:Unknow エンドユーザｰのワークフローのどの時点において問題が発生しているのか?:その他(Other) 最近どのようなイベントが発生したか?:N.A. 顧客は問題について具体的なシステムイベントレポートを作成しているか?:N.A. Engineer_s Comment: _============================== コール内容 操作室からの声が伝わりづらい。良い方法は無いか _============================== What steps are taken during the system diagnosis : 対応後の結果:改善された 故障個所:07その他 診断中にどのようなステップを実施したか?: _================================== 対応 担当FSEに内容を報告。 またアプリ担当者より現状で考えられる対応方法を回答 （マイクから少し距離をおいて話していただく、音量を調整いただく等） _================================== Error # and/or description of error : NA Completed or Recommended Repair Action "External" : 【現象確認】 確認済み 【作業内容詳細】 質問内容を回答 【確認内容】 良好 Completed or Recommended Repair Action "Internal" : --Troubleshooting Action-- 対応後の結果:改善された 故障個所:07その他 診断中にどのようなステップを実施したか?: _================================== 対応 担当FSEに内容を報告。 またアプリ担当者より現状で考えられる対応方法を回答 （マイクから少し距離をおいて話していただく、音量を調整いただく等） _================================== --Repair Action-- --Subject-- 07CS:TM:ﾎｼ様:操作室からの声が伝わりづらい。良い方法は無いか>>担当FSEに内容を報告。またアプリ担当者より現状で考えられる対応方法を回答 --Problem Description-- 対象装置:CT 故障個所:07その他 最初に発生したのはいつか?(~以内):1 Hour 問題発生の頻度は?:Unknow エンドユーザｰのワークフローのどの時点において問題が発生しているのか?:その他(Other) 最近どのようなイベントが発生したか?:N.A. 顧客は問題について具体的なシステムイベントレポートを作成しているか?:N.A. Engineer_s Comment: _============================== コール内容 操作室からの声が伝わりづらい。良い方法は無いか _============================== --4/8Elements-- Q1 報告された問題は確認できましたか? 確認済み Q2 他の問題は確認できましたか? 無し Q3 問題を解決するために何をおこないましたか? 質問内容を回答 Q4 対応後の装置は順調ですか? 良好 Q5 その検査はどの様に終了しましたか? 無し Q6 問題により間違った所見（誤診、誤治療）の有無: 無し Q7 意図しない不要な曝射の有無: 無し Q8 アーチファクトの有無: 無し ----------------------------------------------------------------------------- | 5 |  | Check if the film is stuck on the microphone part. Check if there is a good way | [12 empty] | 117544211 |  | 906473 |  | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Engineering | [6 empty] | FilmStuck |  |  | Completed |  |  | Refresh |  |  | System issue | System | Console-SW | FilmStuck | Console-SW-FilmStuck | Console-SW-FilmStuck | 0 | FilmStuck | Console-SW | Console-SW-FilmStuck | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
900 |  |  | Japan | Maki Hospital | 341017 | 44330 |  | CT:APP:AK:ﾎｼ様:質問 | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Application Support Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer : Q:①頭部アキシャル撮影で、360°と420°があるかと思いますが他のユーザーさんはどちらが多いですか、また見た目で違うのが分かりますでしょうか？ ②16×0.625で撮影していますが40×0.625で撮影されている施設はありますか？ A:①撮影の撮影時間を考えると360°が多く選択されています。また肉眼で画像の大きな違いははっきりわからない程度と思います。 ②16列と40列で使い分けているご施設は有りました。通常は16列ですがOPE後は早く撮影し脳全体を確認すためだけに40列で撮影すると聞いたことがあります。 Safety Questions : Case: 0117544070 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: ATSUSHI KOIZUMI Last updated by: ATSUSHI KOIZUMI 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A OneEMS internal remarks : Q1:報告された問題は確認できましたか？：確認済 Q2：他の問題は確認できましたか？；無 Q3：問題を解決するために何を行いましたか？ １：電話で解決した場合：説明 Q4：対応後の装置は順調ですか？：良好 FSE Notes : This PO is a Reference Number only SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 5 |  | resolved by phone: Explanation | [12 empty] | 117544070 |  | 906407 |  | Closed - Non-Complaint | Closed - Non-Complaint | [14 empty] | QA | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
901 |  |  | Japan | Maki Hospital | 341017 | 44330 |  | 04DM:KK:サジタル画像でZ方向に線状のアーチファクト。アキシャル画像でも感度ムラのようなアーチファクトがでた。 | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 04DM:KK:サジタル画像でZ方向に線状のアーチファクト。アキシャル画像でも感度ムラのようなアーチファクトがでた。 ----------------------------------------------------------------------------- | 4 |  | tbd | [12 empty] | 117544052 |  | 906396 |  | Investigation | Closed - Done |  | TBD-NoEnoughInfo-TBD-NoEnoughInfo | [5 empty] | Engineering | [6 empty] | linear artifact |  |  | Completed |  |  | Refresh |  |  | System issue | System | TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
902 |  |  | Japan | Higashi-Ohmi General Medical C | 341015 | 44329 |  | HY: Cannot install the pediatric attachment. | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : HY: Cannot install the pediatric attachment. Safety Questions:Case: 0117539014 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Hui Yi Last updated by: Wang Hui Yi 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A OneEMS internal remarks: Yamamoto DB SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Support ----------------------------------------------------------------------------- | 5 |  | Cannot install the pediatric attachment. | [12 empty] | 117539014 |  | 902823 |  | Evaluation | Closed - No Further Investigation |  | NotEngineering | [5 empty] | Engineering | [6 empty] | Pediatric |  |  | Completed |  |  | Refresh |  |  | System issue | System | NotEngineering | ShouldbeClinical | NotEngineering-ShouldbeClinical | NotEngineering-ShouldbeClinical | NotEngineering | NotEngineering-ShouldbeClinical | NotEngineering | NotEngineering-ShouldbeClinical | No | y | y | y | y | y | No | y | 0 | y | y | y | y | y | y | 0 | y | y
903 |  |  | Japan | Maki Hospital | 341017 | 44329 |  | 04CS:RS: A shower artifact occurred on head conventional scan images. >> The artifact has a point of converg [Truncated] | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 04CS:RS: A shower artifact occurred on head conventional scan images. >> Safety Questions:Case: 0117538815 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: RUI SATAKE Last updated by: RUI SATAKE 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A OneEMS internal remarks: --Troubleshooting Action-- Result of response: Problem solved. Problem location: 04 Gantry Exclude HV What steps did you perform to diagnose the problem? I asked him to perform re-reconstruction with an FOV of 500 mm, and when I checked the images, there was an artifact with a point of convergence at the left edge of the tabletop. The artifact is the kind that can occur with arcing, but there was no arcing recorded in the log. Also, given that when he performed another conventional scan, a shower artifact occurred in the same place, and he said that he had not repositioned the patient to perform a helical scan, the patient setting is the same in the FOV, but he returned the tilt to zero, so the X-ray irradiation line would have changed. Given the above, it is possible that the tabletop or an accessory such as an additional cushion caused the problem. I told Mr. Matsumoto the above findings and told him that we recommend the use of the head extension for head scans, and he agreed to monitor the situation. --Repair Action-- --Subject-- 04CS:RS:Mr. Matsumoto: A shower artifact occurred on head conventional scan images. >> The artifact has a point of convergence near the tabletop. --Problem Description-- Equipment affected: CT Problem location: 04 Gantry Exclude HV When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user_s work flow did the problem occur?: During reconstruction/post processing What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem?: NA --4/8Elements-- Q1. Was the reported problem confirmed? Confirmed. Q2. Were any other problems confirmed? None. Q3. What did you do to resolve the problem?: I asked him to perform re-reconstruction with an FOV of 500 mm, and when I checked the images, there was an artifact with a point of convergence at the left edge of the tabletop. The artifact is the kind that can occur with arcing, but there was no arcing recorded in the log. Also, given that when he performed another conventional scan, a shower artifact occurred in the same place, and he said that he had not repositioned the patient to perform a helical scan, the patient setting is the same in the FOV, but he returned the tilt to zero, so the X-ray irradiation line would have changed. Given the above, it is possible that the tabletop or an accessory such as an additional cushion caused the problem. I told Mr. Matsumoto the above findings and told him that we recommend the use of the head extension for head scans, and he agreed to monitor the situation. Q4: Is the equipment operating normally following the response? Good. Q5. How did the examination finish? None. Q6. Were there any mistaken findings (misdiagnosis, mistaken treatment) due to the problem? None. Q7. Was there any unintended unnecessary exposure? None. Q8. Were there any artifacts? None. T2 Activities:Problem Description: subject equipment: CT Problem location: 04 Gantry Exclude HV When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user_s work flow did the problem occur?: During reconstruction/post processing What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem?: NA Troubleshooting Action: Result after troubleshooting: Improved Problem location: 04 Gantry Exclude HV What steps did you perform to diagnose the problem? I asked him to perform re-reconstruction with an FOV of 500 mm, and when I checked the images, there was an artifact with a point of convergence at the left edge of the tabletop. The artifact is the kind that can occur with arcing, but there was no arcing recorded in the log. Also, given that when he performed another conventional scan, a shower artifact occurred in the same place, and he said that he had not repositioned the patient to perform a helical scan, the patient setting is the same in the FOV, but he returned the tilt to zero, so the X-ray irradiation line would have changed. Given the above, it is possible that the tabletop or an accessory such as an additional cushion caused the problem. I told Mr. Matsumoto the above findings and told him that we recommend the use of the head extension for head scans, and he agreed to monitor the situation. Repair Action: [Confirmation of problem] Confirmed. Work details: I asked him to perform re-reconstruction with an FOV of 500 mm, and when I checked the images, there was an artifact with a point of convergence at the left edge of the tabletop. The artifact is the kind that can occur with arcing, but there was no arcing recorded in the log. Also, given that when he performed another conventional scan, a shower artifact occurred in the same place, and he said that he had not repositioned the patient to perform a helical scan, the patient setting is the same in the FOV, but he returned the tilt to zero, so the X-ray irradiation line would have changed. Given the above, it is possible that the tabletop or an accessory such as an additional cushion caused the problem. I told Mr. Matsumoto the above findings and told him that we recommend the use of the head extension for head scans, and he agreed to monitor the situation. 【Confirmation content】 Good. ----------------------------------------------------------------------------- | 4 |  | Given the above, it is possible that the tabletop or an accessory such as an additional cushion caused the problem. I told Mr. Matsumoto the above findings and told him that we recommend the use of the head extension for head scans, and he agreed to monitor the situation. | [12 empty] | 117538815 |  | 902670 |  | Closed - Done | Closed - Done |  | NotEngineering | [5 empty] | Engineering | [6 empty] | Artifact |  |  | Completed |  |  | Refresh |  |  | System issue | System | NotEngineering | ShouldbeCustomer | NotEngineering-ShouldbeCustomer | NotEngineering-ShouldbeCustomer | NotEngineering | NotEngineering-ShouldbeCustomer | NotEngineering | NotEngineering-ShouldbeCustomer | No | y | y | y | y | y | No | y | 0 | y | y | y | y | y | y | 0 | y | y
904 |  |  | Japan | Saiseikai Gosu Hospital | 343000 | 44327 |  | HY:ガントリーベルト交換 | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : HY:ガントリーベルト交換 Safety Questions : Case: 0117526745 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Hui Yi Last updated by: Wang Hui Yi 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A OneEMS internal remarks : 河重DB SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support ----------------------------------------------------------------------------- | 5 |  | TBD | [12 empty] | 117526745 |  | 893816 |  | Evaluation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Not clear |  |  | Completed |  |  | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
905 |  |  | Japan | Maki Hospital | 341017 | 44326 |  | CT:APP:NO: Mr. Hoshi: Inquiry. | BU Q&R Feedback: No Reason: Application Support Problem Reported by Customer : Q: Thin slice data from a head axial scan was split into two series. A: It is possible to merge the split series data using the CT Viewer and create cor/sag images. ----------------------------------------------------------------------------- | 5 |  | Kevin investigation | [12 empty] | 117520965 |  | 889502 |  | Evaluation | Closed - non-complaint | [7 empty] | TBD | [6 empty] | conventional |  |  | Completed |  |  | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
906 |  |  | Japan | Maki Hospital | 341017 | 44327 |  | In conventional head exam, when thin slice data read into CT Viewer, the data is displayed separately. |  | WANG, Peter |  | Kevin investigation | [14 empty] | 889106 |  | Investigation | Closed - Done | [7 empty] | TBD | [6 empty] | conventional |  |  | Completed |  |  | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
907 |  |  | Poland | Agencja Rezerw Materiałowych | 341010 | 44330 |  | Artefakty kolorowe | Work Order #: WO-03494007 Robert Kamola Diagnostic performed by Engineer : N/A Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: ----------------------------------------------------------------------------- Work Order #: WO-03495119 Robert Kamola Diagnostic performed by Engineer : N/A Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: ----------------------------------------------------------------------------- | 4 |  | DMS replace | [12 empty] | 117544532 |  | 906631 |  | Closed - Done | Closed - Done |  | CTUDT00632110 | [5 empty] | Engineering | [6 empty] | Artifact |  |  | Completed |  |  | Refresh |  |  | System issue | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | DMS Module | DMS-HW | DMS-HW-DMS Module | No | y | y | y | y | y | No | y | 0 | y | y | y | y | y | y | 0 | y | y
908 |  |  | Poland | Agencja Rezerw Materiałowych | 341010 | 44330 |  | Podpięcie strzykawki | Work Order #: WO-03494014 Robert Kamola Diagnostic performed by Engineer : N/A Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: ----------------------------------------------------------------------------- | 4 |  | Attaching the syringe | [12 empty] | 117544535 |  | 906632 |  | Closed - Non-Complaint | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
909 |  |  | Poland | Uniwersytecki Szpital Kliniczny | 341001 | 44327 |  | Configuration of the DVD robot. | Work Order #: WO-03475733 14-May-2021 Marcin Knyps The system fully meets specification and returned to use Problem Description by Engineer: Change of device configuration, adding a recorder. Resolution: Adding a recorder node. 172.16.18.67 AROBOT1 6001. Recording test - correct result. A change of configuration ended successfully. ----------------------------------------------------------------------------- | 4 |  | Change of device configuration, adding a recorder. | [12 empty] | 117527848 |  | 894609 |  | Closed - Non-Complaint | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
910 |  |  | Portugal | Centro Hospitalar de Vila Nova | 40030 | 44326 |  | Checking the couch cover on the left side of the equipment |  | 3 |  | Checking the table cover on the left side of the equipment | [12 empty] | 117522287 |  | 890531 |  | Evaluation | Evaluation |  | TBD-CouchCover | [5 empty] | Engineering | [6 empty] | Table cover |  |  | Completed |  |  | Refresh |  |  | System issue | System | Couch-HW | Noise | Couch-HW-Noise | Couch-HW-Noise | TBD-CouchCover | Noise | Couch-HW | Couch-HW-Noise | No | y | y | y | y | y | No | y | 0 | y | y | y | y | y | y | 0 | y | y
911 |  |  | Sweden | Ludvika Lasarett | 40049 | 44328 |  | SE5333 - Long time to reconstruct an investigation. |  | 3 |  | Customer will call again if the issue is to reoccur. | [12 empty] | 117534933 |  | 899920 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-NoEnoughInfo-TBD-NoEnoughInfo | [5 empty] | Engineering | [6 empty] | long recon time |  |  | Completed |  |  | Refresh |  |  | System issue | System | TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
912 |  |  | United States | St Barnabas Hospital | 40028 | 44331 | 19 | 0711CFXXX cannot transfer images to the PACS | Problem Description by Engineer : Problem description by engineer : cannot transfer images to the PACS When was the first occurrence? Within... : 15-May-2021 Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: No further information is available as per the customer Patient Impact/Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer Current Software Version: Not provided Resolution : Advised to reset the spooler and the PACS Diagnostic performed by Engineer : advised the customer to reset the spooler in Olevel Internal Comments : Error # and/or description of error : cannot transfer images to the PACS Completed or Recommended Repair Action "Internal" : advised the customer to reset the spooler in Olevel, follow up a few hours later and there was no answer. | 2 |  | reset the spooler, Suggest to install HF3 (FCO72800740). | [12 empty] | 117548905 |  | 907654 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS Transfer |  |  | Completed |  |  | Refresh |  |  | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
913 |  |  | United States | St Barnabas Hospital | 40028 | 44330 |  | 3D cell greyed out when saving as batch | Work Order #: WO-03496169 Hirchak, Michele The system fully meets specification and returned to use Problem description by Clinical Support Specialist : Customer Function/Role: Tech Device Use: In Use Expected and Unexpected Results: Save a 3D image. Patient Impact: None Current Software Version: 4. No rescan, no loss of data, no reinjection. Diagnostic performed by Engineer : Customer already figured it out by the time the call got transferred to Apps. She is all set. MBH ----------------------------------------------------------------------------- | 3 |  | Customer already figured it out by the time the call got transferred to Apps. She is all set. MBH | [12 empty] | 117546527 |  | 904454 |  | Investigation | Closed - Done |  | NotEngineering | [5 empty] | Engineering | [6 empty] | 3D |  |  | Completed |  |  | Refresh |  |  | System issue | System | NotEngineering | ShouldbeCustomer | NotEngineering-ShouldbeCustomer | NotEngineering-ShouldbeCustomer | NotEngineering | NotEngineering-ShouldbeCustomer | NotEngineering | NotEngineering-ShouldbeCustomer | No | y | y | y | y | y | No | y | 0 | y | y | y | y | y | y | 0 | y | y
914 |  |  | United States | Unity Health Systems | 40037 | 44328 |  | Monthly QA for JUne | Work Order #: WO-03483802 Lockwood Thomas Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 5 |  | NAC | [12 empty] | 117535034 |  | 899978 |  | Closed - Non-Complaint | Closed - Non-Complaint | [14 empty] | QA | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
915 |  |  | United States | Unity Health Systems | 40037 | 44328 |  | Monthly QA for May and June | Work Order #: WO-03483746 Lockwood Thomas Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 5 |  | NAC | [12 empty] | 117534992 |  | 899957 |  | Closed - Non-Complaint | Closed - Non-Complaint | [14 empty] | QA | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
916 |  |  | United States | The Regional Medical Center | 40005 | 44328 | 19 | 0614CDY73 Gantry will not come on | Work Order #: WO-03485838 Maier, Donnie The system does not meet specification and is out of use Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: Diag Expected and Unexpected Results: Cant close estop User Impact: No Patient Impact: No Current Software Version: 4.1.10 Diagnostic performed by Engineer : Checked ULV to see following errors. 2021/05/12 14:36:32.132 S_TILT_C_POT_CALIB_KO ESTOP 1.344502E+08 0 0 0 0.00146990740740741 0.0352777777777778 2.11666666666667 2021/05/12 14:36:31.945 S_TILT_C_POT_CALIB_KO WARN 511 90 0 3997697 0.00145833333333333 0.035 2.1 PFSE Doug Holt was onsite and I conferenced in my colleague on NSS side (Don Parr). After having PFSE check connections to pot and belt tension still having issues with calibrating. This is a reoccurring intermittent issue from a couple days ago so new pot assembly was already onsite and PFSE will try installing and calibrating. Will call back if needs further assistance. Follow Up Required : Install new pot assembly and try to recalibrate. Pot assembly is still onsite from previous closed case from a couple days ago. ----------------------------------------------------------------------------- Work Order #: WO-03486157 12-May-2021 Holt Douglas The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: Diag Expected and Unexpected Results: Cant close estop User Impact: No Patient Impact: No Current Software Version: 4.1.10 Tools & Equipment : MULTIMETER,2022-08-30 00:00:00 Follow Up Required : Verify system operation Resolution : Replaced dual pot assembly and calibrated the tilt ----------------------------------------------------------------------------- | 2 |  | Replaced dual pot assembly and calibrated the tilt | [12 empty] | 117536646 |  | 897217 |  | Closed - Done | Closed - Done |  | ECR-077012 | [5 empty] | Engineering | [6 empty] | Dual pot |  |  | Completed |  |  | Refresh |  |  | System issue | System | Gantry-HW | Belt | Gantry-HW-Belt | Gantry-HW-Belt | ECR-077012 | Belt | Gantry-HW | Gantry-HW-Belt | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
917 |  |  | United States | McLaren Macomb | 40021 | 44328 |  | 4.1.1 CTA _s are intermittently not combining results | Work Order #: WO-03484293 Giles Brian The system fully meets specification and returned to use Problem description by Clinical Support Specialist : Customer Function/Role: tech Device Use: In Use Expected and Unexpected Results: CTA _s are intermittently not combining results User Impact: None Patient Impact: None Current Software Version: 4.1.1 Diagnostic performed by Engineer : I followed up with Michael and left a voicemail. that I will be closing the case Diagnostic performed by Engineer : I brought Michele Hrchack (RAS) on board. We had customer look at parameters page. X/Y was off by51mm. Tried recon and it didn_t work. Then we went into CT Viewer and had them merge the disconnected series. That did work and customer was happy. | 5 |  | I brought Michele Hrchack (RAS) on board. We had customer look at parameters page. X/Y was off by51mm. Tried recon and it didn_t work. Then we went into CT Viewer and had them merge the disconnected series. That did work and customer was happy. | [12 empty] | 117535382 |  | 896220 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | ShouldbeClinical | [5 empty] | Engineering | [6 empty] | CTA |  |  | Completed |  |  | Refresh |  |  | System issue | System | Console-SW | CTA | Console-SW-CTA | Console-SW-CTA | ShouldbeClinical | CTA | Console-SW | Console-SW-CTA | No | y | y | y | y | y | No | y | 0 | y | y | y | y | y | y | 0 | y | y
918 |  |  | United States | McLaren Macomb | 40021 | 44327 |  | Getting tube heat message on scans with bolus tracking | Work Order #: WO-03477043 Hirchak, Michele The system fully meets specification and returned to use Diagnostic performed by Engineer : Customer getting tube heat message on scans with bolus tracking. Had her change from 30 trackers to 20. Also, turned off AST and changed the RT to 0.75. Talked about changing the time on the fly if they get the message. Looked at 100kV PE study / Abd Pelvis and the 120kV PE with Abd/Pelvis. Keeping case open. MBH Problem description by Clinical Support Specialist : Called to f/u with Allison. She was busy. Resolution : Spoke to Allison. Tube heat messages have stopped. She is all set. OK to close. MBH Problem description by Clinical Support Specialist : Customer Function/Role: Tech Device Use: In Use Expected and Unexpected Results: Tube heat on scans with bolus tracking. Patient Impact: None Current Software Version: 4.1.10 No rescan, no loss of data, no reinjection | 5 |  | Customer getting tube heat message on scans with bolus tracking. Had her change from 30 trackers to 20. Also, turned off AST and changed the RT to 0.75. | [12 empty] | 117528731 |  | 890964 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | Tube Heat | [5 empty] | Engineering | [6 empty] | Tube Overheat |  |  | Completed |  |  | Refresh |  |  | System issue | System | Gantry-SW | Tube Heat | Gantry-SW-Tube Heat | Gantry-SW-Tube Heat | Tube Heat | Tube Heat | Gantry-SW | Gantry-SW-Tube Heat | No | y | y | y | y | y | No | y | 0 | y | y | y | y | y | y | 0 | y | y
919 |  |  | United States | Advent Health Wesley Chapel | 300099 | 44330 |  | 0711CDX-Console goes black when going to next series of scan | Console goes black when going to next series of scan - Have to start all over when this happens | [15 empty] | 117547277 |  | 905109 |  | Investigation | Opened |  | BlackScreen | [5 empty] | Engineering | [6 empty] | Black screen |  |  | Complete |  |  | Refresh |  |  | Upgrade issue | System | Console-HW | BlackScreen | Console-HW-BlackScreen | Console-HW-BlackScreen | BlackScreen | BlackScreen | Console-HW | Console-HW-BlackScreen | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
920 |  |  | United States | Advent Health Wesley Chapel | 300099 | 44324 |  | 0705CDN84 customer locked out of system, multiple issues | customer locked out of system, multiple issues |  |  | Single case. Replace Video card & Keyboard | [14 empty] | 900856 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | NotEngineering-partreplace | [5 empty] | Engineering | [6 empty] | Video card & Keyboard |  |  | Complete |  |  | Refresh |  |  | Upgrade issue | System | Console-HW | Video card & Keyboard | Console-HW-Video card & Keyboard | Console-HW-Video card & Keyboard | NotEngineering-partreplace | Video card & Keyboard | Console-HW | Console-HW-Video card & Keyboard | No | y | y | y | y | y | No | y | 0 | y | y | y | y | y | y | 0 | y | y
921 |  |  | United States | Advent Health Wesley Chapel | 300099 | 44330 |  | 0711CDX DISPLAYS SYSTEM ERROR WHILE PLOTTING FOR SCAN AFTER SURVIEW | DISPLAYS SYSTEM ERROR WHILE PLOTTING FOR SCAN AFTER SURVIEW | [15 empty] | 117546481 |  | 904438 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | BlackScreen | [5 empty] | Engineering | [6 empty] | Black screen |  |  | Complete |  |  | Refresh |  |  | Upgrade issue | System | Console-HW | BlackScreen | Console-HW-BlackScreen | Console-HW-BlackScreen | BlackScreen | BlackScreen | Console-HW | Console-HW-BlackScreen | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
922 |  |  | United States | Advent Health Wesley Chapel | 300099 | 44322 |  | Lost system functionality post iPatient S upgrade (v4.1.7 to v4.1.10) - unable to utilize Same Series tool | •Feedback Type: Technical Issue •Other Details: NA •Submitter Email: melissa.lyter@philips.com •CFM Form Received Date (Outlook email received date): 05/12/2021 Lost system functionality post iPatient S upgrade- Ability to utilize Same Series tool On the v4.1.7 software, when the customer wanted to save multiple individual ROI images in the calcium scoring application or save multiple individual images in the brain perfusion application, they would use the Same Series option. This option is no longer available when using the same workflow wtih v.4.1.10. The customer_s workflow was to save multiple individual screenshots for calcium scoring and brain perfusion studies. They would use the save image or save display option to do this. When labeling the new image series, as long as they kept the description name the same, the Same series option would appear already checked to keep the images in the same series. Now in v4.1.10, the Same series function no longer appears even though they keep the description name the same. This means when saving this way the images are save separately. This seems to apply to all post-processing applications. Yes the system was in clinical use at the time. The problem occurred during diagnosis. The procedure was continued. The issue occurs every time they try to save multiple images. No impact to the patient. The customer has to use the workaround of saving the individual images to the Film application. From the Filming application, they can choose to save all the images together in the same series. |  |  | Check the behavior on 4.1.10 -- Kevin | [14 empty] | 894938 |  | Investigation | Investigation | [7 empty] | Customer | [6 empty] | Refresh diff |  |  | Complete |  |  | Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
923 |  |  | United States | Eisenhower Medical Center | 68116 | 44328 |  | Table is screaching in in and out | Table is screaching in in and out | [15 empty] | 117532201 |  | 896762 |  | #N/A | Evaluation | [14 empty] | Table noise | [5 empty] | Non-Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | #REF!
924 |  |  | United States | Advent Health Wesley Chapel | 300106 | 44305 |  | After a 4.1.10 upgrade Interventional/Biopsy_s are giving incorrect coordinates | •Feedback Type: Technical Issue •Other Details: NA •Submitter Email: David.Mujica@philips.com •CFM Form Received Date (Outlook email received date): 04/21/2021 After a 4.1.10 upgrade over this past weekend 4/7-4/18, Interventional/Biopsy_s are giving incorrect coordinates as described by CT Technologist Russ Golden. Once it was realized that the Interventional/Biopsy was not working as intended all study_s were stopped and placed on hold and it was reported by the customer to CT Remote Clinical Support and to the Remote Service Engineering group. The device was being used to perform a Biopsy on a patient and was in clinical use. Biopsy_s are frequently performed at this site. There was no negative impact to the user or patient as reported by the customer. The resolve was to create a case and dispatch the Field Service Engineer to fix the problem on site. The device being used was an Ingenuity CT system 4.1.10. Install Product . Reason for delay: Missing event date |  |  | CTUDT00632916 IS CREATED TO COVER CCT wrong position issue | [14 empty] | 905009 |  | Investigation | Opened |  | CTUDT00632916 | [5 empty] | Engineering | [6 empty] | Biopsy |  |  | Complete |  |  | Refresh |  |  | Upgrade issue | System | Console-SW | Biopsy | Console-SW-Biopsy | Console-SW-Biopsy | CTUDT00632916 | Biopsy | Console-SW | Console-SW-Biopsy | No | y | y | y | y | y | No | y | 0 | y | y | y | y | y | y | 0 | y | y
925 |  |  | United States | Advent Health Wesley Chapel | 300106 | 44322 |  | Lost system functionality post iPatient S upgrade (v4.1.7 to v4.1.10) - unable to utilize Same Series tool | •Feedback Type: Technical Issue •Other Details: NA •Submitter Email: melissa.lyter@philips.com •CFM Form Received Date (Outlook email received date): 05/12/2021 Lost system functionality post iPatient S upgrade- Ability to utilize Same Series tool On the v4.1.7 software, when the customer wanted to save multiple individual ROI images in the calcium scoring application or save multiple individual images in the brain perfusion application, they would use the Same Series option. This option is no longer available when using the same workflow wtih v.4.1.10. The customer_s workflow was to save multiple individual screenshots for calcium scoring and brain perfusion studies. They would use the save image or save display option to do this. When labeling the new image series, as long as they kept the description name the same, the Same series option would appear already checked to keep the images in the same series. Now in v4.1.10, the Same series function no longer appears even though they keep the description name the same. This means when saving this way the images are save separately. This seems to apply to all post-processing applications. Yes the system was in clinical use at the time. The problem occurred during diagnosis. The procedure was continued. The issue occurs every time they try to save multiple images. No impact to the patient. The customer has to use the workaround of saving the individual images to the Film application. From the Filming application, they can choose to save all the images together in the same series. |  |  | Check the behavior on 4.1.10 -- Kevin | [14 empty] | 894943 |  | Investigation | Investigation | [7 empty] | Customer | [6 empty] | Refresh diff |  |  | Complete |  |  | Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
926 |  |  | Australia | Maia Financial Pty Limited | 341000 | 44334 | 20 | Scanner cant initialize. | Work Order #: WO-03510806 Ian Fishwick The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: CT Radiographer Device Use: Initialise CT Scan Expected and Unexpected Results: Site indicated that they could not initialise a CT Spine Scan. User Impact: No Patient Impact: None Reported Current Software Version:4.1.10 Diagnostic performed by Engineer : _______________________________________________________________ PROBLEM DESCRIPTION Site indicated that they could not initialise a CT Spine Scan. ________________________________________________________________ DATE THE ISSUE OCCURRED: 19 May 2021 ________________________________________________________________ TROUBLESHOOTING: System SW Version is: 4.1.10 I spoke to CT Radiographer Elizabeth and she has rebooted the system including the CT Host and Reconstructor plus Power cycle of the Gantry, after this action the system completed the Short Tube Conditioning and also a test scan without error. ________________________________________________________________ LOG REVIEW: A Bugreport was saved the previous evening when the original issue occurred, this will be downloaded and reviewed to understand the original issue. Review of Bugreport indicates that the system failed to initialise a Scan at 20:12:14.463 due to the following error. 2021/05/18 20:12:15.021 S_DMC_UHR_COMB_TMOT ERROR_WARNING Description: UHR movement Time Out The previous series and this series are not programed as UHR so UHR should not be in motion. There are also at same time reported from the CIRS an Optical Datapath error. and information in this error indicates that the Optical Channel link was not known reporting 0x0 instead of 0x4321 Error #12 - CheckGantryConnection: ACQ board or Gantry-to-CIRS datapath failure. Gantry Power Cycle has cleared both errors. Review of Gantry on time at time of the event indicates that last Gantry Power Cycle was 53 days ago, Philips Recommends gantry Power cycle once in a 7 day period. I have requested that Philips Engineer attending next Scheduled system PM on the 6th June 2021 to run the sites CT team through the complete User maintenance procedures including CT Host reboot daily, CIRS reboot once per week and Gantry Power Cycle once per week. ________________________________________________________________ CUSTOMER UPDATE: Site have advised that after the Gantry Power cycle that they have been able to complete all new CT Scans for the Clinical Scanning day 19th May 2021, the site will continue to monitor and report any further issues. ________________________________________________________________ RECOMMENDATION: Monitoring is recommended at this stage, further investigation will be required should this issue recur. If you have any further questions don’t hesitate to contact me. Ian Fishwick Philips ANZ Remote Service Engineer. Email, ian.fishwick@philips.com. You may receive a short survey re our service today, and your feedback is gratefully used to drive continuous improvement in our service delivery. ________________________________________________________________ CURRENT SYSTEM STATUS: System is in clinical use. ________________________________________________________________ ----------------------------------------------------------------------------- | 2 |  | RECOMMENDATION: Monitoring is recommended at this stage, further investigation will be required should this issue recur. If you have any further questions don’t hesitate to contact me. | [12 empty] | 117559515 |  | 915891 |  | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Power cycle |  |  | Completed |  |  | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
927 |  |  | Chile | Scotiabank Chile | 40008 | 44335 |  | CL1675-ON SCANNER SCREEN SHOWS A FLASHING ICON. WHEN TAKING AN EXAM, THE SCREEN SHOWS A CT ICON LIKE FLASHING AND NOT CORRECT, IT MAY BE ERRATIC OPERATION | Problem description by engineer: Low imaging disk space. Information to support the complaint handling process Customer Function / Role: BioMed Device Use: Diagnostic Expected and Unexpected results: No disk space errors. Patient Impact / Outcome: Delay in exams. User Impact: Delay in exams. Current Software Version: 4.1.10 It presents 1% of available space in Local disk of images. There is a group of event log files that are taking up a lot of space in the d: \ Pms \ Log \ Data folder. The files are erased and the available disk capacity is good. The client is requested to restart the computer daily. Error # and / or description of error: Space available in local dioco 1%. | 4 |  | Delete event log files that are taking up a lot of space in the d: \ Pms \ Log \ Data folder. The files are erased and the available disk capacity is good. | [12 empty] | 117568913 |  | 924783 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00633269 | [5 empty] | Engineering | [6 empty] | Low disk |  |  | Completed |  |  | Refresh |  |  | System issue | System | Console-SW | DiskFull | Console-SW-DiskFull | Console-SW-DiskFull | CTUDT00632261 | DiskFull | Console-SW | Console-SW-DiskFull | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
928 |  |  | China | 山东省泰安市中医二院 | 345037 | 44338 |  | 02N/Ingenuity Core128/Elite China/add printer | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ Safety Questions : Case: 0117583119 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Qian Wang Last updated by: Qian Wang 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 22.05.2021 13:50:55 UTC Tao Chen (86038373)故障现象: 添加打印机 服务内容: 设置打印机 服务结果: 维修后,设备恢复正常. Next Step Action : chen tao Internal Remarks : 22.05.2021 13:50:58 UTC Tao Chen (86038373) “RDI Accuracy: YES" $Q7: (Confirmed reported problem, add printer) $Q8: ( Get IP, port and AE from hospital IT) $Q9: (Add new DICOM printer) $Q10: ( Test and inspection data refer to: PA. Device returned to full functionality) $Q11: Calibrated Tools name ( N/A) ^Current Tube installed date: N/A ^Current Tube usage: N/A ----------------------------------------------------------------------------- | 5 |  | add printer | [12 empty] | 117583119 |  | 933175 |  | Closed - Non-Complaint | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
929 |  |  | China | 山东安丘市中医医院 | 345017 | 44337 |  | 02NIngenuity Core128/Elite China/image transfer error | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : Ingenuity Core128/Elite China SAP Transfer : Case Transferred to SAP Next Step Action : FSE Yang xingbo 013583235930来电 Safety Questions : Case: 0117578656 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Jianqiang Last updated by: Wang Jianqiang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 图像传输慢 客户问题详细描述或故障现象： 图像传输慢 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 图像传输慢 远程诊断与测试结果： 现场诊断 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 检查网络，检查软件 SAP Transfer : Dispatched for Onsite Support ----------------------------------------------------------------------------- | 5 |  | Intermittent problem. After the exam is completed, the patient image cannot be automatically transmitted to the PACS system and needs to be sent manually. Escalation 0117578777 raised | [12 empty] | 117578656 |  | 931413 |  | Evaluation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS Transfer |  |  | Completed |  |  | Refresh |  |  | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
930 |  |  | China | 陕西省咸阳中心医院 (咸阳市第二人民医院） | 40050 | 44337 |  | Application/Ingenuity CT 728326 | BU Q&R Feedback: No Reason: Application Support Problem Reported by Customer : Application/Ingenuity CT 728326 Next Step Action : 客户要求换个培训老师再次培训，已反馈售前培训经理跟进 Safety Questions : Case: 0117577477 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Bai Jingwen Last updated by: Bai Jingwen 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 5 |  | Application training | [12 empty] | 117577477 |  | 930780 |  | Evaluation | Closed - non-complaint | [7 empty] | TBD | [6 empty] | NAC |  |  | Completed |  |  | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
931 |  |  | China | 济南市第四人民医院 | 40038 | 44337 |  | 02N/Ingenuity CT 728326/filming lost title | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity CT 728326/ T2 Activities : Problem Description: 客户问题概述： 其他类型故障。 客户问题详细描述或故障现象： 主机打印图像时会丢扫描时间title。 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： log无报错。 远程诊断与测试结果： 重启无效。 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议互换扫描时间title，更改字体大小间距排除问题，检查重建柜时区和时间设置。 建议备件12NC及名称： N/A Safety Questions : Case: 0117576901 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Bo Last updated by: Wang Bo 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 21.05.2021 14:40:11 UTC Tao Chen (86038373)故障现象: Film图像上间歇性无患者信息，Film图像间距大 服务内容: 设置layout为全页打印，修改打印排版 服务结果: 维修后,设备恢复正常. Internal Remarks : 22.05.2021 14:20:31 UTC Tao Chen (86038373) “RDI Accuracy: YES" $Q7: (Confirmed reported problem, no patient film title on CT images intermittently) $Q8: ( Check patient film process, found film layout page was single when do film printing) $Q9: ( Adjust film page layout to full page) $Q10: ( Test and inspection data refer to: PA. Device returned to full functionality) $Q11: Calibrated Tools name ( N/A) ^Current Tube installed date: N/A ^Current Tube usage: N/A ----------------------------------------------------------------------------- | 5 |  | Check patient film process, found film layout page was single when do film printing) $Q9: ( Adjust film page layout to full page) | [12 empty] | 117576901 |  | 930474 |  | Closed - Done | Closed - Done | [7 empty] | Application | [6 empty] | Image title |  |  | Completed |  |  | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
932 |  |  | China | 莱阳市人民医院 | 345005 | 44337 |  | 02N/Ingenuity Core128/Rebuild the image failure | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer : 02N/Ingenuity Core128/ Safety Questions : Case: 0117576807 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingsong Last updated by: Lu Mingsong 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: Rebuild the image failure Troubleshooting Action: Remotely instruct the customer to restart the system. Repair Action: Remotely instruct the customer to restart the system. Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- |  |  | Remotely instruct the customer to restart the system. | [12 empty] | 117576807 |  | 930400 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-NoEnoughInfo-TBD-NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Recon |  |  | Completed |  |  | Refresh |  |  | System issue | System | TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
933 |  |  | China | 四川华蓥市华蓥山广能集团总医院 | 345000 | 44336 |  | 02N/Ingenuity Core128/Elite China/Password expired | BU Q&R Feedback: No SAP Customer Symptom Code: HANG Customer Symptom Code: Hang/Crash Reason: Corrective Maintenance Engineer Failure Code: Configuration Error Repair Activity Code: Configuration Change Problem Reported by Customer : 02N/Ingenuity Core128/Elite China Safety Questions : Case: 0117572025 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Bo Last updated by: Wang Bo 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: host software password expired. Troubleshooting Action: Assist customer to reset password Repair Action: Assist customer to reset password Test/Inspection: customer confirm can scan patient and image ok.functional Check PASSED. Device returned to full functionality. Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- |  |  | Assist customer to reset password | [12 empty] | 117572025 |  | 926787 |  | Closed - Non-Complaint | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
934 |  |  | China | 河南省濮阳市人民医院 | 40031 | 44336 |  | 02N/Ingenuity CT 728326/Estop open | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity CT 728326/ T2 Activities : Problem Description: 客户问题概述： 机架倾斜运动故障。 客户问题详细描述或故障现象： 扫描打角度时掉电。 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： S_TILT_C_NO_MOVE 远程诊断与测试结果： 重启恢复。diffPOT12 =-1 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议检查pot皮带张紧度，滑轮六角螺丝固定，limit switch。必要时更换pot并做tilt calibration。 建议备件12NC及名称： PPD 459800441121 SAP Transfer : Case Transferred to SAP Internal Remarks : 23.05.2021 07:34:01 UTC Yongping Luo (86038950) $Q7:Confirmed report problem，the gantry maybe lose power and estop open when tilt movement. $Q8:Perform visiual check on site ,replace the gantry tilt potentiometer, device return to fully function. $Q9:replace the tilt potentiometer,Device return to full function. $Q10: test and inspection data refer to PA. Device returned to full functionality. $Q11:Calibrated Tools refer to PA. External Remarks : 23.05.2021 07:33:58 UTC SUP JCOUSER (86038950)服务内容：机架倾斜时，偶发掉电，无法扫描，更换tilt 电位器。 服务结果：更换tilt电位器，设备恢复正常。 其它重要事项：无 Safety Questions : Case: 0117570465 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Bo Last updated by: Wang Bo 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Dispatched for Onsite Support Next Step Action : PPD 459800441121*1 已安排24H到医院 罗永平-86-15226131133 ----------------------------------------------------------------------------- | 5 |  | Delete event log files that are taking up a lot of space in the d: \ Pms \ Log \ Data folder. The files are erased and the available disk capacity is good. | [12 empty] | 117570465 |  | 925749 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | ECR-077012 | [5 empty] | Engineering | [6 empty] | Dual pot |  |  | Completed |  |  | Refresh |  |  | System issue | System | Gantry-HW | Belt | Gantry-HW-Belt | Gantry-HW-Belt | ECR-077012 | Belt | Gantry-HW | Gantry-HW-Belt | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
935 |  |  | China | 山东省泰安市中医二院 | 345037 | 44336 |  | 02N/Ingenuity Core128/Elite China/workaround for yellow screen | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Calibration/Adjustment Problem Reported by Customer : 02N/Ingenuity Core128/Elite China Safety Questions : Case: 0117569866 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Gaoge Last updated by: Wang Gaoge 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action : Q6 = workaround for yellow screen SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support T2 Activities : Problem Description: 客户问题概述： 计算机软件相关故障。 客户问题详细描述或故障现象： 计算机主机黄屏，间歇性故障 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 主机黄屏，按钮无反应 远程诊断与测试结果： 软件defect PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 1，执行软件升级 Test/Inspection: N/A Test Equipment: N/A ----------------------------------------------------------------------------- |  |  | Yellow screen workaround installation | [12 empty] | 117569866 |  | 925522 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632520 | [5 empty] | Engineering | [6 empty] | New Yellow Screen |  |  | Completed |  |  | Refresh |  |  | System issue | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #REF!
936 |  |  | China | 吉林省松原市中心医院 | 345008 | 44335 |  | 02N/Ingenuity Core128/Elite China/transfer error | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity Core128/Elite China SAP Transfer : Case Transferred to SAP T2 Activities : Problem Description: 客户问题概述： 图像传输失败。 客户问题详细描述或故障现象： 图像老师传输一半之后就不传了 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 传输故障 远程诊断与测试结果： 传输问题，客户要求工程师到场 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 检查网络速度,检查软件 Safety Questions : Case: 0117564872 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Jianqiang Last updated by: Wang Jianqiang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Dispatched for Onsite Support External Remarks : 19.05.2021 14:17:42 UTC Xiaodong Chi (86038306)服务内容： 图像无法传输到pacs，重新启动主机。 服务结果： 维修后，设备恢复正常工作。 其它重要事项： 无。 Internal Remarks : 19.05.2021 14:17:45 UTC Xiaodong Chi (86038306) "RDI Accuracy:No" $Q7: (Confirmed reported problem&the image can not transfer to pacs ) $Q8: (review the system logger，the device need be restart） $Q9: (restart the device) $Q10: (Test and inspection data refer to: PA. Device returned to full functionality) $Q11 Calibrated Tools name (FLUKE287) S/N (16490094) Next Calibration date (10/19/2021) ----------------------------------------------------------------------------- | Gao, Mia |  | review the system logger，the device need be restart） restart the device) | [12 empty] | 117564872 |  | 921576 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS Transfer |  |  | Completed |  |  | Refresh |  |  | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
937 |  |  | China | 江苏省第二中医院 | 345027 | 44335 |  | Wechat:02N/Ingenuity Core128/Elite China/Rack power down | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : Rack power down Safety Questions : Case: 0117563937 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Cheng Zhenhua Last updated by: Cheng Zhenhua 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 机架倾斜运动故障。 客户问题详细描述或故障现象： 打角度掉电 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： " S_TILT_C_NO_MOVE ERROR_FATAL//S_MPS_MAIN_DRIVE_NOT_CLOSING " 远程诊断与测试结果： "Description: A move request is issued but the Tilt is not moving." PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： "Recommendation: Check the Tilt motors, power supply, and see if anything is in the way of the Tilt. Check for any other S_TILT_ errors." 建议备件12NC及名称： DUAL POT ASSEMBLY 459800441121 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : PPD 459800441121 1已发18H至医院FSE齐良东-86-18662702811 Next Step Action : QI LIANGDONG Internal Remarks : 21.05.2021 08:20:14 UTC Liangdong Qi (86036024) "RDI Accuracy: Yes“ $Q7: (Confirmed reported problem, the MDC was opened when tilt gantry occasionally.) $Q8: (Analysis error report, found "A move request is issued but the Tilt is not moving". Checked the potentiometer belt of tilt, it was too tighten.) $Q9: (Adjusted the belt of tilt potentiometer and did tilt calibration.) $Q10: (Test and Inspection data refer to PA & Device return to fully function) $Q11: Calibrated Tools (NA). External Remarks : 21.05.2021 08:20:08 UTC Liangdong Qi (86036024)服务内容：调整倾斜电位器皮带并校正. 服务结果： 维修后，设备恢复正常工作. ----------------------------------------------------------------------------- |  |  | Adjusted the belt of tilt potentiometer and did tilt calibration. | [12 empty] | 117563937 |  | 921008 |  | Closed - Done | Closed - Done |  | ECR-077012 | [5 empty] | Engineering | [6 empty] | Dual pot |  |  | Completed |  |  | Refresh |  |  | System issue | System | Gantry-HW | Belt | Gantry-HW-Belt | Gantry-HW-Belt | ECR-077012 | Belt | Gantry-HW | Gantry-HW-Belt | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
938 |  |  | China | Jizhong Energy jingxing Mining Group Co., Ltd. General Hospital | 345003 | 44334 |  | PRS connect/02N/Ingenuity Core128/Elite China | BU Q&R Feedback: No Reason: Customer Relations Problem Reported by Customer : PRS connect /02N/Ingenuity Core128/Elite China SAP Transfer:Case Transferred to SAP Safety Questions:Case: 0117559119 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Zhang Jing Last updated by: Zhang Jing 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Dispatched for On site Support Internal Remarks: 19.05.2021 07:15:54 UTC Litu Yang (86038906) RDI accuracy Yes $Q7: (Confirmed the reported problem. install prs) $Q8: (There has no error，just need to reinstall PRS system) $Q9: (Install PRS system) $Q10: (Test and inspection data refer to: PA. Device returned to full functionality.) $Q11: Calibrated Tools name (NA) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date: 09/15/2020 ^Current Tube usage: scan second_33217 External Remarks: 19.05.2021 07:15:51 UTC Litu Yang (86038906) Service Details: [No error, just install the PRS device.] Service results: The device returned to normal working conditions.] ----------------------------------------------------------------------------- | 5 |  | PRS config | [12 empty] | 117559119 |  | 917546 |  |  | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
939 |  |  | China | Fujian Provincial Hospital | 345011 | 44334 |  | 02N/Ingenuity Core128/Elite China/add pacs | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 02N/Ingenuity Core128/Elite China Next Step Action: site1 Li Feng 018650369579 called in Safety Questions:Case: 0117558426 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Bo Last updated by: Wang Bo 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer_s problems: Other types of error. Detailed description of Customer_s problems or malfunctions: Add AI workstation. It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: The log has no errors. Remote troubleshoot and test result: Restart but didn_t work. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Suggest to cooperate with customer to add AI workstation. We recommend spare parts 12NC and names: N/A SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support External Remarks: 19.05.2021 02:13:42 UTC SUP JCOUSER (86036224) service details: mobile CT need to add a network AI analysis node.Work on the field to add this network AI node and modify the scan protocol.CT images can be transmitted to this AI analysis node normally. service result: device returns to normal working condition Internal Remarks : 19.05.2021 02:13:44 UTC Feng Li (86036224) RDI Accuracy: Yes $Q7: (Confirmed reported problem. The CT need add an AI node.) $Q8: (On site to add an AI node and change scan protocol.) $Q9: (The image can transfer to AI node.) $Q10: (Test and inspection data refer to: PA.+Device returned to full functionality.) $Q11: Calibrated Tools name (FLUKE 287C Multimeter) S/N (16490128) NextCalibration date (10/20/2021) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date: 09/13/2020 ^Current Tube Usage: Scan Second_10375 ----------------------------------------------------------------------------- | 5 |  | add pacs | [12 empty] | 117558426 |  | 917094 |  |  | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
940 |  |  | China | Shandong Province Chest Hospital | 40025 | 44334 |  | 02N/Ingenuity CT 728326// ASP engineer ask to nemo filter | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity CT 728326 Safety Questions : Case: 0117557592 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Jianqiang Last updated by: Wang Jianqiang 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | Zhai, Lisa |  | NEMO | [12 empty] | 117557592 |  | 916502 |  |  | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
941 |  |  | China | Zhejiang Province Yueqing No. 1 People’s Hospital | 40054 | 44334 |  | WeChat:02N/Ingenuity CT 728326/Frame without electricity | BU Q&R Feedback: No SAP Customer Symptom Code: HANG Customer Symptom Code: Hang/Crash Reason: Corrective Maintenance Engineer Failure Code: Wear-out/Breakage/Mechanical Failure Repair Activity Code: Reboot/Restart Problem Reported by Customer : Frame without electricity T2 Activities : Problem Description: gantry can not boot up Troubleshooting Action: gantry can not boot up Repair Action: Assist the customer to restart the machine Test/Inspection: customer confirm can scan patient and image ok.functional Check PASSED. Device returned to full functionality. Test Equipment: N/A Safety Questions : Case: 0117557549 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Bo Last updated by: Wang Bo 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | Zhai, Lisa |  | Assist the customer to restart the machine | [12 empty] | 117557549 |  | 916466 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD_CouchEstop | [5 empty] | Engineering | [6 empty] | E-Stop can not close |  |  | Completed |  |  | Refresh |  |  | System issue | System | Couch-SW | TBD-bugrep | Couch-SW-TBD-bugrep | Couch-SW-TBD-bugrep | TBD_CouchEstop | TBD_Bugreport | Couch-SW | Couch-SW-TBD-bugrep | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
942 |  |  | China | 南昌大学第一附属医院(井冈山医院) | 345002 | 44333 |  | WeChat：02N/Ingenuity Core128/Elite China/can not run | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Calibration/Adjustment Safety Questions : Case: 0117553262 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Cao Last updated by: Liu Cao 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Problem Reported by Customer : WeChat：02N/Ingenuity Core128/Elite China/can not run T2 Activities : Problem Description: 客户问题概述： 账户锁死，已到场 处理 客户问题详细描述或故障现象： 账户锁死，已到场 处理 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 账户锁死，已到场 处理 远程诊断与测试结果： 记录维修 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 反馈账户锁问题已到场处理，开单记录 Test/Inspection: N/A Test Equipment: N/A SAP Transfer : Dispatched for Onsite Support Next Step Action : yan xuehui SAP Transfer : Case Transferred to SAP External Remarks : 17.05.2021 17:25:21 UTC SUP JCOUSER (86038347)服务内容：账户过期锁定，重置账户；软件卡顿，传输队列堵住，清理队列后恢复。 服务结果：维修后，设备恢复正常工作。 Internal Remarks : 20.05.2021 00:29:23 UTC Xuehui Yan (86038347) RDI Accuracy:Yes $Q7: Confirmed reported problem can not run, additionally it was the account locked due to expired & software running slowly. $Q8: Check the CT and IT user login, indicate that expired; Check the queue, found the queue was stuck. $Q9: Reset Host account and reset spooler. $Q10: Test and inspection data refer to: PA. Device returned to full functionality. $Q11: Calibrated Tools name (N/A) S/N (N/A) Next Calibration date (N/A). ^Current Tube installed date: MM/DD/ YYYY. ^Current Tube usage: scan second_N/A. ----------------------------------------------------------------------------- | Zhai, Lisa |  | 服务内容：账户过期锁定，重置账户；软件卡顿，传输队列堵住，清理队列后恢复。 服务结果：维修后，设备恢复正常工作。 | [12 empty] | 117553262 |  | 913282 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | User account lock + PACS Transfer |  |  | Completed |  |  | Refresh |  |  | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
943 |  |  | China | 内蒙古兴安盟人民医院 | 345030 | 44333 |  | PRS Connection/02n/Ingenuity Core128/Elite China | BU Q&R Feedback: No Reason: Customer Relations Next Step Action : FSE：苏瑞杰-13304719335 call in and disp him Problem Reported by Customer : PRS Connection Safety Questions : Case: 0117550653 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Lili Last updated by: Lu Lili 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 17.05.2021 03:15:41 UTC Ruijie Su (86039043)服务内容： 1.检查华为4G路由的上网情况。 2.重新配置PRS参数并测试。 服务结果： 设备开机正常，曝光正常，PRS正常。 Internal Remarks : 17.05.2021 03:15:44 UTC Ruijie Su (86039043) $Q1: (Y) $Q2: (N) $Q3: (NA) $Q4: (NA) $Q5: (NA) $Q6: (PRS Connection) $Q7: (Confirmed reported problem, no error ，only PRS Connection) $Q8: (check log,no log，PRS Connection) $Q9: (check the router，reconfigure the PRS) $Q10: (Test and inspection data refer to: PA. Device returned to full functionality.) $Calibrated Tools :NA ----------------------------------------------------------------------------- | 5 |  | PRS config | [12 empty] | 117550653 |  | 911488 |  |  | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
944 |  |  | China | DaKuai Town Health, Fengquan District, Xinxiang City, Henan Province | 345009 | 44333 |  | PRS connection/02N/Ingenuity Core128/Elite China | BU Q&R Feedback: No Reason: Customer Relations Problem Reported by Customer:PRS Connection Safety Questions:Case: 0117550359 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Qi Xiaoyu Last updated by: Qi Xiaoyu 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Dispatched for On site Support External Remarks: 18.05.2021 04:05:48 UTC SUP JCOUSER (86038308) Fault Description: There is no malfunction, link PRS device Service Details: There is no malfunction, link PRS device Service result: After the service, the equipment starts as normal other issues: none. Internal Remarks: 19.05.2021 02:24:21 UTC Haibin Guo (86038308) RDI Accuracy: N/A Q7:conform report problem and just PRS connection $Q8:no CM action and just PRS connection $Q9:Replace the 4G router. $Q10: No test and inspection data. Device returned to full functionality. $Q11:Calibrated Tools N/A SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | 5 |  | PRS config | [12 empty] | 117550359 |  | 911320 |  |  | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
945 |  |  | Germany | Gesundheit Nord gGmbH Klinikverbund | 40002 | 44333 |  | Error message: emergency stop switch | Work Order #: WO-03501207 Eduard Berthold The system fully meets specification and returned to use Problem Description by Engineer: Callback to Raik Glogner-Umbeer: Error message emergency stop switch cannot be closed. Resolution: The following activities were carried out for error analysis / troubleshooting: Error log diagnosis:> no indications in the log file It is very likely that a tape switch was pressed. Customer briefed! | 4 |  | It is very likely that a tape switch was pressed. Customer briefed! | [12 empty] | 117551170 |  | 908944 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD_CouchEstop | [5 empty] | Engineering | [6 empty] | E-Stop can not close |  |  | Completed |  |  | Refresh |  |  | System issue | System | Couch-SW | TBD-bugrep | Couch-SW-TBD-bugrep | Couch-SW-TBD-bugrep | TBD_CouchEstop | TBD_Bugreport | Couch-SW | Couch-SW-TBD-bugrep | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
946 |  |  | Japan | Maki Hospital | 341017 | 44336 |  | JS: Investigation of an artifact that appeared on head images. | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: JS: Investigation of an artifact that appeared on head images. Safety Questions:Case: 0117570693 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: SHA Joy Last updated by: SHA Joy 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A OneEMS internal remarks: Yamamoto DB SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Support ----------------------------------------------------------------------------- | 5 |  | TBD | [12 empty] | 117570693 |  | 925897 |  | Investigation | Closed - Done | [7 empty] | TBD | [6 empty] | TBD |  |  | Completed |  |  | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
947 |  |  | Japan | Independent administrative institution National Hospital Organization | 341015 | 44334 |  | 01CS:KKI:After EndStudy, the previous patient_s data remains on the patient registration screen >> I asked t [Truncated] | [Phenomenon confirmation] Confirmed [Details of work] Call content After EndStudy, registration error occurs because the previous patient data remains on the patient registration screen. Correspondence Request that the RIS side handle it or recommend an examination after confirming it when registering a patient. [Confirmation details] Good | 5 |  | Request that the RIS side handle it or recommend an examination after confirming it when registering a patient. | [12 empty] | 117558498 |  | 917153 |  | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | registration |  |  | Completed |  |  | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
948 |  |  | Poland | Government Agency. | 341010 | 44337 |  | Connection of the syringe. | Work Order #: WO-03530711 21-May-2021 Mr. Robert Kamola Diagnostic performed by Engineer : N/A Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: ----------------------------------------------------------------------------- | 4 |  | Connection of the syringe. | [12 empty] | 117578658 |  | 931416 |  | Evaluation | Closed - non-complaint | [7 empty] | TBD | [6 empty] | NAC |  |  | Completed |  |  | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
949 |  |  | Poland | Government Agency. | 341010 | 44336 |  | Parts Supply Only | Work Order #: WO-03526538 Parts Consumed : 1.00000,459800042241,Cushion-Head/ Armrest ----------------------------------------------------------------------------- | 5 |  | Connection of the syringe. | [12 empty] | 117574335 |  | 928259 |  | Closed - Non-Complaint | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
950 |  |  | Spain | S.C.S. HOSP. GRAL. DE FUERTEVENTURA | 40020 | 44336 |  | The images are not transferred. | They confirm to me in later days that the transfers to the Pacs, from the Tac, are working normally. The transfer queue is unlocked remotely to release the shipments. They did not pass the studies towards the Pacs. |  |  | The transfer queue is unlocked remotely to release the shipments. | [12 empty] | 117572758 |  | 927512 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS Transfer |  |  | Completed |  |  | Refresh |  |  | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
951 |  |  | Turkey | ADIYAMANOZELSAGINSTURTASSANVETICA.S | 341006 | 44337 |  | sarı ekran problemi | Work Order #: WO-03530557 Soner Saglam The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Device freezes intermitent SW and gives yellow screen. Follow Up Required: The process from Helpdesk will be applied. Resolution: The process from Helpdesk will be applied. ----------------------------------------------------------------------------- |  |  | Recommended work around procedure for freezing of the device has been carried out. The device has been delivered in working condition. | [12 empty] | 117578514 |  | 926651 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632520 | [5 empty] | Engineering | [6 empty] | New Yellow Screen |  |  | Completed |  |  | Refresh |  |  | System issue | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
952 |  |  | Turkey | ADIYAMANOZELSAGINSTURTASSANVETICA.S | 341006 | 44336 |  | HCOR kalibrasyonu | Work Order #: WO-03524594 Emre Gül The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : The user was given support on the phone and the malfunction was examined. In order to increase the quality of the image, it is necessary to clean the bore blood and opaque material and make the HCOR calibration. An appointment will be made with the customer for this job. The device is operational. Follow Up Required: Bore / Aplane will be cleaned from blood and opaque material and HCOR calibration will be done. ----------------------------------------------------------------------------- |  |  | Bore / Aplane will be cleaned from blood and opaque material and HCOR calibration will be done. | [12 empty] | 117572904 |  | 924254 |  | Evaluation - Pending Translation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | opaque material |  |  | Completed |  |  | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
953 |  |  | Turkey | ADIYAMAN OZEL SAGINSTURTAS SAN VE TIC A.S | 341006 | 44285 |  | No connection between the CT scanner and monitor; Replaced the HDMI cable |  | CFM |  | No connection between the CT scanner and monitor; Replaced the HDMI cable | [14 empty] | 921401 |  | Closed - Done | Closed - Done | [7 empty] | Customer | [6 empty] | DVI Adapter |  |  | Completed |  |  | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
954 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 337000 | 44333 |  | Table movement is difficult and noisy. |  |  |  | The places that needed to be oiled were oiled and the movements were made again, and it was observed that the sound was gone. The device was delivered actively by shooting trials with the Phantom. | [12 empty] | 117553681 |  | 909482 |  | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Couch Noise |  |  | Completed |  |  | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
955 |  |  | United States | Philips Medical Capital | 40028 | 44335 |  | Questions during Biopsy |  | 5 |  | Q&A | [12 empty] | 117568585 |  | 919980 |  | Closed - Non-Complaint | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
956 |  |  | United States | Lackey Memorial Hospital | 40015 | 44245 |  | Found system grounding issue during service; Upsized the ground and installed a separate ground rod to resolve. |  | CFM |  | FSE along with the hospital electrician upsized the ground and installed a separate rod to correct the grounding issue.The device works as specified after the correction. | [14 empty] | 908786 |  | Closed - Done | Closed - Done | [7 empty] | Customer | [6 empty] | Grounding |  |  | Completed |  |  | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
957 |  |  | United States | Advent Health Wesley Chapel | 300106 | 44336 |  | all monitors (3) blanked out during scanning |  | CFM |  | Received bugreport and timestamp of issue happened. System works well after replace Host PC. | [12 empty] | CFM |  | 922906 |  | Closed - Duplicate | Closed - Duplicate | [14 empty] | Black screen | [5 empty] | Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
958 |  |  | Netherlands | Stichting Rijnstate Ziekenhuis | 52101 | 44336 | 20 | Host computer is frozen. | Work Order #: WO-03522547 JOHNNY BASTIAENS The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : The Host computer froze after logging in repeatedly with wrong information. Resolution : Had the user log in with the admin account of the hospital. | 1 |  | Host PC is blocked after repeated login with wrong credentials Let the user log in with the hospital_s admin account | [12 empty] | 117570919 |  | 921414 |  | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Password |  |  | Complete |  |  | Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
959 |  |  | United States | AdventHealth Zephyrhills | 300106 | 44334 |  | Screens go blank and then back on | Work Order #: WO-03512405 Boyd Craig The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: Screens go blank and then back on User Impact: None Patient Impact: None Current Software Version: Diagnostic performed by Engineer : Screens go blank and then back on. FSE is aware of issue and requested customer call in for case number. PRS accessed. ULV: No normal errors seen at time of issue 10:00 AM. Follow Up Required : FSE on site to evaluate Host video connections. Recommendations: Check all video connections. Replace video cables or video card if needed.Display Port to DVI-I cable adapter 455300015471, nVidea Quadro FX580 Display Card 455300004321. ----------------------------------------------------------------------------- Work Order #: WO-03512457 20-May-2021 Sher Mark The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: Screens go blank and then back on User Impact: None Patient Impact: None Current Software Version: Diagnostic performed by Engineer : due to heavy workload no access to the machine was provided BIU you requested different parts be installed Follow Up Required : replace cable Parts Consumed : 2.00000,455300015471,Display Port to DVI-I cable adapter 3.00000,452213232132,CABLE BH-X0 - OCD ----------------------------------------------------------------------------- Work Order #: WO-03526065 22-May-2021 Sher Mark Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: Screens go blank and then back on User Impact: None Patient Impact: None Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 3 |  | Parts Consumed : 2.00000,455300015471,Display Port to DVI-I cable adapter 3.00000,452213232132,CABLE BH-X0 - OCD | [12 empty] | 117561516 |  | 914886 |  | Closed - Done | Closed - Done |  | CTUDT00632173 | [5 empty] | Engineering | [6 empty] | DVI Adapter |  |  | Complete |  |  | Refresh |  |  | Upgrade issue | System | Console-HW | Adaptor | Console-HW-Adaptor | Console-HW-Adaptor | CTUDT00632173 | Adaptor | Console-HW | Console-HW-Adaptor | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
960 |  |  | United States | Christ Hospital | 300205 | 44343 |  | Errors during scan due to xray tube arcing; Advised Biomed to check the candlestick and run long tube conditioning. | Work Order #: WO-03563437 Morrow, Dan The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Problem description by engineer : Errors during scan When was the first occurrence? Within... : 27-May-2021 How often is issue occurring? : First occurrence Malfunction area : Tube&Gen Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected results: Error Patient Impact/Outcome: none User Impact: Retry scan Current Software Version: 4.1.7.25026 Diagnostic performed by Engineer : Reviewed logs. XRT arcing event. Advised Biomed to check the candlestick wells and run LTC. If issue persists replace the XRT. Biomed to perform work on system. Internal Comments : Error # and/or description of error : S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT S_GEN_FAULT_ANODE_TUBE_ARC S_GEN_FAULT_CATHODE_TUBE_ARC Completed or Recommended Repair Action "Internal" : Reviewed logs. XRT arcing event. Advised Biomed to check the candlestick wells and run LTC. If issue persists replace the XRT. Biomed to perform work on system. ----------------------------------------------------------------------------- | 3 |  | Advised Biomed to check the candlestick wells and run LTC. If issue persists replace the XRT. Biomed to perform work on system. | [12 empty] | 117610856 |  | 948774 |  | Closed - Done | Closed - Done |  | CTUDT00630085 | [5 empty] | Engineering | [6 empty] | Tube arc |  |  | Complete |  |  | Refresh |  |  | Upgrade issue | System | IC-HW | Tub Arc | IC-HW-Tub Arc | IC-HW-Tub Arc | CTUDT00630085 | Tub Arc | IC-HW | IC-HW-Tub Arc | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
961 |  |  | United States | Select Specialty Hospitals | 310051 | 44342 |  | how to set contrast time | Work Order #: WO-03556318 Curry, Dyan The system fully meets specification and returned to use Diagnostic performed by Engineer : customer does not have password for ECM . Customer says she will wait to speak to April Panos, Clin Ed for password . OK to close w/o Problem description by Clinical Support Specialist : Customer Function/ Role: technologist Device Use: in use name Expected and Unexpected Results: how to add contrast to exam card. User Impact: none Patient Impact: none Current Software version: 4.1.7 | 4 |  | 44342 | [12 empty] | 117603726 |  | 948676 |  | Closed - Non-Complaint | Closed - Non-Complaint | [14 empty] | Password | [5 empty] | Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
962 |  |  | United States | Select Specialty Hospitals | 310051 | 44341 |  | CT Brain Repeat for motion | Work Order #: WO-03548032 Mujica, David The system fully meets specification and returned to use Problem description by Clinical Support Specialist : Customer Function/Role: Technologist Device Use: In use Expected and Unexpected Results: Repeat scan due to motion User Impact: none Patient Impact: none Current Software Version: 4.1 Diagnostic performed by Engineer : Voncile indicated that she needed to repeat a portion of a CT head scan due to patient motion and wanted to know the quickest way to do so. Repeat was explained, z location, and drag explored. A simulation was done to increase understanding and application. Voncile successfully repeated a mock scan using the repeat function. Voncile had no further questions or need for CT clinical support. | 4 |  | Repeat scan due to motion | [12 empty] | 117595808 |  | 938746 |  | Closed - Done | Closed - Done | [7 empty] | Others | [6 empty] | Patient Motion |  |  | Complete |  |  | Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
963 |  |  | United States | Select Specialty Hospitals | 310051 | 44340 |  | how to get table location | Work Order #: WO-03541643 Curry, Dyan The system fully meets specification and returned to use Problem description by Clinical Support Specialist : Customer Function/ Role: technologist Device Use: in use name Expected and Unexpected Results: how to plan scan location User Impact: none Patient Impact: none Current Software version: 4.4 Diagnostic performed by Engineer : advised customer to scroll though previous images for scan location needed. She can plan new scan and insert needed numbers in start/stop location. | 4 |  | 44340 | [12 empty] | 117589186 |  | 934985 |  |  | Closed - Non-Complaint | [14 empty] | Q&A | [5 empty] | Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
964 |  |  | Chile | BANCO DE CHILE | 341005 | 44342 | 21 | CL1688 THE SCAN IS ABORTED BEFORE FINISHING DUE TO OVEREXPOSURE | CL1688 THE EXAM IS ABORT BY ITSELF BEFORE ENDING BY OVEREXPOSURE Message indicates tube is very hot: change technique. Resolution Description: Error log is reviewed. A message appears on the screen indicating that the tube is very hot. Falla occurs only with large patients. There is an indicator on the screen that shows that the technique used is not recommended for the patient: asterisks appear in planned slices. The protocol is reviewed with an application specialist and changes are requested in the kV, scan time and filter parameters for this type of patient. Error message on screen, does not allow exam. Information to support the complaint handling process Customer Function / Role: Technician / Sonographer Device Use: Diagnostic Expected and Unexpected results: No errors stopping the exam. Patient Impact / Outcome: Delay in exams. User Impact: Delay in exams. | 2 |  | The protocol is reviewed with an application specialist and changes are requested in the kV, scan time and filter parameters for this type of patient. Collect log and scan parameters for eng investigation | [12 empty] | 117603303 |  | 948395 |  | Opened | Closed - Done |  | Tube Heat | [5 empty] | Engineering | [6 empty] | Tube Overheat |  |  | Completed |  |  | Refresh |  |  | System issue | System | Gantry-SW | Tube heat | Gantry-SW-Tube heat | Gantry-SW-Tube heat | Tube Heat | Tube Heat | Gantry-SW | Gantry-SW-Tube heat | No | y | y | y | y | y | No | y | 0 | y | y | y | y | y | y | 0 | y | y
965 |  |  | China | Inner Mongolia Alxa Right Banner People_s Hospital | 345031 | 44344 |  | 02N/Ingenuity Core128/Elite China/UPS fault | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer:UPS fault Safety Questions:Case: 0117617330 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Zixuan Last updated by: Wang Zixuan 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 3 |  | fse suruijie 告知设备已恢复，不用派工 | [12 empty] | 117617330 |  | 955544 |  | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | UPS |  |  | Completed |  |  | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
966 |  |  | China | 苏州广慈肿瘤医院有限公司 | 345019 | 44344 |  | 02N/Ingenuity Core128/Elite China/password failed | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity Core128/Elite China Safety Questions : Case: 0117614148 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wennian Han Last updated by: Wennian Han 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 密码问题 客户问题详细描述或故障现象： 密码问题 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： NA 远程诊断与测试结果： 重启主机后登录提示密码不对，客户告知没有输错密码。工程师指导客户用IT账户也不能登录。 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 用IST权限重置账户；检查是否软件问题。 客户需求备注： 客户着急需要尽快到场 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 28.05.2021 07:12:11 UTC Haotian Xu (86090280) 服务内容：CT账号被锁定，无法登陆，pacs回传配合修改参数 服务结果：设备运行正常。 ----------------------------------------------------------------------------- | 3 |  | Checked the CT and IT account and password Change IT and CT password according to service manual align with MPM and TSE to highlight the Password policy | [12 empty] | 117614148 |  | 954449 |  | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Application | [6 empty] | Password |  |  | Completed |  |  | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
967 |  |  | China | 苏州广慈肿瘤医院有限公司 | 345019 | 44344 |  | 02N/Ingenuity Core128/scan interrupt | BU Q&R Feedback: No SAP Customer Symptom Code: BOOT Customer Symptom Code: Boot up/Power-up Problem Reason: Corrective Maintenance Engineer Failure Code: Electrical/Electronic Failure Repair Activity Code: Reboot/Restart Problem Reported by Customer : 02N/Ingenuity Core128 Safety Questions : Case: 0117613883 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Weng Jinming Last updated by: Weng Jinming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Troubleshooting Action: instruct customer to restart the system Repair Action: instruct customer to restart the system Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: N/A SAP Transfer : Case Transferred to SAP | 3 |  | instruct customer to restart the system align with MPM and TSE to highlight the restart system regularly | [12 empty] | 117613883 |  | 954305 |  | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Scan |  |  | Completed |  |  | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
968 |  |  | China | 青海省海北州第二医院 | 345029 | 44344 |  | Wechat：02N/Ingenuity Core128/Elite China/Black screen | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : Black screen Safety Questions : Case: 0117613609 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Dai Suyang Last updated by: Dai Suyang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 其他类型故障。 客户问题详细描述或故障现象： 右侧显示器黑屏 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 右侧显示器黑屏 远程诊断与测试结果： 重启无效，一直不亮，显示器本身故障 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 更换显示器 建议备件12NC及名称： 459801140821 19" LCD display monitor T2 Activities : Problem Description: 客户问题概述： 其他类型故障。 客户问题详细描述或故障现象： 右侧显示器黑屏 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 右侧显示器黑屏 远程诊断与测试结果： 重启时提示没有视频输入，应该是显卡故障 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 更换显卡 建议备件12NC及名称： 459801536801 NVIDIA Quadro P400 2GB (3)mDP GFX Next Step Action : PPD 459801536801 1数量已核实,SWO产生后发货FSE许永华-86-13327661334 SAP Transfer : Dispatched for Onsite Support SAP Transfer : Case Transferred to SAP External Remarks : 30.05.2021 06:03:47 UTC Yonghua Xu (86001334)放射科CT室 联系人：郭荣寿 联系电话：15309708576 故障现象：右侧显示器黑屏 。 解决方案/维修措施：右侧显示器黑屏,显卡控制面板里面识别不到右侧显示器，和左侧显示器替换数据线及电源线发现显卡接右侧显示器信号输出口数据输出异常 。判断显卡故障。更换显卡。 完成相应校准，完成备份。 服务后设备正常工作。 Internal Remarks : 30.05.2021 06:03:50 UTC Yonghua Xu (86001334) “RDI Accuracy: Yes” $Q7:Confirmation reported problem,The right display is black. $Q8:The right display is black.Replace the data line and power line with the left display and find that the graphics card is connected to the right display signal output data abnormal.Determine graphics card failure. $Q9:Replace the graphics card. $Q10:Test and Inspection data refer to PA , Do a few patient, Observation of a certain period of time. The system is ok. Device returned to full functionality. $Q11:Calibrated tools name（NA）S/N（NA) Next calibration date（NA） ----------------------------------------------------------------------------- | 3 |  | 右侧显示器黑屏,显卡控制面板里面识别不到右侧显示器，和左侧显示器替换数据线及电源线发现显卡接右侧显示器信号输出口数据输出异常 。判断显卡故障。更换显卡。 | [12 empty] | 117613609 |  | 954140 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | NotEngineering-partreplace | [5 empty] | Engineering | [6 empty] | DisplayCard |  |  | Completed |  |  | Refresh | 459801536801 | NVIDIA Quadro P400 2GB (3)mDP GFX | System issue | System | Console-HW | DisplayCard | Console-HW-DisplayCard | Console-HW-DisplayCard | NotEngineering-partreplace | DisplayCard | Console-HW | Console-HW-DisplayCard | No | y | y | y | y | y | No | y | 0 | y | y | y | y | y | y | 0 | y | y
969 |  |  | Japan | Maki Hospital | 341017 | 44343 |  | 02CS:KK: Patient restraint strap order. | BU Q&R Feedback: No Reason: Corrective Maintenance Safety Questions:Case: 0117605937 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: KENICHI KOJIMA Last updated by: KENICHI KOJIMA 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Problem Reported by Customer : 02CS:KK: Patient restraint strap order. T2 Activities:Problem Description: subject equipment: CT Failure location: 02Couch When did the problem first occur (within what period of time)?: Unknown How often does the problem occur?: Unknown At what point in the end user_s work flow did the problem occur?: Other What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem? No.The customer has not created a log file or report. Troubleshooting Action: Result after troubleshooting: Improved Failure location: 02Couch What steps did you perform to diagnose the problem? Patient restraint strap order. OneEMS internal remarks: --Troubleshooting Action-- Result of response: Problem solved. Failure location: 02Couch What steps did you perform to diagnose the problem? Patient restraint strap order. --Repair Action-- Recommended 12Nc: --None-- 453567009951:L-6"PATIENT RESTRAINT STRAP/ 42" QTY:2 SLOC:--None-- --Subject-- 02CS:KK:Person in charge on site: Patient restraint strap order. --Problem Description-- Equipment affected: CT Failure location: 02Couch When did the problem first occur (within what period of time)?: Unknown How often does the problem occur?: Unknown At what point in the end user_s work flow did the problem occur?: Other What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem? No.The customer has not created a log file or report. --4/8Elements-- Q1. Was the reported problem confirmed? Cannot confirm Q2. Were any other problems confirmed? None. Q3. What did you do to resolve the problem?: Patient restraint strap order. Q4: Is the equipment operating normally following the response? Good. Q5. How did the examination finish? None. Q6. Were there any mistaken findings (misdiagnosis, mistaken treatment) due to the problem? None. Q7. Was there any unintended unnecessary exposure? None. Q8. Were there any artifacts? None. SAP Transfer: Case Transferred to SAP SAP Message: Components Added Succesfully to SWO!- SAP Message: LOGI Task added Successfully!- ----------------------------------------------------------------------------- | Yang, Cancan |  | 44343 | [12 empty] | 117605937 |  | 950122 |  |  | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh | 453567009951 | L-6" PATIENT RESTRAINT STRAP/ 42" QTY:2 | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
970 |  |  | United States | Unity Health Systems | 40037 | 44342 |  | 0711CFXXX March 21st issue it initialized for a while until a computer reboot and it is happening again | Work Order #: WO-03555078 Atz, Michael The system fully meets specification and returned to use Problem Description by Engineer : Problem description by engineer : March 21st issue it initialized for a while until a computer reboot and it is happening again When was the first occurrence? Within... : 21-May-2021 How often is issue occurring? : Intermittent Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: No further information is available as per the customer Patient Impact/Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer Current Software Version: Not provided Diagnostic performed by Engineer : I advised the tech to do a daily Host reboot and a weekly CIRS/Gantry reboot. Tech power-cycled the Host/IRS computers and gantry at wall breaker. DAYS_SinceGantryPwrCycle 56.549849537037 I spoke to the FSE tom Lockwood and he has concurred that a weekly Gantry reboot is the issue Resolution : I advised the tech to do a daily Host reboot and a weekly CIRS/Gantry reboot. Tech power-cycled the Host/IRS computers and gantry at wall breaker. DAYS_SinceGantryPwrCycle 56.549849537037 I spoke to the FSE tom Lockwood and he has concurred that a weekly Gantry reboot is the issue Internal Comments : Error # and/or description of error : DAYS_SinceGantryPwrCycle 56.549849537037 Completed or Recommended Repair Action "Internal" : I advised the tech to do a daily Host reboot and a weekly CIRS/Gantry reboot. Tech power-cycled the Host/IRS computers and gantry at wall breaker. DAYS_SinceGantryPwrCycle 56.549849537037 I spoke to the FSE tom Lockwood and he has concurred that a weekly Gantry reboot is the issue ----------------------------------------------------------------------------- | 3 |  | that a weekly Gantry reboot is the issue 459801656021_b the common step by step sheet for customer, make it visible and train customer to follow it | [12 empty] | 117602764 |  | 943369 |  | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Reboot |  |  | Completed |  |  | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
971 |  |  | China | Jiangsu Province No. 2 Chinese Medicine Hospital | 345027 | 44341 |  | Wechat:02N/Ingenuity Core128/Elite China/Gantry stop scanning | BU Q&R Feedback: No SAP Customer Symptom Code: MARF Customer Symptom Code: Monitor Alarm Failure Reason: Corrective Maintenance Engineer Failure Code: Electrical/Electronic Failure Repair Activity Code: Reboot/Restart Problem Reported by Customer : Gantry stop scanning Safety Questions:Case: 0117592397 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Caiqin Zhang Last updated by: Caiqin Zhang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer_s problems: Scanning failed. Detailed description of Customer_s problems or malfunctions: Scan error, scan failed It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: NORMAL ErrNum: S_DMC_COLL_ACOLLZ_TMOT P1:0/P2:0 Remote troubleshoot and test result: The collimator aperture movement timed out. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: 1. After restarting the gantry, the problem disappeared. Test/Inspection: N/A Test Equipment: N/A Next Step Action: Q6 = scan failed and report error SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | After restarting the gantry, the problem disappeared. | [12 empty] | 117592397 |  | 941082 |  | Investigation | Investigation |  | CollimatorTimerOout | [5 empty] | Engineering | [6 empty] | Collimator Timeout |  |  | Completed |  |  | Refresh |  |  | System issue | System | Gantry-SW | CollimatorTimerOout | Gantry-SW-CollimatorTimerOout | Gantry-SW-CollimatorTimerOout | CollimatorTimerOout | CollimatorTimerOout | Gantry-SW | Gantry-SW-CollimatorTimerOout | No | y | y | y | y | y | No | y | 0 | y | y | y | y | y | y | 0 | y | y
972 |  |  | China | Hebei Province Tangshan City Workers Hospital | 40044 | 44341 |  | 02N/Ingenuity CT/Voice interference | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 02N/Ingenuity CT T2 Activities:Problem Description: Summary of customer_s problems: Other types of error. Detailed description of Customer_s problems or malfunctions: The customer talked from CT box, there was a noise from gantry. CT box buttons failed. It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: The log has no errors. Remote troubleshoot and test result: Restart but didn_t work. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Suggest to replace CT box and adjust sound volume appropriately. We recommend spare parts 12NC and names: PPD 459801705022 SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Next Step Action: PPD 459801705022 1 This spare part is out of stock and needs to be ordered FSE Hu Yue-86-13703597273 Safety Questions:Case: 0117590842 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Bo Last updated by: Wang Bo 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 3 |  | 客户从CT BOX说话，机架发出时有杂音，CT box 按键失灵。 建议更换CT box并适当调节音量大小。 建议备件12NC及名称： PPD 459801705022 CT Box 459800367081 - Noise Reduction Gantry Audio Board Assy | [12 empty] | 117590842 |  | 940255 |  | Evaluation - Pending Translation | Closed - No Further Investigation |  | CTUDT00632185 | [5 empty] | Engineering | [6 empty] | CT Box |  |  | Completed |  |  | Refresh | 459801705022 | CT Box | System issue | System | Gantry-HW | CT Box | Gantry-HW-CT Box | Gantry-HW-CT Box | CTUDT00632185 | CT Box | Gantry-HW | Gantry-HW-CT Box | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
973 |  |  | China | 扎兰屯市人民医院 | 40045 | 44342 |  | PRS connection/02N/Ingenuity CT 728326 | BU Q&R Feedback: No Reason: Customer Relations Problem Reported by Customer : PRS connection Safety Questions : Case: 0117598973 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingli Last updated by: Lu Mingli 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action : FSE：苏瑞杰-13304719335 call in and disp him FSE Notes : This PO is a Reference Number only SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 27.05.2021 06:29:08 UTC Ruijie Su (86039043)服务内容： 1.检查华为4G路由的上网情况。 2.重新配置PRS参数并测试。 服务结果： 设备开机正常，曝光正常，PRS正常。 Internal Remarks : 27.05.2021 06:29:11 UTC Ruijie Su (86039043) $Q1: (Y) $Q2: (N) $Q3: (NA) $Q4: (NA) $Q5: (NA) $Q6: (PRS Connection) $Q7: (Confirmed reported problem, no error ，only PRS Connection) $Q8: (check log,no log，PRS Connection) $Q9: (check the router，reconfigure the PRS) $Q10: (Test and inspection data refer to: PA. Device returned to full functionality.) $Calibrated Tools :NA ----------------------------------------------------------------------------- | Lu, Xueying |  | 44342 | [12 empty] | 117598973 |  | 945379 |  |  | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
974 |  |  | Japan | Maki Hospital | 341017 | 44343 |  | Via Nakagawa in Applications.Regarding a band artifact on head images. | ① How should I think about the position setting when shooting plastic surgery (limbs)? (2) When shooting the elbow, there are two shooting methods, raising the arm and lowering the arm. What are the shooting conditions? ③ A black band appears inside the bone in the data of the southern condition when the sinuses are photographed. App person in charge answer (1) Decide which one to position in front and how to position it separately from the trunk. For example, in the case of elbows, the inside is the front and the upper arm is the head. (2) If the part other than the elbow enters the imaging range, the dose will be changed according to the DRI setting. I think that the current setting (DRI25) is okay, but please raise it a little and see the situation at about 26. ③ I checked the image. There is a dark band because the BHC function is not used. Select a function such as UA or UB. ④ Others: It seems that a dark band appears in the Mask and Delay images of the head CTA. It seems that the original data is also used for diagnosis, so I changed it to UB. I explained Dynaminc_Viewer and Brain_Viewer. | 5 |  | NAC? | [12 empty] | 117605901 |  | 950101 |  | Evaluation | Closed - Done | [7 empty] | Application | [6 empty] | Q&A |  |  | Completed |  |  | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
975 |  |  | United States | Philips Medical Capital | 40055 | 44344 |  | 0998CDY73 FSE Rich is aware ring artifact over the weekend there is contrast in the machine | Work Order #: WO-03568810 Atz, Michael The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Problem description by engineer : FSE Rich is aware ring artifact over the weekend there is contrast in the machine, he wants to have it scheduled for Wednesday June 2nd When was the first occurrence? Within... : 28-May-2021 How often is issue occurring? : Multiple times per week Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: No further information is available as per the customer Patient Impact/Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer Current Software Version: Not provided Diagnostic performed by Engineer : The customer states that the ring goes away once the images are processed, FSE Rich is aware ring artifact over the weekend there is contrast in the machine, he wants to have it scheduled for Wednesday June 2nd Internal Comments : Error # and/or description of error : none Completed or Recommended Repair Action "Internal" : The customer states that the ring goes away once the images are processed, FSE Rich is aware ring artifact over the weekend there is contrast in the machine, he wants to have it scheduled for Wednesday June 2nd Follow Up Required : Error # and/or description of error : none Completed or Recommended Repair Action "Internal" : The customer states that the ring goes away once the images are processed, FSE Rich is aware ring artifact over the weekend there is contrast in the machine, he wants to have it scheduled for Wednesday June 2nd ----------------------------------------------------------------------------- Work Order #: WO-03568890 Turley Rich Problem Description by Engineer : Problem description by engineer : FSE Rich is aware ring artifact over the weekend there is contrast in the machine, he wants to have it scheduled for Wednesday June 2nd When was the first occurrence? Within... : 28-May-2021 How often is issue occurring? : Multiple times per week Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: No further information is available as per the customer Patient Impact/Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer Current Software Version: Not provided Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 3 |  | there is contrast in the machine | [12 empty] | 117617633 |  | 952541 |  | Investigation | Investigation |  | ECR-075975 | [5 empty] | Engineering | [6 empty] | artifact |  |  | Completed |  |  | Refresh |  |  | System issue | System | Gantry-HW | Lexan Ring | Gantry-HW-Lexan Ring | Gantry-HW-Lexan Ring | ECR-075975 | Lexan Ring | Gantry-HW | Gantry-HW-Lexan Ring | No | y | y | y | y | y | No | y | 0 | y | y | y | y | y | y | 0 | y | y
976 |  |  | Japan | Independent administrative institution National Hospital Organization | 341015 | 44341 |  | 07CS:TM: ECG monitor lead wire procurement. >> | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer:07CS:TM: ECG monitor lead wire procurement. >> SAP Transfer: Case Transferred to SAP Safety Questions:Case: 0117593113 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: TAKEYUKI MIYAKAWA Last updated by: TAKEYUKI MIYAKAWA 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Message: Components Added Succesfully to SWO!- SAP Message: LOGI Task added Successfully!- ----------------------------------------------------------------------------- | Zhu, Chun |  | 44341 | [12 empty] | 117593113 |  | 941487 |  |  | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
977 |  |  | United States | Philips Medical Capital | 40055 | 44342 |  | 0711CFXXX One study not going to PACS | Work Order #: WO-03556434 Boyd Craig The system fully meets specification and returned to use Diagnostic performed by Engineer : One study not going to PACS. PACS issues on 05-21-21. Patient information manually entered at that time. All other studies going to PACS. Unable to connect to target message when trying to send study. Out of resources. Bring applications, David Mujica, on call. Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: One study not going to PACS User Impact: None Patient Impact: None Current Software Version: Resolution : Found patient header information differed and made changes. Study went to PACS after changes make. ----------------------------------------------------------------------------- Work Order #: WO-03556488 Mujica, David The system fully meets specification and returned to use Problem description by Clinical Support Specialist : Customer Function/Role: Technologist Device Use: In use Expected and Unexpected Results: CT study not going to PACS User Impact: none Patient Impact: none Current Software Version: 4.1 Diagnostic performed by Engineer : Brent indicated that a CT study that was done on May 22nd is not going to PACS from the CT scanner. The error says "resource allocation error, source not mapped" Reconstruction form RAW data performed to possibly send the new reconstructed series to PACS, this was unsuccessful and the error did manifest again. The accession number was adjusted in patient details and changed to the correct accession number. After this was performed the study did successfully sent to PACS. Brent had no further questions or need for further Clinical support. No rescan, no loss of data, no reinjection as stated by the customer. ----------------------------------------------------------------------------- | 3 |  | The accession number was adjusted in patient details and changed to the correct accession number. After this was performed the study did successfully sent to PACS. | [12 empty] | 117603805 |  | 944143 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS Transfer |  |  | Completed |  |  | Refresh |  |  | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
978 |  |  | China | Medical center (contagious disease, He nan, Xu chang) | 40047 | 44341 |  | 02N/Ingenuity CT 728326/stand by | BU Q&R Feedback: No Reason: Customer Relations Problem Reported by Customer: stand by Safety Questions:Case: 0117591335 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Zhang Jing Last updated by: Zhang Jing 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Case Transferred to SAP External Remarks: 26.05.2021 08:45:15 UTC SUP JCOUSER (86038955) Service Details: The hospital conducted an environmental evaluation and the field service engineer is on standby on site. Service result: After the service, the equipment starts as normal Others: None as of now. Internal Remarks:26.05.2021 08:45:18 UTC Yi Zhang (86038955) RDI Accuracy:Yes $Q7: (Confirm the problem, standby . ) $Q8: (Cooperate with hospital environmental impact assessment by standby. ) $Q9: (Cooperate with hospital environmental impact assessment by standby, then the issue was solved! ) $Q10: (Test and Inspection data refer to PA. Device returned to full functionality.) $Q11: (Calibrated Tools :N/A) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date: 00/ 00/ 0000 ^Current Tube usage: SAP Transfer:Dispatched for On site Support ----------------------------------------------------------------------------- | Gao, Mia |  | 44341 | [12 empty] | 117591335 |  | 940507 |  |  | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
979 |  |  | China | Shandong Anqiu City Traditional Chinese Medicine hospital | 345017 | 44343 |  | 02NIngenuity Core128/Elite China/Image reconstruction failure | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 02N Ingenuity Core128/Elite China/ Safety Questions:Case: 0117605512 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingsong Last updated by: Lu Mingsong 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer_s problems: Image reconstruction failed. Detailed description of Customer_s problems or malfunctions: Intermittently unable to reconstruct images. Shut down and reboot, then restore It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: N/A Remote troubleshoot and test result: N/A PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Suspect It is software bug, suggest to collect bugrep and open ESCALATION We recommend spare parts 12NC and names: N/A Memo of Customer request: None SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support ----------------------------------------------------------------------------- | 5 |  | Image reconstruction failure: Unable to rebuild images intermittently after scan complete, normal after shutting down and restarting. Customer temperature: between Yellow and Red? Bugreport: https://www.clouddrive.philips.com.cn/file/iwcqt0xve89fuw15bwxxuzmfejk2f1m7# Password:123 | [12 empty] | 117605512 |  | 949841 |  | Evaluation | Closed - Done |  | CIRS-SW-CannotRecon | [5 empty] | Engineering | [6 empty] | Recon |  |  | Completed |  |  | Refresh |  |  | System issue | System | CIRS-SW | CannotRecon | CIRS-SW-CannotRecon | CIRS-SW-CannotRecon | CIRS-SW-CannotRecon | CannotRecon | CIRS-SW | CIRS-SW-CannotRecon | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
980 |  |  | China | 山东省泰安市中医二院 | 345037 | 44340 |  | 02N/Ingenuity Core128/Elite China/change parameter | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: Instructed Customer Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ Safety Questions : Case: 0117585968 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wennian Han Last updated by: Wennian Han 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: change parameter Troubleshooting Action: need change printer parameter Repair Action: guide customer to change printer parameter Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: NA SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | Gao, Mia |  | 44340 | [12 empty] | 117585968 |  | 936400 |  |  | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
981 |  |  | Japan | Maki Hospital | 341017 | 44340 |  | CT:APP:RK: Question: Mr. Noda. | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Application Support Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer : Q: He wanted to know how to create an MPR image from skull bones series in the same way as for brain parenchyma. A: Please use CT Viewer to load the series you want to create an MPR for. Then set the angle and determine the range as you normally would.Then select Save Preset from the Batch preset and save it, and apply this to the next series, and the angle will be the same.However, please keep in mind that the window from the series will be carried over. OneEMS internal remarks: Q1: Did you confirm the reported problem?: Confirmed Q2. Was there any other problem?None Q3. What did you do to resolve the problem?: 1: If resolved by telephone: explanation Q4: Is the equipment operating normally following the response?: Good Safety Questions:Case: 0117585700 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: RYUSUKE KUJIRAI Last updated by: RYUSUKE KUJIRAI 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | Xu, Xiaoting |  | 44340 | [12 empty] | 117585700 |  | 936213 |  |  | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
982 |  |  | United States | Central Vermont Medical Center | 40024 | 44340 |  | 0308CDY73 Vertical movement noise | Work Order #: WO-03540631 Miller, John M The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Problem description by engineer : Customer states the couch vertical movement is making a lot of noise. When was the first occurrence? Within... : 24-May-2021 How often is issue occurring? : Continuously Malfunction area : Table Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Noise in couch movement Patient Impact/Outcome: None User Impact: Noise in couch movement Current Software Version: Not provided Diagnostic performed by Engineer : Collect information from customer. Follow Up Required : Error # and/or description of error : Vertical noise. Completed or Recommended Repair Action "Internal" : Suggest checking for debris in the Vertical movement, as well as covers binding. Suggest checking for proper lubrication of the the vertical drive screw and debris on the Thompson rods. Inspect the vertical spindle for wear and all parts in the vertical drive system. Refer to the Couch PM manual for proper lubrication and Ball screw inspection. Suggest checking the safety bar switch and its wiring. Check / replace limit switches. On Bariatric Couch, check for leaking piston. Clean any oil residue from the piston and replace them. Suggest checking, repairing, cleaning and/or replacing Vertical Encoder O-Ring. Internal Comments : Error # and/or description of error : Vertical noise. Completed or Recommended Repair Action "Internal" : Suggest checking for debris in the Vertical movement, as well as covers binding. Suggest checking for proper lubrication of the the vertical drive screw and debris on the Thompson rods. Inspect the vertical spindle for wear and all parts in the vertical drive system. Refer to the Couch PM manual for proper lubrication and Ball screw inspection. Suggest checking the safety bar switch and its wiring. Check / replace limit switches. On Bariatric Couch, check for leaking piston. Clean any oil residue from the piston and replace them. Suggest checking, repairing, cleaning and/or replacing Vertical Encoder O-Ring. ----------------------------------------------------------------------------- Work Order #: WO-03540665 24-May-2021 Morgan Russ Van The system fully meets specification and returned to use Problem Description by Engineer : Problem description by engineer : Customer states the couch vertical movement is making a lot of noise. When was the first occurrence? Within... : 24-May-2021 How often is issue occurring? : Continuously Malfunction area : Table Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Noise in couch movement Patient Impact/Outcome: None User Impact: Noise in couch movement Current Software Version: Not provided Diagnostic performed by Engineer : noisy couch , cleaned and lubricated ----------------------------------------------------------------------------- | 3 |  | noisy couch , cleaned and lubricated | [12 empty] | 117588255 |  | 934368 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-Service | [5 empty] | Engineering | [6 empty] | Couch noise |  |  | Completed |  |  | Refresh |  |  | System issue | System | Couch-HW | Vertical spring | Couch-HW-Vertical spring | Couch-HW-Vertical spring | TBD-Service | Vertical spring | Couch-HW | Couch-HW-Vertical spring | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
983 |  |  | Chile | Scotiabank Chile | 40008 | 44350 |  | PROBLEM WITH VERTICAL MOVEMENT, GUIDE PISTON OUT OF ADJUSTMENT | Work Order #: WO-03599146 Alec, Preciado The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Problem description by engineer : PROBLEM WITH VERTICAL MOVEMENT, GUIDE PISTON OUT OF ADJUSTMENT When was the first occurrence? Within... : 03-Jun-2021 12:00:00 How often is issue occurring? : First occurrence Malfunction area : Table Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected results: Restore the patient table_s proper operation. Patient Impact/Outcome: It does not impact. User Impact: It does not impact. Current Software Version: 4.1.10 Diagnostic performed by Engineer: Support is provided with the user, it is identified that the system shows noise in the vertical movement, derived from possible loss of mechanical adjustment in the piston. Internal Comments : Error # and/or description of error : NA Completed or Recommended Repair Action "Internal" : We checked the system with the user and it shares the sound_s audio, which is characteristic of a vertical movement piston not being adjusted. It is in the patient table_s base. Follow Up Required : What are the next Troubleshooting steps : It is necessary to check the patient table pistons. By the sound, it indicates that it is not adjusted. If so, perform activities to adjust again and afterwards grease the patient table according to the maintenance manual. Error # and/or description of error : NA Completed or Recommended Repair Action "Internal" : We checked the system with the user and it shares the sound_s audio, which is characteristic of a vertical movement piston not being adjusted. It is in the patient table_s base. ----------------------------------------------------------------------------- Work Order #: WO-03599244 Osorio, Juan Problem Description by Engineer : Problem description by engineer : PROBLEM WITH VERTICAL MOVEMENT, GUIDE PISTON OUT OF ADJUSTMENT When was the first occurrence? Within... : 03-Jun-2021 12:00:00 How often is issue occurring? : First occurrence Malfunction area : Table Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected results: Restore the patient table_s proper operation. Patient Impact/Outcome: It does not impact. User Impact: It does not impact. Current Software Version: 4.1.10 Diagnostic performed by Engineer: N/A ----------------------------------------------------------------------------- | 3 |  | Planning Notes changed from "DL1 CLIENT TO BE CONFIRMED ¦ FSE J OSORIO ¦ CL1675 ¦ TUESDAY 8 JUN TENT" to "DL1 ASSIGNED ¦ FSE J OSORIO ¦ CL1675 ¦ FRIDAY 11 JUN | [12 empty] | 117648281 |  | 971398 |  | Opened | Closed - No Further Investigation | [7 empty] | Engineering | [6 empty] | Couch Noise |  |  | Completed |  |  | Refresh |  |  | System issue | System | Couch-HW | Table Cover | Couch-HW-Table Cover | Couch-HW-Table Cover | 0 | Table Cover | Couch-HW | Couch-HW-Table Cover | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
984 |  |  | Chile | Scotiabank Chile | 40008 | 44350 |  | CL1675-CIRCULAR ARTIFACT IN THE PICTURE AND NOISE IN THE BED | Work Order #: WO-03597437 Alec, Preciado The system fully meets specification and returned to use Problem Description by Engineer : Problem description by engineer : ARTEFACTO CIRCULAR EN LA IMAGEN When was the first occurrence? Within... : 02-Jun-2021 11:00:00 How often is issue occurring? : First occurrence Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected results: imagenes sin problema de artefactos Patient Impact/Outcome: retraso en atencion User Impact: retraso en carga de trabajo Current Software Version: 4.1.7 Diagnostic performed by Engineer : Se realiza revision de tablas de aire en el sistema, se identifica señal no homogenea, se realiza nuevamente una calibración de aire y se revisa imagen, se corrige artefacto. Internal Comments : Error # and/or description of error : NA Completed or Recommended Repair Action "Internal" : calibracion de aire realizada, se corrige artefacto en sistema. Follow Up Required : PROBLEMA CON MOVIMIENTO VERTICAL, PISTON GUIA FUERA DE AJUSTE Resolution : Se realiza revision de tablas de aire en el sistema, se identifica señal no homogenea, se realiza nuevamente una calibración de aire y se revisa imagen, se corrige artefacto. ----------------------------------------------------------------------------- | 3 |  | A review of air tables is performed in the system, a non-homogeneous signal is identified, an air calibration is performed again and an image is reviewed, an artifact is corrected. | [12 empty] | 117646963 |  | 971330 |  | Opened - Pending Translation | Closed - Done |  | NotEngineering | [5 empty] | Engineering | [6 empty] | Artifact |  |  | Completed |  |  | Refresh |  |  | System issue | System | NotEngineering | ShouldbeCustomer | NotEngineering-ShouldbeCustomer | NotEngineering-ShouldbeCustomer | NotEngineering | NotEngineering-ShouldbeCustomer | NotEngineering | NotEngineering-ShouldbeCustomer | No | y | y | y | y | y | No | y | 0 | y | y | y | y | y | y | 0 | y | y
985 |  |  | Chile | Scotiabank Chile | 40008 | 44313 |  | Burning smell during servicing, verified components in rotor-stator and power sector worked properly | Burning smell during servicing, verified components in rotor-stator and power sector worked properly | CFM |  | After confirmed, It is not from CT, but from air condition | [14 empty] | 962707 |  | Closed - Done | Closed - Done | [7 empty] | Others | [6 empty] | Burning smell |  |  | Completed |  |  | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
986 |  |  | Puerto Rico | METRO MAYAGUEZ INC | 40009 | 44351 |  | Makes noise when tube is spinning | Work Order #: WO-03605095 Caldelas Aldana, Carlos Manuel The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Problem description by engineer : noise when rotating When was the first occurrence? Within... : 04-Jun-2021 11:00:00 How often is issue occurring? : Continuously Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: noise when rotating Patient Impact/Outcome: noise in system User Impact: noise in gantry Current Software Version: Not provided Diagnostic performed by Engineer : validated system is having abnormal noise like something hitting inside the gantry when rotating during scans. no errors appear for the moment. Internal Comments : Error # and/or description of error : noise when rotation Completed or Recommended Repair Action "Internal" : n/a Follow Up Required : What are the next Troubleshooting steps : open and inspect gantry to validate what is loose or causing the noise when rotating Error # and/or description of error : noise when rotation Completed or Recommended Repair Action "Internal" : n/a ----------------------------------------------------------------------------- | 3 |  | During the remote support session, it was identified that a field service engineer visit is required to continue with technical support. The planning team will be in touch with you shortly to coordinate the technical visit. | [12 empty] | 117654113 |  | 975210 |  | Opened | Closed - No Further Investigation |  | TBD-GantryNoise | [5 empty] | Engineering | [6 empty] | Gantry Noise |  |  | Completed |  |  | Refresh |  |  | System issue | System | Gantry-HW | Noise | Gantry-HW-Noise | Gantry-HW-Noise | TBD-GantryNoise | Noise | Gantry-HW | Gantry-HW-Noise | No | y | y | y | y | y | No | y | 0 | y | y | y | y | y | y | 0 | y | y
987 |  |  | United States | Lackey Memorial Hospital | 40015 | 44348 |  | customer says there is a noise inside the gantry when using xray tube | Work Order #: WO-03582737 01-Jun-2021 Denison Chuck The system fully meets specification and returned to use Diagnostic performed by Engineer : Checked the Gantry internal. Made adjustments to front cover. installed plate into the bottom of the Gantry. Problem Description by Engineer : Customer Function/Role: Head of x-ray Dept. Device Use: Philips C.T. Expected and Unexpected Results: Scanner is working. User Impact: None system is up. Patient Impact: System is up. Current Software Version:4.1.10 ----------------------------------------------------------------------------- | 4 |  | Checked the Gantry internal. Made adjustments to front cover. installed plate into the bottom of the Gantry. | [12 empty] | 117632842 |  | 961582 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | NotEngineering | [5 empty] | Engineering | [6 empty] | Gantry Noise |  |  | Completed |  |  | Refresh |  |  | System issue | System | NotEngineering | ShouldbeOPS | NotEngineering-ShouldbeOPS | NotEngineering-ShouldbeOPS | NotEngineering | NotEngineering-ShouldbeOPS | NotEngineering | NotEngineering-ShouldbeOPS | No | y | y | y | y | y | No | y | 0 | y | y | y | y | y | y | 0 | y | y
988 |  |  | China | 河南省濮阳市人民医院 | 40031 | 44349 |  | Wechat:02N/Ingenuity CT 728326/Data transmission failure | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: Calibration/Adjustment Problem Reported by Customer : Data transmission failure Safety Questions : Case: 0117638092 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Bai Jingwen Last updated by: Bai Jingwen 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: Cannot transfer image to ISP Troubleshooting Action: Check ISP free space is 0%, need clean Repair Action: Customer will clean data and try again Test/Inspection: N/A Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | Check ISP free space is 0%, need clean Customer will clean data and try again | [12 empty] | 117638092 |  | 968863 |  | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | ISP |  |  | Completed |  |  | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
989 |  |  | United States | Unity Health Systems | 40037 | 44348 |  | Gantry can not comply due to table hardware issue; Performed hardware reset with a power cycle to resolve. | Work Order #: WO-03581616 Ali, Adam The system fully meets specification and returned to use Problem Description by Engineer : Problem description by engineer : Gantry can not comply When was the first occurrence? Within... : 01-Jun-2021 How often is issue occurring? : Continuously Malfunction area : Table Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No further information is available as per the customer Expected and Unexpected results: Not provided Patient Impact/Outcome: No further information is available as per the customer User Impact: Unable to scan Current Software Version: 4.1 Resolution : System went down with a _Gantry cannot comply error_. Bruce made a bugreport & also power cycled the system. The error has cleared up and tests scanned ran with no issues. I dialed into system, checked logs & also downloaded bugreport. Logs shows a table hardware issue caused an obstruction error which was cleared once hardware reset with a power cycle. Internal Comments : Error # and/or description of error : 2021/06/01 05:36:02.838__02 S_HMC_HOR_MOTOR_RELAY_OPEN NORMAL 0 0 3 0 Hor Motor Relay open 5.71928240740741 137.262777777778 8235.76666666667 2021/06/01 05:34:59.635__01 S_HMC_CTRL_A_POSITION_ERROR_LIMIT_EXCEEDED NORMAL 25 0 3 0 Check for couch obstruction 5.71854166666667 137.245 8234.7 2021/06/01 05:28:47.473__02 S_HMC_AXIS_WRONG_DEST_REACHED NORMAL 50223 50270 3 0 EndStatus: 0 5.71424768518519 137.141944444444 8228.51666666667 ----------------------------------------------------------------------------- | 4 |  | System went down with a _Gantry cannot comply error_. Bruce made a bugreport & also power cycled the system. The error has cleared up and tests scanned ran with no issues. Bruce made a bugreport, collect bugreport for further analysis | [12 empty] | 117631872 |  | 959979 |  | Closed - Done | Closed - Done |  | TBD_Bugreport | [5 empty] | Engineering | [6 empty] | Gantry comply |  |  | Completed |  |  | Refresh |  |  | System issue | System | Console-SW | Bugrep | Console-SW-Bugrep | Console-SW-Bugrep | TBD_Bugreport | TBD_Bugreport | Console-SW | Console-SW-Bugrep | No | y | y | y | y | y | No | y | 0 | y | y | y | y | y | y | 0 | y | y
990 |  |  | Sweden | Falu Lasarett | 40049 | 44351 |  | Application crash after reconstruction. Email attached from cust. | Work Order #: WO-03602652 Knudsen Martin The system fully meets specification and returned to use Problem Description by Engineer : Friday: User experienced yellow screen and needed to reboot system. Have worked since whole week. Diagnostic performed by Engineer : No bugreport was created. Philips asks to close case since there is no bugreport. If error comes back, they will give a new call. ----------------------------------------------------------------------------- | 3 |  | Yellow Screen on HF3, sync with L3 to see if need workaround | [12 empty] | 117652074 |  | 973711 |  | Opened - Pending Translation | Closed - Done |  | CTUDT00632520 | [5 empty] | Engineering | [6 empty] | New Yellow Screen |  |  | Completed |  |  | Refresh |  |  | System issue | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
991 |  |  | China | 浙江省乐清第一人民医院 | 40054 | 44349 |  | Wechat:02N/Ingenuity CT 728326/run slowly | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : run slowly Safety Questions : Case: 0117637652 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Tao Fengming Last updated by: Tao Fengming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 系统卡顿 设备本身技术问题： 是。 Troubleshooting Action: 远程诊断与测试结果： 怀疑软件，显卡。。。问题 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议工程师到场检查主机散热，显卡，重新安装软件 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 03.06.2021 09:34:45 UTC SUP JCOUSER (86036249)服务内容：软件死机黄屏，清洁垃圾数据，清理操作台电脑，修改主机到重建柜网段为10网段。 服务结果：维修后，设备恢复正常工作。 其它重要事项：无 Internal Remarks : 03.06.2021 09:34:51 UTC Yi Huang (86036249) RDI Accuracy:Yes $Q7: (Confirmed reported problem&system yellow screen and hung up) $Q8: (check the hosp network IP is 192.168, same to host to irs) $Q9: (clean the console and change network to 10.168 address) $Q10: (Test and inspection data refer to: PA.Device returned to full functionality.) $Q11: Calibrated tools name(N/A) S/N (N/A) Next calibration date (N/A) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date:N/A ^Current Tube usage: second:N/A ----------------------------------------------------------------------------- | 3 |  | $Q7: (Confirmed reported problem&system yellow screen and hung up) $Q8: (check the hosp network IP is 192.168, same to host to irs) $Q9: (clean the console and change network to 10.168 address) | [12 empty] | 117637652 |  | 968645 |  | Opened - Pending Translation | Closed - No Further Investigation |  | CTUDT00632520 | [5 empty] | Engineering | [6 empty] | New Yellow Screen |  |  | Completed |  |  | Refresh |  |  | System issue | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
992 |  |  | China | 浙江省乐清第一人民医院 | 40054 | 44348 |  | 02N/Ingenuity CT 728326/report error | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: Instructed Customer Safety Questions : Case: 0117631325 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Qi Xiaoyu Last updated by: Qi Xiaoyu 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Problem Reported by Customer : 02N/Ingenuity CT 728326 T2 Activities : Problem Description: report error Troubleshooting Action: report error Repair Action: Customer click off the error, and then scan normally. Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: NA SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | Customer click off the error, and then scan normally. | [12 empty] | 117631325 |  | 963986 |  | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Unknown |  |  | Completed |  |  | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
993 |  |  | China | 浙江省乐清第一人民医院 | 40054 | 44348 |  | Wechat:02N/Ingenuity CT 728326/CIRS boot issue | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer : Correct the error Safety Questions : Case: 0117628950 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Tao Fengming Last updated by: Tao Fengming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: Cannot scan or STC Troubleshooting Action: CIRS disconnection, reboot but still failed. Check logger has no error and CIRS ping sucess. Wait and try again, issue solved Repair Action: Wait for CIRS boot and try again, issue solved Test/Inspection: N/A Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | CIRS disconnection, reboot but still failed. Check logger has no error and CIRS ping sucess. Wait and try again, issue solved Repair Action Wait for CIRS boot and try again, issue solved ask FSE to check if Service node setting changed or not, and see if have implement in OPS | [12 empty] | 117628950 |  | 962567 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | NotEngineering | [5 empty] | Engineering | [6 empty] | Cannot scan |  |  | Completed |  |  | Refresh |  |  | System issue | System | NotEngineering | ShouldbeService | NotEngineering-ShouldbeService | NotEngineering-ShouldbeService | NotEngineering | NotEngineering-ShouldbeService | NotEngineering | NotEngineering-ShouldbeService | No | y | y | y | y | y | No | y | 0 | y | y | y | y | y | y | 0 | y | y
994 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 337000 | 44346 |  | Device does not start shooting | Work Order #: WO-03573165 Can Özgenlik The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : Cihazın verdiği hatalar incelendi. Hataların sonucu olarak slipring ile ilgili hatalar görüldü. Öncelike slipring temizlenmeli ve yağlanmalıdır. ----------------------------------------------------------------------------- Work Order #: WO-03573167 03-Jun-2021 Tunay Uysal The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Cihaz çekime başlamıyor Resolution : Yapılan incelemede slipring data iletim hatası görüldü. Slipring bakımı yapıldı. Sistem klinikte bulunan fantom ile test edilerek kullanıma hazır bir şekilde teslim edildi. ----------------------------------------------------------------------------- | 3 |  | In the examination, slipring data transmission error was observed. Slipring maintenance done. The system was tested with the phantom in the clinic and delivered ready to use. Errors given by the device have been examined. Errors related to slipring were seen as a result of the errors. First of all, the slipring should be cleaned and lubricated. | [12 empty] | 117621514 |  | 957525 |  | Evaluation - Pending Translation | Closed - No Further Investigation |  | NotEngineering-SlipringClean | [5 empty] | Engineering | [6 empty] | Slipring |  |  | Completed |  |  | Refresh |  |  | System issue | System | Gantry-HW | SlipRingClean | Gantry-HW-SlipRingClean | Gantry-HW-SlipRingClean | NotEngineering-SlipringClean | SlipRingClean | Gantry-HW | Gantry-HW-SlipRingClean | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
995 |  |  | Turkey | ADIYAMANOZELSAGINSTURTASSANVETICA.S | 341006 | 44350 |  | There is a data issue | Work Order #: WO-03596130 Ugur,Demirbas The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : There are fluctuations in the ECG signal, BPM is OK. Diagnostic performed by Engineer : possible cable problem between gantry and ECG Follow Up Required: If necessary, the cable between the ECG gantry needs to be changed. otherwise, ekg probes, ekg communication line ibox should be checked. It is useful to go with the ECG simulator. | 3 |  | If necessary, the cable between the ECG gantry needs to be changed. otherwise, ekg probes, ekg communication line ibox should be checked. It is useful to go with the ECG simulator. | [12 empty] | 117645950 |  | 969776 |  | Opened - Pending Translation | Evaluation - Pending Translation | [7 empty] | Others | [6 empty] | ECG |  |  | Completed |  |  | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
996 |  |  | Turkey | ADIYAMAN OZEL SAGINSTURTAS SAN VE TIC A.S | 341006 | 44350 |  | Remote control is out of use. | Work Order #: WO-03596142 Ugur, Demirbas The system partly meets specification and is returned to use with restrictions Diagnostic performed by Engineer : Keypad replacement is required. Problem Description by Engineer: Console down button is out of use. Follow Up Required: Keypad replacement. ----------------------------------------------------------------------------- | 3 |  | down button not working Under failuare part anaysis | [12 empty] | 117645969 |  | 969613 |  | Opened | Evaluation |  | CTUDT00632185 | [5 empty] | Engineering | [6 empty] | CT Box |  |  | Completed |  |  | Refresh |  |  | System issue | System | Gantry-HW | CT Box | Gantry-HW-CT Box | Gantry-HW-CT Box | CTUDT00632185 | CT Box | Gantry-HW | Gantry-HW-CT Box | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
997 |  |  | China | Zhejiang Province Yueqing No. 1 People’s Hospital | 40054 | 44348 |  | 02N/Ingenuity CT 728326/equipment information request | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Environmental Cause Repair Activity Code: Instructed Customer Problem Reported by Customer: 02N/Ingenuity CT 728326 T2 Activities:Problem Description: customer ask equipment information. Troubleshooting Action: customer ask equipment information. Repair Action: check the information, and reply to customer. Test/Inspection: system is ok. Test Equipment: N/A Safety Questions:Case: 0117629563 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | 44348 | [12 empty] | 117629563 |  | 962976 |  |  | Closed - Non-Complaint | [14 empty] | 44348 | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
998 |  |  | Japan | 社会医療法人 母恋 日鋼記念病院 | 341011 | 44347 |  | Since after installation, detector failure 5times . Installed date:Dec-28-2020 |  | CFM |  | Duplicate Check customer temp, and whether need BIU support | [14 empty] | 968521 |  | Investigation | Investigation |  | CTUDT00632110 | [5 empty] | Engineering | [6 empty] | Artifact |  |  | Completed |  |  | Refresh |  |  | System issue | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | 0 |  |  | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | y | y | y | y | y | 0 | y | #REF!
999 |  |  | Japan | 社会医療法人 母恋 日鋼記念病院 | 341011 | 44347 |  | 04DS:KKI: Many DMS module failures | Detector failure frequency is high [Phenomenon confirmation] I was able to confirm. [Details of work] Since the frequency of detector failures from the time of installation is high, countermeasures have been taken. We have replaced the motherboard of the data measurement system (DMS). After the replacement, we made various adjustments and confirmed that there was no problem. We checked the operation and conducted a test scan to confirm that there were no problems. [Other failures / treatments] None [Device status] Equipment can be used [Work status] done [Other] Parts used Part number: 459800263491 / Part name: DMS MOTHER BOARD RIGHT / Serial number: ー / Batch number: ー / Quantity: 1 / List price: ー | 5 |  | We have replaced the motherboard of the data measurement system (DMS). After the replacement, we made various adjustments and confirmed that there was no problem. 459800263491 / Part name: DMS MOTHER BOARD RIGHT | [12 empty] | 117625053 |  | 959140 |  | Evaluation | Closed - No Further Investigation |  | DMSMotherboard | [5 empty] | Engineering | [6 empty] | DMS MB |  |  | Completed |  |  | Refresh | 459800263491 | DMS MOTHER BOARD RIGHT | System issue | System | DMS-HW | DMS Motherboard | DMS-HW-DMS Motherboard | DMS-HW-DMS Motherboard | should be DEFOA | DMS Motherboard | DMS-HW | DMS-HW-DMS Motherboard | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
1000 |  |  | United States | Philips Medical Capital | 40005 | 44348 |  | 4.1.7 Brain Perfusion training | Work Order #: WO-03585212 Giles Brian The system fully meets specification and returned to use Problem description by Clinical Support Specialist : Customer Function/Role: Tech Device Use: In Use Expected and Unexpected Results: Wants Brain Perfusion Training User Impact: None Patient Impact: None Current Software Version: 4.1.7 Diagnostic performed by Engineer : Sent customer to Fisher, Kelly Clinical Services Project Manager Cust Svc Clin Ed - IS Southeast Zone Diagnostic performed by Engineer : Contacted Dianne and she was communicating with Kelly Fisher from scheduling. Case closed. ----------------------------------------------------------------------------- | 5 |  | 44348 | [12 empty] | 117634665 |  | 961938 |  |  | Closed - Non-Complaint | [14 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1001 |  |  | Japan | 社会医療法人 母恋 日鋼記念病院 | 341011 | 44346 |  | 01CS: KKI: Air cal. Does not run >> No hardware error> It seems that it did not run due to the problem of bed position. Wait and see | Original Problem Reported by Customer 01CS: KKI: Air cal. Does not run >> No hardware error> It seems that it did not run due to the problem of bed position. Wait and see Problem Description Problem description by engineer: Target device: CT Failure location: 01Cosole When did it first occur? (Within ~): 1 Hour How often do problems occur ?: First Occurrence At what point in the end-user workflow is the problem occurring ?: Before inspection / measurement starts (Pre-Acquistion / Scan) What happened recently ?: N.A. Does the customer create a specific system event report for the problem ?: No. The customer was unable to create a log file or report. Troubleshooting Action What steps are taken during the system diagnosis: Post-response results: Improved Failure location: 01Cosole What steps were taken during the diagnosis ?: Call content Air cal. Doesn_t run Correspondence As a result of log confirmation, there is no hardware error It seems that he did not run due to the problem of bed position. We have decided that it is not a problem and will wait and see. Error # and / or description of error: N / A Repair Action Completed or Recommended Repair Action "External": [Phenomenon confirmation] Confirmed [Details of work] Call content Air cal. Doesn_t run Correspondence As a result of log confirmation, there is no hardware error It seems that he did not run due to the problem of bed position. We have decided that it is not a problem and will wait and see. [Confirmation details] Good | 4 |  | there is no hardware error It seems that he did not run due to the problem of bed position. We have decided that it is not a problem and will wait and see. | [12 empty] | 117622122 |  | 958191 |  | Closed - No Further Investigation | Closed - No Further Investigation |  | NotEngineering | [5 empty] | Engineering | [6 empty] | Air cal. |  |  | Completed |  |  | Refresh |  |  | System issue | System | NotEngineering | ShouldbeOthers | NotEngineering-ShouldbeOthers | NotEngineering-ShouldbeOthers | NotEngineering | NotEngineering-ShouldbeOthers | NotEngineering | NotEngineering-ShouldbeOthers | No | y | y | y | y | y | No | y | 0 | y | y | y | y | y | y | 0 | y | y
1002 |  |  | Japan | 社会医療法人 母恋 日鋼記念病院 | 341011 | 44346 |  | 04DS:KKI: Ring artifact appeared >> Appeared near M18 or M25 > Did not appear in subsequent test scan | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 01DS: KKI: Artifact appeared Safety Questions:Case: 0117622065 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: KAZUMASA KIMURA Last updated by: KAZUMASA KIMURA 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Support SAP Message: Components Added Succesfully to SWO!- SAP Message: LOGI Task added Successfully!- T2 Activities:Problem Description: Problem description by engineer : Equipment affected: CT Problem location: 04 Gantry Exclude HV When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user_s work flow did the problem occur?: Post-Acquisition/Scan What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem? No.The customer has not created a log file or report. Troubleshooting Action: What steps are taken during the system diagnosis: Result of response: Problem solved. Problem location: 04 Gantry Exclude HV What steps did you perform to diagnose the problem? Call description: Ring artifact occurrence Artifacts appear even when re-scanning immediately. [Response] Artifacts did not appear in subsequent test scans. [Image confirmation results] The problem occurred near M18 or M25. Artifacts are coming from detectors close to where they appeared in March. Error # and/or description of error: N/A Repair Action: Completed or Recommended Repair Action "External": The service engineer for the site visited and worked on the equipment. OneEMS internal remarks : Problem description by engineer : Equipment affected: CT Problem location: 04 Gantry Exclude HV When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user_s work flow did the problem occur?: Post-Acquisition/Scan What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem? No.The customer has not created a log file or report. What steps are taken during the system diagnosis: Result of response: Problem solved. Problem location: 04 Gantry Exclude HV What steps did you perform to diagnose the problem? Call description: Ring artifact occurrence Artifacts appear even when re-scanning immediately. [Response] Artifacts did not appear in subsequent test scans. [Image confirmation results] The problem occurred near M18 or M25. Artifacts are coming from detectors close to where they appeared in March. Error # and/or description of error: N/A Completed or Recommended Repair Action "External" : The service engineer for the site visited and worked on the equipment. Completed or Recommended Repair Action "Internal": --Troubleshooting Action-- Result of response: Problem solved. Problem location: 04 Gantry Exclude HV What steps did you perform to diagnose the problem? Call description: Ring artifact occurrence Artifacts appear even when re-scanning immediately. [Response] Artifacts did not appear in subsequent test scans. [Image confirmation results] The problem occurred near M18 or M25. Artifacts are coming from detectors close to where they appeared in March. --Repair Action-- Recommended 12Nc: PR1 459800911271:TDMS V2 M.Det Class A FRU QTY:2 SLOC: Sapporo (U612) --Subject-- 04DS:KKI: Mr./Ms. Ando: Ring artifact appeared >> Appeared near M18 or M25 > Did not appear in subsequent test scan --Problem Description-- Equipment affected: CT Problem location: 04 Gantry Exclude HV When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user_s work flow did the problem occur?: Post-Acquisition/Scan What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem? No.The customer has not created a log file or report. ----------------------------------------------------------------------------- | 4 |  | The problem occurred near M18 or M25. Artifacts are coming from detectors close to where they appeared in March. --Repair Action-- Recommended 12Nc: PR1 459800911271:TDMS V2 M.Det Class A FRU QTY:2 SLOC: Sapporo (U612) finally system works after replace DMS MB | [12 empty] | 117622065 |  | 958147 |  | Investigation | Closed - Done |  | CTUDT00632110 | [5 empty] | Engineering | [6 empty] | Artifact |  |  | Completed |  |  | Refresh |  |  | System issue | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | 0 |  |  | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | y | y | y | y | y | 0 | y | #REF!
1003 |  |  | Japan | 苫小牧市立病院 | 341014 | 44352 |  | 03CS:KK:Air Calibration中にE-Stop... | Original Problem Reported by Customer 03CS: KK: E-Stop Open occurred during Air Calibration. >> Restored by closing again. >> S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT, S_GEN_FAULT_CATHODE_TUBE_ARC, szInfo: SoftArcReg 0x211 * Problem Description Problem description by engineer: Target device: CT Failure location: 03 High Voltage Energy When did it first occur? (Within ~): 1 Hour How often do problems occur ?: First Occurrence At what point in the end-user workflow is the problem occurring ?: Setting up Protocol What happened recently ?: N.A. Does the customer create a specific system event report for the problem ?: No. The customer was unable to create a log file or report. Troubleshooting Action What steps are taken during the system diagnosis: Post-response results: Improved Failure location: 03 High Voltage Energy What steps were taken during the diagnosis ?: Restored by E-Stop re-Close. It seems that an arc was generated, S_TIMER_WATCHDOG_CAUSED_LAST_RESET was applied, and E-Stop Open was performed. I was allowed to wait and see. | 4 |  | Restored by E-Stop re-Close. It seems that an arc was generated, S_TIMER_WATCHDOG_CAUSED_LAST_RESET was applied, and E-Stop Open was performed. I was allowed to wait and see. | [12 empty] | 117655296 |  | 978314 |  | Evaluation - Pending Translation | Closed - No Further Investigation |  | CTUDT00630085 | [5 empty] | Engineering | [6 empty] | Tube Arc |  |  | Completed |  |  | Refresh |  |  | System issue | System | IC-HW | Tub Arc | IC-HW-Tub Arc | IC-HW-Tub Arc | CTUDT00630085 | Tub Arc | IC-HW | IC-HW-Tub Arc | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
1004 |  |  | Japan | Maki Hospital | 341017 | 44351 |  | The scan stopped. | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : A scan stopped. ----------------------------------------------------------------------------- | 4 |  | No info | [12 empty] | 117650914 |  | 976801 |  | Evaluation | Investigation | [7 empty] | TBD | [6 empty] | TBD |  |  | Completed |  |  | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1005 |  |  | China | Jiangsu Province Yangzhou City No. 1 People_s Hospital, West District Hospital | 345022 | 44347 |  | 02N/Ingenuity Core128/Elite China/dicom transfer problem | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity Core128/Elite China SAP Transfer : Case Transferred to SAP Safety Questions : Case: 0117622669 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 其他类型故障。 客户问题详细描述或故障现象： 图像传输失败 设备本身技术问题： 否。 Troubleshooting Action: 关键故障代码或现象： CT 传PACS有图传不过去 远程诊断与测试结果： 客户要求工程师到场处理 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 查看传输失败的图像，检查pacs设置 SAP Transfer : Dispatched for Onsite Support T2 Activities : Problem Description: 客户问题概述： 其他类型故障。 客户问题详细描述或故障现象： estop 断开 设备本身技术问题： 否。 Troubleshooting Action: 关键故障代码或现象： S_TILT_C_TO_BAD_SERIAL 远程诊断与测试结果： 通讯故障，R2D引起 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 更换R2D 测试 建议备件12NC及名称： 459800467343 R2D Assembly Next Step Action : 013951434175再次来电，联系 qi liangdong回电 External Remarks : 06.06.2021 09:05:07 UTC Liangdong Qi (86036024)服务内容：删除垃圾文件，修改配置参数. 服务结果： 维修后，设备恢复正常工作. Internal Remarks : 06.06.2021 09:05:10 UTC Liangdong Qi (86036024) "RDI Accuracy: Yes“ $Q7: (Confirmed reported problem, the Dose structured Report tansfer failed from CT host to pacs.) $Q8: (No error report, checked the pacs and found it can not accepted the DSR.) $Q9: (Agjusted the image transfer parameter, disable rhe Dose Structure Report option.) $Q10: (Test and Inspection data refer to PA & Device return to fully function) $Q11: Calibrated Tools (NA). ----------------------------------------------------------------------------- | 3 |  | $Q7: (Confirmed reported problem, the Dose structured Report tansfer failed from CT host to pacs.) $Q8: (No error report, checked the pacs and found it can not accepted the DSR.) $Q9: (Agjusted the image transfer parameter, disable rhe Dose Structure Report option.) | [12 empty] | 117622669 |  | 958409 |  | Opened - Pending Translation | Closed - No Further Investigation |  | CTUDT00633122 | [5 empty] | Engineering | [6 empty] | Audit Repository |  |  | Completed |  |  | Refresh |  |  | System issue | System | Console-SW | AuditRepository | Console-SW-AuditRepository | Console-SW-AuditRepository | CTUDT00633122 | AuditRepository | Console-SW | Console-SW-AuditRepository | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y
1006 |  |  | China | Nanjing Medical University Affiliated Yifu Hospital | 345028 | 44349 |  | 02N/Core128/software crash | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 02N/Core128/ SAP Transfer:Dispatched for On site Support SAP Transfer:Case Transferred to SAP T2 Activities:Problem Description: Summary of customer_s problems: Often crash Detailed description of Customer_s problems or malfunctions: Often crash It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: NA Remote troubleshoot and test result: The customer said the machine often crashes, with yellow screen PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: patch according to Yellow screen workaround. Collect bugreport if necessary Safety Questions:Case: 0117636136 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created by: Wennian Han Last updated by: Wennian Han 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 3 |  | Yellow screen workaround | [12 empty] | 117636136 |  | 967619 |  | Opened | Closed - No Further Investigation |  | CTUDT00632520 | [5 empty] | Engineering | [6 empty] | New Yellow Screen |  |  | Completed |  |  | Refresh |  |  | System issue | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #REF!
1007 |  |  | China | 林甸县医院 | 345038 | 44350 |  | Wechat:02N/Ingenuity Core128/Elite China/Consultation questions | BU Q&R Feedback: No SAP Customer Symptom Code: CENV Customer Symptom Code: Customer Environmental Concern Reason: Corrective Maintenance Engineer Failure Code: Environmental Cause Repair Activity Code: Instructed Customer Problem Reported by Customer : Consultation questions Safety Questions : Case: 0117643840 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Zhou Xiaohong Last updated by: Zhou Xiaohong 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Troubleshooting Action: answer customer question Repair Action: answer customer question Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | answer customer question | [12 empty] | 117643840 |  | 972650 |  | Opened - Pending Translation | Closed - non-complaint | [7 empty] | TBD | [6 empty] | NAC |  |  | Completed |  |  | Refresh |  |  | System issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1008 |  |  | United States | Select Medical Corporation | 310050 | 44350 |  | add MWL to CT | Work Order #: WO-03596780 Franklin Rex Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 5 |  | add Worklist to CT | [12 empty] | 117646423 |  | 974268 |  | Opened | Closed - non-complaint | [7 empty] | TBD | [6 empty] | NAC |  |  | Complete |  |  | Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1009 |  |  | United States | AdventHealth Zephyrhills | 300099 | 44350 |  | 1302CFXXX Unable to send to pacs & worklist not updating | Work Order #: WO-03594294 Ali, Adam The system fully meets specification and returned to use Problem Description by Engineer : Problem description by engineer : Unable to send to pacs & worklist not updating When was the first occurrence? Within... : 03-Jun-2021 How often is issue occurring? : Continuously Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No further information is available as per the customer Expected and Unexpected results: Not provided Patient Impact/Outcome: No further information is available as per the customer User Impact: Cannot send images to PACS or pull new patient info from worklist Current Software Version: 4.1.7.25030 Resolution : Restarted Host computer. Worklist is not updating, also images are transferring to PACS. Internal Comments : Error # and/or description of error : Unable to send to pacs & worklist not updating Diagnostic performed by Engineer : Unable to send to pacs & worklist not updating. Restarted Host & CIRS computers. Worklist is updating, also images are transferring to PACS. ----------------------------------------------------------------------------- | 3 |  | Restarted Host & CIRS computers.6/3/2021 collect bugreport if have | [12 empty] | 117643789 |  | 968474 |  | Evaluation | Evaluation |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS Transfer |  |  | Complete |  |  | Refresh |  |  | Upgrade issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | 0 | y | y
1010 |  |  | United States | Select Specialty Hospitals | 310051 | 44347 | 22 | Device shutdown due to power went out and shut down the CIRS causing a disconnect; Verified system_s normal operation. | Work Order #: WO-03578022 01-Jun-2021 Siker Shawn The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : Power went out Friday night and shut down the cirs causing a disconnect, I verified system was working with test scans and handed back to the customer. ----------------------------------------------------------------------------- | 2 |  | Power went out Friday night and shut down the cirs causing a disconnect at Monday, I verified system was working with test scans and handed back to the customer. | [12 empty] | 117627250 |  | 960508 |  | Closed - Done | Closed - Done | [7 empty] | Customer | [6 empty] | CIRS disconnect |  |  | Complete |  |  | Refresh |  |  | Upgrade issue | [6 empty] | 0 |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1011 |  |  | Argentina | INSTITUTO NACIONAL DE | 40027 | 44355 |  | Detector replacement | Work Order #: WO-03636155 Meza, Edgardo Arturo The system partly meets specification and is returned to use with restrictions Problem Description by Engineer: Intermittent rings in the image. Follow Up Required: Replace detector. Diagnostic performed by Engineer: No signal in the detector module. ----------------------------------------------------------------------------- | 5 |  | No signal in the detector module. Replace DMS module | [12 empty] | 117670171 |  | 989472 | TRUE | Investigation | Investigation |  | CTUDT00632110 | [5 empty] | Engineering | [6 empty] | DMS Module | [5 empty] | Refresh |  |  | System issue | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | DMS Module | DMS-HW | DMS-HW-DMS Module | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | y
1012 |  |  | Argentina | INSTITUTO NACIONAL DE | 40027 | 44355 |  | TRAINING ON EQUIPMENT IS REQUESTED | Work Order #: WO-03621433 Alec, Preciado The system fully meets specification and returned to use. Problem Description by Engineer : Problem description by engineer : They request a second visit by Applications department to install a new system. Information to support the complaint handling process Customer Function/Role: BioMed Device Use: No patient involved. Expected and Unexpected results: No further information is available as per the customer Patient Impact/Outcome: It does not impact. User Impact: It does not impact. Current Software Version: 4.1.7 Diagnostic Performed by Engineer: They request a second visit by Applications department for a new system installed on March. The client says that they already contacted the application expert Fernando Dominguez. Internal Comments : Error # and/or description of error : NA Completed or Recommended Repair Action "Internal" : They request a second visit by Applications department for a new system installed on March. The client says that they already contacted the application expert Fernando Dominguez. Follow Up Required: They request a second visit by Applications department for a new system installed on March. The client says that they already contacted the application expert Fernando Dominguez. Resolution: They request a second visit by Applications department for a new system installed on March. The client says that they already contacted the application expert Fernando Dominguez. ----------------------------------------------------------------------------- | 5 |  | They request a second visit by Applications department for a new system installed on March. | [12 empty] | 117669723 |  | 985226 |  |  | Closed - Non-Complaint | [7 empty] | NAC | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1013 |  |  | China | 山东安丘市中医医院 | 345017 | 44360 |  | 02N/Ingenuity Core128/Elite China/couch fault | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : couch fault Next Step Action : 派工 yang xingbo 13583235930 Safety Questions : Case: 0117692556 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Guomei Li Last updated by: Guomei Li 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support ----------------------------------------------------------------------------- | Closed - Done |  | 服务内容：调整tape switch 位置。 服务结果：维修后，设备恢复正常工作。 | [12 empty] | 117692556 |  | 1001027 | TRUE | Closed - Done | Closed - Done |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | Tape switch | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | TabSwitch | Couch-HW-TabSwitch | Couch-HW-TabSwitch | CTUDT00632127 | TabSwitch | Couch-HW | Couch-HW-TabSwitch | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | y
1014 |  |  | China | 浙江省乐清第一人民医院 | 40054 | 44359 |  | 02N/Ingenuity CT 728326/Connect the door machine linkage | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : Connect the door machine linkage Safety Questions : Case: 0117692011 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Zixuan Last updated by: Wang Zixuan 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 12.06.2021 12:40:31 UTC Yi Huang (86036249)服务内容：连接门控开关线。 服务结果：维修后，设备恢复正常工作。 其它重要事项：无 Internal Remarks : 12.06.2021 12:40:34 UTC Yi Huang (86036249) RDI Accuracy:Yes $Q7: (Confirmed reported problem&connect door switch cable) $Q8: (connect door switch cable) $Q9: (connect door switch cable) $Q10: (Test and inspection data refer to: PA.Device returned to full functionality.) $Q11: Calibrated tools name(N/A) S/N (N/A) Next calibration date (N/A) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date:N/A ^Current Tube usage: second:N/A ----------------------------------------------------------------------------- | Closed - No Further Investigation |  | 服务内容：连接门控开关线。 服务结果：维修后，设备恢复正常工作。 | [12 empty] | 117692011 |  | 1000613 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Service | [6 empty] | System Installation | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1015 |  |  | China | 扎兰屯市人民医院 | 40045 | 44358 |  | 02N/Ingenuity CT 728326/cann_t log in | BU Q&R Feedback: No SAP Customer Symptom Code: CENV Customer Symptom Code: Customer Environmental Concern Reason: Corrective Maintenance Engineer Failure Code: Environmental Cause Repair Activity Code: Configuration Change Problem Reported by Customer : 02N/Ingenuity CT 728326 Safety Questions : Case: 0117685942 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Weng Jinming Last updated by: Weng Jinming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Troubleshooting Action: instruct customer to change password Repair Action: instruct customer to change password Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | Closed - No Further Investigation |  | instruct customer to change password | [12 empty] | 117685942 |  | 999974 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Password | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1016 |  |  | China | 内蒙古兴安盟人民医院 | 345030 | 44356 |  | 02N/Ingenuity Core128/Elite China/scan room leakage of rain | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Calibration/Adjustment Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ Safety Questions : Case: 0117672104 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Gaoge Last updated by: Wang Gaoge 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 其他类型故障。 客户问题详细描述或故障现象： 房间漏雨到设备上，需要工程师到场检查 设备本身技术问题： 否。 Troubleshooting Action: 关键故障代码或现象： 设备上漏水了 远程诊断与测试结果： 到场根据实际情况清理设备 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 1，清理水迹，检查设备 Test/Inspection: N/A Test Equipment: N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Internal Remarks : 11.06.2021 01:08:19 UTC Ruijie Su (86039043) $Q7: Confirmed reported problem,Air conditioning leaking。 . $Q8: on sit service ,Air conditioning leaking cpm arc. . need clean apm plug . $Q9: clean apm plug . .system is ok! $Q10:Test and inspection data refer to:PA. Device returned to full functionality. $Q11: Calibrated Toolsname ( Multimeter Fluke 287C) S/N (31060028) Next Calibrationdate (06/1do bed calibration and air calibration31/10/2020) 2( Torque Wrench Kit ) S/N E201602048 Next Calibration date22/03 /2022 For Mainland China CV/ XR, CTengineers only.Please fill in: ^ Current Tubeinstalled date 10/01/2017 ^ Current Tube usage: scan second 6589 External Remarks : 11.06.2021 01:08:16 UTC Ruijie Su (86039043)服务结果： [检查发现打火清洁CPM高压插头后恢复 ，设备恢复正常工作。] 其它重要事项：无 ----------------------------------------------------------------------------- | 5 |  | 服务结果： [检查发现打火清洁CPM高压插头后恢复 ，设备恢复正常工作。] on sit service ,Air conditioning leaking cpm arc. . need clean apm plug . $Q9: clean apm plug . .system is ok! $Q10:Test and inspection data refer to:PA. Device returned to full functionality. | [12 empty] | 117672104 |  | 990986 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Rain leakage | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1017 |  |  | China | Shandong Province Chest Hospital | 40025 | 44355 |  | WeChat:02N/Ingenuity CT 728326/Couch fault | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : Couch fault Safety Questions : Case: 0117665311 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Bo Last updated by: Wang Bo 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 扫描床水平运动故障。 客户问题详细描述或故障现象： 机架掉电。 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： VELOCITY_ENCODER_FAILURE,HOR_MOTION_STOP_ERROR,COUCH IS JOGGING. 远程诊断与测试结果： 重启无效。 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议做床水平校准，检查床水平读数，velocity读数偏差过大，cantrace里床表现卡顿检查床侧板是否变形。检查水平电机皮带松紧。短接tapeswitch。 建议备件12NC及名称： PPD 453566489521，459800712461 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : PPD备件FRU, AMC upgrade KITS 459800712461*1&MCG /PITTMAN MOTOR ASSEMBLY 453566489521*1已安排24H到场地，FSE：黄兆坤-86-18654536122 External Remarks : 08.06.2021 11:59:35 UTC SUP JCOUSER (86038405)服务内容：扫描床水平校正，垂直校正，调节扫描床运动阻力。 服务结果：维修后，设备恢复正常使用。 Internal Remarks : 12.06.2021 08:46:26 UTC Zhaokun Huang (86038405) "RDI Accuracy: YES." $Q7: ( Confirmed reported problem & Couch fault ) $Q8: ( Check error log , report "S_COUCHMGR_HOR_MOTION_STOP_ERROR". Manual push couch , find resistance serious .) $Q9: ( Adjust Couch movement resistance.) $Q10: ( Test and inspection data refer to: PA . Device returned to full functionality.) $Q11: Calibrated Tools name (FLUKE287 ) S/N (23350170) Next Calibration date (04/07/2022) ^Current Tube installed date: N/A ^Current Tube usage: N/A ----------------------------------------------------------------------------- | Closed - No Further Investigation |  | Check error log , report "S_COUCHMGR_HOR_MOTION_STOP_ERROR". Manual push couch , find resistance serious 服务内容：扫描床水平校正，垂直校正，调节扫描床运动阻力。 服务结果：维修后，设备恢复正常使用。 | [12 empty] | 117665311 |  | 986914 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | AssemblyAdjust | [5 empty] | Engineering | [6 empty] | Couch Vertical Movement | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | AssemblyAdjust | Couch-HW-AssemblyAdjust | Couch-HW-AssemblyAdjust | AssemblyAdjust |  | Couch-HW | Couch-HW-AssemblyAdjust | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | y
1018 |  |  | China | Guangxi Province Teng County People’s Hospital | 40036 | 44355 |  | 02N/Ingenuity CT 728326/application | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer:application Next Step Action : yang jun &guo fan follow Safety Questions:Case: 0117664378 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingli Last updated by: Lu Mingli 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | Closed - Non-Complaint |  | NAC | [12 empty] | 117664378 |  | 986558 |  |  | Closed - Non-Complaint | [7 empty] | NAC | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1019 |  |  | United States | Unity Health Systems | 40037 | 44355 |  | Gantry Cannot Initialize | Work Order #: WO-03623534 08-Jun-2021 Taylor, Durrance The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: tech Device Use: patient scan Expected and Unexpected Results: patient scanned/gantry cannot initialize User Impact: delayed Patient Impact: delayed Current Software Version: 4.1.10 Resolution : Found the couch horizontal movement had some resistance when table was inside gantry. Adjusted the horizontal drive screw assy and verified normal operation. | 2 |  | Found the couch horizontal movement had some resistance when table was inside gantry. Adjusted the horizontal drive screw assy and verified normal operation. | [12 empty] | 117671365 |  | 986152 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | AssemblyAdjust | [5 empty] | Engineering | [6 empty] | Couch horizontal drive | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | AssemblyAdjust | Couch-HW-AssemblyAdjust | Couch-HW-AssemblyAdjust | AssemblyAdjust |  | Couch-HW | Couch-HW-AssemblyAdjust | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | y
1020 |  |  | Peru | BANCO BBVA PERU | 40034 | 44356 |  | 0117675674-THE TUBE HEATING IS NOT DONE | Work Order #: WO-03628027 Caldelas Aldana, Carlos Manuel The system does not meet specification and is out of use Problem Description by Engineer : Problem description by engineer : It does not allow them to discharge. When was the first occurrence? Within... : 09-Jun-2021 07:00:00 How often is issue occurring? : First occurrence Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected Results: Discharge is aborted. Patient Impact/Outcome: N/A User Impact: It does not allow scanning. Current Software Version: 4.1.10 Diagnostic Performed by Engineer: It is checked that even after a total shut down, the equipment does not allow scans of any type, not even the X-ray tube heating. We try to access remotely but without success, so we cannot get further details about the diagnosis. Internal Comments : Error # and/or description of error : It does not scan Completed or Recommended Repair Action "Internal" : N/A Follow Up Required : What are the next Troubleshooting steps : Check logs to determine if the cause of the failure is control link, high voltage, etc. Repair and reconfigure PRS connection correctly for this software version and it is working, run functional tests. Error # and/or description of error : It does not scan Completed or Recommended Repair Action "Internal" : N/A ----------------------------------------------------------------------------- Work Order #: WO-03629139 Antonella Haro Problem Description by Engineer : Problem description by engineer : It does not allow them to discharge. When was the first occurrence? Within... : 09-Jun-2021 07:00:00 How often is issue occurring? : First occurrence Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected Results: Discharge is aborted. Patient Impact/Outcome: N/A User Impact: It does not allow scanning. Current Software Version: 4.1.10 Diagnostic performed by Engineer: N/A ----------------------------------------------------------------------------- | 2 |  | Error log is reviewed. Bows are found. High voltage chunks are checked, the one that goes to the generator is very dirty. They are cleaned and lubricated. NO LOAD TEST is performed and the equipment does not generate arcs, generators are discarded. X-ray tube problems, parameters of the well grounded will be reviewed with maintenance personnel from the clinic. Change X-ray tube. After excessive cleaning, lubrication and downtime. The equipment remains operational. Air calibration and scans are performed on patients normally. However, on June 11 artifacts are reported. It is required to measure the wells to earth and diagnose the detectors. Image artifacts. Carry out bad detectors test, check collimator status. | [12 empty] | 117675674 |  | 989268 | TRUE | Closed - Done | Closed - Done |  | ShouldbeService | [5 empty] | Engineering | [6 empty] | Tube | [5 empty] | Refresh |  |  | System issue | System | IC-HW | Tube | IC-HW-Tube | IC-HW-Tube | ShouldbeService |  | IC-HW | IC-HW-Tube | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | y
1021 |  |  | China | Guangxi Province Teng County People’s Hospital | 40036 | 44355 |  | Application/Ingenuity CT 728326/Application Support | BU Q&R Feedback: No Reason: Application Support Problem Reported by Customer : Application Support Next Step Action: Emailed CT Pre-Sales Training Department Manager Yang Jun, Guo Fan to follow up. Safety Questions:Case: 0117664375 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Zhou Xiaohong Last updated by: Zhou Xiaohong 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action: The customer requests a 2nd on-site application training. ----------------------------------------------------------------------------- | Closed - Non-Complaint |  | NAC | [12 empty] | 117664375 |  | 986555 |  |  | Closed - Non-Complaint | [7 empty] | NAC | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1022 |  |  | China | 江苏大学附属医院 | 345023 | 44354 |  | 02N/Ingenuity Core128/Elite China/adjust couch | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ Safety Questions : Case: 0117660478 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wennian Han Last updated by: Wennian Han 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 调床 客户问题详细描述或故障现象： 调床 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： NA 远程诊断与测试结果： 报修床围裙刮擦，客户需要来人调节 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 调节床围裙；围裙是活动部件，需和客户沟通好相关。 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 08.06.2021 07:56:00 UTC SUP JCOUSER (86036528)服务内容：扫描床降到最低的时候有异响，调整扫描床底部弹簧位置。 服务结果：服务后，设备恢复正常 Internal Remarks : 08.06.2021 08:04:32 UTC Yang Liu (86036528) RDI Accuracy: Yes $Q7: (Confirm reported problem,the couch has abnormal noise at the lowest level.) $Q8: ( found there is friction between the spring and the spring post） $Q9: (adjust the spring position according to service manual .) $Q10: (Test and inspection data refer to PA. Device returned to full functionality.) $Q11: Calibrated tools refer to PA. ----------------------------------------------------------------------------- | Closed - No Further Investigation |  | 服务内容：扫描床降到最低的时候有异响，调整扫描床底部弹簧位置。 服务结果：服务后，设备恢复正常 | [12 empty] | 117660478 |  | 983565 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00631058 | [5 empty] | Engineering | [6 empty] | Couch Spring | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | Vertical spring | Couch-HW-Vertical spring | Couch-HW-Vertical spring | CTUDT00631058 |  | Couch-HW | Couch-HW-Vertical spring | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | Couch Spring ECR | y | y
1023 |  |  | China | 湖北省武汉仁爱医院 | 345004 | 44354 |  | 02N/Ingenuity Core128/Elite China/gantry can not close estop | BU Q&R Feedback: No SAP Customer Symptom Code: BOOT Customer Symptom Code: Boot up/Power-up Problem Reason: Corrective Maintenance Engineer Failure Code: Electrical/Electronic Failure Repair Activity Code: Instructed Customer Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ Safety Questions : Case: 0117657807 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Gaoge Last updated by: Wang Gaoge 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 有钥匙框，但是拧钥匙闭合estop没反应，机架不能正常上电 Troubleshooting Action: 指导客户检查紧急开关，找打机架背面的紧急开关被按下，恢复紧急开关后设备恢复正常 Repair Action: 指导客户检查恢复紧急开关 Test/Inspection: N/A Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | Closed - No Further Investigation |  | 指导客户检查紧急开关，找打机架背面的紧急开关被按下，恢复紧急开关后设备恢复正常 | [12 empty] | 117657807 |  | 981894 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | E-Stop button | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1024 |  |  | China | 河北省唐山市工人医院 | 40044 | 44354 |  | 02N/Ingenuity CT/image artifacts | BU Q&R Feedback: No SAP Customer Symptom Code: ACQP Customer Symptom Code: Imaging/Patient Recording/Acquisition Problem Reason: Corrective Maintenance Engineer Failure Code: Connection/Cable/Interconnect Repair Activity Code: Calibration/Adjustment Problem Reported by Customer : 02N/Ingenuity CT/ Safety Questions : Case: 0117657597 Safety Question Status: In Process PSA: false PSE: false PSF Status: NA Created By: Wang Gaoge Last updated by: Wang Gaoge 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Safety Questions : Case: 0117657597 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Gaoge Last updated by: Wang Gaoge 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 图像有一个半环圆形伪影在图像中心位置 Troubleshooting Action: 检查图像发现有探测器坏道引起伪影，指导客户做空气校正后伪影消失 Repair Action: 指导客户做空气校正 Test/Inspection: N/A Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | Investigation |  | 图像有一个半环圆形伪影在图像中心位置 Troubleshooting Action: 检查图像发现有探测器坏道引起伪影，指导客户做空气校正后伪影消失 | [12 empty] | 117657597 |  | 981699 | TRUE | Investigation | Investigation |  | NotEngineering | [5 empty] | Engineering | [6 empty] | Artifact | [5 empty] | Refresh |  |  | System issue | System | NotEngineering | ShouldbeCustomer | NotEngineering-ShouldbeCustomer | NotEngineering-ShouldbeCustomer | NotEngineering |  | NotEngineering | NotEngineering-ShouldbeCustomer | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | y
1025 |  |  | Japan | 苫小牧市立病院 | 341014 | 44358 |  | 03DM:KKI:撮影開始時に、エラー発生＞＞SoftArcReg 0x211* /S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT＞継続して検査可 | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 03DM:KKI:撮影開始時に、エラー発生＞＞SoftArcReg 0x211* /S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT＞継続して検査可 Safety Questions : Case: 0117686469 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: KAZUMASA KIMURA Last updated by: KAZUMASA KIMURA 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Message : LOGI Task added Successfully!- SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support SAP Message : Components Added Succesfully to SWO!- ----------------------------------------------------------------------------- | 4 |  | There are replacement parts 12Nc: 989000086352 Name: PHILIPS MRC 880 QTY: 1 We replaced the high-voltage cable on June 14, but the arc occurred again, so we replaced the X-ray tube, which is thought to be the cause. After replacement, make various adjustments. We have confirmed that there is no problem with the overall performance. 12Nc: 989000005011 Name: HV-ANODECABLE PHILIPS MRC880 CT 1,84m QTY: 1 12Nc: 989000003302 Name: HV CABLE PHILIPS MRC CT-O4S 0,95m QTY: 1 | [12 empty] | 117686469 |  | 1001235 | TRUE | Investigation | Investigation |  | NotEngineering | [5 empty] | Engineering | [6 empty] | Tube | [5 empty] | Refresh |  |  | System issue | System | NotEngineering | ShouldbeOthers | NotEngineering-ShouldbeOthers | NotEngineering-ShouldbeOthers | NotEngineering |  | NotEngineering | NotEngineering-ShouldbeOthers | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | y
1026 |  |  | China | 内蒙古阿拉善右旗人民医院 | 345031 | 44357 |  | 02N/Ingenuity Core128/Elite China/ can not bootup | BU Q&R Feedback: No Reason: Corrective Maintenance Safety Questions : Case: 0117679178 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Zhang Caiqin Last updated by: Zhang Caiqin 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Problem Reported by Customer : 02N/Ingenuity Core128/Elite China T2 Activities : Problem Description: 客户问题概述： 扫描失败。 客户问题详细描述或故障现象： 无法扫描 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： S_MASTER_SEQUENCER_GDC_LOW_AFTER_TIMEOUT 远程诊断与测试结果： "Rail voltage did not go up to at least 500VDC after the soft start contactor needed to be pulled in. The rotor will start to rotate but the Main drive contactor will drop. The E stop GUI will not disappear from the screen. " PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： "Recommendation: Caution: Due to the 480V AC you may have to Kill Main Power wall, in some cases the blue button did not remove the power. 1. Try to close the E stop when the gantry cover switch is midway (cover open) so there is no rotor movement to see if this is caused by rotor movement, it it is, check / clean the slip rings and the brush block. 2. Check the 480V AC fuses to the slipring F351,352,353. 3 .Check the 480V AC fuses on the APM and the KPM F1.2,3. 4. Check the relays on the ACCU. 5. Check CB302. 6.. Check the soft start contactor (SS)and the activation signal (120V) at TB352 / 2: a. If the signal is not there, skip to step 8. b. Check the Aux. Contact side mounted N.C. that is mechanically attached to the right side of the MKD. The signal should be on the bottom of that contactor when the MKD is not pulled in. If the signal is at the top but not at the bottom, replace the Aux contactor. c.. Check the 2 varistors at the bottom of the SS contactor. d. Check the voltage at the bottom of the SS contactor (red wire - voltage, white wire - neutral) if the voltage is there, but the contactor does not pull in, replace the contactor. d. If the SS contactor pulls in check the Soft Start Resistors. 7. If the soft start is OK, with the cover switch midway, check for 480V AC on both the APM and the KPM, if the voltage is not there check the brush block and slip rings, if the voltage is there: a. check the cable from the RHOST to the DMC. (the DMC is reporting to the GHOST on DC level) b. Check / replace the RHOST. c. Check / replace the System controll card. d. Check / replace the DMC. 8. If the soft start signal is missing: a. Check / replace the flat cable from the ACCU to the GHOST. b. Check / replace the ACCU. c. Check / replace the GHSOT Timing (use the Stargate 164 Stationary Power drawings): 1. E stop close + ""START"" relay on will pull in ""PUP"" relay. 2. ""PUP"" relay will pull in ""MK"" which will act as a self holding replay to ""PUP"" check for system LED on ACCU to be on. 3 """"PUP"" relay will send the ""5GNORM"" to the GHOST. 4. The GHOST will send the ""S/SON"" which will pull in the ""S/SSSR"" relay. 5. ""S/SSSR"" relay will send the ""120 S/SON"" to TB 352 / 2 which will pull in the ""S/S (soft start) relay. 6. When the DC voltage is over 500V, the GHOST will send the ""MKDON"" signal. 7. The ""MKDON"" signal will pull in the MKDSSR"" relay. 8. The ""MKDSSR"" relay will send the ""120MKDON"" to TB 352 /1 which will pull in the ""MKD"" (the main 480V AC contactor). 9. ""MKD"" relay will open the ""S/S"" relay by removing the 120 S/SON. (via Aux. Contact side mounted N.C. that is mechanically attached to the right side of the MKD) Experience: 1. The following troubleshooting steps were performed: a. I arrived on site and found 2 out of 3 APM 480 vac input fuses blown. b. I replaced the 30a fuses. I applied power to the gantry and closed the estop. A few seconds into the APM/CPM power charging routine, I heard what I thought might be an arc and the estop opened. The same APM fuses were blown. c. Next, I removed the HSS blue connector and replaced the 30a fuses. The estop closed and the system powered up for about 15 seconds before I heard multiple arcs coming from the APM. No fuses blew this time. d. I replaced the APM. No further problems. " 建议备件12NC及名称： Fuse, 30A/600V 459800066152 1 F1, F2, F3//FUSE,2.50A,500V.406"DIA X1.50" 459800726611//Fuse, 50A, 600VAC, Dual Element TD 459800043431 F351, F352, F353 000090 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 12.06.2021 09:01:23 UTC Sheng Cui (86038219)服务内容：机架无法上电，Softstart 无法吸合，复位交流接触器，更换滑环高压保险和滑环保险。 服务结果：维修后，设备恢复正常工作 Internal Remarks : 12.06.2021 09:01:48 UTC Sheng Cui (86038219) RDI Accuracy（Yes） $Q7: (Confirmation & Description of problem:gantry boot up failed )$Q8: (On site check logeer message GDC low ,need to replace slip ring 50A fuse and HV generator 20A fuse.) $Q9: (Replace slip ring 50A fuse and HV generator 20A fuse. ) $Q10: (Test and inspection data refer to: PA. Device returned to full functionality.) $Q11: Calibrated Tools name (FLUKE287 ) S/N (12470209 ) Next Calibration date (07/12/2021）: For Mainland China CV / XR, CT engineers only. Please fill in: Current Tube installed date Tube installed date: ##09/ 4/ 2020 ^Current Tube usage: scan second_##401202S ----------------------------------------------------------------------------- | 2 |  | 机架无法上电，Softstart 无法吸合，复位交流接触器，更换滑环高压保险和滑环保险。 | [12 empty] | 117679178 |  | 995706 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Operation | [6 empty] | Power Brush assembly WI | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1027 |  |  | Japan | 社会福祉法人恩賜財団済生会 | 343000 | 44357 |  | 07CS:TM:FSE経由。消耗品劣化 | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 07CS:TM:FSE経由。消耗品劣化 SAP Transfer : Case Transferred to SAP SAP Message : Components Added Succesfully to SWO!- SAP Message : LOGI Task added Successfully!- Safety Questions : Case: 0117680995 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: TAKEYUKI MIYAKAWA Last updated by: TAKEYUKI MIYAKAWA 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: Problem description by engineer : 対象装置:CT 故障個所:07その他 最初に発生したのはいつか?(~以内):Unknow 問題発生の頻度は?:Unknow エンドユーザｰのワークフローのどの時点において問題が発生しているのか?:その他(Other) 最近どのようなイベントが発生したか?:N.A. 顧客は問題について具体的なシステムイベントレポートを作成しているか?:N.A. Engineer_s Comment: _============================== コール内容 FSE経由。消耗品劣化 _============================== Troubleshooting Action: What steps are taken during the system diagnosis : 対応後の結果:改善された 故障個所:07その他 診断中にどのようなステップを実施したか?: _================================== 対応 Parts手配 _================================== Error # and/or description of error : NA OneEMS internal remarks : Problem description by engineer : 対象装置:CT 故障個所:07その他 最初に発生したのはいつか?(~以内):Unknow 問題発生の頻度は?:Unknow エンドユーザｰのワークフローのどの時点において問題が発生しているのか?:その他(Other) 最近どのようなイベントが発生したか?:N.A. 顧客は問題について具体的なシステムイベントレポートを作成しているか?:N.A. Engineer_s Comment: _============================== コール内容 FSE経由。消耗品劣化 _============================== What steps are taken during the system diagnosis : 対応後の結果:改善された 故障個所:07その他 診断中にどのようなステップを実施したか?: _================================== 対応 Parts手配 _================================== Error # and/or description of error : NA Completed or Recommended Repair Action "Internal" : --Troubleshooting Action-- 対応後の結果:改善された 故障個所:07その他 診断中にどのようなステップを実施したか?: _================================== 対応 Parts手配 _================================== --Repair Action-- Recommended 12Nc: Out of Stock 453567104591:SINGLE STRAP QTY:5 SLOC:BackOrder --Subject-- 07CS:TM:Yasuhiro Kawashige:FSE経由。消耗品劣化 --Problem Description-- 対象装置:CT 故障個所:07その他 最初に発生したのはいつか?(~以内):Unknow 問題発生の頻度は?:Unknow エンドユーザｰのワークフローのどの時点において問題が発生しているのか?:その他(Other) 最近どのようなイベントが発生したか?:N.A. 顧客は問題について具体的なシステムイベントレポートを作成しているか?:N.A. Engineer_s Comment: _============================== コール内容 FSE経由。消耗品劣化 _============================== --4/8Elements-- Q1 報告された問題は確認できましたか? 確認済み Q2 他の問題は確認できましたか? 無し Q3 問題を解決するために何をおこないましたか? PARTS手配 Q4 対応後の装置は順調ですか? 良好 Q5 その検査はどの様に終了しましたか? 無し Q6 問題により間違った所見（誤診、誤治療）の有無: 無し Q7 意図しない不要な曝射の有無: 無し Q8 アーチファクトの有無: 無し ----------------------------------------------------------------------------- | 5 |  | Deterioration of consumables 453567104591: SINGLE STRAP QTY: 5 SLOC: BackOrder --Subject-- | [12 empty] | 117680995 |  | 996578 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | NotEngineering | [5 empty] | Engineering | [6 empty] | Strap | [5 empty] | Refresh |  |  | System issue | System | NotEngineering | ShouldbeOthers | NotEngineering-ShouldbeOthers | NotEngineering-ShouldbeOthers | NotEngineering |  | ShouldbeOthers | NotEngineering-ShouldbeOthers | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | y
1028 |  |  | Japan | Maki Hospital | 341017 | 44357 |  | Applications request. >> Koizumi will respond. | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Application Support Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer : アプリ依頼>>小泉さん対応 Safety Questions : Case: 0117680088 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: TAKEYUKI MIYAKAWA Last updated by: TAKEYUKI MIYAKAWA 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: Problem description by engineer : 対象装置:CT 故障個所:07その他 最初に発生したのはいつか?(~以内):Unknow 問題発生の頻度は?:Unknow エンドユーザｰのワークフローのどの時点において問題が発生しているのか?:その他(Other) 最近どのようなイベントが発生したか?:N.A. 顧客は問題について具体的なシステムイベントレポートを作成しているか?:N.A. Engineer_s Comment: _============================== コール内容 患者名を間違えたので、修正方法を教えてください。 _============================== Troubleshooting Action: What steps are taken during the system diagnosis : 対応後の結果:改善された 故障個所:07その他 診断中にどのようなステップを実施したか?: _================================== アプリ担当者回答 change Patient Detailから一緒に行っていただきました。 _================================== Error # and/or description of error : NA Repair Action: Completed or Recommended Repair Action "External" : 【現象確認】 確認済み 【作業内容詳細】 質問内容を回答 【確認内容】 良好 OneEMS internal remarks : Problem description by engineer : 対象装置:CT 故障個所:07その他 最初に発生したのはいつか?(~以内):Unknow 問題発生の頻度は?:Unknow エンドユーザｰのワークフローのどの時点において問題が発生しているのか?:その他(Other) 最近どのようなイベントが発生したか?:N.A. 顧客は問題について具体的なシステムイベントレポートを作成しているか?:N.A. Engineer_s Comment: _============================== コール内容 患者名を間違えたので、修正方法を教えてください。 _============================== What steps are taken during the system diagnosis : 対応後の結果:改善された 故障個所:07その他 診断中にどのようなステップを実施したか?: _================================== アプリ担当者回答 change Patient Detailから一緒に行っていただきました。 _================================== Error # and/or description of error : NA Completed or Recommended Repair Action "External" : 【現象確認】 確認済み 【作業内容詳細】 質問内容を回答 【確認内容】 良好 Completed or Recommended Repair Action "Internal" : --Troubleshooting Action-- 対応後の結果:改善された 故障個所:07その他 診断中にどのようなステップを実施したか?: _================================== アプリ担当者回答 change Patient Detailから一緒に行っていただきました。 _================================== --Repair Action-- --Subject-- 07CS:TM:ﾄｶﾞﾜ様:アプリ問い合わせ。患者名を間違えたので、修正方法を教えてください。>>アプリ担当者対応>>change Patient Detailから一緒に行っていただきました。 --Problem Description-- 対象装置:CT 故障個所:07その他 最初に発生したのはいつか?(~以内):Unknow 問題発生の頻度は?:Unknow エンドユーザｰのワークフローのどの時点において問題が発生しているのか?:その他(Other) 最近どのようなイベントが発生したか?:N.A. 顧客は問題について具体的なシステムイベントレポートを作成しているか?:N.A. Engineer_s Comment: _============================== コール内容 患者名を間違えたので、修正方法を教えてください。 _============================== --4/8Elements-- Q1 報告された問題は確認できましたか? 確認済み Q2 他の問題は確認できましたか? 無し Q3 問題を解決するために何をおこないましたか? 質問内容を回答 Q4 対応後の装置は順調ですか? 良好 Q5 その検査はどの様に終了しましたか? 無し Q6 問題により間違った所見（誤診、誤治療）の有無: 無し Q7 意図しない不要な曝射の有無: 無し Q8 アーチファクトの有無: 無し SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | Closed - Non-Complaint |  | NAC | [12 empty] | 117680088 |  | 995964 |  |  | Closed - Non-Complaint | [7 empty] | NAC | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1029 |  |  | Japan | Independent administrative institution National Hospital Organization | 341013 | 44356 |  | 03DS:KKI: An error occurred during a contrast-enhanced tracker scan. >> S_GEN_MGR_MA_OUT_OF_RANGE/S_GEN_FAUL [Truncated] | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 03CS:KKI: An error occurred during a contrast-enhanced tracker scan. >> S_GEN_MGR_MA_OUT_OF_RANGE/S_GEN_FAULT_ANODE_TUBE_ARC > He was able to continue the examination. Safety Questions:Case: 0117673753 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: KAZUMASA KIMURA Last updated by: KAZUMASA KIMURA 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Support External Remarks : 10.06.2021 06:21:07 UTC USHIO TAKADA (81100646) Request: Arcing occurred during a scan. Problem confirmation: I confirmed it. I confirmed it in the log. Work details: Arcing occurred, so today I adjusted the filament output. When I finished, I performed a test scan and confirmed that arcing did not occur and there was no problem with the scanner operation. Other problems/Measures taken: None. Equipment status The equipment can be used. Work status Work finished. 09-Jun-2021 81100646 FSE Handover Tool1.3.4.32612.DB13 20210115 Internal Remarks:11.06.2021 07:03:30 UTC USHIO TAKADA (81100646) Replacement parts: No replacement parts. [Report on mistakes in parts order] No mistakes. Situation monitoring information: Not monitoring the situation TS Request: No parts to return. Additional CT information: Q1. How did the examination finish? Q1. No (not applicable) Q2. Were there any mistaken findings (misdiagnosis, mistaken treatment) due to the problem? None. Q3. Were there any artifacts? None. Q4. Was there any unintended unnecessary exposure? None. Repair sticker: I affixed a Repaired sticker. [Other information to share] 1. Was treatment/quarantine carried out for a patient who has or is suspected of having new coronavirus? No 2. Is it necessary to go into treatment rooms/beds where there are infected patients?(Including workplace and routes of movement) No Did you agree on how to complete the report with the person in charge of devices? Yes. I completed normal reporting that day. 11-Jun-2021 81100646 FSE Handover Tool1.3.4.32612.DB13 20210115 ----------------------------------------------------------------------------- | 4 |  | An arc occurred during the scan. [Phenomenon confirmation] I was able to confirm. I confirmed it in the log. [Details of work] Since an arc was generated, we have adjusted the filament output today. After the work, a test scan is carried out to confirm that no arc is generated and that there is no problem in the operation of the equipment. | [12 empty] | 117673753 |  | 991971 | TRUE | Investigation | Investigation |  | CTUDT00630085 | [5 empty] | Engineering | [6 empty] | Tube Arc | [5 empty] | Refresh |  |  | System issue | System | IC-HW | Tub Arc | IC-HW-Tub Arc | IC-HW-Tub Arc | CTUDT00630085 | Tub Arc | IC-HW | IC-HW-Tub Arc | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | y
1030 |  |  | China | Hebei Province Geriatric Hospital | 345015 | 44355 |  | PRS connection/02N/Ingenuity Core128/Elite China/ | BU Q&R Feedback: No Reason: Customer Relations Problem Reported by Customer : PRS connection Safety Questions : Case: 0117665791 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingli Last updated by: Lu Mingli 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action : FSE 孙建平 018031820728 call in and disp him SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 08.06.2021 06:27:32 UTC Jianping Sun (86039023)服务内容： [连接PRS。] 服务结果： [设备正常工作。] Internal Remarks : 10.06.2021 14:31:35 UTC Jianping Sun (86039023) RDI Accuracy：yes. $Q7:(Confirmed the reported problem. PRS install） $Q8:(PRS install) $Q9:(PRS install ) $Q10:(Test and Inspection data refer to PA.Device returened to full futionnallity.) $Q11:(Calibrated tools name（FLUKE287）S/N（20060013）Next calibration date:(05/10/2022)) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date:N/A ^Current Tube usage:N/A ----------------------------------------------------------------------------- | Zhai, Lisa |  | 44356.5993 | [12 empty] | 117665791 |  | 987202 |  |  | Closed - Non-Complaint | [14 empty] | 44359.0319 | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1031 |  |  | Japan | 独立行政法人国立病院機構 | 341015 | 44355 |  | 07CS:RS:心臓のアプリに関して担当者と連絡希望>>APPサポート時間外の為、明日ご連絡いただけるとのこと | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Customer Information Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer : 07CS:RS:心臓のアプリに関して担当者と連絡希望>>APPサポート時間外の為、明日ご連絡いただけるとのこと Safety Questions : Case: 0117668167 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: RUI SATAKE Last updated by: RUI SATAKE 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | Closed - Non-Complaint |  | NAC | [12 empty] | 117668167 |  | 987965 |  |  | Closed - Non-Complaint | [7 empty] | NAC | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1032 |  |  | Japan | Social Medical Corporation BOKOI, Nikko Memorial Hospital | 341011 | 44354 |  | JY: The bolus tracking was not triggered even when the threshold was exceeded. | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : JY: The bolus tracking was not triggered even when the threshold was exceeded. Safety Questions:Case: 0117658722 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: YAN Joanna Last updated by: YAN Joanna 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A OneEMS internal remarks: Sugahara DB SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Support ----------------------------------------------------------------------------- | 4 |  | The case was reviewed by Console team. Here is the investigation details. Observations: The tracker scan completed successfully. Not observed a big delay between end of tracker to the clinical scan start. It is possible that there could be a extra tracker shot can happen The skip command to gantry also went in the same time. All the series 15, 16 and 17 are successfully completed without any error. Conclusion: Not seen any delay or issue in tracker and subsequent clinical scans System is working as expected. | [12 empty] | 117658722 |  | 982397 | TRUE | Evaluation | Closed - No Further Investigation |  | NotDefect | [5 empty] | Engineering | [6 empty] | Bolus Tracking | [5 empty] | Refresh |  |  | System issue | System | Console-SW | NotDefect | Console-SW-NotDefect | Console-SW-NotDefect | NotDefect |  | Console-SW | Console-SW-NotDefect | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | y
1033 |  |  | Sweden | Falu Lasarett | 40049 | 44355 |  | Application does not work, background turns yellow, but works when restart is done | Work Order #: WO-03618921 Knudsen Martin The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Intermittent user get_s yellow screen and application dissapears. Follow Up Required : Intermittent yellow screen and application disappear. No errors found in logfiles. SW reload is first step. Please book FSE 7 hours for this. No parts. -MK Diagnostic performed by Engineer : Intermittent user get_s yellow screen and application dissapears. Reboot solves issue. Checked logfiles, and was not able to see any errors related. Advised to do a SW reload of the Host computer. ----------------------------------------------------------------------------- | Evaluation |  | Yellow screen issue, SW version is not clear, suggest field to implement new SP1 SR FCO72800758 | [12 empty] | 117667610 |  | 987581 | TRUE | Evaluation | Evaluation |  | CTUDT00632520 | [5 empty] | Engineering | [6 empty] | New Yellow Screen | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | y
1034 |  |  | China | 内蒙古阿拉善右旗人民医院 | 345031 | 44357 |  | 02N/Ingenuity Core128/Elite China/The gantry cann_t power up | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity Core128/Elite China T2 Activities : Problem Description: 客户问题概述： 机架不能上电 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： S_MASTER_SEQUENCER_GDC_LOW_AFTER_TIMEOUT 远程诊断与测试结果： 联系工程师苏瑞杰，怀疑电源刷短路，保险丝问题 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议工程师到场检查清理滑环，检查电源是，保险丝 建议备件12NC及名称： 保险丝：459800043431，459800448931 Safety Questions : Case: 0117679129 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Weng Jinming Last updated by: Weng Jinming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : PPD 459800448931*2 459800043431*3备件已系统下单安排24H到场，FSE：苏瑞杰-86-13304719335 External Remarks : 11.06.2021 07:02:43 UTC Sheng Cui (86038219)"SUAT. Ready for Clinical Use local time" : "2021-06-11 16:00" External Remarks : 12.06.2021 01:25:13 UTC Sheng Cui (86038219)服务内容：机架无法上电，Softstart 无法吸合，复位交流接触器，更换滑环高压保险和滑环保险。 服务结果：维修后，设备恢复正常工作 Internal Remarks : 12.06.2021 01:25:16 UTC Sheng Cui (86038219) RDI Accuracy（Yes） $Q7: (Confirmation & Description of problem:gantry boot up failed )$Q8: (On site check logeer message GDC low ,need to replace slip ring 50A fuse and HV generator 20A fuse.) $Q9: (Replace slip ring 50A fuse and HV generator 20A fuse. ) $Q10: (Test and inspection data refer to: PA. Device returned to full functionality.) $Q11: Calibrated Tools name (FLUKE287 ) S/N (12470209 ) Next Calibration date (07/12/2021）: For Mainland China CV / XR, CT engineers only. Please fill in: Current Tube installed date Tube installed date: ##09/ 4/ 2020 ^Current Tube usage: scan second_##401202S ----------------------------------------------------------------------------- | 2 |  | 机架无法上电，Softstart 无法吸合，复位交流接触器，更换滑环高压保险和滑环保险。 459800066152 FUSE, 30A/600V AJT30 3 24H 459800726611 FUSE, 2.5A/500V FNQ-2-1/2 2 24H 459800043431 Fuse, 50A, 600VAC, Dual Element TD" 3 24H 459800448931 20 AMP FUSE 1 24H已出货 | [12 empty] | 117679129 |  | 995664 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Operation | [6 empty] | Power Brush assembly WI | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1035 |  |  | Turkey | ADIYAMANOZELSAGINSTURTASSANVETICA.S | 341006 | 44358 |  | artifact occurred in cardiac extraction | Work Order #: WO-03641560 Turgay Baylan The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : Case escalation was performed in order to examine the images in the cardiac extraction. ----------------------------------------------------------------------------- | Investigation |  | There is any image problem for 130 HR, but customer complains images for 82 HR. | [12 empty] | 117688189 |  | 998252 | TRUE | Investigation | Investigation - Pending Translation |  | TBD_Bugreport | [5 empty] | Engineering | [6 empty] | Cardiac | [5 empty] | Refresh |  |  | System issue | System | TBD-Bugreport | TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | TBD_Bugreport |  | TBD-Bugreport | TBD-Bugreport-TBD_Bugreport | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | y
1036 |  |  | Turkey | Republic of Turkey Ministry of Health, Bagcilar | 40010 | 44356 |  | The belt is noisy. | Work Order #: WO-03627268 09-Jun-2021 Ugur Altunoz The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : Gantry main drive belt has a crack on it and it should be replaced. Follow Up Required : Gantry main drive belt replacement. ----------------------------------------------------------------------------- | Evaluation |  | There is a crack on the gantry main drive belt and there is a sound. Belt replacement is required. Gantry main drive belt will be replaced. | [12 empty] | 117675108 |  | 989138 | TRUE | Evaluation | Evaluation |  | ECR-077012 | [5 empty] | Engineering | [6 empty] | Belt | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Belt | Gantry-HW-Belt | Gantry-HW-Belt | ECR-077012 | Belt | Gantry-HW | Gantry-HW-Belt | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | y
1037 |  |  | Turkey | Republic of Turkey Ministry of Health, Bagcilar | 40010 | 44355 |  | It interrupts the scanning. | "Work Order #: WO-03620575 Salih Mendi The system does not meet specification and is out of use Problem Description by Engineer : Customer Function/Role: technician Device Use: out of use Expected and Unexpected Results: S_CTRL_LINK_GHOST_CRC_FRAMING_ERROR User Impact: No user impact Patient Impact: no patient impact Current Software Version:NA Diagnostic performed by Engineer : Upon the user_s complaint that ""The device interrupts the recording"", the system was examined by connecting to the system via remote access. Upon inspection, it was observed that the system gave the error S_CTRL_LINK_GHOST_CRC_FRAMING_ERROR. A site visit is required for detailed examination of Slipring, Ghost, DMC and laser systems. Follow Up Required: A site visit is required for detailed examination of Slipring, Ghost, DMC and laser systems. ---------------------------------------------------------------- --------------------------- Work Order #: WO-03621501 09-Jun-2021 Ugur Altunoz The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: technician Device Use: out of use Expected and Unexpected Results: S_CTRL_LINK_GHOST_CRC_FRAMING_ERROR User Impact: No user impact Patient Impact: no patient impact Current Software Version:NA Resolution : Due to CRC frame errors in the device, slipring was cleaned and the images were checked. Trial shots were made on the device and it was delivered in an active condition. | Closed - Done |  | Due to CRC frame errors in the device, slipring was cleaned and the images were checked. Trial shots were made on the device and it was delivered in an active condition. | [12 empty] | 117669046 |  | 984579 | TRUE | Closed - Done | Closed - Done |  | NotEngineering-SlipringClean | [5 empty] | Engineering | [6 empty] | Slipring | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | SlipRingClean | Gantry-HW-SlipRingClean | Gantry-HW-SlipRingClean | NotEngineering-SlipringClean |  | Gantry-HW | Gantry-HW-SlipRingClean | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | y
1038 |  |  | United States | Philips Medical Capital | 40055 | 44360 |  | 0609CDN84 knocking noise when gantry tilted | Work Order #: WO-03647989 Mcginnis, Mike The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Problem description by engineer : knocking noise when gantry tilted When was the first occurrence? Within... : 13-Jun-2021 How often is issue occurring? : First occurrence Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: knocking noise when gantry tilted Patient Impact/Outcome: none User Impact: knocking noise when gantry tilted Current Software Version: Not provided Follow Up Required : Error # and/or description of error : knocking noise when gantry tilted Completed or Recommended Repair Action "Internal" : Check the covers for loose or missing bolts. If none found open the cover and isolate the noise and repair as needed. Internal Comments : Error # and/or description of error : knocking noise when gantry tilted Completed or Recommended Repair Action "Internal" : Check the covers for loose or missing bolts. If none found open the cover and isolate the noise and repair as needed. Diagnostic performed by Engineer : Melissa said when she tilts the gantry posterior she hears a knocking noise from the gantry. She said she feels comfortable using the system but would like it looked at. ----------------------------------------------------------------------------- | Closed - No Further Investigation |  | Tilted the gantry to reproduce the noise experienced by the customer. I heard the covers rubbing against each other and cusing a growling noise. Adjusted the cover to resolve the noises. | [12 empty] | 117693213 |  | 1000880 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-GantryNoise | [5 empty] | Engineering | [6 empty] | Gantry cover misalignment | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Noise | Gantry-HW-Noise | Gantry-HW-Noise | TBD-GantryNoise |  | Gantry-HW | Gantry-HW-Noise | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | y
1039 |  |  | United States | Department of Veterans Affairs | 40018 | 44357 |  | 5-Scheduled Activity | Work Order #: WO-03639047 Powell, Jamie The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : Perform safety clip check/inspection/repair for FSN/FCO 72800752.. Follow Up Required : Perform safety clip check/inspection/repair for FSN/FCO 72800752.. ----------------------------------------------------------------------------- | 5 |  | FCO72800752 | [12 empty] | 117685359 |  | 995412 | TRUE | Evaluation | Closed - non-complaint | [7 empty] | Others | [6 empty] | FCO72800752 | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1040 |  |  | China | Beijing City Hospital of Traditional Chinese and Western Medicine | 40032 | 44354 |  | 02N/Ingenuity CT 728326//ask engineer onsite information to customer about installation | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 02N/Ingenuity CT 728326 Next Step Action: 013811972406 Yu called T2 Activities:Problem Description: Summary of customer_s problems: Other types of error. Detailed description of Customer_s problems or malfunctions: Yu called in, need labor dispatch to the field to give the director an introduction of the machine_s leftovers after installation, and the simple switch on/off. It is a technology problem with the equipment itself: No. Troubleshooting Action: key fault code or symptom: Sales called in, need labor dispatch to the field to give the director an introduction of the machine_s leftovers after installation, and the simple switch on/off. Remote troubleshoot and test result: The customer needs an engineer to go to the field. Contact Director Zhang PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: introduction on site. Safety Questions:Case: 0117660029 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Jianqiang Last updated by: Wang Jianqiang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support External Remarks: 09.06.2021 07:02:13 UTC Guochao Fan (86039456) service details: work with the hospital to organize and clean the scan room, completed service result: device returned to normal working condition after repairing Internal Remarks: 09.06.2021 07:02:18 UTC Guochao Fan (86039456) "RDI Accuracy:Yes" $Q7: Confirm reported problems:Equipment safety hazard detection $Q8: Trouble shooting steps:for Equipment safety hazard detection. Equipment normal $Q9: NA $Q10: Test and Inspection data refer to PA. Device status:Device returned to full functionally $Q11: N/A ----------------------------------------------------------------------------- | Zhai, Lisa |  | 44355.7271 | [12 empty] | 117660029 |  | 983225 |  |  | Closed - Non-Complaint | [14 empty] | 44358.9903 | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1041 |  |  | China | 陕西省咸阳中心医院 (咸阳市第二人民医院） | 40050 | 44354 |  | 02N/Ingenuity CT 728326/asking parameters | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: Instructed Customer Problem Reported by Customer : 02N/Ingenuity CT 728326 FSE Notes : This PO is a Reference Number only Safety Questions : Case: 0117658800 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: customer asking CT parameters Troubleshooting Action: customer asking CT parameters,max KV,max mA, answer it. Repair Action: customer asking CT parameters,max KV,max mA, answer it. Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | Zhai, Lisa |  | 44355.5924 | [12 empty] | 117658800 |  | 982471 |  |  | Closed - Non-Complaint | [14 empty] | 44357.6146 | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1042 |  |  | Japan | Tomakomai City Hospital | 341014 | 44354 |  | He wanted to contact Sugahara about the arcing that occurred the other day. >> Sugahara called him back. | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer: He wanted to contact Sugahara about the arcing that occurred the other day. >> Sugahara called him back. Safety Questions:Case: 0117658774 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: TAKEYUKI MIYAKAWA Last updated by: TAKEYUKI MIYAKAWA 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | Zhu, Chun |  | 44355.5889 | [12 empty] | 117658774 |  | 982452 |  |  | Closed - Non-Complaint | [14 empty] | 44361.4062 | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1043 |  |  | United States | Willis Knighton Hospital | 40004 | 44356 |  | Infant Phantom has Bad Seal | Work Order #: WO-03631273 11-Jun-2021 Williams Spence The system fully meets specification and returned to use Parts Consumed : 1.00000,453567047262,INFANT PHANTOM Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : Phantom leaked. Replaced defective infant phantom ----------------------------------------------------------------------------- | 5 |  | 1.00000,453567047262,INFANT PHANTOM Phantom leaked. Replaced defective infant phantom | [12 empty] | 117678304 |  | 995113 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | NotEngineering-PartReplace | [5 empty] | Engineering | [6 empty] | Infant phantom | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | Phantom | Couch-HW-Phantom | Couch-HW-Phantom | NotEngineering-PartReplace |  | Couch-HW | Couch-HW-Phantom | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | y
1044 |  |  | United States | Odessa Regional Medical Center | 40029 | 44355 |  | The contrast issue, performed air calibration and rebooted the system | Work Order #: WO-03623673 Atz, Michael The system fully meets specification and returned to use Internal Comments : Error # and/or description of error : none Completed or Recommended Repair Action "Internal" : Advised to run an air cal now that the contrast is cleaned off of the gantry and to reboot the system to the wall contrast on the gantry she will call back if further service is needed. Diagnostic performed by Engineer : Advised to run an air cal now that the contrast is cleaned off of the gantry and to reboot the system to the wall contrast on the gantry she will call back if further service is needed. Problem Description by Engineer : Problem description by engineer : contrast on the gantry now a ring artifact air cal fails When was the first occurrence? Within... : 08-Jun-2021 How often is issue occurring? : Unknown Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: No further information is available as per the customer Patient Impact/Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer Current Software Version: Not provided Resolution : system is operational ----------------------------------------------------------------------------- | Closed - Done |  | Advised to run an air cal now that the contrast is cleaned off of the gantry and to reboot the system to the wall contrast on the gantry she will call back if further service is needed. | [12 empty] | 117671425 |  | 985996 | TRUE | Closed - Done | Closed - Done | [7 empty] | Customer | [6 empty] | Contrast | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1045 |  |  | United States | Philips Medical Capital | 40005 | 44354 |  | New EMR Procedure Codes | Work Order #: WO-03612662 Mujica, David The system fully meets specification and returned to use Problem description by Clinical Support Specialist : Customer Function/Role: Technologist/Manager Device Use: In use Expected and Unexpected Results: map accession numbers to new orders on new system User Impact: none Patient Impact: none Current Software Version: 4.1 Diagnostic performed by Engineer : Kelly indicated that she needed help associating newly populated orders to existing protocols/exam cards on their CT system. Kelly was not in front of the CT system and could not be instructed on where to make these changes. A follow up call is required when Kelly is in front of the CT system, so she can be shown where this change to the existing exam cards/protocols can be made. The follow up is scheduled for Thursday 6/10. No further clinical instruction is needed at this time. ----------------------------------------------------------------------------- | 5 |  | NAC | [12 empty] | 117661514 |  | 984359 |  |  | Closed - Non-Complaint | [7 empty] | NAC | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1046 |  |  | Chile | Scotiabank Chile | 40008 | 44361 |  | CL1675-THE COUCH REFERENCE LINE IS NOT SHOWN IN THE SCANS | Work Order #: WO-03653507 Valenzuela Velasquez, Eduardo Ariel The system fully meets specification and returned to use Problem Description by Engineer : Problem description by engineer : No aparece linea de referencia en imágenes de surview. Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected results: Linea de referencia visible en surview. Patient Impact/Outcome: Dificultad para centrar exámenes. User Impact: Dificultad para centrar exámenes. Current Software Version: 4.1.10 Resolution : The options to enable the reference are reviewed with the support of applications. It is activated with the right mouse button on the surview image. Tests are performed and the requested reference is achieved. Internal Comments: Error # and / or description of error: There is no referral line on the surview.--------------- |  |  | The options to enable the reference are reviewed with the support of applications. It is activated with the right mouse button on the surview image. | [12 empty] | 117698098 |  | 1004078 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Application | [6 empty] | Reference line | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1047 |  |  | China | 陕西省咸阳中心医院 (咸阳市第二人民医院） | 40050 | 44364 |  | 02n/Ingenuity CT 728326/Add the PACS | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02n/Ingenuity CT 728326 Safety Questions : Case: 0117716668 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Weng Jinming Last updated by: Weng Jinming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 医院需要添加PACS 客户问题详细描述或故障现象： 用户名，密码一直不断 设备本身技术问题： 是。 Troubleshooting Action: 远程诊断与测试结果： 联系不上工程师，网络不通，不能远程授权 PRS TR状态： 不通，请工程师到场检查。 Repair Action: 建议现场维修步骤： 需要工程师到场设置网络 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support ----------------------------------------------------------------------------- | 5 |  | confirmed report problem,customer want at ingenuity ct add PACS, customer cannot transfer image to pacs $Q8:remort contact the pacs fse ,on the ct login IT account ,and can add pacs dicom parameter , test pacs is normal ,device normal work | [12 empty] | 117716668 |  | 1020727 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | PACS config | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1048 |  |  | China | 河南省濮阳市人民医院 | 40031 | 44364 |  | 02N/Ingenuity CT 728326/ Login interface locked | BU Q&R Feedback: No SAP Customer Symptom Code: CENV Customer Symptom Code: Customer Environmental Concern Reason: Corrective Maintenance Engineer Failure Code: Environmental Cause Repair Activity Code: Instructed Customer Safety Questions : Case: 0117715562 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Guomei Li Last updated by: Guomei Li 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Problem Reported by Customer : Login interface locked T2 Activities : Problem Description: CT账户密码输错三次锁定了 Troubleshooting Action: 指导客户使用IT账户重置密码或等待半小时后重新输入正确密码 Repair Action: 指导客户重置密码或等待一会儿后输入正确密码 Test/Inspection: N/A Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- |  |  | 指导客户使用IT账户重置密码或等待半小时后重新输入正确密码 | [12 empty] | 117715562 |  | 1019849 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Password | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1049 |  |  | China | 河南省濮阳市人民医院 | 40031 | 44364 |  | Wechat:02N/Ingenuity CT 728326/The code is locked | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Software Defect or Virus Repair Activity Code: Instructed Customer Problem Reported by Customer : The code is locked Safety Questions : Case: 0117715480 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Bai Jingwen Last updated by: Bai Jingwen 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: CT password locked Troubleshooting Action: waiting for password unlock. Repair Action: waiting for password unlock. system is ok. Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- |  |  | waiting for password unlock. system is ok. | [12 empty] | 117715480 |  | 1019781 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Password | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1050 |  |  | China | 浙江省乐清第一人民医院 | 40054 | 44362 |  | Wechat:02N/Ingenuity CT 728326/gantry has no power | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : gantry has no power Safety Questions : Case: 0117699800 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Tao Fengming Last updated by: Tao Fengming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 无法加电 客户问题详细描述或故障现象： 无法加电 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： lifesync p1=8, 远程诊断与测试结果： 无法加电，不出钥匙框，CCC无响应，重启无效 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 检查CCC本身及供电，检查接线，检查CCC上power fail是否亮，若亮参考newsletter排查 建议备件12NC及名称： " 453566503742 ASSY, 4 OUTPUT DC SUPPLY； 459800477572 Couch control board with CONFIG-B CPM" SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : PPD 459800477572 1&453566503742 1已安排12H到医院 黄逸-86-15958725282 External Remarks : 16.06.2021 13:57:59 UTC SUP JCOUSER (86036249)服务内容：机架不能上电，lifesync8，测量四输出正常，CCC板在00和22切换，更换AMC。 服务结果：维修后，设备恢复正常工作。 其它重要事项：无 Internal Remarks : 16.06.2021 13:58:02 UTC Yi Huang (86036249) RDI Accuracy:Yes $Q7: (Confirmed reported problem&estop can not close) $Q8: (check errorlog found lifesync 22 and ccc config end bkst parameter is 2) $Q9: (replace the AMC) $Q10: (Test and inspection data refer to: PA.Device returned to full functionality.) $Q11: Calibrated tools name(N/A) S/N (N/A) Next calibration date (N/A) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date:N/A ^Current Tube usage: second:N/A External Remarks : 17.06.2021 01:44:56 UTC Yi Huang (86036249)"SUAT. Ready for Clinical Use local time" : "2021-06-15 21:00" ----------------------------------------------------------------------------- | 2 |  | System works well after replace the AMC, CTUDT00632187 | [12 empty] | 117699800 |  | 1008995 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632187 | [5 empty] | Engineering | [6 empty] | Horizontal Movement | [5 empty] | Refresh |  |  | System issue | System | Couch-SW | Estop | Couch-SW-Estop | Couch-SW-Estop | CTUDT00632187 |  | Couch-SW | Couch-SW-Estop | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | y
1051 |  |  | China | 莱阳市人民医院 | 345005 | 44365 |  | CT systsm stuck, cannot scan. |  | WANG, Peter |  | 44365.6236 | [14 empty] | 1019538 | TRUE | Investigation | Closed - Done | [7 empty] | Engineering | [6 empty] | systsm stuck | [5 empty] | Refresh |  |  | System issue | System | TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | 0 |  |  |  | No | 0 | 0 | 0 | 0 | 0 | No | y | y | y | y | y | y | y | y | y | y | #REF!
1052 |  |  | China | 苏州广慈肿瘤医院有限公司 | 345019 | 44363 |  | Wechat:02N/Ingenuity Core128/Elite China/update PACS | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : update PACS Safety Questions : Case: 0117709352 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Tao Fengming Last updated by: Tao Fengming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 医院PACS系统更新升级，需要设备工程师协助 设备本身技术问题： 是。 Troubleshooting Action: 远程诊断与测试结果： 医院PACS系统更新升级，需要设备工程师协助 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 医院PACS系统更新升级，需要设备工程师协助 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 17.06.2021 07:23:16 UTC Haotian Xu (86090280) 服务内容：主机传输pacs异常，配合pacs工程师修改参数，测试。 服务结果：设备正常运行。 Internal Remarks : 17.06.2021 10:25:45 UTC Haotian Xu (86090280) RDI Accuracy: Yes $Q7: (Confirm reported problem,host delivered image to PACS slowly and sometimes failed.) $Q8: ( check the network and pacs parameter） $Q9: (change the pacs parameter accourding to service manual .) $Q10: (Test and inspection data refer to PA. Device returned to full functionality.) $Q11: Calibrated tools refer to PA. ----------------------------------------------------------------------------- | 5 |  | 服务内容：主机传输pacs异常，配合pacs工程师修改参数，测试。 服务结果：设备正常运行。 | [12 empty] | 117709352 |  | 1015512 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | PACS config | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1053 |  |  | China | 新疆伊犁州友谊医院 | 40048 | 44363 |  | 02N/Ingenuity CT 728326/ISP workflow error | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer : 02N/Ingenuity CT 728326 Safety Questions : Case: 0117707422 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Cao Last updated by: Liu Cao 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: Cannot login ISP LX with workflow error Troubleshooting Action: System is newly installed don_t has network connection Repair Action: Connect network and get IP,then issue solved Test/Inspection: N/A Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | Lu, Xueying |  | System is newly installed don_t has network connection Repair Action: Connect network and get IP,then issue solved | [12 empty] | 117707422 |  | 1014076 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Network setting | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1054 |  |  | China | 河南许昌市中心医院(传染病) | 40047 | 44363 |  | 02N/Ingenuity CT 728326/can not close estop | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: Instructed Customer Problem Reported by Customer : 02N/Ingenuity CT 728326 Safety Questions : Case: 0117707257 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Zhang Caiqin Last updated by: Zhang Caiqin 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: can not close estop Troubleshooting Action: guide customer check emergency switch, it is pressed Repair Action: release this switch, then close estop loop normally Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: NA SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | Gao, Mia |  | Troubleshooting Action: guide customer check emergency switch, it is pressed Repair Action: release this switch, then close estop loop normally | [12 empty] | 117707257 |  | 1013977 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | E-Stop button | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1055 |  |  | China | Hebei Province Tangshan City Workers Hospital | 40044 | 44362 |  | WeChat:02N/Ingenuity CT 728326/Couch fault | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : Couch fault Safety Questions : Case: 0117699667 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Jianqiang Last updated by: Wang Jianqiang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 床偶发报错位置错误 客户问题详细描述或故障现象： 床偶发报错位置错误 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 床偶发报错位置错误 远程诊断与测试结果： 重启可恢复，但是偶发影响使用 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 检查拉线定位器，检查CCC 的控制，如有必要进行更换 建议备件12NC及名称： BRILLIANCE COUCH CNTL PCB ASSY 459800477572// Vertical Transducer 459800623331 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : PPD备件Vertical Transducer 459800623331*1&Couch control board with CONFIG-B CPM 459800477572*1已安排48H到场地，请派工，FSE：徐国梁-86-13869524913 External Remarks : 17.06.2021 11:32:18 UTC Guoliang Xu (86038283) 服务内容：垂直位置故障，更换电位器。 服务结果：维修后观察使用中。 Internal Remarks : 17.06.2021 11:32:21 UTC Guoliang Xu (86038283) $Q1: (Y) $Q2: (N) $Q3: () $Q4: () $Q5: () $Q6: （Couch fault ）RDI Accuracy ：Yes $Q7:Confirmed reported problem，couch position error . $Q8: Couc vertical position fault，.$Q9: Replaced pot transducer. $Q10: Test and inspection data refer to: PA. Device returned to full functionality. $Q11: Calibrated Tools name ( Multimeter Fluke 287C) S/N (15930163) NextCalibration date (04/28/2022) For Mainland China CV / XR, CTengineers only. Please fill in: ^ Current Tube installed date03/13/2021 ^Current Tube usage: scan second_20041 ----------------------------------------------------------------------------- | Lu, Xueying |  | Confirmed reported problem，couch position error . $Q8: Couc vertical position fault，.$Q9: Replaced pot transducer (Vertical Transducer 459800623331*1). $Q10: Test and inspection data refer to: PA. Device returned to full functionality. | [12 empty] | 117699667 |  | 1008884 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | NotEngineering-partreplace | [5 empty] | Engineering | [6 empty] | Vertical Transducer | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | VerticalTransducer | Couch-HW-VerticalTransducer | Couch-HW-VerticalTransducer | NotEngineering-partreplace |  | Couch-HW | Couch-HW-VerticalTransducer | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | y
1056 |  |  | China | 宁安市第二人民医院 | 345039 | 44363 |  | 02N/Ingenuity Core128/Elite China/forgot the admin password | BU Q&R Feedback: No SAP Customer Symptom Code: CENV Customer Symptom Code: Customer Environmental Concern Reason: Corrective Maintenance Engineer Failure Code: Environmental Cause Repair Activity Code: Database Recreation / Recovery Next Step Action : ASP yiyuanji 015604188333 call in Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ Safety Questions : Case: 0117708798 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Gaoge Last updated by: Wang Gaoge 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: olevel_service账户密码忘记了 Troubleshooting Action: 通过远程重置olevel_service账户密码 重置后的密码是 ！1 Repair Action: 重置账户密码 Test/Inspection: N/A Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | Gao, Mia |  | 44364.6896 | [12 empty] | 117708798 |  | 1015072 |  |  | Closed - Non-Complaint | [14 empty] | 44367.4792 | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1057 |  |  | Japan | 紀南病院 | 340000 | 44365 |  | SY:ガントリから異音がする | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : SY:ガントリから異音がする Safety Questions : Case: 0117721677 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Smile YE Last updated by: Smile YE 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A OneEMS internal remarks : 大山DB SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support ----------------------------------------------------------------------------- | 5 |  | There are replacement parts 12Nc: 453566492401 Name: POLY-VEE BELT QTY: 1 | [12 empty] | 117721677 |  | 1024609 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | ECR-077012 | [5 empty] | Engineering | [6 empty] | Belt | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Belt | Gantry-HW-Belt | Gantry-HW-Belt | ECR-077012 | Belt | Gantry-HW | Gantry-HW-Belt | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | y
1058 |  |  | Japan | 社会福祉法人恩賜財団済生会 | 343000 | 44364 |  | 03CS:RS:Short TC時に「Maximum number of arc...」とポップアップし中断>>再実施にて完了>> S_GEN_FAULT_ANODE_TUBE_ARC | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 03CS:RS:Short TC時に「Maximum number of arc...」とポップアップし中断>>再実施にて完了>> ----------------------------------------------------------------------------- | 4 |  | Short TC pops up "Maximum number of arc ..." and interrupts Plan various calibrations at the next inspection. | [12 empty] | 117720230 |  | 1023889 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00630085 | [5 empty] | Engineering | [6 empty] | Tube Arc | [5 empty] | Refresh |  |  | System issue | System | IC-HW | Tub Arc | IC-HW-Tub Arc | IC-HW-Tub Arc | CTUDT00630085 | Tub Arc | IC-HW | IC-HW-Tub Arc | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | y
1059 |  |  | Japan | Independent administrative institution National Hospital Organization | 341013 | 44362 |  | 01CS:KKI:Cannot login to the CT-Host. >> Restarting the CT-Host resolved the problem. | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 01CS:KKI:Cannot login to the CT-Host. >> Restarting the CT-Host resolved the problem. Safety Questions:Case: 0117706588 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: KAZUMASA KIMURA Last updated by: KAZUMASA KIMURA 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 4 |  | Restored by CT host restart | [12 empty] | 117706588 |  | 1013625 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Login | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1060 |  |  | Japan | Tomakomai City Hospital | 341014 | 44362 |  | 03CM:TM: The equipment stopped due to an error during a scan>> SoftArcReg/S_GEN_FAULT_ANODE_TUBE_ARC/S_GEN_F [Truncated] | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: The chest scan stopped due to an error >> Safety Questions:Case: 0117699766 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: TAKEYUKI MIYAKAWA Last updated by: TAKEYUKI MIYAKAWA 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Problem description by engineer : Equipment affected: CT Failure location: 03High Voltage Energy When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user_s work flow did the problem occur?: (During Acquisition/Scan) What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem?: NA Engineer_s Comment: _============================== $version = "v4.1.10.25210" 9:58 Call description: An "X-ray system cannot comply" error message was displayed during a chest scan and the equipment stopped. The necessary range of data was captured, so he didn_t rescan. He asked me to check the log. _============================== Troubleshooting Action: What steps are taken during the system diagnosis: Result of response: Problem solved. Failure location: 03High Voltage Energy What steps did you perform to diagnose the problem? _================================== Results of checking the log: The following error occurred. 2021/06/15_09:48:27.317__03 <*HV*> szInfo: SoftArcReg 0x211* P1:0/P2: 2021/06/15_09:48:27.368 <*HV*> Type: NORMAL ErrNum: S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT szInfo: P1:0/P2: 2021/06/15_09:48:27.376 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_ANALOG_FAULT szInfo:2:5 Analog Fault P1:2/P2:32 2021/06/15_09:48:27.378 [CANCEL] CStateShot:Series 285 cancelled, due to GANTRY_CANCEL_SOURCE_NORMAL_ERROR, in state CStateShot 2021/06/15_09:48:27.415 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_ANODE_TUBE_ARC szInfo: 5:0 APM Arc P1:5/P2: 2021/06/15_09:48:27.418 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_CATHODE_TUBE_ARC szInfo: 5:5 CPM Arc P1:5/P2: 2021/06/15_09:48:29.472 <*ERROR*> Type: NORMAL ErrNum: S_MASTER_SEQUENCER_EXIT_STATE_ERR P1:4096/P2:8369 _================================== Error # and/or description of error: NA Repair Action: Completed or Recommended Repair Action "External": [Confirmation of problem] Confirmed. Work details: I checked the log and reported to the customer. 【Confirmation content】 Good. OneEMS internal remarks : Problem description by engineer : Equipment affected: CT Failure location: 03High Voltage Energy When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user_s work flow did the problem occur?: (During Acquisition/Scan) What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem?: NA Engineer_s Comment: _============================== $version = "v4.1.10.25210" 9:58 Call description: An "X-ray system cannot comply" error message was displayed during a chest scan and the equipment stopped. The necessary range of data was captured, so he didn_t rescan. He asked me to check the log. _============================== What steps are taken during the system diagnosis: Result of response: Problem solved. Failure location: 03High Voltage Energy What steps did you perform to diagnose the problem? _================================== Results of checking the log: The following error occurred. 2021/06/15_09:48:27.317__03 <*HV*> szInfo: SoftArcReg 0x211* P1:0/P2: 2021/06/15_09:48:27.368 <*HV*> Type: NORMAL ErrNum: S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT szInfo: P1:0/P2: 2021/06/15_09:48:27.376 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_ANALOG_FAULT szInfo:2:5 Analog Fault P1:2/P2:32 2021/06/15_09:48:27.378 [CANCEL] CStateShot:Series 285 cancelled, due to GANTRY_CANCEL_SOURCE_NORMAL_ERROR, in state CStateShot 2021/06/15_09:48:27.415 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_ANODE_TUBE_ARC szInfo: 5:0 APM Arc P1:5/P2: 2021/06/15_09:48:27.418 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_CATHODE_TUBE_ARC szInfo: 5:5 CPM Arc P1:5/P2: 2021/06/15_09:48:29.472 <*ERROR*> Type: NORMAL ErrNum: S_MASTER_SEQUENCER_EXIT_STATE_ERR P1:4096/P2:8369 _================================== Error # and/or description of error: NA Completed or Recommended Repair Action "External" : [Confirmation of problem] Confirmed. Work details: I checked the log and reported to the customer. 【Confirmation content】 Good. Completed or Recommended Repair Action "Internal": --Troubleshooting Action-- Result of response: Problem solved. Failure location: 03High Voltage Energy What steps did you perform to diagnose the problem? _================================== Results of checking the log: The following error occurred. 2021/06/15_09:48:27.317__03 <*HV*> szInfo: SoftArcReg 0x211* P1:0/P2: 2021/06/15_09:48:27.368 <*HV*> Type: NORMAL ErrNum: S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT szInfo: P1:0/P2: 2021/06/15_09:48:27.376 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_ANALOG_FAULT szInfo:2:5 Analog Fault P1:2/P2:32 2021/06/15_09:48:27.378 [CANCEL] CStateShot:Series 285 cancelled, due to GANTRY_CANCEL_SOURCE_NORMAL_ERROR, in state CStateShot 2021/06/15_09:48:27.415 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_ANODE_TUBE_ARC szInfo: 5:0 APM Arc P1:5/P2: 2021/06/15_09:48:27.418 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_CATHODE_TUBE_ARC szInfo: 5:5 CPM Arc P1:5/P2: 2021/06/15_09:48:29.472 <*ERROR*> Type: NORMAL ErrNum: S_MASTER_SEQUENCER_EXIT_STATE_ERR P1:4096/P2:8369 _================================== --Repair Action-- --Subject-- 03CM:TM:Mr. Yuki: The equipment stopped due to an error during a scan>> SoftArcReg/S_GEN_FAULT_ANODE_TUBE_ARC/S_GEN_FAULT_CATHODE_TUBE_ARC --Problem Description-- Equipment affected: CT Failure location: 03High Voltage Energy When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user_s work flow did the problem occur?: (During Acquisition/Scan) What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem?: NA Engineer_s Comment: _============================== $version = "v4.1.10.25210" 9:58 Call description: An "X-ray system cannot comply" error message was displayed during a chest scan and the equipment stopped. The necessary range of data was captured, so he didn_t rescan. He asked me to check the log. _============================== --4/8Elements-- Q1. Was the reported problem confirmed? Confirmed. Q2. Were any other problems confirmed? None. Q3. What did you do to resolve the problem?: I checked the log and reported to the customer. Q4: Is the equipment operating normally following the response? Good. Q5. How did the examination finish? None. Q6. Were there any mistaken findings (misdiagnosis, mistaken treatment) due to the problem? None. Q7. Was there any unintended unnecessary exposure? None. Q8. Were there any artifacts? None. ----------------------------------------------------------------------------- | 4 |  | Tube replacement | [12 empty] | 117699766 |  | 1008972 | TRUE | Investigation | Investigation |  | NotEngineering | [5 empty] | Engineering | [6 empty] | Tube | [5 empty] | Refresh |  |  | System issue | System | NotEngineering | ShouldbeOthers | NotEngineering-ShouldbeOthers | NotEngineering-ShouldbeOthers | NotEngineering |  | NotEngineering | NotEngineering-ShouldbeOthers | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | y
1061 |  |  | Japan | Tomakomai City Hospital | 341014 | 44362 |  | 04DS:TM:A dark band ring artifact in head area images>>Improved with AirCal>>There is a single black ring ar [Truncated] | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: A ring artifact occurred SAP Transfer: Dispatched for Onsite Support SAP Message: Components Added Succesfully to SWO!- SAP Message: LOGI Task added Successfully!- SAP Transfer: Case Transferred to SAP External Remarks:15.06.2021 09:36:29 UTC HIDEO KONISHI (81100219) Request: Ring artifact occurrence Problem confirmation: I confirmed it. Work details: I checked where ring artifacts could be found, such as the detector and bowtie filter, but there were no problems. Also, as the ring artifact had disappeared after the second air calibration, an abnormal calibration could have been created during the first air calibration. The second calibration could have corrected to the correct table. There are no problems now, so he can keep using the system as is. Other problems/Measures taken: None. Equipment status The equipment can be used. Work status Work finished. 14-Jun-2021 81100219 FSE Handover Tool1.3.4.32612.DB13 20210415 SUAT: System Up Again Time: June 15, 2021, 13:30" T2 Activities:Problem Description: Problem description by engineer : Equipment affected: CT Problem location: 04 Gantry Exclude HV When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user_s work flow did the problem occur?: (During Acquisition/Scan) What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem?: NA Engineer_s Comment: _============================== $version = "v4.1.10.25210" 9:04 Call description: Ring artifacts were found in head images. They can be clearly seen in the head images but there are no problems with the chest scan. Air calibration was performed in the morning _============================== Troubleshooting Action: What steps are taken during the system diagnosis: Result of response: Problem solved. Problem location: 04 Gantry Exclude HV What steps did you perform to diagnose the problem? _================================== Result I had him do air calibration again and he said that improved the situation. (However, after that, the scan stopped due to arcing.） Image confirmation In head images, there is a single black ring artifact near Module 14 or Module 29. It is not a complete circle, it is an arc artifact, and It is not that data is completely missing either. There were no abnormalities during the morning AirCal, but when I looked at the AirCal images after the problem occurred, there were clear rings on 2 of the 4 images. _================================== Error # and/or description of error: NA Repair Action: Completed or Recommended Repair Action "External": The service engineer for the site visited and worked on the equipment. OneEMS internal remarks : Problem description by engineer : Equipment affected: CT Problem location: 04 Gantry Exclude HV When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user_s work flow did the problem occur?: (During Acquisition/Scan) What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem?: NA Engineer_s Comment: _============================== $version = "v4.1.10.25210" 9:04 Call description: Ring artifacts were found in head images. They can be clearly seen in the head images but there are no problems with the chest scan. Air calibration was performed in the morning _============================== What steps are taken during the system diagnosis: Result of response: Problem solved. Problem location: 04 Gantry Exclude HV What steps did you perform to diagnose the problem? _================================== Result I had him do air calibration again and he said that improved the situation. (However, after that, the scan stopped due to arcing.） Image confirmation In head images, there is a single black ring artifact near Module 14 or Module 29. It is not a complete circle, it is an arc artifact, and It is not that data is completely missing either. There were no abnormalities during the morning AirCal, but when I looked at the AirCal images after the problem occurred, there were clear rings on 2 of the 4 images. _================================== Error # and/or description of error: NA Completed or Recommended Repair Action "External" : The service engineer for the site visited and worked on the equipment. Completed or Recommended Repair Action "Internal": --Troubleshooting Action-- Result of response: Problem solved. Problem location: 04 Gantry Exclude HV What steps did you perform to diagnose the problem? _================================== Result I had him do air calibration again and he said that improved the situation. (However, after that, the scan stopped due to arcing.） Image confirmation In head images, there is a single black ring artifact near Module 14 or Module 29. It is not a complete circle, it is an arc artifact, and It is not that data is completely missing either. There were no abnormalities during the morning AirCal, but when I looked at the AirCal images after the problem occurred, there were clear rings on 2 of the 4 images. _================================== --Repair Action-- Recommended 12Nc: PR1 459800795292:KIT - BR ROHS COLLIMATOR FIELD SERVICE QTY:1 SLOC: Sapporo (U612) PR1 459800911271:TDMS V2 M.Det Class A FRU QTY:1 SLOC: Sapporo (U612) --Subject-- 04DS:TM:A dark band ring artifact in head area images>>Improved with AirCal>>There is a single black ring artifact near Module 14 or Module 29 --Problem Description-- Equipment affected: CT Problem location: 04 Gantry Exclude HV When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user_s work flow did the problem occur?: (During Acquisition/Scan) What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem?: NA Engineer_s Comment: _============================== $version = "v4.1.10.25210" 9:04 Call description: Ring artifacts were found in head images. They can be clearly seen in the head images but there are no problems with the chest scan. Air calibration was performed in the morning _============================== Safety Questions:Case: 0117699673 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: TAKEYUKI MIYAKAWA Last updated by: TAKEYUKI MIYAKAWA 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Internal Remarks : 15.06.2021 09:36:33 UTC HIDEO KONISHI (81100219) Replacement parts: No replacement parts. [Report on mistakes in parts order] No mistakes. Situation monitoring information: Not monitoring the situation TS Request: No parts to return. Additional CT information: Q1. How did the examination finish? As it stopped before X-ray emission, no extra irradiation Q2. Were there any mistaken findings (misdiagnosis, mistaken treatment) due to the problem? None. Q3. Were there any artifacts? None. Q4. Was there any unintended unnecessary exposure? None. Repair sticker: I affixed a Repaired sticker. [Other information to share] 1. Was treatment/quarantine carried out for a patient who has or is suspected of having new coronavirus? Yes 2. Is it necessary to go into treatment rooms/beds where there are infected patients?(Including workplace and routes of movement) No Did you agree on how to complete the report with the person in charge of devices? Yes. I completed normal reporting that day. 14-Jun-2021 81100219 FSE Handover Tool1.3.4.32612.DB13 20210415 ----------------------------------------------------------------------------- | 2 |  | Air calibration was performed | [12 empty] | 117699673 |  | 1008890 | TRUE | Investigation | Investigation |  | NotEngineering | [5 empty] | Engineering | [6 empty] | Artifact | [5 empty] | Refresh |  |  | System issue | System | NotEngineering | ShouldbeCustomer | NotEngineering-ShouldbeCustomer | NotEngineering-ShouldbeCustomer | NotEngineering |  | NotEngineering | NotEngineering-ShouldbeCustomer | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | y
1062 |  |  | Japan | Social Medical Corporation BOKOI, Nikko Memorial Hospital | 341011 | 44358 |  | JY: SR FCO implementation | BU Q&R Feedback: No Reason: Lifecycle Solution Problem Reported by Customer: JY: SR FCO implementation Safety Questions:Case: 0117686447 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: YAN Joanna Last updated by: YAN Joanna 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A OneEMS internal remarks: Hashimoto DB FCO72800737 SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Support External Remarks:15.06.2021 14:25:56 UTC HODAKA HASHIMOTO (81100553)[Details of request] FCO72800737 implemented Problem confirmation: I confirmed it. Work details: I performed FCO72800737. KEYBOARD AND RECON SW UPDATE FOR INGENUITY, BIG BORE AND BR64 Also, I implemented it on February 11th. When I replaced the X-ray tube, I reported that the gantry rpm was about 360,000 rotations. Other problems/Measures taken: None. Equipment status The equipment can be used. Work status Work finished. 15-Jun-2021 81100553 FSE Handover Tool1.3.4.32612.DB13 20210415 Internal Remarks:15.06.2021 14:25:59 UTC HODAKA HASHIMOTO (81100553) Replacement parts: No replacement parts. [Report on mistakes in parts order] No mistakes. Situation monitoring information: Not monitoring the situation TS Request: No parts to return. Additional CT information: Q1. How did the examination finish? Q1. No (not applicable) Q2. Were there any mistaken findings (misdiagnosis, mistaken treatment) due to the problem? None. Q3. Were there any artifacts? None. Q4. Was there any unintended unnecessary exposure? None. Repair sticker: I affixed a Repaired sticker. [Other information to share] 1. Was treatment/quarantine carried out for a patient who has or is suspected of having new coronavirus? No 2. Is it necessary to go into treatment rooms/beds where there are infected patients?(Including workplace and routes of movement) No Did you agree on how to complete the report with the person in charge of devices? Yes. I completed normal reporting that day. 15-Jun-2021 81100553 FSE Handover Tool1.3.4.32612.DB13 20210415 ----------------------------------------------------------------------------- | 5 |  | FCO72800737 | [12 empty] | 117686447 |  | 1006868 |  |  | Closed - Non-Complaint | [7 empty] | Others | [6 empty] | FCO72800737 | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1063 |  |  | United States | Philips Medical Capital | 40005 | 44365 |  | 0402CFXXX no power to gentry | Work Order #: WO-03678282 Ali, Adam The system fully meets specification and returned to use Problem Description by Engineer : Problem description by engineer : No power to system When was the first occurrence? Within... : 18-Jun-2021 How often is issue occurring? : Continuously Malfunction area : Others Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: Not provided Patient Impact/Outcome: No further information is available as per the customer User Impact: System is out of use Current Software Version: Not provided Resolution : Angela walked in to find the power to the gantry & computers was out. She has tried resetting the main breakers. The UPS was beeping earlier but has since stopped. Found the green/blue light on the TEAL is off. Angela will get in touch with the facilities so they can check incoming power from the hospital. I called back later to find they had a flicker in power due to the generator that caused hospital breaker to trip. System is up and running now. Internal Comments : Error # and/or description of error : No power to system ----------------------------------------------------------------------------- | 2 |  | Found the green/blue light on the TEAL is off. | [12 empty] | 117721290 |  | 1020309 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Power | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1064 |  |  | Poland | Rzadowa Agencja | 341008 | 44364 |  | Security of the tomograph during transport | Work Order #: WO-03672267 Paweł Kierzek Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- Work Order #: WO-03672272 Mateusz Ramadan Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 4 |  | Security of the tomograph during transport Launch of a mobile tomograph in a new location - Katowice Głowackiego 10. Preparation from the service side of the tomograph for transport - container relocation. Work completed successfully. | [12 empty] | 117715858 |  | 1020053 | TRUE | Evaluation | Evaluation | [7 empty] | Others | [6 empty] | Cabin CT transport | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1065 |  |  | Turkey | Republic of Turkey Ministry of Health, Bagcilar | 40010 | 44365 |  | spill guard order | A new one was ordered because the previous spill guard was torn immediately. |  |  | A new one was ordered because the previous spill guard was torn immediately. | [12 empty] | 117722467 |  | 1024981 | TRUE | Evaluation | Closed - non-complaint | [7 empty] | Others | [6 empty] | spill guard | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1066 |  |  | United States | Philips Medical Capital | 40055 | 44366 |  | cant log into system its locked the user out | Work Order #: WO-03710811 Boyd Craig The system does not meet specification and is out of use Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: log into system its locked the user out User Impact: None Patient Impact: None Current Software Version: Diagnostic performed by Engineer : User has tried to log in with wrong passwords and is locked out. Customer will call back with PO for service if needed. ----------------------------------------------------------------------------- |  |  | Customer was able to get the emergency IT log on password and log into system. Customer said to close the case. | [12 empty] | 117753160 |  | 1026299 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Password | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1067 |  |  | United States | Philips Medical Capital | 40028 | 44363 |  | How to make sag and coronals | Work Order #: WO-03668359 Hirchak, Michele The system fully meets specification and returned to use Problem description by Clinical Support Specialist : Customer Function/Role: Tech Device Use: In Use Expected and Unexpected Results: Need sagittals and coronals. User Impact: None Patient Impact: None Current Software Version: 4.1.7 Diagnostic performed by Engineer : Customer said she did not get her sag and coronal results at the time of scan. Walked her through recon and they were still in the list, but were not complete. Had her hit start recon, and they reconned without issue. She is all set. MBH ----------------------------------------------------------------------------- | 5 |  | Q&A | [12 empty] | 117712156 |  | 1017705 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Application | [6 empty] | Q&A | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1068 |  |  | United States | Philips Medical Capital | 40055 | 44363 |  | 0703CDN84 Screen goes blank after a exam | Work Order #: WO-03663569 Ali, Adam The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Problem description by engineer : Left screen goes blank after an exam When was the first occurrence? Within... : 16-Jun-2021 How often is issue occurring? : Multiple times per week Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Not provided Patient Impact/Outcome: None User Impact: Cannot restart and have to go a hard shutdown the host pc Current Software Version: 4.1 Diagnostic performed by Engineer : Left screen goes blank after an examination, Leah cannot restart the software due to no visibility to the left screen. This has happened a couple of times now since last week. Unable to download hostlogs due an error _Keyboard-interactive authentication with the SSH2 server failed._ Internal Comments : Error # and/or description of error : Left screen goes blank after an examination Completed or Recommended Repair Action "Internal" : Check host eventlogs for further investigation Inspect/ reseat graphics card, cables. Recommended parts: NVIDIA Quadro P400 2GB (3)mDP GFX 459801536801 CT Gen5.3 Host, Z4 459801493151 Follow Up Required : What are the next Troubleshooting steps : Check host eventlogs for further investigation Inspect/ reseat graphics card, cables. Recommended parts: NVIDIA Quadro P400 2GB (3)mDP GFX 459801536801 CT Gen5.3 Host, Z4 459801493151 Error # and/or description of error : Left screen goes blank after an examination Completed or Recommended Repair Action "Internal" : Check host eventlogs for further investigation Inspect/ reseat graphics card, cables. Recommended parts: NVIDIA Quadro P400 2GB (3)mDP GFX 459801536801 CT Gen5.3 Host, Z4 459801493151 ----------------------------------------------------------------------------- Work Order #: WO-03664336 Donald Vic Problem Description by Engineer : Problem description by engineer : Left screen goes blank after an exam When was the first occurrence? Within... : 16-Jun-2021 How often is issue occurring? : Multiple times per week Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Not provided Patient Impact/Outcome: None User Impact: Cannot restart and have to go a hard shutdown the host pc Current Software Version: 4.1 Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- |  |  | Left screen goes blank after an examination, Leah cannot restart the software due to no visibility to the left screen. This has happened a couple of times now since last week. Unable to download hostlogs due an error _Keyboard-interactive authentication with the SSH2 server failed._ | [12 empty] | 117707968 |  | 1010412 | TRUE | Evaluation | Investigation |  | TBD-Adapter | [5 empty] | Engineering | [6 empty] | Blank screen | [5 empty] | Refresh |  |  | System issue | System | TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-Adapter |  | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | y
1069 |  |  | United States | Philips Medical Capital | 40028 | 44361 |  | 9599CDX- Work List filtering is not setup properly - brand new scanner | Work Order #: WO-03654361 Dangos John The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: Diagnostics Expected and Unexpected Results: Customer having His/Ris lissues where they are unable to use filter option to customize their viewing. User Impact: No Patient Impact: No Current Software Version:4.1 Diagnostic performed by Engineer : Customer having His/Ris lissues where they are unable to use filter option to customize their viewing. Follow Up Required : FSE to communicate with Hospital PACS IT team regarding proper setup of his/ris and match with Scanner HSI/RIS settings thru the LANCONFIG app. ----------------------------------------------------------------------------- |  |  | Customer having His/Ris lissues where they are unable to use filter option to customize their viewing. Work with BU and Carestream to resolve the query. Verify WL Query. | [12 empty] | 117698805 |  | 1003539 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | HIS/RIS | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1070 |  |  | United States | Willis Knighton Hospital | 40004 | 44361 |  | 0308CDN81 Table making noise when moved vertically | Work Order #: WO-03652082 Mcginnis, Mike The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Problem description by engineer : Table making noise when moved vertically When was the first occurrence? Within... : 14-Jun-2021 How often is issue occurring? : First occurrence Malfunction area : Table Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Table making noise when moved vertically Patient Impact/Outcome: none User Impact: hears noise from couch Current Software Version: Not provided Diagnostic performed by Engineer : Spoke to Jennifer. She said the couch is making a terrible noise when they raise or lower it. She said you can hear it from the other scanner. Internal Comments : Error # and/or description of error : Table making noise when moved vertically Completed or Recommended Repair Action "Internal" : Insure the couch is lubricated per the PM manual, check for any cover binding. Check the vertical brake, motor, ball screw and bearings. If defective components found replace parts as necessary Follow Up Required : Error # and/or description of error : Table making noise when moved vertically Completed or Recommended Repair Action "Internal" : Insure the couch is lubricated per the PM manual, check for any cover binding. Check the vertical brake, motor, ball screw and bearings. If defective components found replace parts as necessary ----------------------------------------------------------------------------- Work Order #: WO-03652118 15-Jun-2021 Williams Spence The system fully meets specification and returned to use Problem Description by Engineer : Problem description by engineer : Table making noise when moved vertically When was the first occurrence? Within... : 14-Jun-2021 How often is issue occurring? : First occurrence Malfunction area : Table Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Table making noise when moved vertically Patient Impact/Outcome: none User Impact: hears noise from couch Current Software Version: Not provided Resolution : Lubricated lower couch slides ----------------------------------------------------------------------------- | Dantha, Rajashekar |  | Lubricated lower couch slides | [12 empty] | 117697033 |  | 1002464 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Lubricated | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1071 |  |  | China | Laiyang City People_s Hospital | 345005 | 44369 |  | Wechat:02N/Ingenuity Core128/Transfer Image | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Configuration Error Repair Activity Code: Instructed Customer Problem Reported by Customer : Transfer Image Safety Questions:Case: 0117763367 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Zhou Tingting Last updated by: Zhou Tingting 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: setting images auto save device. Troubleshooting Action: setting images auto save device. Repair Action: instruct customer to set the auto save device, system is ok. Test/Inspection: Functional Check PASSED. The device returned to full functionality. Test Equipment: N/A SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | Troubleshooting Action: setting images auto save device. Repair Action: instruct customer to set the auto save device, system is ok. | [12 empty] | 117763367 |  | 1036037 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Application | [6 empty] | Q&A | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1072 |  |  | Japan | Maki Hospital | 341017 | 44372 |  | SY: CT application support package. | BU Q&R Feedback: No Reason: Application Support Problem Reported by Customer: SY: CT application support package. Safety Questions:Case: 0117781946 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: YE Smile Last updated by: YE Smile 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A OneEMS internal remarks : Otani DB SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Support ----------------------------------------------------------------------------- | 5 |  | Implementation of study sessions [Details of work] --We held a study session on parameters unique to Philips. | [12 empty] | 117781946 |  | 1048773 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Application | [6 empty] | application | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1073 |  |  | Japan | Maki Hospital | 341017 | 44369 |  | SY: Replace patient restraint strap. | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : SY: Replace patient restraint strap. Safety Questions:Case: 0117762965 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Smile YE Last updated by: Smile YE 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A OneEMS internal remarks: Kawazoe DB SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Support ----------------------------------------------------------------------------- | 5 |  | Fixed band replacement | [12 empty] | 117762965 |  | 1035757 | TRUE | Evaluation | Closed - non-complaint | [7 empty] | Service | [6 empty] | patient restraint strap | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1074 |  |  | Japan | Independent administrative institution National Hospital Organization | 341015 | 44370 |  | 01CS:KKI:The CT-Host froze. >> Restarting the CT-Host restored everything. > There was no hardware error. | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 01CS:KKI:The CT-Host froze. >> Restarting the CT-Host restored everything. > There was no hardware error. Safety Questions:Case: 0117768352 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: KAZUMASA KIMURA Last updated by: KAZUMASA KIMURA 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 4 |  | We have not confirmed any hardware errors. Probably the performance has dropped because the reboot hasn_t been done for about a week. We are telling you that it can be alleviated by restarting the CT host every day. No. The customer was unable to create a log file or report. | [12 empty] | 117768352 |  | 1039733 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Reboot | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1075 |  |  | United States | Leesburg Regional Medical Center | 40022 | 44373 |  | 0711CFXXX Locked out of system due to entering the incorrect password | Work Order #: WO-03747860 Miller, John M The system does not meet specification and is out of use Diagnostic performed by Engineer : Collect information from customer. Problem Description by Engineer : Problem description by engineer : Customer states they are locked out of the system due to password incorrect input. When was the first occurrence? Within... : 26-Jun-2021 How often is issue occurring? : First occurrence Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: can not Scan Patient Impact/Outcome: Not provided User Impact: can not Scan Current Software Version: 4.1.10 Internal Comments : Error # and/or description of error : Wrong password entered Completed or Recommended Repair Action "Internal" : Local service will need to reset the password. Customer did not want case sent to local service over the weekend. Suggest customer contact IT department, as they should have a password to reset the CT login. Resolution : Error # and/or description of error : Wrong password entered Completed or Recommended Repair Action "Internal" : Local service will need to reset the password. Customer did not want case sent to local service over the weekend. Suggest customer contact IT department, as they should have a password to reset the CT login. ----------------------------------------------------------------------------- | 2 |  | Local service will need to reset the password. Customer did not want case sent to local service over the weekend. Suggest customer contact IT department, as they should have a password to reset the CT login. Not a defect, customer can reset password by themselves according IFU. | [12 empty] | 117788182 |  | 1049644 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Customer | [6 empty] | Password | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1076 |  |  | China | 山东省胸科医院 | 40025 | 44374 |  | 02N/Ingenuity CT/cannot login | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Calibration/Adjustment Problem Reported by Customer : 02N/Ingenuity CT/ Safety Questions : Case: 0117788721 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Cao Last updated by: Liu Cao 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 无法登录 客户问题详细描述或故障现象： 卡死在Philips桌面 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： FSE 反馈设备为CT5000，异常关机后有时会出现，需要密码狗登录恢复 重装无效 远程诊断与测试结果： 到场恢复系统，需要时收集信息提交escalation PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 到场恢复系统，需要时收集信息提交escalation Test Equipment: N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : Site1 FSE 黄兆坤 18654536122 一直在通话中 External Remarks : 27.06.2021 03:44:26 UTC Zhaokun Huang (86038405)"SUAT. Ready for Clinical Use local time" : "2021-06-27 12:30" External Remarks : 27.06.2021 03:50:39 UTC SUP JCOUSER (86038405)服务内容：用IST权限登录软件后，重新启动Host。 服务结果：维修后，设备恢复正常。 ----------------------------------------------------------------------------- | 2 |  | 服务内容：用IST权限登录软件后，重新启动Host。 服务结果：维修后，设备恢复正常。 收集信息提交escalation Collect issue detail and bugreport, issue happen before and after HF3 https://www.clouddrive.philips.com.cn/file/8uzf62w7mue2ua8ihcovq1cmq23butql# 密码:123 | [12 empty] | 117788721 |  | 1051124 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632553 | [5 empty] | Engineering | [6 empty] | Blue screen | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Blue-Screen | Console-SW-Blue-Screen | Console-SW-Blue-Screen | CTUDT00632553 | Blue-Screen | Console-SW | Console-SW-Blue-Screen | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | y
1077 |  |  | Japan | Independent administrative institution National Hospital Organization | 341015 | 44368 |  | CT:APP:NO:Mr. Hisaki: Question. | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Application Support Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer: Q: (1) When he was editing the ECG, he tried to reconstruct it by shifting the phase for only one heartbeat, but the scan did not run. (2) What is the CT Viewer VIP? A: (1) It is likely that reconstruction was not possible due to data loss. (2) It stands for Volume Intensity Projection, and the maximum value is allocated to the voxel that is closest to the eye point. Safety Questions:Case: 0117756315 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Ootani Natsuki Last updated by: Ootani Natsuki 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 5 |  | Q: (1) When he was editing the ECG, he tried to reconstruct it by shifting the phase for only one heartbeat, but the scan did not run. (2) What is the CT Viewer VIP? A: (1) It is likely that reconstruction was not possible due to data loss. (2) It stands for Volume Intensity Projection, and the maximum value is allocated to the voxel that is closest to the eye point. | [12 empty] | 117756315 |  | 1031070 | TRUE | Investigation | Investigation | [7 empty] | Application | [6 empty] | Q&A | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1078 |  |  | Japan | 苫小牧市立病院 | 341014 | 44371 |  | 01DS: KK: Treatment plan examination shows an artifact from the shoulder to the thyroid gland. >> It seems that even if you check the examination of the past treatment plan, it is out. >> Surveyed on FSE onsite. | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 01DS:KK:治療計画の検査で肩から甲状腺に向かってアーチファクトが発生する。>>過去の治療計画の検査を確認していても出ているそうです。>>FSEオンサイトにて調査。 ----------------------------------------------------------------------------- | 4 |  | It was said that an artifact would appear when shooting the treatment plan. Even if you go back in time, it is said that similar artifacts have occurred in the shooting of the treatment plan. It was a field survey by FSE on-site. No. The customer was unable to create a log file or report. | [12 empty] | 117776693 |  | 1045269 | TRUE | Investigation | Investigation |  | TBD_Bugreport | [5 empty] | Engineering | [6 empty] | Artifact | [5 empty] | Refresh |  |  | System issue | System | TBD-Bugreport | TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | TBD_Bugreport |  | TBD-Bugreport | TBD-Bugreport-TBD_Bugreport | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | y
1079 |  |  | Japan | Tomakomai City Hospital | 341014 | 44369 |  | 03CM: RS: An error occurred during the first scan of the morning >> It was possible to continue with the sca [Truncated] | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 03CM: RS: An error occurred during the first scan of the morning >> It was possible to continue with the scan >> S_MASTER_SEQUENCER_NO_XRAY/S_GEN_FAULT_ANODE_TUBE_ARC Safety Questions:Case: 0117761490 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: RUI SATAKE Last updated by: RUI SATAKE 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 4 |  | You can confirm the occurrence of S_GEN_FAULT_ANODE_TUBE_ARC. It seems that the tube was replaced last Sunday, so I think there is a high possibility of initial discharge. The inspection will continue as it is, and we ask for your understanding by follow-up. | [12 empty] | 117761490 |  | 1034816 | TRUE | Investigation - Pending Translation | Closed - Done |  | CTUDT00630085 | [5 empty] | Engineering | [6 empty] | Tube Arc | [5 empty] | Refresh |  |  | System issue | System | IC-HW | Tub Arc | IC-HW-Tub Arc | IC-HW-Tub Arc | CTUDT00630085 | Tub Arc | IC-HW | IC-HW-Tub Arc | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | y
1080 |  |  | Japan | Independent administrative institution National Hospital Organization | 341013 | 44369 |  | 01CS:RS: The scan screen was closed during a scan. >> He recreated the plan and continued with the scan. >> [Truncated] | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 01CS:RS: The scan screen was closed during a scan. >> He recreated the plan and continued with the scan. >> Safety Questions:Case: 0117762419 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: RUI SATAKE Last updated by: RUI SATAKE 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 4 |  | The scan reads the surface again and continues the inspection. No noticeable error was found in the logs. Judged as temporary application performance degradation. follow-up. | [12 empty] | 117762419 |  | 1035326 | TRUE | Investigation - Pending Translation | Investigation |  | TBD_Bugreport | [5 empty] | Engineering | [6 empty] | scan screen | [5 empty] | Refresh |  |  | System issue | System | TBD-Bugreport | TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | TBD_Bugreport |  | TBD-Bugreport | TBD-Bugreport-TBD_Bugreport | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | y
1081 |  |  | Japan | Social Medical Corporation BOKOI, Nikko Memorial Hospital | 341011 | 44369 |  | Via FSE.Contrast agents disappeared from the list.Contrast agents were listed in the Referring Physician sec [Truncated] | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: Via FSE.Contrast agents disappeared from the list.Contrast agents were listed in the Referring Physician section, but only OMNI300 was listed this morning. Safety Questions:Case: 0117762375 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: TAKEYUKI MIYAKAWA Last updated by: TAKEYUKI MIYAKAWA 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 5 |  | As a result, it seems likely that it disappeared due to the operation | [12 empty] | 117762375 |  | 1035292 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | NotEngineering | [5 empty] | Engineering | [6 empty] | contrast agent | [5 empty] | Refresh |  |  | System issue | System | NotEngineering | ShouldbeService | NotEngineering-ShouldbeService | NotEngineering-ShouldbeService | NotEngineering |  | NotEngineering | NotEngineering-ShouldbeService | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | y
1082 |  |  | China | Jiangsu Xuzhou Medical School Affiliated No. 3 Hospital | 40039 | 44368 |  | 02G/Ingenuity CT/blue screen | BU Q&R Feedback: No SAP Customer Symptom Code: BOOT Customer Symptom Code: Boot up/Power-up Problem Reason: Corrective Maintenance Engineer Failure Code: Electrical/Electronic Failure Repair Activity Code: Reboot/Restart Next Step Action: NTD: This device is not yet under warranty. Please do not add Onsite directly. If we need to dispatch labor, please inform C1 Problem Reported by Customer: 02G/Ingenuity CT/ Next Step Action: Machine install order is 52828478. If need to dispatch, must inform IB to add OP to installation order for FSE Safety Questions:Case: 0117756731 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Weng Jinming Last updated by: Weng Jinming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Troubleshooting Action: instruct customer to restart the system Repair Action: instruct customer to restart the system Test/Inspection: Functional Check PASSED. The device returned to full functionality. Test Equipment: N/A ----------------------------------------------------------------------------- | 3 |  | instruct customer to restart the system | [12 empty] | 117756731 |  | 1031334 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD_Bugreport | [5 empty] | Engineering | [6 empty] | Reboot | [5 empty] | Refresh |  |  | System issue | System | NotEngineering | TBD | NotEngineering-TBD | NotEngineering-TBD | TBD_Bugreport |  | NotEngineering | NotEngineering-TBD | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | y
1083 |  |  | China | 济南市第四人民医院 | 40038 | 44371 |  | 02N/Ingenuity CT 728326/image transfer performance slow | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Calibration/Adjustment Problem Reported by Customer : 02N/Ingenuity CT 728326 Safety Questions : Case: 0117775438 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Gaoge Last updated by: Wang Gaoge 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 图像传输失败。 客户问题详细描述或故障现象： 主机传图慢 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 主机传图慢 远程诊断与测试结果： 重置传输缓存，检查网络环境，如果还是慢，考虑软件defective PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 1，教会客户重置传输缓存 2 检查医院网络是否千兆节点是否正常 Test/Inspection: N/A Test Equipment: N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 25.06.2021 00:24:14 UTC Tao Chen (86038373)故障现象: 软件运行慢，传图慢 服务内容: 重置spooler，清理垃圾文件 服务结果: 维修后,设备恢复正常. Internal Remarks : 25.06.2021 01:05:43 UTC Tao Chen (86038373) “RDI Accuracy: YES" $Q7: (Confirmed reported problem, software work slow) $Q8: ( Reboot CT host, found it was still slow) $Q9: ( Reset spooler and clean useless file for CT host) $Q10: ( Test and inspection data refer to: PA. Device returned to full functionality) $Q11: Calibrated Tools name ( N/A) ^Current Tube installed date: N/A ^Current Tube usage: N/A ----------------------------------------------------------------------------- | 3 |  | $Q7: (Confirmed reported problem, software work slow) $Q8: ( Reboot CT host, found it was still slow) $Q9: ( Reset spooler and clean useless file for CT host) | [12 empty] | 117775438 |  | 1044463 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-SW-Slow | [5 empty] | Engineering | [6 empty] | SW slow | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Performance | Console-SW-Performance | Console-SW-Performance | TBD-SW-Slow | Performance | Console-SW | Console-SW-Performance | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | y
1084 |  |  | United States | Odessa Regional Medical Center | 40029 | 44372 |  | Temperture Control Warning | Work Order #: WO-03746425 McCain, Jimmie Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 5 |  | Parts Consumed Description: 1.00000,459800550141,BR64T RoHS HEATER REPLACEMENT KIT 1.00000,459800550141,BR64T RoHS HEATER REPLACEMENT KIT 1.00000,459800980091,BR64 Air Fan Replacement Kit During PM | [12 empty] | 117786463 |  | 1050377 | TRUE | Evaluation | Evaluation | [7 empty] | Others | [6 empty] | Heater | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1085 |  |  | United States | Philips Medical Capital | 40028 | 44371 |  | 9512CFFB1 4.1 Export Dose Logs | Work Order #: WO-03737120 Giles Brian The system fully meets specification and returned to use Problem description by Clinical Support Specialist : Customer Function/Role: Therapist Device Use: In Use Expected and Unexpected Results: Dose Log archive User Impact: None Patient Impact: None Current Software Version: 4.1 Diagnostic performed by Engineer : Customer states that the CT is not reading the USB. I suggested that she try to insert it again and look for the blinking light with the folder Browser stated the USB is in. She will check the PC she downloads it on. If she doesn_t call back she ask for me to close it out ----------------------------------------------------------------------------- | 5 |  | Customer states that the CT is not reading the USB. I suggested that she try to insert it again and look for the blinking light with the folder Browser stated the USB is in. She will check the PC she downloads it on. If she doesn_t call back she ask for me to close it out | [12 empty] | 117778050 |  | 1041813 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | USB | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1086 |  |  | United States | Lackey Memorial Hospital | 40015 | 44369 |  | down button on control box is not working | Work Order #: WO-03723108 22-Jun-2021 Denison Chuck The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Customer Function/Role: C.T. tech Device Use: Philips C.T. Expected and Unexpected Results: System is up. User Impact: None Patient Impact: None Current Software Version:4.1.10 Diagnostic performed by Engineer : Need to replace the C.T. control box. Down button is not working. Follow Up Required : Need to order parts and install next day. C.T. Box at console. ----------------------------------------------------------------------------- Work Order #: WO-03725579 22-Jun-2021 Denison Chuck Problem Description by Engineer : Customer Function/Role: C.T. tech Device Use: Philips C.T. Expected and Unexpected Results: System is up. User Impact: None Patient Impact: None Current Software Version:4.1.10 Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 5 |  | Need to replace the C.T. control box. Down button is not working. | [12 empty] | 117765203 |  | 1034306 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632185 | [5 empty] | Engineering | [6 empty] | CT Box | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | CT Box | Gantry-HW-CT Box | Gantry-HW-CT Box | CTUDT00632185 | CT Box | Gantry-HW | Gantry-HW-CT Box | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | y
1087 |  |  | Japan | 独立行政法人国立病院機構 | 341015 | 44368 |  | 07CS:KK:中国語のAuto Voiceはありますか？>>Mandarin chineseがあると回答。 | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Application Support Engineer Failure Code: Unknown Cause Repair Activity Code: Instructed Customer Problem Reported by Customer : 07CS:KK:中国語のAuto Voiceはありますか？>>Mandarin chineseがあると回答。 T2 Activities : Problem Description: 対象装置:CT 故障個所:07その他 最初に発生したのはいつか?(~以内):Unknow 問題発生の頻度は?:Unknow エンドユーザｰのワークフローのどの時点において問題が発生しているのか?:その他(Other) 最近どのようなイベントが発生したか?:N.A. 顧客は問題について具体的なシステムイベントレポートを作成しているか?:N.A. Troubleshooting Action: 対応後の結果:その他 その他:お問合せ 故障個所:07その他 診断中にどのようなステップを実施したか?: Mandarin Chineseがあります。と回答。 Repair Action: 【現象確認】 問合せの為、N.A. 【作業内容詳細】 Mandarin Chineseがあります。と回答。 【確認内容】 良好 OneEMS internal remarks : --Troubleshooting Action-- 対応後の結果:その他 その他:お問合せ 故障個所:07その他 診断中にどのようなステップを実施したか?: Mandarin Chineseがあります。と回答。 --Repair Action-- --Subject-- 07CS:KK:ﾋｻｷ様:中国語のAuto Voiceはありますか？>>Mandarin chineseがあると回答。 --Problem Description-- 対象装置:CT 故障個所:07その他 最初に発生したのはいつか?(~以内):Unknow 問題発生の頻度は?:Unknow エンドユーザｰのワークフローのどの時点において問題が発生しているのか?:その他(Other) 最近どのようなイベントが発生したか?:N.A. 顧客は問題について具体的なシステムイベントレポートを作成しているか?:N.A. --4/8Elements-- Q1 報告された問題は確認できましたか? 問合せの為、N.A. Q2 他の問題は確認できましたか? 無し Q3 問題を解決するために何をおこないましたか? Mandarin Chineseがあります。と回答。 Q4 対応後の装置は順調ですか? 良好 Q5 その検査はどの様に終了しましたか? 無し Q6 問題により間違った所見（誤診、誤治療）の有無: 無し Q7 意図しない不要な曝射の有無: 無し Q8 アーチファクトの有無: 無し Safety Questions : Case: 0117756833 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: KENICHI KOJIMA Last updated by: KENICHI KOJIMA 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | Zhu, Chun |  | 44369.7792 | [12 empty] | 117756833 |  | 1031418 |  |  | Closed - Non-Complaint | [14 empty] | 44370.5625 | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1088 |  |  | United States | Mount Grant General Hospital | 40012 | 44370 |  | Pitch factor questions | Work Order #: WO-03732938 Hirchak, Michele The system fully meets specification and returned to use Problem description by Clinical Support Specialist : Customer Function/Role: Physicist Device Use: In Use Expected and Unexpected Results: CTDI measurement for ACR different than what the protocol shows. User Impact: None Patient Impact: None Current Software Version: 4.1.7 Diagnostic performed by Engineer : Customer sent some screen shots with the helical and the axial converted protocol. Customer is missing the mA station, and the scan angle for the axial is 420 and should be 360. Also the collimation is different between the 2. I emailed him the screen shots and told him where to fix that. I also sent that to James Hubenschmidt in CLE so he can reach out to the customer and further assist. Called customer and told him I would close this w/o due to 7 day time limit. Will follow up next week. MBH ----------------------------------------------------------------------------- | 5 |  | Pitch factor questions | [12 empty] | 117773844 |  | 1043465 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Application | [6 empty] | Q&A | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1089 |  |  | United States | Philips Medical Capital | 40005 | 44368 |  | System estop open due to setscrew loose on tilt pot2; Made necessary adjustments, tightened setscrew, ran calibration. | Work Order #: WO-03715902 Mcginnis, Mike The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Problem description by engineer : ESTOP opened When was the first occurrence? Within... : 21-Jun-2021 How often is issue occurring? : First occurrence Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: ESTOP opened Patient Impact/Outcome: none User Impact: had to reset system Current Software Version: Not provided Follow Up Required : Ship TO: The Regional Medical Center 3000 Saint Matthews Rd ORANGEBURG SC 29118-1442 ATT: CT Dept / Philips Error # and/or description of error : ESTOP opened 2021/06/21 09:37:32.705 S_TILT_C_POT1_POT2_KO Completed or Recommended Repair Action "Internal" : Check all connections, replace dual pot assy and calibrate tilt. DUAL POT ASSEMBLY 459800441121 Internal Comments : Error # and/or description of error : ESTOP opened 2021/06/21 09:37:32.705 S_TILT_C_POT1_POT2_KO Completed or Recommended Repair Action "Internal" : Check all connections, replace dual pot assy and calibrate tilt. DUAL POT ASSEMBLY 459800441121 Diagnostic performed by Engineer : Diane said they just completed a scan, Nurse was in the room removing connections to the patient. When she hit the end study ESTOP opened. She power cycled the system, brought it up and ran a test scan with no issues. I downloaded log files and found the tilt error below. She said the Nurse was at the end furthest away from the gantry at the time of the issue. ----------------------------------------------------------------------------- Work Order #: WO-03716017 21-Jun-2021 Beasley Donald The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Problem description by engineer : ESTOP opened When was the first occurrence? Within... : 21-Jun-2021 How often is issue occurring? : First occurrence Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: ESTOP opened Patient Impact/Outcome: none User Impact: had to reset system Current Software Version: Not provided Diagnostic performed by Engineer : checked error logs Follow Up Required : install and calibrate the tilt pot assy Parts Consumed : 1.00000,459800441121,DUAL POT ASSEMBLY ----------------------------------------------------------------------------- Work Order #: WO-03722050 23-Jun-2021 Beasley Donald The system fully meets specification and returned to use Problem Description by Engineer : Problem description by engineer : ESTOP opened When was the first occurrence? Within... : 21-Jun-2021 How often is issue occurring? : First occurrence Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: ESTOP opened Patient Impact/Outcome: none User Impact: had to reset system Current Software Version: Not provided Resolution : found that the setscrew was loose on the tilt pot #2 ----------------------------------------------------------------------------- Work Order #: WO-03725845 23-Jun-2021 Stanczyk, Mark The system fully meets specification and returned to use Problem Description by Engineer : E-stop opens during tilt Resolution : Assisted DB in repair. Removed gantry covers and saw a slight in 90 angle. Jog_d tilt FORWARD ever so slightly and started cal_s. Pot1 was close but could use a nudge. Pot2 however we found the set screw to be loose. made necessary adjustments,tightened setscrew, finished cal. easy peasy ----------------------------------------------------------------------------- | 3 |  | Removed gantry covers and saw a slight in 90 angle. Jog_d tilt FORWARD ever so slightly and started cal_s. Pot1 was close but could use a nudge. Pot2 however we found the set screw to be loose. made necessary adjustments,tightened setscrew, finished cal. easy peasy | [12 empty] | 117758400 |  | 1028247 | TRUE | Closed - Done | Closed - Done | [7 empty] | Others | [6 empty] | Dual pot | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1090 |  |  | Germany | Gesundheit Nord gGmbH Klinikverbund | 40002 | 44369 |  | Gantry briefly turned off and on again twice during operation | Work Order #: WO-03722535 Eduard Berthold The system partly meets specification and is returned to use with restrictions Problem Description by Engineer: Call-back to the customer, Raik Glogner-Umbeer Gantry turned off 2 times Diagnostic performed by engineer: The following actions were performed for error analysis / error correction: Error log diagnosis: S_VMC_AXIS_UNALLOWED_MOVE 12:35, 11:38 Experience: Replacing the Vertical Transducer (String Pot) resolved the error 459800623331 Vertical transducer Follow-up required: CT still in operation Check and clean brake Replacement as first thing: 459800623331 Vertical transducer 2 hours labor The customer wants info tomorrow morning, but I expect a system failure tomorrow! ----------------------------------------------------------------------------- | 3 |  | Replacing the Vertical Transducer (String Pot) resolved the error 459800623331 Vertical transducer | [12 empty] | 117764743 |  | 1032671 | TRUE | Investigation | Investigation |  | NotEngineering-partreplace | [5 empty] | Engineering | [6 empty] | Vertical transducer | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | VerticalTransducer | Couch-HW-VerticalTransducer | Couch-HW-VerticalTransducer | NotEngineering-partreplace |  | Couch-HW | Couch-HW-VerticalTransducer | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | y
1091 |  |  | China | 林甸县医院 | 345038 | 44378 |  | 02N/Ingenuity Core128/Elite China/need push IST | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: Calibration/Adjustment Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ Safety Questions : Case: 0117818159 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wennian Han Last updated by: Wennian Han 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: need push IST Troubleshooting Action: site engineer and customer forget IT account and olevel account password, so PM engineer can not enter in olevel account. want to push ISP Repair Action: Push IST failed, FSE feedback he will reset the password when next time onsite service Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: NA SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 5 |  | Troubleshooting Action: site engineer and customer forget IT account and olevel account password, so PM engineer can not enter in olevel account. want to push IST Repair Action: Push IST failed, FSE feedback he will reset the password when next time onsite service | [12 empty] | 117818159 |  | 1069652 | TRUE | Opened - Pending Translation | Closed - non-complaint | [7 empty] | Customer | [6 empty] | Password | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1092 |  |  | China | 山东省泰安市中医二院 | 345037 | 44378 |  | 02N/Ingenuity Core128/Elite China/timeout | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ Safety Questions : Case: 0117817735 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wennian Han Last updated by: Wennian Han 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 扫描失败。 客户问题详细描述或故障现象： 扫描超时 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： S MASTER SEQUENCER PREP DC READY TIMEOUT 远程诊断与测试结果： 扫描超时，重启后恢复，客户需要工程师到场检查设备 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 检查供电，检查DC相关 SAP Transfer : Case Transferred to SAP External Remarks : 01.07.2021 15:09:28 UTC Tao Chen (86038373)故障现象: CT扫描间歇性报错 服务内容: 重刷Rhost应用程序，做灯丝, 空气校正 服务结果: 维修后,设备恢复正常. Internal Remarks : 01.07.2021 15:16:00 UTC Tao Chen (86038373) “RDI Accuracy: YES" $Q7: (Confirmed reported problem, CT scanning failed intermittently) $Q8: ( Check gantry log, found "S_MASTER_SEQUENCER_PREP_DC_SERIES_READY_TIMEOUT" error) $Q9: ( Do boot flash burner and application flash burner for Rhost, restore NVROM) $Q10: ( Test and inspection data refer to: PA. Device returned to full functionality) $Q11: Calibrated Tools name ( N/A) ^Current Tube installed date: N/A ^Current Tube usage: N/A SAP Transfer : Dispatched for Onsite Support ----------------------------------------------------------------------------- | 3 |  | Do boot flash burner and application flash burner for Rhost, restore NVROM, filament and air calibration | [12 empty] | 117817735 |  | 1069371 | TRUE | Opened - Pending Translation | Closed - Done |  | Timeout | [5 empty] | Engineering | [6 empty] | Timeout | [5 empty] | Refresh |  |  | System issue | System | TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | Timeout |  | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | y
1093 |  |  | China | 正镶白旗医院 | 345033 | 44380 |  | 02N/Core128/reset password | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Core128/ Next Step Action : FSE：苏瑞杰-13304719335 call in FSE Notes : This PO is a Reference Number only Safety Questions : Case: 0117830412 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: reset password Troubleshooting Action: reset password Repair Action: reset password SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support ----------------------------------------------------------------------------- | 5 |  | reset password | [12 empty] | 117830412 |  | 1075696 | TRUE | Opened - Pending Translation | Closed - non-complaint | [7 empty] | Customer | [6 empty] | Password | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1094 |  |  | China | 息烽县中医医院 | 345014 | 44379 |  | Wechat:02N/Ingenuity Core128/Ct box down button stuck | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : Wechat:02N/Ingenuity Core128/Ct box down button stuck T2 Activities : Problem Description: 报修为工作站 EQ87453892已修改 SAP Transfer : Case Transferred to SAP Safety Questions : Case: 0117826032 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wennian Han Last updated by: Wennian Han 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 控制面板按键失灵。 客户问题详细描述或故障现象： ctbox按键问题 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： NA 远程诊断与测试结果： CTBOX下降键不灵，其他键正常。机架面板下降键控制正常。 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 检查CTBOX；检查接线、供电等。 建议备件12NC及名称： " 459801705022 CT BOX ASSEMBLY_BEST" SAP Transfer : Dispatched for Onsite Support Next Step Action : PPD CT BOX ASSEMBLY_BEST 459801705022 1 备件无货，已安排订货，FSE：唐广林-86-18084189777 Next Step Action : tang guanglin ----------------------------------------------------------------------------- | 5 |  | 459801705022 CT BOX ASSEMBLY_BEST Ct box down button stuck | [12 empty] | 117826032 |  | 1074175 | TRUE | Opened - Pending Translation | Evaluation - Pending Translation |  | CTUDT00632185 | [5 empty] | Engineering | [6 empty] | CT Box | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | CT Box | Gantry-HW-CT Box | Gantry-HW-CT Box | CTUDT00632185 | CT Box | Gantry-HW | Gantry-HW-CT Box | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | y
1095 |  |  | Japan | 医療法人清翠会 牧病院 | 341017 | 44380 |  | 01CS: KK: Image transfer to PACS is FAIL, but PACS is able to receive images. >> It seems that it was not sent to a device called Radamés. | BU Q&R Feedback: No SAP Customer Symptom Code: IONW Customer Symptom Code: Interoperability/Network Issue Reason: Application Support Engineer Failure Code: Unknown Cause Repair Activity Code: Instructed Customer Problem Reported by Customer : 01CS:KK:PACSへの画像転送がFAILになるがPACSでは画像受信できている。>>送れていなかったのはRadamesというデバイスに対してだったようです。 T2 Activities : Problem Description: Problem description by engineer : 対象装置:CT 故障個所:01Cosole 最初に発生したのはいつか?(~以内):1 Hour 問題発生の頻度は?:First Occurrence エンドユーザｰのワークフローのどの時点において問題が発生しているのか?:ネットワーク関連(Network) 最近どのようなイベントが発生したか?:N.A. 顧客は問題について具体的なシステムイベントレポートを作成しているか?:いいえ。お客様にログファイル、レポートは作成して頂けませんでした。 Troubleshooting Action: What steps are taken during the system diagnosis : 対応後の結果:改善された 故障個所:01Cosole 診断中にどのようなステップを実施したか?: PACSには画像は送られていました。 Radmesというデバイスへの転送がFailしていたようです。 お客様の勘違いだったそうです。様子見となりました。 Error # and/or description of error : NA Repair Action: Completed or Recommended Repair Action "External" : 【現象確認】 問合せの為、N.A. 【作業内容詳細】 PACSには画像は送られていました。 Radmesというデバイスへの転送がFailしていたようです。 お客様の勘違いだったそうです。様子見となりました。 【確認内容】 良好 OneEMS internal remarks : Problem description by engineer : 対象装置:CT 故障個所:01Cosole 最初に発生したのはいつか?(~以内):1 Hour 問題発生の頻度は?:First Occurrence エンドユーザｰのワークフローのどの時点において問題が発生しているのか?:ネットワーク関連(Network) 最近どのようなイベントが発生したか?:N.A. 顧客は問題について具体的なシステムイベントレポートを作成しているか?:いいえ。お客様にログファイル、レポートは作成して頂けませんでした。 What steps are taken during the system diagnosis : 対応後の結果:改善された 故障個所:01Cosole 診断中にどのようなステップを実施したか?: PACSには画像は送られていました。 Radmesというデバイスへの転送がFailしていたようです。 お客様の勘違いだったそうです。様子見となりました。 Error # and/or description of error : NA Completed or Recommended Repair Action "External" : 【現象確認】 問合せの為、N.A. 【作業内容詳細】 PACSには画像は送られていました。 Radmesというデバイスへの転送がFailしていたようです。 お客様の勘違いだったそうです。様子見となりました。 【確認内容】 良好 Completed or Recommended Repair Action "Internal" : --Troubleshooting Action-- 対応後の結果:改善された 故障個所:01Cosole 診断中にどのようなステップを実施したか?: PACSには画像は送られていました。 Radmesというデバイスへの転送がFailしていたようです。 お客様の勘違いだったそうです。様子見となりました。 --Repair Action-- --Subject-- 01CS:KK:ﾏﾂﾓﾄ様:PACSへの画像転送がFAILになるがPACSでは画像受信できている。>>送れていなかったのはRadamesというデバイスに対してだったようです。 --Problem Description-- 対象装置:CT 故障個所:01Cosole 最初に発生したのはいつか?(~以内):1 Hour 問題発生の頻度は?:First Occurrence エンドユーザｰのワークフローのどの時点において問題が発生しているのか?:ネットワーク関連(Network) 最近どのようなイベントが発生したか?:N.A. 顧客は問題について具体的なシステムイベントレポートを作成しているか?:いいえ。お客様にログファイル、レポートは作成して頂けませんでした。 --4/8Elements-- Q1 報告された問題は確認できましたか? 問合せの為、N.A. Q2 他の問題は確認できましたか? 無し Q3 問題を解決するために何をおこないましたか? PACSには画像は送られていました。 Radmesというデバイスへの転送がFailしていたようです。 お客様の勘違いだったそうです。様子見となりました。 Q4 対応後の装置は順調ですか? 良好 Q5 その検査はどの様に終了しましたか? 無し Q6 問題により間違った所見（誤診、誤治療）の有無: 無し Q7 意図しない不要な曝射の有無: 無し Q8 アーチファクトの有無: 無し Safety Questions : Case: 0117830498 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: KENICHI KOJIMA Last updated by: KENICHI KOJIMA 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 4 |  | Images were sent to PACS. It seems that the transfer to a device called Radmes was failing. It seems that it was a customer_s misunderstanding. It was a wait-and-see. The customer was unable to create a log file or report. | [12 empty] | 117830498 |  | 1075720 | TRUE | Evaluation - Pending Translation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | misunderstanding | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1096 |  |  | Japan | Social Medical Corporation BOKOI, Nikko Memorial Hospital | 341011 | 44377 |  | 03DS:KK: Scanning stopped during a contrast scan.>> He was able to perform the examination with a rescan.>>S [Truncated] | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 03DS:KK: Scanning stopped during a contrast scan.>> He was able to perform the examination with a rescan.>>S_STARTER_FAULT_NOT_UP_TO_SPEED SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Support T2 Activities:Problem Description: Problem description by engineer : Equipment affected: CT Failure location: 03High Voltage Energy When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user_s work flow did the problem occur?: (During Acquisition/Scan) What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem? No.The customer has not created a log file or report. Troubleshooting Action: What steps are taken during the system diagnosis: Result of response: Problem solved. Failure location: 03High Voltage Energy What steps did you perform to diagnose the problem? Stopped during a contrast examination. S_STARTER_FAULT_NOT_UP_TO_SPEED and S_ROTOR_STOP_ANODE_NO_MORE_SPINNING had occurred. The examination was possible with a rescan. Sugawara will respond today. Error # and/or description of error: NA Repair Action: Completed or Recommended Repair Action "External": The service engineer for the site visited and worked on the equipment. OneEMS internal remarks : Problem description by engineer : Equipment affected: CT Failure location: 03High Voltage Energy When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user_s work flow did the problem occur?: (During Acquisition/Scan) What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem? No.The customer has not created a log file or report. What steps are taken during the system diagnosis: Result of response: Problem solved. Failure location: 03High Voltage Energy What steps did you perform to diagnose the problem? Stopped during a contrast examination. S_STARTER_FAULT_NOT_UP_TO_SPEED and S_ROTOR_STOP_ANODE_NO_MORE_SPINNING had occurred. The examination was possible with a rescan. Sugawara will respond today. Error # and/or description of error: NA Completed or Recommended Repair Action "External" : The service engineer for the site visited and worked on the equipment. Completed or Recommended Repair Action "Internal": --Troubleshooting Action-- Result of response: Problem solved. Failure location: 03High Voltage Energy What steps did you perform to diagnose the problem? Stopped during a contrast examination. S_STARTER_FAULT_NOT_UP_TO_SPEED and S_ROTOR_STOP_ANODE_NO_MORE_SPINNING had occurred. The examination was possible with a rescan. Sugawara will respond today. --Repair Action-- Recommended 12Nc: --None--453567560042:80KW High Speed Starter Assembly Qty: 1SLOC:--None-- --Subject-- 03DS:KK:Mr. Tsukakoshi: Scanning stopped during a contrast scan.>> He was able to perform the examination with a rescan.>>S_STARTER_FAULT_NOT_UP_TO_SPEED --Problem Description-- Equipment affected: CT Failure location: 03High Voltage Energy When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user_s work flow did the problem occur?: (During Acquisition/Scan) What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem? No.The customer has not created a log file or report. SAP Message: LOGI Task added Successfully!- Safety Questions:Case: 0117806825 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: KENICHI KOJIMA Last updated by: KENICHI KOJIMA 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Message: Components Added Succesfully to SWO!- Internal Remarks: 30.06.2021 08:30:14 UTC MOTOTSUGU SUGAWARA (81100170) Replacement parts: Replacement parts: As below 12Nc:453567560042 Name: 80KW High Speed Starter Assembly QTY:1 [Report on mistakes in parts order] No mistakes. Situation monitoring information: Not monitoring the situation TS Request: No parts to return. Additional CT information: Q1. How did the examination finish? As it stopped before X-ray emission, no extra irradiation Q2. Were there any mistaken findings (misdiagnosis, mistaken treatment) due to the problem? None. Q3. Were there any artifacts? None. Q4. Was there any unintended unnecessary exposure? None. Repair sticker: I affixed a Repaired sticker. [Other information to share] 1. Was treatment/quarantine carried out for a patient who has or is suspected of having new coronavirus? No 2. Is it necessary to go into treatment rooms/beds where there are infected patients?(Including workplace and routes of movement) No Did you agree on how to complete the report with the person in charge of devices? Yes. I completed normal reporting that day. 30-Jun-2021 81100170 FSE Handover Tool1.3.4.32612.DB13 20210115 External Remarks : 30.06.2021 08:30:11 UTC MOTOTSUGU SUGAWARA (81100170) [Replacement parts] Replacement parts: As below 12Nc:453567560042 Name: 80KW High Speed Starter Assembly QTY:1 [Report on mistakes in parts order] No mistakes. Situation monitoring information: Not monitoring the situation TS Request: No parts to return. Additional CT information: Q1. How did the examination finish? As it stopped before X-ray emission, no extra irradiation Q2. Were there any mistaken findings (misdiagnosis, mistaken treatment) due to the problem? None. Q3. Were there any artifacts? None. Q4. Was there any unintended unnecessary exposure? None. Repair sticker: I affixed a Repaired sticker. [Other information to share] 1. Is treatment or quarantining of confirmed or suspected Covid-19 patients taking place? No 2. Is it necessary to go into treatment rooms/beds where there are infected patients?(Including work area as well as movement corridors.) No Did you agree on how to complete the report with the person in charge of devices? Yes. I completed normal reporting that day. 30-Jun-2021 81100170 FSE Handover Tool1.3.4.32612.DB13 20210115 ----------------------------------------------------------------------------- | 4 |  | Stopped during a contrast examination. S_STARTER_FAULT_NOT_UP_TO_SPEED and S_ROTOR_STOP_ANODE_NO_MORE_SPINNING had occurred. Replacement parts: As below 12Nc:453567560042 Name: 80KW High Speed Starter Assembly QTY:1 | [12 empty] | 117806825 |  | 1064632 | TRUE | Evaluation | Closed - Done |  | NotEngineering | [5 empty] | Engineering | [6 empty] | High Speed Starter Assembly | [5 empty] | Refresh |  |  | System issue | System | NotEngineering | ShouldbeOthers | NotEngineering-ShouldbeOthers | NotEngineering-ShouldbeOthers | NotEngineering |  | NotEngineering | NotEngineering-ShouldbeOthers | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | y
1097 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 337000 | 44379 |  | WS designation | Work Order #: WO-03785420 Soner Saglam The system fully meets specification and returned to use Problem Description by Engineer : User requested to define WS to CT. Resolution : It is understood that the specified WS is already defined.----------------------------------------------------------------------------- | 3 |  | Problem Description by Engineer : User requested to define WS to CT. Resolution : It is understood that the specified WS is already | [12 empty] | 117826878 |  | 1070764 | TRUE | Opened - Pending Translation | Closed - non-complaint | [7 empty] | Customer | [6 empty] | misunderstanding | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1098 |  |  | Japan | 医療法人清翠会 牧病院 | 341017 | 44378 |  | CT:APP:RK：質問：ｲﾇｲさま | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Application Support Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer : Q:brain2DとheadでMPRを作成すると印象が結構変わりますが、プロトコルでAutoで作成したMPRは変化が小さくなっているように感じるのですが、slice厚はどのようになっていますか＞ A:"brain2Dは最小slice厚1.25ｍｍ、headでは最小slice厚0.8mmです。しかしプロトコルでautoで作成されるMPRのslice厚が不明であったため、thin slice settingより確認いただきました。 確認できたslice厚はbrain2D0.625mm,head0.8mmであったため、これがMPRの見た目を変えている可能性があることをお伝えしました。" OneEMS internal remarks : Q1:報告された問題は確認できましたか？：確認済 Q2：他の問題は確認できましたか？；無 Q3：問題を解決するために何を行いましたか？ １：電話で解決した場合：説明 Q4：対応後の装置は順調ですか？：良好 Safety Questions : Case: 0117820388 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: RYUSUKE KUJIRAI Last updated by: RYUSUKE KUJIRAI 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 44379.8236 | [14 empty] | 117820388 |  | 1071024 |  |  | Closed - Non-Complaint | [20 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1099 |  |  | Japan | 医療法人清翠会 牧病院 | 341017 | 44378 |  | CT:APP:RK:質問:ﾎｼさま | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Application Support Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer : Q:頭のscantypeでbrain2Dとheadがありますが、2つは違いますか？ A:はい。それぞれは設定できる項目が異なります。選択できるフィルターや最小slice厚など OneEMS internal remarks : Q1:報告された問題は確認できましたか？：確認済 Q2：他の問題は確認できましたか？；無 Q3：問題を解決するために何を行いましたか？ １：電話で解決した場合：説明 Q4：対応後の装置は順調ですか？：良好 Safety Questions : Case: 0117820343 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: RYUSUKE KUJIRAI Last updated by: RYUSUKE KUJIRAI 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 44379.8201 | [14 empty] | 117820343 |  | 1070998 |  |  | Closed - Non-Complaint | [20 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1100 |  |  | China | Dongguan Donghua Hospital Co., Ltd. | 40059 | 44377 |  | WeChat:02N/Ingenuity CT 728326/Add node | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Calibration/Adjustment Problem Reported by Customer:Add node Safety Questions:Case: 0117808304 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Cao Last updated by: Liu Cao 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer_s problems: Plan to add devices tomorrow and need assistance in adding nodes Detailed description of Customer_s problems or malfunctions: Plan to add devices tomorrow and need assistance in adding nodes It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: NA Remote troubleshoot and test result: Contact the customer to confirm the time to assist in adding nodes PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Contact the customer to confirm the time to assist in adding nodes Test/Inspection: N/A Test Equipment: N/A SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support ----------------------------------------------------------------------------- | 5 |  | 计划明后天添加设备 需协助添加节点 | [12 empty] | 117808304 |  | 1064976 | FALSE | Opened | Closed - non-complaint | [7 empty] | NAC | [6 empty] | PACS Config | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1101 |  |  | United States | Philips Medical Capital | 40055 | 44379 |  | 9507CFFD7 NYN Changing Filter | Work Order #: WO-03788656 Mujica, David The system fully meets specification and returned to use Problem description by Clinical Support Specialist : Customer Function/Role: Technologist Device Use: In use Expected and Unexpected Results: Changing to Bone filter User Impact: none Patient Impact: none Current Software Version: 4.1 Diagnostic performed by Engineer : Chelsey was assisted in changing a filter on their CT Soft Tissue neck Exam Card via Exam Card Manager, show all parameters, Bone Recon. Exploration of filters in show all parameters was explored and changes were made successfully and saved. Chelsey had no further questions or need for clinical support. No rescan, no loss of data, no reinjection per customer. ----------------------------------------------------------------------------- | 5 |  | assisted in changing a filter on their CT Soft Tissue neck Exam Card via Exam Card Manager | [12 empty] | 117829629 |  | 1072793 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | NAC | [6 empty] | Application support | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1102 |  |  | China | 浙江省乐清第一人民医院 | 40054 | 44377 |  | Wechat:02N/Ingenuity CT 728326/scan interrput | BU Q&R Feedback: No SAP Customer Symptom Code: BOOT Customer Symptom Code: Boot up/Power-up Problem Reason: Corrective Maintenance Engineer Failure Code: Electrical/Electronic Failure Repair Activity Code: Replace Part Problem Reported by Customer : scan interrput Safety Questions : Case: 0117811180 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Weng Jinming Last updated by: Weng Jinming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Troubleshooting Action: repeat case0117811185 Repair Action: repeat case0117811185 Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | repeat case0117811185 | [12 empty] | 117811180 |  | 1065963 | FALSE | Evaluation | Closed - Duplicate | [7 empty] | NAC | [6 empty] | Duplicate | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1103 |  |  | China | Zhejiang Province Yueqing No. 1 People’s Hospital | 40054 | 44377 |  | 02N/Ingenuity CT 728326/scan interrupt | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 02N/Ingenuity CT 728326 Safety Questions:Case: 0117811185 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer_s problems: Other types of error. Detailed description of Customer_s problems or malfunctions: Scanning was interrupted. It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: INvalid dfs settings Remote troubleshoot and test result: R2D malfunction. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Replace R2D. 0 R2D Assembly 459800467343 SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Internal Remarks: 01.07.2021 04:05:56 UTC Yi Huang (86036249) RDI Accuracy:Yes $Q7: (Confirmed reported problem&system scan stop ) $Q8: (check errorlog found acq not ok&invalid dfs settings, R2D LED5 is red) $Q9: (replace R2d) $Q10: (Test and inspection data refer to: PA. Device returned to full functionality.) $Q11: Calibrated tools name(N/A) S/N (N/A) Next calibration date (N/A). For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date:N/A ^Current Tube usage: second:N/A External Remarks: 01.07.2021 04:05:45 UTC SUP JCOUSER (86036249) service details: enhanced scan stopped. Checked the log and found DFS setting abnormal. Replace R2D. service result: device returns to normal working condition after service Other significant matters: none Next step Action: PPD parts R2D Assembly 459800467341 1 Have arranged for 12H to the site. (alternative number has been confirmed) FSE: Huang Yi-86-15958725282 ----------------------------------------------------------------------------- | 3 |  | R2D Assembly 459800467341 $Q7: (Confirmed reported problem&system scan stop) $Q8: (check errorlog found acq not ok&invalid dfs settings, R2D LED5 is red) $Q9: (replace R2d) $Q10: (Test and inspection data refer to: PA.Device returned to full | [12 empty] | 117811185 |  | 1065964 | TRUE | Opened | Closed - Done |  | CTUDT00630864 | [5 empty] | Engineering | [6 empty] | R2D | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | R2D | Gantry-HW-R2D | Gantry-HW-R2D | CTUDT00630864 | R2D | Gantry-HW | Gantry-HW-R2D | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | y
1104 |  |  | China | Guangxi Province Teng County People’s Hospital | 40036 | 44378 |  | application/Ingenuity CT 728326 | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer:application Safety Questions:Case: 0117817621 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Zixuan Last updated by: Wang Zixuan 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 3 |  | 杨俊13817647801 & 郭凡 13501285722均未接通，已邮件 | [12 empty] | 117817621 |  | 1069326 | TRUE | Opened | Closed - non-complaint | [7 empty] | Application | [6 empty] | Application | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1105 |  |  | China | 北京市中西医结合医院 | 40032 | 44378 |  | 02N/Ingenuity CT 728326/ECG Loss | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/IntelliSpace IX/LX Workstations/ T2 Activities : Problem Description: 客户问题概述： 新装机，没有安装心电系统 设备本身技术问题： 是。 Troubleshooting Action: 远程诊断与测试结果： 需要工程师到场，安装心电系统 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 需要工程师到场，安装心电系统 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Safety Questions : Case: 0117817822 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Weng Jinming Last updated by: Weng Jinming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A External Remarks : 01.07.2021 13:55:50 UTC SUP JCOUSER (86038892)服务内容：机器ECG无法使用，更换EKG线缆后正常。 服务结果：服务后设备恢复正常。 Internal Remarks : 01.07.2021 13:55:53 UTC Xiang Zhang (86038892) "RDI Accuracy: Yes" $Q7: Confirm reported problems:The system has no ECG signal. $Q8: Trouble shooting steps:The ECG cable is in bad condition,the ECG cable needs to be replaced $Q9: Function returns to normal after the ECG cable is replaced。 $Q10: Test and Inspection data refer to PA& Device status:Device returned to full functionally $Q11: Calibrated Tools refer to PA ----------------------------------------------------------------------------- | 5 |  | The system has no ECG signal. The ECG cable is in bad condition,the ECG cable needs to be replaced | [12 empty] | 117817822 |  | 1069439 | TRUE | Opened - Pending Translation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | ECG | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1106 |  |  | China | Beijing City Hospital of Traditional Chinese and Western Medicine | 40032 | 44376 |  | 02N/Ingenuity CT 728326/CT box fault | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : Ingenuity CT 728326 Safety Questions : Case: 0117798433 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingsong Last updated by: Lu Mingsong 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： CT BOX按键故障 客户问题详细描述或故障现象： CT BOX按键故障 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： N/A 远程诊断与测试结果： N/A PRS TR状态： 无PRS。 Repair Action: 建议现场维修步骤： 更换CT BOX。 建议备件12NC及名称： CT BOX ASSEMBLY_BEST：459801705022 客户需求备注： 无 SAP Transfer : Dispatched for Onsite Support External Remarks : 02.07.2021 12:52:58 UTC Guochao Fan (86039456)服务内容：更换CTbox 服务结果：服务后设备恢复正常。 Internal Remarks : 02.07.2021 12:53:07 UTC Guochao Fan (86039456) "RDI Accuracy: Yes" $Q7: Confirm reported problems: ctbox Out of bed key malfunction $Q8: Trouble shooting steps:for Out of bed key malfunction .The ctbox is in bad condition,the ctbox needs to be replaced $Q9: Function returns to normal after the ctbox is replaced。 $Q10: Test and Inspection data refer to PA& Device status:Device returned to full functionally $Q11: N/A SAP Transfer : Case Transferred to SAP Next Step Action : PPD 459801705022 *1 ,已安排订货，FSE范国超-86-13520729230 ----------------------------------------------------------------------------- | 3 |  | $Q7: Confirm reported problems: ctbox Out of bed key malfunction $Q8: Trouble shooting steps:for Out of bed key malfunction .The ctbox is in bad condition,the ctbox needs to be replaced $Q9: Function returns to normal after the ctbox is replaced CT BOX ASSEMBLY_BEST：459801705022 | [12 empty] | 117798433 |  | 1059417 | TRUE | Opened - Pending Translation | Closed - No Further Investigation |  | CTUDT00632185 | [5 empty] | Engineering | [6 empty] | CT Box | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | CT Box | Gantry-HW-CT Box | Gantry-HW-CT Box | CTUDT00632185 | CT Box | Gantry-HW | Gantry-HW-CT Box | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | y
1107 |  |  | China | Beijing City Hospital of Traditional Chinese and Western Medicine | 40032 | 44376 |  | 05R/Ingenuity CT 728326/password failure | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Calibration/Adjustment SAP Transfer:Case Transferred to SAP Problem Reported by Customer: Password failure Safety Questions:Case: 0117797920 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Tao Fengming Last updated by: Tao Fengming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer_s problems: Other types of error. Detailed description of Customer_s problems or malfunctions: User password has too many error, account is locked It is a technology problem with the equipment itself: No. Troubleshooting Action: key fault code or symptom: Password is locked, fail to log in Remote troubleshoot and test result: No PRS, admin password has also been forgot, need to reset password on site PRS TR status: No PRS. Repair Action: Suggest onsite service procedures: 1. Reset the password. Test/Inspection: N/A Test Equipment: N/A SAP Transfer:Dispatched for On site Support External Remarks: 29.06.2021 05:38:01 UTC HongHong Han (86039436) service details: change login password Service Result: The device returned to working normally. Internal Remarks: 29.06.2021 05:38:11 UTC HongHong Han (86039436) "RDI Accuracy:Yes" $Q7: Confirmed reported problem, need to change password $Q8: need to change password $Q9: change ct and it password $Q10: Test and Inspection data refer to PA & Device status returned to fully functionality. $Q11:N/A Next Step Action: dispatch engineer han honghong ----------------------------------------------------------------------------- | 3 |  | $Q7: Confirmed reported problem, need to change password $Q8: need to change password $Q9: change ct and it password | [12 empty] | 117797920 |  | 1059165 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Password | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1108 |  |  | China | Henan Province Puyang City People’s Hospital | 40031 | 44375 |  | Wechat:02N/Ingenuity CT 728326/tube over heat | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer:tube overheating Safety Questions:Case: 0117791222 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Jianqiang Last updated by: Wang Jianqiang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Case Transferred to SAP T2 Activities:Problem Description: Summary of customer_s problems: The old tube overheated recently, it used to be OK Detailed description of Customer_s problems or malfunctions: The old tube overheated recently, it used to be OK It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: Tube overheat Remote troubleshoot and test result: Remote instruct to modify scan MAS. Customer request engineer_s on site. Contact engineer who said no service had been done yet PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Check mAs, check the minimum/maximum values of the setting, check filter mesh SAP Transfer:Dispatched for On site Support ----------------------------------------------------------------------------- | 5 |  | 最近老报球管过热，以前是好的 收集bugreport，calibration info | [12 empty] | 117791222 |  | 1054371 | TRUE | Opened | Closed - Done |  | Tube Heat | [5 empty] | Engineering | [6 empty] | Tube Overheat | [5 empty] | Refresh |  |  | System issue | System | Gantry-SW | Tube over heat | Gantry-SW-Tube over heat | Gantry-SW-Tube over heat | Tube Heat | Tube Heat | Gantry-SW | Gantry-SW-Tube over heat | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | y
1109 |  |  | United States | Central Vermont Medical Center | 40024 | 44375 |  | 0711CFXXX Storage issue on local | Work Order #: WO-03754268 Morrow, Dan The system fully meets specification and returned to use Problem Description by Engineer : Problem description by engineer : Local database percentage dropping. Only 2 days of studies. How often is issue occurring? : First occurrence Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: Database issue Patient Impact/Outcome: none User Impact: Database issue Current Software Version: 4.1 Diagnostic performed by Engineer : Logged into system and checked database. Orphaned studies found. Advised customer to perform a full database rebuild. Once complete the orphaned studies should populate in the local and be ready for deletion. I will check back with the customer within 24 hours. Internal Comments : Error # and/or description of error : No Error Completed or Recommended Repair Action "Internal" : Logged into system and checked database. Orphaned studies found. Advised customer to perform a full database rebuild. Once complete the orphaned studies should populate in the local and be ready for deletion. I will check back with the customer within 24 hours. Resolution : Database rebuild fixed issue. OK to close case. ----------------------------------------------------------------------------- | 3 |  | Logged into system and checked database. Orphaned studies found. Advised customer to perform a full database rebuild. Once complete the orphaned studies should populate in the local and be ready for deletion. I will check back with the customer within 24 hours. Resolution : Database rebuild fixed issue. OK to close case. | [12 empty] | 117794973 |  | 1056146 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | orphaned studies | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1110 |  |  | Japan | 独立行政法人国立病院機構 | 341015 | 44376 |  | 01CS:KKI:問い合わせ:心臓解析の時間分解能に関して | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Application Support Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer : 01CS:KKI:問い合わせ:心臓解析の時間分解能に関して Safety Questions : Case: 0117799961 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: KAZUMASA KIMURA Last updated by: KAZUMASA KIMURA 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 対象装置:CT 故障個所:01Cosole 最初に発生したのはいつか?(~以内):1 Hour 問題発生の頻度は?:First Occurrence エンドユーザｰのワークフローのどの時点において問題が発生しているのか?:その他(Other) 最近どのようなイベントが発生したか?:N.A. 顧客は問題について具体的なシステムイベントレポートを作成しているか?:いいえ。お客様にログファイル、レポートは作成して頂けませんでした。 Troubleshooting Action: 対応後の結果:改善された 故障個所:01Cosole 診断中にどのようなステップを実施したか?: 【コール内容】 ①時間分解能が悪くなる心拍は？グラフがあったら送ってください。 ②他のサイトの心臓CT造影条件を教えてください。 【対応】 ①71bpm前後が悪くなります。大谷さんからお送りします。 ②調べられる範囲で調べ報告します。 Repair Action: 【現象確認】 確認済み 【作業内容詳細】 【コール内容】 ①時間分解能が悪くなる心拍は？グラフがあったら送ってください。 ②他のサイトの心臓CT造影条件を教えてください。 【対応】 ①71bpm前後が悪くなります。大谷さんからお送りします。 ②調べられる範囲で調べ報告します。 【確認内容】 良好 OneEMS internal remarks : --Troubleshooting Action-- 対応後の結果:改善された 故障個所:01Cosole 診断中にどのようなステップを実施したか?: 【コール内容】 ①時間分解能が悪くなる心拍は？グラフがあったら送ってください。 ②他のサイトの心臓CT造影条件を教えてください。 【対応】 ①71bpm前後が悪くなります。大谷さんからお送りします。 ②調べられる範囲で調べ報告します。 --Repair Action-- --Subject-- 01CS:KKI:ﾋｻｷ様:問い合わせ:心臓解析の時間分解能に関して --Problem Description-- 対象装置:CT 故障個所:01Cosole 最初に発生したのはいつか?(~以内):1 Hour 問題発生の頻度は?:First Occurrence エンドユーザｰのワークフローのどの時点において問題が発生しているのか?:その他(Other) 最近どのようなイベントが発生したか?:N.A. 顧客は問題について具体的なシステムイベントレポートを作成しているか?:いいえ。お客様にログファイル、レポートは作成して頂けませんでした。 --4/8Elements-- Q1 報告された問題は確認できましたか? 確認済み Q2 他の問題は確認できましたか? 無し Q3 問題を解決するために何をおこないましたか? 【コール内容】 ①時間分解能が悪くなる心拍は？グラフがあったら送ってください。 ②他のサイトの心臓CT造影条件を教えてください。 【対応】 ①71bpm前後が悪くなります。大谷さんからお送りします。 ②調べられる範囲で調べ報告します。 Q4 対応後の装置は順調ですか? 良好 Q5 その検査はどの様に終了しましたか? 無し Q6 問題により間違った所見（誤診、誤治療）の有無: 無し Q7 意図しない不要な曝射の有無: 無し Q8 アーチファクトの有無: 無し SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 44377.6813 | [14 empty] | 117799961 |  | 1060260 |  |  | Closed - Non-Complaint | [20 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1111 |  |  | China | Shandong Province Tai_an City T.C.M. No. 2 Hospital | 345037 | 44376 |  | 02N/Ingenuity Core128/Elite China/Add printer | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer:add printer Safety Questions:Case: 0117799692 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Zixuan Last updated by: Wang Zixuan 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Add printer Troubleshooting Action: Add printer Repair Action: Add printer Test/Inspection: N/A Test Equipment: N/A SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | 44377.6507 | [14 empty] | 117799692 |  | 1060079 |  |  | Closed - Non-Complaint | [20 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1112 |  |  | United States | Lackey Memorial Hospital | 40015 | 44379 |  | 0711CDN89 Prompted them to change the password and now she can not get into the exam card manager. Telling her she_s not authorized. Customer requested clinical apps. | Work Order #: WO-03788299 Curry, Dyan The system fully meets specification and returned to use Problem description by Clinical Support Specialist : Customer Function/ Role: technologist Device Use: in use name Expected and Unexpected Results: ECM password expired, customer can not access User Impact: none Patient Impact: none Current Software version: 4.1.10 Diagnostic performed by Engineer : transferred call to Daniel M, RSE to page FSE with dongle. ----------------------------------------------------------------------------- Work Order #: WO-03788362 Morrow, Dan The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Customer Function/Role: Tech Device Use: No Patient involved Expected and Unexpected Results: Customer needs EC password reset User Impact: none Patient Impact: none Current Software Version: 4.1.10 Diagnostic performed by Engineer : Customer prompted to reset EC password. Now locked out of EC manager. Unable to reset remotely. Dispatching FSE. Follow Up Required : Log into service and reset customer EC manager password. ----------------------------------------------------------------------------- | 5 |  | Log into service and reset customer EC manager password. | [12 empty] | 117829310 |  | 1072635 | TRUE | Opened | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Password | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1113 |  |  | Turkey | Republic of Turkey Ministry of Health, Bagcilar | 40010 | 44375 |  | Sound is coming from the table. | Work Order #: WO-03753000 Salih Mendi The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Customer Function/Role: technician Device Use: out of use Expected and Unexpected Results: Noise during patient table vertical and vertical movement. User Impact: No user impact. Patient Impact: No patient impact. Current Software Version:NA Follow Up Required : Noise during vertical and horizontal movements was found. Therefore, the site should be visited to lubricate the table. System is partially operational. Diagnostic performed by Engineer : The user was remotely supported for the noise complaint. In diagnosis, noise during vertical and horizontal movements was found. Therefore, the site should be visited to lubricate the table. System is partially operational. ----------------------------------------------------------------------------- Work Order #: WO-03753317 29-Jun-2021 Ugur Altunoz The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: technician Device Use: out of use Expected and Unexpected Results: Noise during patient table vertical and vertical movement. User Impact: No user impact. Patient Impact: No patient impact. Current Software Version:NA Resolution : The noise problem during table vertical movement was solved after lubricating the side rails. The device is operational. ----------------------------------------------------------------------------- | 3 |  | The noise problem during table vertical movement was solved after lubricating the side rails. | [12 empty] | 117793879 |  | 1052062 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | E-125059 | [5 empty] | Engineering | [6 empty] | Couch Noise | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | Lubricant | Couch-HW-Lubricant | Couch-HW-Lubricant | E-125059 | Lubricant | Couch-HW | Couch-HW-Lubricant | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | y
1114 |  |  | Chile | Scotiabank Chile | 40008 | 44350 |  | Couch noise when reaching the top of table, lubricated vertical guide piston |  |  |  | Couch noise when reaching the top of table, lubricated vertical guide piston | [14 empty] | 1076312 | TRUE | Investigation | Closed - Done |  | E-125059 | [5 empty] | Engineering | [6 empty] | Couch Noise | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | Lubricant | Couch-HW-Lubricant | Couch-HW-Lubricant | E-125059 | Lubricant | Couch-HW | Couch-HW-Lubricant | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | y
1115 |  |  | Japan | 独立行政法人国立病院機構 | 341015 | 44375 |  | CT:APP:NO:ﾋｻｷ様:質問 | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Application Support Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer : Q:メタボ検査用プロトコルを作りたいのですが、どのプロトコルから作ればよいですか？ A:Referenceセットから体幹部コンべのプロトコルを選択いただき、編集して保存していただきました。 Safety Questions : Case: 0117792699 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Ootani Natsuki Last updated by: Ootani Natsuki 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 44376.7125 | [14 empty] | 117792699 |  | 1055327 |  |  | Closed - Non-Complaint | [20 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1116 |  |  | Germany | Gesundheit Nord gGmbH Klinikverbund | 40002 | 44377 |  | After a biopsy, the system does not display a dose | Work Order #: WO-03771553 Schwidrogitz Michaela The system fully meets specification and returned to use Problem description by Clinical Support Specialist: The customer reports that after a puncture, no dose info page was created by the system. She created a bug report and phoned Engineering. Diagnostic performed by Engineer: Asked the customer to perform a restart. The dose info page was not displayed afterwards. It occurred for the first time and so far once. Customer will still contact Ms. Schirmer in the FU appointment next week. The customer will call again if it should occur reoccur. ----------------------------------------------------------------------------- | 3 |  | The customer reports that after a puncture, no dose info page was created by the system. Collect bugreport for engineering investigation. | [12 empty] | 117810739 |  | 1061794 | TRUE | Opened | Closed - No Further Investigation |  | TBD_Bugreport | [5 empty] | Engineering | [6 empty] | Biopsy dose | [5 empty] | Refresh |  |  | System issue | System | TBD-Bugreport | TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | TBD_Bugreport |  | TBD-Bugreport | TBD-Bugreport-TBD_Bugreport | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | y
1117 |  |  | United States | Leesburg Regional Medical Center | 40022 | 44375 |  | 0711CDY73 FSE is aware over the weekend scanner locked out duplicate see case 0117788182 | Work Order #: WO-03753586 Atz, Michael The system partly meets specification and is returned to use with restrictions Internal Comments : Error # and/or description of error : none Completed or Recommended Repair Action "Internal" : FSE is aware over the weekend scanner locked out due to password duplicate see case 0117788182 -- Customer used admin PW as a temporary Problem Description by Engineer : Problem description by engineer : FSE is aware over the weekend scanner locked out due to password duplicate see case 0117788182 When was the first occurrence? Within... : 26-Jun-2021 How often is issue occurring? : Continuously Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: No further information is available as per the customer Patient Impact/Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer Current Software Version: Not provided Diagnostic performed by Engineer : FSE is aware over the weekend scanner locked out due to password duplicate see case 0117788182 Follow Up Required : Error # and/or description of error : none Completed or Recommended Repair Action "Internal" : FSE is aware over the weekend scanner locked out due to password duplicate see case 0117788182 -- Customer used admin PW as a temporary ----------------------------------------------------------------------------- Work Order #: WO-03753629 30-Jun-2021 Le David The system fully meets specification and returned to use Problem Description by Engineer : Problem description by engineer : FSE is aware over the weekend scanner locked out due to password duplicate see case 0117788182 When was the first occurrence? Within... : 26-Jun-2021 How often is issue occurring? : Continuously Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: No further information is available as per the customer Patient Impact/Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer Current Software Version: Not provided Resolution : Instructed customer to re log into system. System only times out for 5 mins after incorrect attempts have been reached. Notified customer to setup Emergency Account as alternate login. ----------------------------------------------------------------------------- | 44375.8868 | [14 empty] | 117794397 |  | 1052169 |  |  | Closed - Duplicate | [20 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1118 |  |  | United States | AdventHealth Zephyrhills | 300099 | 44380 |  | Taking a long time to send images to PACS | Work Order #: WO-03790386 Boyd Craig The system partly meets specification and is returned to use with restrictions Diagnostic performed by Engineer : On going issue. Study taking up to an hour to send over to PACS. PRS accessed. No normal errors seen. Customer requesting service today. PO WC0010001651531. Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: Taking a long time to send images to PACS User Impact: None Patient Impact: None Current Software Version: Follow Up Required : Customer requesting service today. PO WC0010001651531. FSE on site to evaluate DICOM settings and check images going to PACS. Work with IT department at facility. Please request port 80 be opened for remote access. ----------------------------------------------------------------------------- Work Order #: WO-03790399 03-Jul-2021 Garcia Marrero, Jorge The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: Taking a long time to send images to PACS User Impact: None Patient Impact: None Current Software Version: Diagnostic performed by Engineer : Proactively ordering Host computer for networks studies transfer slow. Follow Up Required : Installation of host computer. ----------------------------------------------------------------------------- Work Order #: WO-03790759 03-Jul-2021 Garcia Marrero, Jorge Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: Taking a long time to send images to PACS User Impact: None Patient Impact: None Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 3 | [14 empty] | 117831044 |  | 1075189 | #N/A | #N/A | Opened | [7 empty] | Customer | [6 empty] | Network Switch | [5 empty] | Non-Refresh |  |  | Upgrade issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | #REF!
1119 |  |  | United States | AdventHealth Zephyrhills | 300099 | 44379 |  | Gantry won_t initialize | Work Order #: WO-03788701 Powell, Jamie The system does not meet specification and is out of use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : Gantry won_t initialize.. Customer cycled power multiple times but still having issues. Follow Up Required : Gantry won_t initialize.. Customer cycled power multiple times but still having issues. Internal Comments : Spoke to FSE Mark Sher who was already onsite.. States he blue button the gantry and the system is up and scanning. Mark will look at bugrep later. ----------------------------------------------------------------------------- Work Order #: WO-03788737 02-Jul-2021 Sher Mark The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : system had several issues Blue button the gantry to clear initiation error, customr allowed contrast to get in gantry csusing image artifacts.. rings,, troubleshot pacs transfer issue ----------------------------------------------------------------------------- Work Order #: WO-03789543 03-Jul-2021 Garcia Marrero, Jorge The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : Relieve Mark Sher to continue network troubleshooting. The network cable from the hospital network port to the CT host computer was disconnected and used my laptop as a PACS. Transfer of studies from the CT to my mock PACS went well with normal transfer rate. After 8:00 PM the hospital network normalize and was able to transfer studies to any DICOM node inside the hospital network without issues and release the scanner to clinical usage.. There_s an issue in the hospital network that unable the scanner to send studies to other DICOM nodes. ----------------------------------------------------------------------------- | 2 | [14 empty] | 117829654 |  | 1072828 | #N/A | #N/A | Opened |  | CustomerNetIssue | [5 empty] | Engineering | [6 empty] | Multi issue | [5 empty] | Refresh |  |  | Upgrade issue | System | NotEngineering | ShouldbeCustomer | NotEngineering-ShouldbeCustomer | NotEngineering-ShouldbeCustomer | CustomerNetIssue |  |  |  | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | y | y | y | y | y | 0 | y | #REF!
1120 |  |  | United States | AdventHealth Zephyrhills | 300106 | 44377 |  | 0308CDN8A Couch making noise when moving vertically | Work Order #: WO-03773705 Powell, Jamie The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : Couch making noise when moving vertically Follow Up Required : Couch making noise when moving vertically. Grease vertical drive screw and sliding rails ----------------------------------------------------------------------------- Work Order #: WO-03773715 01-Jul-2021 Sher Mark The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : Greased guide rails and tested system ----------------------------------------------------------------------------- | 3 | [14 empty] | 117815027 |  | 1063029 | #N/A | #N/A | Evaluation | [7 empty] | Others | [6 empty] | lubrecate | [5 empty] | Non-Refresh |  |  | Upgrade issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | #REF!
1121 |  |  | United States | Select Specialty Hospitals | 300037 | 44376 |  | Parts Quote | [16 empty] | 117802415 | 1061640 | 1061640 | #N/A | #N/A | Closed - Non-Complaint | [23 empty] | Upgrade issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1122 |  |  | United States | Select Specialty Hospitals | 300037 | 44371 |  | Cracked compensator message popped out ; Replaced the A-plane collimator and performed calibration. | Work Order #: WO-03738673 Atz, Michael The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Customer Function/Role: tech Device Use: diag Expected and Unexpected Results: N/A User Impact: N/A Patient Impact: N/A Current Software Version: N/A Diagnostic performed by Engineer : CRACKEDWEDGE BR ROHS COLLIMATOR FIELD SERVICE Follow Up Required : On-site recommendations according to the system_s PM manual: Re-confirm the crack, running the Compensator Analyzer Tool. (See the Compensator Analyzer Tool Guide) If confirmed by the tool, replace the compensator/wedge cassette. If not confirmed by the tool, perform compensator visual inspection. (See PM manual, Appendix A: Visual Collimator Compensator Inspection). If visually confirmed, replace the compensator/wedge cassette. After replacement, perform IQ tests. ----------------------------------------------------------------------------- Work Order #: WO-03738802 29-Jun-2021 Beard Michael The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: tech Device Use: diag Expected and Unexpected Results: N/A User Impact: N/A Patient Impact: N/A Current Software Version: N/A Resolution : Arrived onsite. replaced A plane collimator following instructions in DMR 4598 000 72481 Rev X. Performed calibrations following instructions in DMR 4598 015 87311 Rev E. Tested and verified system is performing correctly. Returned back over to customer for clinical use. ----------------------------------------------------------------------------- | 3 | [14 empty] | 117779333 |  | 1042877 | #N/A | #N/A | Closed - Done | [7 empty] | Engineering | [6 empty] | A-Plane | [5 empty] | Non-Refresh |  |  | Upgrade issue | system | Gantry-HW | aplane | Gantry-HW-aplane | NoEnoughInfo | 0 |  | Gantry-HW | Gantry-HW-aplane | No | 0 | 0 | 0 | 0 | 0 | No | y | y | y | 0 | 0 | y | y | y |  |  | y
1123 |  |  | Spain | POLICLINICA NTRA SRA DEL ROSARIO SL | 336190 | 44364 |  | Gantry control panel key is broken | Work Order #: WO-03672241 20-Jun-2021 Miguel Ribot Martínez The system fully meets specification and returned to use. Parts Consumed : 1.00000,453567003451,TENSIONER ASSY 1.00000,459800484501,CONTROL PANEL ASSEMBLY-FRONT LH 2.00000,453567055351,AIR FILTER 1.00000,453566492401,POLY-VEE BELT Diagnostic Performed by Engineer: During the last preventive maintenance, we saw that the key to enter the couch in the gantry on the left control panel is worn out. The air filters have to be replaced with new ones. A new belt has also been ordered for the gantry with its tensor. All these materials are sent to the Polyclinic and will be replaced at the next scheduled visit. Regulatory Questions : N/A Problem Description by Engineer: Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: ----------------------------------------------------------------------------- | [15 empty] | 117716200 | 1020312 | 1020312 | #N/A | #N/A | Closed - No Further Investigation | [20 empty] | Non-Refresh |  |  | Upgrade issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | #REF!
1124 |  |  | United States | Philips Medical Capital | 300057 | 44363 |  | 0699CDY72 System Stuck On Initializing Cannot Scan | Work Order #: WO-03665769 Holeman, David The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Problem description by engineer : System Stuck On Initializing Cannot Scan When was the first occurrence? Within... : 15-Mar-2021 08:00:00 How often is issue occurring? : Intermittent Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: System Stuck On Initializing Cannot Scan Patient Impact/Outcome: None User Impact: System Stuck On Initializing Cannot Scan Current Software Version: 4.1.7.25026 Follow Up Required : Error # and/or description of error : S_HMC_CTRL_A_MOTOR_SHORT_CIRCUIT_OR_OVERTEMP NORMAL (P1 = 26) S_DMC_SHOT_COMM_COLLAPSE NORMAL (P1 = 4098) Error #12 - CheckGantryConnection: ACQ board or Gantry-to-CIRS datapath failure. Unexpected gantry Confirm Resend: 1 (1 = error) Reading resends: 147 CRC error count Chan A: 0 Chan B: 147 Chan C: 0 Chan D: 147 Loss of Signal Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 Line Fault Indicator Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 On-board overflow retry count: 41 Error #12 - : ACQ board was inactive for at least 328 ms Completed or Recommended Repair Action "Internal" : For S_HMC_CTRL_A_MOTOR_SHORT_CIRCUIT_OR_OVERTEMP NORMAL (P1 = 26) 1. Check for an obstruction or mechanical load problem. 2. On a BedRock couch make sure the power to the AMC is between 14VAC and 62VAC. 3. Check / replace the horizontal motor Note: On a BedRock table, It maybe normal to see error For S_DMC_SHOT_COMM_COLLAPSE NORMAL (P1 = 4098) Inspect data transfer between gantry and IRS/Perform Visual Diagnostics. Run ACQ board diagnostics. Connect the optical cable directly from the ACQ PCB to the RCOM. Try to make scans without rotation to see if scanning is possible without above error. Clean optical slipring. Inspect Rotor transmitter and Stator receivers, cabling/connections and incoming power to those modules (i.e. 12vdc from FRC). Repair/replace as necessary. Note: These couch & datapath errors seem to be intermittent, but occurring around the same time, thus one of these errors may be triggering the other error. Diagnostic performed by Engineer : CT Tech reports "System stuck on initializing Cannot scan" and RECON Box error. Tech states she had this error on 6/15 in AM also. Tech is currently power-cycling the Host/IRS/IMR and gantry. Inspected logger which indicates: S_HMC_CTRL_A_MOTOR_SHORT_CIRCUIT_OR_OVERTEMP NORMAL (P1 = 26) and S_DMC_SHOT_COMM_COLLAPSE NORMAL (P1 = 4098) before the power-cycle. Tech is able to scan successfully (no errors noted) post-power cycle. Tech has clinical scans to perform and requested RSE to call back at 9am. RSE called Tech back and Tech reports 2 clinical scans OK, but is concerned about the intermittent issues which she reports are occurring more frequently. RSE bug report analysis from 6/14/21 indicates the same couch and datapath issues around 1412. Internal Comments : Error # and/or description of error : S_HMC_CTRL_A_MOTOR_SHORT_CIRCUIT_OR_OVERTEMP NORMAL (P1 = 26) S_DMC_SHOT_COMM_COLLAPSE NORMAL (P1 = 4098) Error #12 - CheckGantryConnection: ACQ board or Gantry-to-CIRS datapath failure. Unexpected gantry Confirm Resend: 1 (1 = error) Reading resends: 147 CRC error count Chan A: 0 Chan B: 147 Chan C: 0 Chan D: 147 Loss of Signal Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 Line Fault Indicator Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 On-board overflow retry count: 41 Error #12 - : ACQ board was inactive for at least 328 ms Completed or Recommended Repair Action "Internal" : For S_HMC_CTRL_A_MOTOR_SHORT_CIRCUIT_OR_OVERTEMP NORMAL (P1 = 26) 1. Check for an obstruction or mechanical load problem. 2. On a BedRock couch make sure the power to the AMC is between 14VAC and 62VAC. 3. Check / replace the horizontal motor Note: On a BedRock table, It maybe normal to see error For S_DMC_SHOT_COMM_COLLAPSE NORMAL (P1 = 4098) Inspect data transfer between gantry and IRS/Perform Visual Diagnostics. Run ACQ board diagnostics. Connect the optical cable directly from the ACQ PCB to the RCOM. Try to make scans without rotation to see if scanning is possible without above error. Clean optical slipring. Inspect Rotor transmitter and Stator receivers, cabling/connections and incoming power to those modules (i.e. 12vdc from FRC). Repair/replace as necessary. Note: These couch & datapath errors seem to be intermittent, but occurring around the same time, thus one of these errors may be triggering the other error. ----------------------------------------------------------------------------- Work Order #: WO-03666488 Williams John Problem Description by Engineer : Problem description by engineer : System Stuck On Initializing Cannot Scan When was the first occurrence? Within... : 15-Mar-2021 08:00:00 How often is issue occurring? : Intermittent Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: System Stuck On Initializing Cannot Scan Patient Impact/Outcome: None User Impact: System Stuck On Initializing Cannot Scan Current Software Version: 4.1.7.25026 Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 4 | [14 empty] | 117710049 |  | 1011733 | #N/A | #N/A | Evaluation | [7 empty] | Others | [6 empty] | Optic cable | [5 empty] | Non-Refresh |  |  | Upgrade issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | #REF!
1125 |  |  | United States | Upson Regional Medical Center | 310115 | 44363 |  | System error during scanning due to faulty signal brush block and tapeswitch; Replaced signal brush block and tapeswitch | Work Order #: WO-03665959 Miller, John M The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Problem description by engineer : Customer states they had CT System errors last night during scanning. When was the first occurrence? Within... : 15-Jun-2021 How often is issue occurring? : Intermittent Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Treatment Expected and Unexpected results: System errors Patient Impact/Outcome: Delay in Scan User Impact: Delay in Scan Current Software Version: 4.1.10.2508 Follow Up Required : Error # and/or description of error : TmStamp 19548 ERR_TYPE NORMAL ErrNum S_STARTER_FAULT_NOT_UP_TO_SPEED vxwErrNo 3997698 File d_sn/stargate2/gantry/source/rhost/rhostSupport/source/RhostHVPSHssFaultISR.cpp Line 285 Param1 64 Param2 0 Info EalTime 6/16/2021 4:22:26 AM TmStamp 116484 ERR_TYPE NORMAL ErrNum S_CTRL_LINK_GHOST_GOT_CRC_FRAMING_ERROR vxwErrNo 0 File y_bld_main_Bld_sn/stargate2/gantry/source/gaf/Control_Link/src/CntlLinkTask.cpp Line 370 Param1 0 Param2 0 Info EalTime 6/15/2021 5:31:40 PM TmStamp 21769 ERR_TYPE NORMAL ErrNum S_HMC_TAPE_SWITCH_PRESSED vxwErrNo 0 File /gantry/source/galaxy/CouchMgrVertMotion/Src/StoppedToInitVertMoveFailTrans.cpp Line 406 Param1 0 Param2 0 Info EalTime 6/14/2021 3:11:31 PM TmStamp 27631 ERR_TYPE NORMAL ErrNum S_HMC_VELOCITY_ENCODER_FAILURE vxwErrNo 0 File S:\sy_bld_main_Bld_sn/stargate2/gantry/source/galaxy/hmc/src/HMCSMachine.cpp Line 1558 Param1 0 Param2 0 Info DeltaAbs:68544 Rel:68778 Vel:94943 6877 EalTime 6/14/2021 4:49:17 PM Type : Error Date : 6/15/2021 Time : 9:27:11 AM Event : 0 Source : PmsCT Category : None User : N/A Computer : HOST-4CE9110D8L Description: Message :XML parsing: line 135, character 57, illegal xml characterStack Trace : at System.Data.SqlClient.SqlConnection.OnError(SqlException exception, Boolean breakConnection, Action`1 wrapCloseInAction) at System.Data.SqlClient.TdsParser.ThrowExceptionAndWarning(TdsParserStateObject stateObj, Boolean callerHasConnectionLock, Boolean asyncClose) at System.Data.SqlClient.TdsParser.TryRun(RunBehavior runBehavior, SqlCommand cmdHandler, SqlDataReader dataStream, BulkCopySimpleResultSet bulkCopyHandler, TdsParserStateObject stateObj, Boolean& dataReady) at System.Data.SqlClient.TdsParser.Run(RunBehavior runBehavior, SqlCommand cmdHandler, SqlDataReader dataStream, BulkCopySimpleResultSet bulkCopyHandler, TdsParserStateObject stateObj) at System.Data.SqlClient.SqlBulkCopy.RunParser(BulkCopySimpleResultSet bulkCopyHandler) at System.Data.SqlClient.SqlBulkCopy.CopyBatchesAsyncContinuedOnSuccess(BulkCopySimpleResultSet internalResults, String updateBulkCommandText, CancellationToken cts, TaskCompletionSource`1 source) at System.Data.SqlClient.SqlBulkCopy.CopyBatchesAsyncContinued(BulkCopySimpleResultSet internalResults, String updateBulkCommandText, CancellationToken cts, TaskCompletionSource`1 source) at System.Data.SqlClient.SqlBulkCopy.CopyBatchesAsync(BulkCopySimpleResultSet internalResults, String updateBulkCommandText, CancellationToken cts, TaskCompletionSource`1 source) at System.Data.SqlClient.SqlBulkCopy.WriteToServerInternalRestContinuedAsync(BulkCopySimpleResultSet internalResults, CancellationToken cts, TaskCompletionSource`1 source) at System.Data.SqlClient.SqlBulkCopy.WriteToServerInternalRestAsync(CancellationToken cts, TaskCompletionSource`1 source) at System.Data.SqlClient.SqlBulkCopy.WriteToServerInternalAsync(CancellationToken ctoken) at System.Data.SqlClient.SqlBulkCopy.WriteRowSourceToServerAsync(Int32 columnCount, CancellationToken ctoken) at System.Data.SqlClient.SqlBulkCopy.WriteToServer(DataTable table, DataRowState rowState) at Philips.PmsCT.Logger.LogServer.DBInterface.AddLogBuffer(DataTable dataTable, Int32 targetId) Completed or Recommended Repair Action "Internal" : Suggest checking tape switches for proper operation and replace as needed. Test tape switches and check on Couch Control Board leds to determine which tape switch is sticking. Suggest cleaning the signal brush block, slip ring and/or replace brushes. Signal Brush, over time, may loose strength and not hold the slip ring properly. If cleaning helps and then the issue comes back something is contaminating the slip ring. The drive belt on one occasion and misaligned brush block on another, causing a accumulation on the slip rings. Also if there is a Black build-up on the rings, this can come from Covid-19 cleaning process. Make sure the Customer power off the Gantry prior to clean and leave the power off until dry. Recommendation: If this error is seen together with "S_CTRL_LINK_GHOST_GOT_CRC_FRAMING_ERROR" or S_CTRL_LINK_DMC_GOT_CRC_FRAMING_" it is caused by missing AP0, AP90 signals on the GHost.due to the Control Link problems. Mostly cleaning of signal slipring and brushes solves this problem.. On a BB Check / Replace the BI-directional optical control links parts. Check DMC for 33 on 7 segment display. 88 means the DMC is rebooting. Check Ghost for 1 (Normal state), anything else ghost is rebooting. Replace Ghost. Replace DMC Velocity encoder, check for proper couch movement, obstructions and proper lubrication. System errors, possible software reload. Internal Comments : Error # and/or description of error : TmStamp 19548 ERR_TYPE NORMAL ErrNum S_STARTER_FAULT_NOT_UP_TO_SPEED vxwErrNo 3997698 File d_sn/stargate2/gantry/source/rhost/rhostSupport/source/RhostHVPSHssFaultISR.cpp Line 285 Param1 64 Param2 0 Info EalTime 6/16/2021 4:22:26 AM TmStamp 116484 ERR_TYPE NORMAL ErrNum S_CTRL_LINK_GHOST_GOT_CRC_FRAMING_ERROR vxwErrNo 0 File y_bld_main_Bld_sn/stargate2/gantry/source/gaf/Control_Link/src/CntlLinkTask.cpp Line 370 Param1 0 Param2 0 Info EalTime 6/15/2021 5:31:40 PM TmStamp 21769 ERR_TYPE NORMAL ErrNum S_HMC_TAPE_SWITCH_PRESSED vxwErrNo 0 File /gantry/source/galaxy/CouchMgrVertMotion/Src/StoppedToInitVertMoveFailTrans.cpp Line 406 Param1 0 Param2 0 Info EalTime 6/14/2021 3:11:31 PM TmStamp 27631 ERR_TYPE NORMAL ErrNum S_HMC_VELOCITY_ENCODER_FAILURE vxwErrNo 0 File S:\sy_bld_main_Bld_sn/stargate2/gantry/source/galaxy/hmc/src/HMCSMachine.cpp Line 1558 Param1 0 Param2 0 Info DeltaAbs:68544 Rel:68778 Vel:94943 6877 EalTime 6/14/2021 4:49:17 PM Type : Error Date : 6/15/2021 Time : 9:27:11 AM Event : 0 Source : PmsCT Category : None User : N/A Computer : HOST-4CE9110D8L Description: Message :XML parsing: line 135, character 57, illegal xml characterStack Trace : at System.Data.SqlClient.SqlConnection.OnError(SqlException exception, Boolean breakConnection, Action`1 wrapCloseInAction) at System.Data.SqlClient.TdsParser.ThrowExceptionAndWarning(TdsParserStateObject stateObj, Boolean callerHasConnectionLock, Boolean asyncClose) at System.Data.SqlClient.TdsParser.TryRun(RunBehavior runBehavior, SqlCommand cmdHandler, SqlDataReader dataStream, BulkCopySimpleResultSet bulkCopyHandler, TdsParserStateObject stateObj, Boolean& dataReady) at System.Data.SqlClient.TdsParser.Run(RunBehavior runBehavior, SqlCommand cmdHandler, SqlDataReader dataStream, BulkCopySimpleResultSet bulkCopyHandler, TdsParserStateObject stateObj) at System.Data.SqlClient.SqlBulkCopy.RunParser(BulkCopySimpleResultSet bulkCopyHandler) at System.Data.SqlClient.SqlBulkCopy.CopyBatchesAsyncContinuedOnSuccess(BulkCopySimpleResultSet internalResults, String updateBulkCommandText, CancellationToken cts, TaskCompletionSource`1 source) at System.Data.SqlClient.SqlBulkCopy.CopyBatchesAsyncContinued(BulkCopySimpleResultSet internalResults, String updateBulkCommandText, CancellationToken cts, TaskCompletionSource`1 source) at System.Data.SqlClient.SqlBulkCopy.CopyBatchesAsync(BulkCopySimpleResultSet internalResults, String updateBulkCommandText, CancellationToken cts, TaskCompletionSource`1 source) at System.Data.SqlClient.SqlBulkCopy.WriteToServerInternalRestContinuedAsync(BulkCopySimpleResultSet internalResults, CancellationToken cts, TaskCompletionSource`1 source) at System.Data.SqlClient.SqlBulkCopy.WriteToServerInternalRestAsync(CancellationToken cts, TaskCompletionSource`1 source) at System.Data.SqlClient.SqlBulkCopy.WriteToServerInternalAsync(CancellationToken ctoken) at System.Data.SqlClient.SqlBulkCopy.WriteRowSourceToServerAsync(Int32 columnCount, CancellationToken ctoken) at System.Data.SqlClient.SqlBulkCopy.WriteToServer(DataTable table, DataRowState rowState) at Philips.PmsCT.Logger.LogServer.DBInterface.AddLogBuffer(DataTable dataTable, Int32 targetId) Completed or Recommended Repair Action "Internal" : Suggest checking tape switches for proper operation and replace as needed. Test tape switches and check on Couch Control Board leds to determine which tape switch is sticking. Suggest cleaning the signal brush block, slip ring and/or replace brushes. Signal Brush, over time, may loose strength and not hold the slip ring properly. If cleaning helps and then the issue comes back something is contaminating the slip ring. The drive belt on one occasion and misaligned brush block on another, causing a accumulation on the slip rings. Also if there is a Black build-up on the rings, this can come from Covid-19 cleaning process. Make sure the Customer power off the Gantry prior to clean and leave the power off until dry. Recommendation: If this error is seen together with "S_CTRL_LINK_GHOST_GOT_CRC_FRAMING_ERROR" or S_CTRL_LINK_DMC_GOT_CRC_FRAMING_" it is caused by missing AP0, AP90 signals on the GHost.due to the Control Link problems. Mostly cleaning of signal slipring and brushes solves this problem.. On a BB Check / Replace the BI-directional optical control links parts. Check DMC for 33 on 7 segment display. 88 means the DMC is rebooting. Check Ghost for 1 (Normal state), anything else ghost is rebooting. Replace Ghost. Replace DMC Velocity encoder, check for proper couch movement, obstructions and proper lubrication. System errors, possible software reload. Diagnostic performed by Engineer : Collect information from customer and view log files. ----------------------------------------------------------------------------- Work Order #: WO-03666180 18-Jun-2021 Lee John The system fully meets specification and returned to use Problem Description by Engineer : Problem description by engineer : Customer states they had CT System errors last night during scanning. When was the first occurrence? Within... : 15-Jun-2021 How often is issue occurring? : Intermittent Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Treatment Expected and Unexpected results: System errors Patient Impact/Outcome: Delay in Scan User Impact: Delay in Scan Current Software Version: 4.1.10.2508 Resolution : Installed brush block and tape switch Parts Consumed : 1.00000,459800623252,FRU, CBL, HANDLE TAPE SWITCH 1.00000,459800325761,Brush, Signal Block Assembly ----------------------------------------------------------------------------- | 3 | [14 empty] | 117710321 |  | 1011535 | #N/A | #N/A | Closed - Done | [7 empty] | Others | [6 empty] | Signal brush | [5 empty] | Non-Refresh |  |  | Upgrade issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | #REF!
1126 |  |  | United States | AdventHealth Zephyrhills | 300106 | 44362 |  | 0499CFXXX Short TC Failed | Work Order #: WO-03657944 Holeman, David The system fully meets specification and returned to use Diagnostic performed by Engineer : CT Tech reports short TC and Quick IQ Check scan failures. Tech states he has restarted Host and zeroed couch = same issue. Inspected logger which indicates multiple: S_MASTER_SEQUENCER_SUITE_DOOR_OPEN errors. Problem Description by Engineer : Problem description by engineer : Short TC - Scan Failure When was the first occurrence? Within... : 15-Jun-2021 07:00:00 How often is issue occurring? : Continuously Malfunction area : Others Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: Not provided Patient Impact/Outcome: None User Impact: Short TC - Scan Failure Current Software Version: 4.1.7.25030 Internal Comments : Error # and/or description of error : S_MASTER_SEQUENCER_SUITE_DOOR_OPEN Resolution : Advised Tech to inspect/exercise all gantry room doors and re-attempt scans. After inspected gantry room doors, Tech found one door ajar and secured the door. Tech is now able to scan successfully. ----------------------------------------------------------------------------- | 3 | [14 empty] | 117702890 |  | 1005779 | #N/A | #N/A | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Mis-use | [5 empty] | Non-Refresh |  |  | Upgrade issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | #REF!
1127 |  |  | United States | AdventHealth Zephyrhills | 300099 | 44358 |  | Recurring Intermittent Issue Transferring To PACS | Work Order #: WO-03641902 Holeman, David The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Problem description by engineer : Recurring Intermittent Issue Transferring To PACS When was the first occurrence? Within... : 03-May-2021 08:00:00 How often is issue occurring? : Intermittent Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: Recurring Intermittent Issue Transferring To PACS Patient Impact/Outcome: None User Impact: Delayed Transfers To PACS Current Software Version: 4.1.7.25030 Diagnostic performed by Engineer : CT Tech reports slow, then stopped transfer of images to PACS. Tech states around 0633 on 6/11/21 at around 175 images transferred, the transfer just stopped for about 15 minutes, then resumed sending and finished transfer. Tech states this occurs about once a day since Windows 10 upgrade (May 2021). Tech states a work-around to send faster is to access the ED local directory from the Main CT and send to PACS from the Main CT (an transfer is OK). Inspected last Host restart = 6/11/21 at 0509. Inspected LAN info (for PACS): name= PACS InetAddress=10.189.15.12 PortNumber=104 AETitle=WCHWFM1 Guided Tech to login as olevel and perform spooler reset. Communication with PFSE as an FYI. To monitor site and call back customer early next week to determine if issue persists. Internal Comments : Error # and/or description of error : Recurring Intermittent Issue Transferring To PACS Diagnostic performed by Engineer : Called customer back but Russ, CT Tech who placed call is not available. Spoke with another CT Tech who is unaware of any PACS transfer issues. Left message for Russ. RSE to callback on 6/15/21. Diagnostic performed by Engineer : RSE spoke with Russ, CT Tech who reports no further PACS transfer issues since spooler reset. RSE to follow-up with Tech on 6/16/21. Resolution : Called and spoke with Russ, CT Tech. No further issues with study transfers to PACS post-spooler reset. Closing call. ----------------------------------------------------------------------------- | 4 | [14 empty] | 117688366 |  | 997571 | #N/A | #N/A | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS Transfer | [5 empty] | Refresh |  |  | Upgrade issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | y
1128 |  |  | United States | Christ Hospital | 300205 | 44357 |  | 0711CFXXX Q : How to Block EKG in LanConfig. from going to PACS | Work Order #: WO-03637667 Carpenter, Kevin The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Philips CT Clin Apps Onsite [ Rhonda ] Device Use: Diag Expected and Unexpected Results: None User Impact: None Patient Impact: None Current Software Version: IDK Diagnostic performed by Engineer : Assisted CT Clin Apps Onsite [ Rhonda ] reconfig_ing LanConfig such that cardiac studies did not send EKC waveforms to PACS. Resolution : Assisted CT Clin Apps Onsite [ Rhonda ] reconfig_ing LanConfig such that cardiac studies did not send EKC waveforms to PACS. ----------------------------------------------------------------------------- | [15 empty] | 117684427 | 997378 | 997378 | #N/A | #N/A | Closed - Non-Complaint | [23 empty] | Upgrade issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1129 |  |  | United States | Westgate Hospital | 40001 | 44379 |  | Password Locked Out | Work Order #: WO-03787839 Carpenter, Kevin Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 2 |  | Password Locked Out | [12 empty] | 117828844 |  | 1075243 | TRUE | Opened | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Password | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #REF!
1130 |  |  | Chile | Scotiabank Chile | 40008 | 44386 |  | CL1675 - EQUIPMENT NEEDS TO BE CONFIGURED FOR SEVERAL TASKS IN RIS PACS. ON AUGUST 21ST AT 2:00 P.M. LOCAL TIME | Work Order #: WO-03826461 Valenzuela Cariaga, Luciano The system fully meets specification and returned to use. Follow Up Required : What are the next Troubleshooting steps : The customer requests a quote for work during non-working hours (Saturday August 21) for migration of its complete RIS PACS system and local image network. Error # and/or description of error : The customer requests a quote for work during non-working hours (Saturday August 21) for migration of its complete RIS PACS system and local image network. Completed or Recommended Repair Action "Internal" : The customer requests a quote for work during non-working hours (Saturday August 21) for migration of its complete RIS PACS system and local image network. Internal Comments : Error # and/or description of error : The customer requests a quote for work during non-working hours (Saturday August 21) for migration of its complete RIS PACS system and local image network. Completed or Recommended Repair Action "Internal" : The customer requests a quote for work during non-working hours (Saturday August 21) for migration of its complete RIS PACS system and local image network. Diagnostic performed by Engineer: The customer requests a quote for work during non-working hours (Saturday August 21) for migration of its complete RIS PACS system and local image network. Problem Description by Engineer : Problem description by engineer : The customer requests a quote for work during non-working hours (Saturday August 21) for migration of its complete RIS PACS system and local image network. Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected results: Equipment with new RIS PACS configured Patient Impact/Outcome: Not affected. User Impact: Migration of RIS PACS Current Software Version: 4.1.10 ----------------------------------------------------------------------------- | 3 |  | complete RIS PACS system and local image network. | [12 empty] | 117864190 |  | 1096609 | TRUE | Investigation | Investigation | [7 empty] | Service | [6 empty] | RIS PACS Config | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1131 |  |  | China | JX Province Children Hosp. | 40041 | 44389 |  | The sheave has moved,which made the sheave squeeze against the motor | The sheave has moved,which made the sheave squeeze against the motor |  |  | ECR-076490 Optimize torque for pulley bushing | [14 empty] | 1098246 | TRUE | Investigation | Investigation |  | ECR-076490 | [5 empty] | Engineering | [6 empty] | Sheave | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Sheave | Gantry-HW-Sheave | Gantry-HW-Sheave | ECR-076490 |  | Gantry-HW | Gantry-HW-Sheave | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | #N/A
1132 |  |  | China | 息烽县中医医院 | 345014 | 44385 |  | 02N/Ingenuity Core128/Elite China/Rebuild list | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer : 02N/Ingenuity Core128/Elite China Safety Questions : Case: 0117851128 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingsong Last updated by: Lu Mingsong 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: Rebuild list Troubleshooting Action: Remotely instruct the customer to Rebuild list. Repair Action: Remotely instruct the customer to Rebuild list. Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | Troubleshooting Action: Remotely instruct the customer to Rebuild list. Repair Action: Remotely instruct the customer to Rebuild list. Test/Inspection: Functional Check PASSED. Device returned to full functionality. | [12 empty] | 117851128 |  | 1092621 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Application | [6 empty] | Q&A | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1133 |  |  | China | 江西省儿童医院 | 40041 | 44384 | 27 | 02N/Ingenuity CT/Can not scan | T2 Activities : Problem Description: 客户问题概述： 不能扫描 客户问题详细描述或故障现象： 不能扫描，重建柜故障 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： S_ROTOR_SEEK_REF_TIMEOUT_ERROR S_MP_CTRL_DRIVER_SK_FAULT_ISQT_PN_CURRENT_OUT 找不到零位，电机过温 S_DMC_SHOT_COMM_COLLAPSE P1=4098 OVERRUN&CIRS OPTIC CHANNEL 远程诊断与测试结果： 1.检查机架旋转情况，检查刹车 2.确认重建柜情况，重启重建柜 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 1.检查机架旋转情况，检查刹车 2.确认重建柜情况，重启重建柜 Test Equipment: N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 07.07.2021 15:48:34 UTC Jiawei Li (86091272)"SUAT. Ready for Clinical Use local time" : "2021-07-07 22:00“ 服务内容：机架不能旋转，检查发现 服务结果： Internal Remarks : 07.07.2021 16:19:44 UTC Jiawei Li (86091272) $Q1:(Y) $Q2:(N) $Q3:(NA) $Q4:(NA) $Q5:(NA) $Q6:（Unable to scan，CRIS Error ) "RDI Accuracy:Yes" $Q7:(Unable to duplicate reported problem ，gantry cannot rotate) $Q8:( The squeeze against the motor, causing the motor to fail to operate ) $Q9:( Reposition the sheave , solved the problem.) $Q10:(Test and inspection data refer to PA. Device returned to full functionality.) $Q11:(Calibrated Tools: No tools used.) External Remarks : 07.07.2021 16:19:41 UTC SUP JCOUSER (86091272)服务内容：机架不能旋转，检查发现电机滑轮位置移动后与电机摩擦，导致电机无法正常运转。重新调整滑轮位置。 服务结果：维修后，设备恢复正常使用。 ----------------------------------------------------------------------------- | 2 |  | William/Alan investigate the issue is same as issue two years ago from Zhengjiang site, does torque added for motor related parts? Check if system produced before ECR cut in to OPS? Mengya TBIN to OPER. May 13, 2021 at 9:31 AM Partial Gantry SN5047是2020年5月27日完成quality 检验的， ECR-076490 Optimize torque for pulley bushing | [12 empty] | 117846204 |  | 1092188 | TRUE | Closed - Done | Closed - Done |  | ECR-076490 | [5 empty] | Engineering | [6 empty] | Sheave | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Sheave | Gantry-HW-Sheave | Gantry-HW-Sheave | ECR-076490 |  | Gantry-HW | Gantry-HW-Sheave | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | #N/A
1134 |  |  | China | Jiangsu Xuzhou Medical School Affiliated No. 3 Hospital | 40039 | 44383 |  | 02G/Ingenuity CT 728326/yellow screen | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer : 02G/Ingenuity CT 728326 T2 Activities:Problem Description: Summary of customer_s problems: Other types of error. Detailed description of Customer_s problems or malfunctions: Yellow screen. It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: Yellow screen. Remote troubleshoot and test result: Software bug. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Update the patch Safety Questions:Case: 0117838123 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action: confirmed with FSE Zhang Chuanxi, no need to add operation to follow up ----------------------------------------------------------------------------- | 3 |  | Yellow Screen | [12 empty] | 117838123 |  | 1083128 | TRUE | Investigation | Investigation |  | CTUDT00632520 | [5 empty] | Engineering | [6 empty] | Yellow Screen | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1135 |  |  | China | Jiangbei location, Southeast University Affiliated Zhongda Hospital | 345021 | 44383 |  | Wechat:02N/Ingenuity Core128/Elite China/image transfer failed | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: Instructed Customer Problem Reported by Customer : Wechat:02N/Ingenuity Core128/Elite China/image transfer failed Safety Questions:Case: 0117837355 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created by: Wennian Han Last updated by: Wennian Han 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: image transfer failed Troubleshooting Action: image transfer failed Repair Action: guide customer to check and found HOST net cable failed, customer replaced a new net cable, then can transfer Test/Inspection: Functional Check PASSED. The device returned to full functionality. Test Equipment: NA SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | Troubleshooting Action: image transfer failed Repair Action: guide customer to check and found HOST net cable failed, customer replaced a new net cable, then can transfer | [12 empty] | 117837355 |  | 1082495 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Net Cable | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1136 |  |  | China | Lindian County Hospital | 345038 | 44382 |  | 02N/Core128/rotation noise | BU Q&R Feedback: No Reason: Corrective Maintenance T2 Activities:Problem Description: Summary of customer_s problems: NOISE IN ROTATION Detailed description of Customer_s problems or malfunctions: NOISE IN ROTATION It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: High-speed rotation noise. Remote troubleshoot and test result: Opened the cover to check the source of the noise. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Opened the cover to check the source of the noise. Test Equipment: N/A SAP Transfer:Case Transferred to SAP Next Step Action: NTD: Service order opened through Ourmail.cn. Details: Contacted the field service engineer. No need to transfer. 13199082229 Doctor Teacher Li Lindian County Hospital 89753207 There is noise inside when rotating at high speed. Safety problem: Y N Safety Questions:Case: 0117835046 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Cao Last updated by: Liu Cao 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Problem Reported by Customer: 02N/Core128/ T2 Activities:Problem Description: The call didn_t go through. SAP Transfer:Dispatched for On site Support External Remarks: 07.07.2021 14:45:07 UTC SUP JCOUSER (86038210) service details: gantry noise. Replace the pulley and adjust the rotation part cable position. service result: device returns to normal working condition after service Internal Remarks : 07.07.2021 14:57:15 UTC Haijun Du (86038210) RDI Accuracy: no $Q7:(Confirmed reported problem. gantry noise.) $Q8: (the CT gantry have noise when do CT rotation by 0.5s, TENSIONER ASSY failed) $Q9: ( ok after replacement the TENSIONER ASSY and adjust rotor cable position.) $Q10: (Test and inspection data refer to: PA. Device returned to full functionality) $Q11 Calibrated Tools name： refer to PA ----------------------------------------------------------------------------- | 3 |  | $Q7: (Confirmed reported problem，gantry noise) $Q8: (the CT gantry have noise when do CT rotation by 0.5s，TENSIONER ASSY failed) $Q9: ( ok after replacement the TENSIONER ASSY and adjust rotor cable position.) | [12 empty] | 117835046 |  | 1079668 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | ECR-077012 | [5 empty] | Engineering | [6 empty] | Tensioner | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Belt | Gantry-HW-Belt | Gantry-HW-Belt | ECR-077012 | Belt | Gantry-HW | Gantry-HW-Belt | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1137 |  |  | China | 青海省海北州第二医院 | 345029 | 44382 |  | Wechat:02N/Ingenuity Core128/Elite China/check device | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : Wechat:02N/Ingenuity Core128/Elite China Safety Questions : Case: 0117833759 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wennian Han Last updated by: Wennian Han 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 其他类型故障。 客户问题详细描述或故障现象： 检查设备 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： NA 远程诊断与测试结果： 客户告知：CT机房漏水，主机架进水，现已烘干，请工程师前来检测开机 PRS TR状态： 不通，请工程师到场检查。 Repair Action: 建议现场维修步骤： 检查设备，确认有无其他问题 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 06.07.2021 09:41:50 UTC Yonghua Xu (86001334)科室：放射科CT室 联系人：郭荣寿 联系电话：15309708576 故障现象：CT机房漏水，主机架进水，现已烘干，需要工程师前来检测开机 。 解决方案/维修措施：CT机房漏水，主机架进水，现已烘干，需要工程师前来检测开机。检测设备机架、床 、Teal柜还有UPS干燥正常，通电测试设备预热扫描正常。 完成相应校准，完成备份。 服务后设备正常工作。 Internal Remarks : 06.07.2021 09:41:53 UTC Yonghua Xu (86001334) “RDI Accuracy: No” $Q7:Confirmation reported problem,The CT scanning room is leaking, and the gantry is flooded. It has been dried, and an engineer is required to check and start the machine. $Q8:The CT scanning room is leaking, and the gantry is flooded. It has been dried, and an engineer is required to check and start the machine. Check that the equipment gantry, couch, TEAL and UPS are dry properly. $Q9:Check that the equipment gantry, couch, TEAL and UPS are dry properly.Power on test equipment tube warming and scan normal. $Q10:Test and Inspection data refer to PA , Do a few patient, Observation of a certain period of time. The system is ok. Device returned to full functionality. $Calibrated Tools name (NA) S/N (NA) Next Calibration date (NA) ----------------------------------------------------------------------------- | 3 |  | CT机房漏水，主机架进水，现已烘干，需要工程师前来检测开机。检测设备机架、床 、Teal柜还有UPS干燥正常，通电测试设备预热扫描正常。 | [12 empty] | 117833759 |  | 1078731 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1138 |  |  | China | Guangxi Province Teng County People’s Hospital | 40036 | 44382 |  | application/Ingenuity CT 728326 | BU Q&R Feedback: No Reason: Application Support Problem Reported by Customer: Secondary training of high pressure injector Safety Questions:Case: 0117832536 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Zhang Jing Last updated by: Zhang Jing 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 5 |  | Application training | [12 empty] | 117832536 |  | 1077907 | TRUE | Evaluation | Evaluation | [7 empty] | Application | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1139 |  |  | Japan | 独立行政法人国立病院機構 | 341013 | 44386 |  | 01DM: RS: Scan did not proceed in a frozen state after Prep with CTA >> Manual start responded so inspection was possible >> No noticeable hard error | Problem Reported by Customer: 01CS: RS: Scan did not proceed in a frozen state after Prep with CTA >> Manual start responded so inspection was possible >> OneEMS internal remarks: --Troubleshooting Action-- Post-response results: Improved Failure location: 01Cosole What steps were taken during the diagnosis ?: In the subsequent inspection, there was no operation delay and the operation was good. No noticeable error was found in the logs. Judged as application operation delay. FSE compatible. --Repair Action-- --Subject-- 01DM: RS: Mr. Mamamoto: Scan did not proceed in a frozen state after Prep with CTA >> Manual start responded so inspection was possible >> No noticeable hard error --Problem Description-- Target device: CT Failure location: 01Cosole When did it first occur? (Within ~): 1 Hour How often do you have problems ?: Multiple times this Month At what point in the end-user workflow is the problem occurring ?: During inspection / measurement (During Acqustion / Scan) What happened recently ?: N.A. Does the customer create a specific system event report for the problem ?: N.A. SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Support Safety Questions: Case: 0117856986 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: RUI SATAKE Last updated by: RUI SATAKE 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm / alert capability, did it alarm / alert as it should have at the time the issue was discovered? N / A 4. Was this an out of box failure? N / A T2 Activities: Problem Description: Target device: CT Failure location: 01Cosole When did it first occur? (Within ~): 1 Hour How often do you have problems ?: Multiple times this Month At what point in the end-user workflow is the problem occurring ?: During inspection / measurement (During Acqustion / Scan) What happened recently ?: N.A. Does the customer create a specific system event report for the problem ?: N.A. Troubleshooting Action: Result after response: Improved Failure location: 01Cosole What steps were taken during the diagnosis ?: In the subsequent inspection, there was no operation delay and the operation was good. No noticeable error was found in the logs. Judged as application operation delay. FSE compatible. Repair Action: Service representative visit work. -------------------------------------------------- --------------------------- [Replacement parts] No replacement parts [Misorder Parts Report] No mistakes [Waiting information] No wait-and-see [TS Requests] No return parts [CT additional information] Q1 How did the inspection end ?: Not applicable because of Q1 No Q2 Presence or absence of incorrect findings (misdiagnosis, mistreatment) due to problems: None Q3 Presence or absence of artifacts: None Q4 Presence or absence of unintended unnecessary exposure: None [Repair certificate] Repair certificate is pasted [Other information sharing matters] 1. Treatment for patients infected with or suspected of being infected with the new coronavirus / Is it quarantined? Okay 2. Do I need to enter the treatment room or bed where the infected patient is? (Workplace and transfer (Including dynamic path) No Did you agree on how to complete the report with the person in charge of the equipment? YES Normal report completed on the day 09-Jul-2021 81100269 FSE Handover Tool1.3.4.32612.DB13 20210115 16.07.2021 01:54:08 UTC YUKO YAMAMOTO (81100269) [Request details] The left screen display of the console freezes during Tracker scanning (button operation is possible) [Phenomenon confirmation] I couldn_t confirm. [Details of work] Since the manual start button was functioning normally, It is probable that a situation similar to a PC freeze in the screen display area had occurred. As a response, we carried out PC-related work such as deleting unnecessary files on the HOST PC. [Other failures / treatments] None [Device status] Equipment can be used [Work status] done | 4 |  | Since the manual start button was functioning normally, It is probable that a situation similar to a PC freeze in the screen display area had occurred. As a response, we carried out PC-related work such as deleting unnecessary files on the HOST PC. | [12 empty] | 117856986 |  | 1096904 | TRUE | Closed - Done | Closed - Done | [7 empty] | Others | [6 empty] | Disk clean | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1140 |  |  | Poland | Rzadowa Agencja | 341008 | 44384 |  | commissioning and implementation into operation | Resolution Description: Switching on the power supply in the modular computed tomography laboratory. System calibration, Constancy and QuickIQ image quality test in accordance with the manufacturer_s specifications. Due to the lack of a configured computer network, the DICOM settings for external systems have not been entered. The apparatus is technically ready for use. | 4 |  | System acceptance test | [12 empty] | 117846595 |  | 1089244 | TRUE | Evaluation | Evaluation | [7 empty] | Service | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1141 |  |  | Turkey | T.C Saglik Bakanligi Bagcilar | 40010 | 44385 |  | Too much noise coming from the device | "Work Order #: WO-03811957 Soner Saglam The system partially meets specification and is returned to use with restrictions Resolution : Sound was estimated as DMS fan failure, part order was entered. Problem Description by Engineer : It has been reported that the gantry is very loud. Follow Up Required : Sound was estimated as DMS fan failure, parts order entered. | 3 |  | Sound was estimated as DMS fan failure, parts order entered. | [12 empty] | 117851395 |  | 1088133 | TRUE | Evaluation | Evaluation |  | CTUDT00632110 | [5 empty] | Engineering | [6 empty] | DMS fan | [5 empty] | Refresh |  |  | System issue | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | DMS Module | DMS-HW | DMS-HW-DMS Module | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1142 |  |  | China | Jilin Province Zhaonan City TCM Hospital | 345006 | 44383 |  | 02n/Ingenuity Core128/Elite China/need parts | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 02n/Ingenuity Core128/Elite China Safety Questions:Case: 0117839325 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Lili Last updated by: Lu Lili 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action: FSE Hua Tian 018946514633 call in nemo SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | Gao, Mia |  | 44383 | [12 empty] | 117839325 |  | 1083964 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | 44384.7299 | [6 empty] | 44383 | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1143 |  |  | Turkey | Republic of Turkey Ministry of Health, Bagcilar | 40010 | 44383 |  | An error message was displayed, device is up. | Work Order #: WO-03797735 Soner Saglam The system partly meets specification and is returned to use with restrictions Resolution : Log was analyzed. S_MPLTRANSPORTREC_DEFRAG_TOUCAN_ERROR error was found. Slipring and signal brush cleaning is recommended. Problem Description by Engineer : A gantry error was displayed. Follow Up Required : Log was analyzed. S_MPLTRANSPORTREC_DEFRAG_TOUCAN_ERROR error was found. Slipring and signal brush cleaning is recommended. ----------------------------------------------------------------------------- | 3 |  | Log has been examined. S_MPLTRANSPORTREC_DEFRAG_TOUCAN_ERROR is seen. Slipring and signal brush cleaning and control is recommended. | [12 empty] | 117838741 |  | 1081861 | TRUE | Evaluation | Evaluation |  | NotEngineering-SlipringClean | [5 empty] | Engineering | [6 empty] | Slip ring clean | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | SlipRingClean | Gantry-HW-SlipRingClean | Gantry-HW-SlipRingClean | NotEngineering-SlipringClean | SlipRingClean | Gantry-HW | Gantry-HW-SlipRingClean | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1144 |  |  | United Kingdom | STOCKPORT NHS FOUNDATION TRUST | 40016 | 44386 |  | Interventional monitor and table upgrade | Work Order #: WO-03822493 09-Jul-2021 Martin Greenbank The system fully meets specification and returned to use Regulatory Questions : N/A Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : INST - Installation started but not yet completed Installed new extended couch and joystick option. Need to plan another day to complete the ceiling mounted monitor. System tested ok Follow Up Required : replan 2 men for 6h to complete ceiling mounted monitor when parts available ----------------------------------------------------------------------------- Work Order #: WO-03822496 09-Jul-2021 Warren Thorpe The system fully meets specification and returned to use Regulatory Questions : N/A Problem Description by Engineer : CT table needed upgraded from a standard couch to an extended couch. Biopsy monitor and extras needed to be installed. Resolution : New extended table installed successfully with additional biopsy features. Ceiling mounted monitor arm still needs to be fitted but all cables ran and ready to attach. Follow Up Required : Install ceiling mounted monitor arm. Please plan Martin G and myself to install. ----------------------------------------------------------------------------- Product/Part - MONITOR SUPPORT Part number - 453566207542 | 5 |  | New extended table installed successfully with additional biopsy features. Ceiling mounted monitor arm still needs to be fitted but all cables ran and ready to attach. Follow Up Required : Install ceiling mounted monitor arm. Please plan Martin G and myself to install. | [12 empty] | 117861311 |  | 1095469 | TRUE | Evaluation | Evaluation | [7 empty] | Service | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1145 |  |  | United States | Unity Health Systems | 40037 | 44388 |  | 0609CDN84 Clicking noise right side gantry when rotor moving | Work Order #: WO-03828644 Maier, Donnie The system partly meets specification and is returned to use with restrictions Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: Diagnostic Expected and Unexpected Results: Clicking noise from right side of gantry User Impact: No Patient Impact: No Current Software Version: 4.1.10.25210 Diagnostic performed by Engineer : Talked cu thru a complete power cycle of the system as it had been over 10 days for the gantry. After power cycle had cu run a mock scan and he can still hear the noise. No normal errors in ULV. Cu is just going to use the other CT Scanner and wants HTM 7/12 AM Dispatch. I provided case number to the CT Tech. Follow Up Required : Identify the source of the noise and correct. Tie wrap, etc. ----------------------------------------------------------------------------- | 3 |  | Clicking noise from right side of gantry FSE to be follow up Follow Up Required : Identify the source of the noise and correct. Tie wrap, etc. | [12 empty] | 117865876 |  | 1098098 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-GantryNoise | [5 empty] | Engineering | [6 empty] | Gantry Noise | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Noise | Gantry-HW-Noise | Gantry-HW-Noise | TBD-GantryNoise | Noise | Gantry-HW | Gantry-HW-Noise | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1146 |  |  | United States | Odessa Regional Medical Center | 40029 | 44383 |  | Acquiring Dose Info | Work Order #: WO-03799935 Mujica, David The system fully meets specification and returned to use Problem description by Clinical Support Specialist : Customer Function/Role: Technologist Device Use: In use Expected and Unexpected Results: Get dose info for Peds User Impact: none Patient Impact: none Current Software Version: 4.1 Diagnostic performed by Engineer : Maya wanted to report all Pediatric dose information from the last year off of their CT system. It was explained to Maya that unless that patient information is on the directory then that dose information can be acquired from the CT system. Maya indicated that they delete CT exams frequently and this would not be a method to find the dose information for all Pediatrics from the last year. It was suggested to look in their PACS system as it would have the dose information stored. Maya took a look into PACS and successfully found the dose information. Maya had no further questions or need for clinical support. No rescan, no loss of data, no reinjection. ----------------------------------------------------------------------------- | E, Sravanthi |  | 44383 | [12 empty] | 117840721 |  | 1080377 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | 44383.9146 | [6 empty] | 44383 | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1147 |  |  | United States | Odessa Regional Medical Center | 40029 | 44386 |  | Curved MPR | Work Order #: WO-03825534 Curry, Dyan The system fully meets specification and returned to use Problem description by Clinical Support Specialist : Customer Function/ Role: technologist Device Use: in use name Expected and Unexpected Results: curved MPRs of cspine User Impact: none Patient Impact: none Current Software version: 4.1.7 Diagnostic performed by Engineer : helped customer created MPR curved images in planar , CT Viewer. Saved batches of each plane. ----------------------------------------------------------------------------- | 5 |  | helped customer created MPR curved images in planar , CT Viewer. | [12 empty] | 117863525 |  | 1095973 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Application | [6 empty] | Q&A | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1148 |  |  | United States | McLaren Macomb | 40021 | 44384 |  | customer requesting support | Work Order #: WO-03810117 Hirchak, Michele The system fully meets specification and returned to use Problem description by Clinical Support Specialist : Customer Function/Role: TECH Device Use: In Use Expected and Unexpected Results: XR-29 certificate. User Impact: None Patient Impact: None Current Software Version: 4.1.7 Diagnostic performed by Engineer : Customer requesting XR-29 Certificate. Emailed document and uploaded to TEAMS. MBH Resolution : Called Allison to make sure she rec_d the certificate. OK to close. MBH ----------------------------------------------------------------------------- | 5 |  | XR-29 certificate | [12 empty] | 117849487 |  | 1087150 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Customer | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1149 |  |  | United States | Philips Medical Capital | 40028 | 44384 |  | 4.1 Accession number issues | Work Order #: WO-03809653 Mujica, David The system fully meets specification and returned to use Problem description by Clinical Support Specialist : Customer Function/Role: Technologist Device Use: In use Expected and Unexpected Results: Accession numbers not associating User Impact: none Patient Impact: none Current Software Version: 4.1 Diagnostic performed by Engineer : Iris indicated that the accession numbers were not associating to the CT exams performed. The exam cards were explored that were used and after investigation it was noticed that the user was picking all the exams and then launching each protocol which would not associate the proper accession numbers to each exam ordered. Iris was told to only launch the CT studies ordered per scan. Iris will comply and do this going forward. Iris had no further questions or need for clinical support at this time. No rescan, no loss of data, no reinjection per Iris. ----------------------------------------------------------------------------- | 5 |  | user was picking all the exams and then launching each protocol which would not associate the proper accession numbers to each exam ordered. Iris was told to only launch the CT studies ordered per scan. Iris will comply and do this going forward. | [12 empty] | 117849118 |  | 1086689 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Application | [6 empty] | Mis-use | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1150 |  |  | Chile | Scotiabank Chile | 40008 | 44389 |  | CL1675-BURNING SMELL | Integration:Confirmation of order 53875179 saved ----------------------------------------------------------------------------- Work Order #: WO-03834204 González, Carlos The system meets the specification for the performed service and is returned to use. Problem Description by Engineer : Problem description by engineer : Client says that there was a burning smell in the Tomography room. When was the first occurrence? Within... : 12-Jul-2021 12:00:00 How often is issue occurring? : First occurrence Malfunction area : Others Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected results: No further information is available as per the customer Patient Impact/Outcome: No further information is available as per the customer. User Impact: No further information is available as per the customer Current Software Version: 4 Diagnostic Performed by Engineer: User reported a burning smell in Tomography room. Field service engineer Juan Osorio from Chile went previously to check the equipment and no damage was found inside. A preventive maintenance was even done on 08/07 and no damage was found either. Possibly the smell comes from the air conditioning system. Equipment operational. Internal Comments : Error # and/or description of error : NA Repair Action: NA Resolution: Equipment is OK. ----------------------------------------------------------------------------- | 4 |  | User reported a burning smell in Tomography room. Field service engineer Juan Osorio from Chile went previously to check the equipment and no damage was found inside. | [12 empty] | 117871717 |  | 1100496 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | False alarm | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1151 |  |  | China | 北京市中西医结合医院 | 40032 | 44395 |  | 02N/Ingenuity CT 728326/connect PACS | BU Q&R Feedback: No SAP Customer Symptom Code: CENV Customer Symptom Code: Customer Environmental Concern Reason: Corrective Maintenance Engineer Failure Code: Environmental Cause Repair Activity Code: Configuration Change Problem Reported by Customer : connect PACS Safety Questions : Case: 0117907046 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Weng Jinming Last updated by: Weng Jinming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Troubleshooting Action: next Thursday ，connect PACS Repair Action: next Thursday ，connect PACS Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: N/A ----------------------------------------------------------------------------- | 5 |  | PACS Config | [12 empty] | 117907046 |  | 1125199 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Customer | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1152 |  |  | China | 浙江省乐清第一人民医院 | 40054 | 44395 |  | 02N/Ingenuity CT 728326/Cirs fault | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity CT 728326 Safety Questions : Case: 0117906849 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Zhou Xiaohong Last updated by: Zhou Xiaohong 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 主机检查不能连接重建计算机，重建柜需要连接USB键盘，点回车 设备本身技术问题： 是。 Troubleshooting Action: 远程诊断与测试结果： 怀疑重建计算机软件问题 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议工程师到场检查，重新安装重建计算机软件 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : 已派工黄逸 15958725282 ----------------------------------------------------------------------------- | 3 |  | $Q7: (Confirmed reported problem&IRS disconnected) $Q8: (check errorlog found teermal sensor not programmed, all fans works normally) $Q9: (update BIOS version) | [12 empty] | 117906849 |  | 1125038 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | ShouldbeService | [5 empty] | Engineering | [6 empty] | CIRS | [5 empty] | Refresh |  |  | System issue | System | NotEngineering | ShouldbeService | NotEngineering-ShouldbeService | NotEngineering-ShouldbeService | ShouldbeService | ShouldbeService | NotEngineering | NotEngineering-ShouldbeService | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1153 |  |  | China | 江苏大学附属医院 | 345023 | 44393 |  | Wechat:02N/Ingenuity Core128/Elite China/can not scan | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer : can not scan Safety Questions : Case: 0117895563 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingsong Last updated by: Lu Mingsong 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: can not scan Troubleshooting Action: Remotely instruct the customer to restart the system. Repair Action: Remotely instruct the customer to restart the system. Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | Troubleshooting Action: Remotely instruct the customer to restart the system. Repair Action: Remotely instruct the customer to restart the system. Test/Inspection: Functional Check PASSED. Device returned to full functionality. | [12 empty] | 117895563 |  | 1122766 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Reboot | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1154 |  |  | China | 浙江省乐清第一人民医院 | 40054 | 44393 |  | 02N/Ingenuity CT 728326/account is lock | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: Instructed Customer Problem Reported by Customer : 02N/Ingenuity CT 728326 Safety Questions : Case: 0117893670 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wennian Han Last updated by: Wennian Han 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: account is lock Troubleshooting Action: enter wrong password, account is lock Repair Action: wait for some time, then enter wright password, then can enter in account Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: NA SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | Repair Action: wait for some time, then enter wright password, then can enter in account Test/Inspection: Functional Check PASSED. Device returned to full functionality. | [12 empty] | 117893670 |  | 1121897 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Password | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1155 |  |  | China | 福建省医科大学附属协和医院 | 345010 | 44392 |  | 02N/Core128//CT account locked | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Calibration/Adjustment Problem Reported by Customer : 02N/Core128/ Next Step Action : NTD：协助接线同事开单 内容：ct账号被锁定 需要工程师账号解锁18650369579工程师来电要转RSE，中断了 T2 Activities : Problem Description: CT 账户被锁定 Troubleshooting Action: 需要重置密码 Repair Action: 到场重置密码，教育客户保存好密码和使用IT账户重置密码 Test/Inspection: N/A Test Equipment: N/A SAP Transfer : Case Transferred to SAP Safety Questions : Case: 0117888014 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Gaoge Last updated by: Wang Gaoge 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Dispatched for Onsite Support Internal Remarks : 15.07.2021 11:02:10 UTC Feng Li (86036224) RDI Accuracy: Yes $Q7: (Confirmed reported problem.The Ingenuity Core 128 CT user account is locked.) $Q8: (On site use Philips_service account to un-lock the user account and modify the password.) $Q9: (On site use Philips_service account to un-lock the user account and modify the password.) $Q10: (Test and inspection data refer to: PA.+Device returned to full functionality.) $Q11: Calibrated Tools name (FLUKE 287C Multimeter) S/N (16490128) NextCalibration date (10/20/2021) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date: NA ^Current Tube usage: NA External Remarks : 15.07.2021 11:02:07 UTC SUP JCOUSER (86036224)服务内容：CT医生用户账户被锁了。到场使用工程师账户重置医生账户，并修改密码。 服务结果：维修后，设备恢复正常工作。 ----------------------------------------------------------------------------- | 3 |  | Repair Action: 到场重置密码，教育客户保存好密码和使用IT账户重置密码 | [12 empty] | 117888014 |  | 1118020 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | customer | [6 empty] | Password | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1156 |  |  | China | 内蒙古自治区第四医院 | 40023 | 44391 |  | 02N/Ingenuity CT 728326/ECG loss | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : connect fault Safety Questions : Case: 0117881084 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Weng Jinming Last updated by: Weng Jinming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 没有心电信号 客户问题详细描述或故障现象： 拔插USB，没有连接成功 设备本身技术问题： 是。 Troubleshooting Action: 远程诊断与测试结果： 怀疑USB接口问题 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议工程师到场检查USB接口，教客户使用飞利浦微信报修 SAP Transfer : Dispatched for Onsite Support SAP Transfer : Case Transferred to SAP Next Step Action : PPD IBOX2 USB EXTENDER Kit 459801413791 1 无货，替代号咨询RSE不能用，已安排订货，FSE：苏瑞杰-86-13304719335 Next Step Action : PPD 459801731761 替代号已确认可以出货，已安排24H到医院 External Remarks : 16.07.2021 14:19:07 UTC Ruijie Su (86039043)服务结果： [更换I BOX 板和Ibox线，设备恢复正常工作。] 其它重要事项：无 Internal Remarks : 16.07.2021 14:19:10 UTC Ruijie Su (86039043) $Q7: Confirmed reported problem,connect fault $Q8: On site checked ECG connect fault . Need replace Ibox and cable $Q9: replace I-BOX and cable system is ok! $Q10:Test and inspection data refer to: PA. Device returned to not full functionality. $ $Q11: Calibrated Toolsname ( Multimeter Fluke 287C) S/N (31060028) Next Calibrationdate (29/10/2021) 2( Torque Wrench Kit ) S/N E201602090 Next Calibration date10/09 /2021 For Mainland China CV / XR, CT engineers only. Please fill in: ^ Current Tube usage: scan second 53865 External Remarks : 16.07.2021 14:26:37 UTC Zhen Liang (86039057)服务结果： [更换I BOX 板和Ibox线，设备恢复正常工作。] 其它重要事项：无 Internal Remarks : 16.07.2021 14:26:52 UTC Zhen Liang (86039057)$Q7: Confirmed reported problem,connect fault $Q8: On site checked ECG connect fault . Need replace Ibox and cable $Q9: replace I-BOX and cable system is ok! $Q10:Test and inspection data refer to: PA. Device returned to not full functionality. $ $Q11: Calibrated Toolsname ( Multimeter Fluke 287C) S/N (31060028) Next Calibrationdate (29/10/2021) 2( Torque Wrench Kit ) S/N E201602090 Next Calibration date10/09 /2021 For Mainland China CV / XR, CT engineers only. Please fill in: ^ Current Tube usage: scan second 53865 Next Step Action : FSE：苏瑞杰-13304719335 备件DOA申请reopen，40-OP已派梁工 Next Step Action : 梁震 15849133876来电要求转单苏瑞杰 13304719335 Next Step Action : 已转派苏瑞杰 13304719335 External Remarks : 18.07.2021 22:39:41 UTC Ruijie Su (86039043))服务结果： [兼差I BOX 板和Ibox线，设备恢复正常工作。] 其它重要事项：无 Internal Remarks : 18.07.2021 22:39:45 UTC Ruijie Su (86039043)$Q7: Confirmed reported problem,connect fault $Q8: On site checked ECG connect fault . Need adjust Ibox and cable $Q9: adjust I-BOX and cable system is ok! $Q10:Test and inspection data refer to: PA. Device returned to not full functionality. $ $Q11: Calibrated Toolsname ( Multimeter Fluke 287C) S/N (31060028) Next Calibrationdate (29/10/2021) 2( Torque Wrench Kit ) S/N E201602090 Next Calibration date10/09 /2021 For Mainland China CV / XR, CT engineers only. Please fill in: ^ Current Tube usage: scan second 53865 ----------------------------------------------------------------------------- | 5 |  | 更换I BOX 板和Ibox线后IBOX仍然有断开报错 FSE说没有其他设备连接到IBOX， 也没有呼吸门控。 | [12 empty] | 117881084 |  | 1113582 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | TBD-NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Ibox | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | TBD-NoEnoughInfo | Gantry-HW-TBD-NoEnoughInfo | Gantry-HW-TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo | Gantry-HW | Gantry-HW-TBD-NoEnoughInfo | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1157 |  |  | China | 苏州广慈肿瘤医院有限公司 | 345019 | 44391 |  | 02N/Core128/cann_t manual scan | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Core128/ Next Step Action : NTD：协助接线同事开单 内容：扫描的时候检测程序列无法手动触发，不能保障成功率 Safety Questions : Case: 0117880809 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Weng Jinming Last updated by: Weng Jinming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 偶发增强扫描，不能手动触发 客户问题详细描述或故障现象： 不能提供具体时间 设备本身技术问题： 是。 Troubleshooting Action: 远程诊断与测试结果： 很难判断问题，联系工程师徐浩天，到场检查 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议工程师到场检查log，判断问题，教客户使用飞利浦微信报修 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : Xu Haotian External Remarks : 16.07.2021 10:28:44 UTC Haotian Xu (86090280) 服务内容：做增强手动触发偶发无法使用，检查log报错和软件系统，配合培训老师 与客户沟通 服务结果：服务后，设备恢复正常 ----------------------------------------------------------------------------- | 5 |  | $Q7 (Confirmed reported problem , unable to manually scan ) $Q8: (Checked the log and the protocols,and found the problem is a software bug) $Q9: (connect to the trainning teacher and costomers) $Q10: (Test and inspection scan refer to PA . device returned to full functionality) | [12 empty] | 117880809 |  | 1113397 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | CTA | [5 empty] | Refresh |  |  | System issue | System | Console-SW | CTA | Console-SW-CTA | Console-SW-CTA |  | CTA | Console-SW | Console-SW-CTA | No | 0 | 0 | 0 | 0 | 0 | No | y | y | 0 | 0 | 0 | n | y | y | y | y | #N/A
1158 |  |  | China | Shaanxi Province Xianyang Central Hospital (Xianyang City No. 2 People_s Hospital) | 40050 | 44390 |  | 02N/Ingenuity CT/gantry noise | BU Q&R Feedback: No Reason: Corrective Maintenance Next Step Action: NTD: assist the operator to open the service order Shaanxi Province Xianyang City Central Hospital53636103 Teacher Ma 13809141954 gantry has abnormal noise, yes Problem Reported by Customer: 05R/Brilliance CT 64/ Safety Questions:Case: 0117874106 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Bo Last updated by: Wang Bo 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer_s problems: device has abnormal smell. Detailed description of Customer_s problems or malfunctions: The gantry is making an abnormal sound. It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: The log has no errors. Remote troubleshoot and test result: Restart but didn_t work. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: It is recommended to check the noise source, utility transformer, rotor brake part, rotating belt and pressure wheel. We recommend spare parts 12NC and names: N/A SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Next Step Action:GUO YUJIE External Remarks: 13.07.2021 15:20:53 UTC SUP JCOUSER (86038990) Service details: Gantry noise. Check to see if It is new machine belt and motor friction noise. Add talcum powder to restore. service result: device returns to normal working condition after service ----------------------------------------------------------------------------- | 3 |  | 服务内容：机架异响，检查是新机器皮带和电机摩擦声，加滑石粉恢复。 服务结果：服务后，设备正常工作。 | [12 empty] | 117874106 |  | 1108663 | TRUE | Evaluation | Evaluation |  | TBD-GantryNoise | [5 empty] | Engineering | [6 empty] | Gantry Noise | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Noise | Gantry-HW-Noise | Gantry-HW-Noise | TBD-GantryNoise | Noise | Gantry-HW | Gantry-HW-Noise | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1159 |  |  | China | Jinan City No. 4 People_s Hospital | 40038 | 44386 |  | 02n/Ingenuity CT 728326/Check the internal cable machine | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: Check the internal cable machine Next Step Action: FSE chentao called to request for repair Safety Questions:Case: 0117858853 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Lili Last updated by: Lu Lili 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer_s problems: check device. Detailed description of Customer_s problems or malfunctions: check device. It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: N/A Remote troubleshoot and test result: N/A PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: The equipment warranty is about to expire. The customer asked FSE to go to the hospital to check the equipment We recommend spare parts 12NC and names: N/A Memo of Customer request: None SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Next Step Action : Chen Tao External Remarks:09.07.2021 12:52:41 UTC Tao Chen (86038373) Error description: Check the internal cable of the rack Service content: Check and secure the internal cables of the rack Service outcome: Device returns to normal working condition after servicing. Internal Remarks : 09.07.2021 13:18:22 UTC Tao Chen (86038373) “RDI Accuracy: YES" $Q7: (Confirmed reported problem, check cables in the gantry) $Q8: ( Check cables of gantry rotor and stator according to customer requirement) $Q9: ( Fasten cables for the gantry) $Q10: (Test and inspection data refer to: PA. Device returned to full functionality) $Q11: Calibrated Tools name ( N/A) ^Current Tube installed date: N/A ^Current Tube usage: N/A ----------------------------------------------------------------------------- | 5 |  | $Q7: (Confirmed reported problem, check cables in the gantry) $Q8: ( Check cables of gantry rotor and stator according to customer requirement) $Q9: ( Fasten cables for the gantry) | [12 empty] | 117858853 |  | 1105707 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Others | [6 empty] | Gantry Cable | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1160 |  |  | China | Dongguan Donghua Hospital Co., Ltd. | 40059 | 44389 |  | WeChat:02N/Ingenuity CT 728326/adjust CIRS time | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer:adjust CIRS time Safety Questions:Case: 0117867602 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Guoping Last updated by: Liu Guoping 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support External Remarks: 12.07.2021 14:27:31 UTC Zhenhua Liu (86034125) Error description: CIRS time is not in synch with host Service details: synchronizing the time of the CIRS with host Service result: After service, the machine is back to normal. Other important issues: None Internal Remarks : 12.07.2021 14:27:33 UTC Zhenhua Liu (86034125) "RDI Accuracy:Yes" $Q1: () $Q2: () $Q3: () $Q4: () $Q7: (Confirmation & Description of problem) Confirmed reported problem, the time of CIRS is not sync with host $Q8: (Troubleshooting & Results) time of CIRS is not sync with host $Q9: (Repair steps) sync the time of cirs and host $Q10: (Test and inspection data refer to PA . Device returned to full functionality ) $Q11: Calibrated Tools name ( refer to PA ) S/N ( ) Next Calibration date ( refer to PA ) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date: N/A ^Current Tube usage: N/A ----------------------------------------------------------------------------- | 5 |  | $Q7: (Confirmation & Description of problem) Confirmed reported problem, the time of CIRS is not sync with host $Q8: (Troubleshooting & Results) time of CIRS is not sync with host $Q9: (Repair steps) sync the time of cirs and host $Q10: (Test and inspection data refer to PA . Device returned to full functionality ) | [12 empty] | 117867602 |  | 1102201 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Service | [6 empty] | Time Sync | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1161 |  |  | China | Xifeng County Traditional Chinese Medicine Hospital | 345014 | 44389 |  | Wechat:02N/Ingenuity Core128/Elite China/ACQ error | BU Q&R Feedback: No Reason: Corrective Maintenance T2 Activities:Problem Description: Summary of customer_s problems: Scanning failed. Detailed description of Customer_s problems or malfunctions: Report ACQ error It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: DatapathCheck(127): Excessive acquisition Holdoff or stalled buffered stream m_seriesUid = 0x4D timeSinceLastReadingSec = 1.31259 bufStreamWorkquePtr = 0 acquisitorHoldoffMs = 0 m_maxHoldoffTimeSec = 1.28 SafetyCheck(873): 0x4d SafetyCheckThread: HaltAcquisition WorkItem index = 33 SafetyCheck.cpp(469): Error #12 - Unknown error 0x4d SafetyCheckThread: Datapath error. Data Path Diagnostics End Series Series UID 0x4d Real-time stream Number of channels: 4 Physical channel order: 0x4321 Unexpected gantry Confirm Resend: 0 (1 = error) Reading resends: 0 CRC error count Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 Loss of Signal Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 Line Fault Indicator Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 Resend bit-mapped angle Chan A: 0x0000000000000000000000000000000000000000000000000000000000000000 Chan B: 0x0000000000000000000000000000000000000000000000000000000000000000 Chan C: 0x0000000000000000000000000000000000000000000000000000000000000000 Chan D: 0x0000000000000000000000000000000000000000000000000000000000000000 DMS recoverable error count: 0 Duplicate reading count: 0 On-board overflow retry count: 0 On-board overflow max retry period: 0 ms On-board overflow total retry time: 0 ms Remote troubleshoot and test result: Recon removed ACQ_OK signal to stop the scan. There are too many reasons it can either be the Reconstructor or Gantry causing this error. We are still improving the AEL in which we can tackle this error. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: "1. Check for CIRS errors and follow their recommendations. 2. Run Visual Diagnostics/Data Path test. 3. Check during a scan that Enable Acq and Acq Ok Led on the GHOST goes on. If not or any doubt because of intermittent error, run the Acq test with the electrical loopback to verify the Acquisitor PCB (verification of signals comm Ok, enable acq and acq ok). Check / Reseat the internal CRC cable connected to the Acquisitor J4 connector were the cable with D15 connector is mounted. Also check the cable from the CRC rack to the Ghost. 3. Check the Raid at the server 1 for health indication and/or errors on one of the disks. 4. Connect the optical fiber cable directly between the RCOM and ACQ when ACQ_OK is now visible at the RCOM (no S_MASTER_SEQUENCER_ACQ_NOT_OK error in log) you know that there is an issue at the optical slip ring. 5. Check the CIRS.log if you find ""bad dynamic rotor position"" and system has no resends, datapath test runs without errors. The problem was solved after DMC-16 replacement, probably bad control link data caused by the DMC-16 pcb. 6. Check the CIRS. log when you find ""bad table motion"" indicating couch pulses do not match rotor pulses causing the acq error. Check Couch and /or Angulation system. 7. If encountered without a CIRS. error reported and no surrounding Normal errors reported in logger, this indicates a problem with the ACQ_OK signal coming from the ACQ card to the Ghost. Look at the ACQ_OK led on the Ghost and see if it is lit. If it is lit and there are no CTRL link errors then it is a problem with the Ghost. If CTRL link error the check slipring signal brushes and DMC, If it is out then the Ghost never received the ACQ_OK signal from the ACQ card. Check the connections from the ACQ Card to the Ghost, verify cable integrity and replace if bad, replace ACQ card, replace Ghost. Experience: 1. Check the Raid at the server 1 for errors on one of the disks. 2. Check the CIRS.log for ""ERROR - ReadingType changed during shot"" and / or ""ERROR - Number of slices changed during shot"". During a shot the dms is sending temperature data which interrupts the shot. . In the Raw data it shows the X-Ray signal jump to 0 and back to 1 - causing the H.W to send Temperature reading again. So it’s a problem with the “X-Ray On” signal and the parts involved are: Dmc, Rcom, Rhost to DMC Cable 4535-670-25791, and Dmc to Rcom Cable. Replacing DMC and RCOM did not solve, but replacing the RHOST solved the problem. 3. On a TDMS system, got high resends and at times the order of the channels will changed by itself from the normal 4321 (in the CIRS log ""Unexpected"" messages) to another order. Visual Diagnostic reported that an index pulse was received by the RCOM but was not by the Tx box. Found that the wire from pin 12 (AUX_0_INDEX_N) that connects to DMC16 J23 in the Aux. Shut Down Adapter cable 453-567-525791 was broken. This cable adapts to the long cable that connects the DMC to the Tx box. 4. On FRC TX module power connection was problematic. Socket seems seated but it was not working. There were 15V while not connected but it was dropping to 3V while connected. Connection socket seated and sealed fixed the problem. 5. ACQ _NOT _OK with parm 1=0 happened in high speed and High MA scan, no other Xray errors but, in CIRS Errorlog a few lines before the ACQ error, you can see :AcqWorkItem.cpp216ERROR - Max number of consecutive Readings with tube arcs (HV drops) exceeded, BufferSel = 1, ReadingOffset = 0x139bc00. Checked in the Can log and found a lot of soft arcs- Tube was replaced and didn_t fix this problem. The RHOST fixed it. 6. Detector module 10 was swapped with module 2 location to see if ring would move to outside. Tried to perform air calibrations on system and received this error. Error went away when module 10 was replaced with new one and performing calibration procedure for module replacement. Also moved module 2 back to original location. Use FEE Test and view the plots to look for any dead modules that may hang up the datapath." SAP Transfer:Dispatched for On site Support Problem Reported by Customer: Wechat:02N/Ingenuity Core128/Elite China/ACQ error Safety Questions:Case: 0117866756 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Jianqiang Last updated by: Wang Jianqiang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | To be followed with field and get bugreport | [12 empty] | 117866756 |  | 1101522 | TRUE | Closed - Done | Closed - Done |  | CTUDT00630408 | [5 empty] | Engineering | [6 empty] | ACQ Error | [5 empty] | Refresh |  |  | System issue | System | CIRS-SW | ACQ_NOT_OK | CIRS-SW-ACQ_NOT_OK | CIRS-SW-ACQ_NOT_OK | ACQ_NOT_OK | ACQ_NOT_OK | CIRS-SW | CIRS-SW-ACQ_NOT_OK | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1162 |  |  | Japan | Social Welfare Organization Saiseikai Imperial Gift Foundation, Inc. | 343000 | 44391 |  | 01CS:TM:CT-Host power was off although nothing was being done. >> There was normal operation once power was [Truncated] | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer: 01CS:TM:CT-Host power was off. >> Safety Questions:Case: 0117883620 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: TAKEYUKI MIYAKAWA Last updated by: TAKEYUKI MIYAKAWA 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Problem description by engineer : Equipment affected: CT Problem location: 01 Console When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user_s work flow did the problem occur?: Other What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem?: NA Engineer_s Comment: _============================== $version = "v4.1.10.25210" 22:21 Call description: Normally the CT-Host power is left on, but it turned off at some time. The CIRS power was still on. _============================== Troubleshooting Action: What steps are taken during the system diagnosis: Result of response: Problem solved. Problem location: 01 Console What steps did you perform to diagnose the problem? _================================== Results of checking the log: I checked the event log and found that the shutdown was running normally at 16:20 today.There was no error recorded. Also, just before that, the key switch was off.(I cannot identify the problem from the log, but it seems he has also logged out) CIRS power was still on. I think, from the situation, that the user shut down the CT-Host. _================================== Error # and/or description of error: NA Repair Action: [Confirmation of problem] Confirmed. Work details: I checked the log and reported to the customer. 【Confirmation content】 Good. OneEMS internal remarks : Problem description by engineer : Equipment affected: CT Problem location: 01 Console When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user_s work flow did the problem occur?: Other What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem?: NA Engineer_s Comment: _============================== $version = "v4.1.10.25210" 22:21 Call description: Normally the CT-Host power is left on, but it turned off at some time. The CIRS power was still on. _============================== What steps are taken during the system diagnosis: Result of response: Problem solved. Problem location: 01 Console What steps did you perform to diagnose the problem? _================================== Results of checking the log: I checked the event log and found that the shutdown was running normally at 16:20 today.There was no error recorded. Also, just before that, the key switch was off.(I cannot identify the problem from the log, but it seems he has also logged out) CIRS power was still on. I think, from the situation, that the user shut down the CT-Host. _================================== Error # and/or description of error: NA Repair Action "External" : [Confirmation of problem] Confirmed. Work details: I checked the log and reported to the customer. 【Confirmation content】 Good. Repair Action "Internal" : --Troubleshooting Action-- Result of response: Problem solved. Problem location: 01 Console What steps did you perform to diagnose the problem? _================================== Results of checking the log: I checked the event log and found that the shutdown was running normally at 16:20 today.There was no error recorded. Also, just before that, the key switch was off.(I cannot identify the problem from the log, but it seems he has also logged out) CIRS power was still on. I think, from the situation, that the user shut down the CT-Host. _================================== --Repair Action-- --Subject-- 01CS:TM:Mr. Yamaguchi:CT-Host power was off although nothing was being done. >> There was normal operation once power was turned on.It looks like a shutdown was done normally in the evening. --Problem Description-- Equipment affected: CT Problem location: 01 Console When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user_s work flow did the problem occur?: Other What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem?: NA Engineer_s Comment: _============================== $version = "v4.1.10.25210" 22:21 Call description: Normally the CT-Host power is left on, but it turned off at some time. The CIRS power was still on. _============================== --4/8Elements-- Q1. Was the reported problem confirmed? Confirmed. Q2. Were any other problems confirmed? None. Q3. What did you do to resolve the problem?: I checked the log and reported to the customer. Q4: Is the equipment operating normally following the response? Good. Q5. How did the examination finish? None. Q6. Were there any mistaken findings (misdiagnosis, mistaken treatment) due to the problem? None. Q7. Was there any unintended unnecessary exposure? None. Q8. Were there any artifacts? None. SAP Transfer: Case Transferred to SAP ----------------------------------------------------------------------------- | 4 |  | from the situation, that the user shut down the CT-Host. | [12 empty] | 117883620 |  | 1115498 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Host off | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1163 |  |  | Japan | 医療法人清翠会 牧病院 | 341017 | 44389 |  | 01DS: KK: Cannot transfer images to PACS. >> It does not return by restarting the Host. It became possible to transfer after a while. | 01DS: KK: Cannot transfer images to PACS. >> It does not return by restarting the Host. It became possible to transfer after a while. " BU Q & R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 01DS: KK: Unable to transfer images to PACS. >> It does not return by restarting the Host. It became possible to transfer after a while. SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Support T2 Activities: Problem Description: Problem description by engineer: Target device: CT Failure location: 01Cosole When did it first occur? (Within ~): 1 Hour How often do problems occur ?: First Occurrence At what point in the end-user workflow is the problem occurring ?: Network What happened recently ?: N.A. Does the customer create a specific system event report for the problem ?: No. The customer was unable to create a log file or report. Troubleshooting Action: What steps are taken during the system diagnosis: Post-response results: Improved Failure location: 01Cosole What steps were taken during the diagnosis ?: Images are no longer transferred to PACS. I restarted the Host, but it didn_t work. And incoming call. When I checked the log, the PACS queue was accumulated in Waiting. Ping went to PACS. When I contacted the customer with the result of log confirmation, it started to be transferred. Consult with FSE. It was investigated after raising Bugrep. Error # and / or description of error: NA Repair Action: Service representative visit work. OneEMS internal remarks: Problem description by engineer: Target device: CT Failure location: 01Cosole When did it first occur? (Within ~): 1 Hour How often do problems occur ?: First Occurrence At what point in the end-user workflow is the problem occurring ?: Network What happened recently ?: N.A. Does the customer create a specific system event report for the problem ?: No. The customer was unable to create a log file or report. What steps are taken during the system diagnosis: Post-response results: Improved Failure location: 01Cosole What steps were taken during the diagnosis ?: Images are no longer transferred to PACS. I restarted the Host, but it didn_t work. And incoming call. When I checked the log, the PACS queue was accumulated in Waiting. Ping went to PACS. When I contacted the customer with the result of log confirmation, it started to be transferred. Consult with FSE. It was investigated after raising Bugrep. Error # and / or description of error: NA Repair Action "" External "": Visit work for service personnel. Repair Action "" Internal "": --Troubleshooting Action-- Post-response results: Improved Failure location: 01Cosole What steps were taken during the diagnosis ?: Images are no longer transferred to PACS. I restarted the Host, but it didn_t work. And incoming call. When I checked the log, the PACS queue was accumulated in Waiting. Ping went to PACS. When I contacted the customer with the result of log confirmation, it started to be transferred. Consult with FSE. It was investigated after raising Bugrep. --Repair Action-- --Subject-- 01DS: KK: Noda: I can not transfer images to PACS. >> It does not return by restarting the Host. It became possible to transfer after a while. --Problem Description-- Target device: CT Failure location: 01Cosole When did it first occur? (Within ~): 1 Hour How often do problems occur ?: First Occurrence At what point in the end-user workflow is the problem occurring ?: Network What happened recently ?: N.A. Does the customer create a specific system event report for the problem ?: No. The customer was unable to create a log file or report. External Remarks: 14.07.2021 00:49:45 UTC YOSHIYUKI YAMAMOTO (81100236) [Request] PACS transfer is stopped halfway. [Phenomenon confirmation] I was able to confirm. [Details of work] I was able to confirm the phenomenon from the log. About the phenomenon that the image transfer that occurred this time stops in the middle, I got the log file at the time of occurrence. We will send it to the Dutch headquarters for analysis. It is a function that the software has that the transfer resumed and completed without doing anything This is due to the fact that the "function that times out 15 minutes after the transfer stagnation occurs and restarts the transfer" has worked. [Other failures / treatments] Yes When QRing to PACS, all series were searched and frozen, so We have changed the settings so that you can sort by any search. [Device status] Equipment can be used [Work status] done 12-Jul-2021 81100236 FSE Handover Tool1.3.4.32612.DB13 20210415 Internal Remarks: 14.07.2021 00:49:49 UTC YOSHIYUKI YAMAMOTO (81100236) [Replacement parts] No replacement parts [Misorder Parts Report] No mistakes [Waiting information] No wait-and-see [TS Requests] No return parts [CT additional information] Q1 How did the inspection end ?: No extra irradiation due to stop before X-ray irradiation Q2 Presence or absence of incorrect findings (misdiagnosis, mistreatment) due to problems: None Q3 Presence or absence of artifacts: None Q4 Presence or absence of unintended unnecessary exposure: None [Repair certificate] Repair certificate is pasted [Other information sharing matters] 1. Treatment for patients infected with or suspected of being infected with the new coronavirus / Is it quarantined? Okay 2. Do I need to enter the treatment room or bed where the infected patient is? (Workplace and transfer (Including dynamic path) No Did you agree on how to complete the report with the person in charge of the equipment? YES Normal report completed on the day 12-Jul-2021 81100236 FSE Handover Tool1.3.4.32612.DB13 20210415 | 4 |  | About the phenomenon that the image transfer that occurred this time stops in the middle, I got the log file at the time of occurrence. We will send it to the Dutch headquarters for analysis. It is a function that the software has that the transfer resumed and completed without doing anything This is due to the fact that the "function that times out 15 minutes after the transfer stagnation occurs and restarts the transfer" has worked. | [12 empty] | 117866778 |  | 1101553 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS Transfer | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1164 |  |  | Puerto Rico | DORADO HEALTH INC DBA | 40007 | 44392 |  | Noise on Gantry | Work Order #: WO-03867300 Alec, Preciado The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Internal Comments : Error # and/or description of error : belt wear down and tensiones out of spec. Repair Action : Parts requested by the FSE, belt and tension assembly Diagnostic performed by Engineer : Parts requested by the FSE, belt and tension assembly Problem Description by Engineer : Problem description by engineer : The rotor belt and belt tensioner was defective, please order both parts When was the first occurrence? Within... : 16-Jul-2021 10:00:00 How often is issue occurring? : First occurrence Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected results: banda de rotor reemplazada y ajustada Patient Impact/Outcome: no impacta User Impact: no impacta Current Software Version: 4.1.7 Resolution : Belt and tensioner assembly replaced Follow Up Required : Error # and/or description of error : belt wear down and tensiones out of spec. Repair Action : Parts requested by the FSE, belt and tension assembly What are the next Troubleshooting steps : Replace belt and tensioner and adjust as indicated in the repair replacement manual. Technician Name : Alec, Preciado ----------------------------------------------------------------------------- | 3 |  | The rotor belt and belt tensioner was defective, please order both parts | [12 empty] | 117892755 |  | 1120398 | TRUE | Evaluation | Evaluation |  | ECR-077012 | [5 empty] | Engineering | [6 empty] | Belt | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Belt | Gantry-HW-Belt | Gantry-HW-Belt | ECR-077012 | Belt | Gantry-HW | Gantry-HW-Belt | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1165 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 337000 | 44393 |  | Ring artifact problem | Resolution: It was seen that there was a ring artifact in the images taken on the device. When the inside of the gantry was checked, opaque material was seen on the detector sheath, under the mylar and on the collimator. Air calibration was performed by cleaning the opaque material. A test shot was made with the Phantom, no problem was seen, the device was delivered in an active manner. | 3 |  | opaque material was seen on the detector sheath | [12 empty] | 117895561 |  | 1119110 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Contrast intrusion | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1166 |  |  | United States | Philips Medical Capital | 40055 | 44392 | 28 | 0603CDY70 Customer is getting and error when trying to run QC. | Problem Description by Engineer : Problem description by engineer : Spoke with Brent.The customer is unable to scan. They are seeing a CIRS error. Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: No further information is available as per the customer Patient Impact/Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer Current Software Version: 4.1.10 Diagnostic performed by Engineer : Used logger.mdb, UNLV, and AEL to locate the error messages. Follow Up Required : Error # and/or description of error : S_DMC_SHOT_COMM_COLLAPSE. This indicates that no frames are being sent betweenthe DMS and the CIRS. This is a datapath issue. Parts required : Part : 1, 12NC : 453567057167, Description : FRC Assembly, Quantity : 1, Priority : 1, Chance Of Success : 6.1 Part : 2, 12NC : 453567538334, Description : Module - Rotor Electronics TX, Quantity : 1, Priority : 2, Chance Of Success : ? Part : 3, 12NC : 453567538322, Description : Module, Optical RX Stator Moog Velocity, Quantity : 1, Priority : 3, Chance Of Success : ? Repair Action : The S_DMC_SHOT_COMM_COLLAPSE indicates a datapath issue. The CIRS event log shows no line faults or data dropouts so it appears that the frames are never being sent from the DMS/RCOM. Look for power from the FRC to the transmitter to be present. Use the diagnostic LEDs on the TX to assist. Run datapath diagnostics to look for the source of the errors. Bypass the optical slip ring with a quad fibre from the RCOM and do a non rotation scan. What are the next Troubleshooting steps : Look for power from the FRC to the transmitter to be present. Use the diagnostic LEDs on the TX to assist. Run datapath diagnostics to look for the source of the errors. Bypass the optical slip ring with a quad fibre from the RCOM and do a non rotation scan. Technician Name : Mark Finnigan Internal Comments : Error # and/or description of error : S_DMC_SHOT_COMM_COLLAPSE. This indicates that no frames are being sent betweenthe DMS and the CIRS. This is a datapath issue. Repair Action : The S_DMC_SHOT_COMM_COLLAPSE indicates a datapath issue. The CIRS event log shows no line faults or data dropouts so it appears that the frames are never being sent from the DMS/RCOM. Look for power from the FRC to the transmitter to be present. Use the diagnostic LEDs on the TX to assist. Run datapath diagnostics to look for the source of the errors. Bypass the optical slip ring with a quad fibre from the RCOM and do a non rotation scan. ----------------------------------------------------------------------------- | 2 |  | 40055 site is a false alarm. A wall box power recycle fixed he problem. No parts replaced. No more errors on site since day1 of the problem. system is monitoring. | [12 empty] | 117890421 |  | 1115563 | TRUE | Evaluation | Evaluation | [7 empty] | Others | [6 empty] | Power cycle | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1167 |  |  | United States | Philips Medical Capital | 40053 | 44389 |  | Intermittent Issues with EKG Connection | Internal Remarks : FADB: (FSE Aware of Double Booking) Title: Internal Remarks Initials: Twila Actions Taken: Per message sent in Teams from FSE/McDowell Preston - Moved WO from 7/14/21 to 7/21/21 (anytime). I kept the same time. Outcome: Moved WO per FSE_s Request IP: 88993323 Next Steps: None ----------------------------------------------------------------------------- Work Order #: WO-03832660 McDowell Preston Diagnostic performed by Engineer : N/A Problem Description by Engineer : Customer Function/Role: Technologist Device Use: Diagnostic Expected and Unexpected Results: Customer States Intermittent EKG Waveform Not Present on Host PC Monitor. Customer has to Restart Host PC to get EKG Waveform User Impact: Unknown Patient Impact: Unknown Current Software Version:4.1.10 ----------------------------------------------------------------------------- Work Order #: WO-03840684 Weaver David The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Description of Problem Intermittent Issues with EKG Connection CALLER FUNCTION/ROLE Technician/Sonographer/Therapist DEVICE INTENDED USE Diagnostic WHEN WAS THE FIRST OCCURRENCE? WITHIN . . . <1 Week HOW OFTEN IS ISSUE OCCURRING? Intermittent WHERE IN THE END-USER WORKFLOW IS THE ISSUE OCCURRING? Pre-acquisition/scan WHAT RECENT EVENTS HAVE OCCURRED? During normal use NEGATIVE OPERATOR IMPACT EKG not workin sometimes NEGATIVE PATIENT IMPACT/PROCEDURE OUTCOME Delay in scanning HAS CUSTOMER CREATED A SPECIFIC SYSTEM EVENT REPORT OF ISSUE? No 4.1.10 Diagnostic performed by Engineer : Intermittent Issues with EKG Connection Resolution : Intermittent Issues with EKG Connection. ----------------------------------------------------------------------------- Work Order #: WO-03832183 Paine Matthew The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Customer Function/Role: Technologist Device Use: Diagnostic Expected and Unexpected Results: Customer States Intermittent EKG Waveform Not Present on Host PC Monitor. Customer has to Restart Host PC to get EKG Waveform User Impact: Unknown Patient Impact: Unknown Current Software Version:4.1.10 Diagnostic performed by Engineer : Customer states that they have to restart the Host PC everytime the go to use the EKG Monitor. After Restarting the Host PC, the EKG Waveform is now present on the Host PC. Follow Up Required : Issue: Customer States Intermittent Issue with EKG Waveform Not Present on Host PC Monitor, but is Present on the EKG Monitor. Customer States Restart of Host PC Resolves Issue. Recommendation: Check All Cabling to CRC Rack. Use Upgrade Installation Manual for Proper USB Cable Locations. Check to Ensure that Cabling from SBD to Host PC is Correct. Using the wrong USB can cause Intermittent Issue_s if not Installed Correctly. If Cabling is correct, check USB Extender Version. If Upgarde Kit is needed: 459801731761-USB Extender Cable from IBOX To Console Internal Comments : Called and left Primary FSE Preston McDowell a VM. FSE needs to Inspect Cabling for EKG from Ibox to CRC Rack and from CRC Rack to Host PC to ensure proper cabling for USB2 Devices. ----------------------------------------------------------------------------- | 4 |  | Product/Part - USB EXTENDER CABLE FROM IBOX TO CONSOLE Part number - 459801731761 | [12 empty] | 117870146 |  | 1099343 | TRUE | Closed - Done | Closed - Done | [7 empty] | Others | [6 empty] | USB Cable | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | n | y | y | y | y | #N/A
1168 |  |  | Viet Nam | Công ty TNHH Thương mại và | 40046 | 44389 |  | Vert Encoder got defective | Product/Part - VERT ENCODER ASSY Part number - 453566502431 | 4 |  | Product/Part - VERT ENCODER ASSY Part number - 453566502431 | [12 empty] | 117866928 |  | 1101707 | TRUE | Closed - Done | Closed - Done |  | TBD-Encoder | [5 empty] | Engineering | [6 empty] | Vertical Encoder | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | VerticalEncoder | Couch-HW-VerticalEncoder | Couch-HW-VerticalEncoder | TBD-Encoder |  | Couch-HW | Couch-HW-VerticalEncoder | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | #N/A
1169 |  |  | Chile | Scotiabank Chile | 40008 | 44397 |  | CL1675 - ESPEJO - EQUIPMENT NEEDS TO BE CONFIGURED FOR SEVERAL RIS PACS TASKS. ON AUGUST 21ST AT 2:00 P.M. LOCAL TIME |  | 3 |  | RIS/PACS config | [12 empty] | 117917617 |  | 1135736 | TRUE | Opened | Opened | [7 empty] | Customer | [6 empty] | RIS/PACS config | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1170 |  |  | China | 四川华蓥市华蓥山广能集团总医院 | 345000 | 44401 | 29 | 02N/Ingenuity Core128/Elite China/scan fault | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity Core128/Elite China Safety Questions : Case: 0117957832 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Bo Last updated by: Wang Bo 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 扫描失败。 客户问题详细描述或故障现象： 定位像扫描失败。 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： ACQ NOT OK，CRC ABCD=700+ RESEND=500+ fiber optical disconnected 远程诊断与测试结果： 重启无效。 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议现场做ACQ自诊断，测量光强，检查接受发射盒供电和工作状态，清洁光路，bypass Rcom-CIRS端尝试排除故障。 建议备件12NC及名称： PPD 453567538334，453567538322 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : PPD 453567538322*1+453567538334 *1 已安排备件24H到医院 李全飞-86-13699457740 External Remarks : 25.07.2021 07:59:39 UTC Haijun Chen (86001310)"SUAT. Ready for Clinical Use local time" : "2021-07-24 23:00" External Remarks : 25.07.2021 08:02:43 UTC SUP JCOUSER (86001310)服务内容:机器不能扫描，发射盒有问题，更换发射盒。 服务结果: 维修后，设备恢复正常工作。 Internal Remarks : 25.07.2021 08:02:46 UTC Haijun Chen (86001310) "RDI Accuracy:Yes" $Q7:Confirmed reported problem.can not scan. $Q8: Check the logs of CIRS Error: ACQ NOT OK and find the TX have problem. $Q9: Replace TX(12NC:453567538334) $Q10: Test and inspection data refer to: PA. Device returned to full functionality. $Q11: NA. current tube install dated:(N/A) current tube usage(:N/A) | 2 |  | $Q7:Confirmed reported problem.can not scan. $Q8: Check the logs of CIRS Error: ACQ NOT OK and find the TX have problem. $Q9: Replace TX(12NC:453567538334) $Q10: Test and inspection data refer to: PA. Device returned to full functionality. | [12 empty] | 117957832 |  | 1148076 | TRUE | Opened | Opened |  | TBD_OBS | [5 empty] | Engineering | [6 empty] | TX | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Other | Gantry-HW-Other | Gantry-HW-Other | TBD_OBS | TBD-NoEnoughInfo | Gantry-HW | Gantry-HW-Other | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1171 |  |  | China | 北京市中西医结合医院 | 40032 | 44399 |  | Wechat:02N/Ingenuity CT 728326/connect PACS | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : connect PACS Safety Questions : Case: 0117927110 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wennian Han Last updated by: Wennian Han 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 添加PACS 设备本身技术问题： 是。 Troubleshooting Action: 远程诊断与测试结果： 添加PACS PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 添加PACS，教客户使用飞利浦微信报修 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 22.07.2021 11:20:53 UTC Guochao Fan (86039456)服务内容：添加pacs节点。 服务结果：服务后设备恢复正常。 Internal Remarks : 22.07.2021 11:20:56 UTC Guochao Fan (86039456) "RDI Accuracy: Yes" $Q7: Confirm reported problems:Adding pacs $Q8: N/A $Q9:N/A $Q10: Test and Inspection data refer to PA& Device status:Device returned to full functionally $Q11: N/A ----------------------------------------------------------------------------- | 5 |  | RIS/PACS config | [12 empty] | 117927110 |  | 1142280 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | Customer | [6 empty] | PACS Config | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1172 |  |  | China | 东莞市中西医结合医院 | 345013 | 44398 |  | WeChat:02N/Ingenuity Core128/ECG problem | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : error External Remarks : 23.07.2021 13:12:44 UTC Zongxian Li (86001504)服务内容：更换导联线。 服务结果：工作正常。 Safety Questions : Case: 0117920811 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wennian Han Last updated by: Wennian Han 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP Internal Remarks : 23.07.2021 13:12:47 UTC Zongxian Li (86001504) RDI Accuracy:Yes $Q7: (Confirmed the reported problem, ECG signal abnoarmal) $Q8: (Checked the leadset in MP5 broken) $Q9: (Replaced leadset) $Q10: (Test and inspection data refer to PA. Device returned to full functionality.) $Q11: $Calibrated Tools name (N/A) S/N (N/A) Next Calibration date (N/A) SAP Transfer : Dispatched for Onsite Support T2 Activities : Problem Description: 客户问题概述： ECG相关故障。 客户问题详细描述或故障现象： 心电问题 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： NA 远程诊断与测试结果： 客户告知总是出現锯齿样心电图，有时是心电监护仪的心电图与我操作电脑的心电图不一致，检查过电极片，擦拭酒精，排除了心率问题。需要到场检查心电 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 检查电极片，接地，导联线等线缆、IBOX程序，IBOX板、检查IBOX到主机通路，以及转换盒状态，检查心电监护仪。 建议备件12NC及名称： " 3 Leadset, Grabber, IEC, ICU 989803145101" Next Step Action : PPD 989803145101*1 已安排24h到医院 FSE李宗贤-86-13751567086 ----------------------------------------------------------------------------- | 3 |  | 3 Leadset, Grabber, IEC, ICU 989803145101 | [12 empty] | 117920811 |  | 1138158 | TRUE | Investigation | Investigation | [7 empty] | Others | [6 empty] | ECG | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1173 |  |  | China | Shandong Province Chest Hospital | 40025 | 44397 |  | 02N/Ingenuity CT 728326/Unable to start normally | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : Unable to start normally Safety Questions : Case: 0117916251 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Zhou Xiaohong Last updated by: Zhou Xiaohong 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 软件登录故障 客户问题详细描述或故障现象： 软件登录故障 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： N/A 远程诊断与测试结果： N/A PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 软件bug，已开ESCALATION，需要FSE现场用管理员账户登录，恢复CT账号。 建议备件12NC及名称： N/A 客户需求备注： 无 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Internal Remarks : 25.07.2021 00:00:59 UTC Zhaokun Huang (86038405) "RDI Accuracy: No." $Q7: ( Confirmed reported problem & Software can not login in) $Q8: ( Login in Software use smart card and reload software. & It is app bug from TSE and fixed wait for SP2.) $Q9: ( Login in Software use smart card and reload software . ) $Q10: ( Test and inspection data refer to: PA . Device returned to full functionality.) $Q11: Calibrated Tools name (N/A) ^Current Tube installed date: N/A ^Current Tube usage: N/A External Remarks : 21.07.2021 10:00:16 UTC Zhaokun Huang (86038405)服务内容：按照手册，重新安装软件。 服务结果：维修后，设备恢复正常。 ----------------------------------------------------------------------------- | 3 |  | blue screen $Q7: ( Confirmed reported problem & Software can not login in) $Q8: ( Login in Software use smart card and reload software. & It is app bug from TSE and fixed wait for SP2.) $Q9: ( Login in Software use smart card and reload software . ) | [12 empty] | 117916251 |  | 1134828 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | CTUDT00632553 | [5 empty] | Engineering | [6 empty] | Blue screen | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Blue-Screen | Console-SW-Blue-Screen | Console-SW-Blue-Screen | CTUDT00632553 | Blue-Screen | Console-SW | Console-SW-Blue-Screen | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1174 |  |  | China | Nanchang University First Affiliated hospital (Jinggangshan Hospital) | 345002 | 44397 |  | 02N/Ingenuity Core128/Elite China/ | BU Q&R Feedback: No Reason: Corrective Maintenance Safety Questions:Case: 0117915495 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Gao Luqi Last updated by: Gao Luqi 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ ----------------------------------------------------------------------------- | 3 |  | Not clear | [12 empty] | 117915495 |  | 1134208 | TRUE | Evaluation | Evaluation | [7 empty] | TBD | [6 empty] | 44397 | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | n | y | y | y | y | #N/A
1175 |  |  | China | The Second Hospital of Haibei Prefecture in Qinghai Province | 345029 | 44398 |  | 02n/Ingenuity Core128/Elite China/Add printer | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Configuration Error Repair Activity Code: Instructed Customer Problem Reported by Customer: Add printer Safety Questions:Case: 0117921776 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities: Problem Description: add printer Troubleshooting Action: add printer Repair Action: instruct customer add printer.system is ok. Test/Inspection: Functional Check PASSED. The device returned to full functionality. Test Equipment: N/A ----------------------------------------------------------------------------- | Lu, Xueying |  | 44398 | [12 empty] | 117921776 |  | 1138854 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | 44399.6944 | [6 empty] | 44398 | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1176 |  |  | China | 陕西省安康中医医院 | 345041 | 44396 |  | WeChat:02N/Ingenuity Core128/Elite China/Host crash frequently | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : WeChat:02N/Ingenuity Core128/Elite China/Host crash frequently Safety Questions : Case: 0117908321 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Cao Last updated by: Liu Cao 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 其他类型故障。 客户问题详细描述或故障现象： 软件死机 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 软件死机 远程诊断与测试结果： 软件问题 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 重装系统 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 21.07.2021 15:11:12 UTC Yujie Guo (86038990)服务内容：扫描软件间歇性卡顿，重装系统恢复。 服务结果：服务后，设备正常工作。 ----------------------------------------------------------------------------- | 3 |  | 服务内容：扫描软件间歇性卡顿，重装系统恢复。 服务结果：服务后，设备正常工作。 Check if can get bugreport? | [12 empty] | 117908321 |  | 1128823 | TRUE | Closed - Done | Closed - Done |  | TBD-SW-Slow | [5 empty] | Engineering | [6 empty] | SW slow | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Performance | Console-SW-Performance | Console-SW-Performance | TBD-SW-Slow | Performance | Console-SW | Console-SW-Performance | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1177 |  |  | China | Jiangsu Xuzhou Medical School Affiliated No. 3 Hospital | 40039 | 44396 |  | 02G/Ingenuity CT 728326/The bed makes a strange noise | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer: The bed makes a strange noise Safety Questions:Case: 0117908206 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Xianyi Last updated by: Wang Xianyi 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer_s problems: The device makes a strange sound. Detailed description of Customer_s problems or malfunctions: The bed has a strange sound It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: The bed has abnormal sound Remote troubleshoot and test result: The bed has abnormal sound and needs to be checked on the field PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Check the noises in the bed. Look for sound source and troubleshoot Next Step Action: contacted FSE Zhang Chuanxi to confirm, add Operation Next Step Action: mailed IB, add Operation to FSE Zhang Chuanxi Next Step Action : OP added 0030 ----------------------------------------------------------------------------- | 3 |  | Repair Action: Suggest onsite service procedures: Check the noises in the bed. Look for sound | [12 empty] | 117908206 |  | 1128707 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | 44396 | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | n | y | y | y | y | #N/A
1178 |  |  | China | Shanxi YangGao County People_s Hospital | 345025 | 44398 |  | 02N/Ingenuity Core128/Elite China/Print problem | BU Q&R Feedback: No SAP Customer Symptom Code: CENV Customer Symptom Code: Customer Environmental Concern Reason: Corrective Maintenance Engineer Failure Code: Environmental Cause Repair Activity Code: Instructed Customer Problem Reported by Customer: 02N/Ingenuity Core128/Elite China Safety Questions:Case: 0117920393 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Weng Jinming Last updated by: Weng Jinming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Troubleshooting Action: answer customer question Repair Action: answer customer question Test/Inspection: Functional Check PASSED. The device returned to full functionality. Test Equipment: N/A SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | Lu, Xueying |  | 44398 | [12 empty] | 117920393 |  | 1137889 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | 44399.4306 | [6 empty] | 44398 | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1179 |  |  | Germany | Gesundheit Nord gGmbH Klinikverbund | 40002 | 44399 |  | Verzögerte Bildübertragung ins PACS | Customer comment: 0034W000028wmrPQAQ Customer comment: Rs with Mr. Halama, T for 2h ## Appointment with Mr. Glogner-Umbeer for Jul 28, 1:30 p.m. ## -------------------------------------------------- --------------------------- Work Order #: WO-03921170 Halama, Thomas Diagnostic performed by Engineer: N / A Problem Description by Engineer: Radiologists got the impression that the images were from the end of the scan take longer than usual to reach the Pacs. -------------------------------------------------- --------------------------- Work Order #: WO-03912460 Zurstegge, Uwe The system meets the specification for the performed service and is returned to use Diagnostic performed by Engineer: phone call with Raik Glogner-Umbeer Medizintechnik. House IT Mr. Gloistein has commissioned the Pacs company Ashvin analyze the Dicomtransfer CT-Pacs. Follow Up Required: Check Cir_s Recon Times, Delete lan spooler. Contacting the Pacs company Ashvin, whether log files regarding Transfer Ingenuity Pacs have already been evaluated. Telephone number Haus IT H. Gloistein 0421/49781181 Problem Description by Engineer: Radiologists got the impression that the images were from the end of the scan take longer than usual to reach the Pacs. | 4 |  | Delete lan spooler. Contacting the Pacs company Ashvin, whether log files regarding Transfer Ingenuity Pacs have already been evaluated. | [12 empty] | 117949498 |  | 1142449 | TRUE | Evaluation - Pending Translation | Evaluation - Pending Translation | [7 empty] | TBD | [6 empty] | 44399 | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1180 |  |  | Japan | Social Medical Corporation BOKOI, Nikko Memorial Hospital | 341011 | 44398 |  | 02DS:KK: "Tape switch pressed" appeared and scanning is not possible.>> Releasing the tape switch solved the [Truncated] | BU Q&R Feedback: No Reason: Corrective Maintenance T2 Activities:Problem Description: Problem description by engineer : Equipment affected: CT Failure location: 02Couch When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user_s work flow did the problem occur?: (Pre-Acquisition/Scan) What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem? No.The customer has not created a log file or report. Troubleshooting Action: What steps are taken during the system diagnosis: Result of response: Problem solved. Failure location: 02Couch What steps did you perform to diagnose the problem? It seems the tape switch was not released and a scan did not run.The customer said that after he operated the tape switch a number of times to release it, operation returned to normal.An FSE will respond. Error # and/or description of error: NA Repair Action: A service engineer visited the site. OneEMS internal remarks : Problem description by engineer : Equipment affected: CT Failure location: 02Couch When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user_s work flow did the problem occur?: (Pre-Acquisition/Scan) What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem? No.The customer has not created a log file or report. What steps are taken during the system diagnosis: Result of response: Problem solved. Failure location: 02Couch What steps did you perform to diagnose the problem? It seems the tape switch was not released and a scan did not run.The customer said that after he operated the tape switch a number of times to release it, operation returned to normal.An FSE will respond. Error # and/or description of error: NA Repair Action "External" : The service engineer for the site visited and worked on the equipment. Repair Action "Internal" : --Troubleshooting Action-- Result of response: Problem solved. Failure location: 02Couch What steps did you perform to diagnose the problem? It seems the tape switch was not released and a scan did not run.The customer said that after he operated the tape switch a number of times to release it, operation returned to normal.An FSE will respond. --Repair Action-- Recommended 12Nc: --None-- 459800623252:FRU, CBL, HANDLE TAPE SWITCH QTY:2 SLOC:--None-- --Subject-- 02DS:KK:FSE Hashimoto: "Tape switch pressed" appeared and scanning is not possible.>> Releasing the tape switch solved the problem.>> An FSE will respond. --Problem Description-- Equipment affected: CT Failure location: 02Couch When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user_s work flow did the problem occur?: (Pre-Acquisition/Scan) What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem? No.The customer has not created a log file or report. Problem Reported by Customer : 02DS:KK: "Tape switch pressed" appeared and scanning is not possible.>> Releasing the tape switch solved the problem.>> An FSE will respond. SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Support Safety Questions:Case: 0117920858 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: KENICHI KOJIMA Last updated by: KENICHI KOJIMA 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Message: Components Added Succesfully to SWO!- SAP Message: LOGI Task added Successfully!- ----------------------------------------------------------------------------- | 4 |  | It seems the tape switch was not released and a scan did not run.The customer said that after he operated the tape switch a number of times to release it, operation returned to normal.An FSE will respond. | [12 empty] | 117920858 |  | 1138182 | TRUE | Evaluation | Evaluation | [7 empty] | TBD | [6 empty] | Tape switch | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1181 |  |  | Japan | Social Medical Corporation BOKOI, Nikko Memorial Hospital | 341011 | 44398 |  | 01CS:KK: The Exam Summary was not created for a particular examination, and it was also no automatically transferred. | BU Q&R Feedback: No SAP Customer Symptom Code: HANG Customer Symptom Code: Hang/Crash Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: Reboot/Restart Problem Reported by Customer: 01CS:KK: The Exam Summary was not created for a particular examination, and it was also no automatically transferred. T2 Activities:Problem Description: Problem description by engineer : Equipment affected: CT Problem location: 01 Console When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user_s work flow did the problem occur?: Post-Acquisition/Scan What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem? No.The customer has not created a log file or report. Troubleshooting Action: What steps are taken during the system diagnosis: Result of response: Problem solved. Problem location: 01 Console What steps did you perform to diagnose the problem? Could not find any particular error.Data transfer was restored after restarting the host.I had the FSE contact the customer and have them take a wait-and-see approach. Error # and/or description of error: NA Repair Action: [Confirmation of problem] Confirmed. Work details: Could not find any particular error.Data transfer was restored after restarting the host.I had the FSE contact the customer and have them take a wait-and-see approach. 【Confirmation content】 Good. OneEMS internal remarks : Problem description by engineer : Equipment affected: CT Problem location: 01 Console When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user_s work flow did the problem occur?: Post-Acquisition/Scan What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem? No.The customer has not created a log file or report. What steps are taken during the system diagnosis: Result of response: Problem solved. Problem location: 01 Console What steps did you perform to diagnose the problem? Could not find any particular error.Data transfer was restored after restarting the host.I had the FSE contact the customer and have them take a wait-and-see approach. Error # and/or description of error: NA Repair Action "External" : [Confirmation of problem] Confirmed. Work details: Could not find any particular error.Data transfer was restored after restarting the host.I had the FSE contact the customer and have them take a wait-and-see approach. 【Confirmation content】 Good. Repair Action "Internal" : --Troubleshooting Action-- Result of response: Problem solved. Problem location: 01 Console What steps did you perform to diagnose the problem? Could not find any particular error.Data transfer was restored after restarting the host.I had the FSE contact the customer and have them take a wait-and-see approach. --Repair Action-- --Subject-- 01CS:KK: Mr. Nishiya: The Exam Summary was not created for a particular examination, and it was also no automatically transferred. --Problem Description-- Equipment affected: CT Problem location: 01 Console When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user_s work flow did the problem occur?: Post-Acquisition/Scan What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem? No.The customer has not created a log file or report. --4/8Elements-- Q1. Was the reported problem confirmed? Confirmed. Q2. Were any other problems confirmed? None. Q3. What did you do to resolve the problem?: Could not find any particular error.Data transfer was restored after restarting the host.I had the FSE contact the customer and have them take a wait-and-see approach. Q4: Is the equipment operating normally following the response? Good. Q5. How did the examination finish? None. Q6. Were there any mistaken findings (misdiagnosis, mistaken treatment) due to the problem? None. Q7. Was there any unintended unnecessary exposure? None. Q8. Were there any artifacts? None. Safety Questions:Case: 0117920389 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: KENICHI KOJIMA Last updated by: KENICHI KOJIMA 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 4 |  | Could not find any particular error.Data transfer was restored after restarting the host.I had the FSE contact the customer and have them take a wait-and-see approach. Monitoring and collect bugreport/problem details for investigation. | [12 empty] | 117920389 |  | 1137888 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | TBD | [6 empty] | 44398 | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1182 |  |  | Japan | Maki Hospital | 341017 | 44396 |  | 01DM:TM: Data transfer from the CT-Host has stopped.The data is remaining in QM>> Restarting the CT-Host sol [Truncated] | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 01DS:TM: Image transfer stopped>> SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Support Safety Questions:Case: 0117907956 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: TAKEYUKI MIYAKAWA Last updated by: TAKEYUKI MIYAKAWA 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Problem description by engineer : Equipment affected: CT Problem location: 01 Console When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user_s work flow did the problem occur?: Network What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem?: NA Engineer_s Comment: _============================== $version = "v4.1.10.25210" 10:46 Call description: Data transfer is stopped. Restarting the CT-Host did not change the situation. He called us when he was restarting the system after he shut down the CT host. _============================== Troubleshooting Action: What steps are taken during the system diagnosis: Result of response: Problem solved. Problem location: 01 Console What steps did you perform to diagnose the problem? _================================== Result Transfer became possible after restarting. However, the same problem occurs frequently, which interferes with work. _================================== Error # and/or description of error: NA Repair Action: A service engineer visited the site. OneEMS internal remarks : Problem description by engineer : Equipment affected: CT Problem location: 01 Console When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user_s work flow did the problem occur?: Network What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem?: NA Engineer_s Comment: _============================== $version = "v4.1.10.25210" 10:46 Call description: Data transfer is stopped. Restarting the CT-Host did not change the situation. He called us when he was restarting the system after he shut down the CT host. _============================== What steps are taken during the system diagnosis: Result of response: Problem solved. Problem location: 01 Console What steps did you perform to diagnose the problem? _================================== Result Transfer became possible after restarting. However, the same problem occurs frequently, which interferes with work. _================================== Error # and/or description of error: NA Repair Action "External" : The service engineer for the site visited and worked on the equipment. Repair Action "Internal" : --Troubleshooting Action-- Result of response: Problem solved. Problem location: 01 Console What steps did you perform to diagnose the problem? _================================== Result Transfer became possible after restarting. However, the same problem occurs frequently, which interferes with work. _================================== --Repair Action-- --Subject-- 01DM:TM:Mr. Hoshi: Data transfer from the CT-Host has stopped.The data is remaining in QM>> Restarting the CT-Host solved the problem. --Problem Description-- Equipment affected: CT Problem location: 01 Console When did the problem first occur (within what period of time)?: 1 hour How often does the problem occur?: First occurrence At what point in the end user_s work flow did the problem occur?: Network What sort of event has occurred recently?: NA Did the customer create a detailed system event report regarding the problem?: NA Engineer_s Comment: _============================== $version = "v4.1.10.25210" 10:46 Call description: Data transfer is stopped. Restarting the CT-Host did not change the situation. He called us when he was restarting the system after he shut down the CT host. _============================== ----------------------------------------------------------------------------- | 4 |  | Data transfer from the CT-Host has stopped.The data is remaining in QM>> Restarting the CT-Host solved the problem. v4.1.10.25210 Monitoring and collect bugreport/problem details for investigation. | [12 empty] | 117907956 |  | 1128573 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | TBD | [6 empty] | 44396 | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1183 |  |  | Poland | Government Agency. | 341009 | 44398 |  | Starting and implementing it for use. | Integration:VBAKKOM has been processed successfully-VBAPKOM has been processed successfully-KONVKOM has been processed successfully-KONVKOM has been processed successfully-KONVKOM has been processed successfully-KONVKOM has been processed successfully-REQ2RES DMR 0018133501 has been saved-MPG is present on the product code : P7 Integration: Confirmation of order 53913979 saved. ----------------------------------------------------------------------------- Work Order #: WO-03889096 21-Jul-2021 Marcin Knyps The system meets the specification for the performed service and is returned to use Problem Description by Engineer: Connection of a modular CT laboratory. Resolution: Powering on in the module CT laboratory. System calibration, Constancy and QuickIQ image quality test in accordance with the manufacturer_s specifications. Due to the lack of a configured computer network, the DICOM settings for external systems have not been entered. The device is technically ready for use. Tools & Equipment : MULTIMETER,2022-04-14 00:00:00 ----------------------------------------------------------------------------- | 4 |  | Powering on in the module CT laboratory. System calibration, Constancy and QuickIQ image quality test in accordance with the manufacturer_s specifications. Due to the lack of a configured computer network, the DICOM settings for external systems have not been entered. The device is technically ready for use. | [12 empty] | 117922996 |  | 1135626 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | 44398.9097 | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1184 |  |  | Portugal | Centro Hospitalar de Vila Nova | 40030 | 44399 |  | Password access to Exam Card Manager protocols has expired |  | 3 |  | Password access to Exam Card Manager protocols has expired | [12 empty] | 117950281 |  | 1143936 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Password | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | n | y | y | y | y | #N/A
1185 |  |  | Puerto Rico | DORADO HEALTH INC DBA | 40007 | 44400 |  | No esta realizando las reconstrucciones sagitales completas | Work Order #: WO-03923749 Valenzuela Velasquez, Eduardo Ariel Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 3 |  | Not clear | [12 empty] | 117956108 |  | 1147519 | TRUE | Evaluation | Evaluation | [7 empty] | TBD | [6 empty] | 44400 | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1186 |  |  | Puerto Rico | HOSPITAL DR SUSONI INC | 40013 | 44396 |  | the gantry at 0 has angulation, it must be given 2.5 degrees | Work Order #: WO-03878989 23-Jul-2021 Rios, William The system meets the specification for the performed service and is returned to use Resolution : Verify system, found that Gantry was tilt to the front about 3.5 grade, calibrate the Tilt and the problem was fixed, make functional test. Unit OK Problem Description by Engineer : Problem description by engineer : tilt of gantry abnormal When was the first occurrence? Within... : 19-Jul-2021 11:00:00 How often is issue occurring? : Multiple times a day Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: tilt of gantry visually incorrect Patient Impact/Outcome: n/a User Impact: they see tilt is not correct Current Software Version: unknown ----------------------------------------------------------------------------- Work Order #: WO-03875700 Caldelas Aldana, Carlos Manuel The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Problem description by engineer : tilt of gantry abnormal When was the first occurrence? Within... : 19-Jul-2021 11:00:00 How often is issue occurring? : Multiple times a day Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: tilt of gantry visually incorrect Patient Impact/Outcome: n/a User Impact: they see tilt is not correct Current Software Version: unknown Diagnostic performed by Engineer : they visually check that gantry is not correctly in tilt, they chck and when its supposed to be at 0 degrees it has in screen more than 2. and when system is at 0, they gantry looks tilted. No errors io system. Internal Comments : Error # and/or description of error : tilt in gantry Repair Action : n/a Resolution : FSE visit Follow Up Required : Error # and/or description of error : tilt in gantry Repair Action : n/a What are the next Troubleshooting steps : check gantry covers and check gantry tilt Technician Name : Caldelas Aldana, Carlos Manuel ----------------------------------------------------------------------------- | 3 |  | Resolution Description: Verify system, found that Gantry was tilt to the front about 3.5 grade, calibrate the Tilt and the problem was fixed, make functional test. Unit OK | [12 empty] | 117911037 |  | 1128075 | TRUE | Evaluation | Evaluation | [7 empty] | Others | [6 empty] | Tilt Calibration | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | n | y | y | y | y | #N/A
1187 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 337000 | 44397 |  | There are problems in images; Cleaned the opaque substances and performed air calibration. | Integration: Confirmation of order 53907744 has been saved. ----------------------------------------------------------------------------- Work Order #: WO-03881271 20-Tem-2021 Can Ozgenlik The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : Opaque substance leaked inside the device. The opaque substances on the detector sleeve were cleaned and air calibration was completed. Test scan with Phantom was completed successfully. The device was delivered in working condition. ----------------------------------------------------------------------------- | 3 |  | Contrast intrusion • Improved Lexan ring 459801485251 is on stock, need check if Turkey site having contrast intrusion issue because of lexan ring gap and recommend to replace accordingly -- Andy | [12 empty] | 117916099 |  | 1130919 | TRUE | Closed - Done | Closed - Done |  | ECR-075975 | [5 empty] | Engineering | [6 empty] | Lexan Ring | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Lexan Ring | Gantry-HW-Lexan Ring | Gantry-HW-Lexan Ring | ECR-075975 | Lexan Ring | Gantry-HW | Gantry-HW-Lexan Ring | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | n | y | y | 0 | y | #N/A
1188 |  |  | United States | Doctors Hospital LLC | 40014 | 44399 |  | System Removal | Integration : Confirmation of order 53919742 saved ----------------------------------------------------------------------------- Work Order #: WO-03917663 22-Jul-2021 Lynn Matthew The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: n/a Device Use: n/a Expected and Unexpected Results: n/a User Impact: n/a Patient Impact: n/a Current Software Version: n/a Internal Comments : THis work order/case is for a life solutions quote, #LS - 070005 Customer has provided a cashiers check in the amount of $24197.16, per the life solutions quote. Diagnostic performed by Engineer : Customer requested entire system removal. A life solutions quote was generated and accepted by our customer Resolution : Assisted Philips installation/deinstallation personnel with disassembly and removal of CT system. ----------------------------------------------------------------------------- | 5 |  | Assisted Philips installation/deinstallation personnel with disassembly and removal of CT system. | [12 empty] | 117950736 |  | 1145559 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | 44401.184 | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1189 |  |  | United States | Odessa Regional Medical Center | 40029 | 44398 |  | MPR formats | Integration : Confirmation of order 53916107 saved ----------------------------------------------------------------------------- Work Order #: WO-03891810 Curry, Dyan The system fully meets specification and returned to use Problem description by Clinical Support Specialist : Customer Function/ Role: technologist Device Use: in use name Expected and Unexpected Results; help creating cor and sag of crooked elbow User Impact: none Patient Impact: none Current Software version: 4.1.7 Diagnostic performed by Engineer : helped customer use planar to straighten and create sag and cor mprs of twisted elbow. ----------------------------------------------------------------------------- | 5 |  | helped customer use planar to straighten and create sag and cor mprs of twisted elbow. | [12 empty] | 117925122 |  | 1137532 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | 44399.2743 | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1190 |  |  | United States | Philips Medical Capital | 40028 | 44396 |  | ACR questions | Integration : Confirmation of order 53904475 saved ----------------------------------------------------------------------------- Work Order #: WO-03878577 Giles Brian The system meets the specification for the performed service and is returned to use Problem description by Clinical Support Specialist : Customer Function/Role: TECH Device Use: In Use Expected and Unexpected Results: iDose question User Impact: None Patient Impact: None Current Software Version: 4.1.7 Diagnostic performed by Engineer : Customer wanted to know if iDose saved dose to the patient. It does and should be used. ----------------------------------------------------------------------------- Work Order #: WO-03877549 Mujica, David The system meets the specification for the performed service and is returned to use Problem description by Clinical Support Specialist : Customer Function/Role: Technologist Device Use: In use Expected and Unexpected Results: ACR questions User Impact: none Patient Impact: none Current Software Version: 4.1 Diagnostic performed by Engineer : Iris needed help gathering information for the ACR on a Head CT on their Ingenuity CT system. Number of channel, slice, increment, rotation and modulation explored. All of Iris_s questions were answered and the information was gathered for ACR on their Head CT exam card. Iris had no further questions or need for clinical support. No rescan, no loss of data, no reinjection per Iris. ----------------------------------------------------------------------------- | 5 |  | 44396 | [12 empty] | 117912371 |  | 1127499 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | 44397.0521 | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1191 |  |  | Chile | BANCO DE CHILE | 341005 | 44407 |  | CL1688-THE GANTRY WENT OFF TWICE THIS WEEK | Expected and Unexpected results: the gantry is turned off Patient Impact / Outcome: study aborat User Impact: the gantry is turned off when in use Current Software Version: Not provided Diagnostic performed by Engineer: they comment that gantry opens estop when being used in the study, now the equipment is operational at the time of the remote support call. The equipment is accessed remotely to inspect logs from the last 3 days without finding any record or reason for failure. I went to review together with the user team to see the reason for failure to open the estop. Check electrical supply is not having variations. SANDOVAL IS COMMUNICATED THAT NOBODY CAME ON FRIDAY TO THE SITE, FOLLOW-UP MAIL IS SENT | 3 |  | Next step: review together with the user team to see the reason for failure to open the estop. Check electrical supply is not having variations. | [12 empty] | 118006346 | In Progress | 1169625 | TRUE | Opened - Pending Translation | Opened - Pending Translation | [7 empty] | TBD | [6 empty] | Gantry off | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1192 |  |  | China | 东莞市中西医结合医院 | 345013 | 44409 |  | 02N/Ingenuity Core128/Elite China/nemo cable | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity Core128/Elite China Safety Questions : Case: 0118018928 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: nemo part Troubleshooting Action: FSE call in nemo DP to DVI cable. Repair Action: nemo DP to DVI cable. Test/Inspection: N/A Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | FSE call in nemo DP to DVI cable | [12 empty] | 118018928 | In Process | 1175422 | TRUE | Opened - Pending Translation | Opened - Pending Translation |  | CTUDT00632173 | [5 empty] | Engineering | [6 empty] | DVI Adaptor | [5 empty] | Refresh |  |  | System issue | System | Console-HW | Adaptor | Console-HW-Adaptor | Console-HW-Adaptor | CTUDT00632173 | Adaptor | Console-HW | Console-HW-Adaptor | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1193 |  |  | China | 河南省新乡市凤泉区大块镇卫生院 | 345009 | 44408 | 30 | 02n/Ingenuity Core128/Elite China/There is a problem with the program | BU Q&R Feedback: No SAP Customer Symptom Code: BOOT Customer Symptom Code: Boot up/Power-up Problem Reason: Corrective Maintenance Engineer Failure Code: Software Defect or Virus Repair Activity Code: Instructed Customer Problem Reported by Customer : There is a problem with the program Next Step Action : GUO GAIBIN call for,need RSE to deal Priority Safety Questions : Case: 0118015051 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Lili Last updated by: Lu Lili T2 Activities : Problem Description: CT application can not login Troubleshooting Action: CT application can not login, indicate need to create bugrep, instruct customer use olevel mode to login, It is ok, then use CT mode login , this time can login application successfully. system is ok now. Repair Action: instruct customer use olevel mode to login, It is ok, then use CT mode login , this time can login application successfully. system is ok now. Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: N/A SAP Transfer : Case Transferred to SAP | 2 |  | CT application can not login, indicate need to create bugrep, instruct customer use olevel mode to login, It is ok, then use CT mode login , this time can login application successfully. system is ok now. Any screenshot and bugreport for the issue? Check Blue screen defect and if related to this issue? | [12 empty] | 118015051 | Fixed | 1172857 | TRUE | Opened - Pending Translation | Opened - Pending Translation |  | CTUDT00632553 | [5 empty] | Engineering | [6 empty] | Login failed | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Blue-Screen | Console-SW-Blue-Screen | Console-SW-Blue-Screen | CTUDT00632553 | Blue-Screen | Console-SW | Console-SW-Blue-Screen | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1194 |  |  | China | 山东省胸科医院 | 40025 | 44408 |  | 02N/Ingenuity CT 728326/estop open | BU Q&R Feedback: No SAP Customer Symptom Code: BOOT Customer Symptom Code: Boot up/Power-up Problem Reason: Corrective Maintenance Engineer Failure Code: Electrical/Electronic Failure Repair Activity Code: Reboot/Restart Problem Reported by Customer : 02N/Ingenuity CT 728326 Safety Questions : Case: 0118009760 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang T2 Activities : Problem Description: estop open Troubleshooting Action: 1,check error logger, velocity encoder error cause the estop open. 2,instruct customer power off gantry power, then power on ,system is ok. Repair Action: reboot gantry power ,system is ok. Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | 1,check error logger, velocity encoder error cause the estop open. 2,instruct customer power off gantry power, then power on ,system is ok. Repair Action: reboot gantry power ,system is ok. | [12 empty] | 118009760 | Closed | 1172131 | TRUE | Opened - Pending Translation | Opened - Pending Translation |  | TBD-Encoder | [5 empty] | Engineering | [6 empty] | velocity encoder | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | VelocityEncoder | Couch-HW-VelocityEncoder | Couch-HW-VelocityEncoder |  |  | Couch-HW | Couch-HW-VelocityEncoder | No | 0 | 0 | 0 | 0 | 0 | No | y | y | y | 0 | 0 | y | y | y | 0 | y | #N/A
1195 |  |  | China | 河南省新乡市凤泉区大块镇卫生院 | 345009 | 44408 | 30 | 02N/Ingenuity Core128/Elite China/Lifting fault | T2 Activities : Problem Description: 客户问题概述： 其他类型故障。 客户问题详细描述或故障现象： 床降到最低位后升不起来了 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 床降到最低位后升不起来了 远程诊断与测试结果： 怀疑压到限位开关了 PRS TR状态： 不通，请工程师到场检查。 Repair Action: 建议现场维修步骤： 维修模式，把床升起来，重做垂直校正。 External Remarks : 31.07.2021 10:52:11 UTC Haibin Guo (86038308)"SUAT. Ready for Clinical Use local time" : "2021-07-31 16:00" Internal Remarks : 01.08.2021 05:46:24 UTC Haibin Guo (86038308) RDI Accuracy: YES Q7:Confirmed report problem，Lifting fault $Q8:perform visiual check on site and conform the couch can not lift up. when switch to service mode, then the couch can move up normally. then we switch to normal mode, the couch can lift up normaly too $Q9:move up the couch in service mode and do vertical calibration $Q10: test and inspection data refer to PA. Device returned to full functionality. $Q11:Calibrated Tools N/A External Remarks : 01.08.2021 05:46:18 UTC SUP JCOUSER (86038308)故障描述：床在最低位置升不起来 服务内容：使用服务模式升床并重新完成垂直校正 服务结果：服务完成后，设备恢复正常使用 其他事项：无 SAP Transfer : Dispatched for Onsite Support | 2 |  | perform visiual check on site and conform the couch can not lift up. when switch to service mode, then the couch can move up normally. then we switch to normal mode, the couch can lift up normaly too move up the couch in service mode and do vertical calibration Any bugreport and detail for this issue? | [12 empty] | 118008848 | Closed | 1172026 | TRUE | Opened - Pending Translation | Opened - Pending Translation |  | TBD-bugrep | [5 empty] | Engineering | [6 empty] | Couch fault | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | TBD-bugrep | Couch-HW-TBD-bugrep | Couch-HW-TBD-bugrep |  |  | Couch-HW | Couch-HW-TBD-bugrep | No | 0 | 0 | 0 | 0 | 0 | No | y | y | y | 0 | 0 | y | y | y | 0 | y | #N/A
1196 |  |  | China | 东莞市中西医结合医院 | 345013 | 44407 |  | wechat:02n/Ingenuity Core128/Elite China/system error | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : system error Safety Questions : Case: 0118000001 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Lili Last updated by: Lu Lili T2 Activities : Problem Description: 客户问题概述： ECG相关故障。 客户问题详细描述或故障现象： 心电无波形信号。 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： log无报错。 远程诊断与测试结果： 重启无效。 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议现场检查MP5到机架接线，用MP5模拟波形测试MP5是否信号能到IBOX板，如果能到检查firmware IBOX重新尝试连接。 建议备件12NC及名称： N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : li zongxian ----------------------------------------------------------------------------- | 3 |  | 心电无波形信号。建议现场检查MP5到机架接线，用MP5模拟波形测试MP5是否信号能到IBOX板，如果能到检查firmware IBOX重新尝试连接。 Check if USB cable have 5vdc power cable, or Bizlink USB cable. | [12 empty] | 118000001 | In Process | 1171215 | TRUE | Opened - Pending Translation | Opened - Pending Translation | [7 empty] | Others | [6 empty] | ECG | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | n | y | y | y | y | #N/A
1197 |  |  | China | 苏州广慈肿瘤医院有限公司 | 345019 | 44407 |  | 02N/Ingenuity Core128/Elite China | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity Core128/Elite China Next Step Action : FSE xu haotian 018761995321来电报修设备 Safety Questions : Case: 0117998855 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 其他类型故障。 客户问题详细描述或故障现象： 扫描中断 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： S_MASTER_SEQUENCER_TIMER_SETTIME_ERR 远程诊断与测试结果： DMC，RHOST，Ghost通讯故障 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 检查滑环，信号刷，检查高压之间的连线 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : Xu Haotian External Remarks : 01.08.2021 09:22:02 UTC Haotian Xu (86090280) 服务内容：做增强检测时报错，报错代码是sequence_timer_setting_error，检查 hss上的接口，收集bugrep 服务结果：服务后，设备恢复正常 ----------------------------------------------------------------------------- | 3 |  | 远程诊断与测试结果： DMC，RHOST，Ghost通讯故障 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 检查滑环，信号刷，检查高压之间的连线 做增强检测时报错，报错代码是sequence_timer_setting_error，检查 hss上的接口，收集bugrep | [12 empty] | 117998855 | In Process | 1170882 | TRUE | Opened - Pending Translation | Opened - Pending Translation | [7 empty] | TBD | [6 empty] | Data path | [5 empty] | Refresh |  |  | System issue |  |  |  | - | - |  |  | 0 | - | No | y | y | y | 0 | 0 | No | y | y | y | y | 0 | n | y | y | y | y | #N/A
1198 |  |  | China | 息烽县中医医院 | 345014 | 44404 |  | Wechat:02N/Ingenuity Core128/Elite China/ACQ error | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Replace Part Problem Reported by Customer : ACQ error Safety Questions : Case: 0117970209 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Gaoge Last updated by: Wang Gaoge 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 扫描失败。 客户问题详细描述或故障现象： 扫描时报错gantry acquisition system cannot comply 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： S_MASTER_SEQUENCER_ACQ_NOT_OK 远程诊断与测试结果： CIRS日志中显示光路重发数量为0 ，报错时间有ERROR - Invalid DFS setting ，考虑R2D频率导致采集异常 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 1，检查R2D的频率和幅值 ，偶发报错，需要仔细诊断和排查 建议备件12NC及名称： "R2D Assembly 459800467343" Test/Inspection: N/A Test Equipment: N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : PPD 459800467341*1 已确认替代号可出24h到医院 FSE:唐广林-86-18084189777 External Remarks : 29.07.2021 10:03:50 UTC SUP JCOUSER (86029494)服务内容：确认报修故障，扫描中断。现场检查，发现R2D故障，更换R2D后设备恢复正常。 服务结果：维修后设备恢复正常。 ----------------------------------------------------------------------------- | 3 |  | 服务内容：确认报修故障，扫描中断。现场检查，发现R2D故障，更换R2D后设备恢复正常。 William: check if new 459800467344 can improve this issue? | [12 empty] | 117970209 | In Process | 1157518 | TRUE | Opened - Pending Translation | Opened - Pending Translation |  | CTUDT00630864 | [5 empty] | Engineering | [6 empty] | R2D | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | R2D | Gantry-HW-R2D | Gantry-HW-R2D | CTUDT00630864 | R2D | Gantry-HW | Gantry-HW-R2D | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1199 |  |  | China | 哈尔滨市森海医院 | 345007 | 44403 |  | 02N/Ingenuity Core128/Elite China/printer issue | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ Safety Questions : Case: 0117961156 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Cao Last updated by: Liu Cao 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: Cannot print Troubleshooting Action: Print connection failed, temp network issue Repair Action: Temporary network issue , take monitoring Test/Inspection: N/A Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | Problem Description: Cannot print Troubleshooting Action: Print connection failed, temp network issue Repair Action: Temporary network issue , take monitoring | [12 empty] | 117961156 | Fixed | 1151904 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Network Connection | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | n | y | y | y | y | #N/A
1200 |  |  | China | Zhalantun City People_s Hospital | 40045 | 44403 |  | 02N/Ingenuity CT 728326/scan interrupt | BU Q&R Feedback: No SAP Customer Symptom Code: ACQP Customer Symptom Code: Imaging/Patient Recording/Acquisition Problem Reason: Corrective Maintenance Engineer Failure Code: Electrical/Electronic Failure Repair Activity Code: Reboot/Restart Problem Reported by Customer: 02N/Ingenuity CT 728326 Safety Questions:Case: 0117960989 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: scan interrupt Troubleshooting Action: check error logger, ACQ not ok, check eventlogger ,no logger found, this error only occurred once, so need monitor the system_s status. Repair Action: instruct customer monitor the system status. now system is ok. Test/Inspection: Functional Check PASSED. The device returned to full functionality. Test Equipment: N/A SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | Troubleshooting Action: check error logger, ACQ not ok, check eventlogger ,no logger found, this error only occurred once, so need monitor the system_s status. Repair Action: instruct customer monitor the system status. now system is ok. | [12 empty] | 117960989 | Fixed | 1151827 | TRUE | Investigation | Investigation |  | CTUDT00630408 | [5 empty] | Engineering | [6 empty] | ACQ | [5 empty] | Refresh |  |  | System issue | System | CIRS-SW | ACQ_NOT_OK | CIRS-SW-ACQ_NOT_OK | CIRS-SW-ACQ_NOT_OK | ACQ_NOT_OK | ACQ_NOT_OK | CIRS-SW | CIRS-SW-ACQ_NOT_OK | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | n | y | y | 0 | y | #N/A
1201 |  |  | Japan | 社会福祉法人恩賜財団済生会 | 343000 | 44408 |  | 01CS: RS: Axial of the head is noisy when the coronal cross section is created by ISP | Canceled | 5 |  | Canceled | [12 empty] | 118016702 | New | 1173973 | TRUE | Opened | Opened | [7 empty] | Others | [6 empty] | Image Quality | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1202 |  |  | Japan | 医療法人清翠会 牧病院 | 341017 | 44408 |  | SY: CT application support package | Canceled | 5 |  | Canceled | [12 empty] | 118016430 | Canceled | 1173747 | TRUE | Opened | Opened | [7 empty] | Application | [6 empty] | Application | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1203 |  |  | Japan | 社会医療法人 母恋 日鋼記念病院 | 341011 | 44408 |  | 02DS: KK: Tape Switch pressed and cannot scan. >> Recovered by releasing Tape Switch. ＞＞ FSE compatible | 02DS: KK: Tape Switch pressed and cannot scan. >> Recovered by releasing Tape Switch. ＞＞ FSE compatible | 4 |  | Tape Switch | [12 empty] | 118016314 | New | 1173636 | TRUE | Opened | Opened | [7 empty] | Customer | [6 empty] | Tape Switch | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1204 |  |  | Japan | 社会福祉法人恩賜財団済生会 | 343000 | 44408 |  | HY: Gantry belt replacement | HY: Gantry belt replacement | 5 |  | Belt | [12 empty] | 118014821 | New | 1172681 | TRUE | Opened | Opened |  | ECR-077012 | [5 empty] | Engineering | [6 empty] | Belt | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Belt | Gantry-HW-Belt | Gantry-HW-Belt | ECR-077012 | Belt | Gantry-HW | Gantry-HW-Belt | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1205 |  |  | Japan | 社会福祉法人恩賜財団済生会 | 343000 | 44404 |  | 01CS: RS: The image of the coronal cross section of the Axial of the head created by ISP became noisy >> It seems that the image of the soft tissue condition and the bone condition were created together >> The bone condition was recalculated and resolved | 01CS: RS: Moriwaki-sama: The image of the Axial of the head created with the coronal cross section by ISP became noisy >> It seems that the image of the soft tissue condition and the bone condition were created together >> The bone condition is recalculated separately Eliminate --Problem Description-- Target device: CT Failure location: 01Cosole When did it first occur? (Within ~): 1 Hour How often do problems occur ?: First Occurrence At what point in the end-user workflow is the problem occurring ?: Other What happened recently ?: N.A. Does the customer create a specific system event report for the problem ?: N.A. --4 / 8Elements-- Q1 Did you confirm the reported problem? Verified Q2 Did you confirm other problems? None Q3 What did you do to solve the problem? It seems that it was created by combining the image of the soft tissue condition and the bone condition >> Bone condition is divided and recalculated and eliminated Is the device after Q4 compatible? Good Q5 How did the inspection end? None Q6 Presence or absence of incorrect findings (misdiagnosis, mistreatment) due to problems: None Q7 Presence or absence of unintended unnecessary exposure: None Q8 Presence or absence of artifacts: None 与此原文有关的更多信息要查看其他翻译信息，您必须输入相应原文 | 4 |  | It seems that it was created by combining the image of the soft tissue condition and the bone condition >> Bone condition is divided and recalculated and eliminated | [12 empty] | 117977319 | Fixed | 1161155 | TRUE | Closed - No Further Investigation | Closed - No Further Investigation |  | ShouldbeApplication | [5 empty] | Engineering | [6 empty] | Image Quality | [5 empty] | Refresh |  |  | System issue | System | NotEngineering | ShouldbeApplication | NotEngineering-ShouldbeApplication | NotEngineering-ShouldbeApplication | ShouldbeApplication |  | NotEngineering | NotEngineering-ShouldbeApplication | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | #N/A
1206 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 337000 | 44405 |  | Mylar order. | "Work Order #: WO-03944507 Can Ozgenlik The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer: Mylar order. Resolution : The Mylar cover which connects the gantry rear door and front door will be replaced with a newer model. This will prevent liquid leak into the inside, such as blood, opaque substance, etc. Device is operational. Follow Up Required: Mylar replacement. | 3 |  | The Mylar cover which connects the gantry rear door and front door will be replaced with a newer model. This will prevent liquid leak into the inside, such as blood, opaque substance, etc. | [12 empty] | 117979182 | In Progress | 1157769 | TRUE | Evaluation | Evaluation |  | ECR-075975 | [5 empty] | Engineering | [6 empty] | Lexan ring | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Lexan Ring | Gantry-HW-Lexan Ring | Gantry-HW-Lexan Ring | ECR-075975 | Lexan Ring | Gantry-HW | Gantry-HW-Lexan Ring | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1207 |  |  | United States | Westgate Hospital | 40001 | 44407 |  | Add A New Pacs Destination | Internal Remarks : Planner Initials: JJ Action Taken: FSE-Peter Leon called in to create a Case/WO and dispatch to him I created Case#0118006668/WO-03967361 and dispatched as requested Outcome: Dispatched Next Steps: N/A Last Updated: 1:54PM EDT 7/30/21 ----------------------------------------------------------------------------- Work Order #: WO-03967361 Leon Peter Diagnostic performed by Engineer : N/A Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: ----------------------------------------------------------------------------- | 5 |  | Add A New Pacs Destination | [12 empty] | 118006668 | In Progress | 1171663 | TRUE | NAC | NAC | [7 empty] | Customer | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1208 |  |  | Japan | 医療法人清翠会 牧病院 | 341017 | 44404 |  | CT:APP:RK:質問:ﾎｼさま | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Application Support Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer : Q:"①たぶん、体動によるものであると思いますが画質がいつもと若干異なるので確認してください。 ②画像を転送をすると、転送先が前回の設定を覚えてしまうのですが覚えないようにできますか？" A:"①画像を拝見しました。やはり体動が原因で画像が変わっていました。 ②申し訳ありません。前回の転送先を覚えてしまう仕様になっておりまして変更できません。" OneEMS internal remarks : Q1:報告された問題は確認できましたか？：確認済 Q2：他の問題は確認できましたか？；無 Q3：問題を解決するために何を行いましたか？ １：電話で解決した場合：説明 Q4：対応後の装置は順調ですか？：良好 Safety Questions : Case: 0117972537 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: RYUSUKE KUJIRAI Last updated by: RYUSUKE KUJIRAI 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | Zhu, Chun |  | 44404 | [12 empty] | 117972537 | Fixed | 1158116 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | 44405.7771 | [6 empty] | 44404 | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1209 |  |  | United States | Unity Health Systems | 40037 | 44407 |  | Password incorrect. | Integration : Confirmation of order 53957845 saved ----------------------------------------------------------------------------- Work Order #: WO-03966342 30-Jul-2021 Lockwood Thomas The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : Reset IT password. Created updated password for master user to edit Exam cards ----------------------------------------------------------------------------- | 5 |  | Password | [12 empty] | 118004053 | Closed | 1169693 | TRUE | Opened | Opened | [7 empty] | Customer | [6 empty] | Password | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1210 |  |  | United States | Philips Medical Capital | 40028 | 44407 |  | Split Study | Work Order #: WO-03967829 Mujica, David The system meets the specification for the performed service and is returned to use Problem description by Clinical Support Specialist : Customer Function/Role: Technologist Device Use: In use Expected and Unexpected Results: Split Accession numbers User Impact: none Patient Impact: none Current Software Version: 4.1 Diagnostic performed by Engineer : Iris indicated that her multi-study CT were not splitting to their appropriate accession numbers. Her Exam Card was accessed and Iris was shown where to make the association to change her Procedure in the Exam Card/Protocol to associate each accession number to the procedure description and save the Exam Card/Protocol so all accession number associated to each study accordingly. The EC was successfully changed and saved. Iris had no further questions or need for clinical support. No rescan, no loss of data, no reinjection per Iris. ----------------------------------------------------------------------------- | 4 |  | change her Procedure in the Exam Card/Protocol to associate each accession number to the procedure description and save the Exam Card/Protocol so all accession number associated to each study accordingly. The EC was successfully changed and saved. | [12 empty] | 118007049 | In Progress | 1169328 | TRUE | Opened | Opened | [7 empty] | Application | [6 empty] | Exam cards | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1211 |  |  | United States | Philips Medical Capital | 40053 | 44405 |  | 0802CDY72 no ecg signal showing on host monitor | Internal Remarks : FADB: (FSE Aware of Double Booking) Title: Internal Remarks Initials: Twila Actions Taken: Called PFSE/McDowell Preston - He accepted WO for 7/29/21 @ 5hr Labor - No Travel Time I Dispatched WO for 1-6pm EST. Outcome: Dispatched per PFSE_s Request IP: 88993323 Next Steps: None Internal Remarks : Assisted FSE with Flashing replacement iBox 2. Internal Remarks : FADB: (FSE Aware of Double Booking) Title: Internal Remarks Initials: Twila Actions Taken: Dispatched per WO FU Note - Recommended Service Engineer: McDowell Preston Follow Up Date/Time Requested: 8/2/2021 3:00 PM EST. Outcome: Dispatched per WO FU Note IP: 88993323 Next Steps: None ----------------------------------------------------------------------------- Work Order #: WO-03959714 29-Jul-2021 McDowell Preston Problem Description by Engineer : Problem description by engineer : no ecg signal showing on host monitor Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No further information is available as per the customer Expected and Unexpected results: no ecg signal showing on host monitor Patient Impact/Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer Current Software Version: Not provided Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- Work Order #: WO-03950575 29-Jul-2021 McDowell Preston The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Problem description by engineer : no ecg signal showing on host monitor Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No further information is available as per the customer Expected and Unexpected results: no ecg signal showing on host monitor Patient Impact/Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer Current Software Version: Not provided Follow Up Required : replaced patient monitor interface cable. Programmed the Ibox board and verified proper readings at the control desk . At this time, it is working, but will monitor the issue and await other part to arrive from overseas before case closure. Diagnostic performed by Engineer : replaced patient monitor interface cable. Programmed the Ibox board and verified proper readings at the control desk . At this time, it is working, but will monitor the issue and await other part to arrive from overseas before case closure. ----------------------------------------------------------------------------- Work Order #: WO-03959547 29-Jul-2021 McDowell Preston The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : Troubleshot with NSS as well. I box internal tests pass just fine, but it still shows ECG leads disconnected. Will order new cables for Thursday. Customer can utilize scanner for other patients without the patient monitor. ----------------------------------------------------------------------------- Work Order #: WO-03950493 Sheffield, Scott The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Problem description by engineer : no ecg signal showing on host monitor Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No further information is available as per the customer Expected and Unexpected results: no ecg signal showing on host monitor Patient Impact/Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer Current Software Version: Not provided Diagnostic performed by Engineer : Ongoing problem with the ecg waveform not showing on host monitor. Follow Up Required : Error # and/or description of error : none Repair Action : I spoke with primary FSE, he is already talking with NSS about this issue and FSE will be ordering parts himself. Technician Name : Sheffield, Scott Internal Comments : Error # and/or description of error : none Repair Action : I spoke with primary FSE, he is already talking with NSS about this issue and FSE will be ordering parts himself. ----------------------------------------------------------------------------- | 3 |  | replaced patient monitor interface cable. Programmed the Ibox board and verified proper readings at the control desk . At this time, it is working, but will monitor the issue and await other part to arrive from overseas before case closure. | [12 empty] | 117985264 | In Progress | 1160487 | TRUE | Evaluation | Evaluation | [7 empty] | Others | [6 empty] | ECG | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1212 |  |  | Japan | Maki Hospital | 341017 | 44403 |  | SY: CT application support package. | BU Q&R Feedback: No Reason: Application Support Problem Reported by Customer: SY: CT application support package. Safety Questions:Case: 0117961832 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: YE Smile Last updated by: YE Smile 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A OneEMS internal remarks : Otani DB SAP Transfer: Case Transferred to SAP SAP Transfer: Dispatched for Onsite Support Internal Remarks : 29.07.2021 04:27:13 UTC NATSUKI OOTANI (81100565) 1. Are patients with COVID-19 infection or patients who are suspected of having COVID-19 infection being treated or quarantined ? Yes 2. Is it necessary to enter the treatment room or the bed of infected patients?(Including workplace and transfer routes) No External Remarks: 29.07.2021 04:27:10 UTC NATSUKI OOTANI (81100565) [Request details] A study group meeting was held about cardiac CT basics. Work details: I explained about cardiac CT using slides. ----------------------------------------------------------------------------- | Xu, Xiaoting |  | 44403 | [12 empty] | 117961832 | New | 1152384 | TRUE | Closed - Non-Complaint | Closed - Non-Complaint | [7 empty] | 44404.6201 | [6 empty] | 44403 | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1213 |  |  | United States | Unity Health Systems | 40037 | 44405 |  | Monthly QA | Integration : Confirmation of order 53945608 saved ----------------------------------------------------------------------------- Work Order #: WO-03947431 28-Jul-2021 Taylor, Durrance The system meets the specification for the performed service and is returned to use Resolution : Monthly QA completed Problem Description by Engineer : Customer Function/Role: tech Device Use: scheduled activity Expected and Unexpected Results: Monthly QA User Impact: none Patient Impact: none Current Software Version: 4.1.7 | 5 |  | Monthly QA | [12 empty] | 117982135 | Closed | 1160277 | TRUE | NAC | NAC | [7 empty] | Customer | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1214 |  |  | United States | Ascension Seton Medical Center | 40006 | 44404 | 30 | STC Failing Error during scan | Work Order #: WO-03940410 27-Jul-2021 Beato Tom The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: No Patient Involved Expected and Unexpected Results: STC Failed Error during scan User Impact: No Patient Impact: No Current Software Version: 4.1.10 Resolution : Attempted to perform short tube conditioning on system. Scan aborted. Performed complete shutdown of PCs and gantry. Restarted system. Attempted to scan but CIRS not initialized. Waited for CIRS to initialize. Able to scan after 10 minute wait. Performed short tube conditioning successfully. Analyzed bug report. Error occurred because gantry was not restarted in 80 plus days and scan was attempted without CIRS initialized. System is fully operational. ----------------------------------------------------------------------------- Work Order #: WO-03940262 Maier, Donnie The system does not meet the specification for the performed service and is not returned to use Diagnostic performed by Engineer : Cu stated they had already power cycled everything a few time. ULV shows error below before and after I talked Bonnie thru a complete power cycle of the system (Host/CIRS/Gantry) and still unable to perform STC. Unable to pull logs to troubleshoot further, dispatching FSE to analyze CIRS event logs, etc. Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: No Patient Involved Expected and Unexpected Results: STC Failed Error during scan User Impact: No Patient Impact: No Current Software Version: 4.1.10 Follow Up Required : Analyze logs locally as I am unable to pull them remotely. I already talked cu thru a complete power cycle (Host/CIRS/Gantry) but still getting the error during STC. Gantry hadn_t been power cycled in over 80 days according to ULV. ----------------------------------------------------------------------------- | 2 |  | Error occurred because gantry was not restarted in 80 plus days and scan was attempted without CIRS initialized. System is fully operational. Instruct customer to restart whole system weekly | [12 empty] | 117975026 | Closed | 1154684 | TRUE | Evaluation | Evaluation | [7 empty] | Customer | [6 empty] | Reboot | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1215 |  |  | United States | Philips Medical Capital | 40053 | 44403 |  | EKG disconnected intermittently, replaced I-Box | Integration : Confirmation of order 53932202 saved ----------------------------------------------------------------------------- Work Order #: WO-03934417 27-Jul-2021 McDowell Preston The system meets the specification for the performed service and is returned to use Resolution : THe patient monitor readings have been intermittent. Per NSS, replaced the Ibox board today and verified proper operation...though it was tested prior to replacement and found working as well. Will monitor system for issues, and inform customer. Problem Description by Engineer : Description of Event Still have intermittent EKG connectivity problem Ref Case 0117870146 CALLER FUNCTION/ROLE Technician/Sonographer/Therapist DEVICE INTENDED USE Diagnostic <1 Hour Intermittent WHERE IN THE END-USER WORKFLOW IS THE ISSUE OCCURRING? Pre-acquisition/scan WHAT RECENT EVENTS HAVE OCCURRED? During normal use NEGATIVE OPERATOR IMPACT N/A NEGATIVE PATIENT IMPACT/PROCEDURE OUTCOME N/A HAS CUSTOMER CREATED A SPECIFIC SYSTEM EVENT REPORT OF ISSUE? No System Release Version 3.6.8 Parts Consumed : 1.00000,459800166651,I-BOX-2 BASE BOARD ASSY ----------------------------------------------------------------------------- Work Order #: WO-03934258 Weaver David The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Description of Event Still have intermittent EKG connectivity problem Ref Case 0117870146 CALLER FUNCTION/ROLE Technician/Sonographer/Therapist DEVICE INTENDED USE Diagnostic <1 Hour Intermittent WHERE IN THE END-USER WORKFLOW IS THE ISSUE OCCURRING? Pre-acquisition/scan WHAT RECENT EVENTS HAVE OCCURRED? During normal use NEGATIVE OPERATOR IMPACT N/A NEGATIVE PATIENT IMPACT/PROCEDURE OUTCOME N/A HAS CUSTOMER CREATED A SPECIFIC SYSTEM EVENT REPORT OF ISSUE? No System Release Version 3.6.8 Diagnostic performed by Engineer : Still have intermittent EKG connectivity problem Ref Case 0117870146. Follow Up Required : Verify iBox 5v power at board is >5v, suggest 5.15v. IF PART REQUIRED - 12 NC RECOMMENDED: 459800484881 - PNEUMO CHEST BELLOWS 459800484391 - BELLOWS PULMO CABLE This is the Pulmo Digital assembly, attaches to gantry and the Chest bellows attach to this: 459800484361 - PULMO TOOL KIT (SG164) Pulmo Bellows Assy Replacement Kit 455011503261 ----------------------------------------------------------------------------- | 3 |  | THe patient monitor readings have been intermittent. Per NSS, replaced the Ibox board today and verified proper operation...though it was tested prior to replacement and found working as well. Will monitor system for issues, and inform customer. | [12 empty] | 117966766 | Closed | 1150710 | TRUE | Closed - Done | Closed - Done | [7 empty] | Others | [6 empty] | ECG | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1216 |  |  | United States | Philips Medical Capital | 300057 | 44405 |  | Error when trying to do tube warm up | Integration : Confirmation of order 53944768 saved ----------------------------------------------------------------------------- Work Order #: WO-03946309 28-Jul-2021 Hartowicz, Ethan The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: CT will not scan, CIRS error Device Use: CIRS used for reconstruction Expected and Unexpected Results: system would not perform short tube User Impact: system can not scan Patient Impact: NA Current Software Version: 4.1.10.25208 Resolution : Checked errors and logs and found inconsistent 4 channel connection. 0x4321 -> 0x0. I found loose fiber connections on the MOOG. Secured these connections and system returned to normal working condition. ----------------------------------------------------------------------------- Work Order #: WO-03946187 Powell, Jamie The system does not meet the specification for the performed service and is not returned to use Diagnostic performed by Engineer : Recurring system not ready error when scanning. Physical channel order: 0x0 Issue might be with transmitter or power supply.. Check power supply. Check for errors on transmitter.. Check optical.. Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Follow Up Required : Recurring system not ready error when scanning. Physical channel order: 0x0 Issue might be with transmitter or power supply.. Check power supply. Check for errors on transmitter.. Check optical.. MODULE, ROTOR ELECTRONICS TX 453567538334 Internal Comments : Spoke to FSE John Williams who states another FSE Ethan is already on his way to the site to troubleshoot the issue. ----------------------------------------------------------------------------- |  |  | Fiber, check if have fiber cable loose picture. | [12 empty] | 117980984 | Closed | 1158740 |  |  | Evaluation | [20 empty] | Non-Refresh |  |  | Upgrade issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | #N/A
1217 |  |  | United States | Select Specialty Hospitals | 310051 | 44403 |  | QO: need positioning Pads ordered | Internal Remarks : FADB: (FSE Aware of Double Booking) Title: Internal Remarks Initials: Twila Actions Taken: Per message sent in Teams from FSE/Swett, Mike - Create Case and Dispatch WO for 7/27/21 @ 10-11am CST Travel Time 1hr each way Outcome: Created Case and Dispatched WO per FSE_s Request IP: 56915900 Next Steps: None Integration : Confirmation of order 53932767 saved ----------------------------------------------------------------------------- Work Order #: WO-03935087 27-Jul-2021 Swett Mike The system meets the specification for the performed service and is returned to use Diagnostic performed by Engineer : replaced head holder pads Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Parts Consumed : 1.00000,459800033851,Cushion- Small Wedge 1.00000,459800033821,Cushion - Flat 1.00000,459800033841,Cushion- Medium Wedge 1.00000,459800033831,Cushion- Large Wedge ----------------------------------------------------------------------------- |  |  | 44403 | [12 empty] | 117967451 | Closed | 1155705 |  |  | Closed - Non-Complaint | [20 empty] | Non-Refresh |  |  | Upgrade issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | #N/A
1218 |  |  | United States | Select Specialty Hospitals | 310051 | 44403 |  | LS Parts Quote |  |  |  | 44403 | [12 empty] | 117965767 | New | 1154651 |  |  | Evaluation | [20 empty] | Non-Refresh |  |  | Upgrade issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | #N/A
1219 |  |  | United States | Christ Hospital | 300205 | 44397 |  | 0711CDN85 windows error message | Work Order #: WO-03885044 27-Jul-2021 Hanlin Christopher The system meets the specification for the performed service and is returned to use Tools & Equipment : MULTIMETER,2022-06-30 00:00:00 Resolution : 4.1.10 software reload needed. Reloaded 4.1.10 software. Discovered missing FCO patch disk when restoring software and had to return to finalize software configuration. Problem Description by Engineer : Problem description by engineer : windows error message When was the first occurrence? Within... : 20-Jul-2021 How often is issue occurring? : Continuously Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: No further information is available as per the customer Patient Impact/Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer Current Software Version: Not provided ----------------------------------------------------------------------------- Work Order #: WO-03885022 Atz, Michael The system does not meet the specification for the performed service and is not returned to use Follow Up Required : Error # and/or description of error : windows error message Repair Action : windows error message customer has rebooted several times and the issue remains Technician Name : Atz, Michael Internal Comments : Error # and/or description of error : windows error message Repair Action : windows error message customer has rebooted several times and the issue remains Problem Description by Engineer : Problem description by engineer : windows error message When was the first occurrence? Within... : 20-Jul-2021 How often is issue occurring? : Continuously Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: No further information is available as per the customer Patient Impact/Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer Current Software Version: Not provided Diagnostic performed by Engineer : windows error message ----------------------------------------------------------------------------- |  |  | Check with filed to collect log and screen to support investigation | [12 empty] | 117918979 | Incomplete | 1132321 |  |  | Evaluation |  | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | SW error | [5 empty] | Refresh |  |  | Upgrade issue | System | TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo |  |  | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | No | 0 | 0 | 0 | 0 | 0 | No | y | y | y | 0 | 0 | y | y | y | 0 | y | #N/A
1220 |  |  | United States | Select Specialty Hospitals | 300037 | 44393 |  | 0699CD72 getting compensator cracked warnings | Internal Remarks : Assisted FSE with false Cracked Compensator errors during Air Calibrations. Collimator was recently replaced. Suggested FSE submit a CFR and collect a bug report. Integration : Confirmation of order 53897942 saved ----------------------------------------------------------------------------- Work Order #: WO-03869287 24-Jul-2021 Beard Michael The system meets the specification for the performed service and is returned to use Resolution : Inspected system. Performed a compensator analysis. Result show no cracks in the collimator. Will follow up with NSS to find a solution for why the system is giving a crack in the collimator error message during air cal. Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: compensator messages User Impact: System functional Patient Impact: None Current Software Version: 4.8 ----------------------------------------------------------------------------- Work Order #: WO-03869107 Karlen James The system meets the specification for the performed service and is returned to use Follow Up Required : check with TSS on messages and compensator test for Windows 10. These may be false messages Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: compensator messages User Impact: System functional Patient Impact: None Current Software Version: 4.8 Diagnostic performed by Engineer : cracked compensator messages on W10 system. can not remotely test compensator. Collimator was recently replaced. Suggest escalating to TSS ----------------------------------------------------------------------------- |  |  | Check with L3 to sync the latest status of the system. Silva check the required test for collimator replacement Beam and DMS CHECK is used to replace conpenstator crach test | [12 empty] | 117903846 | Closed | 1121245 |  |  | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Compensator | [5 empty] | Refresh |  |  | Upgrade issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | n | y | y | y | y | #N/A
1221 |  |  | United States | Upson Regional Medical Center | 333009 | 44393 |  | 0711CDY73 "Blue Screen" Login Issue Post-Host Restart | Integration : Confirmation of order 53897121 saved ----------------------------------------------------------------------------- Work Order #: WO-03868162 19-Jul-2021 Lee John The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Problem description by engineer : 0711CDY73 "Blue Screen" Login Issue Post-Host Restart When was the first occurrence? Within... : 02-Jul-2021 08:00:00 How often is issue occurring? : Multiple times per month Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: 0711CDY73 "Blue Screen" Login Issue Post-Host Restart - Delayed Login Patient Impact/Outcome: None User Impact: Delayed Login Current Software Version: 4.1.10.25208 Resolution : Loaded software patches ----------------------------------------------------------------------------- Work Order #: WO-03868139 Holeman, David The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Problem description by engineer : 0711CDY73 "Blue Screen" Login Issue Post-Host Restart When was the first occurrence? Within... : 02-Jul-2021 08:00:00 How often is issue occurring? : Multiple times per month Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: 0711CDY73 "Blue Screen" Login Issue Post-Host Restart - Delayed Login Patient Impact/Outcome: None User Impact: Delayed Login Current Software Version: 4.1.10.25208 Diagnostic performed by Engineer : CT Tech reports for last couple of weeks when she performs a Host restart in the AM, when the HP z440 Host boots back up, she only sees "blue screens" and does not see login screen/box. Tech states she hast to proceed with multiple restarts of the Host and finally the login screen is present. Tech states there are 2 Ingenuity CT_s at site and both are exhibiting this issue. Internal Comments : Error # and/or description of error : 0711CDY73 "Blue Screen" Login Issue Post-Host Restart Repair Action : Complete the COPSSH User Activation on the CT Console located in the Ingenuity Family and Brilliance 64 v4.1.10 Software Installation Instructions. Inspect/configure the Host Windows Firewall settings to ensure full PRS/M2M Connectivity. Inspect functionality of restarting z440 Host and logging in. Recommend to proactively install FCO72800740 - FCO TITLE: SOFTWARE UPDATE FOR ING BR64 TO 4.1.10.XX210 Case Number0117119675 WO-03074011 Follow Up Required : Error # and/or description of error : 0711CDY73 "Blue Screen" Login Issue Post-Host Restart Repair Action : Complete the COPSSH User Activation on the CT Console located in the Ingenuity Family and Brilliance 64 v4.1.10 Software Installation Instructions. Inspect/configure the Host Windows Firewall settings to ensure full PRS/M2M Connectivity. Inspect functionality of restarting z440 Host and logging in. Recommend to proactively install FCO72800740 - FCO TITLE: SOFTWARE UPDATE FOR ING BR64 TO 4.1.10.XX210 Case Number0117119675 WO-03074011 Technician Name : Holeman, David ----------------------------------------------------------------------------- |  |  | Link to "Blue screen" defect CTUDT00632553 | [12 empty] | 117903119 | Closed | 1120460 |  |  | Closed - No Further Investigation |  | CTUDT00632553 | [5 empty] | Engineering | [6 empty] | Blue Screen | [5 empty] | Refresh |  |  | Upgrade issue | System | Console-SW | Blue-Screen | Console-SW-Blue-Screen | Console-SW-Blue-Screen | CTUDT00632553 | Blue-Screen | Console-SW | Console-SW-Blue-Screen | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1222 |  |  | United States | Upson Regional Medical Center | 310115 | 44393 |  | 0711CDY73 "Blue Screen" Login Issue Post-Host Restart | Internal Remarks : Caller: Susan Jaudon IP & Issue: 84628582/ When logging in - getting a blue screen Event Type: Incident Service Request: Remote Technical Did Customer Decline Remote Support? No Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Created Remote W/O Routed to RTAC Next Step Action: N/A Verbal PO Number: Verbal PO Approver: Verbal PO CAP: Integration : Confirmation of order 53896790 saved ----------------------------------------------------------------------------- Work Order #: WO-03868080 19-Jul-2021 Lee John The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Problem description by engineer : "Blue Screen" Login Issue Post-Host Restart When was the first occurrence? Within... : 02-Jul-2021 08:00:00 How often is issue occurring? : Multiple times per month Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: "Blue Screen" Login Issue Post-Host Restart Patient Impact/Outcome: None User Impact: Delayed Login Current Software Version: 4.1.10.25208 Resolution : Loaded backups ----------------------------------------------------------------------------- Work Order #: WO-03867221 Holeman, David The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Problem description by engineer : "Blue Screen" Login Issue Post-Host Restart When was the first occurrence? Within... : 02-Jul-2021 08:00:00 How often is issue occurring? : Multiple times per month Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: "Blue Screen" Login Issue Post-Host Restart Patient Impact/Outcome: None User Impact: Delayed Login Current Software Version: 4.1.10.25208 Diagnostic performed by Engineer : CT Tech reports for last couple of weeks when she performs a Host restart in the AM, when the HP z440 Host boots back up, she only sees "blue screens" and does not see login screen/box. Tech states she hast to proceed with multiple restarts of the Host and finally the login screen is present. Tech states there are 2 Ingenuity CT_s at site and both are exhibiting this issue. Follow Up Required : Error # and/or description of error : "Blue Screen" Login Issue Post-Host Restart Repair Action : Complete the COPSSH User Activation on the CT Console located in the Ingenuity Family and Brilliance 64 v4.1.10 Software Installation Instructions. Inspect/configure the Host Windows Firewall settings to ensure full PRS/M2M Connectivity. Inspect functionality of restarting z440 Host and loggin in. Recommend to proactively install FCO72800740 - FCO TITLE: SOFTWARE UPDATE FOR ING BR64 TO 4.1.10.XX210 Case Number 0117119674 WO-03074010 Technician Name : Holeman, David Internal Comments : Error # and/or description of error : "Blue Screen" Login Issue Post-Host Restart Repair Action : Complete the COPSSH User Activation on the CT Console located in the Ingenuity Family and Brilliance 64 v4.1.10 Software Installation Instructions. Inspect/configure the Host Windows Firewall settings to ensure full PRS/M2M Connectivity. Inspect functionality of restarting z440 Host and loggin in. Recommend to proactively install FCO72800740 - FCO TITLE: SOFTWARE UPDATE FOR ING BR64 TO 4.1.10.XX210 Case Number 0117119674 WO-03074010 ----------------------------------------------------------------------------- |  |  | Link to "Blue screen" defect CTUDT00632553 | [12 empty] | 117902663 | Closed | 1120413 |  |  | Evaluation |  | CTUDT00632553 | [5 empty] | Engineering | [6 empty] | Blue Screen | [5 empty] | Refresh |  |  | Upgrade issue | System | Console-SW | Blue-Screen | Console-SW-Blue-Screen | Console-SW-Blue-Screen | CTUDT00632553 | Blue-Screen | Console-SW | Console-SW-Blue-Screen | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1223 |  |  | United States | Christ Hospital | 300205 | 44390 |  | nemo | Integration : Confirmation of order 53882134 saved Integration : Confirmation of order 53882134 saved ----------------------------------------------------------------------------- Work Order #: WO-03843612 Parts Consumed : 1.00000,459801731761,USB EXTENDER CABLE FROM IBOX TO CONSOLE 1.00000,455300027201,CBL,USB A/M TO B/M 0.5M ----------------------------------------------------------------------------- | [15 empty] | 117879514 | Closed | 1112781 |  |  | Closed - Non-Complaint | [20 empty] | Non-Refresh |  |  | Upgrade issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | #N/A
1224 |  |  | United States | Christ Hospital | 300205 | 44390 |  | Gate to CT signal missing | Integration : Confirmation of order 53881648 saved ----------------------------------------------------------------------------- Work Order #: WO-03849843 Paine Matthew The system meets the specification for the performed service and is returned to use Diagnostic performed by Engineer : Customer states that EKG Waveform is Present on Monitor, but EKG Waveform is Not Present on Host PC Monitor. Problem Description by Engineer : Customer Function/Role: Clinical Engineer Device Use: Diagnostic Expected and Unexpected Results: Gate to CT signal missing User Impact: Unknown Patient Impact: Unknown Current Software Version:4.1.10 Resolution : Had Customer observe the Ibox 7-Segment Display_s while Heart Simulator was connected to EKG Monitor to ensure Ibox Functionality. The Displays changed as the Waveform did ruling out issue with Ibox. Had Customer check for connection of USB Hub Extender on Host PC Side. Customer found that USB Hub Extender had come Disconnected from Signal Distribution Board. Customer Connected USB Extender and now EKG Waveform Present on Host PC Monitor. ----------------------------------------------------------------------------- Work Order #: WO-03843020 Weaver David The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, no follow-up requested Diagnostic performed by Engineer : Gate to CT signal missing Problem Description by Engineer : Description of Problem Gate to CT signal missing CALLER FUNCTION/ROLE Biomed DEVICE INTENDED USE Diagnostic WHEN WAS THE FIRST OCCURRENCE? WITHIN . . . Same day HOW OFTEN IS ISSUE OCCURRING? Continuously WHERE IN THE END-USER WORKFLOW IS THE ISSUE OCCURRING? Pre-acquisition/scan WHAT RECENT EVENTS HAVE OCCURRED? During normal use NEGATIVE OPERATOR IMPACT No signal from gate to CT Host NEGATIVE PATIENT IMPACT/PROCEDURE OUTCOME Delay in scanning HAS CUSTOMER CREATED A SPECIFIC SYSTEM EVENT REPORT OF ISSUE? No 4.1.10 Resolution : Gate to CT signal missing ----------------------------------------------------------------------------- | [15 empty] | 117879056 | Closed | 1107668 |  |  | Closed - No Further Investigation | [20 empty] | Non-Refresh |  |  | Upgrade issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | #N/A
1225 |  |  | United States | AdventHealth Zephyrhills | 300106 | 44389 |  | 1302CDY73 problems sending images to pacs | Internal Remarks : I spoke to primary FSE Mark, he can not go today and he said to send this call to secondary FSE Garcia Marrero, Jorge. Integration : Confirmation of order 53874315 saved ----------------------------------------------------------------------------- Work Order #: WO-03833277 22-Jul-2021 Garcia Marrero, Jorge The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Problem description by engineer : problems sending images to pacs Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No further information is available as per the customer Expected and Unexpected results: problems sending images to pacs Patient Impact/Outcome: none User Impact: delay Current Software Version: 4.1.7.25030 Diagnostic performed by Engineer : System is having intermittent network delays when sending studies. Host computer was verify remotely with Philips Remote Service. Host computer is fully operational. ----------------------------------------------------------------------------- Work Order #: WO-03833141 Sheffield, Scott The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Problem description by engineer : problems sending images to pacs Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No further information is available as per the customer Expected and Unexpected results: problems sending images to pacs Patient Impact/Outcome: none User Impact: delay Current Software Version: 4.1.7.25030 Diagnostic performed by Engineer : Russ said they have been having this problem in the past and again today. Russ said that it has been freezing up sending images to pacs, now it is starting to send but is sending images very slow to pacs. He is not having problems sending anywhere else and no other systems are having problems sending to pacs. Internal Comments : Error # and/or description of error : none Repair Action : Reading previous case notes the problem was with customer network. When I spoke with FSE Mark he said it is problem with customer network. Follow Up Required : Error # and/or description of error : none Repair Action : Reading previous case notes the problem was with customer network. When I spoke with FSE Mark he said it is problem with customer network. Technician Name : Sheffield, Scott ----------------------------------------------------------------------------- | [15 empty] | 117870905 | Closed | 1099754 |  |  | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS Transfer | [5 empty] | Refresh |  |  | Upgrade issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1226 |  |  | United States | AdventHealth Zephyrhills | 300106 | 44382 |  | 1302CDN84 DICOM transfer issues sending to PACS | Work Order #: WO-03795988 Mark Finnigan The system partly meets specification and is returned to use with restrictions Follow Up Required : Spoke with Jesse. She states that the DICOM transfers are taking an excessive amount of time to complete. She added that her FSE Jorge worked on this issue this weekend but there is no case showing such activity. Tried to have Jesses log in as olevel but the system was not accepting the username and password. Dispatching the case to the field as there is nothing else that can be done to resolve the issue remotely. According to the customer this system is running 4.1.1. Diagnostic performed by Engineer : Spoke with Jesse. She states that the DICOM transfers are taking an excessive amount of time to complete. She added that her FSE Jorge worked on this issue this weekend but there is no case showing such activity. Tried to have Jesses log in as olevel but the system was not accepting the username and password. Dispatching the case to the field as there is nothing else that can be done to resolve the issue remotely. According to the customer this system is running 4.1.1. Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: The DICOM transfers to PACS are taking an excessive amount of time. User Impact: Patient Impact: Current Software Version: ----------------------------------------------------------------------------- Work Order #: WO-03796009 06-Jul-2021 Sher Mark The system fully meets specification and returned to use Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: The DICOM transfers to PACS are taking an excessive amount of time. User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : customer confirms problem is in their network and has observe lost packets customers network is being handled by their IT dapartment ----------------------------------------------------------------------------- | [15 empty] | 117836591 | Closed | 1077479 |  |  | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS Transfer | [5 empty] | Refresh |  |  | Upgrade issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1227 |  |  | Australia | Maia Financial Pty Limited | 341000 | 44411 |  | Scan cannot initialise | Work Order #: WO-04024450 05-Aug-2021 David BROWN The system meets the specification for the performed service and is returned to use Diagnostic performed by Engineer : Check over system. Carry out replacement of R2D. Carry out full checks and setups. Carry out functional testing. Return to service. Problem Description by Engineer : ________________________________________________________________ CUSTOMER REPORTED PROBLEM OR EVENT DESCRIPTION: 3 Aug 2021 Customer called to report that system would not initialize in preparation for a surview scan, despite shutting down to the wall power. ________________________________________________________________ Parts Consumed : 1.00000,459800467341,R2D Assembly ----------------------------------------------------------------------------- Work Order #: WO-04015383 03-Aug-2021 David BROWN The system does not meet the specification for the performed service and is not returned to use Problem Description by Engineer : ________________________________________________________________ CUSTOMER REPORTED PROBLEM OR EVENT DESCRIPTION: 3 Aug 2021 Customer called to report that system would not initialize in preparation for a surview scan, despite shutting down to the wall power. ________________________________________________________________ Diagnostic performed by Engineer : check over system. all systems seem to be running fine. reset CIRS and full system shut down. system failed during scan. check with Steve Teixera for further instructions. Follow Up Required : N/A Follow Up Required : check angulation of Rotor controller and check r2d signal ----------------------------------------------------------------------------- Work Order #: WO-04021651 04-Aug-2021 David BROWN The system does not meet the specification for the performed service and is not returned to use Diagnostic performed by Engineer : Check over system. Parts have been ordered by RSE. Check system fault with diagnostics by RSE. Order R2D for delivery. Delivery stalled as Border is closed and slowing same day delivery. Return Thursday. Problem Description by Engineer : ________________________________________________________________ CUSTOMER REPORTED PROBLEM OR EVENT DESCRIPTION: 3 Aug 2021 Customer called to report that system would not initialize in preparation for a surview scan, despite shutting down to the wall power. ________________________________________________________________ ----------------------------------------------------------------------------- Work Order #: WO-04015214 Steve, Teixeira The system does not meet the specification for the performed service and is not returned to use Problem Description by Engineer : ________________________________________________________________ CUSTOMER REPORTED PROBLEM OR EVENT DESCRIPTION: 3 Aug 2021 Customer called to report that system would not initialize in preparation for a surview scan, despite shutting down to the wall power. ________________________________________________________________ Diagnostic performed by Engineer : ________________________________________________________________ CUSTOMER REPORTED PROBLEM OR EVENT DESCRIPTION: 3 Aug 2021 Customer called to report that system would not initialize in preparation for a surview scan, despite shutting down to the wall power. ________________________________________________________________ TROUBLESHOOTING: Review of the logs shows Rotor positional WARNING errors (S_MP_POSLOOP_MAX_MIN_ACCEL_ERROR,S_MPS_MOTION_TIMEOUT) on the first attempt at the surview. NORMAL error "S_DMC_SHOT_COMM_COLLAPSE" followed on the 2nd attempt, at 09:21:48.180 3 Aug 2021. Review of the CIRS logs shows numerous resends starting at 09:21:47 and persists throughout the logs. After a shutdown the logs continue to make reference to Rotor positional and Data path errors. ________________________________________________________________ TECHNICAL SUPPORT: 3 Aug 2021 / 4 Aug 2021 Assisted the Engineer on site with troubleshooting. 5 Aug 2021 Assisted the Engineer with handover checks. ________________________________________________________________ ONSITE WORK ORDER: Please refer to onsite work order WO-04015383 for further details of the actions taken to address this issue. ________________________________________________________________ Follow Up Required : TROUBLESHOOTING: Review of the logs shows Rotor positional WARNING errors (S_MP_POSLOOP_MAX_MIN_ACCEL_ERROR,S_MPS_MOTION_TIMEOUT) on the first attempt at the surview. NORMAL error "S_DMC_SHOT_COMM_COLLAPSE" followed on the 2nd attempt, at 09:21:48.180 3 Aug 2021. Review of the CIRS logs shows numerous resends starting at 09:21:47 and persists throughout the logs. After a shutdown the logs continue to make reference to Rotor positional and Data path errors. NEXT STEP ACTION: Engineer required on site to isolate the fault after consultation with the Case RSE. ----------------------------------------------------------------------------- | 2 |  | Carry out replacement of R2D. Carry out full checks and setups. Carry out functional testing. 459800467341,R2D Assembly 可能和供应商水洗工艺相关ECR-077785 | [12 empty] | 118046426 | Closed | 1177833 |  |  | Opened |  | CTUDT00630864 | [5 empty] | Engineering | [6 empty] | R2D | [9 empty] | System | Gantry-HW | R2D | Gantry-HW-R2D | Gantry-HW-R2D | CTUDT00630864 | R2D | Gantry-HW | Gantry-HW-R2D | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1228 |  |  | Chile | BANCO DE CHILE | 341005 | 44409 |  | CL1688 - SYSTEM IS BLOCKED AND WITH IMAGE ERRORS. | Integration:Confirmation of order 53964748 saved ----------------------------------------------------------------------------- Work Order #: WO-03986980 Valenzuela Velázquez, Eduardo Ariel The system meets the specification for the performed service and is returned to use. Problem Description by Engineer : Problem description by engineer : When performing an acquisition, the equipment is not displaying images on the screen. Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected Results: No errors in reconstructed images. Patient Impact/Outcome: Equipment is down. User Impact: Equipment stopped. Current Software Version: 4.1.10 Diagnostic Performed by Engineer: There is a problem with reconstructed images, they are black and without information. Internal Comments : Error # and/or description of error : It does not display acquisition images. Repair Action: Reconstruction tests of previous examinations are performed and it works properly. Air calibration is performed and the equipment is tested, resulting without errors. The equipment is working. Resolution: Reconstruction tests of previous examinations are performed and it works properly. Air calibration is performed and the equipment is tested, resulting without errors. The equipment is working. ----------------------------------------------------------------------------- | 2 |  | Reconstruction tests of previous examinations are performed and it works properly. Air calibration is performed and the equipment is tested, resulting without errors. | [12 empty] | 118019851 | Closed | 1177302 |  |  | Opened |  | TBD-NoEnoughInfo-TBD-NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Recon image | [9 empty] | System | TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1229 |  |  | China | 东莞市中西医结合医院 | 345013 | 44415 |  | Wechat：02N/Ingenuity Core128/Elite China/space is full | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ T2 Activities : Problem Description: 客户问题概述： 计算机软件相关故障。 客户问题详细描述或故障现象： 空间不足 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： NA 远程诊断与测试结果： 主机空间不足，客户已经删除多天病人剩余空间还是很少，告知工程师已经远程指导过，还是不行，需要到场处理。联系工程师怀疑是有异常文件占用空间 PRS TR状态： 不通，请工程师到场检查。 Repair Action: 建议现场维修步骤： 检查是否异常文件占用空间；检查是否需要做full rebuild;检查软件,必要时收集bug report,寻求技术支持；维护PRS，若路由器问题需要备件建议开PRS单订备件。 SAP Transfer : Case Transferred to SAP External Remarks : 08.08.2021 04:42:30 UTC SUP JCOUSER (86001504)服务内容：删除D:\Pms\log\data下的300G日志。 服务结果：工作正常。 Internal Remarks : 08.08.2021 04:42:38 UTC Zongxian Li (86001504) RDI Accuracy:Yes $Q7: (Confirmed the reported problem, D disk was full) $Q8: (Not many patient datas in D disk, the D:\Pms\log\data was too large) $Q9: (Deleted the unnecessariy log files in it) $Q10: (Test and inspection data refer to N/A, no PA required. Device returned to full functionality.) $Q11: $Calibrated Tools name (N/A) S/N (N/A) Next Calibration date (N/A) Safety Questions : Case: 0118076705 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Gao Luqi Last updated by: Gao Luqi 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Dispatched for Onsite Support Next Step Action : Li Zongxian ----------------------------------------------------------------------------- | 3 |  | log is too big, CTUDT00633269 | [12 empty] | 118076705 | In Process | 1196136 |  |  | Opened - Pending Translation |  | CTUDT00632261 | [5 empty] | Engineering | [6 empty] | Disk full | [9 empty] | System | Console-SW | DiskFull | Console-SW-DiskFull | Console-SW-DiskFull | CTUDT00632261 | DiskFull | Console-SW | Console-SW-DiskFull | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1230 |  |  | China | 河北省唐山市工人医院 | 40044 | 44414 |  | 02N/Ingenuity CT 728326/cannot exposure | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Calibration/Adjustment Safety Questions : Case: 0118071432 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Gaoge Last updated by: Wang Gaoge 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 扫描失败。 客户问题详细描述或故障现象： 扫描报错，不能扫描 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 无PRS，看不到具体报错 远程诊断与测试结果： 需要到场检查具体报错 PRS TR状态： 无PRS。 Repair Action: 建议现场维修步骤： 根据报错和故障现象进行维修 Test/Inspection: N/A Test Equipment: N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Problem Reported by Customer : cannot exposure ----------------------------------------------------------------------------- | 3 |  | 服务内容：扫描间温度过热，不让扫描，空调坏了，建议修空调 服务结果：设备恢复正常 | [12 empty] | 118071432 | In Process | 1195213 |  |  | Opened - Pending Translation | [7 empty] | Customer | [6 empty] | Customer | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | n | y | y | y | y | #N/A
1231 |  |  | China | SD Chest Hosp EAST | 40025 | 44415 |  | three issues for 728323 Ingenuity CT 5000 | Problem Describe: Issue 1： Patient get off the couch when scan end，operator use unload or fast out button to drop the couch. At this moment，Estop opened and log report "S_COUCHMGR_HOR_MOTION_STOP_ERROR" . From the Cantrace log no "COUCH_STOP_HOR_ANS and COUCH_STOP_HOR_JOG_ANS" This problem intermediate occur. Issue 2： Host PC reboot,Customer login software in "CT" account intermittent failed. When this issue occur,the PC screen display blue. Need use IST account login can solve it. Issue 3： When print film in CT software , films quantity intermittent lost. In conclusion, Issue 1 and 2 is bug from L3 reply and need time to fix. But the customer temperature is high. | SHU, Tim |  | what is detail? | [14 empty] | 1194004 |  |  | Investigation | [7 empty] | TBD | [6 empty] | 44411 | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1232 |  |  | China | 莱阳市人民医院 | 345005 | 44413 |  | 02N/Ingenuity Core128/Elite China/power off | BU Q&R Feedback: No Reason: Corrective Maintenance External Remarks : 06.08.2021 06:00:42 UTC Xingbo Yang (86038250)服务内容：间歇性掉电，更换FRC电源。 服务结果：维修后，设备恢复正常工作。 Internal Remarks : 06.08.2021 06:00:45 UTC Xingbo Yang (86038250) "RDI Accuracy:yes" $Q7: (Confirmed reported problem & gantry e-stop open) $Q8: ( review error logger _lifesync p1=2 _& indicate FRC assy defect ) $Q9: ( replace FRC assy 453567057167 ) $Q10: (Test and inspection data refer to: PA & Device returned to full functionality) $Q11: Calibrated tools name(NA) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date: 9/10 / 2020 ^Current Tube usage: scan second_246932 Problem Reported by Customer : 02N/Ingenuity Core128/Elite China SAP Transfer : Case Transferred to SAP T2 Activities : Problem Description: 客户问题概述： ESTOP断开。 客户问题详细描述或故障现象： 偶发掉电 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： lifesync p1=2 远程诊断与测试结果： 机器静置时偶发掉电，丢rhost PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 检查供电，RHOST，检查接线 建议备件12NC及名称： " 453567057167 FRC ASSEMBLY" SAP Transfer : Dispatched for Onsite Support Next Step Action : PPD备件453567057167安排24H到医院，杨兴波-86-13583235930 Safety Questions : Case: 0118062481 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wennian Han Last updated by: Wennian Han 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action : YANG XINGBO ----------------------------------------------------------------------------- | 3 |  | $Q7: (Confirmed reported problem & gantry e-stop open) $Q8: ( review error logger _lifesync p1=2 _& indicate FRC assy defect ) $Q9: ( replace FRC assy 453567057167 ) | [12 empty] | 118062481 | Closed | 1191118 |  |  | Opened - Pending Translation | [7 empty] | TBD | [6 empty] | FRC | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1233 |  |  | China | Shandong Province Chest Hospital | 40025 | 44411 |  | 02N/Ingenuity CT/HOST Print failed | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Calibration/Adjustment Internal Remarks : 08.08.2021 02:34:25 UTC Zhaokun Huang (86038405) "RDI Accuracy: YES." $Q7: ( Confirmed reported problem & When Print loss films.) $Q8: ( Ping Printers from Host , passed. Reconfiguration Printers in Lanconfig.) $Q9: ( Reconnected cables between router and printers . Reconfiguration Printers in Lanconfig ) $Q10: ( Test and inspection data refer to: PA . Device returned to fully functionality.) $Q11: Calibrated Tools name (N/A) ^Current Tube installed date: N/A ^Current Tube usage: N/A Problem Reported by Customer : 02N/Ingenuity CT/ T2 Activities : Problem Description: 客户问题概述： 其他类型故障。 客户问题详细描述或故障现象： 打印两张胶片只有一张正常打印 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 打印胶片没有完全打印 远程诊断与测试结果： 考虑软件设置问题 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 1，检查打印机的设置，重新删除添加打印机节点 Test/Inspection: N/A Test Equipment: N/A Safety Questions : Case: 0118048554 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Gaoge Last updated by: Wang Gaoge 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : HUANG ZHAOKUN External Remarks : 03.08.2021 12:38:12 UTC SUP JCOUSER (86038405)服务内容：采集bugreport，重新配置HOST打印机。 服务结果：维修后，设备恢复正常使用。 ----------------------------------------------------------------------------- | 3 |  | 采集bugreport，Software team is under investigation https://teams.microsoft.com/_#/files/%E5%B8%B8%E8%A7%84?threadId=19%3A76cc42a0c8fa4bd7877f7de12141f84c%40thread.tacv2&ctx=channel&context=Escalation%2520%25230118138741%2520new%2520to%2520Queue%2520-%2520customer%2520temperature%2520Yellow!!&rootfolder=%252Fsites%252FBigBoreRTRefreshFOKCoreTeam%252FShared%2520Documents%252FGeneral%252FIngenuity%2520Bugreps%252FChina%2520KM%252FEscalation%2520%25230118138741%2520new%2520to%2520Queue%2520-%2520customer%2520temperature%2520Yellow!! | [12 empty] | 118048554 | Closed | 1183210 |  |  | Opened - Pending Translation |  | CTUDT00633684 | [5 empty] | Engineering | [6 empty] | Printer | [9 empty] | System | Console-SW | Imagemissing-Printer | Console-SW-Imagemissing-Printer | Console-SW-Imagemissing-Printer |  |  | Console-SW | Console-SW-Imagemissing-Printer | No | 0 | 0 | 0 | 0 | 0 | No | y | y | y | 0 | 0 | y | y | y | 0 | y | #N/A
1234 |  |  | China | ZhengXiang White Banner Hospital | 345033 | 44411 |  | Wechat:02n/Ingenuity Core128/Elite China/Cannot log on | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer:cannot log on Safety Questions:Case: 0118046688 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Tao Fengming Last updated by: Tao Fengming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer_s problems: CT login password is locked It is a technology problem with the equipment itself: Yes. Troubleshooting Action: remote diagnosis and testing result: Need an engineer to go to the field to reset the password PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Need an engineer to go to the field to reset the password SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support External Remarks : 04.08.2021 08:36:45 UTC Ruijie Su (86039043)"SUAT. Ready for Clinical Use local time" : "2021-08-04 17:30" External Remarks: 04.08.2021 08:42:05 UTC Ruijie Su (86039043) service outcome: [reset password and it is back to normal. The equipment is back to normal.] Other significant matters: none Internal Remarks: 04.08.2021 08:42:15 UTC Ruijie Su (86039043) $Q7: (Confirmed reported problem: cannot log on ) $Q8: on sit service, Cannot log on, need to reset password. . $Q9: Reset Password . .system is ok! $Q10:Test and inspection data refer to:PA. Device returned to full functionality. $Q11: Calibrated Tools name N/A. Current Tube usage: scan second: 16693 ----------------------------------------------------------------------------- | 2 |  | CT login password is locked Local service communication | [12 empty] | 118046688 | Closed | 1182226 |  |  | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Password | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1235 |  |  | China | 黑龙江省桦川县人民医院 | 345018 | 44410 |  | 02N/Core128estop open | BU Q&R Feedback: No SAP Customer Symptom Code: HANG Customer Symptom Code: Hang/Crash Reason: Corrective Maintenance Engineer Failure Code: Wear-out/Breakage/Mechanical Failure Repair Activity Code: Reboot/Restart Problem Reported by Customer : 02N/Core128/ Safety Questions : Case: 0118021641 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Bo Last updated by: Wang Bo 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 机架掉电。 Troubleshooting Action: 机架掉电，PRS不通，通过与客户沟通怀疑供电不稳造成，重启机架后正常，FSE反馈再观察。 Repair Action: 机架掉电，PRS不通，通过与客户沟通怀疑供电不稳造成，重启机架后正常，FSE反馈再观察。 Test/Inspection: customer confirm can scan patient and image ok.functional Check PASSED. Device returned to full functionality. Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | 机架掉电，PRS不通，通过与客户沟通怀疑供电不稳造成，重启机架后正常，FSE反馈再观察。 | [12 empty] | 118021641 | Fixed | 1178417 |  |  | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Customer | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | n | y | y | y | y | #N/A
1236 |  |  | China | 江西省余干楚东医院(新院） | 337500 | 44410 |  | WeChat:02N/Ingenuity Core128/Elite China/Serious artifact | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Calibration/Adjustment Problem Reported by Customer : Serious artifact Safety Questions : Case: 0118021350 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Cheng Zhenhua Last updated by: Cheng Zhenhua 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 图像伪影。 客户问题详细描述或故障现象： 近期颈胸交界处伪影干扰过大 设备本身技术问题： 否。 Troubleshooting Action: 关键故障代码或现象： 图像有伪影 远程诊断与测试结果： 颈胸间伪影属于特殊部位伪影，只能减弱，不能消除，需要到场检查图像质量，排除设备干扰 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 1，做图像质量校正 2 联合技师尝试优化扫描条件和方式降低伪影影响 Test/Inspection: N/A Test Equipment: N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Internal Remarks : 02.08.2021 11:21:58 UTC Jiawei Li (86091272) "RDI Accuracy:Yes" $Q7:(Confirmed reported problem，Image has artifacts) $Q8:(artifacts of special-site ,cannot to be removed) $Q9:( Instruct the physician to pay attention to patient positioning and scan condition selection .) $Q10:(Test and inspection data refer to PA. Device returned to full functionality.) $Q11:(Calibrated Tools: No tools used.) External Remarks : 02.08.2021 11:21:55 UTC Jiawei Li (86091272)服务内容： 颈肩交接的骨头处有伪影干扰，由于人体生理构造原因，此处属于特殊部位伪影，无法消除。指导医院注意患者摆位及扫描条件选择。 服务结果：设备正常使用。 ----------------------------------------------------------------------------- | 3 |  | $Q8:(artifacts of special-site ,cannot to be removed) $Q9:( Instruct the physician to pay attention to patient positioning and scan condition selection .) $Q10:(Test and inspection data refer to PA. Device returned to full functionality.) 颈肩交接的骨头处有伪影干扰，由于人体生理构造原因，此处属于特殊部位伪影，无法消除。指导医院注意患者摆位及扫描条件选择。 | [12 empty] | 118021350 | Closed | 1178263 |  |  | Investigation | [7 empty] | Application | [6 empty] | Application | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | n | y | y | y | y | #N/A
1237 |  |  | China | Ning_an City No. 2 People_s Hospital | 345039 | 44410 |  | Wechat:02N/Ingenuity Core128/Elite China/X-ray fault | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Replace Part SAP Transfer:Case Transferred to SAP Problem Reported by Customer:X-ray fault Safety Questions:Case: 0118020289 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Bai Jingwen Last updated by: Bai Jingwen 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer_s problems: Scanning failed. Detailed description of Customer_s problems or malfunctions: Scan error, preheat failed It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: " <*HV*> szInfo: 1:1 Grid Comm Fault 2021/08/01 20:32:48.070 <*HV*> szInfo: 8:2 No Fil Current" Remote troubleshoot and test result: Consider the failure of the filament driv and grid control, very likely cathode failure. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: 1. check system status, measure tube light. 2. perform HV diagnosis We recommend spare parts 12NC and names: "Spellman 80 KW RoHS CPM Assy 453567560002" Test/Inspection: N/A Test Equipment: N/A SAP Transfer:Dispatched for On site Support Next Step Action: PPD parts 453567560022 have arranged for 24H to the hospital, Du Haijun -86-18604602671 Next Step Action: Liu Yadong External Remarks : 03.08.2021 14:59:20 UTC Yadong Liu (86039473)"SUAT. Ready for Clinical Use local time" : "2021-08-03 22:00" External Remarks: 04.08.2021 00:36:31 UTC Yadong Liu (86039473) service details: exposure and scan failed. Replace cathode and it is back to normal service result: device returns to normal working condition after service Internal Remarks: 04.08.2021 00:36:34 UTC Yadong Liu (86039473) RDI Accuracy: Yes $Q7:(Confirmed reported problem & X-ray fault,grid comm fault) $Q8: (check error logger, 1:1 Grid Comm Fault and 8:2 No Fil Current, check filament of tube is ok, CPM fault) $Q9: (Replace CPM, cperform relative test, system is ok) $Q10: (Test and inspection data refer to: PA. Device returned to full functionality.) $Q11:Calibrated tools refer to:PA ----------------------------------------------------------------------------- | 2 |  | $Q7:(Confirmed reported problem & X-ray fault,grid comm fault) $Q8: (check error logger, 1:1 Grid Comm Fault and 8:2 No Fil Current, check filament of tube is ok, CPM fault) $Q9: (Replace CPM, cperform relative test, system is ok) | [12 empty] | 118020289 | Closed | 1177704 |  |  | Closed - No Further Investigation |  | TBD-APM/CPM | [5 empty] | Engineering | [6 empty] | CPM | [9 empty] | System | IC-HW | APM/CPM | IC-HW-APM/CPM | IC-HW-APM/CPM | TBD-APM/CPM | APM/CPM | IC-HW | IC-HW-APM/CPM | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1238 |  |  | China | 福建省医科大学附属协和医院 | 345010 | 44409 |  | 02n/Ingenuity Core128/Elite China/Unable to expose | BU Q&R Feedback: No Reason: Corrective Maintenance SAP Transfer : Dispatched for Onsite Support Problem Reported by Customer : Unable to expose Next Step Action : FSE 李枫 018650369579 call in and disp him Safety Questions : Case: 0118019003 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Lili Last updated by: Lu Lili 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP External Remarks : 02.08.2021 05:58:52 UTC SUP JCOUSER (86036224)服务内容：CT不能曝光。到场检查发现重建柜那边接受到的光路顺序为0x0。打开机架前后盖，检测TX和RX，无异常。重启机架后 ，重建柜得到正确的光路顺序为0x1234。 服务结果：维修后，设备恢复正常工作。 Internal Remarks : 02.08.2021 05:58:55 UTC Feng Li (86036224) RDI Accuracy: Yes $Q7: (Confirmed reported problem.The ASP PM Engineer on site do PM find the CT cannot scan.) $Q8: (Philips FSE on site check and check the Gantry TX and RX.) $Q9: (The channel is wrong. 0X0. Restart the gantry to get the right channel, 0X1234.) $Q10: (Test and inspection data refer to: PA.+Device returned to full functionality.) $Q11: Calibrated Tools name (FLUKE 287C Multimeter) S/N (16490128) NextCalibration date (10/20/2021) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date: NA ^Current Tube usage: NA ----------------------------------------------------------------------------- | 3 |  | $Q8: (Philips FSE on site check and check the Gantry TX and RX.) $Q9: (The channel is wrong. 0X0. Restart the gantry to get the right channel, 0X1234.) did customer restart system regulaly | [12 empty] | 118019003 | Closed | 1175573 |  |  | Closed - No Further Investigation | [7 empty] | TBD | [6 empty] | Reboot | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1239 |  |  | Germany | Gesundheit Nord gGmbH Klinikverbund | 40002 | 44413 |  | Soft tissue / sagittal images are not completely recalculated | Internal Remarks : involviert Herr Halama Next Step Action : Anruf wird unter der angegebenen Tel.Nr. nicht beantwortet automatische Ansage: "Der Teilnehmer ist abwesend" Next Step Action : Anruf wird unter der angegebenen Tel.Nr. nicht beantwortet automatische Ansage: "Der Teilnehmer ist abwesend" ----------------------------------------------------------------------------- Work Order #: WO-04037036 Carsten Becker Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- MPR not reconstructed successfully Customer Temperature as Yellow | 3 |  | MPR not reconstructed successfully Customer Temperature as Yellow Reminder sent to the help desk that the escalation is still unprocessed The call will be made on the specified phone number. not answered automatic announcement: "The participant is absent" @Silva, need link to MPR defect | [12 empty] | 118066890 | In Progress | 1192687 |  |  | Opened - Pending Translation |  | CTUDT00632116 | [5 empty] | Engineering | [6 empty] | MPR | [9 empty] | System | CIRS-SW | MPR Recon | CIRS-SW-MPR Recon | CIRS-SW-MPR Recon | CTUDT00632116 | MPR Recon | CIRS-SW | CIRS-SW-MPR Recon | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | n | y | y | 0 | y | #N/A
1240 |  |  | Germany | Gesundheit Nord gGmbH Klinikverbund | 40002 | 44411 |  | It takes a long time until images arrive in PACS when sending | Internal Remarks: Mr. Halama clarifies with Mr. Kraschewski and contact us when he has received the computer. -------------------------------------------------- --------------------------- Work Order #: WO-04018729 Carsten Becker The system meets the specification for the performed service and is returned to use Diagnostic performed by Engineer: LogFile download and analysis Problem Description by Engineer: via FSE Patient data are not automatically sent to the PACS Follow Up Required: after consultation with R. Kraschewski (TSS) Installation of a sniff PC in the customer network is required to clarify the cause. The FSE used (e.g. T. Halama) should contact him so that he can send him the computer and coordinate the further procedure. Internal Comments: No errors listed in the transfer log. All images were successfully transferred The following messages were stored in the default log for the period in question 2021.08.02 10:59:14 TCPLayer: recvCompleted-EndReceive returned with size = 0 2021.08.02 11:00:53 2021.08.02 11:05:21 2021.08.02 11:10:44 2021/08/02 11:14:54 AM DicomSession: DICOM Error. Status = 0110 RemoteDevice: MPPS failed with status 0110 Generic MPPSClient: Exception while sending Create Message to RIS. Failure status returned 0110 DicomSession: servicesClose called but association already disposed ExamManager MPPS Set message sending failed SSDE Computation failed as the WED profile is empty or incomplete Request RR from T. Halama on box Internal Comments: Mr. Glogner-Umbeer has information that there may be a network problem, as a DICOM error message is issued when communicating with the PACS as well as the RIS server (which both have the same IP address and are therefore physically one computer) DICOM Error Code 0110 = Processing failure | 3 |  | 44411 | [12 empty] | 118049952 | In Progress | 1184114 |  |  | Opened - Pending Translation | [7 empty] | TBD | [6 empty] | 44411 | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1241 |  |  | Japan | Maki Hospital | 341017 | 44413 |  | 04DS:KK: A scan stopped yesterday during a contrast-enhanced examination.>>He was able to scan when he retri [Truncated] | Work Order #: WO-04035332 Kojima, Kenichi The system meets the specification for the performed service and is returned to use Problem Description by Engineer : 昨日の造影検査にてスキャン停止。>>再スキャンでスキャン可。>>S_MPLTRANSPORTREC_DEFRAG_TOUCAN_ERROR Diagnostic performed by Engineer : 再スキャンにて検査可。Ctrl linkのカウントにも大きな異常は無し。 造影検査の停止の為FSE対応希望。三宅さんに対応依頼。 Loggerより ============================================================ 2021/07/30 11:13:53.202 [START STUDY] GSE::CState::StateBkst::STUDY_STARTED was sent to GPC 2021/07/30 11:17:21.399__01 <WARN> Type: WARN ErrNum: S_HMC_AXIS_AT_POSITIVE_LIMIT P1:215600/P2:215895 2021/07/30 11:18:40.194__06 [---] ------ Series 13 Load Structure ----- 2021/07/30 11:18:58.382__02 [--o] ----------------- Series 13 Ready ----------------- 2021/07/30 11:19:03.353__01 [-oo] - - - - - - Shot 0 Starts - - - - - - 2021/07/30 11:19:13.387 [ooo] ----------------- Series 13 Complete ----------------- 2021/07/30 11:20:04.216 [---] ------ Series 14 Load Structure ----- 2021/07/30 11:20:22.479__02 [--o] ----------------- Series 14 Ready ----------------- 2021/07/30 11:20:28.991__01 [-oo] - - - - - - Shot 0 Starts - - - - - - 2021/07/30 11:20:45.917 [ooo] ----------------- Series 14 Complete ----------------- 2021/07/30 11:33:56.619__07 [---] ------ Series 15 Load Structure ----- 2021/07/30 11:33:56.693 [CANCEL] Series 0 cancelled, due to GANTRY_CANCEL_SOURCE_NORMAL_ERROR, in state CStateSeriesLoading 2021/07/30 11:33:56.693__01 [CANCEL] CStateSeriesLoading::GantrySeriesCancelAns had GANTRY_CANCEL_SOURCE_NORMAL_ERROR 2021/07/30 11:33:56.703 <*ERROR*> Type: NORMAL ErrNum: S_MPLTRANSPORTREC_DEFRAG_TOUCAN_ERROR P1:3.289968E+07/P2:24576 2021/07/30 11:36:40.546__10 [---] ------ Series 18 Load Structure ----- 2021/07/30 11:36:40.613 [---] ------ Series 19 Load Structure ----- 2021/07/30 11:36:40.707__01 [---] ------ Series 20 Load Structure ----- 2021/07/30 11:36:56.790__02 [--o] ----------------- Series 18 Ready ----------------- 2021/07/30 11:37:46.664__01 [-oo] - - - - - - Shot 0 Starts - - - - - - 2021/07/30 11:37:55.522 [ooo] ----------------- Series 18 Complete ----------------- Internal Comments : [Subject] 04DS:KK:ﾎｼ様:昨日の造影検査にてスキャン停止。>>再スキャンでスキャン可。>>S_MPLTRANSPORTREC_DEFRAG_TOUCAN_ERROR [ディスパッチ] 1 RC -----------------FSE登録情報----------------- -訪問FSE1氏名 三宅 正浩 -FSE1訪問日時 2021/07/3118:00 -FSE1作業時間 2 ----------------------------------------------------------------------------- Work Order #: WO-04035453 06-Aug-2021 Miyake, Masahiro The system meets the specification for the performed service and is returned to use Regulatory Questions : N/A Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : [作業内容詳細] エラーログ等にて故障原因を調査しました結果、ガントリー内部基板内ソフトウェアにて一部通信エラーが初性しており、エラー検知による撮影停止と思われます。 ハードウェアチェックを行いましたが、異常は見られませんでしたので、基板内部ソフトウェア再インストールを実施しております。 作業実施後テストスキャン等実施正常動作を確認しております。 ご迷惑をお掛けしまして誠に申し訳ありませんでした。 Internal Comments : [CX関連] 機器担当者と報告完了方法を合意できましたか？ YES 当日通常報告完了 [作業報告書（紙）提出] 作業報告書は紙で提出しましたか？ YES Q1その検査はどの様に終了しましたか?: ・中断した個所から再撮 Q2 問題により間違った所見（誤診、誤治療）の有無: ・無し Q3 アーチファクトの有無: ・無し Q4 意図しない不要な曝射の有無: ・無し ----------------------------------------------------------------------------- | 4 |  | 44413 | [12 empty] | 118064473 | Fixed | 1188036 |  |  | Investigation - Pending Translation | [7 empty] | TBD | [6 empty] | 44413 | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1242 |  |  | Japan | Social Welfare Organization Saiseikai Imperial Gift Foundation, Inc. | 343000 | 44413 |  | 04DS:KKI:Artifact occurred on Surview images only.There is no problem with the main scan images. | Work Order #: WO-04033110 05-Aug-2021 Kawashige, Yasuhiro Regulatory Questions : N/A Diagnostic performed by Engineer: N/A Problem Description by Engineer: Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: ----------------------------------------------------------------------------- Work Order #: WO-04032952 Kimura, Kazumasa The system meets the specification for the performed service and is returned to use. Problem Description by Engineer: Artifact occurred on Surview images only.There is no problem with the main scan images. Diagnostic performed by Engineer : Artifact occurred on Surview images only.There is no problem with the main scan images. Internal Comments: [Subject] 04DS:KKI:Mr. Moriwaki:Artifact occurred on Surview images only.There is no problem with the main scan images. [Internal Comments] The problem occurs intermittently. Shipped part: 459800911271:TDMS V2 M.Det Class A FRU QTY:2 ZREP/Traceable Okayama (U616) Dispatched: 1 RC FSE Registration Information: Name of FSE 1 who visited the site: Yasuhiro Kawashige FSE 1 visit date and time: 2021/08/0519:00 FSE 1 hours worked: 2 Dispatch instructions: Dispatch instruction #1: Shipping conditions: TDD (date and time specified) Delivery date: Specified date: 2021/08/05 19:00 Shipping destination: Directly to site. To be collected by: Yasuhiro Kawashige Phone number: 070-2196-4606 Other: call mobile phone upon arrival Person making request: Kazumasa Kimura ----------------------------------------------------------------------------- | 4 |  | 44413 | [12 empty] | 118061985 | In Progress | 1186767 |  |  | Investigation | [7 empty] | TBD | [6 empty] | 44413 | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1243 |  |  | Japan | Independent administrative institution National Hospital Organization | 341015 | 44412 |  | 04DS:KKI: Error occurred during a contrast-enhanced scan >> szInfo: SoftArcReg 0x211*/S_RHOST_HVG_HSS_GENERA [Truncated] | Work Order #: WO-04024958 ----------------------------------------------------------------------------- Work Order #: WO-04024933 04-Aug-2021 Yamamoto, Yuko The system meets the specification for the performed service and is returned to use. Problem Description by Engineer: Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Regulatory Questions : N/A Diagnostic performed by Engineer: N/A ----------------------------------------------------------------------------- Work Order #: WO-04025225 Kimura, Kazumasa The system meets the specification for the performed service and is returned to use. Diagnostic performed by Engineer: Call description: Error occurred during a contrast-enhanced scan [Response] He was able to continue with the examination. Problem Description by Engineer: Error occurred during a contrast-enhanced scan >> szInfo: SoftArcReg 0x211*/S_RHOST_HVG_HSS_GENERATOR_ARC_FAULTDevice Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Internal Comments: [Subject] 04DS:KKI: Mr./Ms. Hisaki: Error occurred during a contrast-enhanced scan >> szInfo: SoftArcReg 0x211*/S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT [Internal Comments] [Results of checking the log] An arc has occurred. szInfo: SoftArcReg 0x211* S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT S_GEN_FAULT_ANODE_TUBE_ARC S_GEN_FAULT_CATHODE_TUBE_ARC LongTC and Filament cal. are scheduled to be performed today. Dispatched: 1 RC FSE Registration Information: Name of FSE 1 who visited the site: L1 Engineer Shigeya Yoshimura FSE 1 visit date and time: 2021/08/0417:30 FSE 1 hours worked: 2 - Name of dispatched supervisor Yuko Yamamoto ¦ L2 (Level 2) Dispatched supervisor [Dispatched supervision] Dispatched supervisor case L1 FSE_ Shigeya Yoshimura ¦ L2 L2 FSE_ Yutaka Yamamoto ¦ L2 L1 FSE_ This system has work performed under the supervision of the L2FSE, and it is necessary for the L2FSE (engaged in another operation) to approve the parts order, manage technical escalation, and receive final approval before starting use on the patient. L2 FSE The role of dispatched supervisors is to approve all parts orders, manage all technical escalations, provide guidance and instructions for service work to the L1FSE who has been assigned, and review and approve the work of the FSE. The L2 FSE makes final approval of the of the matters recorded in the case, and Close (CNF) means that the repair of the SWO and the recorded content (including all recorded comments) were officially approved. ----------------------------------------------------------------------------- | 4 |  | 44412 | [12 empty] | 118055041 | In Progress | 1183073 |  |  | Investigation | [7 empty] | TBD | [6 empty] | 44412 | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1244 |  |  | Japan | Kinan Hospital | 340000 | 44412 |  | 01DS: KKI: "Recon Box initialize error" occurred during a scan. Scanning is not possible >>S_DMC_SHOT_COMM_C [Truncated] | Work Order #: WO-04023668 Kimura, Kazumasa The system does not meet the specification for the performed service and is not returned to use Problem Description by Engineer: "Recon Box initialize error" occurred during a scan. Scanning is not possible >> S_DMC_SHOT_COMM_COLLAPSE > Restarting the system did not improve the situation Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer: "Recon Box initialize error" occurred during a scan. Scanning is not possible >> S_DMC_SHOT_COMM_COLLAPSE > Restarting the system did not improve the situation Internal Comments: [Subject] 01DS: KKI: Mr. Kadosaka: "Recon Box initialize error" occurred during a scan. Scanning is not possible >>S_DMC_SHOT_COMM_COLLAPSE > Restarting the system did not improve the situation [Internal Comments] [Response] Even after a system restart, Surview was possible but the main scan was interrupted by an error. The situation has not improved, so I think the system is down. [Results of checking the log] The following error occurred. S_DMC_SHOT_COMM_COLLAPSE The physical channel order was 0 x 0, and a communication problem occurred between the CIRS and the gantry. Shipped part: 453567538332: Module - Rotor Electronics TX QTY: 1 ZREP/Traceable Tokyo (U611) 453567538322: Module, Optical RX Stator Moog Velocity QTY: 1 ZREP/Traceable Tokyo (U611) 453567057167: FRC ASSEMBLY QTY: 1 ZREP/Traceable Osaka (U615) Dispatched: 1 RC FSE Registration Information: Name of FSE 1 who visited the site: Shigeya Yoshimura FSE 1 visit date and time: 2021/08/0413:00 FSE 1 hours worked: 4.0 Dispatch instructions: Dispatch instruction #1: Shipping conditions: ASAP (Arrange shipment) Delivery date: ASAP Shipping destination: Directly to site. To be collected by: Shigeya Yoshimura Phone number: 090-9020-2081 Other: call mobile phone upon arrival Person making request: Kazumasa Kimura ----------------------------------------------------------------------------- | 2 |  | Shipped part: 453567538332: Module - Rotor Electronics TX QTY: 1 ZREP/Traceable Tokyo (U611) 453567538322: Module, Optical RX Stator Moog Velocity QTY: 1 ZREP/Traceable Tokyo (U611) 453567057167: FRC ASSEMBLY QTY: 1 ZREP/Traceable Osaka (U615) Part/Product - ACQ4eG2 Acquisitor FRU Part Number - 459800245334 Actual Quantity - 1.00000 Part/Product - BR64 LC RCOM ASSY Part Number - 459801073611 Actual Quantity - 1.00000 This case was originally reported as system down since the error occurrence was very frequent, but still intermittent actually. After the receiver and ACQ card were replaced, it has not happened for two weeks. | [12 empty] | 118053754 | In Progress | 1182507 |  |  | Investigation |  | ACQCardReplacement | [5 empty] | Engineering | [6 empty] | Recon Box initialize error | [9 empty] | System | Console-HW | ACQCard | Console-HW-ACQCard | Console-HW-ACQCard |  |  | Console-HW | Console-HW-ACQCard | No | 0 | 0 | 0 | 0 | 0 | No | y | y | y | 0 | 0 | y | y | y | 0 | y | #N/A
1245 |  |  | China | ZhengXiang White Banner Hospital | 345033 | 44411 |  | 02N/Ingenuity Core128/Elite China/Login password error | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ Safety Questions:Case: 0118047344 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingsong Last updated by: Lu Mingsong 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Login password error Troubleshooting Action: Case Number : 0118046688 follow up Repair Action: Case Number : 0118046688 follow up Test/Inspection: Functional Check PASSED. The device returned to full functionality. Test Equipment: N/A SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | Gao, Mia |  | 44411 | [12 empty] | 118047344 | Closed | 1182515 |  |  | Closed - Duplicate | [7 empty] | 44412.4854 | [6 empty] | 44411 | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1246 |  |  | Peru | BANCO BBVA PERU | 40034 | 44412 |  | Bad detector module 24 - 459800594931 | Work Order #: WO-04038394 [REDACTED] The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled. Problem Description by Engineer : Customer Function/Role: Technician Device Use: In use Expected and Unexpected Results: Artifacts in image User Impact: Bad quality images Patient Impact: Bad quality images Current Software Version: 4.1.10 Diagnostic Performed by Engineer: Artifacts are displayed in the image due to two damaged detectors in module 24. Due to quality policies, the detector must be replaced. Follow Up Required: Change module 24. Bring long Allen wrench, according to the manual. ----------------------------------------------------------------------------- | 4 |  | 44412 | [12 empty] | 118060169 | In Progress | 1189887 |  |  | Opened | [7 empty] | TBD | [6 empty] | 44412 | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1247 |  |  | Puerto Rico | METRO MAYAGUEZ INC | 40009 | 44413 |  | Head support is broken | Work Order #: WO-04038494 Valenzuela Velasquez, Eduardo Ariel The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Problem description by engineer : Head support is broken. Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Head support good no cracks. Patient Impact/Outcome: No impact. User Impact: No impact. Current Software Version: 4.1.10 Internal Comments : Error # and/or description of error : Head support is broken. Repair Action : Head support replacement. Diagnostic performed by Engineer : The head support is broken and must be replaced. Follow Up Required : Error # and/or description of error : Head support is broken. Repair Action : Head support replacement. Technician Name : Valenzuela Velasquez, Eduardo Ariel ----------------------------------------------------------------------------- | 3 |  | 44413 | [12 empty] | 118068093 | In Progress | 1190730 |  |  | Evaluation | [7 empty] | TBD | [6 empty] | 44413 | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1248 |  |  | Slovakia | S&T Slovakia s.r.o. | 40062 | 44412 |  | CT Ingenuity, SN 40062 |  | 3 |  | 44412 | [12 empty] | 118055255 | New | 1187383 |  |  | Opened | [7 empty] | TBD | [6 empty] | 44412 | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1249 |  |  | Turkey | T.C Saglik Bakanligi Bagcilar | 40010 | 44412 |  | Gantry dönerken ses yapıyor | Integration : Confirmation of order 53977716 saved ----------------------------------------------------------------------------- Work Order #: WO-04027444 05-Aug-2021 Ergun Aydin The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: technician Device Use: in use Expected and Unexpected Results: noise while gantry rotation User Impact: no user impact Patient Impact: no patient impact Current Software Version:NA Diagnostic performed by Engineer : rotor dönüşünü sonlandırırken gürültü geliyor Resolution : Rotor kayış değişiminden sonra gantry den gürültü geldiği söylendi. Kayış kontrol edildi. Problem görülmedi. Bahsi geçen gürültü eski-yeni kayış farkından dolayıdır. Söylenen gürültü sistem normal çalışması sırasında oluşan, mekanik olarak rotoru durdurma sırasındaki fren sesidir. Herhangi bir düzeltmeye gerek yoktur. Sistem çalışır ve hasta alır durumda teslim edildi. ----------------------------------------------------------------------------- Work Order #: WO-04026830 Salih Mendi The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Customer Function/Role: technician Device Use: in use Expected and Unexpected Results: noise while gantry rotation User Impact: no user impact Patient Impact: no patient impact Current Software Version:NA Diagnostic performed by Engineer : Kullanıcının, gantry dönüşü sırasında oluşan gürültü şikayeti üzerine, kullanıcıya uzaktan destek verildi. Yapılan incelemede, gürültü kaynağı tespit edilemediğinden, saha ziyareti gerçekleştirilerek ayrıntılı inceleme yapılması gerekmektedir. Sistem faal durumdadır. Follow Up Required : Yapılan incelemede, gürültü kaynağı tespit edilemediğinden, saha ziyareti gerçekleştirilerek ayrıntılı inceleme yapılması gerekmektedir. Sistem faal durumdadır. ----------------------------------------------------------------------------- | 3 |  | 44412 | [12 empty] | 118057076 | Closed | 1183909 |  |  | Opened - Pending Translation | [7 empty] | TBD | [6 empty] | 44412 | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1250 |  |  | Turkey | Republic of Turkey Ministry of Health, Bagcilar | 40010 | 44410 |  | the table is having difficulty moving | Integration: Confirmation of order 53967425 has been saved. ----------------------------------------------------------------------------- Work Order #: WO-04014224 02-Aug-2021 Serkan Yavaşça The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : In diagnosis, we found that the table had been impacted on the front left and the holder inside the top cover was broken. The broken part was affixed. Table was lubricated. System was delivered in working condition. ----------------------------------------------------------------------------- | 3 |  | 44410 | [12 empty] | 118045756 | Closed | 1177417 |  |  | Investigation | [7 empty] | Customer | [6 empty] | 44410 | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | n | y | y | y | y | #N/A
1251 |  |  | United States | McLaren Macomb | 40021 | 44414 |  | 0399CFXXX Customer Requesting Support - Parts ID. | Internal Remarks : Interaction Type: New Case Caller: Allison Adamczyk IP & Issue: 85380634 Customer Requesting Support - Parts ID. Event Type: Part ID Service Request: Remote Technical Did Customer Decline Remote Support? No Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Remote Technical Next Step Action: N/A Verbal PO Number: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Integration : Confirmation of order 53991395 saved ----------------------------------------------------------------------------- Work Order #: WO-04046517 Boyd Craig The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Problem description by engineer : Customer needs couch restraint and cushion part number. When was the first occurrence? Within... : 06-Aug-2021 How often is issue occurring? : Unknown Malfunction area : Accessories Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: None Patient Impact/Outcome: Not provided User Impact: None Current Software Version: Not provided Resolution : Researched part numbers for customer. Customer already has the table cushion pad for the Brilliance. 453567009951 6" Patient Restraint strap CT Yes 453567009961 14" Patient Restraint strap CT Yes Customer is okay with closing the call. Diagnostic performed by Engineer : Collect information from customer. Internal Comments : Error # and/or description of error : Part numbers Repair Action : Researched part numbers for customer. Customer already has the table cushion pad for the Brilliance. 453567009951 6" Patient Restraint strap CT Yes 453567009961 14" Patient Restraint strap CT Yes Customer is okay with closing the call. ----------------------------------------------------------------------------- | 5 |  | 44414 | [12 empty] | 118074964 | Closed | 1193725 |  |  | Opened | [7 empty] | TBD | [6 empty] | 44414 | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1252 |  |  | United States | Holston Medical Group | 40053 | 44412 |  | When using angle greater than 10 for axial Brain CT scanner does not create complete result |  |  |  | @Silva, link to MPR defect | [14 empty] | 1192574 |  |  | Investigation |  | CTUDT00632116 | [5 empty] | Engineering | [6 empty] | MPR | [9 empty] | System | CIRS-SW | MPR Recon | CIRS-SW-MPR Recon | CIRS-SW-MPR Recon | CTUDT00632116 | MPR Recon | CIRS-SW | CIRS-SW-MPR Recon | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1253 |  |  | United States | Willis Knighton Hospital | 40004 | 44412 |  | Upper Pallet Float Motion Stiff |  | 4 |  | 44412 | [12 empty] | 118060392 | Closed | 1189984 |  |  | Evaluation | [7 empty] | TBD | [6 empty] | 44412 | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1254 |  |  | United States | Philips Medical Capital | 40028 | 44412 |  | ACR phantom scanning | Work Order #: WO-04029582 Hirchak, Michele The system meets the specification for the performed service and is returned to use Diagnostic performed by Engineer : Customer working with Physicist to set up axial protocols to match the helical clinical scans. Customer unsure on how to build an axial protocol. Physicist sent me his ACR worksheet so I can help Iris later to build protocols. I will call her later.> MBH Problem description by Clinical Support Specialist : Customer Function/Role: Tech Device Use: In Use Expected and Unexpected Results:ACR set up axial protocol for Phantom scanning. User Impact: None Patient Impact: None Current Software Version: 4.1.10 Problem description by Clinical Support Specialist : Called Iris multiple times over the last 2 days. Emailed physicist. I was able to review his ACR document and found multiple errors. I sent him an additional email with corrections, and told him we can not be sure on everything since I have never been able to check with the tech regarding the actual protocol parameters. I will keep trying Iris in the morning. MBH ----------------------------------------------------------------------------- | 5 |  | 44412 | [12 empty] | 118059474 | In Progress | 1185766 |  |  | Evaluation | [7 empty] | TBD | [6 empty] | 44412 | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1255 |  |  | United States | Philips Medical Capital | 40053 | 44411 |  | MPR help | Integration : Confirmation of order 53974846 saved ----------------------------------------------------------------------------- Work Order #: WO-04022710 Curry, Dyan The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Diagnostic performed by Engineer : CFR # 1192574 Problem description by Clinical Support Specialist : Customer Function/ Role: technologist Device Use: in use name Expected and Unexpected Results: Axial head not creating complete set of mprs User Impact: none Patient Impact: none Current Software version: 4.1.10 Diagnostic performed by Engineer : customer states cor and sag reformats a partial of needed images for axial heads. Advised customer tilting gantry can conflict with creating MPRS. Customer is able to create thin data set , load thins in Ct Viewer and create full reformat series. Customer also says EC was loaded on scanner from another scanner using EC sharing. Built new EC using reference set. Customer will use this new EC for next exams. Follow up tomorrow afternoon. Diagnostic performed by Engineer : customer says new EC card worked with no issue . Advised customer CFR will be created for science department to take a closer look for updates. CFR created, waiting for reference number ----------------------------------------------------------------------------- | 4 |  | @Silva, link to MPR defect | [12 empty] | 118052925 | Closed | 1181874 |  |  | Duplicate |  | CTUDT00632116 | [5 empty] | Engineering | [6 empty] | MPR | [9 empty] | System | CIRS-SW | MPR Recon | CIRS-SW-MPR Recon | CIRS-SW-MPR Recon | CTUDT00632116 | MPR Recon | CIRS-SW | CIRS-SW-MPR Recon | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1256 |  |  | United States | Unity Health Systems | 40037 | 44410 |  | Monthly QA | Work Order #: WO-04013475 06-Aug-2021 Lockwood Thomas The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : Performed NY State QA testing ----------------------------------------------------------------------------- | 5 |  | 44410 | [12 empty] | 118045329 | In Progress | 1181364 |  |  | Evaluation | [7 empty] | TBD | [6 empty] | 44410 | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1257 |  |  | United States | Philips Medical Capital | 300057 | 44414 |  | Gantry Error. Delayed imaging after injections. | Internal Remarks : Interaction Type: New Case Caller: Michelle Kelso IP & Issue: 54851181 ; Gantry Error. Delayed imaging after injections. Event Type: Incident Service Request: Onsite Did Customer Decline Remote Support? Yes Preferred Service Time: Week of 08/09/2021 Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Create case and onsite work order Next Step Action: N/A Verbal PO Number: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- Work Order #: WO-04048250 Williams John Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 4 |  | 1.00000,459800170271,Upgrade Kit, Commander-SK Brilliance 64 | [12 empty] | 118076167 |  | 1195981 |  |  | Opened | [20 empty] | Non-Refresh | [12 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | #N/A
1258 |  |  | United States | Christ Hospital | 300205 | 44413 |  | Ecg not functioning | Integration : Check the settlement rule-Entry 1 in table OPERATION (method CHANGE) was not found-Error during processing of BAPI methods-MPG is present on the product code : P9-MPG is present on the product code : P9 Integration : VBAKKOM has been processed successfully-VBAPKOM has been processed successfully-VBAPKOM has been processed successfully-VBAPKOM has been processed successfully-VBAPKOM has been processed successfully-KONVKOM has been processed successfully-KONVKOM has been processed successfully-KONVKOM has been processed successfully-KONVKOM has been processed successfully-REQ2RES DMR 18144623 has been saved-MPG is present on the product code : P9-MPG is present on the product code : P9 ----------------------------------------------------------------------------- Work Order #: WO-04040554 Hanlin Christopher Diagnostic performed by Engineer : N/A Problem Description by Engineer : Problem description by engineer : ECG disconnecting When was the first occurrence? Within... : 01-Aug-2021 How often is issue occurring? : Intermittent Malfunction area : Console Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected results: can not do cardiac exams Patient Impact/Outcome: cardiac exams moved to another scanner User Impact: can not do cardiac exams Current Software Version: Not provided ----------------------------------------------------------------------------- Work Order #: WO-04040376 Mcginnis, Mike The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, no follow-up requested Internal Comments : Error # and/or description of error : ECG disconnecting. Biomed has replaced USB extenders, cable from MP5 to gantry, used the new USB cable from i-box to host and replaced i-box. He said the host sometimes sounds like a piano playing with the usb disconnects. He also stated sometimes the host won_t boot up after a power cycle. Repair Action : Check 5v power supply. If good replace host and insure bios setting for usb is correct. Run new usb cable from gantry to console. Cable already on site Diagnostic performed by Engineer : Spoke to Biomed. He said they have been having ecg disconnects since the last upgrade. Problem Description by Engineer : Problem description by engineer : ECG disconnecting When was the first occurrence? Within... : 01-Aug-2021 How often is issue occurring? : Intermittent Malfunction area : Console Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected results: can not do cardiac exams Patient Impact/Outcome: cardiac exams moved to another scanner User Impact: can not do cardiac exams Current Software Version: Not provided Follow Up Required : Error # and/or description of error : ECG disconnecting. Biomed has replaced USB extenders, cable from MP5 to gantry, used the new USB cable from i-box to host and replaced i-box. He said the host sometimes sounds like a piano playing with the usb disconnects. He also stated sometimes the host won_t boot up after a power cycle. Repair Action : Check 5v power supply. If good replace host and insure bios setting for usb is correct. Run new usb cable from gantry to console. Cable already on site What are the next Troubleshooting steps : Check 5v power supply. If good replace host and insure bios setting for usb is correct. Cable already on site Technician Name : Mcginnis, Mike ----------------------------------------------------------------------------- Replaced USB connection with straight through cable recommended by RSE. Connected directly to Host and iBox. Installed, then de-installed iBox and restored system to original iBox. Programmed iBox and tested operation. Will monitor. | 3 |  | Replaced USB connection with straight through cable recommended by RSE. Connected directly to Host and iBox. Installed, then de-installed iBox and restored system to original iBox. Programmed iBox and tested operation. Will monitor. | [12 empty] | 118069547 |  | 1190332 |  |  | Investigation | [20 empty] | Non-Refresh | [12 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | #N/A
1259 |  |  | United States | Christ Hospital | 300205 | 44411 |  | 459801190772 X1 459801190762 X1 459800166652 X1 | Integration : Confirmation of order 53974620 saved Integration : Confirmation of order 53974620 saved ----------------------------------------------------------------------------- Work Order #: WO-04022342 Parts Consumed : 1.00000,459801190772,ICRON USB RG2302 REMOTE EXT MODULE (REX) 1.00000,459801190762,ICRON USB RG2302 LOCAL EXT MODULE (LEX) 1.00000,459800166651,I-BOX-2 BASE BOARD ASSY ----------------------------------------------------------------------------- | 3 |  | Replaced USB connection with straight through cable recommended by RSE. Connected directly to Host and iBox. Installed, then de-installed iBox and restored system to original iBox. Programmed iBox and tested operation. Will monitor. | [12 empty] | 118052759 |  | 1186147 |  |  | Closed - Non-Complaint | [20 empty] | Non-Refresh | [12 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | #N/A
1260 |  |  | United States | Christ Hospital | 300205 | 44411 |  | ECG gating not working | Work Order #: WO-04022239 Karlen James The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: Biomed Device Use: Diagnostic Expected and Unexpected Results: ECG intermittent User Impact: ECG gating is intermittent Patient Impact: gated scans not performed Current Software Version: 4.10 Diagnostic performed by Engineer : assisted Dennis with part numbers for the IBOX2 and USB converter boxes Resolution : assisted Dennis with part numbers for the IBOX2 and USB converter boxes ----------------------------------------------------------------------------- | 3 |  | Replaced USB connection with straight through cable recommended by RSE. Connected directly to Host and iBox. Installed, then de-installed iBox and restored system to original iBox. Programmed iBox and tested operation. Will monitor. | [12 empty] | 118052451 |  | 1181636 |  |  | Closed - Done | [20 empty] | Non-Refresh | [12 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | #N/A
1261 |  |  | United States | Philips Medical Capital | 300057 | 44410 |  | error when doing short tc | Work Order #: WO-04007702 Williams John Diagnostic performed by Engineer : N/A Problem Description by Engineer : Problem description by engineer : error when doing short tc. Message failed to get DMS.. check gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: error when doing short tc Patient Impact/Outcome: no patient involved User Impact: delay Current Software Version: Not provided ----------------------------------------------------------------------------- Work Order #: WO-04007575 Sheffield, Scott The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Problem description by engineer : error when doing short tc. Message failed to get DMS.. check gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: error when doing short tc Patient Impact/Outcome: no patient involved User Impact: delay Current Software Version: Not provided Diagnostic performed by Engineer : Bridget said that the problem started about 7:45 am and the last time was 8:04 am she rebooted computers and still had problem then she power cycled system. After power cycle I had her do a chest scan of table top and system scanned with no errors. I had her try the short tc again and it did the short tc this time. System is working at this time but she wants FSE to check system. I downloaded log and noted error below. Internal Comments : Error # and/or description of error : 2021/08/02 07:57:32.920__004 S_CONT_AXIS_OVERWRITE_TARGET ERROR_NORMAL 2021/08/02 08:01:27.812__001 S_HMC_CTRL_A_MOTOR_SHORT_CIRCUIT_OR_OVERTEMP ERROR_NORMAL Repair Action : Recommendation: 1. Check for an obstruction or mechanical load problem. 2. On a BedRock couch make sure the power to the AMC is between 14VAC and 62VAC. 3. Check / replace the horizontal motor Follow Up Required : Error # and/or description of error : 2021/08/02 07:57:32.920__004 S_CONT_AXIS_OVERWRITE_TARGET ERROR_NORMAL 2021/08/02 08:01:27.812__001 S_HMC_CTRL_A_MOTOR_SHORT_CIRCUIT_OR_OVERTEMP ERROR_NORMAL Repair Action : Recommendation: 1. Check for an obstruction or mechanical load problem. 2. On a BedRock couch make sure the power to the AMC is between 14VAC and 62VAC. 3. Check / replace the horizontal motor Technician Name : Sheffield, Scott ----------------------------------------------------------------------------- Couch motor overtempt errors. Re-flash couch controller (hot fix). | 3 |  | Couch motor overtempt, Hot fix 1 | [12 empty] | 118040308 |  | 1175874 |  |  | Closed - Done |  | CTUDT00630138/630160 | [5 empty] | Engineering | [6 empty] | HF1 | [5 empty] | Refresh |  |  |  | System | Couch-SW | Horizontal Move Block | Couch-SW-Horizontal Move Block | Couch-SW-Horizontal Move Block | CTUDT00630138/630160 | Horizontal Move Block | Couch-SW | Couch-SW-Horizontal Move Block | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1262 |  |  | Argentina | INSTITUTO NACIONAL DE | 40027 | 44422 |  | PRESENT ARTIFACTS IN THE PICTURES | Work Order #: WO-04123291 Meza, Edgardo Arturo The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : white line in survey and thick ring in helix image Diagnostic performed by Engineer : Modulo detector con problemas Follow Up Required : Detector module with problems ----------------------------------------------------------------------------- Product/Part - TDMS V2 M.Det Class B FRU Part number - 459800911281 Requested Qty - 1.0000 Product/Part - TDMS V2 M.Det Class A FRU Part number - 459800911271 Requested Qty - 1.0000 | 3 |  | System year and detail information provided and Silva escalate to DMS team | [12 empty] | 118136979 |  | 1219731 |  |  | Opened - Pending Translation |  | CTUDT00632110 | [5 empty] | Engineering | [6 empty] | DMS Module | [9 empty] | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | DMS Module | DMS-HW | DMS-HW-DMS Module | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1263 |  |  | China | 无锡市江原医院 | 345020 | 44423 |  | 02N/Ingenuity Core128/Elite China/Rack failed to start | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : Rack failed to start Safety Questions : Case: 0118137518 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Jianqiang Last updated by: Wang Jianqiang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 扫描失败。 客户问题详细描述或故障现象： 放好病人点go 后，报错 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： " S_HMC_CTRL_A_POSITION_ERROR_LIMIT_EXCEEDED " 远程诊断与测试结果： ACS Code 25 - External problem is preventing the controller to track the position correctly.设备新设备查看AMC以及接线，伴随ACQ 报错，但是eventlog下载不下来，也连不上S1 ，现场检查 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 检查AMC 电源线等，A broken ACS cable caused this error. In a few instances the motion controller power cable L22455 part number 453566491791 was intermittent. The VP light on the ACS would flicker as the table was moved vertically. Also check 40 VAC to ACS SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support ----------------------------------------------------------------------------- | 3 |  | 453566491791 Cable, | [12 empty] | 118137518 |  | 1219903 |  |  | Opened - Pending Translation |  | TBD-AMC | [5 empty] | Engineering | [6 empty] | AMC Cable | [9 empty] | System | Couch-HW | AMC | Couch-HW-AMC | Couch-HW-AMC | TBD-AMC | AMC | Couch-HW | Couch-HW-AMC | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1264 |  |  | China | 山东省胸科医院 | 40025 | 44421 |  | 02N/Ingenuity CT 728326/Print fault | T2 Activities : Problem Description: 客户问题概述： 打印故障 客户问题详细描述或故障现象： 间隙性打印图像失败 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： N/A 远程诊断与测试结果： N/A PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 收集bug report，上ESCALATION 建议备件12NC及名称： N/A 客户需求备注： 无 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support ----------------------------------------------------------------------------- | 5 |  | bugreport and share to USP team is under investigation | [12 empty] | 118131496 |  | 1217861 |  |  | Opened - Pending Translation |  | CTUDT00633684 | [5 empty] | Engineering | [6 empty] | Print error | [9 empty] | System | Console-SW | Imagemissing-Printer | Console-SW-Imagemissing-Printer | Console-SW-Imagemissing-Printer |  |  | Console-SW | Console-SW-Imagemissing-Printer | No | 0 | 0 | 0 | 0 | 0 | No | y | y | y | 0 | 0 | y | y | y | 0 | y | #N/A
1265 |  |  | China | 吉林省洮南市中医院 | 345006 | 44420 |  | 02N/Ingenuity Core128/Elite China/artifacts | T2 Activities : Problem Description: 客户问题概述： 其他类型故障。 客户问题详细描述或故障现象： 伪影 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 环形伪影 远程诊断与测试结果： 探测器通道故障 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 分析原始数据，找到伪影位置，屏蔽通道，测试伪影是否消失 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 13.08.2021 08:58:06 UTC Tian Hua (86038446)服务内容： 伪影，有通道响应低，校准。 服务结果： 维修后，设备恢复正常工作。 其它重要事项： 无。 Internal Remarks : 13.08.2021 08:58:09 UTC Tian Hua (86038446) RDI Accuracy ：YES $Q7: (Confirmed reported problem&Artifact) $Q8: (Review logger and check the image，The image artifact） $Q9: (Calibration the image and test） $Q10: (Test and inspection data refer to: PA. Device returned to full functionality) $Q11 Calibrated Tools name (FLUKE287C) S/N (27420003) Next Calibration date (03/17/2022) ----------------------------------------------------------------------------- | 3 |  | Keep obseravation | [12 empty] | 118125302 |  | 1213859 |  |  | Pending Review No Further Investigation |  | ShouldbeCustomer | [5 empty] | Engineering | [6 empty] | Ring Artifact | [9 empty] | System | NotEngineering | ShouldbeCustomer | NotEngineering-ShouldbeCustomer | NotEngineering-ShouldbeCustomer | ShouldbeCustomer |  | NotEngineering | NotEngineering-ShouldbeCustomer | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | #N/A
1266 |  |  | China | Jinan City No. 4 People_s Hospital | 40038 | 44414 |  | PRS connection/02N/Ingenuity CT 728326 | Error description: PRS is not through. Service content: Replace mobile network card Service outcome: Device returns to normal working condition after servicing. Internal Remarks : 11.08.2021 03:04:07 UTC Tao Chen (86038373) “RDI Accuracy: YES" $Q7: (Confirmed reported problem, PRS disconnected) $Q8: ( Check 4G signal, found it was low) $Q9: ( Replaced 4G SIM card with new one) $Q10: (Test and inspection data refer to: PA. Device returned to full functionality) $Q11: Calibrated Tools name ( N/A) ^Current Tube installed date: N/A ^Current Tube usage: N/A ----------------------------------------------------------------------------- | 5 |  | ############# | [12 empty] | 118073654 |  | 1204418 |  |  | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | 4G SIM | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1267 |  |  | China | 四川省宜宾市第二中医院 | 345035 | 44417 |  | 02n/Ingenuity Core128/Elite China/Unable | Problem Reported by Customer : Unable to print T2 Activities : Problem Description: Unable to print Troubleshooting Action: Unable to print Repair Action: FSE had guide customer to adjust the network, and then can print Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: NA SAP Transfer : Case Transferred to SAP | 3 |  | ############# | [12 empty] | 118080367 |  | 1199554 |  |  | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Network | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1268 |  |  | Germany | Gesundheit Nord gGmbH Klinikverbund Bremen Klinikum Bremen-Ost | 40002 | 44414 |  | Duplicate of PR#1192687 - Automatic MPR was not finished successfully | Duplicate of PR#1192687 - Automatic MPR was not finished successfully | ZHANG, Yue Xing |  | 44417.8194 | [14 empty] | 1197624 |  |  | Duplicate |  | CTUDT00632116 | [5 empty] | Engineering | [6 empty] | MPR | [9 empty] | System | CIRS-SW | MPR Recon | CIRS-SW-MPR Recon | CIRS-SW-MPR Recon | CTUDT00632116 | MPR Recon | CIRS-SW | CIRS-SW-MPR Recon | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1269 |  |  | Poland | Government Agency. | 341009 | 44420 |  | Repair after PM. | On 08/11/2021, during the review, I found the need to reinstall the software. 08/13/2021 - software reinstallation. After that it works fine. The fault has been removed | 4 |  | software reinstallation. After that it works fine. The fault has been removed | [12 empty] | 118125603 |  | 1214044 |  |  | Opened |  | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Reload SW | [9 empty] | System | TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo |  |  | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | No | 0 | 0 | 0 | 0 | 0 | No | y | y | y | 0 | 0 | y | y | y | 0 | y | #N/A
1270 |  |  | Slovakia | S&T Slovakia s.r.o. | 40062 | 44420 |  | intermittent interrupting scan | Work Order #: WO-04110787 Parts Consumed : null,459800467341,R2D Assembly Resolution : after replacement of R2D system working Diagnostic performed by Engineer : visual diagnostic, log check, discussion with helpdesk Problem Description by Engineer : intermittent interrupting scan | 5 |  | ############# | [12 empty] | 118127842 |  | 1211292 |  |  | Opened |  | CTUDT00630864 | [5 empty] | Engineering | [6 empty] | R2D | [9 empty] | System | Gantry-HW | R2D | Gantry-HW-R2D | Gantry-HW-R2D | CTUDT00630864 | R2D | Gantry-HW | Gantry-HW-R2D | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1271 |  |  | Spain | S.C.S. HOSP. GRAL. DE FUERTEVENTURA | 40020 | 44417 |  | Problems with the image transfer. Francisco Gordillo 677980652 - 928862109 | Integration:Confirmation of order 54002137 saved ----------------------------------------------------------------------------- Work Order #: WO-04058402 [REDACTED] The system meets the specification for the performed service and is returned to use. Diagnostic performed by Engineer: I talk to [REDACTED] and [REDACTED]. Studies are not reaching the station. It is connected to network. I ask them to reconstruct the data base. It seems that it has not worked. Follow Up Required: Check the station and reload the software as required. Regulatory Questions : N/A Problem Description by Engineer: The studies are missing in the station. ----------------------------------------------------------------------------- | 3 |  | ############# | [12 empty] | 118082763 |  | 1201346 |  |  | Duplicate | [7 empty] | TBD | [6 empty] | In progress | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1272 |  |  | Spain | S.C.S. HOSP. GRAL. DE FUERTEVENTURA | 40020 | 44417 |  | The images are not sent to the reconstruction stations. | Reported Problem: Images cannot be reconstructed on workstations. When did the breakdown occur?: Sunday When will the team be available for intervention? What is the best day or time ?: Immediately. Was the equipment in clinical use when the failure occurred ?: Yes Equipment Location: Floor 1 | 2 |  | ############# | [12 empty] | 118082580 |  | 1201218 |  |  | Evaluation | [7 empty] | TBD | [6 empty] | In progress | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1273 |  |  | United States | Philips Medical Capital | 40028 | 44421 |  | 0711CFXXX gantry cannot comply | Integration : Confirmation of order 54022694 saved Integration : According to factory calendar 20210815 (Posting date) is not a workday (check entry) ----------------------------------------------------------------------------- Work Order #: WO-04121033 Atz, Michael The system meets the specification for the performed service and is returned to use Internal Comments : Error # and/or description of error : none Repair Action : gantry cannot comply earlier today looked at the logs no issue found they scanned all day without an issue Diagnostic performed by Engineer : gantry cannot comply earlier today looked at the logs Problem Description by Engineer : Problem description by engineer : gantry cannot comply earlier today When was the first occurrence? Within... : 13-Aug-2021 Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: No further information is available as per the customer Patient Impact/Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer Current Software Version: Not provided Resolution : gantry cannot comply earlier today looked at the logs no issue found they scanned all day without an issue ----------------------------------------------------------------------------- | 3 |  | Seems system power up no having right status | [12 empty] | 118135611 |  | 1217234 |  |  | Opened | [7 empty] | Others | [6 empty] | False alarm | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1274 |  |  | United States | Leesburg Regional Medical Center | 40022 | 44421 |  | 0711CFXXX node needs to be changed | Problem Description by Engineer : Problem description by engineer : node needs to be changed When was the first occurrence? Within... : 13-Aug-2021 How often is issue occurring? : Continuously Malfunction area : Tube&Gen Information to support the complaint handling process Customer Function/Role: BioMed Device Use: No patient involved Expected and Unexpected results: No further information is available as per the customer Patient Impact/Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer Current Software Version: Not provided Diagnostic performed by Engineer : node needs to be changed the new IP 172.25.30.25 dose wise port 1044 Tera recon 172.25.2.155 nighthawk 172.25.30.16 irapid 172.25.2.23 Internal Comments : Error # and/or description of error : none Repair Action : node needs to be changed the new IP 172.25.30.25 dose wise port 1044 Tera recon 172.25.2.155 nighthawk 172.25.30.16 irapid 172.25.2.23 Resolution : node needs to be changed the new IP 172.25.30.25 dose wise port 1044 Tera recon 172.25.2.155 nighthawk 172.25.30.16 irapid 172.25.2.23 ----------------------------------------------------------------------------- | 3 |  | 44422.2014 | [12 empty] | 118135412 |  | 1217035 |  |  | Opened | [7 empty] | TBD | [6 empty] | NAC-PACS SET UP | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | n | y | y | y | NAC | #N/A
1275 |  |  | United States | Philips Medical Capital | 40028 | 44420 |  | 0705CDY71Letter P on Keyboard Not Working | Work Order #: WO-04111366 Paine Matthew The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Customer Function/Role: Technologist Device Use: Diagnostic Expected and Unexpected Results: P Key on Keyboard Does Not Work User Impact: N/A Patient Impact: N/A Current Software Version:N/A Diagnostic performed by Engineer : Clinical Applications Support Remote Michelle Hirchak was working on another case with Customer when issue was determined. Researched Part Number to Order for FSE for FVF. Follow Up Required : Issue-P Key is Not Working On the Keyboard Recommendation to FSE-Replace Defective Keyboard Cherry USB keyboard English Silver 459800918431 Internal Comments : Contacted Primary FSE Mohamed Abbas. Got Contact Information and Verified Site Address with him to Ship Keyboard Directly to Site to have and FSE Install. ----------------------------------------------------------------------------- | 3 |  | 44421.0035 | [12 empty] | 118128240 |  | 1211692 |  |  | Evaluation | [7 empty] | Others | [6 empty] | Keyboard | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1276 |  |  | United States | Unity Health Systems | 40037 | 44418 |  | 4.1.1 CAP questions. | Work Order #: WO-04071844 Mujica, David The system meets the specification for the performed service and is returned to use Problem description by Clinical Support Specialist : Customer Function/Role: Technologist Device Use: In use Expected and Unexpected Results: Accession Numbers split User Impact: none Patient Impact: none Current Software Version: 4.1 Diagnostic performed by Engineer : Andrea indicated that on her CAP CT Exam Card a part of the abdomen (coronals) are getting assigned to the Chest Accession number. The Exam Card manager was accessed and her CAP CT EC was assessed. All description associations were labeled correctly. It was suggested that test orders be submitted so this Exam Card/Protocol could be tested. Andrea got in touch with the PACS administrator to put test orders in and test this Protocol. After the test there was no way to determine if they were associating correctly until real orders are put in. A rebuild from factory EC_s was suggested, also a deletion of the coronals and a copy paste of the recon_s that associated correctly was also suggested. Andrea will make the suggested changes and monitor this issue going forward. Andrea had no further questions or need for clinical support. No rescan, no loss of data, no reinjection per Andrea. ----------------------------------------------------------------------------- | 5 |  | 44419.0417 | [12 empty] | 118092323 |  | 1202759 |  |  | Investigation | [7 empty] | Application | [6 empty] | NAC | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1277 |  |  | United States | Philips Medical Capital | 40005 | 44417 |  | protocol assistance | Integration : Confirmation of order 53999971 saved ----------------------------------------------------------------------------- Work Order #: WO-04062129 Curry, Dyan The system meets the specification for the performed service and is returned to use Problem description by Clinical Support Specialist : Customer Function/ Role: technologist Device Use: in use name Expected and Unexpected Results: ECM password expired User Impact: none Patient Impact: none Current Software version: 4.1.10 Diagnostic performed by Engineer : customer logged out of scanner . When she logged back , scanner gave her option to reset password. she is good to go ----------------------------------------------------------------------------- | 4 |  | 44418.191 | [12 empty] | 118085559 |  | 1199051 |  |  | Closed - Non-Complaint | [7 empty] | 44417 | [6 empty] | 44417 | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1278 |  |  | Spain | POLICLINICA NTRA SRA DEL ROSARIO SL | 336190 | 44420 |  | Colchoneta y reposacabezas deteriorados | Integration : Confirmation of order 54014217 saved ----------------------------------------------------------------------------- Work Order #: WO-04110109 Parts Consumed : 1.00000,459800126661,Cushion-47.24 Inch Mattress Pad 1.00000,453566502591,PATIENT CUSHIONS EXP ----------------------------------------------------------------------------- | 5 | [14 empty] | 118127322 |  | 1215227 |  |  | Opened - Pending Translation | [7 empty] | 44421.8785 | [6 empty] | 44420 | [5 empty] | Non-Refresh | [12 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | #N/A
1279 |  |  | Spain | POLICLINICA NTRA SRA DEL ROSARIO SL | 336190 | 44421 |  | The left gantry buttons are damaged and need to be replaced. | Work Order #: WO-04116786 [REDACTED] The system meets the specification for the performed service and is returned to use. Regulatory Questions : N/A Problem Description by Engineer: The left gantry buttons are damaged. Follow Up Required: Replace the keypad. Diagnostic performed by Engineer: I talked to Beatriz; the left keypad of the gantry is worn out and it is not working well, the membranes are very worn out. ----------------------------------------------------------------------------- | 3 | [14 empty] | 118132597 |  | 1214934 |  |  | Opened | [7 empty] | 44421.8194 | [6 empty] | 44421 | [5 empty] | Non-Refresh | [12 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | #N/A
1280 |  |  | United States | Christ Hospital | 300205 | 44420 |  | intermittent gantry cannot comply | Integration : Confirmation of order 54015456 saved Integration : According to factory calendar 20210815 (Posting date) is not a workday (check entry) ----------------------------------------------------------------------------- Work Order #: WO-04110671 Swain Wesley The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: Biomed Device Use: Morning qa Expected and Unexpected Results: gantry cannot comply User Impact: none Patient Impact: none Current Software Version: n/a Diagnostic performed by Engineer : biomed power cycled gantry using blue button. Afterwards the system is operating normally but would like the logs looked at. Unable to login and download logs via prs. Jeff to email bugrep to me for analysis Internal Comments : talked with Jeff 3:15pm as i still have not received bugrep. he stated he is busy and will send it tomorrow morning. 8/13 Jeff did not send bugrep in and i was unable to reach him. Closing work order as system is operational. Resolution : System operating normally after power cycle using blue button by biomed Jeff. | 4 | [14 empty] | 118127739 |  | 1212520 |  |  | Evaluation | [7 empty] | 44421.1493 | [6 empty] | 44420 | [5 empty] | Non-Refresh | [12 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | #N/A
1281 |  |  | United States | Advent Health Wesley Chapel | 300106 | 44419 |  | Scans taking up to 10 minuets to get start scan button |  | SHU, Tim | [16 empty] | 1212298 |  |  | Investigation | [7 empty] | TBD | [6 empty] | Duplicate | [5 empty] | Refresh | [6 empty] | - | - |  |  | 0 | - | No | y | y | y | 0 | 0 | No | y | y | y | y | 0 | n | y | y | y | y | #N/A
1282 |  |  | United States | AdventHealth Zephyrhills | 300106 | 44420 |  | 0711CDN84 Long wait before scan | Integration : Confirmation of order 54014453 saved Integration : According to factory calendar 20210814 (Posting date) is not a workday (check entry) ----------------------------------------------------------------------------- Work Order #: WO-04109858 12-Aug-2021 Sher Mark The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Problem description by engineer : Long wait before scan would initialize When was the first occurrence? Within... : 11-Aug-2021 How often is issue occurring? : Multiple times per week Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Scan does not start immediately Patient Impact/Outcome: No further information is available as per the customer User Impact: have to wait for a long time to start a scan Current Software Version: 4.1.7.25030 Diagnostic performed by Engineer : reviewed logs in ATP, created CFM and uploaded bug reports ----------------------------------------------------------------------------- Work Order #: WO-04109412 Ali, Adam The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Problem description by engineer : Long wait before scan would initialize When was the first occurrence? Within... : 11-Aug-2021 How often is issue occurring? : Multiple times per week Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Scan does not start immediately Patient Impact/Outcome: No further information is available as per the customer User Impact: have to wait for a long time to start a scan Current Software Version: 4.1.7.25030 Diagnostic performed by Engineer : When customer click the _go_ button, they have to wait about 3 to 10 mins before scan would start. Other times they get _System is still initializing. Please wait?_, another time While running air calibration the system froze and did not complete air calibrations. Downloaded 3 different bugreports since yesterday. Logs do not show a hardware issue, unable to download host logs for further analysis. Internal Comments : Error # and/or description of error : "Problem Description" When clicking go, the scanner freezes for 3 to 10 mins Time of Event" "8/12/2021 6:56:00 AM 2021/08/12 06:54:32.000__001 CirsEvt_ImgNodeMgr_699 CIRS_Warning Pre Series CIRS Event Log Unchecked 2021/08/12 06:57:35.000__100 CirsEvt_ImgNodeMgr_307 CIRS_Warning Pre Series CIRS Event Log Unchecked "Problem Description" Initialization issue "Time of Event" "8/11/2021 11:39:00 AM" 2021/08/11 11:37:47.812__001 HOST_USR_POPUP USR_POPUP Pre Series Usplog.log Unchecked Showing message to user : System is still initializing. Please wait? "Problem Description" While running air calibration the system froze and did not complete air calibrations "Time of Event" "8/11/2021 7:02:00 AM" 2021/08/11 07:00:01.000__001 CirsEvt_ImgNodeMgr_699 CIRS_Warning Pre Series CIRS Event Log Unchecked Host logs fail to download d:\>c:\pms\system\unzip -j D:\PMS\BugRep\packages\*_cdm_test.zip -d D:\RSE\Host_Evts cd d:\ c:\pms\system\unzip -j D:\PMS\BugRep\packages\*_cdm_test.zip -d D:\RSE\Host_Evts unzip: cannot find any matches for wildcard specification "D:/PMS/BugRep/packages/*_cdm_test.zip". No zipfiles found. Repair Action : Check host event logs Run test scans and try to reproduce issue Check disks, host hardware, fans, filters , cooling for host tower Clean database Software reinstall if it software shows errors or running slow. Follow Up Required : Error # and/or description of error : "Problem Description" When clicking go, the scanner freezes for 3 to 10 mins Time of Event" "8/12/2021 6:56:00 AM 2021/08/12 06:54:32.000__001 CirsEvt_ImgNodeMgr_699 CIRS_Warning Pre Series CIRS Event Log Unchecked 2021/08/12 06:57:35.000__100 CirsEvt_ImgNodeMgr_307 CIRS_Warning Pre Series CIRS Event Log Unchecked "Problem Description" Initialization issue "Time of Event" "8/11/2021 11:39:00 AM" 2021/08/11 11:37:47.812__001 HOST_USR_POPUP USR_POPUP Pre Series Usplog.log Unchecked Showing message to user : System is still initializing. Please wait? "Problem Description" While running air calibration the system froze and did not complete air calibrations "Time of Event" "8/11/2021 7:02:00 AM" 2021/08/11 07:00:01.000__001 CirsEvt_ImgNodeMgr_699 CIRS_Warning Pre Series CIRS Event Log Unchecked Host logs fail to download d:\>c:\pms\system\unzip -j D:\PMS\BugRep\packages\*_cdm_test.zip -d D:\RSE\Host_Evts cd d:\ c:\pms\system\unzip -j D:\PMS\BugRep\packages\*_cdm_test.zip -d D:\RSE\Host_Evts unzip: cannot find any matches for wildcard specification "D:/PMS/BugRep/packages/*_cdm_test.zip". No zipfiles found. Repair Action : Check host event logs Run test scans and try to reproduce issue Check disks, host hardware, fans, filters , cooling for host tower Clean database Software reinstall if it software shows errors or running slow. What are the next Troubleshooting steps : Check host event logs Run test scans and try to reproduce issue Check disks, host hardware, fans, filters , cooling for host tower Clean database Software reinstall if it software shows errors or running slow. Technician Name : Ali, Adam ----------------------------------------------------------------------------- | 3 |  | Check the software version, RAM size, and if reinstall software, collect bugreport for analysis | [12 empty] | 118126749 |  | 1210743 |  |  | Evaluation |  | TBD_Bugreport | [5 empty] | Engineering | [6 empty] | Long intialize | [5 empty] | Refresh |  |  |  | System | TBD-Bugreport | TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | TBD-Bugreport-TBD_Bugreport | TBD_Bugreport |  | TBD-Bugreport | TBD-Bugreport-TBD_Bugreport | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | #N/A
1283 |  |  | United States | Select Specialty Hospitals | 310051 | 44418 |  | how to burn CD | The system meets the specification for the performed service and is returned to use Diagnostic performed by Engineer : Went over steps to copy images to crd, open cdr folder and press record. Also went over direct record option. Advised customer to clear CDR folder after each cd burning Problem description by Clinical Support Specialist : Customer Function/ Role: technologist Device Use: in use name Expected and Unexpected Results: how to burn cd User Impact: none Patient Impact: none Current Software version: 4.1.10 | 4 | [14 empty] | 118089837 |  | 1202109 |  |  | Evaluation | [7 empty] | Application | [6 empty] | NAC | [5 empty] | Non-Refresh | [12 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | #N/A
1284 |  |  | United States | Wellstar Kennestone Hospital | 68117 | 44417 |  | the gating not working on EKG | IP & Issue: 78989232 the gating not working on EKG Internal Remarks : Parts ID of IBOX to Console extender cable. Next Step Action : IBox looks that has a problem , approved to order . Next Step Action : FSE is trying to JTAG the new IBox , the file doesn_t show on the dvd. He will take the file from the Host computer . Internal Remarks : Emailed the 14.7 Xilinx zip file to Bill Internal Remarks : Approved TPR for 459800861311 - Xilinx 14.7 service Programming Kit FSE needs the 14.7 DVD, but it is not available for download via InCenter. Advised the FSE to try and get it from Xilinix directly, but It is a 15 GB download. ----------------------------------------------------------------------------- Work Order #: WO-04072829 11-Aug-2021 Davis Paul Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- Work Order #: WO-04062228 10-Aug-2021 Scheyer William Problem Description by Engineer : EKG function not working Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- Work Order #: WO-04072748 10-Aug-2021 Bourne Barrett The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, no follow-up requested Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : Assisted with running a new IBOX cable. ----------------------------------------------------------------------------- Work Order #: WO-04058453 09-Aug-2021 Scheyer William The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : EKG function not working Diagnostic performed by Engineer : Found EKG function intermittent. Ordered replacement cable. Follow Up Required : replace IBOX2 USB cable ----------------------------------------------------------------------------- Product/Part - I-BOX-2 BASE BOARD ASSY Part number - 459800166651 Actual Qty - 1.0000 Product/Part - ECG Monitor Trigger 7M Ca Part number - 455012301901 Actual Qty - 1.0000 Part/Product - USB EXTENDER CABLE FROM IBOX TO CONSOLE Part Number - 459801731761 Actual Quantity - 1.00000 | 3 | [14 empty] | 118082792 |  | 1199078 |  |  | Investigation | [7 empty] | 44418.2014 | [6 empty] | 44417 | [5 empty] | Non-Refresh | [12 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | #N/A
1285 |  |  | China | Shandong Province Chest Hospital | 40025 | 44425 |  | Wechat：02N/Ingenuity CT/Machine fault | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : Machine fault Next Step Action : NTD：微信报修 T2 Activities : Problem Description: 客户问题概述： 其他类型故障。 客户问题详细描述或故障现象： 调整网络 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： NA 远程诊断与测试结果： 客户告知明天全院换PACS，调整网络，不需要远程，需要工程师明天到场配合 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 配合医院调整网络，连PACS， External Remarks : 18.08.2021 08:20:52 UTC SUP JCOUSER (86038405)服务内容：配合医院更换PACS，更改IP等工作。 服务结果：配置后，设备正常使用。 Internal Remarks : 23.08.2021 02:42:15 UTC Zhaokun Huang (86038405) "RDI Accuracy: YES." $Q7: ( Confirmed reported problem & Need Config PACS/Worklist.) $Q8: ( Customer request engineer onsite configuration PACS/Worklist.) $Q9: ( Config PACS/Worklist and backup. ) $Q10: ( Test and inspection data refer to: NO PA Required . Device returned to fully functionality.) $Q11: Calibrated Tools name (N/A) ^Current Tube installed date: N/A ^Current Tube usage: N/A ----------------------------------------------------------------------------- | 5 |  | 客户告知明天全院换PACS，调整网络，不需要远程，需要工程师明天到场配合 | [12 empty] | 118150523 |  | 1229915 |  |  | NAC | [7 empty] | Customer | [6 empty] | PACS Setup | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1286 |  |  | Poland | Government Agency. | 341008 | 44414 |  | Starting the tomography computer and application training. |  | 4 |  | App training | [12 empty] | 118073782 |  | 1221599 |  |  | NAC | [7 empty] | Application | [6 empty] | Application | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1287 |  |  | Japan | Independent administrative institution National Hospital Organization | 341013 | 44427 |  | 01DS:KKI: Cannot log in to the Exam Manager. |  | 4 |  | Reset the IT Password and reset the Exam card editing password. | [12 empty] | 118162854 |  | 1237710 |  |  | Duplicate | [7 empty] | Customer | [6 empty] | Duplicate | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1288 |  |  | Japan | Independent administrative institution National Hospital Organization | 341013 | 44425 |  | 01DS: TM: CIRS won_t turn on.No response even when the power switch is pressed. | Work Order #: WO-04142507 18-Aug-2021 Takada, Ushio Regulatory Questions : N/A Problem Description by Engineer: Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer: N/A ----------------------------------------------------------------------------- Work Order #: WO-04142290 Yamamoto, Yuko Regulatory Questions : N/A Problem Description by Engineer: Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer: N/A ----------------------------------------------------------------------------- Work Order #: WO-04142275 Miyakawa, Takeyuki The system does not meet the specification for the performed service and is not returned to use Problem Description by Engineer: CIRS won_t turn on.No response even when the power switch is pressed. Internal Comments: [Subject] 01DS: TM: Mr. Yamamoto: CIRS won_t turn on.No response even when the power switch is pressed. [Internal Comments] Response I determined that the system went down. Shipped part: 459801493161: CT Gen5.3 IDose CIRS, Z8 QTY: 1 ZREP/Traceable BackOrder 459801647771: HP Power Supply Z8 G4 1125W QTY: 1 ZSID BackOrder 459801633881: 64GB DDR4-2666 (8x8GB) RegRAM QTY: 1 ZSID BackOrder Dispatched: 1 RC FSE Registration Information: Name of FSE 1 who visited the site: Yutaka Yamamoto FSE 1 visit date and time: 2021/08/189:30 FSE 1 hours worked: 2 - Name of FSE 2 who visited the site Shio Takada FSE2 visit date and time 2021/08/189:30 FSE2 hours worked 2 Dispatch instructions: Dispatch instruction #1: Shipping conditions: NBD (mixed loading, next day or thereafter) Delivery date: 19 August 2021. No time specified. Shipping destination: Kyoto Office To be collected by: Yutaka Yamamoto Phone number: 090-1730-2299 Other: If the item is not in stock in Japan , please place an overseas order Person making request: Diagnostic performed by Engineer: A call came in from Mr. Yamamoto at 8:04 Call description: CIRS won_t turn on.Even when the power switch is pressed the system does not respond. CT-Host was on Response I determined that the system went down. ----------------------------------------------------------------------------- | 2 |  | 459801493161: CT Gen5.3 IDose CIRS, Z8 QTY: 1 ZREP/Traceable BackOrder 459801647771: HP Power Supply Z8 G4 1125W QTY: 1 ZSID BackOrder 459801633881: 64GB DDR4-2666 (8x8GB) RegRAM QTY: 1 ZSID BackOrder Product/Part - CT Gen5.3 Hawk CIRS Master Z8 Part number - 459801592003 This was a system down case due to CIRS hardware problem. The CIRS could not be turned on, nothing happened by pressing power button. FSE found the problem was caused by the mother board, but the Z8 computer (459801493161 CT Gen5.3 IDose CIRS, Z8) was out of stock in local warehouse. So while waiting for the delivery from SPS, they installed the one for CT7500 (459801592003 CT Gen5.3 Hawk CIRS Master Z8) instead. So the system is currently available, but they need to replace the computer again. Therefore the case is still open. | [12 empty] | 118153000 |  | 1230114 |  |  | Evaluation | [7 empty] | TBD | [6 empty] | CIRS | [9 empty] | TBD | [8 empty] | No | 0 | y | y | y | y | No | y | y | y | y | y | n | y | y | y | y | #N/A
1289 |  |  | China | Dongguan traditional chinese and western medicine Integrated Hospital Hospital | 345013 | 44425 |  | Duplicate to PR#1223313 default file-host*. XML files in D:\pms\log\data up to 300G+N26:T26 made D disk full in one or two days. |  | ZHANG, Yue Xing |  | 44425 | [14 empty] | 1224189 |  |  | Duplicate |  | CTUDT00633269 | [5 empty] | Engineering | [6 empty] | 44424 | [9 empty] | System | Console-SW | DiskFull | Console-SW-DiskFull | Console-SW-DiskFull | CTUDT00633269 |  | Console-SW | Console-SW-DiskFull | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | #N/A
1290 |  |  | China | 济南市第四人民医院 | 40038 | 44430 |  | 02N/Ingenuity CT 728326/CT BOX Button failure | T2 Activities : Problem Description: 客户问题概述： CT BOX按键故障 客户问题详细描述或故障现象： CT BOX按键故障 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： N/A 远程诊断与测试结果： N/A PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 更换CT BOX 建议备件12NC及名称： "CT BOX ASSEMBLY_BEST : 459801705022" 客户需求备注： 无 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : PPD 459801705022 *1 已安排备件24H手工单出货 陈涛-86-18560023293 | 3 |  | CT BOX ASSEMBLY_BEST : 459801705022 $Q8: ( Do couch down movement with CT box, found it was failed intermittently) $Q9: ( Replaced CT box with new one, part number was 459801705022) | [12 empty] | 118174884 |  | 1243873 |  |  | Opened - Pending Translation |  | CTUDT00632185 | [5 empty] | Engineering | [6 empty] | CT BOX | [9 empty] | System | Gantry-HW | CT Box | Gantry-HW-CT Box | Gantry-HW-CT Box | CTUDT00632185 | CT Box | Gantry-HW | Gantry-HW-CT Box | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1291 |  |  | China | Hebei Province Tangshan City Workers Hospital | 40044 | 44424 |  | 02N/Ingenuity CT 728326/artifact | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : artifact Safety Questions : Case: 0118138744 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingsong Last updated by: Lu Mingsong 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 图像放射状伪影 客户问题详细描述或故障现象： 图像放射状伪影 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： N/A 远程诊断与测试结果： N/A PRS TR状态： 无PRS。 Repair Action: 建议现场维修步骤： 执行Bad Detector Test,确定坏模块位置，更换探测器模块 建议备件12NC及名称： 12NC: 459800594911, Description: TDMS Module Class B FRU 客户需求备注： 无 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : PPD备件TDMS Module Class B FRU 459800594911 1需要TSE加批。批复安排24H到场地，请先派工，FSE：徐国梁-86-13869524913 T2 Activities : Problem Description: 客户问题概述： 图像放射状伪影 客户问题详细描述或故障现象： 图像放射状伪影 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： N/A 远程诊断与测试结果： N/A PRS TR状态： 无PRS。 Repair Action: 建议现场维修步骤： 做原始数据分析，执行Bad Detector Test。 建议备件12NC及名称： N/A 客户需求备注： 要求工程师到现场检查 External Remarks : 19.08.2021 12:55:13 UTC JunLei Wang (86039063) 服务内容：设备伪影，更换探测器 服务结果：设备恢复正常 Internal Remarks : 19.08.2021 12:55:16 UTC JunLei Wang (86039063) "RDI Accuracy: Yes" $Q7: (Confirmation & Description of problem) Confirmed reported problem, have artifact $Q8: (Troubleshooting & Results) check the detector ,need to replace $Q9: (Repair steps) repalace the detector $Q10: (Test and Inspection data & Device status) Test and inspection data refer to: PA. Device returned to full functionality $Q11: Calibrated Tools name to PA ----------------------------------------------------------------------------- | 3 |  | 12NC: 459800594911, Description: TDMS Module Class B FRU | [12 empty] | 118138744 |  | 1222990 |  |  | Investigation |  | CTUDT00632110 | [5 empty] | Engineering | [6 empty] | DMS Module | [9 empty] | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | DMS Module | DMS-HW | DMS-HW-DMS Module | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1292 |  |  | China | Jilin Province Songyuan City Central Hospital | 345008 | 44425 |  | 02N/Ingenuity Core128/Elite China/ scan aborted | T2 Activities:Problem Description: Summary of customer_s problems: Scanning failed. Detailed description of Customer_s problems or malfunctions: can not scan after maintenance. The maintenance engineer called asking to look at the log It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: Scan start failure and reboot to restore. Remote troubleshoot and test result: "no frames were generated by DMS during scan. Watch for more occurrences of this error. It is possible to get infrequent occurrences of this error while the scanner is operating correctly." PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: "1. Check the CIRS log for ACQ status and information. This maybe a data path problem. 2. Connect the optical cable directly from the ACQ PCB to the RCOM. Try to make scans without rotation to see if scanning is possible without above error. 3. Clean optical and slipring. 4. Check data path. Run VD 5. On a BB, check Optical Transceiver Rotor and Stator. T2 Activities : Problem Description: scan aborted Troubleshooting Action: no frames were generated by DMS during scan. Repair Action: reboot gantry, the device return full function Test/Inspection: Functional Check PASSED. The device returned to full functionality. Test Equipment: NA | 3 |  | reboot gantry, the device return full function | [12 empty] | 118149741 |  | 1229404 |  |  | Closed - No Further Investigation |  | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Scan abort | [9 empty] | System | DMS-HW | NoEnoughInfo | DMS-HW-NoEnoughInfo | DMS-HW-NoEnoughInfo |  |  | DMS-HW | DMS-HW-NoEnoughInfo | No | 0 | 0 | 0 | 0 | 0 | No | y | y | y | 0 | 0 | y | y | y | 0 | y | #N/A
1293 |  |  | China | Dongguan City Integrated Traditional Chinese and Western medicine Hospital | 345013 | 44424 |  | 02N/Ingenuity Core128/Elite China/escalation | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 02N/Ingenuity Core128/Elite China/escalation Safety Questions:Case: 0118139355 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Bo Last updated by: Wang Bo 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer_s problems: Other types of error. Detailed description of Customer_s problems or malfunctions: Storage space for disc D can be full in one day. It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: The log has no errors. Remote troubleshoot and test result: Restart but didn_t work.Have communicated with the engineer. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Suggest escalation, ask for TSE help. We recommend spare parts 12NC and names: N/A SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support ----------------------------------------------------------------------------- system is generating more number of logs because Verbose logging is enabled in the system. Because of this, every 2-3 min a single defaultFile is getting generated. This is expected as per the design and the usage of the system. Please change verbosity logging back to event to not to generate more number of logs. This will be enabled to get more logs for debugging purpose. Please follow following steps to revert this change. 1. Login as service user. 2. Launch c:\Pms\System\LogConfigurator.exe 3. Change the value Verbose to Event in the Default Level: Combo box of Source setting. 4. Save and Exit. | 5 |  | Please change verbosity logging back to event to not to generate more number of logs. This will be enabled to get more logs for debugging purpose. CTUDT00633269 fix in SP2 | [12 empty] | 118139355 |  | 1223313 |  |  | Opened |  | CTUDT00633269 | [5 empty] | Engineering | [6 empty] | Verbose logging | [9 empty] | System | Console-SW | DiskFull | Console-SW-DiskFull | Console-SW-DiskFull | CTUDT00632261 | DiskFull | Console-SW | Console-SW-DiskFull | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1294 |  |  | China | Dongguan City Integrated Traditional Chinese and Western medicine Hospital | 345013 | 44424 |  | Wechat:02N/Ingenuity Core128/Elite China/ECG fault | BU Q&R Feedback: No Reason: Corrective Maintenance Safety Questions:Case: 0118139175 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Bai Jingwen Last updated by: Bai Jingwen Problem Reported by Customer:ECG fault T2 Activities:Problem Description: Summary of customer_s problems: ECG related failure. Detailed description of Customer_s problems or malfunctions: ECG waveform interference. It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: The log has no errors. Remote troubleshoot and test result: Restart but didn_t work.Have communicated with the engineer. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: It is recommended to replace the parts to eliminate the interference of ECG waveforms. We recommend spare parts 12NC and names: PPD 455012301901, 453567036891, 459800485201, 455012303732 SAP Transfer:Dispatched for On site Support SAP Transfer:Case Transferred to SAP Next Step Action: PPD 459800485201*1+453567036891*1+455012301901*1 24H FSE: Li Zongxian -86-13751567086 arranged shipment ----------------------------------------------------------------------------- | 3 |  | $Q7: (Confirmed the reported problem, ECG wave abnormal) $Q8:( The wave was normal in ECG monitor, considered the cable to ibox and the PS of ibox) $Q9: (Replaced the 5V PS and ecg cable) 更换电源和信号线 | [12 empty] | 118139175 |  | 1223188 |  |  | Opened | [7 empty] | Others | [6 empty] | ECG | [12 empty] | - | - |  |  | 0 | - | No | y | y | y | 0 | 0 | No | y | y | y | y | 0 | n | y | y | y | y | #N/A
1295 |  |  | China | 黑龙江省桦川县人民医院 | 345018 | 44429 |  | 02N/Ingenuity Core128/Elite China/Add node | Problem Reported by Customer : 02N/Ingenuity Core128/Elite China T2 Activities : Problem Description: 客户问题概述： 添加相机节点 客户问题详细描述或故障现象： 添加相机节点 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 远程指导客户用olevel_service账号登录失败 远程诊断与测试结果： N/A PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 现场恢复olevel_service账号密码，添加节点 建议备件12NC及名称： N/A 客户需求备注： 客户很着急，接到工单后及时联系客户，CT.Dr.Zhang：15246304333 Safety Questions : Case: 0118172888 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingsong Last updated by: Lu Mingsong | 3 |  | 现场恢复olevel_service账号密码，添加节点 | [12 empty] | 118172888 |  | 1242806 |  |  | Opened - Pending Translation | [7 empty] | Customer | [6 empty] | PACS Setup | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1296 |  |  | China | 黑龙江省桦川县人民医院 | 345018 | 44427 |  | 02N/Ingenuity Core128/Elite China/Spare parts need to be replaced | BU Q&R Feedback: No SAP Customer Symptom Code: CENV Customer Symptom Code: Customer Environmental Concern Reason: Corrective Maintenance Engineer Failure Code: Environmental Cause Repair Activity Code: Configuration Change Problem Reported by Customer : Spare parts need to be replaced. Safety Questions : Case: 0118161518 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Lili Last updated by: Lu Lili 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Troubleshooting Action: wrong case ,should be workstation,They didn_t have Philips workstation Repair Action: wrong case ,should be workstation,They didn_t have Philips workstation Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | wrong case ,should be workstation,They didn_t have Philips workstation | [12 empty] | 118161518 |  | 1236679 |  |  | Opened - Pending Translation | [7 empty] | Service | [6 empty] | NAC | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1297 |  |  | China | 苏州广慈肿瘤医院有限公司 | 345019 | 44426 |  | WeChat：02N/Ingenuity Core128/Elite China/gantry noise | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : gantry noise SAP Transfer : Case Transferred to SAP Next Step Action : PPD 453566492401*1 453567003451*1 12H FSE:徐昊天-86-18761995321 已安排发货 External Remarks : 20.08.2021 09:40:00 UTC Haotian Xu (86090280) 服务内容： 旋转异响，更换皮带 服务结果： 服务后，设备恢复正常使用。 External Remarks : 20.08.2021 11:08:11 UTC Yang Liu (86036528)服务内容： 旋转异响，更换皮带 服务结果： 服务后，设备恢复正常使用。 Safety Questions : Case: 0118153519 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Zixuan Last updated by: Wang Zixuan 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 设备异响。 客户问题详细描述或故障现象： 机架旋转时异响。 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： log无报错。 远程诊断与测试结果： 重启无效。 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议检查异响来源，重点检查机架压轮和rotor皮带。 建议备件12NC及名称： PPD 453567003451,453566492401 SAP Transfer : Dispatched for Onsite Support Next Step Action : XU HAOTIAN Internal Remarks : 20.08.2021 11:37:43 UTC Yang Liu (86036528) RDI Accuracy: Yes $Q7 (Confirmed reported problem , The problem is the gantry has abnormal sound) $Q8: ( check the gantry ，found the belt has abnormal sound, ) $Q9: (replace the belt according to service manual) $Q10: (Test and inspection scan refer to PA . device returned to full functionality $Q11: Calibrated Tools name (refer to PA ) Internal Remarks : 21.08.2021 06:32:15 UTC Haotian Xu (86090280)RDI Accuracy: Yes $Q7 (Confirmed reported problem , The problem is the gantry has abnormal sound) $Q8: ( check the gantry ，found the belt has abnormal sound, ) $Q9: (replace the belt according to service manual) $Q10: (Test and inspection scan refer to PA . device returned to full functionality $Q11: Calibrated Tools name (refer to PA ) ----------------------------------------------------------------------------- | 3 |  | $Q8: ( check the gantry ，found the belt has abnormal sound, ) $Q9: (replace the belt according to service manual) Check with Yiming to collect customer feedback of the issue | [12 empty] | 118153519 |  | 1232193 |  |  | Opened - Pending Translation |  | ECR-077012 | [5 empty] | Engineering | [6 empty] | Belt | [9 empty] | System | Gantry-HW | Belt | Gantry-HW-Belt | Gantry-HW-Belt | ECR-077012 | Belt | Gantry-HW | Gantry-HW-Belt | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1298 |  |  | China | 正镶白旗医院 | 345033 | 44427 |  | 02N/Ingenuity Core128/Elite China/consulting scan parameter | BU Q&R Feedback: No SAP Customer Symptom Code: MARF Customer Symptom Code: Monitor Alarm Failure Reason: Corrective Maintenance Engineer Failure Code: Incompatible Use Repair Activity Code: Instructed Customer Safety Questions : Case: 0118163251 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Jianqiang Last updated by: Wang Jianqiang 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Problem Reported by Customer : 02N/Ingenuity Core128/Elite China T2 Activities : Problem Description: consulting scan slice Troubleshooting Action: information to customer how to change slice Repair Action: instruct to customer Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: NA SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | instruct customer how to change slice | [12 empty] | 118163251 |  | 1238100 |  |  | Opened - Pending Translation | [7 empty] | Application | [6 empty] | NAC | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1299 |  |  | China | 陕西省安康中医医院 | 345041 | 44426 |  | WeChat:02N/Ingenuity Core128/Elite China/noisy | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : noisy Safety Questions : Case: 0118154307 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Weng Jinming Last updated by: Weng Jinming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 飞利浦64排CT的机 是8月12号 做过一次保养 但是目前做病人 升降床的时候 有噪声 类似于有东西卡在里面的感觉 设备本身技术问题： 是。 Troubleshooting Action: 远程诊断与测试结果： 怀疑滑轨，丝杠问题 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议工程师到场检查床丝杠，滑轨，外壳。。。 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 19.08.2021 15:27:11 UTC Jun Liu (86039454)服务原因：床水平运动阻力大 服务内容：修正被撞部分水平 服务结果：设备工作正常 Internal Remarks : 19.08.2021 16:18:21 UTC Jun Liu (86039454) RDI Accuracy :Yes $Q7: (Confirmed reported problem. & The couch has unexpected noise when in and out) $Q8: (1.Check the couchhorizontal lead screw and found the . & We need mechanical repair to so that the patient table horizontal move normally .) $Q9: (1. Mechanical repair and calibraiton the CT. 2. Test scan for CT was OK. ) $Q10: (Test and inspection data refer to PA. & Device returned to full functionality.) $Q11：Calibrated Tools name (Multimeter Fluke 287) S/N (22940012) Next Calibration date 06/17/2022 Internal Remarks : 19.08.2021 16:21:49 UTC Jun Liu (86039454)RDI Accuracy :Yes $Q7: (Confirmed reported problem. & The couch has unexpected noise when in and out) $Q8: (1.Check the couch horizontal lead screw and found the left LH subframe cover support was knocked out of shape. & We need mechanical repair to so that the patient table horizontal move normally .) $Q9: (1. Mechanical repair and calibraiton the CT. 2. Test scan for CT was OK. ) $Q10: (Test and inspection data refer to PA. & Device returned to full functionality.) $Q11：Calibrated Tools name (Multimeter Fluke 287) S/N (22940012) Next Calibration date 06/17/2022 ----------------------------------------------------------------------------- | 3 |  | $Q8: (1.Check the couch horizontal lead screw and found the left LH subframe cover support was knocked out of shape. & We need mechanical repair to so that the patient table horizontal move normally .) | [12 empty] | 118154307 |  | 1232608 |  |  | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Couch movement | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1300 |  |  | China | 宣城市骨科医院 | 345042 | 44427 |  | 02N/Ingenuity Core128/PACS issue | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Safety Questions : Case: 0118162391 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Cao Last updated by: Liu Cao 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Problem Reported by Customer : 02N/Ingenuity Core128/PACS issue T2 Activities : Problem Description: PACS issue, transfer stuck Troubleshooting Action: PACS issue, transfer stuck Repair Action: Adjust setting with pacs engineer Test/Inspection: N/A Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | Troubleshooting Action: PACS issue, transfer stuck Repair Action: Adjust setting with pacs engineer | [12 empty] | 118162391 |  | 1237263 |  |  | Opened - Pending Translation | [7 empty] | Others | [6 empty] | PACS Vendor | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1301 |  |  | China | 鄄城马爱云妇儿医院 | 345044 | 44427 |  | 02N/Ingenuity Core128/Elite China/Change the information | BU Q&R Feedback: No SAP Customer Symptom Code: HANG Customer Symptom Code: Hang/Crash Reason: Corrective Maintenance Engineer Failure Code: Configuration Error Repair Activity Code: Configuration Change Problem Reported by Customer : Change the information Safety Questions : Case: 0118161710 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Qian Wang Last updated by: Qian Wang T2 Activities : Problem Description: 客户需要修改医院抬头 Troubleshooting Action: 远程告诉客户如何修改医院名称在软件上。 Repair Action: 远程告诉客户如何修改医院名称在软件上。 Test/Inspection: customer confirm can scan patient and image ok.functional Check PASSED. Device returned to full functionality. Test Equipment: N/A | 3 |  | 远程告诉客户如何修改医院名称在软件上。 | [12 empty] | 118161710 |  | 1236792 |  |  | Opened - Pending Translation | [7 empty] | Application | [6 empty] | NAC | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1302 |  |  | Japan | 独立行政法人国立病院機構 | 341013 | 44428 |  | 01DS: KKI: Cannot log in to Examcard Manager | Work Order #: WO-04157913 20-Aug-2021 Yamamoto, Yuko The system meets the specification for the performed service and is returned to use Resolution: Reset the IT Password and reset the Exam card editing password. | 5 |  | Reset the IT Password and reset the Exam card editing password. | [12 empty] | 118167220 |  | 1240804 |  |  | Evaluation - Pending Translation | [7 empty] | Customer | [6 empty] | Password | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1303 |  |  | Japan | Independent administrative institution National Hospital Organization | 341013 | 44427 |  | 03CS:TM:Via FSE.It stopped due to an error during SortTC. >> Can be re-implemented. >> SoftArcRe/"S_GEN_FAUL [Truncated] | Integration : Confirmation of order 54043997 saved ----------------------------------------------------------------------------- Work Order #: WO-04152996 Miyakawa, Takeyuki The system meets the specification for the performed service and is returned to use. Problem Description by Engineer : via FSE.It stopped due to an error during SortTC. >> Can be re-implemented. >> SoftArcRe/"S_GEN_FAULT_ANODE_TUBE_ARC/S_GEN_FAULT_CATHODE_TUBE_ARC" Diagnostic performed by Engineer : via FSE.It stopped due to an error during SortTC. >> Can be re-implemented. >> SoftArcRe/"S_GEN_FAULT_ANODE_TUBE_ARC/S_GEN_FAULT_CATHODE_TUBE_ARC" Internal Comments: [Subject] 03CS:TM:Via FSE:Via FSE.It stopped due to an error during SortTC. >> Can be re-implemented. >> SoftArcRe/"S_GEN_FAULT_ANODE_TUBE_ARC/S_GEN_FAULT_CATHODE_TUBE_ARC" [Internal Comments] Results of checking the log: 2021/08/19_07:53:32.059 <*HV*> szInfo: SoftArcReg 0x211* P1:0/P2: 2021/08/19_07:53:32.106 <*HV*> Type: NORMAL ErrNum: S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT szInfo: P1:0/P2: 2021/08/19_07:53:32.124 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_ANALOG_FAULT szInfo:2:5 Analog Fault P1:2/P2:32 2021/08/19_07:53:32.126 [CANCEL] CStateShot:Series 3 cancelled, due to GANTRY_CANCEL_SOURCE_NORMAL_ERROR, in state CStateShot 2021/08/19_07:53:32.134__03 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_ANODE_TUBE_ARC szInfo: 5:0 APM Arc P1:5/P2: 2021/08/19_07:53:32.139 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_CATHODE_TUBE_ARC szInfo: 5:5 CPM Arc P1:5/P2: 2021/08/19_07:53:34.164__01 <*ERROR*> Type: NORMAL ErrNum: S_MASTER_SEQUENCER_EXIT_STATE_ERR P1:4096/P2:8369 [4/8 Elements] Q5. How did the examination end? None. Q6. Did any mistaken findings (mistaken diagnosis or treatment) result from the problem? No. Q7. Was there any unintended unnecessary exposure? No Q8 Artifacts? None ----------------------------------------------------------------------------- | 4 |  | It stopped due to an error during SortTC. >> Can be re-implemented. | [12 empty] | 118161118 |  | 1234044 |  |  | Closed - No Further Investigation |  | CTUDT00630085 | [5 empty] | Engineering | [6 empty] | Tube Arc | [9 empty] | System | IC-HW | Tub Arc | IC-HW-Tub Arc | IC-HW-Tub Arc | CTUDT00630085 | Tub Arc | IC-HW | IC-HW-Tub Arc | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1304 |  |  | Japan | 独立行政法人国立病院機構 | 341015 | 44431 |  | 01CS: RS: The operation became slow, such as updating the tracker image on the Bolus Track around 11:30 on the 18th. | "Work Order #: WO-04172697 Satake, Rui The system meets the specification for the performed service and is returned to use Problem Description by Engineer: The operation became slow, such as updating the tracker image on the Bolus Track around 11:30 on the 18th. Resolution: There was no default log left, but there were no noticeable errors in the Logger or Host events. In the back grant, the tracking was triggered and the scan was performed. Consider the possible cause of poor host performance. Please note that there is no reproducibility after the 18th, so please follow up. Internal Comments: [Subject] 01CS: RS: Mr. Hisaki: The operation became slow, such as updating the tracker image on the Bolus Track around 11:30 on the 18th. [Internal Comments] Error code: None [4/8 Elements] Q5 How did the test end ?: None Q6 Presence or absence of incorrect findings (misdiagnosis, mistreatment) due to problems: None Q7 Presence or absence of unintended unnecessary exposure: None Q8 Presence / absence of artifacts: None | 4 |  | Consider the possible cause of poor host performance. Please note that there is no reproducibility after the 18th, so please follow up. Keep monitoring | [12 empty] | 118180196 |  | 1244172 |  |  | Opened - Pending Translation |  | TBD-SW-Slow | [5 empty] | Engineering | [6 empty] | System slow | [9 empty] | System | Console-SW | Performance | Console-SW-Performance | Console-SW-Performance | TBD-SW-Slow | Performance | Console-SW | Console-SW-Performance | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1305 |  |  | Japan | Maki Hospital | 341017 | 44427 |  | 01DS:KKI: When two orders are raised for the same patient on the MWM, both are selected. | Work Order #: WO-04152337 20-Aug-2021 Yamamoto, Yoshiyuki The system meets the specification for the performed service and is returned to use. Regulatory Questions : N/A Problem Description by Engineer: Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : Work details: When data was imported from the worklist, two items of data were automatically selected and after one of those was selected, if it was double-clicked without moving the cursor, with the first click the data was reselected, and with the second click it was retrieved. This is why two items of data were retrieved. As the data is automatically selected at the point when it appears on the worklist, please select the data that is not necessary and then retrieve the necessary data. Internal Comments: [CX related] Did you agree on how to complete the report with the person in charge of devices? Yes. I completed normal reporting that day. [Work report (paper) submission] Did you submit a work report on paper? <Yes> Q1. How did the examination finish? ・As it stopped before X-ray emission, no extra irradiation Q2. Were there any mistaken findings (misdiagnosis, mistaken treatment) due to the problem? None. Q3. Were there any artifacts? None. Q4. Was there any unintended unnecessary exposure? None. ----------------------------------------------------------------------------- | 4 |  | When data was imported from the worklist, two items of data were automatically selected and after one of those was selected, if it was double-clicked without moving the cursor, with the first click the data was reselected, and with the second click it was retrieved. This is why two items of data were retrieved. Try to reproduce the issue | [12 empty] | 118162946 |  | 1237062 |  |  | Evaluation |  | TBD-WorklistUsage | [5 empty] | Engineering | [6 empty] | Cursor | [9 empty] | System | Console-SW | WorklistUsage | Console-SW-WorklistUsage | Console-SW-WorklistUsage |  |  | Console-SW | Console-SW-WorklistUsage | No | 0 | 0 | 0 | 0 | 0 | No | y | y | y | 0 | 0 | y | y | y | 0 | y | #N/A
1306 |  |  | Japan | Maki Hospital | 341017 | 44424 |  | AF: An artifact appears on images. | Work Order #: WO-04125507 Ohyama, Tetsuo Regulatory Questions : N/A Problem Description by Engineer: Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer: N/A ----------------------------------------------------------------------------- | 4 |  | Calibration and image data is available in One EMS, to be check with engeineering | [12 empty] | 118138756 |  | 1223001 |  |  | Investigation |  | Image Artifact | [5 empty] | Engineering | [6 empty] | Image Artifact | [9 empty] | System | TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | Image Artifact |  | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | #N/A
1307 |  |  | Peru | BANCO BBVA PERU | 40034 | 44426 |  | 40034-PROBLEMS TO ACCESS LAN AND REMOTE CONFIGURATION - REINSTALL SOFTWARE | 40034-PROBLEMS TO ACCESS LAN AND REMOTE CONFIGURATION - REINSTALL SOFTWARE | 5 |  | PROBLEMS TO ENTER THE LAN AND REMOTE SETUP - REINSTALL SOFTWARE need check with local service if it is really software issue | [12 empty] | 118160599 |  | 1236139 |  |  | Opened |  | CTUDT00631055 | [5 empty] | Engineering | [6 empty] | LAN Config | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Lan Config | Console-SW-Lan Config | Console-SW-Lan Config | CTUDT00631055 | Lan Config | Console-SW | Console-SW-Lan Config | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | 0 | 0 | y | #N/A
1308 |  |  | Sweden | Falu Lasarett | 40049 | 44427 |  | software and application hangup, bug rep done 09:00 today. | Work Order #: WO-04152219 BERG NIKLAS Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 3 |  | @Silva, need engineering investigation on the software hangup issue https://teams.microsoft.com/_#/files/%E5%B8%B8%E8%A7%84?threadId=19%3A76cc42a0c8fa4bd7877f7de12141f84c%40thread.tacv2&ctx=channel&context=118162766_software%2520and%2520application%2520hangup%2520bug%2520rep%2520done%25209AM%2520today&rootfolder=%252Fsites%252FBigBoreRTRefreshFOKCoreTeam%252FShared%2520Documents%252FGeneral%252FIngenuity%2520Bugreps%252F118162766_software%2520and%2520application%2520hangup%2520bug%2520rep%2520done%25209AM%2520today | [12 empty] | 118162766 |  | 1237645 |  |  | Opened - Pending Translation |  | CTUDT00632520 | [5 empty] | Engineering | [6 empty] | Software hang | [9 empty] | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1309 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 337000 | 44423 |  | E-stop is not turning off. | Work Order #: WO-04124886 17-Aug-2021 Can Özgenlik The system meets the specification for the performed service and is returned to use Problem Description by Engineer : estop does not close Resolution: All estop buttons on the device were checked, no problem was found. The front and rear sides of the gantry were examined and all cards and leds were found to be normal. In the errors, it was seen that there was a ct box problem in detail. When the CTbox was opened, a socket was found to be loose. After the socket was inserted, it was seen that the key was turned without any problem and the estop circuit was closed. The test shot was taken with Pahntom and it was delivered in an active way. | 3 |  | In the errors, it was seen that there was a ct box problem in detail. When the CTbox was opened, a socket was found to be loose. After the socket was inserted, it was seen that the key was turned without any problem and the estop circuit was closed. | [12 empty] | 118138256 |  | 1222547 |  |  | Evaluation - Pending Translation | [7 empty] | Others | [6 empty] | E-Stop | [12 empty] | - | - |  |  | 0 | - | No | y | y | y | 0 | 0 | No | y | y | y | y | 0 | n | y | y | y | y | #N/A
1310 |  |  | United States | McLaren Macomb | 40021 | 44427 |  | Tech support | Work Order #: WO-04155045 Powell, Jamie The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Diagnostic performed by Engineer : Issue with getting EKG wave on the CT.. It works fine on the monitors. Customer has ran ibox diag (which passes), checked ekg monitor and ohmed out cables, reseated usb extenders but still having issues. Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Follow Up Required : Issue with getting EKG wave on the CT.. It works fine on the monitors. Customer has ran ibox diag (which passes), checked ekg monitor and ohmed out cables, reseated usb extenders but still having issues. | 3 |  | Check I-Box firmware & software. Program and update I-Box | [12 empty] | 118165167 |  | 1235433 |  |  | Investigation |  | TBD-NoEnoughInfo | [5 empty] | Engineering | [6 empty] | IBOX | [9 empty] | System | Gantry-HW | TBD-NoEnoughInfo | Gantry-HW-TBD-NoEnoughInfo | Gantry-HW-TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo | Gantry-HW | Gantry-HW-TBD-NoEnoughInfo | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1311 |  |  | United States | Unity Health Systems | 40037 | 44427 |  | 0699CFXXX gantry cannot comply shut down the system | Problem Description by Engineer : Problem description by engineer : gantry cannot comply shut down the system Work Order #: WO-04154954 19-Aug-2021 Shultz Bob The system does not meet the specification for the performed service and is not returned to use Diagnostic performed by Engineer : Acq system can’t comply Problem Description by Engineer : Problem description by engineer : gantry cannot comply shut down the system When was the first occurrence? Within... : 19-Aug-2021 How often is issue occurring? : Unknown Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: No further information is available as per the customer Patient Impact/Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer Current Software Version: Not provided Follow Up Required : Replace xmitter or receiver ----------------------------------------------------------------------------- Work Order #: WO-04154724 Atz, Michael The system meets the specification for the performed service and is returned to use Resolution : customer rans test scans and the system is now operational Internal Comments : Error # and/or description of error : S_DMC_SHOT_COMM_COLLAPSE NORMAL Explanation no frames were generated by DMS during scan. Watch for more occurrences of this error. It is possible to get infrequent occurrences of this error while the scanner is operating correctly. Repair Action : gantry cannot comply shut down the system to the wall ran a STC Diagnostic performed by Engineer : gantry cannot comply shut down the system to the wall Problem Description by Engineer : Problem description by engineer : gantry cannot comply shut down the system When was the first occurrence? Within... : 19-Aug-2021 How often is issue occurring? : Unknown Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: No further information is available as per the customer Patient Impact/Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer Current Software Version: Not provided Follow Up Required : CCB they would like for a FSE to come onsite ----------------------------------------------------------------------------- | 3 |  | Part/Product - MODULE, ROTOR ELECTRONICS TX Part Number - 453567538334 Part/Product - Module, Optical RX Stator Moog Velocity Part Number - 453567538322 | [12 empty] | 118164944 |  | 1235004 |  |  | Investigation |  | TBD_OBS | [5 empty] | Engineering | [6 empty] | TX & RX | [9 empty] | System | Gantry-HW | Other | Gantry-HW-Other | Gantry-HW-Other | TBD_OBS | TBD-NoEnoughInfo | Gantry-HW | Gantry-HW-Other | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1312 |  |  | United States | Philips Medical Capital | 40055 | 44425 |  | 0609CDN83 Noise coming from gantry | Work Order #: WO-04139719 20-Aug-2021 Turley Rich The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: Noise coming from gantry User Impact: None Patient Impact: None Current Software Version: Resolution : Checked system found warning label detached and rubbing rotor. Had issues restarting system had to shutdown to wall wait 5 minutes allow cirs and host to come fully up before starting gantry ----------------------------------------------------------------------------- Work Order #: WO-04139671 Boyd Craig The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: Noise coming from gantry User Impact: None Patient Impact: None Current Software Version: Follow Up Required : FSE needed on site to evaluate noise coming from gantry. Diagnostic performed by Engineer : Clicking noise coming from gantry when rotating. Started this morning. FSE, Rich Turley, is aware of issue. ----------------------------------------------------------------------------- | 3 |  | Checked system found warning label detached and rubbing rotor. Had issues restarting system had to shutdown to wall wait 5 minutes allow cirs and host to come fully up before starting gantry | [12 empty] | 118151226 |  | 1226084 |  |  | Evaluation | [7 empty] | Operation | [6 empty] | Gantry Noise | [18 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1313 |  |  | United States | Lakeland Regional Medical Center | 300143 | 44423 |  | Delays in scanner in regards to timing. | Internal Remarks : Interaction Type: New Case Caller: Mark Germany IP & Issue: 55151787. Delays in scanner in regards to timing. Event Type: Incident Service Request: Onsite Did Customer Decline Remote Support? No Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Created case and onsite workorder. Next Step Action: N/A Verbal PO Number: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- |  |  | Per customer issue resolved with hard reboot- cancel call | [12 empty] | 118138318 |  | 1222770 |  |  | Investigation | [20 empty] | Non-Refresh | [12 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y |  |  | #N/A
1314 |  |  | China | 河南许昌市中心医院(传染病) | 40047 | 44435 |  | 02N/Ingenuity CT 728326/couch can not move up and down | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Replace Part Problem Reported by Customer : 02N/Ingenuity CT 728326 SAP Transfer : Dispatched for Onsite Support Safety Questions : Case: 0118234247 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Gaoge Last updated by: Wang Gaoge 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 扫描床垂直运动故障。 客户问题详细描述或故障现象： 床不能正常上升 设备本身技术问题： 是 Troubleshooting Action: 关键故障代码或现象： 按生降床按钮床没有反应，水平移动正常 远程诊断与测试结果： invert 故障 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 1，建议客户先重启一下机架尝试 2 更换vertical controller 建议备件12NC及名称： vertical control （459800623432） Test/Inspection: N/A Test Equipment: N/A SAP Transfer : Case Transferred to SAP Next Step Action : PPD 459800623432 *1, 已安排24H发货，FSE黄勇-86-18039573296 Internal Remarks : 29.08.2021 11:58:23 UTC Yong Huang (86038368) RDI Accuracy: Yes $Q7: (Confirm reported problem,the couch cannot up and down ) $Q8: (Performed visual inspection ,found CCC display 02， need to reset the CCC ) $Q9: (Reset the CCC ) $Q10: (Test and inspection data refer to: PA. Device returned to full functionality.) $Q11: Calibrated Tools (N/A); External Remarks : 29.08.2021 11:58:20 UTC SUP JCOUSER (86038368)服务内容：[床不能升降，检查发现床控制板故障，重刷床控制板] 服务结果：[维修后，设备恢复正常工作] 其他重要事项：[无] ----------------------------------------------------------------------------- | 3 |  | 服务内容：[床不能升降，检查发现床控制板故障，重刷床控制板] 服务结果：[维修后，设备恢复正常工作] check with field to get log for analysis | [12 empty] | 118234247 |  | 1264588 |  |  | Opened - Pending Translation |  | NotEngineering | [5 empty] | Engineering | [6 empty] | Vertical movement | [9 empty] | System | Couch-SW | Fwupdate | Couch-SW-Fwupdate | Couch-SW-Fwupdate | NotEngineering |  | Couch-SW | Couch-SW-Fwupdate | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | n | y | y | 0 | y | #N/A
1315 |  |  | China | 山东省泰安市中医二院 | 345037 | 44435 |  | 02N/Ingenuity Core128/Elite China/scan error | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : scan error Next Step Action : FSE 陈涛 018560023293 call in and disp him SAP Transfer : Dispatched for Onsite Support External Remarks : 28.08.2021 12:20:51 UTC Tao Chen (86038373)故障现象: CT扫描间歇性报错 服务内容: 检查并重启CIRS和CT host计算机 服务结果: 维修后,设备恢复正常. Next Step Action : chen tao Safety Questions : Case: 0118232585 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Qian Wang Last updated by: Qian Wang 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 扫描失败。 客户问题详细描述或故障现象： 机架扫描中断。 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 已和TSE确认。Error data path 远程诊断与测试结果： 重启无效。 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议根据TSE建议维修故障。 建议备件12NC及名称： N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 5 |  | 故障现象: CT扫描间歇性报错 服务内容: 检查并重启CIRS和CT host计算机 服务结果: | [12 empty] | 118232585 |  | 1264013 |  |  | Opened - Pending Translation |  | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Reboot | [9 empty] | System | TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo |  |  |  |  | No | 0 | 0 | 0 | 0 | 0 | No | y | y | y | y | y | n | y | y | 0 | y | #N/A
1316 |  |  | China | 陕西省咸阳中心医院 (咸阳市第二人民医院） | 40050 | 44435 |  | wechat:02N/Ingenuity CT 728326/Abnormal sound is heard in the rack | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : Abnormal sound is heard in the rack Safety Questions : Case: 0118232297 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Zixuan Last updated by: Wang Zixuan 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 机架旋转有噪声，一个月前，报修同样问题 设备本身技术问题： 是。 Troubleshooting Action: 远程诊断与测试结果： 联系工程师郭宇杰，需要调整皮带 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议工程师到场检查皮带，教客户使用飞利浦微信报修 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 30.08.2021 01:19:09 UTC yifei zhang (86039086)服务原因：机架异响 服务内容：调整张紧轮位置 服务结果：设备工作正常 Internal Remarks : 30.08.2021 01:19:12 UTC yifei zhang (86039086) $Q7: (Confirmed reported problem. & The bearing have hardware problem ) $Q8: (1.Check Belt& We need to remove the tensioning wheel) $Q9: (1. remove the tensioning wheel) . 2. Test scan for CT was OK. ) $Q10: (Test and inspection data refer to PA. & Device returned to full functionality.) $Q11：Calibrated Tools name (N.A.） ----------------------------------------------------------------------------- | 3 |  | $Q7: (Confirmed reported problem. & The bearing have hardware problem ) $Q8: (1.Check Belt& We need to remove the tensioning wheel) $Q9: (1. remove the tensioning wheel) . 2. Test scan for CT was OK. ) | [12 empty] | 118232297 |  | 1263862 |  |  | Opened - Pending Translation | [7 empty] | Engineering | [6 empty] | Belt ajustment | [9 empty] | System | Gantry-HW | Belt | Gantry-HW-Belt | Gantry-HW-Belt | ECR-077012 | Belt | Gantry-HW | Gantry-HW-Belt | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | n | y | y | y | y | #N/A
1317 |  |  | China | Dongguan City Integrated Traditional Chinese and Western medicine Hospital | 345013 | 44433 |  | WeChat：02N/Ingenuity Core128/Elite China/Failure to connect | T2 Activities:Problem Description: Summary of customer_s problems: There is no ECG signal. Detailed description of Customer_s problems or malfunctions: There is no ECG signal. It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: N/A Remote troubleshoot and test result: An I-Box malfunction is suspected. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Check the I-Box input power supply and the I-Box board. Replace if necessary. We recommend spare parts 12NC and names: I-Box-2 Base board assy：459800166651 Memo of Customer request: None SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Next Step Action: PPD part: 459800166651 arranged, 24H to Dongguan City Hospital of Integrated Chinese and Western Medicine, Li Zongxian - 86-13751567086; please dispatch, thanks. External Remarks: 27.08.2021 04:46:10 UTC Zongxian Li (86001504) service details: replace ECG board. Service_s result: device operating normally Internal Remarks:27.08.2021 04:46:13 UTC Zongxian Li (86001504) RDI Accuracy:Yes $Q7: (Confirmed the reported problem, ibox disconnected) $Q8: (reboot the ibox, it didn_t respond to host) $Q9: (Replaced the ibox) $Q10: (Test and inspection data refer to PA. Device returned to full functionality._ $Q11: $Calibrated Tools name (N/A) S/N (N/A) Next Calibration date (N/A | 3 |  | I-Box-2 Base board assy：459800166651 check how FSE diagnost the IBOX issue? Have 5v dc cable? | [12 empty] | 118206661 |  | 1256414 |  |  | Opened |  | TBD-NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Ibox | [9 empty] | System | Gantry-HW | TBD-NoEnoughInfo | Gantry-HW-TBD-NoEnoughInfo | Gantry-HW-TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo | Gantry-HW | Gantry-HW-TBD-NoEnoughInfo | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | n | y | y | 0 | y | #N/A
1318 |  |  | China | 扎兰屯市人民医院 | 40045 | 44433 |  | 02N/Ingenuity CT 728326/report errors | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: Instructed Customer Problem Reported by Customer : report errors Safety Questions : Case: 0118204929 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wennian Han Last updated by: Wennian Han 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: report errors Troubleshooting Action: report that tube is too hot, room temperature is high, also use it with continuous high-dose scan Repair Action: advise customer to reduce room temperature, optimize scanning method, keep monitoring for a long time, CT is normal Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: NA SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | report that tube is too hot, room temperature is high, also use it with continuous high-dose scan Repair Action: advise customer to reduce room temperature, optimize scanning method, keep monitoring for a long time, CT is normal | [12 empty] | 118204929 |  | 1255447 |  |  | Closed - No Further Investigation |  | Not Engineering | [5 empty] | Engineering | [6 empty] | Room temp | [9 empty] | System | NotEngineering | ShouldnotEngineering | NotEngineering-ShouldnotEngineering | NotEngineering-ShouldnotEngineering | Not Engineering |  | NotEngineering | NotEngineering-ShouldnotEngineering | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | n | y | y | 0 | y | #N/A
1319 |  |  | China | Shaanxi Province Ankang T.C.M. Hospital | 345041 | 44432 |  | 02N/Ingenuity Core128/Elite China/couch can not move | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity Core128/Elite China SAP Transfer : Dispatched for Onsite Support SAP Transfer : Case Transferred to SAP Next Step Action : PPD 床拐角 CORNER, PEDESTAL COVERS, TOP 453567027541 1 已发货 张翼飞-86-17621702391 24H External Remarks : 30.08.2021 02:08:05 UTC yifei zhang (86039086)服务原因：偶发病床不能移动 服务内容：查看日志，后续观察使用 服务结果：设备工作正常 Internal Remarks : 30.08.2021 02:08:08 UTC yifei zhang (86039086) $Q7: (Confirmed reported problem. & couch prblem ) $Q8: (1.Check log& We need to observation use) $Q9: (1.Check log，observation use) . 2. Test scan for CT was OK. ) $Q10: (Test and inspection data refer to PA. & Device returned to full functionality.) $Q11：Calibrated Tools name (N.A.） Safety Questions : Case: 0118194524 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Dai Suyang Last updated by: Dai Suyang 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： ESTOP断开。 客户问题详细描述或故障现象： 床不能动，水平垂直都不行，重启后已恢复正常 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： PRS远程用户名密码登陆错误，尝试了remoteservice Philips1,philips_service service_only 都无法登陆 远程诊断与测试结果： 设备重启后已恢复正常，但是客户要求工程师到场查看原因。 PRS TR状态： bu通。 Repair Action: 建议现场维修步骤： 按照PRS手册重新配置PRS,重置密码remoteservice Philips1 。查看log信息，排除设备隐患，与客户做好良好沟通。 ----------------------------------------------------------------------------- | 3 |  | ESTOP断开。 客户问题详细描述或故障现象： 床不能动，水平垂直都不行，重启后已恢复正常,后续观察使用 PPD 床拐角 CORNER, PEDESTAL COVERS, TOP 453567027541 1 已发货 张翼飞-86-17621702391 24H check field to get bugreport/log | [12 empty] | 118194524 |  | 1250776 |  |  | Opened - Pending Translation | [7 empty] | Others | [6 empty] | Couch movement | bar: No; No
1320 |  |  | China | Fujian Province Medical University Affiliated No. 1 Hospital | 345012 | 44431 |  | 02N/Ingenuity Core128/Elite China/add AI | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : /add AI Safety Questions:Case: 0118178045 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Tao Fengming Last updated by: Tao Fengming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action: When requesting service on behalf of the customer, please make sure to provide: Device Number: 87219903 Contact Person: Radiology Department, Dr. Chen Ruiquan 18959128263 Hospital Name: The First Affiliated Hospital of Fujian Medical University Fault Details: Work together to add an AI node. Safety issue: None, there is no safety issue. Li Feng T2 Activities:Problem Description: Summary of customer_s problems: Other types of error. Detailed description of Customer_s problems or malfunctions: Add an AI node. It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: Add an AI node. Remote troubleshoot and test result: The customer_s olevel account can not be accessed, version 4.1.10 PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Reset the olevel_service account and record for later service use. Do the related node configuration according to the manual. SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Internal Remarks : 25.08.2021 06:42:19 UTC Feng Li (86036224) RDI Accuracy: Yes $Q7: (Confirmed reported problem.The Ingenuity Core128/Elite China need update time and add AI node.) $Q8: (On site to update Host & CIRS time. And Lan-config to add AI node.) $Q9: (On site to update Host & CIRS time. And Lan-config to add AI node.) $Q10: (Test and inspection data refer to: PA.+Device returned to full functionality.) $Q11: Calibrated Tools name (FLUKE 287C Multimeter) S/N (16490128) NextCalibration date (10/20/2021) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date: 09/16/2020 ^Current Tube usage: 36536 External Remarks: 25.08.2021 06:42:16 UTC SUP JCOUSER (86036224) Service Details: The on-board CT needed to have the time synchronized and add an AI node.On site, synchronized the time of the host computer and the reconstruction cabinet computer.Added the AI node transmission, testing is normal. service result: device returns to normal working condition after service ----------------------------------------------------------------------------- | 3 |  | The on-board CT needed to have the time synchronized and add an AI node.On site, synchronized the time of the host computer and the reconstruction cabinet computer.Added the AI node transmission, testing is normal. | [12 empty] | 118178045 |  | 1247274 |  |  | Closed - Non-Complaint | [14 empty] | Time sync | bar: No; No
1321 |  |  | China | Jiangxi Province Yugan Chu East Hospital (New Branch) | 337500 | 44431 |  | Wechat:02N/Ingenuity Core128/Elite China/Couch fault | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : Couch fault Safety Questions : Case: 0118177211 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Bai Jingwen Last updated by: Bai Jingwen 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Dispatched for Onsite Support T2 Activities : Problem Description: 客户问题概述： 床升降噪声 客户问题详细描述或故障现象： 床升降噪声 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： N/A 远程诊断与测试结果： N/A PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 1.确认噪音来源；2.润滑两侧滑轨、检查丝杆 建议备件12NC及名称： N/A 客户需求备注： 无 SAP Transfer : Case Transferred to SAP External Remarks : 24.08.2021 15:52:33 UTC Jiawei Li (86091272)服务内容： 床升降异响，检查床丝杆及两侧滑轨，按照维修手册，完成床两侧滑轨和丝杆润滑。异响消失。 服务结果：维修后，设备恢复正常使用。 Internal Remarks : 24.08.2021 15:52:35 UTC Jiawei Li (86091272) $Q1:(N) $Q2:(N) $Q3:(NA) $Q4:(NA) $Q5:(NA) $Q6:(Big noise during couch vertical move.) "RDI Accuracy:Yes" $Q7:(Confirmed reported problem.Big noise during couch vertical move.) $Q8:(Checked and found couch slide noise.) $Q9:(Lubricated couch slides solved the problem.) $Q10:(Test and inspection data refer to PA. Device returned to full functionality.) $Q11:(Calibrated Tools: No tools used.) ----------------------------------------------------------------------------- | 3 |  | 服务内容： 床升降异响，检查床丝杆及两侧滑轨，按照维修手册，完成床两侧滑轨和丝杆润滑。异响消失。 459801813061 - FRU, LUBRICANT, GADUS S3 V460D 2 is available in SPS, old stock will use up | [12 empty] | 118177211 |  | 1247067 |  |  | Opened |  | E-125059 | [5 empty] | Engineering | [6 empty] | Lubrication | [9 empty] | System | Couch-HW | Lubricant | Couch-HW-Lubricant | Couch-HW-Lubricant | E-125059 | Lubricant | Couch-HW | Couch-HW-Lubricant | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | n | y | y | 0 | y | #N/A
1322 |  |  | Italy | ASST SPEDALI CIVILI | 40026 | 44438 |  | Gantry non emette raggi. | Internal Remarks : Allego richiesta cliente per completamento pratica ----------------------------------------------------------------------------- Work Order #: WO-04239661 CENTEMERO, LORENZO The system does not meet the specification for the performed service and is not returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Customer reports scanning error. After running stop, the closing and stop mask no longer appears. From the initially error logs of GENERATOR ALE ERROR and after restarting RDTH HC TIMEOUT ERROR with RHOST_STARTUP_UNKNOWN__GENERATOR_ID. Suspected problem with RHOST or system control board, I order both. System stopped Follow Up Required: Customer reports scanning error. After running stop, the closing and stop mask no longer appears. From the initially error logs of GENERATOR ALE ERROR and after restarting RDTH HC TIMEOUT ERROR with RHOST_STARTUP_UNKNOWN__GENERATOR_ID. Suspected problem with RHOST or system control board, I order both. System stopped Diagnostic performed by Engineer: Scan problem confirmed, intervention will follow System cntrl board replaced and successful operation tests performed Product/Part - 80 KW System Control Board Part number - 453567560022 | 3 |  | System cntrl board replaced and successful operation tests performed Product/Part - 80 KW System Control Board Part number - 453567560022 | [12 empty] | 118243669 |  | 1267033 |  |  | Opened - Pending Translation |  | NotEngineering | [5 empty] | Engineering | [6 empty] | SCB | [9 empty] | System | Couch-HW | SCB | Couch-HW-SCB | Couch-HW-SCB | NotEngineering |  | Couch-HW | Couch-HW-SCB | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | n | y | y | 0 | y | #N/A
1323 |  |  | Japan | 社会福祉法人恩賜財団済生会 | 343000 | 44438 |  | 07CS:TM:use previous surviewの表示がされるようになった>>担当FSEへ報告。どのような操作による挙動か調査を行います | Work Order #: WO-04240829 Miyakawa, Takeyuki The system meets the specification for the performed service and is returned to use Diagnostic performed by Engineer : 担当FSEへ報告。どのような操作による挙動か調査を行います Problem Description by Engineer : use previous surviewの表示がされるようになった>>担当FSEへ報告。どのような操作による挙動か調査を行います Internal Comments : [Subject] 07CS:TM:モリワキ様:use previous surviewの表示がされるようになった>>担当FSEへ報告。どのような操作による挙動か調査を行います [Internal Comments] どのような操作による挙動か調査を行います [4/8 Elements] Q5 その検査はどの様に終了しましたか?:無しQ6 問題により間違った所見（誤診、誤治療）の有無:無し Q7 意図しない不要な曝射の有無:無し Q8 アーチファクトの有無:無し ----------------------------------------------------------------------------- | 5 |  | Keep monitoring | [12 empty] | 118241370 |  | 1267344 |  |  | Opened - Pending Translation | [7 empty] | TBD | [6 empty] | Surview | bar: No; No
1324 |  |  | Japan | Kinan Hospital | 340000 | 44434 |  | 01DS:RS: He set a high threshold and tried to start a bolus tracking scan manually but it did not work. | Work Order #: WO-04219242 Satake, Rui The system meets the specification for the performed service and is returned to use. Resolution: I did not find any hardware error in the log. It is possible that the application was not working properly. Monitoring the situation. Problem Description by Engineer: Set the threshold value high and tried to start BolusTrack manually, but there was no response. Internal Comments: [Subject] 01CS:RS:Mr./Ms. Kadosaka: Set the threshold value high and tried to start BolusTrack manually, but there was no response. [Internal Comments] Error code: None [4/8 Elements] Q5. How did the examination end? None. Q6. Did any mistaken findings (mistaken diagnosis or treatment) result from the problem? No. Q7. Was there any unintended unnecessary exposure? No Q8 Artifacts? None | 3 |  | Check with field to get bugreport for analysis | [12 empty] | 118226641 |  | 1260704 |  |  | Evaluation |  | Shouldbecustomer | [5 empty] | Engineering | [6 empty] | Bolus tracking | [9 empty] | System | Console-SW | ShouldbeCustomer | Console-SW-ShouldbeCustomer | Console-SW-ShouldbeCustomer | Shouldbecustomer |  | Console-SW | Console-SW-ShouldbeCustomer | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | n | y | y | 0 | y | #N/A
1325 |  |  | Japan | Maki Hospital | 341017 | 44435 |  | •What were they expecting the device to do? Get patient data with MWM connection | Check CFM detail with Peter | SHU, Tim |  | Kevin share the introduction to Peter/Tim and communicate with submitter and check if have detail step in IFU | [14 empty] | 1260215 |  |  | Investigation | [7 empty] | Application | [6 empty] | Worklist | bar: No; No
1326 |  |  | Japan | Independent administrative institution National Hospital Organization | 341015 | 44431 |  | 01CS:RS: Cannot transfer raw NMRS data to an external device. >> Restarting the host solved the problem. | cancel CASE due to a mistake in the site. | 4 |  | Mistake | [12 empty] | 118181278 |  | 1248107 |  |  | Closed - Non-Complaint | [14 empty] | NAC | bar: No; No
1327 |  |  | Japan | Tomakomai City Hospital | 341014 | 44432 |  | 02DS:KK: The error "Patient table cannot comply" appears frequently.>> It is possible to close the error and [Truncated] | Work Order #: WO-04179994 ----------------------------------------------------------------------------- Work Order #: WO-04179865 Kojima, Kenichi The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : The error "Patient table cannot comply" appears frequently.>> It is possible to close the error and continue the examination.>>S_HMC_ABSOLUTE_POSITION_ENCODER_FAILURE Internal Comments: [Subject] 02DS:KK:Mr. Masuda: The error "Patient table cannot comply" appears frequently.>> It is possible to close the error and continue the examination.>>S_HMC_ABSOLUTE_POSITION_ENCODER_FAILURE Shipped part: 453567399251:Assy, Absolute Encoder QTY:1 ZSID --None-- 453567399231:Assy, Position Encoder QTY:1 ZSID --None-- 453567099401:COUPLING/SPIDER QTY:1 ZSID --None-- Dispatched: 1 RC FSE Registration Information: Name of FSE 1 who visited the site: Hodaka Hashimoto FSE 1 visit date and time: 2021/08/2417:00 FSE 1 hours worked: 2 Diagnostic performed by Engineer : "S_HMC_ABSOLUTE_POSITION_ENCODER_FAILURE" had occurred. On-site reposnse by FSE. ----------------------------------------------------------------------------- | 4 |  | An error occurred during horizontal operation of the bed. I lost sight of the location information because the coupling connecting the units viewing the horizontal position information was cracked and slipped. We replaced the coupling and calibrated the location information. After the replacement, we checked the operation and confirmed that there was no problem. We performed a test scan and confirmed that there were no problems. [Parts used] Part number: <453567099401> / Part name: <COUPLING / SPIDER> / Serial number: <ー> / Batch number: <-> / Quantity: <1> | [12 empty] | 118195806 |  | 1246979 |  |  | Evaluation |  | NotEngineering | [5 empty] | Engineering | [6 empty] | Absolute encoder | [9 empty] | System | Couch-HW | Encoder | Couch-HW-Encoder | Couch-HW-Encoder | NotEngineering |  | Couch-HW | Couch-HW-Encoder | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | n | y | y | 0 | y | #N/A
1328 |  |  | Puerto Rico | METRO MAYAGUEZ INC | 40009 | 44438 |  | When trying to expose, user gets an error “acquisition cannot comply” and then “gantry can not comply” | Work Order #: WO-04243624 Caldelas Aldana, Carlos Manuel The system does not meet the specification for the performed service and is not returned to use Problem Description by Engineer : Problem description by engineer : Aborting aquisitions When was the first occurrence? Within... : 30-Aug-2021 12:00:00 How often is issue occurring? : First occurrence Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: aborting study Patient Impact/Outcome: cant perform estudy User Impact: cant perform aquisition Current Software Version: 4.1.10 Diagnostic performed by Engineer : validated that system is having error of aquisition cant comply. user mentions probably liquid or something went inside rotor. Analisis with remote conection of logs shows errors of refence detector errors. Internal Comments : Error # and/or description of error : S_DMC_REFERENCE_CONV_DATA S_DMC_DMS_VOLTAGE_OUT_OF_SPEC Repair Action : N/A Resolution : FSE Visit Follow Up Required : Error # and/or description of error : S_DMC_REFERENCE_CONV_DATA S_DMC_DMS_VOLTAGE_OUT_OF_SPEC Repair Action : N/A What are the next Troubleshooting steps : clean inside DMS and collimator to discard short circuit. Check if milar or any cover needs to be replaced (new follow up case if needed) Physically look at the RCOM LED’s reading, and/or measuring the PS’s output. Measure the (Vicors) PS output and Check the respective DMS Power Supply cabling. Replace as needed. Run (DMS) VD and check/follow-up accordingly with the results/outcome. If all the above does not address the issue, check / replace a motherboard Replace reference converter if needed replace reference detector if needed Check / Replace the RCOM. Technician Name : Caldelas Aldana, Carlos Manuel ----------------------------------------------------------------------------- Product/Part - REF DETECTOR ASSY Part number - 455012004371 Actual Qty - 1.0000 Product/Part - BR. REFERENCE CONVERTER ASSY ROHS Part number - 459800258961 Actual Qty - 1.0000 Visual inspection revealed liquid contras from the patient felt inside the gantry. As the gantry rotated, the contrast felt inside the gantry components damaging the Reference detector and the Reference detector converter. These parts have to be replaced. | 2 |  | The Reference Detector and the Reference Detector Converter were ordered. The parts were delivered to FedEx. The parts were installed. The CT was tested and left fully operational. | [12 empty] | 118246886 |  | 1268612 |  |  | Opened |  | NotEngineering | [5 empty] | Engineering | [6 empty] | REF DETECTOR | [9 empty] | System | Gantry-HW | Reference-Detector | Gantry-HW-Reference-Detector | Gantry-HW-Reference-Detector | NotEngineering |  | Gantry-HW | Gantry-HW-Reference-Detector | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | n | y | y | 0 | y | #N/A
1329 |  |  | Puerto Rico | DORADO HEALTH INC DBA | 40007 | 44432 |  | Noise from the couch during vertical movement | Work Order #: WO-04192386 Rios, William Problem Description by Engineer : Problem description by engineer : Noise during vertical movement. Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: No noise during vertical movement of the couch. Patient Impact/Outcome: Noise during movement. User Impact: Noise during movement. Current Software Version: 4.1.7 Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- Work Order #: WO-04191193 Valenzuela Velasquez, Eduardo Ariel The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Problem description by engineer : Noise during vertical movement. Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: No noise during vertical movement of the couch. Patient Impact/Outcome: Noise during movement. User Impact: Noise during movement. Current Software Version: 4.1.7 Follow Up Required : Error # and/or description of error : Noise from the patient support. Repair Action : - Check/lubricate couch vertical movement. Technician Name : Valenzuela Velasquez, Eduardo Ariel Internal Comments : Error # and/or description of error : Noise from the patient support. Repair Action : - Check/lubricate couch vertical movement. Diagnostic performed by Engineer : Customer reports internal noise during vertical movement of the patient support. The case will be sent to planning. ----------------------------------------------------------------------------- | 4 |  | Lubricate all patient support, perform functional test to the system. Unit OK 459801813061 - FRU, LUBRICANT, GADUS S3 V460D 2 is available in SPS, old stock will use up | [12 empty] | 118198669 |  | 1249298 |  |  | Evaluation |  | E-125059 | [5 empty] | Engineering | [6 empty] | Lubrication | [9 empty] | System | Couch-HW | Lubricant | Couch-HW-Lubricant | Couch-HW-Lubricant | E-125059 | Lubricant | Couch-HW | Couch-HW-Lubricant | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | n | y | y | 0 | y | #N/A
1330 |  |  | United States | Lackey Memorial Hospital | 40015 | 44431 |  | 0602CDN84 the gentry randomly turns off, lowering and raising the table/ stopped scanning during scan | Internal Remarks : Interaction Type: Customer Call Back Caller: Lane Sulton for Michea Mclemore IP & Issue: 85117259 ccb- requesting for service asap, pls call CT SCAN DEPT-5707703056 for ETA Event Type: Incident Service Request: Remote Technical Did Customer Decline Remote Support? No Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: case Next Step Action: N/A Verbal PO Number: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A Internal Remarks : FADB: (FSE Aware of Double Booking) Title: Internal Remarks Initials: Twila Actions Taken: Called Recommended Service Engineer: Denison Chuck - He accepted WO for today (8/23/21 @ 10am CST). Travel Time 1hr to Site Outcome: Dispatched per FSE_s Request IP: 85117259 Next Steps: None Next Step Action : Looks like Estop opened. In the bugrep it didn_t show why. The customer is scanning . FSe will monitor for errors. Internal Remarks : FADB Planner Initials: LV Actions Taken: Received call FSE Richard Turley, WO scheduled with FSE for 8/23at 2p CST, customer is aware Outcome: Dispatched Next Steps: N/A Last Updated: 5:58PM EST Internal Remarks : FADB: (FSE Aware of Double Booking) Title: Internal Remarks Initials: Twila Actions Taken: Dispatched per WO FU Note - Recommended Service Engineer: Denison Chuck Follow Up Date/Time Requested: 9/1/2021 2:00 AM CST Outcome: Dispatched per WO FU Note IP: 85117259 Next Steps: None ----------------------------------------------------------------------------- Work Order #: WO-04224112 25-Aug-2021 Denison Chuck Diagnostic performed by Engineer : N/A Problem Description by Engineer : Problem description by engineer : Customer states over the weekend the gantry shut off intermittently. When was the first occurrence? Within... : 22-Aug-2021 How often is issue occurring? : Intermittent Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Gantry turning off. Patient Impact/Outcome: Delay in having Scan User Impact: Delay in Scan Current Software Version: Not provided ----------------------------------------------------------------------------- Work Order #: WO-04197673 25-Aug-2021 Denison Chuck The system meets the specification for the performed service and is returned to use Diagnostic performed by Engineer : Wrong Part > not the correct Vertical Motor. Will ship back. Problem Description by Engineer : Problem description by engineer : Customer states over the weekend the gantry shut off intermittently. When was the first occurrence? Within... : 22-Aug-2021 How often is issue occurring? : Intermittent Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Gantry turning off. Patient Impact/Outcome: Delay in having Scan User Impact: Delay in Scan Current Software Version: Not provided Follow Up Required : Will try to order the correct Vertical motor for the patient support. ----------------------------------------------------------------------------- Work Order #: WO-04216194 26-Aug-2021 Donald Vic The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, no follow-up requested Problem Description by Engineer : Customer Function/Role: Tech Device Use: normal use Expected and Unexpected Results: the table is making a noise when raising the table User Impact: no impact Patient Impact: no impact Current Software Version: 4.1.7 Diagnostic performed by Engineer : The system is working and just making noise. We are sourcing a new motor for the vertical drive. ----------------------------------------------------------------------------- Work Order #: WO-04177875 25-Aug-2021 Turley Rich The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : assisted with diagnostics ----------------------------------------------------------------------------- Work Order #: WO-04177965 24-Aug-2021 Denison Chuck The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Problem description by engineer : Customer states over the weekend the gantry shut off intermittently. When was the first occurrence? Within... : 22-Aug-2021 How often is issue occurring? : Intermittent Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Gantry turning off. Patient Impact/Outcome: Delay in having Scan User Impact: Delay in Scan Current Software Version: Not provided Diagnostic performed by Engineer : Waited to get the C.T. scanner. Could not get work done to the Patient support because of covid patient. Follow Up Required : Need to change out the Patient support vertical motor assembly. ----------------------------------------------------------------------------- Work Order #: WO-04174624 23-Aug-2021 Denison Chuck The system meets the specification for the performed service and is returned to use Diagnostic performed by Engineer : Worked with Ronnie Dober NSS On logs. Gantry Error. We could not come up on what caused the error. Also worked on the Patient couch support. Needs motor replaced. Problem Description by Engineer : Problem description by engineer : Customer states over the weekend the gantry shut off intermittently. When was the first occurrence? Within... : 22-Aug-2021 How often is issue occurring? : Intermittent Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Gantry turning off. Patient Impact/Outcome: Delay in having Scan User Impact: Delay in Scan Current Software Version: Not provided Follow Up Required : Need to come back to site to replace the Vertical Motor and worm drive. ----------------------------------------------------------------------------- Work Order #: WO-04174128 Miller, John M The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Problem description by engineer : Customer states over the weekend the gantry shut off intermittently. When was the first occurrence? Within... : 22-Aug-2021 How often is issue occurring? : Intermittent Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Gantry turning off. Patient Impact/Outcome: Delay in having Scan User Impact: Delay in Scan Current Software Version: Not provided Follow Up Required : Error # and/or description of error : TmStamp 106389 ERR_TYPE NORMAL ErrNum S_MASTER_SEQUENCER_SHOT_OUT_OF_SEQUENCE_OR_ERR vxwErrNo 0 File S:\sy_bld_main_Bld_sn/stargate2/gantry/source/ghost/Sequencer/src/ShotTrans.cpp Line 774 Param1 128 Param2 0 Info EalTime 8During Shot state Param1 has a response bit map: Bit 0 (1) 0x01 XOFF_GUARD - if 1 the RHOST did not respond on a request from GHOST to Switch off X-ray. Bit 1 (2) 0x02 GRANT_GUARD - if 1 the RHOST did not respond on a request from GHOST to go to Grant State Bit 2 (4) 0x04 ARMED_GUARD - if 1 the RHOST did not respond on a request from GHOST to go to Armed State. Bit 3 (8) 0x08 FILAMENT_GUARD - if 1 the RHOST did not respond on a request from GHOST to preheat the filament or Filament timer was set twice. Bit 4 (16) 0x10 EXECUTING_GUARD. Bit 5 (32) 0x20 PRE_SHOT_DELAY_GUARD - If 1 Could mean that the RHOST did not responded for GEN_GRANT or GEN ARM. Bit 6 (64) 0x40 XON_GUARD - If 1 RHOST did not detect Xray ON Bit 7(128) 0x80 XARM_GUARD - If 1 the RHOST ARMED state timed out. Bit 8 (256) 0x100 DC_TAKESHOT_GUARD - If 1 the DMC did not start the shot. Bit 9 (512) 0x200 COUCH_GO_HOR_MOVE_GUARD - If 1 the couch did not move (correctly) during surview or Helix scan. 8/21/2021 11:27:55 PM TmStamp 204122 ERR_TYPE FATAL ErrNum S_MPL_VX_SEND_ERROR vxwErrNo 3997697 File S:\sy_bld_main_Bld_sn/stargate2/gantry/source/gaf/mpl/src/MPL.cpp Line 598 Param1 -1 Param2 429062 Info 0 EalTime 8/23/2021 2:38:40 AM ERR_TYPE NORMAL ErrNum S_MP_MPSHELL_SERIAL_LIFE_SYNC_ERROR vxwErrNo 3997697 File S:\sy_bld_main_Bld_sn/stargate2/gantry/source/mp/MPShell/src/MPShell.cpp Line 336 Param1 36700168 Param2 3997697 Info MP Latest Fatal Error Data. EalTime 8/23/2021 2:38:41 AM Repair Action : Life Ping between GHOST and GMP lost for some time (>2 seconds). Check laser cabling to center laser. Can be shorting out causing the Ghost to restart. Customer stated the couch was popping at the same time. Not sure they are related. Suggest checking for debris in the horizontal movement as well as covers binding. If the couch is hard to move at various points in horizontal travel, very possible the Kerk Screw and bearings are going bad and need to be replaced. Disconnect the linkage between the Kerk Screw drive and the table assembly to halve the problem. If it is the lead screw it could be the Kerk nut, the brake/clutch dragging/ or the thrust bearings needing replacement. If the table trolley is the issue it could be the linear bearings needing lubrication or replacement. Also check the Horizontal encoders for proper operation, tight mechanical connections and calibrations. Check the motor coupling between the encoders. Lubrication of Horizontal drive per PM manual. Check 5v switching power for at least 4.85v measured on the Couch Control Board. The power supply is adjustable. Refer to the Couch troubleshooting manual for further assistance. Technician Name : Miller, John M Internal Comments : Error # and/or description of error : TmStamp 106389 ERR_TYPE NORMAL ErrNum S_MASTER_SEQUENCER_SHOT_OUT_OF_SEQUENCE_OR_ERR vxwErrNo 0 File S:\sy_bld_main_Bld_sn/stargate2/gantry/source/ghost/Sequencer/src/ShotTrans.cpp Line 774 Param1 128 Param2 0 Info EalTime 8During Shot state Param1 has a response bit map: Bit 0 (1) 0x01 XOFF_GUARD - if 1 the RHOST did not respond on a request from GHOST to Switch off X-ray. Bit 1 (2) 0x02 GRANT_GUARD - if 1 the RHOST did not respond on a request from GHOST to go to Grant State Bit 2 (4) 0x04 ARMED_GUARD - if 1 the RHOST did not respond on a request from GHOST to go to Armed State. Bit 3 (8) 0x08 FILAMENT_GUARD - if 1 the RHOST did not respond on a request from GHOST to preheat the filament or Filament timer was set twice. Bit 4 (16) 0x10 EXECUTING_GUARD. Bit 5 (32) 0x20 PRE_SHOT_DELAY_GUARD - If 1 Could mean that the RHOST did not responded for GEN_GRANT or GEN ARM. Bit 6 (64) 0x40 XON_GUARD - If 1 RHOST did not detect Xray ON Bit 7(128) 0x80 XARM_GUARD - If 1 the RHOST ARMED state timed out. Bit 8 (256) 0x100 DC_TAKESHOT_GUARD - If 1 the DMC did not start the shot. Bit 9 (512) 0x200 COUCH_GO_HOR_MOVE_GUARD - If 1 the couch did not move (correctly) during surview or Helix scan. 8/21/2021 11:27:55 PM TmStamp 204122 ERR_TYPE FATAL ErrNum S_MPL_VX_SEND_ERROR vxwErrNo 3997697 File S:\sy_bld_main_Bld_sn/stargate2/gantry/source/gaf/mpl/src/MPL.cpp Line 598 Param1 -1 Param2 429062 Info 0 EalTime 8/23/2021 2:38:40 AM ERR_TYPE NORMAL ErrNum S_MP_MPSHELL_SERIAL_LIFE_SYNC_ERROR vxwErrNo 3997697 File S:\sy_bld_main_Bld_sn/stargate2/gantry/source/mp/MPShell/src/MPShell.cpp Line 336 Param1 36700168 Param2 3997697 Info MP Latest Fatal Error Data. EalTime 8/23/2021 2:38:41 AM Repair Action : Life Ping between GHOST and GMP lost for some time (>2 seconds). Check laser cabling to center laser. Can be shorting out causing the Ghost to restart. Customer stated the couch was popping at the same time. Not sure they are related. Suggest checking for debris in the horizontal movement as well as covers binding. If the couch is hard to move at various points in horizontal travel, very possible the Kerk Screw and bearings are going bad and need to be replaced. Disconnect the linkage between the Kerk Screw drive and the table assembly to halve the problem. If it is the lead screw it could be the Kerk nut, the brake/clutch dragging/ or the thrust bearings needing replacement. If the table trolley is the issue it could be the linear bearings needing lubrication or replacement. Also check the Horizontal encoders for proper operation, tight mechanical connections and calibrations. Check the motor coupling between the encoders. Lubrication of Horizontal drive per PM manual. Check 5v switching power for at least 4.85v measured on the Couch Control Board. The power supply is adjustable. Refer to the Couch troubleshooting manual for further assistance. Diagnostic performed by Engineer : Collect information from customer and view log files. ----------------------------------------------------------------------------- Product/Part - FRU,ASSY, VERTICLE MOTOR/BRAKE 20:1 Part number - 459801403271 Actual Qty - 1.0000 Parts Consumed Description: 1.00000,459801402611,BALL SCREW ASSEMBLY SERVICE KIT | 3 |  | Vertical drive check with field to get bugreport | [12 empty] | 118182951 |  | 1244897 |  |  | Investigation |  | NotEngineering | [5 empty] | Engineering | [6 empty] | Vertical movement | [9 empty] | System | Couch-HW | VerticalDriver | Couch-HW-VerticalDriver | Couch-HW-VerticalDriver | NotEngineering |  | Couch-HW | Couch-HW-VerticalDriver | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | n | y | y | 0 | y | #N/A
1331 |  |  | United States | AdventHealth Zephyrhills | 300106 | 44433 |  | Applaction issue with Brain Perfusion program not pulling up some of windows | Internal Remarks : Interaction Type: New Case Caller: James Golden IP & Issue: 54648842 Applaction issue with Brain Perfusion program not pulling up some of windows Event Type: Incident Service Request: Remote Clinical Did Customer Decline Remote Support? No Preferred Service Time: Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: Created Clinical W/O routed to Clinical Next Step Action: N/A Verbal PO Number: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- Work Order #: WO-04203601 26-Aug-2021 Garcia Marrero, Jorge The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, no follow-up requested Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: Applaction issue with Brain Perfusion program not pulling up some of windows User Impact: None Patient Impact: None Current Software Version: Diagnostic performed by Engineer : Collecting data from Brain Perfusion study to create a Customer Feedback case about Brain Perfusion application that fail to create a report correctly. Follow Up Required : Follow up to monitor performance of application ----------------------------------------------------------------------------- Work Order #: WO-04203560 Boyd Craig The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: Applaction issue with Brain Perfusion program not pulling up some of windows User Impact: None Patient Impact: None Current Software Version: Diagnostic performed by Engineer : Applications spoke with customer regarding issue and covered all bases related to it. PRS accessed. ULV: no normal errors seen. Follow Up Required : FSE needed on site to evaluate issue. Recommendations: Reload software. ----------------------------------------------------------------------------- Work Order #: WO-04203451 Mujica, David The system meets the specification for the performed service and is returned to use Problem description by Clinical Support Specialist : Customer Function/Role: Technologist Device Use: In use Expected and Unexpected Results: Color Maps Issue for BP User Impact: none Patient Impact: none Current Software Version: Diagnostic performed by Engineer : Russ indicated that the color maps is intermittently failing when post processing a CT Jog Brain Perfusion study. No errors were given on the scan portion, no errors given when loading the scan into the BP post processing software. The CT system_s local is optimal at 71% availability and daily shutdown_s with weekly complete power cycles are performed per Russ. RSE group involved, bug report captured. Case transferred to Tech for dispatch for remedy onsite. No further clinical support needed. No rescan, no reinjection per Russ. ----------------------------------------------------------------------------- | 5 |  | Collecting data from Brain Perfusion study to create a Customer Feedback case about Brain Perfusion application that fail to create a report correctly. TW PR ID 1254071 collect bugreport and detail information for engineering analysis | [12 empty] | 118222828 |  | 1254071 |  |  | Evaluation |  | CTUDT00633806 | [5 empty] | Engineering | [6 empty] | Brain Perfusion | [5 empty] | Refresh |  |  |  | System | Console-SW | Perfusion | Console-SW-Perfusion | Console-SW-Perfusion | CTUDT00633806 |  | Console-SW | Console-SW-Perfusion | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | n | y | y | 0 | y | #N/A
1332 |  |  | United States | AdventHealth Zephyrhills | 300099 | 44431 |  | GANTRY WAS TILTED AND E STOP OPENED - NOW UNABLE TO CLOSE E STOP | Internal Remarks : Interaction Type: Customer Call Back Caller: vinny IP & Issue: calling for FSE Dispatch tonight 8/23 Event Type: Incident Service Request: Onsite Did Customer Decline Remote Support? Yes Preferred Service Time: today 8/23 Entitlement Options: Entitled Reason for Entitlement Claim: Action Taken: created remote wo Next Step Action: N/A Verbal PO Number: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- Work Order #: WO-04197124 25-Aug-2021 Stillwell Philip The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: Technician Device Use: Clinical Expected and Unexpected Results: None User Impact: None Patient Impact: None Current Software Version: Resolution : Did not realize Jorge Garcia made it to site and got scanner working. Ran T/C, axial, & helical. ----------------------------------------------------------------------------- Work Order #: WO-04177610 24-Aug-2021 Garcia Marrero, Jorge The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: Diagnostics Expected and Unexpected Results: Tilted gantry to clean bottom and E Stop opened. Now unable to close E Stop even after reboot from wall. User Impact: No Patient Impact: No Current Software Version:4.0 Diagnostic performed by Engineer : Tilt limit switch in the right leg of the gantry was activated. Gantry was move out of the tilt limit switch manualy. System is fully operational. ----------------------------------------------------------------------------- Work Order #: WO-04177552 Dangos John The system does not meet the specification for the performed service and is not returned to use Diagnostic performed by Engineer : Tilted gantry to clean bottom and E Stop opened. Now unable to close E Stop even after reboot from wall. Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: Diagnostics Expected and Unexpected Results: Tilted gantry to clean bottom and E Stop opened. Now unable to close E Stop even after reboot from wall. User Impact: No Patient Impact: No Current Software Version:4.0 Follow Up Required : Left/Right Tilt Safety switch is activated and now unable to close E Stop. Release safety switch in service mode first. ----------------------------------------------------------------------------- | 2 |  | Tilt limit switch in the right leg of the gantry was activated. Gantry was move out of the tilt limit switch manualy. System is fully operational. | [12 empty] | 118193757 |  | 1245727 |  |  | Investigation | [14 empty] | Tilt limit switch | [5 empty] | Non-Refresh | bar: No; No
1333 |  |  | United States | Lakeland Regional Medical Center | 300143 | 44431 |  | The gantry is clicking when rotating | Integration : Confirmation of order 54061215 saved Integration : Confirmation of order 54061215 saved ----------------------------------------------------------------------------- Work Order #: WO-04175076 26-Aug-2021 Garcia Marrero, Jorge The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Problem description by engineer : The gantry is clicking when rotating When was the first occurrence? Within... : 23-Aug-2021 How often is issue occurring? : Continuously Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: No further information is available as per the customer Patient Impact/Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer Current Software Version: 4.1.7.25030 Diagnostic performed by Engineer : FOV mylar was creating tic noise inside the gantry. The FOV mylar was adjusted. System is fully operational. ----------------------------------------------------------------------------- Work Order #: WO-04175059 Atz, Michael The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, no follow-up requested Problem Description by Engineer : Problem description by engineer : The gantry is clicking when rotating When was the first occurrence? Within... : 23-Aug-2021 How often is issue occurring? : Continuously Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: No further information is available as per the customer Patient Impact/Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer Current Software Version: 4.1.7.25030 Diagnostic performed by Engineer : The gantry is clicking when rotating Internal Comments : Error # and/or description of error : none Repair Action : The gantry is clicking when rotating Follow Up Required : Error # and/or description of error : none Repair Action : The gantry is clicking when rotating Technician Name : Atz, Michael ----------------------------------------------------------------------------- | 3 |  | FOV mylar was creating tic noise inside the gantry. The FOV mylar was adjusted. System is fully operational. | [12 empty] | 118192581 |  | 1245021 |  |  | Evaluation | [14 empty] | Mylar | [5 empty] | Non-Refresh | bar: No; No
1334 |  |  | United States | Select Specialty Hospitals | 300037 | 44468 |  | warning message about cracked compensator | Internal Remarks : Title: Internal Remarks Initials: Twila Actions Taken: I called and left FSE a VM inquiring if he_s able to accept WO Outcome: N/A Next Steps: 9/29/21 @ 11:55am EST. Please call FSE/Beard Michael back in 20 min (12:15pm). Internal Remarks : FADB: (FSE Aware of Double Booking) Title: Internal Remarks Initials: Twila Actions Taken: Per message sent to me in Teams from FSE/Beard, Michael - Dispatch for 10/7/21. I Dispatched for 9am EST. Outcome: Dispatched WO per FSE_s Request IP: 53340864 Next Steps: None Internal Remarks : Phantom cals stopped with error. Had FSe reboot everything and it is working. Integration : Confirmation of order 54230698 saved ----------------------------------------------------------------------------- Work Order #: WO-04446472 05-Oct-2021 Ostapchenko Patrick The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Problem description by engineer : warning message about cracked compensator during air calibration Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: warning message about cracked compensator during air calibration Patient Impact/Outcome: none User Impact: none Current Software Version: 4.1.7.25037 Diagnostic performed by Engineer : moved adm mod 22 to slot 26, and 41 to 26. Perform cals, clean cone ring and test ok. Tools & Equipment : MULTIMETER,2022-03-31 00:00:00 ----------------------------------------------------------------------------- Work Order #: WO-04445475 Sheffield, Scott The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Follow Up Required : Error # and/or description of error : cracked compensator. warning Repair Action : There are 2 previous cases for this problem 0117903846 7/16/2021 0117779333 6/24/2021 Recommendation: The Auto bad det Signal test plot shows that detetor module 22 is low on the plot compaired to modules 21 and 23. I will attach the plots to the case. Use the compensator analizer tool and check compensator for crack when onsite. Swap detector module 26 with detector module 22 and do necessary calibrations. Refer to calibration matrix. Consult with NSS Sam Pierce for more information on repair. I was consulting with NSS Sam Pierce for the recommendations. Technician Name : Sheffield, Scott Internal Comments : Error # and/or description of error : cracked compensator. warning Repair Action : There are 2 previous cases for this problem 0117903846 7/16/2021 0117779333 6/24/2021 Recommendation: The Auto bad det Signal test plot shows that detetor module 22 is low on the plot compaired to modules 21 and 23. I will attach the plots to the case. Swap detector module 26 with detector module 22 and do necessary calibrations. Refer to calibration matrix. Consult with NSS Sam Pierce for more information on repair. I was consulting with NSS Sam Pierce for the recommendations. Diagnostic performed by Engineer : Brian said that during air cal he got message about cracked compensator. He is not seeing any artifacts on phantom images. They have not had any patients to scan, he is scanning the phantom with Clinical Applications support that is onsite. System is working at this time. I downloaded logs and plots and reviewed them. Problem Description by Engineer : Problem description by engineer : warning message about cracked compensator during air calibration Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: warning message about cracked compensator during air calibration Patient Impact/Outcome: none User Impact: none Current Software Version: 4.1.7.25037 ----------------------------------------------------------------------------- |  |  | may related to Beam and DMS check | [12 empty] | 118453627 |  | 1368962 |  |  | Evaluation |  | ECR-076862 | [5 empty] | Engineering | [6 empty] | Compensator | [5 empty] | Refresh |  |  | Upgrade | System | DMS-SW | Beam-DMSCheck | DMS-SW-Beam-DMSCheck | DMS-SW-Beam-DMSCheck | ECR-076862 |  | DMS-SW | DMS-SW-Beam-DMSCheck | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | #N/A
1335 |  |  | United States | Christ Hospital | 300205 | 44460 |  | 0798CDY72 Calcium Score Case - Intermittently Cannot Initialize | Work Order #: WO-04388474 22-Sep-2021 Hanlin Christopher The system meets the specification for the performed service and is returned to use Problem Description by Engineer : DESCRIPTION OF PROBLEM (EXPECTED/UNEXPECTED RESULTS) Calcium Score Case - Intermittently Cannot Initialize CALLER FUNCTION/ROLE Biomed DEVICE INTENDED USE Diagnostic WHEN WAS THE FIRST OCCURRENCE? WITHIN . . . <1 Week HOW OFTEN IS ISSUE OCCURRING? Intermittent WHERE IN THE END-USER WORKFLOW IS THE ISSUE OCCURRING? Post acquisition/scan WHAT RECENT EVENTS HAVE OCCURRED? During normal use NEGATIVE OPERATOR IMPACT Delay in scanning NEGATIVE PATIENT IMPACT/PROCEDURE OUTCOME Delay in being scanned HAS CUSTOMER CREATED A SPECIFIC SYSTEM EVENT REPORT OF ISSUE? Yes DEVICE SOFTWARE VERSION 4.1.10 Resolution : Examined error logger and parsed BUGREP. Issue caused by intermittent horizontal motor relay errors. Resubmitting plan works to work around issue. Issue traced to K103 Horizontal Motor Relay. ----------------------------------------------------------------------------- Work Order #: WO-04388306 Weaver David The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : DESCRIPTION OF PROBLEM (EXPECTED/UNEXPECTED RESULTS) Calcium Score Case - Intermittently Cannot Initialize CALLER FUNCTION/ROLE Biomed DEVICE INTENDED USE Diagnostic WHEN WAS THE FIRST OCCURRENCE? WITHIN . . . <1 Week HOW OFTEN IS ISSUE OCCURRING? Intermittent WHERE IN THE END-USER WORKFLOW IS THE ISSUE OCCURRING? Post acquisition/scan WHAT RECENT EVENTS HAVE OCCURRED? During normal use NEGATIVE OPERATOR IMPACT Delay in scanning NEGATIVE PATIENT IMPACT/PROCEDURE OUTCOME Delay in being scanned HAS CUSTOMER CREATED A SPECIFIC SYSTEM EVENT REPORT OF ISSUE? Yes DEVICE SOFTWARE VERSION 4.1.10 Diagnostic performed by Engineer : Calcium Score Case - Intermittently Cannot Initialize Follow Up Required : Calcium Score Case - Intermittently Cannot Initialize Cannot access system via PRS, must page out to FSE. Encouraged FSE to reach out to NSS for advice. Possible CT Host software reload. ----------------------------------------------------------------------------- |  |  | Resolution : Examined error logger and parsed BUGREP. Issue caused by intermittent horizontal motor relay errors. Resubmitting plan works to work around issue. Issue traced to K103 Horizontal Motor Relay. | [12 empty] | 118366055 |  | 1340567 |  |  | Investigation |  | NotEngineering-PartReplacement | [5 empty] | Engineering | [6 empty] | Hori Motor Relay | [5 empty] | Refresh |  |  | Upgrade | System | Couch-HW | Relay | Couch-HW-Relay | Couch-HW-Relay | NotEngineering-PartReplacement | Relay | Couch-HW | Couch-HW-Relay | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1336 |  |  | United States | Christ Hospital | 300205 | 44459 |  | 0699CFXXX Calcium Score Case - No Initialize [ PRS Fail ] | Integration : Confirmation of order 54186844 saved ----------------------------------------------------------------------------- Work Order #: WO-04381465 Carpenter, Kevin The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: Untrained biomed Cory Device Use: Diag Expected and Unexpected Results: Gantry would not Init for CS exam, until power cycled User Impact: None Patient Impact: None Current Software Version: IDK Diagnostic performed by Engineer : PRS/ FTP - problem - unable to get BugRep All ports scan green / open, and the FTP client does launch but the CT side of the FTP dialog never populates. Called biomed back ~ offered to dispatch, but he declined. Resolution : Untrained biomed declined to have FSE onsite to review bugrep as CT is now working. Also there are other biomed who are trained who could likely review the bugrep. ----------------------------------------------------------------------------- |  |  | Resolution : Examined error logger and parsed BUGREP. Issue caused by intermittent horizontal motor relay errors. Resubmitting plan works to work around issue. Issue traced to K103 Horizontal Motor Relay. | [12 empty] | 118360983 |  | 1337625 |  |  | Closed - No Further Investigation |  | NotEngineering-PartReplacement | [5 empty] | Engineering | [6 empty] | Hori Motor Relay | [5 empty] | Refresh |  |  | Upgrade | System | Couch-HW | Relay | Couch-HW-Relay | Couch-HW-Relay | NotEngineering-PartReplacement | Relay | Couch-HW | Couch-HW-Relay | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1337 |  |  | United States | AdventHealth Zephyrhills | 300106 | 44448 |  | Unable to reset password/ logged in correctly | Integration : Confirmation of order 54139504 saved Internal Remarks : Interaction Type: New Case Caller: Jessie Ros IP & Issue: 54648842 Unable to reset password/ log in correctly Event Type: Incident Service Request: Remote Clinical Did Customer Decline Remote Support? No Preferred Service Time: Regular Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: CASE Next Step Action: N/A Verbal PO Number: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- Work Order #: WO-04311218 Curry, Dyan The system meets the specification for the performed service and is returned to use Problem description by Clinical Support Specialist : Customer Function/ Role: technologist Device Use: in use name Expected and Unexpected Results: locked out of scanner User Impact: none Patient Impact: none Current Software version: 4.1. Diagnostic performed by Engineer : advised customer to wait 15 mins . Scanner will unlock and then enter correct password. Also advised to use Emergency Customer Function/ Role: technologist Device Use: in use name Expected and Unexpected Results: locked out of scanner User Impact: none Patient Impact: none Current Software version: 4.1. password to scan if needed until it unlocks ----------------------------------------------------------------------------- |  |  | Password | [12 empty] | 118304183 |  | 1302199 |  |  | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Password | [5 empty] | Refresh |  |  | Upgrade | bar: No; No
1338 |  |  | United States | AdventHealth Zephyrhills | 300099 | 44442 |  | 0711CDY73 Stuck in init. | Integration : Confirmation of order 54117038 saved ----------------------------------------------------------------------------- Work Order #: WO-04276721 03-Sep-2021 Sher Mark The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Problem description by engineer : Customer states the system was stuck in init state and wouldn_t scan. When was the first occurrence? Within... : 03-Sep-2021 How often is issue occurring? : First occurrence Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Stuck in init Patient Impact/Outcome: Delay in having Scan User Impact: Delay in Scan Current Software Version: 4.1.7.25030 Resolution : moved dvd ram option from lanconfig ----------------------------------------------------------------------------- Work Order #: WO-04276521 Miller, John M The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Problem description by engineer : Customer states the system was stuck in init state and wouldn_t scan. When was the first occurrence? Within... : 03-Sep-2021 How often is issue occurring? : First occurrence Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Stuck in init Patient Impact/Outcome: Delay in having Scan User Impact: Delay in Scan Current Software Version: 4.1.7.25030 Diagnostic performed by Engineer : Collect information from customer and view log files. Internal Comments : Error # and/or description of error : Type : Critical Date : 9/2/2021 Time : 10:26:53 PM Event : 41 Source : Microsoft-Windows-Kernel-Power Category : (63) User : \SYSTEM Computer : 128-CT-ED Description: The description for Event ID ( 41 ) in Source ( Microsoft-Windows-Kernel-Power ) could not be found. Either the component that raises this event is not installed on the computer or the installation is corrupted.You can install or repair the component or try to change Description Server. The following information was included with the event (insertion strings): 0 00000000 00000000 00000000 00000000 0 132751131434068928 0 0 0 0 0 No Gantry errors Repair Action : Suggest running the Dell or HP Diagnostic (Dell press F12 during bootup, HP press Esc then F2) to determine the health of the PC. Also check the health of the local drive. Also suggest re-seating memory modules. Suggest replacing hard drive and reload software on new drive. 459801652541 - HP TurboDrive 512GB TLC Z4 SSD w/adapter 1- Long Time to initialize: Is this problem fixed? How? If not, have you checked LAN config and removed DVDRAM from Local devices? Follow Up Required : Error # and/or description of error : Type : Critical Date : 9/2/2021 Time : 10:26:53 PM Event : 41 Source : Microsoft-Windows-Kernel-Power Category : (63) User : \SYSTEM Computer : 128-CT-ED Description: The description for Event ID ( 41 ) in Source ( Microsoft-Windows-Kernel-Power ) could not be found. Either the component that raises this event is not installed on the computer or the installation is corrupted.You can install or repair the component or try to change Description Server. The following information was included with the event (insertion strings): 0 00000000 00000000 00000000 00000000 0 132751131434068928 0 0 0 0 0 No Gantry errors Repair Action : Suggest running the Dell or HP Diagnostic (Dell press F12 during bootup, HP press Esc then F2) to determine the health of the PC. Also check the health of the local drive. Also suggest re-seating memory modules. Suggest replacing hard drive and reload software on new drive. 459801652541 - HP TurboDrive 512GB TLC Z4 SSD w/adapter 1- Long Time to initialize: Is this problem fixed? How? If not, have you checked LAN config and removed DVDRAM from Local devices? Technician Name : Miller, John M ----------------------------------------------------------------------------- |  |  | Long Time to initialize | [12 empty] | 118274392 |  | 1286066 |  |  | Closed - No Further Investigation |  | CTUDT00633734 | [5 empty] | Engineering | [6 empty] | Long Time to initialize | [5 empty] | Refresh |  |  | Upgrade | System | Console-SW | DVDRAM | Console-SW-DVDRAM | Console-SW-DVDRAM | CTUDT00633734 |  | Console-SW | Console-SW-DVDRAM | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | #N/A
1339 |  |  | China | Jiangsu Province Yangzhou City No. 1 People_s Hospital, West District Hospital | 345022 | 44469 |  | 02N/Ingenuity Core128/Elite China/key fault | T2 Activities:Problem Description: Summary of customer_s problems: Control panel button is not working. Detailed description of Customer_s problems or malfunctions: CTBOX buttons are not responsive. It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: NA Remote troubleshoot and test result: The customer said Down and Exit bed buttons on ctbox are not working properly. Down and Exit bed buttons on Gantry panel are normal PRS TR status: It is not through, ask engineer for maintenance on site. Repair Action: Suggest onsite service procedures: Check the CT BOX. We recommend spare parts 12NC and names: " 459801705022 CT BOX ASSEMBLY_BEST" Internal Remarks : 02.10.2021 04:53:01 UTC (86036024) "RDI Accuracy:NA" $Q7: (Confirmed reported problem: CTbox can not move couh down and back.) $Q8: (No error report. The down key and back key of ctbox are defective. Need to be changed.) $Q9: (Changed CTbox assembly_best and 12NC is 459801705022.) | 3 |  | Changed CTbox assembly_best and 12NC is 459801705022 | [12 empty] | 118462384 |  | 1375863 |  |  | Opened |  | CTUDT00632185 | [5 empty] | Engineering | [6 empty] | CT Box | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | CT Box | Gantry-HW-CT Box | Gantry-HW-CT Box | CTUDT00632185 | CT Box | Gantry-HW | Gantry-HW-CT Box | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1340 |  |  | China | Xifeng TCM Hospital | 345014 | 44469 |  | CT box of the CT5000 is of poor quality. | CT box of the CT5000 is of poor quality. 12NC:459801705022 A month ago,the down button of the CT box is out of order, so FSE replaced the CT box . But now,the old problem popped up again. At the moment ,we had to replace it again. So the customer complained about the poor quality of the product. The first broken CT box SN :107 The second broken CT box SN:313 |  |  | CT box of the CT5000 is of poor quality | [14 empty] | 1373289 |  |  | Investigation |  | CTUDT00632185 | [5 empty] | Engineering | [6 empty] | CT Box | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | CT Box | Gantry-HW-CT Box | Gantry-HW-CT Box | CTUDT00632185 | CT Box | Gantry-HW | Gantry-HW-CT Box | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1341 |  |  | China | Sichuan Province Yibin City No.2 T.C.M. Hospital | 345035 | 44468 |  | 02N/Ingenuity Core128/Elite China/ | T2 Activities : Problem Description: Scan interrupted by ARC Troubleshooting Action: Check logger APM ARC 5:0. No more repeat, do filament calibration and take monitoring. Repair Action: No more repeat, do filament calibration and take monitoring. | 3 |  | Perform filament calibration to resolve tube arc issue and keep monitoring. | [12 empty] | 118451184 |  | 1371429 |  |  | Evaluation |  | CTUDT00630085 | [5 empty] | Engineering | [6 empty] | Tube Arc | [5 empty] | Refresh |  |  | System issue | System | IC-HW | Tub Arc | IC-HW-Tub Arc | IC-HW-Tub Arc | CTUDT00630085 | Tub Arc | IC-HW | IC-HW-Tub Arc | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1342 |  |  | China | Laiyang City People_s Hospital | 345005 | 44468 |  | 05L/Ingenuity Core128/Elite China/worklist fault | External Remarks: 30.09.2021 01:31:08 UTC (86038250) Service details: worklist fail to refresh, reconfigure server node. service result: device returns to normal working condition after service Internal Remarks : 30.09.2021 01:31:11 UTC (86038250) "RDI Accuracy:yes" $Q7: (Confirmed reported problem & failed to refresh worklist) $Q8: ( check lan configuration & found worklist node lost) $Q9: ( reconfig His/Ris server node by lanconfig app ) $Q10: (Test and inspection data refer to: PA & Device returned to full functionality) | 3 |  | reconfig His/Ris server node via lanconfig | [12 empty] | 118447052 |  | 1370663 |  |  | Opened | [7 empty] | Others | [6 empty] | HIS/RIS Configuration | [5 empty] | Refresh |  |  | System issue | bar: No; No
1343 |  |  | Germany | Gesundheit Nord gGmbH Klinikverbund | 40002 | 44468 |  | The system sporadically makes a strange USB sound like Connect / disconnect | Problem Description by Engineer: intermittent disconnect USB can be heard. Probably IBOX CT console USB connection. Resolution: Voltage on the power supply for the I-Box was increased to 5.33VDC. In the tests performed, the system corresponds to the manufacturer_s specifications and is ready for clinical use. Tools & Equipment : MULTIMETER,2022-05-25 00:00:00 | 4 |  | intermittent disconnect USB | [12 empty] | 118451431 |  | 1367341 |  |  | Evaluation |  | TBD-NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Ibox USB | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | TBD-NoEnoughInfo | Gantry-HW-TBD-NoEnoughInfo | Gantry-HW-TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo | Gantry-HW | Gantry-HW-TBD-NoEnoughInfo | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1344 |  |  | Japan | Social Medical Corporation BOKOI, Nikko Memorial Hospital | 341011 | 44468 |  | 01CS: RS: When creating an MPR from an Axail image of the head, it is recognized that there are 2 series of data. | Problem Description by Engineer : When MPR view is reconstructed from head axial images, the data is recognized as being two series. Resolution : I asked him to use the Merge function to process the data that was recognized as being split into two series. I asked him to change the scan type to 2D axial or perform a helical scan. Internal Comments: [Subject] 01CS:RS:Mr. Sato: When MPR view is reconstructed from head axial images, the data is recognized as being two series. [Internal Comments] None [4/8 Elements] Q5. How did the examination end? None. Q6. Did any mistaken findings (mistaken diagnosis or treatment) result from the problem? No. Q7. Was there any unintended unnecessary exposure? No Q8 Artifacts? None | 4 |  | Q&A | [12 empty] | 118453258 |  | 1372503 |  |  | NAC | [7 empty] | Application | [6 empty] | Q&A - MPR | [5 empty] | Refresh |  |  | System issue | bar: No; No
1345 |  |  | Japan | Social Medical Corporation BOKOI, Nikko Memorial Hospital | 341011 | 44468 |  | 02CS: KKI: E-Stop open occurs when the bed is horizontally operated >> S_HMC_TAPE_SWITCH_PRESSED> Restored by system OFF / ON | Problem Description by Engineer: The E-stop opened when he was operating the bed horizontally. Resolution : Response Turning the system off/on achieved recovery. Internal Comments: [Subject] 02CS:KKI:Mr. Ando: The E-stop opened when he was operating the bed horizontally>>S_HMC_TAPE_SWITCH_PRESSED> It recovered when he turned the system off/on. [Internal Comments] [Results of checking the log] The following error occurred. S_HMC_TAPE_SWITCH_PRESSED It seems that the tape switch, Down button and Up button were pressed at the same time, and the emergency stop happened. I think it was a temporary problem, so I asked him to monitor the situation. [4/8 Elements] Q5. How did the examination end? None. Q6. Did any mistaken findings (mistaken diagnosis or treatment) result from the problem? No. Q7. Was there any unintended unnecessary exposure? No Q8 Artifacts? None | 4 |  | t seems that the tape switch, Down button and Up button were pressed at the same time, and the emergency stop happened. | [12 empty] | 118447376 |  | 1369140 |  |  | Investigation | [7 empty] | Others | [6 empty] | Misoperation | [5 empty] | Refresh |  |  | System issue | bar: No; No
1346 |  |  | Portugal | Centro Hospitalar de Vila Nova | 40030 | 44468 |  | Faulty southern philips TV equipment does not perform topograms or exams | Work Order #: WO-04443053 29-Sep-2021 Jaime Granja The system does not meet the specification for the performed service and is not returned to use Follow Up Required: Replacement of the optical transmitter on the rotating part of the gantry. Eventually the problem may be in the associated optical fibers. Diagnostic performed by Engineer: Detected malfunction in the detector data communication circuit (DMS) for the stationary part of the gantry. It is necessary to replace the transmitter and/or associated optical fibers. Problem Description by Engineer : Customer Function/Role: Maintenance Department Device Use: Diagnostics Expected and Unexpected Results: Unable to acquire images. User Impact: Examinations to be done on another system on the same hospital or postponed Patient Impact: Examinations to be done on another system on the same hospital or postponed Current Software Version: 4.1.10 ----------------------------------------------------------------------------- | 3 |  | Product/Part - Module, Optical RX Stator Moog Velocity Part number - 453567538322 Product/Part - Harness-Rotor TX Fiber Assy Part number - 453567538381 Product/Part - Stator Optics Assembly Part number - 453567538341 Product/Part - MODULE, ROTOR ELECTRONICS TX Part number - 453567538334 | [12 empty] | 118451509 |  | 1369206 |  |  | Evaluation |  | TBD_OBS | [5 empty] | Engineering | [6 empty] | TX/RX | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Other | Gantry-HW-Other | Gantry-HW-Other | TBD_OBS | TBD-NoEnoughInfo | Gantry-HW | Gantry-HW-Other | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1347 |  |  | China | DaKuai Town Health, Fengquan District, Xinxiang City, Henan Province | 345009 | 44467 |  | 02N/Ingenuity Core128/Elite China/Blue screen | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure T2 Activities:Problem Description: Blue screen Troubleshooting Action: Remotely instruct the customer to restart the system. Repair Action: Remotely instruct the customer to restart the system. Test/Inspection: Functional Check PASSED. The device returned to full functionality. Test Equipment: N/A SAP Transfer:Case Transferred to SAP | 3 |  | Blue screen reported from customer and system worked after restart the system. AI: Check with TSE and collect the bluescreen photo and bugreport for eng analysis | [12 empty] | 118439242 |  | 1366085 |  |  | Closed - No Further Investigation |  | CTUDT00632553 | [5 empty] | Engineering | [6 empty] | Blue Screen | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Blue-Screen | Console-SW-Blue-Screen | Console-SW-Blue-Screen | CTUDT00632553 |  | Console-SW | Console-SW-Blue-Screen | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | #N/A
1348 |  |  | Japan | Social Welfare Organization Saiseikai Imperial Gift Foundation, Inc. | 343000 | 44467 |  | 01DS: KKI: Contrast-enhanced tracker CT host freezes and scan does not proceed >> S_MASTER_SEQUENCER_ACQ_NOT_OK / Error # 12 --Unknown error … | Integration : Confirmation of order 54220738 saved ----------------------------------------------------------------------------- Work Order #: WO-04429715 Kimura, Kazumasa The system meets the specification for the performed service and is returned to use. Internal Comments: [Subject] 01DS:KKI:Mr. Moriwaki: After a contrast trucker scan, the CT host froze and the scan did not progress.>>S_MASTER_SEQUENCER_ACQ_NOT_OK/ Error #12 - Unknown error 0x12> Restarting the CT host restored the scan. [Internal Comments] [Results of checking the log] The following error occurred. S_MASTER_SEQUENCER_ACQ_NOT_OK From S1 Error #12 - Unknown error 0x12 I did not find an error that allowed me to identify the cause of the problem. Dispatched: 1 RC FSE Registration Information: Name of FSE 1 who visited the site: Yasuhiro Kawashige FSE 1 visit date and time: 2021/09/2816:00 FSE 1 hours worked: 2 Diagnostic performed by Engineer : Response: Logging out/in to the CT host solved the problem. Problem Description by Engineer: After a contrast trucker scan, the CT host froze and the scan did not progress. ----------------------------------------------------------------------------- Work Order #: WO-04431295 28-Sep-2021 Kawashige, Yasuhiro The system meets the specification for the performed service and is returned to use. Internal Comments: [CX related] Did you agree on how to complete the report with the person in charge of devices? <Select one of the following> Yes. I completed normal reporting that day. [Work report (paper) submission] Did you submit a work report on paper? NO Resolution : [Work details] I investigated the problem of a contrast scan examination not proceeding, but it was not reproduced. I checked the log entries for both systems (HOST PC, CIRS) when the problem occurred, but there was no particular clear error. However, the CIRS had not been restarted for more than a month, so it is possible that the CIRS operation had become sluggish. Therefore, I restarted the Host PC and CIRS onsite, and I carried out an IQ test and a contrast scan test. I carried out the above and confirmed normal system operation. I apologize for the inconvenience caused. | 4 |  | the CIRS had not been restarted for more than a month, so it is possible that the CIRS operation had become sluggish. Therefore, I restarted the Host PC and CIRS onsite, and I carried out an IQ test and a contrast scan test. | [12 empty] | 118439293 |  | 1362781 |  |  | Investigation | [7 empty] | Others | [6 empty] | Reboot - CIRS | [5 empty] | Refresh |  |  | System issue | bar: No; No
1349 |  |  | Japan | 独立行政法人国立病院機構 | 341013 | 44467 |  | 01CS: RS: Scan application crashed while shooting >> Scan was running as planned | Problem Description by Engineer: The scanning application crashed while shooting Resolution: The scan itself ran with the planned content. No noticeable hardware errors were found in the various logs. From the situation, it is assumed that the scanning application has crashed temporarily. Request a restart of the host computer and complete. Internal Comments: [Subject] 01CS: RS: Mr. Mamamoto: The scanning application crashed during shooting >> The scan was running as planned [Internal Comments] none [4/8 Elements] Q5 How did the test end ?: None Q6 Presence or absence of incorrect findings (misdiagnosis, mistreatment) due to problems: None Q7 Presence or absence of unintended unnecessary exposure: None Q8 Presence / absence of artifacts: None | 4 |  | "The scan itself ran with the planned content. No noticeable hardware errors were found in the various logs. From the situation, it is assumed that the scanning application has crashed temporarily. Request a restart of the host computer and complete | [12 empty] | 118440496 |  | 1362177 |  |  | Closed - No Further Investigation |  | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | App crash | [5 empty] | Refresh |  |  | System issue | System | TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | NoEnoughInfo |  | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | #N/A
1350 |  |  | United States | Philips Medical Capital | 40053 | 44467 |  | 0711CDN8A FSE req_d for DICOM problem | Internal Remarks : FADB: (FSE Aware of Double Booking) Title: Internal Remarks Initials: Twila Actions Taken: Dispatched to FSE/Fairchild, Andrew - (9/28/21 @ 3pm EST). Outcome: Dispatched WO per PFSE/McDowell Preston to FSE/Fairchild IP: 88993323 Next Steps: None **************************** Work Order #: WO-04433863 28-Sep-2021 Fairchild, Andrew The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Diagnostic performed by Engineer : Cust says FSE Andrew Fairchild is aware They decline T2 Remote Tech Support ~ FSE req_ed. Diagnostic performed by Engineer : Troubleshooted issues with PACS, can query and send but not retrieve. Customer is checking with IT to confirm their side is good. If issue persists, will contact tech support for assistance. Follow Up Required : Additional PACS troubleshooting ----------------------------------------------------------------------------- | 3 |  | Customer is checking with IT to confirm their side is good. If issue persists, will contact tech support for assistance. | [12 empty] | 118443199 |  | 1365766 |  |  | Evaluation | [7 empty] | TBD | [6 empty] | PACS Retrieve | [5 empty] | Refresh |  |  | System issue | bar: No; No
1351 |  |  | China | Yangfangdang Hospital | 40035 | 44466 |  | Wechat:02N/Ingenuity CT 728326/transfer fault | T2 Activities:Problem Description: Summary of customer_s problems: ISP cannot be used to transfer back and forth with CT Detailed description of Customer_s problems or malfunctions: ISP cannot be used to transfer back and forth with CT It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: Test CT connection in ISP, but configuration test failed Parameter record is ISP 172.20.110.25 CT 172.20.110.149 AE MXV Remote troubleshoot and test result: 1. confirm CT and ISP parameter setup, see if It is correct 2. Check network, check node type, re-add node SAP Transfer:Dispatched for On site Support External Remarks: 28.09.2021 07:53:59 UTC (86039460) Service details: ISP cannot transfer images with CT host to each other Service Result: The device returned to working normally. Internal Remarks : 28.09.2021 07:54:02 UTC (86039460) "RDI Accuracy:Yes" $Q7: (Confirmation & Description of problem) Confirmed reported problem: ISP can not deliver image to host. $Q8: (Troubleshooting & Result Check it, found have lots of image delivered with failure, $Q9: (Repair steps) Do reset spooler, and add ISP network node again, device work normal, | 3 |  | Do reset spooler, and add ISP network node again, device work normal, ISP can deliver image to host. AI: Check with TSE if It is CT or ISP issue. | [12 empty] | 118431322 |  | 1361845 |  |  | Closed - No Further Investigation | [7 empty] | TBD | [6 empty] | ISP Transfer | [5 empty] | Refresh |  |  | System issue | bar: No; No
1352 |  |  | Poland | Government Agency. | 341008 | 44466 |  | Turn off gantry when tilting. In addition, the need to configure the PACS server. | The gantry turns off during tilting. Additionally, it is necessary to configure the PACS server. |  |  | Duplicate | [14 empty] | 1384552 |  |  | Duplicate | [7 empty] | Duplicate | [6 empty] | Duplicate | [5 empty] | Refresh |  |  | System issue | bar: No; No
1353 |  |  | Poland | Government Agency. | 341008 | 44466 |  | Turn off gantry when tilting. In addition, the need to configure the PACS server. | PACS server configuration - correct. Tilting the gantry - the fuse was off. Enabled. Works properly. | 3 |  | unexpect fuse disconnect/E-stop open when tilting; config PACS | [12 empty] | 118435526 |  | 1363518 |  |  | Investigation |  | NotEngineering | [5 empty] | Engineering | [6 empty] | E-Stop - Tilting | [5 empty] | Refresh |  |  | System issue | System | NotEngineering | ShouldbeCustomer | NotEngineering-ShouldbeCustomer | NotEngineering-ShouldbeCustomer | NotEngineering | NotEngineering-ShouldbeCustomer | NotEngineering | NotEngineering-ShouldbeCustomer | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1354 |  |  | Puerto Rico | METRO MAYAGUEZ INC | 40009 | 44466 |  | When performing calibration of the machine indicates that there is a strange artifact, it may be that the contrast... | Problem Description by Engineer : Problem description by engineer : "When performing calibration of the machine indicates that there is a strange artifact, it may be that the contrast medium" When was the first occurrence? Within... : 27-Sep-2021 09:00:00 Diagnostic performed by Engineer : The customer reported artifacts when doing patient studies. The image calibration failed to complete. The gantry cover was open. I visually inspected the area. I found contras found under the plastic center cover. This contrast is what is causing the image artifact. The contrast has to be removed Resolution : The contrast spots were removed from the Millar. The calibration was completed successfully. The CT was returned to the customer. | 3 |  | The contrast spots were removed from the Millar | [12 empty] | 118436926 |  | 1360478 |  |  | Investigation | [7 empty] | Others | [6 empty] | Contrast - Mylar | [5 empty] | Refresh |  |  | System issue | bar: No; No
1355 |  |  | China | 江苏省第二中医院 | 345027 | 44465 |  | Wechat:02N/Ingenuity Core128/Elite China/Couch fault | BU Q&R Feedback: No Reason: Corrective Maintenance Safety Questions : Case: 0118428950 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Problem Reported by Customer : Couch fault T2 Activities : Problem Description: 客户问题概述： 其他类型故障。 客户问题详细描述或故障现象： 床卡死 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 床卡死动不了 远程诊断与测试结果： 客户说断电情况下床也推不动，应该是被撞的 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 检查床水平运动 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support External Remarks : 26.09.2021 11:26:17 UTC (86036024)服务内容：调整床的外壳机械结构。 服务结果：维修后，设备恢复正常工作。 Internal Remarks : 26.09.2021 11:26:20 UTC (86036024) "RDI Accuracy:Yes" $Q7: (Confirmed reported problem. The couch can not move in or out.) $Q8: (Mannually move couch in or out and found it is hard to move. The cover of couch was out of shape, due to be knocked by patient table.) $Q9: (Adjusted the mechanism of couch cover.) $Q10: (Test and inspection data refer to PA. Device returned to full functionality.) $Q11: Calibrated Tools (NA). ----------------------------------------------------------------------------- | 3 |  | $Q8: (Mannually move couch in or out and found it is hard to move. The cover of couch was out of shape, due to be knocked by patient table.) | [12 empty] | 118428950 |  | 1358336 |  |  | Investigation |  | TBD-CouchCover | [5 empty] | Engineering | [6 empty] | Couch cover | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | Noise | Couch-HW-Noise | Couch-HW-Noise | TBD-CouchCover | Noise | Couch-HW | Couch-HW-Noise | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1356 |  |  | Japan | Social Medical Corporation BOKOI, Nikko Memorial Hospital | 341011 | 44465 |  | 04DS: KK: The scan does not run after the rolling blackouts are restored. >> S_DMC_SHOT_COMM_COLLAPSE | Work Order #: WO-04417624 Kojima, Kenichi The system does not meet the specification for the performed service and is not returned to use Problem Description by Engineer: Unable to scan following restoration of power after a planned power outage.>> S_DMC_SHOT_COMM_COLLAPSE Diagnostic performed by Engineer: A S_DMC_SHOT_COMM_COLLAPSE error occurred. Physical channel order is 0 x 0. Internal Comments: [Subject] 04DS: KK: Mr. Ando: Unable to scan following restoration of power after a planned power outage.>> S_DMC_SHOT_COMM_COLLAPSE Shipped part: 453567538334: MODULE, ROTOR ELECTRONICS TX QTY: 1 ZREP/Traceable --None-- 453567538322: Module, Optical RX Stator Moog Velocity QTY: 1 ZREP/Traceable --None-- 453567057167:FRC ASSEMBLY QTY:1 ZREP/Traceable --None-- Dispatched: 1 RC FSE Registration Information: Name of FSE 1 who visited the site: Hodaka Hashimoto FSE 1 visit date and time: 2021/09/2616:00 FSE 1 hours worked: 4 ----------------------------------------------------------------------------- | 2 |  | [Details of work] After the rolling blackout, X-rays could not be emitted, so support was provided by telephone. Since it was completed by telephone support, we are returning during the visit. | [12 empty] | 118429061 |  | 1357741 |  |  | Closed - No Further Investigation | [7 empty] | TBD | [6 empty] | Unable to Scan | [5 empty] | Refresh |  |  | System issue | bar: No; No
1357 |  |  | China | 苏州广慈肿瘤医院有限公司 | 345019 | 44464 |  | 02N/Ingenuity Core128/Elite China/equipment error | BU Q&R Feedback: No SAP Customer Symptom Code: HANG Customer Symptom Code: Hang/Crash Reason: Corrective Maintenance Engineer Failure Code: Wear-out/Breakage/Mechanical Failure Repair Activity Code: Instructed Customer Problem Reported by Customer : 02N/Ingenuity Core128/Elite China T2 Activities : Problem Description: 客户报修思创工作站电脑故障。 Troubleshooting Action: 客户报修的工作站不属于CT机电脑或星云工作站维修范围内，已联系客户去找思创工作站部门报修。 Repair Action: 客户报修的工作站不属于CT机电脑或星云工作站维修范围内，已联系客户去找思创工作站部门报修。 Test/Inspection: customer confirm can scan patient and image ok.functional Check PASSED. Device returned to full functionality. Test Equipment: N/A Safety Questions : Case: 0118427823 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Bo Last updated by: Wang Bo 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 3 |  | 客户报修的工作站不属于CT机电脑或星云工作站维修范围内，已联系客户去找思创工作站部门报修。 | [12 empty] | 118427823 |  | 1357795 |  |  | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | ISP | [5 empty] | Refresh |  |  | System issue | bar: No; No
1358 |  |  | China | 江西省儿童医院 | 40041 | 44463 |  | Wechat:02N/Ingenuity CT 728326/Remove PRS | BU Q&R Feedback: No SAP Customer Symptom Code: CENV Customer Symptom Code: Customer Environmental Concern Reason: Corrective Maintenance Engineer Failure Code: Environmental Cause Repair Activity Code: Configuration Change Problem Reported by Customer : Remove PRS Safety Questions : Case: 0118421861 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Bai Jingwen Last updated by: Bai Jingwen 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Troubleshooting Action: answer customer question Repair Action: answer customer question Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: N/A ----------------------------------------------------------------------------- | [15 empty] | 118421861 |  | 1355688 |  |  | Closed - Non-Complaint | [20 empty] | Refresh |  |  | System issue | bar: No; No
1359 |  |  | China | 林甸县医院 | 345038 | 44463 |  | 02N/Ingenuity Core128/Elite China/Printing failed | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ Safety Questions : Case: 0118421668 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingsong Last updated by: Lu Mingsong 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: Printing failed Troubleshooting Action: Remotely instruct the customer to Select printer settings. Repair Action: Remotely instruct the customer to Select printer settings. Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: N/A ----------------------------------------------------------------------------- | 3 |  | Remotely instruct the customer to Select printer settings. | [12 empty] | 118421668 |  | 1355556 |  |  | Closed - No Further Investigation | [7 empty] | Application | [6 empty] | Print setting | [5 empty] | Refresh |  |  | System issue | bar: No; No
1360 |  |  | Japan | 社会医療法人 母恋 日鋼記念病院 | 341011 | 44463 |  | 02DS: KK: Last night, when I tried to shoot, E-Stop … | Work Order #: WO-04408495 Konishi, Hideo Regulatory Questions : N/A Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- Work Order #: WO-04408430 Kojima, Kenichi The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : 昨晩、撮影しようとしたらE-Stop Openが発生し、その後Closeできなかった。>>TEALからの再起動にて復帰。>>S_MC_DEVICE_VERT_POT_OPEN、S_VMC_AXIS_UNALLOWED_MOVE Diagnostic performed by Engineer : S_MC_DEVICE_VERT_POT_OPEN、S_VMC_AXIS_UNALLOWED_MOVEが発生しE-Stop Openしていました。 再Closeしようとしていましたが_MPS_MAIN_DRIVE_NOT_CLOSINGとなりCloseできなかったようです。 TEALからの再起動にて復旧。本日検査終了後対応。 Internal Comments : [Subject] 02DS:KK:ｻﾄｳ様:昨晩、撮影しようとしたらE-Stop Openが発生し、その後Closeできなかった。>>TEALからの再起動にて復帰。>>S_MC_DEVICE_VERT_POT_OPEN、S_VMC_AXIS_UNALLOWED_MOVE [出庫部品] 459800623331:Vertical Transducer QTY:1 ZSID --None-- [ディスパッチ] 1 RC -----------------FSE登録情報----------------- -訪問FSE1氏名 小西 英生 -FSE1訪問日時 2021/09/2416:00 -FSE1作業時間 4 [出荷指示] -----------------出荷指示#1----------------- -配送条件 CPU(倉庫引取り・輸送手配無し) -納期 2021/09/24 13:00 必着 -配送先名 札幌倉庫止め・FSEが引取り -受取者名 小西 英生 -電話番号 080-1605-0255 -その他 -依頼者名 小島健一 ----------------------------------------------------------------------------- | 4 |  | Part/Product - Vertical Transducer Part Number - 459800623331 | [12 empty] | 118421645 |  | 1351234 |  |  | Closed - No Further Investigation |  | NotEngineering-partreplace | [5 empty] | Engineering | [6 empty] | Vertical Transducer | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | VerticalTransducer | Couch-HW-VerticalTransducer | Couch-HW-VerticalTransducer | NotEngineering-partreplace |  | Couch-HW | Couch-HW-VerticalTransducer | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | #N/A
1361 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 337000 | 44463 |  | Gantry dönerken sürtme sesi geliyor | Integration: Confirmation of order 54206872 has been saved. ----------------------------------------------------------------------------- Work Order #: WO-04409960 24-Sep-2021 Can Ozgenlik The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Rubbing noise in the gantry. Resolution : We visited the site for the gantry noise and we found that one of the fan cables was a little higher than where it should be. The cable was fixed in its position. No other problems were found. Test scans with phantom were completed successfully. Device is up. ----------------------------------------------------------------------------- Work Order #: WO-04409565 Soner Saglam The system does not meet the specification for the performed service and is not returned to use Problem Description by Engineer : Rubbing noise in the gantry. Diagnostic performed by Engineer : In the video sent, rubbing noise was found. We asked the user to stop using the device in this condition. Site visit is required. Follow Up Required : In the video sent, rubbing noise was found. We asked the user to stop using the device in this condition. Site visit is required. ----------------------------------------------------------------------------- | 2 |  | The system was examined for the problem of sound coming from the gantry, and it was seen that one of the fan cables was a little higher than it should be. The cable is fixed in place. No other adverse events were observed. Test shots were made with the Phantom, no problem was seen, the device is in active condition. Check Ali to collect picture of issue | [12 empty] | 118422825 |  | 1352051 |  |  | Closed - No Further Investigation | [7 empty] | Operation | [6 empty] | Fan cable | [5 empty] | Refresh |  |  | System issue | bar: No; No
1362 |  |  | China | 四川省宜宾市第二中医院 | 345035 | 44462 |  | 02N/Ingenuity Core128/Elite China/information consultation | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer : information consultation Next Step Action : 邮件ASP chengduyunjia跟进 维护保养体系的相关规章制度要求 Safety Questions : Case: 0118416849 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Zhou Xiaohong Last updated by: Zhou Xiaohong 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | [15 empty] | 118416849 |  | 1352299 |  |  | Closed - Non-Complaint | [20 empty] | Refresh |  |  | System issue | bar: No; No
1363 |  |  | United States | Philips | 40021 | 44462 |  | Ingenuity 5000 has had continued issues with Tube heat | •Feedback Type: Technical Issue •Other Details: Other •Submitter Email: rhonda.swain@philips.com •CFM Form Received Date (Outlook email received date): 09/25/2021 Ingenuity 5000 installed October 2021 has had continued issues with Tube heat. The customer has placed multiple calls to the CCSC and me for help. We have adjusted the exam cards, and the issue continues. This week, I was onsite for a follow-up visit and learned that the tube heat issues had not been resolved. The hospital is a level two trauma center, and the exam cards can be complex. We see no reason in the exam cards that would be causing the tube heat issues. This does cause a delay in patient care. The staff sometimes get messages that they have to wait three to four minutes for the tube to cool before they can scan. All pertinent details: The reported issue did not cause harm or injury to a patient,but does sometimes cause a delay in patient care. |  |  | We have adjusted the exam cards, and the issue continues | [14 empty] | 1355116 |  |  | Investigation |  | CTD00005656 | [5 empty] | Engineering | [6 empty] | Tube Overheat | [5 empty] | Refresh |  |  | System issue | System | Gantry-SW | Tube over heat | Gantry-SW-Tube over heat | Gantry-SW-Tube over heat | CTD00005656 | Tube Heat | Gantry-SW | Gantry-SW-Tube over heat | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1364 |  |  | United States | McLaren Macomb | 40021 | 44462 |  | 4.1.10 Tube heat again | Integration : Confirmation of order 54202832 saved ----------------------------------------------------------------------------- Work Order #: WO-04404140 29-Sep-2021 Koren Moshe The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: Diagnostic Expected and Unexpected Results: Ongoing Tube Heat Warnings User Impact: No Patient Impact: No Current Software Version: 4.1.10 Diagnostic performed by Engineer : Issue will be resolved by chief TECH and APPS. ----------------------------------------------------------------------------- Work Order #: WO-04412476 27-Sep-2021 Mona Tony The system meets the specification for the performed service and is returned to use Resolution : Ran Filament Cals, copied Exam-card on external drive & Bug-report. Send information for analysis. David Gaffney the Hospital Radiology manager will provide a technologist contact information to work with Philips applications to tweak the Exam-Cards. Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: ----------------------------------------------------------------------------- Work Order #: WO-04403795 Maier, Donnie The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Follow Up Required : Several calls with tube heat issues in last 6 months. Create escalation, customer complaint, as necessary. Clin Apps is onsite and Remote Clin Apps has been involved. Protocols have been adjusted as far as lowering DRI, etc. already and cu is still getting tube heat warnings. May want to look into if protocols were converted for example as part of an upgrade. Diagnostic performed by Engineer : Clinical Applications is onsite and remote clin apps has also assisted on this site. Protocols have been adjusted (Bolus Tracker cycle times decreased, lower DRI, etc.) according to earlier cases. Dispatching for FSE to escalate and create Customer Complaint if necessary, etc. Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: Diagnostic Expected and Unexpected Results: Ongoing Tube Heat Warnings User Impact: No Patient Impact: No Current Software Version: 4.1.10 ----------------------------------------------------------------------------- Work Order #: WO-04403695 Hirchak, Michele The system meets the specification for the performed service and is returned to use Resolution : Called Rhonda Swain, apps onsite, to let her know the FSE will be there tomorrow at 2pm, according to the notes in the case. Closing clinical work order. MBH Problem description by Clinical Support Specialist : Customer Function/Role: Apps onsite. Device USE: IN USE Expected and Unexpected Results:Tube heat, happening very often, protocols are set very low. User Impact: NONE Patient Impact: NONE Current Software Version: 4.1.10 No rescan, no re-injection, no patient harm. Diagnostic performed by Engineer : Rhonda Swain, CES, is onsite. She observed multiple tube heat messages yesterday. Called today to review protocol changes from previous customer calls. PE Abd/Pelvis reviewed. Auto Scan time off, set RT to .75. DRI settings are low. Protocols should not be giving tube heat. Customer requesting FSE to come and troubleshoot. I will have an RSE look at logs. Case yesterday 9/22/21 @ 3:30pm EST. MBH ----------------------------------------------------------------------------- | 5 |  | We have adjusted the exam cards, and the issue continues | [12 empty] | 118418154 |  | 1349067 |  |  | Closed - No Further Investigation |  | Tube Heat | [5 empty] | Engineering | [6 empty] | Tube Overheat | [5 empty] | Refresh |  |  | System issue | System | Gantry-SW | Tube over heat | Gantry-SW-Tube over heat | Gantry-SW-Tube over heat | Tube Heat | Tube Heat | Gantry-SW | Gantry-SW-Tube over heat | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1365 |  |  | Chile | BANCO DE CHILE | 341005 | 44461 |  | CL1688-SOLICITA APOYO PARA REALIZAR AJUSTE DE HORA EN EL EQUIPO | Integration:Confirmation of order 54198333 saved ----------------------------------------------------------------------------- Work Order #: WO-04397274 [REDACTED] The system meets the specification for the performed service and is returned to use. Problem Description by Engineer : Problem description by engineer : Set time in the equipment. Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected results: No time problems in the system. Patient Impact/Outcome: No impact. User Impact: No impact. Current Software Version: 4.1.10 Resolution: The time is adjusted in the reconstructor. Test is performed with an exam and it is verified that now the images have the correct time. Internal Comments : Error # and/or description of error : Incorrect time in reconstructor. Repair Action: The time is adjusted in the reconstructor. Test is performed with an exam and it is verified that now the images have the correct time. Diagnostic performed by Engineer: There is a problem with the time of the reconstructor. The images contain at a different time than the one in the main computer. ----------------------------------------------------------------------------- | 3 |  | Resolution: The time is adjusted in the reconstructor. Test is performed with an exam and it is verified that now the images have the correct time. | [12 empty] | 118412909 |  | 1349710 |  |  | Evaluation |  | TBD-TimeSync | [5 empty] | Engineering | [6 empty] | Time sync | [5 empty] | Refresh |  |  | System issue | System | Console-SW | TimeSync | Console-SW-TimeSync | Console-SW-TimeSync | TBD-TimeSync |  | Console-SW | Console-SW-TimeSync | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | #N/A
1366 |  |  | Japan | Maki Hospital | 341017 | 44461 |  | 03CS:KK:検査しようとしたらエラーが出て停止。>>Anode Arc発生。>>再スキャンで検査可。 | Integration : Confirmation of order 54194725 saved Integration:According to factory calendar 20210923 (Posting date) is not a workday (check entry) ----------------------------------------------------------------------------- Work Order #: WO-04392955 Kojima, Kenichi The system meets the specification for the performed service and is returned to use. Resolution: Anode arc had occurred. It seemed to be a sporadic problem, so he agreed to keep an eye on it. I reported the issue to an FSE who is scheduled to visit the site for inspection. Problem Description by Engineer : An error occurred when he tried to run a scan, and it stopped.>> Anode arcing occurred.>> He was able to perform the examination with a rescan. Internal Comments: [Subject] 03CS:KK: Mr. Togawa: An error occurred when he tried to run a scan, and it stopped.>> Anode arcing occurred.>> He was able to perform the examination with a rescan. [4/8 Elements] Q5. How did the examination end? None. Q6. Did any mistaken findings (mistaken diagnosis or treatment) result from the problem? No. Q7. Was there any unintended unnecessary exposure? No Q8 Artifacts? None ----------------------------------------------------------------------------- | 4 |  | wait and see. | [12 empty] | 118369358 |  | 1342796 |  |  | Investigation | [7 empty] | TBD | [6 empty] | TBD | [5 empty] | Refresh |  |  | System issue | bar: No; No
1367 |  |  | China | 息烽县中医医院 | 345014 | 44460 |  | Wechat：02N/Core128/The key of couch fault | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : The key of couch fault Safety Questions : Case: 0118363153 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Pan Chuchu Last updated by: Pan Chuchu 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action : 客户：CT BOX降床键故障 已通知值班工程师 刘国平18066898085 先跟进客户判断是否需要派工，短信已发 Next Step Action : Site1 FSE 唐广林 18084189777 正在通话中 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : 唐工反馈他明天去正安的单子维修，这个场地去不了 联系SITE2 FSE 胡克 18585068636 告知可以派给他，已发派工短信通知到工程师 ----------------------------------------------------------------------------- | 3 |  | CT Box button | [12 empty] | 118363153 |  | 1342767 |  |  | Closed - No Further Investigation |  | CTUDT00632185 | [5 empty] | Engineering | [6 empty] | CT Box | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | CT Box | Gantry-HW-CT Box | Gantry-HW-CT Box | CTUDT00632185 | CT Box | Gantry-HW | Gantry-HW-CT Box | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1368 |  |  | Japan | Independent administrative institution National Hospital Organization | 341013 | 44460 |  | 04CS:KKI:撮影開始時に、エラー発生＞＞S_DMC_SHOT_READY_STATE_TIMEOUT | Work Order #: WO-04385479 Kimura, Kazumasa Problem Description by Engineer: Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer: N/A ----------------------------------------------------------------------------- | 4 |  | TBD | [12 empty] | 118363767 |  | 1343092 |  |  | Closed - No Further Investigation | [7 empty] | TBD | [6 empty] | TBD | [5 empty] | Refresh |  |  | System issue | bar: No; No
1369 |  |  | Puerto Rico | HOSPITAL DR SUSONI INC | 40013 | 44460 |  | Table noise when going up and down | Integration : Confirmation of order 54193468 saved ----------------------------------------------------------------------------- Work Order #: WO-04390999 28-Sep-2021 Ramos Rivera, Pablo The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Problem description by engineer : noise in vertical movement When was the first occurrence? Within... : 21-Sep-2021 13:00:00 How often is issue occurring? : Intermittent Malfunction area : Table Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: abnormal noise Patient Impact/Outcome: n/a User Impact: abnormal noise Current Software Version: unknown Resolution : Lubricated the vertical movement mechanism of the patient support. Now there is no noise when it moves vertically. ----------------------------------------------------------------------------- Work Order #: WO-04390455 Caldelas Aldana, Carlos Manuel The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Problem description by engineer : noise in vertical movement When was the first occurrence? Within... : 21-Sep-2021 13:00:00 How often is issue occurring? : Intermittent Malfunction area : Table Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: abnormal noise Patient Impact/Outcome: n/a User Impact: abnormal noise Current Software Version: unknown Diagnostic performed by Engineer : abnomral noise detected in vertical movement of patient support. no errors reported while in use. Follow Up Required : Error # and/or description of error : abnormal noise Repair Action : n/a What are the next Troubleshooting steps : inspect, clean and lubricate inf necesary vertical elements of patient suport. Technician Name : Caldelas Aldana, Carlos Manuel Resolution : FSE Visit Internal Comments : Error # and/or description of error : abnormal noise Repair Action : n/a ----------------------------------------------------------------------------- | 3 |  | inspect, clean and lubricate inf necesary vertical elements of patient suport. | [12 empty] | 118367512 |  | 1342002 |  |  | Closed - No Further Investigation |  | E-125059 | [5 empty] | Engineering | [6 empty] | Table Noise | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | Lubricant | Couch-HW-Lubricant | Couch-HW-Lubricant | E-125059 | Lubricant | Couch-HW | Couch-HW-Lubricant | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1370 |  |  | United States | Philips Medical Capital | 40053 | 44460 |  | 1301CDN - Customer requested a new destination added to the system | Integration : Confirmation of order 54193152 saved ----------------------------------------------------------------------------- Work Order #: WO-04390599 22-Sep-2021 Fairchild, Andrew The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: Diagnostics Expected and Unexpected Results: Add a new destination User Impact: No Patient Impact: No Current Software Version:4.1 Resolution : Customer was unable to communicate to a newly installed site. Changed AE-Title in PACS to allow for transmission of data. System restored to full functionality without restriction. ----------------------------------------------------------------------------- Work Order #: WO-04390532 Dangos John The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: Diagnostics Expected and Unexpected Results: Add a new destination User Impact: No Patient Impact: No Current Software Version:4.1 Diagnostic performed by Engineer : Customer requested a new destination added but does not have all the network information needed to add remotely. Customer will get information from IT dept and will have it ready for local FSE when he arrives Follow Up Required : Customer requested a new destination added but does not have all the network information needed to add remotely. Customer will get information from IT dept and will have it ready for local FSE when he arrives. The only Network info provided is the AE Title: MVODC-CT Customer will still need to provide IP address:? Port:?, Name:? ----------------------------------------------------------------------------- | 5 |  | Customer requested a new destination added | [12 empty] | 118367566 |  | 1341765 |  |  | NAC | [7 empty] | NAC | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | bar: No; No
1371 |  |  | United States | Westgate Hospital | 40001 | 44459 |  | 82488278 FAILED CALIBRATION DUE TO TUBE ARCING | Internal Remarks : **CUST REQUESTED ONSITE SUPPORT Interaction Type: New Case Caller: jim dodson - IPCR 7300 N 99TH AVE GLENDALE AZ IP & Issue: 82488278 FAILED CALIBRATION DUE TO TUBE ARCING Event Type: Incident Service Request: Onsite Did Customer Decline Remote Support? Yes Preferred Service Time: N/A Entitlement Options: Quote Reason for Entitlement Claim: N/A Action Taken: CREATED CASE AND QUOTE Next Step Action: Awaiting Quote Verbal PO Number: 51-982518 Verbal PO Approver: JIM DODSON Verbal PO CAP: Integration : VBAKKOM has been processed successfully-VBAPKOM has been processed successfully-VBAPKOM has been processed successfully-KONVKOM has been processed successfully-KONVKOM has been processed successfully-KONVKOM has been processed successfully-KONVKOM has been processed successfully-KONVKOM has been processed successfully-KONVKOM has been processed successfully-KONVKOM has been processed successfully-KONVKOM has been processed successfully-REQ2RES DMR 0018218618 has been saved-MPG is present on the product code : P9-MPG is present on the product code : P9 Integration : Confirmation of order 54187107 saved ----------------------------------------------------------------------------- Work Order #: WO-04381399 27-Sep-2021 Leon Peter The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : Performed long tube conditioning. Performed filament calibrations. Performed short tube conditioning. Ran test scans on system, all tests passed￼. ----------------------------------------------------------------------------- | 4 |  | Actions Taken: Dispatched per fse request Outcome: scheduled for 9/22 | [12 empty] | 118360881 |  | 1341234 |  |  | Closed - No Further Investigation |  | CTUDT00630085 | [5 empty] | Engineering | [6 empty] | Tube Arc | [5 empty] | Refresh |  |  | System issue | System | IC-HW | Tub Arc | IC-HW-Tub Arc | IC-HW-Tub Arc | CTUDT00630085 | Tub Arc | IC-HW | IC-HW-Tub Arc | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1372 |  |  | China | 四川省宜宾市第二中医院 | 345035 | 44458 |  | 02N/Ingenuity Core128/Elite China/ | BU Q&R Feedback: No Reason: Corrective Maintenance Next Step Action : 故障：添加打印机 RSE:LUMINGSONG Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ Safety Questions : Case: 0118356269 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Gao Luqi Last updated by: Gao Luqi 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : Xu Zehong ----------------------------------------------------------------------------- | [15 empty] | 118356269 |  | 1336226 |  |  | Closed - Non-Complaint | [20 empty] | Refresh |  |  | System issue | bar: No; No
1373 |  |  | China | Shandong Province Chest Hospital | 40025 | 44457 |  | 02N/Ingenuity CT 728326/Software login failure | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 02N/Ingenuity CT 728326 Safety Questions:Case: 0118355210 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingsong Last updated by: Lu Mingsong 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer_s problems: can not log on to the scan interface. Detailed description of Customer_s problems or malfunctions: can not log on to the scan interface. It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: N/A Remote troubleshoot and test result: N/A PRS TR status: Only M2M is not connected, ask FSE check onsite. Repair Action: Suggest onsite service procedures: can not log in to the scan interface. It is a software bug; use the "smart card" to reactivate the user. We recommend spare parts 12NC and names: N/A Memo of Customer request: None SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Next Step Action: huang zhaokun External Remarks: 19.09.2021 01:21:26 UTC (86038405) Fault Description: The CT account can not log in to the software. It has a blue screen after logging in. Service Details: Used IST privileges to log in to the configuration and acquire the Bug report.Replaced the Hosp-Lan network cable. Swapped the switch interface position where the network cable is located. Set the network card_s speed and changed the camera Association time. Service result: device returns to normal working condition after service.Need follow-up to monitor usage. Internal Remarks : 23.09.2021 13:35:47 UTC (86038405) "RDI Accuracy: YES." $Q7: ( Confirmed reported problem: Software can not login) ) $Q8: ( Login in Software use smart card and configuration. & L3 support reply this is Windows 10 Bug and need to Fixed future.) $Q9: ( Login Software use smart card and configuration.) $Q10: ( Test and inspection data refer to: PA . Device returned to full functionality.) $Q11: Calibrated Tools name (N/A) ^Current Tube installed date: N/A ^Current Tube usage: N/A ----------------------------------------------------------------------------- | 3 |  | Blue Screen | [12 empty] | 118355210 |  | 1335501 |  |  | Closed - No Further Investigation |  | CTUDT00632553 | [5 empty] | Engineering | [6 empty] | Blue Screen | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Blue-Screen | Console-SW-Blue-Screen | Console-SW-Blue-Screen | CTUDT00632553 | Blue-Screen | Console-SW | Console-SW-Blue-Screen | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1374 |  |  | China | Fujian Provincial Hospital | 345011 | 44457 |  | 02N/Ingenuity Core128/Elite China//system performance low | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Calibration/Adjustment Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ SAP Transfer:Dispatched for On site Support Safety Questions:Case: 0118354932 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Gaoge Last updated by: Wang Gaoge 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer_s problems: system is running slow or crashed. Detailed description of Customer_s problems or malfunctions: The system is slow and is stuck and needs debugging It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: System is slow. Remote troubleshoot and test result: Consider mcafee causing slow operation PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Replace virus database + close Mcafee update Test/Inspection: N/A Test Equipment: N/A SAP Transfer:Case Transferred to SAP Next Step Action: LI FENG Internal Remarks : 19.09.2021 02:15:48 UTC (86036224) RDI Accuracy: Yes $Q7: (Confirmed reported problem. The CT App run is slowly.) $Q8: (On site check and find the McAfee version is ok.) $Q9: (Delete the rubbish file and do disk clean up.) $Q10: (Test and inspection data refer to: PA.+Device returned to full functionality.) $Q11: Calibrated Tools name (FLUKE 287C Multimeter) S/N (16490128) NextCalibration date (10/20/2021) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date: 09/13/2020 ^Current Tube usage: 17832 External Remarks: 19.09.2021 02:15:38 UTC (86036224) Service details: CT software runs slowly and is stuck.Check on the site to make sure that the virus database of the Mcafee anti-virus software is normal.Clean junk files and clean disk. service result: device returns to normal working condition after service ----------------------------------------------------------------------------- | 3 |  | System is slow. Delete the rubbish file and do disk clean up. | [12 empty] | 118354932 |  | 1335350 |  |  | Closed - No Further Investigation |  | CTUDT00632261 | [5 empty] | Engineering | [6 empty] | System slow-DISK FULL | [5 empty] | Refresh |  |  | System issue | System | Console-SW | DiskFull | Console-SW-DiskFull | Console-SW-DiskFull | CTUDT00632261 | DiskFull | Console-SW | Console-SW-DiskFull | No | #REF! | #REF! | #REF! | #REF! | #REF! | No | #REF! | #REF! | #REF! | #REF! | #REF! | y | y | y | #REF! | System slow | #REF!
1375 |  |  | China | Yangfangdang Hospital | 40035 | 44456 |  | Wechat:02N/Ingenuity CT 728326/print information display setting | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Incompatible Use Repair Activity Code: Instructed Customer Problem Reported by Customer: wechat:02N/Ingenuity CT 728326 Safety Questions:Case: 0118349098 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Bai Jingwen Last updated by: Bai Jingwen 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: print information display setting Troubleshooting Action: the time and date not display on head images print. instruct customer change the information setting. system is ok. Repair Action: Functional Check PASSED. The device returned to full functionality. Test/Inspection: Functional Check PASSED. The device returned to full functionality. Test Equipment: N/A SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | the time and date not display on head images print. instruct customer change the information setting. system is ok. | [12 empty] | 118349098 |  | 1333533 |  |  | Closed - No Further Investigation | [7 empty] | Application | [6 empty] | Print setting | [5 empty] | Refresh |  |  | System issue | bar: No; No
1376 |  |  | China | Medical center (contagious disease, He nan, Xu chang) | 40047 | 44456 |  | System prompted e-stop problem, replaced front control panel | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 02n/Ingenuity CT 728326 Safety Questions:Case: 0118348031 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer_s problems: Other types of error. Detailed description of Customer_s problems or malfunctions: The customer reported for the system prompting the emergency switch problem It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: The customer reported for the system prompting the emergency switch problem Remote troubleshoot and test result: The customer rotated emergency switch and it restore. The panel has an emergency switch problem PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Replace the gantry left side front and back panels 0 "459800484501 CONTROL PANEL ASSEMBLY-FRONT LH 459800484531 OPERATOR PANEL ASSEMBLY-REAR RH" SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Next Step Action: PPD 459800484501*1 459800484531*1, 24H, FSE: Huang yong -86-18039573296 arranged shipment External Remarks: 18.09.2021 08:16:23 UTC (86038368) service details: [emergency switch on the panel is faulty. Replace the front panel.] Service result: [The device returns to normal working condition after maintenance.] Other important issues: [None] Internal Remarks : 18.09.2021 08:16:26 UTC (86038368) RDI Accuracy: Yes $Q7: (Confirm reported problem,estop button broken ) $Q8: (Performed visual inspection, found the estop button broken, need to replace the front control panel ) $Q9: (Replace the front control panel ) $Q10: (Test and inspection data refer to: PA. Device returned to full functionality.) $Q11:Calibrated Tools(N/A). ----------------------------------------------------------------------------- | 3 |  | estop button broken, Replace the front control panel 459800484501*1 | [12 empty] | 118348031 |  | 1333173 |  |  | Investigation |  | NotEngineering | [5 empty] | Engineering | [6 empty] | E-Stop button | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Control Panel | Gantry-HW-Control Panel | Gantry-HW-Control Panel | NotEngineering |  | Gantry-HW | Gantry-HW-Control Panel | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | #N/A
1377 |  |  | China | Sichuan Province Yibin City No.2 T.C.M. Hospital | 345035 | 44456 |  | 02N/Ingenuity Core128/Elite China/artifact | BU Q&R Feedback: No SAP Customer Symptom Code: QUAL Customer Symptom Code: Quality Problem with Image or Signal Reason: Corrective Maintenance Engineer Failure Code: Performance or Calibration Drift/Out-of-Specification Repair Activity Code: Calibration/Adjustment Problem Reported by Customer: 02N/Ingenuity Core128/Elite China Safety Questions:Case: 0118347991 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Case Transferred to SAP T2 Activities:Problem Description: artifact Troubleshooting Action: customer call service, after air calibration, scan with artifact images. instruct customer perform air calibration again, system is ok. Repair Action: perform air calibration. system is ok. Test/Inspection: Functional Check PASSED. The device returned to full functionality. Test Equipment: N/A ----------------------------------------------------------------------------- | 3 |  | scan with artifact images. instruct customer perform air calibration again, system is ok. | [12 empty] | 118347991 |  | 1333135 |  |  | Investigation | [7 empty] | Application | [6 empty] | Air Cal. | [5 empty] | Refresh |  |  | System issue | bar: No; No
1378 |  |  | China | Medical center (contagious disease, He nan, Xu chang) | 40047 | 44455 |  | 02N/Ingenuity CT 728326/gantry cannot power up | BU Q&R Feedback: No SAP Customer Symptom Code: BOOT Customer Symptom Code: Boot up/Power-up Problem Reason: Corrective Maintenance Engineer Failure Code: Electrical/Electronic Failure Repair Activity Code: Instructed Customer Problem Reported by Customer: 02N/Ingenuity CT 728326 Safety Questions:Case: 0118343291 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Gaoge Last updated by: Wang Gaoge 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities: Problem Description: The gantry can not start up. Troubleshooting Action: Guided the customer to check the E-stops on the front and back of the gantry. The equipment is back to normal use. Repair Action: Guided the customer to check the emergency switch. The E-stop on the back of the gantry was pressed by mistake. After restoring the emergency switch, the equipment is back to normal. Test/Inspection: N/A Test Equipment: N/A SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | Guided the customer to check the emergency switch. The E-stop on the back of the gantry was pressed by mistake. After restoring the emergency switch, the equipment is back to normal. | [12 empty] | 118343291 |  | 1329507 |  |  | Closed - No Further Investigation | [7 empty] | Application | [6 empty] | False alert | [5 empty] | Refresh |  |  | System issue | bar: No; No
1379 |  |  | China | Ning_an City No. 2 People_s Hospital | 345039 | 44455 |  | Wechat:02N/Ingenuity Core128/Elite China/Required Operation Instructions | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: Instructed Customer Safety Questions:Case: 0118343310 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Zixuan Last updated by: Wang Zixuan 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Problem Reported by Customer : Required Operation Instructions T2 Activities : Problem Description: Required Operation Instructions Troubleshooting Action: Required Operation Instructions Repair Action: Give customer IFU and TRG file Test/Inspection: Functional Check PASSED. The device returned to full functionality. Test Equipment: NA SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | [15 empty] | 118343310 |  | 1329523 |  |  | Closed - Non-Complaint | [20 empty] | Refresh |  |  | System issue | bar: No; No
1380 |  |  | Turkey | ADIYAMAN OZEL SAGINSTURTAS SAN VE TIC A.S | 341006 | 44455 |  | Hava kalibrasyonu yapılamıyor, konstrat uyarısı veriyor | Integration: Confirmation of order 54172012 has been saved. ----------------------------------------------------------------------------- Work Order #: WO-04360531 23-Sep-2021 Turgay Baylan The system meets the specification for the performed service and is returned to use Resolution : We opened the gantry covers to clean any contrasting agents. The system was returned to use in functional condition after testing with a patient. Problem Description by Engineer : Error displayed during air calibration. ----------------------------------------------------------------------------- Work Order #: WO-04360428 Soner Saglam The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, no follow-up requested Diagnostic performed by Engineer : Contrasting agent leak. There are artifacts in images. The leak should be cleaned with a site visit. Problem Description by Engineer : Error displayed during air calibration. Follow Up Required : Contrasting agent leak. There are artifacts in images. The leak should be cleaned with a site visit. ----------------------------------------------------------------------------- | 3 |  | Contrasting agent leak. There are artifacts in images. The leak should be cleaned with a site visit. | [12 empty] | 118344743 |  | 1325909 |  |  | Evaluation | [7 empty] | Customer | [6 empty] | Contrast | [5 empty] | Refresh |  |  | System issue | bar: No; No
1381 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 337000 | 44455 |  | CT Box Arızası | Integration: Confirmation of order 54170625 has been saved. ----------------------------------------------------------------------------- Work Order #: WO-04358387 27-Sep-2021 Can Ozgenlik The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Some of the keys on the CTBox are out of use. Parts Consumed : 1.00000,459801705022,CT BOX ASSEMBLY_BEST Resolution : The problem was solved after replacing the faulty CTbox. Device is operational. ----------------------------------------------------------------------------- Work Order #: WO-04358334 Soner Saglam The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, no follow-up requested Resolution : CTBox replacement. Problem Description by Engineer : Some of the keys on the CTBox are out of use. Follow Up Required : CTBox replacement. ----------------------------------------------------------------------------- | 3 |  | Product/Part - CT BOX ASSEMBLY_BEST Part number - 459801705022 | [12 empty] | 118342763 |  | 1324559 |  |  | Closed - No Further Investigation |  | CTUDT00632185 | [5 empty] | Engineering | [6 empty] | CT Box | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | CT Box | Gantry-HW-CT Box | Gantry-HW-CT Box | CTUDT00632185 | CT Box | Gantry-HW | Gantry-HW-CT Box | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1382 |  |  | United States | Philips Medical Capital | 40055 | 44455 |  | Backup media damaged, replaced USB | Internal Remarks : FADB: (FSE Aware of Double Booking) Title: Internal Remarks Initials: Twila Actions Taken: Per message sent to me in Teams from FSE/Turley, Rich - Create Case and Dispatch WO for 9/17/21 (afternoon). I Dispatched for 1pm. Outcome: Created Case and Dispatched WO per FSE_s Request IP: 88192637 Next Steps: None Integration : Confirmation of order 54174928 saved Integration : According to factory calendar 20210919 (Posting date) is not a workday (check entry) ----------------------------------------------------------------------------- Work Order #: WO-04363444 17-Sep-2021 Turley Rich The system meets the specification for the performed service and is returned to use Resolution : Replaced bad missing backup media Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Parts Consumed : 1.00000,459801128162,64GB USB BACKUP AND RESTORE ----------------------------------------------------------------------------- | 5 |  | 1.00000,459801128162,64GB USB BACKUP AND RESTORE | [12 empty] | 118346771 |  | 1332491 |  |  | Closed - Done |  | NotEngineering | [5 empty] | Engineering | [6 empty] | USB | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | USBExtended | Gantry-HW-USBExtended | Gantry-HW-USBExtended | NotEngineering |  | Gantry-HW | Gantry-HW-USBExtended | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | #N/A
1383 |  |  | China | 河南许昌市中心医院(传染病) | 40047 | 44454 |  | 02N/Ingenuity CT 728326 | BU Q&R Feedback: No SAP Customer Symptom Code: BOOT Customer Symptom Code: Boot up/Power-up Problem Reason: Corrective Maintenance Engineer Failure Code: Software Defect or Virus Repair Activity Code: Instructed Customer Problem Reported by Customer : 02N/Ingenuity CT 728326 Safety Questions : Case: 0118336170 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Guoping Last updated by: Liu Guoping 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: cannot login Troubleshooting Action: need change password Repair Action: guide customer change password Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: N/A ----------------------------------------------------------------------------- | 3 |  | guide customer change password | [12 empty] | 118336170 |  | 1324704 |  |  | Closed - No Further Investigation | [7 empty] | Application | [6 empty] | Password | [5 empty] | Refresh |  |  | System issue | bar: No; No
1384 |  |  | China | Zhejiang Province Yueqing No. 1 People’s Hospital | 40054 | 44454 |  | 02n/Ingenuity CT 728326/ECG wave transfer to host failed | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: Chip monitor is not connected to CT Safety Questions:Case: 0118336105 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Lili Last updated by: Lu Lili 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: ECG wave transfer to host failed Troubleshooting Action: this is intermittent issue, ECG monitor display wave is normal, but transfer to host failed. maybe cables connection problem. nemo ECG monitor to gantry_s cable. Repair Action: nemo cable Test/Inspection: N/A Test Equipment: N/A T2 Activities:Problem Description: Summary of customer_s problems: Other types of error. Detailed description of Customer_s problems or malfunctions: ECG signal cannot be transferred to host It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: ECG signal cannot be transferred to host Remote troubleshoot and test result: ECG display is normal, host is frequently disconnected, It is possible that the connection of the middle wire causing signal transmission failure PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Replace ECG to rack connection cable to test 0 CBL, ECG Monitor trigger 4M Cable 455012301901 SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Next Step Action: PPD455012301901*1 have arranged for 18h delivery, please dispatch FSE: Huang Yi-86-15958725282 Next Step Action:huang yi External Remarks: 16.09.2021 13:03:41 UTC (86036249) Service details: ECG signal loss. Checked and tested, found cable from IBOX to HOST is faulty. Re plug in and adjust, then restore. service result: device returns to normal working condition after service Other significant matters: none Internal Remarks : 16.09.2021 13:03:44 UTC (86036249) RDI Accuracy:Yes $Q7: (Confirmed reported problem, ECG signal lost) $Q8: (check the connection found ibox disconnect with host) $Q9: (replug and insert the related cables) $Q10: (Test and inspection data refer to: PA. Device returned to full functionality.) $Q11: Calibrated tools name (N/A). S/N (N/A). Next calibration date (N/A); For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date:N/A ^Current Tube usage: second:N/A ----------------------------------------------------------------------------- | 3 |  | ECG signal loss. Checked and tested, found cable from IBOX to HOST is faulty. Re plug in and adjust, then restore. | [12 empty] | 118336105 |  | 1324651 |  |  | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Cable loose | [5 empty] | Refresh |  |  | System issue | bar: No; No
1385 |  |  | China | Dongguan Donghua Hospital Co., Ltd. | 40059 | 44454 |  | Disk space was insufficient, replaced hard disk, reinstalled software | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: Out of memory Safety Questions:Case: 0118335062 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Zixuan Last updated by: Wang Zixuan 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer_s problems: failure related to computer software. Detailed description of Customer_s problems or malfunctions: Space is insufficient It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: NA Remote troubleshoot and test result: When scanning, it prompts that local space is insufficient. Checked the local and there is still 20% space left. The customer informed that it has also occurred before. PRS TR status: No PRS. Repair Action: Suggest onsite service procedures: Check and clean the C disk and D disk space, whether there are junk files; delete patient data and do a full rebuild; check the software. SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Next Step Action:LIN ZHITAO External Remarks : 17.09.2021 08:43:35 UTC (86034034) Service Details: The scan is intermittently interrupted. Prompted that there was a hard disk buffer space issue. Replaced the hard disk and reinstalled the software. Service Result: The device returned to working normally. Other significant matters: [none] Internal Remarks : 17.09.2021 08:43:38 UTC (86034034) RDI Accuracy:YES $Q7: (Confirmed reported problem, Out of memory ) $Q8: (out of disk memory when write data, the hard disk was not well) $Q9: (replace a new SSD) $Q10: (Test and inspection data refer to: PA. Device returned to full functionality) Q11: Calibrated Tools name ( NA ) S/N ( NA ) Next Calibration date ( NA ) For Mainland China CV / XR, CT engineers only. Please fill in: Current Tube installed date:NA Current Tube usage: NA ----------------------------------------------------------------------------- | 3 |  | The scan is intermittently interrupted. Prompted that there was a hard disk buffer space issue. Replaced the hard disk and reinstalled the software. | [12 empty] | 118335062 |  | 1324092 |  |  | Closed - Done |  | CTUDT00632261 | [5 empty] | Engineering | [6 empty] | Disk full | [5 empty] | Refresh |  |  | System issue | System | Console-SW | DiskFull | Console-SW-DiskFull | Console-SW-DiskFull | CTUDT00632261 | DiskFull | Console-SW | Console-SW-DiskFull | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1386 |  |  | China | 莱阳市人民医院 | 345005 | 44454 |  | Blue screen and "reboot the device" error occurred during Brain scanning | Blue screen and "reboot the device" error occurred during Brain scanning |  |  | System is out of warranty, Link to "Blue screen" defect CTUDT00632553 | [14 empty] | 1324012 |  |  | Investigation |  | CTUDT00632553 | [5 empty] | Engineering | [6 empty] | Blue Screen | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Blue-Screen | Console-SW-Blue-Screen | Console-SW-Blue-Screen | CTUDT00632553 | Blue-Screen | Console-SW | Console-SW-Blue-Screen | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | #REF! | y | #N/A
1387 |  |  | United States | Unity Health Systems | 40037 | 44454 |  | Monthly QA | Integration : Confirmation of order 54169063 saved ----------------------------------------------------------------------------- Work Order #: WO-04355570 16-Sep-2021 Taylor, Durrance The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: tech Device Use: scheduled activity Expected and Unexpected Results: Monthly QA User Impact: none Patient Impact: none Current Software Version: 4.1.10 Resolution : Monthly QA completed ----------------------------------------------------------------------------- | [15 empty] | 118340430 |  | 1326668 |  |  | Closed - Non-Complaint | [20 empty] | Refresh |  |  | System issue | bar: No; No
1388 |  |  | China | Shaanxi Province Ankang T.C.M. Hospital | 345041 | 44453 |  | WeChat：02N/Ingenuity Core128/Elite China/reconstruct image | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: WeChat：02N/Ingenuity Core128/Elite China Safety Questions:Case: 0118330314 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: mo xueke Last updated by: mo xueke 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer_s problems: For a certain patient, there is a problem with the reconstructed image It is a technology problem with the equipment itself: Yes. Troubleshooting Action: remote diagnosis and testing result: Suspect software, raw data problem PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Suggest engineer to check raw data and reconstruction card on the field.。。。 SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Next Step Action: zhang yifei Internal Remarks : 22.09.2021 03:37:56 UTC (86039086) $Q7: (Confirmed reported problem. & The system IRS have error problem ) $Q8: ( 1.Check IRS system. We need to restart system ) $Q9: ( 1. change system time). 2. Test scan for CT was OK. $Q10: (Test and inspection data refer to: PA. & Device returned to full functionality.) $Q11: Calibrated Tools name (N.A.) External Remarks: 22.09.2021 03:37:53 UTC (86039086) service reason: rebuild failed Service details: remote instruct customer to reboot CIRS, restore Sevice result: Device work normal. ----------------------------------------------------------------------------- | 3 |  | $Q7: (Confirmed reported problem. & The system IRS have error problem ) $Q8: (1.Check IRS system& We need to restart system ) $Q9: (1. change system time) . 2. Test scan for CT was OK. ) | [12 empty] | 118330314 |  | 1320767 |  |  | Investigation |  | CTUDT00632116 | [5 empty] | Engineering | [6 empty] | CIRS error | [5 empty] | Refresh |  |  | System issue | System | CIRS-SW | MPR Recon | CIRS-SW-MPR Recon | CIRS-SW-MPR Recon | CTUDT00632116 | MPR Recon | CIRS-SW | CIRS-SW-MPR Recon | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1389 |  |  | Chile | Scotiabank Chile | 40008 | 44452 |  | CL1675-HUELE A QUEMADO | Integration:Confirmation of order 54153944 saved ----------------------------------------------------------------------------- Work Order #: WO-04351965 28-Sep-2021 [REDACTED] The system meets the specification for the performed service and is returned to use. Problem Description by Engineer : Problem description by engineer : Burning smell in examination room. Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected Results: No problems with burning smell. Patient Impact/Outcome: No impact. User Impact: No impact. Current Software Version: 4.1.10 Diagnostic performed by Engineer: Visual and functional check-ups are performed in the equipment. - Isolation transformer is opened to check its components. - Internal components (power supplies, transformers, fuses and electronic cards) are checked in the equipment. In both cases, no overheating signs or visually affected components are detected. Functional tests are performed with X-rays and image quality is assessed without problems. Equipment is left under observation, event has not occurred again and it has been working normally for over a week so there are doubts that the problem comes from the scanner. Resolution: Equipment is under monitoring, no damage is detected in components and the equipment works normally. ----------------------------------------------------------------------------- Work Order #: WO-04336506 [REDACTED] The system meets the specification for the performed service and is returned to use. Diagnostic performed by Engineer: Intermittent burning smell in the examination room. The room is checked with the cusstomer_s help, but we are unable to identify the cause of the problem. It is asked to check the electrical board and air conditioning, but the customer reports that no problems were found. Case is sent to Planning to schedule a visit. Internal Comments : Error # and/or description of error : Burning smell in examination room. Repair Action: - Check system due to burning smell in examination room. Follow Up Required: Error Number and/or Description: Burning smell in examination room. Repair Action: - Check system due to burning smell in examination room. Technician_s Name: [REDACTED] Problem Description by Engineer : Problem description by engineer : Burning smell in examination room. Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected Results: No problems with burning smell. Patient Impact/Outcome: No impact. User Impact: No impact. Current Software Version: 4.1.10 ----------------------------------------------------------------------------- | 2 |  | Burning smell in examination room, unable to identify the cause of the problem remotely, onsite visit is planned at 28 Sep (Formal PR962707 at April, After confirmed, It is not from CT, but from air condition) | [12 empty] | 118324751 |  | 1318359 |  |  | Investigation | [7 empty] | TBD | [6 empty] | Burning | [5 empty] | Refresh |  |  | System issue | bar: No; No
1390 |  |  | China | Xifeng County Traditional Chinese Medicine Hospital | 345014 | 44452 |  | 05R/Ingenuity Core128/Elite China/cannot scan | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 05R/Ingenuity Core128/Elite China SAP Transfer:Dispatched for On site Support T2 Activities:Problem Description: Summary of customer_s problems: Scanning failed. Detailed description of Customer_s problems or malfunctions: Gantry acquisition system cannot comply。 It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: "event logger display Get DoseMODrefref, did not retrieve valid data " Remote troubleshoot and test result: Normal after restart PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Analyze reference detector data and confirm if it is the problem of converting or reference detector We recommend spare parts 12NC and names: BB Reference Detector FRU 459801702431 459801531261 DMS Ref. Converter FRU SAP Transfer:Case Transferred to SAP External Remarks: 27.09.2021 16:17:38 UTC (86039448) failure: scan failure. Failure analysis: Check event logger, display no data resend.Optical path testing is normal. Occasional failure. Failure handling: observe during use. Internal Remarks : 27.09.2021 16:17:41 UTC (86039448) $Q1: () $Q2: () $Q3: () $Q4: () "RDI Accuracy:YES" $Q7: (Confirmed reported problem.the scan failed. $Q8: (Troubleshooting,On site, Check evnt loger for no data resend. The optical path test is normal, and the occasional failure occurs. $Q9: (Repair steps. Using observation.12nc: ) $Q10: (Test and inspection data refer to: PA. Device returned to full functionality._ $Q11: Calibrated Tools name:N/A For Mainland China CV / XR, CT engineers only. Please fill in: tangguanglin Current Tube installed date: N/A Current Tube usage: scan second_N/A Safety Questions:Case: 0118321509 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Guoping Last updated by: Liu Guoping 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action: PPD 459801702431 ordered 459801702431 Sub no. 459800258961 available, FSE confirm when to ship. FSE: Tang Guanglin-86-18084189777 ----------------------------------------------------------------------------- | 3 |  | PPD 459801702431 已经订货 459801702431 替代号459800258961 有货 等FSE 确定何时发货 FSE:唐广林-86-18084189777 | [12 empty] | 118321509 |  | 1315760 |  |  | Closed - No Further Investigation |  | NotEngineering | [5 empty] | Engineering | [6 empty] | Ref. Converte | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Reference Converter | Gantry-HW-Reference Converter | Gantry-HW-Reference Converter | NotEngineering |  | Gantry-HW | Gantry-HW-Reference Converter | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | #N/A
1391 |  |  | Japan | Independent administrative institution National Hospital Organization | 341013 | 44452 |  | 01DS:TM:撮影途中でCT-Hostがフリーズ>>9秒ほどしてから回復するが撮影タイミングが遅れてしまった>>エラーコード無し | Integration : Confirmation of order 54148695 saved ----------------------------------------------------------------------------- Work Order #: WO-04329894 13-Sep-2021 Yamamoto, Yuko The system meets the specification for the performed service and is returned to use. Problem Description by Engineer: Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Regulatory Questions : N/A Resolution: There may have been a delay due to a software malfunction. A new patch (SP1) has been released, so I installed it.(V4.1.10.xxx210 to V4.1.10.xxx220) Internal Comments: [CX related] Did you agree on how to complete the report with the person in charge of devices? <Select one of the following> Yes. I completed normal reporting that day. [Work report (paper) submission] Did you submit a work report on paper? <YES/ Q1. How did the examination finish? Select from the following: -Only the scan in question was retaken ・The customer resumed the examination from the point at which it had stopped. Q2. Were there any mistaken findings (misdiagnosis, mistaken treatment) due to the problem? None. Q3. Were there any artifacts? None. Q4. Was there any unintended unnecessary exposure? None. ----------------------------------------------------------------------------- Work Order #: WO-04329886 Miyakawa, Takeyuki The system meets the specification for the performed service and is returned to use. Internal Comments: [Subject] 01DS:TM: Mr. Yamamoto: The CT-Host froze during a scan>>It recovered after about 9 seconds, but the timing of the scan was delayed>>No error codes. [Internal Comments] Results of checking the log: There were no hardware errors on logger.mdb. There was no record of any events on the CT-Host or the CIRS side. There were no problems with the amount of disk usage, etc.Both the CT and CIRS restarted this morning. Dispatched: 1 RC FSE Registration Information: Name of FSE 1 who visited the site: Yutaka Yamamoto FSE 1 visit date and time: 2021/09/1316:30 FSE 1 hours worked: 2 Diagnostic performed by Engineer: $version = "v4.1.10.25210" 9:24 Call description: After doing Prep, the CT-Host froze and stopped working. After about 9 seconds, the operation started but the timing of the scan was delayed. The similar problem occurred before, and we have reinstalled the CT-Host software. Problem Description by Engineer: The CT-Host froze during a scan>>It recovered after about 9 seconds, but the timing of the scan was delayed>>No error codes. ----------------------------------------------------------------------------- | 4 |  | The CT-Host froze during a scan>>It recovered after about 9 seconds, The similar problem occurred before, A new patch (SP1) has been released, so I installed it.(V4.1.10.xxx210 to V4.1.10.xxx220) | [12 empty] | 118318777 |  | 1312698 |  |  | Investigation |  | CTUDT00632520 | [5 empty] | Engineering | [6 empty] | Yellow Screen | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1392 |  |  | United States | Lackey Memorial Hospital | 40015 | 44452 |  | Q/O could not get ECG to work on CT | Integration : Confirmation of order 54153532 saved Integration : Confirmation of order 54153532 saved ----------------------------------------------------------------------------- Work Order #: WO-04336542 14-Sep-2021 Denison Chuck The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: C.T. tech. Device Use: Philips C.T. Expected and Unexpected Results: ECG. not working on host monitor. User Impact: C.T. still worked. Patient Impact: C.T. ECG. not showing. Current Software Version:4.1.10 Diagnostic performed by Engineer : Checked all connections for the ECG monitor. Not showing on the Host console. Found connection on i-box was loose. Tested the performance of the unit. Worked as should. ----------------------------------------------------------------------------- | 3 |  | Checked all connections for the ECG monitor. Not showing on the Host console. Found connection on i-box was loose. Tested the performance of the unit. Worked as should | [12 empty] | 118324775 |  | 1314504 |  |  | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Cable loose | [5 empty] | Refresh |  |  | System issue | bar: No; No
1393 |  |  | United States | Philips Medical Capital | 40053 | 44452 |  | 4.1 Reformats on aixials are being cut off. | Integration : Confirmation of order 54152659 saved ----------------------------------------------------------------------------- Work Order #: WO-04334832 Giles Brian The system meets the specification for the performed service and is returned to use Problem description by Clinical Support Specialist : Customer Function/Role: Tech Device Use: In Use Expected and Unexpected Results: Cut off reformats User Impact: None Patient Impact: None Current Software Version: 4.1 Diagnostic performed by Engineer : Brought David Mujica (RAS) onto the call call to help her out ----------------------------------------------------------------------------- Work Order #: WO-04334814 Mujica, David The system meets the specification for the performed service and is returned to use Problem description by Clinical Support Specialist : Customer Function/Role: Technologist Device Use: In use Expected and Unexpected Results: User Impact: none Patient Impact: none Current Software Version: 4.1 Diagnostic performed by Engineer : Alyssa was aided to merge images in CT viewer to create coronal and sagittal images on a CT head exam. Merge was explored in CT viewer. All images were reconstructed and saved successfully. No further clinical support needed. No rescan, no loss of data, no reinjection, per Alyssa. ----------------------------------------------------------------------------- | 5 |  | Alyssa was aided to merge images in CT viewer to create coronal and sagittal images on a CT head exam. Merge was explored in CT viewer. All images were reconstructed and saved successfully. | [12 empty] | 118323508 |  | 1313180 |  |  | Closed - No Further Investigation | [7 empty] | Application | [6 empty] | Merge image | [5 empty] | Refresh |  |  | System issue | bar: No; No
1394 |  |  | China | Zhalantun City People_s Hospital | 40045 | 44451 |  | wechat:02N/Ingenuity CT 728326/Scan, image transmission and reconstruction failure | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: Scan, image transmission and reconstruction failure Safety Questions:Case: 0118318000 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer_s problems: Other types of error. Detailed description of Customer_s problems or malfunctions: Continuous scan can cause overheating. Image transmission can crash. Reconstructed images are not recognized in the clinic It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: Continuous scan can cause overheating. Image transmission can crash. Reconstructed images are not recognized in the clinic Remote troubleshoot and test result: Scan protocol setup need to be adjusted PRS TR status: It is not through, ask engineer for maintenance on site. Repair Action: Suggest onsite service procedures: Adjust scan protocol. Communicated with trainer. Setup again and check image transmission issue SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support External Remarks: 15.09.2021 04:33:31 UTC (86038219) service details: scan software freezes, restore host software. service result: device returns to normal working condition after repairing Internal Remarks : 15.09.2021 04:33:34 UTC (86038219) RDI Accuracy（Yes）Q7: (Confirmation & Description of problem. Host software error. )$Q8: (On site check Host software failed, Need to install Host software.) $Q9: (Install Host software.) $Q10: (Test and inspection data refer to: PA. Device returned to full functionality.) $Q11: Calibrated tools name (FLUKE287) S/N (12470209) Next calibration date (06/30/2021) For Mainland China CV / XR, CT engineers only. Please fill in: Current Tube installed date Tube installed date: ##5/ 19/ 2021 ^Current Tube usage: scan second_##35126 ----------------------------------------------------------------------------- | 3 |  | Q7: (Confirmation & Description of problem. Host software error. )$Q8: (On site check Host software failed, Need to install Host software.) $Q9: (Install Host software.) $Q10: (Test and inspection data refer to: PA. Device returned to full functionality.) | [12 empty] | 118318000 |  | 1312034 |  |  | Investigation |  | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Recon failure | [5 empty] | Refresh |  |  | System issue | System | CIRS-SW | NoEnoughInfo | CIRS-SW-NoEnoughInfo | CIRS-SW-NoEnoughInfo | NoEnoughInfo |  | CIRS-SW | CIRS-SW-NoEnoughInfo | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | #N/A
1395 |  |  | China | Zhalantun City People_s Hospital | 40045 | 44451 |  | 02N/Ingenuity CT 728326/account locked | BU Q&R Feedback: No SAP Customer Symptom Code: MARF Customer Symptom Code: Monitor Alarm Failure Reason: Corrective Maintenance Engineer Failure Code: Incompatible Use Repair Activity Code: Instructed Customer Problem Reported by Customer: 02N/Ingenuity CT 728326 Safety Questions:Case: 0118317948 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: CT account locked Troubleshooting Action: wait half hour , then reuse CT account , input password, system is ok. Repair Action: wait half hour , then reuse CT account , input password, system is ok. Test/Inspection: N/A Test Equipment: N/A SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | wait half hour , then reuse CT account , input password, system is ok. | [12 empty] | 118317948 |  | 1311866 |  |  | Closed - No Further Investigation | [7 empty] | Application | [6 empty] | Password | [5 empty] | Refresh |  |  | System issue | bar: No; No
1396 |  |  | China | Caofeidian District Hospital | 345016 | 44449 |  | Wechat:02N/Ingenuity Core128/Elite China/Estop open | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer:Estop open Safety Questions:Case: 0118308389 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Bai Jingwen Last updated by: Bai Jingwen 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Case Transferred to SAP T2 Activities:Problem Description: Summary of customer_s problems: ESTOP disconnected. Detailed description of Customer_s problems or malfunctions: The gantry loses power continuously. It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: No PRS. Remote troubleshoot and test result: Restart but didn_t work. PRS TR status: No PRS. Repair Action: Suggest onsite service procedures: Suggest troubleshooting based on log on the site. We recommend spare parts 12NC and names: N/A SAP Transfer:Dispatched for On site Support External Remarks: 15.09.2021 02:40:08 UTC (86039063) service details: replace R2D Service Result: The device returned to normal. Internal Remarks : 15.09.2021 02:40:11 UTC (86039063) "RDI Accuracy:Yes" $Q7: (Confirmation & Description of problem) Confirmed reported problem, sometimes gantry power down $Q8: (Troubleshooting & Results) Check the error log, need to replace R2D $Q9: (Repair steps) Replace the R2D. $Q10: (Test and Inspection data & Device status) Test and inspection data refer to: PA. Device returned to full functionality. $Q11: Calibrated Tools name to PA ----------------------------------------------------------------------------- | 3 |  | Confirmed reported problem, sometimes gantry power down $Q8: (Troubleshooting & Results) Check the error log, need to replace R2D $Q9: (Repair steps) Replace the R2D. | [12 empty] | 118308389 |  | 1309146 |  |  | Closed - No Further Investigation |  | CTUDT00630864 | [5 empty] | Engineering | [6 empty] | R2D | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | R2D | Gantry-HW-R2D | Gantry-HW-R2D | CTUDT00630864 | R2D | Gantry-HW | Gantry-HW-R2D | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1397 |  |  | China | 黑龙江省桦川县人民医院 | 345018 | 44449 |  | 02N/Ingenuity Core128/Elite China/Mouse and keyboard faults | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Replace Part Problem Reported by Customer : Mouse and keyboard faults Safety Questions : Case: 0118307510 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Gaoge Last updated by: Wang Gaoge 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 键盘鼠标泡水了，坏了 Troubleshooting Action: nemo键盘鼠标给客户，客户自行更换 Repair Action: nemo 459800918431 Cherry USB keyboard English Silver 459800918441 Microsoft Optical Mouse 200 Test/Inspection: N/A Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | nemo键盘鼠标给客户，客户自行更换 | [12 empty] | 118307510 |  | 1308883 |  |  | NAC | [7 empty] | NEMO | [6 empty] | NEMO | [5 empty] | Refresh |  |  | System issue | bar: No; No
1398 |  |  | United States | Unity Health Systems | 40037 | 44449 |  | Errors during study with large patient; Inspected system and confirmed that system_s operating normally without issues. | Integration : Confirmation of order 54146605 saved ----------------------------------------------------------------------------- Work Order #: WO-04324718 16-Sep-2021 Taylor, Durrance The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: Diagnostic Expected and Unexpected Results: Errors during study with large patient User Impact: No Patient Impact: No Current Software Version: 4.1.10 Resolution : This all happened 9/10/21 00:39- 00:49 hours involving the patient support. Technician reported the patient 520 lbs rubbing against the sides of the gantry during the scan. Technician stated, "The AP scout was completed and there was an error after and the lateral scout was not completed. They moved the table in and zeroed again and the same thing happened with the second attempt". Customer also stated that after the patient was taken to another hospital that, "They completed the QA and IQ check after they removed the patient from the room. Everything went fine with that and has been working consistently since". Error logs showed that the table was not reaching the requested position in time. While I was doing the patient support PM I inspected the tables horizontal resistance and no issue was found there. I did perform a horizontal calibration as well. They also reported they could not run a bug report. I had them shut down the system and try again and they still could not run one. During the PM the system was blue button_d and all computers powered down. After power down the scanner was able to perform a bug report. This event occurred this one time and has not presented itself since 9/10. ----------------------------------------------------------------------------- Work Order #: WO-04324574 Maier, Donnie The system meets the specification for the performed service and is returned to use Follow Up Required : Customer explained they had some errors on 9/10 just after midnight during a study with a heavy patient. System is functioning properly now. She has already talked to PFSE Taylor Durrance and asked to have a case created and dispatched for 9/15 @1:00PM local time as he was already going to be onsite. CHA only had 3 cases with this particular soft limit error. States to calibrate couch horizontal and if continues to replace limit switches. Customer also stated that this was a heavy patient and only occurred this one time so may be a result of heavy patient. May calibrate and try to reproduce issue. Diagnostic performed by Engineer : PFSE Taylor Durrance is aware and asked for customer to call in and create a case. Also would like it scheduled for 9/15 1:00PM according to customer. System doesn_t appear to be in radar yet. Cu states that they received an error just after midnight this morning 9/10 on this system during a study with a larger patient. I will attempt to pull logs and add case notes if possible. Dispatching as requested. I pulled logs/analyzed and seen S_HMC_AXIS_WRONG_DEST_REACHED normal error @00:41:36 and then S_MC_DEVICE_IN_SOFT_LIMIT_ON Normal Error followed by S_MANMOT_HORIZONTAL_START_ERROR. Only 3 cases in CHA and they elude to performing Couch Horizontal Calibration first and then proceed to checking/replacing limit switches. However cu indicated this was a large patient so I would run thru couch calibration and try to reproduce error as it only occurred once according to cu. I cannot verify however as system doesn_t appear to be in Radar yet. Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: Diagnostic Expected and Unexpected Results: Errors during study with large patient User Impact: No Patient Impact: No Current Software Version: 4.1.10 ----------------------------------------------------------------------------- | 3 |  | After power down the scanner was able to perform a bug report. This event occurred this one time and has not presented itself since 9/10. | [12 empty] | 118315169 |  | 1308351 |  |  | Closed - Done | [7 empty] | Customer | [6 empty] | Over load | [5 empty] | Refresh |  |  | System issue | bar: No; No
1399 |  |  | Poland | University Clinical Hospital in | 341001 | 44448 |  | UPS do CT | Integration: Confirmation of order 54140370 saved. ----------------------------------------------------------------------------- Work Order #: WO-04312462 22-Sep-2021 Marcin Knyps The system meets the specification for the performed service and is returned to use Problem Description by Engineer: Parts were collected for diagnostics. Resolution: Collection and writing of spare parts. Correct operation of the device. Parts Consumed : 1.00000,459801652561,HP Z Turbo Drive 2TB TLC Z8 SSD Kit 1.00000,459801652551,HP Z Turbo Drive 256GB TLC Z8 SSD Module 2.00000,459801583761,8GB DDR4-2666 RegRAM 1.00000,459800306392,EXTERNAL BATT PACK, 3KVA UPS, 72VDC 1.00000,459800306382,UPS,NA,230V 50/60HZ,3KVA ----------------------------------------------------------------------------- | 5 |  | 1.00000,459801652561,HP Z Turbo Drive 2TB TLC Z8 SSD Kit 1.00000,459801652551,HP Z Turbo Drive 256GB TLC Z8 SSD Module 2.00000,459801583761,8GB DDR4-2666 RegRAM 1.00000,459800306392,EXTERNAL BATT PACK, 3KVA UPS, 72VDC 1.00000,459800306382,UPS,NA,230V 50/60HZ,3KVA | [12 empty] | 118305028 |  | 1307285 |  |  | Closed - No Further Investigation | [7 empty] | TBD | [6 empty] | TBD | [5 empty] | Refresh |  |  | System issue | bar: No; No
1400 |  |  | United States | Odessa Regional Medical Center | 40029 | 44448 |  | True Coronal/Sagittal on Scoliosis Patient | Work Order #: WO-04313269 Mujica, David The system meets the specification for the performed service and is returned to use Problem description by Clinical Support Specialist : Customer Function/Role: Technologist Device Use: In use Expected and Unexpected Results: Scoliosis patient reconstruction User Impact: none Patient Impact: none Current Software Version: 4.1 Diagnostic performed by Engineer : Maya indicated that she may need help straightening a CT scan in post processing of a scoliosis patient. It was confirmed that this patient was already diagnosed with a scoliosis condition, therefore the orthopod/radiologist would wish to see the condition of the scoliosis and therefore straightening the spine would defeat the conditional diagnosis. Maya will inquire from the radiologist if additional straightening views are needed. Maya said she could construct true coronal and sagittal by herself with the data given. Maya had no further questions or need for clinical support. No rescan, no loss of data, no reinjection per Maya. ----------------------------------------------------------------------------- | [15 empty] | 118305606 |  | 1303468 |  |  | Closed - Non-Complaint | [20 empty] | Refresh |  |  | System issue | bar: No; No
1401 |  |  | Japan | Maki Hospital | 341017 | 44448 |  | 01CS:RS: Inquiry about changing a patient name. >> I explained the "Change patient details" function. | Work Order #: WO-04309241 Satake, Rui The system meets the specification for the performed service and is returned to use. Problem Description by Engineer : Inquiry about changing a patient name. Resolution : I explained the "Change patient details" function. Internal Comments: [Subject] 01CS:RS:Mr. Togawa: Inquiry about changing a patient name. >> I explained the "Change patient details" function. [Internal Comments] None [4/8 Elements] Q5. How did the examination end? None. Q6. Did any mistaken findings (mistaken diagnosis or treatment) result from the problem? No. Q7. Was there any unintended unnecessary exposure? No Q8 Artifacts? None ----------------------------------------------------------------------------- | [15 empty] | 118300918 |  | 1300880 |  |  | Closed - Non-Complaint | [20 empty] | Refresh |  |  | System issue | bar: No; No
1402 |  |  | China | 山东安丘市中医医院 | 345017 | 44447 |  | 02N/Ingenuity Core128/Elite China/gantry noise | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity Core128/Elite China SAP Transfer : Case Transferred to SAP Safety Questions : Case: 0118296659 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Bo Last updated by: Wang Bo 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 设备异响。 客户问题详细描述或故障现象： 机架旋转时异响。 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： log无报错。 远程诊断与测试结果： 重启无效。已和工程师确认rotor皮带裂损。 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议检查异响来源，必要时更换皮带。 建议备件12NC及名称： PPD 453566492401 SAP Transfer : Dispatched for Onsite Support Next Step Action : PPD 453566492401*1 24H FSE:杨兴波-86-13583235930 已安排出货 External Remarks : 09.09.2021 08:02:48 UTC Xingbo Yang (86038250)服务内容：旋转噪声，更换Rotor驱动皮带。 服务结果：维修后，设备恢复正常工作。 Internal Remarks : 09.09.2021 08:02:50 UTC Xingbo Yang (86038250) "RDI Accuracy:yes" $Q7: (Confirmed reported problem & make noise when the rotor totation) $Q8: ( check rotor drive belt & found belt broken ) $Q9: ( replace poly_belt 453566492401 ) $Q10: (Test and inspection data refer to: PA & Device returned to full functionality) $Q11: Calibrated tools name(NA) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date: 9/ 30/ 2020 ^Current Tube usage: scan second_250394 ----------------------------------------------------------------------------- | 3 |  | $Q7: (Confirmed reported problem & make noise when the rotor totation) $Q8: ( check rotor drive belt & found belt broken ) $Q9: ( replace poly_belt 453566492401 ) | [12 empty] | 118296659 |  | 1301514 |  |  | Closed - No Further Investigation |  | ECR-077012 | [5 empty] | Engineering | [6 empty] | Belt | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Belt | Gantry-HW-Belt | Gantry-HW-Belt | ECR-077012 | Belt | Gantry-HW | Gantry-HW-Belt | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1403 |  |  | China | Shaanxi Province Ankang T.C.M. Hospital | 345041 | 44447 |  | WeChat：02N/Ingenuity Core128/Elite China | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: WeChat: 02N/Ingenuity Core128 Safety Questions:Case: 0118296354 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: mo xueke Last updated by: mo xueke 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Case Transferred to SAP T2 Activities:Problem Description: Summary of customer_s problems: system is running slow or crashed. Detailed description of Customer_s problems or malfunctions: The computer host system crashing. It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: The log has no errors. Remote troubleshoot and test result: Restart but didn_t work. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Suggest to add memory and observe usage according to TSE suggestion. We recommend spare parts 12NC and names: PPD 459801633881 Internal Remarks : 26.09.2021 02:25:58 UTC (86039086) $Q7: (Confirmed reported problem. & The system software have error ) $Q8: ( 1.Check system software. We need to replace RAM) $Q9: ( 1. replace RAM). 2. Test scan for CT was OK. $Q10: (Test and inspection data refer to: PA. The device returned to full functionality. $Q11: Calibrated Tools name (N.A.) SAP Transfer:Dispatched for On site Support External Remarks: 26.09.2021 02:25:56 UTC (86039086) service reason: the system is not working smoothly Service details: add 64g ram Sevice result: Device work normal. ----------------------------------------------------------------------------- | 5 |  | 服务原因：系统运行卡顿 服务内容：加装64g内存条 服务结果：设备工作正常 | [12 empty] | 118296354 |  | 1301293 |  |  | Closed - No Further Investigation |  | AddRAM | [5 empty] | Engineering | [6 empty] | System slow disk full | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Performance | Console-SW-Performance | Console-SW-Performance | AddRAM | Performance | Console-SW | Console-SW-Performance | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | System slow | #N/A
1404 |  |  | China | Medical center (contagious disease, He nan, Xu chang) | 40047 | 44447 |  | 02n/Ingenuity CT 728326/images transfer failed | BU Q&R Feedback: No SAP Customer Symptom Code: DIDM Customer Symptom Code: Patient or Image Data Management/Post Processing Data Reason: Corrective Maintenance Engineer Failure Code: Software Defect or Virus Repair Activity Code: Install Update Problem Reported by Customer: 02n/Ingenuity CT 728326 SAP Transfer:Case Transferred to SAP Next Step Action: 15637449981 repeat call in, transfer yang Safety Questions:Case: 0118296297 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: images transfer from CT to ISP failed Troubleshooting Action: customer call service, images transfer from CT to ISP failed. FSE has perform FCO72800740, this issue not solved. FSE call back to customer, will perform FCO72800758. close this case. Repair Action: FSE will perform FCO72800758,. Test/Inspection: N/A Test Equipment: N/A ----------------------------------------------------------------------------- | 3 |  | FSE will perform FCO72800758 to resolve image transfer issue | [12 empty] | 118296297 |  | 1301237 |  |  | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS transfer | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1405 |  |  | Poland | University Clinical Hospital in | 341001 | 44447 |  | Assistance during installation of Stellant CT DUAL injector. | Integration: Confirmation of order 54133709 saved. ----------------------------------------------------------------------------- Work Order #: WO-04303316 26-Sep-2021 Gerard Dabrowka The system meets the specification for the performed service and is returned to use Problem Description by Engineer: integration of the contrast injector. Diagnostic performed by Engineer: integration of the contrast injector Stellant Bayer. ----------------------------------------------------------------------------- | [15 empty] | 118297498 |  | 1302123 |  |  | Closed - Non-Complaint | [20 empty] | Refresh |  |  | System issue | bar: No; No
1406 |  |  | Germany | Gesundheit Nord gGmbH Klinikverbund | 40002 | 44446 |  | Probleme bei der Konstanzprüfung | Internal Remarks: 459801415161 Customer comment: # Arranged appointment with Ms. Schichtl for 9/16, 2 PM # Integration:Confirmation of order 54127560 saved ----------------------------------------------------------------------------- Work Order #: WO-04294491 16-Sep-2021 Halama, Thomas The system meets the specification for the performed service and is returned to use Problem Description by Engineer: System Phantom Head defective. Parts Consumed: 1.00000,459801415161,HARMONIZED PHANTOM HEAD SECTION Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution: Performed constancy test. In the tests performed, the system corresponds to the manufacturer_s specifications and is ready for clinical use. ----------------------------------------------------------------------------- | 3 |  | System Phantom Head defective. Parts Consumed: 1.00000,459801415161,HARMONIZED PHANTOM HEAD SECTION | [12 empty] | 118289659 |  | 1297009 |  |  | Evaluation |  | NotEngineering-PartReplace | [5 empty] | Engineering | [6 empty] | Head Phantom | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | Phantom | Couch-HW-Phantom | Couch-HW-Phantom | NotEngineering-PartReplace |  | Couch-HW | Couch-HW-Phantom | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | #N/A
1407 |  |  | China | Medical center (contagious disease, He nan, Xu chang) | 40047 | 44446 |  | 02N/Ingenuity CT 728326/change dicom setting | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 02N/Ingenuity CT 728326 Safety Questions:Case: 0118289121 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Dispatched for On site Support SAP Transfer:Case Transferred to SAP External Remarks: 07.09.2021 10:32:18 UTC Yong Huang (86038368) Service details: [customer request to add dicom node] Service result: [The device returns to normal working condition after maintenance.] Other important issues: [None] Internal Remarks: 07.09.2021 10:32:20 UTC Yong Huang (86038368) RDI Accuracy: Yes $Q7: (Confirm reported problem, change dicom set ) $Q8: (Performed visual inspection, need to change dicom set) $Q9: (change dicom set ) $Q10: (Test and inspection data refer to: PA. Device returned to full functionality.) $Q11: Calibrated Tools (N/A) T2 Activities:Problem Description: Summary of customer_s problems: Other types of error. Detailed description of Customer_s problems or malfunctions: Modify DICOM setting It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: Modify DICOM setting Remote troubleshoot and test result: FSE modify dicom setting onsite PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Modify DICOM setting ----------------------------------------------------------------------------- | [15 empty] | 118289121 |  | 1296651 |  |  | Closed - Non-Complaint | [20 empty] | Refresh |  |  | System issue | bar: No; No
1408 |  |  | United States | Clay County Medical Center | 40003 | 44446 |  | Front Left Control Panel Up Button is intermittent. | Next Step Action : left panel doesn_t work right , FSE approved to order. ----------------------------------------------------------------------------- Work Order #: WO-04297146 10-Sep-2021 Stauffer Jeffrey The system meets the specification for the performed service and is returned to use Resolution : Replaced front left control panel. Diagnostic performed by Engineer : Looked at front left control panel and saw that it was damaged. Pressed up button and it worked but had to be pressed hard to get it to work. Problem Description by Engineer : Customer Function/Role: Philips FSE Device Use: Diagnostic CT exams Expected and Unexpected Results: Expected - up button to be easy to press. Unexpected - up button on left front control panel is damaged and hard to press. User Impact: None Patient Impact: None Current Software Version: 4.1.10 Parts Consumed : 1.00000,459801484991,CONTROL PANEL ASSEMBLY-FRONT LEFT Tools & Equipment : MULTIMETER,2023-06-30 00:00:00 ----------------------------------------------------------------------------- | 3 |  | Replaced front left control panel. 1.00000,459801484991,CONTROL PANEL ASSEMBLY-FRONT LEFT | [12 empty] | 118292376 |  | 1299156 |  |  | Closed - No Further Investigation |  | tbd | [5 empty] | Engineering | [6 empty] | Control panel | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Control Panel | Gantry-HW-Control Panel | Gantry-HW-Control Panel | Not a Defect | Control Panel | Gantry-HW | Gantry-HW-Control Panel | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1409 |  |  | Chile | BANCO DE CHILE | 341005 | 44445 |  | CL1688 - TIME SET IS REQUIRED. | Integration:Confirmation of order 54123296 saved ----------------------------------------------------------------------------- Work Order #: WO-04288476 Valenzuela Velázquez, Eduardo Ariel The system meets the specification for the performed service and is returned to use. Problem Description by Engineer : Problem description by engineer : Adjust time. Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected Results: The time is set. Patient Impact/Outcome: No impact. User Impact: No impact. Current Software Version: Not provided. Diagnostic performed by Engineer: Due to a new time in the country, a time set is requested. Internal Comments : Error # and/or description of error : Adjust time. Repair Action: - The time is updated. Resolution: The equipment_s time is updated. ----------------------------------------------------------------------------- | 4 |  | Customer request setting | [12 empty] | 118285065 |  | 1290972 |  |  | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Time setting | [5 empty] | Refresh |  |  | System issue | bar: No; No
1410 |  |  | China | Henan Province Puyang City People’s Hospital | 40031 | 44445 |  | 02N/Ingenuity CT 728326/software hang up | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 02N/Ingenuity CT 728326 Safety Questions:Case: 0118284008 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Dispatched for On site Support SAP Transfer:Case Transferred to SAP External Remarks : 07.09.2021 06:48:56 UTC Yongping Luo (86038950) Service details: Software crash, slow response. Service result: Reset the software, after cleaned the disk, the equipment returned to normal. Continue to monitor. Other significant matters: none Internal Remarks : 07.09.2021 14:29:08 UTC Yongping Luo (86038950) $Q7:Confirmed report problem: the APP no response. Q8:Perform visual check on site, find the software running slowly, reset the software and clean the disk device return to full function. $Q9:reset the software and clean the disk, Device return to full function. $Q10: test and inspection data refer to PA. Device returned to full functionality. $Q11:Calibrated Tools refer to PA. Next Step Action: Luo Yongping 015226131133 call in and disp him Informed that the customer got drunk, and could not report for repair T2 Activities:Problem Description: Summary of customer_s problems: Other types of error. Detailed description of Customer_s problems or malfunctions: The software is crashing. It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: The software is crashing. Remote troubleshoot and test result: FSE reports for repair on behalf of the customer that software crash, yellow screen PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Re-install the system, and conduct FCO ----------------------------------------------------------------------------- | 3 |  | $Q7:Confirmed report problem: the APP no response. Q8:Perform visual check on site, find the software running slowly, reset the software and clean the disk device return to full function. FCO72800758 performed | [12 empty] | 118284008 |  | 1292924 |  |  | Investigation |  | CTUDT00632520 | [5 empty] | Engineering | [6 empty] | SP1 FCO | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1411 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 337000 | 44445 |  | spill guard order | Integration: Entry 1 in table OPERATION (method CHANGE) was not found. - Error occurred during the processing of the BAPI methods. Integration: Confirmation of order 54122643 has been saved. ----------------------------------------------------------------------------- Work Order #: WO-04288143 Parts Consumed : 1.00000,453566136722,APRON PACKING-SPILL GUARD ----------------------------------------------------------------------------- Work Order #: WO-04287967 Can Ozgenlik The system meets the specification for the performed service and is returned to use Resolution : The spill guard which is laid on the gantry cover to prevent opaque material leak was ordered. Device is operational. Problem Description by Engineer : Spill guard order. Follow Up Required : Spill guard order. ----------------------------------------------------------------------------- | 3 |  | 1.00000,453566136722,APRON PACKING-SPILL GUARD @William, need check what is the kit include, and the propose of the kit. | [12 empty] | 118284526 |  | 1290640 |  |  | Evaluation | [7 empty] | Others | [6 empty] | NEMO | [5 empty] | Refresh |  |  | System issue | bar: No; No
1412 |  |  | Portugal | Centro Hospitalar de Vila Nova | 40030 | 44445 |  | Repair of the couch top of the Emergency CT that is broken | [16 empty] | 118286513 |  | 1294456 |  |  | Closed - Duplicate | [20 empty] | Refresh |  |  | System issue | bar: No; No
1413 |  |  | China | Juancheng Ma Aiyun WCH | 345044 | 44445 |  | 02N/Ingenuity Core128/Elite China/config exam data | BU Q&R Feedback: No SAP Customer Symptom Code: MARF Customer Symptom Code: Monitor Alarm Failure Reason: Corrective Maintenance Engineer Failure Code: Incompatible Use Repair Activity Code: Instructed Customer Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ Safety Questions:Case: 0118284164 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Jianqiang Last updated by: Wang Jianqiang 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: consulting the exam data how to config Troubleshooting Action: information customer can config in preference Repair Action: instruct to customer Test/Inspection: Functional Check PASSED. The device returned to full functionality. Test Equipment: NA SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | [15 empty] | 118284164 |  | 1293042 |  |  | Closed - Non-Complaint | [20 empty] | Refresh |  |  | System issue | bar: No; No
1414 |  |  | United States | Philips Medical Capital | 40028 | 44445 |  | 0599CDN83 EStop Opens when couch motion fails | Next Step Action : Cust to call back if Holiday service wanted. If they do call back then either a PO# will be provided, of the name of the HCF Mgr approving billable holiday OT. Also agreed if they did not call back by 5pm, that case would be dispatched HTM Tuesday AM. Customer has not called back, but Mike Blair says to use Jerry Prendergast _ as PO# Integration : Confirmation of order 54123438 saved ----------------------------------------------------------------------------- Work Order #: WO-04290212 06-Sep-2021 Plewa Michael The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: CTTech Device Use: Diag Expected and Unexpected Results: Estop opens when couch motion fails User Impact: None Patient Impact: None Current Software Version:IDK Diagnostic performed by Engineer : Reset gantry. Checked error logs and estop circuit. Ran test scans . Tested OK. ----------------------------------------------------------------------------- Work Order #: WO-04288565 Carpenter, Kevin The system does not meet the specification for the performed service and is not returned to use Problem Description by Engineer : Customer Function/Role: CTTech Device Use: Diag Expected and Unexpected Results: Estop opens when couch motion fails User Impact: None Patient Impact: None Current Software Version:IDK Diagnostic performed by Engineer : Site PRS / RSN not working ~ possible misconfig - CT RSE unable to get logs. Cust to call back if they want billable onsite service ~ if they don_t call by 5PM, then they_ve agree RSE is to dispatch for HTM Tuesday 9/7 Diagnostic performed by Engineer : Cust never called back for onsite service ~ but Mike Blair says to use "Jerry Prendergast" as PO# and to dispatch. Follow Up Required : PRS / RSN to site is not working properly. Cust was to call back if Holiday service was wanted ~ CT RSE did not receive call from site. FSE please review logs onsite as customer reports EStop opens each time couch motion is attemoted. ----------------------------------------------------------------------------- | 2 |  | Reset gantry. Checked error logs and estop circuit. Ran test scans . Tested OK. | [12 empty] | 118284998 |  | 1290809 |  |  | Investigation |  | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Gantry reboot | [5 empty] | Refresh |  |  | System issue | System | TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | NoEnoughInfo |  | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | #N/A
1415 |  |  | United States | Odessa Regional Medical Center | 40029 | 44444 |  | Spiral artifact, unknown for customer cancelled onsite service | Integration : Confirmation of order 54119947 saved ----------------------------------------------------------------------------- Work Order #: WO-04281297 Carpenter, Kevin The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Diagnostic performed by Engineer : CT RSE advised customer to LOCK patients scan which have this artifact ( she says there are 3 cases ) ~ She says she has already wiped the bore to clear any specks of contrast and has then run Air Cals but the artifact persists. Site only has Remote Console on PRS / PRS - CT RSE unable to pull logs, images nor cal data. Problem Description by Engineer : Customer Function/Role: CTTech Device Use: Diag Expected and Unexpected Results: Spiral Artifact User Impact: None Patient Impact: None Current Software Version: 4 Follow Up Required : Cust reports Spiral Artifact on Helical scans ~ even after wipe bore and running air cals. Service is requested Tuesday AM. FSE run Bad Dets and inspect inner DMS cover to contrast, etc. ----------------------------------------------------------------------------- | 3 |  | cx stated no servie is needed issue is resolved. | [12 empty] | 118277997 |  | 1289541 |  |  | Closed - Done | [7 empty] | Others | [6 empty] | Artifact | [5 empty] | Refresh |  |  | System issue | bar: No; No
1416 |  |  | China | 江西省儿童医院 | 40041 | 44443 |  | 02N/Ingenuity CT/Gantry error | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: Instructed Customer Problem Reported by Customer : 02N/Ingenuity CT/ Next Step Action : CT Application 刘俊强 15314108698 来电 Safety Questions : Case: 0118276017 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wennian Han Last updated by: Wennian Han 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action : 应用老师在现场培训时，演示重启过程，后开机发现机器报错。建议应用老师关机长些时间，再重启，后反馈机器扫描正常。 Next Step Action : 也通知了FSE相关情况 T2 Activities : Problem Description: Gantry error Troubleshooting Action: Application specialist onsite training, demonstrated the reboot process to customer, and then power on and found that can not scan. Repair Action: Advise application specialist power down for a longer time, then reboot, and then feedback that the scan normally. Inform FSE related situation. Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: NA SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | Advise application specialist power down for a longer time, then reboot, and then feedback that the scan normally. Inform FSE related situation. | [12 empty] | 118276017 |  | 1289199 |  |  | Closed - No Further Investigation | [7 empty] | Application | [6 empty] | Training | [5 empty] | Refresh |  |  | System issue | bar: No; No
1417 |  |  | China | 河南省新乡市凤泉区大块镇卫生院 | 345009 | 44443 |  | 02N/Core128/need reset password | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Software Defect or Virus Repair Activity Code: Configuration Change Problem Reported by Customer : 02N/Core128/ Safety Questions : Case: 0118276372 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Guoping Last updated by: Liu Guoping 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: need reset password Troubleshooting Action: need reset password Repair Action: guide customer reset password Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | [15 empty] | 118276372 |  | 1289297 |  |  | Closed - Non-Complaint | [20 empty] | Refresh |  |  | System issue | bar: No; No
1418 |  |  | China | 鄄城马爱云妇儿医院 | 345044 | 44442 |  | 02N/Ingenuity Core128/Elite China/add printer | BU Q&R Feedback: No SAP Customer Symptom Code: MARF Customer Symptom Code: Monitor Alarm Failure Reason: Corrective Maintenance Engineer Failure Code: Incompatible Use Repair Activity Code: Instructed Customer Problem Reported by Customer : 02N/Ingenuity Core128/Elite China SAP Transfer : Case Transferred to SAP Safety Questions : Case: 0118270896 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Jianqiang Last updated by: Wang Jianqiang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: add printer Troubleshooting Action: information customer can add the printer under olevel modl Repair Action: instruct to customer add the printer by themselvs Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: NA ----------------------------------------------------------------------------- | [15 empty] | 118270896 |  | 1287404 |  |  | Closed - Non-Complaint | [20 empty] | Refresh |  |  | System issue | bar: No; No
1419 |  |  | United States | Leesburg Regional Medical Center | 40022 | 44442 |  | Screen froze; Rebooted host and CIRS to resolve. | Integration : Confirmation of order 54117367 saved Integration : According to factory calendar 20210905 (Posting date) is not a workday (check entry) ----------------------------------------------------------------------------- Work Order #: WO-04276916 Morrow, Dan The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Problem description by engineer : Screen froze and customer rebooted system. Raw data missing. When was the first occurrence? Within... : 03-Sep-2021 How often is issue occurring? : First occurrence Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Missing raw data Patient Impact/Outcome: none User Impact: missing raw data Current Software Version: 4.1.10 Diagnostic performed by Engineer : Consulted RAS Dyan Curry. Raw data was lost because system was rebooted before raw data was complete. Reviewed logs and no errors reported. Customer preformed a test scan and system is functional after Host/CIRS reboot. Internal Comments : Error # and/or description of error : No errors Repair Action : Consulted RAS Dyan Curry. Raw data was lost because system was rebooted before raw data was complete. Reviewed logs and no errors reported. Customer preformed a test scan and system is functional after Host/CIRS reboot. Resolution : Consulted RAS Dyan Curry. Raw data was lost because system was rebooted before raw data was complete. Reviewed logs and no errors reported. Customer preformed a test scan and system is functional after Host/CIRS reboot. OK to close. ----------------------------------------------------------------------------- | 3 |  | Raw data was lost because system was rebooted before raw data was complete. Reviewed logs and no errors reported. Customer preformed a test scan and system is functional after Host/CIRS reboot. | [12 empty] | 118274663 |  | 1286204 |  |  | Closed - Done | [7 empty] | Others | [6 empty] | Reboot - Host&CIRS | [5 empty] | Refresh |  |  | System issue | bar: No; No
1420 |  |  | Australia |  | 341000 | 44441 |  | Cant complete short tube condtioning | Integration : Confirmation of order 54112963 saved ----------------------------------------------------------------------------- Work Order #: WO-04300038 09-Sep-2021 Ruben MARQUEZ The system meets the specification for the performed service and is returned to use Diagnostic performed by Engineer : Checked the CIRS - unable to connect remotely but functional. Checked system logs and carried out visual diagnostics. Resolution : Assisted David. Replaced the rotor transmitter. Cold started the system and carried out tube warm up and test scans - system ok. Carried out constancy, IQ and stand by during procedures- system ok. Problem Description by Engineer : Customer Function/Role: Radiographer Device Use: not in use Expected and Unexpected Results: Gantry/CIRS error. User Impact: NA Patient Impact: NA Current Software Version:4.1.10 ----------------------------------------------------------------------------- Work Order #: WO-04271467 27-Sep-2021 David BROWN The system meets the specification for the performed service and is returned to use Problem Description by Engineer : System unable to carry out short Tube conditioning. Resolution : Carry out system inspection. Carry full diagnostics. Reseat all PCB’s in CIRS. Carry out functional testing. Return to service. ----------------------------------------------------------------------------- Work Order #: WO-04300042 09-Sep-2021 David BROWN The system does not meet the specification for the performed service and is not returned to use Problem Description by Engineer : System not working , failing on tube conditioning. Diagnostic performed by Engineer : Check over system RSE reported problem with CIRS AQU board. Send Bug reports to Engineer for checking. ----------------------------------------------------------------------------- Work Order #: WO-04300901 05-Sep-2021 Dan Quoc Le The system does not meet the specification for the performed service and is not returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : Could be due to the optical TX/RX in the gantry ----------------------------------------------------------------------------- Work Order #: WO-04281235 05-Sep-2021 David BROWN The system does not meet the specification for the performed service and is not returned to use Problem Description by Engineer : System will not complete a short tube conditioning. Diagnostic performed by Engineer : Check over system. Check all boards on CIRS and reseat. Reset system. Problem with CIRS computer and boards. Order Aquisitor PCB and fit ASAP ----------------------------------------------------------------------------- Work Order #: WO-04270951 Brian Donnelly The system does not meet the specification for the performed service and is not returned to use Follow Up Required : Error # and/or description of error : vBufferBytes = 547356672 AcqBufferByteSize=268435456 NumBuffers=2 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 Preparing for test execution... AcqGem_Diag::DiagRegisterTest Acq board#: 0 PASSED Diagnostic Data Path Info after DIAG_REGISTER_TEST : Real-time stream Physical channel order: 0x0 CommOK Status: 0 Loss of Signal: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 Line Fault Indicator: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 AcqGem_Diag::OdlMemoryDataBusTest Acq board#: 0, Gem board#: 0 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp945:AcqGem_Diag::PollOnAcqRecvBytes: AcqPutPointer(0x0), acqGetPtr(0x0), Bytes Received(0x0) != Bytes Expected(0x5580) ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp1457: AcqGem_Diag::OdlMemoryDataBusTest PollOnAcqRecvBytes FAILED on indx: 0 AcqGem_Diag::OdlMemoryDataBusTest FAILED. Diagnostic Data Path Info after ODL_MEM_DATA_BUS_TEST : Real-time stream Physical channel order: 0x0 CommOK Status: 0 Loss of Signal: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 Line Fault Indicator: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 AcqGem_Diag::SpiFlashIdVerifyTest PASSED Diagnostic Data Path Info after FPGA_FLASH_ID_CODE_VERIFY_TEST : Real-time stream Physical channel order: 0x0 CommOK Status: 0 Loss of Signal: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 Line Fault Indicator: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp945:AcqGem_Diag::PollOnAcqRecvBytes: AcqPutPointer(0x0), acqGetPtr(0x0), Bytes Received(0x0) != Bytes Expected(0xffd500) ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp1259: AcqGem_Diag::DmaGem2Acq PollOnAcqRecvBytes FAILED. AcqGem_Diag::DmaTest FAILED. Diagnostic Data Path Info after DMA_TEST : Real-time stream Physical channel order: 0x0 CommOK Status: 0 Loss of Signal: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 Line Fault Indicator: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp945:AcqGem_Diag::PollOnAcqRecvBytes: AcqPutPointer(0x0), acqGetPtr(0x0), Bytes Received(0x0) != Bytes Expected(0xffd500) ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp1259: AcqGem_Diag::DmaGem2Acq PollOnAcqRecvBytes FAILED. AcqGem_Diag::DmaTest FAILED. Diagnostic Data Path Info after DMA_TEST : Real-time stream Physical channel order: 0x0 CommOK Status: 0 Loss of Signal: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 Line Fault Indicator: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp945:AcqGem_Diag::PollOnAcqRecvBytes: AcqPutPointer(0x0), acqGetPtr(0x0), Bytes Received(0x0) != Bytes Expected(0xffd500) ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp1259: AcqGem_Diag::DmaGem2Acq PollOnAcqRecvBytes FAILED. AcqGem_Diag::DmaTest FAILED. Diagnostic Data Path Info after DMA_TEST : Real-time stream Physical channel order: 0x0 CommOK Status: 0 Loss of Signal: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 Line Fault Indicator: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp945:AcqGem_Diag::PollOnAcqRecvBytes: AcqPutPointer(0x0), acqGetPtr(0x0), Bytes Received(0x0) != Bytes Expected(0xffd500) ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp1259: AcqGem_Diag::DmaGem2Acq PollOnAcqRecvBytes FAILED. AcqGem_Diag::DmaTest FAILED. Diagnostic Data Path Info after DMA_TEST : Real-time stream Physical channel order: 0x0 CommOK Status: 0 Loss of Signal: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 Line Fault Indicator: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiagThread.cpp55:Acquire error: WaitForStartOfSeries Gantry data error found while wait for start of series. ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp2609: Acquire GetThreadStatus(1) unsuccessful. AcqGem_Diag::GantryDataChecksTest : subtest(1) : Normal Acquisition Test FAILED. GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiagThread.cpp55:Acquire error: WaitForStartOfSeries Gantry data error found while wait for start of series. ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp2609: Unexpected WorkItem WorkType Error bit(s) set: ErrorBitsFound: 0x1, ErrorBitsExpected: 0x10 AcqGem_Diag::GantryDataChecksTest : subtest(2) : Invalid Reading Number within a shot Test FAILED. GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiagThread.cpp55:Acquire error: WaitForStartOfSeries Gantry data error found while wait for start of series. ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp2609: Unexpected WorkItem WorkType Error bit(s) set: ErrorBitsFound: 0x1, ErrorBitsExpected: 0x400 AcqGem_Diag::GantryDataChecksTest : subtest(3) : Measurement Reading not in a shot Test FAILED. GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiagThread.cpp55:Acquire error: WaitForStartOfSeries Gantry data error found while wait for start of series. ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp2609: Unexpected WorkItem WorkType Error bit(s) set: ErrorBitsFound: 0x1, ErrorBitsExpected: 0x200 AcqGem_Diag::GantryDataChecksTest : subtest(4) : Idle Reading within a shot Test FAILED. GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiagThread.cpp55:Acquire error: WaitForStartOfSeries Gantry data error found while wait for start of series. ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp2609: Unexpected WorkItem WorkType Error bit(s) set: ErrorBitsFound: 0x1, ErrorBitsExpected: 0x40 AcqGem_Diag::GantryDataChecksTest : subtest(5) : Reading UID changed in a shot Test FAILED. GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiagThread.cpp55:Acquire error: WaitForStartOfSeries Gantry data error found while wait for start of series. ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp2609: Unexpected WorkItem WorkType Error bit(s) set: ErrorBitsFound: 0x1, ErrorBitsExpected: 0x20 AcqGem_Diag::GantryDataChecksTest : subtest(6) : Invalid DFS Setting In Reading Header Test FAILED. GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiagThread.cpp55:Acquire error: WaitForStartOfSeries Gantry data error found while wait for start of series. ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp2609: Unexpected WorkItem WorkType Error bit(s) set: ErrorBitsFound: 0x1, ErrorBitsExpected: 0x8 AcqGem_Diag::GantryDataChecksTest : subtest(7) : Non-zero Reading number at start of shot Test FAILED. GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiagThread.cpp55:Acquire error: WaitForStartOfSeries Gantry data error found while wait for start of series. ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp2609: Unexpected WorkItem WorkType Error bit(s) set: ErrorBitsFound: 0x1, ErrorBitsExpected: 0x2 AcqGem_Diag::GantryDataChecksTest : subtest(8) : Non-zero Shot number at start of Series Test FAILED. GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiagThread.cpp55:Acquire error: WaitForStartOfSeries Gantry data error found while wait for start of series. ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp2609: Unexpected WorkItem WorkType Error bit(s) set: ErrorBitsFound: 0x1, ErrorBitsExpected: 0x800 AcqGem_Diag::GantryDataChecksTest : subtest(9) : Reading slice size changed in series Test FAILED. GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiagThread.cpp55:Acquire error: WaitForStartOfSeries Gantry data error found while wait for start of series. Diagnostic Data Path Info after TIMEOUT_WAITING_FOR_NEXT_READING_TEST : Real-time stream Physical channel order: 0x0 CommOK Status: 0 Loss of Signal: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 Line Fault Indicator: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 AcqGem_Diag::GantryDataChecksTest : subtest(10) : Timeout waiting for next reading Test PASSED GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiagThread.cpp55:Acquire error: WaitForStartOfSeries Gantry data error found while wait for start of series. ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp2609: Unexpected WorkItem WorkType Error bit(s) set: ErrorBitsFound: 0x1, ErrorBitsExpected: 0x80 AcqGem_Diag::GantryDataChecksTest : subtest(11) : Maximum consecutive bad readings Test FAILED. GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiagThread.cpp55:Acquire error: WaitForStartOfSeries Gantry data error found while wait for start of series. ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp2609: Unexpected WorkItem WorkType Error bit(s) set: ErrorBitsFound: 0x1, ErrorBitsExpected: 0x1000 AcqGem_Diag::GantryDataChecksTest : subtest(12) : System memory overflow Test FAILED. GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiagThread.cpp55:Acquire error: WaitForStartOfSeries Gantry data error found while wait for start of series. ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp2609: Unexpected WorkItem WorkType Error bit(s) set: ErrorBitsFound: 0x1, ErrorBitsExpected: 0x2000 AcqGem_Diag::GantryDataChecksTest : subtest(13) : WorkQueue overflow Test FAILED. GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 Diagnostic Cycle Report ------------------------ Tests completed: 20 Tests remaining: 0 Total tests queued: 20 Cycles completed: 1 Cycles remaining: 0 Tests passed this cycle: 3 Tests failed this cycle: 17 Total tests passed: 3 Total tests failed: 17 Repair Action : Downloaded bugreport and could see h2R errors. Opened CIRS logs and could see that the physical channel for the ACQ was 0x0 Logged into the CIRS remotely and ran ACQ diagnostics. 17 of 20 tests failed. Technician Name : Brian Donnelly Internal Comments : Error # and/or description of error : vBufferBytes = 547356672 AcqBufferByteSize=268435456 NumBuffers=2 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 Preparing for test execution... AcqGem_Diag::DiagRegisterTest Acq board#: 0 PASSED Diagnostic Data Path Info after DIAG_REGISTER_TEST : Real-time stream Physical channel order: 0x0 CommOK Status: 0 Loss of Signal: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 Line Fault Indicator: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 AcqGem_Diag::OdlMemoryDataBusTest Acq board#: 0, Gem board#: 0 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp945:AcqGem_Diag::PollOnAcqRecvBytes: AcqPutPointer(0x0), acqGetPtr(0x0), Bytes Received(0x0) != Bytes Expected(0x5580) ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp1457: AcqGem_Diag::OdlMemoryDataBusTest PollOnAcqRecvBytes FAILED on indx: 0 AcqGem_Diag::OdlMemoryDataBusTest FAILED. Diagnostic Data Path Info after ODL_MEM_DATA_BUS_TEST : Real-time stream Physical channel order: 0x0 CommOK Status: 0 Loss of Signal: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 Line Fault Indicator: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 AcqGem_Diag::SpiFlashIdVerifyTest PASSED Diagnostic Data Path Info after FPGA_FLASH_ID_CODE_VERIFY_TEST : Real-time stream Physical channel order: 0x0 CommOK Status: 0 Loss of Signal: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 Line Fault Indicator: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp945:AcqGem_Diag::PollOnAcqRecvBytes: AcqPutPointer(0x0), acqGetPtr(0x0), Bytes Received(0x0) != Bytes Expected(0xffd500) ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp1259: AcqGem_Diag::DmaGem2Acq PollOnAcqRecvBytes FAILED. AcqGem_Diag::DmaTest FAILED. Diagnostic Data Path Info after DMA_TEST : Real-time stream Physical channel order: 0x0 CommOK Status: 0 Loss of Signal: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 Line Fault Indicator: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp945:AcqGem_Diag::PollOnAcqRecvBytes: AcqPutPointer(0x0), acqGetPtr(0x0), Bytes Received(0x0) != Bytes Expected(0xffd500) ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp1259: AcqGem_Diag::DmaGem2Acq PollOnAcqRecvBytes FAILED. AcqGem_Diag::DmaTest FAILED. Diagnostic Data Path Info after DMA_TEST : Real-time stream Physical channel order: 0x0 CommOK Status: 0 Loss of Signal: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 Line Fault Indicator: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp945:AcqGem_Diag::PollOnAcqRecvBytes: AcqPutPointer(0x0), acqGetPtr(0x0), Bytes Received(0x0) != Bytes Expected(0xffd500) ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp1259: AcqGem_Diag::DmaGem2Acq PollOnAcqRecvBytes FAILED. AcqGem_Diag::DmaTest FAILED. Diagnostic Data Path Info after DMA_TEST : Real-time stream Physical channel order: 0x0 CommOK Status: 0 Loss of Signal: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 Line Fault Indicator: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp945:AcqGem_Diag::PollOnAcqRecvBytes: AcqPutPointer(0x0), acqGetPtr(0x0), Bytes Received(0x0) != Bytes Expected(0xffd500) ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp1259: AcqGem_Diag::DmaGem2Acq PollOnAcqRecvBytes FAILED. AcqGem_Diag::DmaTest FAILED. Diagnostic Data Path Info after DMA_TEST : Real-time stream Physical channel order: 0x0 CommOK Status: 0 Loss of Signal: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 Line Fault Indicator: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiagThread.cpp55:Acquire error: WaitForStartOfSeries Gantry data error found while wait for start of series. ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp2609: Acquire GetThreadStatus(1) unsuccessful. AcqGem_Diag::GantryDataChecksTest : subtest(1) : Normal Acquisition Test FAILED. GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiagThread.cpp55:Acquire error: WaitForStartOfSeries Gantry data error found while wait for start of series. ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp2609: Unexpected WorkItem WorkType Error bit(s) set: ErrorBitsFound: 0x1, ErrorBitsExpected: 0x10 AcqGem_Diag::GantryDataChecksTest : subtest(2) : Invalid Reading Number within a shot Test FAILED. GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiagThread.cpp55:Acquire error: WaitForStartOfSeries Gantry data error found while wait for start of series. ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp2609: Unexpected WorkItem WorkType Error bit(s) set: ErrorBitsFound: 0x1, ErrorBitsExpected: 0x400 AcqGem_Diag::GantryDataChecksTest : subtest(3) : Measurement Reading not in a shot Test FAILED. GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiagThread.cpp55:Acquire error: WaitForStartOfSeries Gantry data error found while wait for start of series. ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp2609: Unexpected WorkItem WorkType Error bit(s) set: ErrorBitsFound: 0x1, ErrorBitsExpected: 0x200 AcqGem_Diag::GantryDataChecksTest : subtest(4) : Idle Reading within a shot Test FAILED. GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiagThread.cpp55:Acquire error: WaitForStartOfSeries Gantry data error found while wait for start of series. ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp2609: Unexpected WorkItem WorkType Error bit(s) set: ErrorBitsFound: 0x1, ErrorBitsExpected: 0x40 AcqGem_Diag::GantryDataChecksTest : subtest(5) : Reading UID changed in a shot Test FAILED. GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiagThread.cpp55:Acquire error: WaitForStartOfSeries Gantry data error found while wait for start of series. ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp2609: Unexpected WorkItem WorkType Error bit(s) set: ErrorBitsFound: 0x1, ErrorBitsExpected: 0x20 AcqGem_Diag::GantryDataChecksTest : subtest(6) : Invalid DFS Setting In Reading Header Test FAILED. GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiagThread.cpp55:Acquire error: WaitForStartOfSeries Gantry data error found while wait for start of series. ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp2609: Unexpected WorkItem WorkType Error bit(s) set: ErrorBitsFound: 0x1, ErrorBitsExpected: 0x8 AcqGem_Diag::GantryDataChecksTest : subtest(7) : Non-zero Reading number at start of shot Test FAILED. GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiagThread.cpp55:Acquire error: WaitForStartOfSeries Gantry data error found while wait for start of series. ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp2609: Unexpected WorkItem WorkType Error bit(s) set: ErrorBitsFound: 0x1, ErrorBitsExpected: 0x2 AcqGem_Diag::GantryDataChecksTest : subtest(8) : Non-zero Shot number at start of Series Test FAILED. GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiagThread.cpp55:Acquire error: WaitForStartOfSeries Gantry data error found while wait for start of series. ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp2609: Unexpected WorkItem WorkType Error bit(s) set: ErrorBitsFound: 0x1, ErrorBitsExpected: 0x800 AcqGem_Diag::GantryDataChecksTest : subtest(9) : Reading slice size changed in series Test FAILED. GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiagThread.cpp55:Acquire error: WaitForStartOfSeries Gantry data error found while wait for start of series. Diagnostic Data Path Info after TIMEOUT_WAITING_FOR_NEXT_READING_TEST : Real-time stream Physical channel order: 0x0 CommOK Status: 0 Loss of Signal: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 Line Fault Indicator: Chan A: 0 Chan B: 0 Chan C: 0 Chan D: 0 AcqGem_Diag::GantryDataChecksTest : subtest(10) : Timeout waiting for next reading Test PASSED GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiagThread.cpp55:Acquire error: WaitForStartOfSeries Gantry data error found while wait for start of series. ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp2609: Unexpected WorkItem WorkType Error bit(s) set: ErrorBitsFound: 0x1, ErrorBitsExpected: 0x80 AcqGem_Diag::GantryDataChecksTest : subtest(11) : Maximum consecutive bad readings Test FAILED. GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiagThread.cpp55:Acquire error: WaitForStartOfSeries Gantry data error found while wait for start of series. ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp2609: Unexpected WorkItem WorkType Error bit(s) set: ErrorBitsFound: 0x1, ErrorBitsExpected: 0x1000 AcqGem_Diag::GantryDataChecksTest : subtest(12) : System memory overflow Test FAILED. GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiagThread.cpp55:Acquire error: WaitForStartOfSeries Gantry data error found while wait for start of series. ..\..\..\sft\AcqGem\AcqGemDiag\src\AcqGemDiag.cpp2609: Unexpected WorkItem WorkType Error bit(s) set: ErrorBitsFound: 0x1, ErrorBitsExpected: 0x2000 AcqGem_Diag::GantryDataChecksTest : subtest(13) : WorkQueue overflow Test FAILED. GEM Stream is Real-time stream Number of GEM Data buffers supported by the Acquisitor: 1 Diagnostic Cycle Report ------------------------ Tests completed: 20 Tests remaining: 0 Total tests queu | 2 |  | Replaced the rotor transmitter. Cold started the system and carried out tube warm up and test scans - system ok. | [12 empty] | 118269593 |  | 1282954 |  |  | Closed - No Further Investigation |  | TBD_OBS | [5 empty] | Engineering | [6 empty] | TX | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Other | Gantry-HW-Other | Gantry-HW-Other | TBD_OBS | TBD-NoEnoughInfo | Gantry-HW | Gantry-HW-Other | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1421 |  |  | China | 鄄城马爱云妇儿医院 | 345044 | 44441 |  | 02N/Ingenuity Core128/Elite China//consult | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ SAP Transfer : Case Transferred to SAP Safety Questions : Case: 0118266107 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Cao Last updated by: Liu Cao 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: Customer consult issue Troubleshooting Action: Customer not in working time. may call back tomorrow Repair Action: Customer not in working time. may call back tomorrow Test/Inspection: N/A Test Equipment: N/A ----------------------------------------------------------------------------- | [15 empty] | 118266107 |  | 1284519 |  |  | Closed - Non-Complaint | [20 empty] | Refresh |  |  | System issue | bar: No; No
1422 |  |  | Australia | Maia Financial Pty Limited | 341000 | 44440 |  | Bad Detector | Work Order #: WO-04262865 David BROWN Diagnostic performed by Engineer : N/A Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: ----------------------------------------------------------------------------- | 3 |  | Product/Part - TDMS V2 M.Det Class A FRU Part number - 459800911271 | [12 empty] | 118262778 |  | 1282716 |  |  | Evaluation |  | CTUDT00632110 | [5 empty] | Engineering | [6 empty] | DMS Module | [5 empty] | Refresh |  |  | System issue | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | DMS Module | DMS-HW | DMS-HW-DMS Module | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1423 |  |  | China | Medical center (contagious disease, He nan, Xu chang) | 40047 | 44440 |  | 02N/Ingenuity CT 728326/Add node | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer: 02N/Ingenuity CT 728326 Safety Questions:Case: 0118259071 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Gao Luqi Last updated by: Gao Luqi 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: Add node and test network Troubleshooting Action: Customer add node but test failed Delete Retrieve and query service, then test pass Repair Action: Instruct customer. Test/Inspection: N/A Test Equipment: N/A SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | Add node and test network | [12 empty] | 118259071 |  | 1279895 |  |  | NAC | [7 empty] | NAC | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | bar: No; No
1424 |  |  | United States | Mount Grant General Hospital | 40012 | 44440 |  | 0698CDN - Gantry noise and also one image will not transfer to PACS | Gantry noise and also one image will not transfer to PACS |  |  | Need CHU to collect more probelm detail for investigation@Miley | [14 empty] | 1326147 |  |  | Closed - No Further Investigation |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS transfer | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1425 |  |  | United States | Mount Grant General Hospital | 40012 | 44440 |  | 0698CDN - Gantry noise and also one image will not transfer to PACS | Integration : Confirmation of order 54105234 saved ----------------------------------------------------------------------------- Work Order #: WO-04315202 13-Sep-2021 Schuler Bruce The system meets the specification for the performed service and is returned to use Resolution : Replaced belt and checked image transfer status. No other issues found Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: Diagnostics Expected and Unexpected Results: Gantry noise and image will not transfer to PACS- FSE is aware User Impact: Patient Impact: Current Software Version: Parts Consumed : 1.00000,453566492401,POLY-VEE BELT ----------------------------------------------------------------------------- Work Order #: WO-04260727 09-Sep-2021 Schuler Bruce The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Diagnostic performed by Engineer : Found belt to squeak every revolution. Will Replace belt Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: Diagnostics Expected and Unexpected Results: Gantry noise and image will not transfer to PACS- FSE is aware User Impact: Patient Impact: Current Software Version: Follow Up Required : Replace Belt ----------------------------------------------------------------------------- Work Order #: WO-04260647 Dangos John The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: Diagnostics Expected and Unexpected Results: Gantry noise and image will not transfer to PACS- FSE is aware User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : Gantry noise and image will not transfer to PACS- FSE is aware Follow Up Required : Gantry noise and image will not transfer to PACS- FSE is aware ----------------------------------------------------------------------------- | 3 |  | 1.00000,453566492401,POLY-VEE BELT | [12 empty] | 118261133 |  | 1277051 |  |  | Closed - No Further Investigation |  | ECR-077012 | [5 empty] | Engineering | [6 empty] | Belt | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Belt | Gantry-HW-Belt | Gantry-HW-Belt | ECR-077012 | Belt | Gantry-HW | Gantry-HW-Belt | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1426 |  |  | United States | CENTRAL VERMONT MEDICAL CENTER | 40024 | 44440 |  | Ring artifact on images; Replaced detector module,cleaned contrast out of ultra high Res filter, adjusted/calibrated DMS | Internal Remarks : Provided approval for a detector module. Also recommended checking the UHR comb to see if there is contrast in it along with on the Collimator. ----------------------------------------------------------------------------- Work Order #: WO-04263504 02-Sep-2021 Sangster, Robert The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Problem description by engineer : ring artifacts on temporal bone head scans Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: ring artifacts on temporal bone head scans Patient Impact/Outcome: delay, moved patient to another scanner and rescanned patient. User Impact: delay Current Software Version: Not provided Parts Consumed : 1.00000,459800911271,TDMS V2 M.Det Class A FRU Resolution : Replace detector module 23. Clean contrast out of ultra high Res filter. Adjust and calibrate the DMS as needed. ----------------------------------------------------------------------------- Work Order #: WO-04260353 01-Sep-2021 Sangster, Robert The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, no follow-up requested Problem Description by Engineer : Problem description by engineer : ring artifacts on temporal bone head scans Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: ring artifacts on temporal bone head scans Patient Impact/Outcome: delay, moved patient to another scanner and rescanned patient. User Impact: delay Current Software Version: Not provided Follow Up Required : Replace detector module Diagnostic performed by Engineer : Trouble shot issue to bad detector ordered parts ----------------------------------------------------------------------------- Work Order #: WO-04259603 Sheffield, Scott The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Problem description by engineer : ring artifacts on temporal bone head scans Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: ring artifacts on temporal bone head scans Patient Impact/Outcome: delay, moved patient to another scanner and rescanned patient. User Impact: delay Current Software Version: Not provided Follow Up Required : Error # and/or description of error : none Repair Action : I downloaded bad dets, air cal preview images and delta air. No bad dets were found during air cals Air cal images were black I did not see anything in the scan field logs attached to case for NSS if needed Recommended: Open gantry and check dms and collimator mylar covers for contrast or other fluids that may be causing artifacts. Also check the gantry mylar that it is clean Run bad detector test Consult with NSS for further assistance Technician Name : Sheffield, Scott Internal Comments : Error # and/or description of error : none Repair Action : I downloaded bad dets, air cal preview images and delta air. No bad dets were found during air cals Air cal images were black I did not see anything in the scan field logs attached to case for NSS if needed Recommended: Open gantry and check dms and collimator mylar covers for contrast or other fluids that may be causing artifacts. Also check the gantry mylar that it is clean Run bad detector test Consult with NSS for further assistance Diagnostic performed by Engineer : Chad said that he got the ring artifacts on the temporal head scan only on all images. He did a regular scan and did not have artifacts. He scanned the phantom with the temporal protocol and it had the same artifacts. I had him copy the phantom images to cdr phantom and I downloaded the images and saw the artifacts. I will attached them and logs to case. I had Chad run air cals, it gave him a message saying something was in scan field. He cleaned gantry mylar then air cals ran. After air cals he scanned the phantom again and still has the artifacts for temporal heads. He is not seeing the artifacts on any other type scans at this time. He answered yes to safety question # 2 he said because he had to re-scan the patient and the patient received extra radiation because of the re-scan. I_m dispatching FSE. ----------------------------------------------------------------------------- | 3 |  | Product/Part - TDMS V2 M.Det Class A FRU Part number - 459800911271 | [12 empty] | 118260423 |  | 1276409 |  |  | Closed - Done |  | CTUDT00632110 | [5 empty] | Engineering | [6 empty] | DMS Module | [5 empty] | Refresh |  |  | System issue | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | DMS Module | DMS-HW | DMS-HW-DMS Module | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1427 |  |  | China | Xinjiang Province Yili Prefecture Friendship Hospital | 40048 | 44439 |  | 02N/Ingenuity CT 728326/ESTOP opened | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer:ESTOP opened SAP Transfer:Dispatched for On site Support SAP Transfer:Case Transferred to SAP External Remarks: 01.09.2021 03:44:12 UTC Jianjun Sang (86038298) Service Details: The gantry couldn_t power on. Instructed the customer to check all the emergency switches, and they found that an emergency switch was pressed in. They reset the emergency switch, powered the gantry back on and the system is working properly. service result: device returns to normal working condition after repairing Internal Remarks:01.09.2021 03:44:15 UTC Jianjun Sang (86038298) $Q7: (Confirmed reported problem,the system e_stop open, can not powere on) $Q8: ( ask the customer to check e_stop switch find one switch was pressed ) $Q9: (ask the customer to reset the e_stop switch and restat the system , the system works normal ) $Q10: (Test and inspection data refer to: PA. Device returned to full functionality) $Q11: N/A Next Step Action:SANG Jianjun Safety Questions:Case: 0118254164 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Gao Luqi Last updated by: Gao Luqi 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- |  |  | ask the customer to check e_stop switch find one switch was pressed | [12 empty] | 118254164 |  | 1276688 |  |  | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Misoperation | [5 empty] | Refresh |  |  | System issue | bar: No; No
1428 |  |  | Spain | S.C.S. HOSP. GRAL. DE FUERTEVENTURA | 40020 | 44439 |  | Images are not sent to PACS, RIS or local. | Integration:Confirmation of order 54097057 saved ----------------------------------------------------------------------------- Work Order #: WO-04248201 Carlos Saumell Lasala The system meets the specification for the performed service and is returned to use. Diagnostic performed by Engineer: The transfer queue is deleted, but the studies continue to transfer very slowly to the PACS. Problem Description by Engineer: They cannot send studies to the PACS. Regulatory Questions : N/A Resolution: Finally, the IT technicians of the Hospital found that the cause of the problem was in the PACS itself, and they have already solved it. ----------------------------------------------------------------------------- |  |  | IT technicians of the Hospital found that the cause of the problem was in the PACS itself, and they have already solved it. | [12 empty] | 118251247 |  | 1275013 |  |  | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | PACS Vendor | [5 empty] | Refresh |  |  | System issue | bar: No; No
1429 |  |  | China | Inner Mongolia Hinggan League People_s Hospital | 345030 | 44439 |  | 02N/Ingenuity Core128/Elite China/software stuck | BU Q&R Feedback: No Reason: Corrective Maintenance Safety Questions:Case: 0118249760 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Dai Suyang Last updated by: Dai Suyang 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Case Transferred to SAP External Remarks: 01.09.2021 08:33:00 UTC Zhen Liang (86039057) Service Result: [The customer said the equipment can crash when in use. Implement the database replacement and standby according to the customer_s request.device returns to normal working condition.] Other significant matters: none Internal Remarks: 01.09.2021 08:33:03 UTC Zhen Liang (86039057) $Q7: Confirmed reported problem,the system was slowly and system crashed。 $Q8: on sit service，data base might be damaged. $Q9: do backup and restore .system is ok! $Q10:Test and inspection data refer to: PA. Device returned to not full functionality. $Q11: Calibrated Toolsnamen/a Current Tube usage: scan second 47825. SAP Transfer:Dispatched for On site Support Problem Reported by Customer: 02N/Ingenuity Core128/Elite China T2 Activities:Problem Description: Summary of customer_s problems: system is running slow or crashed. Detailed description of Customer_s problems or malfunctions: The software crashed. It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: The log has no errors. Remote troubleshoot and test result: Restart but didn_t work. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: It is recommended to first reinstall the software to resolve the fault. If it is still not resolved, it is recommended to collect the bug report and upload to the technical support engineer (TSE). We recommend spare parts 12NC and names: N/A ----------------------------------------------------------------------------- |  |  | $Q7: Confirmed reported problem,the system was slowly and system crashed。 $Q8: on sit service，data base might be damaged. $Q9: do backup and restore .system is ok! | [12 empty] | 118249760 |  | 1274263 |  |  | Closed - No Further Investigation |  | NoEnoughInfo | [5 empty] | engineering | [6 empty] | System slow | [5 empty] | Refresh |  |  | System issue | System | TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | NoEnoughInfo |  | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | #N/A
1430 |  |  | China | The Second Hospital of Haibei Prefecture in Qinghai Province | 345029 | 44439 |  | 02N/Ingenuity Core128/Elite China/need onsite support | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Clean / Lubricate Problem Reported by Customer: 02N/Ingenuity Core128/Elite China Safety Questions:Case: 0118247962 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Bai Jingwen Last updated by: Bai Jingwen 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer_s problems: Other types of error. Detailed description of Customer_s problems or malfunctions: The top of the machine room is leaking water onto the machine. Need to go to the field to check if it can be turned on It is a technology problem with the equipment itself: No. Troubleshooting Action: key fault code or symptom: The machine room is leaking water. The equipment needs to be checked to see if it can be turned on normally Remote troubleshoot and test result: need on site support PRS TR status: It is not through, ask engineer for maintenance on site. Repair Action: Suggest onsite service procedures: 1. Check the water leaking in the frame. Clean and dry it, and try to start the machine Test/Inspection: N/A Test Equipment: N/A SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Internal Remarks: 31.08.2021 08:14:41 UTC Yonghua Xu (86001334) RDI Accuracy: no $Q7:Confirmation reported problem,The CT scanning room is leaking, and the gantry is flooded. It has been dried, and an engineer is required to check and start the machine. $Q8:The CT scanning room is leaking, and the gantry is flooded. It has been dried, and an engineer is required to check and start the machine. Check that the equipment gantry, couch, TEAL and UPS are dry properly. $Q9:Check that the equipment gantry, couch, TEAL and UPS are dry properly.Power on test equipment tube warming and scan normal. $Q10:Test and Inspection data refer to PA , Do a few patient, Observation of a certain period of time. The system is ok. Device returned to full functionality. $Calibrated Tools name (NA) S/N (NA) Next Calibration date (NA) External Remarks: 31.08.2021 08:14:38 UTC Yonghua Xu (86001334) Department: Radiology Department_s CT Room. Contact Person: Guo Rongshou. Contact Phone: 15309708576 Fault Description: The CT machine room leaks water. The main frame was in water, now It is dried. Ask the field service engineer to check and start the machine. Solution/Maintenance measures: The CT machine room leaks water. The main frame was in water, now It is dried. Ask the field service engineer to check and start the machine.Checked the device_s gantry, couch, Teal cabinet and UPS and they are normal. Powered on to test the device, preheating and scanning are normal. Complete corresponding calibration, complete backing up. Device returns to normal working condition after service. ----------------------------------------------------------------------------- |  |  | $Q9:Check that the equipment gantry, couch, TEAL and UPS are dry properly.Power on test equipment tube warming and scan normal. | [12 empty] | 118247962 |  | 1273769 |  |  | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Water | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1431 |  |  | China | Jiangsu Province Yangzhou City No. 1 People_s Hospital, West District Hospital | 345022 | 44438 |  | 06J: Relocation/Ingenuity Core128/Elite China | BU Q&R Feedback: No Reason: Lifecycle Solution Problem Reported by Customer: Relocation Field Service Engineer Notes:This PO is a Reference Number only SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Safety Questions:Case: 0118241254 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Tao Fengming Last updated by: Tao Fengming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | [15 empty] | 118241254 |  | 1269145 |  |  | Closed - Non-Complaint | [20 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1432 |  |  | China | Medical center (contagious disease, He nan, Xu chang) | 40047 | 44438 |  | 02N/Ingenuity CT 728326/need edit examcard | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 02N/Ingenuity CT 728326 Safety Questions:Case: 0118241119 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Guoping Last updated by: Liu Guoping 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | [15 empty] | 118241119 |  | 1269083 |  |  | Closed - Non-Complaint | [20 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1433 |  |  | China | Zhalantun City People_s Hospital | 40045 | 44438 |  | WeChat：02N/Ingenuity CT 728326/The tube overheating | BU Q&R Feedback: No SAP Customer Symptom Code: BOOT Customer Symptom Code: Boot up/Power-up Problem Reason: Corrective Maintenance Engineer Failure Code: Electrical/Electronic Failure Repair Activity Code: Instructed Customer Problem Reported by Customer:The tube overheating Safety Questions:Case: 0118241109 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Zixuan Last updated by: Wang Zixuan 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Troubleshooting Action: answer customer question Repair Action: answer customer question Test/Inspection: Functional Check PASSED. The device returned to full functionality. Test Equipment: N/A SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- |  |  | answer customer question | [12 empty] | 118241109 |  | 1269070 |  |  | Closed - No Further Investigation | [7 empty] | Application | [6 empty] | Q&A - Tube Overheat | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1434 |  |  | China | Dongguan Donghua Hospital Co., Ltd. | 40059 | 44438 |  | WeChat: 02N/Ingenuity CT 728326/scan fault | BU Q&R Feedback: No Reason: Corrective Maintenance Safety Questions:Case: 0118241024 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Zixuan Last updated by: Wang Zixuan 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Problem Reported by Customer: scan fault T2 Activities:Problem Description: Summary of customer_s problems: Scan malfunction Detailed description of Customer_s problems or malfunctions: Scan malfunction It is a technology problem with the equipment itself: Yes. Troubleshooting Action: summary of customer_s problem: Scan malfunction Detailed description of Customer_s problems or malfunctions: Scan malfunction It is a technology problem with the equipment itself: Yes. Repair Action: Suggest onsite service procedures: Check the system log file and service further according to the log file. We recommend spare parts 12NC and names: N/A Memo of Customer request: None SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support External Remarks:07.09.2021 05:51:52 UTC Zhitao Lin (86034034) Fault description: Scan interruption, prompted error Service details: Checked the equipment, analyzed error file, did not find hardware error. Continue to monitor Service result: After service, the machine is back to normal. Other important issues: None Internal Remarks:07.09.2021 05:51:55 UTC Zhitao Lin (86034034) $Q7: (Confirmed reported problem, scan abort ) $Q8: (check log and system hardware, find nothing serious. ) $Q9: (inspect the system and restart the system) $Q10: (Test and inspection data refer to PA . Device returned to full functionality ) $Q11: Calibrated Tools name ( refer to PA ) S/N ( ) Next Calibration date ( refer to PA ) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date: N/A ^Current Tube usage: N/A ----------------------------------------------------------------------------- |  |  | inspect the system and restart the system, did not find hardware error. Continue to monitor | [12 empty] | 118241024 |  | 1268985 |  |  | Investigation | [7 empty] | Others | [6 empty] | Reboot - System | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1435 |  |  | China | 东莞东华医院有限公司 | 40059 | 44500 |  | Wechat:02N/Ingenuity CT 728326/Connection monitoring | T2 Activities : Problem Description: 客户问题概述： 连接MP5 客户问题详细描述或故障现象： 连接MP5监护 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 机架后侧ECG口连接不上 远程诊断与测试结果： 确认心电接口连接是否正常 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 确认心电接口连接是否正常 Test Equipment: N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : FSE林志桃 13570009796 T2 Activities : Problem Description: PPD 24H: 455012301901 ECG Monitor Trigger 7M Ca 1 补发备件 Next Step Action : PPD ECG Monitor Trigger 7M Ca 455012301901 1 已安排24H到医院，FSE：林志桃-86-13570009796 | 3 |  | 服务内容：心电图仪连接不到机架，线和接头有问题，更换两根连接线 服务结果：设备恢复正常工作 | [12 empty] | 118633745 |  | 1475840 |  |  | Opened - Pending Translation | [7 empty] | Others | [6 empty] | ECG | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1436 |  |  | China | 山东省胸科医院 | 40025 | 44499 |  | Wechat:01/Ingenuity CT 728326/Blue screen | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer : Blue screen Safety Questions : Case: 0118631654 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: han susu Last updated by: han susu 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: Cannot login in and stuck at desktop Troubleshooting Action: Cannot login in and stuck at desktop Repair Action: Login Philips account and run workflow, than issue fixed. Consider as Win10 bug. Test Equipment: N/A ----------------------------------------------------------------------------- | 3 |  | Login Philips account and run workflow, than issue fixed. Consider as Win10 bug. Link to Blue screen defect which in SP2 scope, | [12 empty] | 118631654 |  | 1475209 |  |  | Opened - Pending Translation |  | CTUDT00632553 | [5 empty] | Engineering | [6 empty] | Blue Screen | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Blue-Screen | Console-SW-Blue-Screen | Console-SW-Blue-Screen | CTUDT00632553 | Blue-Screen | Console-SW | Console-SW-Blue-Screen | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1437 |  |  | China | 黑龙江省桦川县人民医院 | 345018 | 44498 |  | 02N/Ingenuity Core128/nemo couch cushion | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Replace Part Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ Safety Questions : Case: 0118624895 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Cao Last updated by: Liu Cao 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: Need couch cushion Troubleshooting Action: Need couch cushion Repair Action: Nemo 72h 453567517311 Cushion, Patient Support iCT 1 Test/Inspection: N/A Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | NEMO | [12 empty] | 118624895 |  | 1473233 |  |  | Opened - Pending Translation | [7 empty] | TBD | [6 empty] | NEMO | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | n | y | y | y | y | #N/A
1438 |  |  | China | Wuxi Jiangyuan Hospital | 345020 | 44498 |  | Oil pump has abnormal noise，anode not up to speed. | Oil pump has abnormal noise，anode not up to speed. |  |  | $Q7: (Confirm reported problem. Gantry make noise.) $Q8: (Open gantry cover find Anode make noise.) $Q9: (Replace the tube12NC:989000086352.) | [14 empty] | 1469113 |  |  | Evaluation |  | PartReplacement | [5 empty] | Engineering | [6 empty] | Tube | [5 empty] | Refresh |  |  | System issue | System | IC-HW | TubeReplacement | IC-HW-TubeReplacement | IC-HW-TubeReplacement | PartReplacement |  | IC-HW | IC-HW-TubeReplacement | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | #N/A
1439 |  |  | China | Jilin Province Zhaonan City TCM Hospital | 345006 | 44497 |  | 01/Ingenuity Core128/Elite China/The bed makes a strange noise | Next Step Action:Transfer to sales handling, this case is valid within 90 days starts from this case created, once beyond of this period without any further steps, this case is considered invalid, customer refuses our service T2 Activities:Problem Description: Summary of customer_s problems: The device makes a strange sound. Detailed description of Customer_s problems or malfunctions: The bed is noisy when moving horizontally It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: The bed is noisy when moving Remote troubleshoot and test result: Need an engineer to go to the site to check. PRS TR status: It is not through, ask engineer for maintenance on site. Repair Action: Suggest onsite service procedures: 1. need to check if It is related to coupling replacement | 3 |  | related to coupling replacement, customer refuses our service | [12 empty] | 118619750 |  | 1469974 |  |  | Opened |  | E-125059 | [5 empty] | Engineering | [6 empty] | Couch Noise | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | Lubricant | Couch-HW-Lubricant | Couch-HW-Lubricant | E-125059 | Lubricant | Couch-HW | Couch-HW-Lubricant | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1440 |  |  | China | Sichuan Province Yibin City No.2 T.C.M. Hospital | 345035 | 44497 |  | Application/Ingenuity Core128/Elite China//Enhanced scanning issues | Next Step Action: newly installed device, respond, during enhanced scan, encounter trigger and dosing issue. Need to transfer to app teacher to follow up ASAP, confirm if protocol setup or operation issue.The result needs a receipt, thanks. | 5 |  | Application Support 新装机客户，反应增强扫描，触发和打药出问题，需转应用老师尽快跟进，确认是否协议设置或者操作问题。 | [12 empty] | 118618420 |  | 1469146 |  |  | Opened | [7 empty] | Application | [6 empty] | Application Support | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1441 |  |  | United States | Lackey Memorial Hospital | 40015 | 44496 |  | Table Noises | Integration : Confirmation of order 54360580 saved Integration : According to factory calendar 20211030 (Posting date) is not a workday (check entry) ----------------------------------------------------------------------------- Work Order #: WO-04644426 28-Oct-2021 Denison Chuck The system meets the specification for the performed service and is returned to use Diagnostic performed by Engineer : The Customer said that sometimes there is a noise when lowering the patient support. Checked the couch covers. Made small adjustment to the covers. will monitor the performance of the couch support. Problem Description by Engineer : Customer Function/Role: C.T. Tech Device Use: Philips C.T. Expected and Unexpected Results: None System Up. User Impact: None System up. Patient Impact: None System up. Current Software Version: 4.1.8 | 5 |  | Made small adjustment to the covers. will monitor the performance of the couch support. | [12 empty] | 118616928 |  | 1466558 |  |  | Evaluation | [7 empty] | Others | [6 empty] | Couch Cover | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1442 |  |  | United States | Unity Health Systems | 40037 | 44496 |  | 0306CFX84 Gantry cannot comply | The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Problem description by engineer : Gantry cannot comply When was the first occurrence? Within... : 27-Oct-2021 How often is issue occurring? : First occurrence Malfunction area : Table Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Not provided Patient Impact/Outcome: None User Impact: Customer had to restart Current Software Version: Not provided Diagnostic performed by Engineer : System went down and customer power cycled and did a bugreport now. Biomed will come in the morning and run checks. Br`uce want us to check the bugreport and let him know what may have caused this issue. Internal Comments : Error # and/or description of error : 2021/10/27 02:54:56.520__001 S_HMC_CTRL_A_POSITION_ERROR_LIMIT_EXCEEDED ERROR_NORMAL Series Exit Logger.mdb Checked ACS Code 25 (P1=25) - External problem is preventing the controller to track the position correctly. Not enough (relative) encoder pulses reached the ACS in time Repair Action : Check for an obstruction or mechanical load problem, check for any obstructions that might have caused this matter (Couch friction, Caught mattress, etc.). Resolution : This may have been due to an obstruction on the couch or a hardware glitch. Customer power cycled and system is running as normal now | 2 |  | This may have been due to an obstruction on the couch or a hardware glitch. Customer power cycled and system is running as normal now | [12 empty] | 118610869 |  | 1460632 |  |  | Investigation |  | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Reboot | [5 empty] | Refresh |  |  | System issue | System | Gantry-SW | NoEnoughInfo | Gantry-SW-NoEnoughInfo | Gantry-SW-NoEnoughInfo | NoEnoughInfo | TBD-NoEnoughInfo | Gantry-SW | Gantry-SW-NoEnoughInfo | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1443 |  |  | China | 江苏省徐州医学院附属第三医院 | 40039 | 44496 |  | 02G/Ingenuity CT 728326/add dicom node | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Connection/Cable/Interconnect Repair Activity Code: Instructed Customer Problem Reported by Customer : 02G/Ingenuity CT 728326 Safety Questions : Case: 0118610545 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: add dicom node Troubleshooting Action: instruct customer add dicom node, modify exam card. system is ok, close case. Repair Action: Functional Check PASSED. Device returned to full functionality. Test/Inspection: N/A Test Equipment: N/A ----------------------------------------------------------------------------- |  |  | 44496 | [12 empty] | 118610545 |  | 1464913 |  |  | Closed - Non-Complaint | [7 empty] | 44496 | [6 empty] | 44497.6201 | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1444 |  |  | United States | Leesburg Regional Medical Center | 40022 | 44496 |  | HIS RIS issue | Expected and Unexpected results: Unable to connect to HIS/RIS Patient Impact/Outcome: none User Impact: Manual patient data entry Current Software Version: 4.1.10 Diagnostic performed by Engineer : Able to ping system, Able to connect to CIRS, Customer able to send to other destinations. Unable to log into olevel due to password change. Checked worklist setting = ok. Advised customer to contact PACS admin. Will leave case open if customer needs further assistance. Internal Comments : Error # and/or description of error : No error Repair Action : Able to ping system, Able to connect to CIRS, Customer able to send to other destinations. Unable to log into olevel due to password change. Checked worklist setting = ok. Advised customer to contact PACS admin. Will leave case open if customer needs further assistance. Resolution : No further needs from customer. Closing case. | 3 |  | No further needs from customer. customer to contact PACS admin | [12 empty] | 118609634 |  | 1459632 |  |  | Evaluation | [7 empty] | Others | [6 empty] | PACS Vendor | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1445 |  |  | China | Dongguan Donghua Hospital Co., Ltd. | 40059 | 44495 |  | wechat：02N/Ingenuity CT 728326/Panel fault | T2 Activities:Problem Description: Summary of customer_s problems: Other types of error. Detailed description of Customer_s problems or malfunctions: Gantry Right Panel button is not working. It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: Confirmed the problem with the engineer. Remote troubleshoot and test result: Restart but didn_t work. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Suggest to replace with a new right panel of the frame to eliminate the problem. We recommend spare parts 12NC and names: PPD 459800484511 Next Step Action:PPD CONTROL PANEL ASSEMBLY-FRONT RH 459800484511 1 parts are shipped, FSE: Lin Zhitao-86-13570009796 24H SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support ----------------------------------------------------------------------------- | 3 |  | 服务内容：机架右侧面板按键破损，更换新的面板 服务结果：设备恢复正常工作 | [12 empty] | 118597517 |  | 1460488 |  |  | Opened - Pending Translation |  | CTUDT00632181 | [5 empty] | Engineering | [6 empty] | GCP | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Gantry Panel | Gantry-HW-Gantry Panel | Gantry-HW-Gantry Panel | CTUDT00632181 | Control Panel | Gantry-HW | Gantry-HW-Gantry Panel | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1446 |  |  | Spain | Comercial Centro Medico | 40042 | 44494 |  | Unable to turn off the computer in the daily shutdown | Work Order #: WO-04615355 [REDACTED] The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled. Regulatory Questions : N/A Problem Description by Engineer: We talk with Vanessa, the CT technician. In the mornings, when trying to do tube heating, or start acquisitions, the equipment displays a message indicating a reconstructor failure (CIRS). They have to reset the equipment completely to be able to work. Diagnostic performed by Engineer: The host error logs do not show any detailed information about the cause of the problem; simply the gantry cancels the acquisitions due to a host order. But it is not possible to access the CIRS Event Viewer since there is no direct connection to this computer. Follow Up Required: We should go to the site and check the CIRS event file when acquisitions or tube heating are blocked (see first thing on 10/25). See if the communication between the gantry and CIRS shows then any problem. ----------------------------------------------------------------------------- On Monday 10/25/2021, when launching the tube heating, the process did not start. They rebooted the system and after the acquisition, it gave a CIRS error. In the system error log, there is a main contactor error, which was resolved with the first reboot. Then the IMR rebuilder did not start correctly. They rebooted again and it already worked properly. During my diagnosis, they performed several patients, and the system worked correctly. It is not necessary to intervene in the system because they are specific errors that have not been reproduced the rest of the week. | 3 |  | In the system error log, there is a main contactor error, which was resolved with the first reboot. Then the IMR rebuilder did not start correctly. They rebooted again and it already worked properly. | [12 empty] | 118589840 |  | 1455731 |  |  | Evaluation - Pending Translation |  | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Reboot | [5 empty] | Refresh |  |  | System issue | System | CIRS-SW | NoEnoughInfo | CIRS-SW-NoEnoughInfo | CIRS-SW-NoEnoughInfo |  |  |  |  | No | 0 | 0 | 0 | 0 | 0 | No | y | y | y | y | y | y | y | y | 0 | y | #N/A
1447 |  |  | China | Wuxi City Jiangyuan Hospital | 345020 | 44494 |  | 02N/Ingenuity Core128/Elite China/have noise | T2 Activities:Problem Description: Summary of customer_s problems: Other types of error. Detailed description of Customer_s problems or malfunctions: There is a loud noise in the rack It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: There is a loud noise in the rack Remote troubleshoot and test result: It should be a loud fan noise PRS TR status: It is not through, ask engineer for maintenance on site. Repair Action: Suggest onsite service procedures: Check the fans, SAP Transfer:Dispatched for On site Support ----------------------------------------------------------------------------- | 3 |  | $Q7: (Confirm reported problem. Gantry make noise.) $Q8: (Open gantry cover find Anode make noise.) $Q9: (Replace the tube12NC:989000086352.) | [12 empty] | 118587501 |  | 1455637 |  |  | Opened - Pending Translation |  | PartReplacement | [5 empty] | Engineering | [6 empty] | Tube | [5 empty] | Refresh |  |  | System issue | System | IC-HW | TubeReplacement | IC-HW-TubeReplacement | IC-HW-TubeReplacement | PartReplacement |  | IC-HW | IC-HW-TubeReplacement | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | #N/A
1448 |  |  | Japan | Independent administrative institution National Hospital Organization | 341013 | 44494 |  | HY: Memory replacement work | Work Order #: WO-04610607 27-Oct-2021 TE 005 The system meets the specification for the performed service and is returned to use. Memory replacement operation check Reported to the on-duty technician for approval Parts shipping Part number (12nc) Part 1: 459801633881 Part name Part 1: 64GB DDR4-2666 (8x8GB) RegRAM Number of parts 1: 1 | 5 |  | Replace Memory Set | [12 empty] | 118585611 |  | 1455379 |  |  | Evaluation |  | MemoryReplacement | [5 empty] | Engineering | [6 empty] | Memory | [5 empty] | Refresh |  |  | System issue | System | Console-HW | MemoryReplacement | Console-HW-MemoryReplacement | Console-HW-MemoryReplacement | MemoryReplacement | MemoryReplacement | Console-HW | Console-HW-MemoryReplacement | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1449 |  |  | United States | McLaren Macomb | 40021 | 44494 |  | Have a question about a study | Internal Remarks : CSA INTERACTION TEMPLATE : Interaction Type: New Case Caller: Nadia Lewis - phone # 586 945 4530 IP & Issue: 85380634 - Have a question about a study Event Type: Incident Service Request: Remote Clinical Did Customer Decline Remote Support? No Preferred Service Time: ASAP Entitlement Options: Claims Contract (with reason) Reason for Entitlement Claim: N/A Action Taken: Created case & locked it for entitlement verification Next Step Action: Awaiting Entitlement Verbal PO Number: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- Work Order #: WO-04616049 Maier, Donnie The system meets the specification for the performed service and is returned to use Diagnostic performed by Engineer : Called and left VM @11:30AM EST for contact in case to call back in for assistance. Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: Diagnostic Expected and Unexpected Results: Question about a study User Impact: No Patient Impact: No Current Software Version: 4.1.10 | 5 |  | No clinical applications support is needed | [12 empty] | 118589569 |  | 1453484 |  |  | Evaluation | [7 empty] | Application | [6 empty] | Application Support | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1450 |  |  | Puerto Rico | HOSPITAL DR SUSONI INC | 40013 | 44494 |  | Device does not accepts Login and password | Work Order #: WO-04615609 25-Oct-2021 Ramos Rivera, Pablo The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Problem description by engineer : User is disabled. Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected results: User CT enabled. Patient Impact/Outcome: System is down User Impact: System is down Current Software Version: 4.1.10 Resolution : CT user was disabled. Reseted the CT user password. Now it is working properly. Ran abdomen scans, system is working properly. ----------------------------------------------------------------------------- Work Order #: WO-04614430 Valenzuela Velasquez, Eduardo Ariel The system does not meet the specification for the performed service and is not returned to use Internal Comments : Error # and/or description of error : CT user disabled. Repair Action : - Reset password. - Enable emergency user. - GIve access to customer for the IT user Diagnostic performed by Engineer : CT user is disabled, no login possible. Remote access is limited and reset password is not available. Problem Description by Engineer : Problem description by engineer : User is disabled. Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected results: User CT enabled. Patient Impact/Outcome: System is down User Impact: System is down Current Software Version: 4.1.10 Follow Up Required : Error # and/or description of error : CT user disabled. Repair Action : - Reset password. - Enable emergency user. - GIve access to customer for the IT user | 2 |  | - Reset password. - Enable emergency user. - GIve access to customer for the IT user | [12 empty] | 118589165 |  | 1453243 |  |  | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Password | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1451 |  |  | China | Henan Province Puyang City People’s Hospital | 40031 | 44493 |  | wechat:02N/Ingenuity CT 728326/system crash | T2 Activities:Problem Description: Summary of customer_s problems: Other types of error. Detailed description of Customer_s problems or malfunctions: Computer is crashing. It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: Computer is crashing. Remote troubleshoot and test result: It returned to normal after restart. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Collect bug report, escalation SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Internal Remarks : 27.10.2021 12:57:39 UTC (86038950) Q7:Confirmed report problem: the host crash. Q8:Perform visual check on site, clean the disk and continue to watch. $Q9:clean the disk and continue to watch, Device return to full function. $Q10: test and inspection data refer to PA. Device returned to full functionality. $Q11:Calibrated Tools refer to PA. External Remarks: 27.10.2021 12:57:36 UTC (86038950) service details: host machine crash. Click on any area without response. Clean the disc and junk on the site. Continue to monitor. Service outcome: clean the disc and junk on the field. The equipment is back to normal. Continue to monitor. Other significant matters: none | 5 |  | 服务内容： 主机死机，点击任何地方无响应，到场清理磁盘和垃圾，继续观察。 服务结果： 到场清理磁盘和垃圾，设备恢复正常，继续观察。 Need bugreport for investigation | [12 empty] | 118582823 |  | 1451593 |  |  | Closed - No Further Investigation | [7 empty] | Engineering | [6 empty] | SW Crash | [5 empty] | Refresh |  |  | System issue | System | Console-SW | DiskFull | Console-SW-DiskFull | Console-SW-DiskFull |  |  |  |  | No | 0 | 0 | #REF! | 0 | 0 | No | y | y | y | y | y | y | y | y | y | y | #N/A
1452 |  |  | Poland | Independent Public Health Maintenance Organization of the Ministry of Interior Affairs and Administration in Katowice. | 341008 | 44491 |  | DICOM configuration and inspection of sanitary protection fixed guards. | DICOM configuration and inspection of sanitary protection fixed guards. Work related to the preparation of the tomograph for commissioning in a new location: 20/10/2021 - visit, DICOM configuration + assistance for Sanepid 28/10/2021 - remote service configuration visit. | 4 |  | NAC_DICOM Configuration | [12 empty] | 118578700 |  | 1450273 |  |  | Evaluation - Pending Translation | [7 empty] | NAC | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1453 |  |  | China | Jilin Province Zhaonan City TCM Hospital | 345006 | 44491 |  | 01/Ingenuity Core128/Elite China/Estop opened | Service details: The horizontal movement of the bed loses power. The coupling position of the encoder is damaged, replace. Service outcome: device returns to normal working condition after service Other significant matters: None. Internal Remarks : 24.10.2021 07:15:53 UTC (86038446) RDI Accuracy ：No $Q7: (Confirmed reported problem: E-stop opened when the couch Horizontal motion) $Q8: (Review logger and check the couch, The couch Horizontal coupling damaged) $Q9: (Replaced the part and calibration) $Q10: (Test and inspection data refer to: PA. The device returned to full functionality. $Q11 Calibrated Tools name (FLUKE287C) S/N (27420003) Next Calibration date (03/17/2022) | 3 |  | Replace Horizontal coupling, Adjust the cover and test for the noise issue | [12 empty] | 118576982 |  | 1449651 |  |  | Investigation - Pending Translation |  | CTUDT00632179 | [5 empty] | Engineering | [6 empty] | Coupling | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | Motor Coupling | Couch-HW-Motor Coupling | Couch-HW-Motor Coupling | CTUDT00632179 | Motor Coupling | Couch-HW | Couch-HW-Motor Coupling | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1454 |  |  | China | Jiangxi Province Children’s Hospital | 40041 | 44491 |  | 02N/Ingenuity CT 728326/Gantry nosie | SAP Transfer:Dispatched for On site Support T2 Activities:Problem Description: Summary of customer_s problems: The device makes a strange sound. Detailed description of Customer_s problems or malfunctions: There is a problem with the couch_s cover. It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: NA Remote troubleshoot and test result: The gantry_s rotation is noisy when scanning but not when It is at rest. It is recommended for the customer to consider not using it for now.Contacted the field service engineer, who says that It is a problem with the position of the slide rail pulley that happened before. PRS TR status: It is not through, ask engineer for maintenance on site. Repair Action: Suggest onsite service procedures: Check the rail pulley components. Check for scraping, large belt, etc. We recommend spare parts 12NC and names: " 459800445141 SPLIT TAPER BUSHING； 453566563321 293,76POLY-V-SHEAVE (T84A-1007)" Next Step Action: PPD have arranged for 24h delivery, please dispatch labor. Field service engineer: Li Jiawei - 86-18579112310 External Remarks: 24.10.2021 15:53:30 UTC (86091272)"SUAT. Ready for Clinical Use local time" : "2021-10-23 16:30" Service description: The customer reported that the machine is very noisy. Checked and found that the motor wheel is moving around, causing friction. Replaced the roller cone block and adjusted the mechanical position. Service result: The device returns to normal working condition after service. External Remarks: 25.10.2021 02:18:10 UTC (86091296) Service details: The customer reported that the machine is very noisy. Checked and found that the motor wheel is moving around, causing friction. Replaced the roller cone block and adjusted the mechanical position. Service result: The device returns to normal working condition after service. Internal Remarks : 27.10.2021 10:39:34 UTC (86091272) "RDI Accuracy:Yes" $Q7:(Confirmed reported problem. gantry noise.) $Q8:( Split Taper bushing position moved, The belt is friction with the machine. ) $Q9:( Replace POLY-V-SHEAVE solved the problem.) $Q10:(Test and inspection data refer to PA. Device returned to full functionality.) $Q11:Calibrated Tools: No tools used. Internal Remarks : 27.10.2021 14:51:05 UTC (86091296)"RDI Accuracy:Yes" $Q7:(Confirmed reported problem. gantry noise.) $Q8:( Split Taper bushing position moved, The belt is friction with the machine. ) $Q9:( Replace POLY-V-SHEAVE solved the problem.) $Q10:(Test and inspection data refer to PA. Device returned to full functionality.) $Q11:Calibrated Tools: No tools used. | 2 |  | Split Taper bushing position adjusted. Replace POLY-V-SHEAVE 453566563321 solved the problem. | [12 empty] | 118575835 |  | 1449198 |  |  | Closed - No Further Investigation |  | ECR-076490 | [5 empty] | Engineering | [6 empty] | Sheave | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Sheave | Gantry-HW-Sheave | Gantry-HW-Sheave | ECR-076490 |  | Gantry-HW | Gantry-HW-Sheave | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | #N/A
1455 |  |  | Japan | Independent administrative institution National Hospital Organization | 341013 | 44491 |  | 03DS: RS: Tracker behavior becomes heavy and then E-STOP OPEN >> Reboot and return >> S_GEN_FAULT_ANODE_TUBE_ARC / S_GEN_FAULT_CATHODE_TUBE_ARC | Work Order #: WO-04598981 27-Oct-2021 Yamamoto, Yuko The system meets the specification for the performed service and is returned to use. Resolution: Performed witnessing. Internal Comments: [CX related] Did you agree on how to complete the report with the person in charge of devices? <Select one of the following> Yes. I completed normal reporting that day. [Work report (paper) submission] Did you submit a work report on paper? Yes. Work Order #: WO-04598349 22-Oct-2021 Kawazoe, Junichi The system meets the specification for the performed service and is returned to use. Regulatory Questions : N/A Resolution: E tc [gibberish] Internal Comments: [CX related] Did you agree on how to complete the report with the person in charge of devices? YES On-site the next day [Work report (paper) submission] Did you submit a work report on paper? YES Resolution : Work details: I carried out Long TC and Filament Cal.I acquired the bug report.There is no problem with the operation of the equipment.I reported to Technician Yamamoto that the work was complete. ----------------------------------------------------------------------------- Work Order #: WO-04598090 Satake, Rui The system meets the specification for the performed service and is returned to use. Diagnostic performed by Engineer: After he had switched the key switch on/off and restarted the host, it recovered and scanning was possible. I confirmed a record of an X-ray tube discharge from the log. S_GEN_FAULT_ANODE_TUBE_ARC/S_GEN_FAULT_CATHODE_TUBE_ARC However, I could not determine from the log what had slowed down the host immediately before. I asked the customer to create a bug report.An FSE will respond. Problem Description by Engineer: During a tracker scan, the host_s behavior became sluggish and then the E-Stop opened. >>> "S_GEN_FAULT_ANODE_TUBE_ARC/S_GEN_FAULT_CATHODE_TUBE_ARC" Internal Comments: [Subject] 03DS:RS:Mr. Yamamoto:During a tracker scan, the host_s behavior became sluggish and then the E-Stop opened. >> Following a restart, it recovered. >> "S_GEN_FAULT_ANODE_TUBE_ARC/S_GEN_FAULT_CATHODE_TUBE_ARC" [Internal Comments] None Dispatched: 1 RC FSE Registration Information: Name of FSE 1 who visited the site: Junichi Kawazoe [L2 confirmation] FSE 1 visit date and time: 2021/10/2216:00 FSE 1 hours worked: 4 ----------------------------------------------------------------------------- | 4 |  | I carried out Long TC and Filament Cal. I acquired the bug report.There is no problem with the operation of the equipment. | [12 empty] | 118575342 |  | 1444759 |  |  | Evaluation |  | CTUDT00630085 | [5 empty] | Engineering | [6 empty] | Tube Arc | [5 empty] | Refresh |  |  | System issue | System | IC-HW | Tub Arc | IC-HW-Tub Arc | IC-HW-Tub Arc | CTUDT00630085 | Tub Arc | IC-HW | IC-HW-Tub Arc | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1456 |  |  | Turkey | Republic of Turkey Ministry of Health, Bagcilar | 40010 | 44490 |  | 0115747107/WO-01714930 Torque tool calibration process | Integration: VBAKKOM has been processed successfully-VBAPKOM has been processed successfully-KONVKOM has been processed successfully-KONVKOM has been processed successfully-KONVKOM has been processed successfully-KONVKOM has been processed successfully-REQ2RES DMR 0018257989 has been saved-MPG is present on the product code: P9. ----------------------------------------------------------------------------- | 3 |  | NAC_Torque tool calibration | [12 empty] | 118571566 |  | 1446313 |  |  | Evaluation | [7 empty] | NAC | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1457 |  |  | China | Shaanxi Province Xianyang Central Hospital (Xianyang City No. 2 People_s Hospital) | 40050 | 44490 |  | 02N/Ingenuity CT 728326/Rack power down | T2 Activities:Problem Description: Summary of customer_s problems: Frequent power failure of rack Detailed description of Customer_s problems or malfunctions: Frequent power failure of rack It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: S_MC_DEVICE_VERT_POT_OPEN Remote troubleshoot and test result: N/A PRS TR status: It is not through, ask engineer for maintenance on site. Repair Action: Suggest onsite service procedures: On site, check the couch_s vertical encoder. If necessary, replace it. We recommend spare parts 12NC and names: Vertical Transducer : 459800623331 Memo of Customer request: None SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Next Step Action: PPD parts 459800623331 arranged for 24H to the hospital, Guo Yujie - 86-15158180650 External Remarks: 24.10.2021 02:52:01 UTC (86038990) service details: the machine frequently loses power. Replace the bed_s height POT and it restore. service result: device returns to normal working condition after service ----------------------------------------------------------------------------- | 3 |  | Vertical Transducer : 459800623331 | [12 empty] | 118571505 |  | 1446274 |  |  | Investigation |  | NotEngineering-partreplace | [5 empty] | Engineering | [6 empty] | Vertical Transducer | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | VerticalTransducer | Couch-HW-VerticalTransducer | Couch-HW-VerticalTransducer | NotEngineering-partreplace |  | Couch-HW | Couch-HW-VerticalTransducer | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | #N/A
1458 |  |  | China | Fujian Provincial Hospital | 345011 | 44490 |  | 02N/Ingenuity Core128/Worklist and PACS | Unable to connect PACS and WORKLIST. Troubleshooting found network cable error. Replace the network cable with The Hospital_s IT dept. Service result (required): Device returns to normal working condition after service Other significant matters: none Internal Remarks : 22.10.2021 03:56:29 UTC (86036497) “RDI Accuracy: Yes” $Q7: confirmed the reported problem, Pacs and worklist are disconnected. $Q8: on site Reviewed the error log, Performed visual inspection and found Network cables are faulty. $Q9: Cooperate with hospital information department to replace network cables. $Q10: Test and Inspection data refer to:PA.Device returned to full functionality. $Q11: Calibrated Tools refer to:PA ^Current Tube installed date: NA ^Current Tube usage: NA | 3 |  | Network cables are faulty. | [12 empty] | 118569398 |  | 1444853 |  |  | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Net cable | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1459 |  |  | Japan | Maki Hospital | 341017 | 44489 |  | HY: Arc countermeasure work | Work Order #: WO-04583810 27-Oct-2021 Yamamoto, Yoshiyuki The system meets the specification for the performed service and is returned to use. Regulatory Questions : N/A Problem Description by Engineer: Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : Work details: Arching has been occurring occasionally on the X-ray tube while it is warming up (Short TC), and given the frequency of the problem, I thought it could be due to accumulated impurities (causing arcing) in the X-ray tube, so I performed aging (Long TC). Internal Comments: [CX related] Did you agree on how to complete the report with the person in charge of devices? Yes. I completed normal reporting that day. [Work report (paper) submission] Did you submit a work report on paper? <Yes> Q1. How did the examination finish? ・As it stopped before X-ray emission, no extra irradiation Q2. Were there any mistaken findings (misdiagnosis, mistaken treatment) due to the problem? None. Q3. Were there any artifacts? None. Q4. Was there any unintended unnecessary exposure? None. ----------------------------------------------------------------------------- | 5 |  | it could be due to accumulated impurities (causing arcing) in the X-ray tube, so I performed aging (Long TC). Also replace Part/Product - ECG Monitor Trigger 7M Ca Part Number - 455012301901 | [12 empty] | 118564450 |  | 1441305 |  |  | Closed - No Further Investigation |  | CTUDT00630085 | [5 empty] | Engineering | [6 empty] | Tube Arc | [5 empty] | Refresh |  |  | System issue | System | IC-HW | Tub Arc | IC-HW-Tub Arc | IC-HW-Tub Arc | CTUDT00630085 | Tub Arc | IC-HW | IC-HW-Tub Arc | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1460 |  |  | China | 北京市中西医结合医院 | 40032 | 44489 |  | Wechat:02N/Ingenuity CT 728326/Screen fault | T2 Activities : Problem Description: 客户问题概述： 其他类型故障。 客户问题详细描述或故障现象： 显示器不亮 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 显示器不亮 远程诊断与测试结果： 数据线连接问题 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 检查显示器连线 SAP Transfer : Dispatched for Onsite Support | 3 |  | found DVI cable broken,after change backup cable,device work normal. | [12 empty] | 118562409 |  | 1439956 |  |  | Closed - No Further Investigation |  | CTUDT00632173 | [5 empty] | Engineering | [6 empty] | DVI Cable | [5 empty] | Refresh |  |  | System issue | System | Console-HW | Adaptor | Console-HW-Adaptor | Console-HW-Adaptor | CTUDT00632173 | Adaptor | Console-HW | Console-HW-Adaptor | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1461 |  |  | China | 陕西省咸阳中心医院 (咸阳市第二人民医院） | 40050 | 44489 |  | wechat:02N/Ingenuity CT 728326/tube can not expose | SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Problem Reported by Customer : tube can not expose T2 Activities : Problem Description: 客户问题概述： 扫描失败。 客户问题详细描述或故障现象： 机架无法扫描。 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： PRS不通。 远程诊断与测试结果： 重启无效。 PRS TR状态： 不通，请工程师到场检查。 Repair Action: 建议现场维修步骤： 建议现场根据log或event信息排查机器故障。 建议备件12NC及名称： N/A External Remarks : 20.10.2021 03:07:37 UTC (86038990)"SUAT. Ready for Clinical Use local time" : "2021-10-20 11:00" 服务内容：不能曝光，重启后恢复。 服务结果：服务后，设备正常工作。 ----------------------------------------------------------------------------- | 2 |  | 服务内容：不能曝光，重启后恢复。 服务结果：服务后，设备正常工作。 review error logs ,found CIRS event log report error ,ACQ not ok, check the CIRS optical fiber is connect, perform the VD test, not found error, so to perform the restart the CIRS computer, device runing normal, so think customer observation and use. $Q9:maybe is CIRS computer not normal start, restart CIRS device return normal., | [12 empty] | 118561943 |  | 1439686 |  |  | Closed - No Further Investigation |  | CTUDT00630408 | [5 empty] | Engineering | [6 empty] | ACQ_NOT_OK | [5 empty] | Refresh |  |  | System issue | System | CIRS-SW | ACQ_NOT_OK | CIRS-SW-ACQ_NOT_OK | CIRS-SW-ACQ_NOT_OK | ACQ_NOT_OK | ACQ_NOT_OK | CIRS-SW | CIRS-SW-ACQ_NOT_OK | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | Reboot | #N/A
1462 |  |  | Japan | Maki Hospital | 341017 | 44489 |  | 作業日程に関して>>FSEより回答 |  |  |  | 44489 | [12 empty] | 118561986 |  | 1439708 |  |  | Closed - Non-Complaint | [7 empty] | 44489 | [6 empty] | 44490.3875 | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1463 |  |  | Australia | Maia Financial Pty Limited | 341000 | 44490 |  | Quality Control Computer Network Configuration | Integration : Confirmation of order 54329236 saved Integration : According to factory calendar 20211023 (Posting date) is not a workday (check entry) ----------------------------------------------------------------------------- Work Order #: WO-04589670 Steve, Teixeira The system meets the specification for the performed service and is returned to use Problem Description by Engineer : ________________________________________________________________ CUSTOMER REPORTED PROBLEM OR EVENT DESCRIPTION: 21 Oct 2021 Customer called requesting assistance to change the AE Titles for the Quality Control Computers, in the CT network configuration. ________________________________________________________________ Diagnostic performed by Engineer : ________________________________________________________________ CUSTOMER REPORTED PROBLEM OR EVENT DESCRIPTION: 21 Oct 2021 Customer called requesting assistance to change the AE Titles for the Quality Control Computers, in the CT network configuration. ________________________________________________________________ CONFIGURATION: 21 Oct 2021 Assisted the Customer with the configuration of the RAD600 and RAD600Vidar network nodes. Customer to test new configuration. ________________________________________________________________ FOLLOW UP: 22 Oct 2021 Contacted site and confirmed that transfers successful. Customer had to do further configuring, as IP Addresses also required amending. The AE Title and IP address changed to: RA600 - AE Title: NC_MUR_RA_02 - 10.64.65.143 RA600Vidar - AE Title : NC_MUR_RA_01 - 10.64.65.142. ________________________________________________________________ Resolution : ________________________________________________________________ CASE CLOSED: As the transfers to the Quality Control computers are successful after the CT network configuration, as agreed with the Customer, no further service related action will be taken on this case. ________________________________________________________________ ----------------------------------------------------------------------------- |  |  | 44490 | [12 empty] | 118568654 |  | 1439684 |  |  | Closed - Non-Complaint | [7 empty] | 44490 | [6 empty] | 44490.3806 | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1464 |  |  | Japan | 医療法人清翠会 牧病院 | 341017 | 44488 |  | 03CS: KKI: Short TC stopped at 76% >> S_GEN_FAULT_ANODE_TUBE_ARC / S_RHOST_HVG_HSS_GENERATOR_ARC_FAUL | Work Order #: WO-04576935 Kimura, Kazumasa The system meets the specification for the performed service and is returned to use Problem Description by Engineer : ShortTCが76%で止まっていた＞＞S_GEN_FAULT_ANODE_TUBE_ARC/S_RHOST_HVG_HSS_GENERATOR_ARC_FAUL Resolution : 【対応】 継続して検査可 Internal Comments : [Subject] 03CM:KKI:ﾄｶﾞﾜ様:ShortTCが76%で止まっていた＞＞S_GEN_FAULT_ANODE_TUBE_ARC/S_RHOST_HVG_HSS_GENERATOR_ARC_FAUL [Internal Comments] 【ログ確認結果】 S_GEN_FAULT_ANODE_TUBE_ARC S_RHOST_HVG_HSS_GENERATOR_ARC_FAUL 陽極側にてアークが発生しておりました。 様子見とさせて頂いております。 [4/8 Elements] Q5 その検査はどの様に終了しましたか?:無しQ6 問題により間違った所見（誤診、誤治療）の有無:無し Q7 意図しない不要な曝射の有無:無し Q8 アーチファクトの有無:無し ----------------------------------------------------------------------------- | 4 |  | Duplicate with 118564450 | [12 empty] | 118558926 |  | 1436681 |  |  | Closed - No Further Investigation | [7 empty] | TBD | [6 empty] | Tube Arc | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | n | y | y | y | y | #N/A
1465 |  |  | China | 山东省泰安市中医二院 | 345037 | 44488 |  | 02N/Ingenuity Core128/Elite China/report error | T2 Activities : Problem Description: 客户问题概述： 扫描失败。 客户问题详细描述或故障现象： 无法扫描 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： S_COUCHMGR_HOR_MOTION_STOP_ERROR；S_HMC_VELOCITY_ENCODER_FAILURE ；S_MPS_WAITING_FOR_MAIN_DRIVE_TO_CLOSE 远程诊断与测试结果： 床水平运动掉电一次，后main contactor不吸合。客户告知大关过一次无效。晚上没人上班，无法尝试再重启，沟通后先派工。 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 尝试大关机器稍久再开机；检查gullwing switch等main contactor相关；检查床水平阻力，考虑观察掉电故障是否偶发一次 SAP Transfer : Dispatched for Onsite Support Next Step Action : JIN XIAO External Remarks : 19.10.2021 10:53:12 UTC (86038373)"SUAT. Ready for Clinical Use local time" : "2021-10-19 18:30" External Remarks : 19.10.2021 11:19:59 UTC (86038373)故障现象: CT扫描失败 服务内容: 按压前盖，重启机架 服务结果: 维修后,设备恢复正常. Internal Remarks : 19.10.2021 11:54:05 UTC (86038373) “RDI Accuracy: YES" $Q7: (Confirmed reported problem, CT scanning failed) $Q8: ( Check gantry log, found "S_MPS_MAIN_DRIVE_NOT_CLOSING" error) $Q9: ( Press gantry front cover and reboot CT gantry) $Q10: ( Test and inspection data refer to: PA. Device returned to full functionality) $Q11: Calibrated Tools name ( N/A) ^Current Tube installed date: N/A ^Current Tube usage: N/A ----------------------------------------------------------------------------- | 2 |  | Press gantry front cover and reboot CT gantry, system back to normal | [12 empty] | 118557515 |  | 1436279 |  |  | Investigation | [7 empty] | Engineering | [6 empty] | Front Cover | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Misalignment | Gantry-HW-Misalignment | Gantry-HW-Misalignment | 0 | Noise | Gantry-HW | Gantry-HW-Misalignment | No | 0 | 0 | 0 | 0 | 0 | No | y | y | 0 | 0 | 0 | y | y | y | y | y | #N/A
1466 |  |  | China | 广东省广州市第八人民医院（嘉禾院区） | 345032 | 44487 |  | Wechat:02N/Ingenuity Core128/Elite China/air conditioning temperature is too high | BU Q&R Feedback: No SAP Customer Symptom Code: CENV Customer Symptom Code: Customer Environmental Concern Reason: Corrective Maintenance Engineer Failure Code: Environmental Cause Repair Activity Code: Instructed Customer Problem Reported by Customer : Wechat:02N/Ingenuity Core128/Elite China Safety Questions : Case: 0118550159 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Bai Jingwen Last updated by: Bai Jingwen T2 Activities : Troubleshooting Action: air conditioning，contact third party Repair Action: air conditioning，contact third party Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: N/A | 5 |  | NAC_air conditioning problem | [12 empty] | 118550159 |  | 1430864 |  |  | Closed - No Further Investigation | [7 empty] | NAC | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1467 |  |  | United Kingdom | Philips | 40016 | 44487 |  | adaptor was faulty, replaced the adaptor cable | Whilst installing the monitors for the first time, one of the minitors would not come up correctly. After swapping ports on the back of the Host PC and changing DP adaptors we found that the adaptor was faulty. This is also the case with the iQon on the same site. I believe other sites are also having the same issue during install. Not able to change the date in the boxes below. Retrospective complaint. 14th Oct 2021. FOR UKI REFERENCE ONLY: Feedback type: Technical issue UKI KMO: Paul Rothwell Submitter email: warren.thorpe@philips.com |  |  | DVI adapter | [14 empty] | 1428290 |  |  | Closed - Done |  | CTUDT00632173 | [5 empty] | Engineering | [6 empty] | DVI adapter | [5 empty] | Refresh |  |  | System issue | System | Console-HW | Adaptor | Console-HW-Adaptor | Console-HW-Adaptor | CTUDT00632173 | Adaptor | Console-HW | Console-HW-Adaptor | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1468 |  |  | Japan | Maki Hospital | 341017 | 44486 |  | 04CS: KKI: Bed does not work >> S_MPS_MAIN_DRIVE_NOT_CLOSING szInfo: MP: Waiting for MDC to close ..> Restored by TEAL OFF / ON | 04CS: KKI: Bed does not work >> S_MPS_MAIN_DRIVE_NOT_CLOSING szInfo: MP: Waiting for MDC to close ..> Restored by TEAL OFF / ON . [Subject] 04CS: KKI: Noda: Bed does not work >> S_MPS_MAIN_DRIVE_NOT_CLOSING szInfo: MP: Waiting for MDC to close ..> Restored by TEAL OFF / ON [Internal Comments] [Log confirmation result] The following error has occurred. S_MPS_MAIN_DRIVE_NOT_CLOSING szInfo: MP: Waiting for MDC to close .. Since the key switch was turned off twice, the gantry power supply did not turn off, so I turned the key switch twice. It is probable that the gantry side was hung because the gantry did not start up normally after that. It seems to be a temporary phenomenon, so I will wait and see. [4/8 Elements] Q5 How did the test end ?: None Q6 Presence or absence of incorrect findings (misdiagnosis, mistreatment) due to problems: None Q7 Presence or absence of unintended unnecessary exposure: None Q8 Presence / absence of artifacts: None | 4 |  | Since the key switch was turned off twice, the gantry power supply did not turn off, so I turned the key switch twice. It is probable that the gantry side was hung because the gantry did not start up normally after that. It seems to be a temporary phenomenon, so I will wait and see. | [12 empty] | 118548042 |  | 1429673 |  |  | Closed - No Further Investigation |  | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Reboot | [5 empty] | Refresh |  |  | System issue | System | Gantry-SW | NoEnoughInfo | Gantry-SW-NoEnoughInfo | Gantry-SW-NoEnoughInfo |  |  |  |  | No | 0 | 0 | 0 | 0 | 0 | No | y | y | y | y | y | y | y | y | 0 | y | #N/A
1469 |  |  | China | Jilin Province Zhaonan City TCM Hospital | 345006 | 44484 |  | 01/Ingenuity Core128/Elite China/Mouse fault | N/A Next Step Action: ASP jilinyiyang follow Transfer to sales handling, this case is valid within 90 days starts from this case created, once beyond of this period without any further steps, this case is considered invalid, customer refuses our service | 3 |  | out of warranty, customer refuses our service | [12 empty] | 118541561 |  | 1424965 |  |  | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Mouse | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1470 |  |  | China | 河南省新乡市凤泉区大块镇卫生院 | 345009 | 44483 |  | 02N/Ingenuity Core128/Elite China/gantry noise | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity Core128/Elite China T2 Activities : Problem Description: 客户问题概述： 设备异响。 客户问题详细描述或故障现象： 设备异响，静止时也响 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 设备异响，静止时也响 远程诊断与测试结果： 怀疑探测器风扇 PRS TR状态： 不通，请工程师到场检查。 Repair Action: 建议现场维修步骤： 检查噪音来源，根据发现做相关的troubleshooting 建议备件12NC及名称： "BR64 Air Fan Replacement Kit 459800980091 Replaces BR AIR FAN REPLACEMENT KIT 455011509231; 000130" SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : PPD placement Kit 459800980091 1 已安排 36H到医院，FSE：郭海滨-86-18738187087 Safety Questions : Case: 0118534604 | 3 |  | $Q8:perform visiual and conform the noise was from the fan of the DMS. and the fan was wore out $Q9:replace the defective fans | [12 empty] | 118534604 |  | 1421695 |  |  | Closed - No Further Investigation |  | CTUDT00632110 | [5 empty] | Engineering | [6 empty] | DMS Fan | [5 empty] | Refresh |  |  | System issue | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | DMS Module | DMS-HW | DMS-HW-DMS Module | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1471 |  |  | China | 曹妃甸区医院 | 345016 | 44483 |  | Wechat:02N/Ingenuity Core128/Elite China/CT box bed lowering button cannot be used | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : CT box bed lowering button cannot be used Safety Questions : Case: 0118534137 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: mo xueke Last updated by: mo xueke 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 其他类型故障。 客户问题详细描述或故障现象： CT box的降床的键不好用 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： CT box的降床的键不好用 远程诊断与测试结果： CT box的降床的键不好用 PRS TR状态： 无PRS。 Repair Action: 建议现场维修步骤： 检查CT box，如有必要进行更换 建议备件12NC及名称： 453567027195 CT box T2 Activities : Problem Description: 客户问题概述： 其他类型故障。 客户问题详细描述或故障现象： CT box的降床的键不好用 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： CT box的降床的键不好用 远程诊断与测试结果： CT box的降床的键不好用 PRS TR状态： 无PRS。 Repair Action: 建议现场维修步骤： 检查CT box，如有必要进行更换 建议备件12NC及名称： "CT BOX ASSEMBLY_BEST 459801705022" SAP Transfer : Dispatched for Onsite Support Next Step Action : PPD CT BOX ASSEMBLY_BEST 459801705022 1 已安排24H到医院，FSE：张翔-86-18600120460 SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | CT BOX ASSEMBLY_BEST 459801705022 | [12 empty] | 118534137 |  | 1421355 |  |  | Closed - No Further Investigation |  | CTUDT00632185 | [5 empty] | Engineering | [6 empty] | CT Box | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | CT Box | Gantry-HW-CT Box | Gantry-HW-CT Box | CTUDT00632185 | CT Box | Gantry-HW | Gantry-HW-CT Box | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1472 |  |  | China | 东南大学附属中大医院江北院区 | 345021 | 44483 |  | 02n/Ingenuity Core128/Elite China/exposure failed | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02n/Ingenuity Core128/Elite China Safety Questions : Case: 0118532093 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 其他类型故障。 客户问题详细描述或故障现象： 不能曝光 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： 不能曝光 远程诊断与测试结果： 不能曝光，机架报错，重启后还是不行 PRS TR状态： 不通，请工程师到场检查。 Repair Action: 建议现场维修步骤： 查看报错信息，指导医院正常使用注意事项 SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support ----------------------------------------------------------------------------- | 2 |  | $Q7: (Confirmed reported problem,CIRS boot up failed and system report CIRS connection error) $Q8: (on site check CIRS statues，found during boot up，system can not detect CPU fans ) $Q9: (reconnect CPU fans and computer cover) | [12 empty] | 118532093 |  | 1420172 |  |  | Investigation |  | ShouldbeService | [5 empty] | TBD | [6 empty] | CPU Fan | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | n | y | y | 0 | y | #N/A
1473 |  |  | Germany | Gesundheit Nord gGmbH Klinikverbund | 40002 | 44483 |  | In the case of the thorax, some of the lung windows are not sent to the PACS | Internal Remarks: Mr. Deerberg, question on voicemail, can he pull the error log here? Customer Comment: Did not reach the customer. Customer comment: Appointment discussed with Ms. Schichtl. Friday, November 5, 2021, approx. 9:00 am Problem Description by Engineer: Error description: Sniffer from Mr. Hallermann was attached to the customer During the thorax examinations, in part, the lung windows are not sent to the PACS Error has already been known for approx. 1 year. Please check remotely Error message: - Error occurred on: Yesterday and today 14.10.21 Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- Work Order #: WO-04546390 Sandner, Michael The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Follow-up required: Request for a technical service to copy the data from the laptop; no system time required; colleague should only copy the data from the laptop since that does not work remotely; then consultation with R. Kraschweski and send him the data; copy data regarding the following periods: 12.10. Bollhagen, Doris 11:09 am 14.10. Böttcher, Gisela 10:03 am Mureia, Correia 10:35 am Problem Description by Engineer: Error description: Sniffer from Mr. Hallermann was attached to the customer During the thorax examinations, in part, the lung windows are not sent to the PACS Error has already been known for approx. 1 year. Please check remotely Error message: - Error occurred on: Yesterday and today 14.10.21 Diagnostic performed by Engineer: return call Ms. Schichtl, discussion of the problem; call to Th. Halama - only voicemail; 3 cases since the day before yesterday, sometimes it works automatically, sometimes not, in approx. 30% of the case, there is no automatic transmission to the PACS. ----------------------------------------------------------------------------- | 4 |  | In the case of the thorax, some of the lung windows are not sent to the PACS Error has been known for about 1 year. 3 cases since the day before yesterday, sometimes it works automatically, sometimes not, about 30% does not go into PACS automatically | [12 empty] | 118534170 |  | 1417111 |  |  | Evaluation |  | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | PACS Transfer_Lung | [5 empty] | Refresh |  |  | System issue | System | Console-SW | NoEnoughInfo | Console-SW-NoEnoughInfo | Console-SW-NoEnoughInfo | CTUDT00632185 |  | Console-SW | Console-SW-NoEnoughInfo | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | #N/A
1474 |  |  | China | Jilin Province Songyuan City Central Hospital | 345008 | 44482 |  | 02n/Ingenuity Core128/TheGantry power down | BU Q&R Feedback: No Reason: Corrective Maintenance Internal Remarks : 16.10.2021 15:28:30 UTC (86038306) "RDI Accuracy:Yes" $Q7: (Confirmed reported problem&estop open) $Q8: (review the system logger，the R2D error，the R2D damaged） $Q9: (replaced the R2D） $Q10: (Test and inspection data refer to: PA. Device returned to full functionality) $Q11 Calibrated Tools name (FLUKE287) S/N (16490094) Next Calibration date (10/19/2021) SAP Transfer:Case Transferred to SAP Problem Reported by Customer: 02n/Ingenuity Core128/Elite China SAP Transfer:Dispatched for On site Support T2 Activities:Problem Description: Summary of customer_s problems: Gantry intermittently power off. It is a technology problem with the equipment itself: Yes. Troubleshooting Action: remote diagnosis and testing result: It is very difficult to determine the problem. We suspect the couch and signal brush.。。problem PRS TR status: It is not through, ask engineer for maintenance on site. Repair Action: Suggest onsite service procedures: It is recommended to have a field service engineer go on site to check the error log and determine the problem. Memo of Customer request: contact hospitalm when arrive Safety Questions:Case: 0118526792 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Weng Jinming Last updated by: Weng Jinming Next Step Action: PPD called to contact the field service engineer to confirm the alternate number 459800467341*1, shipped 24H, Chi Xiaodong - 86-18943690010 External Remarks: 14.10.2021 11:31:59 UTC (86038306) Service details: The device intermittently powers off, R2D malfunction, replace. Service outcome: device returns to normal working condition after service Other significant matters: None. ----------------------------------------------------------------------------- | 3 |  | replaced the R2D 459800467341 | [12 empty] | 118526792 |  | 1416266 |  |  | Investigation |  | CTUDT00630864 | [5 empty] | Engineering | [6 empty] | R2D | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | R2D | Gantry-HW-R2D | Gantry-HW-R2D | CTUDT00630864 | R2D | Gantry-HW | Gantry-HW-R2D | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1475 |  |  | China | Dongguan Donghua Hospital Co., Ltd. | 40059 | 44481 |  | Image trasferring gets stuck; Had the customer reset the spooler to resolve. | BU Q&R Feedback: No SAP Customer Symptom Code: HANG Customer Symptom Code: Hang/Crash Reason: Corrective Maintenance Engineer Failure Code: Configuration Error Repair Activity Code: Configuration Change Problem Rep T2 Activities: Problem Description: Automatic sending of images gets stuck. Troubleshooting Action: Tried to have the customer reset spooler to reset. The customer said he will reset it at work on October 13th, then observe again. Repair Action: Tried to have the customer reset spooler to reset. The customer said he will reset it at work on October 13th, then observe. Test/Inspection: customer confirm can scan patient and image ok.functional Check PASSED. The device returned to full functionality. Test Equipment: N/A SAP Transfer:Case Transferred to SAP | 3 |  | reset spooler | [12 empty] | 118520957 |  | 1412285 |  |  | Closed - Done |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS Transfer | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1476 |  |  | China | Shandong Province Chest Hospital | 40025 | 44481 |  | WeChat：02N/Ingenuity CT 728326/software hang up | T2 Activities:Problem Description: Summary of customer_s problems: Other types of error. Detailed description of Customer_s problems or malfunctions: SW stuck. It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: SW stuck. Remote troubleshoot and test result: There is still not enough memory. FSE requires more memory PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Increase memory 0 459801813461 16GB DDR4-2933 (2x8GB) RegRAM SAP Transfer:Case Transferred to SAP External Remarks: 13.10.2021 07:50:51 UTC (86038405) error: blue screen when turned on to login, software is crashing. Service details: add memory to host Z4. Service result: device returns to normal working condition after service. Next Step Action: huang zhaokun SAP Transfer:Dispatched for On site Support Next Step Action: PPD 459801813461*1 shipped, 24H Huang Zhaokun-86-18654536122 | 3 |  | Increase memory 459801813461 16GB DDR4-2933 (2x8GB) | [12 empty] | 118519183 |  | 1411148 |  |  | Investigation |  | Performance Slow | [5 empty] | Engineering | [6 empty] | Performance Slow | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Performance Slow | Console-SW-Performance Slow | Console-SW-Performance Slow | Performance Slow | Performance Slow | Console-SW | Console-SW-Performance Slow | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1477 |  |  | China | No.8 People_s Hospital of Guangzhou City, Guangdong Province (Jiahe Branch) | 345032 | 44481 |  | 02N/Ingenuity Core128/Elite China/IRS disconnected | T2 Activities:Problem Description: Customer consult the machine life span. Troubleshooting Action: inform customer the machine life span. Device is normal. Repair Action: Inform customer the machine life span, and the machine is normal. Test/Inspection: customer confirm can scan patient and image ok.functional Check PASSED. The device returned to full functionality. Test Equipment: N/A T2 Activities:Problem Description: Summary of customer_s problems: The host cannot be connected to CIRS Detailed description of Customer_s problems or malfunctions: Instruct customer, unplug and replug in hard disc, still the same It is a technology problem with the equipment itself: Yes. Troubleshooting Action: remote diagnosis and testing result: Suspect CIRS, hardware issue PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Suggest engineer to go to the field to check CIRS, and reinstall the software.Teach customer to use Philips WeChat to request service SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support External Remarks: 12.10.2021 05:28:32 UTC (86034061)"SUAT. Ready for Clinical Use local time" : "2021-10-12 14:30" Service content: CIRS can not start. Memory failure. Replace the memory and test, back to normal Service outcome: Device returns to normal working condition after service. ----------------------------------------------------------------------------- | 2 |  | replace memory and finish recommended test 服务内容: 重建计算机无法开机，内存故障，更换内存，测试恢复正常 服务结果: 服务完成后,设备恢复正常工作 | [12 empty] | 118518739 |  | 1410844 |  |  | Closed - No Further Investigation |  | MemoryReplacement | [5 empty] | Engineering | [6 empty] | Memory | [5 empty] | Refresh |  |  | System issue | System | Console-HW | MemoryReplacement | Console-HW-MemoryReplacement | Console-HW-MemoryReplacement | MemoryReplacement | MemoryReplacement | Console-HW | Console-HW-MemoryReplacement | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1478 |  |  | China | Fujian Provincial Hospital | 345011 | 44481 |  | 02N/Ingenuity Core128/Elite China | BU Q&R Feedback: No Reason: Corrective Maintenance External Remarks: 15.10.2021 02:31:02 UTC (86036497) Service details: CT software occasionally lagging.Check on the field to confirm It is the application software which is abnormal.Replace hard disc, install OS and app, import backup file, and complete corresponding configuration service result: device returns to normal working condition after service Internal Remarks : 15.10.2021 02:31:05 UTC (86036497)RDI Accuracy: Yes $Q7: (Confirmed reported problem. The Host Apps run slowly.) $Q8: (On site check and find the OS and Apps are abnormal.) $Q9: (Replace the Host hard disk and re-install OS & Apps. And then do restore.) $Q10: (Test and inspection data refer to: PA.+Device returned to full functionality.) $Q11: Calibrated Tools name (FLUKE 287C Multimeter) S/N (16490128) NextCalibration date (10/08/2022) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date: NA ^Current Tube usage: NA Problem Reported by Customer: 02N/Ingenuity Core128/Elite China Safety Questions:Case: 0118518507 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Guoping Last updated by: Liu Guoping 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer_s problems: SW stuck. Detailed description of Customer_s problems or malfunctions: SW stuck. It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: SW stuck. " Remote troubleshoot and test result: Reboot no avail. Need to replace the hard disc and re install the software PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Replace the hard disk and reinstall the software. We recommend spare parts 12NC and names: HP TurboDrive 512GB TLC Z4 SSD w/adapter 459801652541 SAP Transfer:Case Transferred to SAP External Remarks: 15.10.2021 00:49:15 UTC (86036224) Service details: CT software occasionally lagging.Check on the field to confirm It is the application software which is abnormal.Replace hard disc, install OS and app, import backup file, and complete corresponding configuration service result: device returns to normal working condition after service Internal Remarks : 15.10.2021 00:49:18 UTC (86036224) RDI Accuracy: Yes $Q7: (Confirmed reported problem. The Host Apps run slowly.) $Q8: (On site check and find the OS and Apps are abnormal.) $Q9: (Replace the Host hard disk and re-install OS & Apps. And then do restore.) $Q10: (Test and inspection data refer to: PA.+Device returned to full functionality.) $Q11: Calibrated Tools name (FLUKE 287C Multimeter) S/N (16490128) NextCalibration date (10/08/2022) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date: NA ^Current Tube usage: NA SAP Transfer:Dispatched for On site Support Next Step Action: PPD 459801652541*1 shipped, 24H Ding Shi duo-86-13509388496 Next Step Action: Ding shiduo | 3 |  | Replace the Host hard disk(459801652541) and re-install OS & Apps. And then do restore. | [12 empty] | 118518507 |  | 1410696 |  |  | Closed - No Further Investigation |  | HardDiskReplacement | [5 empty] | Engineering | [6 empty] | SW slow | [5 empty] | Refresh |  |  | System issue | System | Console-HW | HardDiskReplacement | Console-HW-HardDiskReplacement | Console-HW-HardDiskReplacement | HardDiskReplacement | HardDiskReplacement | Console-HW | Console-HW-HardDiskReplacement | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1479 |  |  | Peru | BANCO BBVA PERU | 40034 | 44481 |  | Boton de CTbox no funciona | Work Order #: WO-04534498 [REDACTED] The system meets the specification for the performed service and is returned to use. Problem Description by Engineer: Customer Function/Role: Field Service Engineer Device Use: Diagnostic equipment Expected and Unexpected Results: CTBox button does not allow to move the table down. User Impact: NA Patient Impact: NA Current Software Version: 4.1.10 Diagnostic performed by Engineer: CTBox button to move the table down does not work. CTBox needs to be replaced. Follow Up Required: [REDACTED] Contact: [REDACTED] ----------------------------------------------------------------------------- | 5 |  | Product/Part - CT BOX ASSEMBLY_BEST Part number - 459801705022 | [12 empty] | 118524225 |  | 1410393 |  |  | Closed - No Further Investigation |  | CTUDT00632185 | [5 empty] | Engineering | [6 empty] | CT Box | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | CT Box | Gantry-HW-CT Box | Gantry-HW-CT Box | CTUDT00632185 | CT Box | Gantry-HW | Gantry-HW-CT Box | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1480 |  |  | Japan | Maki Hospital | 341017 | 44479 |  | 03DM: RS: Error occurred during test scan >> Rescan is possible >> S_GEN_FAULT_ANODE_TUBE_ARC | Work Order #: WO-04520283 Satake, Rui Problem Description by Engineer: Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer: N/A ----------------------------------------------------------------------------- | 4 |  | LTC is required | [12 empty] | 118512793 |  | 1406421 |  |  | Closed - No Further Investigation |  | CTUDT00630085 | [5 empty] | Engineering | [6 empty] | Tube Arc | [5 empty] | Refresh |  |  | System issue | System | IC-HW | Tub Arc | IC-HW-Tub Arc | IC-HW-Tub Arc | CTUDT00630085 | Tub Arc | IC-HW | IC-HW-Tub Arc | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1481 |  |  | United States | Odessa Regional Medical Center | 40029 | 44479 |  | 1001CDN81 Issues with QC. | The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: System will not initiate User Impact: System unavailable for use Patient Impact: Current Software Version: Diagnostic performed by Engineer : checked system and verified issue. ----------------------------------------------------------------------------- Work Order #: WO-04519987 Mark Finnigan The system meets the specification for the performed service and is returned to use Follow Up Required : Spoke with Mia on the customer callback. She is not reporting any issues with artifacts but she is stating that she is having s tough time getting QC to run. It sounds like the phantom holder needs to be adjusted. Make sure that it is level and centered in bith axis and not skewed from the top perspective. Internal Comments : The customer originally called in Sunday with a complaint about a possible ring. I asked that they do an air calibration but they were too busy at that time. I was not able to get a hold of them latter in the day. Called back on Monday and spoke with Mia. She believes that an air calibration was never done and that she was too busy to do one today. She added that they have been having issues running QC as it fails. No other details are provided. I would start by checking the phantom holder alignment and leveling. Their PRS is not working so no further remore assistance can be rendered. The site is connected but all ports are blocked preventing any access. Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: System will not initiate User Impact: System unavailable for use Patient Impact: Current Software Version: Diagnostic performed by Engineer : Spoke with the CT tech. They could not get the system to initiate. Had them power cycle the machine and it worked properly after the restart. Diagnostic performed by Engineer : Spoke with Mia on the customer callback. She is not reporting any issues with artifacts but she is stating that she is having s tough time getting QC to run. It sounds like the phantom holder needs to be adjusted. Make sure that it is level and centered in bith axis and not skewed from the top perspective. ----------------------------------------------------------------------------- | 2 |  | power cycle the machine and it worked properly after the restart. | [12 empty] | 118512628 |  | 1406026 |  |  | Closed - No Further Investigation |  | ShouldbeCustomer | [5 empty] | Engineering | [6 empty] | Reboot | [5 empty] | Refresh |  |  | System issue | System | Console-SW | ShouldbeCustomer | Console-SW-ShouldbeCustomer | Console-SW-ShouldbeCustomer | ShouldbeCustomer |  | Console-SW | Console-SW-ShouldbeCustomer | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | y | 0 | 0 | y | y | y | 0 | y | #N/A
1482 |  |  | China | 北京市中西医结合医院 | 40032 | 44478 |  | 02N/Ingenuity CT 728326/network change | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity CT 728326 SAP Transfer : Dispatched for Onsite Support SAP Transfer : Case Transferred to SAP Safety Questions : Case: 0118510577 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 其他类型故障。 客户问题详细描述或故障现象： worklist 修改参数 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： worklist 修改参数 远程诊断与测试结果： 需要工程师到场帮忙修改worklist参数 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 修改worklist参数 ----------------------------------------------------------------------------- | 3 |  | NAC_Change network configuration | [12 empty] | 118510577 |  | 1403271 |  |  | Closed - No Further Investigation | [7 empty] | NAC | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1483 |  |  | China | Wuhan Province Wuhan Ren’ai Hospital | 345004 | 44478 |  | 05K FOR IB/Ingenuity Core128/Elite China | T2 Activities:Problem Description: Summary of customer_s problems: Images transmission failure. Detailed description of Customer_s problems or malfunctions: The images can not be transmitted to Shenfei Yun and PACS It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: "The images can not be transmitted to Shenfei Yun and PACS " Remote troubleshoot and test result: PACS is not through, reset invalid PRS TR status: It is not through, ask engineer for maintenance on site. Repair Action: Suggest onsite service procedures: Check network, reset spooler SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Next Step Action: Customer called in again to push, informed FSE Hu to respond to customer External Remarks: 10.10.2021 15:01:36 UTC (86036472) Service details: [Image transmission to Shenfei Yun and PACS failed. Re-setup IP, restore after connecting to hospital network. After service, device is restored] service result: [device returns to normal working condition after service] Internal Remarks : 10.10.2021 15:11:48 UTC (86036472) RDI Accuracy: Yes $Q7: (Confirm reported problem & Image cannot be transmitted to shenfeiyun and PACS） $Q8: (The scene inspection found that the image could not be transmitted. Due to the doctor_s feedback that the film could be printed, the data was suspected to be stuck. Spooler reset was executed, but the transmission still failed after the restart, and the LAN was connected The config test network was also abnormal. The network IP was detected and it was found that the second IP was set by the hospital, and the second IP was also lost. The information department was asked to re-assign the address, and the Settings were restored） $Q9: ( Reset the IP address to connect to the hospital network then the system recovery) $Q10: (Test and inspection data refer to PA ,Device return to full functionality) $Q11: Calibrated Tools name :N/A For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date:N/A ^Current Tube usage: scan second_N/A ----------------------------------------------------------------------------- | 3 |  | Reset the IP address to connect to the hospital network then the system recovery | [12 empty] | 118510268 |  | 1403172 |  |  | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | PACS Config | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | n | y | y | y | y | #N/A
1484 |  |  | China | Harbin City Sen Hai Hospital | 345007 | 44476 |  | Wechat:05K for IB/Ingenuity Core128/Elite China/High pressure injector failure | Problem Reported by Customer : Fan inoperative Next Step Action: The console enhanced small computer_s fan is not working. Remote Service Engineer (RSE) Weng Jinming Next Step Action: Remote Service Engineer (RSE) feedback: need to dispatch labor Next Step Action: Remote service engineer (RSE) feedback: Sorry, the hospital_s description is wrong, It is a high-pressure injector issue, contact the third party. Next Step Action: Called Teacher Zhu back at the hospital to consult about the manufacturer_s information. Teacher Zhu informed that the manufacturer has been contacted, and now it has been fixed. | 3 |  | NAC_Injector problem | [12 empty] | 118496929 |  | 1396890 |  |  | Evaluation | [7 empty] | NAC | [6 empty] | NAC | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1485 |  |  | Poland | Government Agency. | 341002 | 44477 |  | Przeniesienie aparatu z kontenera do budynku głównego | Integration:Check the settlement rule-Entry 1 in table OPERATION (method CHANGE) was not found-Error during processing of BAPI methods-MPG is present on the product code : P7 Integration: Confirmation of order 54273802 saved. Integration:According to factory calendar 20211023 (Posting date) is not a workday (check entry) ----------------------------------------------------------------------------- Work Order #: WO-04511999 15-Oct-2021 Mr. Mariusz Pachla The system meets the specification for the performed service and is returned to use Tools & Equipment : MULTIMETER,2022-07-24 00:00:00 Problem Description by Engineer: Moving the device from the container to the main building. Resolution: Moving the device from the container to the main building. ----------------------------------------------------------------------------- Work Order #: WO-04511990 22-Oct-2021 Gerard Dabrowka The system meets the specification for the performed service and is returned to use Tools & Equipment : TORQUE WRENCH,2021-10-26 00:00:00 TORQUE WRENCH,2021-10-26 00:00:00 Problem Description by Engineer: order to transfer the device. Diagnostic performed by Engineer: Transfer supervision, disassembly of the device, installation in a new location. ----------------------------------------------------------------------------- |  |  | 44477 | [12 empty] | 118506501 |  | 1402011 |  |  | Closed - Non-Complaint | [7 empty] | 44477 | [6 empty] | 44478.7653 | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1486 |  |  | China | Wuhan Province Wuhan Ren’ai Hospital | 345004 | 44477 |  | 05K for IB/Ingenuity Core128/Elite China/ | BU Q&R Feedback: No Reason: Corrective Maintenance External Remarks: 15.10.2021 06:57:03 UTC (86036493) Service details: [Doctor forgot password, failed to log in to scan software. Restored to normal after resetting ct account password. After service, device is restored] service result: [device returns to normal working condition after service] Internal Remarks : 15.10.2021 06:57:06 UTC (86036493)RDI Accuracy: Yes $Q7: (Confirm reported problem: The doctor forgot his password and couldn_t log into the scanning software) $Q8: (When I tried to log in with IT account, I found that the password was also wrong, so I went in to reset the password in maintenance mode. However, I happened to enter the account incorrectly for many times, and the account was locked. After unlocking, I reset the CT account password） $Q9: ( Reset the CT account password then the system recovery) $Q10: (Test and inspection data refer to PA ,Device return to full functionality) $Q11: Calibrated Tools name :N/A For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date:N/A ^Current Tube usage: scan second_N/A Problem Reported by Customer: 05K for IB/Ingenuity Core128/Elite China Safety Questions:Case: 0118505272 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Guoping Last updated by: Liu Guoping 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer_s problems: Other types of error. Detailed description of Customer_s problems or malfunctions: Forgot password and failed to log in It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: Forgot password and failed to log in " Remote troubleshoot and test result: PRS is not through, customer doesn_t know IT password PRS TR status: It is not through, ask engineer for maintenance on site. Repair Action: Suggest onsite service procedures: 1. reset password; 2. replace IT password, and ask customer to keep the account safe SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Next Step Action: hu fangtao External Remarks: 09.10.2021 09:43:01 UTC (86036472)"SUAT. Ready for Clinical Use local time" : "2021-10-08 23:00" External Remarks: 10.10.2021 15:15:18 UTC (86036472) Service details: [Doctor forgot password, failed to log in to scan software. Restored to normal after resetting ct account password. After service, device is restored] service result: [device returns to normal working condition after service] Internal Remarks : 10.10.2021 15:18:21 UTC (86036472) RDI Accuracy: Yes $Q7: (Confirm reported problem & The doctor forgot his password and couldn_t log into the scanning software） $Q8: (When I tried to log in with IT account, I found that the password was also wrong, so I went in to reset the password in maintenance mode. However, I happened to enter the account incorrectly for many times, and the account was locked. After unlocking, I reset the CT account password） $Q9: ( Reset the CT account password then the system recovery) $Q10: (Test and inspection data refer to PA ,Device return to full functionality) $Q11: Calibrated Tools name :N/A For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date:N/A ^Current Tube usage: scan second_N/A ----------------------------------------------------------------------------- |  |  | 44477 | [12 empty] | 118505272 |  | 1401508 |  |  | Closed - Non-Complaint | [7 empty] | 44477 | [6 empty] | 44478.616 | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1487 |  |  | Australia | Maia Financial Pty Limited | 341000 | 44476 |  | PACS Changes | Work Order #: WO-04522710 20-Oct-2021 Ruben MARQUEZ The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: Radiographer Device Use: in use Expected and Unexpected Results: Change Network DICOM nodes. User Impact: NA Patient Impact: NA Current Software Version: 4.1.10 Resolution : Assisted David in creating new DICOM nodes and amending local AET changes on the system. Tested all destinations - all passed. ----------------------------------------------------------------------------- |  |  | 44476 | [12 empty] | 118499748 |  | 1398575 |  |  | Closed - Non-Complaint | [7 empty] | 44476 | [6 empty] | 44477.859 | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1488 |  |  | Turkey | ADIYAMAN OZEL SAGINSTURTAS SAN VE TIC A.S | 341006 | 44475 |  | Image artifact, cleaned the UHR filter | Resolution : The system was returned to the use in functional condition after cleaning the UHR filter and air calibration. ----------------------------------------------------------------------------- Work Order #: WO-04495201 Salih Mendi The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Problem description by engineer : Artefact on images Information to support the complaint handling process Customer Function/Role: Not provided Device Use: Not provided Expected and Unexpected results: Not provided Patient Impact/Outcome: Not provided User Impact: Not provided Current Software Version: Not provided. Diagnostic performed by Engineer : There are artefacts on the images using the UHR filter due to contrasting agent leak/contamination. Internal Comments : Error # and/or description of error : Poor image quality caused from contrast on UHR filter. Repair Action : The problem occurs only with UHR filter imaging. The contrasting agent should be removed from the UHR filter. Resolution : The user was remotely supported for the image quality problem. In diagnosis, we found that the artefact was caused by contrasting agent leak. The site should be visited to clean the UHR filter. If replacement doesn_t solve the problem, UHR filter should be ordered for replacement. The system is partially up. Follow Up Required : Error # and/or description of error : Poor image quality caused from contrast on UHR filter. Repair Action : The problem occurs only with UHR filter imaging. The contrasting agent should be removed from the UHR filter. | 3 |  | artefact was caused by contrasting agent leak, Poor image quality caused from contrast on UHR filter. | [12 empty] | 118493421 |  | 1389887 |  |  | Closed - Done | [7 empty] | Others | [6 empty] | Contrast | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1489 |  |  | Japan | Maki Hospital | 341017 | 44475 |  | 02CS:KK:固定帯劣化。 | Work Order #: WO-04491841 ----------------------------------------------------------------------------- Work Order #: WO-04491836 Kojima, Kenichi The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Resolution : Restraint strap is deteriorated and he asks for a new one. Problem Description by Engineer : Deteriorated restraint strap. ----------------------------------------------------------------------------- | 4 |  | Parts Consumed Description: 2.00000,453567009951,L-6" PATIENT RESTRAINT STRAP/ 42" | [12 empty] | 118490725 |  | 1387963 |  |  | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Strap | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1490 |  |  | China | Shandong Province Tai_an City T.C.M. No. 2 Hospital | 345037 | 44474 |  | 02N/Ingenuity Core128/Elite China/can not rebuild | Problem Reported by Customer: can not rebuild Next Step Action: Remote Service Engineer (RSE) Cao Liu Next Step Action: Calling back was successful. SAP Transfer:Case Transferred to SAP External Remarks : 06.10.2021 13:44:58 UTC (86038373) Fault Description: There are fewer images reconstructed Service content: Clean and unplug GPU graphic card Service outcome: Device returns to normal working condition after servicing. Internal Remarks : 07.10.2021 13:43:33 UTC (86038373) “RDI Accuracy: YES" $Q7: (Confirmed reported problem, CT image reconstruction failed ) $Q8: (Do CT image reconstruction manually, found it was also failed) $Q9: ( Clean and reseat GPU card) $Q10: (Test and inspection data refer to: PA. Device returned to full functionality) $Q11: Calibrated Tools name ( N/A) ^Current Tube installed date: N/A ^Current Tube usage: N/A Next Step Action:FSE Tao Chen 018560023293 call in and disp him Unable to rebuild. ----------------------------------------------------------------------------- | 3 |  | Clean and reseat GPU card | [12 empty] | 118484594 |  | 1388065 |  |  | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | GPU | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1491 |  |  | Japan | Maki Hospital | 341017 | 44473 |  | 02CS:KK:固定帯劣化。 | Work Order #: WO-04474261 ----------------------------------------------------------------------------- Work Order #: WO-04474233 Kojima, Kenichi The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Deteriorated restraint strap. Resolution: The restraint strap is deteriorated, so he wants us to send a new one. ----------------------------------------------------------------------------- | 4 |  | 453566451921,14 INCH RESTRAINT STRAP | [12 empty] | 118477364 |  | 1381004 |  |  | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | Strap | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1492 |  |  | United States | Unity Health Systems | 40037 | 44474 |  | Monthly QA | Integration : Confirmation of order 54257308 saved ----------------------------------------------------------------------------- Work Order #: WO-04487915 05-Oct-2021 Taylor, Durrance The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: tech Device Use: scheduled event Expected and Unexpected Results: monthly qa User Impact: none Patient Impact: none Current Software Version: 4.1.10 Resolution : Monthly QA completed ----------------------------------------------------------------------------- |  |  | 44474 | [12 empty] | 118487789 |  | 1387210 |  |  | Closed - Non-Complaint | [7 empty] | 44474 | [6 empty] | 44475.1368 | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1493 |  |  | United States | Leesburg Regional Medical Center | 40022 | 44472 |  | 0602CDN84 Left side UP/Down button hard to push. | Work Order #: WO-04473095 05-Oct-2021 Mayhew Bruce The system meets the specification for the performed service and is returned to use Resolution : Left UP/Down buttons hard to push. No Patient Involved Expected and Unexpected Results:Left UP/Down buttons hard to push. User Impact: can not scan Patient Impact: None Current Software Version: 4.1.10 Resolution : Replaced left panel on front of gantry and tested successfully. Problem Description by Engineer : Front left panel vertical up button is difficult to press. Parts Consumed : 1.00000,459801484991,CONTROL PANEL ASSEMBLY-FRONT LEFT ----------------------------------------------------------------------------- Work Order #: WO-04490158 05-Oct-2021 Denning, Dale The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : Helped replace the controller and tested out the system. ----------------------------------------------------------------------------- Work Order #: WO-04473086 Miller, John M The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Follow Up Required : Errors: Control Panel Right side UP/DN hard to press. Suggest checking cabling connections and/or replace Control Panel. 459800484511 - CONTROL PANEL ASSEMBLY-FRONT RH Internal Comments : Errors: Control Panel Right side UP/DN hard to press. Suggest checking cabling connections and/or replace Control Panel. 459800484511 - CONTROL PANEL ASSEMBLY-FRONT RH Diagnostic performed by Engineer : Collect information from customer and view log files. Problem Description by Engineer : Left UP/Down buttons hard to push. No Patient Involved Expected and Unexpected Results:Left UP/Down buttons hard to push. User Impact: can not scan Patient Impact: None Current Software Version: 4.1.10 ----------------------------------------------------------------------------- | 3 |  | 459800484511 - CONTROL PANEL ASSEMBLY-FRONT RH | [12 empty] | 118476542 |  | 1380784 |  |  | Closed - No Further Investigation |  | CTUDT00632181 | [5 empty] | Engineering | [6 empty] | GCP | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Gantry Panel | Gantry-HW-Gantry Panel | Gantry-HW-Gantry Panel | CTUDT00632181 | Control Panel | Gantry-HW | Gantry-HW-Gantry Panel | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1494 |  |  | Chile | BANCO DE CHILE | 341005 | 44473 |  | CL1688-FALTA CABLE SG PARA MONITOR | Integration:Confirmation of order 54251519 saved ----------------------------------------------------------------------------- Work Order #: WO-04480200 Parts Consumed : 1.00000,455012301901,ECG Monitor Trigger 7M Ca ----------------------------------------------------------------------------- Work Order #: WO-04479707 [REDACTED] The system meets the specification for the performed service and is returned to use. Internal Comments : Error # and/or description of error : NA Repair Action: NA Follow Up Required: Error Number and/or Description: NA Repair Action: NA Technician_s Name: [REDACTED] Dispatch to: CLINICA AVANSALUD PROVIDENCIA SA Address: Salvador 100 Providencia Contact Person: [REDACTED] Diagnostic performed by Engineer: The customer does not find the interconnection cable between the ECG monitor and the Gantry. Problem Description by Engineer : Problem description by engineer : The customer does not find the connection cable between the ECG monitor and the Gantry. When was the first occurrence? Within... : 04-Oct-2021 09:30:00 How often is issue occurring? : Continuously Malfunction area : Accessories Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected results: Cardiology studies cannot be performed. Patient Impact/Outcome: Cardiology studies cannot be performed. User Impact: Cardiology studies cannot be performed. Current Software Version: 4 ----------------------------------------------------------------------------- |  |  | 44473 | [12 empty] | 118481906 |  | 1382414 |  |  | Closed - Non-Complaint | [7 empty] | 44473 | [6 empty] | 44473.9667 | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1495 |  |  | United States | Philips Medical Capital | 40053 | 44470 |  | 0699CDN80 Squeak noise during gantry rotation | Internal Remarks : FADB: (FSE Aware of Double Booking) Title: Internal Remarks Initials: Twila Actions Taken: Dispatched WO to FSE/Fairchild, Andrew for 10/4/21 @ 1pm EST. Outcome: Dispatched WO per PFSE_s Request (McDowell Preston) IP: 88993323 Next Steps: None **************************** Note: I called PFSE/McDowell Preston - He stated to Dispatch WO to FSE/Fairchild, Andrew - Being they_re already aware of Customer_s issue Integration : Confirmation of order 54243944 saved Integration : Confirmation of order 54254068 saved ----------------------------------------------------------------------------- Work Order #: WO-04483938 05-Oct-2021 Parts Consumed : 1.00000,453566492401,POLY-VEE BELT ----------------------------------------------------------------------------- Work Order #: WO-04466982 05-Oct-2021 Fairchild, Andrew The system meets the specification for the performed service and is returned to use Follow Up Required : Belt replacement next PM Resolution : Replaced tensioner, recommend a new belt to be installed next PM Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: 4.1.10 Parts Consumed : 1.00000,453567003451,TENSIONER ASSY ----------------------------------------------------------------------------- Work Order #: WO-04483594 05-Oct-2021 McDowell Preston The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : replaced the belt tensioner and adjusted the bracket for smooth operation. verified no noise, and also performed gantry rotor balance. System fully operational at this time. ----------------------------------------------------------------------------- | 3 |  | replaced the belt tensioner 453567003451,TENSIONER ASSY and adjusted the bracket for smooth operation | [12 empty] | 118471860 |  | 1379945 |  |  | Closed - No Further Investigation |  | ECR-077012 | [5 empty] | Engineering | [6 empty] | Belt Tensioner | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Belt | Gantry-HW-Belt | Gantry-HW-Belt | ECR-077012 | Belt | Gantry-HW | Gantry-HW-Belt | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1496 |  |  | Japan | Maki Hospital | 341017 | 44470 |  | 03CM:KK:Short TC中に停止。再スタートで完了。>>S_GEN_FAULT_ANODE_TUBE_ARC | Integration : Confirmation of order 54242329 saved Integration:According to factory calendar 20211009 (Posting date) is not a workday (check entry) ----------------------------------------------------------------------------- Work Order #: WO-04465091 Kojima, Kenichi The system meets the specification for the performed service and is returned to use. Resolution: S_GEN_FAULT_ANODE_TUBE_ARC had occurred. He restarted it and it completed. Problem Description by Engineer : Short TC stopped partway through.Completed after restart.>> S_GEN_FAULT_ANODE_TUBE_ARC Internal Comments: [Subject] 03CM:KK:Mr. Hoshi: Short TC stopped partway through.Completed after restart.>> S_GEN_FAULT_ANODE_TUBE_ARC [4/8 Elements] Q5. How did the examination end? None. Q6. Did any mistaken findings (mistaken diagnosis or treatment) result from the problem? No. Q7. Was there any unintended unnecessary exposure? No Q8 Artifacts? None ----------------------------------------------------------------------------- | 4 |  | Short TC stopped partway through.Completed after restart. | [12 empty] | 118470518 |  | 1379514 |  |  | Closed - No Further Investigation |  | CTUDT00630085 | [5 empty] | Engineering | [6 empty] | Tube Arc | [5 empty] | Refresh |  |  | System issue | System | IC-HW | Tub Arc | IC-HW-Tub Arc | IC-HW-Tub Arc | CTUDT00630085 | Tub Arc | IC-HW | IC-HW-Tub Arc | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1497 |  |  | Puerto Rico | HOSPITAL METROPOLITANO SAN GERMAN | 40017 | 44470 |  | It does not perform studies | 07-Oct-2021 Hernandez, Alvin The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Problem description by engineer : Gantry error. Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: No gantry errors. Patient Impact/Outcome: System is down. User Impact: System is down. Current Software Version: 4.1.10 Diagnostic performed by Engineer : The filament circuit on the cathode power module is defective. This is a complete assembly and has to be replaced. Parts Consumed : 1.00000,453567560002,Spellman 80 KW RoHS CPM Assy Resolution : The cathode power module was delivered to the hospital facilities. The CPM module was installed and calibrated. All the calibrations concerning the generator replacement were performed. A new system backup was made. The CT was tested and left fully operational. ----------------------------------------------------------------------------- Work Order #: WO-04480516 06-Oct-2021 Pinto-Rodriguez, Felipe The system meets the specification for the performed service and is returned to use Resolution : The Cathode Power Module has been replaced. Problem Description by Engineer : Filament errors Diagnostic performed by Engineer : There are Cathode Power Module filament errors. ----------------------------------------------------------------------------- Work Order #: WO-04468193 01-Oct-2021 Hernandez, Alvin The system does not meet the specification for the performed service and is not returned to use Tools & Equipment : MULTIMETER,2022-04-30 00:00:00 Follow Up Required : The Cathode Power Module needs to be replaced. The Cathode has a filament problem. Diagnostic performed by Engineer : The following actions were done: 1. I create a Bug report from the error events data. 2. Using the Can-Trace tool I evaluate the problem in real-time. Can-Trace reported filament fault. 3. The ATP reports three errors: - Filament fault. - No filament current. - No filament feedback. 4. I did an X-ray tube inspection: -The Cathode module had an electric spike mark. - The high tension cable was visually verified. The Cables are ok. 5. The X-ray tube filament was measured. - The Filament is not open. - Filament measures 0.29 ohms. The filament is OK. - The four-pin connections on each side of the high tension cable are not internally open. They pass the conductivity test. 6. The No-Load Test was done. The test failed. This means that the X-ray tube is not the problem. The conclusion is that the Cathode Power Module is defective. This error is very likely to be caused by many incoming power failures during the week. The Cathode Power Module needs to be replaced. Recommendations: It is greatly advised to connect the Brilliance Ingenuity CT to a UPS system. These will prevent the CT from damage and downtimes caused by the power failures disruptions. Resolution : The Power Cathode Module was ordered. It will be delivered to the hospital facilities. The Part information is the following: Part Desc: Spellman 80 KW RoHS CPM Assy. Part Num: 453567560002. Deliver to Hospital Metro San German. As soon as the part arrives I will be there to install it. Problem Description by Engineer : Problem description by engineer : Gantry error. Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: No gantry errors. Patient Impact/Outcome: System is down. User Impact: System is down. Current Software Version: 4.1.10 ----------------------------------------------------------------------------- Work Order #: WO-04467269 Valenzuela Velasquez, Eduardo Ariel The system does not meet the specification for the performed service and is not returned to use Follow Up Required : Error # and/or description of error : S_RHOST_HVG_HSS_GENERATOR_NON_ARC_FAULT Repair Action : - Check log file in detail Technician Name : Valenzuela Velasquez, Eduardo Ariel Problem Description by Engineer : Problem description by engineer : Gantry error. Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: No gantry errors. Patient Impact/Outcome: System is down. User Impact: System is down. Current Software Version: 4.1.10 Diagnostic performed by Engineer : A gantry error is present and the tube warm up is not possible. Remote access is not working properly. The case is sent to planning. Internal Comments : Error # and/or description of error : S_RHOST_HVG_HSS_GENERATOR_NON_ARC_FAULT Repair Action : - Check log file in detail ----------------------------------------------------------------------------- | 2 |  | The filament circuit on the cathode power module is defective. This is a complete assembly and has to be replaced. Parts Consumed : 1.00000,453567560002,Spellman 80 KW RoHS CPM Assy | [12 empty] | 118472055 |  | 1378117 |  |  | Closed - No Further Investigation |  | TBD-APM/CPM | [5 empty] | Engineering | [6 empty] | CPM | [5 empty] | Refresh |  |  | System issue | System | IC-HW | APM/CPM | IC-HW-APM/CPM | IC-HW-APM/CPM | TBD-APM/CPM | APM/CPM | IC-HW | IC-HW-APM/CPM | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1498 |  |  | Japan | Independent administrative institution National Hospital Organization | 341013 | 44470 |  | AF:固定帯交換 | Integration : Confirmation of order 54241525 saved ----------------------------------------------------------------------------- Work Order #: WO-04464183 06-Oct-2021 Yamamoto, Yuko The system meets the specification for the performed service and is returned to use. Regulatory Questions : N/A Parts Consumed: 2.00000,453567009961,L-14" PATIENT RESTRAINT STRAP/ 42" 2.00000,453567009951,L-6" PATIENT RESTRAINT STRAP/ 42" Internal Comments: [CX related] Did you agree on how to complete the report with the person in charge of devices? <Select one of the following> Yes. I completed normal reporting that day. [Work report (paper) submission] Did you submit a work report on paper? Yes. Q1. How did the examination finish? Select from the following: *Because the answer is No for Q1, not applicable Q2. Were there any mistaken findings (misdiagnosis, mistaken treatment) due to the problem? None. Q3. Were there any artifacts? None. Q4. Was there any unintended unnecessary exposure? None. Problem Description by Engineer: Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : I replaced the patient restraint strap. October 6, 2021 / I checked with the person in charge Mr. Yamamoto. ----------------------------------------------------------------------------- |  |  | 44470 | [12 empty] | 118469809 |  | 1379329 |  |  | Closed - Non-Complaint | [7 empty] | 44470 | [6 empty] | 44471.6819 | [5 empty] | Refresh |  |  | System issue | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1499 |  |  | United States | Adventhealth Zephyrhills | 300106 | 44498 |  | ECM password expired | CT SWare Ver = 4.1.10 Clin Apps req_ed RSE support ~ Exam Card Mgr password expired and cust cant reset password. RSE advised to login with HCF define Admin password, then reset ECMgr password. Cust says they do not have Admin password. FSE please assist customer reseting ECMgr password ~ educate and assist customer so they understand password system and have their HCF specific password noted where they can be found. Problem Description by Engineer : Customer Function/Role: Case from CT Apps Device Use: Diag Expected and Unexpected Results: Reset Exam Card Mgr password User Impact: None Patient Impact: None Current Software Version: 4.1.10 Diagnostic performed by Engineer : Accepted case from CT Clin Apps ~ As customer did not know HCF defined Admin password ~ we were not able to assist them resetting their Exam Card Mgr password. FSE dispatched. | 4 |  | Reset password according Refresh policy | [12 empty] | 118629203 |  | 1471670 |  |  | Evaluation | [7 empty] | Customer | [6 empty] | Password | [5 empty] | Refresh |  |  | Upgrade | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1500 |  |  | United States | BIBB MEDICAL CENTER | 310175 | 44497 |  | Can CT system perform DEXA scans | Laura indicated that her old DEXA system is down and was inquiring if the CT system would be capable of gathering Bone Mineral Density images. After some investigation their CT system is capable of gathering this information, however a Bone Mineral Post processing packaged must be purchased to take these measurements after a CT scan. I had inquired if she needed me to get a quote but Laura said she would reach out to her Account Manager. A follow up call was performed to make sure this quote was generated for Laura. No rescan, no loss of data, no reinjection Per Laura. Laura had no further questions or needed for clinical support. | 5 |  | NAC_Q&A | [12 empty] | 118621855 |  | 1466633 |  |  | Evaluation | [7 empty] | NAC | [6 empty] | NAC | [5 empty] | Non-Refresh |  |  | Upgrade | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1501 |  |  | United States | Christ Hospital | 300205 | 44495 |  | Intermittent problem with cardiac gating. | Integration : Confirmation of order 54353188 saved Integration : According to factory calendar 20211030 (Posting date) is not a workday (check entry) ----------------------------------------------------------------------------- Work Order #: WO-04634273 Mark Finnigan The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: Biomed Device Use: No patient involved. Expected and Unexpected Results: The cardiac gating waveform has had intermittent disconnects. User Impact: The customer has to power cycle the gantry when this happens. Patient Impact: Current Software Version: Diagnostic performed by Engineer : Spoke with Corey in biomed. He staes that the cardiac signal has experienced intermittent drop-outs requiring a power cycle to recover. The new CAT 6 cable has already been installed in place of the USB extenders. The cable between the ECG monitor and gantry has already been replaced, I recommended that Corey check and adjust the iBox 5VDC power supply to read 5.2 at the iBox. The next step would be to flash the iBox. If that does not help then the iBox should be replaced. Corey will take the first two steps and let us know if that helped. Diagnostic performed by Engineer : Called the customer back looking for an update. Call went to VM. Will try again later today. Diagnostic performed by Engineer : Corey called back in . The PS was at 5.185 volts. He flashed the iBox. He is going to proactively order a iBox board and needed the part number. The PN is 459800166652 Resolution : Provided the customer with te part number requested. He will complete the repair. The customer does not report to Philips so this case is closed as a resolution and complete. ----------------------------------------------------------------------------- | 4 |  | The PS was at 5.185 volts. He flashed the iBox. He is going to proactively order a iBox board and needed the part number. The PN is 459800166652 | [12 empty] | 118606839 |  | 1457901 |  |  | Investigation | [20 empty] | Non-Refresh |  |  | Upgrade | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1502 |  |  | United States | Philips Medical Capital | 300057 | 44495 |  | 0999CFXXX Error message: Artifact stating a crack compensator | Work Order #: WO-04632649 Powell, Jamie The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Problem description by engineer : Error message: Artifact stating a crack compensator while running air cals Information to support the complaint handling process Customer Function/Role: Not provided Device Use: Not provided Expected and Unexpected results: Not provided Patient Impact/Outcome: Not provided User Impact: Not provided Current Software Version: Not provided Diagnostic performed by Engineer : Spoke to customer who states that she saw the message yesterday but she_s not seeing any artifacts or any issues. Internal Comments : Error # and/or description of error : Down loading calibration files and no issues found. Ran cracked compensator tool and no indication found. Could be a false alarm or contrast in the field. Repair Action : Down loading calibration files and no issues found. Ran cracked compensator tool and no indication found. Could be a false alarm or contrast in the field. Resolution : No issues found with image quality and no indication of cracked comp. ----------------------------------------------------------------------------- | 4 |  | No issues found with image quality and no indication of cracked comp. | [12 empty] | 118605667 |  | 1457221 |  |  | Closed - No Further Investigation | [7 empty] | Others | [6 empty] | False Alarm | [5 empty] | Non-Refresh |  |  | Upgrade | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1503 |  |  | United States | Select Medical Corporation | 32261 | 44490 |  | 0599CFXXX E-stop not closing | Integration : Confirmation of order 54332955 saved ----------------------------------------------------------------------------- Work Order #: WO-04594414 Boyd Craig The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: E-stop not closing User Impact: None Patient Impact: None Current Software Version: Diagnostic performed by Engineer : E-stop not closing on bootup. Power cycle system PRS accessed. ULV: no normal errors seen. System operational. Test scans are good. Resolution : Power cycle system PRS accessed. ULV: no normal errors seen. System operational. Test scans are good. ----------------------------------------------------------------------------- | 2 | [14 empty] | 118572701 |  | 1442695 |  |  | Investigation | [20 empty] | Non-Refresh |  |  | Upgrade | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1504 |  |  | United States | Adventhealth Zephyrhills | 300106 | 44489 |  | Q/O L3 Support requested a video from Brain Perfusion application, color map issue | Integration : Confirmation of order 54326617 saved ----------------------------------------------------------------------------- Work Order #: WO-04585496 20-Oct-2021 Garcia Marrero, Jorge The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : Capture video of Brain Perfusion application from host computer monitor. ----------------------------------------------------------------------------- | 5 |  | Make a video from Brain Perfusion application, color map issue for engineering investigation | [12 empty] | 118565695 |  | 1439474 |  |  | Investigation |  | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Brain Perfusion | [5 empty] | Refresh |  |  | Upgrade | System | Console-SW | NoEnoughInfo | Console-SW-NoEnoughInfo | Console-SW-NoEnoughInfo | NoEnoughInfo | NoEnoughInfo | Console-SW | Console-SW-NoEnoughInfo | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1505 |  |  | United States | AdventHealth Zephyrhills | 300099 | 44488 |  | Black ring artifat on images; Performed air calibraiton to resolve the issue. | Integration : Confirmation of order 54320397 saved ----------------------------------------------------------------------------- Work Order #: WO-04578554 20-Oct-2021 Sher Mark The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: Ring artifact User Impact: None Patient Impact: None Current Software Version: Diagnostic performed by Engineer : customer cleared issue wiyh aircals. ran bad detectors and fee test.. checked system ----------------------------------------------------------------------------- Work Order #: WO-04577520 Boyd Craig The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: Ring artifact User Impact: None Patient Impact: None Current Software Version: Diagnostic performed by Engineer : Black ring artifact. Air calibrations performed. PRS accessed. Physics Analysis shows M26 bad. Possible contrast on window. Follow Up Required : FSE needed on site to replace detector module 26. Recommendations: Clean gantry window. Check for anything in the field of view. Perform air calibrations. Perform Test scans. Analyze images. Run bad detector test and FEE Detector Tests. Replace M26 if needed. ----------------------------------------------------------------------------- | 3 | [14 empty] | 118559312 |  | 1433008 |  |  | Closed - Done | [20 empty] | non-Refresh |  |  | Upgrade | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1506 |  |  | Netherlands | Stichting Rijnstate Ziekenhuis locatie Arnhem | 52101 | 44488 |  | Additional CT imaging due to occurring interference. | "Submitter name: Frank Nieuwenhuijse Submitter email: frank.nieuwenhuijse@philips.com Other contact information: Tel.: 0648109620 Best method of contact: e-mail Form submit date: October 19 2021 Attachments: 1 Reason for delay: There were 2 days delay between the occurrence of the event and reporting to philips. This report was made to an email address (philips employee works in-house in Rijnstate) of the Rijnstate organization. This email was only available on October 18. (this email is only available in the Rijnstate building) Response required for submitter/customer: No Additional notes: No Safety Impact: Other details: The patient received an extra dose of X-ray radiation (272.5 mGy) Feedback type: Technical Issue Equipment ID: ---- -------------------------------------------------- ----------------- Problem definition: A malfunction has been reported with the CT scanner. This malfunction has made it necessary to repeat part of the scan protocol. This has added extra radiation. ng dose delivered to patient. This patient also received an extra dose of intravenous contrast agent. Description of incident The patient was referred to radiology for CT examination. Indication AAA size and indications for endoleak. An adapted version of scan protocol 4.2 was protocolled by the radiologist for this (in which, in addition to the abdomen, the thorax had to be scanned in both acquisitions and a blank CT thorax also had to be made). The patient was scanned on an Ingenuity CT scanner (NZC109, Philips, 2018) in Zevenaar. While performing bolus tracking for correct contrast timing, the speed of reconstructions visible on scanner was so low that an additional delay occurred causing the arterial phase acquisition to start too late. MBB aborted protocol after this scan so that the 2nd acquisition was cancelled. After this consultation took place with the intervention radiologist and the examination with regard to the contrast series was repeated. What now appears to be facts: - The patient has received an extra dose. From dose report in PACS it can be concluded that the total dose of this study yields a dose length product of 840.6 mGy.cm. Of this, 272.5 mGy.cm is due to additional CT imaging due to interference occurring. - The cause of the bolus tracking problem appears to be a stagnant reconstruction process while performing bolus track. Event/occurence date: October 11, 2021 Philips aware date: October 13, 2021 Ability to obtain the information: October 18, 2021" |  |  | Under Engineering investigation, need raise defect to tracking | [14 empty] | 1431481 |  |  | Investigation |  | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Bolus Tracking | [5 empty] | Refresh |  |  | Upgrade | System | Console-SW | NoEnoughInfo | Console-SW-NoEnoughInfo | Console-SW-NoEnoughInfo | NoEnoughInfo | NoEnoughInfo | Console-SW | Console-SW-NoEnoughInfo | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1507 |  |  | United States | Select Medical Corporation | 32261 | 44487 |  | 1205CFXXX Gantry not coming on. | Integration : Confirmation of order 54314443 saved ----------------------------------------------------------------------------- Work Order #: WO-04569963 Morrow, Dan The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Problem description by engineer : System left off over the weekend. Gantry will not turn on. Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: Unable to use Patient Impact/Outcome: none User Impact: Unable to use Current Software Version: 4.1.7 Diagnostic performed by Engineer : Reviewed logs. Rhost not communicating. I walked the customer through a power cycle of the system. After power cycle system is up. Customer able to run STC and test images. OK to close. Internal Comments : Error # and/or description of error : S_RTDH_HC_TIMEOUT_ERROR Repair Action : Reviewed logs. Rhost not communicating. I walked the customer through a power cycle of the system. After power cycle system is up. Customer able to run STC and test images. OK to close. Resolution : Reviewed logs. Rhost not communicating. I walked the customer through a power cycle of the system. After power cycle system is up. Customer able to run STC and test images. OK to close. ----------------------------------------------------------------------------- | 2 | [14 empty] | 118553179 |  | 1428814 |  |  | Closed - No Further Investigation | [20 empty] | Non-Refresh |  |  | Upgrade | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1508 |  |  | United States | Upson Regional Medical Center | 310115 | 44480 |  | Gantry error, replaced DMS fan and heaters, GPU, cleaned slip ring | Integration : Confirmation of order 54284159 saved ----------------------------------------------------------------------------- Work Order #: WO-04526142 12-Oct-2021 Lee John The system meets the specification for the performed service and is returned to use Resolution : Replaced dms heater fan assemblies, gpu in CRC, and cleaned the slip rings. Problem Description by Engineer : Customer Function/Role: CTTech Device Use: Diag Expected and Unexpected Results: Gantry Error User Impact: None Patient Impact: None Current Software Version: IDK Parts Consumed : 1.00000,459800550141,BR64T RoHS HEATER REPLACEMENT KIT 1.00000,459800550141,BR64T RoHS HEATER REPLACEMENT KIT 1.00000,459800980091,BR64 Air Fan Replacement Kit 1.00000,459801622381,NVidia P2000 GPU Card ----------------------------------------------------------------------------- Work Order #: WO-04531921 Sheffield, Scott The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Internal Comments : Error # and/or description of error : 2021/10/12 09:15:52.000__010 CirsEvt_GPUMonitor_177 CIRS_Error Pre Series GPUMonitor177GPUMonitor::RunConfigTests() FAILED (device=0), error details: CudaDevice::CheckConfig failed: GPU current PCIe link width is 8 which is less than the maximum value of 16, which may result in slower performance. Failed device info: Name: Quadro P2000 Serial: 1560420600473 PCI bus ID: 0000:03:00.0 2021/10/12 09:15:57.000__010 CirsEvt_ProxyServer_90 CIRS_Error Series Prep ProxyServer90Exception occurred (type: class Cirs::BasicExceptStream<class Cirs::CirsExcept>) at SeriesProcess.cpp(1442): Error #12 - CheckGantryConnection: ACQ board or Gantry-to-CIRS datapath failure. CIRS Event Log Checked 2021/10/12 09:15:57.961__002 S_DMC_SHOT_COMM_COLLAPSE ERROR_NORMAL Series Exit CStateStudyExit Logger.mdb Checked Type : Error Date : 10/12/2021 Time : 9:33:45 AM Event : 0 Source : Supervisor Category : Error User : N/A Computer : CIRSPHI-27FF54D Description: ImageStudyElmt(342): Error #5 - #61658D43:000 (ACQ #3696) - Failed to start properly 2021/10/12 09:33:46.036 S_DMC_SHOT_COMM_COLLAPSE ERROR_NORMAL Series Exit CStateStudyExit Logger.mdb Checked 2021/10/12 09:43:56.000__001 CirsEvt_GPUMonitor_177 CIRS_Error Series Exit GPUMonitor177GPUMonitor::RunConfigTests() FAILED (device=0), error details: CudaDevice::CheckConfig failed: GPU current PCIe link width is 8 which is less than the maximum value of 16, which may result in slower performance. Failed device info: Name: Quadro P2000 2021/10/12 09:43:56.000__079 CirsEvt_ProxyServer_90 CIRS_Error Series Exit ProxyServer90Exception occurred (type: class Cirs::BasicExceptStream<class Cirs::Os::SEExcept>) at OsSETranslator.cpp(215): Error #12 - ACCESS_VIOLATION occurred at 0x00007FF787C341B1 attempting to read from address 0x0000000000000010 Repair Action : I called FSE John and spoke to him about the GPU errors. John is about to be onsite to work on system. John said that he saw GPU errors in the bug report he reviewed yesterday and he ordered a GPU board already. Diagnostic performed by Engineer : Katy said that they had a power outage this morning and now system is down. They have rebooted a couple times. I was able to log into cirs with remote desktop and troubleshoot problem. Problem Description by Engineer : Problem description by engineer : cirs can not initialize Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: cirs can not initialize Patient Impact/Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer Current Software Version: Not provided Resolution : FSE John is going to resolve the problem for onsite wo WO-04526142. ----------------------------------------------------------------------------- Work Order #: WO-04526121 Carpenter, Kevin The system does not meet the specification for the performed service and is not returned to use Problem Description by Engineer : Customer Function/Role: CTTech Device Use: Diag Expected and Unexpected Results: Gantry Error User Impact: None Patient Impact: None Current Software Version: IDK Diagnostic performed by Engineer : CT RSE remotely collected error logs and reviewed same. Added details and possible CHA parts correlations to dispatch follow-up notes for FSE. Follow Up Required : 2021/10/11 10:28:45.733 S_MPROC_DEVICE_NO_DATA_FROM_GMP_SERIAL P1=5 P2=5 xMany 10:28:45.217 S_MP_CTRL_NO_MORE_AT_SPEED P1=8000 MP-Ctrl No-More-AT-Speed. IP: 8045 10:28:44.328 S_RMC_INVALID_POSITION 10:28:44.286__08 S_CTRL_LINK_GHOST_GOT_CRC_FRAMING_ERROR Last gantry power cycle ~ 14 days ago CHA not working at this time ~ no specific parts recommendations. FSE consult NSS as needed. ----------------------------------------------------------------------------- | 3 | [14 empty] | 118517645 |  | 1406037 |  |  | Closed - Done | [20 empty] | Non-Refresh |  |  | Upgrade | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1509 |  |  | United States | AdventHealth Zephyrhills | 300099 | 44478 |  | Estop not closing; Replaced couch relay and tilt actuators, performed tilt calibrations. | Integration : Confirmation of order 54278587 saved ----------------------------------------------------------------------------- Work Order #: WO-04518522 13-Oct-2021 Garcia Marrero, Jorge The system meets the specification for the performed service and is returned to use Diagnostic performed by Engineer : Installation of left and right Linear Actuator on gantry legs and calibration of gantry Tilt. System is fully operational. Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: E-stop not closing User Impact: None Patient Impact: None Current Software Version: Parts Consumed : 1.00000,453566502541,LINEAR ACTUATOR SKF LH 1.00000,453567114552,ACTUATOR RH SKF 1.00000,453566496941,RELAY,24VDC PLUGIN W/COVER ----------------------------------------------------------------------------- Work Order #: WO-04533050 12-Oct-2021 Watson Adam The system meets the specification for the performed service and is returned to use Resolution : Replaced gantry tilt actuators, performed tilt calibrations, and backed up NVRAM. Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: ----------------------------------------------------------------------------- Work Order #: WO-04518316 09-Oct-2021 Watson Adam The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Diagnostic performed by Engineer : Troubleshot system and found the right side tilt actuator has lifted off of its switches causing estop to open when the customer tilted the gantry. Engineer engaged the actuator switches to move the gantry to the 0 degree position and instructed the customers not to tilt the gantry until he new actuators arrive. Also found horizontal motor relay open errors when the gantry would attempt to scan and move the table. Replaced the horizontal motor relay. System has been turned back over to the customer for clinical use. Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: E-stop not closing User Impact: None Patient Impact: None Current Software Version: Follow Up Required : Replace both tilt actuators ----------------------------------------------------------------------------- Work Order #: WO-04518271 Boyd Craig The system does not meet the specification for the performed service and is not returned to use Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: E-stop not closing User Impact: None Patient Impact: None Current Software Version: Diagnostic performed by Engineer : Patient vomited in gantry. When gantry tilted the E-stop tripped. System power cycled. All E-stops cycled. E-stop not closing. Unable to retrieve error logs. Follow Up Required : PO WC001-0001651531. Customer requesting service today. FSE needed on site to evaluate E-stop circuit and error logs. ----------------------------------------------------------------------------- | 2 | [14 empty] | 118511359 |  | 1402522 |  |  | Closed - Done | [20 empty] | Non-Refresh |  |  | Upgrade | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1510 |  |  | United States | Upson Regional Medical Center | 333009 | 44475 |  | 4.1.10 EC Manager password expired | Integration : Confirmation of order 54274523 saved ----------------------------------------------------------------------------- Work Order #: WO-04512872 Hirchak, Michele The system meets the specification for the performed service and is returned to use Problem description by Clinical Support Specialist : Customer Function/Role: TECH Device USE: IN USE Expected and Unexpected Results: Reset EC Manager password. User Impact: NONE Patient Impact: NONE Current Software Version: 4.1.10 No rescan, no re-injection, no patient harm. Diagnostic performed by Engineer : Customer unable to get in EC Manager. Message tells her password has expired. I told her their IT Admin should be able to reset it. However, their IT Admin is out of the office. This is not an emergency, she will let me know if she wants the FSE to come out to reset. MBH Resolution : Customer was in contact with FSE. John Lee will be going there for another issue, he will reset the passwords when he is onsite. Spoke to John to confirm. MBH Problem description by Clinical Support Specialist : Called to f/u. No answer, no voicemail.MBH ----------------------------------------------------------------------------- | 5 |  | Reset password according Refresh policy | [12 empty] | 118496068 |  | 1396246 |  |  | Closed - No Further Investigation | [7 empty] | Customer | [6 empty] | Password | [5 empty] | Refresh |  |  | Upgrade | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1511 |  |  | United States | Upson Regional Medical Center | 310115 | 44474 |  | Temperature control error. | duplicate case to 0118517645 | 5 | [14 empty] | 118489673 |  | 1391704 |  |  | Investigation | [20 empty] | Non-Refresh |  |  | Upgrade | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1512 |  |  | United States | Philips Medical Capital | 300057 | 44474 |  | Image artifact, inreased Idose level and centrered patient | Work Order #: WO-04489209 Curry, Dyan The system meets the specification for the performed service and is returned to use Problem description by Clinical Support Specialist : Customer Function/ Role: technologist Device Use: in use name Expected and Unexpected Results: artifact questions User Impact: none Patient Impact: none Current Software version: 4.1 Diagnostic performed by Engineer : FSE emailed images to me. MAS is low and patient is off centered. Advised FSE to increase Idose level and center patient. check MAS profile before scanning | 4 | [14 empty] | 118488817 |  | 1386699 |  |  | Closed - Done | [20 empty] | Non-Refresh |  |  | Upgrade | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1513 |  |  | United States | Philips Medical Capital | 300057 | 44473 |  | 0999CDN83 Star Artifact on Thins | Work Order #: WO-04496328 08-Oct-2021 Williams John The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: CTTech Device Use: Diag Expected and Unexpected Results: Star Artifact User Impact: None Patient Impact: None Current Software Version: IDK Resolution : Review images with technologist . Passed on information applications gave me. ----------------------------------------------------------------------------- Work Order #: WO-04482035 05-Oct-2021 Williams John The system meets the specification for the performed service and is returned to use Follow Up Required : Observe and convey applications recommendations regarding image quality. Will work with apps while onsite. Problem Description by Engineer : Customer Function/Role: CTTech Device Use: Diag Expected and Unexpected Results: Star Artifact User Impact: None Patient Impact: None Current Software Version: IDK Diagnostic performed by Engineer : Download and send images to applications and NSS. Work via phone with applications. ----------------------------------------------------------------------------- Work Order #: WO-04481960 Carpenter, Kevin The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Follow Up Required : Star Artifact on Thins ~ FSE please review artifact images onsite. RSE asked cust to copy artifact to CDR for easy download ~ but images not found there. RSE also had cust _ lock _ Pt exam contain artifact ~ see series with _ thins_ Problem Description by Engineer : Customer Function/Role: CTTech Device Use: Diag Expected and Unexpected Results: Star Artifact User Impact: None Patient Impact: None Current Software Version: IDK Diagnostic performed by Engineer : Star Artifact on Thins RSE asked cust to copy artifact to CDR for easy download ~ but images not found there. RSE also had cust _ lock _ Pt exam contain artifact ~ see series with _ thins_ ----------------------------------------------------------------------------- | 3 | [14 empty] | 118483434 |  | 1383083 |  |  | Investigation | [20 empty] | Non-Refresh |  |  | Upgrade | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1514 |  |  | Spain | POLICLINICA NTRA SRA DEL ROSARIO SL | 336190 | 44473 |  | The UPS has failed. | Work Order #: WO-04486742 05-Oct-2021 Parts Consumed : 1.00000,459800306392,EXTERNAL BATT PACK, 3KVA UPS, 72VDC ----------------------------------------------------------------------------- Work Order #: WO-04477685 05-Oct-2021 [REDACTED] The system meets the specification for the performed service and is returned to use. Resolution: Battery is requested. Problem Description by Engineer: The UPS is damaged. Regulatory Questions : N/A Follow Up Required: Order battery to send directly to the customer. ----------------------------------------------------------------------------- Work Order #: WO-04477048 [REDACTED] The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled. Diagnostic performed by Engineer: I talk to [REDACTED]. The UPS has problems with the battery, the equipment is in bypass mode, need to replace the battery. Problem Description by Engineer: The UPS is damaged. Regulatory Questions : N/A Follow Up Required: Replace the battery. ----------------------------------------------------------------------------- | 3 | [14 empty] | 118479908 |  | 1381591 |  |  | Closed - No Further Investigation | [20 empty] | Non-Refresh |  |  | Upgrade | [9 empty] | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | #N/A
1515 |  |  | Australia | Maia Financial Pty Limited | 341000 | 44529 | 48 | Recording Voice Messages Assistance | Work Order #: WO-04866036 Steve, Teixeira Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 5 |  | Application Support | [12 empty] | 118802178 | In Progress | 1568366 |  |  | Opened | Opened | [6 empty] | Application | [6 empty] | Application Support | [5 empty] | Refresh |  |  | System issue | bar: No; No
1516 |  |  | Australia | Maia Financial Pty Limited | 341000 | 44525 | 47 | Surview scan stopped <CT> | Work Order #: WO-04849511 Steve, Teixeira The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : CUSTOMER REPORTED PROBLEM OR EVENT DESCRIPTION: 26 Nov 2021 The Customer called to report on behalf of a colleague that during a Chest Abdomen study the surview scan stopped 3/4 of the way through. Spoke to the Radiographer scanning at the time and informed that the first surview stopped 3/4 of the way through and produced a message asking if the User wanted to proceed to the 2nd surview. On the 2nd atempt the scan appeared to complete but no image appeared. Whilst on the phone, on the 3rd attempt the User was able to complete a dual durview and the next scans successfully. Study completed. Bug report generated. ________________________________________________________________ Diagnostic performed by Engineer : CUSTOMER REPORTED PROBLEM OR EVENT DESCRIPTION: 26 Nov 2021 The Customer called to report on behalf of a colleague that during a Chest Abdomen study the surview scan stopped 3/4 of the way through. Spoke to the Radiographer scanning at the time and was informed that the first surview stopped 3/4 of the way through and produced a message asking if the User wanted to proceed to the 2nd surview. On the 2nd attempt, the scan appeared to complete but no image appeared. Whilst on the phone, on the 3rd attempt the User was able to complete a dual durview and the next scans successfully. Study completed. Bug report generated. SOFTWARE REVISION OF THE SYSTEM: Ver. 4.1.10 WAS A BUG REPORT SAVED? Yes. LOG REVIEW: A review of the logs confirms that during the 2 attempts the surview scans stopped (at 10:09:05 AM and 10:11:19 AM 26 Nov 2021) due to arcing (anode and cathode). TROUBLESHOOTING: 26 Nov 2021 Assisted the Customer with running Filament calibration and starting the long tube conditioning. CUSTOMER UPDATE: 26 Nov 2021 Sent an email update to the Customer summarising the events. Email extract reads: " HI XXXXX, As discussed today the surview scans stopped due to arcing. ZZZZ successfully completed the Filament calibration and started a long tube conditioning which I believe is complete by now. ZZZZ (hoping I have the spelling correct) wanted to know where to find the tube heat status. The CT Application Specialist YYYY confirmed that It is found under Utilities (Spanner icon) -> Show tube heat. Generally we want the tube heat to be around 10%, and if the system is idle for extended periods, a short tube conditioning prior to the study is good practice. Given the arcing today please be mindful of the tube heat and running a short tube conditioning when the system is idle for more than 2 hours, or when called out to do after hour scans. This is purely a recommendation as appreciate that different Users are involved and there are other work demands involved. I will keep the case open until next week to ensure that there are no further signs of arcing, as we may need to run another long tube conditioning should the arcing persist. As a gentle reminder please restart the Host computer daily and the CIRS Server/s weekly, although not related to the arcing it does minimize the chance of lockups, performance issues etc. Please feel free to contact me should you have any further queries in this regard. ..." FOLLOW UP: 2 Dec 2021 Called the Customer and confirmed that no further reports of arcing-related messages or scans stopping prematurely. ________________________________________________________________ Internal Comments : LOG REVIEW: *********** ________________________________________________________________ Filament cal last done 16 Oct 2020. ________________________________________________________________ LOGGER: * 2021/11/26 10:19:34.226__01 ERROR_BLOB CANMS2 ------ Series 24 Load Structure ----- 2021/11/26 10:19:34.137__01 ERROR_BLOB CANMS2 ------ Series 23 Load Structure ----- 2021/11/26 10:19:34.058__05 ERROR_BLOB CANMS2 ------ Series 22 Load Structure ----- 2021/11/26 10:19:00.245 EVENT_MAJOR GSE ----------------- Series 21 Complete ----------------- 2021/11/26 10:18:58.655__01 EVENT_MAJOR GSE - - - - - - Shot 0 Starts - - - - - - 2021/11/26 10:18:46.763__05 ERROR_BLOB CANMS2 ------ Series 21 Load Structure ----- 2021/11/26 10:17:22.136 EVENT_MAJOR GSE ----------------- Series 20 Complete ----------------- 2021/11/26 10:17:13.604__01 EVENT_MAJOR GSE - - - - - - Shot 0 Starts - - - - - - 2021/11/26 10:16:53.503 EVENT_MAJOR GSE ----------------- Series 19 Complete ----------------- 2021/11/26 10:16:43.973__01 EVENT_MAJOR GSE - - - - - - Shot 0 Starts - - - - - - 2021/11/26 10:16:01.178__01 ERROR_BLOB CANMS2 ------ Series 20 Load Structure ----- 2021/11/26 10:16:01.113__09 ERROR_BLOB CANMS2 ------ Series 19 Load Structure ----- 2021/11/26 10:15:49.450__02 EVENT_MINOR GSE GantryStartStudyBkst was sent to Gantry 2021/11/26 10:15:49.450 EVENT_MINOR GSE GSE::CState::StateBkst::STUDY_STARTED was sent to GPC 2021/11/26 10:15:30.288 EVENT_MINOR GSE Series 18 cancelled, due to GANTRY_CANCEL_SOURCE_HOST, in state CStateSeriesReady 2021/11/26 10:15:30.269 EVENT_MINOR GSE GSE::CState::StateBkst::STUDY_ENDED was sent to GPC 2021/11/26 10:15:19.183 ERROR_WARNING Gantry CStateSeriesReady 0 ERROR 165295 WARN S_MASTER_SEQUENCER_SERIES_DELAY_LESS_THAN_MIN 0 4123 2021/11/26 10:15:12.994 ERROR_BLOB CANMS2 ------ Series 18 Load Structure ----- 2021/11/26 10:14:36.903 EVENT_MINOR GSE Series 17 cancelled, due to GANTRY_CANCEL_SOURCE_CTBOX, in state CStateSeriesReady 2021/11/26 10:14:34.271 ERROR_WARNING Gantry CStateSeriesReady 0 ERROR 165250 WARN S_MASTER_SEQUENCER_SERIES_DELAY_LESS_THAN_MIN 0 4123 2021/11/26 10:14:13.695__06 ERROR_BLOB CANMS2 ------ Series 17 Load Structure ----- 2021/11/26 10:11:19.982 ERROR_WARNING Gantry CStateStudyExit 1 ERROR 165056 WARN S_GEN_MGR_SYNC_OFF_OUT_OF_SEQUENCE 0 0 2021/11/26 10:11:19.976 ERROR_WARNING Gantry CStateStudyExit 1 ERROR 165056 WARN S_GEN_FAULT_CATHODE_TUBE_ARC 5 33 5:5 CPM Arc 2021/11/26 10:11:19.934 ERROR_WARNING Gantry CStateStudyExit 1 ERROR 165056 WARN S_GEN_FAULT_ANODE_TUBE_ARC 5 33 5:0 APM Arc 2021/11/26 10:11:19.932 ERROR_WARNING Gantry CStateStudyExit 1 ERROR 165056 WARN S_GEN_FAULT_ANALOG_FAULT 2 32 2:5 Analog Fault 2021/11/26 10:11:19.887__15 ERROR_NORMAL Gantry CStateStudyExit 0 ERROR 165056 NORMAL S_MASTER_SEQUENCER_NO_XRAY 0 1 2021/11/26 10:11:19.885__03 EVENT_MINOR GSE CStateShot:Series 15 cancelled, due to GANTRY_CANCEL_SOURCE_NORMAL_ERROR, in state CStateShot 2021/11/26 10:11:18.102__01 EVENT_MAJOR GSE - - - - - - Shot 0 Starts - - - - - - 2021/11/26 10:10:58.646__01 ERROR_BLOB CANMS2 ------ Series 16 Load Structure ----- 2021/11/26 10:10:58.576__09 ERROR_BLOB CANMS2 ------ Series 15 Load Structure ----- 2021/11/26 10:10:11.632 EVENT_MINOR GSE Series 14 cancelled, due to GANTRY_CANCEL_SOURCE_CTBOX, in state CStateSeriesPrep 2021/11/26 10:09:35.927__04 ERROR_BLOB CANMS2 ------ Series 14 Load Structure ----- 2021/11/26 10:09:19.417 EVENT_MINOR GSE Series 13 cancelled, due to GANTRY_CANCEL_SOURCE_CTBOX, in state CStateSeriesPrep 2021/11/26 10:09:15.693__06 ERROR_BLOB CANMS2 ------ Series 13 Load Structure ----- 2021/11/26 10:09:07.087__01 ERROR_NORMAL Gantry CStateStudyExit 0 ERROR 164923 NORMAL S_MASTER_SEQUENCER_EXIT_STATE_ERR 4096 8369 2021/11/26 10:09:05.071 ERROR_WARNING Gantry CStateStudyExit 1 ERROR 164921 WARN S_GEN_MGR_SYNC_OFF_OUT_OF_SEQUENCE 0 0 2021/11/26 10:09:05.068 ERROR_WARNING Gantry CStateStudyExit 1 ERROR 164921 WARN S_GEN_FAULT_CATHODE_TUBE_ARC 5 33 5:5 CPM Arc 2021/11/26 10:09:05.032__06 ERROR_WARNING Gantry CStateStudyExit 1 ERROR 164921 WARN S_GEN_FAULT_ANODE_TUBE_ARC 5 33 5:0 APM Arc 2021/11/26 10:09:05.029 EVENT_MINOR GSE CStateShot:Series 11 cancelled, due to GANTRY_CANCEL_SOURCE_NORMAL_ERROR, in state CStateShot 2021/11/26 10:09:05.027 ERROR_WARNING Gantry CStateShot 1 ERROR 164921 WARN S_GEN_FAULT_ANALOG_FAULT 2 32 2:5 Analog Fault 2021/11/26 10:09:05.021 ERROR_NORMAL Gantry CStateShot 1 ERROR 164921 NORMAL S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT 0 0 2021/11/26 10:09:00.335 EVENT_MAJOR GSE - - - - - - Shot 0 Starts - - - - - - 2021/11/26 10:08:38.732 ERROR_BLOB CANMS2 ------ Series 12 Load Structure ----- 2021/11/26 10:08:38.651__07 ERROR_BLOB CANMS2 ------ Series 11 Load Structure ----- 2021/11/26 10:08:27.554__02 EVENT_MINOR GSE GantryStartStudyBkst was sent to Gantry 2021/11/26 10:08:27.554 EVENT_MINOR GSE GSE::CState::StateBkst::STUDY_STARTED was sent to GPC * ________________________________________________________________ Resolution : CASE CLOSED: As the system is operational after running the relevant calibrations and monitoring, as agreed with the Customer, no further service-related action will be taken on this case. This issue will continue to be monitored via Customer calls and alerts over time, to identify trends if any, that may warrant further action. ________________________________________________________________ ----------------------------------------------------------------------------- | 3 |  | As the system is operational after running the relevant calibrations and monitoring, as agreed with the Customer, no further service-related action will be taken on this case. This issue will continue to be monitored via Customer calls and alerts over time, to identify trends if any, that may warrant further action. | [12 empty] | 118788208 | Incomplete | 1558982 |  |  | Investigation | Investigation | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Scan error | [5 empty] | Refresh |  |  | System issue | System | Console-SW | NoEnoughInfo | Console-SW-NoEnoughInfo | Console-SW-NoEnoughInfo | NoEnoughInfo | NoEnoughInfo | bar: No; No
1517 |  |  | Australia | Murwillumbah Hospital | 341000 | 44525 | 47 | Phantom is hard to use when performing IQ test on an Ingenuity system; |  | #N/A |  | Phantom is hard to use when performing IQ test on an Ingenuity system; | [14 empty] | 1554916 |  |  | Investigation | Investigation | H-Phantom | [5 empty] | Engineering | [6 empty] | H-Phantom RFC | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | H-Phantom | Couch-HW-H-Phantom | Couch-HW-H-Phantom | H-Phantom | H-Phantom | bar: No; No
1518 |  |  | China | 江苏大学附属医院 | 345023 | 44529 | 48 | wechat：02N/Ingenuity Core128/Elite China/Deleting Network Node | BU Q&R Feedback: No SAP Customer Symptom Code: HANG Customer Symptom Code: Hang/Crash Reason: Corrective Maintenance Engineer Failure Code: Configuration Error Repair Activity Code: Configuration Change Problem Reported by Customer : wechat：02N/Ingenuity Core128/Elite China/ Safety Questions : Case: 0118802918 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Bo Last updated by: Wang Bo 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户需要删除节点。 Troubleshooting Action: 远程协助客户删除添加的节点。 Repair Action: 远程协助客户删除添加的节点。机器正常。 Test/Inspection: customer confirm can scan patient and image ok.functional Check PASSED. Device returned to full functionality. Test Equipment: N/A ----------------------------------------------------------------------------- | 3 | [14 empty] | 118802918 | Fixed | 1568804 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [19 empty] | Refresh |  |  | System issue | bar: No; No
1519 |  |  | Chile | BANCO DE CHILE | 341005 | 44523 | 47 | CL1688 EL GANTRY SE APAGA | Work Order #: WO-04840371 26-Nov-2021 [REDACTED] The system meets the specification for the performed service and is returned to use. Diagnostic performed by Engineer: Equipment stops once with errors: S_ROTOR_MCO S_RMC_Error_MCO Problem Description by Engineer : Problem description by engineer : Gantry shuts off Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected Results: No gantry errors. Patient Impact/Outcome: Delay in studies. User Impact: Delay in studies. Current Software Version: 4.1.10 Resolution: Optical and electrical slip ring inspection is done, firmware is loaded to GMP. ----------------------------------------------------------------------------- Work Order #: WO-04829286 [REDACTED] The system meets the specification for the performed service and is returned to use. Internal Comments : Error # and/or description of error : S_MPROC_DEVICE_NO_DATA_FROM_GMP_SERIAL S_MPS_WAITING_FOR_MAIN_DRIVE_TO_CLOSE Repair Action: - Check GMP connections and cables. - Reconfigure GMP. Diagnostic performed by Engineer: The gantry turns off. In the error log we observed an interruption in the communication of the gantry and an error in a controller. Gantry check-up is needed. Problem Description by Engineer : Problem description by engineer : Gantry shuts off Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected Results: No gantry errors. Patient Impact/Outcome: Delay in studies. User Impact: Delay in studies. Current Software Version: 4.1.10 Follow Up Required: Error Number and/or Description: S_MPROC_DEVICE_NO_DATA_FROM_GMP_SERIAL S_MPS_WAITING_FOR_MAIN_DRIVE_TO_CLOSE Repair Action: - Check GMP connections and cables. - Reconfigure GMP. Technician_s Name: [REDACTED] ----------------------------------------------------------------------------- | 3 |  | - Check GMP connections and cables. - Reconfigure GMP | [12 empty] | 118772061 | Closed | 1553335 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | TBD-GMP | [5 empty] | Engineering | [6 empty] | GMP | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | GMP | Gantry-HW-GMP | Gantry-HW-GMP | TBD-GMP | GMP | Gantry-HW | Gantry-HW-GMP | bar: No; No
1520 |  |  | China | Medical center (contagious disease, He nan, Xu chang) | 40047 | 44507 | 44 | 02N/Ingenuity CT 728326/host have no free disk space | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Calibration/Adjustment Problem Reported by Customer: 02N/Ingenuity CT 728326 Safety Questions:Case: 0118675828 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Gaoge Last updated by: Wang Gaoge 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer_s problems: failure related to computer software. Detailed description of Customer_s problems or malfunctions: Host system failed to start It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: The host computer failed to start software Remote troubleshoot and test result: There is not enough storage space for the host PRS TR status: It is not through, ask engineer for maintenance on site. Repair Action: Suggest onsite service procedures: 1. re-format and re-install; 2. delete redundant data Test/Inspection: N/A Test Equipment: N/A SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support ----------------------------------------------------------------------------- | 2 |  | $Q7: (Confirmed reported problem ,low disk space ) $Q8:( low disk space ,software can not run ) Q9 (delete log files and other no used files , then issue solved !) | [12 empty] | 118675828 | In Process | 1498106 |  |  | Evaluation - Pending Translation | Evaluation - Pending Translation | CTUDT00632261 | [5 empty] | Engineering | [6 empty] | Disk Full | [5 empty] | Refresh |  |  | System issue | System | Console-SW | DiskFull | Console-SW-DiskFull | Console-SW-DiskFull | CTUDT00632261 | DiskFull | Console-SW | Console-SW-DiskFull | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1521 |  |  | Chile | BANCO DE CHILE | 341005 | 44515 | 46 | CL1688 ERROR EN EL ESTUDIO POST-CARDIACO | Work Order #: WO-04766671 [REDACTED] The system meets the specification for the performed service and is returned to use. Problem Description by Engineer: Customer Function/Role: Technician Device Use: Diagnostic Expected and Unexpected Results: No error messages. User Impact: Delay in studies. Patient Impact: Delay in studies. Current Software Version: 11.1.3 Internal Comments: RSW not available Diagnostic performed by Engineer: A warning error is displayed for access to reserved programs. After accepting the message, they can continue working. The equipment software needs to be reviewed. Follow Up Required: Configuration check-up and software check-up. Evaluate possible software reloading. ----------------------------------------------------------------------------- | 4 |  | CL1688 ERROR IN POST-CARDIAC STUDY I get a warning error for accessing reserved programs. After accepting the message, you can continue working. Equipment software revision is required. | [12 empty] | 118718057 | In Progress | 1526933 |  |  | Evaluation | Evaluation | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | SW error | [5 empty] | Refresh |  |  | System issue | System | Console-SW | NoEnoughInfo | Console-SW-NoEnoughInfo | Console-SW-NoEnoughInfo | NoEnoughInfo | NoEnoughInfo | bar: No; No
1522 |  |  | China | Sichuan Province Yibin City No.2 T.C.M. Hospital | 345035 | 44530 | 48 | 02N/Ingenuity Core128/Elite China/couch noise | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 02N/Ingenuity Core128/Elite China Safety Questions:Case: 0118811414 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang T2 Activities:Problem Description: Summary of customer_s problems: Other types of error. Detailed description of Customer_s problems or malfunctions: noise with the bed It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: noise with the bed Remote troubleshoot and test result: Not sure about the specifics. Need to go to the field to check PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Check to see if It is caused by the shell, then add lube and try SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support ----------------------------------------------------------------------------- | 3 |  | noise with the bed $Q7: (Confirmation the report problem.Couch noise when the couch up from low position.) $Q8: (Couch noise when the couch up from low position.check and found the cover was rub.adjust and clean the cover,lubricate the cover.) $Q9: (Couch noise when the couch up from low position.check and found the cover was rub.adjust and clean the cover,lubricate the cover.) | [12 empty] | 118811414 | Closed | 1573362 |  |  | Opened - Pending Translation | Opened - Pending Translation | E-125059 | [5 empty] | Engineering | [6 empty] | Couch Noise | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | Lubricant | Couch-HW-Lubricant | Couch-HW-Lubricant | E-125059 | Lubricant | Couch-HW | Couch-HW-Lubricant | bar: No; No
1523 |  |  | China | Lindian County Hospital | 345038 | 44529 | 48 | 02N/Ingenuity Core128/Elite China/Yellow screen | T2 Activities:Problem Description: Summary of customer_s problems: The host has a yellow screen. Detailed description of Customer_s problems or malfunctions: The host has a yellow screen. It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: N/A Remote troubleshoot and test result: N/A PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: It is recommended to do a workaround upgrade. We recommend spare parts 12NC and names: N/A Memo of Customer request: None Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support ----------------------------------------------------------------------------- | 3 |  | $Q7: (Confirmed reported problem & host down) $Q8: (host down & host down) $Q9: (restart host) $Q10: (functional test pass. & Device returned to full functionality) | [12 empty] | 118802673 | Closed | 1568658 |  |  | Investigation - Pending Translation | Investigation - Pending Translation | CTUDT00632520 | [5 empty] | Engineering | [6 empty] | Yellow Screen | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow Screen | Console-SW | Console-SW-Yellow-Screen | bar: No; No
1524 |  |  | China | 扎兰屯市人民医院 | 40045 | 44529 | 48 | 02N/Ingenuity CT 728326/Ist authorization | BU Q&R Feedback: No SAP Customer Symptom Code: HANG Customer Symptom Code: Hang/Crash Reason: Corrective Maintenance Engineer Failure Code: Configuration Error Repair Activity Code: Configuration Change Problem Reported by Customer : 2N/Ingenuity CT 728326/ Next Step Action : ASP yiyuanji 015604188333 call in T2 Activities : Problem Description: 保养工程师需要IST授权。 Troubleshooting Action: 远程协助保养工程师进行授权。 Repair Action: 远程协助保养工程师进行授权。授权失败，PRS报错连接失败。无法进行授权。 Test/Inspection: customer confirm can scan patient and image ok.functional Check PASSED. Device returned to full functionality. Test Equipment: N/A Safety Questions : Case: 0118802526 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Bo Last updated by: Wang Bo 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 3 |  | 保养工程师需要IST授权。 | [12 empty] | 118802526 | Fixed | 1568572 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Service | [6 empty] | IST Push | [5 empty] | Refresh |  |  | System issue | bar: No; No
1525 |  |  | China | Shaanxi Province Xianyang Central Hospital (Xianyang City No. 2 People_s Hospital) | 40050 | 44529 | 48 | 02N/Ingenuity CT 728326/gantry noise when rotating | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Replace Part Safety Questions:Case: 0118801707 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Gaoge Last updated by: Wang Gaoge 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer_s problems: The device makes a strange sound. Detailed description of Customer_s problems or malfunctions: Abnormal sound can be heard when the gantry rotates. It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: An abnormal sound occurs on the tensioner. Remote troubleshoot and test result: Consider that the belt is causing the sound or that the tensioner is causing the sound. PRS TR status: No PRS. Repair Action: Suggest onsite service procedures: 1. Replace the belt or tensioner. We recommend spare parts 12NC and names: "453567003451 TENSIONER ASSY 453566492401 POLY-VEE BELT" Test/Inspection: N/A Test Equipment: N/A Next Step Action: PPD 453566492401*1 453567003451*1 all have been shipped. field service engineer: Guo Yujie - 86-15158180650 24H Problem Reported by Customer: 02N/Ingenuity CT 728326 SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support ----------------------------------------------------------------------------- | 3 |  | $Q7:confirmed report problem,ct device has noise when gantry rotation, customer can not use device $Q8:review error logs ,not found any error ,to check gantry rotor motor,belt and belt Tensioner , found is Tensioner has noise, belt is normal, so replace the new tensioner, gantry rotation not has noise, device runing normal, so think customer observation and use. $Q9:I replace the new tensioner, gantry rotation not has noise, device return normal., $Q10:Test and inspection data refer to PA ，Device returned to full functionality 453567003451 TENSIONER ASSY | [12 empty] | 118801707 | Closed | 1568109 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | ECR-077012 | [5 empty] | Engineering | [6 empty] | TENSIONER ASSY | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Belt | Gantry-HW-Belt | Gantry-HW-Belt | ECR-077012 | Belt | Gantry-HW | Gantry-HW-Belt | bar: No; No
1526 |  |  | China | Inner Mongolia Alxa Right Banner People_s Hospital | 345031 | 44529 | 48 | 02N/Ingenuity Core128/Elite China/power on failed | BU Q&R Feedback: No SAP Customer Symptom Code: BOOT Customer Symptom Code: Boot up/Power-up Problem Reason: Corrective Maintenance Engineer Failure Code: Incompatible Use Repair Activity Code: Instructed Customer Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ T2 Activities:Problem Description: power on failed Troubleshooting Action: customer call service, system can not power on. instruct customer check power on, customer not power on teal transform, instruct customer power on it, then power on UPS, system is ok. Repair Action: nothing to do, close case. Test/Inspection: Functional Check PASSED. The device returned to full functionality. Test Equipment: N/A SAP Transfer:Case Transferred to SAP Safety Questions:Case: 0118800659 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 3 |  | customer call service, system can not power on. instruct customer check power on, customer not power on teal transform, instruct customer power on it, then power on UPS, system is ok. | [12 empty] | 118800659 | Closed | 1567498 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Others | [6 empty] | System Reboot | [5 empty] | Refresh |  |  | System issue | bar: No; No
1527 |  |  | China | Jinan City No. 4 People_s Hospital | 40038 | 44529 | 48 | Intermittent power failure, closed keyswitch and rebooted CT gantry | BU Q&R Feedback: No Reason: Corrective Maintenance T2 Activities:Problem Description: Summary of customer_s problems: The gantry intermittently loses power. It is a technology problem with the equipment itself: Yes. Troubleshooting Action: remote diagnosis and testing result: Engineer Chen Tao is reporting repair for the customer. It is very difficult to determine the problem PRS TR status: It is not through, ask engineer for maintenance on site. Repair Action: Suggest onsite service procedures: Suggest engineer to go to the hospital to check the error log and determine the problem. Teach the customer to use Philips WeChat to report for repair Safety Questions:Case: 0118800493 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Weng Jinming Last updated by: Weng Jinming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Problem Reported by Customer:Intermittent power failure SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support ----------------------------------------------------------------------------- | 3 |  | $Q7: (Confirmed reported problem, E-stop open intermittently) $Q8: (Close keyswitch and it was normal) $Q9: ( Reboot CT gantry) | [12 empty] | 118800493 | Fixed | 1567319 |  |  | Closed - Done | Closed - Done | [6 empty] | Others | [6 empty] | System Reboot | [5 empty] | Refresh |  |  | System issue | bar: No
1528 |  |  | China | Xuancheng City Orthopedic Hospital | 345042 | 44528 | 47 | WeChat：02N/Ingenuity Core128/Elite China/instal patch | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Calibration/Adjustment Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ Safety Questions:Case: 0118799473 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Cao Last updated by: Liu Cao 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer_s problems: running stuck Detailed description of Customer_s problems or malfunctions: running stuck It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: Need to install patch Remote troubleshoot and test result: Install patch PRS TR status: NA Repair Action: Suggest onsite service procedures: Install patch SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Next Step Action:Zhou Bing ----------------------------------------------------------------------------- | 5 |  | $Q7: (Confirmed reported problem,CT host software crash） $Q8: ( check the system logger，found the software issue ) $Q9: (reinstall the host software , trouble disppeared.) | [12 empty] | 118799473 | Closed | 1565023 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | incorrect installation | [5 empty] | Others | [6 empty] | SW Crash | [5 empty] | Refresh |  |  | System issue |  |  |  | - | - | incorrect installation | 0 | bar: No
1529 |  |  | China | Medical center (contagious disease, He nan, Xu chang) | 40047 | 44528 | 47 | 02N/Ingenuity CT 728326/ | BU Q&R Feedback: No Reason: Corrective Maintenance Safety Questions:Case: 0118799206 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Guoping Last updated by: Liu Guoping 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer_s problems: ESTOP disconnected. Detailed description of Customer_s problems or malfunctions: Twice Estop disconnected It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: Twice Estop disconnected Remote troubleshoot and test result: It is normal after re-keying PRS TR status: No PRS. Repair Action: Suggest onsite service procedures: Check the logger and confirm the problem. Problem Reported by Customer: 02N/Ingenuity CT 728326 Next Step Action: HUANG YONG SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support ----------------------------------------------------------------------------- | 3 |  | $Q7: (Confirm reported problem, estop open) $Q8: (On site checked the logger ,found relay open error， need to replace the relay) $Q9: (Replace the relay ) | [12 empty] | 118799206 | Closed | 1564625 |  |  | Investigation - Pending Translation | Investigation - Pending Translation | [6 empty] | Engineering | [6 empty] | Relay | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | Relay | Couch-HW-Relay | Couch-HW-Relay | NotEngineering-PartReplacement | Relay | Couch-HW | Couch-HW-Relay | bar: No
1530 |  |  | China | 济南市第四人民医院 | 40038 | 44526 | 47 | 02N/Ingenuity CT 728326/add node | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: Instructed Customer Problem Reported by Customer : 02N/Ingenuity CT 728326/ Safety Questions : Case: 0118789101 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wennian Han Last updated by: Wennian Han 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: add node Troubleshooting Action: add workstation node Repair Action: guide customer to add workstation node, feedback transmission is normal Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: NA ----------------------------------------------------------------------------- | 5 | [14 empty] | 118789101 | Fixed | 1562265 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [19 empty] | Refresh |  |  | System issue | bar: No
1531 |  |  | China | Zhejiang Province Yueqing No. 1 People’s Hospital | 40054 | 44525 | 47 | Application/Ingenuity CT 728326 | BU Q&R Feedback: No Reason: Application Support Safety Questions:Case: 0118782324 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingsong Last updated by: Lu Mingsong 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Problem Reported by Customer: 02N/Ingenuity CT 728326 ----------------------------------------------------------------------------- | 5 | [14 empty] | 118782324 | In Process | 1559229 |  |  | Evaluation | Evaluation | [6 empty] | Application | [6 empty] | Application Support | [5 empty] | Refresh |  |  | System issue | bar: No
1532 |  |  | China | Jiangsu University Affiliated Hospital | 345023 | 44506 | 44 | 02N/Ingenuity Core128/Elite China/Can not boot | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : Can not boot SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Safety Questions:Case: 0118675308 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Zhao Shixin Last updated by: Zhao Shixin 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A External Remarks: 07.11.2021 13:50:04 UTC (86036528)"SUAT. Ready for Clinical Use local time" : "2021-11-07 10:45" External Remarks: 07.11.2021 14:18:05 UTC (86036528) Service details: Device fail to boot up normally. Onsite find CIRS is not turned on. Re do switch on and off procedure and train the operator on machine start up procedure, then restore. Service Result: The device returned to normal. Internal Remarks : 07.11.2021 14:21:15 UTC (86036528) RDI Accuracy: Yes $Q7: (Confirm reported problem, the gantry can not power on.) $Q8: (check the console found CIRS did not turned on.) $Q9:Turn on the CIRS. $Q10: (Test and inspection data refer to PA. Device returned to full functionality.) $Q11: Calibrated tools refer to PA. ----------------------------------------------------------------------------- | 2 |  | (86036528) Service details: Device fail to boot up normally. Onsite find CIRS is not turned on. Re do switch on and off procedure and train the operator on machine start up procedure, then restore. | [12 empty] | 118675308 | New | 1497474 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Others | [6 empty] | CIRS Power on | [5 empty] | Refresh |  |  | System issue | bar: No
1533 |  |  | China | Heilongjiang Province Huachuan County People_s Hospital | 345018 | 44524 | 47 | Application/Core128 | BU Q&R Feedback: No Reason: Application Support Problem Reported by Customer : Application Safety Questions:Case: 0118775224 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Pan Chuchu Last updated by: Pan Chuchu Next Step Action: Customer request second clinical training. Yang Jun 13817647801 yang.yj.jun@philips.com Guo Fan 13501285722 fan.guo@philips.com, follow up. ----------------------------------------------------------------------------- | 5 |  | Application Training 客户申请第二次临床培训， 杨俊 13817647801 yang.yj.jun@philips.com 郭凡 13501285722 fan.guo@philips.com 跟进 | [12 empty] | 118775224 | New | 1555527 |  |  | NAC | Evaluation | [6 empty] | Application | [6 empty] | APP Training | [5 empty] | Refresh |  |  | System issue | bar: No
1534 |  |  | China | Jiangsu Xuzhou Medical School Affiliated No. 3 Hospital | 40039 | 44523 | 47 | 02G/Ingenuity CT 728326/Adapter failure | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer: 02G/Ingenuity CT 728326/Adapter failure Safety Questions:Case: 0118772508 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: shao jianan Last updated by: shao jianan 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 3 |  | 02G/Ingenuity CT 728326/Adapter failure | [12 empty] | 118772508 | Closed | 1553677 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632173 | [5 empty] | Engineering | [6 empty] | Adapter | [5 empty] | Refresh |  |  | System issue | System | Console-HW | Adaptor | Console-HW-Adaptor | Console-HW-Adaptor | CTUDT00632173 | Adaptor | Console-HW | Console-HW-Adaptor | bar: No
1535 |  |  | China | Jiangsu Province No. 2 Chinese Medicine Hospital | 345027 | 44523 | 47 | 02N/Ingenuity Core128/Elite China/tube overheating | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ Next Step Action: Caller: Dr Fu: 013914749786 T2 Activities:Problem Description: Summary of customer_s problems: Other types of error. Detailed description of Customer_s problems or malfunctions: The customer said the tube is frequently running in overheat protection. It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: Log has no error. Oil cooling is normal, around 40 degrees all day. Remote troubleshoot and test result: Restart but didn_t work. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: It is recommended to check if the customer set up too many slices in scan protocol. Verify if the protocol dose setting is too high in mAs or scan time is too long causing the tube heat capacity to reach the limit. We recommend spare parts 12NC and names: N/A Safety Questions:Case: 0118763700 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Bo Last updated by: Wang Bo 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Next Step Action:Qi liangdong ----------------------------------------------------------------------------- | 3 |  | $Q7: (Confirmed reported problem: Heat capacity warnimg sometimes.) $Q8: (Checked the error log, the temperature of oil flow is normal. Checked scan protocol, found mA value of one scan protocol is too higher.) $Q9: (Adjusted the scan parameter, keep monitoring.) | [12 empty] | 118763700 | In Process | 1552109 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Customer | [6 empty] | Tube Overheat | [5 empty] | Refresh |  |  | System issue | bar: No
1536 |  |  | China | Henan Province Puyang City People’s Hospital | 40031 | 44523 | 47 | 02N/Ingenuity CT 728326/The tube overheating | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Safety Questions:Case: 0118762961 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingsong Last updated by: Lu Mingsong 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: The tube overheating Troubleshooting Action: The room temperature is too high. Advise customer to lower room temperature. Repair Action: The room temperature is too high.Advise customer to lower room temperature. Test/Inspection: Functional Check PASSED. The device returned to full functionality. Test Equipment: N/A Problem Reported by Customer: 02N/Ingenuity CT 728326 SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | Troubleshooting Action: The room temperature is too high. Advise customer to lower room temperature. Repair Action: The room temperature is too high.Advise customer to lower room temperature. Test/Inspection: Functional Check PASSED. The device returned to full functionality. | [12 empty] | 118762961 | Closed | 1551612 |  |  | Investigation | Investigation | [6 empty] | Customer | [6 empty] | Tube Overheat | [5 empty] | Refresh |  |  | System issue | bar: No
1537 |  |  | China | Inner Mongolia Autonomous Region IV Hospital | 40023 | 44521 | 46 | 05K for IB/Ingenuity CT 728326/password expired | BU Q&R Feedback: No SAP Customer Symptom Code: MARF Customer Symptom Code: Monitor Alarm Failure Reason: Corrective Maintenance Engineer Failure Code: Incompatible Use Repair Activity Code: Instructed Customer Problem Reported by Customer: 05K for IB/Ingenuity CT 728326/ Field Service Engineer Notes:This PO is a Reference Number only Safety Questions:Case: 0118754238 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Password expired Troubleshooting Action: instruct customer change password, close case. Repair Action: instruct customer change password, close case. Test/Inspection: Functional Check PASSED. The device returned to full functionality. Test Equipment: N/A ----------------------------------------------------------------------------- | 3 |  | Repair Action: instruct customer change password, close case. | [12 empty] | 118754238 | Fixed | 1544053 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Others | [6 empty] | Password | [5 empty] | Refresh |  |  | System issue | bar: No
1538 |  |  | China | Jilin Province Zhaonan City TCM Hospital | 345006 | 44521 | 46 | 05L/Ingenuity Core128/Elite China/artifact | BU Q&R Feedback: No Reason: Corrective Maintenance Safety Questions:Case: 0118754020 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer_s problems: Other types of error. Detailed description of Customer_s problems or malfunctions: Artifact It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: Ring artifact. Remote troubleshoot and test result: Detector channel failure PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Analyze data, find artifact location, mask channel SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support External Remarks: 22.11.2021 00:48:45 UTC (86039473) Service Details: There are artifacts in the images. Work on the bad detector and seal the bad channel. service result: device returns to normal working condition after service Problem Reported by Customer : 05L/Ingenuity Core128/Elite China/ Internal Remarks : 22.11.2021 00:48:48 UTC (86039473) RDI Accuracy: Yes $Q7:(Confirmed reported problem & the image has ring artifacts) $Q8: (Do the bad detector and analyze data,have bad channel) $Q9: (Shielding bad channels,cperform relative test,system is ok) $Q10: (Test and inspection data refer to: PA. Device returned to full functionality.) $Q11: Calibrated Tools name: refer to PA Next Step Action: Liu Yadong ----------------------------------------------------------------------------- | 3 |  | $Q7:(Confirmed reported problem & the image has ring artifacts) $Q8: (Do the bad detector and analyze data,have bad channel) $Q9: (Shielding bad channels,cperform relative test,system is ok) | [12 empty] | 118754020 | Closed | 1543949 |  |  | Investigation | Investigation | CTUDT00632110 | [5 empty] | Engineering | [6 empty] | Ring Artifact | [5 empty] | Refresh |  |  | System issue | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | DMS Module | DMS-HW | DMS-HW-DMS Module | bar: No
1539 |  |  | China | Heilongjiang Province Huachuan County People_s Hospital | 345018 | 44524 | 47 | 05K For IB/Core128/Consult warranty | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer:Consult warranty Safety Questions:Case: 0118775223 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Pan Chuchu Last updated by: Pan Chuchu 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action: Case#0118775223 Heilongjiang Province Huachuan County People_s Hospital, Ingenuity Core128 EQ#87552930, customer consult warranty expiration date Teacher Zhang 15246304333 Next Step Action: Screenshot of warranty period has been sent to Dark Blue Trading Ltd. Co to call back customer to inform Next Step Action: Case# 0118775224 follow up, Yang, Jun <yang.yj.jun@philips.com> feedback: have arranged two trainings. I asked the trainer to explain to the hospital. Next Step Action: Case# 0118776997 followed up, Yang, Jun <yang.yj.jun@philips.com> already arranged for trainer to contact DEALER, follow up with 2nd follow up visit ----------------------------------------------------------------------------- | 5 | [14 empty] | 118775223 | Closed | 1555529 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [19 empty] | Refresh |  |  | System issue | bar: No
1540 |  |  | China | Jiangxi Province Children’s Hospital | 40041 | 44520 | 46 | 02N/Ingenuity CT 728326/rotation noise | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 02N/Ingenuity CT 728326 SAP Transfer:Case Transferred to SAP T2 Activities:Problem Description: Summary of customer_s problems: The device makes a strange sound. Detailed description of Customer_s problems or malfunctions: NOISE IN ROTATION It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: NA Remote troubleshoot and test result: The customer informed the rack scan when rotating has noise, not when stationary, contacted the engineer, suspected tension wheel problems PRS TR status: It is not through, ask engineer for maintenance on site. Repair Action: Suggest onsite service procedures: Check the tension wheel, belt-related mechanical structure We recommend spare parts 12NC and names: " 453567003451 TENSIONER ASSY: 453566502101 SGL MTR SPINDLE ALIGN TOOL" Next Step Action: PPD 453566502101 1 & 453567003451 1 scheduled 72H to hospital, FSE Li Jiawei -86-18579112310 Safety Questions:Case: 0118752935 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created by: Wennian Han Last updated by: Wennian Han 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Dispatched for On site Support ----------------------------------------------------------------------------- | 3 |  | 453566502101 1 & 453567003451 $Q7:(Confirmed reported problem，gantry noise) $Q8:( Check rotation system, found the noise from tension pulley ) $Q9:( Replace the tension pulley solved the problem.) | [12 empty] | 118752935 | Closed | 1543530 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | ECR-077012 | [5 empty] | Engineering | [6 empty] | Tensioner | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Belt | Gantry-HW-Belt | Gantry-HW-Belt | ECR-077012 | Belt | Gantry-HW | Gantry-HW-Belt | bar: No
1541 |  |  | China | Jiangsu Xuzhou Medical School Affiliated No. 3 Hospital | 40039 | 44520 | 46 | 02G/Ingenuity CT 728326/Gantry nosie | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 01/Ingenuity CT 728326 Safety Questions:Case: 0118752470 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created by: Wennian Han Last updated by: Wennian Han 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Gantry noise Troubleshooting Action: The gantry is noisy when rotating. Repair Action: Check for scraping: the belt, tension pulley, etc. ----------------------------------------------------------------------------- | 3 |  | T2 Activities:Problem Description: Gantry noise Troubleshooting Action: The gantry is noisy when rotating. Repair Action: Check for scraping: the belt, tension pulley, etc. | [12 empty] | 118752470 | In Process | 1543427 |  |  | Evaluation | Evaluation | ECR-077012 | [5 empty] | Engineering | [6 empty] | Tensioner | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Belt | Gantry-HW-Belt | Gantry-HW-Belt | ECR-077012 | Belt | Gantry-HW | Gantry-HW-Belt | bar: No
1542 |  |  | China | Heilongjiang Province Huachuan County People_s Hospital | 345018 | 44519 | 46 | Cannot hear operator voice in Gantry room, | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Replace Part SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support External Remarks: 21.11.2021 07:09:55 UTC (86039442) Service details: check ct-box cable, adjust sound volume Service Result: The device returned to normal. Internal Remarks : 21.11.2021 07:09:57 UTC (86039442) "RDI Accuracy:Yes" $Q7: (Confirmed reported problem: The horn doesn_t sound ) $Q8: (check cable ok) $Q9:(adjust voice) $Q10: (Test and inspection data refer to: PA. The device returned to full functionality. $Q11 Calibrated Tools name (no) Problem Reported by Customer: The horn doesn_t sound Next Step Action: PPD451213440782*1 have arranged for 24h delivery, please dispatch FSE: Jin Jun-86-13796075213 Safety Questions:Case: 0118746499 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Gaoge Last updated by: Wang Gaoge 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer_s problems: Other types of error. Detailed description of Customer_s problems or malfunctions: There was no sound in the rack when speaking It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: Use CT BOX to speak, no sound. Auto audio playback is normal Remote troubleshoot and test result: Consider CT BOX failure. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: 1. check audio playback; 2. replace CT box We recommend spare parts 12NC and names: "459801705022 CT BOX ASSEMBLY_BEST " Test/Inspection: N/A Test Equipment: N/A ----------------------------------------------------------------------------- | 3 |  | $Q7: (Confirmed reported problem&The horn doesn_t sound ) $Q8: (check cable ok） $Q9:(adjust voice) 459801705022 CT BOX ASSEMBLY_BEST | [12 empty] | 118746499 | Closed | 1541375 |  |  | Investigation | Investigation | CTUDT00632185 | [5 empty] | Engineering | [6 empty] | CT Box | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | CT Box | Gantry-HW-CT Box | Gantry-HW-CT Box | CTUDT00632185 | CT Box | Gantry-HW | Gantry-HW-CT Box | No: No; y: 0; y: 0; y: 0; y: 0; y: 0; No: No
1543 |  |  | China | Shandong Province Tai_an City T.C.M. No. 2 Hospital | 345037 | 44519 | 46 | System opens estop intermittently; Adjusted the gantry gullwing switch position to resolve. | BU Q&R Feedback: No Reason: Corrective Maintenance Safety Questions:Case: 0118746354 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingsong Last updated by: Lu Mingsong 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer_s problems: The gantry lost power. Detailed description of Customer_s problems or malfunctions: The gantry lost power. It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: MCO open Remote troubleshoot and test result: N/A PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Adjust Gullwing switch position. We recommend spare parts 12NC and names: N/A Memo of Customer request: None Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ Next Step Action: Customer said to call back Teacher Chen: 18560023293 Next Step Action: power on failure intermittently SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Internal Remarks : 19.11.2021 11:57:21 UTC (86038373) “RDI Accuracy: YES" $Q7: (Confirmed reported problem, close E-stop failed intermittently $Q8: ( Check gantry log, found "MCO open" error. Push gantry front cover and it can be closed normally) $Q9: ( Adjust the position of gullwing) $Q10: (Test and inspection data refer to: PA. Device returned to full functionality) $Q11: Calibrated Tools name ( N/A) ^Current Tube installed date: N/A ^Current Tube usage: N/A External Remarks : 19.11.2021 11:30:32 UTC (86038373) Fault Description: Power on by turning key intermittently failed Service content: Adjust gull wing switch position Service outcome: Device returns to normal working condition after servicing. ----------------------------------------------------------------------------- | 3 |  | $Q7: (Confirmed reported problem, close E-stop failed intermittently $Q8: ( Check gantry log, found "MCO open" error. Push gantry front cover and it can be closed normally) $Q9: ( Adjust the position of gullwing) | [12 empty] | 118746354 | Closed | 1541213 |  |  | Closed - Done | Closed - Done | AssemblyAdjust | [5 empty] | Others | [6 empty] | Gullwing | [5 empty] | Refresh |  |  | System issue |  |  |  | - | - | AssemblyAdjust | 0 | bar: No
1544 |  |  | China | Guangxi Province Teng County People’s Hospital | 40036 | 44518 | 46 | 01/Ingenuity CT 728326/The image is not very clear | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : The image is not very clear Next Step Action: OB Zhangyuanyuan emailed to inform that Teacher Tan_s feedback during the follow-up visit is as follows: 1. After the scan is finished, there is no one-button Couch In operation. 2. The diagnosing doctor said the images are not very clear and a bit coarse; they are not as bright as Siemens. 3. We need a clinical teacher to come out and train them again. There was no medical case when I came. I still don_t understand it after saying that. Please help with the feedback, and please ask the online engineer to check if it is possible to instruct the customer remotely.Thanks! Next Step Action: I called and texted CT Application Specialist Deng Mingming at 18929200533 to follow up and reply to the customer. Next Step Action: Regarding the third point, I emailed ASP YingLian to follow up with Business. Safety Questions:Case: 0118739639 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Zhou Xiaohong Last updated by: Zhou Xiaohong 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action: Deng Mingming (18929200533) said that the customer rejected Teacher Deng_s callback to handle it remotely and requested to have a clinical teacher on site. Teacher Deng has confirmed with the training manager, the site has had on-site training twice and there is no need to go on site again. Next Step Action: ASP YingLian replied: This device is still within the Philips warranty period. Already followed up by Manager Lin Ruixin. T2 Activities:Problem Description: Summary of customer_s problems: Other types of error. Detailed description of Customer_s problems or malfunctions: "During the follow-up visit, Teacher Tan said the problem is as follows: 1. After the scan is finished, there is no one-button Couch In operation. 2. The diagnosing doctor said the images are not very clear and a bit coarse; they are not as bright as Siemens. 3. We need a clinical teacher to come out and train them again. There was no medical case when I came. I still don_t understand it after saying that." It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: Follow-up feedback Remote troubleshoot and test result: "1. Inform the customer that for the CT box from the last century, the one Couch In button doesn_t exist, but on both sides of the gantry there is one Couch In button. 2. Regarding the feedback_s second item, transfer to operations, the application engineer will call back to process it. 3. For the third item, please transfer to Business to follow up." PRS TR status: It is not through, ask engineer for maintenance on site. ----------------------------------------------------------------------------- | 3 | [14 empty] | 118739639 | In Process | 1538261 |  |  | Evaluation | Evaluation | [6 empty] | Application | [6 empty] | Application Support | [5 empty] | Refresh |  |  | System issue | bar: No
1545 |  |  | Spain | Comercial Centro Medico | 40042 | 44502 | 44 | Error de reconstructor | Work Order #: WO-04686646 05-Nov-2021 [REDACTED] The system meets the specification for the performed service and is returned to use. Problem Description by Engineer: Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Regulatory Questions : N/A Diagnostic performed by Engineer: Reconstructor error message is reproduced. It stops sending images during reconstruction. Resolution: After checking the error log of the different components, I do not find an element in the hardware that could be causing the error, so I re-installed the software on both servers of the reconstruction system. After reloading the software, I check the correct operation of the system. The system works correctly and is ready for use. ----------------------------------------------------------------------------- Work Order #: WO-04683286 02-Nov-2021 [REDACTED] The system meets the specification for the performed service and is returned to use. Regulatory Questions : N/A Problem Description by Engineer: Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Follow Up Required: Check operation in S2 IMR Diagnostic performed by Engineer: When launching a patient study, the error "restart reconstruction system" is displayed. The S2 reconstructor of the IMR crashes. I reset it and works correctly. After performing a patient it crashes again. There is no network connection from the HOST. It is restarted and fails again. I do not find any evident error in the error logs of the system or the reconstructors. The monitor and keyboard are connected, awaiting for the error to reappear, but it does not. The system is tested several times and does not fail. ----------------------------------------------------------------------------- | 2 |  | Diagnostic performed by Engineer: When launching a patient study, the error "restart reconstruction system" is displayed. The S2 reconstructor of the IMR crashes. I reset it and works correctly. After performing a patient it crashes again. There is no network connection from the HOST. It is restarted and fails again. I do not find any evident error in the error logs of the system or the reconstructors. The monitor and keyboard are connected, awaiting for the error to reappear, but it does not. The system is tested several times and does not fail. | [12 empty] | 118650957 | Closed | 1480937 |  |  | Investigation | Investigation | part replace | [5 empty] | Others | Net Switch | [5 empty] | Net Switch | [5 empty] | Refresh |  |  | System issue |  |  |  | - | - | part replace | 0 | bar: No
1546 |  |  | China | 浙江省乐清第一人民医院 | 40054 | 44517 | 46 | wechat：02N/Ingenuity CT 728326/software running slowly | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : wechat：02N/Ingenuity CT 728326 Safety Questions : Case: 0118730781 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: software running slowly Troubleshooting Action: software running slowly ,after customer add menory for test, speed will change fast. so customer request upgrade memory. nemo it . Repair Action: nemo memory . Test/Inspection: N/A Test Equipment: N/A SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | software running slowly ,after customer add menory for test, speed will change fast. so customer request upgrade memory. nemo it . | [12 empty] | 118730781 | In Process | 1534771 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | MemoryReplacement | [5 empty] | Engineering | [6 empty] | Memory | [5 empty] | Refresh |  |  | System issue | System | Console-HW | MemoryReplacement | Console-HW-MemoryReplacement | Console-HW-MemoryReplacement | MemoryReplacement | MemoryReplacement | Console-HW | Console-HW-MemoryReplacement | bar: No
1547 |  |  | China | Nanjing Medical University Affiliated Yifu Hospital | 345028 | 44516 | 46 | 02N/Ingenuity Core128/Elite China/couch noise | BU Q&R Feedback: No Reason: Corrective Maintenance Safety Questions:Case: 0118720842 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingsong Last updated by: Lu Mingsong 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer_s problems: The vertical movement of the couch is noisy. Detailed description of Customer_s problems or malfunctions: The feedback: sounded like friction It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: N/A Remote troubleshoot and test result: N/A PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Check the screw and horizontal glide on the site. If necessary add lube. We recommend spare parts 12NC and names: N/A Memo of Customer request: None SAP Transfer:Dispatched for On site Support SAP Transfer:Case Transferred to SAP Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ Next Step Action:YAO Liu Bin ----------------------------------------------------------------------------- | 3 |  | $Q7: (Confirm reported problem. Patient table has noise when moves vertically.) $Q8: (Reviewed the error log and found the ball screw and slide rails are leak of oil.) $Q9: (Lubricate the ball screw and slide rails.) | [12 empty] | 118720842 | In Process | 1528788 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | E-125059 | [5 empty] | Engineering | [6 empty] | Lubrication | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | Lubricant | Couch-HW-Lubricant | Couch-HW-Lubricant | E-125059 | Lubricant | Couch-HW | Couch-HW-Lubricant | bar: No
1548 |  |  | China | Heilongjiang Province Huachuan County People_s Hospital | 345018 | 44522 | 47 | praise/Ingenuity Core128/Elite China/engineer | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer: praise Safety Questions:Case: 0118755415 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Qian Wang Last updated by: Qian Wang 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action: Praise engineer Jin Jun, each time when report for repair, he arrive on site timely, and can solve the problem. he is very professional with good attitude. Email jingwen to follow up ----------------------------------------------------------------------------- | 3 | [14 empty] | 118755415 | Closed | 1546758 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [19 empty] | Refresh |  |  | System issue | bar: No
1549 |  |  | China | Jiangxi Province Yugan Chu East Hospital (New Branch) | 337500 | 44516 | 46 | Wechat:02N/Ingenuity Core128/Elite China/Unable to log in after restarting the system | BU Q&R Feedback: No SAP Customer Symptom Code: BOOT Customer Symptom Code: Boot up/Power-up Problem Reason: Corrective Maintenance Engineer Failure Code: Electrical/Electronic Failure Repair Activity Code: Reboot/Restart Problem Reported by Customer: Unable to log in after restarting the system Safety Questions:Case: 0118720501 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Weng Jinming Last updated by: Weng Jinming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Troubleshooting Action: instruct customer to restart the system Repair Action: instruct customer to restart the system Test/Inspection: Functional Check PASSED. The device returned to full functionality. Test Equipment: N/A ----------------------------------------------------------------------------- | 3 |  | Troubleshooting Action: instruct customer to restart the system Repair Action: instruct customer to restart the system | [12 empty] | 118720501 | Fixed | 1528562 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632553 | [5 empty] | Engineering | [6 empty] | System Reboot | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Blue-Screen | Console-SW-Blue-Screen | Console-SW-Blue-Screen |  | Blue-Screen | Console-SW | Console-SW-Blue-Screen | bar: No
1550 |  |  | China | Heilongjiang Province Huachuan County People_s Hospital | 345018 | 44515 | 46 | 05K FOR IB/Ingenuity Core128/Elite China/can not burn | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 05K FOR IB/Ingenuity Core128/Elite China Next Step Action: Wait for customer to confirm host model in the afternoon Safety Questions:Case: 0118712859 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created by: Wennian Han Last updated by: Wennian Han 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | 05K FOR IB/Ingenuity Core128/Elite China/can not burn | [12 empty] | 118712859 | Fixed | 1523118 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Service | [6 empty] | Burn | [5 empty] | Refresh |  |  | System issue |  |  |  | - | - | 0 | 0 | bar: No
1551 |  |  | China | Guangxi Province Teng County People’s Hospital | 40036 | 44515 | 46 | 02N/Ingenuity CT/has prodigious sound | BU Q&R Feedback: No Reason: Corrective Maintenance SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Next Step Action: Manual Case#PCC2020111411 Caller: CT Teacher Tan. 15878407675 Hospital: Guangxi Teng County People_s Hospital Device model: Ingenuity CT 728326 EQ# 87185445 SN#40036 Fault Description: Q1&Q2: Follow up contact: CT Teacher Tan. 15878407675 Warranty info: 02N SITE1:FSE Huang Yao pei 18128388600 Dispatch FSE/ASP: FSE Huang Yao pei 18128388600 RSE Wang Jianqiang: Customer issues: After the maintenance, ding dong noise is often heard. Detailed description of Customer_s problems or malfunctions: After the maintenance, ding dong noise is often heard. It is a technology problem with the equipment itself: Yes. Key error code or symptom: After the maintenance, ding dong noise is often heard. Remote troubleshoot and test result: Suspect a cable is poorly in contact. Check cause of the abnormal sound. The cable needs to be plugin tight. Also the cable from the back of the rack to the monitor is not plugged in securely. PRS TR status: It is not through, ask engineer for maintenance on site. Suggest to go to the site to repair: Check the wiring and unplug and re plug tightly. Problem Reported by Customer: Has prodigious sound Safety Questions:Case: 0118712610 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Pan Chuchu Last updated by: Pan Chuchu 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 3 |  | $Q6: (Gantry noise and ECG signal lost) $Q7: (Confirm reported problem & Found ibox ECG signal problem.) $Q8: (Found ibox convert box connected problem & Need to adjust ibox cable.) $Q9: (Readjust ibox cable.) | [12 empty] | 118712610 | New | 1522889 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | IBOX | [5 empty] | Engineering | [6 empty] | IBOX | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | IBOX | Gantry-HW-IBOX | Gantry-HW-IBOX | IBOX | IBOX | Gantry-HW | Gantry-HW-IBOX | bar: No
1552 |  |  | China | Jiangsu University Affiliated Hospital | 345023 | 44511 | 45 | 02N/Ingenuity Core128/Elite China/spooler reset | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: No Fault Found: Unconfirmed Problem / Failure Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ T2 Activities:Problem Description: Need reset spooler Troubleshooting Action: Need reset spooler Repair Action: Remotely run script Safety Questions:Case: 0118697950 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Cao Last updated by: Liu Cao 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 3 |  | Troubleshooting Action: Need reset spooler Repair Action: Remotely run script | [12 empty] | 118697950 | Fixed | 1513961 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Reset Spooler | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | Relay | Couch-HW-Relay | Couch-HW-Relay | NotEngineering-PartReplacement | Relay | Couch-HW | Couch-HW-Relay | bar: No
1553 |  |  | China | Jiangsu University Affiliated Hospital | 345023 | 44511 | 45 | Wechat:02N/Ingenuity Core128/Elite China/Unable to transfer | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: Instructed Customer Problem Reported by Customer: Unable to transfer Safety Questions:Case: 0118697789 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Bai Jingwen Last updated by: Bai Jingwen 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities: Problem Description: Unable to transfer Troubleshooting Action: Unable to transfer image Repair Action: reboot HOST computer, then transmission is normal Test/Inspection: Functional Check PASSED. The device returned to full functionality. Test Equipment: NA SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | Troubleshooting Action: Unable to transfer image Repair Action: reboot HOST computer, then transmission is normal | [12 empty] | 118697789 | Fixed | 1513926 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | Image Transfer | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | bar: No
1554 |  |  | China | Inner Mongolia Alxa Right Banner People_s Hospital | 345031 | 44511 | 45 | 02N/Ingenuity Core128/Elite China/estop close failed | BU Q&R Feedback: No SAP Customer Symptom Code: BOOT Customer Symptom Code: Boot up/Power-up Problem Reason: Corrective Maintenance Engineer Failure Code: Electrical/Electronic Failure Repair Activity Code: Instructed Customer Problem Reported by Customer :Ingenuity Core128/Elite China Safety Questions:Case: 0118697661 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: estop close failed Troubleshooting Action: check errorlogger, no error found, instruct customer check emergency button, and close estop, system is ok. close case. Repair Action: nothing to do, close case. Test/Inspection: N/A Test Equipment: N/A ----------------------------------------------------------------------------- | 3 |  | T2 Activities:Problem Description: estop close failed Troubleshooting Action: check errorlogger, no error found, instruct customer check emergency button, and close estop, system is ok. close case. Repair Action: nothing to do, close case. | [12 empty] | 118697661 | Fixed | 1513877 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Others | [6 empty] | System Reboot | [5 empty] | Refresh |  |  | System issue | bar: No
1555 |  |  | China | Caofeidian District Hospital | 345016 | 44510 | 45 | 05K For IB/Ingenuity Core128/Elite China/Unable to burn | BU Q&R Feedback: No Reason: Corrective Maintenance T2 Activities:Problem Description: Summary of customer_s problems: Other types of error. Detailed description of Customer_s problems or malfunctions: Customer problem It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: NA Remote troubleshoot and test result: The customer refused to transfer and did not answer the phone.Contacted the engineer to tell him that there is a disc burning recognition issue. The customer needs on site assistance. PRS TR status: No PRS. Repair Action: Suggest onsite service procedures: Confirm if the customer_s disc burning procedure is correct; check Rom drive, software; rule out disc issue Safety Questions:Case: 0118692235 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created by: Wennian Han Last updated by: Wennian Han 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Internal Remarks : 11.11.2021 15:07:20 UTC (86038892) "RDI Accuracy:Yes" $Q7: Confirm reported problems:system need to change configuration. $Q8: Trouble shooting steps:change the film configuration. $Q9: Function returns to normal after the of configuration is changed. $Q10: Test and Inspection data refer to PA. Device status:Device returned to full functionally $Q11: N/A Problem Reported by Customer: Unable to burn External Remarks: 11.11.2021 10:22:45 UTC (86038892) service details: machine burn failure. After the process, it is back to normal. service result: device returned to normal working condition after repairing ----------------------------------------------------------------------------- | 3 |  | $Q7: Confirm reported problems:system need to change configuration. $Q8: Trouble shooting steps:change the film configuration. $Q9: Function returns to normal after the of configuration is changed. | [12 empty] | 118692235 | Closed | 1510753 |  |  | Evaluation | Evaluation | [6 empty] | Service | [6 empty] | Burn | [5 empty] | Refresh |  |  | System issue |  |  |  | - | - | 0 | 0 | bar: No
1556 |  |  | China | Jiangsu Xuzhou Medical School Affiliated No. 3 Hospital | 40039 | 44518 | 46 | 02G/Ingenuity CT 728326/add dicom node | BU Q&R Feedback: No SAP Customer Symptom Code: CENV Customer Symptom Code: Customer Environmental Concern Reason: Corrective Maintenance Engineer Failure Code: Environmental Cause Repair Activity Code: Instructed Customer Problem Reported by Customer: 02G/Ingenuity CT 728326 Safety Questions:Case: 0118740003 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Gaoge Last updated by: Wang Gaoge 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Customer need to add a printer Troubleshooting Action: Instructed the customer to add printers. Need to find the device administrator account and password. Repair Action: Instructed the customer to add the printer. Test/Inspection: N/A Test Equipment: N/A ----------------------------------------------------------------------------- | 3 | [14 empty] | 118740003 | Fixed | 1538350 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [19 empty] | Refresh |  |  | System issue | bar: No
1557 |  |  | China | Medical center (contagious disease, He nan, Xu chang) | 40047 | 44510 | 45 | 02N/Ingenuity CT 728326/application hanged up | BU Q&R Feedback: No SAP Customer Symptom Code: HANG Customer Symptom Code: Hang/Crash Reason: Corrective Maintenance Engineer Failure Code: Environmental Cause Repair Activity Code: Instructed Customer SAP Transfer:Case Transferred to SAP Safety Questions:Case: 0118690732 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Gaoge Last updated by: Wang Gaoge 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Problem Reported by Customer: 02N/Ingenuity CT 728326 T2 Activities:Problem Description: Software stuck when boot up, fail to enter interface Troubleshooting Action: Instruct customer to log off current user, log in again, then restore Repair Action: Instruct customer to log in again Test/Inspection: N/A Test Equipment: N/A ----------------------------------------------------------------------------- | 3 |  | T2 Activities:Problem Description: Software stuck when boot up, fail to enter interface Troubleshooting Action: Instruct customer to log off current user, log in again | [12 empty] | 118690732 | Fixed | 1509836 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632553 | [5 empty] | Engineering | [6 empty] | SW Stuck | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Blue-Screen | Console-SW-Blue-Screen | Console-SW-Blue-Screen |  | Blue-Screen | Console-SW | Console-SW-Blue-Screen | bar: No
1558 |  |  | China | Fujian Province Medical University Affiliated No. 1 Hospital | 345012 | 44509 | 45 | 02N/Ingenuity Core128/Elite China/right panel key fault | BU Q&R Feedback: No Reason: Corrective Maintenance T2 Activities:Problem Description: Summary of customer_s problems: Control panel button is not working. Detailed description of Customer_s problems or malfunctions: Right control panel buttons are not good It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: Right control panel buttons are not good Remote troubleshoot and test result: Buttons are not usable. The customer needs to replace it PRS TR status: No PRS. Repair Action: Suggest onsite service procedures: 1. Check the right control panel.2. Replace the right control panel according to the manual We recommend spare parts 12NC and names: "CONTROL PANEL ASSEMBLY-FRONT RH 459800484511" Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ Next Step Action: FSE Li Feng 018650369579 call in Next Step Action: Calling back was successful. Safety Questions:Case: 0118685700 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Jianqiang Last updated by: Wang Jianqiang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Next Step Action : PPD 459800484511 *1, arranged 24H delivery, Field Service Engineer Li Feng - 86-18650369579 ----------------------------------------------------------------------------- | 3 |  | CONTROL PANEL ASSEMBLY-FRONT RH 459800484511 | [12 empty] | 118685700 | In Process | 1506266 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632181 | [5 empty] | Engineering | [6 empty] | GCP | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Gantry Panel | Gantry-HW-Gantry Panel | Gantry-HW-Gantry Panel | CTUDT00632181 | Gantry Panel | Gantry-HW | Gantry-HW-Gantry Panel | No: No; y: 0; y: 0; y: 0; y: 0; y: 0; No: No
1559 |  |  | China | Suzhou Guang ci Tumor Hospital Co., Ltd | 345019 | 44509 | 45 | The images has artifacts. Performed the air calibration again to resolve. | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Calibration/Adjustment Internal Remarks : 11.11.2021 16:01:02 UTC (86090280) RDI Accuracy: Yes $Q7: (Confirm reported problem,the images has artifacts.) $Q8: (Check the air calibration table and found has error.) $Q9: (perform the air calibration according to service manual) $Q10: (Test and inspection data refer to PA. Device returned to full functionality.) $Q11: Calibrated tools refer to PA. SAP Transfer:Case Transferred to SAP Problem Reported by Customer:image fault Safety Questions:Case: 0118684258 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Zixuan Last updated by: Wang Zixuan 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Dispatched for On site Support T2 Activities:Problem Description: Summary of customer_s problems: Image density is nonuniform. Detailed description of Customer_s problems or malfunctions: present in all parts, air calibration has been done this morning It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: The compensator tool did not prompt for cracks Remote troubleshoot and test result: Confirm raw data and calibrate PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Confirm raw data and calibrate Next Step Action: Xu Haotian External Remarks : 09.11.2021 11:23:38 UTC (86090280) Service details: There are artifacts to the images. Re-run air calibration and it restore. Service Result: The device returned to normal. ----------------------------------------------------------------------------- | 3 |  | $Q7: (Confirm reported problem,the images has artifacts.) $Q8: (Check the air calibration table and found has error.) $Q9: (perform the air calibration according to service manual) | [12 empty] | 118684258 | In Process | 1505450 |  |  | Closed - Done | Closed - Done | ShouldbeCustomer | [5 empty] | Engineering | [6 empty] | Image Artifact | [5 empty] | Refresh |  |  | System issue | System | Console-SW | ShouldbeCustomer | Console-SW-ShouldbeCustomer | Console-SW-ShouldbeCustomer | ShouldbeCustomer | ShouldbeCustomer | bar: No
1560 |  |  | United States | Westgate Hospital | 40001 | 44518 | 46 | 0699CFXXX customer requesting support - part id - standoff | Internal Remarks : Interaction Type: New Case Caller: Justin Bertino IP & Issue: 82488278 Event Type: Part ID Service Request: Remote Technical Did Customer Decline Remote Support? No, not entitled Preferred Service Time: n/a Entitlement Options: Free Remote Reason for Entitlement Claim: part id Action Taken: created case, wo and routed call to rtech for part id callback Next Step Action: N/A Verbal PO Number: n/a Verbal PO Approver: n/a Verbal PO CAP: n/a ----------------------------------------------------------------------------- Work Order #: WO-04798258 Miller, John M The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Problem description by engineer : Customer needs part number of the stand off for the Footswitch I/F board. When was the first occurrence? Within... : 18-Nov-2021 How often is issue occurring? : First occurrence Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: BioMed Device Use: No patient involved Expected and Unexpected results: None Patient Impact/Outcome: Not provided User Impact: None Current Software Version: N/A Diagnostic performed by Engineer : Collect information from customer and view log files. Internal Comments : Error # and/or description of error : Research part number. Repair Action : Researched parts number of the Stand off for the Foot Switch Interface board. The customer is okay with closing the case. 459800088051 - Standoff-Hex M/F 8-32 X2 Resolution : Researched parts number of the Stand off for the Foot Switch Interface board. The customer is okay with closing the case. 459800088051 - Standoff-Hex M/F 8-32 X2 ----------------------------------------------------------------------------- | 5 | [14 empty] | 118744103 | Incomplete | 1536147 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [19 empty] | Refresh |  |  | System issue | bar: No
1561 |  |  | China | Guangxi Province Teng County People’s Hospital | 40036 | 44507 | 44 | 02N/Ingenuity CT 728326/power off | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer:power off Safety Questions:Case: 0118675996 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Qian Wang Last updated by: Qian Wang 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action: huang yaopei SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support ----------------------------------------------------------------------------- | 3 |  | $Q6: (Power up problem) $Q7: (Confirm reported problem & Found local power problem.) $Q8: (Found power up problem . & Need to check power resource.) $Q9: (Check power input and restart system.) | [12 empty] | 118675996 | New | 1498428 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Others | [6 empty] | System Reboot | [5 empty] | Refresh |  |  | System issue | bar: No
1562 |  |  | United States | Philips Medical Capital | 40028 | 44504 | 44 | 0199CDN84 Recon box error | Work Order #: WO-04699813 10-Nov-2021 Abbas Mohamed Abbas The system meets the specification for the performed service and is returned to use Tools & Equipment : MULTIMETER,2023-05-31 00:00:00 Resolution : Reset all boards for recon. Verify recon booting up. verify acquisition. verify QC Problem Description by Engineer : Problem description by engineer : Recon box error When was the first occurrence? Within... : 04-Nov-2021 How often is issue occurring? : Continuously Malfunction area : Recon Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: Not provided Patient Impact/Outcome: No further information is available as per the customer User Impact: Cannot use system to scan Current Software Version: Not provided ----------------------------------------------------------------------------- Work Order #: WO-04699406 Ali, Adam The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Internal Comments : Error # and/or description of error : RECON box error Repair Action : Check / repair power button on CIRS Run CIRS diagnostics Diagnostic performed by Engineer : Recon box error on the system. Customer power cycled the system, found as soon as the power button on CIRS is released it will turn itself off. Problem Description by Engineer : Problem description by engineer : Recon box error When was the first occurrence? Within... : 04-Nov-2021 How often is issue occurring? : Continuously Malfunction area : Recon Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: Not provided Patient Impact/Outcome: No further information is available as per the customer User Impact: Cannot use system to scan Current Software Version: Not provided Follow Up Required : Error # and/or description of error : RECON box error Repair Action : Check / repair power button on CIRS Run CIRS diagnostics Technician Name : Ali, Adam ----------------------------------------------------------------------------- | 2 |  | Reset all boards for recon. Verify recon booting up. verify acquisition. verify QC | [12 empty] | 118663947 | Incomplete | 1487783 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | ReconBox | [5 empty] | Engineering | [6 empty] | CIRS HW | [5 empty] | Refresh |  |  | System issue | System | Console-HW | ReconBox | Console-HW-ReconBox | Console-HW-ReconBox | ReconBox | ReconBox | bar: No
1563 |  |  | China | Shaanxi Province Ankang T.C.M. Hospital | 345041 | 44507 | 44 | 02N/Ingenuity Core128/Elite China/estop open | BU Q&R Feedback: No Reason: Corrective Maintenance Internal Remarks : 15.11.2021 04:27:45 UTC (86039086) Q7: (Confirmed reported problem. & The Gantry shut down. $Q8: (Check the error log was not found. Check the table & We need to calibrate table) $Q9: (1. calibrate table2. Test use for CT was OK. $Q10: (Test and inspection data refer to: PA. & Device returned to full functionality.) $Q11: NA External Remarks: 15.11.2021 04:27:42 UTC (86039086) Service Reason: The gantry powered off. Service Details: Instructed the customer to calibrate. Result of service: device operating normally Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ T2 Activities:Problem Description: Summary of customer_s problems: ESTOP disconnected. Detailed description of Customer_s problems or malfunctions: The frame lost power when moving the bed board horizontally. It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: Confirmed the fault with the customer.PRS Password expiredPlease reset SSH2 FTP password onsite Remote troubleshoot and test result: Restart but didn_t work. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: It is recommended to check scan bed horizontal position readings. Re perform horizontal position calibration. Based on the error report, determine if clutch is suction closed well. If necessary replace the absolute and relative encoder to eliminate the problem. We recommend spare parts 12NC and names: PPD 453566500911, 45356500901 SAP Transfer:Case Transferred to SAP Next Step Action: PPD 453566500901 *1 +453566500911* pending to confirm the replacement number, Zhang Yifei-86-17621702391 SAP Transfer:Dispatched for On site Support Safety Questions:Case: 0118675529 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: han susu Last updated by: han susu 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 3 |  | The frame lost power when moving the bed board horizontally. Q7: (Confirmed reported problem. & The Gantry shut down. ) $Q8: ( Check the error log was not found. Check the table & We need to calibrate table) 453566500901 *1+453566500911* 待确认替待号 | [12 empty] | 118675529 | In Process | 1497632 |  |  | Investigation | Investigation | PartReplacement | [5 empty] | Others | [6 empty] | E-Stop | [5 empty] | Refresh |  |  | System issue |  |  |  | - | - | PartReplacement | 0 | bar: No
1564 |  |  | Japan | Social Medical Corporation BOKOI, Nikko Memorial Hospital | 341011 | 44511 | 45 | 04DS: KK: Ring artifact occurred during coronary shooting. | Work Order #: WO-04741237 11-Nov-2021 Konishi, Hideo The system meets the specification for the performed service and is returned to use. Resolution: [Work details] </br> <br> signals were not received at all from the 2 channels of Detector Module 25. I replaced Detector Module 25 and that improved the situation. After the replacement, I ran a test scan and confirmed that there were no problems. Internal Comments: [CX related] Did you agree on how to complete the report with the person in charge of devices? <Select one of the following> Yes. I completed normal reporting that day. [Work report (paper) submission] Did you submit a work report on paper? Yes Q1. How did the examination finish? Select from the following: *Because the answer is No for Q1, not applicable Q2. Were there any mistaken findings (misdiagnosis, mistaken treatment) due to the problem? None. Q3. Were there any artifacts? None. Q4. Was there any unintended unnecessary exposure? None. Resolution: "Dead channel" was confirmed in Mod25 during a bad detector test. Parts Consumed : 1.00000,459800911271,TDMS V2 M.Det Class A FRU Problem Description by Engineer: A coronary scan produced a ring artifact. Regulatory Questions : N/A ----------------------------------------------------------------------------- Work Order #: WO-04741195 Kojima, Kenichi The system does not meet the specification for the performed service and is not returned to use Internal Comments: [Subject] 04DS: KK: Mr. Nishiya: A coronary scan produced a ring artifact. Shipped part: 459800911271:TDMS V2 M.Det Class A FRU QTY:1 ZREP/Traceable --None-- Dispatched: 1 RC FSE Registration Information: Name of FSE 1 who visited the site: Hideo Konishi FSE 1 visit date and time: 2021/11/1112:00 FSE 1 hours worked: 4 Dispatch instructions: Dispatch instruction #1: Shipping conditions: CPU (remove from warehouse, no shipping arrangements) Delivery date: Must arrive by 10:00 on November 11, 2021. Shipping destination: Sapporo warehouse. FSE will collect. To be collected by: Hideo Konishi Phone number: 080-1605-0255 Other: Person making request: Kenichi Kojima Problem Description by Engineer: A coronary scan produced a ring artifact. Diagnostic performed by Engineer: The ring artifact appears to have been caused by a problem with the module. An FSE will respond. Follow Up Required : * ----------------------------------------------------------------------------- | 2 |  | 459800911271:TDMS V2 M.Det Class A FRU QTY:1 | [12 empty] | 118697409 | Closed | 1509826 |  |  | Investigation | Investigation | CTUDT00632110 | [5 empty] | Engineering | [6 empty] | Ring Artifact | [5 empty] | Refresh |  |  | System issue | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | DMS Module | DMS-HW | DMS-HW-DMS Module | bar: No
1565 |  |  | China | Suzhou Guang ci Tumor Hospital Co., Ltd | 345019 | 44505 | 44 | 02N/Ingenuity Core128/Elite China/Image contrast | BU Q&R Feedback: No Reason: Corrective Maintenance Safety Questions:Case: 0118668075 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Jianqiang Last updated by: Wang Jianqiang 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Problem Reported by Customer: Image contrast T2 Activities:Problem Description: Summary of customer_s problems: Other types of error. Detailed description of Customer_s problems or malfunctions: Image quality is low and contrast test exceeds range. Need an engineer to calibrate It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: Image contrast is low, failed. Remote troubleshoot and test result: There is a small problem with image quality PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: On-site calibration of images, air, and module. SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support ----------------------------------------------------------------------------- | 3 |  | 服务内容：设备低分测试没过，重新水膜校正 服务结果: 设备正常使用 | [12 empty] | 118668075 | In Process | 1494923 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Others | [6 empty] | Calibration | [5 empty] | Refresh |  |  | System issue | bar: No
1566 |  |  | China | Guangxi Province Teng County People’s Hospital | 40036 | 44504 | 44 | 02N/Ingenuity CT 728326/Replace the head mop pad | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : Replace the head mop pad Safety Questions:Case: 0118664106 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: mo xueke Last updated by: mo xueke 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | 5 |  | head mop pad | [12 empty] | 118664106 | Closed | 1491986 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Others | [6 empty] | Head pad | [5 empty] | Refresh |  |  | System issue | bar: No
1567 |  |  | China | Laiyang City People_s Hospital | 345005 | 44515 | 46 | 05/Ingenuity Core128/Elite China/Add a node | BU Q&R Feedback: No SAP Customer Symptom Code: MARF Customer Symptom Code: Monitor Alarm Failure Reason: Corrective Maintenance Engineer Failure Code: Incompatible Use Repair Activity Code: Instructed Customer Problem Reported by Customer : 05/Ingenuity Core128/Elite China/ Next Step Action: 20211114sjn15 Safety Questions:Case: 0118713549 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: liu xiaohu Last updated by: liu xiaohu 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: ask to add a dicom node Troubleshooting Action: instruct to customer to add the node under advanced account Repair Action: instruct the customer to add the node by themselves Test/Inspection: Functional Check PASSED. The device returned to full functionality. Test Equipment: NA ----------------------------------------------------------------------------- | 3 | [14 empty] | 118713549 | Fixed | 1523602 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [19 empty] | Refresh |  |  | System issue | bar: No
1568 |  |  | China | Shandong Province Chest Hospital | 40025 | 44502 | 44 | Login failed | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Calibration/Adjustment Safety Questions:Case: 0118651053 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Gaoge Last updated by: Wang Gaoge 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Dispatched for On site Support Problem Reported by Customer: 05L/Ingenuity CT 728326 T2 Activities:Problem Description: Summary of customer_s problems: failure related to computer software. Detailed description of Customer_s problems or malfunctions: System shows blue screen It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: The Windows blue screen requires IST privileges before it can be used again. Remote troubleshoot and test result: Software fault PRS TR status: It is not through, ask engineer for maintenance on site. Repair Action: Suggest onsite service procedures: Need to reinstall the Windows system and software. Test/Inspection: N/A Test Equipment: N/A SAP Transfer:Case Transferred to SAP Next Step Action: huang zhaokun External Remarks: 02.11.2021 10:01:12 UTC (86038405) Service Details: Used IST privileges to log in to the software, then rebooted the Host. Service result: The device returned to normal working conditions after being repaired. Internal Remarks : 02.11.2021 10:01:15 UTC (86038405) "RDI Accuracy: YES." $Q7: ( Confirmed reported problem: Software can not login) ) $Q8: ( Login in Software use smart card and configuration. & L3 support reply this is Windows 10 Bug and need to Fixed future.) $Q9: ( Login Software use smart card and configuration.) $Q10: ( Test and inspection data refer to: PA . Device returned to fully functionality.) $Q11: Calibrated Tools name (N/A) ^Current Tube installed date: N/A ^Current Tube usage: N/A ----------------------------------------------------------------------------- | 3 |  | $Q7: ( Confirmed reported problem: Software can not login) ) $Q8: ( Login in Software use smart card and configuration. & L3 support reply this is Windows 10 Bug and need to Fixed future.) $Q9: ( Login Software use smart card and configuration.) | [12 empty] | 118651053 | Closed | 1483386 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632553 | [5 empty] | Engineering | [6 empty] | Blue screen login | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Blue-Screen | Console-SW-Blue-Screen | Console-SW-Blue-Screen | CTUDT00632553 | Blue-Screen | bar: No
1569 |  |  | China | Zhejiang Province Yueqing No. 1 People’s Hospital | 40054 | 44502 | 44 | 02N/Ingenuity CT/ECG signal lost sometimes | BU Q&R Feedback: No Reason: Corrective Maintenance Next Step Action NTD: Customer has contacted site1, refused to transfer to RSE Problem Reported by Customer: 02N/Ingenuity CT/ SAP Transfer:Case Transferred to SAP Safety Questions:Case: 0118649526 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer_s problems: Other types of error. Detailed description of Customer_s problems or malfunctions: Sometimes the ECG signal is lost It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: Sometimes the ECG signal is lost Remote troubleshoot and test result: ECG signal transmission problems. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Replace ECG signal_s direct connection, no adapter. Lower to the fault point 0 459801731761 USB EXTENDER CABLE FROM IBOX TO CONSOLE SAP Transfer:Dispatched for On site Support Next Step Action:huang yi Next Step Action: PPD part USB EXTENDER CABLE FROM IBOX TO CONSOLE 459801731761 1, has arranged 12H to the site, FSE: Huang Yi -86-15958725282 Internal Remarks : 03.11.2021 05:12:04 UTC (86036249) RDI Accuracy:Yes $Q7: (Confirmed reported problem: ECG signal intermittent lost) $Q8: (check found usb often disconnect from ibox to host) $Q9: (replace the USB cable from ibox to host) $Q10: (Test and inspection data refer to: PA, Device returned to full functionality.) $Q11: Calibrated tools name(N/A) S/N (N/A) Next calibration date (N/A) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date:N/A ^Current Tube usage: second:N/A External Remarks: 03.11.2021 05:12:02 UTC (86036249) service details: ECG signal is disconnected intermittently. Suspect the cable from IBOX to HOST is unstable. Replace the cable. service result: device returns to normal working condition after service Other significant matters: none ----------------------------------------------------------------------------- | 3 |  | $Q7: (Confirmed reported problem: ECG signal intermittent lost) $Q8: (check found usb often disconnect from ibox to host) $Q9: (replace the USB cable from ibox to host) | [12 empty] | 118649526 | Closed | 1482588 |  |  | Investigation | Investigation | IBOX | [5 empty] | Engineering | [6 empty] | IBOX USB | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | IBOX | Gantry-HW-IBOX | Gantry-HW-IBOX | IBOX | IBOX | Gantry-HW | Gantry-HW-IBOX | bar: No
1570 |  |  | China | Guangxi Province Teng County People’s Hospital | 40036 | 44501 | 44 | 02N/Ingenuity CT 728326/pacs received images quality is poor | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 02N/Ingenuity CT 728326 Safety Questions:Case: 0118638063 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer_s problems: Other types of error. Detailed description of Customer_s problems or malfunctions: The image quality received by pacs is reduced It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: The image quality received by pacs is reduced Remote troubleshoot and test result: Pacs compressed the image after receiving it PRS TR status: Everything passed through. Repair Action: Suggest onsite service procedures: Check PACS node setup. Work with PACS to find cause SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support ----------------------------------------------------------------------------- | 5 |  | $Q6: (Image problem) $Q7: (Confirm reported problem & Found image problem.) $Q8: (Found image problem . & Need to do image calibration.) $Q9: (Do QA calibration.) | [12 empty] | 118638063 | Closed | 1479208 |  |  | Investigation | Investigation | [6 empty] | Others | [6 empty] | Calibration | [5 empty] | Refresh |  |  | System issue | bar: No
1571 |  |  | China | Shaanxi Province Ankang T.C.M. Hospital | 345041 | 44501 | 44 | WeChat：02N/Ingenuity Core128/Elite China/couch fault | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Clean / Lubricate Problem Reported by Customer: couch fault SAP Transfer:Case Transferred to SAP Safety Questions:Case: 0118637504 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Gaoge Last updated by: Wang Gaoge 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer_s problems: There is a fault with the horizontal movement of the scan bed. Detailed description of Customer_s problems or malfunctions: There is an abnormal sound when the couch is moving horizontally. It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: There is an abnormal sound when the couch is moving horizontally. Remote troubleshoot and test result: Consider drive shaft friction or a different shaft causing the abnormal sound. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: 1. Check the source of the abnormal sound in the horizontal drive. Lubricate or adjust to eliminate the abnormal sound. Test/Inspection: N/A Test Equipment: N/A SAP Transfer:Dispatched for On site Support T2 Activities : Repair Action: replace Motor Coupling 453567099401 External Remarks: 08.11.2021 02:48:32 UTC (86039086) service reason: patient bed is noisy. Service Contents: replace coupling Result of service: device operating normally Internal Remarks : 08.11.2021 02:48:35 UTC (86039086) Q7: (Confirmed reported problem. & The Gantry couple have problem. $Q8: (Check the error log was not found. Check the couple & We need to replace the couple) $Q9: ( 1. Replace the couple2. Test use for CT was OK. $Q10: (Test and inspection data refer to: PA. The device returned to full functionality. $Q11: NA ----------------------------------------------------------------------------- | 3 |  | replace Motor Coupling 453567099401 | [12 empty] | 118637504 | Closed | 1478839 |  |  | Investigation | Investigation | CTUDT00632179 | [5 empty] | Engineering | [6 empty] | Coupling | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | Motor Coupling | Couch-HW-Motor Coupling | Couch-HW-Motor Coupling | CTUDT00632179 | Motor Coupling | Couch-HW | Couch-HW-Motor Coupling | bar: No
1572 |  |  | Germany | Gesundheit Nord gGmbH Klinikverbund | 40002 | 44529 | 48 | Kunde bemängelt die Bildqualität des System bei Axialen Scans, Biopsien und Kopfaufnahmen | Internal Remarks: See email attachments => Ms. Bösemann /Application will still clarify ----------------------------------------------------------------------------- | 3 |  | Customer complains about the image quality of the system for axial scans, biopsies and head images | [12 empty] | 118803627 | New | 1569285 |  |  | Investigation | Investigation | [6 empty] | Application | [6 empty] | Application | [5 empty] | Refresh |  |  | System issue | bar: No
1573 |  |  | Japan | Independent administrative institution National Hospital Organization | 341015 | 44519 | 46 | CT:APP:FN | ①冠動脈の後にAoを撮影したい。冠動脈はBTでAoはTimedで組みsaveできますが、大丈夫でしょうか？②胸服部単純+CS+BT+冠動脈+胸腹部2相を撮影するとTube_hotが出ます。#_of_cycleを少なくしたり、Cycle_Timeを長くすることで効果ありますか？③最近2件ほど起きた現象ですが、CTViewerで作成したMPRが真っ黒になりました。他のサイトでもありますか？ |  |  | Split CFM for MPR and Bolus Track issue Case Number# 0118748208 Description ① I want to take a picture of Ao after the coronary artery. The coronary arteries can be assembled with BT and Ao with Timed, but is that okay? ② (2) Tube_hot appears when you take a picture of Chest Hattori Simple + CS + BT + Coronary Arteries + Chest and Abdomen Phase 2. Is it effective to reduce #_of_cycle or increase Cycle_Time? ③ (3) This is a phenomenon that has occurred about two times recently, but the MPR created by CTViewer has turned black. Is it another site? Response from Remote Clinical Support Nakagawa, Futoshi: ① Basically, BT and Timed cannot be used together. Even if you can save, some settings may have been changed automatically. Please do a test scan once. ② I think that there will be some effect by reviewing the BT conditions. The number of times is 30 → 20, the cycle time is 1.3sec → 1.7sec, etc. ③ I have hardly heard of it. If the phenomenon occurs repeatedly, please leave a history such as taking a picture. | [14 empty] | 1567977 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632116 | [5 empty] | Engineering | Application | [5 empty] | MPR | [5 empty] | Refresh |  |  | System issue | System | CIRS-SW | MPR Recon | CIRS-SW-MPR Recon | CIRS-SW-MPR Recon | CTUDT00632116 | MPR Recon | CIRS-SW | CIRS-SW-MPR Recon | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y | #N/A
1574 |  |  | Japan | Independent administrative institution National Hospital Organization | 341015 | 44519 | 46 | CT:APP:FN | ①冠動脈の後にAoを撮影したい。冠動脈はBTでAoはTimedで組みsaveできますが、大丈夫でしょうか？②胸服部単純+CS+BT+冠動脈+胸腹部2相を撮影するとTube_hotが出ます。#_of_cycleを少なくしたり、Cycle_Timeを長くすることで効果ありますか？③最近2件ほど起きた現象ですが、CTViewerで作成したMPRが真っ黒になりました。他のサイトでもありますか？ |  |  | Split CFM for MPR and Bolus Track issue | [14 empty] | 1567972 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | Application | [5 empty] | Bolus Track | [5 empty] | IB Exist |  |  | System issue | system | cosnole-sw | bolustracking | cosnole-sw-bolustracking | cosnole-sw-bolustracking | 0 | bolustracking | bar: No
1575 |  |  | Japan | 医療法人清翠会 牧病院 | 341017 | 44528 | 47 | 03CS:KKI:ShortTC時に、アーク発生＞＞継続してShortTC実施＞szInfo: SoftArcReg 0x211*/S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT | Work Order #: WO-04862714 Kimura, Kazumasa Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 4 |  | Continue to implement Short TC | [12 empty] | 118798971 | In Progress | 1564481 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00630085 | [5 empty] | Engineering | [6 empty] | Generator Arc | [5 empty] | Refresh |  |  | System issue | System | IC-HW | Tub Arc | IC-HW-Tub Arc | IC-HW-Tub Arc | CTUDT00630085 | Tub Arc | IC-HW | IC-HW-Tub Arc | bar: No
1576 |  |  | Japan | Tomakomai City Hospital | 341014 | 44525 | 47 | I will cancel because I made a mistake | Work Order #: WO-04843034 Miyakawa, Takeyuki Problem Description by Engineer: Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer: N/A ----------------------------------------------------------------------------- | 4 |  | I will cancel because I made a mistake Therefore, it cannot be canceled. Thank you for the closing process. | [12 empty] | 118782466 | In Progress | 1559261 |  |  | Evaluation | Evaluation | [6 empty] | Others | [6 empty] | Mistake | [5 empty] | Refresh |  |  | System issue | bar: No
1577 |  |  | Japan | Tomakomai City Hospital | 341014 | 44525 | 47 | 01DS: TM: When burning a CD, the data can be written, but failed is displayed. | Work Order #: WO-04843151 Miyakawa, Takeyuki The system meets the specification for the performed service and is returned to use. Diagnostic performed by Engineer: When he burns a CD, data can be written, but a "Failed" message was displayed Problem Description by Engineer: When he burns a CD, data can be written, but a "Failed" message was displayed Internal Comments: [Subject] 01DS:TM:Mr. Hanada: When he burns a CD, data can be written, but a "Failed" message was displayed [Internal Comments] FSE will respond. Shipped part: 459801631381:HPE HP 9.5mm Slim SuperMulti DVD Writer QTY:1 ZSID BackOrder Dispatched: 1 RC FSE Registration Information: Name of FSE 1 who visited the site: Mototsugu Sugahara FSE 1 visit date and time: 2021/11/3018:00 FSE 1 hours worked: 2 Dispatch instructions: Dispatch instruction #1: Shipping conditions: NBD (mixed loading, next day or thereafter) Delivery date: November 26, 2021, no time specified Shipping destination: Sapporo Office To be collected by: Mototsugu Sugahara Phone number: Other: If there is no stock in Japan, please order it from overseas. Person making request: Follow Up Required : * ----------------------------------------------------------------------------- Work Order #: WO-04843316 Sugawara, Mototsugu Problem Description by Engineer: When he burns a CD, data can be written, but a "Failed" message was displayed Regulatory Questions : N/A Diagnostic performed by Engineer: N/A ----------------------------------------------------------------------------- | 4 |  | 1.00000,459801631381,HPE HP 9.5mm Slim SuperMulti DVD Writer | [12 empty] | 118782728 | In Progress | 1556046 |  |  | Evaluation | Evaluation | TBD-DVDDrive | [5 empty] | Engineering | [6 empty] | ShouldnotEngineering | [5 empty] | Refresh |  |  | System issue | System | Console-HW | ShouldnotEngineering | Console-HW-DVDDrive | Console-HW-DVDDrive | TBD-DVDDrive | DVDDrive | Console-HW | Console-HW-DVDDrive | bar: No
1578 |  |  | Japan | Social Welfare Organization Saiseikai Imperial Gift Foundation, Inc. | 343000 | 44524 | 47 | 01CS: TM: Message "Missing ..." was displayed after shooting >> All images were created by closing the window >> No hardware error | Work Order #: WO-04836280 Miyakawa, Takeyuki Problem Description by Engineer: Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer: N/A ----------------------------------------------------------------------------- | 4 |  | Description: $ version = "v4.1.10.25210" 9:49 Call content The message "Missing ..." was displayed twice after the Axial shooting of the head. Click OK to complete all the images. I am creating a Bug Report. correspondence Confirm that there are no errors in the Event log for both CIRS and CT-Host. It was judged that there was an error on the CT-Host side, and we asked you to restart the CT-Host when the time was up, and we decided to continue using it. https://philipsb2bsc.my.salesf | [12 empty] | 118775297 | In Progress | 1555574 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | SW Error | [5 empty] | Refresh |  |  | System issue | System | Console-SW | NoEnoughInfo | Console-SW-NoEnoughInfo | Console-SW-NoEnoughInfo | NoEnoughInfo | NoEnoughInfo | bar: No
1579 |  |  | Japan | Social Medical Corporation BOKOI, Nikko Memorial Hospital | 341011 | 44523 | 47 | 04CS: KKI: 11/23 An error occurred when the key switch was turned on >> Continued use | Work Order #: WO-04833729 Kimura, Kazumasa Problem Description by Engineer: Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer: N/A ----------------------------------------------------------------------------- | 4 |  | 04CS: KKI: Tsukakoshi: 11/23 An error occurred when the key switch was turned on >> Continued use [Internal Comments] [Log confirmation result] It is probably caused by the following error. S_TILT_C_MANUAL_TILT_OFF It was said that a tilt error occurred. It seems that it was used continuously because the tilt operation was confirmed and there was no problem. After that, when the key switch was turned off / on, the same error did not occur. We have decided that this is a temporary phenomenon and will wait and see. | [12 empty] | 118774968 | In Progress | 1555475 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Tilt | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | NoEnoughInfo | Gantry-HW-NoEnoughInfo | Gantry-HW-NoEnoughInfo | NoEnoughInfo | NoEnoughInfo | bar: No
1580 |  |  | Japan | Independent administrative institution National Hospital Organization | 341015 | 44519 | 46 | CT:APP:FN | Work Order #: WO-04802999 Nakagawa, Futoshi The system meets the specification for the performed service and is returned to use. Resolution: (1) Basically speaking, BT and Timed cannot be used together.Even if saving can be done, it is possible that some settings have been automatically changed.I asked him to try a test scan.(2) I think it will be somewhat effective if the BT conditions are reviewed.Reduce the number of cycles from 30 to 20, and then change the cycle time from 1.3sec to 1.7sec, etc.(3) I_ve hardly ever heard of it.If the problem occurs repeatedly, please take photos and leave a record. Problem description by Clinical Support Specialist : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer: NA ----------------------------------------------------------------------------- | 5 | [14 empty] | 118748208 | Closed | 1538262 |  |  | Investigation | Investigation | [6 empty] | Application | [6 empty] | Application | [5 empty] | Refresh |  |  | System issue | bar: No
1581 |  |  | Japan | Maki Hospital | 341017 | 44510 | 45 | JS:warranty期間中のフォローアップ | Work Order #: WO-04732805 Ootani, Natsuki Regulatory Questions : N/A Problem Description by Engineer: Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer: N/A ----------------------------------------------------------------------------- | 5 | [14 empty] | 118690608 | In Progress | 1509798 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [19 empty] | Refresh |  |  | System issue | bar: No
1582 |  |  | Japan | Maki Hospital | 341017 | 44518 | 46 | An error occurred during head shooting, performed long tube calibration | Work Order #: WO-04788786 Kimura, Kazumasa The system meets the specification for the performed service and is returned to use. Diagnostic performed by Engineer: Examinations can continue Problem Description by Engineer: An error occurred during a head scan. Internal Comments: [Subject] 03DM:KKI:Mr. Noda: An error occurred during a head scan. >> SoftArcReg 0x211*/S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT >> He was able to continue the examination. [Internal Comments] [Results of checking the log] An arc occurred. SoftArcReg 0x211* S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT I think it was probably the initial arcing after X-ray tube replacement. I think it was a temporary problem, so I asked him to monitor the situation. Dispatched: 1 RC FSE Registration Information: Name of FSE 1 who visited the site: Oyama Tetsuo FSE 1 visit date and time: 2021/11/1822:00 FSE 1 hours worked: 2 Follow Up Required : * ----------------------------------------------------------------------------- Work Order #: WO-04789955 Ohyama, Tetsuo Regulatory Questions : N/A Problem Description by Engineer: An error occurred during a head scan. Diagnostic performed by Engineer: N/A ----------------------------------------------------------------------------- | 4 |  | I think it was probably the initial arcing after X-ray tube replacement. I think it was a temporary problem, so I asked him to monitor the situation. | [12 empty] | 118735894 | In Progress | 1533185 |  |  | Closed - Done | Closed - Done | CTUDT00630085 | [5 empty] | Engineering | [6 empty] | Generator Arc | [5 empty] | Refresh |  |  | System issue | System | IC-HW | Tub Arc | IC-HW-Tub Arc | IC-HW-Tub Arc | CTUDT00630085 | Tub Arc | IC-HW | IC-HW-Tub Arc | bar: No
1583 |  |  | Japan | 医療法人清翠会 牧病院 | 341017 | 44516 | 46 | 01CS: RS: CT indicator at the bottom left of the screen is blinking >> Remove and eliminate the SR file that was transferred by mistake | Work Order #: WO-04768640 Satake, Rui The system meets the specification for the performed service and is returned to use Problem Description by Engineer : 画面左下のCTインジケータが点滅している Internal Comments : [Subject] 01CS:RS:ﾄｶﾞﾜ様:画面左下のCTインジケータが点滅している>>誤って転送したSRファイルをRemoveし解消 [Internal Comments] なし [4/8 Elements] Q5 その検査はどの様に終了しましたか?:無しQ6 問題により間違った所見（誤診、誤治療）の有無:無し Q7 意図しない不要な曝射の有無:無し Q8 アーチファクトの有無:無し Diagnostic performed by Engineer : 誤って転送したSRファイルをRemoveし解消 ----------------------------------------------------------------------------- | 4 |  | Diagnostic performed by Engineer: Removes and resolves erroneously transferred SR files | [12 empty] | 118719656 | In Progress | 1522964 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | TBD-DoseReport | [5 empty] | Engineering | [6 empty] | DSR | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Dose SR | Console-SW-Dose SR | Console-SW-Dose SR | TBD-DoseReport | Dose SR | Console-SW | Console-SW-Dose SR | bar: No
1584 |  |  | Japan | Tomakomai City Hospital | 341014 | 44511 | 45 | 02DS: KKI: Abnormal noise is generated when the bed is horizontally operated. | Work Order #: WO-04743178 Sugawara, Mototsugu Regulatory Questions : N/A Problem Description by Engineer : The couch makes a rattling noise when it is moved horizontally. Diagnostic performed by Engineer: N/A ----------------------------------------------------------------------------- Work Order #: WO-04742802 Kimura, Kazumasa The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Internal Comments: [Subject] 02DS:KKI:Mr. Hisano: The couch makes a rattling noise when it is moved horizontally. [Internal Comments] None Shipped part: 453566489521:MCG /PITTMAN MOTOR ASSEMBLY QTY:1 ZSID Sapporo (U612) 453566494823:CLUTCH - HORIZONTAL DRIVE QTY:1 ZSID/Traceable Sapporo (U612) Dispatched: 1 RC FSE Registration Information: Name of FSE 1 who visited the site: Mototsugu Sugahara FSE 1 visit date and time: 2021/11/1220:00 FSE 1 hours worked: 2 Dispatch instructions: Dispatch instruction #1: Shipping conditions: ASAP (Arrange shipment) Delivery date: ASAP Shipping destination: The FSE will pick it up at Shin-Chitose Airport. To be collected by: Mototsugu Sugahara Phone number: 080-1053-9342 Other: Person making request: Kazumasa Kimura Follow Up Required : * Problem Description by Engineer : The couch makes a rattling noise when it is moved horizontally. Diagnostic performed by Engineer: Examinations can be performed without any problems.No error. ----------------------------------------------------------------------------- | 3 |  | Description: [Details of work] There was a misalignment of the coupling (elastic metal cylinder) that connects the horizontal drive shaft of the sleeper and the encoder for position detection, causing noise. The alignment has been readjusted to eliminate the abnormal noise. November 11, 2021 Approved by Mr. Fujiki. | [12 empty] | 118699016 | In Progress | 1511119 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632179 | [5 empty] | Engineering | [6 empty] | Coupling | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | Motor Coupling | Couch-HW-Motor Coupling | Couch-HW-Motor Coupling | CTUDT00632179 | Motor Coupling | Couch-HW | Couch-HW-Motor Coupling | bar: No
1585 |  |  | United States | Lackey Memorial Hospital | 40015 | 44509 | 45 | Clinical support | Internal Remarks : Interaction Type: New Case Caller: Michea McLemore IP & Issue: 85117259 & Protocol Event Type: Incident Service Request: Remote Clinical Did Customer Decline Remote Support? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Created case & xfer Clinapps Next Step Action: N/A Verbal PO Number: Verbal PO Approver: Verbal PO CAP: Integration : Confirmation of order 54417555 saved ----------------------------------------------------------------------------- Work Order #: WO-04731324 Curry, Dyan The system meets the specification for the performed service and is returned to use Problem description by Clinical Support Specialist : Customer Function/ Role: technologist Device Use: in use Expected and Unexpected Results: ped EC questions User Impact: none Patient Impact: none Current Software version: 4.1 Diagnostic performed by Engineer : answered dose tools questions for pediatric ecs. advised customer dose tools should not be used of head and spine work for infants. ----------------------------------------------------------------------------- | 5 | [14 empty] | 118689619 | Closed | 1504587 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [19 empty] | Refresh |  |  | System issue | bar: No
1586 |  |  | Japan | 医療法人清翠会 牧病院 | 341017 | 44510 | 45 | Recon error, confirmed system could work properly | Integration : Confirmation of order 54419806 saved ----------------------------------------------------------------------------- Work Order #: WO-04734347 Kimura, Kazumasa The system meets the specification for the performed service and is returned to use Internal Comments : [Subject] 01CS:KKI:ﾄｶﾞﾜ様:”Not all images for (Patient Name) were created..”のポップアップ発生＞＞継続して検査可 [Internal Comments] 【ログ確認結果】 ハード的なエラーは確認できておりません。 撮影途中でキャンセルを行い撮影を終了しておりました。 おそらく、その影響でポップアップが発生したものと思われます。 一時的な現象だと思われますので、様子見とさせて頂いております。 [4/8 Elements] Q5 その検査はどの様に終了しましたか?:無しQ6 問題により間違った所見（誤診、誤治療）の有無:無し Q7 意図しない不要な曝射の有無:無し Q8 アーチファクトの有無:無し Resolution : 継続して検査可 Problem Description by Engineer : ”Not all images for (Patient Name) were created..”のポップアップ発生 ----------------------------------------------------------------------------- | 4 |  | Work Order #: WO-04734347 Kimura, Kazumasa The system meets the specification for the performed service and is returned to use Internal Comments: [Subject] 01CS: KKI: Togawa: "Not all images for (Patient Name) were created .." pop-up occurred >> Continued inspection possible [Internal Comments] [Log confirmation result] We have not confirmed any hardware errors. I canceled the shooting in the middle of shooting and finished shooting. Perhaps that_s what caused the pop-up. It seems to be a temporary phenomenon, so I will wait and see. [4/8 Elements] Q5 How did the test end ?: None Q6 Presence or absence of incorrect findings (misdiagnosis, mistreatment) due to the problem: None Q7 Presence or absence of unintended unnecessary exposure: None Q8 Presence / absence of artifacts: None Resolution: Can be continuously inspected Problem Description by Engineer: "Not all images for (Patient Name) were created .." pop-up occurred | [12 empty] | 118692096 | Closed | 1510375 |  |  | Closed - Done | Closed - Done | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Popup message | [5 empty] | Refresh |  |  | System issue | System | CIRS-SW | NoEnoughInfo | CIRS-SW-NoEnoughInfo | CIRS-SW-NoEnoughInfo | NoEnoughInfo | NoEnoughInfo | bar: No
1587 |  |  | Peru | BANCO BBVA PERU | 40034 | 44508 | 45 | 40034-THERE IS AN ERROR IN THE SYSTEM, IT DOES NOT ALLOW TO DO CALIBRATION OR HEATING THE TUBE | Work Order #: WO-04720573 [REDACTED] The system does not meet the specification for the performed service and is not returned to use. Problem Description by Engineer : Problem description by engineer : Gantry error Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected Results: No gantry errors. Patient Impact/Outcome: Equipment is down. User Impact: Equipment stopped. Current Software Version: 4.1.10 Diagnostic performed by Engineer: There is an error in the gantry and the device does not allow tube heating. The remote access is not working. They were able to extract the error file and the customer sent them for review. Cathode arcing error occurs in the high voltage stage. The cathode tank will be ordered. Internal Comments : Error # and/or description of error : S_GEN_FAULT_ANALOG_FAULT 2:5 Analog Fault S_GEN_FAULT_CATHODE_TUBE_ARC 5:5 CPM Arc Repair Action: - Check high voltage contacts. - Perform HV diagnostics - If necessary, change the cathode module. Follow Up Required: Error Number and/or Description: S_GEN_FAULT_ANALOG_FAULT 2:5 Analog Fault S_GEN_FAULT_CATHODE_TUBE_ARC 5:5 CPM Arc Repair Action: - Check high voltage contacts. - Perform HV diagnostics - If necessary, change the cathode module. Technician_s Name: [REDACTED] ----------------------------------------------------------------------------- Work Order #: WO-04730012 [REDACTED] Problem Description by Engineer : Problem description by engineer : Gantry error Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected Results: No gantry errors. Patient Impact/Outcome: Equipment is down. User Impact: Equipment stopped. Current Software Version: 4.1.10 Diagnostic performed by Engineer: N/A ----------------------------------------------------------------------------- | 2 |  | Resolution Description: Inspection of high voltage connectors in X-ray tube and generators was carried out, finding evidence of arcs in cathode, cleaning and lubrication of connectors was carried out, long conditioning tests of X-ray tube were continued without inconveniences . Satisfactory functional tests were performed. | [12 empty] | 118680961 | Closed | 1503168 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | HV Connector | [5 empty] | Engineering | [6 empty] | HV Connector | [5 empty] | Refresh |  |  | System issue | System | IC-HW | HV Connector | IC-HW-HV Connector | IC-HW-HV Connector | HV Connector | HV Connector | bar: No
1588 |  |  | Japan | Nikko Memorial Hospital | 341011 | 44505 | 44 | Unable to recon coronal/sagittal with tilt angle during non-helical scan, CTUDT00632116 addressed investigation |  | #N/A |  | CTUDT00632116 | [14 empty] | 1491718 |  |  | Closed - Done | Closed - Done | CTUDT00632116 | [5 empty] | Engineering | [6 empty] | coronal/sagittal | [5 empty] | Refresh |  |  | System issue | System | CIRS-SW | MPR Recon | CIRS-SW-MPR Recon | CIRS-SW-MPR Recon | CTUDT00632116 | MPR Recon | CIRS-SW | CIRS-SW-MPR Recon | bar: No
1589 |  |  | Japan | 医療法人清翠会 牧病院 | 341017 | 44502 | 44 | 03DM: RS: X-ray system cannot comply at the time of Survuew and stop scanning >> Continue scanning after replanning >> S_GEN_FAULT_ANODE_TUBE_ARC | Work Order #: WO-04768072 16-Nov-2021 Ohyama, Tetsuo The system meets the specification for the performed service and is returned to use Internal Comments : [CX関連] 機器担当者と報告完了方法を合意できましたか？ ＜以下のいずれかを選択＞ YES 当日通常報告完了 [作業報告書（紙）提出] 作業報告書は紙で提出しましたか？ YES Q1その検査はどの様に終了しましたか?: ＜以下から選択＞ ・中断した箇所から再撮影 Q2 問題により間違った所見（誤診、誤治療）の有無: ・無し Q3 アーチファクトの有無: ・無し Q4 意図しない不要な曝射の有無: ・無し Regulatory Questions : N/A Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : WO-04705797にて対応、報告済み。翌日、立会い分WOになります。 ----------------------------------------------------------------------------- Work Order #: WO-04716475 16-Nov-2021 Kawazoe, Junichi The system meets the specification for the performed service and is returned to use Resolution : Case#0118649575報告済み Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Regulatory Questions : N/A ----------------------------------------------------------------------------- Work Order #: WO-04705797 15-Nov-2021 Ono, Junpei The system meets the specification for the performed service and is returned to use Resolution : [使用部品] 部品番号:453567560022 /部品名:Spellman 80 KW RoHS CPM Assy /個数:1 Tools & Equipment : TORQUE WRENCH,2022-05-11 00:00:00 Parts Consumed : 1.00000,453567560002,Spellman 80 KW RoHS CPM Assy 1.00000,989000086352,PHILIPS MRC 880 Follow Up Required : N/A Resolution : [作業内容詳細] アークが頻発する件につきまして、Ｘ線管球の交換および高圧発生装置の 交換作業を実施致しました。 作業後、装置の動作確認を行い正常に動作する旨確認致しました。 X線管球の交換後、焦点のキャリブレーションにて一部の項目が基準値内に 収まらないために、調査およびＸ線管球の再交換により予定時間を大幅に 超過しご迷惑をおかけし大変申し訳ありませんでした。 Problem Description by Engineer : Survuewの際にX-ray system cannot complyとでてスキャン停止 Regulatory Questions : N/A ----------------------------------------------------------------------------- Work Order #: WO-04682011 Satake, Rui The system meets the specification for the performed service and is returned to use Diagnostic performed by Engineer : 再プラン後スキャン継続. S_GEN_FAULT_ANODE_TUBE_ARCの発生を確認。 FSE対応。 Problem Description by Engineer : Survuewの際にX-ray system cannot complyとでてスキャン停止 Follow Up Required : * Internal Comments : [Subject] 03DM:RS:ﾄｶﾞﾜ様:Survuewの際にX-ray system cannot complyとでてスキャン停止>>再プラン後スキャン継続>>S_GEN_FAULT_ANODE_TUBE_ARC [Internal Comments] なし [出庫部品] 989000086352:PHILIPS MRC 880 QTY:1 ZREP/Traceable 大阪(U615) 453566495091:1XRT SERVICE TOOL ASSY EXP QTY:1 ZSID 大阪(U615) 459800089031:Tool, XRT Strap Pad Assy QTY:1 ZSID 大阪(U615) 459800911271:TDMS V2 M.Det Class A FRU QTY:1 ZREP/Traceable 大阪(U615) [ディスパッチ] 1 RC -----------------FSE登録情報----------------- -訪問FSE1氏名 小野 純平 -FSE1訪問日時 2021/11/1410:00 -FSE1作業時間 10 -訪問FSE2氏名 川添 純一 L1 Engineer -FSE2訪問日時 2021/11/1410:00 -FSE2作業時間 10 [出荷指示] -----------------出荷指示#1----------------- -配送条件 TDD(日時指定) -納期 2021/11/14 10:00 指定 -配送先名 現地直送 -受取者名 小野 純平 -電話番号 080-9880-9929 -その他 同時引き取り有り:待機5h -依頼者名 佐竹類 ----------------------------------------------------------------------------- | 4 |  | [Details of work] Regarding the frequent occurrence of arcs, we replaced the X-ray tube and the high-voltage generator. After the work, we checked the operation of the device and confirmed that it works normally. After replacing the X-ray tube, some items did not fit within the standard values ​​in the focus calibration, so the scheduled time was significantly exceeded due to the investigation and re-replacement of the X-ray tube, which caused inconvenience. I_m very sorry. 989000086352:PHILIPS MRC 880 QTY:1 ZREP/Traceable 大阪(U615) 453566495091:1XRT SERVICE TOOL ASSY EXP QTY:1 ZSID 大阪(U615) 459800089031:Tool, XRT Strap Pad Assy QTY:1 ZSID 大阪(U615) 459800911271:TDMS V2 M.Det Class A FRU QTY:1 ZREP/Traceable 大阪(U615) | [12 empty] | 118649575 | Fixed | 1482624 |  |  | Investigation | Investigation | CTUDT00630085 | [5 empty] | Engineering | [6 empty] | Tube Arc | [5 empty] | Refresh |  |  | System issue | System | IC-HW | Tub Arc | IC-HW-Tub Arc | IC-HW-Tub Arc | CTUDT00630085 | Tub Arc | IC-HW | IC-HW-Tub Arc | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y
1590 |  |  | Japan | Independent administrative institution National Hospital Organization | 341013 | 44501 | 44 | 01DS: KKI: Freezes when shooting Prep >> Recovers after 5 seconds and can shoot | Work Order #: WO-04698358 02-Nov-2021 Yamamoto, Yuko Regulatory Questions : N/A Problem Description by Engineer: Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer: N/A ----------------------------------------------------------------------------- Work Order #: WO-04667805 02-Nov-2021 Yamamoto, Yuko The system meets the specification for the performed service and is returned to use. Regulatory Questions : N/A Problem Description by Engineer: Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution: I will retrieve and analyze the log. November 1, 2021 I reported to the person in charge Mr. Takumi Yamamoto. Internal Comments: [CX related] Did you agree on how to complete the report with the person in charge of devices? <Select one of the following> Yes. I completed normal reporting that day. [Work report (paper) submission] Did you submit a work report on paper? <YES/ Q1. How did the examination finish? Select from the following: *Because the answer is No for Q1, not applicable Q2. Were there any mistaken findings (misdiagnosis, mistaken treatment) due to the problem? None. Q3. Were there any artifacts? None. Q4. Was there any unintended unnecessary exposure? None. ----------------------------------------------------------------------------- Work Order #: WO-04667773 Kimura, Kazumasa The system meets the specification for the performed service and is returned to use. Diagnostic performed by Engineer: It recovered after 5 seconds and a scan could be performed. Problem Description by Engineer: The system froze during a Prep scan. Internal Comments: [Subject] 01DS:KKI:Mr. Yamamoto: The system froze during a Prep scan >> 5 seconds later it recovered and he was able to run a scan. [Internal Comments] The log showed no hardware errors. The CT host was restarted first thing in the morning, but the cause could not be identified. Dispatched: 1 RC FSE Registration Information: Name of FSE 1 who visited the site: Yutaka Yamamoto FSE 1 visit date and time: 2021/11/0117:00 FSE 1 hours worked: 2 ----------------------------------------------------------------------------- | 4 |  | CT application was freeze about 5sec after SAS injection. | [12 empty] | 118636379 | In Progress | 1475816 |  |  | Investigation | Investigation | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | SW Freeze | [5 empty] | Refresh |  |  | System issue | System | Console-SW | NoEnoughInfo | Console-SW-NoEnoughInfo | Console-SW-NoEnoughInfo | NoEnoughInfo | NoEnoughInfo | bar: No
1591 |  |  | Peru | BANCO BBVA PERU | 40034 | 44523 | 47 | 40034 WHEN A SECOND IP IS CONFIGURED FOR REMOTE CONNECTION PURPOSES, YOU CANNOT ENTER THE LAN CONFIG, NOR … |  | 5 |  | Known issue, link to LAN Config can not open defect | [12 empty] | 118773692 | New | 1554584 |  |  | Opened | Opened | CTUDT00631055 | [5 empty] | Engineering | [6 empty] | LAN Config | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Lan Config | Console-SW-Lan Config | Console-SW-Lan Config | CTUDT00631055 | Lan Config | Console-SW | Console-SW-Lan Config | bar: No
1592 |  |  | Chile | Scotiabank Chile | 40008 | 44519 | 46 | CL1675-ERROR DE TUBO | Work Order #: WO-04815858 23-Nov-2021 [REDACTED] The system meets the specification for the performed service and is returned to use. Resolution: 30A/600V fuse replacement since it is open in cathode. Functional tests are conducted with no issues. Diagnostic performed by Engineer: Remote diagnosis is performed and a blown cathode fuse is found. Problem Description by Engineer : Problem description by engineer : Gantry error. Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected Results: No gantry errors. Patient Impact/Outcome: Equipment is down. User Impact: Equipment stopped. Current Software Version: 4.1.10 Parts Consumed : 2.00000,459800066152,FUSE, 30A/600V AJT30 ----------------------------------------------------------------------------- Work Order #: WO-04806757 [REDACTED] The system does not meet the specification for the performed service and is not returned to use. Problem Description by Engineer : Problem description by engineer : Gantry error. Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected Results: No gantry errors. Patient Impact/Outcome: Equipment is down. User Impact: Equipment stopped. Current Software Version: 4.1.10 Diagnostic performed by Engineer: An error appears in phase voltage in the gantry. With the customer_s help, a blown fuse is identified in the gantry. Fuse must be replaced. Internal Comments : Error # and/or description of error : S_GEN_FAULT_KPM_480_LINE_PHASES_MISSING Repair Action: - Replace fuse. - Set input voltage for the gantry. Follow Up Required: Error Number and/or Description: S_GEN_FAULT_KPM_480_LINE_PHASES_MISSING Repair Action: - Replace fuse. - Set input voltage for the gantry. Technician_s Name: [REDACTED] ----------------------------------------------------------------------------- | 2 |  | - Replace fuse. - Set input voltage for the gantry. 2.00000,459800066152,FUSE, 30A/600V AJT30 | [12 empty] | 118751009 | Closed | 1543055 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | Not a Defect | [5 empty] | Engineering | [6 empty] | FUSE | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Other | Gantry-HW-Other | Gantry-HW-Other | Not a Defect | TBD-NoEnoughInfo | Gantry-HW | Gantry-HW-Other | bar: No
1593 |  |  | China | 河南许昌市中心医院(传染病) | 40047 | 44505 | 44 | 02N/Ingenuity CT 728326/add the printer | BU Q&R Feedback: No SAP Customer Symptom Code: CENV Customer Symptom Code: Customer Environmental Concern Reason: Corrective Maintenance Engineer Failure Code: Environmental Cause Repair Activity Code: Configuration Change Problem Reported by Customer : 02N/Ingenuity CT 728326 Safety Questions : Case: 0118670402 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Weng Jinming Last updated by: Weng Jinming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Troubleshooting Action: add the printer Repair Action: add the printer Test/Inspection: Functional Check PASSED. Device returned to full functionality. Test Equipment: N/A ----------------------------------------------------------------------------- | 5 | [14 empty] | 118670402 | Fixed | 1495869 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [19 empty] | Refresh |  |  | System issue | bar: No
1594 |  |  | China | 广西藤县人民医院 | 40036 | 44505 | 44 | Application/Ingenuity CT 728326/ | BU Q&R Feedback: No Reason: Application Support Problem Reported by Customer : Application/Ingenuity CT 728326/ Next Step Action : 邓老师015177135531 来电 Next Step Action : 已联系应用丰川13004507022 跟进 Safety Questions : Case: 0118669635 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Gao Luqi Last updated by: Gao Luqi 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 3 | [14 empty] | 118669635 | New | 1495565 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [19 empty] | Refresh |  |  | System issue | bar: No
1595 |  |  | Poland | University Clinical Hospital in | 341001 | 44523 | 47 | problem with the appearance of research on the worklist | Work Order #: WO-04827835 Gerard Dabrowka The system meets the specification for the performed service and is returned to use Diagnostic performed by Engineer: In consultation with an IT specialist and an IT specialist from the hospital, we found that despite the computer was turned off, the host_s IP address responded to Ping. It means that there is another computer with the same IP in the hospital network that has been assigned to the Hospital LAN card. Problem Description by Engineer : Problem description by engineer : The worklist does not appear, patient data was entered manually. After performing the exam, it is not sent to PACS or is sent partially. When was the first occurrence? Within... : 22-Nov-2021 18:00:00 How often is issue occurring? : Intermittent Information to support the complaint handling process Customer Function/Role: Not provided Device Use: Not provided Expected and Unexpected results: Not provided Patient Impact/Outcome: Not provided User Impact: Not provided Current Software Version: Not provided Internal Comments : Error # and/or description of error : There is no evident error, the system is not connected to PRS. Repair Action : During the next PM visit planned for 30 Nov 2021, the IT specialist will assign a new IP for the CT scanner from the restricted list of addresses. Extra Internal comment field : Continuation during the PM visit. Resolution: IT technician from the hospital will locate and turn the computer off with duplicate IP address. ----------------------------------------------------------------------------- | 4 |  | IT technician from the hospital will locate and turn the computer off with duplicate IP address. | [12 empty] | 118769322 | Closed | 1548856 |  |  | Evaluation | Evaluation | [6 empty] | Others | [6 empty] | Worklist | [5 empty] | Refresh |  |  | System issue | bar: No
1596 |  |  | Poland | Government Agency. | 341010 | 44503 | 44 | Recognizable artifact on images due to faulty detector module;Replaced detector module and calibrated | Work Order #: WO-04719373 09-Nov-2021 Mr. Robert Kamola The system meets the specification for the performed service and is returned to use Problem Description by Engineer: Replacement of the damaged module in the detector. Resolution: Replacement of the damaged module in detector no. 34. Calibrations and tests after the repair. The fault was eliminated, the device works correctly. ----------------------------------------------------------------------------- Work Order #: WO-04700546 09-Nov-2021 Mr. Robert Kamola The system does not meet the specification for the performed service and is not returned to use Problem Description by Engineer: The report concerns image artifacts. Resolution: The report concerns artifacts on the image. Test scans and detector quality tests were performed - incorrect signal was detected on module no. 34. The detector module needs to be replaced. The device is out of order. ----------------------------------------------------------------------------- | 4 |  | Resolution: The report concerns artifacts on the image. Test scans and detector quality tests were performed - incorrect signal was detected on module no. 34. The detector module needs to be replaced. The device is out of order. 1.00000,453567528582,Kit, Torque Wrench (0-100FT-LBS) 1.00000,459800911271,TDMS V2 M.Det Class A FRU 1.00000,452250066171,TOOL KIT, COMPACT | [12 empty] | 118658881 | In Progress | 1488513 |  |  | Closed - Done | Closed - Done | CTUDT00632110 | [5 empty] | Engineering | [6 empty] | dms module | [5 empty] | Refresh |  |  | System issue | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | DMS Module | DMS-HW | DMS-HW-DMS Module | bar: No
1597 |  |  | Puerto Rico | HOSPITAL DR SUSONI INC | 40013 | 44530 | 48 | Patient support lubrication | Work Order #: WO-04881573 Ramos Rivera, Pablo The system meets the specification for the performed service and is returned to use Follow Up Required : lubricate patient support Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: patient support lubrication needed User Impact: none Patient Impact: none Current Software Version:4.1.10 Diagnostic performed by Engineer : Order lubrication kit, lubricate patient support ----------------------------------------------------------------------------- | 5 | [14 empty] | 118815263 | In Progress | 1571447 |  |  | Evaluation | Evaluation | E-125059 | [5 empty] | Engineering | [6 empty] | Lubrication | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | Lubricant | Couch-HW-Lubricant | Couch-HW-Lubricant | E-125059 | Lubricant | Couch-HW | Couch-HW-Lubricant | bar: No
1598 |  |  | Spain | Comercial Centro Medico | 40042 | 44511 | 45 | Rebuilder error appears again | Work Order #: WO-04743628 12-Nov-2021 [REDACTED] The system meets the specification for the performed service and is returned to use. Diagnostic performed by Engineer: Intermittent connection error with reconstructors. Problem Description by Engineer: Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Regulatory Questions : N/A Resolution: The connection switch is changed between the 3 computers of the CT console. I perform an operation test and the system is working. Parts Consumed: 1.00000,459800457501,SMC SMCGS803 GB Network Switch 8 port ----------------------------------------------------------------------------- | 3 |  | Resolution: The connection switch is changed between the 3 computers of the CT console. I perform an operation test and the system is working. Parts Consumed: 1.00000,459800457501,SMC SMCGS803 GB Network Switch 8 port | [12 empty] | 118699798 | Fixed | 1515241 |  |  | Closed - Done | Closed - Done | part replace | [5 empty] | Others | [6 empty] | Net Switch | [5 empty] | Refresh |  |  | System issue |  |  |  | - | - | part replace | 0 | bar: No
1599 |  |  | Spain | S.C.S. HOSP. GRAL. DE FUERTEVENTURA | 40020 | 44503 | 44 | Streaks appear in images when making skulls | Work Order #: WO-04691426 [REDACTED] The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled. Regulatory Questions : N/A Problem Description by Engineer: In some cranial protocol, a horizontal line shaped artifact is showing up across the image. Follow Up Required: We would need to go by and check with the phantom and by using the same cranial protocol if the artifacts are repeating. Check the Raw Data for these images to find out if it the issue is coming from the gantry. Diagnostic performed by Engineer: The images of the relevant patient are checked, and the artifact is observed in one of the series; this artifact is moving vertically along the series images. It seems to be caused by something external to the reconstruction circle. ----------------------------------------------------------------------------- | 3 |  | Check the Raw Data for these images to find out if it the issue is coming from the gantry. Diagnostic performed by Engineer: The images of the relevant patient are checked, and the artifact is observed in one of the series; this artifact is moving vertically along the series images. It seems to be caused by something external to the reconstruction circle. | [12 empty] | 118657714 | Closed | 1484487 |  |  | Investigation | Investigation | TBD-NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Image Artifact | [5 empty] | Refresh |  |  | System issue | System | TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo | bar: No
1600 |  |  | China | Fujian Provincial Hospital | 345011 | 44518 | 46 | 02N/Ingenuity Core128/stop 0x00000ed | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 02N/Ingenuity Core128/Elite China SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support Safety Questions:Case: 0118734289 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Weng Jinming Last updated by: Weng Jinming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer_s problems: Host computer blue screen Detailed description of Customer_s problems or malfunctions: Service requester is not on site. It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: Stop 0x00000ed Remote troubleshoot and test result: Contact engineer Ding Shi duo, suspect software, memory.。。problem, hard disc has been replaced before PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Suggest engineer to check host computer memory, graphic card, reinstall software on site, teach customer to use Philips WeChat to report for repair Next Step Action: Ding shiduo ----------------------------------------------------------------------------- | 2 |  | $Q7:Confirmed reported problem. the host blue screen. $Q8: on site Reviewed the error log , Performed visual inspection and During the BIOS hardware test, found the memory is abnormal. $Q9: Clean Computer host, unplug Computer host memory. | [12 empty] | 118734289 | In Process | 1537109 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Others | [6 empty] | Memory | [5 empty] | Refresh |  |  | System issue |  |  |  | - | - | 0 | 0 | bar: No
1601 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 337000 | 44510 | 45 | Hava kalibrasyonunu geçmiyor | Work Order #: WO-04734527 16-Nov-2021 Osman Alpay The system meets the specification for the performed service and is returned to use Diagnostic performed by Engineer : Air Calibration fails with error. Problem Description by Engineer : Problem description by engineer : Air cal could be completed. Information to support the complaint handling process Customer Function/Role: Not provided Device Use: Not provided Expected and Unexpected results: Not provided Patient Impact/Outcome: Not provided User Impact: Not provided Current Software Version: Not provided. Resolution : We opened the gantry cover to check DMS. We found contrasting agent deposits on the UHR Grid and detectors. System was returned to use in functional condition after cleaning the contrasting agent. Also, gantry room ar conditioners were faulty and the ambient temperature was high. We advised the user that they should have the a/c unit repaired. Tools & Equipment : MULTIMETER,2022-02-05 00:00:00 ----------------------------------------------------------------------------- Work Order #: WO-04734382 Salih Mendi The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Internal Comments : Error # and/or description of error : Air cal could be completed. Repair Action : Air cal could be completed. site visit needed for further inspection. There is no artefact on images also. Diagnostic performed by Engineer : Air cal could be completed. Problem Description by Engineer : Problem description by engineer : Air cal could be completed. Information to support the complaint handling process Customer Function/Role: Not provided Device Use: Not provided Expected and Unexpected results: Not provided Patient Impact/Outcome: Not provided User Impact: Not provided Current Software Version: Not provided. Resolution : We remotely supported the user for the air calibration problem. Mylar was cleaned. When starting the air calibration, object warning inside the imaging area error occurred. Site visit is needed for a detailed inspection. The user reported that there were no artefacts and the system was up and running. Follow Up Required : Error # and/or description of error : Air cal could be completed. Repair Action : Air cal could be completed. site visit needed for further inspection. There is no artefact on images also. Technician Name : Salih Mendi ----------------------------------------------------------------------------- | 3 |  | Resolution : We remotely supported the user for the air calibration problem. Mylar was cleaned. When starting the air calibration, object warning inside the imaging area error occurred. Site visit is needed for a detailed inspection. The user reported that there were no artefacts and the system was up and running. | [12 empty] | 118692132 | Closed | 1506220 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Others | [6 empty] | Contrast Intrusion | [5 empty] | Refresh |  |  | System issue | bar: No
1602 |  |  | United Kingdom | STOCKPORT NHS FOUNDATION TRUST | 40016 | 44514 | 45 | Replace harmonized phantom with the old blue style system phantom. | Internal Remarks : Email from Warren Thorpe dated the 12th Nov 2021 Hospital Name: Stepping Hill Hospital, Stockport Customer Contact Name (not FSE or Philips Employee) : Karen Andrew Contact Telephone number:01614194782 Serial/Pin/Tech ID: 40016 System out of action? No System Priority: 4 - Intermittent Problem Detail of work/Problem Description: Replace harmonized phantom with the old blue style system phantom after the customer complained about alignment issues with the new ones. Replace phantom and run calibrations. Customer Temperature: Green No. of FSE/s to fit: 1 Hours to fit (excluding travel time): 3 Required FSE skill level: 3 - Adequate/Experienced Q1 Was the equipment in ‘Clinical use’ at the time the issue was discovered?: No Q2 Was any patient or user harmed?: No Parts TPR Review needed: Choose an item. TPR Reviewer: Click here to enter text. Part number: 453566502022 Description: Phantom Holder Assy, Carbon Top Quantity: 1 FCO part: Choose an item. PUDO: Click here to enter text. TPR Review needed: Choose an item. TPR Reviewer: Click here to enter text. Part number: 453567135961 Description: EXTENDED PHANTOM PERFORMAN KIT Quantity: 1 FCO part: Choose an item. PUDO:Click here to enter text. Delivery detail if to site: Click here to enter text. Any other relevant info: Customer has complained, so knows that we are ordering a new system phantom. Date and time needs to be arranged. Best Regards, Warren Thorpe Customer comment : Warren to attend on 7/12 - Confirmed with Karen Andrew ----------------------------------------------------------------------------- Work Order #: WO-04758834 Warren Thorpe Regulatory Questions : N/A Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 4 | [14 empty] | 118711553 | In Progress | 1521512 |  |  | Evaluation | Evaluation | H-Phantom | [5 empty] | Engineering | [6 empty] | H-Phantom | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | H-Phantom | Couch-HW-H-Phantom | Couch-HW-H-Phantom | H-Phantom | H-Phantom | bar: No
1603 |  |  | United States | Central Vermont Medical Center | 40024 | 44524 | 47 | Images not sending to PACs | Work Order #: WO-04882078 02-Dec-2021 Sangster, Robert The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Diagnostic performed by Engineer : Made adjustments to the network settings, set two of the queue nodes (AMBRA and iSite) to series level instead of image. Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : Working with site and national support to resolve network issues. ----------------------------------------------------------------------------- Work Order #: WO-04841061 29-Nov-2021 Sangster, Robert The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : Issue with pacs reviewed. Bug report gathered and sent for review. Follow Up Required : Adjust dicom settings to send series to pacs ----------------------------------------------------------------------------- | 4 |  | Diagnostic performed by Engineer : Issue with pacs reviewed. Bug report gathered and sent for review. Follow Up Required : Adjust dicom settings to send series to pacs under engineering investigation | [12 empty] | 118781092 | In Progress | 1558724 |  |  | Evaluation | Evaluation | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS Transfer | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y
1604 |  |  | China | Dongguan City Integrated Traditional Chinese and Western medicine Hospital | 345013 | 44502 | 44 | 02N/Ingenuity Core128/Elite China/Check equipment | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ Safety Questions:Case: 0118651564 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingsong Last updated by: Lu Mingsong 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer_s problems: The equipment_s warranty is about to expire. The customer asked the field service engineer to go on site to check the equipment_s condition. Detailed description of Customer_s problems or malfunctions: The equipment_s warranty is about to expire. The customer asked the field service engineer to go on site to check the equipment_s condition. It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: N/A Remote troubleshoot and test result: N/A PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Go on site and check the equipment_s condition. We recommend spare parts 12NC and names: N/A Memo of Customer request: None SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support External Remarks: 03.11.2021 11:53:23 UTC (86034125) Service details: inspection before machine warranty expires. Service outcome: the machine is checked and found no abnormality. The machine is working normally. Other important issues: None Internal Remarks : 05.11.2021 15:42:16 UTC (86034125) "RDI Accuracy:Yes" $Q1: () $Q2: () $Q3: () $Q4: () $Q7: (Confirmation & Description of problem) Confirmed reported problem, customer requires check of the machine situation before the warranty expires. $Q9: (Repair steps) check the machine as the customer required, no abnormality is found. $Q10: (Test and inspection data refer to PA . Device returned to full functionality ) $Q11: Calibrated Tools name ( refer to PA ) S/N ( ) Next Calibration date ( refer to PA ) For Mainland China CV / XR, CT engineers only. Please fill in: ^Current Tube installed date: N/A ^Current Tube usage: N/A External Remarks: 08.11.2021 02:25:39 UTC (86001504) Service details: inspection before machine warranty expires. Service outcome: the machine is checked and found no abnormality. The machine is working normally. Other important issues: None Internal Remarks: 08.11.2021 02:25:42 UTC (86001504)RDI Accuracy:Yes $Q7: (Confirmed the reported problem, the sales asked to checked the device status) $Q8: (no fault found by other FSE, double confirmed) $Q9: (no action required) $Q10: (Test and inspection data refer to N/A, PA by other FSE. Device returned to full functionality.) $Q11: $Calibrated Tools name (N/A) S/N (N/A) Next Calibration date (N/A) ----------------------------------------------------------------------------- | 5 | [14 empty] | 118651564 | In Process | 1483697 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [19 empty] | Refresh |  |  | System issue | bar: No
1605 |  |  | United States | UF Health Leesburg Hospital | 40022 | 44504 | 44 | Result would jump when planning a scan with multiple results |  | #N/A |  | Known issue: result jump | [14 empty] | 1524653 |  |  | Investigation | Investigation | CTUDT00632986 | [5 empty] | Engineering | [6 empty] | Result Jump | [5 empty] | Refresh |  |  | System issue | System | Console-SW | ResultSelectionJump | Console-SW-ResultSelectionJump | Console-SW-ResultSelectionJump | CTUDT00632986 | ResultSelectionJump | Console-SW | Console-SW-ResultSelectionJump | bar: No
1606 |  |  | China | Hebei Province Geriatric Hospital | 345015 | 44502 | 44 | 05K For IB/Core128/Consult Manufacture date | BU Q&R Feedback: No Reason: Customer Information Problem Reported by Customer:consult Manufacture date Safety Questions:Case: 0118649876 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Pan Chuchu Last updated by: Pan Chuchu 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action: Case# 0118649876 Hebei Province Geriatrics Hospital, Ingenuity Core128, EQ# 84722645, customer want to query the device_s manufacturer date The caller is teacher Zhang at equipment dept 031189892511 Next Step Action: Notified ASP Jiunuo Business to follow up with customer ----------------------------------------------------------------------------- | 3 | [14 empty] | 118649876 | Closed | 1482740 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [19 empty] | Refresh |  |  | System issue | bar: No
1607 |  |  | United States | Philips Medical Capital | 40053 | 44508 | 45 | PACS troubleshooting | Work Order #: WO-04725430 16-Nov-2021 Fairchild, Andrew The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Diagnostic performed by Engineer : Unable to log into Portal system with service user names/passwords. Follow-up created with Portal site ID. Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: 4.1.10 ----------------------------------------------------------------------------- Work Order #: WO-04723289 08-Nov-2021 Fairchild, Andrew The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: 4.1.10 Diagnostic performed by Engineer : Troubleshooted multiple configurations and ran tests with logging turned up to debug layer. Data and logs gathered to be forwarded to other support. Equipment downtime needs to be scheduled for additional troubleshooting when support available. Follow Up Required : Continue troubleshooting with all involved parties that work with the PACS system. ----------------------------------------------------------------------------- | 5 |  | Troubleshooted multiple configurations and ran tests with logging turned up to debug layer. Data and logs gathered to be forwarded to other support. Equipment downtime needs to be scheduled for additional troubleshooting when support available. Continue troubleshooting with all involved parties that work with the PACS system. | [12 empty] | 118682975 | Closed | 1501066 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS Transfer | [5 empty] | Refresh |  |  | System issue | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y
1608 |  |  | United States | Leesburg Regional Medical Center | 40022 | 44506 | 44 | 0711CFXXX HIS/RIS not populating | Integration : Confirmation of order 54403986 saved ----------------------------------------------------------------------------- Work Order #: WO-04714094 Mcginnis, Mike The system meets the specification for the performed service and is returned to use Internal Comments : Error # and/or description of error : his/ris only showing yesterdays patients Repair Action : I had him check the by date filter and it was set for yesterday. He cleared that filter and it is now pulling current scheduled patients Diagnostic performed by Engineer : Angelo said the his/ris only shows yesterdays patient list. I had him check the by date filter and it was set for yesterday. He cleared that filter and it is now pulling current scheduled patients Problem Description by Engineer : Problem description by engineer : HIS/RIS only showing yesterdays patients When was the first occurrence? Within... : 06-Nov-2021 How often is issue occurring? : First occurrence Malfunction area : Console Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: HIS/RIS only showing yesterdays patients Patient Impact/Outcome: none User Impact: Could not use his/ris Current Software Version: Not provided Resolution : I had him check the by date filter and it was set for yesterday. He cleared that filter and it is now pulling current scheduled patients ----------------------------------------------------------------------------- | 3 |  | Repair Action : I had him check the by date filter and it was set for yesterday. He cleared that filter and it is now pulling current scheduled patients Diagnostic performed by Engineer : Angelo said the his/ris only shows yesterdays patient list. I had him check the by date filter and it was set for yesterday. He cleared that filter and it is now pulling current scheduled patients Problem Description by Engineer : Problem description by engineer : | [12 empty] | 118675011 | Closed | 1496538 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Application | [6 empty] | HIS/RIS | [5 empty] | Refresh |  |  | System issue | bar: No
1609 |  |  | China | Panjin City DaWa District People_s Hospital | 345040 | 44502 | 44 | 02N/Ingenuity Core128/Elite China/add node | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 02N/Ingenuity Core128/Elite China Safety Questions:Case: 0118649455 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Jinbo Yang Last updated by: Jinbo Yang 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Dispatched for On site Support T2 Activities:Problem Description: Summary of customer_s problems: Other types of error. Detailed description of Customer_s problems or malfunctions: Add node It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: Add node Remote troubleshoot and test result: Add nodes. Need an engineer to go to the field PRS TR status: It is not through, ask engineer for maintenance on site. Repair Action: Suggest onsite service procedures: Add node SAP Transfer:Case Transferred to SAP External Remarks: 02.11.2021 11:33:43 UTC (86038273) service details: need to add network nodes. Add nodes. service result: device returns to normal working condition after service Internal Remarks : 02.11.2021 11:33:45 UTC (86038273) RDI Accuracy:Yes $Q7: (Confirmed reported problem: add dicom device) $Q8: (add dicom device) $Q9:(add dicom device) $Q10: (Test and Inspection data refer to pa & Device returned to full functionality) $Q11:Calibrated Tools name (NA) ----------------------------------------------------------------------------- | 3 | [14 empty] | 118649455 | Closed | 1482538 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [19 empty] | Refresh |  |  | System issue | bar: No
1610 |  |  | United States | Clay County Medical Center | 40003 | 44503 | 44 | Physics phantom wires have lost tension. | Work Order #: WO-04695066 05-Nov-2021 Stauffer Jeffrey The system meets the specification for the performed service and is returned to use Resolution : Ordered new phantom and picked up at UPS. Phantom will be delivered to customer and used on next PM in December 2021. Diagnostic performed by Engineer : Looked at phantom during last service call and saw that tension on wires in physics phantom are loose. Problem Description by Engineer : Customer Function/Role: Philips FSE Device Use: General CT Exams Expected and Unexpected Results: Unexpected - tension wire in physics phantom is not tight. User Impact: None Patient Impact: None Current Software Version:4.10 Parts Consumed : 1.00000,459801415171,HARMONIZED PHANTOM PHYSICS SECTION ----------------------------------------------------------------------------- | 3 |  | Resolution : Ordered new phantom and picked up at UPS. Phantom will be delivered to customer and used on next PM in December 2021. Diagnostic performed by Engineer : Looked at phantom during last service call and saw that tension on wires in physics phantom are loose. 1.00000,459801415171,HARMONIZED PHANTOM PHYSICS SECTION | [12 empty] | 118660267 | Closed | 1489660 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | Physics Phantom | [5 empty] | Engineering | [6 empty] | Physics Phantom | [5 empty] | Refresh |  |  | System issue | System | Couch-HW | Physics Phantom | Couch-HW-Physics Phantom | Physics Phantom-Couch-HW-Physics Phantom | Physics Phantom | Physics Phantom | bar: No
1611 |  |  | China | Beijing City Hospital of Traditional Chinese and Western Medicine | 40032 | 44501 | 44 | 02N/Ingenuity CT 728326/stand by | BU Q&R Feedback: No Reason: Customer Relations SAP Transfer:Dispatched for On site Support Problem Reported by Customer: stand by Next Step Action: 013811972406 Yu called Safety Questions:Case: 0118640561 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Qian Wang Last updated by: Qian Wang 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action : fan guochao SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- | 5 | [14 empty] | 118640561 | New | 1479550 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [19 empty] | Refresh |  |  | System issue | bar: No
1612 |  |  | China | DaKuai Town Health, Fengquan District, Xinxiang City, Henan Province | 345009 | 44524 | 47 | 02N/Ingenuity Core128/Elite China/Gantry Unable to start | BU Q&R Feedback: No Reason: Corrective Maintenance SAP Transfer:Dispatched for On site Support Problem Reported by Customer: Gantry Unable to start Safety Questions:Case: 0118779357 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Zhang Caiqin Last updated by: Zhang Caiqin 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Case Transferred to SAP External Remarks: 25.11.2021 11:28:11 UTC (86038950)"SUAT. Ready for Clinical Use local time" : "2021-11-25 15:00" External Remarks: 26.11.2021 12:26:25 UTC (86038950) Service details: No key box when turned on, scan failed. On-site check found MKD_s NC port open. Service outcome: after replacing MKD contactor, the equipment is back to normal. Other significant matters: none ----------------------------------------------------------------------------- | 2 |  | Service details: No key box when turned on, scan failed. On-site check found MKD_s NC port open. Service outcome: after replacing MKD contactor, the equipment is back to normal. | [12 empty] | 118779357 | Closed | 1557847 |  |  | Investigation - Pending Translation | Investigation - Pending Translation | PartReplacement | [5 empty] | Others | [6 empty] | MKD Contactor | [5 empty] | Refresh |  |  | System issue |  |  |  | - |  | PartReplacement | 0 | bar: No
1613 |  |  | United States | Clay County Medical Center | 40003 | 44503 | 44 | Left Control Panel on Gantry needs adjustment. | Integration : Confirmation of order 54389280 saved ----------------------------------------------------------------------------- Work Order #: WO-04693042 03-Nov-2021 Stauffer Jeffrey The system meets the specification for the performed service and is returned to use Diagnostic performed by Engineer : Pushed up button and it did not work. Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: General CT exams Expected and Unexpected Results: Left control panel does not work properly. User Impact: None Patient Impact: None Current Software Version:4 Resolution : Replaced left control panel. ----------------------------------------------------------------------------- | 3 | [14 empty] | 118658916 | Closed | 1486176 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632181 | [5 empty] | Engineering | [6 empty] | GCP | [5 empty] | Refresh |  |  | System issue | System | Gantry-HW | Gantry Panel | Gantry-HW-Gantry Panel | Gantry-HW-Gantry Panel | CTUDT00632181 | Gantry Panel | Gantry-HW | Gantry-HW-Gantry Panel | No: No; y: 0; y: 0; y: 0; y: 0; y: 0; No: No
1614 |  |  | China | 河南省新乡市凤泉区大块镇卫生院 | 345009 | 44501 | 44 | 02N/Ingenuity Core128/Elite China/Customer consultation | BU Q&R Feedback: No SAP Customer Symptom Code: HANG Customer Symptom Code: Hang/Crash Reason: Corrective Maintenance Engineer Failure Code: Wear-out/Breakage/Mechanical Failure Repair Activity Code: Instructed Customer Problem Reported by Customer : 02N/Ingenuity Core128/Elite China T2 Activities : Problem Description: 客户咨询是否能够把64排物理结构探测器暂时锁住两排。 Troubleshooting Action: 告诉客户探测器物理结构无法锁住，并通知工程师郭海滨跟进。 Repair Action: 告诉客户探测器物理结构无法锁住，并通知工程师郭海滨跟进。机器无故障。 Test/Inspection: customer confirm can scan patient and image ok.functional Check PASSED. Device returned to full functionality. Test Equipment: N/A Safety Questions : Case: 0118637059 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Bo Last updated by: Wang Bo 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A ----------------------------------------------------------------------------- | 3 | [14 empty] | 118637059 | Closed | 1478558 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [19 empty] | Refresh |  |  | System issue | bar: No
1615 |  |  | Argentina | INSTITUTO NACIONAL DE | 40027 | 44501 | 44 | SE BLOQUEO USUARIO DEBIDO A INTRODUCIR MAL LA CONTRASEÑA | Integration:Confirmation of order 54378031 saved ----------------------------------------------------------------------------- Work Order #: WO-04667653 [REDACTED] The system meets the specification for the performed service and is returned to use. Problem Description by Engineer : Problem description by engineer : THE USER GOT BLOCKED BECAUSE THE PASSWORD WAS ENTERED WRONGLY. When was the first occurrence? Within... : 01-Nov-2021 08:00:00 How often is issue occurring? : First occurrence Malfunction area : Console Information to support the complaint handling process Customer Function/Role: BioMed Device Use: Diagnostic Expected and Unexpected Results: It is working properly. Patient Impact/Outcome: Delay in attention User Impact: Delay in work flow. Current Software Version: 4.1.10 Diagnostic performed by Engineer: We contacted the user, they said that a field service engineer visited the site and helped them to restore the password. Equipment functioning correctly. Internal Comments : Error # and/or description of error : Blocked password. Repair Action: We contacted the user, and they said that a field service engineer visited the site and helped them to restore the password. Equipment functioning correctly. Resolution: Serviced remotely. ----------------------------------------------------------------------------- | 2 | [14 empty] | 118636312 | Closed | 1476930 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [19 empty] | Refresh |  |  | System issue | bar: No
1616 |  |  | United States | Mount Grant General Hospital | 40012 | 44503 | 44 | surview image questions | Integration : Confirmation of order 54391354 saved ----------------------------------------------------------------------------- Work Order #: WO-04696344 Curry, Dyan The system meets the specification for the performed service and is returned to use Problem description by Clinical Support Specialist : Customer Function/ Role: technologist Device Use: in use Expected and Unexpected Results: surview image labels User Impact: none Patient Impact: none Current Software version: 4.1.10 Diagnostic performed by Engineer : FSE calling for assistance. Dual surview labeling both images as 100-1. PACS not accepting 2 images labeled the same. Did test scan on KIVM 4.1. Surviews are both labeled 100-1. Advised FSE that is the way system is set up. There are no options to change from user /application side. Advised FSE to check with PACS admin , there could be a setting. Also to contact his NSS for advice. ----------------------------------------------------------------------------- | 4 | [14 empty] | 118661233 | Closed | 1486082 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Application | [6 empty] | Q&A | [5 empty] | Refresh |  |  | System issue | bar: No
1617 |  |  | United States | Adventhealth Zephyrhills | 300106 | 44530 | 48 | 1302CDY72 slow transferring images to pacs | Work Order #: WO-04881472 02-Dec-2021 Watson Adam The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Problem description by engineer : slow transferring images to pacs Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: slow transferring images to pacs Patient Impact/Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer Current Software Version: 4.1.7.25030 Resolution : When engineer arrived onsite, images were transferring normally. Customer said the problem happens on larger studies. Checked bug report that customer has made, and nothing showed in the logs. Ran Tracert on both hospital CT’s and both had identical issues. When tracert was sent to the default gateway, the trace would happen immediately. When the trace was ran to the PACS and BackUp PACS, there were time out errors before successfully completing the trace. Engineer talked with next level support and they agreed that the problem appears to be a hospital switch. Instructed customer reach out to their IT department to troubleshoot network switches. No confirmed problem with the Philips CT. ----------------------------------------------------------------------------- Work Order #: WO-04879925 Sheffield, Scott The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Problem description by engineer : slow transferring images to pacs Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: No patient involved Expected and Unexpected results: slow transferring images to pacs Patient Impact/Outcome: No further information is available as per the customer User Impact: No further information is available as per the customer Current Software Version: 4.1.7.25030 Diagnostic performed by Engineer : Russ said that this has been an ongoing problem for about 4 months. It is intermittent problem does not happen all the time. I guided Russ thru olevel to reset spooler and reboot host computer. The images transferred fine afterwards. I reviewed the bug report and did not find any problems. Follow Up Required : Error # and/or description of error : none Repair Action : Previous cases the FSE has noted problem is with the customer network. Russ told me the pacs ip address is 10.189.15.12 I compared the lanconfig settings for pacs for this scanner and the other scanner that Mark told me and they looked the same. system1 site # 5468842 transfers slow C:\>ping 10.189.15.12 ping 10.189.15.12 Pinging 10.189.15.12 with 32 bytes of data: Reply from 10.189.15.12: bytes=32 time<1ms TTL=254 Reply from 10.189.15.12: bytes=32 time<1ms TTL=254 Reply from 10.189.15.12: bytes=32 time<1ms TTL=254 Reply from 10.189.15.12: bytes=32 time<1ms TTL=254 Ping statistics for 10.189.15.12: Packets: Sent = 4, Received = 4, Lost = 0 (0% loss), Approximate round trip times in milli-seconds: Minimum = 0ms, Maximum = 0ms, Average = 0ms C:\>tracert 10.189.15.12 tracert 10.189.15.12 Tracing route to 10.189.15.12 over a maximum of 30 hops 1 * * * Request timed out. 2 <1 ms <1 ms <1 ms 10.189.15.12 Trace complete. scanner 54648530 transfers fine C:\>tracert 10.189.15.12 tracert 10.189.15.12 Tracing route to 10.189.15.12 over a maximum of 30 hops 1 * * * Request timed out. 2 <1 ms <1 ms <1 ms 10.189.15.12 Trace complete. Technician Name : Sheffield, Scott Internal Comments : Error # and/or description of error : none Repair Action : Previous cases the FSE has noted problem is with the customer network. Russ told me the pacs ip address is 10.189.15.12 I compared the lanconfig settings for pacs for this scanner and the other scanner that Mark told me and they looked the same. system1 site # 5468842 transfers slow C:\>ping 10.189.15.12 ping 10.189.15.12 Pinging 10.189.15.12 with 32 bytes of data: Reply from 10.189.15.12: bytes=32 time<1ms TTL=254 Reply from 10.189.15.12: bytes=32 time<1ms TTL=254 Reply from 10.189.15.12: bytes=32 time<1ms TTL=254 Reply from 10.189.15.12: bytes=32 time<1ms TTL=254 Ping statistics for 10.189.15.12: Packets: Sent = 4, Received = 4, Lost = 0 (0% loss), Approximate round trip times in milli-seconds: Minimum = 0ms, Maximum = 0ms, Average = 0ms C:\>tracert 10.189.15.12 tracert 10.189.15.12 Tracing route to 10.189.15.12 over a maximum of 30 hops 1 * * * Request timed out. 2 <1 ms <1 ms <1 ms 10.189.15.12 Trace complete. scanner 54648530 transfers fine C:\>tracert 10.189.15.12 tracert 10.189.15.12 Tracing route to 10.189.15.12 over a maximum of 30 hops 1 * * * Request timed out. 2 <1 ms <1 ms <1 ms 10.189.15.12 Trace complete. ----------------------------------------------------------------------------- | 3 |  | Customer said the problem happens on larger studies SW Version: SP1 FCO72800754 Russ said that this has been an ongoing problem for about 4 months. It is intermittent problem does not happen all the time. I guided Russ thru olevel to reset spooler and reboot host computer. The images transferred fine afterwards. I reviewed the bug report and did not find any problems. Engineer talked with next level support and they agreed that the problem appears to be a hospital switch. Instructed customer reach out to their IT department to troubleshoot network switches. No confirmed problem with the Philips CT. | [12 empty] | 118814165 | Fixed | 1571314 |  |  | Evaluation |  | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS Transfer | [5 empty] | Refresh |  |  | Upgrade | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No | 0 | 0 | 0 | 0 | 0 | No | y | 0 | 0 | 0 | 0 | y | y | y | 0 | y
1618 |  |  | United States | Lakeland Regional Medical Center | 300143 | 44520 | 46 | CT Box flashing. | [16 empty] | 118753770 | Fixed | 1543896 |  |  | Closed - No Further Investigation | [20 empty] | Non-Refresh |  |  | Upgrade | bar: No
1619 |  |  | United States | Lakeland Regional Medical Center | 300143 | 44520 | 46 | CT box flashing; Replaced CT box buttons and enabled button. | Work Order #: WO-04811947 22-Nov-2021 Watson Adam The system meets the specification for the performed service and is returned to use Parts Consumed : 1.00000,453567027202,Keypad assy, CT Box 1.00000,453566496941,RELAY,24VDC PLUGIN W/COVER Resolution : Replaced CT box buttons and enable button issue cleared. Attempted to perform a scan and logs showed horizontal motor relay errors. Replaced relay and performed test scans with no further issues. Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: ----------------------------------------------------------------------------- Work Order #: WO-04811517 21-Nov-2021 Sher Mark The system does not meet specification and is out of use Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: CT Box flashing User Impact: None Patient Impact: None Current Software Version: Diagnostic performed by Engineer : parts delay expedited Adam Watson will be taking over call Follow Up Required : replace button board ----------------------------------------------------------------------------- Work Order #: WO-04810589 21-Nov-2021 Sher Mark The system does not meet specification and is out of use Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: CT Box flashing User Impact: None Patient Impact: None Current Software Version: Diagnostic performed by Engineer : found system reports stuck command button ordered button board Follow Up Required : Install ordered parts ----------------------------------------------------------------------------- Work Order #: WO-04810495 Boyd Craig The system does not meet the specification for the performed service and is not returned to use Problem Description by Engineer : Customer Function/Role: Tech Device Use: Diagnostic Expected and Unexpected Results: CT Box flashing User Impact: None Patient Impact: None Current Software Version: Diagnostic performed by Engineer : PRS accessed. ULV: S_COMMAND_BUTTON_STUCK_ERROR, AEnb Follow Up Required : FSE needed on site to evaluate CT Box Buttons. S_COMMAND_BUTTON_STUCK_ERROR , AEnb. Recommendations; Try to unstick Scan Enable button on CT Box. Replace parts if needed. CT Box 453567027195, CRC- Power Distribution Module 459801032622 ----------------------------------------------------------------------------- | [15 empty] | 118753804 | Closed | 1543400 |  |  | Closed - Done | [20 empty] | Non-Refresh |  |  | Upgrade | bar: No
1620 |  |  | Netherlands | Stichting Rijnstate Ziekenhuis | 52101 | 44509 | 45 | Systeem wil niet scannen ganrtry can not initialize. | Work Order #: WO-04726691 09-Nov-2021 EDDY ELFRING The system meets the specifications for the performed service and was returned to use. Follow Up Required: send parts attn Frank Nieuwenhuijs Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution: Relay 24 DC replaced. ----------------------------------------------------------------------------- Work Order #: WO-04726599 JOHNNY BASTIAENS The system does not meet the full specifications for the performed service and was returned to use with limitations as accepted by the Customer, Follow-up work scheduled. Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnosis performed by Engineer : A restart did not solve the problem. A field service engineer will go on site. Follow-up Required : Diagnose. ----------------------------------------------------------------------------- Work Order #: WO-04734867 09-Nov-2021 Parts Consumed : 2.00000,453566496941,RELAY,24VDC PLUGIN W/COVER 10.00000,300006737161,ECG adaptor FRU ----------------------------------------------------------------------------- | [15 empty] | 118686034 | Closed | 1501932 |  |  | Evaluation | [20 empty] | Non-Refresh |  |  | Upgrade | bar: No
1621 |  |  | Netherlands | Stichting Rijnstate Ziekenhuis | 52101 | 44509 | 45 | Gantry can not initialize | Integration:Confirmation of order 54413401 saved ----------------------------------------------------------------------------- Work Order #: WO-04726624 JOHNNY BASTIAENS The system meets the specifications for the performed service and was returned to use. Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer: Duplicate case 0118686034 ----------------------------------------------------------------------------- | [15 empty] | 118686038 | Closed | 1501912 |  |  | Closed - Duplicate | [20 empty] | Non-Refresh |  |  | Upgrade | bar: No
1622 |  |  | United States | Lakeland Regional Medical Center | 320474 | 44504 | 44 | Head Arm Rest needs replacing | Internal Remarks : approved: per the attachment added to the case the customers contract should include accessory/consumables. ----------------------------------------------------------------------------- Work Order #: WO-04702339 08-Nov-2021 Watson Adam The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Resolution : Delivered pad to customer Parts Consumed : 1.00000,459800042241,Cushion-Head/ Armrest ----------------------------------------------------------------------------- | [15 empty] | 118665982 | Incomplete | 1493402 |  |  | Closed - Non-Complaint | [20 empty] | Non-Refresh |  |  | Upgrade | bar: No
1623 |  |  | United States | Philips Medical Capital | 40055 | 44536 |  | Needs cable for in room monitor replaced | Work Order #: WO-04917752 Maier, Donnie The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Follow Up Required : Customer explained that the trigger cable for in room physiological monitor needs replaced (intermittently working). Looked up trigger cable part number in CHA 455012301901 - ECG Monitor Trigger 7M Ca (10 on Shelf in Parts-O-Pedia) after confirming with PFSE Rich Turley that is indeed what he is needing ordered to the site. I will order to site (verified address and POC with PFSE) and create onsite for PFSE Turley as requested. Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: Diagnostic Expected and Unexpected Results: Trigger cable for in room monitor needs replaced User Impact: No Patient Impact: No Current Software Version: 4.1.10.25210 Diagnostic performed by Engineer : Customer explained that the trigger cable for in room physiological monitor needs replaced (intermittently working). Looked up trigger cable part number in CHA 455012301901 - ECG Monitor Trigger 7M Ca after confirming with PFSE Rich Turley that is indeed what he is needing ordered to the site. I will order to site (verified address and POC with PFSE) and create onsite for PFSE Turley as requested. ----------------------------------------------------------------------------- Work Order #: WO-04917909 Turley Rich Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: Diagnostic Expected and Unexpected Results: Trigger cable for in room monitor needs replaced User Impact: No Patient Impact: No Current Software Version: 4.1.10.25210 Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 4 | [14 empty] | 118844487 | Closed | 1590033 |  |  | Opened | Opened | [6 empty] | Engineering | [6 empty] | ECG trigger cable | [5 empty] | IB Exist | 455012301901 | ECG Monitor Trigger 7M Ca | bar: No
1624 |  |  | China | 息烽县中医医院 | 345014 | 44533 |  | 05k for IB/Ingenuity Core128/Elite China/CT box | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : 05k for IB/Ingenuity Core128/Elite China/ Next Step Action : FSE Hu Ke 018585068636 call in, SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support Next Step Action : 已回访成功 Safety Questions : Case: 0118832558 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Weng Jinming Last updated by: Weng Jinming 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 工程师胡克替客户报修，需要更换CT box 设备本身技术问题： 是。 Troubleshooting Action: 远程诊断与测试结果： 工程师胡克替客户报修，需要更换CT box PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议工程师到场检查更换CT box 建议备件12NC及名称： CT box：459801704602 ----------------------------------------------------------------------------- | 3 | [14 empty] | 118832558 | Closed | 1585998 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | Engineering | [6 empty] | CT-Box | [5 empty] | IB Exist | 459801705022 | CT BOX ASSEMBLY_BEST | bar: No
1625 |  |  | United States | Philips Medical Capital | 40005 | 44533 |  | Error Gantry cant comply. | Work Order #: WO-04906131 Greer Daniel The system does not meet the specification for the performed service and is not returned to use Follow Up Required : Error # and/or description of error : gantry can not comply error. Repair Action : Check tape switch_s and foot switch_s. Then slip ring data flow. Technician Name : Greer Daniel Internal Comments : Error # and/or description of error : gantry can not comply error. Repair Action : Check tape switch_s and foot switch_s. Then slip ring data flow. Problem Description by Engineer : Problem description by engineer : Gantry cannot comply errors. When was the first occurrence? Within... : 03-Dec-2021 How often is issue occurring? : First occurrence Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Gantry cannot comply errors. Patient Impact/Outcome: Not provided User Impact: System down Current Software Version: 4.1.10 Diagnostic performed by Engineer : Could not get logs to download. Customer went through shut down and power up. System still has gantry can not comply error. ----------------------------------------------------------------------------- Work Order #: WO-04906385 Womack Stephen Problem Description by Engineer : Problem description by engineer : Gantry cannot comply errors. When was the first occurrence? Within... : 03-Dec-2021 How often is issue occurring? : First occurrence Malfunction area : Gantry Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: Gantry cannot comply errors. Patient Impact/Outcome: Not provided User Impact: System down Current Software Version: 4.1.10 Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- Work Order #: WO-04907443 Dhaher Zaid Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 2 |  | Check tape switch_s and foot switch_s. Then slip ring data flow. Could not get logs to download. Customer went through shut down and power up. System still has gantry can not comply error. | [5 empty] | Gantry_HW | TR/RX | [5 empty] | 118835316 | In Progress | 1584225 |  |  | Opened | Opened | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Gantry can not comply | [5 empty] | Refresh |  |  | System issue | System | Gantry-SW | NoEnoughInfo | Gantry-SW-NoEnoughInfo | Gantry-SW-NoEnoughInfo | NoEnoughInfo | NoEnoughInfo | Gantry-SW | Gantry-SW-NoEnoughInfo | bar: No
1626 |  |  | China | Wuhan Province Wuhan Ren’ai Hospital | 345004 | 44532 |  | 05K for IB/Ingenuity Core128/Elite China/Slow image reconstruction | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 05K for IB/Ingenuity Core128/Elite China Safety Questions:Case: 0118826256 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Lu Mingsong Last updated by: Lu Mingsong 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer_s problems: Image reconstruction is slow. Image transmission often fails Detailed description of Customer_s problems or malfunctions: Image reconstruction is slow. Image transmission often fails It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: N/A Remote troubleshoot and test result: N/A PRS TR status: No PRS. Repair Action: Suggest onsite service procedures: Clean the reconstruction cabinet filter mesh. Check the reconstruction card cooling fan and software system. We recommend spare parts 12NC and names: N/A Memo of Customer request: The customer wants an engineer to go on site to do maintenance SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support ----------------------------------------------------------------------------- | 3 | [14 empty] | 118826256 | In Process | 1582414 |  |  | Opened | Opened | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | Recon slow, image transfer fail | [5 empty] | Refresh |  |  |  | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | bar: No
1627 |  |  | Japan | Independent administrative institution National Hospital Organization | 341015 | 44533 |  | X-ray system cannot complyと出てスキャンが途中で停止した。>>再撮影でスキャン可。>>S_GEN_FAULT_ANODE_TUBE_ARC | Work Order #: WO-04901733 Kojima, Kenichi The system meets the specification for the performed service and is returned to use. Problem Description by Engineer : The error message "X-ray system cannot comply" appeared and a scan stopped partway through.>> Another attempt was made and scanning was possible.>> S_GEN_FAULT_ANODE_TUBE_ARC Resolution : [Subject] 03CS:KK:Mr. Hisaki: The error message "X-ray system cannot comply" appeared and a scan stopped partway through.>> Another attempt was made and scanning was possible.>> S_GEN_FAULT_ANODE_TUBE_ARC [4/8 Elements] Q5. How did the examination end? Rescanned from where scan had stopped. Q6. Did any mistaken findings (mistaken diagnosis or treatment) result from the problem? No. Q7. Was there any unintended unnecessary exposure? No Q8 Artifacts? None Resolution: Arcing occurred.It seems to have been a once-off problem. Scanning was possible after a rescan.Suggested the client to wait-and-see. From the logger, 2021/12/03 11:53:27.320__02 [--o] ----------------- Series 76 Ready ----------------- 2021/12/03 11:53:30.090__01 [-oo] - - - - - - Shot 0 Starts - - - - - - 2021/12/03 11:53:38.481 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_ANALOG_FAULT szInfo:2:5 Analog Fault P1:2/P2:32 2021/12/03 11:53:38.523 <*HV*> Type: WARN ErrNum: S_GEN_FAULT_ANODE_TUBE_ARC 2021/12/03 11:53:38.527 <*HV*> Type: NORMAL ErrNum: S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT szInfo: P1:0/P2: 2021/12/03 11:53:38.533__02 [CANCEL] CStateShot:Series 76 cancelled, due to GANTRY_CANCEL_SOURCE_NORMAL_ERROR, in state CStateShot 2021/12/03 11:53:38.536__02 [***] - - - - - - Study Exit - - - - - - 2021/12/03 11:53:38.536__04 <WARN> Type: WARN ErrNum: S_GEN_MGR_SYNC_OFF_OUT_OF_SEQUENCE P1:0/P2:0 2021/12/03 11:53:40.581 <WARN> Type: WARN ErrNum: S_FDOM_PROFILE_FRAME_ISR_TBL_END_DEFAULT_MA_USED P1:0/P2:0 2021/12/03 11:53:40.586__02 <*ERROR*> Type: NORMAL ErrNum: S_MASTER_SEQUENCER_EXIT_STATE_ERR P1:4096/P2:8369 2021/12/03 11:53:40.605__01 <WARN> Type: WARN ErrNum: S_FDOM_PROFILE_FRAME_ISR_TBL_END_DEFAULT_MA_USED P1:0/P2:0 ----------------------------------------------------------------------------- | 4 | [14 empty] | 118831795 | Closed | 1581622 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Arc | [5 empty] | IB Exist | bar: No
1628 |  |  | Japan | 医療法人清翠会 牧病院 | 341017 | 44532 |  | PACS側へ画像転送が止まっている＞＞CTホストの再起動にて復旧 | Work Order #: WO-04892792 Kimura, Kazumasa Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 4 | [14 empty] | 118824423 | Closed | 1581418 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS transfer stop, back to normal after Host reboot | [5 empty] | Refresh |  |  |  | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | bar: No
1629 |  |  | United States | Odessa Regional Medical Center | 40029 | 44531 |  | Changing Table Top | Work Order #: WO-04890835 Brian, Kluever, Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- Work Order #: WO-04890855 Zeitouni Fadi Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 5 |  | Adjust the table cover | [12 empty] | 118822892 | Closed | 1580458 |  |  | Evaluation | Evaluation | [6 empty] | Other | [6 empty] | Noise with table movement | [5 empty] | IB Exist | bar: No
1630 |  |  | China | 内蒙古阿拉善右旗人民医院 | 345031 | 44531 |  | 02N/Core128/artifact | BU Q&R Feedback: No Reason: Corrective Maintenance Next Step Action : PPD 459800911281*1 24H FSE:梁震-86-15849133876 备件已安排出货 Safety Questions : Case: 0118820007 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wennian Han Last updated by: Wennian Han 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 图像伪影。 客户问题详细描述或故障现象： 环形伪影 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： NA 远程诊断与测试结果： 环形伪影，大关重启无效 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 检查模块，屏蔽通道，结合场地实际情况，必要时更换模块 建议备件12NC及名称： " 459800911281 TDMS V2 M.Det Class B FRU" Next Step Action : 派工先联系su ruijie SAP Transfer : Dispatched for Onsite Support SAP Transfer : Case Transferred to SAP Problem Reported by Customer : 02N/Core128/ Next Step Action : NTD:备件组接到客户报修电话 报修2 ：345031 内蒙古阿拉善右旗人民医院 Ingenuity Core128/Elite China（客户给的SN查到的是128排，但她报修说是64排，请确认） 故障：图像有环形伪影 联系人： 王老师- 13948833710 已转RSE：hanwennian Next Step Action : FSE 苏瑞杰 13304719335 电话没接，稍后继续联系 Next Step Action : FSE 苏瑞杰 13304719335告知可以派给他，已发派工短信通知到工程师 ----------------------------------------------------------------------------- | 3 |  | Replace M30 | [12 empty] | 118820007 | Closed | 1578394 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | Engineering | [6 empty] | Artifact | [5 empty] | IB Exist | 459800911281 | TDMS V2.X M.Det Class B FRU | bar: No
1631 |  |  | Japan | 医療法人清翠会 牧病院 | 341017 | 44532 |  | sub_selectionの使用方法ですが、2つのシリーズを同時に読み込むと転送時にWW/WLが同じになってしまいます。 | Work Order #: WO-04894151 Nakagawa, Futoshi The system meets the specification for the performed service and is returned to use Resolution : どちらもシリーズ番号が小さい方のWW/WLになってしまいます。基本的には2つ読み込んで処理するアプリケーションではないため、ご希望に添えない仕様となっています。 Problem description by Clinical Support Specialist : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Diagnostic performed by Engineer : NA ----------------------------------------------------------------------------- | 5 | [14 empty] | 118825781 | Closed | 1577665 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Others | [6 empty] | WW/WL | [5 empty] | Refresh | bar: No
1632 |  |  | China | 中国医科大学附属盛京医院大连分院 | 345001 | 44531 |  | 05R/Ingenuity Core128/Elite China/estop open | BU Q&R Feedback: No Reason: Corrective Maintenance Next Step Action : PPD POWER SUPPLY 120 VAC DUAL OUTPUT ROHS 453567374153 1，24H场地，FSE于志龙-86-18640955881 Problem Reported by Customer : 05R/Ingenuity Core128/Elite China Safety Questions : Case: 0118817570 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Bo Last updated by: Wang Bo 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Dispatched for Onsite Support T2 Activities : Problem Description: 客户问题概述： ESTOP断开。 客户问题详细描述或故障现象： 机架掉电。 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： LIFESYNC P1=2&4 远程诊断与测试结果： 重启无效。 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议测量vicor和FRC供电，必要时先更换VICOR排除掉电故障。 建议备件12NC及名称： PPD POWER SUPPLY 120 VAC DUAL OUTPUT 453567374153 SAP Transfer : Case Transferred to SAP ----------------------------------------------------------------------------- | 3 |  | Replace the Dual Output Vicor PS | [12 empty] | 118817570 | Closed | 1577280 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | Engineering | [6 empty] | E-stop open | [5 empty] | IB Exist | 453567374153 | POWER SUPPLY 120 VAC DUAL OUTPUT ROHS | bar: No
1633 |  |  | China | 中国医科大学附属盛京医院大连分院 | 345001 | 44531 |  | 05R/Ingenuity Core128/Elite China/scan fault | BU Q&R Feedback: No Reason: Corrective Maintenance Next Step Action : PPD MODULE, ROTOR ELECTRONICS TX 453567538334 1，24H场地，FSE于志龙-86-18640955881 Problem Reported by Customer : 05R/Ingenuity Core128/Elite China T2 Activities : Problem Description: 客户问题概述： 扫描失败。 客户问题详细描述或故障现象： 机架扫描中断。 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： S_DMC_SHOT_COMM_COLLAPSE physical address 0x0 远程诊断与测试结果： 重启无效。工程师已到场诊断。 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 建议测量TX输入，必要时更换TX。 建议备件12NC及名称： "PPD MODULE, ROTOR ELECTRONICS TX 453567538334 " Safety Questions : Case: 0118816819 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Bo Last updated by: Wang Bo 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support ----------------------------------------------------------------------------- | 3 |  | Scan cannot initialize, TX no output, replace the TX | [12 empty] | 118816819 | Closed | 1576930 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | Engineering | [6 empty] | Scan cannot initialize | [5 empty] | IB Exist | 453567538334 | ROTOR ELECTRONICS TX | bar: No
1634 |  |  | China | Xifeng County Traditional Chinese Medicine Hospital | 345014 | 44531 |  | 05K For IB/Ingenuity Core128/Elite China/Belt cracking | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer: 05K for IB/Ingenuity Core128/Elite China Safety Questions:Case: 0118816542 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Bai Jingwen Last updated by: Bai Jingwen 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Next Step Action: Tang Guanglin found the belt is broken during maintenance and needs to be replaced. Dispatch hu ke to RSE, ask RSE PPD part T2 Activities:Problem Description: Summary of customer_s problems: BELT BROKEN Detailed description of Customer_s problems or malfunctions: The rotating motor belt of the frame cracked It is a technology problem with the equipment itself: Yes. Troubleshooting Action: key fault code or symptom: N/A Remote troubleshoot and test result: N/A PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Replace the gantry_s rotation motor belt. We recommend spare parts 12NC and names: POLY-VEE BELT: 453566492401 Memo of Customer request: None Next Step Action:PPD POLY-VEE BELT 453566492401 1.24H on site, FSE Tang Guanglin -86-18084189777 SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support ----------------------------------------------------------------------------- | 3 |  | Belet crack found during PM, replace the belt | [12 empty] | 118816542 | Closed | 1576707 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Belt crack | [5 empty] | IB Exist | 453566492401 | POLY-VEE BELT | bar: No
1635 |  |  | Japan | 社会医療法人 母恋 日鋼記念病院 | 341011 | 44543 |  | ShortTC時にアーク発生＞＞継続してShortTCを行い正常終了＞S_GEN_FAULT_ANODE_TUBE_ARC/S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT | Work Order #: WO-04970689 Kimura, Kazumasa The system meets the specification for the performed service and is returned to use Follow Up Required : * Internal Comments : [Subject] 03DS:KKI:アンドウ:ShortTC時にアーク発生＞＞継続してShortTCを行い正常終了＞S_GEN_FAULT_ANODE_TUBE_ARC/S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT [Internal Comments] ログ確認結果 以下のエラーが発生しております。 S_GEN_FAULT_ANODE_TUBE_ARC S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT [出庫部品] 453567560012:Spellman 80 KW RoHS APM Assy# QTY:1 ZREP/Traceable 東京(U611) 989000005011:HV-ANODECABLE PHILIPS MRC880 CT 1,84m QTY:1 ZSID 札幌(U612) 453566495091:1XRT SERVICE TOOL ASSY EXP QTY:1 ZSID 札幌(U612) 453566495101:GENERATOR SERVICE TOOL QTY:1 ZSID 札幌(U612) 459800089031:Tool, XRT Strap Pad Assy QTY:1 ZSID 札幌(U612) 989000086352:PHILIPS MRC 880 QTY:1 ZREP/Traceable 札幌(U612) [ディスパッチ] 1 RC -----------------FSE登録情報----------------- -訪問FSE1氏名 橋本 穂高 -FSE1訪問日時 2021/12/141:00 -FSE1作業時間 2 [出荷指示] -----------------出荷指示#1----------------- -配送条件 ASAP(輸送手配) -納期 ASAP -配送先名 札幌倉庫止め・FSEが引取り -受取者名 橋本 穂高 -電話番号 070-3876-3812 -その他 -依頼者名 木村 和正 Problem Description by Engineer : ShortTC時にアーク発生 Diagnostic performed by Engineer : 継続してShortTCを行い正常終了 ----------------------------------------------------------------------------- Work Order #: WO-04971129 Hashimoto, Hodaka Regulatory Questions : N/A Problem Description by Engineer : ShortTC時にアーク発生 Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | 4 | [14 empty] | 118887945 | Closed | 1612666 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | Engineering | [6 empty] | Arc | [5 empty] | IB Exist | 453567560012 | Spellman 80 KW RoHS APM Assy | bar: No
1636 |  |  | Turkey | ADIYAMANOZELSAGINSTURTASSANVETICA.S | 341006 | 44543 |  | CT belt siparişi | Work Order #: WO-04969812 Salih Mendi The system meets the specification for the performed service and is returned to use Diagnostic performed by Engineer : CT belt siparişi Problem Description by Engineer : Problem description by engineer : Main belt noise issue Information to support the complaint handling process Customer Function/Role: Not provided Device Use: Not provided Expected and Unexpected results: Not provided Patient Impact/Outcome: Not provided User Impact: Not provided Current Software Version: Not provided Internal Comments : Error # and/or description of error : Main belt noise issue Repair Action : Main belt noise issue Extra Internal comment field : Belt değişiminden sonra Tool "unused- good and sealed" olarak raporlanmalıdır. Follow Up Required : Error # and/or description of error : Main belt noise issue Repair Action : Main belt noise issue Extra Internal comment field : Belt değişiminden sonra Tool "unused- good and sealed" olarak raporlanmalıdır. Technician Name : Salih Mendi Resolution : Main belt_in rotasyon sırasında gürültü çıkardığı şikayeti üzerine, main belt siparişi yapılacaktır. Sistem faal durumdadır. ----------------------------------------------------------------------------- | 5 | [14 empty] | 118887293 | Closed | 1612234 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | Engineering | [6 empty] | Belt noise | [5 empty] | IB Exist | 453566492401 | POLY-VEE BELT | bar: No
1637 |  |  | Japan | 独立行政法人国立病院機構 | 341013 | 44543 |  | 撮影時にカタンカタンという異音が発生 | Work Order #: WO-04967241 Miyakawa, Takeyuki The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Follow Up Required : * Diagnostic performed by Engineer : a Problem Description by Engineer : a ----------------------------------------------------------------------------- Work Order #: WO-04967265 Takada, Ushio Problem Description by Engineer : a Regulatory Questions : N/A Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- | [15 empty] | 118884943 | Closed | 1611472 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | Engineering | [6 empty] | DMS Cover noise | [5 empty] | IB Exist | 459800624062 | TDMS Cover RoHS FRU | bar: No
1638 |  |  | China | 河南省濮阳市人民医院 | 40031 | 44541 |  | Wechat:05K for IB/Ingenuity CT 728326/Stop scanning | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : Stop scanning Safety Questions : Case: 0118882346 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Zhou Xiaohong Last updated by: Zhou Xiaohong 1. Was the device in clinical use at the time the issue was discovered? Yes 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: 客户问题概述： 扫描中断 客户问题详细描述或故障现象： 扫描中断 设备本身技术问题： 是。 Troubleshooting Action: 关键故障代码或现象： S_MASTER_SEQUENCER_ACQ_NOT_OK；S_DMC_SHOT_COMM_COLLAPSE 远程诊断与测试结果： 不能远程连接重建主机，eventlog不能下载 PRS TR状态： 全通。 Repair Action: 建议现场维修步骤： 现场检查数据通路，做可视化诊断，查看eventlog，检查TX,RX,需要时更换 建议备件12NC及名称： N/A 客户需求备注： 无 Next Step Action : Luo Yongping SAP Transfer : Case Transferred to SAP SAP Transfer : Dispatched for Onsite Support ----------------------------------------------------------------------------- |  |  | ACQ not ok, resend, replace both TX and RX | [12 empty] | 118882346 | Closed | 1610864 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | Engineering | [6 empty] | ACQ not ok | [5 empty] | IB Exist | Not provided | TX and RX | bar: No
1639 |  |  | United States | Philips Medical Capital | 40053 | 44540 |  | EKG isn_t connecting to the CT | Internal Remarks : FSE did PM lastnight and now ekg not connecting to the CT. ----------------------------------------------------------------------------- |  |  | Conducted IBOX test, checked all cable connections and full reboot of the gantry. IBOX is functioning normally. | [12 empty] | 118881127 | Closed | 1610581 |  |  | Opened | Opened | [6 empty] | Other | [6 empty] | ECG isn_t connecting to CT | [5 empty] | IB Exist | bar: No
1640 |  |  | Japan | Maki Hospital | 341017 | 44540 |  | PACS側へ画像転送が止まっている | Work Order #: WO-04955542 Kawazoe, Junichi Regulatory Questions : N/A Diagnostic performed by Engineer: N/A Problem Description by Engineer: Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: ----------------------------------------------------------------------------- | 5 | [14 empty] | 118876751 | In Progress | 1609141 |  |  | Evaluation | Evaluation | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | Manual PACS transfer stuck | [5 empty] | Refresh |  |  |  | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | bar: No
1641 |  |  | United States | Odessa Regional Medical Center | 40029 | 44540 |  | IQ is failing temperature control malfunction errors | Internal Remarks : Planner Initials: JJ Actions Taken: FSE-Brian Kluever called in to have WO-04958358 dispatch to him. FSE said he will contact the cx I put FADB in the planning notes I dispatched the WO per FSE request Next Steps: N/A Last Updated: 12:40PM EST 12/10/21 ----------------------------------------------------------------------------- Work Order #: WO-04958263 Maier, Donnie The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Follow Up Required : Customer states they are unable to perform the scheduled IQ test they are trying to run. They are getting temperature control malfunction errors. I asked if gantry room feels overly warm and she states no, normal temperatures. I asked about any weird noises like fans going bad and she stated no. system doesn_t appear to have telnet of FTP ports open to pull logs. I asked if customer has tried to perform a normal scan to see if there are artifacts and they have not. I let them know that maybe they can try putting phantom up on table and running belly scan to check for artifacts. I will dispatch FSE to try and duplicate and perform further troubleshooting onsite as we don_t have access remotely. Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: No Patient Involved Expected and Unexpected Results: Cant pass the normal IQ Check giving Temperatuere Control Malfunction Error User Impact: No Patient Impact: No Current Software Version: 4.1.10 Diagnostic performed by Engineer : Customer states they are unable to perform the scheduled IQ test they are trying to run. They are getting temperature control malfunction errors. I asked if gantry room feels overly warm and she states no, normal temperatures. I asked about any weird noises like fans going bad and she stated no. system doesn_t appear to have telnet of FTP ports open to pull logs. I asked if customer has tried to perform a normal scan to see if there are artifacts and they have not. I let them know that maybe they can try putting phantom up on table and running belly scan to check for artifacts. I will dispatch FSE to try and duplicate and perform further troubleshooting onsite as we don_t have access remotely. ----------------------------------------------------------------------------- Work Order #: WO-04958358 Brian, Kluever, Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: No Patient Involved Expected and Unexpected Results: Cant pass the normal IQ Check giving Temperatuere Control Malfunction Error User Impact: No Patient Impact: No Current Software Version: 4.1.10 Diagnostic performed by Engineer : N/A ----------------------------------------------------------------------------- |  |  | With occasion heating error on IQ only, can pass the IQ test, ask customer to check if will happen again. | [12 empty] | 118879015 | Closed | 1606802 |  |  | Opened | Opened | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Heating error | [5 empty] | Refresh |  |  |  | System | Console-SW | NoEnoughInfo | Console-SW-NoEnoughInfo | Console-SW-NoEnoughInfo | NoEnoughInfo | NoEnoughInfo | Console-SW | Console-SW-NoEnoughInfo | bar: No
1642 |  |  | China | Jinan City No. 4 People_s Hospital | 40038 | 44539 |  | 02N/Ingenuity CT 728326/Delete data | BU Q&R Feedback: No SAP Customer Symptom Code: OTHR Customer Symptom Code: Other Symptom; No Symptom Reason: Corrective Maintenance Engineer Failure Code: Unknown Cause Repair Activity Code: Database Recreation / Recovery Problem Reported by Customer: 02N/Ingenuity CT 728326 Next Step Action:FSE Tao Chen 018560023293 call in and disp him Safety Questions:Case: 0118859120 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Liu Guoping Last updated by: Liu Guoping 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities : Problem Description: need delete data Troubleshooting Action: need delete date via PRS Repair Action: delete date via PRS Test/Inspection: Functional Check PASSED. The device returned to full functionality. Test Equipment: N/A SAP Transfer:Case Transferred to SAP ----------------------------------------------------------------------------- |  |  | Delete data via PRS | [12 empty] | 118859120 | Closed | 1604616 |  |  | Opened | Opened | [6 empty] | TBD | [6 empty] | Delete data | [5 empty] | Refresh | bar: No
1643 |  |  | China | Hebei Province Tangshan City Workers Hospital | 40044 | 44539 |  | 02N/Ingenuity CT 728326/barcode reader missing | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Replace Part SAP Transfer:Case Transferred to SAP Problem Reported by Customer: 02N/Ingenuity CT 728326 Safety Questions:Case: 0118858991 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Gaoge Last updated by: Wang Gaoge 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: The person who reported the repair said the scan code gun that came with the machine was lost. Need to send another code scan gun Resolver Action: nemo code scan gun to the customer, the customer replaced it by themselves Repair Action: nemo 455018005052 Test/Inspection: N/A Test Equipment: N/A ----------------------------------------------------------------------------- | [15 empty] | 118858991 | Closed | 1604511 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Customer | [6 empty] | Barcoder scanner missing | [5 empty] | Non-refresh | 455018005052 | BARCODE SCANNER LS2208 USB | bar: No
1644 |  |  | United States | Odessa Regional Medical Center | 40029 | 44539 |  | Need to remove the word _draft_ off of the calcium scoring report. | Internal Remarks : Interaction Type: New Case Caller: Liliana Echeverria IP & Issue: 84812213; Customer requesting clinical support: Need to remove the word _draft_ off of the calcium scoring report. Event Type: Incident Service Request: Remote Clinical Did Customer Decline Remote Support? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: Created Case & Work Order Next Step Action: N/A Verbal PO Number: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- Work Order #: WO-04950411 Mujica, David The system meets the specification for the performed service and is returned to use Problem description by Clinical Support Specialist : Customer Function/Role: Technologist Device Use: In use Expected and Unexpected Results: Draft on Report Removal User Impact: none Patient Impact: none Current Software Version: 4.1 Diagnostic performed by Engineer : Liliana required support for removing the word "draft" off of their Calcium Score reports. After final approval of the the reporting process the system prompts you where you would like to send the final report to. At this phase you must select printer and then the final report will not say "draft". Liliana had no further questions or need for clinical support. No rescan, no loss of data, no reinjection per Liliana. ----------------------------------------------------------------------------- | 5 |  | After final approval of the the reporting process the system prompts you where you would like to send the final report to. At this phase you must select printer and then the final report will not say "draft". | [12 empty] | 118872043 | Closed | 1603038 |  |  | Opened | Opened | [6 empty] | Customer | [6 empty] | Remove "draft" in the calcium scoring report | [5 empty] | Non-refresh | bar: No
1645 |  |  | Puerto Rico | METRO MAYAGUEZ INC | 40009 | 44539 |  | Throw message Gantry cannot comply and the table gets stuck | Work Order #: WO-04950634 09-Dec-2021 Valenzuela Velasquez, Eduardo Ariel The system does not meet the specification for the performed service and is not returned to use Problem Description by Engineer : Problem description by engineer : Gantry cannot comply Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: NO gantry errors. Patient Impact/Outcome: System is down User Impact: System is down Current Software Version: 4.1.10 Internal Comments : Error # and/or description of error : S_DMC_REFERENCE_CONV_DATA Repair Action : - Power supply replacement. Follow Up Required : Error # and/or description of error : S_DMC_REFERENCE_CONV_DATA Repair Action : - Power supply replacement. Technician Name : Valenzuela Velasquez, Eduardo Ariel Diagnostic performed by Engineer : Gantry was checked and a power supply is bad. The part will be ordered. ----------------------------------------------------------------------------- Work Order #: WO-04948963 09-Dec-2021 Hernandez, Alvin The system does not meet the specification for the performed service and is not returned to use Follow Up Required : The Reference detector Power Supply for the Ingenuity CT will be Arrived tomorrow by FedEx. The part was ordered by Ing Edwardo Valenzuela. I will be here to install it. Resolution : The power supply was ordered by FSE Edwardo Valenzuela. It will be installed as soon as I received it. Problem Description by Engineer : Problem description by engineer : Gantry cannnot comply. Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: NO gantry error. Patient Impact/Outcome: System is down. User Impact: System is down Current Software Version: 4.1.10 Diagnostic performed by Engineer : The problem was caused by the -5.0/+5.0 Vdc defectie Reference Detector power supply. this power supply needs to be replaced. ----------------------------------------------------------------------------- Work Order #: WO-04948201 Valenzuela Velasquez, Eduardo Ariel The system does not meet the specification for the performed service and is not returned to use Problem Description by Engineer : Problem description by engineer : Gantry cannnot comply. Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: NO gantry error. Patient Impact/Outcome: System is down. User Impact: System is down Current Software Version: 4.1.10 Diagnostic performed by Engineer : System is showing an error with the reference converter. Gantry mus be checked for power supply and reference converter. Internal Comments : Error # and/or description of error : S_DMC_REFERENCE_CONV_DATA Repair Action : Check POWER SUPPLY, TRIPLE OUTPUT 459800359381 - Check power supply - Check reference converter and cables - Check DMS fuses - Check contrast media Follow Up Required : Error # and/or description of error : S_DMC_REFERENCE_CONV_DATA Repair Action : Check POWER SUPPLY, TRIPLE OUTPUT 459800359381 - Check power supply - Check reference converter and cables - Check DMS fuses - Check contrast media Technician Name : Valenzuela Velasquez, Eduardo Ariel | 2 | [14 empty] | 118870499 | Closed | 1602219 |  |  | Opened | Opened | [6 empty] | Engineering | [6 empty] | Gantry cannot comply | [5 empty] | IB Exist | 459800723941 | PSA5V SUPPLY ASS_Y | bar: No
1646 |  |  | Turkey | ADIYAMANOZELSAGINSTURTASSANVETICA.S | 341006 | 44539 |  | Kontrat madde temizliği (Contrast matter cleanup ) | Work Order #: WO-04948632 Turgay Baylan The system meets the specification for the performed service and is returned to use Resolution : CT cihazı bakım için program yapılması gerekmektedir, bakım sırasında kontrast temizliğide yapılacaktır. Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: ----------------------------------------------------------------------------- | 3 | [14 empty] | 118870833 | Closed | 1602064 |  |  | Evaluation - Pending Translation | Evaluation - Pending Translation | [6 empty] | TBD | [6 empty] | Contrast matter cleanup | [5 empty] | IB Exist | bar: No
1647 |  |  | Spain | S.C.S. HOSP. GRAL. DE FUERTEVENTURA | 40020 | 44539 |  | Salen las imágenes con rayas El fallo ha comenzado ayer En uso clínico Ubicación: Planta 1era | Work Order #: WO-04937851 [REDACTED] The system does not meet the specification for the performed service and is not returned to use. Follow Up Required: We need to go and check the central detector modules. Exchange with external module to see if the artifact moves. Replace module as needed. Regulatory Questions : N/A Problem Description by Engineer: I talk to [REDACTED]. A very noticeable artifact can be observed in all the acquired images. Straight lines focused in the center. Diagnostic performed by Engineer: Images from different protocols (cranial and abdomen studies) are checked, and we observe what seems to be a failure of either a detector or the central detectors. They cause lines in the shape of stretch marks. ----------------------------------------------------------------------------- | 2 |  | Se sustituye el módulo 21 del detector que no pasaba los test de detectores. Una vez sustituido se realizan todos los test indicados en el manual de usuario tras sustitución de módulos incluyendo comprobaciones de calidad de imagen para comprobar que no se reproducen los artefactos. | [12 empty] | 118860492 | Closed | 1601184 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Artifact | [5 empty] | IB Exist | 459800911271 | TDMS V2 M.Det Class A FRU | bar: No
1648 |  |  | China | The Second Hospital of Haibei Prefecture in Qinghai Province | 345029 | 44538 |  | 02N/Ingenuity Core128/Elite China/Replace worklist parameters | BU Q&R Feedback: No Reason: Corrective Maintenance Problem Reported by Customer : Replace worklist parameters Safety Questions:Case: 0118854890 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Zixuan Last updated by: Wang Zixuan 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A SAP Transfer:Case Transferred to SAP SAP Transfer:Dispatched for On site Support T2 Activities:Problem Description: Summary of customer_s problems: Change the PACS system, It is not in a network segment, a field service engineer is needed on site to work with them. It is a technology problem with the equipment itself: Yes. Troubleshooting Action: remote diagnosis and testing result: Change the PACS system, It is not in a network segment, a field service engineer is needed on site to work with them. PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: Change the PACS system, It is not in a network segment, a field service engineer is needed on site to work with them. Memo of Customer request: Contact the person who requested service and confirm the time. Next Step Action: Site1 Field Service Engineer Xu Yonghua 13327661334 Inform OK to dispatch him, a dispatch text message has been sent to the engineer ----------------------------------------------------------------------------- | 5 | [14 empty] | 118854890 | Closed | 1601133 |  |  | Closed - Non-Complaint | Opened | [6 empty] | Customer | [6 empty] | Support for hospital PACS installation | [5 empty] | IB Existing | bar: No
1649 |  |  | China | Anhui Tongling City Bo_ai hospital new campus | 345034 | 44538 |  | 02N/Ingenuity Core128/Elite China/nstall the high-pressure syringe | BU Q&R Feedback: No Reason: Corrective Maintenance Safety Questions:Case: 0118854624 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Bo Last updated by: Wang Bo 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A Problem Reported by Customer : 02N/Ingenuity Core128/Elite China/ T2 Activities:Problem Description: Customer need to dispatch labor to install high pressure injector. Ask C1 to follow up. Troubleshooting Action: Customer need to dispatch labor to install high pressure injector. Ask C1 to follow up. Repair Action: Customer need to dispatch labor to install high pressure injector. Ask C1 help to follow up. ----------------------------------------------------------------------------- | [15 empty] | 118854624 | In Process | 1600931 |  |  | Closed - Non-Complaint | Opened | [6 empty] | Customer | [6 empty] | Do injector installation | [5 empty] | IB Existing | bar: No
1650 |  |  | Japan | Independent administrative institution National Hospital Organization | 341015 | 44539 |  | X-Ray system cannot complyとメッセージが出てスキャンが止まった。>>ログを確認しましたがエラーは見受けられず。 | Work Order #: WO-04937298 Kojima, Kenichi The system meets the specification for the performed service and is returned to use. Problem Description by Engineer : The message "X-ray system cannot comply" appeared and a scan stopped.>> I checked the log, but there was no error. Resolution: I checked the log but did not find any error, such as arcing. It seems "X-ray system cannot comply" appeared, so a high-voltage system error may have occurred. I had the FSE contact the customer, and he agreed to keep an eye on it. Internal Comments: [Subject] 03CS:KK:Mr. Hisaki: "X-ray system cannot comply" appeared and a scan stopped.>> I checked the log, but there was no error. [4/8 Elements] Q5. How did the examination end? Rescanned from where scan had stopped. Q6. Did any mistaken findings (mistaken diagnosis or treatment) result from the problem? No. Q7. Was there any unintended unnecessary exposure? No Q8 Artifacts? None ----------------------------------------------------------------------------- | 4 |  | ログを確認しましたがアークなどのエラーは確認できず。 X-Ray system cannot complyと出たらしいので何かしら高圧系のエラーが起きたと思われます。 FSEより連絡を入れていただき様子見とさせていただきました。 | [12 empty] | 118859965 | Closed | 1600567 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | XRAY system cannot comply | [5 empty] | IB Exist |  |  |  | System | Gantry-SW | bar: No
1651 |  |  | Japan | Maki Hospital | 341017 | 44537 |  | Long TC作業 | Work Order #: WO-04922289 08-Dec-2021 Kawazoe, Junichi The system meets the specification for the performed service and is returned to use. Internal Comments: [CX related] Did you agree on how to complete the report with the person in charge of devices? <Select one of the following> Yes. I completed normal reporting that day. [Work report (paper) submission] Did you submit a work report on paper? Yes Q1. How did the examination finish? Select from the following: *Because the answer is No for Q1, not applicable Q2. Were there any mistaken findings (misdiagnosis, mistaken treatment) due to the problem? None. Q3. Were there any artifacts? None. Q4. Was there any unintended unnecessary exposure? None. Resolution : Work details: I performed a tube adjustment (Long Tube Conditioning, Filament Calibration). There is no problem with the operation of the equipment. Problem Description by Engineer: Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Regulatory Questions : N/A ----------------------------------------------------------------------------- | 5 |  | Long Tube condition and filament calibration | [12 empty] | 118848277 | Closed | 1596654 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Engineering | [6 empty] | PM to do LTC and Filament | [5 empty] | IB Existing | bar: No
1652 |  |  | China | Hebei Province Geriatric Hospital | 345015 | 44536 |  | 01/Ingenuity Core128/Elite China/add dicom node | BU Q&R Feedback: No Reason: Corrective Maintenance Repair Activity Code: Calibration/Adjustment Next Step Action: Dispatched Engineer Liang Longse to follow 159 3113 3573 up Problem Reported by Customer: 01/Ingenuity Core128/Elite China Next Step Action : Transfer to sales handling, this case is valid within 90 days starts from this case created, once beyond of this period without any further steps, this case is considered invalid, customer refuses our service Next Step Action: contacted JiuNuo ASP to follow up with Business Safety Questions:Case: 0118840146 Safety Question Status: Submitted PSA: false PSE: false PSF Status: NA Created By: Wang Gaoge Last updated by: Wang Gaoge 1. Was the device in clinical use at the time the issue was discovered? No 2. Was any patient or user harmed? No 3. If the device has alarm/alert capability, did it alarm/alert as it should have at the time the issue was discovered? N/A 4. Was this an out of box failure? N/A T2 Activities:Problem Description: Summary of customer_s problems: Other types of error. Detailed description of Customer_s problems or malfunctions: Need to add nodes. Admin password doesn_t help It is a technology problem with the equipment itself: No. Troubleshooting Action: key fault code or symptom: Admin password is missing Remote troubleshoot and test result: Need an engineer to reset admin password on the site. PRS reset failed PRS TR status: All passed through. Repair Action: Suggest onsite service procedures: 1. Reset admin password Test/Inspection: N/A Test Equipment: N/A ----------------------------------------------------------------------------- | [15 empty] | 118840146 | In Process | 1590950 |  |  | Opened | Opened | [6 empty] | Customer | [6 empty] | Forget IT password | [5 empty] | Refresh | bar: No
1653 |  |  | China | 河北省唐山市工人医院 | 40044 | 44549 |  | 02N/Ingenuity CT 728326/cable failed |  | 3 | [14 empty] | 118924822 | Closed | 1635285 |  |  | Opened - Pending Translation | Opened - Pending Translation | CTD00006847 | [5 empty] | Service | [6 empty] | Video cable | [5 empty] | IB Exist |  |  |  | System | CIRS-SW | ACQ_NOT_OK | CIRS-SW-ACQ_NOT_OK | CIRS-SW-ACQ_NOT_OK | CTD00006847 | ACQ_NOT_OK | CIRS-SW | CIRS-SW-ACQ_NOT_OK | bar: No
1654 |  |  | China | 河北省唐山市工人医院 | 40044 | 44549 |  | 02N/Ingenuity CT 728326/dispaly problem |  |  |  | 机器无法使用，处理视频线后正常。 服务结果：服务后设备恢复正常。 | [12 empty] | 118924795 | Closed | 1635248 |  |  | Opened - Pending Translation | Opened - Pending Translation | CTD00006847 | [5 empty] | Service | [6 empty] | Video cable | [5 empty] | IB Exist |  |  |  | System | CIRS-SW | ACQ_NOT_OK | CIRS-SW-ACQ_NOT_OK | CIRS-SW-ACQ_NOT_OK | CTD00006847 | ACQ_NOT_OK | =BE1656 | CIRS-SW-ACQ_NOT_OK | bar: No
1655 |  |  | China | HEBEI Tangshan Staff Hosp. | 40044 | 44550 |  | System can not scan. Err code ACQ NOT OK. |  | CFM | [14 empty] | CFM |  | 1635224 |  |  | Opened | Opened | CTUDT00630408 | [5 empty] | Engineering | [6 empty] | ACQ NOT OK | [5 empty] | Refresh |  |  |  | System | CIRS-SW | ACQ_NOT_OK | CIRS-SW-ACQ_NOT_OK | CIRS-SW-ACQ_NOT_OK | CTUDT00630408 | ACQ_NOT_OK | CIRS-SW | CIRS-SW-ACQ_NOT_OK | bar: No
1656 |  |  | China | 扎兰屯市人民医院 | 40045 | 44549 |  | 02N/Ingenuity CT/cannot scan |  | 2 | [14 empty] | 118924630 | In Process | 1635172 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | Engineering | [6 empty] | Gantry initialization | [5 empty] | IB Exist | bar: No
1657 |  |  | China | 内蒙古兴安盟人民医院 | 345030 | 44548 |  | 02N/Ingenuity Core128/Elite China/ring artifact |  | 3 |  | Image artifact, M12 replaced. | [12 empty] | 118923872 | Closed | 1634692 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | Engineering | [6 empty] | Artifact | [5 empty] | IB Exist | 459800911281 | TDMS V2 M.Det Class B FRU | bar: No
1658 |  |  | China | 河南省新乡市凤泉区大块镇卫生院 | 345009 | 44547 |  | 05K for IB/Ingenuity Core128/Elite China/Customer consultation | [16 empty] | 118918764 | Closed | 1633115 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | Customer | [6 empty] | Consult for Service Contract and report for the workstation issue | [5 empty] | Non-refresh | bar: No
1659 |  |  | Puerto Rico | DORADO HEALTH INC DBA | 40007 | 44547 |  | It throws an error that asks to delete data since there is not enough space but there is only 5% occupation |  |  |  | Customer store the patient data to local hard disk | [12 empty] | 118921581 | Closed | 1631997 |  |  | Opened | Opened | CTUDT00632261 | [5 empty] | Engineering | [6 empty] | Disk full | [5 empty] | Refresh |  |  |  | System | Console-SW | DiskFull | Console-SW-DiskFull | Console-SW-DiskFull | CTUDT00632261 | DiskFull | Console-SW | Console-SW-DiskFull | bar: No
1660 |  |  | Japan | 独立行政法人国立病院機構 | 341015 | 44547 |  | 撮影したデータが短くExtendしました。シンスライスも結合されていますが、MPRを作成するためにCTViewerで開くと片方しか表示されません。 |  | 5 |  | Start up the CView and open the siries tab and split into two. If you select two, right-click, Marge will appear, so choose and combine | [12 empty] | 118919064 | Closed | 1629595 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Application | [6 empty] | Clinical consult: marge combine | [5 empty] | Refresh | bar: No
1661 |  |  | United Kingdom | STEPPING HILL HOSPITAL | 40016 | 44545 |  | at 13:30pm on 14/12/21 during a diagnostic colonoscopy the scanner reported at reconstruction error Customer did not try an off-line recon nor did they do a bugreport as reported that they were very busy The customer rescanned the patientok but resulting in an addition DLP and dose Since the incident the scanner has been in use without any reported errors after FSE Warren Thorpe spoke with customer a bugreport was generated today at 12:15 on 15/12/21 FOR UKI REFERENCE ONLY: Feedback type: potential incident UKI KMO: LisaHollis Submitter email: doug.anderson_2@philips.com |  | CFM | [14 empty] | CFM |  | 1625818 |  |  | Investigation | Investigation | [6 empty] | TBD | [6 empty] | Recon error | [5 empty] | Refresh | bar: No
1662 |  |  | Japan | Maki Hospital | 341017 | 44546 |  | Quick viewerを開こうとしたらBugrep作成するようポップアップが出た。>>Host再起動後正常動作。ハード的なエラーは無し。 |  | 4 |  | Restart | [12 empty] | 118905403 | Closed | 1624861 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Popup create Bugreport request, after click the quick view. | [5 empty] | Refresh |  |  |  | System | Console-SW | NoEnoughInfo | Console-SW-NoEnoughInfo | Console-SW-NoEnoughInfo | NoEnoughInfo | NoEnoughInfo | Console-SW | Console-SW-NoEnoughInfo | bar: No
1663 |  |  | China | 息烽县中医医院 | 345014 | 44545 |  | WeChat：05K for IB/Ingenuity Core128/Elite China/Mat breakage |  | Zhai, Lisa | [14 empty] | 118899967 | Closed | 1624597 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Mat breakage | [5 empty] | IB Exist | bar: No
1664 |  |  | China | 河北省唐山市工人医院 | 40044 | 44545 |  | 02N/Ingenuity CT 728326/ACQ issue |  |  |  | Reload CIRS software, re-plug the ACQ board | [12 empty] | 118898185 | In Process | 1623633 |  |  | Opened - Pending Translation | Opened - Pending Translation | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | ACQ error | [5 empty] | Refresh |  |  |  | System | Console-HW | NoEnoughInfo | Console-HW-NoEnoughInfo | Console-HW-NoEnoughInfo | NoEnoughInfo | NoEnoughInfo | Console-HW | Console-HW-NoEnoughInfo | bar: No
1665 |  |  | United Kingdom | STOCKPORT NHS FOUNDATION TRUST | 40016 | 44544 |  | came up with system error during reconstruction - occurred at 13:30 |  | 4 | [14 empty] | 118895711 | In Progress | 1621949 |  |  | Investigation | Investigation | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | reconstruction failed during a colonoscopy | [5 empty] | Refresh |  |  |  | System | CIRS-SW | NoEnoughInfo | CIRS-SW-NoEnoughInfo | CIRS-SW-NoEnoughInfo | NoEnoughInfo | NoEnoughInfo | CIRS-SW | CIRS-SW-NoEnoughInfo | bar: No
1666 |  |  | China | Lindian County Hospital | 345038 | 44544 |  | 02N/Ingenuity Core128/Elite/yellow China |  | 5 |  | Reload Host software | [12 empty] | 118892511 | In Process | 1619715 |  |  | Opened | Opened | CTUDT00632520 | [5 empty] | Engineering | [6 empty] | Console stuck and yellow screen after scan complete. | [5 empty] | Refresh |  |  |  | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow-Screen | Console-SW | Console-SW-Yellow-Screen | bar: No
1667 |  |  | Japan | Maki Hospital | 341017 | 44544 |  | ガントリー回転時に異音が発生する |  | 5 |  | Replace the main belt | [12 empty] | 118891082 | Closed | 1618885 |  |  | Evaluation | Evaluation | [6 empty] | Engineering | [6 empty] | Noise when gantry rotation | [5 empty] | IB Exist | 453566492401 | POLY-VEE BELT | bar: No
1668 |  |  | Japan | Tomakomai City Hospital | 341014 | 44544 |  | 画面がワンモニターになってしまった （The screen has become one monitor ） |  | Zhu, Chun |  | One of the DisplayPort cabinets on the console side was missing from the signal. Cable replacement | [12 empty] | 118897534 | In Progress | 1618706 |  |  | Evaluation - Pending Translation | Evaluation - Pending Translation | ReplaceCable | [5 empty] | Engineering | [6 empty] | 1 Video signal missing | [5 empty] | Refresh |  | Cable |  | System | Console-HW | ReplaceCable | Console-HW-ReplaceCable | Console-HW-ReplaceCable | ReplaceCable | ReplaceCable | Console-HW | Console-HW-ReplaceCable | bar: No
1669 |  |  | United States | Lackey Memorial Hospital | 40015 | 44544 |  | Gantry wont turn on. |  | 2 | [14 empty] | 118892135 | Fixed | 1616223 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Gantry cannot initialize | [5 empty] | IB Exist | 453567017751 | RHOST SG164 ASSY | bar: No
1670 |  |  | Portugal | Centro Hospitalar de Vila Nova | 40030 | 44543 |  | O TC do SU aquando a programação de duas aquisições sequenciadas arterial e venosa, dá erro na altura de fazer a aquisição arterial e é abortada pelo (The SU CT when programming two sequential acquisitions in blood and vein, gives error at the time of making the arterial acquisition and is aborted by the equipment) equipamento. ao fazer ok ao erro e retirando a contagem do tempo é possível fazer as aquisições. Problema tem vindo a ser frequente radiologia Convencional - U1 SN: 40030 PO: OT 2021120349 | [16 empty] | 118886493 | Closed | 1615521 |  |  | Evaluation | Evaluation | [6 empty] | TBD | [6 empty] | System error inbetween two acquisition series | [5 empty] | TBD | bar: No
1671 |  |  | China | Hebei Province Tangshan City Workers Hospital | 40044 | 44543 |  | WeChat：02N/Ingenuity CT 728326/report error |  | Cai, Cathy |  | Replace GPU | [12 empty] | 118884797 | Closed | 1614437 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | Replace-GPU | [5 empty] | Engineering | [6 empty] | System rebuild slowly-GPU | [5 empty] | Refresh | 459801622381 | NVidia P2000 GPU Card |  | System | Console-HW | GPU | Console-HW-GPU | Console-HW-GPU | Replace-GPU | GPU | Console-HW | Console-HW-GPU | bar: No
1672 |  |  | United States | Philips Medical Capital | 300057 | 44533 |  | Fan is noisy, replaced 453566503142,Fan, Tube Axial | PR1584749 investigates the ring artifact issue, PR1644755 investigates the noisy fan issue. Ring artifact. #40 replacement. 459800594811,TDMS Module Class A | Kanzeg, Mike | [14 empty] | 118836197 | Closed | 1644755 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | Fan noise | [5 empty] | IB Exist | 453566503142 | Fan, Tube Axial | bar: No
1673 |  |  | Mexico | DE LAGE LANDEN SA DE CV | 10007 | 44549 |  | EL EQUIPO ESTA BLOQUEADO. | Work Order #: WO-05015036 [REDACTED] The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled. Problem Description by Engineer : Problem description by engineer : The system opens the e-Stop. When was the first occurrence? Within... : December 18 2021, 6:00:00 PM How often is issue occurring? : Intermittent Malfunction area : Table Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: The gantry shuts down. Patient Impact/Outcome: Scan is aborted. User Impact: Scans cannot be done and starting device is hard. Current Software Version: 4.1.10 Diagnostic performed by Engineer: The system is checked because it has been failing related to aborted acquisitions or e-Stop opening and being unable to be restored until several full resets. Bug report is checked remotely and communication failure logs are found in canbus because of the table controller. Case is derived to Planning area with some suggestions for parts for visit. Internal Comments : Error # and/or description of error : S_LIFESYNCMGRTASK_LIFESYNC_FAIL P1=8 CRC RESENDS Repair Action: Resends were also found in channels A and C mainly. Check. Follow Up Required: Error Number and/or Description: S_LIFESYNCMGRTASK_LIFESYNC_FAIL P1=8 CRC RESENDS Repair Action: Resends were also found in channels A and C mainly. Check. What are the next Troubleshooting steps : DC power supply and DC-Couch Control board cable proactive replacement. Check data path to check resends detected mainly in channel A and C. Technician_s Name: [REDACTED] ----------------------------------------------------------------------------- Work Order #: WO-05022597 [REDACTED] Problem Description by Engineer : Problem description by engineer : The system opens the e-Stop. When was the first occurrence? Within... : December 18 2021, 6:00:00 PM How often is issue occurring? : Intermittent Malfunction area : Table Information to support the complaint handling process Customer Function/Role: Technician/Sonographer Device Use: Diagnostic Expected and Unexpected results: The gantry shuts down. Patient Impact/Outcome: Scan is aborted. User Impact: Scans cannot be done and starting device is hard. Current Software Version: 4.1.10 Diagnostic performed by Engineer: N/A ----------------------------------------------------------------------------- | 2 | [14 empty] | 118925159 | Fixed | 1635717 |  |  | Opened | Opened | [6 empty] | Engineering | [6 empty] | E-stop open | [5 empty] | IB Exist | 453566503742 and 453567133673 | ASSY, 4 OUTPUT DC SUPPLY and CBL, DC POWER | bar: No
1674 |  |  | United States | Upson Regional Medical Center | 310115 | 44546 |  | System error during scan, e-stop opened - Replaced faulty 24V couch relay | Work Order #: WO-05005844 18-Dec-2021 Davis Paul The system meets the specification for the performed service and is returned to use Resolution : 24V couch relay is faulty. Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: Diagnostic Expected and Unexpected Results: System Error Please recover sequence error User Impact: No Patient Impact: No Current Software Version: 4.1.10.25220 Parts Consumed : 1.00000,453566496941,RELAY,24VDC PLUGIN W/COVER 1.00000,453566496941,RELAY,24VDC PLUGIN W/COVER ----------------------------------------------------------------------------- Work Order #: WO-05000975 18-Dec-2021 Davis Paul The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: Diagnostic Expected and Unexpected Results: System Error Please recover sequence error User Impact: No Patient Impact: No Current Software Version: 4.1.10.25220 Resolution : 24V couch relay faulty. Follow Up Required : Ran bug report. Error not recurring. ----------------------------------------------------------------------------- Work Order #: WO-05000882 Maier, Donnie The system does not meet the specification for the performed service and is not returned to use Follow Up Required : Customer states system is giving multiple errors. Most recent was a system error during scan please reconstruct. Customer states they have already power cycled completely down to wall breaker multiple times. She believes they were having other errors as well last night as far as couch free floating as if estop opened. Problem Description by Engineer : Customer Function/Role: CT Tech Device Use: Diagnostic Expected and Unexpected Results: System Error Please recover sequence error User Impact: No Patient Impact: No Current Software Version: 4.1.10.25220 Diagnostic performed by Engineer : Customer states system is giving multiple errors. Most recent was a system error during scan please reconstruct. Customer states they have already power cycled completely down to wall breaker multiple times. She believes they were having other errors as well last night as far as couch free floating as if estop opened. I will try to pull logs, analyze and add case notes if possible. ----------------------------------------------------------------------------- | Mulhern, Melissa |  | Replace the couch 24V relay | [12 empty] | 118914162 | Closed | 1626503 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Gantry wouldn_t initialize Couch Free Floating | [5 empty] | IB Exist | 453566496941 | RELAY, 24VDC PLUG-IN W/COVER | bar: No
1675 |  |  | United States | Christ Hospital | 300205 | 44544 |  | 0903CFXXX parts ID request | Internal Remarks : Interaction Type: New Case Caller: Corey Claxton IP & Issue: 65842604 & parts id// Event Type: Incident Service Request: Remote Technical Did Customer Decline Remote Support? No Preferred Service Time: ASAP Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: created case and WO and routed customer to tech support Next Step Action: N/A Verbal PO Number: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- Work Order #: WO-04982116 Morrow, Dan The system meets the specification for the performed service and is returned to use Resolution : Customer wants part number and part ordered for class A V2 module. 459800911271 - TDMS V2 M.Det Class A FRU. NEMO ordered part for customer. Internal Comments : Error # and/or description of error : No error. Repair Action : Customer wants part number and part ordered for class A V2 module. 459800911271 - TDMS V2 M.Det Class A FRU. NEMO ordered part for customer. Diagnostic performed by Engineer : Customer wants part number and part ordered for class A V2 module. 459800911271 - TDMS V2 M.Det Class A FRU. NEMO ordered part for customer. Follow Up Required : Please ship requested parts to the following location for Next Day Arrival (NDA): The Christ Hospital-Ft Wright Attention: CT/Cory Claxton 1955 Dixie Hwy FT WRIGHT KY 41011-2792 Problem Description by Engineer : Problem description by engineer : Parts ID request When was the first occurrence? Within... : 14-Dec-2021 How often is issue occurring? : First occurrence Malfunction area : Detector Sys Information to support the complaint handling process Customer Function/Role: BioMed Device Use: No patient involved Expected and Unexpected results: Part ID Patient Impact/Outcome: none User Impact: Artifact Current Software Version: Not provided ----------------------------------------------------------------------------- Work Order #: WO-04982201 ----------------------------------------------------------------------------- | Shai, Gal | [14 empty] | 118897024 | In Progress | 1618436 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Customer order the part | [5 empty] | IB Exist | 459800911271 | TDMS V2 M.Det Class A FRU | bar: No
1676 |  |  | United States | Christ Hospital | 300205 | 44544 |  | 65842604 RING ARTIFACTS | Internal Remarks : **CUST REQUESTED TECH SUPPORT Interaction Type: New Case Caller: CORI CLAXTON IP & Issue: 65842604 RING ARTIFACTS Event Type: Incident Service Request: Remote Technical Did Customer Decline Remote Support? No Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: CREATED CASE AND REMOTE WO; RTAC Next Step Action: N/A Verbal PO Number: Verbal PO Approver: Verbal PO CAP: ----------------------------------------------------------------------------- Work Order #: WO-05003253 22-Dec-2021 Hanlin Christopher The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Customer Function/Role: Technician Device Use: during normal scanning Expected and Unexpected Results: Customer reports artifacts in image; not quite a complete circle at times. User Impact: Affecting image quality, but scans are diagnostic. Customer is continuing scanning Patient Impact: NA except for potential image quality issue. Current Software Version: Resolution : Unable to replace detector module in question ; part was not ordered. There is no artifact in images, but module is scheduled for replacement due to noise issue found during calibrations. Customer will order module and reschedule replacement at their convenience. ----------------------------------------------------------------------------- Work Order #: WO-04981237 16-Dec-2021 Hanlin Christopher The system meets the specification for the performed service and is returned to use Tools & Equipment : MULTIMETER,2022-06-30 00:00:00 Problem Description by Engineer : Customer Function/Role: Technician Device Use: during normal scanning Expected and Unexpected Results: Customer reports artifacts in image; not quite a complete circle at times. User Impact: Affecting image quality, but scans are diagnostic. Customer is continuing scanning Patient Impact: NA except for potential image quality issue. Current Software Version: Follow Up Required : Replace detector module 24 due to marginal noise spikes in module. Resolution : Ran Ring ID to locate potential artifact; verified bad detector elements and module by running Bad Detector diagnostics. Replaced detector module and recalibrated system per recommendation. ----------------------------------------------------------------------------- Work Order #: WO-04981056 Powell, Jamie The system does not meet the specification for the performed service and is not returned to use Diagnostic performed by Engineer : customer states that they_re having ring artifacts during scans.. Customer ran air cals bust still having the issue.. No contrast that customer knows of. No ftp access to download logs. Problem Description by Engineer : Customer Function/Role: Device Use: Expected and Unexpected Results: User Impact: Patient Impact: Current Software Version: Follow Up Required : No ftp access to download logs. customer ran air cals but still having issues. Check for contrast. Run bad detector test. Run fee test. Internal Comments : No ftp access to download logs. customer ran air cals but still having issues. Check for contrast. Run bad detector test. Run fee test. ----------------------------------------------------------------------------- | Shai, Gal |  | Replace module 24 | [12 empty] | 118896346 | Closed | 1617837 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Artifact | [5 empty] | IB Exist | No infor | bar: No
1677 |  |  | United States | Select Medical Corporation | 32261 | 44536 |  | reference account is currently locked | Internal Remarks : Interaction Type: New Case Caller: Jesse Nieves IP & Issue: 7688812 currently locked out of system Event Type: Incident Service Request: Remote Technical Did Customer Decline Remote Support? No Preferred Service Time: N/A Entitlement Options: Quote Reason for Entitlement Claim: N/A Action Taken: Created case and quote Next Step Action: N/A Verbal PO Number: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- | [15 empty] | 118843629 | New | 1593453 |  |  | Evaluation | Evaluation | [6 empty] | Customer | [6 empty] | Customer account lock | [5 empty] | Refresh | bar: No
1678 |  |  | Australia | CITISCAN RADIOLOGY PTY LTD | 10491 | 44536 |  | Gantry turned of using the UNLOAD option | Work Order #: WO-04935704 17-Dec-2021 Richard Wade The system meets the specification for the performed service and is returned to use Problem Description by Engineer : CUSTOMER REPORTED PROBLEM OR EVENT DESCRIPTION: 6 Dec 2021 The Customer called to report on behalf of a colleague that the Gantry turned off using the UNLOAD option. The issue occurred at approx. 10:05 AM and was resolved by turning the gantry back on. Resolution : Performed patient support cable checks - all ok. Performed numerous load/ unload motions all without error. Checked and tested system ok. ----------------------------------------------------------------------------- Work Order #: WO-04912325 Steve, Teixeira The system meets the specification for the performed service and is returned to use Problem Description by Engineer : CUSTOMER REPORTED PROBLEM OR EVENT DESCRIPTION: 6 Dec 2021 The Customer called to report on behalf of a colleague that the Gantry turned off using the UNLOAD option. The issue occurred at approx. 10:05 AM and was resolved by turning the gantry back on. Diagnostic performed by Engineer : ________________________________________________________________ CUSTOMER REPORTED PROBLEM OR EVENT DESCRIPTION: 6 Dec 2021 The Customer called to report on behalf of a colleague that the Gantry turned off using the UNLOAD option. The issue occurred at approx. 10:05 AM and was resolved by turning the gantry back on. SOFTWARE REVISION OF THE SYSTEM: Ver. 4.1.10 WAS A BUG REPORT SAVED? No TROUBLESHOOTING - 6 Dec 2021: Informed the Customer that would download and review the logs. Logs confirm that the UNLOAD button was pressed at 09:46:17.145300 AM 6 Dec 2021 and released at 09:46:24.975473 AM. The logs indicate that the Gantry turned off shortly after at 09:46:26.996509 AM due to ESTOP error "S_COUCHMGR_HOR_MOTION_STOP_ERROR " at 09:46:27.183 AM followed by Normal type Error "S_MANMOT_HORIZONTAL_STOP_ERROR" at 09:46:30.087 AM. The errors appear to indicate that the Couch did not stop in time. Informed the Customer of the findings. The Customer confirmed with the colleague unloading the patient, that the patient was not moving unexpectedly at the time. CUSTOMER UPDATE: 9 Dec 2021 Sent an email update to the Customer advising of the action plan. Email extract reads: "Hi XXXX, Regarding the case logged on the 6th Dec 2021 for the gantry turning off whilst using the unload – Case 0118839759. As discussed, a review of the logs shows the User released the UNLOAD button just before the bed reached the out limit, it then appears that the bed although requested to stop continued to move slightly which was not expected, and as consequence turned the gantry off due to an unexpected movement. A Planned Maintenance service is scheduled for Thursday the 16th at 12:30 PM, so the attending Engineer as part of the service will inspect the integrity of the Horizontal couch assemblies and action where required. Please feel free to contact me should you have any further queries in this regard. ..." ONSITE WORK ORDER: Please refer to the onsite work order WO-04935704 for further details of the actions taken to address this issue. ________________________________________________________________ Internal Comments : LOG REVIEW: *********** ________________________________________________________________ LOGGER: * 2021/12/06 10:37:38.038 EVENT_MINOR GSE GSE::CState::StateBkst::STUDY_ENDED was sent to GPC 2021/12/06 10:35:39.042 EVENT_MAJOR GSE ----------------- Series 12 Complete ----------------- 2021/12/06 10:29:46.870 EVENT_MAJOR GSE ----------------- Series 11 Complete ----------------- 2021/12/06 10:29:20.504 EVENT_MAJOR GSE ----------------- Series 10 Complete ----------------- 2021/12/06 10:28:40.889 EVENT_MINOR GSE GSE::CState::StateBkst::STUDY_STARTED was sent to GPC 2021/12/06 09:47:18.789 ERROR_WARNING Gantry CStateConfigured 3 ERROR 522593 WARN S_HMC_COUCH_24V_IS_ON 3 1073938000 2021/12/06 09:47:18.740__02 EVENT_MINOR GSE GantryEstopClosedBkst in CStateConfigured 2021/12/06 09:47:18.646 ERROR_WARNING Gantry CStateConfigured 0 ERROR 522593 WARN S_MPS_WAITING_FOR_MAIN_DRIVE_TO_CLOSE 0 0 MP: Waiting for MDC to close.. 2021/12/06 09:46:30.087 ERROR_NORMAL Gantry CStateConfigured 0 ERROR 522544 NORMAL S_MANMOT_HORIZONTAL_STOP_ERROR 0 0 2021/12/06 09:46:27.183 ERROR_FATAL Gantry CStateConfigured 3 ERROR 522541 ESTOP S_COUCHMGR_HOR_MOTION_STOP_ERROR 0 0 2021/12/06 09:46:27.039 ERROR_FATAL Gantry GantryEstopOpenedBkst(0) in CStateSeriesNotLoaded 2021/12/06 09:46:09.711 EVENT_MINOR GSE GSE::CState::StateBkst::STUDY_ENDED was sent to GPC 2021/12/06 09:45:27.792 EVENT_MAJOR GSE ----------------- Series 9 Complete ----------------- * ________________________________________________________________ CANTRACE: * 2021/12/06 09:46:17.145300 [PANEL] PANEL_UNLOAD_PANEL_REQ [PRESS] 2021/12/06 09:46:17.201735 COUCH_START_HOR_JOG_REQ 2021/12/06 09:46:17.259700 COUCH_REPORT_HOR_POS_BKST 5136 1(OUT) 0 0 2021/12/06 09:46:17.512603 COUCH_REPORT_HOR_POS_BKST 5193 1(OUT) 4276 0 2021/12/06 09:46:17.763763 COUCH_REPORT_HOR_POS_BKST 5376 1(OUT) 9269 0 2021/12/06 09:46:18.015409 COUCH_REPORT_HOR_POS_BKST 5683 1(OUT) 14245 0 2021/12/06 09:46:18.265495 COUCH_REPORT_HOR_POS_BKST 6111 1(OUT) 18433 0 2021/12/06 09:46:18.516199 COUCH_REPORT_HOR_POS_BKST 6575 1(OUT) 18539 0 2021/12/06 09:46:18.768578 COUCH_REPORT_HOR_POS_BKST 7041 1(OUT) 18488 0 2021/12/06 09:46:19.019183 COUCH_REPORT_HOR_POS_BKST 7505 1(OUT) 18506 0 2021/12/06 09:46:19.269510 COUCH_REPORT_HOR_POS_BKST 7968 1(OUT) 18510 0 2021/12/06 09:46:19.520522 COUCH_REPORT_HOR_POS_BKST 8432 1(OUT) 18499 0 2021/12/06 09:46:19.772574 COUCH_REPORT_HOR_POS_BKST 8898 1(OUT) 18491 0 2021/12/06 09:46:20.022802 COUCH_REPORT_HOR_POS_BKST 9362 1(OUT) 18476 0 2021/12/06 09:46:20.274589 COUCH_REPORT_HOR_POS_BKST 9826 1(OUT) 18493 0 2021/12/06 09:46:20.525580 COUCH_REPORT_HOR_POS_BKST 10292 1(OUT) 18493 0 2021/12/06 09:46:20.777598 COUCH_REPORT_HOR_POS_BKST 10757 1(OUT) 18499 0 2021/12/06 09:46:21.028266 COUCH_REPORT_HOR_POS_BKST 11222 1(OUT) 18495 0 2021/12/06 09:46:21.278869 COUCH_REPORT_HOR_POS_BKST 11686 1(OUT) 18494 0 2021/12/06 09:46:21.530088 COUCH_REPORT_HOR_POS_BKST 12150 1(OUT) 18492 0 2021/12/06 09:46:21.781585 COUCH_REPORT_HOR_POS_BKST 12614 1(OUT) 18498 0 2021/12/06 09:46:22.032298 COUCH_REPORT_HOR_POS_BKST 13080 1(OUT) 18551 0 2021/12/06 09:46:22.283489 COUCH_REPORT_HOR_POS_BKST 13544 1(OUT) 18497 0 2021/12/06 09:46:22.534097 COUCH_REPORT_HOR_POS_BKST 14008 1(OUT) 18508 0 2021/12/06 09:46:22.786107 COUCH_REPORT_HOR_POS_BKST 14473 1(OUT) 18501 0 2021/12/06 09:46:23.036215 COUCH_REPORT_HOR_POS_BKST 14937 1(OUT) 18516 0 2021/12/06 09:46:23.287252 COUCH_REPORT_HOR_POS_BKST 15400 1(OUT) 18495 0 2021/12/06 09:46:23.538756 COUCH_REPORT_HOR_POS_BKST 15865 1(OUT) 18480 0 2021/12/06 09:46:23.789504 COUCH_REPORT_HOR_POS_BKST 16330 1(OUT) 18472 0 2021/12/06 09:46:24.040424 COUCH_REPORT_HOR_POS_BKST 16794 1(OUT) 18493 0 2021/12/06 09:46:24.291803 COUCH_REPORT_HOR_POS_BKST 17258 1(OUT) 18508 0 2021/12/06 09:46:24.543007 COUCH_REPORT_HOR_POS_BKST 17724 1(OUT) 18490 0 2021/12/06 09:46:24.792994 COUCH_REPORT_HOR_POS_BKST 18184 1(OUT) 17834 129 2021/12/06 09:46:24.975473 [PANEL] PANEL_UNLOAD_PANEL_REQ [RELEASE] 2021/12/06 09:46:24.984781 COUCH_STOP_HOR_JOG_REQ 2021/12/06 09:46:25.002961 COUCH_STOP_VERT_JOG_REQ 2021/12/06 09:46:25.044888 COUCH_REPORT_HOR_POS_BKST 18561 1(OUT) 12924 129 2021/12/06 09:46:25.296010 COUCH_REPORT_HOR_POS_BKST 18811 1(OUT) 7962 129 2021/12/06 09:46:25.547028 COUCH_REPORT_HOR_POS_BKST 18936 1(OUT) 2988 129 2021/12/06 09:46:25.798474 COUCH_REPORT_HOR_POS_BKST 18952 1(OUT) 3 133 2021/12/06 09:46:26.049129 COUCH_REPORT_HOR_POS_BKST 18952 1(OUT) 3 133 2021/12/06 09:46:26.300214 COUCH_REPORT_HOR_POS_BKST 18952 1(OUT) 6 133 2021/12/06 09:46:26.551292 COUCH_REPORT_HOR_POS_BKST 18952 1(OUT) 3 133 2021/12/06 09:46:26.802750 COUCH_REPORT_HOR_POS_BKST 18952 1(OUT) 9 133 2021/12/06 09:46:26.996509 **RHOST_ESTOP_OPENED_BKST** 2021/12/06 09:46:26.996799 **GHOST_ESTOP_OPENED_BKST** 2021/12/06 09:46:26.997096 **DMC_ESTOP_OPENED_BKST** 2021/12/06 09:46:26.998104 **GANTRY_ESTOP_OPENED_BKST** 2021/12/06 09:46:27.006350 **COUCH_ESTOP_OPENED_BKST** 2021/12/06 09:46:27.008788 COUCH_REPORT_HOR_POS_BKST 18952 1(OUT) 9 389 2021/12/06 09:46:27.023419 COUCH_REPORT_VERT_POS_BKST 1066 1(DOWN) 1 2021/12/06 09:46:27.163575 ==COUCH_ESTOP_ERROR_BKST== S_COUCHMGR_HOR_MOTION_STOP_ERROR 2021/12/06 09:46:27.176639 ==GANTRY_ESTOP_ERROR_BKST== S_COUCHMGR_HOR_MOTION_STOP_ERROR 2021/12/06 09:46:30.051295 <<NORMAL_ERROR>> S_MANMOT_HORIZONTAL_STOP_ERROR 2021/12/06 09:46:30.064170 <<NORMAL_ERROR>> S_MANMOT_HORIZONTAL_STOP_ERROR 2021/12/06 09:46:30.350342 <<GANTRY_EAL_REPORT_ANS>> GHOST EAL_NORMAL S_MANMOT_HORIZONTAL_STOP_ERROR 2021/12/06 09:47:06.703867 COUCH_REPORT_HOR_POS_BKST 18951 0(IN) 0 260 2021/12/06 09:47:06.955585 COUCH_REPORT_HOR_POS_BKST 18948 0(IN) 0 260 2021/12/06 09:47:07.207334 COUCH_REPORT_HOR_POS_BKST 18779 0(IN) 0 260 ________________________________________________________________ Follow Up Required : CUSTOMER REPORTED PROBLEM OR EVENT DESCRIPTION: 6 Dec 2021 The Customer called to report on behalf of a colleague that the Gantry turned off using the UNLOAD option. The issue occurred at approx. 10:05 AM and was resolved by turning the gantry back on. TROUBLESHOOTING - 6 Dec 2021: Informed the Customer that would download and review the logs. Logs confirm that the UNLOAD button was pressed at 09:46:17.145300 AM 6 Dec 2021 and released at 09:46:24.975473 AM. The logs indicate that the Gantry turned off shortly after at 09:46:26.996509 AM due to ESTOP error "S_COUCHMGR_HOR_MOTION_STOP_ERROR " at 09:46:27.183 AM followed by Normal type Error "S_MANMOT_HORIZONTAL_STOP_ERROR" at 09:46:30.087 AM. The errors appear to indicate that the Couch did not stop in time. Informed the Customer of the findings. The Customer confirmed with the colleague unloading the patient, that the patient was not moving unexpectedly at the time. NEXT STEP ACTION: An onsite work order is to be dispatched to coincide with PM work order WO-02519912 at Citiscan at 12:30 16th Dec 2021. Requested the attending Engineer to please inspect the couch mechanics in terms of horizontal travel as mentioned in emails. Included the AEL recommendations associated with error S_COUCHMGR_HOR_MOTION_STOP_ERROR that may help the Engineer focus on the areas of concern. The Customer has been made aware of the action plan via email. ----------------------------------------------------------------------------- | Howard, Amanda |  | Performed patient support cable checks - all ok | [12 empty] | 118839759 | Closed | 1588236 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Gantry off when upload | [5 empty] | IB Exist | bar: No
1679 |  |  | United States | Philips Medical Capital | 300057 | 44533 |  | Ring artifacts appear in images, replaced 459800594811 TDMS Module Class A FRU | Internal Remarks : approved: onsite is covered by risk sharing model. Internal Remarks : Interaction New Case Caller: John Williams IP & Issue: 54851181 FSE callback for NSS Event Type: Incident Service Request: Onsite Did Customer Decline Remote Support? Yes Preferred Service Time: N/A Entitlement Options: Entitled Reason for Entitlement Claim: N/A Action Taken: warm transferred to NSS Next Step Action: N/A Verbal PO Number: N/A Verbal PO Approver: N/A Verbal PO CAP: N/A ----------------------------------------------------------------------------- Work Order #: WO-04927213 14-Dec-2021 Williams John The system meets the specification for the performed service and is returned to use Problem Description by Engineer : Problem description by engineer : Ring artifacts on scans Information to support the complaint handling process Customer Function/Role: Not provided Device Use: Not provided Expected and Unexpected results: Not provided Patient Impact/Outcome: Not provided User Impact: Not provided Current Software Version: Not provided Parts Consumed : 1.00000,459800594811,TDMS Module Class A FRU Resolution : Replace detector module ----------------------------------------------------------------------------- Work Order #: WO-04910786 08-Dec-2021 Williams John The system meets the specification for the performed service and is returned to use Resolution : noisy fan Problem Description by Engineer : Problem description by engineer : Ring artifacts on scans Information to support the complaint handling process Customer Function/Role: Not provided Device Use: Not provided Expected and Unexpected results: Not provided Patient Impact/Outcome: Not provided User Impact: Not provided Current Software Version: Not provided Parts Consumed : 1.00000,453566503142,Fan, Tube Axial ----------------------------------------------------------------------------- Work Order #: WO-04910787 07-Dec-2021 Williams John The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Follow Up Required : Replace module 40 Problem Description by Engineer : Problem description by engineer : Ring artifacts on scans Information to support the complaint handling process Customer Function/Role: Not provided Device Use: Not provided Expected and Unexpected results: Not provided Patient Impact/Outcome: Not provided User Impact: Not provided Current Software Version: Not provided Diagnostic performed by Engineer : Bad detector module ----------------------------------------------------------------------------- Work Order #: WO-04907471 03-Dec-2021 Williams John The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Problem Description by Engineer : Problem description by engineer : Ring artifacts on scans Information to support the complaint handling process Customer Function/Role: Not provided Device Use: Not provided Expected and Unexpected results: Not provided Patient Impact/Outcome: Not provided User Impact: Not provided Current Software Version: Not provided Follow Up Required : Replace module 37 Follow Up Required : Replace Noisy Fan Diagnostic performed by Engineer : Rings in image ----------------------------------------------------------------------------- Work Order #: WO-04907434 Powell, Jamie The system does not meet full specification for the performed service and is returned to use with limitation as accepted by the Customer, Follow-up work scheduled Internal Comments : Looked at image and used ring ID.. Looks like M6 or M37 is bad. Internal Comments : Error # and/or description of error : Ring artifacts on scans Repair Action : Ring artifacts on scans Follow Up Required : Error # and/or description of error : Ring artifacts on scans Repair Action : Ring artifacts on scans Looked at image and used ring ID.. Looks like M6 or M37 is bad. Technician Name : Powell, Jamie Problem Description by Engineer : Problem description by engineer : Ring artifacts on scans Information to support the complaint handling process Customer Function/Role: Not provided Device Use: Not provided Expected and Unexpected results: Not provided Patient Impact/Outcome: Not provided User Impact: Not provided Current Software Version: Not provided Diagnostic performed by Engineer : Ring artifacts on scans.. Customer states that the last air cal was ran on monday.. Suggested customer run another but she has scans coming up and will run one when she can. ----------------------------------------------------------------------------- | Kanzeg, Mike |  | Rplace module | [12 empty] | 118836197 | Closed | 1584749 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | Artifact | [5 empty] | IB Exist | 459800594811 | TDMS Module Class A FRU | bar: No
1680 |  |  | China | 阳高县医疗集团(阳高县人民医院 | 345025 | 44556 |  | 05K For IB/Core128/Gantry noise |  | 3 |  | Replace the main belt | [12 empty] | 118957042 | Closed | 1654215 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | Engineering | [6 empty] | Gantry noise when rotation | [5 empty] | IB Exist | 453566492401 | POLY-VEE BELT | bar: No
1681 |  |  | China | 四川省宜宾市第二中医院 | 345035 | 44556 |  | 02N/Ingenuity Core128/Elite China/couch vertical problem |  | 3 |  | Replace vertical transducer, Couch vertical re-calibration | [12 empty] | 118957019 | Closed | 1654142 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | Engineering | [6 empty] | S_Master_sequencer_vert_couch_pos_err | [5 empty] | IB Exist | 459800623331 | FRU, Vertical Transducer | bar: No
1682 |  |  | China | 四川省宜宾市第二中医院 | 345035 | 44556 |  | 02N/Ingenuity Core128/Elite China/noise with rotation |  | 3 |  | Replae the main belt | [12 empty] | 118956787 | Closed | 1654084 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | Engineering | [6 empty] | Gantry noise when rotation | [5 empty] | IB Exist | 453566492401 | POLY-VEE BELT | bar: No
1683 |  |  | United States | Willis Knighton Hospital | 40004 | 44556 |  | Local drive is full |  | 2 | [14 empty] | 118957431 | Fixed | 1653957 |  |  | Opened | Opened | CTUDT00632261 | [5 empty] | Engineering | [6 empty] | Disk full | [5 empty] | Refresh |  |  |  | System | Console-SW | DiskFull | Console-SW-DiskFull | Console-SW-DiskFull | CTUDT00632261 | DiskFull | Console-SW | Console-SW-DiskFull | bar: No
1684 |  |  | Portugal | Centro Hospitalar de Vila Nova | 40030 | 44552 |  | TAC do SU está a dar erros de reconstrução. Está a trabalhar mas tem dado erros. (It is making reconstruction errors. ) Radiologia Convencional- U1 |  | 3 |  | Reconstructor software error when preparing the data receipt. Indication to the operator to restart the CIRS server. Trouble disappeared after CIRS restarted. There_s no report of a new occurrence. No follow-up required. | [12 empty] | 118944241 | Closed | 1649021 |  |  | Opened - Pending Translation | Opened - Pending Translation | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | CIRS Reconstruction error | [5 empty] | Refresh |  |  |  | System | CIRS-SW | NoEnoughInfo | CIRS-SW-NoEnoughInfo | CIRS-SW-NoEnoughInfo | NoEnoughInfo | NoEnoughInfo | CIRS-SW | CIRS-SW-NoEnoughInfo | bar: No
1685 |  |  | Argentina | INSTITUTO NACIONAL DE | 40027 | 44552 |  | NO SE PUEDE INICIAR EL CUADRO DE RECONSTRUCCIÓN (cirs). CAMBIE LOS PARÁMETROS DE EXPLORACIÓN |  | 2 | [14 empty] | 118945072 | Closed | 1648159 |  |  | Opened - Pending Translation | Opened - Pending Translation | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | CIRS rebuild table fail | [5 empty] | Refresh |  |  |  | System | CIRS-SW | NoEnoughInfo | CIRS-SW-NoEnoughInfo | CIRS-SW-NoEnoughInfo | NoEnoughInfo | NoEnoughInfo | CIRS-SW | CIRS-SW-NoEnoughInfo | bar: No
1686 |  |  | China | 莱阳市人民医院 | 345005 | 44552 |  | WeChat:05L/Ingenuity Core128/Elite China/No heart rate display |  | 3 |  | instruct to customer the rate can see in scan | [12 empty] | 118941802 | Closed | 1647343 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | Customer | [6 empty] | no heart rate display | [5 empty] | Refresh | bar: No
1687 |  |  | Japan | 社会福祉法人恩賜財団済生会 | 343000 | 44553 |  | Loadフットスイッチを操作した瞬間E-STOP OPENとなった |  | 4 WANG LIN, LV |  | Suspect cables block the Moog light | [12 empty] | 118946610 | Closed | 1646610 |  |  | Investigation - Pending Translation | Investigation - Pending Translation | [6 empty] | Engineering | [6 empty] | S_CTRL_LINK_DMC_GOT_CRC_FRAMING_ERROR | [5 empty] | Non-Refresh | bar: No
1688 |  |  | China | 河南省濮阳市人民医院 | 40031 | 44552 |  | WeChat:05K for IB/Ingenuity CT 728326/Scan module damaged |  | 3 |  | Replace module | [12 empty] | 118940433 | Closed | 1646592 |  |  | Opened - Pending Translation | Opened - Pending Translation | CTUDT00632110 | [5 empty] | Engineering | [6 empty] | dms module | [5 empty] | Refresh | 459800911281 | TDMS V2.X M.Det Class B FRU |  | System | DMS-HW | DMS Module | DMS-HW-DMS Module | DMS-HW-DMS Module | CTUDT00632110 | DMS Module | DMS-HW | DMS-HW-DMS Module | bar: No
1689 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 337000 | 44552 |  | Masa hareketlerinde ses geliyor (Sounds on the desk movements) |  | 3 |  | When the mechanical parts of the table were checked, the sound was detected from the bed motion brake. Replacing brake cleanup sets the problem on the device | [12 empty] | 118941867 | Fixed | 1643701 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | Engineering | [6 empty] | Noise on couch vertical movement | [5 empty] | Non-Refresh | bar: No
1690 |  |  | Poland | Government Agency. | 341009 | 44551 |  | Integracja TK z systemem HIS |  | 4 | [14 empty] | 118935452 | In Progress | 1643293 |  |  | Opened | Opened | [6 empty] | TBD | [6 empty] | TBD | [5 empty] | TBD | bar: No
1691 |  |  | Japan | Maki Hospital | 341017 | 44551 |  | 本日午前中にパスワードがロックされた件に関して(With the password locked in the morning ) |  | 4 Yang, Cancan | [14 empty] | 118934845 | Closed | 1642194 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Customer | [6 empty] | Password locked | [5 empty] | Refresh | [6 empty] | - | - | 0 | 0 | 0 | - | bar: No
1692 |  |  | United States | Philips Medical Capital | 40053 | 44550 |  | slow application when doing coronary scans. |  | 3 WANG LIN, LV | [14 empty] | 118932562 | In Progress | 1641742 |  |  | Investigation | Investigation | [6 empty] | TBD | [6 empty] | TBD | [5 empty] | TBD | bar: No
1693 |  |  | Puerto Rico | METRO MAYAGUEZ INC | 40009 | 44551 |  | Table Joy stick does not work |  | 3 |  | The Joystick was not the problem. The problem was caused by the cable JOYSTICK PANEL MUT CONN. HARNESS. PN: 453567542042. | [12 empty] | 118937100 | Closed | 1641475 |  |  | Evaluation | Evaluation | [6 empty] | Engineering | [6 empty] | Joystick not work | [5 empty] | IB Exist | 453567542042 | CABLE, JOYSTICK PANEL MNT CONN. HARNESS | bar: No
1694 |  |  | Poland | Rzadowa Agencja | 341008 | 44550 |  | Integracja TK z systemem HIS |  | 4 | [14 empty] | 118929865 | Closed | 1639653 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Other | [6 empty] | Other | [5 empty] | Non-Refresh | bar: No
1695 |  |  | France | CLINIQUE DE FLANDRE | 40019 | 44551 |  | problème de mot de passe, quand il faut entrer le nouveau, ça a bloqué au bout de plusieurs essais. (Password problem. When a new password needs to be entered, it is blocked after multiple attempts ) |  | 3 | [14 empty] | 118934988 | Closed | 1639157 |  |  | Opened | Opened | [6 empty] | Customer | [6 empty] | Password locked | [5 empty] | Refresh | [6 empty] | - | - | 0 | 0 | 0 | - | bar: No
1696 |  |  | Turkey | ADIYAMAN OZEL SAGINSTURTAS SAN VE TIC A.S | 341006 | 44551 |  | Mouse siparişi |  | 3 | [14 empty] | 118935314 | Closed | 1638834 |  |  | Evaluation | Evaluation | [6 empty] | Customer | [6 empty] | Customer order the part | [5 empty] | IB Exist | 459801378751 | White 1000 DPI 3 Button Mouse | bar: No
1697 |  |  | China | 江苏大学附属医院 | 345023 | 44550 |  | PRS Connection/02N/Ingenuity Core128/Elite China |  | 5 Wang, Elaine |  | Add dual IP adress | [12 empty] | 118926310 | Closed | 1637996 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Service | [6 empty] | PRS not connect | [5 empty] | Refresh | [6 empty] | - | - | 0 | 0 | 0 | - | bar: No
1698 |  |  | Japan | 医療法人清翠会 牧病院 | 341017 | 44550 |  | パスワードを間違えロックされてしまった。>>ユーザーitでログインいただきロック解除していただきました。(The password has been mistaken.) > > I locked it with the user it) |  | 4 Yang, Cancan | [14 empty] | 118932927 | Closed | 1637976 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Customer | [6 empty] | Password locked | [5 empty] | Refresh | [6 empty] | - | - | 0 | 0 | 0 | - | bar: No
1699 |  |  | China | DaKuai Town Health, Fengquan District, Xinxiang City, Henan Province | 345009 | 44550 |  | 05K for IB/Ingenuity Core128/Elite China/consult information display |  | 3 Wang, Elaine |  | customer consult print information display issue, answer customer_s question | [12 empty] | 118926167 | Closed | 1637881 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Customer consult printer setting | [5 empty] | Non-Refresh | bar: No
1700 | Date0103 |  | China | 陕西省咸阳中心医院 (咸阳市第二人民医院） | 40050 | 44562 |  | wechat:05R/Ingenuity CT 728326 |  | 3 | [14 empty] | 118990920 | New | 1669486 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | TBD | bar: No
1701 |  |  | United States | Philips Medical Capital | 40055 | 44561 |  | 0711CDN83Gantry Initialize Errors/Please Restart the Gantry |  | 2 | No bugreport yet | None | [12 empty] | 118990606 | Closed | 1669474 |  |  | Opened | Opened | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Intermittent gantry initial error | [5 empty] | Refresh |  |  |  | System | Gantry-SW | NoEnoughInfo | Gantry-SW-NoEnoughInfo | Gantry-SW-NoEnoughInfo | NoEnoughInfo | NoEnoughInfo | Gantry-SW | Gantry-SW-NoEnoughInfo | bar: No
1702 |  |  | United States | Philips Medical Capital | 40028 | 44562 |  | 0505CFXXX Gantry not coming on |  | 2 |  | Recycle the E-stop button on CT BOX | [12 empty] | 118991931 | Closed | 1669309 |  |  | Evaluation | Evaluation | [6 empty] | Customer | [6 empty] | E-stop closing | [5 empty] | IB Existing | bar: No
1703 |  |  | France | CLINIQUE DE FLANDRE | 40019 | 44561 |  | Le Host ne redémarre plus | Display failure when host restarts. After stopping the host on the "stop" button on the front panel, the host will not restart | 2 |  | FSE identified the display displaying issue when restar the host, no further information about what action done for the display. | [12 empty] | 118987426 | Closed | 1667428 |  |  | Opened | Opened | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Didplaying | [5 empty] | Refresh |  |  |  | System | Console-SW | NoEnoughInfo | Console-SW-NoEnoughInfo | Console-SW-NoEnoughInfo | NoEnoughInfo | NoEnoughInfo | Console-SW | Console-SW-NoEnoughInfo | bar: No
1704 |  |  | China | 宁安市第二人民医院 | 345039 | 44560 |  | 02N/Ingenuity Core128/add the PACS |  | 5 |  | FSE to add the PACS as requested | [12 empty] | 118975961 | Closed | 1667101 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Add PACS | [5 empty] | Non-refresh | bar: No
1705 |  |  | United States | Philips Medical Capital | 40005 | 44559 |  | Q/O Thumping and squeaking noise generated by drive belt |  | 5 |  | Replace the main belt and tensioner | [12 empty] | 118972284 | Closed | 1665316 |  |  | Evaluation | Evaluation | [6 empty] | Engineering | [6 empty] | Main belt noise | [5 empty] | IB Existing | 453567003451, 453566492401 | TENSIONER ASSY, POLY-VEE BELT | bar: No
1706 |  |  | China | 河南许昌市中心医院(传染病) | 40047 | 44559 |  | 02N/Ingenuity CT 728326/disk space problem |  | 3 |  | Performed visual inspection ,found useless files, delete the useless file | [12 empty] | 118970175 | Closed | 1663990 |  |  | Opened - Pending Translation | Opened - Pending Translation | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Disk full | [5 empty] | Refresh |  |  |  | System | TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | bar: No
1707 |  |  | China | DaKuai Town Health, Fengquan District, Xinxiang City, Henan Province | 345009 | 44559 |  | 02N/Ingenuity Core128/Elite China |  | 3 Wang, Elaine |  | N/A | [12 empty] | 118969785 | Closed | 1663825 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Customer | [6 empty] | 神飞云工作站 | [5 empty] | Non-refresh | bar: No
1708 |  |  | United States | Unity Health Systems | 40037 | 44558 |  | MONTHLY QA |  | 5 Dantha, Rajashekar |  | N/A | [12 empty] | 118968008 | Closed | 1662882 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Monthly QA | [5 empty] | Non-refresh | bar: No
1709 |  |  | United States | Odessa Regional Medical Center | 40029 | 44558 |  | infant EC questions |  | 4 WANG LIN, LV |  | advised customer to hand inject and scan with 5 sec delay. Follow up with customer. Scan went well, but she will scan immediately after injection next time | [12 empty] | 118966789 | Closed | 1662008 |  |  | Closed - Done | Closed - Done | [6 empty] | Customer | [6 empty] | Clinical consult of the injection delay for infant | [5 empty] | Non-refresh | bar: No
1710 |  |  | China | 苏州广慈肿瘤医院有限公司 | 345019 | 44558 |  | 02N/Ingenuity Core128/Elite China/Quality Control |  | 5 |  | Do the image calibration | [12 empty] | 118965430 | Closed | 1661057 |  |  | Opened - Pending Translation | Opened - Pending Translation | Calibration | [5 empty] | Engineering | [6 empty] | Low contract resolution failure | [5 empty] | Refresh |  |  |  | System | NotEngineering | Calibration | NotEngineering-Calibration | NotEngineering-Calibration | Calibration | Calibration | NotEngineering | NotEngineering-Calibration | bar: No
1711 |  |  | China | Sichuan Province Yibin City No.2 T.C.M. Hospital | 345035 | 44558 |  | 02N/Ingenuity Core128/Elite China |  | 5 Wang, Elaine |  | N/A | [12 empty] | 118964516 | Closed | 1660536 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Add the node | [5 empty] | Non-refresh | bar: No
1712 |  |  | China | Shaanxi Province Ankang T.C.M. Hospital | 345041 | 44558 |  | WeChat:02N/Ingenuity Core128/Elite China/Scan fault |  | 3 WANG LIN, LV |  | delete patient data, and do full rebuild | [12 empty] | 118963411 | Closed | 1660062 |  |  | Evaluation | Evaluation | SystemStuck | [5 empty] | Engineering | [6 empty] | Scan fault, software stuck | [5 empty] | Refresh |  |  |  | System | Console-SW | SystemStuck | Console-SW-SystemStuck | Console-SW-SystemStuck | SystemStuck | SystemStuck | Console-SW | Console-SW-SystemStuck | bar: No
1713 | Date0110 |  | China | 内蒙古兴安盟人民医院 | 345030 | 44570 |  | 05K for IB/Ingenuity Core128/Elite China/gantry initialize failed |  | 3 |  | On site check the scan parameter, change the scan parameter | [12 empty] | 119025216 | Closed | 1690655 |  |  | Opened - Pending Translation | Opened - Pending Translation | NotEngineering | [5 empty] | Engineering | [6 empty] | Gantry initialize failure | [5 empty] | Refresh |  |  |  | System | NotEngineering | ParameterSetting | NotEngineering-ParameterSetting | NotEngineering-ParameterSetting | NotEngineering | ParameterSetting | NotEngineering | NotEngineering-ParameterSetting | bar: No
1714 |  |  | Australia | Maia Financial Pty Limited | 341000 | 44570 |  | Gantry Acquisition cannot comply failure, system is down. |  | 2 |  | Replace 453567538334,MODULE, ROTOR ELECTRONICS TX replacement | [12 empty] | 119026083 | Closed | 1690554 |  |  | Opened | Opened | [6 empty] | Engineering | [6 empty] | Gantry acquisition cannot comply | [5 empty] | IB Existing | 453567538334 | MODULE, ROTOR ELECTRONICS TX | bar: No
1715 |  |  | China | 哈尔滨市森海医院 | 345007 | 44569 |  | 01/Ingenuity Core128/Elite China/password was locked |  | 3 |  | instruct customer wait more than 20 min to log in again. make a change password | [12 empty] | 119023787 | Closed | 1690098 |  |  | Opened - Pending Translation | Opened - Pending Translation | CTUDT00632236 | [5 empty] | Engineering | [6 empty] | Password locked | [5 empty] | Refresh |  |  |  | System | Console-SW | User Account | Console-SW-User Account | Console-SW-User Account | CTUDT00632236 | User Account | Console-SW | Console-SW-User Account | bar: No
1716 |  |  | Japan | 医療法人清翠会 牧病院 | 341017 | 44569 |  | データ転送が途中で停止。Logout/Loginでは解消されずコンピュータの再起動で復旧した |  | 4 Zhu, Chun |  | Reboot the computer | [12 empty] | 119023916 | Closed | 1688848 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | Data transfer stop | [5 empty] | Refresh |  |  |  | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | bar: No
1717 |  |  | Germany | Gesundheit Nord gGmbH Klinikverbund | 40002 | 44568 |  | FM: Fehlfunktion an der Temperaturregelung , bitte an FSE wenden (Error message with Air Calibration ) |  | 3 |  | FRU replaced | [12 empty] | 119019245 | Closed | 1688749 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | Engineering | [6 empty] | DMS Temperature | [5 empty] | IB Existing | 459800325761, 453567488921, 459800624062, | Brush, Signal Block Assembly; Power Supply Assy / Single; TDMS Cover RoHS FRU | bar: No
1718 |  |  | China | 四川省宜宾市第二中医院 | 345035 | 44568 |  | 02N/Ingenuity Core128/Elite China/couch vertical error |  | 4 |  | Replace the couch vertical moter brake | [12 empty] | 119018734 | Closed | 1688529 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | Engineering | [12 empty] | IB Existing | bar: No
1719 |  |  | China | 无锡市江原医院 | 345020 | 44567 |  | 05K for IB/Ingenuity Core128/Elite China |  | 3 |  | None | [12 empty] | 119013713 | Closed | 1685160 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | TBD | [6 empty] | TBD | [5 empty] | TBD | bar: No
1720 |  |  | United States | Central Vermont Medical Center | 40024 | 44567 |  | Error stating the recon is not initialized. |  | 3 |  | Flashed the DMC | [12 empty] | 119014761 | Closed | 1682144 |  |  | Evaluation | Evaluation | NotEngineering | [5 empty] | Engineering | [6 empty] | Recon is not initialized | [5 empty] | Refresh |  |  |  | System | NotEngineering | DMC | NotEngineering-DMC | NotEngineering-DMC | NotEngineering | DMC | NotEngineering | NotEngineering-DMC | bar: No
1721 |  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 337000 | 44567 |  | E- stop ilerlemiyor (E-stop not progressing) |  | 2 |  | Slip ring clean up | [12 empty] | 119014667 | Closed | 1682141 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | Engineering | [6 empty] | E-stop cannot be closed | [5 empty] | IB Existing | bar: No
1722 |  |  | China | Xinjiang Province Yili Prefecture Friendship Hospital | 40048 | 44566 |  | 02N/Ingenuity CT 728326/gantry fault |  | 3 |  | Reboot gantry and host | [12 empty] | 119006027 | Closed | 1680588 |  |  | Opened | Opened | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Gantry cannot scan | [5 empty] | Refresh |  |  |  | System | Gantry-SW | NoEnoughInfo | Gantry-SW-NoEnoughInfo | Gantry-SW-NoEnoughInfo | NoEnoughInfo | NoEnoughInfo | Gantry-SW | Gantry-SW-NoEnoughInfo | bar: No
1723 |  |  | Japan | 独立行政法人国立病院機構 | 341015 | 44567 |  | 04DS:KK:リングアーチファクト発生。 |  | 3 Yang, Cancan |  | Replace the module | [12 empty] | 119012031 | Closed | 1680466 |  |  | Evaluation - Pending Translation | Evaluation - Pending Translation | [6 empty] | Engineering | [6 empty] | Ring artifact | [5 empty] | IB Existing | 459800074772 | BR.64 TDMS V2 (RoHS FRU) | bar: No
1724 |  |  | China | Shaanxi Province Ankang T.C.M. Hospital | 345041 | 44565 |  | Application/Ingenuity Core128/Elite China |  | 5 |  | N/A | [12 empty] | 118998684 | New | 1675936 |  |  | Opened | Opened | [6 empty] | TBD | [6 empty] | TBD | [5 empty] | TBD | bar: No
1725 |  |  | United States | Philips Medical Capital | 40055 | 44564 |  | Abnormal Sound from Gantry |  | 3 Dantha, Rajashekar |  | Abnormal noise at high speed/ discolored track, found 3 foot section on belt track of rotor (V-Groove) has foreign material, cleaned 6 times, tested unable to hear afterwaards | [12 empty] | 118997843 | Closed | 1671626 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Gantry rotation noise | [5 empty] | IB Existing | bar: No
1726 |  |  | Peru | BANCO BBVA PERU | 40034 | 44564 |  | 40034-ERROR DE CALIBRACION Y DE IQ COLOCANDO TIEMPO DE ESPERA |  | 3 |  | The IQ check error message was due to the detector temperature. We provide instructions to perform this procedure in the future. | [12 empty] | 118996183 | Closed | 1671002 |  |  | Opened | Opened | NotEngineering | [5 empty] | Engineering | [6 empty] | DMS Temp error during IQ check | [5 empty] | Refresh |  |  |  | System | DMS-HW | Instruction | DMS-HW-Instruction | DMS-HW-Instruction | NotEngineering | Instruction | DMS-HW | DMS-HW-Instruction | bar: No
1727 | Date0117 |  | China | 河南省新乡市凤泉区大块镇卫生院 | 345009 | 44575 |  | 05K For IB/Ingenuity Core128/Elite China/Crash |  | 5 |  | just remote service . instruct the customer to delete patient data and restart computer | [12 empty] | 119059752 | Closed | 1711914 |  |  | Opened - Pending Translation | Opened - Pending Translation | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Host software stuck | [5 empty] | Refresh |  |  |  | System | Console-SW | NoEnoughInfo | Console-SW-NoEnoughInfo | Console-SW-NoEnoughInfo | NoEnoughInfo | NoEnoughInfo | Console-SW | Console-SW-NoEnoughInfo | bar: No
1728 |  |  | China | 中国医科大学附属盛京医院大连分院 | 345001 | 44575 |  | 05R/Ingenuity Core128/Elite China/xray error |  | 3 |  | Clean the HV connector, filament cal, LTC | [12 empty] | 119059781 | Closed | 1711928 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | Engineering | [6 empty] | ARC | [5 empty] | IB Existing | bar: No
1729 |  |  | China | 羊坊店医院 | 40035 | 44575 |  | 05K for IB/Ingenuity CT 728326/ |  | 3 |  | delete data and restart host | [12 empty] | 119059347 | Fixed | 1711720 |  |  | Opened - Pending Translation | Opened - Pending Translation | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Software system run slowly | [5 empty] | Refresh |  |  |  | System | Console-SW | NoEnoughInfo | Console-SW-NoEnoughInfo | Console-SW-NoEnoughInfo | NoEnoughInfo | NoEnoughInfo | Console-SW | Console-SW-NoEnoughInfo | bar: No
1730 |  |  | United States | Clay County Medical Center | 40003 | 44575 |  | Research FCO requirements for system. |  | 5 |  | Researched current software level (is V4.1.10.25202) and future implementation of FCO72800737 and FCO72800740 | [12 empty] | 119065701 | Closed | 1710741 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Engineering | [6 empty] | Implement FCO | [5 empty] | Refresh | bar: No
1731 |  |  | United States | Unity Health Systems | 40037 | 44575 |  | Gantry speaker not working |  | 3 |  | Two cables for intercom and speaker were disconnected. Reconnected them and verified normal operation of system | [12 empty] | 119064442 | Closed | 1709721 |  |  | Opened | Opened | [6 empty] | Other | [6 empty] | Gantry audio cables not connected | [5 empty] | IB Existing | bar: No
1732 |  |  | China | 林甸县医院 | 345038 | 44574 |  | 02N/Ingenuity Core128/Elite China/Change IP |  | 3 |  | New PACS/RIS has different IP, need change HOST and ISP IP | [12 empty] | 119051002 | Closed | 1707110 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Customer | [6 empty] | Change IP | [5 empty] | IB Existing | bar: No
1733 |  |  | China | Harbin City Sen Hai Hospital | 345007 | 44574 |  | 01/Ingenuity Core128/Elite China/Consulting the warranty |  | 3 |  | N/A | [12 empty] | 119050862 | Closed | 1706990 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Consolut Warranty | [5 empty] | IB Existing | bar: No
1734 |  |  | United States | Leesburg Regional Medical Center | 40022 | 44573 |  | 0705CDY73 Water spilled on keyboard. |  | 2 |  | Replace the keyboard | [12 empty] | 119049434 | Closed | 1705871 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Water spilled on the keyboard | [5 empty] | Non-refresh | bar: No
1735 |  |  | United States | Leesburg Regional Medical Center | 40022 | 44574 |  | Need XR-29 certificate |  | 5 |  | created PDF certificate . emailed to customer. Uploaded to TEAMS files | [12 empty] | 119055716 | Closed | 1704854 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Engineering | [6 empty] | XR 29 Request | [5 empty] | Refresh | [6 empty] | - | - | 0 | 0 | 0 | - | bar: No
1736 |  |  | United Kingdom | STOCKPORT NHS FOUNDATION TRUST | 40016 | 44574 |  | The table is not coming on, scanner down: error message saying that the system is not in a working state |  | 2 Suggest to close, not Refresh SW issue. |  | Reboot gantry, Host, CIRS | [12 empty] | 119054406 | Closed | 1704206 |  |  | Opened | Opened | NotEngineering | [5 empty] | Engineering | [6 empty] | System is not in a working state | [5 empty] | Refresh |  |  |  | System | NotEngineering | SeemsHWProblem | NotEngineering-SeemsHWProblem | NotEngineering-SeemsHWProblem | NotEngineering | SeemsHWProblem | NotEngineering | NotEngineering-SeemsHWProblem | bar: No
1737 |  |  | China | 济南市第四人民医院 | 40038 | 44573 |  | Appliaction/Ingenuity CT 728326 |  | 5 |  | TBD | [12 empty] | 119039999 | New | 1702553 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | TBD | [6 empty] | TBD | [5 empty] | TBD | bar: No
1738 |  |  | China | Nanchang University First Affiliated hospital (Jinggangshan Hospital) | 345002 | 44572 |  | 01/Ingenuity Core128/Elite China/add printer |  | 3 Zhai, Lisa |  | None | [12 empty] | 119033110 | Closed | 1698344 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Add printer | [5 empty] | IB Existing | bar: No
1739 |  |  | China | Hebei Province Tangshan City Workers Hospital | 40044 | 44572 |  | 02N/Ingenuity CT 728326/Maintenance work order content |  | 5 |  | TBD | [12 empty] | 119032969 | New | 1698263 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | TBD | [6 empty] | TBD | [5 empty] | TBD | bar: No
1740 |  |  | Australia | Maia Financial Pty Limited | 341000 | 44573 |  | Ring artifact |  | 2 WANG LIN, LV |  | Replace Module 31 | [12 empty] | 119039779 | Closed | 1698250 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Ring Artifact | [5 empty] | IB Existing | 459800911271 | TDMS V2 M.Det Class A FRU | bar: No
1741 |  |  | China | Shaanxi Province Ankang T.C.M. Hospital | 345041 | 44572 |  | 02N/Ingenuity Core128/Elite China/change print header |  | 3 Zhai, Lisa |  | N/A | [12 empty] | 119032918 | Closed | 1698226 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Change printer header | [5 empty] | IB Existing | bar: No
1742 |  |  | Japan | Maki Hospital | 341017 | 44569 |  | DICOM image transfer stopped |  | Lu, Xueying | [14 empty] | CTUDT00632127, CTUDT00631602. Wait for SP2 deploy.Ask for the detail SW version of the customer, if HF3 and SP1 have been implemented. |  | 1698189 |  |  | Investigation | Investigation | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | PACS transfer | [5 empty] | Refresh |  |  |  | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | bar: No
1743 |  |  | United States | Philips Medical Capital | 40028 | 44572 |  | Artifact in the field of view. Customer requesting onsite service in the morning. |  | 2 De la Zerda, Yotam |  | Replace Module 17 | [12 empty] | 119032657 | Closed | 1698133 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Artifact | [5 empty] | IB Existing | 459800911271 | TDMS V2 M.Det Class A FRU | bar: No
1744 |  |  | China | Hebei Province Tangshan City Workers Hospital | 40044 | 44572 |  | Wechat:02N/Ingenuity CT 728326/intercom fault |  | 3 |  | Reconnect the intercom cables | [12 empty] | 119032464 | Closed | 1698003 |  |  | Opened | Opened | [6 empty] | Others | [6 empty] | Gantry audio cables not connected | [5 empty] | IB Existing | bar: No
1745 |  |  | China | Zhalantun City People_s Hospital | 40045 | 44571 |  | 02N/Ingenuity CT/gantry rotor brake fault |  | 3 |  | REPLACE BRAKE ASSY FOR AC | [12 empty] | 119027577 | Closed | 1694155 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | Engineering | [6 empty] | Main motor problem | [5 empty] | IB Existing | 453566502551 | SPARE BRAKE ASSY FOR AC MOTOR | bar: No
1746 |  |  | China | Beijing City Hospital of Traditional Chinese and Western Medicine | 40032 | 44571 |  | 05K for IB/Ingenuity CT 728326/Gantry noise |  | 3 |  | Performed visual inspection, DMS fan | [12 empty] | 119026617 | Fixed | 1693379 |  |  | Opened | Opened | [6 empty] | Engineering | [6 empty] | Gantry noise | [5 empty] | IB Existing |  | DMS Fan | bar: No
1747 |  |  | United States | Philips Medical Capital | 40028 | 44571 |  | Partial Ring Artifact |  | 3 Smith-Kifus, Vikki |  | None | [12 empty] | 119031824 | Closed | 1692909 |  |  | Investigation | Investigation | [6 empty] | TBD | [6 empty] | Partial ring artifact | [5 empty] | IB Existing | bar: No
1748 |  |  | China | Jiangsu University Affiliated Hospital | 345023 | 44570 |  | Wechat:05K for IB/Ingenuity Core128/Elite China/Images cannot be automatically transferred to PACS |  | 3 Zhai, Lisa |  | check the auto storage configuration ，found the auto storage configuration changed, re-config the auto storage | [12 empty] | 119025632 | Closed | 1691281 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | Image cannot auto transferred to PACS | [5 empty] | Refresh |  |  |  | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | bar: No
1749 |  |  | United States | Ashland Hospital Corporation | 40011 | 44535 |  | E-stop can not be closed due to vertical potentiometer connection issue, FSE cleaned potentiometer connection |  | Suggest to closed this case. 2 Cohen, Shmuel |  | Cleaned the contacts and the string of the potentiometer. Ran the couch up and down multiple times and performed tests scans | [12 empty] | 118839314 | Closed | 1588094 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | E-stop | [5 empty] | IB Existing | bar: No
1750 | Date0124 |  | United States | Philips Medical Capital | 40055 | 44584 |  | Streak Artifact on Heads and Face studies |  | 2 |  | Cleaned all surfaces within the FOV and ran air calibration. The image artifact was eliminated and the system returned to the customer ready for use | [12 empty] | 119105972 | Closed | 1735649 |  |  | Opened | Opened | [6 empty] | Others | [6 empty] | Streak Artifact | [5 empty] | IB Existing | bar: No
1751 |  |  | China | 东南大学附属中大医院江北院区 | 345021 | 44583 |  | 05K/Ingenuity Core128/Elite China/report errors |  | 3 |  | on site check the couch statues，found the top-right cover was distorted，need to adjust | [12 empty] | 119104489 | Closed | 1735612 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | Engineering | [6 empty] | Couch movement stop | [5 empty] | IB Existing | bar: No
1752 |  |  | China | Shandong Province Tai_an City T.C.M. No. 2 Hospital | 345037 | 44583 |  | 02N/Ingenuity Core128/Elite China/ct box key pressed not ok |  | 3 |  | Replace CT Box | [12 empty] | 119103684 | Closed | 1735508 |  |  | Opened | Opened | [6 empty] | Engineering | [6 empty] | CT Box Button | [5 empty] | IB Existing | 459801705022 | CT BOX ASSEMBLY_BEST | bar: No
1753 |  |  | Japan | Independent administrative institution National Hospital Organization | 341015 | 44581 |  | WY:ディテクター交換作業 |  | 5 Xu, Xiaoting |  | Replace module | [12 empty] | 119090993 | Closed | 1730053 |  |  | Closed - Cancelled | Closed - Cancelled | [6 empty] | Engineering | [6 empty] | Rplace module | [5 empty] | IB Existing | 459800074772 | BR.64 TDMS V2 (RoHS FRU) | bar: No
1754 |  |  | China | 林甸县医院 | 345038 | 44581 |  | WeChat：application/Ingenuity Core128/Elite China |  | 3 |  | Clinical engineer feedback the customer | [12 empty] | 119089517 | In Progress | 1729229 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Customer sonsult for the head scan protocol | [5 empty] | IB Existing | bar: No
1755 |  |  | United States | Clay County Medical Center | 40003 | 44580 |  | Knee wedge pad and head/neck pad are coming apart. |  | 5 | [14 empty] | 119088067 | Closed | 1728208 |  |  | Evaluation | Evaluation | CushionBreakdown | [5 empty] | Engineering | [6 empty] | Head / Armreset Cushion break down | [5 empty] | Refresh | 453566502591, 459800042241 | Cushion-Head/ Armrest, PATIENT CUSHIONS EXP |  | System | Couch-HW | CushionBreakdown | Couch-HW-CushionBreakdown | Couch-HW-CushionBreakdown | CushionBreakdown | CushionBreakdown | Couch-HW | Couch-HW-CushionBreakdown | bar: No
1756 |  |  | United States | Clay County Medical Center | 40003 | 44580 |  | Replace CT Slicker on Couch |  | 5 Wedemeier, Tracy |  | Replace the couch cushion | [12 empty] | 119087728 | Closed | 1727994 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Table cushion | [5 empty] | IB Existing | 459801350221 | STANDARD CUSHION-47.24 INCH MATTRESS | bar: No
1757 |  |  | United States | Clay County Medical Center | 40003 | 44580 |  | Pads/cushions are already failing |  | CFM SHU, Tim | [14 empty] | CFM |  | 1726643 |  |  | Investigation | Investigation | CTD00006748 | [5 empty] | Engineering | [6 empty] | pad/ cushion | [5 empty] | IB Existing |  |  |  | System | Couch-HW | Accessories | bar: No
1758 |  |  | China | 陕西省安康中医医院 | 345041 | 44580 |  | wechat：02N/Ingenuity Core128/Elite China/report errors |  | 3 Cai, Cathy |  | Reboot gantry | [12 empty] | 119083073 | Closed | 1724894 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | X-ray system cannot comply | [5 empty] | Refresh |  |  |  | System | Gantry-HW | NoEnoughINfo | Gantry-HW-NoEnoughINfo | Gantry-HW-NoEnoughINfo | NoEnoughInfo | NoEnoughINfo | Gantry-HW | Gantry-HW-NoEnoughINfo | bar: No
1759 |  |  | United States | Central Vermont Medical Center | 40024 | 44580 |  | Q/O: Error stating the recon is not initialized. |  | 3 Wedemeier, Tracy |  | None, 1 time occur, no available bug report | [12 empty] | 119082878 | Closed | 1724656 |  |  | Investigation | Investigation | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Recon is not initialized | [5 empty] | Refresh |  |  |  | System | CIRS-SW | NoEnoughInfo | CIRS-SW-NoEnoughInfo | CIRS-SW-NoEnoughInfo | NoEnoughInfo | NoEnoughInfo | CIRS-SW | CIRS-SW-NoEnoughInfo | bar: No
1760 |  |  | United States | Clay County Medical Center | 40003 | 44580 |  | Operation Manual is not installed on system. |  | 5 |  | Installed Operation Manual on application software from CD disk | [12 empty] | 119088210 | Closed | 1724191 |  |  | Evaluation | Evaluation | [6 empty] | Service | [6 empty] | Operation Manual is not installed on system. | [5 empty] | Refresh | bar: No
1761 |  |  | United States | Unity Health Systems | 40037 | 44579 |  | 0304CDN8A Table making a rubbing sound when moving in |  | 3 |  | Located the noise coming from the clutch on the rear of the table. When the table moves in. the clutch seems to be wobbling. Since the system is not down and the noise can barely be heard, it will be kept monitored for now | [12 empty] | 119080646 | Closed | 1723393 |  |  | Evaluation | Evaluation | [6 empty] | Engineering | [6 empty] | Couch movement noise | [5 empty] | IB Existing | bar: No
1762 |  |  | China | 福建省立医院 | 345011 | 44579 |  | 05K For IB/Ingenuity Core128/Elite China/Check the machine |  | 5 Wang, Elaine |  | N/A | [12 empty] | 119075426 | Closed | 1720326 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Support customer for the quality check | [5 empty] | IB Existing | bar: No
1763 |  |  | China | Inner Mongolia Hinggan League People_s Hospital | 345030 | 44579 |  | 05K for IB/Ingenuity Core128/Elite China/image artifact |  | 3 |  | Add M12, channel 181 into the bad detector list | [12 empty] | 119075313 | Closed | 1720265 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | Engineering | [6 empty] | Ring artifact | [5 empty] | IB Existing | bar: No
1764 |  |  | China | Xinjiang Province Yili Prefecture Friendship Hospital | 40048 | 44578 |  | 02N/Ingenuity CT 728326/change HIS/RIS config |  | 3 Zhai, Lisa |  | None | [12 empty] | 119071417 | Closed | 1716866 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Change HIS/RIS config | [5 empty] | IB Existing | bar: No
1765 |  |  | Japan | Tomakomai City Hospital | 341014 | 44578 |  | SF:ガントリー回転時に異音 |  | 5 Xu, Xiaoting |  | Replace main belt and tensioner bracket | [12 empty] | 119070533 | Imcomplete | 1716355 |  |  | Evaluation | Evaluation | [6 empty] | Engineering | [6 empty] | Gantry rotation noise | [5 empty] | IB Existing | 453567178301, 453566492401, | Bracket, Motor to Tensioner; POLY-VEE BELT | bar: No
1766 |  |  | Japan | Maki Hospital | 341017 | 44578 |  | 03CS:KKI:撮影中にエラー発生＞＞szInfo: SoftArcReg 0x211* /S_RHOST_HVG_HSS_GENERATOR_ARC_FAULT |  | 4 WANG LIN, LV |  | Longtc (removal of inert gas in X-ray tube) was carried out for the error caused by the discharge in the X-ray tube. | [12 empty] | 119069836 | Closed | 1715961 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Arc | [5 empty] | IB Existing | bar: No
1767 |  |  | China | Le qing City People_s Hospital | 40054 | 44578 |  | 02N/Ingenuity CT 728326/Dual display |  | 4 |  | Replace video adapter cable, clean software junk files. | [12 empty] | 119069836 |  | 1715767 |  |  | Opened | Opened | CTUDT00632173 | [5 empty] | Engineering | [6 empty] | Dual monitor didplay | [5 empty] | Refresh |  |  |  | System | Console-HW | Monitor | Console-HW-Monitor | Console-HW-Monitor | CTUDT00632173 | Monitor | Console-HW | Console-HW-Monitor | bar: No
1768 |  |  | United States | Central Vermont Medical Center | 40024 | 44567 |  | Error stating the recon is not initialized. CFM and bug report are required to check the detail of the recon error. |  | CFM E, Sravanthi |  | TBD | [12 empty] | CFM |  | 1715357 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | TBD | [6 empty] | TBD | [5 empty] | TBD | bar: No
1769 |  |  | Portugal | Centro Hospitalar de Vila Nova | 40030 | 44578 |  | O TAC da Urgência sobreaquece sem motivo aparente, sendo que este tipo de anomalia foi já identificado várias vezes |  | 3 |  | User messages stating that the acquisition parameters exceeded the generator and/or tube capacity. Customer has to decrease acquisition parameters in order to continue | [12 empty] | 119072657 | Closed | 1715261 |  |  | Evaluation | Evaluation | [6 empty] | Engineering | [6 empty] | User message: exceed the tube capacity | [5 empty] | IB Existing | bar: No
1770 |  |  | China | DaKuai Town Health, Fengquan District, Xinxiang City, Henan Province | 345009 | 44577 |  | 05K For IB/Core128/ESTOP opened |  | 3 Cai, Cathy |  | Ask customer to check the foot switch | [12 empty] | 119068537 | Closed | 1714527 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | E-stop open | [5 empty] | IB Existing | bar: No
1771 |  |  | Japan | 医療法人清翠会 牧病院 | 341017 | 44578 |  | 【CT:APP:NO】 |  | 5 Zhu, Chun |  | Since the parameter cannot be changed, replace the protocol | [12 empty] | 119071342 | Closed | 1714420 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Consolut for protocol parameter change | [5 empty] | IB Existing | bar: No
1772 |  |  | China | Shandong Province Chest Hospital | 40025 | 44577 |  | 05L/Ingenuity CT 728326/blue screen Silva will check for the similar blue screen during CT login case. CTUDT00632553 |  | 2 |  | Login in Software use smart card and configuration | [12 empty] | 119068325 | Closed | 1714155 |  |  | Opened | Opened | CTUDT00632553 | [5 empty] | Engineering | [6 empty] | Blue screen when login CT. | [5 empty] | Refresh |  |  |  | System | Console-SW | Blue-Screen | Console-SW-Blue-Screen | Console-SW-Blue-Screen | CTUDT00632553 | Blue-Screen | Console-SW | Console-SW-Blue-Screen | bar: No
1773 |  |  | China | Inner Mongolia Hinggan League People_s Hospital | 345030 | 44577 |  | 05K for IB/Ingenuity Core128/Elite China/wrong password |  | 3 Zhai, Lisa |  | Notify customer wait for half hour, then re-enter password. | [12 empty] | 119067917 | Closed | 1713864 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Customer | [6 empty] | Password lock | [5 empty] | Refresh | bar: No
1774 |  |  | United States | Unity Health Systems | 40037 | 44575 |  | MONTHLY QA |  | 5 Dantha, Rajashekar |  | N/A | [12 empty] | 119066415 | Closed | 1713419 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Monthly QA | [5 empty] | Non-refresh | bar: No
1775 |  |  | China | China Medical University Affiliated Shengjing Hospital Dalian Branch | 345001 | 44575 |  | 05R/Ingenuity Core128/Elite China/cushion |  | 3 |  | None | [12 empty] | 119062453 | Closed | 1712212 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Customer | [6 empty] | Customer need cushion | [5 empty] | Non-refresh | 453567517311 | Cushion, Patient Support iCT 2 | bar: No
1776 | Date0103U |  | United States | Select Specialty Hospitals | 310051 | 44559 |  | in the middle of procedure , needing clinical support ASAP |  | 1 Mohanta, Deepranjan |  | N/A | [12 empty] | 118972702 | Closed | 1662461 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Need clinical support | [5 empty] | Non-refresh | bar: No
1777 | Date0110U |  | United States | Adventhealth Zephyrhills | 300106 | 44568 |  | Issue with perfusion mapping software |  | 5 |  | Russ indicated that post processing absolute value graph was not acquiring the necessary data per the Radiologist on CT Brain Perfusion studies. It was explained to Russ that the interscan delay is too long to acquire these values due to the size of the detector array and the interscan delay of the jog cycle. | [12 empty] | 119021901 | Closed | 1687244 |  |  | Opened | Opened | NotEngineering | [5 empty] | Engineering | [6 empty] | CT Brain Perfusion study post processing absolute value graph was not acquiring the necessary data per the Radiologist | [5 empty] | Refresh |  |  |  | System | NotEngineering | ParameterSetting | NotEngineering-ParameterSetting | NotEngineering-ParameterSetting | NotEngineering | ParameterSetting | NotEngineering | NotEngineering-ParameterSetting | bar: No
1778 |  |  | United States | Lakeland Regional Medical Center | 320474 | 44566 |  | Q/O Customer Visit |  | 5 |  | Customer visit. No problems noted | [12 empty] | 119011183 | Closed | 1680115 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | None | [6 empty] | None | [5 empty] | TBD | bar: No
1779 |  |  | United States | Adventhealth Zephyrhills | 300106 | 44564 |  | Needs password reset - to change protocols |  | 5 E, Sravanthi | [14 empty] | 118996944 | Closed | 1671122 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Customer request for password reset | [5 empty] | Non-refresh | bar: No
1780 | Date0117U |  | Mexico | DE LAGE LANDEN SA DE CV | 10007 | 44575 |  | EMPEZO MARCAR LEYENDA: EL SISTEMA DE ADQUISICION DEL GANTRY NO PUEDE CUMPLIR CON LOS REQUISITOS. (Legend of start mark: gantry procurement system cannot meet the requirements) |  | Wait for case close, then close this one. 2 |  | TBD | [12 empty] | 119058058 | In Progress | 1711239 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | TBD | [6 empty] | TBD | [5 empty] | TBD | bar: No
1781 |  |  | Netherlands | Stichting Rijnstate Ziekenhuis | 52101 | 44575 |  | NZC109 Terugkomende ringartefacten met CT hersenen |  | 3 |  | Reboot gantry | [12 empty] | 119062380 | Closed | 1708651 |  |  | Opened | Opened | [6 empty] | Engineering | [6 empty] | On 12-1 and 14-1 there has been an error with the aircal | [5 empty] | IB Existing | bar: No
1782 |  |  | United States | Select Specialty Hospitals | 310051 | 44572 |  | Unable to burn a CD |  | 5 Reja, Afroj |  | None | [12 empty] | 119037460 | Closed | 1701199 |  |  | Evaluation | Evaluation | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Unable to burn CD | [9 empty] | System | TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo-TBD-NoEnoughInfo | bar: No
1783 | Date0124U |  | United States | Adventhealth Zephyrhills | 300099 | 44582 |  | Detector Noise |  | 3 |  | Replace the module | [12 empty] | 119102132 | Imcomplete | 1734864 |  |  | Opened | Opened | [6 empty] | Engineering | [6 empty] | Contrast inside, image artifact | [5 empty] | IB Existing | 459800594811 | TDMS Module Class A FRU | bar: No
1784 |  |  | United States | Christ Hospital | 300205 | 44582 |  | 65842604 IP System will not power on |  | 2 |  | Replace tape switch cable | [12 empty] | 119103253 | Closed | 1733242 |  |  | Opened | Opened | [6 empty] | Engineering | [6 empty] | E-stop cannot be closed | [5 empty] | IB Existing | 459800623252 | FRU, CBL, HANDLE TAPE SWITCH | bar: No
1785 |  |  | Netherlands | Stichting Rijnstate Ziekenhuis | 52101 | 44580 |  | Geen ECG signaal op host computer |  | 3 |  | Replace the ECG cable | [12 empty] | 119086967 | Closed | 1722972 |  |  | Opened | Opened | [6 empty] | Engineering | [6 empty] | ECG | [5 empty] | IB Existing | 455012301901 | ECG Monitor Trigger 7M Ca | bar: No
1786 |  |  | United States | Select Medical Corporation | 32261 | 44578 |  | error code and unit is not working properly |  | 2 |  | Ran air calibrations on system and found no problems with the system. | [12 empty] | 119074698 | Closed | 1718985 |  |  | Evaluation | Evaluation | [6 empty] | TBD | [6 empty] | TBD | [5 empty] | TBD | bar: No
1787 | Date0214 |  | China | 新疆伊犁州友谊医院 | 40048 | 44605 |  | 02N/Ingenuity CT 728326/add node | CT has no problem. Customer forget OLEVEL password, can not add node now. Customer feedback no need add it now. Inform FSE that set password when next time onsite service | 5 |  | Reset olevel password in next visit | [12 empty] | 119220768 | Closed | 1802238 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Customer | [6 empty] | Forget olevel password | [5 empty] | Refresh | [12 empty] | No: Yes
1788 |  |  | China | 四川省宜宾市第二中医院 | 345035 | 44605 |  | 05K For IB/Ingenuity Core128/Elite China/software crash | An error was been found when open the application.restore the Backup,the problem was same as before | 3 |  | Reload the SW | [12 empty] | 119220747 | Closed | 1802200 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Error when open the application | [5 empty] | Refresh |  |  |  | System | Console-SW | NoEnoughInfo | Console-SW-NoEnoughInfo | Console-SW-NoEnoughInfo | NoEnoughInfo | NoEnoughInfo | Console-SW | Console-SW-NoEnoughInfo | bar: No
1789 |  |  | China | Fujian Provincial Hospital | 345011 | 44603 |  | 05K for IB/Ingenuity Core128/Elite China/image artifact |  | 3 |  | Bad detector test, add the bad detector in the list | [12 empty] | 119212899 | Closed | 1799824 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Ring artifact | [5 empty] | IB Existing | [12 empty] | No: Yes
1790 |  |  | United States | Philips Medical Capital | 40055 | 44602 |  | 0699CDN84 can not hear auto voice commands at gantry since last service |  | 3 |  | Removed the gantry cover and reconnected the gantry speaker and microphone. Tested the auto voice and everything was working as it should. | [12 empty] | 119209686 | Closed | 1797588 |  |  | Investigation | Investigation | [6 empty] | Service | [6 empty] | Gantry rear intercom cable not connected. | [5 empty] | IB Existing | [12 empty] | No: Yes
1791 |  |  | China | Jiangxi Province Children’s Hospital | 40041 | 44602 |  | 02N/Ingenuity CT 728326/artifact |  | 3 |  | FSE onsite and check compenstor pass. After air calibration, no more artifact | [12 empty] | 119206501 | Closed | 1795803 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Artifact | [5 empty] | IB Existing | [12 empty] | No: Yes
1792 |  |  | Peru | BANCO BBVA PERU | 40034 | 44601 |  | CADA SEMANA PIDE CALIBRACION Y ENVIA MENSAJE DE ERROR DE TEMPERATURA | Temperature error when doing calibration | 3 |  | None | [12 empty] | 119205728 | Closed | 1795273 |  |  | Evaluation | Evaluation | [6 empty] | Customer | [6 empty] | Temperature error | [5 empty] | Refresh | [6 empty] | - | - | 0 | 0 | 0 | - | No: Yes
1793 |  |  | United States | Lackey Memorial Hospital | 40015 | 44601 |  | Calcium Scoring off/ not communicating - not tracking the heartrate |  | 3 ZHANG, Yue Xing |  | None | [12 empty] | 119205544 | Closed | 1795081 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | No track of heart rate | [5 empty] | IB Existing | [12 empty] | No: Yes
1794 |  |  | Germany | Gesundheit Nord gGmbH Klinikverbund | 40002 | 44601 |  | Artefakte bei sequentiellen Scans untersuchen |  | 3 |  | A-plane aligned and necessary calibrations carried out. Artifacts still present. Data compiled and forwarded to the factory. System operational. | [12 empty] | 119203651 | Closed | 1793755 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Artifact | [5 empty] | IB Existing | [12 empty] | No: Yes
1795 |  |  | China | 江西省儿童医院 | 40041 | 44600 |  | 02N/Ingenuity CT 728326/ Artifact |  | 3 ZHANG, Yue Xing |  | has ring artifact, customer did no do air calibration for a long time Repair Action: FSE feedback that no need dispatch, do air calibration and keep monitoring. Test/Inspection: Functional Check PASSE | [12 empty] | 119193681 | Closed | 1787314 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Artifact | [5 empty] | IB Existing | [12 empty] | No: Yes
1796 |  |  | United States | Philips Medical Capital | 40053 | 44599 |  | 0802CDN85 Recurring issue EKG monitor not connecting to CT |  | 3 Cohen, Shmuel |  | Per RSE checked 5V power supply to iBox. Voltage slightly under 5V, adjusted to between 5 and 5.1 as recommended | [12 empty] | 119191364 | Closed | 1785655 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | EKG Monitoring disconnect | [5 empty] | IB Existing | [12 empty] | No: Yes
1797 |  |  | Japan | 医療法人清翠会 牧病院 | 341017 | 44600 |  | 【CT:APP:RK】 |  | 5 Zhu, Chun |  | None | [12 empty] | 119195776 | Closed | 1784109 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | N/A | [5 empty] | Non-refresh | [12 empty] | No: Yes
1798 |  |  | Japan | Maki Hospital | 341017 | 44600 |  | If the Patient Orientation set when entering patient information is incorrectly set and scanned, it cannot be corrected. RFC |  | CFM Lu, Xueying |  | N/A | [12 empty] | CFM |  | 1783058 |  |  | Closed - Done | Closed - Done | [6 empty] | Customer | [6 empty] | Patient Orientation set incorrectly | [5 empty] | Non-refresh | [12 empty] | No: Yes
1799 |  |  | China | Laiyang City People_s Hospital | 345005 | 44591 |  | ADR: Error: Cantry acquisition system cannot comply NMPA, Sean to check with Chinal Local TSE |  | CFM Lu, Xueying |  | None | [12 empty] | NMPA |  | 1782945 |  |  | Closed - Done | Closed - Done | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Gantry acquisition system cannot comply | [5 empty] | Refresh |  |  |  | System | Gantry-SW | NoEnoughINfo | Gantry-SW-NoEnoughINfo | Gantry-SW-NoEnoughINfo | NoEnoughInfo | NoEnoughINfo | Gantry-SW | Gantry-SW-NoEnoughINfo | bar: No
1800 |  |  | United States | Philips Medical Capital | 40053 | 44599 |  | Scanner locks up during gated studies | EC has more than 15 results. also has Cal Score series with addition results. Did test scan and hour glass appeared with a wait until customer can continue planning. Did another test scan with reference EC, no hourglass or delay. | 4 Mohanta, Deepranjan |  | Advised customer to remove some functional results , leaving 0 and 40% or to build new EC from refence set. Also local disc at 25% remaining, deleted old exams from local. also advised customer to keep disk space at 50-60% available | [12 empty] | 119192014 | Closed | 1781846 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Customer | [6 empty] | Scanner lock up. | [5 empty] | Refresh | [6 empty] | - | - | 0 | 0 | 0 | - | No: Yes
1801 |  |  | Puerto Rico | METRO MAYAGUEZ INC | 40009 | 44599 |  | Throws message: gantry can not comply |  | 2 ZHANG, Yue Xing |  | The TX was replaced to resolve the issue | [12 empty] | 119190214 | Imcomplete | 1781674 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Gantry cannot comply | [5 empty] | IB Existing | 453567560151 | Kit, Moog Velocity Laser Power | [10 empty] | No: Yes
1802 |  |  | United States | Lackey Memorial Hospital | 40015 | 44599 |  | Getting clicking noise and error message during scan; Could not find any events caused problems. |  | 5 XU, Miley |  | Checked the CIRS Event Logs. Nothing showing as errors. Could not find any events caused problems | [12 empty] | 119191523 | Closed | 1781469 |  |  | Closed - Done | Closed - Done | Should be assembly issue. so change to Others | [5 empty] | Others | [6 empty] | Clicking noise and error message during scan | [5 empty] | Refresh | [6 empty] | - | - | Should be assembly issue. so change to Others | 0 | 0 | - | No: Yes
1803 |  |  | United States | Central Vermont Medical Center | 40024 | 44599 |  | 0699CFXXX IQ Check and Air Cal Failed |  | 4 |  | Reboot Console and gantry | [12 empty] | 119188616 | Closed | 1780192 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | NoEnoughINfo | [5 empty] | Engineering | [6 empty] | IQ check and Air cal fail | [5 empty] | Refresh |  |  |  | System | Console-SW | NoEnoughINfo | Console-SW-NoEnoughINfo | Console-SW-NoEnoughINfo | NoEnoughINfo | NoEnoughINfo | Console-SW | Console-SW-NoEnoughINfo | bar: No
1804 |  |  | China | Lindian County Hospital | 345038 | 44597 |  | Wechat:02N/Ingenuity Core128/Elite China/There is noise |  | 3 |  | Power on the UPS | [12 empty] | 119184777 | Closed | 1779223 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Customer | [6 empty] | Noise | [5 empty] | Non-refresh | [12 empty] | No: Yes
1805 |  |  | United States | Philips Medical Capital | 40055 | 44597 |  | 0304CDN84 Unable to run table in/out |  | 2 Wedemeier, Tracy |  | R2D or CGMP were replaced to resolve the issue | [12 empty] | 119184325 | Closed | 1777502 |  |  | Closed - Duplicate | Closed - Duplicate | [6 empty] | Engineering | [6 empty] | Unable to run table in/out | [5 empty] | IB Existing |  | R2D, CGMP | [10 empty] | No: Yes
1806 |  |  | Argentina | INSTITUTO NACIONAL DE | 40027 | 44596 |  | LA PRIMERA ES QUE EN EL MOVIMIENTO VERTICAL DE LA CAMILLA SE GENERA UN SONIDO FUERA DE LO HABITUAL | Lubrication of vertical mechanism | 3 |  | Vertical sliding lubrication mechanism | [12 empty] | 119182792 | Closed | 1776768 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Lubrication | [5 empty] | IB Existing | [12 empty] | No: Yes
1807 |  |  | Argentina | INSTITUTO NACIONAL DE | 40027 | 44596 |  | SE REALIZAN LOS ESTUDIOS CARDIACOS HAY QUE REINICIAR TODO EL EQUIPO PORQUE NO RECONOCE EL MONITOR MULTIPAREMETRICO | Cardiac examination must restart the whole device because it cannot recognize the multi area monitor | 3 |  | Replace IBOX-2 | [12 empty] | 119182814 | In Progress | 1776772 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Cardiac examination | [5 empty] | IB Existing | 459800166651 | I-BOX-2 BASE BOARD ASSY | [10 empty] | No: Yes
1808 |  |  | United States | Leesburg Regional Medical Center | 40022 | 44595 |  | Customer Wants To Update the Worklist and Store |  | 5 Mohanta, Deepranjan | [14 empty] | 119176029 | Closed | 1773219 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Customer want to update the worklist and store | [5 empty] | Non-refresh | [12 empty] | No: Yes
1809 |  |  | United States | Unity Health Systems | 40037 | 44595 |  | Couch is not moving horizontally Gantry Cant Init |  | 2 Temzin, Maria |  | None. System itself back to normal. | [12 empty] | 119175579 | Closed | 1772876 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Couch can move horizontal, suspect horizontal motor relay | [5 empty] | IB Existing | [12 empty] | No: Yes
1810 |  |  | China | Jinan City No. 4 People_s Hospital | 40038 | 44594 |  | 05K For IB/Ingenuity CT 728326/gantry power down | intermittent E-Stop open | 3 WANG, Peter |  | Replace R2D | [12 empty] | 119160348 | Closed | 1769792 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | E-stop open | [5 empty] | IB Existing | 459800467343 | R2D | [10 empty] | No: Yes
1811 |  |  | United States | Philips Medical Capital | 40055 | 44594 |  | Gantry acquisition system can not comply error due to R2D failure; FSE replaced the R2D |  | 4 Shai, Gal |  | Replace R2D and CGMP | [12 empty] | 119170019 | Closed | 1769485 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | Gantry acquisition system cannot comply | [5 empty] | IB Existing | 459800467343, 459800440404 | R2D, CGMP | [10 empty] | No: Yes
1812 |  |  | United States | Philips Medical Capital | 40053 | 44593 |  | Log collection for slow coronary scans | Obtained log information on two scans using different protocols for review and comparison | 3 |  | N/A | [12 empty] | 119159286 | Closed | 1769168 |  |  | Evaluation | Evaluation | [6 empty] | Service | [6 empty] | Collect Log | [5 empty] | Non-refresh | [12 empty] | No: Yes
1813 |  |  | Chile | Scotiabank Chile | 40008 | 44594 |  | TIENE FISURADA LA CORREA DEL GANTRY POLI 5 |  | 4 |  | Replace main belt | [12 empty] | 119166236 | Incomplete | 1768688 |  |  | Evaluation | Evaluation | [6 empty] | Engineering | [6 empty] | Main belt has broken | [5 empty] | IB Existing | 453566492401 | POLY-VEE BELT | [10 empty] | No: Yes
1814 |  |  | Chile | Scotiabank Chile | 40008 | 44592 |  | CL1675-PROBLEMA CON LA HORA |  | 4 |  | Se ingresa por via remota al reconstructor y se ajusta la hora para coincidir con la hora del computador principal. Se comprueba que las imágenes ahora aparecen con la hora correcta | [12 empty] | 119152878 | Incomplete | 1763938 |  |  | Investigation | Investigation | time not sync, belongs to service | [5 empty] | Service | [6 empty] | Image display the incorrect time | [5 empty] | Refresh | [6 empty] | - | - | time not sync, belongs to service | 0 | 0 | - | bar: No
1815 |  |  | United States | Philips Medical Capital | 40055 | 44592 |  | QO: material for install |  | 5 |  | None | [12 empty] | 119150998 | Closed | 1763841 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | N/A | [5 empty] | Non-refresh | bar: No
1816 |  |  | Chile | BANCO DE CHILE | 341005 | 44593 |  | CL1688-FALLA DE SOFTWARE, NO LEVANTA EL SOFTWARE |  | 2 |  | Reload host SW | [12 empty] | 119156334 | Imcomplete | 1763457 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632553 | [5 empty] | Engineering | [6 empty] | Login CT, cannot open application, blue screen | [5 empty] | Refresh |  |  |  | System | Console-SW | Blue-Screen | Console-SW-Blue-Screen | Console-SW-Blue-Screen | CTUDT00632553 | Blue-Screen | Console-SW | Console-SW-Blue-Screen | bar: No
1817 |  |  | Turkey | ADIYAMAN OZEL SAGINSTURTAS SAN VE TIC A.S | 341006 | 44592 |  | bazı sekans isimlerinde yazım yanlışlığı var. | There_s a mistake in writing in some random names. | 3 Dogan, Meltem |  | It was described to the user how to change the filming protocols. System is running | [12 empty] | 119151455 | Closed | 1761966 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Customer | [6 empty] | Write in some random names | [5 empty] | Non-refresh | bar: No
1818 |  |  | Italy | ASST SPEDALI CIVILI | 40026 | 44592 |  | ARTEFATTO SULLE IMMAGINI | Image artifact | 3 Lu, Xueying |  | This problem is already under management with case 0119150388 . This report is closed as double | [12 empty] | 119150706 | Closed | 1760129 |  |  | N/A, Duplicated | Investigation | [6 empty] | N/A, Duplicated | bar: No
1819 |  |  | Italy | ASST SPEDALI CIVILI | 40026 | 44592 |  | Artefatto concentrico | Concentric artefact | 3 Lu, Xueying | [14 empty] | 119150388 | Closed | 1760109 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Image Artifact | [5 empty] | IB Existing | 459800911271 | TDMS V2 M.Det Class A FRU | bar: No
1820 |  |  | United States | Lackey Memorial Hospital | 40015 | 44591 |  | 0308CDN84 3-System Restricted |  | 3 XU, Miley |  | Replace vertical transducer | [12 empty] | 119143848 | Closed | 1758669 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | E-stop open | [5 empty] | IB Existing | 459800623331 | Vertical Transducer | bar: No
1821 |  |  | Japan | Maki Hospital | 341017 | 44590 |  | 01CS:RS:朝のホスト再起動後画面がフリーズしている>>リモートから再起動し解消>>顕著なハードのエラーなし | The screen is frozen after the morning restart | 4 Zhu, Chun |  | Restart and reboot from remote. The event log shows no marked hardware error. Judging from malfunction on temporary applications. Follow-up | [12 empty] | 119142546 | Closed | 1758190 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Screen frozen after restart monitor | [5 empty] | Refresh |  |  |  | System | TBD | NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo | NoEnoughInfo | NoEnoughInfo | TBD | TBD-NoEnoughInfo | bar: No
1822 |  |  | Poland | Government Agency. | 341010 | 44589 |  | Unstable temperature in the detector, replaced heater modules, power supply cable and common DMS power |  | 3 Lu, Xueying |  | Replace the heater, 48V power supply, CDMP | [12 empty] | 119134833 | Closed | 1757509 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | DMS temperature is not stable | [5 empty] | IB Existing | 459800559353, 459800550141, 453567434781, | CDMP RoHS with pack; BR64T RoHS HEATER REPLACEMENT KIT; CBL, Heater 48VDC Power Supply | bar: No
1823 |  |  | Poland | Government Agency. | 341002 | 44589 |  | Pękaja zaczepy | Hooks are breaking | 4 |  | Replace the strap | [12 empty] | 119134588 | Fixed | 1757423 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Hooks are breaking | [5 empty] | IB Existing | 453566451921 | 14 INCH RESTRAINT STRAP | bar: No
1824 |  |  | Japan | Maki Hospital | 341017 | 44590 |  | 01DS:RS:キーボードが反応しない>>延長ケーブルを介さず、HOSTに直接接続し復帰 |  | 4 Zhu, Chun |  | The keyboard doesn_t react to me Confirmed reproducibility of the phenomenon I can not. From the situation that occurred, USB Extended cable and I decided to exchange extension cable. After work, confirm the operation of the device I confirmed that it works. | [12 empty] | 119141839 | Closed | 1756825 |  |  | Evaluation | Evaluation | [6 empty] | Engineering | [6 empty] | Keyboard not working | [5 empty] | IB Existing | 452213250062 | USB EXTENSION CABLE A-A 3 METERS | bar: No
1825 |  |  | China | Laiyang City People_s Hospital | 345005 | 44589 |  | 01/Ingenuity Core128/Elite China/scan interrupt |  | 3 WANG, Peter |  | None | [12 empty] | 119131954 | In Process | 1756475 |  |  | Investigation | Investigation | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | ACQ not OK | [5 empty] | Refresh |  |  |  | System | Gantry-SW | NoEnoughInfo | Gantry-SW-NoEnoughInfo | Gantry-SW-NoEnoughInfo | NoEnoughInfo | NoEnoughInfo | Gantry-SW | Gantry-SW-NoEnoughInfo | bar: No
1826 |  |  | China | Jiangxi Province Children’s Hospital | 40041 | 44589 |  | 02N/Ingenuity CT 728326/Login account locked |  | 3 Wang, Elaine |  | 客户多次输错密码后，主机登录账号被锁，告知账户被锁半个小时后可以解锁，等待半个小时后登录正常。 | [12 empty] | 119131739 | Fixed | 1756388 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Customer | [6 empty] | Account locked | [5 empty] | Refresh | [6 empty] | - | - | 0 | 0 | 0 | - | bar: No
1827 |  |  | Poland | Government Agency. | 341002 | 44588 |  | sprawdzenie paska rotora gantry | Loud gantry rotor belt operation | 4 |  | Replace the main belt | [12 empty] | 119127889 | Closed | 1753636 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Gantry rotate main belt loud | [5 empty] | IB Existing | 453566492401 | POLY-VEE BELT | bar: No
1828 |  |  | China | ZJ Yueqing No.1 people Hosp | 40054 | 44587 |  | Duplicate to PR1738861_Intermittent ring artifact in head images. |  | CFM XU, Miley |  | Swap module position, and re-calibrate | [12 empty] | CFM |  | 1752690 |  |  | Closed - Duplicate | Closed - Duplicate | [6 empty] | Engineering | [6 empty] | Intermittent ring artifact | [5 empty] | IB Existing | bar: No
1829 |  |  | United States | Philips Medical Capital | 40055 | 44588 |  | Ring on image due to faulty detector modules; |  | 4 XU, Miley |  | Replaced detector modules and calibrated DMS | [12 empty] | 119129904 | Incomplete | 1751910 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | Image Artifact | [5 empty] | IB Existing | 459800911271 | TDMS V2 M.Det Class A FRU | bar: No
1830 |  |  | Portugal | Centro Hospitalar de Vila Nova | 40030 | 44587 |  | Paragens sucessivas do software entre exames no tc do SU sendo necessário fazer shutdown do equipamento | Successive software stops between exams in the CT of the SU requiring shutdown of the equipment | 3 |  | The reported problem could not be reproduced and the analysis of the bug reports was inconclusive. System software update expected to be available soon | [12 empty] | 119123395 | Closed | 1750826 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | PendingBugRep | [5 empty] | Engineering | [6 empty] | Successive software stops between exams | [5 empty] | Refresh |  |  |  | System | TBD | NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo | PendingBugRep | NoEnoughInfo | TBD | TBD-NoEnoughInfo | bar: No
1831 |  |  | Puerto Rico | DORADO HEALTH INC DBA | 40007 | 44588 |  | No video on monitor, replaced host computer |  | 3 SHU, Tim |  | Replae Host | [12 empty] | 119128366 | Fixed | 1749805 |  |  | Closed - Done | Closed - Done | DEFOA | [5 empty] | Engineering | [6 empty] | No video on the monitor | [5 empty] | Refresh | 459801493151 | CT Gen5.3 Host, Z4 |  | System | Console-HW | Monitor | Console-HW-Monitor | Console-HW-Monitor | DEFOA | Monitor | Console-HW | Console-HW-Monitor | bar: No
1832 |  |  | China | Zhalantun City People_s Hospital | 40045 | 44587 |  | 02N/Ingenuity CT 728326/Transport node change |  | 3 Wang, Elaine | [14 empty] | 119120122 | Closed | 1748217 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | N/A | [5 empty] | Non refresh | bar: No
1833 |  |  | China | Zhalantun City People_s Hospital | 40045 | 44587 |  | 02N/Ingenuity CT 728326/Image artifacts |  | 2 WANG, Peter |  | Module M39 and M40 were replaced | [12 empty] | 119120114 | Closed | 1748210 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Image Artifact | [5 empty] | IB Existing | 459800911281 | TDMS Module V2 | bar: No
1834 |  |  | United States | Leesburg Regional Medical Center | 40022 | 44587 |  | Trying fill out the ACR paperwork to get accredited. Needs info on how to change scan field of view, number of... |  | 3 E, Sravanthi |  | N/A | [12 empty] | 119121843 | Closed | 1745330 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | N/A | [5 empty] | Non refresh | bar: No
1835 |  |  | Japan | 苫小牧市立病院 | 341014 | 44587 |  | 【CT:APP:NO】 | A tube hot can be seen by a protocol to shoot the chest after taking a neck (surview hand grater). Would you like to cut DRI? | 5 Zhu, Chun |  | I think it is one countermeasure to switch to fixed MAS. If the maximum MAS is input to the protocol, the generation is reduced. | [12 empty] | 119120771 | Fixed | 1744325 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | NoEnoughINfo | [5 empty] | Engineering | [6 empty] | Tube is too hot | [5 empty] | Refresh |  |  |  | System | Gantry-SW | TubTooHot | Gantry-SW-TubTooHot | Gantry-SW-TubTooHot | NoEnoughINfo | TubTooHot | Gantry-SW | Gantry-SW-TubTooHot | bar: No
1836 |  |  | China | 四川省宜宾市第二中医院 | 345035 | 44586 |  | application/Ingenuity Core128/Elite China/ |  | 3 Zhai, Lisa |  | N/A | [12 empty] | 119112938 | New | 1743398 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | N/A | [5 empty] | Non refresh | bar: No
1837 |  |  | China | Shaanxi Province Xianyang Central Hospital (Xianyang City No. 2 People_s Hospital) | 40050 | 44586 |  | wechat:02N/Ingenuity CT 728326/disk space release failed |  | 3 Wang, Elaine |  | Clean D disk | [12 empty] | 119112742 | In Process | 1743261 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632261 | [5 empty] | Engineering | [6 empty] | Disk full | [5 empty] | Refresh |  |  |  | System | Console-SW | DiskFull | Console-SW-DiskFull | Console-SW-DiskFull | CTUDT00632261 | DiskFull | Console-SW | Console-SW-DiskFull | bar: No
1838 |  |  | China | Le qing City People_s Hospital | 40054 | 44585 |  | Image artifact - Swapped module locations and calibrated system |  | 3 Mulhern, Melissa |  | Swap module position, and re-calibrate | [12 empty] | 119106843 | Closed | 1738861 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | Intermittent ring artifact | [5 empty] | IB Existing | bar: No
1839 |  |  | China | Inner Mongolia Hinggan League People_s Hospital | 345030 | 44585 |  | 05K for IB/Ingenuity Core128/Elite China/artifact |  | 3 WANG, Peter |  | Replace module | [12 empty] | 119106582 | Closed | 1738702 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Image Artifact | [5 empty] | IB Existing | 459800911281 | TDMS Module V2 | bar: No
1840 |  |  | China | Zhalantun City People_s Hospital | 40045 | 44585 |  | Wechat:02N/Ingenuity CT 728326/Images fault |  | 3 Zhai, Lisa |  | Replace module | [12 empty] | 119106242 | Closed | 1738506 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Image Artifact | [5 empty] | IB Existing |  | Module | bar: No
1841 |  |  | Turkey | Yıldırım Beyazıt Universitesi Yenimahalle Eğitim Ve Araştırma Hastanesi | 337000 |  |  | The device is too slow and fails quickly |  | CFM WANG, Peter |  | N/A | [12 empty] | CFM |  | 1736840 |  |  | Investigation | Investigation | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Device is too slow, and fails quickly | [5 empty] | Refresh |  |  |  | System | Console-SW | Performance | Console-SW-Performance | Console-SW-Performance | NoEnoughInfo | Performance | Console-SW | Console-SW-Performance | bar: No
1842 |  |  | China | Fujian Province Medical University Affiliated No. 1 Hospital | 345012 | 44584 |  | E-stops will not close, adjusted left and right tilt switch assembly |  | 3 Kanzeg, Mike |  | Re-adjust the tilt switch | [12 empty] | 119105682 | Closed | 1736474 |  |  | Closed - Done | Closed - Done | [6 empty] | 车载CT | [6 empty] | E-stop open | [5 empty] | IB Existing | bar: No
1843 |  |  | Portugal | Centro Hospitalar de Vila Nova | 40030 | 44578 |  | Iberia-PT-Centro Hospitalar de Vila Nova-Ingenuity CT 728326 |  | CFM ZHANG, Yue Xing |  | Modify scan parameter | [12 empty] | CFM |  | 1736433 |  |  | Investigation | Investigation | CTD00007125 | [5 empty] | Engineering | [6 empty] | Tube is too hot, increase KV or decrease mAs | [5 empty] | Refresh |  |  |  | System | Gantry-SW | TubTooHeat | Gantry-SW-TubTooHeat | Gantry-SW-TubTooHeat | CTD00007125 | TubTooHeat | Gantry-SW | Gantry-SW-TubTooHeat | bar: No
1844 | Date0221 |  | China | 济南市第四人民医院 | 40038 | 44610 |  | 05K For IB/Ingenuity CT/Can not scan |  | 2 |  | Replace 2pcs of 453566503142 tube axial fan. | [12 empty] | 119250866 | Closed | 1824106 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | CT Cannot scan | [5 empty] | IB Existing | 453566503142 | tube axial fan | [10 empty] | No: Yes
1845 |  |  | Japan | 苫小牧市立病院 | 341014 | 44610 |  | WY:モニタケーブル交換 | Monitor cable exchange | 5 |  | Replace the DVI-Mini DP adaptor | [12 empty] | 119249023 | In Progress | 1823424 |  |  | Evaluation | Evaluation | DEFOA | [5 empty] | Engineering | [6 empty] | Monitor cable exchange | [5 empty] | Refresh | 459801495951 | Mini-DisplayPort To DVI Adapter Cable |  | System | Console-HW | Monitor | Console-HW-Monitor | Console-HW-Monitor | DEFOA | Monitor | Console-HW | Console-HW-Monitor | bar: No
1846 |  |  | United States | Philips Medical Capital | 40028 | 44610 |  | 0705CDY70 Keyboard numbers not working |  | 3 |  | Replace the keyboard | [12 empty] | 119253236 | In Progress | 1822728 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Keyboard not working | [5 empty] | IB Existing | 459800918431 | Cherry USB keyboard English Silver | [10 empty] | No: Yes
1847 |  |  | Turkey | ADIYAMANOZELSAGINSTURTASSANVETICA.S | 341006 | 44609 |  | Gantry çekimi durdurdu | Recon performance is bad. Recon is slow | 3 |  | User is connected to the system by remote access, request for slow access to the recon process. Run disk testing done. The result was successfully completed. Raw data disc formatted. After the work, the problem was confirmed. The system is active | [12 empty] | 119245308 | Closed | 1820022 |  |  | Investigation | Investigation | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Recon is slow | [5 empty] | Refresh |  |  |  | System | TBD | NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo | NoEnoughInfo | NoEnoughInfo | TBD | TBD-NoEnoughInfo | bar: No
1848 |  |  | Turkey | ADIYAMANOZELSAGINSTURTASSANVETICA.S | 341006 | 44610 |  | görüntüde çizgi çıkıyor | Poor image quality | 3 |  | UHR filter cleanup done, air calibration bad detector and beam and dms check and pass. The device has been successfully ejected | [12 empty] | 119249529 | In Progress | 1819889 |  |  | Investigation | Investigation | E-125059 | [5 empty] | Engineering | [6 empty] | Poor IQ/ Flurian | [5 empty] | Refresh |  |  |  | System | Couch-HW | Cleaning | Couch-HW-Cleaning | Couch-HW-Cleaning | E-125059 | Cleaning | Couch-HW | Couch-HW-Cleaning | bar: No
1849 |  |  | China | 北京市中西医结合医院 | 40032 | 44609 |  | 05K for IB/Ingenuity CT 728326/add node |  | 5 |  | N/A | [12 empty] | 119241757 | Fixed | 1818896 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | N/A | [5 empty] | Non refresh | [12 empty] | No: Yes
1850 |  |  | Japan | 独立行政法人国立病院機構 | 341015 | 44610 |  | 画面のCTのインジケータが黄色で点滅を繰り返している | Screen CT indicator repeats yellow blinking | 4 Yang, Cancan |  | Delete the left of the quota on the screen | [12 empty] | 119247922 | Closed | 1818850 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632520 | [5 empty] | Engineering | [6 empty] | Yellow screen blinking | [5 empty] | Refresh |  |  |  | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow-Screen | Console-SW | Console-SW-Yellow-Screen | bar: No
1851 |  |  | United States | Leesburg Regional Medical Center | 40022 | 44609 |  | ACR question |  | 4 |  | N/A | [12 empty] | 119245183 | Closed | 1816755 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | N/A | [5 empty] | Non refresh | [12 empty] | No: Yes
1852 |  |  | Chile | Scotiabank Chile | 40008 | 44607 |  | CL1675-CUANDO SE TOMA UN ANGIOGRAMA, SE PIERDE EL ROI | When you do angiography, you lose ROI | 4 |  | He checked the incident record of submitting questions during the exam. It is verified that the pause key is pressed in the ROI positioning stage in the workflow. The detailed work done during the review will be reviewed in the application. Steps to enhance the customer_s angiographic procedures. After verification, the problem was not raised subsequently. | [12 empty] | 119234542 | In Progress | 1813872 |  |  | Opened | Opened | [6 empty] | Customer | [6 empty] | ROI loose during angiography | [5 empty] | Non refresh | [12 empty] | No: Yes
1853 |  |  | Poland | Government Agency. | 341009 | 44607 |  | Ustawienie automatycznego restaru |  | 4 |  | N/A | [12 empty] | 119232729 | Closed | 1812501 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | N/A | [5 empty] | Non refresh | [12 empty] | No: Yes
1854 |  |  | China | Jiangxi Province Children’s Hospital | 40041 | 44607 |  | 02N/Ingenuity CT 728326/artifact |  | 3 |  | Replace the module | [12 empty] | 119229022 | In Process | 1810251 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Ring Artifact | [5 empty] | IB Existing | [12 empty] | No: Yes
1855 |  |  | China | Inner Mongolia Hinggan League People_s Hospital | 345030 | 44606 |  | 05K for IB/Ingenuity Core128/Gantry belt | Gantry main belt worn out. | 3 Zhai, Lisa |  | Replace the main belt. | [12 empty] | 119222032 | Closed | 1805144 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Main belt worn out | [5 empty] | IB Existing | 453566492401 | POLY-VEE BELT | [10 empty] | No: Yes
1856 |  |  | United States | Philips Medical Capital | 40055 | 44606 |  | No ECG signal on Host; RSE instructed customer to restart the Host and Gantry |  | 3 Shai, Gal |  | Reboot Console and Gantry | [12 empty] | 119227201 | Closed | 1804528 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | No ECG Signal in Host | [5 empty] | IB Existing | [12 empty] | No: Yes
1857 |  |  | Chile | BANCO DE CHILE | 341005 | 44606 |  | FALLA DE SOFTWARE EN EL ANGIOTAC | After reinstalling the software, the device is not allowed to have a cardiac examination | 3 |  | Check the USB connection of the electrocardiograph (Ibox), detect that the host does not detect the card, restart the Ibox card and re-establish communication. Use ECG monitor for function test | [12 empty] | 119225449 | In Progress | 1804145 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | IBOX | [5 empty] | Engineering | [6 empty] | device is not allowed to have a cardiac examination | [5 empty] | Refresh |  |  |  | System | Gantry-HW | IBOX | Gantry-HW-IBOX | Gantry-HW-IBOX | IBOX | IBOX | Gantry-HW | Gantry-HW-IBOX | bar: No
1858 |  |  | France | CLINIQUE DE FLANDRE | 40019 | 44606 |  | Probleme de table |  | 3 |  | Replace the table switch cable | [12 empty] | 119225913 | Closed | 1803910 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Table problem | [5 empty] | IB Existing | 459800623252 | FRU, CBL, HANDLE TAPE SWITCH | [10 empty] | No: Yes
1859 | Date0228 |  | Chile | Scotiabank Chile | 40008 | 44618 |  | CORTA LOS EXAMENES |  | 4 |  | TBD | [12 empty] | 119297882 | New | 1849697 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | TBD | [6 empty] | TBD | [5 empty] | TBD | [12 empty] | No: Yes
1860 |  |  | China | 内蒙古自治区第四医院 | 40023 | 44618 |  | 05L/Ingenuity CT 728326/have noise |  | 3 |  | TBD | [12 empty] | 119297746 | New | 1849628 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | TBD | [6 empty] | TBD | [5 empty] | TBD | [12 empty] | No: Yes
1861 |  |  | China | Hebei Province Tangshan Workers Hospital | 40044 | 44617 |  | wechat:02n/Ingenuity CT 728326/system report error |  | 3 |  | TBD | [12 empty] | 119294102 | New | 1848864 |  |  | Opened | Opened | [6 empty] | TBD | [6 empty] | TBD | [5 empty] | TBD | [12 empty] | No: Yes
1862 |  |  | China | Laiyang City People_s Hospital | 345005 | 44617 |  | pay job：01/Ingenuity Core128/Elite China/SWO#54874484/Buy parts |  | 3 |  | Replace R2D | [12 empty] | 119291069 | New | 1847697 |  |  | Opened | Opened | [6 empty] | Engineering | [6 empty] | ACQ not OK | [5 empty] | IB Existing | 459800467343 | R2D | [10 empty] | No: Yes
1863 |  |  | Japan | National Hospital Organization | 341013 | 44617 |  | 03CS:KK:10時30分ごろスキャン停止。再撮影でスキャン可。>> S_GEN_FAULT_ANODE_TUBE_ARC |  | 4 WANG LIN, LV |  | None | [12 empty] | 119291623 | Closed | 1844720 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Tube ARC | [5 empty] | IB Existing | [12 empty] | No: Yes
1864 |  |  | China | 铜陵博爱医院有限公司 | 345034 | 44616 |  | 05K for IB/Ingenuity Core128/Elite China/couch horizontal movement error |  | 3 |  | Replace the encoder | [12 empty] | 119283825 | In Process | 1844335 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | Engineering | [6 empty] | couch horizontal movement error | [5 empty] | IB Existing |  | Couch horizontal encoders | [10 empty] | No: Yes
1865 |  |  | China | 福建省立医院 | 345011 | 44616 |  | 05K For IB/Core128/Need move CT |  | 5 |  | 车载CT需要到场配合调整输入电。到场完成配合工作。 | [12 empty] | 119283034 | In Process | 1843954 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | Customer | [6 empty] | 配合车载CT调整输入电 | [5 empty] | Non refresh | [12 empty] | No: Yes
1866 |  |  | China | 广西藤县人民医院 | 40036 | 44616 |  | 01/Ingenuity CT/Customer consultation |  | 3 |  | 客户咨询此报修的设备是否是能谱CT。告知客户Ingenuity CT 不属于能谱CT。 | [12 empty] | 119282074 | Fixed | 1843710 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | Customer | [6 empty] | 咨询ING是否是能谱CT | [5 empty] | Non refresh | [12 empty] | No: Yes
1867 |  |  | United States | Unity Health Systems | 40037 | 44615 |  | Monthly QA |  | 5 |  | N/A | [12 empty] | 119278798 | In Progress | 1842888 |  |  | Opened | Opened | [6 empty] | Customer | [6 empty] | Monthly QA | [5 empty] | Non refresh | [12 empty] | No: Yes
1868 |  |  | Poland | Rzadowa Agencja | 341002 | 44615 |  | Głośna praca w czasie podnoszenia stołu | Loud work while lifting the table | 4 |  | Lubrication of the running elements in the base of the table | [12 empty] | 119273732 | Closed | 1840540 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | Engineering | [6 empty] | Couch vertical movement noise | [5 empty] | IB Existing | [12 empty] | No: Yes
1869 |  |  | China | 陕西省咸阳中心医院 (咸阳市第二人民医院） | 40050 | 44615 |  | 02N/Ingenuity CT 728326/artifact | Image artifact | 3 |  | TBD | [12 empty] | 119273253 | In Process | 1840352 |  |  | Opened - Pending Translation | Opened - Pending Translation | [6 empty] | TBD | [6 empty] | TBD | [5 empty] | TBD | [12 empty] | No: Yes
1870 |  |  | United States | Odessa Regional Medical Center | 40029 | 44614 |  | CT couch is making a creaking noise. FSE required to examine. |  | 4 |  | tested the table just found some normal table noise | [12 empty] | 119266879 | Incomplete | 1835864 |  |  | Evaluation | Evaluation | [6 empty] | Engineering | [6 empty] | Couch movement noise | [5 empty] | IB Existing | [12 empty] | No: Yes
1871 |  |  | China | 陕西省咸阳中心医院 (咸阳市第二人民医院） | 40050 | 44614 |  | 02N/Ingenuity CT 72832/Image quality issue |  | 3 |  | 客户对图像质量不满意，建议联系应用培训专家跟进 | [12 empty] | 119265591 | In Process | 1834798 |  |  | Opened - Pending Translation | Opened - Pending Translation | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | IQ satisfaction | [5 empty] | Refresh |  |  |  | System | Console-SW | NoEnoughInfo | Console-SW-NoEnoughInfo | Console-SW-NoEnoughInfo | NoEnoughInfo | NoEnoughInfo | Console-SW | Console-SW-NoEnoughInfo | bar: No
1872 |  |  | China | 江苏省徐州医学院附属第三医院 | 40039 | 44614 |  | 02N/Ingenuity CT 728326/acq not ok |  | 2 Zhai, Lisa |  | Replace TX Box | [12 empty] | 119263416 | Closed | 1833485 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | ACQ not OK | [5 empty] | IB Existing |  | TX Box | [10 empty] | No: Yes
1873 |  |  | China | 江苏省扬州市第一人民医院西区医院 | 345022 | 44613 |  | 05K/Ingenuity Core128/Elite China/reset password |  | 3 Zhai, Lisa |  | N/A | [12 empty] | 119259328 | In Process | 1830248 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | N/A | [5 empty] | Non refresh | [12 empty] | No: Yes
1874 |  |  | China | Laiyang City People_s Hospital | 345005 | 44613 |  | 05L/Ingenuity Core128/Elite China/report error | 冠脉扫描失败 | 3 |  | Replace R2D | [12 empty] | 119258068 | Closed | 1829389 |  |  | Opened | Opened | [6 empty] | Engineering | [6 empty] | Scan error | [5 empty] | IB Existing | 459800467343 | R2D | [10 empty] | No: Yes
1875 |  |  | China | 宁安市第二人民医院 | 345039 | 44613 |  | 02N/Ingenuity Core128/Elite China/Yellow screen |  | 3 |  | 升级软件至4.1.10.25220 | [12 empty] | 119258026 | Closed | 1829353 |  |  | Opened - Pending Translation | Opened - Pending Translation | CTUDT00632520 | [5 empty] | Engineering | [6 empty] | Yellow screen | [5 empty] | Refresh |  |  |  | System | Console-SW | Yellow-Screen | Console-SW-Yellow-Screen | Console-SW-Yellow-Screen | CTUDT00632520 | Yellow-Screen | Console-SW | Console-SW-Yellow-Screen | bar: No
1876 |  |  | China | DaKuai Town Health, Fengquan District, Xinxiang City, Henan Province | 345009 | 44613 |  | Application/Ingenuity Core128/Elite China | 客户咨询应用问题 | 5 |  | 已联系Application工程师跟进 | [12 empty] | 119257970 | New | 1829296 |  |  | Opened | Opened | [6 empty] | Customer | [6 empty] | consolut for applicable usage | [5 empty] | Non refresh | [12 empty] | No: Yes
1877 | Date0214U |  | Mexico | DE LAGE LANDEN SA DE CV | 10007 | 44603 |  | HACE RUIDO AL GIRAR EL ROTOR | The equipment made noise on the gantry | 3 |  | Replace the main belt. | [12 empty] | 119218089 | In Progress | 1799517 |  |  | Evaluation | Evaluation | [6 empty] | Engineering | [6 empty] | Gantry rotation noise | [5 empty] | IB Existing | 453566492401 | POLY-VEE BELT | [10 empty] | No: Yes
1878 |  |  | United States | Wake Med Cary Hospital | 9591 | 44602 |  | NEMO |  | 3 |  | N/A | [12 empty] | 119210589 | Closed | 1797366 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | N/A | [5 empty] | Non refresh | [12 empty] | No: Yes
1879 |  |  | United States | Wake Med Cary Hospital | 9591 | 44601 |  | NEMO |  | 3 Mukherjee, Susmita |  | N/A | [12 empty] | 119205537 | Closed | 1793202 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | N/A | [5 empty] | Non refresh | [12 empty] | No: Yes
1880 |  |  | United States | Adventhealth Zephyrhills | 300106 | 44600 |  | Cirs errors; Customer performed a reboot | Reboot CIRS | 3 Squires, Ciera | [14 empty] | 119199380 | Fixed | 1791295 |  |  | Closed - Done | Closed - Done | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | CIRS Error | [5 empty] | Refresh |  |  |  | System | CIRS-SW | NoEnoughInfo | CIRS-SW-NoEnoughInfo | CIRS-SW-NoEnoughInfo | NoEnoughInfo | NoEnoughInfo | CIRS-SW | CIRS-SW-NoEnoughInfo | bar: No
1881 |  |  | United States | Wake Med Cary Hospital | 9591 | 44601 |  | DMC shot comm collapse errors - Reboot to resolve |  | 3 Mulhern, Melissa |  | Reboot the system | [12 empty] | 119205423 | Closed | 1790791 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | DMC shot comm collapse error | [5 empty] | IB Existing | [12 empty] | No: Yes
1882 |  |  | United States | Adventhealth Zephyrhills | 300106 | 44599 |  | 1203CDY70 x-ray system can not comply errors |  | 2 Wedemeier, Tracy |  | CPM replaced | [12 empty] | 119190362 | Closed | 1780811 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | X-ray system cannot comply | [5 empty] | IB Existing |  | CPM Generator | [10 empty] | No: Yes
1883 |  |  | United States | Adventhealth Zephyrhills | 300099 | 44598 |  | 0199CDN84 Vertical lines through images. |  | 3 LI, Xue |  | Clean the some contrast on the mylar ring | [12 empty] | 119186133 | Closed | 1780145 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | some contrast on the mylar ring | [5 empty] | IB Existing | [12 empty] | No: Yes
1884 |  |  | United States | Christ Hospital | 300205 | 44592 |  | LIFECYCLE |  | 2 E, Sravanthi |  | N/A | [12 empty] | 119152147 | Closed | 1763052 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | N/A | [5 empty] | Non refresh | bar: No
1885 |  |  | United States | Lakeland Regional Medical Center | 320474 | 44592 |  | Q/O |  | E, Sravanthi |  | N/A | [12 empty] | 119152117 | Closed | 1761010 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | N/A | [5 empty] | Non refresh | bar: No
1886 | Date0228U |  | United States | Select Medical Corporation | 32261 | 44616 |  | Linear artifacts showing up when doing CT_s of head |  | 3 |  | None | [12 empty] | 119287186 | Incomplete | 1846310 |  |  | Opened | Opened | [6 empty] | Engineering | [6 empty] | Line artifact | [5 empty] | IB Existing | [12 empty] | No: Yes
1887 |  |  | Australia | CITISCAN RADIOLOGY PTY LTD | 10491 | 44614 |  | Left Hand monitor power switch broken, monitor is currently on, system operational. |  | 3 |  | Replace the monitor and the Graphic card | [12 empty] | 119269085 | In Progress | 1833267 |  |  | Evaluation | Evaluation | [6 empty] | Engineering | [6 empty] | Monittor broken | [5 empty] | IB Existing | 459801140821, 459801536801 | 19" LCD display monitor EA193Mi, NVIDIA Quadro P400 2GB (3)mDP GFX | [10 empty] | No: Yes
1888 |  |  | United States | Adventhealth Zephyrhills | 300106 | 44614 |  | 4.1.10 laser QA failed 2mm, and no SMPTE pattern. |  | 5 Das, Rituparna |  | TBD | [12 empty] | 119268772 | In Progress | 1832961 |  |  | Evaluation | Evaluation | [6 empty] | TBD | [6 empty] | TBD | [5 empty] | TBD | [12 empty] | No: Yes
1889 |  |  | United States | Lakeland Regional Health Systems | 300143 | 44613 |  | Q/O: Right panel touch panel in switch works but is getting harder to press |  | 5 |  | Replace the control panel | [12 empty] | 119260732 | Closed | 1831229 |  |  | Evaluation | Evaluation | [6 empty] | Engineering | [6 empty] | Panel button hard to press | [5 empty] | IB Existing | 459800484511 | CONTROL PANEL ASSEMBLY-FRONT RH | [10 empty] | No: Yes
| Date0307 |  | China | Zhalut Banner Mein medical hospital | 345036 | 44624 |  | 05k for IB/ Ingenuity Core128/Elite China/Microphone interference |  | 3 |  | information customer to adjust the button of control voice, adjust the voice button | [12 empty] | 119332804 | Closed | 1871894 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Customer | [6 empty] | CT Box noise too loud | [5 empty] | IB Existing | [12 empty] | No: Yes
|  |  | Poland | Rare Agency. | 341010 | 44624 |  | Artefakty |  | 4 |  | M11, channel 161 low signal, do air calibration, systm test | [12 empty] | 119332181 | Closed | 1871648 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | Artifact | [5 empty] | IB Existing | [12 empty] | No: Yes
|  |  | United States | Philips Medical Capital | 40055 | 44625 |  | 1302CFXXX Images not going to Pacs |  | 3 |  | Customer states they can archive to other locations. Suggested they check with Pacs admin/IT department to determine if Pacs is having problems. Contacted customer and archiving is working after Pacs Admin helped. | [12 empty] | 119336448 | Closed | 1871610 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Customer | [6 empty] | PACS | [5 empty] | Non refresh | [12 empty] | No: Yes
|  |  | United States | Odessa Regional Medical Center | 40029 | 44624 |  | Voice Manager |  | 4 |  | The Voice Manager was selected under the Wrench - the English language was selected - Add Phrase was selected - Maya recorded a new breathing instruction and save the new instruction successfully. Maya had no further questions or need for clinical support | [12 empty] | 119333815 | Closed | 1869393 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Engineering | [6 empty] | Clinical support - Voice Manager | [5 empty] | Refresh | [12 empty] | No: Yes
|  |  | China | Zhalut Banner Mein medical hospital | 345036 | 44623 |  | 05K for IB/Ingenuity Core128/Elite China/tube condition failed |  | 3 |  | None | [12 empty] | 119324806 | Closed | 1867415 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Rotor seek reference time out | [5 empty] | IB Existing | [12 empty] | No: Yes
|  |  | China | Fujian Provincial Hospital | 345011 | 44623 |  | 05K for IB/Ingenuity Core128/Elite China/panel failed |  | 3 |  | Replace the control panel | [12 empty] | 119324547 | Closed | 1867232 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Control panel malfunction | [5 empty] | IB Existing | 459800484501 | CONTROL PANEL ASSEMBLY-FRONT LH | [10 empty] | No: Yes
|  |  | China | Laiyang City People_s Hospital | 345005 | 44621 |  | Scan abort due to Gantry acquisition system cannot comply error reported |  | WANG LIN, LV |  | None. No bug report for this case, FSE was not required to visit this site. | [14 empty] | 1866490 |  |  | Closed - Done | Closed - Done | CTUDT00630408 | [5 empty] | Engineering | [6 empty] | ACQ not OK | [5 empty] | Refresh |  |  |  | System | CIRS-SW | ACQ_NOT_OK | CIRS-SW-ACQ_NOT_OK | CIRS-SW-ACQ_NOT_OK | CTUDT00630408 | ACQ_NOT_OK | CIRS-SW | CIRS-SW-ACQ_NOT_OK | No: Yes
|  |  | China | Lindian County Hospital | 345038 | 44622 |  | 02N/Core128/The belt fault |  | 3 |  | FSE cannot on site for inspect, since the covid 19. Suspect belt or tension Assy | [12 empty] | 119322503 | In Process | 1866164 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Gantry rotation noise | [5 empty] | IB Existing | [12 empty] | No: Yes
|  |  | United States | Philips Medical Capital | 40005 | 44622 |  | LS Quote |  | 5 |  | New case, the customer ask for quote | [12 empty] | 119321776 | New | 1865692 |  |  | Evaluation | Evaluation | [6 empty] | Customer | [6 empty] | Quote | [5 empty] | Non refresh | [12 empty] | No: Yes
|  |  | United States | Lackey Memorial Hospital | 40015 | 44622 |  | Q/O unit is making gantry noise |  | 5 |  | FSE onsite check, no more noise of the gantry. | [12 empty] | 119321317 | Closed | 1864385 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Gantry noise | [5 empty] | IB Existing | [12 empty] | No: Yes
|  |  | China | Inner Mongolia Autonomous Region No. 4 Hospital | 40023 | 44622 |  | Appliaction/Ingenuity CT 728326 |  | 5 |  | Clinical training for enhanced scan | [12 empty] | 119319786 | New | 1864097 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Ehanced scan | [5 empty] | Non refresh | [12 empty] | No: Yes
|  |  | China | Inner Mongolia Autonomous Region No. 4 Hospital | 40023 | 44622 |  | pay job：01/Ingenuity CT 728326/swo#54899233 buy parts |  | 3 |  | Replace main belt | [12 empty] | 119316866 | Closed | 1862246 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Engineering | [6 empty] | Gantry rotation noise | [5 empty] | IB Existing | 453566492401 | POLY-VEE BELT | [10 empty] | No: Yes
|  |  | China | Beijing City Hospital of Integrated Chinese and Western Medicine | 40032 | 44622 |  | 05K for IB/Ingenuity CT 728326/ECG issue |  | 3 |  | Replace IBOX and IBOX PS | [12 empty] | 119316239 | Closed | 1861940 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | ECG issue | [5 empty] | IB Existing | 459800166651 | I-BOX-2 BASE BOARD ASSY | [10 empty] | No: Yes
|  |  | China | Sxi feng County T.C.M. Hospital | 345014 | 44621 |  | 01/Ingenuity Core128/Elite China/add node |  | 3 Zhai, Lisa |  | Provide account and password, instruct customer to add node | [12 empty] | 119311089 | Closed | 1858161 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Add node | [5 empty] | Non refresh | [12 empty] | No: Yes
|  |  | China | Sxi feng County T.C.M. Hospital | 345014 | 44621 |  | 01/Ingenuity Core128/Elite China/ADD dicom node |  | 3 Cai, Cathy | [14 empty] | 119310704 | Closed | 1857875 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Add Dicom node | [5 empty] | Non refresh | [12 empty] | No: Yes
|  |  | China | Shandong Province Chest Hospital | 40025 | 44621 |  | 05L/Ingenuity CT 728326/add the printer |  | 5 Zhai, Lisa | [14 empty] | 119310649 | Fixed | 1857812 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Add printer | [5 empty] | Non refresh | [12 empty] | No: Yes
|  |  | United States | Philips Medical Capital | 40055 | 44621 |  | QO: Add new ct to scanner |  | 5 |  | Added new ct as network node verified reachable | [12 empty] | 119315231 | Closed | 1857317 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Engineering | [6 empty] | Add new CT as network node | [5 empty] | Non refresh | [12 empty] | No: Yes
|  |  | China | Xinjiang YiliuZhou Friendship Hospital | 40048 | 44621 |  | wechat：02N/Ingenuity CT 728326/The administrator account and password cannot be logged in |  | 3 Zhai, Lisa | [14 empty] | 119309814 | Closed | 1857286 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Engineering | [6 empty] | Administrator account cannot login | [5 empty] | Refresh | [12 empty] | No: Yes
|  |  | China | Inner Mongolia Alasan Right Banner People_s Hospital | 345031 | 44620 |  | 01/Ingenuity Core128/Elite China/consult for password |  | 3 Zhai, Lisa | [14 empty] | 119304377 | Fixed | 1854070 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Consult for password | [5 empty] | Non refresh | [12 empty] | No: Yes
|  |  | United States | Ashland Hospital Corporation | 40011 | 44620 |  | PM WAS DONE ON THURSDAY 2.24.22 AND NOW ITS RUNNING SLOW |  | 4 Dantha, Rajashekar |  | Recycled the power to the HOST and CIRS computers. Performed Visual Diagnostics on the CIRS and Database Cleanup. | [12 empty] | 119308478 | Closed | 1853132 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | Performance Slow | [5 empty] | Service | [6 empty] | System running slow | [5 empty] | Refresh |  |  |  | System |  |  | - | - | Performance Slow |  |  |  | No: Yes
|  |  | China | Jinan City No. 4 People_s Hospital | 40038 | 44620 |  | Application/05K for IB/Ingenuity CT 728326 |  | 5 |  | None | [12 empty] | 119301166 | Closed | 1852848 |  |  | Investigation | Investigation | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Vessel enhance scan protocol issue | [5 empty] | Refresh |  |  |  | System | TBD | NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo | bar: No
|  |  | China | Guangdong Province Guangzhou City No.8 People_s Hospital (Jiahe Hospital District) | 345032 | 44620 |  | 05K for IB/Ingenuity Core128/Elite China/Scan error |  | 3 |  | Replace the TX box | [12 empty] | 119300552 | Closed | 1852646 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | Scan abort | [5 empty] | IB Existing |  | TX Box | [10 empty] | No: Yes
|  |  | United States | Philips Medical Capital | 40053 | 44620 |  | 0699CFXXX Received a temperature control malfunction error when running QC |  | 4 Mohanta, Deepranjan |  | Power cycle the system | [12 empty] | 119306348 | Closed | 1851328 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | DMStemperature error | [5 empty] | IB Existing | [12 empty] | No: Yes
| Date0307U |  | United States | Upson Regional Medical Center | 310115 | 44622 |  | 84628582 TABLE WILL NOT COME OUT WITH FOOT SWITCH OR GANTRY SWITCH - HAS TO BE DONE MANUALLY |  | 5 |  | performed horizontal table cals | [12 empty] | 119320599 | Closed | 1864840 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | Couch can only move manually | [5 empty] | IB Existing | [12 empty] | No: Yes
|  |  | Mexico | DE LAGE LANDEN SA DE CV | 10007 | 44621 |  | EL EQUIPO MUESTRA ERROR NO SE PUEDE INICIAR LA RECONTRUCCION |  | 3 |  | Replace the FRC | [12 empty] | 119314864 | Closed | 1856898 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | No power to TX | [5 empty] | IB Existing | 453567057167 | FRC ASSEMBLY | [10 empty] | No: Yes
|  |  | Netherlands | Stichting Rijnstate Ziekenhuis | 52101 | 44621 |  | Bolus tracking scan start niet vanwege trage recon delay. |  | 3 |  | Reload server 1 software | [12 empty] | 119313407 | Closed | 1855542 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Slow recon delay | [5 empty] | Refresh |  |  |  | System | Console-SW | BolusTracking | Console-SW-BolusTracking | Console-SW-BolusTracking | 0 |  |  |  | No: Yes
|  |  | United States | Wake Med Cary Hospital | 9591 | 44621 |  | 0399CFXXX Need part number for Mattress Pad |  | 5 Mohanta, Deepranjan | [14 empty] | 119313234 | Closed | 1855472 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Customer consult for the pad PN | [5 empty] | Non refresh | [12 empty] | No: Yes
| Date0314U |  | United States | BIBB MEDICAL CENTER | 310175 | 44630 |  | QO: cracked lexan ring |  | 5 |  | Replace the Lexan ring | [12 empty] | 119365623 | Closed | 1889691 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | Lexan ring crack | [5 empty] | IB Existing | 459800048742 | RING-LEXAN GN5 | [10 empty] | No: Yes
|  |  | United States | Adventhealth Zephyrhills | 300099 | 44628 |  | Q/O |  | 3 |  | None | [12 empty] | 119351998 | Closed | 1880130 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | TBD | [6 empty] | TBD | [5 empty] | TBD | [12 empty] | No: Yes
| Date0314 |  | China | Jiangxi Province Yugan Chudong Hospital (New Branch) | 337500 | 44633 |  | wechat:05K/Ingenuity Core128/Elite China/Button failure |  | 3 |  | Replace CT Box | [12 empty] | 119374121 | Closed | 1897118 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632185 | [5 empty] | Engineering | [6 empty] | CT Box button malfunction | [5 empty] | TBD |  | CT Box |  | System | Gantry-HW | CT Box | Gantry-HW-CT Box | Gantry-HW-CT Box | CTUDT00632185 | CT Box | Gantry-HW | Gantry-HW-CT Box | No: Yes
|  |  | United States | Leesburg Regional Medical Center | 40022 | 44632 |  | 85536292 the A-STOP WILL NOT BOOT |  | 2 |  | Replace the left control panel | [12 empty] | 119373696 | Closed | 1896221 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | E-stop cannot close | [5 empty] | IB Existing | 459800484501 | CONTROL PANEL ASSEMBLY-FRONT LH | [10 empty] | No: Yes
|  |  | China | Lindian County Hospital | 345038 | 44631 |  | 02N/Ingenuity Core128/Elite China/ECG signal lost |  | 3 |  | None, access constraint due to Covid19 | [12 empty] | 119366570 | In Process | 1894532 |  |  | Investigation | Investigation | [6 empty] | TBD | [6 empty] | TBD | [5 empty] | TBD | [12 empty] | No: Yes
|  |  | China | Lindian County Hospital | 345038 | 44631 |  | 02N/Ingenuity Core128/Elite China/ECG fault |  | 3 |  | Replace the usb cable between IBOX and Host | [12 empty] | 119367244 | Closed | 1894802 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Engineering | [6 empty] | ECG issue | [5 empty] | IB Existing | 459801731761 | USB EXTENDER CABLE FROM IBOX TO CONSOLE | [10 empty] | No: Yes
|  |  | China | Inner Mongolia Xing_an People_s Hospital | 345030 | 44631 |  | 05k for IB/ Ingenuity Core128/Elite China/Artifact |  | 3 |  | Air calibration | [12 empty] | 119367368 | Closed | 1894884 |  |  | Investigation | Investigation | air calibration | [5 empty] | Customer | [6 empty] | Artifact | [5 empty] | Refresh | [12 empty] | No: Yes
|  |  | China | Inner Mongolia Xing_an People_s Hospital | 345030 | 44631 |  | 05k for IB/ Ingenuity Core128/couch cushion |  | 5 |  | Provide couch curshion | [12 empty] | 119368315 | Closed | 1895326 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Customer need couch curshion | [5 empty] | Non refresh | [12 empty] | No: Yes
|  |  | Portugal | Centro Hospitalar de Vila Nova | 40030 | 44631 |  | Substituição de colchão da TC do SU que se encontra danificado e embebido em sangue |  | 3 |  | TBD | [12 empty] | 119370548 | New | 1896165 |  |  | Investigation | Investigation | [6 empty] | TBD | [6 empty] | TBD | [5 empty] | TBD | [12 empty] | No: Yes
|  |  | Chile | BANCO DE CHILE | 341005 | 44630 |  | SE CORTO DISPARO EN EXAMEN CAT |  | 3 |  | None | [12 empty] | 119365128 | Closed | 1889756 |  |  | Investigation | Investigation | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Short shot in CAT | [5 empty] | Refresh |  |  |  | System | TBD | NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo |  |  |  |  | No: Yes
|  |  | Japan | Kinan Hospital | 340000 | 44630 |  | CX:ガントリベアリンググリスアップ |  | 5 Yang, Cancan |  | Apply grease | [12 empty] | 119359819 | Closed | 1890185 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Grease grill up | [5 empty] | IB Existing | [12 empty] | No: Yes
|  |  | China | Suzhou Guang ci Tumor Hospital Co., Ltd | 345019 | 44630 |  | 05K for IB/Ingenuity Core128/Elite China/gantry noise |  | 3 |  | Replace the tensioner | [12 empty] | 119361326 | Closed | 1891110 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Gantry noise | [5 empty] | IB Existing | 453567003451 | TENSIONER ASSY | [10 empty] | No: Yes
|  |  | China | Shaanxi Ankang T.C.M | 345041 | 44629 |  | WeChat：02N/Ingenuity Core128/Recon cann_t initialize |  | 3 Gao, Mia |  | Restart the system | [12 empty] | 119353376 | Closed | 1885516 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | ReconNotInitialize | [5 empty] | Engineering | [6 empty] | Recon cannot initialize | [5 empty] | Refresh |  |  |  | System | CIRS-SW | ReconNotInitialize | CIRS-SW-ReconNotInitialize | CIRS-SW-ReconNotInitialize |  |  |  |  | No: Yes
|  |  | China | Shandong Anchu City T.C.M | 345017 | 44629 |  | 02N/Ingenuity Core128/Elite China/keyboard failure |  | 3 Zhai, Lisa |  | Replace the keyboard | [12 empty] | 119354837 | Closed | 1886511 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Keyboard not working | [5 empty] | IB Existing | 459800918431 | Cherry USB keyboard English Silver | [10 empty] | No: Yes
|  |  | United States | Unity Health Systems | 40037 | 44629 |  | Monthly QA |  | 5 |  | N/A | [12 empty] | 119357528 | Fixed | 1888805 |  |  | Evaluation | Evaluation | [6 empty] | Customer | [6 empty] | Monthly QA | [5 empty] | Non refresh | [12 empty] | No: Yes
|  |  | United States | Philips Medical Capital | 40028 | 44628 |  | 1202CDN84 Maximum number of hard arcs |  | 2 |  | Ran a Long Tube Conditioning and Filament calibrations before and after to eliminate arcing in the system | [12 empty] | 119348890 | Closed | 1877958 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | Tube Arc | [5 empty] | IB Existing | [12 empty] | No: Yes
|  |  | United States | Philips Medical Capital | 40055 | 44628 |  | 0704CFXXX Local full |  | 3 Boddolla, Ramu |  | Deleted orphaned files. Customer checked local space and It is back to 70% free | [12 empty] | 119351031 | Closed | 1879783 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632261 | [5 empty] | Engineering | [6 empty] | Local disk is full | [5 empty] | Refresh |  |  |  | System | Console-SW | DiskFull | Console-SW-DiskFull | Console-SW-DiskFull | CTUDT00632261 | DiskFull | Console-SW | Console-SW-DiskFull | No: Yes
|  |  | China | Zhalantun City People_s Hospital | 40045 | 44628 |  | 02N/Ingenuity CT 728326/IST authorize |  | 3 Gao, Mia |  | N/A | [12 empty] | 119348087 | Closed | 1881830 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | Non refresh | [12 empty] | No: Yes
|  |  | United States | Philips Medical Capital | 40053 | 44628 |  | QO: Blood pressure cuff order |  | 5 |  | Unable to order blood pressure cuff directly, provided customer with contact to order | [12 empty] | 119351842 | Closed | 1884627 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Order | [5 empty] | Non refresh | [12 empty] | No: Yes
|  |  | Japan | National Hospital Organization | 341015 | 44627 |  | SF:ディテクターモジュール交換 |  | 5 Xu, Xiaoting |  | Replace the module | [12 empty] | 119339688 | Closed | 1876826 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Artifact | [5 empty] | IB Existing | 459800911271 | TDMS V2 M.Det Class A FRU | [10 empty] | No: Yes
|  |  | China | Suzhou Guang ci Tumor Hospital Co., Ltd | 345019 | 44627 |  | 05K for IB/Ingenuity Core128/Elite China/Control panel key failure |  | 3 |  | Replace the control panel | [12 empty] | 119340409 | Closed | 1877313 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Control panel malfunction | [5 empty] | IB Existing | 459801484981 | CONTROL PANEL ASSEMBLY-FRONT RIGHT | [10 empty] | No: Yes
| Date0321U |  | Spain | POLICLINICA NTRA SRA DEL ROSARIO SL | 336190 | 44636 |  | No han podido arrancarlo, se queda bloqueado al intentar encenderlo The password expire date will be in 90 days. |  | 2 |  | By entering with the administrator password, finally we were able to check and update the usual user passwords, so that the client enters the application of the host computer. | [12 empty] | 119396901 | Closed | 1910144 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632236 | [5 empty] | Engineering | [6 empty] | Password | [5 empty] | Refresh |  |  |  | System | Console-SW | User Account | Console-SW-User Account | Console-SW-User Account | CTUDT00632236 | User Account | Console-SW | Console-SW-User Account | No: Yes
| Date0321 |  | China | 河南许昌市中心医院(传染病) | 40047 | 44640 |  | 05K for IB/Ingenuity CT 728326/ECG problem |  | 3 |  | Reboot | [12 empty] | 119417707 | Fixed | 1919941 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | ECG issue | [5 empty] | Refresh |  |  |  | System | TBD | NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo | NoEnoughInfo |  |  |  | No: Yes
|  |  | China | Henan Xuchang City Central Hospital (Infectious Diseases) | 40047 | 44638 |  | 05k for IB/ Ingenuity CT 728326/Head rest failure |  | 3 |  | None | [12 empty] | 119409162 | Closed | 1917647 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | HeadRest | [5 empty] | Engineering | [6 empty] | Head rest failure | [5 empty] | Refresh |  |  |  | System | Couch-HW | HeadRest | Couch-HW-HeadRest | Couch-HW-HeadRest | HeadRest |  |  |  | No: Yes
|  |  | China | Suzhou Guang ci Tumor Hospital Co., Ltd | 345019 | 44638 |  | 01/Ingenuity Core128/Elite China/High pressure syringe product fault |  | 3 |  | TBD | [12 empty] | 119409498 | In Process | 1917737 |  |  | Evaluation | Evaluation | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | High pressure syring fail | [5 empty] | Refresh |  |  |  | System | TBD | NotEngineering | TBD-NotEngineering | TBD-NotEngineering |  |  |  |  | No: Yes
|  |  | Spain | S.C.S. HOSP. GRAL. DE FUERTEVENTURA | 40020 | 44638 |  | Defectos estéticos encontrados durante el mantenimiento preventivo |  | 5 |  | It is required to use consumable materials during the next preventive maintenance. The customer was informed that although the patient pad is considered to be an accessory not covered by the maintenance contract, in this case, it is ordered and provided free of charge | [12 empty] | 119412660 | Closed | 1918478 |  |  | Evaluation | Evaluation | [6 empty] | Engineering | [6 empty] | Crushion | [5 empty] | Non refresh | 459801350221 | STANDARD CUSHION-47.24 INCH MATTRESS PAD | [10 empty] | No: Yes
|  |  | Poland | Rzadowa Agencja | 341009 | 44638 |  | W trakcie skanowania pojawiają się pierścienie pochodzące najprawdopodobniej z detektora. |  | 3 |  | Replace the module | [12 empty] | 119413147 | Closed | 1918663 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Ring Artifact | [5 empty] | IB Existing | 459800911271 | TDMS V2 M.Det Class A FRU | [10 empty] | No: Yes
|  |  | United States | Odessa Regional Medical Center | 40029 | 44638 |  | recon images |  | 4 |  | N/A | [12 empty] | 119414279 | Canceled | 1919097 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A | [12 empty] | No: Yes
|  |  | Chile | BANCO DE CHILE | 341005 | 44638 |  | NO HAY SEÑAL ECG EN ANGIOTAC |  | 3 |  | TBD | [12 empty] | 119415052 | In Progress | 1919280 |  |  | Evaluation | Evaluation | [6 empty] | Engineering | [6 empty] | No ECG Signal | [5 empty] | IB Existing | [12 empty] | No: Yes
|  |  | China | Henan Xuchang City Central Hospital (Infectious Diseases) | 40047 | 44637 |  | 05K/Ingenuity CT 728326/Check warranty |  | 3 |  | N/A | [12 empty] | 119405527 | Closed | 1915270 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A | [12 empty] | No: Yes
|  |  | United States | Unity Health Systems | 40037 | 44637 |  | Monthly QA |  | 5 |  | N/A | [12 empty] | 119407021 | Closed | 1916477 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A | [12 empty] | No: Yes
|  |  | China | Zhalantun City People_s Hospital | 40045 | 44636 |  | 02N/Ingenuity CT/CT box&noise |  | 3 |  | Replace the main belt | [12 empty] | 119395480 | Closed | 1909316 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Gantry rotation noise | [5 empty] | IB Existing | 453566492401 | BELT, 5-RIB POLY-VEE, 198" | [10 empty] | No: Yes
|  |  | China | Inner Mongolia Xing_an People_s Hospital | 345030 | 44636 |  | 05K for IB/Core128/couch cushion |  | 3 Zhai, Lisa |  | None | [12 empty] | 119395938 | Closed | 1909506 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Engineering | [6 empty] | Crushion | [5 empty] | Non refresh | [12 empty] | No: Yes
|  |  | China | Zhalantun City People_s Hospital | 40045 | 44636 |  | 02N/Ingenuity CT/CT box |  | 3 |  | Replace CT Box | [12 empty] | 119395988 | Closed | 1909533 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | CT Box button malfunction | [5 empty] | IB Existing | 453567027195 | CTBOX ASSEMBLY RAFI/TILTED (RoHS) | [10 empty] | No: Yes
|  |  | China | Le qing City People_s Hospital | 40054 | 44636 |  | 02N/Ingenuity CT 728326/scan complete and image not recon |  | 3 |  | Replace GPU P2000 | [12 empty] | 119396185 | Closed | 1909676 |  |  | Investigation | Investigation | GPU | [5 empty] | Engineering | [6 empty] | Not recon | [5 empty] | Refresh | 459801622381 | NVidia P2000 GPU Card |  | System | Console-HW | GPU | Console-HW-GPU | Console-HW-GPU |  |  |  |  | No: Yes
|  |  | China | Shandong Anchu City T.C.M | 345017 | 44636 |  | 02N/Ingenuity Core128/Password |  | 5 Zhai, Lisa |  | N/A | [12 empty] | 119396596 | Closed | 1909951 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A | [12 empty] | No: Yes
|  |  | China | Jiangsu Province No. 2 T.M. Hospital | 345027 | 44636 |  | 05K/Ingenuity Core128/The time was not correct |  | 5 |  | Remote guide customer correct the time | [12 empty] | 119397038 | Closed | 1910246 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | ShouldbeService | [5 empty] | Service | [6 empty] | Host time incorrect | [5 empty] | Refresh | [6 empty] | - | - | ShouldbeService |  |  |  | No: Yes
|  |  | China | Henan Xuchang City Central Hospital (Infectious Diseases) | 40047 | 44636 |  | 02N/Ingenuity CT 728326/Hung up |  | 3 |  | N/A | [12 empty] | 119397426 | Closed | 1910487 |  |  | Closed - Done | Closed - Done | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Software hung up | [5 empty] | Refresh |  |  |  | System | TBD | NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo | NoEnoughInfo |  |  |  | No: Yes
|  |  | China | Fujian Provincial Hospital | 345011 | 44635 |  | 05K for IB/Ingenuity Core128/Elite China/ACCU ERROR |  | 3 |  | Replace ACCU | [12 empty] | 119385119 | Closed | 1905156 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | F301 burn | [5 empty] | IB Existing | 459800298762 | AC CONTROL UNIT PCB | [10 empty] | No: Yes
|  |  | Japan | Kinan Hospital | 340000 | 44635 |  | MN:寝台部グリスアップ |  | 5 Xu, Xiaoting |  | Bed section grease up was implemented. There is no problem with the operation of the device. | [12 empty] | 119385423 | Fixed | 1905400 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Other | [6 empty] | Do lubrication | [5 empty] | Non refresh | [12 empty] | No: Yes
|  |  | China | Zhalut Banner Mein medical hospital | 345036 | 44635 |  | 05k for ib/Ingenuity Core128/Elite China/add dicom node |  | 3 |  | N/A | [12 empty] | 119385732 | Closed | 1905560 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A | [12 empty] | No: Yes
|  |  | China | Yang gao County Medical Group (Yang gao County People_s Hospital). | 345025 | 44635 |  | Application/Ingenuity Core128/Elite China |  | 5 |  | N/A | [12 empty] | 119386003 | New | 1905747 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A | [12 empty] | No: Yes
|  |  | China | Jiangxi Province Children_s Hospital | 40041 | 44635 |  | 02N/Ingenuity CT 728326/gantry noise |  | 3 |  | Adjust the POLY-V SHEAVE position | [12 empty] | 119386538 | Closed | 1906082 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Gangry rotation noise | [5 empty] | IB Existing | [12 empty] | No: Yes
|  |  | Turkey | ADIYAMANOZELSAGINSTURTASSANVETICA.S | 341006 | 44634 |  | Görüntülerde artefak var |  | 3 WANG LIN, LV |  | the App Support Team investigated the imagination of artefact and reported that the images were normal with the exact residue of the shooting. Specifies whether the lines in the word view will not appear when maximising the remainder. The hospital will be informed about this. No field visits will be performed. The system is active. | [12 empty] | 119380685 | Closed | 1898350 |  |  | Investigation | Investigation | [6 empty] | Customer | [6 empty] | Artifact with residue shooting | [5 empty] | Non refresh | [12 empty] | No: Yes
|  |  | United States | Philips Medical Capital | 40028 | 44634 |  | 0902CDY73 x6558 QA Low contrast failed |  | 3 Mukherjee, Susmita |  | Customer resolved the issue themselves, cancel the call | [12 empty] | 119381869 | Closed | 1898820 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Customer | [6 empty] | Low contrast failed | [5 empty] | Non refresh | [12 empty] | No: Yes
|  |  | Chile | BANCO DE CHILE | 341005 | 44634 |  | EL GANTRY SE APAGA |  | 2 |  | Checked the equipment, error log but could not find a specific cause of the problem, the problem could be caused by a power drop in the power supply of the equipment, equipment is operational and under observation | [12 empty] | 119381360 | Closed | 1898857 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Customer | [6 empty] | System Power Off | [5 empty] | Non refresh | [12 empty] | No: Yes
|  |  | China | China Medical University Affiliated Shengjing Hospital Dalian Campus | 345001 | 44634 |  | 05R/Ingenuity Core128/Elite China/Unable to scan |  | 2 |  | check error log, no filament current, test tube filament&filament open. Replace the tube. | [12 empty] | 119376368 | Closed | 1900666 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Filament fault | [5 empty] | IB Existing |  | Tube | [10 empty] | No: Yes
|  |  | China | Henan Xuchang City Central Hospital (Infectious Diseases) | 40047 | 44633 |  | 02N/Ingenuity CT 728326/couch rear cover fault |  | 3 WANG LIN, LV |  | Replace the broken cover | [12 empty] | 119374636 | Closed | 1897687 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Couch cover broken | [5 empty] | IB Existing | 459800808451 | COVER ASSY, REAR SUBFRAME, GREY | [10 empty] | No: Yes
|  |  | Chile | Clinica Redsalud Rancagua | 40008 | 44594 |  | Noise detected when gantry was rotating |  | CFM |  | TBD | [14 empty] | 1916699 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Gantry rotation noise | [5 empty] | IB Existing | [12 empty] | No: Yes
| Date0328U |  | United States | Adventhealth Zephyrhills | 300099 | 44645 |  | ring artifact |  | 3 |  | Air calibration | [12 empty] | 119450770 | Closed | 1939430 |  |  | Closed - Done | Closed - Done | air calibration | [5 empty] | Customer | [6 empty] | Ring Artifact | [5 empty] | Refresh | [6 empty] | - | - | air calibration |  |  |  | No: Yes
| Date0328 |  | United States | Mclaren Macomb | 40021 | 44645 |  | Injector not connecting |  | 3 |  | System reboot | [12 empty] | 119450002 | Closed | 1938803 |  |  | Closed - Done | Closed - Done | CTD00007194 | [5 empty] | Engineering | [6 empty] | Injector not connecting | [5 empty] | Refresh |  |  |  | System | Console-SW | InjectorLose | Console-SW-InjectorLose | Console-SW-InjectorLose | CTD00007194 | Console-SW | InjectorLose | Console-SW-InjectorLose | No: Yes
|  |  | Argentina | INSTITUTO NACIONAL DE | 40027 | 44645 |  | NO DEJA REALIZAR CALIBRACIONES Y NO NO NOS DEJA HACER ESTUDIOS |  | 3 |  | TBD, reboot can restore the system, but will occur again, ask FSE to check on site | [12 empty] | 119449156 | In Progress | 1939031 |  |  | Opened | Opened | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | System cannot do calibration | [5 empty] | Refresh |  |  |  | System | TBD | NoEnoughInfo | TBD-NoEnoughInfo | TBD-NoEnoughInfo |  |  |  |  | No: Yes
|  |  | Japan | Social Welfare Corporation 3-beekai | 343000 | 44644 |  | 01CS:RS:問い合わせ:画面左下のCTインジケータでのImage Transfer failedというメッセージついて |  | 5 Yang, Cancan |  | None | [12 empty] | 119438961 | Fixed | 1931674 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | Image transfer fail | [5 empty] | Refresh |  |  |  | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No: Yes
|  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 337000 | 44644 |  | Pacs sistemine bağlanmasını istiyorlar |  | 3 |  | N/A | [12 empty] | 119441298 | In Progress | 1933474 |  |  | Evaluation | Evaluation | [6 empty] | Customer | [6 empty] | customer want to connect to the PACS | [5 empty] | Non refresh | [12 empty] | No: Yes
|  |  | Japan | Maki Hospital | 341017 | 44644 |  | WY:warranty期間中のフォローアップ |  | 5 Yang, Cancan |  | N/A | [12 empty] | 119439229 | Closed | 1936163 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A | [12 empty] | No: Yes
|  |  | Portugal | Centro Hospitalar de Vila Nova | 40030 | 44644 |  | Faz ruído ao deslizar a mesa e ao angular a gantry |  | 3 |  | Lubrication | [12 empty] | 119442567 | Closed | 1937305 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Noise when couch movement and gantry tilt | [5 empty] | IB Existing | [12 empty] | No: Yes
|  |  | Japan | Maki Hospital | 341017 | 44643 |  | 01CS:KK:AZEへの画像転送が止まっている。>>Unable to Connect with the target device.となっていましたがPingを確認するとPingは通りました。 |  | 4 Zhu, Chun |  | Hospital Lan issue | [12 empty] | 119432861 | Closed | 1928195 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Customer | [6 empty] | Image transfer fail | [5 empty] | Non refresh | [12 empty] | No: Yes
|  |  | Spain | S.C.S. HOSP. GRAL. DE FUERTEVENTURA | 40020 | 44643 |  | Las imagenes salen mal, se piensa que es el detector. |  | 2 |  | Replace the M15 | [12 empty] | 119434326 | Incomplete | 1928833 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Artifact | [5 empty] | IB Existing | 459800911271 | TDMS V2 M.Det Class A FRU | [10 empty] | No: Yes
|  |  | United States | Philips Medical Capital | 40028 | 44643 |  | failed quality control test |  | 3 |  | Replace HP phantom head section | [12 empty] | 119435041 | Closed | 1929017 |  |  | Evaluation | Evaluation | Hphantom Head | [5 empty] | Engineering | [6 empty] | IQ fail | [5 empty] | Refresh | 459801415161 | HARMONIZED PHANTOM HEAD SECTION |  | System | Couch-HW | H-phantom | Couch-HW-H-phantom | Couch-HW-H-phantom |  |  |  |  | No: Yes
|  |  | China | Zhalantun City People_s Hospital | 40045 | 44643 |  | 02N/Ingenuity CT 728326/Automatic voice volume is high |  | 3 Zhai, Lisa |  | Remote guide customer to adjust the voice volumn | [12 empty] | 119433356 | Closed | 1932264 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Customer | [6 empty] | Voice volume high | [5 empty] | Non refresh | [12 empty] | No: Yes
|  |  | China | Henan Xuchang City Central Hospital (Infectious Diseases) | 40047 | 44643 |  | 05K For IB/Ingenuity CT 728326/consult node information |  | 5 Cai, Cathy |  | N/A | [12 empty] | 119433727 | Closed | 1932493 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A | [12 empty] | No: Yes
|  |  | Chile | BANCO DE CHILE | 341005 | 44642 |  | PEDIR PARTES POSTERIOR EVALUACION |  | 5 SHU, Tim |  | TBD | [12 empty] | 119429129 | In Progress | 1926494 |  |  | Investigation | Investigation | [6 empty] | TBD | [6 empty] | TBD | [5 empty] | TBD | [12 empty] | No: Yes
|  |  | China | Le qing City People_s Hospital | 40054 | 44642 |  | 02N/Ingenuity CT 728326/report errors |  | 3 WANG LIN, LV |  | Reinstall Host and CIR SW | [12 empty] | 119426425 | Closed | 1927824 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | SW intermittent hung up doing CTA scan | [5 empty] | Refresh |  |  |  | System | Console-SW | Installation | Console-SW-Installation | Console-SW-Installation | 0 | Installation | Console-SW | Installation-Console-SW | No: Yes
|  |  | Puerto Rico | METRO MAYAGUEZ INC | 40009 | 44642 |  | Device is showing lines at imaging on Pilot |  | 3 |  | Clean the contrast | [12 empty] | 119429197 | Closed | 1929795 |  |  | Closed - Done | Closed - Done | [6 empty] | Customer | [6 empty] | Image Artifact | [5 empty] | Non refresh | [12 empty] | No: Yes
|  |  | United States | Clay County Medical Center | 40003 | 44641 |  | Patient Database on Local drive only shows last 2 days |  | 3 Boddolla, Ramu |  | Performed a Full Rebuild of the patient Local database | [12 empty] | 119423356 | Closed | 1921974 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | Bugrep | [5 empty] | Engineering | [6 empty] | Only last 2 days patient data show in Local | [5 empty] | Refresh |  |  |  | System | Console-SW | LocalDBLost | Console-SW-LocalDBLost | Console-SW-LocalDBLost |  |  |  |  | No: Yes
|  |  | China | Suzhou Guang ci Tumor Hospital Co., Ltd | 345019 | 44641 |  | 05K for IB/Core128/Blue screen |  | 3 |  | Close the back programs | [12 empty] | 119419075 | Closed | 1922520 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632553 | [5 empty] | Engineering | [6 empty] | Intermitten blue screen after reboot Host | [5 empty] | Refresh |  |  |  | System | Console-SW | Blue-Screen | Console-SW-Blue-Screen | Console-SW-Blue-Screen | CTUDT00632553 | Blue-Screen | Console-SW | Console-SW-Blue-Screen | bar: No
|  |  | China | Fujian Province Medical University No. 1 Affiliated Hospital | 345012 | 44641 |  | SA:05K for IB/Ingenuity Core128/Elite China/recon speed |  | 3 Cai, Cathy |  | 车载CT，delete the stuck queries to AI terminal | [12 empty] | 119419492 | Closed | 1922791 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Recon speed slow | [5 empty] | IB existing |  |  |  | System | CIRS-SW | [7 empty] | No: Yes
|  |  | China | Shaanxi Ankang T.C.M | 345041 | 44641 |  | Wechat:02N/Ingenuity Core128/Elite China/time not accuray |  | 5 |  | Modify the host time | [12 empty] | 119419569 | Closed | 1922832 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | TBD-TimeSync | [5 empty] | Engineering | [6 empty] | Host time is not accuracy | [5 empty] | TBD |  |  |  | System | Console-SW | Time sync | Console-SW-Time sync | Console-SW-Time sync | TBD-TimeSync |  | Console-SW | Console-SW-Time sync | No: Yes
|  |  | Poland | Uniwersytecki Szpital Kliniczny w | 341001 | 44641 |  | Awaria aparatu |  | 4 WANG LIN, LV |  | Correct the FRC input plug | [12 empty] | 119422936 | Closed | 1925481 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Fail to active scan after power outage | [5 empty] | IB Existing | [12 empty] | No: Yes
|  |  | United States | Westgate Hospital | 40001 | 44641 |  | Have a question about longitude condition |  | 5 |  | N/A | [12 empty] | 119423646 | Closed | 1926199 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A | [12 empty] | No: Yes
|  |  | United States | Clay County Medical Center | 40003 | 44640 |  | During scans, a gantry error was populated. |  | 5 WANG LIN, LV | [14 empty] | 119418853 | Closed | 1922262 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | Gantry intermittent error | [5 empty] | IB Existing | 453566496941 | RELAY,24VDC PLUGIN W/COVER | [10 empty] | No: Yes
| Date0406 |  | China | Xinjiang YiliuZhou Friendship Hospital | 40048 | 44648 |  | 05K for IB/Ingenuity CT 728326/RH panel |  | 3 Ma, Julita |  | Replace the right control panel | [12 empty] | 119455567 | Closed | 1945840 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Right control panel malfunction | [5 empty] | Refresh | 459800484511 | CONTROL PANEL ASSEMBLY-FRONT RIGHT |  | System | Gantry-HW | Control Panel | Gantry-HW-Control Panel | Gantry-HW-Control Panel |  | Control Panel | Gantry-HW | Gantry-HW-Control Panel | No: Yes
|  |  | China | Heilongjiang Province Huachuan County People_s Hospital | 345018 | 44648 |  | 01/Ingenuity Core128/Elite China/Image failure |  | 3 |  | TBD, suspect the reference detector and converter | [12 empty] | 119456018 | In Process | 1946071 |  |  | Evaluation | Evaluation | [6 empty] | Engineering | [6 empty] | Image contrast artifact | [5 empty] | IB Existing | [12 empty] | No: Yes
|  |  | China | Xinjiang YiliuZhou Friendship Hospital | 40048 | 44648 |  | 05K for IB/Ingenuity CT 728326/transfer slow |  | 3 Zhai, Lisa |  | Guide customer change the IP | [12 empty] | 119457234 | Closed | 1946882 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Customer | [6 empty] | Guide customer change IP | [5 empty] | Non refresh
|  |  | United States | Mount Grant General Hospital | 40012 | 44650 |  | 0711CDY72 not able to login or change the user//customer requested tech support |  | 3 |  | None. Customer cancelled the call, issue resolved by themselves | [12 empty] | 119474726 | Closed | 1953285 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Cannot login or change user | [5 empty] | Refresh |  |  |  | System | Console-SW | Not a defect | Console-SW-Not a defect | Console-SW-Not a defect |  | Not a defect | Console-SW | Console-SW-Not a defect | No: Yes
|  |  | United States | Philips Medical Capital | 40028 | 44650 |  | phantom was dropped |  | 3 |  | None | [12 empty] | 119472930 | Closed | 1956133 |  |  | Closed - Duplicate | Closed - Duplicate | [6 empty] | N/A | [6 empty] | Phantom drop down | [5 empty] | N/A
|  |  | United States | Philips Medical Capital | 40028 | 44650 |  | someone dropped fantam bar broke off inside |  | 5 |  | N/A | [12 empty] | 119472976 | Closed | 1956164 |  |  | Closed - Duplicate | Closed - Duplicate | [6 empty] | N/A | [6 empty] | Phantom drop down | [5 empty] | N/A
|  |  | China | Shaanxi Province Xianyang Central Hospital (Xianyang City No.2 People_s Hospital) | 40050 | 44650 |  | 02N/Ingenuity CT 728326/consult |  | 5 Zhai, Lisa |  | N/A | [12 empty] | 119473389 | Closed | 1956502 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A
|  |  | Poland | Uniwersytecki Szpital Kliniczny w | 341001 | 44651 |  | Konfiguracja nowego węzła |  | 4 |  | N/A | [12 empty] | 119484334 | Closed | 1960992 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A
|  |  | United States | Philips Medical Capital | 40055 | 44653 |  | 0711CFXXX Local Full |  | 3 |  | Full Rebuild, deleting, and restarting host | [12 empty] | 119496938 | Closed | 1963391 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632261 | [5 empty] | Engineering | [6 empty] | Local disk is full | [5 empty] | Refresh |  |  |  | System | Console-SW | DiskFull | Console-SW-DiskFull | Console-SW-DiskFull | CTUDT00632261 | DiskFull | Console-SW | Console-SW-DiskFull | No: Yes
|  |  | China | Suzhou Guang ci Tumor Hospital Co., Ltd | 345019 | 44652 |  | 01/Ingenuity Core128/Elite China/transfer image fault |  | 3 |  | PACS SW issue, customer contact with PACS service | [12 empty] | 119488761 | Closed | 1963428 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Customer | [6 empty] | Image transfer fail | [5 empty] | Non refresh
|  |  | United States | Unity Health Systems | 40037 | 44653 |  | Couch would not move in and out. |  | 3 |  | None. Customer did not request the service, since the error did not happen again. | [12 empty] | 119498086 | Closed | 1964804 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | Couch cannot move in/out | [5 empty] | IB Existing | [12 empty] | No: Yes
|  |  | United States | Philips Medical Capital | 40053 | 44652 |  | Left side cover popping when tilting |  | 3 |  | Replace the gantry cover | [12 empty] | 119495228 | Closed | 1965222 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | Left side cover poping when tilting | [5 empty] | IB Existing | 459801485211 | ASSEMBLY-UPPER SIDE COVER LH | [10 empty] | No: Yes
|  |  | Spain | S.C.S. HOSP. GRAL. DE FUERTEVENTURA | 40020 | 44655 |  | Ha salido un artefacto en la radiografía. |  | 3 |  | Replace defective module, and re-calibration. | [12 empty] | 119502367 | Closed | 1967257 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Image Artifact | [5 empty] | IB Existing | 459800911271 | TDMS V2 M.Det Class A FRU | [10 empty] | No: Yes
|  |  | Germany | Gesundheit Nord gGmbH Klinikverbund | 40002 | 44655 |  | Koronare und Sagittale Rekonstruktion vom Kopf werden nicht richtig berechnet. |  | 3 |  | None | [12 empty] | 119501530 | Closed | 1969415 |  |  | Investigation | Investigation | CTUDT00632116 | [5 empty] | Engineering | [6 empty] | Coronary and sagittal reconstruction of the head are not calculated correctly. | [5 empty] | Refresh |  |  |  | System | CIRS-SW | MPR Recon | CIRS-SW-MPR Recon | CIRS-SW-MPR Recon | CTUDT00632116 | MPR Recon | CIRS-SW | CIRS-SW-MPR Recon | No: Yes
| Date0407U |  | Mexico | DE LAGE LANDEN SA DE CV | 10007 | 44650 |  | VERIFICACION DE CALIBRACION DE AIRE Y MONITOR SIGNOS VITALES. |  | 5 |  | Replace the trigger cable | [12 empty] | 119469722 | Closed | 1954700 |  |  | Evaluation | Evaluation | [6 empty] | Engineering | [6 empty] | Check monitor vita signal | [5 empty] | IB Existing | 455012301901 | ECG Monitor Trigger 7M Cable | [10 empty] | No: Yes
|  |  | United States | Select Medical Corporation | 32261 | 44649 |  | Service Call |  | 5 Butler, Daria |  | replaced the Mylar ring | [12 empty] | 119467155 | Closed | 1953028 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | N/A | [5 empty] | IB Existing | 459800048742 | RING-LEXAN GN5 | [10 empty] | No: Yes
|  |  | United States | Philips Medical Capital | 300057 | 44648 |  | 0799CFXXX too many login attempts |  | 3 Datta, Avinaba |  | Call sustomer, and ask for wait for 30min, then try again. | [12 empty] | 119458842 | Closed | 1944077 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632236 | [5 empty] | Engineering | [6 empty] | System locked due to wrong password login too many times | [5 empty] | Refresh |  |  |  | System | Console-SW | User Account | Console-SW-User Account | Console-SW-User Account | CTUDT00632236 | User Account | Console-SW | Console-SW-User Account | No: Yes
| Date0411 |  | China | Jiangxi Province Children_s Hospital | 40041 | 44660 |  | 05K for IB/Ingenuity CT 728326/ left monitor display fault |  | 3 |  | Replace the left DP-DVI adaptor | [12 empty] | 119533401 | Closed | 1987919 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | DEFOA | [5 empty] | Engineering | [6 empty] | Left monitor display fault | [5 empty] | Refresh |  |  |  | System | Console-HW | Monitor | Console-HW-Monitor | Console-HW-Monitor | DEFOA | Monitor | Console-HW | Console-HW-Monitor | No: Yes
|  |  | United States | UNITY HEALTH SYSTEMS | 40037 | 44659 |  | Couch would not move in and out. |  | 3 |  | Lubricant the horizontal table rails | [12 empty] | 119529894 | Closed | 1986880 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | Couch cannot move in/out | [5 empty] | IB Existing | [12 empty] | No: Yes
|  |  | Argentina | INSTITUTO NACIONAL DE | 40027 | 44659 |  | TIRA DISTINTOS ERRORES EN CONSOLA DE COMANDOS |  | 4 |  | TBD | [12 empty] | 119529955 | New | 1986893 |  |  | Evaluation | Evaluation | [6 empty] | TBD | [6 empty] | TBD | [5 empty] | TBD
|  |  | Poland | Rzadowa Agencja | 341008 | 44659 |  | Uruchomienie aparatu |  | 4 |  | None | [12 empty] | 119529421 | Closed | 1986698 |  |  | Evaluation | Evaluation | [6 empty] | N/A | [6 empty] | Support for perform CT test | [5 empty] | N/A
|  |  | United States | Mclaren Macomb | 40021 | 44659 |  | 0711CFXXX protocol password expired. |  | 3 |  | Introduce to customer to reset the password via IT admin or call local service to reset the password. | [12 empty] | 119532290 | Closed | 1985569 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Engineering | [6 empty] | Protocol password expired | [5 empty] | Refresh |  |  |  | System | Console-SW | User Account | Console-SW-User Account | Console-SW-User Account |  | User Account | Console-SW | Console-SW-User Account | No: Yes
|  |  | Chile | BANCO DE CHILE | 341005 | 44659 |  | LOS EXAMENES APARECEN CON HORARIO CAMBIADO |  | 5 |  | Rebuilding time. | [12 empty] | 119530859 | Closed | 1984528 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Wrong schedule appers in the patient image. | [5 empty] | Refresh |  |  |  | System | Console-SW | NotEnoughInformation | Console-SW-NotEnoughInformation | Console-SW-NotEnoughInformation |  | NotEnoughInformation | Console-SW | Console-SW-NotEnoughInformation | No: Yes
|  |  | Spain | S.C.S. HOSP. GRAL. DE FUERTEVENTURA | 40020 | 44638 |  | Defectos estéticos encontrados durante el mantenimiento preventivo |  |  |  | N/A | [12 empty] | CFM |  | 1983711 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A
|  |  | Chile | BANCO DE CHILE | 341005 | 44658 |  | SOLICITA EL CAMBIO DE HORA DEL EQUIPO |  | 3 |  | Update time in computer. | [12 empty] | 119524895 | Closed | 1980925 |  |  | Evaluation | Evaluation | [6 empty] | Customer | [6 empty] | Request to change device time | [5 empty] | Non refresh
|  |  | United States | Central Vermont Medical Center | 40024 | 44658 |  | There is a bad detector. The FSE asked the customer to place a call so that he can order a part. |  | 4 |  | Replace the defective module | [12 empty] | 119524843 | Closed | 1980820 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | Image artifact | [5 empty] | IB Existing | 459800911281 | TDMS V2 M.Det Class B FRU | [10 empty] | No: Yes
|  |  | United States | Philips Medical Capital LLC | 40028 | 44657 |  | the vertical motion buttons on the ct control box do not work |  | 5 |  | Replace the CT Box | [12 empty] | 119519238 | Closed | 1980773 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | Vertical button on CT Box malfunction. | [5 empty] | IB Existing | 453567027195 | CTBOX ASSEMBLY RAFI/TILTED (RoHS) | [10 empty] | No: Yes
|  |  | China | Jiangxi Province Children_s Hospital | 40041 | 44657 |  | 05k for IB/Ingenuity CT 728326/Maintenance record |  | 3 |  | TBD | [12 empty] | 119514884 | New | 1978387 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | TBD | [6 empty] | TBD | [5 empty] | TBD
|  |  | China | Heilongjiang Province Huachuan County People_s Hospital | 345018 | 44657 |  | Wechat：01/Ingenuity Core128/Elite China/Double screen to single screen |  | 3 Zhai, Lisa |  | None | [12 empty] | 119513280 | Closed | 1977326 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | N/A | [6 empty] | Double screen to single screen | [5 empty] | N/A
|  |  | Germany | Gesundheit Nord gGmbH Klinikverbund Bremen Klinikum Bremen-Ost Chirurgische Klinik | 40002 | 44655 |  | MPR reconstruction incomplete |  | WANG, Peter |  | None | [12 empty] | CFM |  | 1976231 |  |  | Investigation | Investigation | CTUDT00632116 | [5 empty] | Engineering | [6 empty] | MPR reconstruction incomplete | [5 empty] | Refresh |  |  |  | System | CIRS-SW | MPR Recon | CIRS-SW-MPR Recon | CIRS-SW-MPR Recon | CTUDT00632116 | MPR Recon | CIRS-SW | CIRS-SW-MPR Recon | No: Yes
|  |  | Viet Nam | Công ty TNHH Thương mại và | 40046 | 44656 |  | CT BOX got defective |  | 3 |  | Replace the CT Box | [12 empty] | 119508962 | Closed | 1974461 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | CT Box defective | [5 empty] | IB Existing | 459801705022 | CT BOX ASSEMBLY_BEST | [10 empty] | No: Yes
| Date0411U |  | Mexico | DE LAGE LANDEN SA DE CV | 10007 | 44659 |  | EL EQUIPO ESTA ESTA MUY LENTO PARA LA ADQUISICION DE IMAGEN, PRESENTA CUENTA REGRESIVA PARA EXPLORACIÓN. SE TIENE... |  | 3 |  | Clean the filter, and re-cal EP | [12 empty] | 119532439 | Closed | 1987775 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | NoEnoughInfo | [5 empty] | Engineering | [6 empty] | Image acq speed low | [5 empty] | Refresh |  |  |  | System | CIRS-SW | ReconSlow | CIRS-SW-ReconSlow | CIRS-SW-ReconSlow | TBD-ReconSlow | ReconSlow | CIRS-SW | CIRS-SW-ReconSlow | No: Yes
|  |  | United States | Adventhealth Zephyrhills | 300106 | 44659 |  | 0611CDY70 Gantry shutoff |  | 3 |  | Replace the dual output PS | [12 empty] | 119531173 | Closed | 1984817 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | gantry shut down | [5 empty] | IB Existing | 453567374153 | POWER SUPPLY 120 VAC DUAL OUTPUT | [10 empty] | No: Yes
|  |  | United States | Lakeland Regional Health Systems | 300143 | 44658 |  | Parts Needed |  | 3 |  | Deliver new pads and straps to customer. | [12 empty] | 119524606 | Incomplete | 1984484 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Customer order parts. | [5 empty] | IB Existing | 455300002861, 459800033821, 459800042241, 453567009961, 453567111051 | Mouse pad, Flat Cushion, Head/armrest cushion, L-14"PATIENT RESTRAINT STRAP/ 42", HEADREST STRAP ASSY
|  |  | United States | Lakeland Regional Health Systems | 320474 | 44658 |  | Parts Needed |  | 3 |  | Deliver new pads and straps to customer. | [12 empty] | 119524552 | Closed | 1984449 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Customer order parts. | [5 empty] | IB Existing | 455300002861, 459800033821, 459800042241, 453567009961, 453567111051 | Mouse pad, Flat Cushion, Head/armrest cushion, L-14"PATIENT RESTRAINT STRAP/ 42", HEADREST STRAP ASSY
| Date0418 |  | United States | UNITY HEALTH SYSTEMS | 40037 | 44665 |  | Monthly QA |  | 5 |  | N/A | [12 empty] | 119559615 | Closed | 2006339 |  |  | Evaluation | Evaluation | [6 empty] | N/A | [6 empty] | Monthly QA | [5 empty] | IB Existing
|  |  | China | Zhalut Banner Mein medical hospital | 345036 | 44665 |  | 05k for ib/Ingenuity Core128/Elite China/monitor |  | 5 |  | Cancelled, system no error. | [12 empty] | 119558576 | Fixed | 2005650 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A
|  |  | China | Zhalut Banner Mein medical hospital | 345036 | 44665 |  | 05k for ib/Ingenuity Core128/Elite China/Stand by |  | 5 |  | Cancelled, system no error. | [12 empty] | 119558405 | Fixed | 2005525 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A
|  |  | China | Shandong Province Chest Hospital | 40025 | 44665 |  | 05L/Ingenuity CT 728326/blue screen |  | 3 |  | N/A | [12 empty] | 119557528 | Closed | 2004906 |  |  | Closed - Done | Closed - Done | CTUDT00632553 | [5 empty] | Engineering | [6 empty] | Bule screen | [5 empty] | Refresh |  |  |  | System | Console-SW | Blue-Screen | Console-SW-Blue-Screen | Console-SW-Blue-Screen | CTUDT00632553 | Blue-Screen | Console-SW | Console-SW-Blue-Screen | No: Yes
|  |  | Japan | Nikko Memorial Hospital | 341011 | 44666 |  | CT:APP:AK |  | 5 Xu, Xiaoting |  | None | [12 empty] | 119562703 | Closed | 2004742 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A
|  |  | Japan | Nikko Memorial Hospital | 341011 | 44665 |  | 01DS:KKI:心臓検査にて、バンディングアーチファクト発生 |  | 4 Yang, Cancan |  | Operation mistake, thickness were changed. | [12 empty] | 119561669 | Closed | 2004570 |  |  | Investigation | Investigation | [6 empty] | Customer | [6 empty] | Cardiac testing results in banding artifacts | [5 empty] | Non refresh
|  |  | China | Shaanxi Province Xianyang Central Hospital (Xianyang City No.2 People_s Hospital) | 40050 | 44665 |  | Wechat:02N/Ingenuity CT 728326/Controller key failure |  | 3 |  | Replace the CT Box | [12 empty] | 119556787 | Closed | 2004465 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632185 | [5 empty] | Engineering | [6 empty] | CT Box cannot control the patient support movement | [5 empty] | TBD | 459801705022 | CT BOX ASSEMBLY_BEST |  | System | Gantry-HW | CT Box | Gantry-HW-CT Box | Gantry-HW-CT Box | CTUDT00632185 | CT Box | Gantry-HW | Gantry-HW-CT Box | No: Yes
|  |  | United States | UNITY HEALTH SYSTEMS | 40037 | 44665 |  | Q/O: Foot switch needs adjustment |  | 5 |  | Adjust the footswitch | [12 empty] | 119560819 | Closed | 2003678 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Service | [6 empty] | Footswitch need to be adjusted | [5 empty] | Non refresh
|  |  | Japan | Yonago Medical Center | 341015 | 44665 |  | Slight ring artifact, performed air calibration |  | 4 SHU, Tim |  | Air calibration | [12 empty] | 119557344 | Closed | 2001278 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | Ring artifact | [5 empty] | IB Existing | [12 empty] | No: Yes
|  |  | China | Henan Xuchang City Central Hospital (Infectious Diseases) | 40047 | 44664 |  | 05K for IB/Ingenuity CT 728326/auto storage setting |  | 3 Cai, Cathy |  | Instructure customer to set the auto storage device. | [12 empty] | 119550210 | Closed | 2000495 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Customer reqest for set the auto storage device | [5 empty] | Non refresh
|  |  | United States | Philips Medical Capital LLC | 40028 | 44663 |  | ext 4203* iq check fail and air calibration (foreign object was found error) |  | 3 |  | Clean the foreign object | [12 empty] | 119545758 | Closed | 1997580 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | IQ Check fail | [5 empty] | IB Existing | [12 empty] | No: Yes
|  |  | China | Henan Xuchang City Central Hospital (Infectious Diseases) | 40047 | 44663 |  | 05K for IB/Ingenuity CT 728326/ADD PACS |  | 3 Ma, Julita |  | Add node of HIS and PACS | [12 empty] | 119545428 | Closed | 1997331 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | Support connect HIS and PACS | [5 empty] | Non refresh
|  |  | United States | Odessa Regional Medical Center | 40029 | 44663 |  | 1302CDY73 Images not transferring to PACS |  | 3 |  | checked network configurations, pacs and all destinations, no problem found no issues detected | [12 empty] | 119548945 | Closed | 1995673 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | Image transfer to PACS fail | [5 empty] | Refresh |  |  |  | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No: Yes
|  |  | United States | Philips Medical Capital LLC | 40053 | 44662 |  | equipment is making a squeaking noise when gantry is titled/please refer to 0119495228 |  | 5 |  | Adjusted CT side covers to eliminate noise when tilting | [12 empty] | 119541962 | Closed | 1995243 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | Squeaking noise when gantry tilting | [5 empty] | Refresh |  |  |  | System | Gantry-HW | Noise | Gantry-HW-Noise | Gantry-HW-Noise | TBD-GantryNoise | Noise | Gantry-HW | Gantry-HW-Noise | No: Yes
|  |  | Japan | Tomakomai City Hospital | 341014 | 44662 |  | SF:寝台カバーの歪み |  | 5 Xu, Xiaoting |  | Replace the BRACKET-PEDESTAL | [12 empty] | 119538282 | Closed | 1992463 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Couch noise | [5 empty] | IB Existing | 459800721902 | BRACKET-PEDESTAL | [10 empty] | No: Yes
|  |  | United States | Mclaren Macomb | 40021 | 44662 |  | 9799CFXXX need to chg protocol |  | 3 |  | Reset IT password, and introduce how to reset the CT user password. | [12 empty] | 119539383 | Closed | 1989961 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Cannot login the exam card management | [5 empty] | Refresh
| Date0418U |  | United States | Lakeland Regional Health Systems | 320474 | 44665 |  | MP5 monitor mounting bracket needs to be replaced |  | 3 |  | Replace the MP5 bracket | [12 empty] | 119560751 | Closed | 2007115 |  |  | Evaluation | Evaluation | [6 empty] | Engineering | [6 empty] | Need to replace the MP5 install bracket | [5 empty] | Non refresh | 451261019401 | IV-MP5 MET Mounting Adapter | [10 empty] | No: Yes
|  |  | United States | Upson Regional Medical Center | 310115 | 44665 |  | Remote Alert: CT-PRO-GANTRY_ControlLink |  | 5 |  | Clean the slip ring | [12 empty] | 119560813 | Closed | 2003574 |  |  | Evaluation | Evaluation | [6 empty] | Engineering | [6 empty] | gantry control link alert | [5 empty] | IB Existing | [12 empty] | No: Yes
|  |  | United States | Adventhealth Zephyrhills | 300099 | 44663 |  | No able to send images to PACS |  | 3 |  | Customer properly seating CAT5 Cable | [12 empty] | 119547377 | Closed | 1997732 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Cannot transfor images to PACS | [5 empty] | Non refresh | [12 empty] | No: Yes
|  |  | United States | Adventhealth Zephyrhills | 300099 | 44663 |  | Scanner not sending to PACS |  | 3 |  | Reseated Nic cable. And cold started system | [12 empty] | 119546605 | Closed | 1994977 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Cannot transfor images to PACS | [5 empty] | Non refresh | [12 empty] | No: Yes
|  |  | Spain | POLICLINICA NTRA SRA DEL ROSARIO SL | 336190 | 44663 |  | No deja disparar, error intermitente |  | 3 |  | Replace the anode module | [12 empty] | 119544364 | Closed | 1992306 |  |  | Evaluation | Evaluation | [6 empty] | Engineering | [6 empty] | Anode fail | [5 empty] | IB Existing | 453567560012 | Spellman 80 KW RoHS APM Assy | [10 empty] | No: Yes
| Date0425 |  | Japan | Yonago Medical Center | 341015 | 44676 |  | 01CS:KK:1件目の患者様撮影開始時に患者情報登録画面が出てこなかった。 |  | 4 |  | None | [12 empty] | 119602286 | Closed | 2030664 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632553 | [5 empty] | Engineering | [6 empty] | When logout / logged in the customer judgment, it seems to have frozen in the blue screen | [5 empty] | Refresh |  |  |  | System | Console-SW | Blue-Screen | Console-SW-Blue-Screen | Console-SW-Blue-Screen | CTUDT00632553 | Blue-Screen | Console-SW | Console-SW-Blue-Screen | No: Yes
|  |  | China | Shandong Province Chest Hospital | 40025 | 44675 |  | Wechat:05L/Ingenuity CT /Couch cann_t up/down |  | 3 |  | 重新连接绝对编码器接线，进行扫描床水平校正 | [12 empty] | 119600669 | Closed | 2030650 |  |  | Opened | Opened | TBD | [5 empty] | Engineering | [6 empty] | Couch cannot move up/down | [5 empty] | Refresh |  |  |  | System | Couch-HW | VerticalDriver | Couch-HW-VerticalDriver | Couch-HW-VerticalDriver | TBD | VerticalDriver | Couch-HW | Couch-HW-VerticalDriver | No: Yes
|  |  | China | Shandong Province Chest Hospital | 40025 | 44665 |  | System software can not log in. |  | Cai, Cathy |  | None | [12 empty] | CFM |  | 2030325 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632553 | [5 empty] | Engineering | [6 empty] | System software can not log in. blue screen | [5 empty] | Refresh |  |  |  | System | Console-SW | Blue-Screen | Console-SW-Blue-Screen | Console-SW-Blue-Screen | CTUDT00632553 | Blue-Screen | Console-SW | Console-SW-Blue-Screen | No: Yes
|  |  | Poland | Rzadowa Agencja | 341002 | 44673 |  | Wielokrotne wyłączanie się gantry. |  | 4 |  | Replace the dual output PS and the signal brush | [12 empty] | 119596804 | Closed | 2029435 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Gantry shut down automatically | [5 empty] | IB Existing | POWER SUPPLY 120 VAC DUAL OUTPUT ROHS, Brush, Signal Block Assembly | [11 empty] | No: Yes
|  |  | China | Shaanxi Province Xianyang Central Hospital (Xianyang City No.2 People_s Hospital) | 40050 | 44673 |  | 02N/Ingenuity CT 728326/The image has artifacts |  | 5 |  | TBD | [12 empty] | 119594347 | In Process | 2028495 |  |  | Opened | Opened | [6 empty] | TBD | [6 empty] | Image artifact | [5 empty] | TBD
|  |  | China | Shaanxi Ankang T.C.M | 345041 | 44671 |  | Wechat：02N/Ingenuity Core128/Elite China/Transfer work order |  | 5 |  | N/A | [12 empty] | 119582264 | New | 2020990 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A
|  |  | United States | Ascension Seton Medical Center | 40006 | 44671 |  | 0304CFXXX Parts ID Encoder failure |  | 5 |  | Provided customer with part numbers for the horizontal couch encoders and the gantry speake | [12 empty] | 119585643 | Closed | 2019538 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Ask for the part part number | [5 empty] | Non Refresh
|  |  | Chile | Scotiabank Chile | 40008 | 44671 |  | EL GANTRY NO INICIA |  | 2 |  | Replace the R2D | [12 empty] | 119583336 | In Progress | 2018170 |  |  | Opened | Opened | [6 empty] | Engineering | [6 empty] | System start-up error | [5 empty] | IB Existing | 459800467343 | R2D ASSEMBLY | [10 empty] | No: Yes
|  |  | China | Shandong Province Chest Hospital | 40025 | 44670 |  | WeChat：05L/Ingenuity CT 728326/Blue screen cannot start |  | 3 |  | None | [12 empty] | 119575391 | Closed | 2016595 |  |  | Evaluation | Evaluation | CTUDT00632553 | [5 empty] | Engineering | [6 empty] | Blue screen after restart host | [5 empty] | Refresh |  |  |  | System | Console-SW | Blue-Screen | Console-SW-Blue-Screen | Console-SW-Blue-Screen | CTUDT00632553 | Blue-Screen | Console-SW | Console-SW-Blue-Screen | No: Yes
|  |  | Japan | Maki Hospital | 341017 | 44670 |  | Scan aborted |  | 4 ZHANG, Yue Xing |  | None | [12 empty] | 119574541 | Closed | 2016040 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Tube soft arc | [5 empty] | IB Existing | [12 empty] | No: Yes
|  |  | United States | Lackey Memorial Hospital | 40015 | 44669 |  | 0711CDN85 unit stating that a critical error has occurred |  | 2 Butler, Daria |  | restored the backup and exam cards | [12 empty] | 119572105 | Closed | 2010830 |  |  | Evaluation | Evaluation | [6 empty] | Engineering | [6 empty] | system errors after closing exam cards | [5 empty] | Refresh |  |  |  | System | Console-SW | System-Error | Console-SW-System-Error | Console-SW-System-Error | 0 | Blue-Screen | Console-SW | Console-SW-System-Error | No: Yes
|  |  | Puerto Rico | DORADO HEALTH INC DBA | 40007 | 44669 |  | Ring artifact in images; Cleaned contrast on the DMS detectors and ran calibrations. |  | 3 XU, Miley |  | contrast liquid was found in multiple detectors, do the deep clean. | [12 empty] | 119571019 | Closed | 2010713 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | Image artifact | [5 empty] | IB Existing | [12 empty] | No: Yes
| Date0425U |  | United States | Adventhealth Zephyrhills | 300099 | 44675.2229 |  | IT password reset needed |  | 3 |  | The password has no issue when FSE on site. | [12 empty] | 119598751 | Closed | 2030080 |  |  | Opened | Opened | [6 empty] | Customer | [6 empty] | Password need reset | [5 empty] | Refresh
|  |  | United States | Adventhealth Zephyrhills | 300099 | 44675.1076 |  | Add PACS node |  | 3 |  | Add PACS node | [12 empty] | 119598242 | Closed | 2029911 |  |  | Opened | Opened | [6 empty] | Customer | [6 empty] | Add the PACS node | [5 empty] | Non Refresh
|  |  | United States | Adventhealth Zephyrhills | 300099 | 44672.834 |  | 1301CDN8A Need to add multiple nodes |  | 5 |  | Add the nodes | [12 empty] | 119583154 | Closed | 2021621 |  |  | Evaluation | Evaluation | [6 empty] | Customer | [6 empty] | Need to add multi nodes | [5 empty] | Non Refresh
|  |  | United States | Adventhealth Zephyrhills | 300106 | 44672.8306 |  | 1301CDN8A Need to add multiple nodes |  | 5 |  | Dicom node add remotely. | [12 empty] | 119583138 | Closed | 2021605 |  |  | Evaluation | Evaluation | [6 empty] | Customer | [6 empty] | Need to add multi nodes | [5 empty] | Non Refresh
|  |  | United States | Select Medical Corporation | 32261 | 44671.9139 |  | Yesterday the device had a ring artifact. Today they were unable to perform QA testing. Device won_t power on and... |  | 2 Miller, Steve |  | Add bad detector in the compliance list | [12 empty] | 119577235 | Closed | 2017931 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Ring artifact | [5 empty] | IB Existing | [12 empty] | No: Yes
|  |  | United States | Philips Medical Capital LLC | 300057 | 44670.0174 |  | has a question about info not crossing over to the scanner Requested tech support |  | 5 Dantha, Rajashekar |  | RIS/PACS system is the rootcause, not the CT system. | [12 empty] | 119571779 | Closed | 2010682 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | RIS/PACS | [6 empty] | Infor not corss over to the scanner | [5 empty] | Non Refresh
| Date0516 |  | China | Hebei Province Geriatric Hospital | 345015 | 44694 |  | 01/Ingenuity Core128/Elite China/ct value shift |  |  |  | TBD | [12 empty] | 119705678 | In Process | 2093780 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | CT number not accuracy | [5 empty] | Refresh |  |  |  | System | DMS-SW | CT number | DMS-SW-CT number | CT number-DMS-SW-CT number | 0 | Cover | Couch-HW |  | No: Yes
|  |  | China | Dongguan City Hospital of Integrated Chinese and Western Medicine | 345013 | 44693 |  | Wechat:05R/Ingenuity Core128/Elite China/Gantry noise |  |  |  | Replace the main driver belt | [12 empty] | 119694749 | Closed | 2089138 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Gantry rotation noise | [5 empty] | IB Existing | 453566492401 | BELT, 5-RIB POLY-VEE, 198"
|  |  | China | Nanchang University No. 1 Affiliated Hospital (Jinggangshan Hospital) | 345002 | 44692 |  | 01/Ingenuity Core128/Elite China/Query tube exposure record |  | Zhai, Lisa |  | N/A | [12 empty] | 119687327 | Closed | 2084269 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A
|  |  | China | Shaanxi Ankang T.C.M | 345041 | 44692 |  | 05k for IB/ Ingenuity Core128/Elite China/Repair and maintenance work order |  |  |  | TBD | [12 empty] | 119686970 | New | 2084065 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Engineering | [6 empty] | Routine PM | [5 empty] | Non Refresh
|  |  | China | Shandong Anchu City T.C.M | 345017 | 44692 |  | 05K For IB/Core128/ct box key pressed error |  |  |  | Replace the CT Box | [12 empty] | 119686924 | In Process | 2084023 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632185 | [5 empty] | Engineering | [6 empty] | CT Box button malfunction | [5 empty] | Refresh | 459801705022 | CT BOX ASSEMBLY_BEST |  | System | Gantry-HW | CT Box | Gantry-HW-CT Box | Gantry-HW-CT Box | CTUDT00632185 | CT Box | Gantry-HW | Gantry-HW-CT Box | No: Yes
|  |  | United States | Lackey Memorial Hospital | 40015 | 44692 |  | Doesn_t know how to Backup Protocols on Exam Cards |  | Das, Sayan |  | Work with customer on loading the EC onto the Thumb Drive to do the store. | [12 empty] | 119692201 | Closed | 2083728 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Customer do not know how to do the EC protocol backup. | [5 empty] | Non Refresh
|  |  | China | Sxi feng County T.C.M. Hospital | 345014 | 44691 |  | 01/Ingenuity Core128/Elite China/CT box damage |  |  |  | TBD | [12 empty] | 119682291 | In Process | 2080786 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632185 | [5 empty] | Engineering | [6 empty] | CT Box button malfunction | [5 empty] | Refresh |  |  |  | System | Gantry-HW | CT Box | Gantry-HW-CT Box | Gantry-HW-CT Box | CTUDT00632185 | CT Box | Gantry-HW | Gantry-HW-CT Box | No: Yes
|  |  | China | Harbin City Senhai Hospital | 345007 | 44691 |  | 01/Core128/consulting |  | Zhai, Lisa |  | N/A | [12 empty] | 119680749 | Closed | 2079879 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A
|  |  | Puerto Rico | METRO MAYAGUEZ INC | 40009 | 44691 |  | It does not expose, gantry can not comply |  | Howard, Amanda |  | N/A. When FSE got to the site, the system was working already. | [12 empty] | 119683764 | Closed | 2077715 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | Arc | [5 empty] | IB Existing
|  |  | Japan | YAMAGATA SAISEI HOSPITAL | 343000 | 44690 |  | AF:寝台下降時に異音がする |  | Yang, Cancan |  | Adjust the couch cover | [12 empty] | 119673828 | Closed | 2075696 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | NA | [5 empty] | Engineering | [6 empty] | Couch noise | [5 empty] | Refresh |  |  |  | System | Couch-HW | Cover | Couch-HW-Cover | Cover-Couch-HW-Cover | NA | Cover | Couch-HW |  | No: Yes
|  |  | China | Hebei Province Tangshan Workers Hospital | 40044 | 44690 |  | Wechat：01/Ingenuity CT 728326/gantry abnormal noise |  |  |  | TBD | [12 empty] | 119673359 | In Process | 2075407 |  |  | Evaluation | Evaluation | [6 empty] | Engineering | [6 empty] | Gantry rotation noise | [5 empty] | IB Existing
|  |  | Chile | Scotiabank Chile | 40008 | 44690 |  | EL GANTRY NO PUEDE INICIAR, LA FALLA REINCIDIO |  |  |  | Reload CGMP software, do calibration and image test | [12 empty] | 119678645 | Closed | 2074804 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | CGMP software corrupted due to sysetm unexpected shut down | [5 empty] | IB Existing
|  |  | China | Harbin City Senhai Hospital | 345007 | 44688 |  | 01/Ingenuity Core128/Elite China/artifact |  |  |  | TBD | [12 empty] | 119669631 | In Process | 2071983 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Artifact | [5 empty] | IB Existing
|  |  | United States | UNITY HEALTH SYSTEMS | 40037 | 44687 |  | Monthly QA |  |  |  | None | [12 empty] | 119666980 | In Progress | 2071270 |  |  | Evaluation | Evaluation | [6 empty] | Customer | [6 empty] | Monthly QA | [5 empty] | Non Refresh
|  |  | Japan | Maki Hospital | 341017 | 44687 |  | SF:Long TC作業 |  | Zhu, Chun |  | Long tube condition | [12 empty] | 119664625 | Closed | 2070323 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | N/A | [6 empty] | None | [5 empty] | IB Existing
|  |  | Germany | Gesundheit Nord gGmbH Klinikverbund Bremen Klinikum Bremen-Ost | 40002 |  |  | Coronal and Sagital MPR were not complet reconstructed |  | Lu, Xueying |  | CTUDT00632116 MPR, in SP2 scope | [14 empty] | 2069481 |  |  | Investigation | Investigation | CTUDT00632116 | [5 empty] | Engineering | [6 empty] | MPR incomplete reconstructed | [5 empty] | Refresh |  |  |  | System | CIRS-SW | MPR Recon | CIRS-SW-MPR Recon | CIRS-SW-MPR Recon | CTUDT00632116 | MPR Recon | CIRS-SW | CIRS-SW-MPR Recon | No: Yes
|  |  | United Kingdom | STOCKPORT NHS FOUNDATION TRUST | 40016 | 44684 |  | *FSE* Replace bottom front cover. Damage was picked up during a routine service. |  |  |  | Replace the lower front cover | [12 empty] | 119649287 | Closed | 2060215 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Cover damage | [5 empty] | IB Existing | 459801483431 | COVER-FRONT LOWER
|  |  | China | Jilin Province Songwong City Central Hospital | 345008 | 44684 |  | 01/Ingenuity Core128/Elite China |  | Cai, Cathy |  | None (Customer contact the sales themselves) | [12 empty] | 119646638 | Closed | 2058588 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Artifact | [5 empty] | IB Existing
|  |  | China | Zhalantun City People_s Hospital | 40045 | 44684 |  | wechat:01/Ingenuity CT 728326 |  | Zhai, Lisa |  | None | [12 empty] | 119646627 | Closed | 2058578 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | Account locked | [5 empty] | Refresh |  |  |  | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No: Yes
|  |  | United States | Clay County Medical Center | 40003 | 44684 |  | Change Philips VPN Encryption to SHA256 |  | Mohanta, Deepranjan |  | N/A | [12 empty] | 119649010 | Closed | 2057325 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A
|  |  | Japan | Kinan Hospital | 340000 | 44683 |  | SF:システムファントム破損 |  | Zhu, Chun |  | Replace the H phantom | [12 empty] | 119641966 | Closed | 2055336 |  |  | Evaluation | Evaluation | [6 empty] | Engineering | [6 empty] | H phantom broken | [5 empty] | Refresh | 459801608821 | HARMONIZED PHANTOM BODY SECTION-40MM |  | System | Couch-HW | H-Phantom | Couch-HW-H-Phantom | Couch-HW-H-Phantom |  | H-Phantom | Couch-HW | Couch-HW-H-Phantom | No: Yes
|  |  | United States | Clay County Medical Center | 40003 | 44683 |  | need to change the Encryption of the vpn tunnel |  | Dantha, Rajashekar |  | N/A | [12 empty] | 119646122 | Closed | 2054930 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A
|  |  | China | Jiangxi Province Children_s Hospital | 40041 | 44682 |  | 05K for IB/Ingenuity CT 728326/Reconstruction tank failure |  | Cai, Cathy |  | Ajust the CIRS time | [12 empty] | 119640240 | Closed | 2053025 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | CT image time is incorrect | [5 empty] | Refresh |  |  |  | System | Console-SW | TimeSync | Console-SW-TimeSync | Console-SW-TimeSync |  | TimeSync | Console-SW | Console-SW-TimeSync | No: Yes
|  |  | China | Sxi feng County T.C.M. Hospital | 345014 | 44680 |  | WeChat：01/Ingenuity Core128/Elite China/CT BOX Key failure |  | Cai, Cathy |  | TBD | [12 empty] | 119633696 | In Process | 2051531 |  |  | Evaluation | Evaluation | CTUDT00632185 | [5 empty] | Engineering | [6 empty] | CT Box couch down button malfunction | [5 empty] | Refresh |  |  |  | System | Gantry-HW | CT Box | Gantry-HW-CT Box | Gantry-HW-CT Box | CTUDT00632185 | CT Box | Gantry-HW | Gantry-HW-CT Box | No: Yes
|  |  | China | Qinghai Province HaibeiZhou No. 2 Hospital | 345029 | 44680 |  | 01/Ingenuity Core128/Elite China/Abnormal noise of machine |  | Zhai, Lisa |  | TBD | [12 empty] | 119630840 | In Process | 2050604 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Abnormal noise in system | [5 empty] | IB Existing
|  |  | United States | Odessa Regional Medical Center | 40029 | 44680 |  | Vertical Noisy Couch |  |  |  | Lubricant the couch mechanical system | [12 empty] | 119635055 | Closed | 2049862 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Couch vertical movement noise | [5 empty] | IB Existing
|  |  | Portugal | Centro Hospitalar de Vila Nova | 40030 | 44679 |  | Erro no reconstrutor |  |  |  | None. FSE on site, the system working fine. | [12 empty] | 119627324 | Closed | 2048272 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Rebuilder problem | [5 empty] | Refresh |  |  |  | System | Console-SW | Rebuilder | Console-SW-Rebuilder | Console-SW-Rebuilder |  | Rebuilder | Console-SW | Console-SW-Rebuilder | No: Yes
|  |  | China | Inner Mongolia Xing_an People_s Hospital | 345030 | 44679 |  | Appliaction/Ingenuity Core128/Elite China |  | Zhai, Lisa |  | N/A | [12 empty] | 119624588 | New | 2046983 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A
|  |  | United States | Philips Medical Capital LLC | 40055 | 44678 |  | 0711CDN89 Phantom files are full. FSE Vic Donald to be onsite for other issue. |  | Mohanta, Deepranjan |  | TBD | [12 empty] | 119622593 | Incomplete | 2044590 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Host HD full | [5 empty] | Refresh |  |  |  | System | Console-SW | Dispk full | Console-SW-Dispk full | Console-SW-Dispk full |  | Dispk full | Console-SW | Console-SW-Dispk full | No: Yes
|  |  | Japan | Yonago Medical Center | 341015 | 44679 |  | CT:APP:AK |  | Zhu, Chun |  | Introduce the correct protocol to scan | [12 empty] | 119626458 | Closed | 2043699 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Customer | [6 empty] | None | [5 empty] | Non Refresh
|  |  | Japan | Maki Hospital | 341017 | 44678 |  | 01CS:KKI:PACSへの画像転送が止まる |  | Xu, Xiaoting |  | None | [12 empty] | 119617884 | Closed | 2042952 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Image transfer to PACS stop | [5 empty] | Refresh |  |  |  | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection |  | Remote Connection | Console-SW | Console-SW-Remote Connection | No: Yes
|  |  | China | Jiangsu University Affiliated Hospital | 345023 | 44678 |  | 05k/Ingenuity Core128/Elite China/ password missing |  | Zhai, Lisa |  | Guide customer modify the password and add printer | [12 empty] | 119617411 | Closed | 2042670 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Consult for password and add printer | [5 empty] | Refresh |  |  |  | System | NA | NA | NA-NA | NA-NA |  | NA | NA | NA-NA | No: Yes
|  |  | Spain | S.C.S. HOSP. GRAL. DE FUERTEVENTURA | 40020 | 44678 |  | Problemas en el envio de estudios. Juan Pedro 607941742 - Wendy 928862092 |  |  |  | N/A | [12 empty] | 119619331 | Closed | 2040818 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A
|  |  | China | Jiangxi Province Children_s Hospital | 40041 | 44677 |  | 05K FOR IB/Ingenuity CT 728326/estop open |  | Zhai, Lisa |  | Reboot the CT system | [12 empty] | 119608972 | Closed | 2037869 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Estop Open | [5 empty] | IB Existing
|  |  | Turkey | Sezin Tibbi Goruntuleme Ve Kalp | 337000 | 44677 |  | Gantry fan arızası - yerinde tespit |  |  |  | Replace the fan | [12 empty] | 119613139 | Closed | 2037770 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Gantry fan error | [5 empty] | IB Existing | 459800980091 | BR64 Air Fan Replacement Kit
|  |  | United States | Leesburg Regional Medical Center | 40022 | 44676 |  | Line/Ring artifacts on images during clinical use; Detector module replaced |  | Sanchez, Josiah |  | Replace detector module | [12 empty] | 119605142 | Closed | 2035217 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | Ring artifact | [5 empty] | IB Existing
|  |  | Japan | Tomakomai City Hospital | 341014 | 44677 |  | 01CS:KK:17時45分頃HostからTSHSTSRV02に画像転送ができなかった。>>自然復旧したようですがログ確認依頼。関連は不明ですがFaulting application name:... |  | Zhu, Chun |  | None | [12 empty] | 119611808 | Closed | 2035048 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Host image transfer issue | [5 empty] | IB Existing
| Date0516U |  | United States | Wake Med Cary Hospital | 9591 | 44692 |  | NEMO |  | 3 Dantha, Rajashekar |  | N/A | [12 empty] | 119690502 | Closed | 2085763 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A
|  |  | United States | Wake Med Cary Hospital | 9591 | 44692 |  | X-ray system can not comply |  | 3 Reiner, Eyal |  | None | [12 empty] | 119691968 | Closed | 2083280 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | XRAY system can not comply | [5 empty] | IB Existing
|  |  | United States | Wake Med Cary Hospital | 9591 | 44687 |  | NEMO |  | 3 Boddolla, Ramu |  | N/A | [12 empty] | 119668360 | Closed | 2071720 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A
|  |  | United States | Adventhealth Zephyrhills | 300099 | 44687 |  | Gantry turned off itself and error message occurred, re-seated and cleaned all affected connections |  | 2 Lu, Xueying |  | Constrast inside gantry, clean and reset rotor connectors. | [12 empty] | 119669106 | Closed | 2069794 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | Constrast leakage | [5 empty] | IB Existing
|  |  | United States | Wake Med Cary Hospital | 9591 | 44687 |  | 0308CFXXX Customer Requesting Support - Parts ID |  | 5 Mohanta, Deepranjan |  | N/A | [12 empty] | 119668177 | Closed | 2069214 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A
|  |  | United States | Adventhealth Zephyrhills | 300106 | 44686 |  | Customer may have deleted the some DICOM node information accidentally and needs help restore those nodes. FSE required. |  | 4 |  | Looks like a policy in the hospital network isn’t allowing to send studies between them | [12 empty] | 119661608 | Closed | 2065670 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Customer | [6 empty] | Dicom node communications | [5 empty] | Refresh
|  |  | United States | Wake Med Cary Hospital | 9591 | 44684 |  | Table making knocking noise when moving vertically; No correction performed |  | 3 Squires, Ciera |  | None | [12 empty] | 119650530 | Closed | 2057885 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | Couch horizontal movement knock noise | [5 empty] | IB Existing
|  |  | United States | Wake Med Cary Hospital | 9591 | 44683 |  | NEMO |  | 3 Boddolla, Ramu |  | N/A | [12 empty] | 119643824 | Closed | 2056374 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A
|  |  | United States | Wake Med Cary Hospital | 9591 | 44683 |  | High Voltage Errors - NEMO - Cathode Power Module & Control PCA |  | 3 Furtado, Peter | [14 empty] | 119643586 | Closed | 2053605 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | Grid feedback imbalance | [5 empty] | IB Existing | 453567028521; 453566502911 | - CH/* CATHODE POWER MODULE; - CHPRINTED CIRCUIT CNTRL BD
| Date0523 |  | China | Nanjing Medical University Affiliated Yif Hospital | 345028 | 44702 |  | WeChat：05K/Ingenuity Core128/Elite China/Image artifact |  | 3 |  | Isolate the bad detector and do air-cal | [12 empty] | 119750638 | Closed | 2116980 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Image artifact | [5 empty] | IB Existing
|  |  | United States | UNITY HEALTH SYSTEMS | 40037 | 44701 |  | cluncking noise during rotation |  | 2 |  | Re-align the inner ring cover | [12 empty] | 119750421 | Fixed | 2116931 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Cluncking noise during rotation | [5 empty] | IB Existing
|  |  | China | 鄄 City Ma_aiyun Women and Children_s Hospital | 345044 | 44701 |  | 05K for IB/Ingenuity Core128/Elite China/mouse damage |  | 3 |  | Replace the mouse | [12 empty] | 119746432 | Closed | 2115795 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Mouse broken | [5 empty] | IB Existing
|  |  | China | Fujian Province Medical University Affiliated United Hospital | 345010 | 44701 |  | 05K for IB/Ingenuity Core128/Elite China/panel damage |  | 3 |  | Replace the control panel | [12 empty] | 119746266 | Closed | 2115726 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Panel damage | [5 empty] | IB Existing | 459800484511 | CONTROL PANEL ASSEMBLY-FRONT RH
|  |  | China | Fujian Province Medical University Affiliated United Hospital | 345010 | 44701 |  | 05k for IB/Ingenuity Core128/Elite China/To connect to the Internet |  | 3 |  | N/A | [12 empty] | 119745619 | Closed | 2115470 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A
|  |  | United States | Philips Medical Capital LLC | 40028 | 44700 |  | Need help exporting files |  | 5 |  | N/A | [12 empty] | 119742823 | Closed | 2113803 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A
|  |  | Japan | Tomakomai City Hospital | 341014 | 44700 |  | 01CS:KKI:右側の画面にてDirectoryが表示されない＞＞CTホストの再起動にて復旧 |  | 4 Yang, Cancan |  | Restart the host computer | [12 empty] | 119737839 | Closed | 2110967 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Directory is not displayed on the right screen | [5 empty] | Refresh |  |  |  | System | Console-HW | Mornitor | Console-HW-Mornitor | Console-HW-Mornitor |  | Mornitor | Console-HW | Console-HW-Mornitor | No: Yes
|  |  | China | Jiangsu Province No. 2 T.M. Hospital | 345027 | 44700 |  | Wechat：05K/Core128/image fault |  | 3 |  | Re-install the SW | [12 empty] | 119737788 | Closed | 2110922 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | Quick film malfunction | [5 empty] | Refresh |  |  |  | System | Console-SW | Film | Console-SW-Film | Console-SW-Film |  | Film | Console-SW | Console-SW-Film | No: Yes
|  |  | China | Jiangsu Province No. 2 T.M. Hospital | 345027 | 44699 |  | Wechat：05K/Core128/No image |  | 3 WANG LIN, LV |  | Guide customer use ISP or execute 1 serial once for film | [12 empty] | 119732920 | Closed | 2107756 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | No image in film view | [5 empty] | Refresh |  |  |  | System | Console-SW | Film | Console-SW-Film | Console-SW-Film |  | Film | Console-SW | Console-SW-Film | No: Yes
|  |  | China | Zhalantun City People_s Hospital | 40045 | 44699 |  | Wechat：01/Ingenuity CT/Consult |  | 3 Ma, Julita |  | N/A | [12 empty] | 119732077 | Closed | 2107338 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A
|  |  | China | Beijing City Hospital of Integrated Chinese and Western Medicine | 40032 | 44699 |  | 05K for IB/Ingenuity CT 728326/problem not confirm |  | 3 Zhai, Lisa |  | FSE contact with customer, state the system is workin gnormal. | [12 empty] | 119731038 | Closed | 2106637 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A
|  |  | Spain | S.C.S. HOSP. GRAL. DE FUERTEVENTURA | 40020 | 44699 |  | Tiene la colchoneta y almohadilla rotas |  | 3 | [14 empty] | 119732364 | Closed | 2103292 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Cushion broken | [5 empty] | Refresh | 459800126661 | Cushion-47.24 Inch Mattress Pad |  | System | Couch-HW | Cushion broken | Couch-HW-Cushion broken | Couch-HW-Cushion broken |  | Cushion broken | Couch-HW | Couch-HW-Cushion broken | No: Yes
|  |  | China | Fujian Province Medical University No. 1 Affiliated Hospital | 345012 | 44698 |  | SA:05K for IB/Ingenuity Core128/Elite China |  | 3 Zhai, Lisa |  | Guide customer to modify the password. | [12 empty] | 119724396 | Closed | 2102120 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | CTUDT00632127 | [5 empty] | Engineering | [6 empty] | Account locked | [5 empty] | Refresh |  |  |  | System | Console-SW | Remote Connection | Console-SW-Remote Connection | Console-SW-Remote Connection | CTUDT00632127 | Remote Connection | Console-SW | Console-SW-Remote Connection | No: Yes
|  |  | China | Shaanxi Ankang T.C.M | 345041 | 44698 |  | 05k for IB/Ingenuity Core128/Elite China/Does not automatically rebuild the image |  | 3 |  | None | [12 empty] | 119724336 | Closed | 2102069 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | Does not rebuild the image | [5 empty] | Refresh |  |  |  | System | Console-SW | Image recon | Console-SW-Image recon | Console-SW-Image recon |  | Image recon | Console-SW | Console-SW-Image recon | No: Yes
|  |  | China | Hebei Province Tangshan Workers Hospital | 40044 | 44697 |  | 01/Ingenuity CT 728326/buy parts |  | 3 |  | TBD | [12 empty] | 119719634 | New | 2098933 |  |  | Evaluation | Evaluation | [6 empty] | TBD | [6 empty] | TBD | [5 empty] | TBD
|  |  | China | Shandong Province Chest Hospital | 40025 | 44697 |  | Wechat：01/Ingenuity CT 728326/equipment failure |  | 3 |  | TBD | [12 empty] | 119719131 | In Process | 2098624 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Tube too hot, and cannot cool down | [5 empty] | IB Existing
|  |  | China | Henan Xuchang City Central Hospital (Infectious Diseases) | 40047 | 44697 |  | 05K For IB/ Ingenuity CT/Image fault |  | 5 |  | TBD | [12 empty] | 119718022 | New | 2097894 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | first 2 slice show left half image, last 2 slice show right half image. | [5 empty] | TBD |  |  |  | System | Console-SW | Image recon | Console-SW-Image recon | Console-SW-Image recon |  | Image recon | Console-SW | Console-SW-Image recon | No: Yes
|  |  | China | Beijing City Hospital of Integrated Chinese and Western Medicine | 40032 | 44697 |  | Scan failed with gantry cannot comply error, replaced Resolver to Digital Assembly |  | 3 ZHANG, Yue Xing |  | Replace R2D. | [12 empty] | 119717557 | Closed | 2097550 |  |  | Closed - Done | Closed - Done | [6 empty] | Engineering | [6 empty] | ACQ not OK | [5 empty] | IB Existing |  | R2D
| Date0523U |  | Spain | POLICLINICA NTRA SRA DEL ROSARIO SL | 336190 | 44700 |  | Uno de los laser no se enciende |  | 3 |  | Replace Laser | [12 empty] | 119740345 | Closed | 2108759 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Laser not on | [5 empty] | IB Existing | 453567033441 | LASER ASSY, RH FRONT COVER
|  |  | United States | Wake Med Cary Hospital | 9591 | 44697 |  | NEMO |  | 3 |  | N/A | [12 empty] | 119723011 | Closed | 2101400 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A
|  |  | United States | Adventhealth Zephyrhills | 300099 | 44697 |  | 0599CDY72 Unable to turn gantry on |  | 2 Boddolla, Ramu |  | Replace the victor PS and the brush | [12 empty] | 119723253 | Closed | 2096961 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [13 empty] | Unable to turn gantry on | [5 empty] | IB Existing | 453567374153 | Victor PS, Brush block
|  |  | United States | Adventhealth Zephyrhills | 300099 | 44697 |  | Q/O |  | 3 |  | N/A | [12 empty] | 119723118 | Fixed | 2096877 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A
| Date0530 |  | China | Jiangsu Southeast University Affiliated China University Hospital | 345021 | 44709 |  | SA:05K/Ingenuity Core128/Elite China/Add network node |  | 3 |  | Add the network node | [12 empty] | 119796474 | Closed | 2139615 |  |  | Opened | Opened | [6 empty] | Customer | [6 empty] | Customer ask for add network node | [5 empty] | Non Refresh
|  |  | United Kingdom | STOCKPORT NHS FOUNDATION TRUST | 40016 | 44708 |  | images haven’t reconstructed on, we have done them from the raw data |  | 3 |  | TBD | [12 empty] | 119790751 | In Progress | 2138101 |  |  | Evaluation | Evaluation | [6 empty] | Engineering | [6 empty] | Image not reconstructed | [5 empty] | Refresh |  |  |  | System | Console-SW | Image recon | Console-SW-Image recon | Console-SW-Image recon |  | Image recon | Console-SW | Console-SW-Image recon | No: Yes
|  |  | China | Dongguan Donghua Hospital Co., Ltd | 40059 | 44708 |  | WeChat：SA:05K FOR IB/Ingenuity CT /Reconstruct image |  | 3 |  | Add RAM | [12 empty] | 119789082 | Closed | 2137518 |  |  | Opened | Opened | [6 empty] | Engineering | [6 empty] | Image reconstruction slow | [5 empty] | Refresh |  |  |  | System | Console-SW | Image recon | Console-SW-Image recon | Console-SW-Image recon |  | Image recon | Console-SW | Console-SW-Image recon | No: Yes
|  |  | United States | Philips Medical Capital LLC | 40005 | 44708 |  | advising on product alert letter 9599CFXXX |  | 5 |  | N/A | [12 empty] | 119794227 | Closed | 2137110 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A
|  |  | Japan | Yonago Medical Center | 341013 | 44708 |  | 05CS:KK:検査室内の除湿器が排水されず満水になっている。>>装置には影響ないことから、そのままにしていただき、後日FSEが確認となりました。 |  | 4 Xu, Xiaoting |  | N/A | [12 empty] | 119791065 | Closed | 2134847 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A
|  |  | China | 山东省胸科医院 | 40025 | 44707 |  | 01/Ingenuity CT 728326/add printer |  | 3 |  | N/A | [12 empty] | 119777665 | Closed | 2134230 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | N/A | [6 empty] | N/A | [5 empty] | N/A
|  |  | China | Ning_an City No. 2 People_s Hospital | 345039 | 44707 |  | 01/Ingenuity Core128/add the printer |  | 3 |  | TBD | [12 empty] | 119776203 | In Process | 2133996 |  |  | Opened | Opened | [6 empty] | Customer | [6 empty] | Add the printer | [5 empty] | Non Refresh
|  |  | Portugal | Centro Hospitalar de Vila Nova | 40030 | 44704 |  | Com barulhos anormais vindos da gantry |  | 3 |  | TBD | [12 empty] | 119756950 | New | 2122651 |  |  | Evaluation | Evaluation | [6 empty] | Engineering | [6 empty] | Gantry abnormal noise | [5 empty] | IB Existing
|  |  | Spain | S.C.S. HOSP. GRAL. DE FUERTEVENTURA | 40020 | 44704 |  | Se cambió uno de los sensores pero el problema continua, no funciona |  | 2 WANG LIN, LV |  | Replace the module, do the cal. | [12 empty] | 119755457 | In Progress | 2118388 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Artifact | [5 empty] | IB Existing
| Date0530U |  | United States | Upson Regional Medical Center | 333009 | 44709 |  | After Power Outage/Unable to Scan |  | 2 |  | Customer restart the entire system | [12 empty] | 119795620 | Closed | 2137815 |  |  | Investigation | Investigation | [6 empty] | Engineering | [6 empty] | Unable to scan | [5 empty] | IB Existing
|  |  | United States | Adventhealth Zephyrhills | 300099 | 44707 |  | 1301CDY73 Network changes needed |  | 3 |  | N/A | [12 empty] | 119782475 | Closed | 2132768 |  |  | Closed - Non-Complaint | Closed - Non-Complaint | [6 empty] | Customer | [6 empty] | Need to change network | [5 empty] | IB Existing
|  |  | United States | Adventhealth Zephyrhills | 300099 | 44704 |  | Gantry will not turn on. |  | 2 |  | Repalce the main fused, inspection brush block, test syttem. | [12 empty] | 119759920 | Closed | 2120325 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Engineering | [6 empty] | Gantry cannot turn on | [5 empty] | IB Existing
|  |  | United States | Wellstar Kennestone Hospital | 68117 | 44704 |  | CIRS isn_t powering on; Customer resolved |  | 2 |  | Customer had a power outage. Ask customer press the start contactor on the wall. | [12 empty] | 119756967 | Closed | 2118801 |  |  | Closed - No Further Investigation | Closed - No Further Investigation | [6 empty] | Customer | [6 empty] | Cannot power on the system | [5 empty] | IB Existing

## brilliance1
id | FOK_Action_Item_ID | Status_2 | Country | Site_Name | S_or_N | Date_Reported | wk | Field_Issues_Heading | Field_Issue_Description___Field_reported | Customer_Impact_or_Effect | Bugreport | CA | Target_to_Close___in_5_Workday | PA | Target_to_Define_action_plan___in_10_Workday | Picture | During__or__After_Installation | Lebel2 | Level3 | Root_Cause_ | Countermeasures | Action___PA | Owner | Timing | Viper_or_One_EMS_No | Case_Status | Complaint___PR__No | tru | status | PR_Status | Complaint_Status | Defect___CTUDT_ | Application_Specialist__or__Clinical_Science_Comments | Customer_Feedback_ | Highlight | Assign_Owner | Customer_Facing_Issues | Funct_Owner_1 | Investigation_Results | Correction_Plan | Action | Status_1 | Failure_mode | Categoly | SSS | Status_comment | Start_Clinical_Date | Between_failure | Fix_Priority_ | Mitigation_or__Workaround | IB_Exist_or_New_Found | Related_FRU_P_or_N | FRU_Description | Details__or__Comments | Level1 | Level2_ | Level3_1 | Level2_3 | System_Type_0 | silva | check_rename_Level3_Type | Type1_level2 | type2_silva | Month_Rolling | level1_check_rename | level2_check_rename | level3_check_rename | level2__3_check_rename | System_Type_0_check_rename | Month_Rolling_check_rename | ifequal | silva_check_rename___compared_with_BI | check_rename_Level3_Type__BJ | Type1_level2__BK | type2_silva__BL | Funct_Owner | PR_status_check_rename | IB_Existing_check_rename | CTUDTcheck_rename | SSS_check_rename | check_rename | end_rename
1 | 43993 | TRUE | US | The_Regional_Medical_Center | 40005 | 43993 | 24 | Cable_Harness_been_squeezed | Cable_Harness_Rear_Cover_Close_Up_been_pressed_by_screw_hole,_see_picture | No |  | Adjust_and_fixing_cable | HFA_investigation | Drawing_update_and_supplier_quality_control ECR_075587_submitted_to_fix_drawing_update | ECR_075587 | 0 | During |  |  |  | Short_term | Long_term |  |  | CFM |  | 152152 | TRUE | Closed__Non_Complaint | Closed___Non_Complaint | Valid | CTUDT00630134 |  |  | N | Wu_ZhiMing | Installation_Covers | Engineering | Cable_Harness_Rear_Cover_Close_Up_been_pressed_by_screw_hole The_cable_harness_was_assemblied_in_supplier_side | Supplier_Quality_Issue | need_PE_or_SQE_to_check_rename_and_adjust_the_cable_tie_position_away_from_the_screw_hole | Closed | Manufacturing_Issue | System_Hardware | Supplier_Quality | Ongoing,_ZhiMing_is_aligning_with_supplier_to_add_additional_cable_clip_for_above_the_screw_hole 6_or_16_update: drawing_will_update_to_add_more_cable_clip |  |  | Low | N_or_A,_Cable_work_with_function,_continue_monitoring |  | N_or_A | N_or_A | System_issue | system | Gantry_HW | gantry_cover | Gantry_HW_gantry_cover | Gantry_HW_gantry_cover | CTUDT00630134 | gantry_cover |  |  | No | y | y | y | y | y | No | y | y | y | y | y | y | y | y | y | y | y